## Sheet1
Recall Number | Recall Event ID | Device Type | Product Code | Regulation Description | Medical Specialty | Review Panel | Submission Type | Trade Name/Product | Main Name | Recall Class | Date Posted | Year Posted | Recalling Manufacturer | Reason for Recall | FDA Determined Cause | Action | Consumer Instructions | Clean Quantity | Quantity in Commerce | Distribution | Recall Status | Termination Date | Time to Terminate | Merged Quantities | Fault_Class
Z-0321-2007 | 36854 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | RYM-5001 InVision-PlusÂż Neutral Displacement I.V. Connector, packaged 100 units per box, 5 boxes per case (500 units/case); labeled as Sterile-Nonpyrogenic; assembled in Mexico for Rymed Technologies, Inc., Austin, TX. |  |  | May 25, 2006 | 2006 | Rymed Technologies, Inc. | Weld failures on IV connector (area of female luer and spike body) | Other | Consignees were notified by e-mail and/or phone on 05/25/06 regarding the recall. Distributors were asked to return product. |  | 28479 | 28,479 units. | Nationwide Recall | Terminated | March 26, 2007 | 305 | 79328 | Undetermined
Z-1061-2007 | 37778 | No | MLY |  |  | Physical Medicine | 510(k) | Gebauer''s Pain Ease Medium Stream -Topical Anesthetic Skin Refrigerant is a vapocoolant (skin refrigerant) packed in an aerosol container. Rx only. 103.5ml (3.5 fl. oz.) can-(Product Number-0386-0008-03) |  |  | April 17, 2007 | 2007 | Gebauer Company | Mold contamination | Other | The firm sent a recall letter dated April 17, 2007 to all customers. |  | 9732 | 9,732 cans of product | Nationwide | Terminated | November 30, 2009 | 958 | 41499 | Undetermined
Z-1065-2007 | 37881 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | ALFA II Modular Femoral Left Modular Stem Porous Coated (Bowed); Size 6 Left; Length: 250mm; Distal Diameter: 18.0mm; Catalog #652-06-180L; component of the ENCORE Hip System; Single Use, Sterile; distributed by Encore Medical, L.P., Austin, TX 78758. |  |  | April 12, 2007 | 2007 | Encore Medical, Lp | Sterility/Package Integrity: Device packaging may potentially be unable to ensure sterility when exposed to shipping conditions. | Other | Urgent Device Recall Notification letters distributed to consignees on 04/14/07. Consignees requested to return product to Encore Medical, L.P. |  | 2 | 2 units. | Product distributed to surgical centers in CT, KS and MO. | Terminated | July 25, 2007 | 104 | 8 | Undetermined
Z-1117-2007 | 38194 | No | FRW | Patient scale. | General Hospital | General Hospital | 510(K) Exempt | Liko Universal Sling Bar 450 Quick Release Hook (QRH) |  |  | June 13, 2007 | 2007 | Liko, Inc. | Sling bar may have a center bolt break at the adapter while in use | Component design/selection | Liko Inc notified distributors and end users by letter on 6/13/07. |  | 22 | 22 | Nationwide | Terminated | October 20, 2015 | 3051 | 25 | Undetermined
Z-1090-2007 | 38154 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Performer CPB-Universal R2100375. |  |  | June 04, 2007 | 2007 | Medtronic Perfusion Systems | Device failure requiring use of back up systems (malfunctioning printed circuit board (PCB)) | Other | There are two separate letters dated June 4, 2007 for the customers, based on the configuration of the unit being inspected/updated in the field. These letters explains the situaiton and informs the customers that Medtronic will be visiting each customer site to inspect/update the units. Medtronic Service Technicians will complete a Field Action Certicate acknowledging the successful completion of the update. |  | 24 | 24 units | Worldwide Distribution: USA including states of PA, ND, MN, CA, AL, CO, MI, NY, AZ, and OH and countries of Sweden, Iceland, Latvia, Germany, Italy, Austria, Belgium, and France. | Terminated | March 02, 2008 | 272 | 24 | Undetermined
Z-1089-2007 | 38128 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | BCI Advisor Vital Signs Monitor (Model/Catalog Number: 9200) |  |  | June 05, 2007 | 2007 | Smiths Medical PM, Inc. | Turn off unexpectedly/ Not turn on : monitors turned off unexpectedly sounding the critical failure alarm. Advisor Vital Signs monitors that were off, would not turn back on. | Other | Consignees will be contacted via a Safety Action Bulletin 07-SAB02 dated June 5, 2007. Technical Worksheet, 07-TW01 including a kit will be sent to the International consignees whom Technical Service has deemed qualified to perform the repair. The repair will be performed at the consignee site. International consignees who require or request the repair be done at the Smiths Medical PM, Inc. Technical Service department will return their monitors for repair. |  | 57 | 57 | International Distribution Only: Czech Republic, Peru, Argentina, Pakistan, Greece, Romania, The Netherlands, Germany, Puerto Rico, United Kingdom, Brazil, Uruguay, Spain, Australia, and Israel | Terminated | October 29, 2008 | 512 | 57 | Undetermined
Z-1116-2007 | 38105 | No | DYG | Flow-directed catheter. | Cardiovascular | Cardiovascular | 510(k) | Thermodilution Right Heart Catheters, Catalog Numbers 604028 and 604165. The product is shipped 5 units per carton.-B. Braun Medical, Allentown, Pa. |  |  | June 05, 2007 | 2007 | B. Braun Medical, Inc. | Cracks at the Catheter Tip | Other | On 6/5/07, a Product Removal Notice was issued to the only distributor that received the product in Japan |  | 4695 | 4,695 | International ONLY- Japan | Terminated | October 31, 2007 | 148 | 4695 | Undetermined
Z-1115-2007 | 38055 | No | MCW | Intraluminal artery stripper. | Cardiovascular | Cardiovascular | 510(k) | ThromCat Thrombectomy Catheter System. Catalog number 63000-01 |  |  | May 21, 2007 | 2007 | Kensey Nash Corp | Defective component (power supply) | Other | New power supplies were sent directly to Sales Representatives who personally replaced the defective power supplies and returned them to Kensey Nash. A voice mail to all affected KNC Sales Representatives went out on 5/21/07 instructing them to contact the affected accounts immediately. |  | 11 | 11 units | Nationwide:The products were shipped to medical facilities in FL, MO, NJ, and TX. | Terminated | August 29, 2007 | 100 | 11 | Undetermined
Z-1056-2007 | 37728 | No | LAO |  |  | Toxicology | 510(k) | TDx/TDxFLx Methotrexate II Reagent-List Number 07A12-60 - a reagent system for In Vitro diagnostic use. This reagent system is composed of one bottle of Wash Solution, one bottle of Methotrexate Antibody, one bottle of Methotrexate Fluorescein Tracer and one bottle of Pretreatment Solution. |  |  | March 20, 2007 | 2007 | Abbott Diagnostic International, Ltd. | Incorrect configuration: kits contain reagent bottles in the order S-W-T-P instead of the correct order, W-S-T-P. | Other | The Product Correction Letter dated 3/20/2007, was sent to all customers who received the affected lots. The Customers were instructed to: Determine if they had the lot listed in the Product Correction letter. Inspect each kit of the affected lot and discard any kits incorrectly configured. If they had forwarded the lots listed in the Product Recall letter to other laboratories, they were instructed to inform them of this Product Recall and provide a copy of this letter to them and retain Product Correction letter for their laboratory records. |  | 1026 | 1,026 | World wide Distribution : USA including states of AR, CA, CO, CT, DE, FL,GA, ID, IL, IA, KS, KY, LA, MA, ME, MD, MI, MN, MS, MO, NJ, NY, NC,NV, OH, SC, PA, RI, TN, TX, UT, WA, WI, WV, Hawaii, Washington DC, and Puerto Rico. Countries of :Canada, Mexico, Venezuela, Columbia, Chile, Brazil, Germany, Japan, Hong Kong, Australia and New Zealand | Terminated | June 10, 2011 | 1543 | 1026 | Undetermined
Z-1019-2007 | 38220 | No | GAM | Absorbable poly(glycolide/l-lactide) surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Various Sizes of Sterile Polyglycolic Acid Synthetic Surgical Suture-NDR, POLYGLY 2-0 C-2/CST43 2N36, Catalog #833-135 |  |  | June 06, 2007 | 2007 | Teleflex Creek Dr | Suture Diameter/Strength specification failure: The products may not meet their final specifications for suture diameter and/or strength which could lead to premature breakage. | Other | Consignees were notified by letter on/about 06/07/2007. They were advised to cease use, remove and quarantine all affected product immediately. Affected products are to be returned to Teleflex Medical. An Acknowledgement &amp; Stock form was included to be completed and returned via fax to Teleflex Medical. |  | 1920 | 1920 units total for all types of sutures | Worldwide-, USA including states of CO, FL, IL, IN, MA, and OH, and Countries of Canada, France and Germany | Terminated | March 14, 2008 | 282 | 1920 | Undetermined
Z-1042-2007 | 38016 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | OneTouch Ultra Test strip, a component of the following OneTouch Ultra Glucose monitoring systems: a)OneTouch Ultra, b) OneTouch InDuo, c) OneTouch Ultra2, d) OneTouch UltraMini, e) OneTouch UltraSmart. 25 strips per vial. Product comes in packs of 100 and 50. (Part numbers: 020-245-07, 020-244-07, 020-963-01) |  |  | March 22, 2007 | 2007 | Lifescan Inc | Damaged Vials/Inaccurate test results: a manufacturing equipment problem on one of the packaging lines had the potential to puncture vials, causing a single hole in the side of the vial below the rim. This result of this defect may produce inaccurately high or low blood glucose test results or error message readings. | Other | March 2007 via letters to consumers and consignees and web postings. The firm has also set up verbal communication plans for customers who contact Lifescan by phone. |  | 33000000 | ~33 million vials (25 strips per vial in the US, ~19 million worldwide. | Worldwide; USA, Europe, Middle East, Africa, Latin America and Canada. | Terminated | August 15, 2008 | 512 | 33000000 | Undetermined
Z-1077-2007 | 38169 | No | EAZ | Dental operating light. | Dental | Dental | 510(K) Exempt | DCI Equipment Track Mount Light Assembly, 110v |  |  | May 01, 2007 | 2007 | Dental Equipment LLC | Potential for light shield to melt due to light assemblies not properly rated to handle the VAC. | Other | On 5/1/07 the dental office contacted the manufacturer via telephone regarding the incident. On 5/2/07 the firm's service representative visited the dental office and replaced the lights. |  | 4 | 4 dual track lights | The lights were distributed to one distributor in NY who in turn distributed the lights to one dental office. | Terminated | July 12, 2007 | 72 | 4 | Undetermined
Z-0989-2007 | 38339 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GEHC OEC 9900 ESP-15 Elite C-Arm fluoroscopic X-ray system, GE Healthcare, Salt Lake City, UT. |  |  | July 06, 2007 | 2007 | GE OEC Medical Systems, Inc | The Air Kerma Rate displayed value and cumulative air kerma displayed value could exceed the allowable error under certain imaging conditions. | Other | Consignees were notified by certified letter on 07/06/2007. They were told that a service rep would contact them to provide a label and an addendum to their operator's manual. A permanent software solution is under development and will be provided when it becomes available. |  | 22 | 22 units | Worldwide, including USA, France, Singapore, Sweden, and Switzerland. | Terminated | February 01, 2013 | 2037 | 26 | Undetermined
Z-1074-2007 | 38048 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Medical Systems Clinac Medical Linear Accelerator. Clinac 600 C &C;/D, Clinac 2100 C, C/D &amp; EX and Clinac 2300 C &C;/D |  |  | May 02, 2007 | 2007 | Varian Medical Systems Inc | The wedge tray may slip out of the Clinac interface mount during gantry rotation. | Process control | The firm sent out a Product Notification Letter on 5/2/07 to customers who have ordered this guide block through a Varian Service Center. A Varian Field Engineer will visually inspect guide blocks at-site that have been installed from Service stock for all affected customers, and replace if necessary. |  | 209 | 209 units | Product was distributed and manufactured between 1999 and 2006. Distributed worldwide, including USA, Ecuador, Columbia, Taiwan, Austrailia, Austria, Brazil, Canada, Chile, China, Finland, France, Germany, India, Italy, Japan, New Zealand, Poland, Portugal, Serbia and Montenegro, Singapore, Spain, Switzerland, UAE, and UK. | Terminated | November 27, 2010 | 1305 | 209 | Undetermined
Z-1071-2007 | 37921 | No | JIT | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | NEFA Standard Solution Calibrator, Oleic Acid, 1 mmol/L, 4x10mL bottles per carton, Cat. 276-76491, labeled in part ***For in vitro diagnostic use ***Distributed by: Wako Chemicals, USA, Inc. 1600 Bellwood Road Richmond, VA 23237*** |  |  | March 30, 2007 | 2007 | Wako Chemicals USA, Inc. | Diagnostic reagent for clinical laboratory testing was incorrectly labeled. | Other | The recalling firm notified consignees by letter 03/30/07 and advised of labeling error. Consignees were advised that they could return the product for relabeling or affix correct label enclosed with the notification. Distributors were requested to forward notification to end users. The firm subsequently notified consignees on 04/26/07 by fax that product should be returned for label correction. |  | 94 | 94 cartons | Worldwide, including USA, Germany and Canada. | Terminated | July 24, 2007 | 116 | 94 | Undetermined
Z-1075-2007 | 38062 | No | BZD | Noncontinuous ventilator (IPPB). | Anesthesiology | Anesthesiology | 510(k) | AirLife nCPAP System Driver; a nasal continuous positive airway pressure device, which is part of a larger system that includes an nCPAP generator, fixation device and patient interface (prongs and.or mask); Cardinal Health, McGaw Park, IL 60085; cat. 006900 |  |  | June 01, 2007 | 2007 | Cardinal Health | Under isolated conditions, the device may experience issues with O2 Calibration, which may result in either a nuisance alarm or a delivered FiO2 of up to 5% below the specified tolerance. | Other | Cardinal Health sent the accounts recall letters dated 6/1/07 via UPS next day delivery, which listed the catalog number and affected serial numbers of the CPAP Driver, and informed them that under certain isolated conditions, the CPAP Driver may experience issues with O2 Calibration, which may result in either a nuisance alarm or a delivered FiO2 of up to 5% below the specified tolerance. Customers have experienced this under the following conditions: 1) air inlet pressure is greater than O2 pressure by 5 to 10 psi and 2) FiO2 settings less than 45%. Separate recall letters were sent to those accounts who had filed complaints of O2 Calibration Errors, stating that their device will be replaced by the end of June, 2007; those accounts who purchased the devices, but had not reported O2 Calibration Errors, stating that a Cardinal Health representative will upgrade the firmware on the devices at their facility by the end of August, 2007; and those accounts who had had demonstration units in their possession, advising them to notify their sales representative if any of the conditions mentioned in the letter occurred. The accounts were requested to complete and return via fax at 847-689-9101, the enclosed acknowledgment form, indicating the number of affected drivers in their possession, and were advised that they would be contacted by a Cardinal Health representative to arrange for the replacement/service of their drivers. Any questions were directed to Quality Systems-Professional Services at 1-800-292-9332 or 1-800-327-7503. |  | 48 | 48 units | Nationwide. | Terminated | December 12, 2007 | 194 | 48 | Undetermined
Z-1076-2007 | 38115 | joint, knee, external limb component | ISY | External limb prosthetic component. | Physical Medicine | Physical Medicine | 510(K) Exempt | Ossur Total Knee Model Number 2100, external knee prosthesis. |  |  | April 02, 2007 | 2007 | Ossur | Some units of the Total Knee Model 2100 produced from August 1, 2006 through October 4, 2006 may contain a faulty linkage component. | Other | On 4/2/07 a letter was mailed to all affected customers via certified mail, return receipt requested. All mailing will be followed by telephone calls from an Ossur Customer Service Representative. Customers were notified that Ossur will replace the affected knees with units carrying a two-year warranty from the date of replacement.Ossur will undertake a sub-recall by notifying the one affected distributor by letter on 4/2/07. Ossur will follow up with this distributor by telephone and request a list of consignees that the distributor sold the product to so that Ossur can directly contact each of those consignees. Upon receipt of this list, Ossur will then communicate the recall information directly to these consignees. |  | 44 | 44 | Nationwide. | Terminated | May 28, 2008 | 422 | 44 | Undetermined
Z-1070-2007 | 37920 | No | JHI | Human chorionic gonadotropin (HCG) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Total B-hCG Speciment Diluent is composed of 1 bottle of Total B-hCG Specimen Diluent containing animal serum. |  |  | April 10, 2007 | 2007 | Abbott Diagnostic International, Ltd. | An increase in background counts elevates the results of low level samples up to approximately 6mIU/mL. This is not enough to shift a negative result to positive on an undiluted sample. If low level samples are diluted then the background effect is magnified potentially yelding false positive results for pregnancy and or Dilution Inaccuracy. | Other | The 4/10/2007 Product Recall Letter was sent to all customers who received the affected lots. |  | 1853 | 1853 | Worldwide, including USA, PR, USVI, CANADA, MEXICO, GUATEMALA, COSTA RICA, COLOMBIA, ECUADOR, PERU, CHILE, URUGUAY, ARGENTINA, BRAZIL, DOMINICAN REPUBLIC, EL SALVADOR, GERMANY, JAPAN, SINGAPORE, HONG KONG, TAIWAN, THAILAND, SOUTH KOREA, AUSTRALIA, NEW ZEALAND, HONDURAS, TRINIDAD &amp; TOBAGO, GRENADA, ST. VINCENT, ST. LUCIA, ANTIGUA, BARBADOS, CURACAO, BRITISH VIRGIN ISLANDS, BAHAMAS, PANAMA, and CAYMAN ISLANDS. | Terminated | February 14, 2012 | 1771 | 30275 | Undetermined
Z-0992-2007 | 38052 | No | DXE | Embolectomy catheter. | Cardiovascular | Cardiovascular | 510(k) | Amplatz Goose Neck Snare Kit, 6 Fr, Lot 1650307, REF: GN3000, Use By 2009-09-01, Sterile EO |  |  | May 18, 2007 | 2007 | Ev3, Inc | Amplatz Goose Neck Snare Kit's catheter contained in the kit may be 4 Fr instead of the intended 6 Fr. | Other | Letter May 18, 2007. |  | 23 | 23 units | Worldwide, including USA, Belgium, France, Germany, Italy, Jordan, Spain, United Kingdom, and Yemen. | Terminated | October 17, 2008 | 518 | 23 | Undetermined
Z-1008-2007 | 37977 | No | HBB | Pneumatic cranial drill motor. | Neurology | Neurology | 510(k) | Anspach Single Use, Sterile Bone Cutting Burrs |  |  | May 18, 2007 | 2007 | The Anspach Effort, Inc. | The white film portion of the sterile pouch may become brittle prior to the expiry date compromising the sterility of the product. | Nonconforming Material/Component | The firm issued a recall letter on May 18, 2007, to its consignee, advising them of the recall. Consignees were asked to screen their inventory for the recalled devices and complete an accompanying product replacement form. The firm will replace all returned product within 30 days. |  | 1962123 | 1,962,123 | World wide-Distributed throughout the United States and to Puerto Rico, Canada, Mexico, Japan, Lithuania, Cyprus, Austrailia, England, Sweden, Bahamas, Korea, Germany, Switzerland, Malaysia, Iceland, Portugal, Singapore, Poland, Hungary, Thailand, Austria, Israel, Belgium, Egypt, South Africa, Netherlands, India, Syria, Kenya, Italy, Czech Republic, Jamaica, Columbia, Argentina, Honduras, Saudi Arabia, Greece, Finland, Taiwan, Republic of China, Lebanon, Denmark, France, Hong Kong, Philippines, Venezuela, Argintina, and the United Kingdom. | Terminated | January 17, 2009 | 610 | 1962123 | Undetermined
Z-1023-2007 | 37887 | No | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | AxSYM Troponin-I ADV Reagent Pack; list 02J44; 100 test pack; in-vitro diagnostic; Abbott Laboratories, Abbott Park, IL 60064 |  |  | April 23, 2007 | 2007 | Abbott Laboratories | Failure of the AxSYM instrument to open (actuate) some reagent packs from certain lots. The defects can result in probe crashes if not detected prior to placement of the reagent kit on the AxYM instrument. | Other | Abbott initiated a field correction recall by letter dated 4/23/07 to all customers receiving the affected reagents. The accounts were informed of the failure of the AxSYM to open (actuate) some reagent packs with the affected lots. The accounts were requested to inspect each reagent pack for broken or improperly sealed caps; manually open each bottle to ensure the flipper bar raises the reagent cap correctly; close each flipper bar prior to loading the reagent pack on the instrument; and discard any reagent packs that do not open properly. If the user observes a probe crash, prior to testing samples again, the user must perform a probe crash recovery following the procedure in the Operations Manual. If necessary, they need to replace the probe. If a probe crash recovery is not performed, subsequent results may be impacted. The customers were requested to retain a copy of the letter for their records. Customers requiring additional information were directed to contact Abbott Customer Service at 1-877-422-2688. |  | 9277 | 9,277 packs | Worldwide, including USA, Puerto Rico, Argentina, Australia, Brazil, Canada, Cayman Islands, Chile, Colombia, Ecuador, Germany, Hong Kong, Mexico, New Zealand, Singapore, South Korea, Taiwan, Thailand and Venezuela. | Terminated | May 18, 2008 | 391 | 28061 | Undetermined
Z-1054-2007 | 37769 | Yes | NIK |  |  | Cardiovascular | PMA | VITALITY DR HE Implantable Cardioverter Defibrillator (model T180). Sterile EO. Guidant Corporation, Cardiac Rhythm Management, 4100 Hamline Avenue North, St. Paul, MN 55112-5798 USA. This ICD is designed to detect and terminate ventricular tachycardia (VT) and ventricular fibrillation (VF) and provide bradycardia therapy (atrial and ventricular pacing). Therapies include both low- and high-eneergy shocks using either a biphasic or monophasic waveform. Bradycardia pacing, including adaptive-rate features, is available to detect and treat bradyarrhythmias and to support the cardiac rhythm after defibrillation therapy. The device offers dual-chamber bradycardia features (atrial and/or ventricular pacing and sensing). |  |  | March 23, 2007 | 2007 | Boston Scientific CRM Corp | Boston Scientific has recently identified low-voltage capacitors from a former supplier that may be subject to degradation at a low rate of occurrence. These capacitors may cause accelerated battery depletion and may reduce the time between elective replacement indicator (ERI) and end of life (EOL) to less than 3 months. | Other | An Urgent Medical Devivce Information letter dated April 05, 2007 was sent to implanting and following physicians. The letter describes the issue, identifies patients with affected devices and provides recommendations. |  | 3041 | 3,041 | Worldwide, including USA, Argentina, Aruba, Australia, Austria, Barbados, Belgium, Bermuda, Brazil, Canada, Chile, Costa Rica, Cyprus, Czech Republic, Czechoslovakia, Denmark, Dominican Republic, Ecuador, Egypt, Finland, France, Germany, Greece, Guam, Hong Kong, Hungary, Iceland, India, Iran, Ireland, Israel, Italy, Japan, Jordan, Korea, Kuwait, Lebanon, Luxembourg, Malaysia, Mexico, Netherlands, New Zealand, Norway, Pakistan, Polynesia, Portugal, Puerto Rico, Qatar, Saint Lucia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Syrian Arab Republic, Taiwan R O C, Thailand, Tunisia, Turkey, U.S. Virgin Islands, United Arab Emirates, United Kingdom, United States, Uruguay, Venezuela, and West Indies. | Terminated | March 02, 2008 | 345 | 73008 | Undetermined
Z-0885-2007 | 37996 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Terumo TenderFlow Pediatric Arterial Cannulae, 5.3mm (16 Fr) OD, with 1/4 inch non-vented connector, 9 inch (22.9 cm) long; Catalog no. 813578. |  |  | May 22, 2007 | 2007 | Terumo Cardiovascular Systems Corp | The bond may fail and a leak or separation of the device occur at the distal end (where this device connects to the arterial line). | Other | Consignees were notified via phone beginning on 5/22/07 and by follow-up letters sent between 5/23/07 and 6/14/07. |  | 39 | 39 | Nationwide. | Terminated | July 07, 2009 | 777 | 912 | Undetermined
Z-0888-2007 | 37870 | No | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Patrol enteral pump set with piercing pin, list #520540. |  |  | May 02, 2007 | 2007 | Ross Products Division Abbott Laboratories | The enteral pump set may have a 1-3 mm slit in the silicone tubing located between the blue connector and the stretch band Safe-T-Valve. | Other | The recalling firm sent a recall letter dated April 27, 2007 to all customers. |  | 1042050 | 1,042,050 units | Nationwide | Terminated | March 09, 2012 | 1773 | 1042050 | Undetermined
Z-1043-2007 | 37733 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Innova 2100 IQ Digital Fluoroscopic Imaging System (Versatile cardiac and vascular). |  |  | March 26, 2007 | 2007 | General Electric Med Systems LLC | Caution on the non-detection of small invasive objects/devices, which project out from the skin surfaces (e.g. biopsy or other needles) while using the Innova Sense software option designed for patient contouring. Usage of Innova Sense patient contouring for such devices could potentially lead to patient injury. | Other | A "GE Healthcare Urgent: Safety Notice" dated January 24, 2007 was sent to consignees describing the product, problem, safety instructions, product correction and contact information. |  | 205 | 205 | Worldwide, including USA, Guam, Canada, Australia, China, India, Japan, Korea, Malaysia, New Zealand, Algeria, Czech Republic, Denmark, Finland, France, Germany, Greece, Hungary, Italy, Jordan, Morocco, Norway, Poland, Russia, Saudi Arabia, Spain, Sweden, Switzerland, Syria, Turkey, and United Kingdom. | Terminated | October 21, 2008 | 575 | 373 | Undetermined
Z-1007-2007 | 37944 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | Medtronic MiniMed Paradigm Insulin Infusion Pumps for the following Paradigm models: MMT-511, MMT-512, MMT-712, MMT-515, MMT-715, MMT-522 and MMT-722 |  |  | April 27, 2007 | 2007 | Medtronic MiniMed | Exposure to Magnetic Resonance Imaging (MRI) has resulted in damage to the component that monitors and controls movement of the motor in the MiniMed Paradigm insulin infusion pump. Although there were alarms as a result of the damage, some users cleared these alarms and continued using the pump. Under such conditions, the pump will significantly over-deliver; potentially causing severe hypoglycemi | Other | Firm mailed Medical Device Safety letters on April 24, 2007, to healthcare professionals, existing pump users and MRI facilities reiterating the existing warning in the pump User Guide to avoid exposing the pump to MRI (or similar high strength electromagnetic fields) and strengthening previous warning by specifically mentioning the potential for over-delivery and severe hypoglycemia. Firm is also including an insert with this information with any Paradigm infusion pumps shipped to new customers. |  | 334000 | 334,000 Paradigm insulin pumps (approx) | Worldwide: USA, US Virgin Islands, Antigua, Australia, Austria, Bahamas, Bahrain, Bermuda, Brazil, Canada, Chile, China, Columbia, Czech Republic, Denmark, Dominican Republic, Egypt, France, Germany, Greece, Hong Kong, India, Ireland, Israel, Italy, Japan, Korea, Kuwait, Lebanon, Malaysia, Mexico, Netherlands, New Zealand, Norway, Panama, Poland, Puerto Rico, Qatar, Romania, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, United Arab Emirates, United Kingdom, Uruguay, Venezuela. | Terminated | September 19, 2012 | 1972 | 334000 | Undetermined
Z-1005-2007 | 37367 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Perfusion System 8000 Roller Pump, 1040V; Model 8000; Catalog No. 16411. (Not distributed within the United States). |  |  | July 01, 2004 | 2004 | Terumo Cardiovascular Systems Corp | Due to a circuit board malfunction the roller pumps may stop during use without displaying an error message, may stop during use with an overspeed error posted on the local display, or may fail to start/restart resulting in the pump(s) being non-functional. | Other | As the complaints were received (between 07/04 and 01/06), the firm visited the customer filing the complaint and repaired the pumps demonstrating the problem. The firm will correct other units as complaints are received. The firm modified its recall strategy and issued a Device Correction Letter dated 6/22/07 notifying all users of the problem |  | 728 | 728 pumps of all models | Worldwide: USA and Bangladesh, Belgium, Bolivia, Canada, Chile, Colombia, Costa Rica, Egypt, Equador, Hong Kong, India, Indonesia, Japan, Latin America, Malaysia, Nicaragua, Philippines, Saudi Arabia, Taipe, Taiwan, United Arab Emirates, Uruguay, and Viet Nam. | Terminated | December 12, 2007 | 1259 | 728 | Undetermined
Z-0993-2007 | 38116 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | Graft Preparation System, sterile, single use, disposable, Distributed by Biomet Biologics, Inc., Warsaw, Indiana; REF 800-0300. |  |  | June 11, 2007 | 2007 | Biomet, Inc. | The syringe assembly has an obstruction in the tube that prevents the plunger from being depressed. | Other | Consignees were notified by letter dated either 6/11/07 or 6/13/07 and requested to return the product. |  | 46 | 46 to U.S. and international | Worldwide, including USA, Chile, Italy, Spain and United Kingdom. | Terminated | December 05, 2007 | 177 | 51 | Undetermined
Z-0966-2007 | 37929 | No | ITI | Powered wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Invacare Storm TDX Power Wheelchair, Models: #TDX3-PS, TDX3-SE-PS, TDX4-PS. |  |  | March 19, 2007 | 2007 | Invacare Corporation | Software error may permit the chair to recline too far. | Other | The recalling firm sent a recall letter dated March 19, 2007. The recalling firm also posted information regarding the recall on their website. |  | 227 | 227 devices | Nationwide and Canada. | Terminated | September 27, 2012 | 2019 | 227 | Undetermined
Z-0991-2007 | 37968 | Yes | BTL | Powered emergency ventilator. | Anesthesiology | Anesthesiology | 510(k) | Pneupac VR1 Standard and VR1 Responder is a hand held portable, time cycled, gas powered, flow generator ventilatory resuscitators that are intended for emergency resuscitation by medical personnel, paramedics and ambulance technicians inside and outside hospital and for ventilation by medical personnel inside and outside hospital in emergency situations and for intra- and inter- hospital transport. |  |  | May 16, 2007 | 2007 | Smiths Medical PM, Inc. | Smiths Medical is conducting a recall of 93 Pneupac VR1 Emergency &amp; Transport Ventilators following the identification of a problem which may exist with the demand valve that could cause the ventilator to stop cycling &amp; to deliver continuous flow output to the patient. | Other | Distriblutors and end users were contacted by phone (May 10 - May 24, 2007) and end users were additionally sent a Safety Action Bulletin May 21, 2007. The Bulletin was used as a script. |  | 93 | 93 | Nationwide. | Terminated | October 10, 2008 | 513 | 93 | Undetermined
Z-0887-2007 | 37692 | No | MJK | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | Batteries used with the LIFEPAK 20 defibrillator/monitor |  |  | March 30, 2007 | 2007 | Medtronic Emergency Response Systems, Inc. | Defibrillator internal battery cell may contain a short which could reduce battery capacity and cause the device to shut down without warning while operating on backup (DC) battery power. | Other | Firm sent letters on 03/30/2007 dated 03/2007 via certified return receipt to consignees with affected LIFEPAK 20 devices. The firm advised consignees to use the Operator's Checklist in the Operating Instructions to perform daily user tests and function checks while operating the defibrillator on battery power. In addition consignees are advised to replace their battery every 1 to 2 years. |  | 3002 | 3002 | Distribution is worldwide. The device is designed to be used in healthcare provider facilities such as hospitals and clinic settings. | Terminated | May 04, 2011 | 1496 | 29420 | Undetermined
Z-1119-2007 | 38354 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Biomedicare Implantium Fixtures, Model numbers : FX3408, FX3410, FX3412, FX3414, FX3808, FX3810, FX3812, FX3814, FX4308, FX4310, FX4312, FX4314, FX4808, FX4810, FX4812, FX4814, FX4814W |  |  | June 29, 2007 | 2007 | Biomedicare, Inc. | Lack of validation for sterilization cycle: This recall was initiated because the validation of the sterilization process for these devices has not been completed. As a result, Biomedicare cannot assure the sterility of fixtures distributed to date. | Other | The firm initiated the recall notification with telephone calls to all consignees on 06/30/2007. Recall letters with attached Recall Response Form were mailed to all customers on 06/30/2007 by certified US mail. Customers were asked to examine their inventory and immediately remove all Implantium Fixtures from their inventory and quarantine until arrangements could be made for product return. |  | 8844 | 8,844 units | In US to AZ, CA, IL, TX &amp; WA. Internationally to Turkey | Terminated | July 25, 2012 | 1853 | 8844 | Undetermined
Z-1099-2007 | 37982 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Power Linear Cutter Reusable Digital Loading Units with Reloads. Common name - Implantable staple. Model number PLC60 |  |  | May 12, 2007 | 2007 | Power Medical Interventions, Inc | Malfunction - improper staple formation | Other | The field sales representatives were telephoned 5/14/2007, and instructed to return the affected devices from the customer sites. |  | 9 | 9 units | The products were shipped to medical facilities in DC, FL, GA, and NY. The product was also shipped to Germany | Terminated | August 29, 2007 | 109 | 9 | Undetermined
Z-1010-2007 | 38170 | holder, head, radiographic | IWY | Radiographic head holder. | Radiology | Radiology | 510(K) Exempt | BrainLAB Target Positioner for Leksell Headring used with the iPlan RT Dose (versions 3.0, 3.0.1 and 3.0.2) and BrainSCAN (version 5.31) Radiotherapy Treatment Planning Software, ; BrainLab AG, Ammerthalstrasse 8, 85551 Heimstetten, Germany; catalog/version number 40700-3A |  |  | June 05, 2007 | 2007 | Brainlab AG | Target Alignment Error : A mechanical component/software incompatibility caused by a specific combination of the BrainLAB Target Positioner version 40700-3A for Leksell headrings and BrainLAB planning software resulted in a 1.25 mm shift in target area alignment during radiosurgery treatment. | Other | BrainLAB immediately notified all customers with the BrainSCAN or iPlan RT Dose software in combination with the BrainLAB Target Positioner for Leksell Headring, via e-mail on June 5, 2007. The Urgent Medical Device Product Notification letter dated June 4, 2007 included in the e-mail informed the customers that both the BrainSCAN and iPlan RT Dose treatment planning software calculates the isocenter position printed on the Target Positioner Overlays shifted 1.25 mm in cranial direction from the intended position, when the BrainLAB Target Positioner for Leksell Headring Revision A is used. The customers were instructed to cease using the BrainSCAN or iPlan RT Dose software in combination with the BrainLAB Target Positioner for Leksell Headring for treatment planning and patient set-up. An updated target positioner will be provided to the customers to correct the problem, estimated to be available in September 2007. Any questions were directed to their local BrainLAB Customer Support Representative or CAPA Manager Markus Hofman at 49-89-99-15-68-44 or 1-800-597-5911 in the U.S. |  | 7 | 7 units | Ohio, Washington and France | Terminated | February 24, 2008 | 264 | 7 | Undetermined
Z-1101-2007 | 38075 | No | FTA | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Tri-Star Surgical Exam Light MR-16 ( Lamp Holder) used in the following: Model Numbers: TS2015W (wall mount) , TS2015M (mobile) , TS2015W (wall mount) , TS2035M (mobile) |  |  | February 13, 2007 | 2007 | Sunnex Inc. | Surgical and/or exam lamp holder may malfunction, discolor, melt and smolder. | Other | Sunnex notifiied distributors by letter March 9, 2007. Sunnex is providing replacement product to end users |  | 30 | 30 lamps | Nationwide | Terminated | January 06, 2009 | 693 | 376 | Undetermined
Z-1162-2007 | 38283 | No | EZB | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Cook PTFE Wire Guide with 3 cm flexible tip, to establish a tract for ureteral/renal access, .038", 145 cm; sterile; Order/Ref number 638413. |  |  | June 12, 2007 | 2007 | Cook Urological, Inc. | Lack of assurance of sterility, in that the product is labeled as sterile but the package may not have been sealed. | Process control | Consignees were notified via letter dated 6/12/07 and instructed to cease using the product and to inform the firm if they have product to return. |  | 100 | 100 | Nationwide and Canada. | Terminated | November 25, 2009 | 897 | 100 | Undetermined
Z-1161-2007 | 38117 | Yes | NIM |  |  | Cardiovascular | PMA | Boston Scientific NexStent Monorail 5F Carotid Stent and Monorail Delivery System, Manufactured by: Endotex International Systems, Inc. 10231 Bubb Road Cupertino, CA 95014 USA |  |  | May 01, 2007 | 2007 | Boston Scientific Corporation | Boston Scientific NexStent Carotid Stent device failed to deploy when the outer catheter (proximal outer and distal sheath) did not pull back and expose the self-expanding stent. | Other | On May 1, 2007, Boston Scientific initiated a Customer Notification worldwide. On May 21, 2007 Boston Scientific expanded this field action to a recall/removal status for Germany. Affected customers were notified in Germany on May 24, 2007. |  | 2487 | 2487 | Worldwide, including USA, France, Italy, Germany, Belgium, Cyprus, Austria, Spain, Finland, Netherlands, Luxembourg, Romania , Czech Republic, Poland, Israel, Hungary, New Zealand, and Switzerland. | Terminated | March 02, 2008 | 306 | 2487 | Undetermined
Z-1160-2007 | 38112 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific Impulse Flextrusion Shaft 5F FL4 16391-22 H74916391222, (5-Pack Carton Label), Angiographic Catheters. |  |  | June 04, 2007 | 2007 | Boston Scientific Corporation | Labels on the carton from two batches/lots of Impulse Angiographic Catheters may indicate a different curve style than the device that is actually in the carton. | Process control | Sales Force were contacted of the product and problem via voice mail and a memo dated June 4, 2007. Customers were contacted via a letter on June 4, 2007. |  | 233 | 233 | Worldwide, including USA, Austria, Australia, Belgium, Chile, Germany, Spain, Finland, France, Great Britain, Greece, Ireland, Panama, Poland, Sweden, Thailand, and South Africa. | Terminated | December 19, 2007 | 198 | 468 | Undetermined
Z-1157-2007 | 38005 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | The ClearGlide EVH small with Scissors. |  |  | May 07, 2007 | 2007 | Datascope Corporation | Datascope has experienced a higher than usual number of reports involving the Precision Bipolar Device. Reports include the locking or sticking of the buttons of the device which operate the jaws, particularly when a side load is placed on the shaft of the device. | Other | Letters were sent to all hospitals, (contact: Hospital Administrators) the week of May 7, 2007 (certified mail, Return Receipt requested) with Urgent Medical Device Field Correction. The letter requests that all potential users of the devices at the institution be mdae aware of the situation. The letter instructs clinicians what to do if they experience problems with locking buttons of the Precision Device and also if the device is clamped closed on a vein branch and cannot be opened.Institutions are also requested to sign and return a postal card acknowledging receipt of the Field Correct notification and Datascope provides telephone numbers for customer service. |  | 549 | 549 | Worldwide, including USA, Belgium, the Netherlands, and Germany. | Terminated | March 04, 2008 | 302 | 1671 | Undetermined
Z-1073-2007 | 37999 | No | FFK | Electrohydraulic lithotriptor. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Gyrus CyberWand Ultrasonic Lithotripter System includes: |  |  | May 23, 2007 | 2007 | Gyrus ACMI Corporation | Sterilization Instructions not adequate for sterilizing. | Other | Gyrus has informed their sales force to visit each customer that has a CyberWand system(s) and hand deliver a letter, dated May 23, 2007, that explains the situation and requests that the customer cease using the unit and segregate it to prevent accidental use. The letter also explains that the firm is collecting any sterilization trays with printed sterilization instructions and replacing them. The rep will assist the customer in complying. The customer will also be required to complete a reply form to track compliance. |  | 53 | 53 units | Nationwide and Canada. | Terminated | May 25, 2010 | 1098 | 53 | Undetermined
Z-1072-2007 | 37952 | No | GGP | Factor deficiency test. | Hematology | Hematology | 510(k) | STA Liatest Free Protein S |  |  | May 08, 2007 | 2007 | Diagnostica Stago, Inc. | Specific lots of STA Liatest Free Protein S may result in low results in the assigned value ranges and sometimes outside the lower limit. | Other | Customers were sent a Product Recall Notification on May 8, 2007 by first class mail labeled " Recall notice, Open Immediately" |  | 365 | 365 units | Medical facilities nationwide | Terminated | October 13, 2010 | 1254 | 365 | Undetermined
Z-1143-2007 | 38226 | Yes | BTR | Tracheal tube. | Anesthesiology | Anesthesiology | 510(k) | Kimberly-Clark Microcuff Endotracheal Tube, Pediatric Oral/Nasal Magill, sizes 3.0mm, 3.5mm, 4.0mm, and 4.5mm, Rx Only, Sterile, Manufacturer: Kimberly-Clark NV Belgicastraat 13, 1930 Zaventem, Belgium, Made in Malaysia |  |  | June 04, 2007 | 2007 | Kimberly-Clark Corporation | The tubes may kink during use which may impede or prevent ventilation. | Other | Consignees were notified by letter sent via overnight mail on/about June 4, 2007. A follow-up letter was sent to consignees via fax on June 5, 2007. Consignees were instructed to cease distributing the product and return all affected product in inventory to Kimberly-Clark. Distributors were instructed to notify their accounts of the recall. A Product Recall Response Sheet (Packing Slip) was attached to the recall notice. |  | 4376 | 4376 cases (43760 units/total for both products) | Worldwide, including USA, Canada, Europe, Australia, Chile, Hong Kong, Singapore, Thailand, India, Malaysia and Saudi Arabia. | Terminated | June 12, 2008 | 374 | 4376 | Undetermined
Z-1149-2007 | 38289 | component, external, limb, ankle/foot | ISH | External limb prosthetic component. | Physical Medicine | Physical Medicine | 510(K) Exempt | VS-1000 Senator Prosthetic Foot, (Categories 5 and 6, sizes 22-30 cm, both Left and Right) Applied Composite Technology, Gunnison, UT. |  |  | June 20, 2007 | 2007 | Applied Composite Technology | Prosthetic foot device may loosen or separate during use; due to structural inconsistency that may cause a pyramid failure. | Other | On 06/20 and 06/21/2007 all recipients of the affected Senators were notified by telephone. They were asked to return all suspect prostheses for replacement. Follow-up emails and fax letters (6/21/07) were also sent. |  | 97 | 97 units | Worldwide Distribution- USA including states of CA, CO, FL, IL, MI, MN, MS, NC, NY, OR, PA, SC, TN, TX, VA, WA. VA facilities in CA and MI. Foreign distribution to Australia, Canada, Germany, Russia, South Africa, South Korea, and UK. | Terminated | June 05, 2008 | 351 | 97 | Undetermined
Z-1129-2007 | 38341 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Curlin Medical Administration Sets: Model Numbers 340-4111, IV Administration Sets. |  |  | February 09, 2007 | 2007 | Curlin Medical Llc | The Administration Sets may leak at the reservior bag. | Other | A Field Correction Advisory (Recall Letter) was issued on February 9, 2007 to the firm's one US Customer (distributor): B. Braun Medical Inc. recalling Curlin Medical, Inc. Administration Sets 340-4111- Lot # D628223 and D627616 because the affected lots do not meet acceptable performance levels.A second Field Correction Advisory (Recall Letter) was issued on March 15, 2007 to B. Braun Medical Inc. recalling Curlin Medical, Inc. Administration Sets: 340-4111- Lot # D628223, D6276l6 ,D620544, D630503; 340-4127- Lot# D620104; 340-4119- Lot# D623716 because the affected lots do not meet acceptable performance levels. |  | 25280 | 25, 280 | One US Customer (distributor): B. Braun Medical Inc. 901 Marcon Blvd. Allentown, PA 18109 | Terminated | May 28, 2008 | 474 | 25280 | Undetermined
Z-1128-2007 | 38239 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips Medical HeartStartXL Defibrillator/Monitor |  |  | June 11, 2007 | 2007 | Philips Medical Systems | The ECG signal may become unusable to assess patient condition and need for therapy. | Nonconforming Material/Component | Philips Medical contacted accounts by letter dated June 11, 2007. Users are advised to remove product from service, if possible. If not, a back up unit should be available to use. Service reps will contact accounts to replace units. |  | 312 | 312 units | Worldwide, including USA, Canada, Australia, Brazil, China, Czech Republic, France, Germany, Hong Kong, India, Ireland, Italy, Japan, Korea, Kuwait, Macau, Poland, Portugal, Romania, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, UK, and Venezuela. | Open, Classified |  |  | 312 | Undetermined
Z-1130-2007 | 38373 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Cleaning probe included as an accessory to aspirators that are sterile packed single use disposable accessories to the GE InstaTrak 3500 Plus x-ray system. Probes are packaged and shipped together with the following aspirators: Straight Aspirator - 1000370-NAV, 7 French Aspirator - 100936, Keane - Moses Extended Straight Aspirator - 1000748, 15 degree aspirator - 1003895, Fried 45 Degree Aspirator - 1000541, Kuhn 90 degree Aspirator - 1000543-NAV. |  |  | May 18, 2007 | 2007 | GE OEC Medical Systems, Inc | Cleaning probes used to unclog disposable aspirators may shed small particles of nylon material which may be transferred into the nasal cavity. | Other | Consignees were notified by certified mail on 05/18/2007 and were told to stop using and to dispose of the affected probes. |  | 67625 | 67,625 units | Worldwide, including USA, Canada and France. | Terminated | August 22, 2011 | 1557 | 67625 | Undetermined
Z-1127-2007 | 38234 | No | GKP | Coagulation instrument. | Hematology | Hematology | 510(k) | Instrumentation Laboratory ACL TOP Coagulation Analyzer (Base Model only). Blood analyzer. |  |  | June 19, 2007 | 2007 | Instrumentation Laboratory Co. | ACL TOP unit, when in emergency stop may release liquid during recovery sequence and contaminate sample tube. | Other | Instrumentation Laboratory notiifed the distributor Beckman Coulter on 6/19/2007. Beckman Coulter notified their customers and provided a copy with the Instructions for Recovery Procedure on how to avoid contamination from probe leaks. |  | 94 | 94 | Nationwide. | Terminated | May 23, 2012 | 1800 | 94 | Undetermined
Z-1151-2007 | 38121 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Biomet Straight Magnum Inserter Handle, a preparatory instrument used for acetabular component insertion; REF # or Catalog Number: S313141. (systems affected by this action include M2a-Magnum, Regenerex, Acetabular Revision Shell and M2a-38) |  |  | June 14, 2007 | 2007 | Biomet, Inc. | The instrument was manufactured incorrectly and use could result in the instrument binding to the acetabular component and cause implant to not release following inpaction. | Other | On 6/13/07, the firm sent a Urgent Medica Device Correction Notice to its sales force instructing them to retrieve affected units from their hospital accounts and to return them to the firm. International accounts were notified via email on 6/14/07. |  | 94 | 94 | World wide: USA, Australia, and Finland. | Terminated | December 12, 2007 | 181 | 94 | Undetermined
Z-1137-2007 | 37878 | No | LKH | Chlamydia serological reagents. | Microbiology | Microbiology | 510(k) | Femoral Speed Block Size 6, Part/Catalog #800-01-367; positioning device included as part of the M.I.K.A. (Minimally Invasive Knee Arthroplasty) surgical kit; distributed by Encore Medical, L.P., Austin, TX 78758. |  |  | August 15, 2006 | 2006 | Encore Medical, Lp | Complaints received that device used to prepare femur to receive femoral implant would split during surgery. | Other | Firm sent notification letter to consignees on 08/25/2006, along with replacement speedblocks. Consignees asked to trade out old speedblocks with the replacements. |  | 24 | 24 units. | Product distributed to direct accounts in USA and one consignee in Japan. | Terminated | March 27, 2008 | 590 | 24 | Undetermined
Z-1132-2007 | 37640 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | GE Medical Systems Mac-Lab System, cardiac catheterization data system. Software Version 6.5 2021423-015 Revision D GE Medical Systems, Information Technologies, Inc., 8200 West Tower Avenue, Milwaukee, WI 53223 USA. |  |  | February 16, 2007 | 2007 | General Electric Medical Systems Information Technology | GE Healthcare has identified conditions that could cause a potential loss of real-time monitoring with the CardioLab/Mac-Lab/ComboLab Electrophysiology and Hemodynamic Monitoring Systems. These issues are associated with an audio driver, and the use of the application's log and imaging windows. | Other | Consignees were sent a letter dated 2/16/07. The letter described the product and problem, made recommendations and gave information on software update. |  | 1513 | 1513 (1067 US and 446 OUS) | Worldwide, including USA, Argentina, Australia, Bangladesh, Belgium, Brazil, Canada, Chile, China, Czech Republic, Denmark, Egypt, France, Germany, Greece, Hong Kong, India, Indonesia, Ireland, Israel, Italy, Japan, Republic of Korea, Kuwait, Malaysia, Mexico, Netherlands, New Zealand, Nicaragua, Pakistan, Panama, Puerto Rico, Saudi Arabia, Singapore, South Africa, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Turkey, United Kingdom, and Venezuela. | Terminated | December 11, 2011 | 1759 | 2909 | Undetermined
Z-1148-2007 | 38232 | No | MKT |  |  | Microbiology | PMA | biphasic LIFEPAK 500 automated external defibrillators (AEDs) with software version 4.4 or lower; Part Nos. 3011790-XXXXXX, D3100790-XXXXXX, and U3011790-XXXXXX . (XXXX indicates various non-contiguous numbers) |  |  | June 13, 2007 | 2007 | Medtronic Emergency Response Systems, Inc. | Reduced Shock- the device may deliver 100 Joule (J) to a patient, which is less than the recommended factory minimum default setting of 200 J. | Other | On June 13, 2007 the recalling firm sent a letter dated June 11, 2007, titled URGENT - MEDICAL DEVICE CORRECTIVE ACTION. The letter advises of the situation, recommends to the consignee to set defibrillator to FIXED Sequence Protocol and to contact the firm for a FLEXIBLE Sequence Energy Protocol. The letter provides an instruction sheet and a response sheet to be faxed back to the company. |  | 59996 | 59,996 devices | Worldwide | Terminated | March 20, 2011 | 1376 | 59996 | Undetermined
Z-1078-2007 | 38215 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris Medical Systems IV Administration Set, Low Sorbing Extension Set, Model C20014, Rx Only, Alaris Medical Systems, Lot Number 07036013 |  |  | June 05, 2007 | 2007 | Cardinal Health 303 Inc DBA Alaris Products | The sterility of the product may be compromised as the packaging used for the sterility barrier was not designed for the sterility process used by the firm. | Other | Recall Notification consists of a phone call and Recall Notification letters to consignees (distributors) as well as direct customers. On June 5, 2007, Cardinal Health began calling customers and distributors to inform them of a recall affecting IV Administration Set, Model C20014, Lot Number 07036013 (Low Sorbing Extension Set). A follow-up Recall Notification Letter was sent on June 11, 2007.The customers and distributors were notified by registered return receipt mail. For direct customers, letters were sent directly to the point of contact at the hospital. For distributors, they were sent multiple copies of the Recall Notification; along with a cover letter outlining their responsibilities to complete the Recall plan and notify their customers as requested. |  | 4248 | 4,248 sets (59 boxes) | Nationwide. | Terminated | February 10, 2012 | 1711 | 4248 | Undetermined
Z-1021-2007 | 38230 | No | HNO | Keratome. | Ophthalmic | Ophthalmic | 510(k) | Bausch &amp; Lomb Zyoptix XP Epi Separator, packaged 1 per sterile pouch, 10 per carton, Reorder No. 507-0068, RX, for single use only. Ophthalmic surgical accessory. The product is labeled as Manufactured by: Bausch &amp; Lomb Incorporated, St. Louis, MO. |  |  | June 08, 2007 | 2007 | Bausch &amp; Lomb Inc | Reports have been received of stromal incursions after use of the product. | Other | The firm notified their customers via telephone between 6/8-12/07. During the call, the customer was informed reports had been received of stromal incursions and they were to quit using the product and return it. |  | 243 | 243/10-pouch cartons | Distribution was limited to "key" physicians in MN, SC, MD, and NJ. Foreign distribution was made to the firm's entities in The Netherlands, Singapore, Mexico, Canada, Hong Kong, and India. | Terminated | April 28, 2009 | 690 | 544 | Undetermined
Z-1020-2007 | 38225 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | All IV Administration Infusion Sets intended for use with the Gemini Infusion Pump and Alaris Pump module (a.k.a. Medley Pump module) including: Alaris SmartSite Low Sorbing Set, Alaris Vented Syringe Adapter Set, Alaris Burette Set, Cardinal Health SmartSite Infusion Set, Alaris Infusion Set, Alaris SmartSite Burette Set, Cardinal Health SmartSite Burette Set, Alaris SmartSite Administration Set, Alaris Infusion Burette Set, Alaris SmartSite Blood Set, Cardinal Health SmartSite Combination Infusion Set, Cardinal Health Blood Set, Alaris Medley/Gemini Infusion Set, Alaris Medley/Gemini Blood Set, Alaris Latex-Free Infusion Set, Alaris Light Sensitive Solution Set, Alaris Medley/Gemini Enteral Set, Imed Gemini 20 Vented/Nonvented Primary Administration Set, Imed Vented/Nonvented Gemini Prime-Saver Set, Imed Vented/Nonvented Gemini Short Set, Alaris Half Set, Cardinal Health Half Set, Alaris Latex-Free Burette Set, Alaris Opaque Set, Alaris Medley/Gemini Combination Infusion Set |  |  | June 18, 2007 | 2007 | Cardinal Health 303 Inc DBA Alaris Products | It has been determined that with low probability, over-infusion may occur as a result of an open safety clamp fitment (a.k.a. Flo-Stop fitment) on the pumping segment, whether opened intentionally in the course of expected clinical practice, or as a result of inadvertent action. The over-infusion occurrence is dependent on the user not closing the roller clamp first, as required by clinical practi | Other | The field action will consist of a mailing and will be conducted through consignees (distributors) in the affected countries as well as direct customers for all IV Administration sets intended for use with Gemini Infusion Pumps and Alaris/Medley Pump modules. The customers and distributors will be notified by registered return receipt mail. For direct customers, separate letters will be sent to the Director of Nursing, Director of Materials and Director of Risk Management at each facility. Distributors will be sent multiple copies of the Safety Alert Notification and Tip Sheets; along with a cover letter outlining their responsibilities to complete the Safety Alert plan and notify their customers as requested. The domestic letters (direct customers and distributors) were mailed beginning on June 18, 2007 to all of the affected accounts via US Postal Service Certified Mail. |  | 2000000 | &gt;2,000,000 units produced annually | Worldwide including USA, Argentina, Australia, Bahrain, Belgium, Canada, China, Cyprus, Czech Republic, Denmark, Finland, France, Germany, Hong Kong, India, Indonesia, Ireland, Israel, Italy, Japan, Korea, Kuwait, Lebanon, Malta, Mexico, Myanmar, Netherlands, New Zealand, Norway, Oman, Paraguay, Peru, Philippines, Portugal, Puerto Rico, Qatar, Saudi Arabia, Serbia/Montenegro, Singapore, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Trinidad &Tobago;, United Kingdom, United Arab Emirates, Uruguay &amp; Venezuela. | Terminated | March 28, 2012 | 1745 | 2000000 | Undetermined
Z-1106-2007 | 37963 | No | JJF | Micro chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Amplilink Software versions 1.1, 1.3, 1.4, 2.41. |  |  | May 21, 2007 | 2007 | Roche Molecular Systems, Inc. | For select COBAS AMPLICOR Tests run on the COBAS AMPLICOR Analyzer in conjunction with AMPLILINK software, a discrepancy has been identified between the onboard working reagent stability information reported by AMPLILINK Software (versions 1.1, 1.3, 1.4 and 2.41) and the stability information provided in the test kit package inserts/method manual. | Other | Urgent Medical Device Correction letters were mailed on 5/21/07 by first class mail. Letters describe the discrepancy between software reagent expiration dates and package insert expiration dates. |  | 5806 | 5,806 total | Nationwide to hospitals and laboratories. | Terminated | February 18, 2010 | 1004 | 5806 | Undetermined
Z-1105-2007 | 37614 | No | DWQ | External vein stripper. | Cardiovascular | Cardiovascular | 510(k) | Smith &amp; Nephew TRIVEX System Pneumatic Footswitch, Part #7209791; distributed by Smith &amp; Nephew, Inc. - Endoscopy, Andover, MA 01810. |  |  | October 12, 2004 | 2004 | Smith &amp; Nephew, Inc Endoscopy Division | Build-up of static charge in Pneumatic Footswitch tubing may cause control unit to reset to default settings in the operating room during surgery. | Other | Firm sent notification letters to consignees via overnight FedEx on 10/12-13/04. International customers were notified via letter on 11/10/04. Customers were asked to return footswitches for a replacement. |  | 274 | 274 units. | Product sent to distributors nationwide and to the following countries: Australia, Austria, Canada, China, Germany, Italy, Japan, and Taiwan. | Terminated | August 02, 2007 | 1024 | 274 | Undetermined
Z-1112-2007 | 38221 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Flo-Gard 6301 Dual Channel Volumetric Infusion Pump; Baxter Healthcare Corporation, Deerfield, IL 60015; Product code # 2M8064 |  |  | June 21, 2007 | 2007 | Baxter Healthcare Corp. | Falsified Repair/Test and Inspection records: Pumps submitted for repair or service may not have been properly repaired or serviced before being returned to the customer. | Other | Baxter telephoned the direct accounts starting on 6/21/07, informing them that functional verification testing may not have been performed on their Flo-Gard and/or Colleague infusion pumps before they were returned to them. Since Baxter has no evidence that this quality control testing was conducted on these pumps, the pumps may not function as designed. The accounts were requested to locate the affected pumps and take them out of service as soon as possible, and send them back to Baxter for the required functional verification testing.*** Baxter expanded the recall and telephoned the direct accounts starting on 7/30/07, informing them that the electrical safety testing may not have been performed on their Colleague infusion pumps before they were returned to them. Since Baxter has no evidence that this quality control testing was conducted on these pumps, the pumps may not function as designed. The accounts were requested to locate the affected pumps and take them out of service as soon as possible, and send them back to Baxter for the required electrical safety and all functional verification testing. Follow-up recall letters were sent to the accounts on 8/7/07 reiterating the telephone conversations and providing a list affected pump serial numbers to be removed from service and returned to Baxter for testing. They were also requested to report any malfunction or adverse event related to the Colleague or Flo-Gard Infusion Pumps to Baxter Corporate Product Surveillance at 1-800-437-5176. ****Recall 2 Extension******* Baxter expanded the recall for a second time via letter on 8/20/07 and telephone calls starting on 8/21/07, informing the direct accounts that Baxter's investigation found documentation inconsistencies by three more of their service technicians, representing the possibility of incomplete testing and/or documentation by these technicians who serviced Colleague infusion pumps at the account's facility. Without complete and properly do |  | 144 | 144 pumps | Nationwide: Alaska, Arizona, California, Colorado, Florida, Georgia, Hawaii, Idaho, Illinois, Iowa, Kansas, Kentucky, Louisiana, Maryland, Michigan, Mississippi, Montana, Nebraska, Nevada, New Jersey, New Mexico, New York, North Dakota, Ohio, Oklahoma, Oregon, Pennsylvania, South Carolina, Tennessee, Texas, Utah, Washington and Wyoming | Terminated | March 16, 2010 | 999 | 2013 | Undetermined
Z-1108-2007 | 38325 | No | DSQ |  |  | Cardiovascular | PMA | Thoratec TLC-II Portable VAD Driver. Component (pneumatic) driver of Thoratec Ventricular Assist Device system |  |  | June 08, 2007 | 2007 | Thoratec Corp | Component Failure; Device Failure without warning; due to premature wear-out of the compressor motor, VAD support for the patient's circulatory system may fail without warning. | Nonconforming Material/Component | Urgent Medical Device Recall letters, dated June 08, 2007, were issued to consignees, with instructions to take the following actions: 1) check the current number of service hours on all units. 2) do not use TLC-II drivers exceeding 1500 hours until they are serviced by thoratec. 3) note the indicated hours on the product inventory form and return the form to Thoratec. 4) ensure that all patients have a back up driver and hand pumps at all times and have been trained in backup procedures. The firm is also arranging for the return and servicing of all affected drivers, with highest priority to those drivers already exceeding 1500 hours of life. A July 19, 2007, Press Release was issued by the firm. |  | 378 | 378 drivers | Worldwide: USA including states of AL, AR, AZ, CA, CO, CT, FL, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, NC, NY, OH, OR, PA, TN, TX, UT, VA, and WA, as well as, countries of Australia, Austria, Belgium, Canada, Czech Republic, Denmark, France, Germany, Ireland, Israel, Italy, Netherlands, Qatar, Singapore, Switzerland, Taiwan, UAE, and UK. | Terminated | July 17, 2008 | 405 | 378 | Undetermined
Z-1110-2007 | 37273 | Yes | DTZ | Cardiopulmonary bypass oxygenator. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Intersept Filtered Cardiotomy Reservoir, Model 1351, Non-pyrogenic, Sterile, Manufacturer: Medtronic, Inc., Minneapolis, MN |  |  | April 18, 2007 | 2007 | Medtronic Perfusion Systems | Pyrogen Levels detected above specification: The product contains a pyrogen level above Medtronic''s internal specifications while the product is labeled as nonpyrogenic. | Other | The consignees were notified by personal contact and requested to stop using the devices in question. A customer letter (April 24, 2007) will be sent by e-mail which explains the situation and requests that the impacted product be returned to Medtronic for disposal. The customer will be requested to complete a Recall Certificate acknowledging the receipt of this information. |  | 11 | 11 units | Nationwide: Distributed to states of CO and PA | Terminated | November 01, 2007 | 197 | 11 | Undetermined
Z-1104-2007 | 38313 | No | MID | Multiple autoantibodies immunological test system. | Immunology | Immunology | 510(k) | american diagnostica inc. IgG Conjugate, 15mL; component of the american diagnostica, inc. IMUCLONEÂż aPL IgG ELISA Kit, Product #649G; manufactured by Louisville APL Diagnostics, Inc., Seabrook, TX 77586. |  |  | August 10, 2004 | 2004 | Louisville Apl Diagnostics Inc | Product mislabeled; conjugate IgG placed in bottles labeled as conjugate IgM and vice versa. | Other | Firm notified consignee of recall via phone on 08/10/04. Consignee asked to return conjugates for properly labeled replacements. |  | 45 | 45 | Product distributed to a single direct account in CT. | Terminated | November 27, 2007 | 1204 | 80 | Undetermined
Z-1086-2007 | 38267 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | SPI EASY Abutment. Endosseous dental implant abutment. |  |  | April 19, 2007 | 2007 | Thommen Medical, AG | The undercut on the SPI EASY abutments which engages the stabilization ring of the implant is missing. Although the abutments can still be placed on the implant by the user the parts do not fulfill the quality expectations of the firm. | Other | Letter 5/14/2007. Customers are requested to check their current inventory for the recall product. If found they are to return the product for replacement. If they have implanted the product they are to check patient x-rays for the defect. If defects are found they are to do a restoration. |  | 84 | 84 | Nationwide. | Terminated | May 12, 2008 | 389 | 84 | Undetermined
Z-1102-2007 | 38206 | No | NQI | Immunomagnetic circulating cancer cell selection and enumeration system. | Immunology | Pathology | 510(k) | Cell Search Circulating Tumor Cell Kit (Epithelial) |  |  | April 17, 2007 | 2007 | Veridex, LLC | An increased incidence of aggregates was observed in Lot 0036 and 0036A. The presence of aggregates will cause customers to invalidate patient results. | Other | All Direct customers and customers of distributors received a customer notification informing them of the increased incidence of aggregates in cartridges on April 30, 2007 by Fed Ex overnight. This letter instructed them to discontinue use and to discard any remaining inventory. This was followed by a notification on May 24, 2007 by Fed Ex overnight to all US Distributors which also requested that they immediately stop using and/or distributing kits and to destroy any remaining inventory. Distributors were also asked to notifiy customers who received product. All foreign affiliate consignees were also informed on this date by e-mail and instructed to notify their customers. |  | 665 | Lot 0036: 468 kits; Lot 0036A: 197 kits | The kits were distributed to hospitals and laboratories nationwide. They were also sent to 7 J&J; affiliated companies in France, Japan, Canada, England, Germany, Italy and Spain. | Terminated | December 06, 2007 | 233 | 665 | Undetermined
Z-1080-2007 | 38261 | No | GDF | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Allez Medical Applications, Inc. 4Closure Surgical Fascia Closure System , Catalog ref. 4-10GPS. Sterile, One Time Use, Contents 4 sets 1- suture passer instrument, 1-guide rod, 10 mm, 250 mm long. |  |  | June 19, 2007 | 2007 | Synovis Surgical Inovation Div. of | There is a potential for the pouch to become damaged during shipping, compromising the sterility of the device inside. | Other | Consignees were sent a letter on June 19, 2007. The letter identified the product and problem. The letter also included a form for inventory purposes which is to be completed and faxed to Synovis. |  | 2 | 2 boxes (8 units) | Nationwide. | Terminated | May 08, 2012 | 1785 | 63 | Undetermined
Z-1098-2007 | 37961 | No | LOM |  |  | Microbiology | PMA | ARCHITECT AUSAB Controls list number 1L82-10 have two bottles with the following components: 1) Negative Control and 2) Positive Control. in vitro diagnostic. |  |  | May 14, 2007 | 2007 | Abbott Diagnostic International, Ltd. | During an internal study, the Architect AUSAB assay was found to read preparations of the WHO standard up to 53% higher than the WHO assigned value. This has the potential to cause elevated results and may impact the interpretation for a small percentage of results with grayzone or low-level reactive concentrations and higher proficiency sample results. | Other | All ARCHITECT AUSAB U.S. customers will be contacted. New customers will be contacted through a Product Information Letter that will be included in all ARCHITECT AUSAB Reagent Kits until assay is restandardized. All ARCHITECT AUSAB U.S. customers will be contacted. The following communication types were used: Product Correction Letter Customer Reply Form &gt; Product Information Letter to be included in new kits to be distributed. Letters were sent to customers by Federal Express Priority mail. The Customers were instructed to: Please follow their laboratory procedures for evaluation of these results Disregard the automated interpretation of results for low-level reactive samples.A customer reply form will accompany the Product Correction letter. Effectiveness Check Depth: 100% of customers receiving the Product Correction letter. The start date of communication was May 14,2007. |  | 393 | 393 | Nationwide and Puerto Rico | Terminated | June 10, 2011 | 1488 | 2134 | Undetermined
Z-0988-2007 | 37954 | No | MWM |  |  | Obstetrics/Gynecology | PMA | MediSpectra Luma Cervical Imaging System |  |  | May 15, 2007 | 2007 | Medispectra Inc. | Power supply may fail and render the unit inoperable. | Other | MediSpectra notified users by telephone on 5/15/07 and follow-up letter. Users are requested to select a date for field servce to replace the power supply. |  | 8 | 8 units | AZ, CA, FL, &amp; NY. | Terminated | July 09, 2007 | 55 | 8 | Undetermined
Z-0717-2007 | 38248 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Kodak Direct View DR3000 System, Stationary X-ray System, Catalog #s 655-1345, 655-1386. |  |  | March 06, 2007 | 2007 | Carestream Health, Inc. | The x-ray systems do not indicate the field size when the system's detector is turned off and the collimator is in manual mode. This results in the systems being noncompliant with the field indication performance standards of section 1020.31(e)(2)&amp; (3) of 21Code of Federal Regulations (CFR). | Other | On 3/06/07, Eastman Kodak notified CDRH of its proposed Corrective Action Plan (CAP). The CAP was approved by CDRH under the following conditions: (1) Kodak will visit each user facility, and apply dial indicator labels to the collimator knobs, apply a field size label to the collimator housing, and verify that each system complies with the field indication performance standard. (2) Kodak will prepare adequate instructions for the assembly/installation and testing of the dial indicators labels and field size labels. These instructions are to be furnished to all owners of the Kodak Direct View DR3000 Systems. (3) Kodak will supplement the Kodak DR3000 system product report with a copy of the adequate instructions. (4) Kodak will perform 100% verification of the recall effectiveness. (5) Kodak will not charge for the corrections and will reimburse the owner for any out-of-pocket expenses. (6) Kodak will notify the manufacturer of the Ralco 302 ACS collimators, Ralco, that the collimator is noncompliant with the performance standard. |  | 9 | Domestic: 7 units, Canada: 2 units | FL, NY, MO, MI, TX, and Canada | Terminated | July 02, 2007 | 118 | 9 | Undetermined
Z-0820-2007 | 37873 | No | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | VITROSÂż Immunodiagnostic Products Troponin I Reagent Pack, REF/Catalog No. 194 9882. A Troponin Immunoassay Method, IVD. Ortho-Clinical Diagnostics, Amersham UK |  |  | April 27, 2007 | 2007 | Ortho-Clinical Diagnostics | Quality control and patient results can be lower than expected when the true Troponin I value is less than or equal to 0.2 ng/mL which may result in the device failing to detect myocardial injury. | Nonconforming Material/Component | On 4/27/07, Ortho-Clinical Diagnostics in Rochester, NY, notified the end users of the recall of the product by phone, fax, and FedEx overnight letter. Clinical labs in possession of these lots have been contacted by OCD, and were instructed to (a) discontinue use of this product and (b) notify healthcare providers who ordered the test in recent weeks. Replacement product has been provided to clinical labs with the affected product lots. Clinical laboratories with questions may contact the company at 1-800-421-3311. On 4/27/07, OCD foreign affiliates were informed of this action by e-mail and instructed to inform their customers who received these reagent packs of this issue. Ortho-Clinical Diagnostics in Raritan, NJ, issued a nationwide press release on 5/04/07. |  | 5066 | Lot 3151: 1,750 units, Lot 3170: 3,316 units | Worldwide-including USA, Canada, and Bermuda | Terminated | September 07, 2012 | 1960 | 5066 | Undetermined
Z-0936-2007 | 37869 | Yes | MVK |  |  | Cardiovascular | PMA | Electrode Belt, component of the LifeVest Wearable Defibrillator, Product number 10A0889-A01. Specifically the 3 Therapy Electrode Pads attached to the Electrode Belt. |  |  | April 27, 2007 | 2007 | Zoll Lifecor Corporation | The electrodes do not properly release conductive gel | Other | Patients currently using the product should continue to do so while the product removal is underway. Product replacements will begin with the belts currently in the possession of the sales reps. By replacing these first any new patients will be receiving new or refurbished belts with new Therapy Electrodes that have not been exposed to methanol. As patients complete their use of the LifeVest and return their device to the firm the belt's refurbishment process will include replacement of all 3 Therapy Electrodes. The firm expects this process to be completed in approximately 3 months based on the average use of time and current production capabilities. The process may take an additional month for devices distributed overseas. Physicians with active patients were notified via letter sent certified mail on 5/3/07. Long term use patients (patients using the device longer than 3 months) will be notified of the product removal via telephone and follow up letter. Each long term patient will be provided with a replacement belt. They will be instructed to return their existing belt in the same shipping box. Due to the low residual risk no notification is planned for short term use patients. Prescribing physicians have the option to notify their patients if they choose to do so. |  | 944 | 944 belts | Worldwide: The product was shipped to patients in AZ, CA, CT, FL, GA, IL, IN, KY, LA, MA, MI, MN, MO, MT, NC, NE, NJ, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WI, and WV. The product was also shipped to Germany, United Kingdom, Italy, Spain, France, and Denmark. | Terminated | March 06, 2008 | 314 | 944 | Undetermined
Z-0960-2007 | 37671 | No | FTM | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Synovis Peri-Strips Dry with Veritas Collagen Matrix Circular Staple Line Reinforcement,( PSD 21-E-V), 21mm, Ethicon Circular stapler. Each package contains the following components: Cartridge and anvil assemblies, Cartridge cone, PSD Gel, &amp; C-clip. |  |  | March 16, 2007 | 2007 | Synovis Surgical Inovation Div. of | Package Integrity/Sterility Compromised: The foil pouch in the Peri-Strips Dry with Veritas Collagen Matrix Circular Staple Line Reinforcement has a potential for minute defects which may compromise the sterility of the outside of the tray in which the device is packaged. | Other | Recall notification letters were sent to hospitals via Federal Express on 03/16/2007. This letter describes the product being recalled and problem. It requests consignees to complete and fax back the product return form. EXPANDED Another recall letter dated 04/23/2007 was sent to inform customers of 17 aditional lot numbers being recalled. A third recall letter dated 05/05/2007 was set to hospitals to inform them of one additional Lot/Serial number that was affected. This letter also requests a product return form be completed and faxed to Synovis Surgical Innovations. |  | 1722 | 1722 | Nationwide including states of AZ, CA, CT, FL, IA, IL, IN, KY, MA, MI, MN, MO, NJ, NY, OH, OK, PA, TN, TX, VA, and WI | Terminated | February 24, 2008 | 345 | 3902 | Undetermined
Z-0739-2007 | 37767 | No | IZW | Diagnostic x-ray beam-limiting device. | Radiology | Radiology | 510(K) Exempt | Philips Collimators, Model numbers: 9896 010 00614, 9896 010 00615, 9804 607 00301, and 9804 602 83501. Product is a Collimator/ Beam Limiting Device and are replacement components for the Bucky Diagnost TH X-ray system. |  |  | March 15, 2007 | 2007 | Philips Medical Systems North America Co. Phillips | The Collimator field-defining light beam does not meet the light field contrast ratio requirement of 4:1. System is non-compliant with performance standard (1020.31 (d)(2)(iii) 21 CFR | Other | Firm issued letter on 03/20/2007 to consignees advising them that firm representatives will contact consignee to replace collimators at no charge. |  | 6 | 6 | Nationwide; Units were distributed to hospitals, radiologists, and health care clinics in IA, LA, ND, PA, and TX, and Washington DC. | Terminated | November 15, 2010 | 1341 | 6 | Undetermined
Z-0855-2007 | 37570 | No | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ARCHITECT STAT Troponin-I Reagent; list numbers 2K41-20 (400 test kit), 2K41-25 (100 test kit) and 2K41-30 (2000 test kit); each kit contains bottles of microparticles, conjugate and assay diluent; Abbott Laboratories, Abbott Park, IL 60064 |  |  | March 08, 2007 | 2007 | Abbott Laboratories MPG | False Negative Results : The analytical sensitivity claim of less than or equal to 0.01 ng/mL (ug/L) might not be met for all lots of Architect STAT Troponin-I reagents, thus false negative results may occur with negative results less than 0.1 ng/mL | Other | Abbott sent a Product Correction letter on 3/8/07 to all Troponin-I customers who received reagent lots in the last 12 months. The letters informed the accounts that the analytical sensitivity claim of |  | 9438 | 9,438 kits | Worldwide: USA including Puerto Rico, and internationally to Canada, Brazil, Germany, Japan, Taiwan, Hong Kong, Thailand, South Korea, Australia, New Zealand, Mexico, Colombia, Chile, Argentina and Trindad &amp; Tobago. | Terminated | July 16, 2008 | 496 | 9438 | Undetermined
Z-0933-2007 | 37818 | No | MCX |  |  | Cardiovascular | PMA | St. Jude Medical-CPS Direct SL- slittable CRT lead delivery tool. -Model Numbers 410110, 410111, 410112, 410113, 410114, 410115, 410116, 410120,410121. 410122, 410123, 410124, 410125 and 410126. |  |  | March 20, 2007 | 2007 | St Jude Medical CRMD | Device Separation- product may separate at the mapping collar transition joint located at the proximal (light blue) marker band end of the tool during the implant procedure. | Other | By Letter dated March 20, 2007, customers were requested to returned affected product immediately to St. Jude Medical. St. Judes Sales Representative will be assisting customers in retrieving the subject inventory from their shelves.The firm instructed customers that if they should require further information regarding this issue, to contact the firm's Technical Service department at (800) 722-3774 or customers can contact local St. Jude Medical representative. |  | 14000 | 14,000 | Worldwide-including states of AZ, CA DC, FL GA, IN, KS, LA, MA, MD, ME, MI, MN, MS, NC, NH, NY, OH, PA, RI, TN, VA, and WI and countries of Austria, Belgium, Denmark, Finland, Netherlands, Sweden, France, Germany, Italy, UK, Spain, Portugal, Poland, Switzerland, Australia, New Zealand, Thailand, Hong Kong, Korea, Malaysia and Singapore. | Terminated | February 29, 2012 | 1807 | 14000 | Undetermined
Z-0922-2007 | 36616 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | EP Healing abutment- Product Code-ITHA52-a pre-manufactured prosthetic component directly connected to an endosseous dental implant. |  |  | October 16, 2006 | 2006 | Biomet 3i | Package Integrity; sterility compromised-Incomplete seals on the nylon bags in which the product is packed. | Other | Each consignee will be contacted by email or telephone followed by a faxed letter. Each consignee was instructed to check their respective inventory for the recalled part and corresponding lot number. They were directed to return the product to the recalling firm. |  |  |  | Worldwide. | Terminated | December 30, 2008 | 806 |  | Undetermined
Z-0925-2007 | 37864 | No | MJK | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | LIFEPAK 20 defibrillator/monitor. The device is an AC powered defibrillator device with battery backup. The defibrillators are primarily used by healthcare providers in device user facilities such as hospitals and clinics. |  |  | April 27, 2007 | 2007 | Medtronic Emergency Response Systems, Inc. | White Screen Lock up: LifePak 20 defibrillators with software version 48 or version 52 may experience an intermittent timing issue during the power on sequence. This may lock up the device resulting in the monitor screen going white, and may cause a potential delay or prevention of shock treatment. | Other | On 04/27/2007 firm sent letters via return receipt to all domestic and foreign customers. The letter recommends for customers to turn the defibrillator power off and back on to restore normal operation in the event of a white screen condition. In addition, the letter advises that a representative from the firm will contact customer to make arrangements to provide a software upgrade to correct the issue. |  | 859 | 859 units | Distribution is Worldwide. Devices are used by healthcare providers in hospital or healthcare clinic settings. | Terminated | March 23, 2012 | 1792 | 859 | Undetermined
Z-0926-2007 | 37984 | No | FMK | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | NeatNick Sweeping Action Heel Lancet -Full-Term- |  |  | May 11, 2007 | 2007 | Hawaii Medical LLC | Lancet may misfire and an inadequate blood sample may result; additionally, the blade may not retract after use and result in a needle stick to the user/patient | Other | Hawaii Medical notified the initial distributor and end users by telephone and follow-up Recall Notification Letter on May 15, 2007. |  | 286376 | 286,376 units | Nationwide | Terminated | July 22, 2008 | 438 | 344434 | Undetermined
Z-0920-2007 | 37796 | No | LEG | Neuroleptic drugs radioreceptor assay test system. | Toxicology | Toxicology | 510(k) | Roche VALP2, Online TDM Valproic acid , for use with the Roche/Hitachi Cobas c 501 analyzer; Catalog No.04491041190. |  |  | April 30, 2007 | 2007 | Roche Diagnostics Corp. | Probe Wash: Reported results may have a positive bias because inadequate analyzer probe wash instructions may result in carry-over from the following assays: Creatinine Kinase L and Creatine Kinase-MBL. | Other | Consignees were issued a recall letter dated 4/30/07 instructing them to perform a R2 reagent probe wash when using CKL and CKMBL with the vancomycin or valproic acid assays. |  | 992 | 992 | Nationwide. | Terminated | December 17, 2007 | 231 | 1654 | Undetermined
Z-0917-2007 | 37795 | No | KLB | Tobramycin test system. | Toxicology | Toxicology | 510(k) | Roche Online TDM Tobramycin, for use on Roche/Hitachi analyzers 911,912, 917 and Modular P; Catalog No. 04642511190. (Note: Not distributed within the United States) |  |  | April 05, 2007 | 2007 | Roche Diagnostics Corp. | Reported results may have a positive bias of up to 113% because inadequate cleaning instructions may result in carry-over of the following reagents: Online TDM carbamazepine, Online TDM gentamycin, Online TDM phenytoin, Online DAT cocaine II, Online DAT methadone II, Online DAT opiates II and Online DAT THC II. | Other | Foreign affiliates were notified via email letter dated 4/5/07 requesting that to notify their customers of the need to implement a special wash cycle to prevent reagent carry over. |  | 474 | 474 total (Z-0917-0918-2007) | International Distribution ONLY-Australia, Canada and Germany. | Terminated | April 30, 2008 | 391 | 474 | Undetermined
Z-0916-2007 | 37788 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | AcuMatch M-Series Straight Fluted Stem Segment and AcuMatch M-Series Curved Fluted Stem Segment. Catalog numbers 150-07-19 and 150-07-44. |  |  | July 21, 2006 | 2006 | Exactech, Inc. | The packaging was mislabeled: Exterior and interior packaging on the device was labeled with catalog number 150-07-19 15mm x 200mm, however, the actual part inside the packaging was a 21mm x 200mm (catalog number 150-07-44) device. | Other | The consignees were contacted by fax and email (dated July 21, 2006) to inform them of the recall requesting return of product |  | 6 | 6 | World wide-Domestically distributed to one account in Alabama and internationally to Columbia and Australia. | Terminated | May 06, 2012 | 2116 | 6 | Undetermined
Z-0913-2007 | 37775 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Innova 2100 IQ Digital Fluoroscopic Imaging System (Versatile cardiac and vascular imaging system.). |  |  | February 05, 2007 | 2007 | General Electric Med Systems LLC | Unintended Gantry Movements : GE Healthcare has recently become aware of accidental, operated-initiated activation of gantry movements associated with the new, metallic Smart Box controller, that may impact patient or staff safely. | Other | Consignees were contacted via a letter "Urgent: Safety Notice" dated 2/5/07. The letter described the safety issue, affected product, safety instructions and product correction contact information. -This is the short-term fix, GE is working on a correction that will be installed at each site as the long term correction |  | 201 | 201 | Class II Recall - Worldwide Distribution --- USA including states of WV, WI, WA, VA, UT, TX, TN, SD, SC, PA, OR, OK,OH, NY, NV, NJ, NE, NC, MT, MO, MN, MI, ME, MD, MA, LA, KY, KS, IN, IL, ID, HI, GU, GA, FL, DE, CT, CO, CA, AZ, AL, AK and countries of UNITED KINGDOM, TURKEY, SYRIA, SWITZERLAND, SWEDEN, SPAIN, SAUDI ARABIA, RUSSIA, POLAND, NORWAY, NEW ZEALAND, THE NETHERLANDS, MOROCCO, MALAYSIA, KOREA, JORDAN, JAPAN, ITALY, ISREAL, IRELAND, INDIA, HUNGARY, GREECE, GERMANY, FRANCE, FINLAND , DENMARK, CZECH REPUBLIC, CHINA, CANADA, BELGIUM, AUSTRALIA, and ALGERIA. | Terminated | December 17, 2011 | 1776 | 446 | Undetermined
Z-0868-2007 | 37826 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | ProfemurÂż R, Revision Hip system, Proximal Body (part), STD 4, REF: PPW38354, Cementless, Titanium Plasma Spray Coated, Sterile, Rx only. Hip implant. |  |  | March 27, 2007 | 2007 | Wright Medical Technology Inc | The titanium plasma coating was found to have missing fragments. | Other | The firm sent out recall notices to implanting surgeons by letter dated 03/27/2007 (via FedEx) to notify the surgeons of the issue. The hospitals and distributors were also notified of the issue and given instructions for return of the product. |  | 4 | 4 units | Nationwide, Puerto Rico, &amp; Japan | Terminated | June 30, 2008 | 461 | 22 | Undetermined
Z-0840-2007 | 37849 | No | FFS | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Smith &amp; Nephew, 420XL Xenon Light Source, 180 Watt; for endoscopic use. Catalog/Part #72200510; manufactured for Smith &amp; Nephew Endoscopy, Andover, MA 01810. |  |  | April 09, 2007 | 2007 | Smith &amp; Nephew, Inc Endoscopy Division | Several units were manufactured using a power supply that contains a 1 amp fuse where a 5 amp fuse is required. When the unit's power supply draws more that 1A, the fuse will fail causing the light to cease operating. The problem can occur anytime: from the initial turning on (powering up), to anytime during use. | Nonconforming Material/Component | Notification sent 04/09/07 via FedEx to distribution centers, sales reps and user facilities. Users asked to send light source to Smith &amp; Nephew GMBH Service Center in Germany for repair (contact Ute Dieling at +49 (0) 7462-208-313). |  | 67 | 67 units. | Product sent to distributors in Argentina, Austria, Chile, Denmark, Dubai, France, Greece, Hong Kong, India, Italy, Japan, Korea, Malaysia, Norway, South Africa, Spain, Sweden and United Kingdom. No US distribution. | Terminated | September 02, 2008 | 512 | 67 | Undetermined
Z-0839-2007 | 37821 | [6 empty] | Silver Speed Hydrophilic Guidewire, for placement of vascular catheters. Model # 103-0601-200. |  |  | April 05, 2007 | 2007 | EV3 Neurovascular | Two customer complaints were received from physicians in Japan indicating the hydrophilic coating was observed flaking from the guidewire while immersed in hydrating media following use during aneurysm embolization procedures. | Other | A recall plan was developed for 100% recovery of guidewires within affected lot numbers. The plan includes notification to U.S., European, and Rest-of-World customers and distributors. A recall notification letter was prepared 04/05/2007 which summarizes the coating problem and potential risks, identifies the affected lot numbers and quantities shipped to the customer, and provides instructions for return of unused guidewires. |  | 2120 | 2120 | Worldwide, including USA, Italy, United Kingdom, Switzerland, Germany, Belgium, South Africa, Canada, Slovenia, China, Austria, France, Spain, Jordan, Mexico, Portugal, Colombia, Greece, Brazil, Venezula, Ireland, Australia, Malaysia, Saudi Arabia, South Korea, Lithuania, Netherlands, Hungary, and United Arab Emirates | Terminated | August 23, 2012 | 1967 | 2120 | Undetermined
Z-0862-2007 | 37782 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Smith + Nephew, Birmingham Hip Resurfacing, Acetabular Cup W/ Impactor, 46 mm cup Ref.: 74120146 |  |  | March 13, 2007 | 2007 | Smith &amp; Nephew Inc | The carton label reads BHR Acetabular Cup 60mm contained a 58mm cup. | Other | The consignees (hospital CEOs) were notified of the problem and the recall by letter (Fed Ex) on 03/13/2007. |  | 1 | 1 unit | Europe. No domestic distribution. | Terminated | October 28, 2011 | 1690 | 5 | Undetermined
Z-0861-2007 | 37665 | No | BZD | Noncontinuous ventilator (IPPB). | Anesthesiology | Anesthesiology | 510(k) | Infant Nasal CPAP Cannula, Size 4; an Rx infant nasal prong continuous positive airway pressure (CPAP) cannula one inspiratory and one expiratory elbow connector; 10 units per case; Teleflex Medical, Research Triangle Park, NC 27709 U.S.A., Made in Mexico; The cannulas are sold as components in the following configuration: catalog No. 1694: Infant CPAP Set intended for use with mechanical ventilators, which includes one cannula with one inspiratory and one expiratory elbow connector, two 10mm to 7.5mm adapters, one knit cap and two 6 inch sections of Velcro securing tape; |  |  | March 23, 2007 | 2007 | Teleflex Medical | The luer-port in the expiratory elbow may be occluded, which interferes with a ventilator's ability to detect an accurate pressure reading. | Other | Teleflex Medical sent Urgent Medical Device Recall letters dated 3/23/07 via UPS 2nd Day Air to the direct accounts, informing them that the luer-port in the expiratory elbow may be occluded, which interferes with a ventilator's ability to detect an accurate pressure reading. The accounts were requested to cease use of the affected units, remove them from use/stock, quarantine and destroy them. If the accounts further distributed the product, they were requested to conduct a sub-recall of the affected units from their customers. The accounts were instructed to complete and return by fax the enclosed Acknowledgement &amp; Stock Status Form, indicating the number of units destroyed and if they are requesting replacement product. Any questions were directed to Linda Todd at 1-800-334-9751, ext. 4951. |  | 13530 | 13,530 for all products | Worldwide, including USA, Puerto Rico, Australia, Bolivia, Brazil, Canada, Chile, China, Colombia, Dominican Republic, El Salvador, Germany, Guatemala, Hong Kong, India, Israel, Kuwait, Malaysia, Mexico, New Zealand, Paraguay, Peru, Saudi Arabia, Singapore, South Africa, South Korea, Taiwan, and Venezuela | Terminated | November 16, 2009 | 969 | 13530 | Undetermined
Z-0856-2007 | 37538 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Precision Xtra Blood Glucose monitoring system. Individual test meter and test strip kit. |  |  | March 05, 2007 | 2007 | Abbott Diabetes Care, Inc. | The front button covers may detach from the meter. | Other | The firm intends to send letters to all customers that have registered their units under the firm's warranty database and the firm's first consignees. In addition, this notification will be placed on the company website, with e-mail and phone contacts as needed. Letter dated March 5, 2007. |  | 300 | approximately 300 lots | Product was distributed worldwide. | Terminated | April 04, 2008 | 396 | 300 | Undetermined
Z-0841-2007 | 37876 | No | MIE |  |  | Cardiovascular | Contact ODE | CryoValve, Pulmonary Valve &amp; Conduit |  |  | April 17, 2007 | 2007 | CryoLife, Inc. | Donor tissue was released prior to information received regarding a pericardial effusion. | Other | Both physicians were contacted by telephone on 4/17 &amp; 18/2007. Both tissues were returned to CryoLife |  | 1 | 1 tissue | AZ, CO | Terminated | December 31, 2007 | 258 | 1 | Undetermined
Z-0872-2007 | 37792 | No | NBH | Arthroscope. | Orthopedic | Orthopedic | 510(K) Exempt | Well Leg Holder With Boot, part of the Supine Universal Hip Distractor System; Ref (Part) #72200632; Manufactured for Smith &amp; Nephew, Inc., Endoscopy Divsion, Andover, MA 01810. |  |  | November 01, 2006 | 2006 | Smith &amp; Nephew, Inc Endoscopy Division | The device was not manufactured to correct specification causing interference when used in conjunction with Patient Transfer Board. | Other | Firm notified consignees by letter and phone call on 11/1/2006. Consignees asked to return out-of-spec Well Leg Holder for a replacement. |  | 12 | 12 | Nationwide Distribution: USA including states of WI, CA, MO, IL, TX, MA, WA, NM, FL, NH, and SC and the UK | Terminated | June 16, 2008 | 593 | 12 | Undetermined
Z-0891-2007 | 37309 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare CIC Pro Clinical Information Center, Central Station, Software Version 4.1. , GE Medical Systems Information Technologies, Inc., 8200 W. Tower Ave. Milwaukee, Wisconsin USA. (System hardware consists of a processing unit and a monitor) |  |  | January 24, 2007 | 2007 | General Electric Medical Systems Information Technology | Alarm and Reboot problems: Two separate issues could occur with the CIC Pro system when used with Telemetry. Each issue occurs in software version v4.1 only; A) the system may experience a loss of audible alarms when operating system memory resources reach 170MB and/or B) the system may intentionally or unexpectedly reboot . | Other | GE Healthcare sent a certified letter titled "Urgent Medical Device Correction" to customers on 3/13/07 that explained the short and long term solutions and that the company would be providing a software update that will rectify the issue specific to CIC Pro v4.1 software. |  | 5190 | 5,190 Units | Nationwide | Terminated | December 10, 2011 | 1781 | 5190 | Undetermined
Z-0906-2007 | 37857 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | c.cam Gamma Camera, Model 9VIR1200; Siemens Medical Solutions USA, Inc., Molecular Imaging, 2501 N. Barrington Road, Hoffman Estates, IL 60195-5203; Made in Denmark; |  |  | April 30, 2007 | 2007 | Siemens Medical Solutions USA, Inc. | Chair Breakage: Under extreme load, the chair back of the c.cam emission computed tomography system may bend and eventually break. This could be caused by the accumulated stress from heavy patients seating or stretching themselves in the chair or simply from long time use. | Other | Important Safety Notice letters dated 4/30/07 on Siemens letterhead were sent to the end users, informing them of the potential risk that the c.cam chair may bend or break under heavy load. They were told that a Siemens service representative will contact them within the next few days to arrange the repair of the chair at no cost to the end user. To ensure the chair does not become damaged, SIemens recommended that the c.cam not be used for patient scanning until the update has been performed on the chair. Any questions regarding the Important Safety Notice were directed to their service representative or Siemens' office at: 1-800-888-7436 (USA). |  | 413 | 413 cameras | Worldwide-USA, Canada, France, Germany, Israel, The Netherlands and Portugal. | Terminated | January 09, 2009 | 620 | 413 | Undetermined
Z-0896-2007 | 37806 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Surge Medical Solutions 15 Fr. (5.0mm) Retrograde Cardioplegia Cannula, self-inflating balloon 18 mm, handle stylet, sterile, REF CODE: RSH-M014S. |  |  | April 17, 2007 | 2007 | Surge Medical Solutions, LLC | The device may leak between the two lumens or at the stop cock, which could result in inaccurate coronary sinus pressure readings. | Other | Consignees were notified and requested to return the recalled products via recall letter dated 4/17/07. |  | 2211 | 2,211 for all Recalls Z-0892-0904-2007 | Worldwide, USA, Argentina, Germany, Ireland, Italy, Lebanon, Saudi Arabia, Turkey, and the United Kingdom. | Terminated | March 10, 2009 | 693 | 2211 | Undetermined
Z-0873-2007 | 37793 | [6 empty] | Light Source 300 Watt Xenon Lamp used in the following models: |  |  | April 05, 2007 | 2007 | Vision Systems Group, A Div of Viking Systems | Incorrect power entry module assembled into unit may cause burning to insulation, compromised electrical emissions and premature light source failure | Nonconforming Material/Component | Viking notified 2 distributors and 1 end user by telephone on 3/15/07. Viking representatives performed on site visit to perform inspection and replace units if needed. |  | 4 | 4 | Nationwide to CA and NC | Terminated | June 07, 2007 | 63 | 4 | Undetermined
Z-0890-2007 | 37997 | No | LMD | Medical image communications device. | Radiology | Radiology | 510(K) Exempt | Sienet MagicStore. Radiological Digital Image Communications System |  |  | May 02, 2007 | 2007 | Siemens Medical Solutions USA, Inc | System could crash and result in data loss. | Other | The recalling firm issued a Safety Advisory letter to their customers (May 2, 2007) to inform them of the problem. A company representative will be contacting the customers to arrange for replacement of the RAID component. |  | 19 | 19 units | The products were shipped to medical facilities in CA, CO, FL, MS, NH, NJ, NY, and WA. | Terminated | September 17, 2008 | 504 | 19 | Undetermined
Z-0889-2007 | 37886 | No | LJE |  |  | Gastroenterology/Urology | 510(k) | Boston Scientific PCNL (Percutaneous Tract) Kit with Nephromax Balloon; |  |  | April 11, 2007 | 2007 | Boston Scientific Corporation | Mislabeled: Outer box unit has correct part number but product is incorrectly identiified as a Percutaneous Tract Kit with Nephromax Balloon. The box actually contains a Percutaneous Tract Kit with Amplatz Dilator and Sheath Kit | Other | Boston Scientific notified customers on 4/11/07 by letter via Federal Express mail. Users are requested to return product in inventory. |  | 15 | 15 kits | Worldwide Distribution; USA including states of OH, MS, NM, MD, and TX and Foreign country of Trinidad | Terminated | December 12, 2007 | 245 | 15 | Undetermined
Z-0854-2007 | 37605 | No | LZS |  |  | Ophthalmic | PMA | LADAR6000 Excimer Laser System |  |  | February 21, 2007 | 2007 | Alcon Refractive Horizons, Inc. | Observed 'central islands' in patients following custom myopia laser procedures using the LADAR6000 Excimer Laser. (On May 11, 2007, Alcon notified consignees that the device software programs for CustomCornea Myopia and CustomCornea Myopia with Astigmatism of the device would be deactivated) | Other | Consignees were notified of the recall via two letters. On Feb. 21, 2007, consignees were requested to immediately stop performing all custom cornea myopia procedures and to return an acknowledgement card indicating they have received the alert and understand the required actions. On May 11, 2007, Alcon notified consignees that the device software programs for CustomCornea Myopia and CustomCornea Myopia with Astigmatism of the device would be deactivated. After meeting with CDRH's Office of Compliance on August 27, 2007, the recalling firm committed to total product removal of the device. Notification will begin in Sept 2007 and will be made by personal visits to each consignee. |  | 102 | 102 | Worldwide Distribution-including USA, Czech Republic, Spain, Australia, Italy, Hong Kong, Vietnam, Singapore, South Korea, and Canada | Terminated | March 21, 2017 | 3681 | 102 | Undetermined
Z-0869-2007 | 37723 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Innova 4100/4100 IQ. Digital Fluoroscopic Imaging System (Vascular Angiography System) |  |  | March 26, 2007 | 2007 | General Electric Med Systems LLC | Measurement errors during the computation of vessel diameters while using the Stenosis Analysis software (also known as Quantitative Coronary Analysis function-QCA) using the auto-calibration function on Innova Digital (located in control room) or in-room InnovaCentral/Touch-Screen | Other | A letter dated Jan. 18, 2007, was sent to consignees describing the problem, product ,safety instructions and contact information. Software upgrades were promised for correction. |  | 70 | 70 | Worldwide Distribution: USA states including-AZ,CA,CO,DE,FL,GA,IL,IN, MD,MN,MO,NJ,NV,NY,OH,OK,OR,PA,SC,SD,TX, UT, VA,WA,WI, and countries of Canada, Belgium, China, Czech Republic, Finland, France, Germany, Hungary, India, Isreal, Italy, Japan, Korea, New Zealand, Norway, Russia, Spain, and United Kingdom. | Terminated | October 21, 2008 | 575 | 288 | Undetermined
Z-0847-2007 | 37902 | No | FMG | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Custom Kit, Left Heart Kit, REF/CAT No.: K09T-09576, Rx only, Merit Medical Systems Inc., South Jordan, Utah. |  |  | April 17, 2007 | 2007 | Merit Medical Systems, Inc. | Male luer in fluid administration sets may crack due to inadequate molding process | Other | Consignees were notified by letter on 04/17/2007 and instructed to quarantine and return any unused product. |  | 17 | 17 units | CA, IL, NJ, NY. No military, government or foreign distribution. | Terminated | May 05, 2008 | 384 | 41251 | Undetermined
Z-0818-2007 | 37856 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Ultrasite Horizon Pump IV Set. The product is shipped in cases containing 24 sets. |  |  | April 18, 2007 | 2007 | B. Braun Medical, Inc. | Secondary infusion backs up into primary container. | Other | The recalling firm issued Product Removal Notices on 4/20/07 via certified priority mail. The notice informed the customer of the problem and the need to return the product. |  | 2500000 | 2.5 million for all products | Nationwide and Canada. | Terminated | April 22, 2008 | 370 | 2500000 | Undetermined
Z-0934-2007 | 37847 | No | MHY |  |  | Neurology | PMA | Medtronic DBS Lead Kit for Deep Brain Stimulation, 3389S-40, Medtronic, Inc. Minneapolis, MN 55432-5604 |  |  | April 16, 2007 | 2007 | Medtronic Neuromodulation | Mislabeling: The inner package labeling incorrectly specified the model number of the lead as 3387-40. The outer package labeling is correct and the package actually contains the correct lead (3389S-40). | Other | Medtronic field personnel visited affected customers and retrieved product from this lot, unless the product was needed for an imminent scheduled surgery. Telephone scripts and leave behind letters to each consignee were initiated in April 2007. |  | 31 | 31 kits | Nationwide including states of CA, CO, IL, IN, MD, NC, OH, PA, and WV. | Terminated | February 24, 2008 | 314 | 31 | Undetermined
Z-0921-2007 | 37774 | No | FRW | Patient scale. | General Hospital | General Hospital | 510(K) Exempt | IMS Patient Scale used on Sunrise Medical's Hoyer Presence and Stature Patient Lifts (Cradle), Models: HOY-4pt-WSC and HOY-StatureWSC |  |  | January 30, 2007 | 2007 | Integrated Measurement Systems, Inc. | Patient Scale may become detached, which can cause a patient to fall and result in injury. | Other | The scale manufacturer, IMS, in conjunction with the lift manufacturer, Joerns Healthcare (previously Sunrise Medical), will initiate a field fix strategy consisting of shipping replacement 4-point cradles without scales, for use until a replacement can be designed, tested and validated. Joerns will utilize their Customer Service Help Line (1-800-826-0270) on behalf of IMS, to inform customers of the voluntary field correction, address any questions regarding the IMS scale field fix strategy and to provide additional information on return of affected IMS Patient Scales. Customers were notified by telephone beginning on January 30, 2007. |  | 149 | 149 units | Worldwide | Terminated | March 02, 2008 | 397 | 149 | Undetermined
Z-0987-2007 | 37888 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | HD11XE Diagnostic Ultrasound System |  |  | April 25, 2007 | 2007 | Philips Medical Systems | Potential for erroneous measurements which may lead to incorrect diagnosis based upon the presence of an inaccurate, undetectable error in clinical measurement or calculation. If the user exits an imaging mode while Active Annotation is in use, the system can enter an undefined state where the image scale calibration may not match the scale of the on-screen measurements and graphics and possib | Other | On 04/25/2007 firm sent letters dated 04/23/2007 to consignees informing them of the issue and that a new software release to correct the issue became available 04/19/2007. The firm's Field Service Representatives will contact consignees to install the new software at no charge to the consignee. |  | 566 | 219 units in US and 347 units distributed internationally | Worldwide. | Terminated | January 24, 2011 | 1370 | 566 | Undetermined
Z-0937-2007 | 37905 | No | KNQ | Esophageal dilator. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | ConMed Eliminator Esophageal PET Balloon Dilator; Outer Diameter (OD) size 8mm (24 Fr), length 8cm; product code/end item #000341; Single Use, Non-Sterile; distributed by ConMed Endoscopic Technologies, Billerica, MA 01821. |  |  | April 16, 2007 | 2007 | Conmed Endoscopic Technologies, Inc. | Mislabeling issues including; incorrect balloon diameters, lot number discrepancies, and incorrect type of dilator. | Other | Notification Letters were sent to domestic consignees on 04/16/07 and to international consignees on 04/17/07. Consignees instructed to discontinue product use, sub recall to accounts sold product, and to contact Stericycle, Inc. at 1-800-668-4391 for instructions on returning product. |  | 68 | 68 units. | Worldwide-Devices distributed to medical facilities and distributors Nationwide (USA) and to distributors in the following countries: Australia, Belgium, Canada, Croatia, Denmark, Germany, India, Italy, Japan, Mexico, Poland, Spain and Sweden. | Terminated | July 30, 2010 | 1201 | 12185 | Undetermined
Z-1001-2007 | 38118 | No | LFR | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Ascensia Contour Blood Glucose Monitoring System ( Bayer )Diabetes Care System; Product 7152A.---NO USA Units are affected---. |  |  | June 06, 2007 | 2007 | Bayer Healthcare LLC | Incorrect Unit of Measure displayed : The meters report results in the wrong unit of measurement for Canadian users. | Other | The firm began calling Canadian customers on or about 6/06/07 to inform them of the problem, to tell them that a meter with the correct measurement units will be sent to them, and to request that they return the meter currently in their possession. |  | 16 | 16-Canadian Units only are affected | Canada. | Terminated | November 30, 2007 | 177 | 16 | Undetermined
Z-1006-2007 | 37832 | No | BZD | Noncontinuous ventilator (IPPB). | Anesthesiology | Anesthesiology | 510(k) | ResMed S8 Continuous Positive Airway Pressure (CPAP) Flow generator with the following model names and product codes: S8 AUTOSET VANTAGE, #33112; S8 COMPACT, #33030; S8 ELITE, #33021; S8 ESCAPE, #33007. |  |  | April 23, 2007 | 2007 | Resmed Corporation | S8 devices manufactured between July 2004 and May 15, 2006 are susceptible to fatigue in the connection of the AC appliance inlet connector and the power supply (power cord) leading to possible failure (short circuit, electrical failure, sparks, smoke). | Other | ResMed began to broadly distribute information regarding the recall on 4/23/2007. A press release was issued on 4/23/2007 with information regarding the recall and information was posted at ww.resmed.com/s8program. The firm is providing a Consignee notification package and press release to include(1) specific device names subject to the recall and affected serial numbers; specific information to be provided in each package to the end-customer. The specific information to the end user customers will include: (1) specific device names subject to the recall; (2) affected serial numbers; (3)continuing use and safety information; and (4) contact information for requesting additional details. Patients are advised that they may continue to use their S8 flow generators until a replacement unit is provided. ResMed is instructing patients, sleep labs, and hospitals to continue using the device, placing it on a hard clean surface, with a clear area around the device. ResMed intends to replace the affected products with either a new device or a device in which the affected components have been replaced. |  | 300000 | 300,000 | Worldwide, USA, Canada, and Mexico. | Terminated | December 06, 2010 | 1323 | 300000 | Undetermined
Z-0958-2007 | 37957 | No | LER | Theophylline test system. | Toxicology | Toxicology | 510(k) | IMx Theophylline Reagent Pack (List 1A81-20) In Vitro test. |  |  | April 11, 2007 | 2007 | Abbott Diagnostic International, Ltd. | A decrease in mP values (calibration curve Net Polarization value) has been observed with this reagent lot at all calibrator levels and may produce the following results: Calibration Error Code 148 (Polarization Too Small), Calibration Error Code 172 (Calibration Not Stored or Calibration Not Accepted) Controls out of range (when using a stored calibration curve). | Other | An April 11, 2007, Product Recall Letter was sent to all customers who received the affected lot. All customers who received the affected lot will be contacted. The following communication types were used: Product Recall Letter &amp; Customer Reply Form. Letters were sent to customers by Federal Express Priority mail. The Customers were instructed the following: To discontinue use of and destroy IMx Theophylline reagent lot 41610QI00. If they had forwarded the lots listed in the Product Recall letter to other laboratories, they were instructed to inform them of this Product Recall and provide a copy of this letter to them. They also we instructed to retain a copy of the communication for their records |  | 17 | 17 Kits | Worldwide: CA, IL, MA, NY, OR, PA, VA, and WV and countries of Germany and Singapore | Terminated | June 10, 2011 | 1521 | 17 | Undetermined
Z-1041-2007 | 37969 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ACCELERATOR APS System Input/Output Module (IOM); List Number: 7L01-01; manufactured by Inpecco SpA, Segrate (MI), Italy. |  |  | May 15, 2007 | 2007 | Abbott Laboratories, Inc | Product's hinged lid can fall, if not fully opened as designed, representing a potential injury hazard to the operator. | Other | Firm notified consignee via letter on 05/15/07. Consignee informed of potential for injury and advised that a label would be affixed to the device lid by an Abbott representative. |  | 1 | 1 unit. | Device distributed to a hospital in UT. | Terminated | November 27, 2007 | 196 | 1 | Undetermined
Z-0979-2007 | 37834 | No | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | STA Neoplastine CI Plus (10), (cat. 0667). |  |  | March 30, 2007 | 2007 | Diagnostica Stago, Inc. | Based on the confirmed clotting times in the upper part of the assigned ranges of STA-Coag Control ABN control plasmas when used with STA-Neoplastine CI Plus, use of this lot may cause difficulty during the validation of the STA-Coag Control ABN. | Other | Recall notification letters were sent on March 30, 2007 by first class mail. The letter advises customers that lot 061781 of STA-Neoplastine Cl Plus may cause difficulty during the vailidation process of STA-Coag Control ABN control plasmas, and asks that the lot number be discontinued. |  | 170 | 170 units | Nationwide. | Terminated | November 30, 2009 | 976 | 170 | Undetermined
Z-0853-2007 | 37746 | No | GGP | Factor deficiency test. | Hematology | Hematology | 510(k) | STA Liatest VWF:Ag, Product Catalog Number : 0518. This is a Immuno-Turbidimetric Assay of von Willebrand Factor. |  |  | March 13, 2007 | 2007 | Diagnostica Stago, Inc. | A dose-hook effect is observed starting from VWF:AG levels of 300%, while the package insert indicates that no dose-hook effect has been observed with VWF:Ag levels up to 800%. | Other | Customers were notified on March 13, 2007 by letter, first class mail. |  | 1098 | 1,098 units | Hospitals and laboratories Nationwide | Terminated | January 16, 2009 | 675 | 1098 | Undetermined
Z-0975-2007 | 37526 | No | GIO | Capillary blood collection tube. | Hematology | Hematology | 510(K) Exempt | SafeClinitubes, REF 942-892, D957P-70-100x1, contained in kits containing 250 SafeClinitubes, insert, mixing wires and end caps. Capillary samplers for drawing small amounts of blood for blood gas analysis. |  |  | June 22, 2006 | 2006 | Radiometer America Inc | The clinitubes are not able to be filled due to contamination with silicone during production. | Other | The firm notified their customers of this recall with a letter dated 6/22/2006. The letter informs the customers of the problem with the affected lots of SafeCLINITUBES and asks that the customers discard any remaining boxes of affected product on-site. The recall letter is accompanied by a FAX FORM (response form) which the customer is asked to complete in order to request replacement product to be sent to their individual locations. |  | 175 | 175 boxes | Nationwide and Canada. | Terminated | May 08, 2012 | 2147 | 175 | Undetermined
Z-0928-2007 | 38272 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | OEC Miniview 6800 mobile fluoroscopic x-ray system, GE Healthcare, Surgery, Salt Lake City, Utah. |  |  | May 24, 2007 | 2007 | OEC Medical Systems, Inc | X-ray systems did not have the required labeling to indicate Air Kerma Rate and cumulative Air Kerma during fluoroscopic procedures | Other | Consignees were notified by letter on 05/24/2007 and advised that a service representative would be contacting them for installation of the required label and to provide them with an update to their user's manuals. |  | 811 | 811 systems | Worldwide; USA including VA facilities in AK, CA, ID, MD, MI, NY and OR. Military distribution to CA, CO, CT, FL, GA, KS, LA, MS, OH, OK, TX, VA and Iraq. Foreign distribution to Austria, Australia, Belgium, Canada, China, Denmark, Finland, France, Germany, Hong Kong, Israel, Italy, Japan, Korea, Malaysia, Netherlands, Norway, Portugal, Singapore, South Africa, Spain, Sweden, Switzerland, Taiwan and UK. | Terminated | June 02, 2008 | 375 | 811 | Undetermined
Z-0738-2007 | 38233 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Kodak Direct View DR7500 System, Catalog #s 855-1046, 879-1345, 161-9139 --- A stationary X-Ray System. Device Listing # E222512. |  |  | March 06, 2007 | 2007 | Carestream Health, Inc. | The x-ray systems do not indicate the field size when the system's detector is turned off and the collimator is in manual mode. This results in the systems being noncompliant with the field indication performance standards of section 1020.31(e)(2)&amp; (3) of 21Code of Federal Regulations (CFR). | Other | On 2/21/07, Eastman Kodak notified CDRH of its proposed Corrective Action Plan CAP). The CAP was approved by CDRH under the following conditions: (1) Kodak will visit each user facility, and apply dial indicator labels to the collimator knobs, apply a field size label to the collimator housing, and verify that each system complies with the field indication performance standard. (2) Kodak will prepare adequate instructions for the assembly/installation and testing of the dial indicators labels and field size labels. These instructions are to be furnished to all owners of the Kodak Direct View DR 7500 Systems. (3) Kodak will supplement the Kodak DR 7500 system product report with a copy of the adequate instructions. (4) Kodak will perform 100% verification of the recall effectiveness. (5) Kodak will not charge for the corrections and will reimburse the owner for any out-of-pocket expenses. (6) Kodak will notify the manufacturer of the Ralco 302 ACS collimators, Ralco, that the collimator is noncompliant with the performance standard. Kodak has reported that the CAP was completed as of 5/07/07. |  | 28 | Domestic: 28 units | Nationwide: CO, IL, MI, MN, MS, NY, OK, PA, SC, TN, WA | Terminated | June 29, 2007 | 115 | 28 | Undetermined
Z-0206-2007 | 36690 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | One Touch Ultra Blood Glucose test strips, LifeScan, a Johnson &amp; Johnson Co., 50 test strips, 2 vials of 25 test strips |  |  | October 19, 2006 | 2006 | Discount Diabetic Supply | Counterfeit-blood glucose test strips (manufacturer unknown) | Other | The recalling firm notified its consignees of the problem and the recall by telephone on 10/19/2006. The firm followed this with a written notice on 10/30/2006. |  | 240 | 240 units | Nationwide distribution --- States of CA, TX, TN, GA, AL, and OK | Terminated | December 22, 2009 | 1160 | 240 | Undetermined
Z-0932-2007 | 37242 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1, 4 inch diameter Roller Pump; Model 801040. |  |  | August 01, 2005 | 2005 | Terumo Cardiovascular Systems Corp | The tube clamp mechanism on the roller pump may not open when the rotation knob is turned, making it difficult to remove and/or insert tubing. | Other | The firm is visiting the consignees as complaints are received and making corrections. Corrections were made in August 2005 and April 2006 to date. Other units will be repaired if complaints are received and the firm issued a Device Correction Letter dated 6/6/07 notifying all users of the problem and appropriate action to take should it occur. |  | 1982 | 1982 total for both models distributed | Worldwide; USA, Australia, Belgium, Canada, Dominician Republic, Egypt, Germany, Hong Kong, India, Indonesia, Japan, Korea, Malaysia, Mexico, Pakistan, Republic of Georgia, Singapore, South Africa, Taiwan, Thailand and United Arab Emirates. | Terminated | February 23, 2008 | 936 | 1982 | Undetermined
Z-0924-2007 | 37244 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1, 6 inch diameter Roller Pump; Model 801041. |  |  | April 01, 2004 | 2004 | Terumo Cardiovascular Systems Corp | A belt slip message may be displayed on the roller pump due to low belt tension or the presence of grease. | Other | The firm is visiting consignees and making corrections as complaints are received. These corrections, to date, have been made between 04/04 and 09/06. The firm will correct other units as complaints are received and issued a Device Correction Letter dated 6/6 /07 notifying all users of the problem and appropriate responses should it occur. |  | 525 | 525 total for both models | Worldwide: USA, Belgium, Canada, Egypt, Germany, Hong Kong, India, Indonesia, Japan, Korea, Malaysia, Singapore and United Arab Emirates. | Terminated | June 11, 2008 | 1532 | 525 | Undetermined
Z-0929-2007 | 37241 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1; 100V -120V, 15 A (circuit breaker), 50/60 Hz (20A power source required) base; Model 801763. |  |  | February 01, 2005 | 2005 | Terumo Cardiovascular Systems Corp | System 1 power supply may electrically short to the main console, disrupting normal system operation by causing unexpected re-setting or shutdown of pumps and modules, false alarms and/or overheating and failure of the central control monitor. | Other | The firm corrected the complaint units at the time of the complaints in 2005. The firm issued a Device Correction Letter dated 6/6/07 notifying all users of the problem and advising them that their system will be corrected when replacement components become available. |  | 416 | 416 for both models | Worldwide: USA, Australia, Belgium, Canada, Dominican Republic, Egypt, Germany, Hong Kong, India, Indonesia, Japan, Malaysia, Mexico, Pakistan and Singapore. | Terminated | July 07, 2009 | 1617 | 416 | Undetermined
Z-0980-2007 | 37872 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Proximate PPH, Procedure for Prolapse and Hemmorroids Set, REF PPH03. Each set contains: 1 hemorrhoidal circular stapler, 1 suture threader, 1 circular anal dilator, and 1 purse string suture anoscope. |  |  | April 02, 2007 | 2007 | Ethicon Endo-Surgery Inc | The adhesive used to bond the casing of the device may have migrated to the staple driver. While the device still fires, it may result in audible and tactile feedback at the midpoint of the firing cycle which may be interpreted by some surgeons as a completed firing cycle, which may cause the surgeon to prematurely stop firing the device and result in incomplete staple formation or cut line. | Other | The recalling firm sent a recall letter dated 04/02/07 to foreign consignees and a second letter dated 04/03/07 to their United States consignees. The letter asks that the customers: 1) Examine their inventories for the presence of the affected product and remove it from inventory if found; 2) Complete the Business Reply Form and fax it to Stericycle their representative and 3) Return all affected product in order to received credit, along with the completed Business Reply Form in the box identified with a pre-printed UPS shipping label supplied by the recalling firm. |  | 23966 | 23,966 units | Worldwide, including USA and Canada. | Terminated | December 09, 2009 | 982 | 23966 | Undetermined
Z-0976-2007 | 37651 | No | IZH | Mammographic x-ray system. | Radiology | Radiology | 510(k) | Diamond Mammographic X-Ray System, type MGX-2000, model 5111771. GE Healthcare P.O. Box 414, Milwaukee, Wisconsin 53201 U.S.A. |  |  | November 16, 2005 | 2005 | General Electric Med Systems LLC | GE Healthcare has become aware of the absence of thread-locking compound associated with the C-arm fasteners of your Diamond mammographic x-ray system that may impact patient safety. | Other | A Product Safety Notification, dated 05/24/07, was sent to Hospital Administrators, Managers of Radiology and Radiologists. This letter describes the safety issue, instructions to continue the use of the product and states that the device will be corrected by GE Healthcare service engineers. |  | 106 | 106 | Worldwide, including USA, Germany, France, Italy, South Africa, Finland, Norway, Sweden, Venezuela, Benelux, Japan, China, and the UK. | Terminated | February 24, 2008 | 830 | 106 | Undetermined
Z-0977-2007 | 37764 | No | NHG | Cortisol (hydrocortisone and hydroxycorticosterone) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Roche Cortisol reagent for use in conjunction with Roche cobas e 411, cobas e 601, Elecsys 1010, Elecsys 2010 and Modular E 170 analyzers; Catalog no. 11875116122. in vitro diagnostic. |  |  | May 15, 2007 | 2007 | Roche Diagnostics Corp. | The test is not as sensitive, or as precise between runs, as claimed in the labeling. Low concentrations are observed only in saliva samples. | Other | The firm notified consignees via letter dated 5/15/07 and notifed them of the changes to functional sensitivity and between-run precision claims. |  | 6051 | 6051 | Nationwide. | Terminated | June 11, 2008 | 393 | 6051 | Undetermined
Z-0740-2007 | 38120 | No | KYG | Ocular surgery irrigation device. | Ophthalmic | Ophthalmic | 510(K) Exempt | Disposable Bimanuals Irrigation and Aspiration Handpiece Set: Item #1719 Disposable Bimanuals Irrigation/Aspiration Set 19 Gauge, Dual Port; Item #1719-A - Disposable Bimanual Aspiration Handpiece 19g; Item #1719-I Disposable Bimanuals Irrigation/Aspiration Set, 19g, Open End; Item #1719-S Disposable Bimanuals Irrigation/Aspiration Set, 19g, Dual Port, Smooth. For use in cataract surgical procedures. |  |  | April 09, 2007 | 2007 | Oasis Medical Inc | Plastic shavings are coming from the producthandle during irrigation. Use of this product may leave plastic shavings in the patient's eyes. | Other | An April 9, 2007, notification of the recall will be sent to distributors and users. They are requested to notify down to the user (physician) level. Included in the RecallNotification Letter are instructions: "In addition, if you may have further distributed this product,please identify your customers and notify them at once of thisproduct recall. Your notification to your customers may beenhanced by induding a copy of this recall notification letter.This recall should be carried out to the user level. Yourassistance is appreciated and necessary to prevent anyinconvenience." |  | 3050 | 3050 total for all item numbers | Worldwide, including USA, United Kingdom, Canada, Spain, Turkey, Kuwait, Colombia, Italy, Greece, Australia, Switzerland, South Africa, Germany, Portugal, and France. | Terminated | October 06, 2011 | 1641 | 3050 | Undetermined
Z-0985-2007 | 37965 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Toshiba Aplio 80 Digital Ultrasound System, SSA-770A with version 6 software |  |  | February 02, 2007 | 2007 | Toshiba American Med Sys Inc | When using the LV measurement function and making numerous measurements the value displayed on the image screen may on occasion not update correctly. The values that are stored in MDA (measurement display area) are correct and no issue has been reported with these values. | Other | Notification letters were sent to customers on February 2, 2007. Customers were informed of the safety issue regarding the possibility of incorrect measurement values being displayed on the active image screen while performing multiple measurements and instructed to verify actual results against the values stored in the MDA (measurement display area). Customers were also informed that a Toshiba Customer Service Engineer will contact them to schedule a software upgrade when software becomes available. If customers have any questions they were instructed to contact their Toshiba Customer Service Engineer or call (800) 421-1968 and ask for the Regulatory Affairs Department. Customers were also instructed to distribute this information to all users of the product and also with the Bio-medical Department at their facility. |  | 260 | 260 | Nationwide | Terminated | April 18, 2012 | 1902 | 305 | Undetermined
Z-0968-2007 | 37300 | No | FZP | Implantable clip. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Medtronic U-CLIP Anastomotic Device Removal Tool, model REM1860FW. The U-Clip Anastomotic Device Removal Tool is designed to remove U-Clip Anastomotic Devices. |  |  | February 08, 2007 | 2007 | Medtronic Cardiac Surgery Technologies | Six lots of U-Clip Removal Tools, catalog REM1860FW were incorrectly assembled. The nose cone attachment to the body of the removal tool does not contain sufficient adhesive which can cause the nose to separate from the body of the device. | Other | Customers to be notified via a letter on 2/8/07. The letter will explain the situation and requests that the impacted product be returned to Medtronic for disposal. The customer will be requested to complete a "Recall Certificate" acknowledging the receipt of this information. |  | 187 | 187 (US 155, OUS 32) | Nationwide and Canada. | Terminated | February 24, 2008 | 381 | 187 | Undetermined
Z-0981-2007 | 37950 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Gyroscan MRI |  |  | April 30, 2007 | 2007 | Philips Medical Systems North America Co. Phillips | Potential for gaseous helium to improperly vent, resulting in rupture of the magnet vessel and venting of helium into the MRI site. | Other | On 4/30/07 Philips sent an Urgent Device Correction letter dated April 25, 2007 via certified mail to customers. The letter advises customers of the issue and that that a representative will visit the sites, inspect and correct the problem if indicated. |  | 293 | 293 units | Units were distributed throughout the US to 1,078 hospitals and medical centers. | Terminated | September 21, 2011 | 1605 | 1372 | Undetermined
Z-0971-2007 | 36955 | No | BZD | Noncontinuous ventilator (IPPB). | Anesthesiology | Anesthesiology | 510(k) | Fisher &amp; Paykel FlexiFix Nasal Mask, Part Numbers: HC405A, HC405SL, HC405U, HCHC406A, HC406SL, HC407A, HC407SL &amp; HC407U. CPAP masks. |  |  | November 24, 2006 | 2006 | Fisher &amp; Paykel Healthcare Inc | Plastic tabs incorporated into the connector components of these CPAP masks may break off if they are not cleaned in accordance with the Instructions for Use.In the event of a tab breaking off, it may enter the CPAP system air path and there is an extremely remote possibility that the broken tab may then enter the patient's nasal cavity or mouth. Masks manufactured since April 2006 do not feat | Other | Fisher &amp; Paykel Healthcare Inc. commenced sending out letters by mail to consignees the week commencing 4 December 2006 advising consignees to quarantine and destroy affected product. A faxback form was also be provided for consignees to advise Fisher &amp; Paykel Healthcare of quarantined and destroyed stock.The firm will supply consignees with replacement product as determined from completed faxback information.The firm will collate responses from consignees and will follow up with a letter to consignees who do not respond within a month of initial mail-out. The firm will then make telephone contact with any consignees who do not respond within a month of the second mail-out.Fisher &amp; Paykel Healthcare Inc. estimated the completion of this voluntary recall by April 2007. |  | 815239 | 815,239 total for all models, all lots | Nationwide | Terminated | March 26, 2012 | 1949 | 815239 | Undetermined
Z-0956-2007 | 37948 | No | NPQ | Factor V Leiden DNA mutation detection systems. | Hematology | Hematology | 510(k) | MagNA Pure LC Total Nucleic Acid Isolation Kit - Large Volume; Catalog # 03264793001. |  |  | May 14, 2007 | 2007 | Roche Diagnostics Corp. | False Positive Results: Product labeling states cross contamination is not an issue for high viral load titer samples, but the firm has determined that a potential sample contamination may not be detected at these levels and could result in false positive results being reported. | Other | Consignees were sent an Urgen Medical Device Correction letter dated 5/14/07 which instructs them to define and establish a lower detection limit, to use UNG to prevent carryover whenever possible, and to confirm low positive results through an independent experiment. A revised device correction letter dated 6/15/07 was subsequently sent to each consignee updating the reason for recall and associated risks. |  | 4201 | 4,201 | Nationwide. | Terminated | December 17, 2009 | 948 | 4201 | Undetermined
Z-0967-2007 | 37493 | No | MED | Liquid chemical sterilants/high level disinfectants. | General Hospital | General Hospital | 510(k) | Steris QFC1728 and QFC1729 Quick Connects used in connections and processing endoscopes. The QFC1728 is used to process the Fujinon 250/270/450/470/490 Group 5 and 530/590 series GI Endoscopes w/o Water Jet. The QFC1729 is used to process 250/450 Group 5 and 530/590 series Endoscopes w/ Water Jet. |  |  | February 23, 2007 | 2007 | Steris Corp | The center insert of the adaptor #6 restrictor could become loose and/or come out of the housing and may affect the sterilization process of the attached endoscope | Process control | On February 22, 2007, customers were contacted by phone and were notified by recall letter dated February 23, 2007. The recall notifications indicate that Steris Account managers will visit each consignee and replace affect product. |  | 146 | 146 | Worldwide-The product was distributed to USA domestic consignees located in MT, CA, NJ, IA, MA, WY, SC, MO, SC, TN, ID, MT, CO, PA, NV, and TX, and to foreign consignees located in Great Britian and Qatar. | Terminated | March 01, 2012 | 1833 | 146 | Undetermined
Z-0957-2007 | 37953 | [6 empty] | BD Microtainer Tube with Glycolytic Inhibitor, Catalog Number 365992, Non Sterile, 200 (4 x 50), 25 degrees C, Becton Dickinson and Company, Franklin Lakes, NJ |  |  | May 17, 2007 | 2007 | Becton Dickinson &amp; Company | Mislabeling : box and bag labels are incorrectly labeled to indicate the product contains Sodium Fluoride/Potassium Oxalate, however the instructions for use and the product itself are correct in the additive indicated and used (Sodium Fluoride/Na2EDTA). | Other | Product Recall letters dated May 1, 2007. were sent to all distributors by first class mail. The letter requested all distributors to examine inventory for specific catalog and lot numbers. It also requested distributors to send a copy of the enclosed customer letter to all customers as a subrecall. BD has contracted a third party (West Inc., Appleton, Wisconsin) to coordinate the mailing to all customers, if the firm chooses to use the third party, or the distributor can send the letters themselves. |  | 142800 | 142,800 worldwide | Worldwide; USA, Canada, Australia, Chile, Columbia, Japan, Malaysia, New Zealand, Singapore, Bangladesh and Indonesia | Terminated | August 15, 2007 | 90 | 142800 | Undetermined
Z-0955-2007 | 37919 | No | DXE | Embolectomy catheter. | Cardiovascular | Cardiovascular | 510(k) | Angiojet XMI Rapid Exchange, catheters, 135 cm, 4 F, Single Use Only, Sterilized with Ethylene Oxide, POSSIS MEDICAL, INC. 9055 Evergreen Boulevard NW Minneapolis, MN 55433-8003 USA-- (XMI - Rapid Exchange Rheolytic Thrombectomy Catheter) is one component of the AngioJet Rheolytic Thrombectomy System (AngioJet System). |  |  | May 01, 2007 | 2007 | Possis Medical, Inc | Mislabeling: XMI (OTW) catheters were mislabeled as XMI-RX units and AVX catheters were mislabeled as Spiroflex catheters. | Other | Consignees were notified of the problem on May 1, 2007 via a Urgent Medical Device Recall Letter. The consignees were instructed to quarantine the product until a Possis Representative contacts them. The representative will examine the units and determine which ones to be returned to Possis. They were also asked to fill and return the attached form. |  | 45 | 45 | Nationwide including states of AK, AR, AZ, CA, CT, FL,IL, IN, KY, MD, ME, MI, MN, MS, NC, ND, NJ, NY, OH, OR, PA, SC, TN, TX, WA, and WI. | Terminated | July 15, 2007 | 75 | 114 | Undetermined
Z-0836-2007 | 36484 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | LIFEPAK CR Plus defibrillator |  |  | October 05, 2006 | 2006 | Medtronic Emergency Response Systems, Inc. | Premature battery depletion due to electrical component failure | Nonconforming Material/Component | Per request from Japan, the firm initiated action on 10/5/06 replacing the devices in Japan with new devices. On 3/2/07 the firm issued a Safety Alert letter dated February 2007 to consignees (with the exception of Japan), advising that the defibrillators have an identified, lead-free component that is susceptible to failure within the first 250 days after manufacture. Failure of this component might cause a short that will drain the battery and eventually render the device inoperative. The letter further states that the firm strongly recommends that the consignee immediately and on a regular basis, inspect their defibrillator(s) per the LIFEPAK CR Plus/EXPRESS Operating Instructions - Section 5 Caring for the Defibrillator. The letter also provides a website to allow consignees to enter their device serial number(s) to determine if their LIFEPAK CR Plus/LIFEPAK EXPRESS are included in the action. This recall strategy is subject to modification pending the performance of devices remaining on the market. |  | 13741 | Total of CR Plus 13,741 units worldwide; amount does not include units distributed to Japan. | Distribution is worldwide. The device is designed for non-professional users and are primarily used in public access defibrillator environments such as airports, schools, and public libraries. | Terminated | August 08, 2013 | 2499 | 14747 | Undetermined
Z-0935-2007 | 37863 | No | BTT | Respiratory gas humidifier. | Anesthesiology | Anesthesiology | 510(k) | HUMIDIFLOW Respiratory Gas Humidifier; Manufactured by: Porous Media Corporation 1350 Hammond Road St Paul, MN 55110, Part Number HMK-RM |  |  | April 26, 2007 | 2007 | Porous Media Corporation | The Intake filter of the Humidiflow kits may became detached from the module and fall onto the compressor cooling fan during patient treatment. (If the filter were to come in contact with the fan blades this could stop the fan from operating and cause the concentrator to overheat and interrupt the flow of oxygen to the patient) | Other | Customers were notified via a letter dated April 26, 2007. The letter describes the problem and product. Included with the letter, is a correction kit and installation instructions. The letter also urges the consignee to take an action to fill the attached form to confirm that the new intake filter restaint has been installed. |  | 178 | 178 | WORLDWIDE DISTRIBUTION: USA including states of GA, MO, FL, UT, CO, TX, MT, MN, PA, OH, AL, SC, VA, KY, IA, MA, TN, IL, NM, UT, NJ, IN, and WI, and countries of CANADA, MEXICO, and JAPAN. | Terminated | December 10, 2011 | 1689 | 178 | Undetermined
Z-0959-2007 | 37994 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Biomet 3 i, Disposable Twist Drills, Model Number DT2715. |  |  | May 14, 2007 | 2007 | Biomet 3i | The depth indicator laser line markings on the drills are too faint for clinicians to read during use. | Other | Each domestic consignee was contacted on 5/14/07 by phone using a prepared script. A follow up letter was then faxed to each consignee. International consignees were contacted by email followed by a faxed letter. Each consignee was instructed to examine their inventory for the recalled product. They were requested to complete and fax back a response form to Biomet and to return any affected product. |  | 147 | 147 | World Wide: USA states including: NY, MA, MD, NJ, IL, DE, RI, CA, MI, VA, KS and fL. Internationally to Brazil, Germany, France, Great Briton Japan, Greece and Italy. | Terminated | September 11, 2009 | 851 | 147 | Undetermined
Z-0950-2007 | 37918 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ARCHITECT i2000 Processing Module, for in vitro diagnostics; List Number 8C89-01; v2.10, v2.11, v2.12, v2.20, v2.20DB, v2.30, v2.60 and v3.00; distributed by Abbott Laboratories, Irving, TX. |  |  | April 30, 2007 | 2007 | Abbott Laboratories, Inc | One or more of four software issues : 1) calculation error for Architect cSystem assays using the Spline math model. 2) The system configuration option, Run controls onboard reagents by Kit, does not function correctly for assays requiring a standard sample dilution. 3) Infrequent data may be missing from automatically or manually printed reports generated on the Architect system. 4) Clinical Chem | Other | Notification letters sent to consignees on 04/30/07. Consignees provided with work around instructions for the software glitches and asked to respond via fax using Customer Reply form acknowledging receipt, understanding and completion of correction letter. This is the short term fix pending software updates, validation,and full root cause determination. |  | 170 | 170 units. | Worldwide distribution, including: USA, Argentina, Australia, Bermuda, Brazil, Canada, Chile, China, Columbia, Costa Rica, Dominican Republic, El Salvador, Germany, Guetamala, Honduras, Hong Kong, Jamaica, Japan, Korea, Mexico, New Zealand, Panama, Peru, Puerto Rico, Singapore, South Korea, Taiwan, Thailand, Trinidad/Tobago, Uruguay, and Venezuela. | Terminated | March 19, 2008 | 324 | 3344 | Undetermined
Z-1131-2007 | 38244 | No | DXE | Embolectomy catheter. | Cardiovascular | Cardiovascular | 510(k) | Possis AngioJet XVG Thrombectomy Sets, |  |  | June 19, 2007 | 2007 | Possis Medical, Inc | Mislabeling : Possis AngioJet XVG Catheter Thrombectomy Sets display incorrect information on the outer packaging box and also on the inner sterile tray. The labeling incorrectly states that the unit is indicated for use in coronary vessels. These units should indicated use only for peripheral arteries. | Other | A Urgent Medical Device recall letter dated 19 June 2007, was sent to consignees advising that all unused product be removed and returned to Possis. An acknowledgement of the disposiiton of the product will be requested from consignees. |  | 194 | 194 | AR, AZ, CA, FL, GA, IL, IN, KS, KY, MA, ME, MI, MN, MO, MS, NC, NH, NJ, NY, OH, OR, PA, SC, TN, TX, WA, WI, WV | Terminated | March 02, 2008 | 257 | 194 | Undetermined
Z-1163-2007 | 38309 | No | ITI | Powered wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Sunrise Medical Quickie Power Wheelchairs: Three models -Rhythm, Groove, and Zippy Z-Bop (with Quickie IQ Enhanced Displays) |  |  | April 23, 2007 | 2007 | Sunrise Medical Inc | The wire harness connected to the display housing (Quickie IQ Enhanced Displays) may separate and cause sparks from a short circuit. | Component design/selection | Firm contacted consignees by phone and mailed May 1, 2007, Urgent Medical Device notification letters. |  | 229 | 229 Enhanced Display units (affecting 220 wheelchairs) | Product was distributed Worldwide. | Terminated | March 25, 2008 | 337 | 229 | Undetermined
Z-1150-2007 | 38458 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Bayonet Electrosurgical Pencil. P/N 9560575. The product is shipped in cases containing 100 blades/pencils. |  |  | July 12, 2007 | 2007 | Zinnanti Surgical Design LLC | Bubble Leak Test Failure of sterility barrier, therefore sterility of the device may be compromised | Other | The recalling firm issued an Urgent Medical Device Recall Letter dated 7/24/07 to its single wholesaler requesting recall and sub-recall to all their consignees. |  | 8860 | 8, 860 units | Nationwide | Terminated | March 07, 2008 | 239 | 8860 | Undetermined
Z-0171-2008 | 37912 | No | LYJ |  |  | Neurology | PMA | Cyberonics Neuro Cyberonics Prosthesis System NCP Bipolar Lead, Model #300; Sterile, Single use only; Distributed by Cyberonics, Inc., Houston, TX 77058. |  |  | August 02, 2007 | 2007 | Cyberonics, Inc | Dissolution/Fractures to the leads of the VNS Therapy System | Process control | Cyberonics issued a correction notice (Safety Alert) to physicians in June of 2007, informing them of the present labeling and including updated labeling to provide instructions on how to detect lead fractures and how to minimize risk to patients. |  | 21419 | 21,419 units. | Worldwide Distribution - USA and countries of Argentina, Australia, Austria, Belgium, Canada, Costa Rica, Cyprus, Czech Rep, Denmark, Finland, France, Germany, Greece, Hong Kong, Hungary, Iceland, India, Ireland, Israel, Italy, Jordan, Korea, Malta, Mexico, Netherlands, Northern Ireland, Norway, Poland, Slovenia, South Africa, Spain, Sweden, Switzerland, Turkey, United Kingdom, and Venezuela. | Terminated | October 05, 2009 | 795 | 56189 | Undetermined
Z-0291-2008 | 45526 | No | HWR | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Pioneer Quantum Spinal Rod System, MIS Yoke Manipulator, Long Slot, non-sterile; Cat # 51-YOKEMAN-L, Pioneer Surgical Technology, Marquette, MI 49855 |  |  | October 10, 2007 | 2007 | Pioneer Surgical Technology | The pins may come out of the instrument and, if this happens, they could fall into/remain in the patient. | Component change control | Consignees were notified via a Product Recall letter dated 10/10/07 to remove these instruments from their instrument cases and to return them to the firm. |  | 38 | 38 | Nationwide: including states of California, Florida, Michigan and North Carolina. | Terminated | March 31, 2008 | 173 | 38 | Undetermined
Z-0245-2008 | 44894 | No | JHI | Human chorionic gonadotropin (HCG) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Clearview hCG Combo Waived for urine &amp; Moderate for serum, Wampole Laboratories, ABI List No/Product No. 6026KCV; Contents 30 Individually pouched test cassettes with disposable pipettes, 1 Directional Insert; Inverness Medical Professional Diagnostics Group; Marketed by Wampole Laboratories, 2 Research Way, Princeton NJ 08540 USA |  |  | September 28, 2006 | 2006 | Applied Biotech Inc | Sensitivity: Kits contain test devices which may not meet the label claims for sensitivity, therefore use of such devices may produce false negative results with patient samples at or near the limit of detection of the product. | Process design | A Product Recall Notification dated 9/28/2006, was provided via fax, email or direct mail to all primary consignees. Customers were instructed to block all stock of affected lots in their warehouses, contact their primary consignees and advise them to discontinue use immediately and dispose of any unused materials; using a fax back form to communicate to the recalling firm the quantity blocked and disposed of so that arrangements could be made for providing replacement product |  | 3238 | 3,238 Kits (97,140 tests) | Worldwide: USA and Canada | Terminated | August 03, 2012 | 2136 | 3638 | Undetermined
Z-0175-2008 | 37880 | No | INI | Motorized three-wheeled vehicle. | Physical Medicine | Physical Medicine | 510(k) | Rascal Scooters with Electric seat lifts, Model 600 Series, distributed between 10/27/2005 - 3/31/2006, Electric Mobility Corporation (EMC), Sewell, New Jersey 08080 |  |  | March 29, 2007 | 2007 | Electric Mobility Corp | Seat and/or lift actuator problems: 1) Models equipped with a Seat Lift, the mast screw and washer holding the Seat Lift Actuator in place will loosen break off or fall out; 2) Models assembled with Seat Lift Actuator or Seat Post Mounting Bolts, the seat may wobble or the seat will break off from the vehicle; 3) Models using Seat-On Lifting brackets with a hoist to raise an EMC vehicle equipped w | Process change control | Consumers were sent an Urgent Medical Device Recall Notification on October 12, 2007, requesting that they immediately respond to the Medical Device Recall by calling EMC at 1-800-814-4365, or by returning an enclosed Recall Response Form, or by responding on-line at www.EMCIT.NET/801B. The Notification address three important safety issues on 1) Seat Lift Actuator Mast Screws and Washers; 2) Seat Lift Actuator or Seat Post Mounting Bolts and; 3) Incorrect Seat-On Lift Bracket Installation. Corrections to the devices will be made by the nearest Service Representatives, by scheduled in-home repairs, or return of the rear section of the scooter with EMC providing shipping material and freight payment. |  | 1756 | 1,756 | Nationwide: to consumers via direct purchase from Electric Mobility Corporation. | Terminated | April 08, 2014 | 2567 | 1756 | Undetermined
Z-0174-2008 | 37879 | No | INI | Motorized three-wheeled vehicle. | Physical Medicine | Physical Medicine | 510(k) | Electric Mobility Corporation (EMC)- Scooters and Turnabout Powerchairs with either an Electric Seat Lift Actuator or a Seat Post to attach the seat to the vehicle; Model Number(s) 300F, 600C, 600T, and 600F. Electronic Mobility Corporation, 599 Mantua Blvd, One Mobility Plaza, Sewell, NJ 08080 |  |  | March 22, 2007 | 2007 | Electric Mobility Corp | Seat and/or lift actuator problems: 1) Models equipped with a Seat Lift, the mast screw and washer holding the Seat Lift Actuator in place will loosen break off or fall out; 2) Models assembled with Seat Lift Actuator or Seat Post Mounting Bolts, the seat may wobble or the seat will break off from the vehicle; 3) Models using Seat-On Lifting brackets with a hoist to raise an EMC vehicle equipped | Process change control | Consumers were sent an Urgent Medical Device Recall Notification on October 12, 2007, requesting that they immediately respond to the Medical Device Recall by calling EMC at 1-800-814-4365, or by returning an enclosed Recall Response Form, or by responding on-line at www.EMCIT.NET/801B. The Notification address three important safety issues on 1) Seat Lift Actuator Mast Screws and Washers; 2) Seat Lift Actuator or Seat Post Mounting Bolts and; 3) Incorrect Seat-On Lift Bracket Installation. Corrections to the devices will be made by the nearest Service Representatives, by scheduled in-home repairs, or return of the rear section of the scooter with EMC providing shipping material and freight payment. |  | 12625 | 12,625 | Nationwide | Terminated | April 08, 2014 | 2574 | 12625 | Undetermined
Z-0219-2008 | 37214 | No | FPO | Wheeled stretcher. | General Hospital | General Hospital | 510(K) Exempt | Stryker Zoom (Atlas) powered patient transport stretcher, Model 660, Stryker Medical, Portage, MI 49002 |  |  | November 15, 2004 | 2004 | Stryker Medical Div. of Stryker Corporation | The brake cams may crack or develop excessive wear, resulting in the brakes not locking the wheels, or the stretcher not being locked firmly in position. | Component change control | The firm issued an Urgent Medical Device Correction Letter dated 9/24/07, to each consignee informing them of the problems, that the brakes on each affected stretcher needed to be inspected/tested and removed from service if brake problems were discovered, and that a Stryker service representative would be in contact to replace the brakes. |  | 81367 | 81,367 (61,708 U.S. and 19,569 ex-U.S.) for all models. | Worldwide-USA, Argentina, Australia, Brazil, Canada, Chile, China, Dubai, France, India, Italy, Japan, Germany, Greece, Hong Kong, Korea, Latin America, Malaysia, Mexico, Netherlands, Poland, Portugal, Singapore, South Africa, Spain, Sweden, Switzerland, Taiwan, and United Kingdom. | Terminated | October 27, 2010 | 2172 | 81367 | Undetermined
Z-0181-2008 | 37661 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE OEC, InstaTrak 3500 Plus FluoroTrak and Fluoro CAT Spinal Applications used with Regular and Super 9800 and 9600 Arm Systems. GE Health Medical Systems, Inc, Salt Lake City, UT 84116-2862 |  |  | March 23, 2007 | 2007 | GE OEC Medical Systems, Inc | Tracking Accuracy - Image intensifier supplier's change on the C-Arm could lead to patient injury, due to incorrect anatomical location of the instrument during a procedure. | Component change control | Consignees were sent an Urgent Recall Notice letter, via mail dated 10/11/07, advising users to stop usage of the FluoroTrak and FluoroCAT spinal applications on the InstaTrak 3500 Plus System. Once a solution is identified, a representative will contact consignees to arrange for installation of the solution at no charge. |  | 61 | 61 units | Class II Recall - Worldwide Distribution - USA including states of AZ, CA, FL, GA, IA, IL, KS, LA, MA, MD, MI, NC, NY, OK, PA, SC, TX, UT, VA, and WI, and countries of Canada, China, Germany, Hong Kong, Italy, Japan, Mexico, and Spain. | Terminated | April 16, 2012 | 1851 | 61 | Undetermined
Z-0218-2008 | 44826 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Texium Closed Male Luer (CML), Model Number 10012241, Cardinal Health, SanDiego, CA. |  |  | August 30, 2007 | 2007 | Cardinal Health 303 Inc DBA Alaris Products | Leaks: there is a risk of leakage at the female end of the Texium closed male luer (CML) when attached to a spin luer. (NOTE: users who connect the female portion of the Texium closed male luer to fixed (one-piece) male luers on tubing or on syringes are not affected by this issue) | Finished device change control | Recall notification was made via a Safety Alert Notification Letter dated August 30, 2007. Recall will be conducted through consignees (distributors) in the affected countries, as well as, direct customers for the Texium closed male luer. The customers and distributors will be notified by registered return receipt mail. For direct customers, separate letters will be sent to the Director of Nursing, Director of Materials and Director of Risk Management, and Director of Pharmacy at each facility. Distributors were sent multiple copies of the Safety Alert Notification and Tip Sheets; along with a cover letter outlining their responsibilities to complete the Safety Alert plan and notify their sub-account customers. The domestic letters (direct customers and distributors) were mailed on August 30, 2007 to all of the affected accounts. The international notifications will be coordinated out of the offices located in the affected countries. The appropriate international regulatory authorities will be notified of this issue, which will work with the country managers at each location. |  | 76300 | 76, 300 units | Worldwide: USA including states of Arkansas, California, Colorado, Hawaii, Illinois, Indiana, Louisiana, Massachusetts, Maryland, Maine, Michigan, Mississippi, North Carolina, New England, New Hampshire, New Jersey, New York, Ohio, Oregon, Pennsylvania, South Carolina, Texas, Utah, Virginia, Washington, and Wisconsin. and countries of Australia, Canada, Japan and Switzerland | Terminated | March 28, 2012 | 1672 | 76300 | Undetermined
Z-0100-2008 | 44792 | Yes | JOH | Tracheostomy tube and tube cuff. | Anesthesiology | Anesthesiology | 510(k) | Portex Perfit Percutaneous Tracheostomy Tube Replacement Kit, Sterile |  |  | September 18, 2007 | 2007 | Smiths Medical ASD, Inc. | Mislabeled. The product is labeled as a size 9.0 mm tracheostomy tube, but the package contains a 7.0 mm tracheostomy tube. | Process control | Smiths Medical contacted direct customers and US distributors by letter sent by UPS on Septmeber 18, 2007. US Distributors instructed to return their inventory and provide a customer list to Smiths Medical. |  | 40 | 40 units | Nationwide. | Terminated | May 08, 2012 | 1694 | 40 | Undetermined
Z-0129-2008 | 45476 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Welch Allyn AED 10, Automated External Defibrillator, a portable 12V internal battery powered defibrillator; Model AED10; Part Numbers: 970302E,970308E,970309E, 970310E and 970311E; |  |  | October 26, 2007 | 2007 | Welch Allyn Protocol, Inc | Failure to Deliver Shock; a defective capacitor may cause the delay or non-delivery of the defibrillating shock which may result in failure to resuscitate the patient | Nonconforming Material/Component | The firm issued a press release on 10/30/07, and has set up a call center for customers, and plans to send three (3) certified mailings to notify affected customers, alert them of the issue, and provide instructions on how to return their device. The first letter dated 10/26/07 was sent via certified mail to customers on 10/26/07. The letter informs customers of the firm recalling the product to correct faulty capacitors on the circuit board, and which could result in a malfunction that prevents or delays delivery of a defibrillating shock. The letter indicates the faulty AED10 devices were manufactured between 3/29/07 and 8/9/07. The letter also indicates the serial numbers affected, and where the serial number is located - upper left hand corner of the back panel of the device. In addition the letter includes a response form, the firm's contact information, and indicates that the firm will exchange the recalled defibrillator with a replacement and new five (5) year warranty. |  | 3588 | 1,794 distributed - 1176 in US and 618 internationally | Worldwide. | Terminated | November 09, 2011 | 1475 | 3588 | Undetermined
Z-0133-2008 | 39354 | [6 empty] | Harleco |  |  | July 30, 2007 | 2007 | EMD Chemicals Inc | Out of Spec -- The Batch (Lot) was incorrectly released after being rejected for out of specification of high azure and eosin absorbance at the end of the run and particulate matter at the end of the run. | Employee error | EMD Chemicals via fax and mailed an Important Product Recall letter dated 7/30/07. The letter notified end use consignee that after conducting an investigation it has determined that Harleco May and Grunwald's Stain; 0.17% w/v Item #660/75 % 660/85 has unacceptable levels of azure and eosin absorbance and undissolved dye matter. Users were requested to discontinue shipping and return product. Wholesalers were advised this is a Customer Level Product Recall and requested to sub recall product sold after 7/5/07. |  | 26 | 26 units (1 liter), 1 unit (4 liter) | Product was released to 2 distributors; one in the US, one in Canada | Terminated | December 26, 2007 | 149 | 26 | Undetermined
Z-0080-2008 | 38775 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | ReUnion Fracture Stem Humeral Component - PureFix HA; Shoulder prosthesis. |  |  | May 09, 2007 | 2007 | Stryker Howmedica Osteonics Corp. | The plasma spray may delaminate from the proximal portion of the humeral stem. | Process control | Market Withdrawal letters were sent on May 9, 2007 via Federal Express. Letters to implanting physicians will be mail on or about mailed November 1, 2007. Stryker is also reissuing original letters indicating that this action is a Recall and advising of possible health hazards. |  | 232 | 232 units | One hospital in Akron OH and 12 Stryker branches/agencies | Terminated | March 07, 2008 | 303 | 232 | Undetermined
Z-0108-2008 | 44929 | No | LQC | Mechanical lithotriptor. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Olympus Single Use Mechanical Lithotriptor, Model number BML-201Q, Olympus America, Inc. |  |  | September 20, 2007 | 2007 | Olympus America Inc. | Slider knob does not move smoothly causing the user difficulty in advancing the coil sheath | Process control | The recalling firm issued a recall letter dated 9/20/07 to the medical facilities informing them of the problem and the need to return the product. The letter also informed the hospitals that replacement products would be sent. |  | 37 | 37 units | Nationwide | Terminated | February 15, 2008 | 148 | 482 | Undetermined
Z-0117-2008 | 44736 | No | FFF | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Boston Scientific Radial Jaw Large Capacity Biopsy Forceps with Needle, Sterile, 240 mm. Box of 20 |  |  | August 28, 2007 | 2007 | Boston Scientific Corporation | Device lacked manufacturing finishing process resulting in endoscope damage. | Process control | Boston Scientific notified accounts by letter on 8/28/07 by Federal Express Priority mail. Accounts are requested to remove product from inventory and return product. |  | 640 | 640 boxes | Worldwide, including USA, Sweden, and South Africa. | Terminated | July 10, 2008 | 317 | 1519 | Undetermined
Z-0104-2008 | 44901 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific mach 1 Guide Catheter, VL3.5 SH Voda Left 3.5, REF 34358-78, UPN H74934358780, Sterile EO, Made in Mexico Avenida Norske Edicficio G1, Local B y H Parque Industrial La Mesa, Fracc. Rubio La Mesa, Tijuana, B.C. Mexico C.P. 22550 |  |  | September 11, 2007 | 2007 | Boston Scientific | Missing holes: Catheters are missing the catheter's side holes. | Process control | A voice mail was left to affected members of the domestic sales force on September 11, 2007 and an interoffice memo with the recall implementation instructions was also sent to the Sales Force. An Urgent Voluntary Medical Device Recall letter dated September 11, 2007, was sent to affected consignees describing the problem and products affected. The letter included the customer steps for recall, a Reply Verification Tracking Form and a return shipping label. |  | 7 | 7 | Worldwide: NY, PA, and Japan | Terminated | February 24, 2008 | 166 | 13 | Undetermined
Z-0102-2008 | 44812 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Biomet Orthopedics Apical Hole Plug, 3/8 - 24 threaded, TI-6AL-4V alloy, sterile; REF 123741, Biomet Orthopedics, Inc., Warsaw, IN 46581 |  |  | August 09, 2007 | 2007 | Biomet, Inc. | Incorrect or lack of threading: The plug component was manufactured incorrectly and does not fit with the screw hole as intended. | Process control | The distributor was notified via phone on or about 8/9/07 and by letter dated 8/31/07 and requested to return the product. Returns included subaccounts |  | 5 | 5 | Nationwide | Terminated | June 26, 2009 | 687 | 5 | Undetermined
Z-0072-2008 | 44780 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Trident PSL HA Cluster Acetabular Shell; |  |  | August 30, 2007 | 2007 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics had identified that specific lots of Trident PSL Acetabular shells may have a dimensional discrepancy. The deviation regarding the difference in wall thickness will increase the gap between the shell and liner on one side and will decrease the gap between shell and liner on the opposing side, resulting in interference. | Process control | Recall notification letters were sent out on August 30, 2007 with return receipt. |  | 142 | 142 units | Product was distributed nationwide and one unit to Brazil. | Terminated | September 05, 2008 | 372 | 142 | Undetermined
Z-0101-2008 | 44810 | No | MED | Liquid chemical sterilants/high level disinfectants. | General Hospital | General Hospital | 510(k) | Steris QFC1728 and QFC1729 Quick Connects;(used to connect gastrointestinal endoscopes to processing trays) The malfunctioning piece of the quick connect component is adaptor #6, STERIS part number 201158. Adaptor #6 attaches to the suction port of the endoscope. |  |  | September 12, 2007 | 2007 | Steris Corporation Hopkins Facility | Loose or missing adaptor: The center insert of the adaptor #6 restrictor could become loose and/or come out of the adaptor's housing. The absence of adaptor #6's center insert may affect the sterilization process of the attached endoscope. | Component design/selection | STERIS has contacted affected direct accounts by phone to notify direct accounts of the recall. On 9/28/2007, STERIS will mail a recall notification to all affected direct accounts via certified mail. The notification informs direct accounts that STERIS representatives will be visiting direct accounts to retrieve affected products and provide replacements. Where direct accounts further distribute, STERIS representatives will inform the direct accounts to contact their sub-accounts and facilitate exchange of affected product for replacement product. Replacement product will be available beginning November 1, 2007. |  | 391 | 391 Quick Connects | Worldwide: United States, Australia, Bahrain, Canada, Germany, Great Britain, Hungary, Italy, Qatar, and Spain | Terminated | September 28, 2010 | 1112 | 391 | Undetermined
Z-0058-2008 | 44938 | No | EXF | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Remington Medical, Inc, (RMI) -- Urinary Drainage Bags: 500ml Disposable Urinary Leg Bag (Order number: 500-D) and 1000ml Disposible Urinary Leg Bag (Order number 1000-D) -- Sterile Fluid Path, For Single Use Only, Remington Medical, Inc., Alpharetta, GA 30005. |  |  | July 23, 2007 | 2007 | Remington Medical Inc. | Sterility, inadequate -- Pathway may not be sterile as labeled. | Process control | Consignees were notified by Urgent Recall Letter (dated 07/23/07) via fax from 07/24/2007 through 07/27/2007. They were instructed to quarantine all affected product in inventory. If product was distributed they were instructed to identify their customers and notify them of the recall via copy of the notification letter. |  | 45000 | 45,000 units | Nationwide Distribution -- including states of AL, AR, AZ, CA, CO, CT, GA, HI, IL, IN, KS, KY, MD, MI, MT, NB, NC, OH, NY, PA, RI, TN, WA, WV, and WY. | Terminated | April 06, 2009 | 623 | 45000 | Undetermined
Z-0066-2008 | 38739 | No | HMY | Keratome. | Ophthalmic | Ophthalmic | 510(k) | BD Microkeratome K3000 and K4000;Catalog Number: 378660, Becton Dickinson &amp; Company (BD) |  |  | August 09, 2007 | 2007 | Becton Dickinson and Company | Microkeratome blade may stop oscillating and transverse the eye causing injury to the cornea | Component design/selection | BD Medical notified customers by Advisory Letter dated 8/09/07, advising users to return heads and bearing cartridges for replacement of the o-ring. |  | 256 | 256 units | Worldwide | Terminated | February 29, 2012 | 1665 | 256 | Undetermined
Z-0046-2008 | 44940 | [6 empty] | Microtainer Safety Flow Lancets; Blue |  |  | September 26, 2007 | 2007 | Becton Dickinson &amp; Company | Potential for the blade and holder assembly to become loose and then separate from the main body of the lancet during the process of disposal. | Process design | Recall communication issued on or about 9/28/2007. |  | 4349000 | 4,349,000 units | Product distributed worldwide to Australia, Brazil, Canada, China, Colombia, Guatemala, Hong Kong, India, Indonesia, Mexico, Malaysia, New Zealand, Pakistan, Philippines, Singapore, Thailand, and United States. | Terminated | January 15, 2008 | 111 | 7803000 | Undetermined
Z-0043-2008 | 44909 | No | DOF | Lead test system. | Toxicology | Toxicology | 510(k) | ESA LeadCare II Blood Test Kit, 6 test Kit containing Level 1 and Level 2 Control Vials with lot number: 0704A |  |  | September 20, 2007 | 2007 | Esa Biosciences, Inc. | Level 1 and Level 2 Control values were incorrectly assigned. | Process design | ESA notifed Distributors and End Users by telephone and letter on 9/20/07. Distributors and End Users will be informed of the Level 1 and 2 Control value assignment change. End Users will be sent a new Level 1 and 2 Control Viallabel for each vial they originally received. Distributors will be asked to return all test kits in their inventory to ESA Biosciences. |  | 2 | 2 kits | Nationwide and United Kingdom. | Terminated | March 25, 2008 | 187 | 367 | Undetermined
Z-0038-2008 | 38511 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Pump waste tube for blood gas analysis kit. One yellowish/white rubber tube per plastic bag. Used on ABL800 Flex Series blood gas analyzers. |  |  | June 22, 2007 | 2007 | Radiometer America Inc | Tube may break. The firm discovered that the yellowish/white tubes used in their blood gas analyzers (ABL800 Flex Series) were manufactured incorrectly. These tubes have the potential to break off at the black plastic connectors. The breakage may occur during the installation of the tubes or after a short time in use. | Process design | By telephone call or service representative visit beginning July 18, 2007, Consumers/End Users are asked to locate the analyzers and yearly service kits concerned and replace all pump tubes with new ones. |  | 289 | 289 pieces | Product was distributed throughout the United States and Canada. | Terminated | April 25, 2012 | 1769 | 289 | Undetermined
Z-0039-2008 | 39419 | No | GKP | Coagulation instrument. | Hematology | Hematology | 510(k) | STA Compact and STA Compact CT, Haemostasis Analyzer, in vitro diagnostic. |  |  | July 30, 2007 | 2007 | Diagnostica Stago, Inc. | Incorrect results. Internal investigation determined the presence of a defective measurement head which can impact the results of clotting tests by either yielding no result or rendering an error message. | Process design | Recall notification letters were mailed to consignees on 7/30/07 by first class mail. |  | 41 | 41 units installed | Nationwide. | Terminated | December 17, 2007 | 140 | 41 | Undetermined
Z-0011-2008 | 39367 | No | JSO | Culture medium for antimicrobial susceptibility tests. | Microbiology | Microbiology | 510(k) | Brain Heart Infusion Agar (BHIA) + Vancomycin 6 MCG/ML |  |  | August 17, 2007 | 2007 | Pml Microbiologicals Inc. | Marketed a regulated product without a 510(k). | PMA | Firm notified consignees by phone and fax on August 17, 2007, instructing them to destroy the product. |  | 2245 | 1205 plates distributed in US; 1040 distributed in Canada | US, Puerto Rico, &amp; Canada | Terminated | June 07, 2011 | 1390 | 2245 | Undetermined
Z-0014-2008 | 39382 | No | JSO | Culture medium for antimicrobial susceptibility tests. | Microbiology | Microbiology | 510(k) | Brain Heart Infusion Agar labeled BHIA + Vancomycin 6 MCG/ML For Screening for Vancomycin Resistant Enterococci PML microbiologicals Wilsonville, Oregon 97070 USA |  |  | April 23, 2007 | 2007 | Pml Microbiologicals Inc. | Plates manufactured without vancomycin. Possibility for false resistant results. | Other | On 4/23/07 the firm began contacting customers via fax and telephone. |  | 204800 | 640 total plates, 320 per lot (lot 229957-1 was dist only in Canada; lot 229958-1 was dist in the US and Canada) | lot #229957-1: 120 plates were distributed to 9 Canadian consignees; 200 plates were shipped, quarantined, and destroyed at the Canadian distribution center.lot #229958-1: 110 plates were distributed to 9 US consignees; 210 plates were shipped, quarantined, and destroyed at the Canadian distribution center. | Terminated | June 07, 2011 | 1506 | 204800 | Undetermined
Z-0057-2008 | 44933 | No | MED | Liquid chemical sterilants/high level disinfectants. | General Hospital | General Hospital | 510(k) | ON-Q PainBuster with OnDemand (270 ml, 2 ml/hr + 5 ml bolus /60 min refill), Model PMB01, I-Flow Corporation, Lake Forest, CA 92630 U.S.A. |  |  | March 27, 2007 | 2007 | I-Flow Corporation | Misassembly: The ON-Q PainBuster PMB01 package may contain a Soaker catheter instead of a Non-Soaker catheter. As a result, the bolus button may remain latched in the down position which may cause the the device to deliver continuously at 7 ml/hr instead of 2 ml/hr. | Process control | An Important Recall Notification Letter was faxed to affected customers on March 27, 2007. Affected customers were notified of the recall, the reason for recall, and instructed to immediately quarantine any remaining inventory. A Customer Response (and Fax back) form was included with the Recall Notification. Affected customers were requested to return the FAX back form to I-Flow within 48 hours of receipt. Replacement product will be sent out as applicable. |  | 2587 | 2,587 units | Worldwide, USA and Australia | Terminated | October 07, 2009 | 925 | 2587 | Undetermined
Z-0034-2008 | 39412 | No | MID | Multiple autoantibodies immunological test system. | Immunology | Immunology | 510(k) | APhLÂż HRP IgG &amp; IgM ELISA Kit; Ref. #LAPL-K-HRP-00GM; in vitro diagnostic. Distributed by Louisville APL Diagnostics, Inc., Doraville, GA 30360 (address now in Seabrook, TX 77856). |  |  | July 18, 2006 | 2006 | Louisville Apl Diagnostics Inc | Product mislabeled; outer container labeled with incorrect lot number. | Process design | Firm sent replacement labels to consignee via FedEx on 07/18/06 with instructions for consignee to change the labels on the kit boxes before shipping kits to end users. |  | 3 | 3 test kits (each kit containing 1 IgG and 1 IgM test kit) | Product was shipped to a distributor in Canada. | Terminated | October 24, 2007 | 463 | 3 | Undetermined
Z-0006-2008 | 38484 | No | BSF | Carbon dioxide absorber. | Anesthesiology | Anesthesiology | 510(K) Exempt | GE Disposable Multi Absorber Canister used with EZchange Module for the Aisys, Avance, Aespire Anesthesia systems. The EZchange Module is an optional accessory for these anesthesia systems. |  |  | July 24, 2007 | 2007 | Datex - Ohmeda, Inc | Elevated inspired CO2 levels related to the use of the Datex-Ohmeda GE Healthcare Disposable Multi Absorber Canister with the EZchange Module. | Vendor change control | Because this issue only occurs when the Disposable Multi Absorber is used in conjunction with EZchange Module, they have contacted all EZchange Module customers. It is critical to contact these customers specifically, rather than all customers shipped Disposable Multi Absorbers as this may cause unwarranted confusion with customers who do not have the EZchange Module. Consignees were sent an letter dated 7/19/2007 describing the product and problem. The letter lists a number of interim actions to implement with the anesthesia systems. |  | 389070 | 389,070 | Worldwide, including USA, Australia, Belgium, Brazil, Canada, Chile, China, Denmark, Egypt, Estonia, Finland, France, Germany, India, Ireland, Italy, Jordan, Republic of Korea, Kuwait, Mexico, Netherlands, New Zealand, Norway, Peru, Poland, Qatar, Singapore, Slovak Republic, South Africa, Spain, Sweden, Switzerland, United Arab Emirates, United Kingdom, and Venezuela. | Terminated | October 16, 2015 | 3006 | 389070 | Undetermined
Z-0007-2008 | 38494 | No | JJX | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | 4 glass vials per box. 2/4ml vials of Sample Diluting Solution &amp; 2/2ml vials of Calibration Verification Material. BetaHCG Calibration Verification Test Set, in vitro diagnostic. |  |  | June 22, 2007 | 2007 | Tosoh Bioscience Inc | Firm discovered that a deterioration value of approximately 20% was found when the product of this lot was measured. Tosoh AIA confirmed that deterioration of the product occurred over time and resulted in a deviation from the assigned value. | Nonconforming Material/Component | Letter, dated July 2, 2006. Immediately examine inventory and quarantine product. If product was further distributed, the direct accounts are to notify their customers including a copy of the original recall letter and Notification of Recall response card. |  | 47 | 47 Boxes | Nationwide. | Terminated | October 07, 2010 | 1203 | 47 | Undetermined
Z-0296-2008 | 45529 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Xoran MiniCAT for ENT, Xoran Technologies, Incorporated, Ann Arbor, MI 48103. This device is a cone beam CT scan for sinue, temporal bones and maxillofacial imaging. |  |  | July 13, 2006 | 2006 | Xoran Technologies, Inc. | The hard drive may crash, resulting in loss of data from patient CT scans. | Finished device change control | Consignee accounts were visited between 7/25/06 and 8/10/07 and the hard drives on their systems were replaced. |  | 5 | 5 | Nationwide Distribution: California, South Dakota, Tennessee, Utah and West Virginia. | Terminated | March 13, 2008 | 609 | 5 | Undetermined
Z-0179-2008 | 45436 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Edwards Lifesciences LifeStent FlexStar Bilary Stent System Application, Model EX080801CD; 8 MM x 80 MM stent, 80 CM Delivery System, Lot FR6K0319, Edwards Lifesciences, LLC., Irvine, CA 92614 |  |  | March 19, 2007 | 2007 | Edwards Lifesciences, Llc | Failure to deploy. | Process control | This recall was executed by an Urgent-Product Recall letter dated March 15, 2007, which was delivered via overnight service to the affected accounts on 3/19/07. In addition, the affected sales force was notified by Email and given a copy of the recall letter. The sales force was asked to assist in retrieval of the product. The letter advises that one lot of the Edwards LifeStent FlexStar Bilary Stent System -- Model EX080801CD; 8 MM x 80 MM stent, 80 CM delivery system, Lot FR6K0319 may exhibit incidences of failure to deploy and is being recalled. Customers are advised to check their inventory for lot #FR6K0319 and remove all unused product from stock. Product is to be returned or it will be retrieved. The letter also advises that if the product has been implanted, there is no need to inform the patient of this notification as this is solely related to the delivery system and the characteristics of the stent are not affected in any way. Either the customer or the sales representatives will return written notification of the status of product from this lot which had been shipped. |  | 12 | 12 units | Nationwide including states of OK, GA, CA WI, MO, and LA | Terminated | March 26, 2012 | 1834 | 12 | Undetermined
Z-1166-2007 | 38209 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Orbit 90 Subcutaneous Infusion Set, 9 mm Catheter, 42" Tubing - Blue, List #P4291 |  |  | September 26, 2006 | 2006 | ICU Medical, Inc. | The product contains a female luer which has been reported to crack in use allowing fluid to leak from the device. | Other | Firm mailed a letter dated September 26, 2006 to notify customers of the recall. The letter advises the product contains a female luer which has been reported to crack in use allowing fluid to leak from the device. Customers are asked that if they have inventory of the lots identified to quarantine these lots and contact your consignees of this recall and return the product immediately for destruction. They are asked to complete and return the attached product inventory summary sheet in the enclosed self-addressed stamped envelope and return product to ICU Medical. |  | 26361 | 26,361 units | VA, OH, TX, CA, GA | Terminated | March 22, 2012 | 2004 | 26361 | Undetermined
Z-0177-2008 | 44926 | No | DQA | Oximeter. | Cardiovascular | Anesthesiology | 510(k) | Capnostream20 Monitor with Nellcor Pulse Oximetry Module--Part Number: CS 08657- Oridion Medical |  |  | July 28, 2007 | 2007 | Oridion Capnography Inc | Handle Cracking - Cracks in the handle will lead to disconnection of the handle from the monitor, causing the monitor to fall. | Process design | Oridion sent Dear Customer letters dated, August 21, 2007, informing users that a new monitor with the improved handle will be shipped by August 31, 2007. |  | 21 | 21 | Nationwide Distribution- including states of AZ, CA, FL, MA, MO, OH, TX, WA, and WI, | Terminated | September 08, 2008 | 408 | 27 | Undetermined
Z-0456-2008 | 45442 | No | JTR | Quality control kit for culture media. | Microbiology | Microbiology | 510(K) Exempt | MicroBiologics; P. vulgaris, DuoPacksl Product Catalog Number: 0691P; ARCC 0427; Exp 2008-9; MicroBiologics, St. Cloud, MN 56303. |  |  | October 12, 2007 | 2007 | Microbiologics Inc | Mislabeled: DuoPacks of Proteus vulgaris contained Listeria monocytogenes instead of the labeled Proteus vulgaris. | Employee error | Consignees were initially notified of this recall by a telephone survey. An Urgent Product Recall letter was sent on 10/17/07 explaining the risk to patients. The firm sent replacements of the recalled product on 10/17/07 or 10/18/07, with request for acknowledgement of Recall Notification, and certification of disposal of incorrect product. |  | 20 | 20 | Worldwide Distribution - USA including states of OR, PA, NY, RI, GA, SC, IL, PR, CA, AZ, and NJ, and countries of Japan, France, Colombia, and Brazil. | Terminated | June 26, 2008 | 258 | 20 | Undetermined
Z-0211-2008 | 45426 | No | CAF | Nebulizer. | Anesthesiology | Anesthesiology | 510(k) | HUDSON RCIÂż, Variable Concentration, Large Volume Nebulizer, Catalog Number (REF) 1770 and 41770, Rx Only, Teleflex Medical, Research Triangle Park, NC 27709. |  |  | August 22, 2007 | 2007 | Teleflex Creek Dr | Decreased in flow output; through the nebulizer -(when used at higher oxygen concentration settings, this decrease in flow, may lower the oxygen concentration inspired by the patient) | Release of Material/Component prior to receiving test results | Customers were notified by an Urgent Medical Device Recall letter on/about 08/22/2007. They were instructed to return the affected products to Teleflex Medical or acknowledge that the products no longer exist. In addition, each were instructed to forward the letter if they have further distributed into institutional settings or to homecare patients and retrieve relevant product from those locations. This recall is to the consumer/user level via sub-recall from distributors or direct accounts as per the Recalling firm. |  | 215656 | 215,656 units | Worldwide: USA, Brazil, Canada, Chile, Colombia, Costa Rica, Ecuador, France, Germany, Guatemala, India, Mexico, Panama, Peru, Philippines, South Korea, Thailand, United Arab Emirates, and The United Kingdom. | Terminated | August 13, 2012 | 1818 | 215656 | Undetermined
Z-0149-2008 | 45404 | No | LJT | Subcutaneous, implanted, intravascular infusion port and catheter. | General Hospital | General Hospital | 510(k) | VortexÂż VX, Port System, 9.6FR, REF/Part Number : P5455K, Catalog Number: P5455, Latex Free, Sterile, AngioDynamics Incorporated, Manchester, GA. 31816 |  |  | October 03, 2007 | 2007 | Rita Medical Systems, Inc. | Incorrect Size Catheter: The product may contain an incorrect size catheter. | Employee error | Notification of the AngioDynamics, Inc. sales force and the domestic and international distributors began on 10/3/2007. Starting on 10/09/2007, Urgent Medical Device Recall letters were sent to all customers requesting that device useage cease and immediate return begin as well as the completion of a fax back form (number of devices used, number to be returned). |  | 150 | 150 kits | Worldwide : USA including states of AL, AR, CA, CO, FL, GA, MD, NC, OH, TX, and VT, and countries of Australia, Korea and The Netherlands. | Terminated | July 14, 2008 | 285 | 150 | Undetermined
Z-0153-2008 | 45435 | No | FYA | Surgical apparel. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Precept Medical Products, Tape Fog Shield Surgical Mask Green Fiberglass Free, REORDER #: 65-3322, non-sterile, 50/box, 5 box/case, Precept Medical Products, Inc., Arden, NC 28704, 1-800-438-5827, Mexico |  |  | September 14, 2007 | 2007 | Precept Medical Products, Inc. | Small slits under the folds of the masks; thus they may not provide the protection necessary between the patient and surgical staff | Process control | The firm notified its consignees on 09/14/2007 by Urgent Voluntary Medical Device Recall letter (USPS mail, certified, return-receipt requested). The recall was expanded on 10/04/2007 by Urgent Voluntary Medical Device Recall--Expansion Letter to include additional lots of product. Recall letters instruct users to cease use of devices and return them as well as conduct sub-recalls to the users level for any devices they have sub-distributed. |  | 11546350 | 11,546,350 (both products-Z-0153/0154-2008) | Worldwide; USA, Australia and Canada | Terminated | September 11, 2009 | 728 | 11546350 | Undetermined
Z-0533-2008 | 44734 | Yes | BYE | Positive end expiratory pressure breathing attachment. | Anesthesiology | Anesthesiology | 510(k) | Ambu, Positive End Expiratory Pressure (PEEP) Valves, PEEP Valve 10, 30 mm, REF 000 137 000, Manufacturer: Ambu International A/S, DK2750 Ballerup USA: Ambu Inc. (Ambu USA) Linthicum, MD 21090-1356 USA, Made in Denmark |  |  | August 21, 2007 | 2007 | Ambu Incorporated | Lower pressure than expected: Medical device for respiratory care may leak and not register accurate flow settings that affect patient respiration. | Process design | Ambu notified consignees with a letter dated 08/21/07 and flagged as Trouble Shooting Guide. The letter advised that some valves do not hold the set PEEP level under certain conditions and consignees were further advised to consider the guideline as an addendum to existing directions for use. Consignees were instructed to return (to Ambu) for replacement the identified lots listed in the notification and confirm receipt of the advisory by e-m or fax. Ambu subsequently identified additional lots of valves and replacement valve parts for recall and notified all consignees by "Urgent Device Recall" letter dated 11/29/07. The letter advised healthcare practitioners of the impact on patient treatment that may result in less than effective treatment. Consignees were further advised to cease use of the device and return for replacement. |  | 197 | 197 total | Nationwide. | Terminated | June 16, 2008 | 300 | 197 | Undetermined
Z-0137-2008 | 44928 | No | IOR | Mechanical wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Freedom Designs TriPod w/Anti tipper Manual Wheelchair, Models Numbers: TriPod, Tripod LT and TriPod Pediatric LT with Trailing Anti-Tipper-Caster-Axle Link, Freedom Designs Incorporated, Simi Valley, CA 93065 |  |  | May 14, 2007 | 2007 | Freedom Designs Inc | Tip-over hazard; the Anti tipper may fail/break and the wheel chair could tip over. | Process control | Urgent Field Correction Notification letters were sent to customers on May 15, 2007 via US Postal return receipt requested. Consignees were notified via delivery of a field correction package consisting of:(1) A letter to the dealer (2) FAQ sheet clarifying their responsibility (3) A corrective/preventive action tracking sheet specific for the consignees affected units(4) Replacements of anti-tipper assembly for each unit shipped with an instruction sheet.A Notice was also posted on the Freedom Designs website.Correction will be made to the consumer level |  | 41 | 41 units | Nationwide: including states of AL, CA, FL, CO, GA, KS, LA, MI, MS, NC, OH, TN, TX, and VA | Terminated | July 14, 2008 | 427 | 41 | Undetermined
Z-0145-2008 | 44958 | No | MDK | Streptococcus spp. serological reagents. | Microbiology | Microbiology | 510(K) Exempt | Accuprobe Group B Streptococcus Culture Identification Test, Catalog Number 2820B, Batch Number 531379, Gen-Probe, Incorporated, San Diego, CA. 92121 |  |  | May 02, 2007 | 2007 | Gen-Probe Inc | Mispackaging: Kits may contain Haemophilus Influenza probe pouches in addition to Group B Streptococcus probe pouches | Other | Gen-Probe notified customers by phone on May 2nd and 3rd, 2007; a recall notice letter was faxed to each customer on May 3rd, 2007 and mailed on May 4th, 2007 via certified mail. The letter describes the packaging error and the remote health hazard risk, and provides instructions for users to return all kits of this lot number. |  | 106 | 106 kits | Nationwide: including states of AZ, CA, CO, FL, KY, MD, MO, NC, NJ, TX, and UT | Terminated | May 03, 2010 | 1097 | 106 | Undetermined
Z-0185-2008 | 44800 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | EncoreÂż ReverseÂż Shoulder Prosthesis (RSP) System, Standard Size 40 RSP Humeral Socket Insert; Catalog #: 508-00-040; Sterile; Manufactured by Encore, Medical, L.P., Austin, TX 78758. |  |  | August 08, 2007 | 2007 | Encore Medical, Lp | Devices Mislabeled; one lot of Size 36 Standard RSP Socket Inserts were labeled as Size 40 Standard RSP Sockets and vice versa. | Other | Encore notified their sales reps (direct accounts) via telephone and overnight letter on 08/08/07. Encore sales reps were asked to return all outstanding devices, for which they would be given replacements. Notification of implanting surgeons was done by phone and letter (Urgent : Device Recall) between 08/07 and 09/12/07. Implanting surgeons notified of recall and asked to closely monitor patients who had received and already been implanted with the affected lots. If explant was required decision of the surgeons, they were asked to notify the recalling firm. |  | 25 | 25 units. | Nationwide:Direct accounts and surgeons in the following states: CA, GA, IA, MA, MD, MI, MO, SC, TX and UT. | Terminated | March 10, 2008 | 215 | 49 | Undetermined
Z-0136-2008 | 44883 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | DRL System Trays comprised of DRLock Bone Fixation Plates, Pegs, and Screw Systems: Four Plate Styles include: Part Numbers: DRV-002-0L Fixed Short Left Plate, DRV-002-0R , Fixed Short Right Plate, DRV-002-1L, Fixed Standard Left Plate, and DRV-002-1R, Fixed Standard Right Plate; Distal Screw Part Numbers: DRV-021-24-10, -12, -14, -16, -18, -20, -22 and -24; Distal Peg Part Number: DRV-121-20-10, -12, -14, -16, -18, -20, -22, and -24; All screw caddies and any distal screw and peg inventory not yet added to screw caddies, Ortho Helix Surgical Designs, Inc., Akron, OH |  |  | September 04, 2007 | 2007 | OrthoHelix Surgical Designs Inc | Seat/Fit problems: The heads of the distal screws/pegs rest above the surface of the distal plate after the screws/pegs are fully tightened. | Component change control | On 9/4/2007, OrthoHelix mailed an Important Recall Notification letters via certified mail to all direct accounts. The notification instructed all accounts to examine their inventory and return all recalled product, identified within the notification, to OrthoHelix for replacement. |  | 13884 | 267 plates/52 DRL System Trays | Nationwide; including states of IL, CA, VA, TX, MD, UT, GA, CO, NJ, OH, &amp; WI. | Terminated | May 20, 2011 | 1354 | 13884 | Undetermined
Z-0317-2008 | 38468 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Zimmer Minimally Invasive Solutions Technology Offset Rasp Handle, Right, 45 degrees, non-sterile; REF 7712-35-02, Zimmer U.K. Ltd., SN3, 4FP, UK |  |  | June 28, 2007 | 2007 | Zimmer Inc. | Clip Detachment: The retainer c-clip can detach from the instrument during use and may fall into the surgical wound | Process control | Zimmer notified their distributors via Product Recall Notification Letter dated 6/30/07 and instructed them to notify their customers of the recall and to pick up the product and return it to Zimmer. |  | 62 | 62 (total for Z-0316-0317-2008) | Worldwide: USA including states of Arizona, California, Florida, Georgia, Indiana, Michigan, New York, North Carolina, and Virginia, and countries of Germany and Japan. | Terminated | March 17, 2008 | 263 | 62 | Undetermined
Z-0475-2008 | 45982 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Hipstar TMZF Cementless Hip Stem Protocol, Product Number: LSP60. |  |  | September 28, 2007 | 2007 | Stryker Howmedica Osteonics Corp. | Mis-labeling: An error on page 4 of the Hipstar protocol was found. In the table for compatible Heads, the 40 mm diameter heads with a Neck Length of +12 should have been omitted. | Other | Urgent Product Correction Letters, dated 10/17/2007 were sent on 10/17/07by Federal Express to all Stryker Branches/Agencies who had received the protocol. The branches were instructed to quarantine and destroy the protocols. |  | 192 | 192 protocols | Nationwide:The protocol was distributed to 9 Stryker branches/agencies in NJ, MD, IN, MI, MN, IL and TX. | Terminated | September 05, 2008 | 343 | 192 | Undetermined
Z-0458-2008 | 45483 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Hypoguard Advance Micro-draw Blood Glucose Monitoring System, Product Number : 411100, For In Vitro Diagnostic Use, Manufactured for ARKRAY USA, Inc., Minneapolis, MN 55439 USA. (Package Contents -Advance Micro-draw Blood Glucose Meter, 10 test Strips, Codeing chip and Control Solution (3mL), Lancing device, 10 Lancets, 3V Lithium Battery, Logbook, User Instruction Manual, Quick Reference Guide) |  |  | October 18, 2007 | 2007 | ARKRAY USA INC. | Mis-labeled/coded strips : Meter kits from this lot contain mislabeled 10-count test strip vials. The code number printed on the Advance Micro-draw 10-count test strip bottle is incorrect. The code number on the vial is 867 when it should be 687. The 687-labeled code chip contains the correct calibration information for the "867" test strips, so the meter will function properly and give accurat | Process control | Distributors were sent a letter on 10/18/07 asking that product be quarantined. On 10/26/07, an Urgent Voluntary Medical Device Recall letter was sent asking that they return affected product and contact their customers to the retail level. |  | 2005 | 2,005 Kits | Nationwide: USA including states of CA, MI, MO, MS, NE, OH and TX. | Terminated | March 02, 2008 | 136 | 2005 | Undetermined
Z-0459-2008 | 38897 | No | JJF | Micro chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Roche Diagnostics LightCycler Capillaries (100uL) for use in conjunction with the LightCycler analyzers; Catalog #: 03337090001, plus any capillaries received at the time of installation of the LightCycler instruments or systems **, Roche Diagnostics Corporation, Indianapolis 46250.(**LightCycler 2.0 System, Catalog Number: 04391039001; Batch Number Range: 1415471-1415933; LightCycler 2.0 Instrument-Laptop; Catalog Number: 03603946001; Batch Number Range; 1415434-1415734; LightCycler 2.0 Instrument-Desktop; Catalog Number: 03531414201; Batch Number Range: 1415271-1415732 |  |  | August 30, 2007 | 2007 | Roche Diagnostics Corp. | False Negative Results: Air bubbles in the lens or the area surrounding the lens of the LightCycler 100 uL capillaries may cause false negative results. The affected capillaries showed up to a 40% reduction in signal | Process design | Consignees were notified by Urgent Medical Device Recall letter dated 8/30/07, to discontinue use of the 100uL capillaries, plus all of those received at installation, and to disposed of them per local guidelines. |  | 11424 | 119 boxes of 96 capillaries each | Nationwide | Terminated | September 08, 2008 | 375 | 11424 | Undetermined
Z-0454-2008 | 45887 | No | HNM | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | BD Visitec Nucleus Hydrodissector, .40 x 22 mm, (27 G x 7/8 in), Catalog #/REF #: 585793, Becton, Dickinson and Company, Waltham, MA 02452 |  |  | November 20, 2007 | 2007 | Becton Dickinson and Company | No Fluid Flow or Restricted Fluid Flow: Nucleus Hydrodissector(s) were found to have reduced or no fluid flow | Process control | The recalling firm notified consignees of the recall through a Voluntary Medical Device Recall letter and phone call on 11/20/2007. Instructions for product return were provided and a telephone number for questions was given (Customer Service 1-866-906-8080 Option 2) |  | 43000 | 43,000 | Nationwide-USA including the states of CA, CO, CT, FL, GA, HI, IL, IN, KS, LA, MA, MD, MO, NC, NJ, NY, OH, PA, SC, TN, VA, and WA | Terminated | February 29, 2012 | 1562 | 83450 | Undetermined
Z-0452-2008 | 45867 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Conmed Accessory Electrode, Universal Electrodes, Part/REF Numbers : 138100 through 138-105, 138107, 138110, 138112 and 138113, and Ultra Clean Accessory Electrode, Part/REF Numbers: 139100, 139107, 139108, 139321, 139112, 139104 EXT. 139105EXT, 139110EXT, and 139112EXT, Sterile, Rx Only, Conmed Corporation, Utica, New York 13502. |  |  | November 02, 2007 | 2007 | ConMed Electrosurgery | Sterility (package integrity) Compromised: packages of the electrosurgical blades and needles have an inadequate packaging seal | Process design | Consignees were notified by Urgent Voluntary Device Recall letter on 11/02/2007, from ConMed Corporation. They were asked to return affected product for replacement. |  | 11319 | 11,319 cases (476,160 units) | Worldwide: USA and countries of Australia, Canada, China, Belgium, Chile, Denmark, Ecuador, France, Greece, India, Israel, Italy, Japan, Korea, Mexico, New Zealand, Norway, Philippines, Portugal, Saudi Arabia, South Africa, Spain, Thailand, Trinidad, and Uruguay. | Terminated | March 31, 2010 | 880 | 11319 | Undetermined
Z-0054-2008 | 45525 | No | HIS | Condom. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Lubricated Latex Condom. Labeled in part: "COMING TOGETHER IN DC", Distributed by Global Protection Corp., #12 Channel Street, Boston, MA |  |  | September 11, 2007 | 2007 | Global Protection Corp. | Condoms fail the water leak test and package integrity | Process control | Global Protection Corp. recalled product by email on 9/11/07 |  | 12700 | 127,00 condoms | The District of Columbia Health Department | Terminated | March 18, 2008 | 189 | 12700 | Undetermined
Z-0363-2008 | 45523 | No | MNI | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Ascent POCTS (Posterior Occipital Cervical Thoracic System) Part Numbers: 65-33 XX Series (65-3310,-3312, -3314, -3318, -3320-, -3322, -3324, -3326, -3328, and -3330), Multi-Axial Screw, 4.0mm x 26mm;Implant grade Titanium Alloy; Blacksone Medical Inc., 90 Brookdale Drive, Springfield, MA 01104. |  |  | May 15, 2007 | 2007 | Blackstone Medical, Inc. | Manufacturing Defect: The saddle of the screw's body relative to the keyway exceeds the tolerance outlined in the device specifications. The extent to which the tolerance is exceeded varies from part to part, in the extreme the tolerance is exceeded by (+) 50%. As a consequence, there is a potential for set-screw cross-threading during implantation. Cross-threading could in turn lead to a comprom | Nonconforming Material/Component | Blackstone Medical, Inc., field representatives were notified via phone, e-mail or fax on 4/30/07. Attention: Notification of Product Removal Letters were sent on May 15, 2007, to distributors and implanting surgeons requesting return of non-implanted stock. |  | 319 | 319 screws ( 212 domestic, 107 international) | Worldwide: USA and countries of China, Cyprus, Turkey and Switzerland. | Terminated | June 06, 2008 | 388 | 319 | Undetermined
Z-0362-2008 | 45484 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | VKS Tibial Trial, Right/Left, Prep Guide; Model Number: 1214331, 1214332, 1214333, 1214334, 1214335: ( Model Number: 1214331 VKS Tibial Prep Guide, Right/Left, Size 1, Lot Code EM. 1 and EM.6; 1214332 VKS Tibial Prep Guide, Right/Left, Size 2 Lot Code EM.l and EM.6, 1214333 VKS Tibial Prep Guide, Right/Left, Size 3 Lot Code EM.1; 1214334 VKS Tibial Prep Guide, Right/Left, Size 4 Lot Code EM.5; 1214335 YKS Tibial Prep Guide, Right/Left, Size 5 Lot Code EM), Plus Orthopedics |  |  | November 27, 2006 | 2006 | Plus Orthopedics USA | The VKS "Tibial Trial" was found to have an incorrect fixation pin hole diameter. | Process control | The initial depth of the withdrawal action taken November 2006 was to the distributor (agent) level. On 11/27/2006 Plus Orthopedics USA contacted individual agent sites showing inventory of the affected article numbers on hand and requested that those items be returned to Plus USA. Agents were notified by telephone to check their inventory and return any affected units. A telephone script was used. |  | 19 | 19 | Nationwide to : AZ, CA, CO, FL, ID, NE, NY, SC, TX, UT, VA, and WA | Terminated | December 13, 2007 | 381 | 19 | Undetermined
Z-0357-2008 | 45445 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Stryker inTouch Zoom Critical Care bed, Model 2140, Stryker Medical, Portage, MI 49002 |  |  | November 02, 2007 | 2007 | Stryker Medical Div. of Stryker Corporation | Multiple defects: 1) beds exhibit reverse motions when the Fowler, gatch or foot section is activated 2) a battery charging defect which may result in a depletion of battery power 3) the touch screen on the footboard freezing, resulting in controls to powered, pressure-relief mattresses being difficult to access 4) beds equipped with zoom drive systems may have an overdrive problem which could re | Process control | Consignees were notified via an Urgent Medical Device Correction letter dated 11/2/07, and instructed to remove the beds from service if they display any unintended motions and informed that one of the firm's field service representatives will visit the facility to make corrections to all beds under recall. |  | 329 | 329 units total for both models. | Worldwide-USA and Canada | Terminated | October 27, 2010 | 1090 | 329 | Undetermined
Z-0359-2008 | 45465 | No | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | ACIST Multi-Use Syringe Kit, Model A2000; SKU # : 014612; labeled as STERILE; Manufactured: Acist Medical Systems, Inc., 7450 Flying Cloud Drive, Suite 150, Eden Prairie, MN 55344, USA. Catalogue No: 800061-003. (The multiple use A20000 syringe kit is comprised of a terminally sterilized syringe barrel and a contrast spike assembly) |  |  | October 03, 2007 | 2007 | Acist Medical Systems | Syringe Fracture: the syringe may be fractured, causing contrast to be sprayed from syringe fracture. | Process design | The Urgent Medical Device Recall letter was sent by certified mail to all U.S. base customers on 10/3,4/07. Customers were advised to stop distribution of the product and immediately cease use of the product. Recalled product is to be returned using the FedEx Airbill that was enclosed with the recall letter. Customers were to complete the enclosed "Device Recall Response" form even if they did not have any impacted kits in inventory. |  | 177050 | 177,050 Kits | Nationwide Distribution - including states of AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, and WV. | Terminated | October 06, 2008 | 369 | 177050 | Undetermined
Z-0360-2008 | 45475 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Plus Orthopedics Gliding Nail 17mm proximal / 11mm distal, 125 degrees, 180mm; Model Number: 0132073, Plus Orthopedics USA, San Diego, CA 92121 |  |  | June 30, 2003 | 2003 | Plus Orthopedics USA | The device was produced without the required distal locking hole | Reprocessing Controls | The initial depth of the withdrawal action taken in June, 2003 ,was to the distributor (agent) level. Plus USA contacted individual agent sites showing inventory of the affected article numbers on hand and requested that those items be returned to Plus USA. |  | 3 | 3 | Nationwide: AZ and ID | Terminated | December 13, 2007 | 1627 | 3 | Undetermined
Z-0361-2008 | 45480 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | VKS PE Tibia Insert Ultra-Congruent; Left, Size 2, 12mm; Model Number: 0214612; Plus Orthopedics, San Diego, CA 92121 |  |  | November 13, 2006 | 2006 | Plus Orthopedics USA | Labeling Error: The VKS PE Ultra-congruent insert package, may actually contain a Standard PE insert. | Employee error | Consignees were notified by telephone on 11/13/2006 to check their inventory and return any affected units. On 11/28/06 Physician Notifications and Patient Assessment forms were sent to the two physicians who had implanted a total of 4 devices. |  | 10 | 10 | Nationwide: USA including states of ID, NY, VA &amp; WI | Terminated | December 13, 2007 | 395 | 10 | Undetermined
Z-0328-2008 | 38471 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific ACCUSTICK II Introducer System with .038 inch Stainless Steel J Tip and .018 inch Nitinol Guidewire, sterile, REF/Catalog No. 20-710; UPN/Material Number: M001207100, Boston Scientific Corporation, Natick, MA 01760 |  |  | July 19, 2007 | 2007 | Boston Scientific Corp | Marker Detachment: Out of specifications for pull strength, thus the marker may detach from the introducer during use. | Process control | Consignees were notified via an Urgent Medical Device Recall Letter sent federal express on 7/19/07, which requested them to check their stocks, remove product from inventory, and return the product. |  | 249 | 249 (199 U.S.A. plus 50 international) | Worldwide: USA and The Netherlands. | Terminated | February 17, 2008 | 213 | 249 | Undetermined
Z-0122-2008 | 38310 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific Target, Transend EX Soft Tip Guide Wire, .014in /205cm, UPN: M001468060, Catalog Number: 46-806, |  |  | June 14, 2007 | 2007 | Boston Scientific Target | Mislabeling: Product is labeled with the wrong expiration date: 2009-12. The correct expiration date is 2008-12. | Process control | Boston Scientific issued a letter to all consignees by mail on June 14, 2007, titled Urgent Medical Device Recall or Correction - Immediate Action Required. The letter describes that the expiration date is incorrect. Discontinue use of the product, segregate it, complete the Reply Verification Tracking Form and fax it to Boston Scientific. Call 888-272-1001 to obtain a Returned Goods Authorization (RGA) number. |  | 40 | 40 | Nationwide. | Terminated | December 21, 2007 | 190 | 40 | Undetermined
Z-0288-2008 | 45333 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Abbott Clinical Chemistry Aeroset/Architect (c Systems)Bilirubin Calibrator, List Number (LN) 1E66-03 &amp; List Number (LN) 1E66-04, Abbott Laboratories, Abbott Park, IL 60064 |  |  | August 10, 2007 | 2007 | Abbott Laboratories Inc. | Incorrect Calibrator values: changes in a raw material used in the Bilirubin Calibrator value assignment process contributed to variability in the preparation of the standard. | Material/Component Contamination | An August 10, 2007, Product Correction letter was provided to all ARCHITECT cSystem and AEROSET System customers. The letter advises that changes in a raw material used in the Bilirubin Calibrator value assignment process contributed to variability in the preparation of the NIST standard 916a. Revised Bilirubin values for the calibrator lot numbers in this recall are attached to the recall notice. Six necessary actions were listed in the recall notice as: 1) Identify the lot numbers of Clinical Chemistry Bilirubin Calibrator(s) currently in use. 2) Discard the value sheet and replace with the appropriate values provided with the recall letter. 3) Enter the revised Bilirubin values. 4) Calibrate the assays. 5) Evaluate your quality control. 6) Evaluate your reference range and any shifts in patient results. |  | 7033 | 7,033 units total worldwide | Worldwide: USA, Canada, Mexico, Guatemala, Puerto Rico, Germany, Hong Kong, Singapore, Australia, New Zealand, Jamaica, Venezuela, Columbia, Ecuador, Chile, Uruguay, Argentina, Brazil, Thailand, Trinidad &amp; Tobago, Bahamas, Panama, Cayman Islands, and El Salvador | Terminated | April 19, 2012 | 1714 | 7033 | Undetermined
Z-0289-2008 | 44752 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | A1CNOW+ with sampler 10-Test Kits (Product No: 3022), Metrika, Inc., Sunnyvale, CA 94085. (This test provides quantitative measurement of the percent of glycated hemoglobin (%A1C) levels in capillary (fingerstick) or venous whole blood samples. The test is used by professional healthcare providers to monitor long-term glycemic control in people with diabetes. The product is sold as a kit with a disposable monitor and reagents for ten A1C tests and is suitable for use at the point-of-care. |  |  | August 31, 2007 | 2007 | Metrika Inc | Incorrect Test Results: Due to manufacturing errors regarding temperature compensation factors, the device will give erroneously high results if the test is performed below 70 degrees Fahrenheit. | Process control | A September 14, 2007, Urgent Device Recall letter was issued to all Distribution Center Managers to return recalled product for replacement. Customer notification was made via a Dear Healthcare Professional letter dated Sept. 14, 2007, also titled Urgent Medical Device Recall. |  | 4286 | 4,286 | Nationwide including states of AK, AL, CA, CO, CT, DC, FL, GA, HI, IL, IN, KS, KY, MA, MD, MI, MN, MS, NC, NJ, NV, NY, OH, OK, PA, SC, TN, TX, UT, VA, WA, WI and Puerto Rico. | Terminated | March 25, 2008 | 207 | 4286 | Undetermined
Z-0321-2008 | 38466 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System, 1 Roller Pump, 4 inch diameter; Catalog No. 801040, Terumo Cardiovascular Systems, Ann Arbor, MI 48103 |  |  | September 10, 2007 | 2007 | Terumo Cardiovascular Systems Corp | Failure to operate; The pump may fail to operate, and be difficult to hand crank, due to an internal electrical short condition. | Process control | Consignees were notified via an Urgent Medical Device Correction letter dated 9/10/07, which provided them guidance in the event the problem occurs and advised them they would be contacted to schedule replacement of the circuit boards. |  | 16 | 16 | Worldwide Distribution-USA including states of Alabama, Arizona, Florida, Illinois, Indiana, Mississippi, Nebraska, North Carolina, Pennsylvania, and Virginia, and countries of Belgium, Hong Kong and United Arab Emirates. | Terminated | February 23, 2008 | 166 | 184 | Undetermined
Z-0323-2008 | 38467 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1 Integrated Centrifugal System Control Unit; Catalog Number 801046, Terumo Cardiovascular Systems, Ann Arbor, MI 48103 |  |  | September 10, 2007 | 2007 | Terumo Cardiovascular Systems Corp | Failure to operate: The unit display may go blank and reset when the start button is pressed, due to an internal electrical short, resulting in a centrifugal pump that fails to start as expected. | Component change control | Consignees were notified via an Urgent Medical Device Correction letter dated 9/10/07, given instructions to follow if the problem were to manifest itself, and informed the firm would contact them to schedule an upgrade. |  | 120 | 120 | Nationwide. | Terminated | July 07, 2009 | 666 | 120 | Undetermined
Z-0266-2008 | 45515 | No | HTD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | PillingÂż Center-Action Forceps with Insulated Shafts- Serrated Grasping Jaws, Catalog Number: 505015, Non Sterile, Rx Only, (ALLIG GRASPING FCPS 30 CM INSUL), Manufactured for : Teleflex Medical, Research Triangle Park, NC 27709 USA |  |  | October 31, 2007 | 2007 | Telefelx Medical | The specified insulation is not present on the shaft of the devices, leaving a risk of arcing and/or damage to surrounding tissue. | Process control | Consignees were notified by an Urgent Medical Device Recall letter on/about 11/1/2007. They were instructed to immediately cease use and distribution of the affected product and to return them to Teleflex Medical, Durham, NC. In addition, each were instructed to forward the letter if they had further distribution into institutional settings or to homecare patients and retrieve relevant product form those locations. An Acknowledgement &amp; Stock Status Form was enclosed to be completed and faxed back to Teleflex Medical. |  | 73 | 73 | Worldwide: USA including states of AL, AR, AZ, CA, FL, GA, HI, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NH, NJ, NM, NY, NV, OH, OR, PA, SC, TN, TX, UT, VA, WA, WI, and WV, and countries of Canada, France New Zealand and Singapore. | Terminated | August 31, 2009 | 670 | 235 | Undetermined
Z-1265-2007 | 44735 | No | LZK | Chin prosthesis. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Spectrum Designs Medical Anatomical Chin Implant; Catalog Number 250-134S; Lot # 31140; Serial Numbers BAHO45 through BAH144 |  |  | July 24, 2007 | 2007 | Spectrum Designs Inc | Mispackaging/Mislabeing: a portion of the lot may be labeled as containing nasal implants | Other | The firm initiated the recall on 07/24/2007 by contacting the customer (distributor) who filed the original complaint of mislabeling by email. The firm followed-up by sending recall letters (dated July 24, 2007) via FedEx or certified mail with signature confirmation to the one other direct account who received the recalled product on 08/03/2007. Customers were instructed to check their inventory for mislabeled products and call Spectrum at (800)239-6399 to report the number of this product being returned. They were given an address for product return. The firm indicated they will examine returned products, and scrap after examination. |  | 100 | 100 | Nationwide: California and China | Terminated | July 29, 2011 | 1466 | 230 | Undetermined
Z-0036-2008 | 44827 | No | KLE | Resin tooth bonding agent. | Dental | Dental | 510(k) | Heraeus Kulzer GLUMA Comfort Bond + Desensitizer light curing one component adhesive, Assortment Value Pack 3 x 4 ml, dental bonding agent. Distributed by Heraeus Kulzer, Inc., Armonk, NY; Reorder No. 66018222. |  |  | July 09, 2007 | 2007 | Heraeus Kulzer Inc., Dental Products Division | The material may solidify in the bottle, and thus be rendered unusable. | Process design | Consignees were notified via phone or undated letter sent via fax beginning on or about 7/9/07 and instructed to contact the firm for return authorization/instructions. |  | 736 | 736 | Nationwide and Canada. | Terminated | March 12, 2008 | 247 | 3944 | Undetermined
Z-1088-2007 | 44936 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Innova 2100-IQ Digital Fluoroscopic Imaging System, Versatile cardiac and vascular imaging system. GE Healthcare Americas P.O.Box 414, Milwaukee, Wisconsin 53201, U.S.A. |  |  | June 06, 2006 | 2006 | GE Healthcare | The systems do not contain a certification label on the C2 Cabinet (Beam-limiting control). | Other | Starting on June 6, 2007, a GE service engineer will go to each affected site to affix certification labels on the affected devices. |  | 161 | 161 | World wide Distribution includes: AL, AR, AZ, CA,, CO, CT FL, GA, HI, ID, IL, IN, KS, MA, MD, MN, MO, NC, NE, NJ, NY, OH, OR, PA, SC, TN, TX, WA, WI,and WV, and countries of ; Algeria, Australia, China, Denmark, France, Germany, Hungary, India, Indonesia, Italy, Japan, Korea, Kuwait, Latvia, New Zealand, Norway, Poland, Saudi Arabia, Singapore, Switzerland, Syrian Arab Republic, Turkey, and United Kingdom. | Terminated | December 08, 2015 | 3472 | 161 | Undetermined
Z-0676-2007 | 37498 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | OEC 9900 Elite image-intensified fluoroscopic x-ray system, GE Healthcare, Salt Lake City, UT. |  |  | March 01, 2007 | 2007 | GE OEC Medical Systems, Inc | During routine service, primary and/or secondary collimators may not have been properly reinstalled. | Other | Customers were notified by letter on 03/01/2007. They were told to continue using the machines and that an inspection would be scheduled immediately to verify that collimators had been correctly installed. |  | 19 | 19 units | Nationwide and Germany. | Terminated | May 01, 2012 | 1888 | 1975 | Undetermined
Z-1234-2007 | 44721 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | The product is an isolation transformer, an accessory to Computed Tomography Scanner. It consists of electronic circuitry enclosed in a metal case. According to Philips Medical Systems, the metal case is manufactured from self-extinguishing materials. The unit is labeled in part, "Power Conditioner *** Part Number 4450028-XX *** Frequency: *** S/N: *** Rev: *** MFG date: *** INPUT RATINGS *** Model: MCT 100/480 ISOTRAN PLUS *** OUTPUT RATINGS *** MEDICAL POWER SUPPLY UL *** TEAL Electronics Corporation, San Diego, California ***". |  |  | April 27, 2007 | 2007 | Philips Medical Systems (Cleveland) Inc | The TEAL Power Conditioning units may fail and result in the unit emitting fire and/or smoke. Note: the TEAL Power Conditioning Unit has the potential to fail when the scanner that is attached to the TEAL is powered and in use as well as when not in use. | Component design/selection | On 9/04/2007, Philips Medical Systems (PMS) sent a letter via courier to all affected customers. The letter notified customers about the failures of the TEAL power conditioner and the possible emisson of fire and/or smoke from the units. In addition, the letter notes that PMS implemented a Field Change Order in May, 2007 that directs onsite repair of all affected units by PMS service technicians. |  | 491 | 491 | Worldwide, including, USA, Canada, China, Italy, Russia, Brazil, Belgium, Germany, Austria, Australia, Netherlands, France, Korea, Denmark, Israel, England, Argentina, Mexico, Spain, Yemen, Greece, Egypt, Japan, UAE, &amp; Chile. | Terminated | March 02, 2012 | 1771 | 491 | Undetermined
Z-1203-2007 | 38428 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | Kion Anesthesia Workstation |  |  | July 10, 2007 | 2007 | MAQUET Inc. | Overcharging of batteries and the release of acid fumes | Other | Safety Alert letters were sent out by FedEx tracking.A Service Express letter will be sent out emphasizing that only qualified batteries must be used in KION equipment and will list all qualified battery types.A Field Action - update of existing operating manuals will include warning on battery charging on KION. This action is currently scheduled to start 9/1/07 and be completed in US by 10/1/07. After FDA consultation the recalling firm issued a 9/5/2007 URGENT DEVICE CORRECTION Letter stating that KION will visit consignees and replace batteries older than three years, verify that the batteries are ones of the approved type, and verify battery condition. |  | 9 | 9 units in US | Nationwide: 2 units at the National Institute of Health in Bethesda, MD and 7 units at Oconee Regional Medical Center, Milledgeville, GA | Terminated | December 07, 2007 | 150 | 9 | Undetermined
Z-1182-2007 | 38400 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Intellivue Central Station Software for the: Intellivue Information Center; Philips Information Center, Central Station |  |  | July 16, 2007 | 2007 | Philips Medical Systems | Incorrect Patient Assignment: the information center will in most cases be associated with the "incorrect" patient. This occurs on both standalone and networked systems running fixed mode monitoring. (This does not occur on systems that run flex mode monitoring) | Other | Philips Medical Systems issued a letter notification dated July 2007 " Urgent Device Correction". Customers with networked information centers will be corrected through a database server. Stand-Alone units will have a software correction. Software version J.00.25 will correct the problem. Users are advised not to use th 12 Lead Analysis and Export Option in fixed or flex mode monitoring unitl the software correction is made. |  | 76 | 76 units | Worldwide Distribution: USA including states of NJ, NY, OK, MI, and UT and Foreign: Australia, Finland, France, Netherlands, and the UK | Open, Classified |  |  | 76 | Undetermined
Z-1168-2007 | 38273 | No | FAD | Ureteral stent. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Cook Urological Black Silicone Filiform Double Pigtail Ureteral Stent Set, 7.0Fr/24cm, sterile, Order Number (REF) 133724. |  |  | June 29, 2007 | 2007 | Vance Products Incorporated | The radiographic band may detach from the stent and remain in the patient after removal of the stent. | Other | Consignees were notified of the recall via letter dated 6//29/07, which instructed them to quarantine stock on hand and to fax number on hand back to Cook for return authorization. |  | 75 | 75 | Worldwide, including USA, Canada, and Japan. | Terminated | June 11, 2008 | 348 | 75 | Undetermined
Z-1207-2007 | 38308 | No | GZJ | Transcutaneous electrical nerve stimulator for pain relief. | Neurology | Neurology | 510(k) | Avazzia Tennant BiomodulatorÂż; Model: BEST AV-1; P/N #TB04303-103; manufactured by Avazzia, Inc., Dallas, TX 75240. |  |  | June 29, 2007 | 2007 | Avazzia, Inc | Plastic Cases may crack due to manufacturing by the supplier not meeting specifications | Other | Firm notified consignees by letter on June 29, 2007 requesting that all affected product be returned to the firm for repair or replacement. The firm further asked for distributors who had shipped affected product to notify receiving patients so they could arrange for repair or replacement. |  | 134 | 134 units. | Nationwide including states of Texas, Pennsylvania, Louisiana, Wyoming, and California | Terminated | April 04, 2008 | 280 | 556 | Undetermined
Z-1215-2007 | 38306 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Smiths CADD Medication Cassette Reservoir, With clamp and female Luer, Nonvented stopper included, 100 ml PC |  |  | June 18, 2007 | 2007 | Smiths Medical MD, Inc. | Smith Medical has become aware that an increase in pH may occur with some medications when instilled in the Affected Cassettes. Also, a small number of catheter occlusions have been reported from one customer when these cassettes have been used for delivering Flolan. | Other | All customers who have received the CADD Medication Cassette reservoirs have been sent the Product Recall Notification Letter on June 18, 2007 making them aware of the potential issue. Customers will be required to return the affected devices to Smiths Medical. All returned recall affected product will be quarantined and then destroyed. The consignees are to fill and fax "Attachment 1" form in regard to affected product inventory to Smiths Medical MD. |  | 113143 | 113,143 | Worldwide, including USA, Canada, Mexico, Sweden, Germany, Japan, Australia, Taiwan, Poland, Jordan, Czech Republic, Denmark, Paraguay, Hong Kong, Singapore, Argentina, Italy, Venezuela, Greece, Norway, Pakistan, Puerto Rico, Thailand, United Arab Emirates, Netherlands Antilles, Malaysia, Panama, Hungary, Saudi Arabia, Portugal, South Africa, Denmark, the Netherlands, France, United Kingdom, Great Britain, Spain, Finland, Guam, Virgin Islands, China, Israel,Austria, and Switzerland. | Terminated | February 10, 2011 | 1333 | 153798 | Undetermined
Z-1211-2007 | 39147 | No | MAF |  |  | Cardiovascular | PMA | Multi-Link Vision RX Coronary Stent System, 3.0 x 15mm, Part number 1007848-15, lot number 7052951 |  |  | July 16, 2007 | 2007 | Abbott Vascular-Cardiac Therapies dba Guidant Corporation | The affected single lot did not meet an internal specification. One unit failed the test criteria for stent dislodgment testing. | Other | The firm initiated the recall on 07/16/2007 by sending recall letters with attached Customer Reply form via FedEx to each direct account that received the recalled product. Customers were informed of the issue and given instructions to work with their local account representative to review their inventory, remove the recalled product, complete the Customer Reply form and return the recalled product to Abbott Vascular. |  | 43 | 43 units | Nationwide | Terminated | June 14, 2011 | 1429 | 43 | Undetermined
Z-1206-2007 | 38898 | No | LZK | Chin prosthesis. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Spectrum Designs Medical Anatomical I (Silhouette) Malar Facial Implant, Catalog Number 120-224S. |  |  | June 15, 2007 | 2007 | Spectrum Designs Inc | The product is being recalled because it was discovered the at least one unit in the lot contained a mismatched pair (two rights and no left). Since the lot was produced with equal numbers of pairs, it is assumed that one other unit was packaged with two lefts and no right. This was discovered through a physician complaint. | Other | The firm initiated the recall on 06/15/2007 by contacting their customers by telephone. The firm followed-up by sending recall letters via FedEx or certified mail with signature confirmation to each direct account who received the recalled product on 06/18/2007. |  | 15 | 15 pairs (left/right) | Nationwide. | Terminated | March 07, 2012 | 1727 | 15 | Undetermined
Z-1205-2007 | 39374 | No | KDD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Assorted non- sterile Tracecarts (with sterile component(s)); |  |  | July 03, 2007 | 2007 | DeRoyal Lafollette | Packaging failure. Lack of sterility assurance | Other | Written notification of the recall was sent (via UPS Ground) beginning 7/03/2007 to consignees to advise them of problem and the recall. The letter stated that the recall was to the User level (which in this case, are the hospitals) and requested return of the product. |  | 112 | 112 cases | Nationwide | Terminated | June 12, 2008 | 345 | 2135 | Undetermined
Z-1200-2007 | 38322 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Abbott CELL-DYN 3700 Hematology Analyzer |  |  | May 11, 2007 | 2007 | Abbott Laboratories | The product failed electrostatic discharge testing, which could result in product failure. | Other | Mandatory Technical Service Bulletins (TSB) are being issued to Field Service to upgrade all instruments in the field. All customers will be contacted by Field Service to schedule this mandatory upgrade. The upgrades are expected to take 12 months. |  | 3429 | 3,429 units, all models | Worldwide, including USA, Puerto Rico, Canada, Mexico, Argentina, Australia, Bahamas, Brazil, Chile, China, Colombia, Costa Rica, Dominican Republic, El Salvador, Ecuador, Germany, Guatemala, Hong Kong, Jamaica, Japan, New Zealand, Panama, Paraguay, Peru, Singapore, South Korea, Taiwan, Thailand, Uruguay, and Venezuela. | Terminated | June 19, 2008 | 405 | 3429 | Undetermined
Z-1218-2007 | 38512 | No | IZX | Diagnostic x-ray beam-limiting device. | Radiology | Radiology | 510(K) Exempt | PCS125 Swivel Collimator, Model F100; Manual Radiographic Collimator; Summit Industries, Inc., Chicago, Illinois |  |  | August 10, 2007 | 2007 | Summit Industries Inc. | The F100 Swivel Collimator may come loose from the tubestand and fall onto the x-ray table. | Other | The affected dealers were notified of the field correction recall by letter dated 8/3/07, informing them of the problem with the stripped screws and requesting that they notify their affected end user locations. The end user notification letters inform them of the problem and state that the dealer will be contacting them soon to arrange a time to visit the end user location and conduct the rework on the collimator, which should take less than 30 minutes. In the mean time the end user was instructed to examine their swivel collimator for looseness. If it is found to be loose, the end user was instructed to take the unit out of service and call the dealer. |  | 3034 | 3034 for all units | Worldwide, including USA, Australia, Cambodia, Canada, Chile, Colombia, Dominican Republic, Ecuador, France, Germany, Haiti, Mexico, Nicaragua, Singapore, and South Africa. | Terminated | March 18, 2010 | 951 | 3034 | Undetermined
Z-1201-2007 | 38350 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ABL8O FLEX Software Version 1.10 included in the ABL8O Flex Analyzer, Model Number 914863 |  |  | June 04, 2007 | 2007 | Sendx Medical Inc | The ABL8O FLEX software version 1.10 contains errors in the derived parameter calculation for cBase(B,ox) and cBase(Ecf,ox). These incorrect calculations result in a mild to moderate lowering of the base excess value, the severity depending on the concentration of hemoglobin and the oxygen saturation in the sample. Elevated hemoglobin concentrations have the greatest impact on this calculation err | Other | A total of eight (8) distributors and the Radiometer main office in Denmark were sent the recall package (Field Action Notice) by email on 6-6-07. The distributors were provided background information and specific instruction on the steps to perform to correct the problem. The Recall package included a Management Cover for Field Action Notes and the Field Action Notice.Radiometer has released a subsequent software upgrade (v.1.11) to correct the two problems noted in the Recall Notice. |  | 165 | 165 software upgrade CDs | In US to Cleveland, OH and Worldwide to Belgium, Denmark, France, Germany, The Netherlands, Sweden &amp; United Kingdom | Terminated | February 29, 2012 | 1731 | 165 | Undetermined
Z-1202-2007 | 37904 | No | KXT | Digoxin test system. | Toxicology | Toxicology | 510(k) | Abbott AxSYM System Digoxin III Reagent Pack |  |  | March 30, 2007 | 2007 | Abbott Diagnostic International, Ltd. | Customers are receiving error codes (1062, 1063, 1113, &amp; 1118) when running patient samples on the AxSYM Digoxin III assay. | Other | On Abbott letterhead dated March 30, 2007 - The instructions are to provide a copy of the accompanying Product Correction Letter to the laboratory manager/supervisor responsible for AxSYM Digoxin II or AxSYM Digoxin III testing. The laboratory manager/supervisor should complete the information on the letter acknowledging receipt of the product Correction Letter and Fax the Customer Reply Form, prior April 12, 2007 to Fax number 1-800-777-0051. |  | 2579 | 2,579 kits | Worldwide, including USA, Puerto Rico, Canada, Uruguay, Germany, Japan, Hong Kong, South Korea, Australia, and New Zealand. | Terminated | June 10, 2011 | 1533 | 2579 | Undetermined
Z-1197-2007 | 38382 | No | HMX | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | Atkinson Retrobulbar Anesthesia Needles, 25 gauge 1 1/2". Catalog Number 8065420920. The product is sold in boxes that contain 10 needles each. |  |  | July 11, 2007 | 2007 | Alcon Research, Ltd | Mislabeled: boxes actually contained Tri-Fact needles | Process control | The recalling firm telephoned the hospital on 7/11/07 to inform them of the problem and the need to return the product. As a follow up to the telephone call, the recalling firm issued a recall letter via FedEx on 7/11/07. The firm's representative is making arrangements to pick up the recalled product. |  | 2 | 2 boxes | CA. | Terminated | August 28, 2007 | 48 | 2 | Undetermined
Z-1196-2007 | 38172 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ABL80 FLEX Analyzer, Model #393-839 |  |  | March 27, 2007 | 2007 | Sendx Medical Inc | There have been reported experiences where waste fluids have sprayed out from the waste drain at the front of the ABL80 analyzer when excessive pressure builds within the waste line. | Other | The recall was initiated 3/27/2007. All distributors were notified via the Radiometer FAN (field action notice) system which provides a communication mechanism for these types of actions. The notice advises that there have been a few reported experiences where waste fluids have sprayed out from the waste drain at the front of the ABL80 analyzer which occurs when excessive pressure builds within the waste line. To prevent any spraying of fluids from the waste drain a check valve has been added to the waste tubing, between the waste drain and the waste pump which allows fluids to flowin one direction thus making it impossible to spray fluids out the waste drain. The letter further advises that all analyzers shipped from SenDx on or after March 16, 2007 will include this check valve. All analyzers already in the field must be updated to include this check valve. Distributors are to distribute a copy of the Customer Advisory Letter associated with this case to all their customers and update all ABL80 analyzers to include the waste drain check valve as per the instructions that accompany the check valve. The Advisory Letter informs the customer of the problem and the new waste drain check valve and provides precautionary instructions such as appropriate hand, eye, face and clothing shields. In addition, it advises that the analyzer should never be used if there is any indication of a waste drain occlusion or blockage. The most typical indication of an occlusion or blockage is fluid overflowing out of the waste drain. |  | 334 | 334 units worldwide | Class II Recall - Worldwide Distribution --- USA including state of Ohio, and countries of Austria, Belgium, Denmark, Finland, France, Germany, Italy, Japan, The Netherlands, Norway, Poland, Spain, Sweden, Switzerland, United Kingdom, Australia, New Zealand, Canada, and Denmark. | Terminated | October 13, 2011 | 1661 | 334 | Undetermined
Z-1190-2007 | 38324 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | Volcano s5 Imaging system, Ultrasound pulsed echo imaging system. Product consists of the CPU, control console, monitor, printer, and patient interface module (PIM) |  |  | April 02, 2007 | 2007 | Volcano Corporation | During emissions testing, it was discovered that the product exceeds emissions limits for electromagnetic compatibility. | Other | The firm notified its customers by letter beginning April 2, 2007, and started replacing the monitor with appropriate replacement parts. |  | 351 | 351 units worldwide | Product was distributed to health care facilities nationwide and to distributors in Europe (via subsidiary), Japan (via subsidiary), Latin America and Asia (third party distributors). | Terminated | June 03, 2008 | 428 | 351 | Undetermined
Z-1194-2007 | 38462 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer Anatomical Shoulder Prosthesis, Glenoid, Cemented, sterile; REF 01.04214.370. |  |  | June 25, 2007 | 2007 | Zimmer Inc. | The package actually contains a knee patellar implant. | Other | The firm notifed the distributor via letter dated 6/25/07, and sent by email, to retrieve the product from their customers and to return the product. |  | 10 | 10 | Maryland. | Terminated | December 05, 2007 | 163 | 10 | Undetermined
Z-1193-2007 | 38031 | No | FLE | Steam sterilizer. | General Hospital | General Hospital | 510(k) | Amsco Century Medium Steam Sterilizer (26" x 37.5") (660 x 950mm). |  |  | March 22, 2007 | 2007 | Steris Corp | The cables inside the pump conduit box may be exposed (not covered with wire nut and/or electrical tape) that could cause a potential electrical shock hazard. | Employee error | By telephone 3/22/2007, customers were instructed to cease using the product and a service technician would be dispatched immediately to repair the unit. |  | 78 | 78 | Worldwide, including USA, Denmark, Belgium, Greece, India, and Russia.There are no Canadian accounts. | Terminated | September 28, 2010 | 1286 | 78 | Undetermined
Z-1189-2007 | 38433 | No | MBD |  |  | Orthopedic | PMA | BacT/ALERTÂż FN Culture Bottles, Product Number 259793 |  |  | February 16, 2007 | 2007 | bioMerieux, Inc. | Microbial Contamination: Bacillus sp. contamination was detected in inoculated bottles. | Other | Consignees were notified by Urgent Product Recall letter on 02/19/2007. They were instructed to destroy any unused inventory of this product lot that they have on hand. |  | 2403 | 2,403 cases (100 bottles per case) | Worldwide; USA, Austria, Bulgaria, Caledonia, China, Croatia, Curacao, Finland, France, Germany, Greece, Guadeloupe, Italy, Latvia, Mexico, Netherlands, Poland, Portugal, Spain, Saudi Arabia, Sultanate of Oman, Switzerland, Sweden, and United Kingdom | Terminated | December 08, 2009 | 1026 | 2403 | Undetermined
Z-1174-2007 | 38384 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | EasyLink Informatics System, software version 3.0, Part Number 1000034941, DADE BEHRING INC., Newark, DE. |  |  | July 03, 2007 | 2007 | Dade Behring, Inc. | Sample ID and Patient results could be associated with the wrong patient | Other | The recalling firm issued a Field Corrective Action letter dated 7/3/07 to inform their customers of the problem and that a software patch would be sent for them to download. The recalling firm plans to telephone the customers after the patch has been released to verify that it has been installed. |  | 72 | 72 units | Worldwide: USA including states of AZ, CA, FL, IL, MA. MI, MO, NC, OH, PA, TX, and WV. The product was also shipped to other Dade facilities in MO, DE, IL, CT, and Belgium. | Terminated | November 07, 2007 | 127 | 72 | Undetermined
Z-1178-2007 | 38251 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | PhaSeal, Protector 14, Item No. P14, US Order No. 21320 |  |  | June 26, 2007 | 2007 | Carmel Pharma, Inc. | Risk of Leakage: When the PhaSeal Protector and Injector are used with drug vials with thick rubber stoppers which have a concave upper surface and a flat bottom surface there is a potential for leakage. These rubber stoppers are made of a harder rubber formulation and may be more difficult to penetrate. | Other | On or about 6/15/2007, the firm sent advisory notice letters to the end users and the instructions for use have been updated by adding a precaution statement regarding the use of thick rubber stoppers. The firm is also updating the customer training program to include how to handle drug vials with thick concave rubber stoppers. |  | 6100000 | 6.1 million for Z-1178-1181-2007 recalls | Product was distributed throughout the US and Canada | Terminated | March 05, 2013 | 2079 | 6100000 | Undetermined
Z-1187-2007 | 38167 | No | LGI | Neuroleptic drugs radioreceptor assay test system. | Toxicology | Toxicology | 510(k) | XSYSTEMS Quinidine Calibrators consists of six bottles of Cal A-F (2.5 mL each).-Product List Number-9506-01 |  |  | May 31, 2007 | 2007 | Abbott Diagnostic International, Ltd. | Controls Out of Range/Specification. | Other | May 31, 2007, Product Correction Letter to Customers were instructed to: "Revise the expiration date for all lots identified in the customer letter. "Discontinue use of the expired lots, and discard the material according to their laboratory procedures. "An additional option is to use the Abbott Immunoassay Multiconstituent Control (List Number 06E20), manufactured by BIO-RAD. |  | 250 | 250 | Worldwide: USA including states of AR, AZ, CA, CT, FL, GA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, NC, NJ, NY, NV, OH, OK, OR, SC, TN, TX, VA, WA, WI, WV, Hawaii and Puerto Rico. International: CANADA, GERMANY, JAPAN, SINGAPORE, AUSTRALIA, and BARBADOS | Terminated | June 10, 2011 | 1471 | 729 | Undetermined
Z-1184-2007 | 38409 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Arizant Healthcare Ranger Standard Flow Blood/Fluid Warming Disposable Sets with injection port and extension set, Part Number 24250. For use with a Ranger warming unit. |  |  | June 26, 2007 | 2007 | Arizant Healthcare, Inc | Incomplete seal on the product pouch; with the seal incomplete, the sterility of the product inside the pouch cannot be guaranteed. | Other | All consignees will be contacted by Urgent Medical Device Recall letter on 6/26/07. |  | 12000 | 12,000 | Worldwide Distribution --- USA including states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL., IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY, and countries of United Kingdom, France, Switzerland, Germany, India, Canada, Australia, Netherlands, Italy, Sweeden, South Korea, and Greece. | Terminated | March 02, 2008 | 250 | 97710 | Undetermined
Z-1173-2007 | 38359 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ABL77 Cal Pack, Model Number 944-069, Lot Number 17323, SenDx Medical Inc., A Radiometer Company |  |  | June 19, 2007 | 2007 | Sendx Medical Inc | Error in the Ca12 calibration value for pCO2 recorded in the cal pack barcode for lot 17323. This error will result in falsely elevated pCO2 results at the high end of the measuring range. | Other | A total of four (4) distributors were sent the recall package (Field Action Notice) by email on 6-29-07. The distributors were provided background information and specific instruction that all Cal packs from lot 17323 should be removed from customer sites and destroyed. The Recall package included a Management Cover for Field Action Notes, a Customer Letter and the Field Action Notice #915-261.All communications regarding field actions associated with this recall were distributed using the Radiometer Field Action Note (FAN) system. |  | 69 | 69 | Worldwide Distribution --- USA including state of OH and countries of China, Denmark Latvia, Mexico &amp; Spain. | Terminated | November 07, 2011 | 1602 | 69 | Undetermined
Z-1171-2007 | 38042 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Cell-Dyn 1700CS (Closed Sampler) Hematology Analyzer, List Numbers 03H57-01 and 03H57-03, software controlled medical device for diagnostic use and CELL DYN Control Assay Disk, LIst Numbers 01H91-01 and 01H92-01 when used with above CELL DYN 1700, Abbott Laboratories Diagnostics Division |  |  | May 16, 2007 | 2007 | Abbott Laboratories | Abnormal results may not be "flagged" appropriately because of a software anomaly. (Flags for the Hemoglobin (HGB), MCH and MCHC parameters when the upper assay paramter limit has been exceeded) | Other | The firm initially sent reports (Product Information Letters) of this issue to the field on May 7, 2007. The firm sent a notice with customer notification response form on May 16, 2007 (Product Correction Letter). In addition, product information letters are being packed with the instruments accessory kit and quality control products used with the Cell-Dyn 1700 system. |  | 4197 | 4,197 units total (both models) | WORLDWIDE: Product was distributed nationwide to all 50 states, as well as worldwide to Anguilla, Argentina, Australia, Bermuda, Bolivia, Brazil, Canada, Cayman Islands, Chile, Columbia, Curacao, Dominican Republic, Ecuador, El Salvador, Germany, Guatemala, Honduras, Hong Kong, Jamaica, Japan, Malaysia, Mexico, Panama, Paraguay, Peru, Puerto Rico, Singapore, South Korea, Taiwan, Thailand, Uruguay, and Venezuela. | Terminated | January 11, 2011 | 1336 | 4197 | Undetermined
Z-1175-2007 | 38396 | No | MID | Multiple autoantibodies immunological test system. | Immunology | Immunology | 510(k) | APhLÂż HRP IgG &amp; IgM ELISA Kit; Ref. #LAPL-K-HRP-00GM; distributed by Louisville APL Diagnostics, Inc., Doraville, GA 30360- Quality Control Fact Sheets for the Louisville APL Diagnostics, Inc. (address now in Seabrook, TX 77586). |  |  | October 06, 2005 | 2005 | Louisville Apl Diagnostics Inc | Mislabeing : Errors (control values) identified in the QC Fact Sheet for a specific lot of APhL HRP IgG and IgM test kits. | Other | Firm notified consignees of recall via fax on 05/05-06/2005. Consignees asked to replace Quality Control Fact Sheets shipped with test kits with the corrected Quality Control Fact Sheets provided with fax. |  | 29 | 29 test kits (ea. kit containing 1 IgG and 1 IgM test kit) | World Wide: Product distributed to direct accounts in CA, FL and GA. Foreign consignees located in Canada and Turkey. | Terminated | October 11, 2007 | 735 | 70 | Undetermined
Z-1169-2007 | 38352 | No | JQW | Pipetting and diluting system for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Microlab F.A.M.E. reagent containers, Part number 148482, box containing 48 containers each. For in vitro diagnostic use. |  |  | March 29, 2007 | 2007 | Hamilton Co | Some product containers may be leaking. | Other | The firm contacted consignees via certified mail and e-mail, initated on March 29, 2007. |  | 797 | 797 boxes | Product was distributed to 3 US consignees in NJ, WA and TX, with one consignee in France. | Terminated | December 21, 2007 | 267 | 797 | Undetermined
Z-1165-2007 | 38035 | No | FMK | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | One Touch SureSoft Lancing Device in carton. Part Number 021-140-01, aka OT SureSoft Hospital, Gentle 200 |  |  | March 30, 2007 | 2007 | Lifescan Inc | Lancet tip may not fully retract and cause needle stick and exposure to used lancets. | Nonconforming Material/Component | The firm has conducted a health hazard evaluation, and has initiated a corrective action plan. The firm initated recall on March 30, 2007, and notifications were sent to hospital end users known to Lifescan and its direct accounts. In addition, a customer service script has been developed to address any questions from incoming users. The firm has requested direct accounts to perform subrecalls on Lifescan's behalf to notify customers of the recall as needed. |  | 632 | 632 cartons (200 units per carton) | Product was distributed nationwide to distributors and wholesale dealers. | Terminated | December 03, 2007 | 248 | 632 | Undetermined
Z-1167-2007 | 38217 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROSÂż 5,1 FS (Fusion Series) Chemistry System Software Version 2.2.1, only when using VITROSÂż Chemistry Products THC (Cannabinoids) or AMPH (Amphetamine) Reagents, Cat. No. 680 1375. |  |  | May 17, 2007 | 2007 | Ortho-Clinical Diagnostics | Software design. The firm received a customer complaint regarding a qualitative positive THC (Cannabinoids) result obtained from a THC quality control fluid that was known to be negative when using VITROSÂż Chemistry Products THC Reagent. A second customer complaint was received a month later. OCD's investigation concluded that under specific circumstances, incorrect qualitative results for Cannabi | Other | On 5/17/07, OCD informed its foreign affiliates of the field correction by e-mail, and instructed them to inform their customers who received the VITROS 5,1 FS Chemistry Systems as well as the VITROS Chemistry Products THC (Cannabinoids) and/or AMPH (Amphetamine) Reagents of this issue. On 5/18/07, all US OCD direct customers and end-use customers of consignees that received the VITROS 5,1 FS Chemistry Systems as well as the VITROS Chemistry Products THC (Cannabinoids) and/or AMPH (Amphetamine) Reagents, were informed of the "field correction" by US Postal Service letter, provided with a new Assay Data Disk (ADD #5465), and were instructed to load it and recalibrate these assays, if necessary. |  | 1202 | Domestic: 604, International: 598 | Nationwide and Worldwide | Terminated | May 02, 2012 | 1812 | 1202 | Undetermined
Z-1232-2007 | 38635 | No | CHI | Lipase test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Clinical Chemistry Lipase, List Numbers: 7D80-20 and 7D80-30 |  |  | June 01, 2007 | 2007 | Abbott Laboratories Inc. | Lipase results are falsely elevated when a MULTIGENT Acetaminophen test is pipetted immediately before a Lipase test from the same sample cup or tube on the ARCHITECT c8000. | Other | The firm initiated the recall on 06/01/2007 by sending recall letters with attached Customer Reply form via FedEx to each direct account that received the recalled product. Customers were informed of the issue and given instructions for modification to the SmartWash feature in the Lipase Assay Configuration in order to prevent Acetaminophen from contaminating Lipase. |  | 7260 | 7260 | Worldwide, including USA, Australia, Brazil, Canada, Chile, El Salvador, Dominican Republic, Ecuador, Germany, Hong Kong, Japan, Mexico, New Zealand, Panama, Puerto Rico, Singapore, South Korea, Thailand, and Venezuela. | Terminated | April 05, 2012 | 1770 | 7602 | Undetermined
Z-1222-2007 | 38439 | No | JJX | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ARCHITECT Prolactin MasterCheck; a single analyte control - each kit contains 1 - 2 mL bottle of Architect Prolactin MasterCheck Level 0 containing HEPES buffer with protein (bovine) stabilizers, and 4 - 2 mL bottles of Architect Prolactin MasterCheck Levels 1, 2, 3 and 4 containing prolactin (human) prepared in HEPES buffer with protein (bovine) stabilizers; Abbott Laboratories, Abbott Park, IL 60064, list 6C27-05 |  |  | July 06, 2007 | 2007 | Abbott Laboratories | The Architect Prolactin MasterChecks may not be stable at the recommended storage conditions of 2 to 8 degrees C through the expiration dating (15 months). The MasterChecks may exceed the suggested tolerance limits listed in the Reportable Range/Calibration Verification section in Appendix B of the Architect iSystem Operations Manual. | Other | Abbott sent a Product Correction letter dated 7/6/07, to all customers who received the affected product lot, informing them that the ARCHITECT Prolactin MasterChecks may not be stable at the recommended storage condition of 2 to 8 degrees Celsius (C) through expiration dating (15 months). Labeling is being permanently updated to indicate that ARCHITECT Prolactin MasterChecks are to be stored frozen until first use. After thawing, ARCHITECT Prolactin MasterChecks can be stored unopened at 2-8 degrees C for thirty days or until expiration which ever occurs first. All future shipments will contain updated storage and handling conditions of the ARCHITECT Prolactin MasterCheck. They were requested to check their inventory to see if they have any remaining inventory of lot number 82900M. If they know the date the lot was received or the date it was thawed, they may continue to use the ARCHITECT Prolactin MasterChecks if it has been stored at 2 to 8 degrees C unopened for less than 30 days. If the material has been stored at 2 to 8 degrees C for greater than 30 days or they do not know the number of days it has been thawed, they were instructed to discard the material. Reimbursement for discarded kits will be provided based on information provided in the Customer Reply form. Until labeling is updated, it will be necessary to use an alternate material for verification of sensitivity, calibration linearity, and the reportable range of the ARCHITECT Prolactin assay. They were requested to forward a copy of the letter to any laboratory to whom they may have forwarded the kits, and to retain a copy of the letter for their records. Any questions were directed to Abbott Customer Support at 1-877-4ABBOTT in the U.S. The accounts were requested to complete and fax back the customer reply form, acknowledging the receipt and understanding of the correction letter. Customers outside of the U.S., were requested to contact their local customer support representative. |  | 57 | 57 kits | Nationwide and Canada. | Terminated | May 18, 2008 | 317 | 57 | Undetermined
Z-1220-2007 | 38348 | No | MQK | Assisted reproduction labware. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | IVF Pasteur Pipettes, glass, cotton plugged, length 9", in boxes of 90 consisting of 30 nylon tube packages of 3 pipettes each. Item # PP-9-90PL. Boxes are labeled in part ***Humagen Fertility Diagnostics, Inc. 2400 Hunter's Way Charlottesville, VA 22911*** |  |  | May 08, 2007 | 2007 | Humagen Fertility Diagnostics, Inc. | Pasteur pipettes for laboratory use were labeled as sterile but may have become contaminated. | Other | The recalling firm notified consignees by phone and e-mail on 05/07/07. International consignees were advised to discard product for replacement. |  | 160 | 160 boxes | The recalled product was distributed to 2 direct accounts in IL and to 6 international distributors in United Kingdom, Greece, Australia, United Arab Emirates, the Netherlands, and Israel. | Terminated | October 29, 2007 | 174 | 160 | Undetermined
Z-1124-2007 | 44899 | [6 empty] | Laser Distance and Level Transmitters, Model No. LT200, ASI, Housotn, TX 77024 |  |  | November 03, 2006 | 2006 | K Tek | Failure to comply with Federal Laser Product Performance Standards (21 CFR 1040.10, 1040.11, 1002.10-improper classification, incomplete labeling, incomplete user information and failure to submit required product reports) | Other | The firm notified its consignee of problem and the field correction by letter dated 11/03/2006. |  | 21 | 21 units | Nationwide | Terminated | December 10, 2008 | 768 | 1008 | Undetermined
Z-1238-2007 | 38372 | No | NEP | Intravascular administration set. | General Hospital | General Hospital | 510(K) Exempt | Exacta-Mix 2400, OS v1.07, Model No. 8300-0073, Pharmacy Compounding System, Baxa Corporation, Englewood, CO 80112. |  |  | June 27, 2007 | 2007 | Baxa Corporation | A software defect could allow up to 50mL of extra volume being added to a TPN solution. | Other | All customers were notified via telephone on 06/27/2007 with follow-up of a faxed/emailed Urgent Product Recall letter dated 27 June 2007, detailing the software defect. End users were given a work around procedure to use until conversion back to version 1.06 could be scheduled. An additional copy of the recall letter was sent by FedEx. |  | 30 | 30 units | Class I Recall-Worldwide Distribution: USA including states of CA, CO, FL, KY, MI, NY, OH, OR, and TX and the UK | Terminated | December 07, 2007 | 163 | 30 | Undetermined
Z-1245-2007 | 38301 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | daVinci S Surgical system, patient side cart Auxiliary power board (APB) (printed circuit board assembly), endoscopic instrument control system. |  |  | May 14, 2007 | 2007 | Intuitive Surgical, Inc. | Product may malfunction and fail to start up on AC power. | Process control | As of 9/12/2007, the firm has conducted a field inspection of any involved units via service visit. No other notice is planned or indicated by the firm's submission. The firm feels that the risk associated with this issue is minimal, and therefore, any advanced notification is unneccessary. If any system failed one or more of the inspection points, its APB (3510404-04) was replaced with a verified APB (351040-04 or 351040-05). |  | 38 | 38 units | Worldwide, including USA, France, China, Japan, Mexico, Spain, Italy, Belgium and Singapore. | Terminated | December 03, 2007 | 203 | 38 | Undetermined
Z-1263-2007 | 44702 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Altaire Shoulder Coil, an accessory to the Altaire MRI system. The shoulder coil is a two-channel, receive only coil that functions at 29.641 MHz band. It consists of a cable and a coil (imaging) body. A portion of the cable is referred to as the balun assembly or balun trap. The shoulder coil assembly is packaged in a cardboard box with foam inserts. |  |  | July 30, 2007 | 2007 | Hitachi Medical Systems America Inc | Overheating; A portion (balun assembly) of the cable that connects the shoulder coil to the actual MRI unit may malfunction and result in overheating of the balun assembly. The overheating of the balun assembly could result in injury to the patient or operator if either came into contact with the assembly. | Other | On 7/30/07, Hitachi sent a Device Correction letter via courier with receipt tracking to all affected customers outlining the risk of overheating of the cable balun assembly and subsequent risk of patient/operator injury. The letter notifies customers that (1) Hitachi will be contacting them in the future to schedule an on-site visit for replacement of the Altaire Shoulder Coil; (2) the replacement will begin in September, 2007 and continue until all affected coils are replaced; and (3) customers that have noticed overheating of the cable balun assembly are urged to notify Hitachi for immediate replacement service. |  | 273 | 273 | Nationwide: U.S. A. to include: AL, AR, AZ, CA, CO, CT, DE, FL, HI, ID, IL, IN, KS, KY, LA, MD, MI, MN, MO, MS, MT, NC, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WI, &amp; WY. | Terminated | October 05, 2010 | 1163 | 273 | Undetermined
Z-0555-2007 | 39378 | No | BRY | Anesthetic cabinet, table, or tray. | Anesthesiology | Anesthesiology | 510(K) Exempt | Pyxis Anesthesia System 3500, Cardinal Health, Clinical Technologies and Services, San Diego, CA |  |  | August 13, 2007 | 2007 | Cardinal Health 303 Inc DBA Alaris Products | System lock up; mini-drawers -under certain circumstances during use, are not accessable, delaying the user from accessing critical medications. | Other | The field action consisted of phone calls and a mailing conducted to all customers for the Pyxis Anesthesia System 3500 beginning on 08/13/2007. The customers will be notified by registered return receipt mail by Recall Notifications dated August 14, 2007. Separate letters will be sent to the Director of Pharmacy, Chief of Anesthesia, Director of Risk Management, and Director of Surgery at each facility. Recommended actions will be provided as a short-term fix and a long-term fix of the software will be issued shortly. |  | 369 | 369 | Nationwide to Arizona, California, Colorado, Florida, Georgia, Hawaii, Iowa, Illinois, Indiana, Louisiana, Massachusetts, Missouri, Mississippi, Montana, New Jersey, Ohio, Oklahoma, Pennsylvania, Rhode Island, South Carolina, Tennessee, Texas, Virginia, Vermont, and Wyoming. | Terminated | March 28, 2012 | 1689 | 369 | Undetermined
Z-1261-2007 | 38815 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Trident PSL Acetabular Shell;Arc Deposited;Hydroxylapatite Coated;Cluster Screw Holes;Howmedica Osteonics Corp., |  |  | March 13, 2006 | 2006 | Stryker Howmedica Osteonics Corp. | Dimensional anomaly | Other | Recall Notification letters were sent on March 13, 2006 to all hospitals, branches/agencies via Federal Express. |  | 1220 | 1,220 units | Worldwide -USA, Romania, Chile and Brazil. | Terminated | January 18, 2008 | 676 | 1220 | Undetermined
Z-0014-2007 | 37567 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Brilliance(a) CT 6 slice, Model #4535 670 73191. |  |  | November 06, 2006 | 2006 | Philips Medical Systems (Cleveland) Inc | A potential anomaly was identified on the CT subsytem, which could result in the radiation dose being different than the calculated dose reported to the operator. | Other | The customers were sent Public Safety Notification November 3, 2006, indicating there was a problem with the software. However, since their unit was recently upgraded the problem did not impact their unit. |  | 20 | 20 | Worldwide, including USA, Australia, Canada, China, England, Finland, France, Germany, Ireland, Israel, Italy, Japan, Jordan, Korea, Netherlands, Poland, Puerto Rico, Spain, and Taiwan. | Terminated | October 05, 2010 | 1429 | 90 | Undetermined
Z-1242-2007 | 38459 | No | JXK | Cranial electrotherapy stimulator. | Neurology | Neurology | 510(k) | HealthPax Cranial Electrotherapy Stimulator |  |  | July 19, 2007 | 2007 | Health Directions, Inc | Peak voltage test result not in acceptable limits. | Other | The recalling firm sent a recall letter 07/30/2007 to all their customers. The letter indicated that it is necessary to return the product for testing and the device will either be returned or if necessary replaced. Each unit will be tested with an oscilloscope device to determine whether it meets or fails design specifications. The recalling firm will reimburse shipping costs. |  | 212 | 212 units | Worldwide, including USA, Canada, Ireland, and India. | Terminated | June 03, 2008 | 320 | 212 | Undetermined
Z-0764-2007 | 37844 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Practix 300 Mobile X-ray systems. The systems are used in general purpose diagnostic imaging, and are specifically used as mobile X-ray systems for bedside radiological examinations. |  |  | March 23, 2007 | 2007 | Philips Medical Systems North America Co. Phillips | X-ray systems lack product report certification demonstrating compliance, and were illegally introduced into the United States commerce. | Other | The firm sent letters out via certified mail from 04/03/2007 through 04/09/2007 to consignees informing them of the issue. The letter indicates that the firm plans to replace consignee's systems with the Practix Convenio Mobile X-ray System at no charge by the end of May 2007. |  | 9 | 9 units | Nationwide: Distributed to Puerto Rico. | Terminated | September 16, 2010 | 1273 | 11 | Undetermined
Z-0911-2007 | 44910 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Philips Omni Diagnost Eleva |  |  | July 23, 2007 | 2007 | Philips Medical Systems North America Co. Phillips | Units do not comply with the requirement to provide automatic beam limitation within +/- 3 degrees of horizontal, vertical, and perpendicular to the plane of the image receptor during radiographic procedures. | Other | On 7/23/07 Philips mailed a letter dated July 9, 2007 titled Electronic Product Radiation Warning. The letter advises customers of the issue and asks that customers apply the enclosed cautionary label to the device. The letter also explains that there will be a system modification to be installed. On 8/20/07 the firm sent the same letter, along with a letter dated 8/20/07 to customers thought to have either scrapped their units and the customer had previously indicated they were no longer using the unit. The 8/20/07 letter requested the customer contact Philps if they are currently using the device. |  | 14 | 14 units | Units were distributed to 1543 hospitals and medical centers throughout the US. No international distribution from this firm. | Terminated | September 17, 2012 | 1883 | 3244 | Undetermined
Z-1087-2007 | 38633 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Proteus XR/a General Purpose Radiographic System, Models 2259973-4 and 2380179 containing Proteus Operator Console model# 2259976. |  |  | May 31, 2006 | 2006 | General Electric Med Systems LLC | Labeling; failure to properly label per 21 CFR 1010.2 and 1010.3 -The Operator Console of the Proteus XR/a General Purpose Radiographic Systems were not properly labeled. | Other | Consignees were notified of the problem by a visit. To address the Proteus System shipped to the field, a Field Modification (FMI) was created to correct units in the field. |  | 22 | 22 units | Nationwide; including states of WA,KS, CO, NM, NC, FL, MT, AZ, TX, IL, NY, MD, and GA | Terminated | December 11, 2007 | 559 | 22 | Undetermined
Z-1212-2007 | 39420 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | The Pinnacle3 Radiation Therapy Planning (RTP) system and its associated software; composed of several modules including the core Pinnacle3 functionality, Syntegra (TM), P3IMRT, P3MD and AcQSim3 (TM). The Pinnacle3 RTP system is composed of a Sun UNIX workstation (or UNIX-complaint computer) running the Solaris operating system and software, and a UNIX terminal emulation package on a personal computer, which allows the user to enter patient data into the system, use that data to construct a plan for radiation therapy, and evaluate the plan- Philips Medical Systems-Model #9200-0652A-ENG, 9200-0700A-ENG, 9200-0700B-ENG, 9200-0700C-ENG (with software versions 7.6C, 8.0d and 8.0h) |  |  | August 21, 2007 | 2007 | Philips Nuclear Medicine, Inc. | Incorrect Absolute Marking Coordinates when used with PQ-Series scanners due to a software malfunction | Other | All customers will receive an August 21, 2007, Customer Advisory Notification letter notifying them of the issue and the steps required to avoid the issue. Return receipts will be tracked. This is a short-term fix, a long term software fix will be issued shortly by the firm. |  | 1407 | 1,407 | Worldwide Distribution; USA including states of AL, AK, AZ, AR, CA, CO, CT, DE, DC, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MT, NE, NV, NH, NJ, NM, NY, NC, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VT, VA, WA, WV, and WI and Country of CANADA | Terminated | December 17, 2011 | 1579 | 1407 | Undetermined
Z-1227-2007 | 38485 | No | DWE | Cardiopulmonary bypass pump tubing. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Intersept 3608 Sterile tubing (8') 1/2 x 3/32 in. Quantity 10, Lot 7626621, Use by 2009-07-31. Manufacturer: Medtronic, Inc. Minneapolis, MN 55432-5604 USA. |  |  | August 01, 2007 | 2007 | Medtronic Perfusion Systems | Medtronic Custom Tubing Packs might have open header bag seals. | Other | Certified letters dated 8/1/07 were sent to customers. The letter explains the situation and requests the impacted product to be returned to Medtronic for replacement or credit. The customers are being requested to complete a Recall Certificate acknowledging the receipt of the information. |  | 853 | 853 packs | Nationwide. | Terminated | October 10, 2008 | 436 | 853 | Undetermined
Z-1231-2007 | 38624 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Triathlon X3 |  |  | July 31, 2007 | 2007 | Stryker Howmedica Osteonics Corp. | Stabilizer posts, packaged with the Triathlon TS Plus Tibial inserts have a specification anomaly related to length, resulting in a potential for the post to slightly protrude above the top surface of the plastic post. | Other | Recall notification letters and product acknowledgment forms were send via Federal Express on August 2, 2007. |  | 14 | 14 | Product distributed to 5 hospitals in GA, OH, MN, TX. | Terminated | December 05, 2007 | 127 | 14 | Undetermined
Z-1256-2007 | 37874 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Reverse Shoulder Prosthesis (RSP) device component; RSP Stem; Device manufactured by Encore Medical, L.P., Austin, TX 78758. |  |  | March 22, 2007 | 2007 | Encore Medical, Lp | Marketed without approval: Investigational devices were implanted during an IDE study prior to receiving approval for the devices in the study; patients were implanted with investigational devices outside of the IDE study; and devices implanted as "custom devices" by the manufacturer did not meet the definition of "custom devices" and were, therefore, unapproved devices. | Other | Letters were sent via Certified Mail on 03/22/07 to the surgeons and their associated IRBs. Each letter included a list of the patients to be notified and a form letter that the surgeons could use to notify the patients. Surgeons asked to notify their patients and to provide Encore with documented evidence that patience were notified. |  | 120 | 120 units. | Nationwide: Devices sold to and implanted by physicians in CA, CO, FL, GA, NY, SC and TX. | Terminated | March 27, 2008 | 371 | 678 | Undetermined
Z-1192-2007 | 38630 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Newport HT50 Ventilator (all Models) |  |  | July 31, 2007 | 2007 | Newport Medical Instruments Inc | Battery Life: With use over time, the internal battery begins losing its ability to hold a charge resulting in shortened ventilator operation time. Specifically, Newport has received a number of reports concerning the internal battery not lasting as long as expected followed by ventilator shut down with less advance warning than when the battery was new. | Other | An Important Medical Device Correction Letter was issued on September 14, 2007 as a short term fix. The firm has submitted a long term correction as a supplement to its 510 K which is pending review by FDA. |  | 8160 | 8,160 | Worldwide including USA, Algeria, Argentina, Armenia, Australia, Belgium, Bolivia, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Cote D' Ivoire, Croatia, Cyprus, Dominican Republic, Ecuador, Egypt, El Salvador, England, France, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Kenya, Korea, Kuwait, Liechtenstein, Lithuania, Malaysia, Mexico, Netherlands, New Caldonia, New Zealand, Oman, Paraguay, Peru, Philippines, Poland, Portugal, Qatar, Russia, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Switzerland, Syria, Taiwan, Thailand, Turkey, UAE, Ukraine, United Kingdom, Venuzuela, Yemen, and Zambia | Terminated | March 21, 2012 | 1695 | 8160 | Undetermined
Z-1239-2007 | 38406 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Stryker Radius Locking Caps, Catalog Number: 486610000, Manufactured by Stryker Spine SAS Z.I. Marticot - 33610 Cestas France.,Distributed by Howmedica Osteonics Corp. 325 Corporate Drive, Mahway, NJ 07430 |  |  | July 10, 2007 | 2007 | Stryker Spine | Product dissembled in the tray or prior to being loaded onto the inserter or tightening on the rod | Other | Product recall letters were sent to branches/agencies and hospitals on 7/10/2007 requesting product return. |  | 1423 | 1,300 in US, 123 in Italy | Worldwide- USA: Distributed to Stryker Branches/Agencies in MA, NY, NC. IN, MI, OK, TX,CO, AZ, WI, KS, CA, OR and hospitals in NC IN, MI, MO. KS, TZ, AZ, CA. OR and WA and Country of Italy | Terminated | December 11, 2007 | 154 | 1423 | Undetermined
Z-1241-2007 | 38814 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Duracon Total Knee Modular Femoral Component; |  |  | October 17, 2006 | 2006 | Stryker Howmedica Osteonics Corp. | Loose fixation pegs: One lot of femoral components may have loose fixation pegs. | Other | Notification letters, dated 10/17/2006, were sent via Federal Express on 10/17/2006 to all consignees requesting return of product. |  | 3 | 3 total, 2 in USA | Worldwide : One Hospital in Georgia (USA), Stryker Branch in Indiana and Stryker Branch in Sweden. | Terminated | July 07, 2009 | 994 | 3 | Undetermined
Z-1240-2007 | 38472 | No | DQA | Oximeter. | Cardiovascular | Anesthesiology | 510(k) | BCI Autocorr Plus Pulse Oximeter/ECG/Respiration Monitor, Catalog Numbers, 3404001 (including S) and 3404002 (including E,F, G, and S) Smiths Medical- a part of Smiths Group plc |  |  | July 18, 2007 | 2007 | Smiths Medical PM, Inc. | Phantom pacer spikes caused by a damaged resistor. | Other | Customers will be notified via a Customer Information Bulletin CIB Tracking Number: 07-CIB 12, July 18,2007 requesting them to evalute the functionality of their monitor (s). If the monitor does not exhibit any false pace detect makers when pace detect is enabled and respiration is disabled, no further action is required. If it does exhibit false pace detect they will be instructed to contact the Smiths Medical PM, Inc service department. At that time Technical Worksheet, 07-TW02, will be provided. The consignee will need to verify they are capable of perfoming the Technical Worksheet provided. The Smiths Medical PM. Inc. Service Manger will veirfy if the consignee is capable of performing the rework. |  | 178 | 178 | World wide: USA including states of KY, MS, PA, LA, and PR and countries of Poland, Mexico, Kuwait, Taiwan, Singapore, Hon Kong, India, Bolivia, Colombia, Chile, Argentia, Brunei Darssalam, Israel, Greece, Romania, Germany, The Netherlands, France, Great Britain, Spain, Brazil, Russia, China. | Terminated | October 31, 2008 | 471 | 178 | Undetermined
Z-1009-2007 | 37978 | No | FPO | Wheeled stretcher. | General Hospital | General Hospital | 510(K) Exempt | Stryker Medical -Wheeled Electric Stretcher, Model 1550 (with and without scales) |  |  | May 07, 2007 | 2007 | Stryker Medical Div. of Stryker Corporation | The internal power cord may become damaged due to wear and, if this occurs, a shock hazard exists. | Other | Consignees were notified of the problem via letter dated 5/7/07 instructing them to immediately unplug the affected stretchers with scales, affix a warning label enclosed with the letter to the power cord of stretchers with scales, and that they will be contacted so that the firm can schedule a visit to upgrade all affected stretchers. A second notification letter was issued on August 7, 2007 to cover all 1550 stretchers with and without scales. |  | 1048 | 1,048 (1009 in the U.S. plus 39 international) | Worldwide: USA, Latin America, Mexico and Spain. | Terminated | June 11, 2008 | 401 | 1048 | Undetermined
Z-1229-2007 | 38621 | Yes | MPB | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Guidewire, 70cm x 1mm, REF 5583705, Sterile, Single Use, BARD Access Systems, Hemodialysis Catheter, Salt Lake City, UT 84116. |  |  | August 10, 2007 | 2007 | C.R. Bard Access Systems, Inc | Metal guidewires used in Hemodialysis Catheters may have been reverse loaded within the dispenser. Potential exists for the stiff end of the guidewire to be inserted into the patient instead of the flexible end. | Other | Consignees were notified by letter on 08/10/2007 and asked to remove and return all unused affected product to Bard. |  | 7225 | 7225 units | Worldwide, including USA, Canada, Latin America, and Europe (Austria, Belgium, Finland, France, Germany, Great Britain, Holland, Italy, Norway, Spain, and Sweden). | Terminated | April 25, 2008 | 259 | 7225 | Undetermined
Z-1237-2007 | 39413 | No | MRG | Vitamin D test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | 25-Hydroxyvitamin D 125 I RIA Kit, in vitro diagnostic. REF:68100E DiaSorin Stillwater, Minnesota 55082-0285, U.S.A. |  |  | August 21, 2007 | 2007 | Diasorin Inc. | 25 Hydroxyvitamin D RIA kits may contain the incorrect lot of tracer. | Other | Consignees were notified by phone on 8/21/07 followed by a written notification by Fax or email. The letter described the product and problem and recommended to inspect the kits and to destroy all tracer vials of lot number 726023. They also requested the consignee to return the findings by fax or mail. |  | 117 | 117 kits | Worldwide, including USA, CANADA, BRAZIL, INDIA, QATAR, SWEDEN, URUGUAY, ARGENTINA, and JORDAN. | Terminated | October 17, 2008 | 423 | 117 | Undetermined
Z-1230-2007 | 38623 | No | MUU | Suction lipoplasty system. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | 5mm x 30cm Hollow "Golf" Tip Ultrasonic Aspiration Probe, Model 5620-530, Catalog Number 5620-530. |  |  | July 27, 2007 | 2007 | Mentor Texas, Inc | Mis-identified and mis-labeled surgical probes were distributed. 5mm x 30cm Hollow "Golf" Tip Ultrasonic Aspiration Probe, Model 5620-530, Catalog Number 5620-530 was shipped, instead of LySonix 5mm x 30cm Hollow Bullet Tip Ultrasonic Probe, Catalog Number 5610-530. | Other | Firm notified consignees by letter on July 27, 2007. |  | 4 | 4 units | Nationwide, including VA, IN, CA, and IL. | Terminated | April 04, 2008 | 252 | 4 | Undetermined
Z-1236-2007 | 39386 | No | NGT | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Normal Saline 10ml in 12ml Syringe. The product is shipped in 120 units per carton. |  |  | July 30, 2007 | 2007 | B. Braun Medical, Inc. | Particulate matter in product. | Other | The recalling firm issued a Product Removal Notice to the distributors on 7/31/07 informing them of the problem and the need to contact their customers. Product should be returned to the recalling firm for destruction. On 9/14/07, the recalling firm issued a press release to alert any consumer that may have the product of the particulate problem. |  | 1292520 | 1,292,520 units | The products were shipped to distributors nationwide. | Terminated | April 01, 2008 | 246 | 1325760 | Undetermined
Z-1225-2007 | 38424 | No | GWQ | Electroencephalograph. | Neurology | Neurology | 510(k) | BMSI "MicroJack" PN 672-510000 patient interface electrode connection device used during electroencephalography (EEG) /Epilepsy Monitoring evaluations with various EEG monitoring systems. The MicroJack is used with BMSI Models 2000, 3000, 4000, 5000 &amp; 6000, and is used with the BMSI kit upgraded with NicoletOne. Nicolet Biomedical, Inc. |  |  | July 17, 2007 | 2007 | Nicolet Biomedical Div of Viasys Healthcare | The BMSI "MicroJack" patient interface electrode connection device used with the EEG monitor systems was found to be mis-wired. A complainant alleged that during depth electrode EEG study of a patient in preparation for brain surgery there was an observation that was perceived as non-logical data output. | Other | VIASYS will initiate a field correction for units that have been distributed. The correction consists of testing the units and replacing those that have been found to contain the defect. Method of communication will be by certiified mail to identify consignees and receipt of the communication will be tracked. Letters "Viasys HealthCare Urgent Medical Device Field Correction (Recall)" were sent 7/17/07 to both domestic and international customers. |  | 2102 | 2102 | Worldwide, including USA, ISRAEL, LITHUANIA, RUSSIA, FRANCE, IRELAND, UNITED KINGDOM , SPAIN, YEMEN, JAPAN, TAIWAN, KOREA, CHINA, SAUDI ARABIA, HONG KONG, INDIA, COLOMBIA, BRAZIL, GERMANY, CANADA, DENMARK, ITALY, MEXICO, and THAILAND. | Terminated | December 10, 2011 | 1607 | 2102 | Undetermined
Z-1228-2007 | 38546 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | NxStage Cartridge Express, CAR-170, for hemodialysis. |  |  | July 27, 2007 | 2007 | NxStage Medical, Inc. | Potential for dialysate leaks from the cartridge during treatment. | Process change control | NxStage initiated the recall by letter on 8/24/07 advising users of the specific lots to be removed from use. Prior notifications were issued on July 27, 2007, July 30, 2007 and July 31, 2007 notifying users of the potential for leaks and instructions to detect leaks prior to use. |  | 12085 | 12,085 cases | Nationwide | Terminated | September 13, 2010 | 1144 | 12085 | Undetermined
Z-1226-2007 | 38399 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | ReliOn 0.5cc 30 g x 8mm Blister Pak U-100 syringe |  |  | June 21, 2007 | 2007 | Becton Dickinson &amp; Company | Complaints received indicating open seals on blister paks, potentially compromising the sterility of the product. | Other | Urgent Product Recall notifications were sent on 6/21/07 asking them to discontinue use of product, remove from inventory and complete Tracking Verification form. |  | 644500 | 644,500 units | Nationwide. | Terminated | June 03, 2008 | 348 | 644500 | Undetermined
Z-1223-2007 | 39368 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Welch Allyn AED 20 Automated External Defibrillator; a portable, 12V internal battery powered defibrillator; Manufactured by MRL Inc., A Welch Allyn Company, Buffalo Grove, Illinois; Model AED 20, Catalog Numbers 972213, 972214, 972215, 972216 (Part Numbers 972200E, 972211E, 972212E, 972213E, 972214E, 972215E, 972216E) |  |  | August 23, 2007 | 2007 | Welch Allyn Protocol, Inc | Failure to deliver therapy: The Welch Allyn AED20 Defibrillators may display a "Defib Comm" error message on the device display during use which may result in failure of the device to analyze the patient's ECG and deliver the appropriate therapy. | Process control | On 8/23/07 the firm sent a press release to the Associated Press for release. On 8/29/07 the firm began sending URGENT MEDICAL DEVICE RECALL letters to their consignees informing them of the possibility of the defibrillator to fail to resuscitate the patient. The accounts were asked to respond to the notification within 5 working days by faxing or mailing an enclosed response. End users will be provided with a loaner device while their unit is being serviced. Distributors were sent a separate letter request they provide their customer names to the firm so that the firm can send letters to those customers. |  | 1637 | 1,637 devices | Worldwide to include USA, Australia, New Zealand, South Africa, China, Hong Kong, Taiwan, Malaysia, Singapore, Phillippines, South Korea, Viet Nam, Egypt, France, Germany, Ireland, India, Israel, Pakistan, Poland, Russia, Saudi Arabia, Switzerland, United Arab Emirates, United Kingdom, Canada, Ecuador, Puerto Rico, and South America. | Terminated | September 26, 2011 | 1495 | 1637 | Undetermined
Z-1224-2007 | 39415 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Colleague CX Single Channel Volumetric Infusion Pumps; Baxter Healthcare Corporation, Medication Delivery Division, Deerfield, IL 60015 U.S.A., Made in Singapore; product code 2M8161 |  |  | August 15, 2007 | 2007 | Baxter Healthcare Corp. | Some of the pumps were upgraded by one service technician who failed to replace prisms and batteries as required in the upgrade. | Other | The affected hospitals were sent an Urgent Colleague Deployment Correction letter dated 8/15/07 via overnight mail. The hospitals were informed that some of the Colleague pumps at their facility may not have been completely upgraded to include all the necessary components required as part of Baxter's Corrective Action Plan (CAP). The hospitals were provided a list of affected pump serial numbers and were requested to remove the pumps from clinical use and notify all end users of the action. The accounts were requested to complete the enclosed customer reply form, acknowledging the receipt and understanding of the letter and fax it back to Baxter at 1-847-270-5457. The hospitals were also requested to report any malfunction or adverse event related to the Colleague Infusion Pumps to Baxter Corporate Product Surveillance at 1-800-437-5176. |  | 42 | 42 pumps | Florida and California | Terminated | February 21, 2008 | 190 | 42 | Undetermined
Z-0834-2007 | 37701 | No | FRO |  |  | General &amp; Plastic Surgery | 510(k) | FX Spare Shieldwire Inflators (3.0 - 5.0 mm), Catalog number 60042-02. The product is sold 5 per box. |  |  | March 28, 2007 | 2007 | Kensey Nash Corp | Insufficient wire sealing | Other | On 3/26/07, the recalling firm left voice mail message with all sales representatives instructing the sales representatives to contact all TriActiv FX accounts immediately. Recall notices were faxed on 3/28/07 to all accounts which received the product. Recall notices instructed the hospitals to quarantine all FX and to contact their sales representative to arrange for retrieval of the product if they had not already been contacted. |  | 109 | 109 units | Worldwide, including USA and Germany. | Terminated | August 30, 2007 | 155 | 547 | Undetermined
Z-0852-2007 | 38279 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE OEC 2800 mobile Fluoroscopic X-ray System, GE Healthcare, Salt Lake City, Utah. |  |  | June 12, 2007 | 2007 | GE OEC Medical Systems, Inc | X-ray systems were not configured to display Air Kerma Rate and Cumulative Air Kerma. | Other | Consignees were notified by letter on 06/12/2007 and told that a service representative would arrange for an inspection and would modify their systems. |  | 25 | 25 units | Worldwide Distribution: USA including states of FL, GA, PA. CO, MD, VA and WA. and countries of Australia, Austria, Belgium, Brazil, Canada, Colombia, Finland, France, Germany, Hong Kong, India, Israel, Italy, Japan, Korea, Netherlands, New Zealand, Norway, Portugal, Switzerland, Taiwan, and the UK. | Terminated | November 01, 2011 | 1603 | 575 | Undetermined
Z-0827-2007 | 37903 | Yes | DTZ | Cardiopulmonary bypass oxygenator. | Cardiovascular | Cardiovascular | 510(k) | Endura No-React Dural Substitute 2 cm x 10 cm |  |  | May 04, 2007 | 2007 | Integra LifeSciences Corp. | Integra LifeSciences Corporation d/b/a Integra NeuroSciences, the distributor of EnDura No-React Dural Substitute, manufactured by Shelhigh, Inc., Union, NJ, is recalling this product based on an FDA Public Health Notification and FDA requested recall letter to Shelhigh citing sterility and other manufacturing concerns regarding this product manufactured by Shelhigh. Integra is a distributor of t | Other | On May 4, 2007, an Urgent Recall Notice and accompanying recall acknowledgement form was sent to all customers that had purchased or been supplied EnDura No-React Dural Substitute product since sale of the product was initiated in 2003. Recall Coordinators at each facility are asked to review inventory for EnDura products and to call Integra Customer Service for an authorization number for return of product. |  | 12736 | 12,736 for all recalls Z-0827-0832-2007 | Nationwide Distribution to hospitals and medical centers | Terminated | September 24, 2008 | 509 | 12736 | Undetermined
Z-0501-2007 | 37002 | No | NRP | Sirolimus test system. | Toxicology | Toxicology | 510(k) | IMx Sirolimus Reagent Pack; list 5C91-21; 100 test pack containing 1 bottle (9.7 mL) Anti-Sirolimus Antibody Coated Microparticles, 1 bottle (9.7 mL) Sirolimus Alkaline Phosphatase Conjugate, and 1 bottle (10 mL) 4-Methylumbelliferyl Phosphate, 1.2 mM; Produced by Axis-Shield Diagnostics, Ltd., Dundee, UK for Abbott Diagnostics Division, Distributed by Abbott Laboratories, Abbott Park, IL 60064 USA |  |  | August 11, 2006 | 2006 | Abbott Laboratories MPG | The sirolimus values may shift higher after storage at 2-8ÂżC or after one freeze/thaw cycle of specimens. | Other | Abbott initiated a field correction recall by letter dated 8/11/06 to all customers receiving the affected reagent. The accounts were informed of the shift in sample values after storage at 2-8*C or with freeze thaw of specimens with the affected lots of reagents. The accounts were requested to review the revised sample handling information included in the attachment to the letter with their laboratory personnel and retain a copy of the letter and attachment for their records. Customers requiring additional information were directed to contact Abbott Customer Service at 1-877-422-2688. |  | 2800 | 2,800 kits | Worldwide, including USA, Puerto Rico, Poland, Croatia, Latvia, Estonia, Serbia &amp; Montenegro, Japan, Turkey, Singapore, China, India, Taiwan, Thailand, South Africa, South Korea, Australia, New Zealand, Norway, Italy, Portugal, Spain, Greece, Panama, Hungary, Slovakia, Russia, Romania, Cyprus, Germany, Lithuania, Saudi Arabia, Jordan, Israel, Lebanon, Kuwait, Oman, Qatar, UAE, Egypt, Mexico, Guatemala, Venezuela, Colombia, Peru, Chile, Argentina, Brazil, El Salvador, Canada, Ireland, United Kingdom, Sweden, Netherlands, Belgium, France, Austria, Czech Republic, and Switzerland | Terminated | January 18, 2008 | 525 | 2800 | Undetermined
Z-0520-2007 | 37213 | No | KNQ | Esophageal dilator. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Quantum TTC Biliary Balloon Dilation Catheter, Wilson Cook Medical, GI Endoscopy, 4900 Bethania Station Road, Winston-Salem, NC 27105 |  |  | January 24, 2007 | 2007 | Cook Endoscopy | The graphic label is incorrect, it lists 6mm/ 18FR. The product label is correct, it lists 10mm/ 30FR. | Other | Consignees were notified by UPS certified mail and email on/about 01/31/2007. They were instructed to return any affected product to Cook Endoscopy. A response form was attached to be completed and returned to Cook Endoscopy. |  | 12 | 12 units | Worldwide, including NY and TX in USA, Australia, Canada and India. | Terminated | July 26, 2012 | 2010 | 12 | Undetermined
Z-0525-2007 | 37120 | No | FTD | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Skytron Infinity Series (also called Astro series) minor surgery light with dual lightheads, 3 bulbs in each of 2 - 19' diameter lightheads, Manufactured by Dai-ichi Shomei Co., Ltd., Made in Japan; Models IN1919 and IN1919EL and IN1919EL/PT. |  |  | January 08, 2007 | 2007 | Skytron, Div. The KMW Group, Inc | If the mounting plug is improperly adjusted, the unit could separate, come apart, and fall from the mount. | Other | The firm initially notified consignees of this problem on 11/23/1999 and again on 11/01/2002 and most lights have already been corrected. Remaining consignees with lights in need of correction were notified of the problem via letter dated 1/8/07 and instructed to contact the distributor to have their lights inspected and modified. |  | 5000 | 5,000 for all models. | Nationwide and Canada. | Terminated | November 18, 2009 | 1045 | 5000 | Undetermined
Z-0838-2007 | 37900 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | BD AutoShield Pen Needle (Re order Number : 329300) |  |  | April 23, 2007 | 2007 | Becton Dickinson &amp; Company | Needle Sticks: Complaints were received regarding Needle Stick Injuries while administering insulin to patients. (Labeling will be revised to provide clear direction on product usage) | Other | April 23, 2007, Recall letters were sent via UPS 2nd day air and/or hand delivered to all direct consignees and potential customers of the product. |  | 496200 | 496,200 units | Nationwide :Distribution centers in NJ, MO, TX, OH, NM, IL, ND, FL, KS, VA, LA, and GA. | Terminated | April 16, 2008 | 359 | 496200 | Undetermined
Z-0423-2007 | 36974 | No | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | OsSatura TCP (pure tricalcium phosphate) 3mm granules packaged in glass jars, Part Numbers: 05-6010-050 5cc, 05-6010-100 10cc, 05-6010-150 15cc &amp; 05-6010-300 30cc |  |  | August 31, 2006 | 2006 | Teknimed SA | The firm has determined an incorrect Instructions for Use (IFU) was packaged within the product box. | Other | Teknimed notified IsoTis OrthoBiologics via electronic recall letter dated August 31, 2006. IsoTis was instructed to notify their consignees of this field correction and they were asked to be responsible for identifying and making any corrections necessary at the medical user level. Teknimed supplied the correct IFU to IsoTis so that they could perform a field correction of the affected product. October 1, 2006 IsoTis OrthoBiologics sent their consignees a recall notification that included the IsoTis notification letter, a copy of the Teknimed recall letter, IsoTis Field Correction Instruction Sheet and replacement TCP IFU's for the affected lots. |  | 1681 | 1681 units | Nationwide | Terminated | April 23, 2012 | 2062 | 1681 | Undetermined
Z-0422-2007 | 36722 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | iE33 Ultrasound System (System, imaging, pulsed doppler, ultrasonic and system, imaging, pulsed echo, ultrasonic) |  |  | November 02, 2006 | 2006 | Philips Ultrasound, Inc. | Acoustic output intensity may exceed maximum specified limits and the patient contact temperature may exceed 43 degree C (when the system is used with the S12-4 transducer in CW-mode) | Other | On 11/2/06 the firm sent a letter dated October 31, 2006 to their customers. The letter explains the problem and states that new software is being released and will be available sometime in November. Customer service representatives will visit consignees to upgrade their units. The firm later determined they sent letters out to consignees that did not have affected units. Those consignees will be notified the week of 11/20/06 to disregard the previous letter. |  | 1325 | 619 domestic units, 706 international units | Worldwide | Terminated | July 14, 2009 | 985 | 1325 | Undetermined
Z-0498-2007 | 37215 | [6 empty] | Anaerotest for microbiology |  |  | August 17, 2006 | 2006 | EMD Chemicals Inc | Lot # OC553145 of Anaerotest failed to remain in specification. QC indicator shows an anaerobic environment even if oxygen is present. | Other | A Recall Letter was issued by EMD Chemicals on 8/17/2006 to discontinue product usage and for product return. |  | 8 | 8 rolls | Product was only sent to one US customer: Becton Dickinson, Cayey. Pureto Rico. | Terminated | June 06, 2008 | 659 | 40008 | Undetermined
Z-0505-2007 | 36866 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | LATERAL BIOPSY PLATE - For use with the Liberty 9000, 8-Channel Breast Coil. The Lateral Biopsy plate is sent to customers in the following configurations: 1) The plate is purchased in sets of 5 as part number E8800BG 8 Channel Breast Biopsy Lateral Grids (pack of 5 grids) or 2) the plate is purchased in pairs with the following part numbers: M3087JG, which contains Liberty 9000 8-Channel Breast Coil, 2 plates and a manual; G3087JH, which contains 2 plates and a manual; M3335LA, which contains a 3T HD 8 channel Vibrant Breast Array, 2 plates and a manual. |  |  | November 21, 2006 | 2006 | USA Instruments Incorporated | The packaging used for the Breast Biopsy Plate may exhibit small holes or tears. | Other | Customers were notifed by certified letter, dated 11/28/06. |  | 250 | 250 units | Nationwide. | Terminated | April 13, 2012 | 1970 | 250 | Undetermined
Z-0517-2007 | 37263 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | The recalled product is the Lateral biopsy plate, part #2414346. The biopsy plate is labeled in part: Lateral Biopsy Plate, part 2414346, For Use with the Liberty 9000 8-Channel Breast Coil, USA Instruments, Inc., Aurora, OH, sterile. The breast biopsy plate is used with GE coils during Magnetic Resonance breast biopsy procedures. The lateral biopsy plate is packaged and sold as a part of catalogue E8800BG, M3087JG, G3087JH, M3335LA. |  |  | January 31, 2007 | 2007 | USA Instruments Incorporated | The packaging used for the Breast Biopsy plate may exhibit small holes or tears | Other | The recalling firm notified consignees via certified letter, dated 1/31/07. |  | 2200 | Approximately 2200 units for all products being recalled | Worldwide, including USA, Canada, Mexico, Argentina, Australia, Austria, Bhutan, Brazil, Chile, China, Denmark, Finland, Germany, Great Britain, Greece, Guatemala, Honduras, Hong Kong, Indonesia, Ireland, Italy, Japan, Korea, Malaysia, Netherlands, New Zealand, Panama, Puerto Rico, Singapore, South Africa, Spain, Switzerland, Taiwan, Thailand, and Turkey. | Terminated | April 13, 2012 | 1899 | 2200 | Undetermined
Z-0429-2007 | 37009 | No | MFK | Intraocular lens. | Ophthalmic | Ophthalmic | PMA | AcrySofÂż ReSTORÂż Intraocular Lens |  |  | December 14, 2006 | 2006 | Alcon Laboratories, Inc | Intraocular lenses exposed to extreme temperatures while in storage. | Other | The recall of the fifteen re-shipped lenses was conducted by telephone calls to the nine customers that had received the lenses. These calls were made on December 14, 2006. |  | 4 | 4 units | Nationwide. | Terminated | April 09, 2008 | 482 | 15 | Undetermined
Z-0502-2007 | 37192 | No | NCX | Thermal regulating system. | Cardiovascular | Neurology | 510(k) | Alsius Cool Line Catheter, Model CL-2295A |  |  | January 01, 2007 | 2007 | Alsius Corporation | Testing by Alsius has identified that 6-10% of catheters from the three affected lots may exhibit a significant reduction in their heat exchange capacity. | Other | Customer letters were sent starting January 1, 2006 via FEDEX and DHL. Verification of receipt of the notification letter to level A will be undertaken by using tracked delivery requiring signature upon receipt . The Customer letters will notify consignees of the nature of the problem and provide instructions to return the product. |  | 77 | 77 | Worldwide including USA, Austria, Canada, France, Greece, Japan, Norway, Spain &amp; Switzerland | Terminated | June 22, 2007 | 172 | 202 | Undetermined
Z-0428-2007 | 37182 | No | GBW | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Taut Intraducer Peritoneal Catheter, 12 Fr x 10 Fr x 20.3 cm; a sterile, Rx peritoneal catheter for single use, individually packaged in a Tyvek/clear pouch; Taut Inc., 2561 Kaneville Ct., Geneva, IL 60134; the individually packaged, sterilirzed catheter was sold in two configurations: |  |  | January 19, 2007 | 2007 | Taut, Inc. | Taut, Inc. has become aware that some package seals are not intact, compromising the sterility of the catheters. | Other | Taut sent Urgent Medical Device Recall letters dated 1/19/07, via UPS over night, on the same date, informing them that the two lots of catheters may have package seals that are not intact. Customers were instructed to check their stocks for the affected lots, remove any found and return all affected product immediately. The PI-128 is also a component of the System One Comprehensive Lap CBDE Kist, catalog number 50000, which is a convenience kit. The customer is being instructed to remove the PI-128, Intraducer Peritoneal Catheter contained in the System Once Comprehensive Lap CBDE Kit in the eight kit lots listed, and return the PI-128 catheters. All other components of the kit may be used. As the PI-128, Intraducer Peritoneal Catheter, can also be ordered as a replacement component, the customer is also being instructed to check all opened System One Comprehensive lap CBDE kits, regardless of lot number, to ensure that it does not contain the impacted PI-128 lots. The customers were requested to complete and return the enclosed Acknowledgment &amp; Stock Status Form, which includes an option to obtain replacement product for the PI-128, immediately faxing it back to Linda Todd at 919-433-4986. Any questions were directed to Linda Todd at 1-800-334-9751, ext. 4951. |  | 5769 | 5,769 catheters | Worldwide, including USA, Guam, Australia, Iceland and Italy | Terminated | August 12, 2009 | 936 | 5769 | Undetermined
Z-0515-2007 | 37191 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | OEC Uroview 2800 fluoroscopic x-ray system with 3 Phase Power Distribution Box option (Wonder Box) , Model Numbers: 882082-01, 882082-02, 886765-01, GE Healthcare, Surgery, Salt Lake City, UT. |  |  | January 15, 2007 | 2007 | GE OEC Medical Systems, Inc | Uroview 2800 system with 3 Phase Power Distribution Box (Wonder Box) option can present a shock hazard to the user if the access door to the power box is not completely closed when the unit is powered on. . | Other | Consignees were notified by letter on 01/15/2007. Instructions were given to position and not move the dual monitors until inspection and correction by GE personnel. For the affected units with a 3 phase power distribution box, users were cautioned to ensure that the access door to the power box is completely closed when the unit is powered on to prevent a shock hazard. |  | 277 | 277 units | Worldwide, including USA, Canada, China, France, Germany, Hong Kong, Italy, Japan, Singapore, South Africa, Switzerland, and Venezuela. | Terminated | December 09, 2011 | 1789 | 702 | Undetermined
Z-0427-2007 | 37113 | No | LPM | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic | PMA | OÂżOptixÂż (lotrafilcon B) Soft Contact Lenses, Rx only, Ciba Vision. The product is distributed in 3 packs (3 lenses per package) and 6 packs (6 lenses per package). |  |  | January 11, 2007 | 2007 | Ciba Vision Corporation | Reduced Ion Permeability | Other | Consignees were notified by traceable mail on 01/12/2007. Consignees were advised to locate any of the affected lots and return them to CIBA Vision in accordance with the attached instructions. Returned lenses will be destroyed. |  | 11400000 | 11.4 million lenses | Worldwide, Including USA, Canada, Australia, Benelux, UK and Germany | Terminated | March 21, 2008 | 435 | 11400000 | Undetermined
Z-0509-2007 | 37180 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Candela Corporation Vbeam Aesthetica Laser System, Model/Catalog # 9914-08-0320 Vbeam Aesthetica w/ gray enclosure and gray display |  |  | January 15, 2007 | 2007 | Candela Laser Corporation | Internal component may overheat causing the possibility of risk of fire or smoke | Other | Candela Corporation notified US consignees by letter (w/return receipt) identifying the potential problem on 1/17/2007. Candela services will contact the customers and schedule upgrades. International packages were sent to each distributor via Fed-x/DHL tracking on 1/17/2007. Distributors were asked to forward individual serial number packages to their customers. Upgrade material will be forwarded to distributors upon which they will schedule upgrades with their customers. |  | 474 | 474 laser systems total (all models) | Worldwide including USA, Europe, Far East, Canada, Japan, Australia, Middle East, Mexico | Terminated | February 03, 2014 | 2576 | 474 | Undetermined
Z-0507-2007 | 37028 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | The LightSpeed VCT System is intended for head and whole body X-ray Computed Tomography applications. |  |  | December 08, 2006 | 2006 | GE Healthcare | When the manual cradle release feature on the table is used to move a patient out of the gantry, a discrepancy of up to 6mm between the numerical display on the gantry and the actual position of the table could be introduced. | Other | A Product Safety Advisory Notice, dated December 8, 2006, was issued to all customers informing them of the safety issue and production correction. Customers are advised to operate the CT cradle as intended with the system''s electronic controls. If a critical need arises to release the cradle during a radiation therapy simulation scan, they are advised to independently verify patient position prior to marking the patient for therapy to avoid the potential inaccuracy. A software update will be installed to correct this problem by a GE representative. |  | 1256 | 1,256 units | Worldwide, including USA, Puerto Rico, Albania, Argentina, Austria, Australia, Belarus, Belgium, Bulgaria, Brazil, Canada, Chile, China, Columbia, Denmark, Finland, France, Germany, Great Britain, Greece, Guam, Hong Kong, Honduras, India, Indonesia, Ireland, Israel, Italy, Japan, Korea, Kuwait, Lebanon, Malaysia , Martinique, Mexico, Morocco, Netherlands, New Zealand, Norway, Panama, Philippines, Poland, Portugal, Reunion Island, Russia, Saudi Arabia, Singapore, Slovenia, South Africa, Spain , Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, United Arab Emirates, Venezuela, Vietnam, and Yugoslavia. | Terminated | November 30, 2007 | 357 | 1278 | Undetermined
Z-0554-2007 | 37349 | No | GDC | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Cmax Surgical Table, Model No. 2182625, Manufactured by STERIS Corporation, 2720 Gunter Park East, Montgomery, AL 36109 USA, TEL: 334 277 6660, FAX: 334 271 5450 |  |  | February 06, 2007 | 2007 | Steris Corporation | A variation in suppliers welding processes, at the inner section of the hydraulic lift column, of the surgical table, may result in the table top becoming unstable, which could result in injury to patients or staff. | Other | A certified letter was sent to all affected consignees on February 6, 2007 notifying them of the issue and informing them of the recall and the replacement of the affected tables with new tables. |  | 46 | 46 tables | USA and Canada | Terminated | September 11, 2009 | 948 | 46 | Undetermined
Z-0426-2007 | 37085 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Sara Active Sling - Standard; item KK53210.0 - model KA1210 (with safety); manufactured by Medibo Medical Products NV, Hanmont, Achel, Belgium, distributed by Arjo Inc., Roselle, IL 60172 |  |  | January 09, 2007 | 2007 | Arjo, Inc. | There is the potential for the slings to come apart, potentially allowing injury to the patient. | Other | Arjo sent Urgent Device Field Correction Recall Notice letters dated 1/9/07 via FedEx to the end user accounts who received the affected slings, advising them of the potential for the slings to come apart. They were instructed to remove the affected slings from use and replace them with the slings enclosed with the letter, complete the enclosed customer response form, and return the form and the original recalled slings, using the pre-paid call tag provided. If they have any questions or require additional, they were instructed to contact the Arjo Quality Department at 800-323-1245, ext. 6140. |  | 13 | 13 slings | Nationwide | Terminated | September 02, 2008 | 602 | 95 | Undetermined
Z-0504-2007 | 36855 | No | IXR | Radiologic table. | Radiology | Radiology | 510(K) Exempt | Stille Sonesta Video Fluoroscopy Table 6210; manufactured by Stille AB, Solna, Sweden. |  |  | November 09, 2006 | 2006 | Stille-Sonesta, Inc | Bolts may back out of their joints causing the back section of the table to tilt. | Other | Notifications letters sent to consignees on 11/20/06. Users are instructed to check the bolts located under the table. The Notice includes instructions with pictures for checking the bolts and tightening them if necessary. Stille AB has developed a kit for replacing the bolt that is involved and will send a service engineer out to the site to do the replacement. |  | 66 | 66 units. | Nationwide. | Terminated | June 20, 2007 | 223 | 66 | Undetermined
Z-0500-2007 | 36925 | Yes | CAC | Autotransfusion apparatus. | Anesthesiology | Anesthesiology | 510(k) | Atrium Express Dry Seal Chest Drain with Sterile Fluid Path Package |  |  | November 08, 2006 | 2006 | Atrium Medical Corporation | Chest drainage tubing incorrectly assembled and may disconnect from patient during use | Other | Atrium Medical notiifed accounts by fax to hospitals and phone/email to Distributors on 11/08/06. Product is requested to be returned. |  | 49 | 49 cases ( 294 units) | Nationwide, including CA, and PA. | Terminated | November 27, 2007 | 384 | 49 | Undetermined
Z-0380-2007 | 36241 | No | GWQ | Electroencephalograph. | Neurology | Neurology | 510(k) | NicoletOne version 5.20 software (also known as NicoletOne Version 5.2 Software) used with the NicoletOne LTM, Sleep, nEEG, vEEG and ICU monitoring Systems. |  |  | August 25, 2006 | 2006 | Nicolet Biomedical Div of Viasys Healthcare | A software anomaly exists in the 5.20.1038 NicoletOne LTM, Sleep, nEEG, vEEG and ICU Monitor systems using the M, and/or C series amplifiers. If this anomaly occurs, the system will display the data for channel one in all channels on NicoletOne LTM, Sleep, nEEG, vEEG or ICU Monitor systems. | Other | An Advisory Notice, dated August 25, 2006, was sent to all affected customers. The notice described the issues, states to discontinue use of 5.20.1038 software until updated 5.21.1039c software is installed. |  | 75 | 75 | Worldwide, including USA, Belgium, Chile, Denmark, Italy, Japan, Poland, and Spain. | Terminated | December 10, 2011 | 1933 | 75 | Undetermined
Z-0477-2007 | 36932 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | One Touch Ultra Blood Glucose Test Strips-Lot # 2691191, LifeScan |  |  | October 09, 2006 | 2006 | Matrix Distributors, Inc. | Counterfeit Test Strips (manufacturer unknown) | Other | Matrix Distributors contacted by telephone those pharmacies to which they had shipped the affected lots. These call were conducted from 10/9 to 10/13/2006. A follow-up letter was issued to the pharmacies extending the recall to the user level. |  | 800 | 800 units | Product was distributed to 96 pharmacies throughout the United States. | Terminated | January 07, 2008 | 455 | 800 | Undetermined
Z-0437-2007 | 36492 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1; 6 inch diameter Roller Pump; Model 801041. |  |  | October 16, 2006 | 2006 | Terumo Cardiovascular Systems Corp | The large roller pump may experience false "overspeed' error message when starting rotation, which will prevent the pump from operating; until the condition is cleared by power cycling the roller pump. | Other | Consignees were notified of the problem via letter dated 11/10/06 and informed that they would be contacted to schedule a service call to install software upgrades. Installation of software upgrades began at customer locations on 10/16/06. |  | 1951 | 1,951 units | Worldwide distribution ---- including USA and countries of Australia, Belgium, Canada, China, Dominican Republic, Egypt, Germany, India, Indonesia, Japan, Malaysia, Mexico, Pakistan, Republic of China, Singapore, South Africa, South Korea, Thailand and United Arab Emirates. | Terminated | July 07, 2009 | 995 | 1951 | Undetermined
Z-0494-2007 | 36498 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1; Air Bubble Detection Module; Catalog number 802110. |  |  | October 16, 2006 | 2006 | Terumo Cardiovascular Systems Corp | The air bubble detection system may falsely continue to sound the air alarm after the user has disabled it. | Other | Consignees were notified of the problem via letter dated 11/10/06 and informed that they would be contacted to schedule a service call to install software upgrades. Installation of software upgrades began at customer locations on 10/16/06. |  | 806 | 806 | Worldwide distribution --- including USA and countries of Australia, Belgium, Canada, China, Dominican Republic, Egypt, Germany, India, Indonesia, Japan, Malaysia, Mexico, Pakistan, Republic of China, Singapore, South Africa, South Korea, Thailand and United Arab Emirates. | Terminated | July 07, 2009 | 995 | 806 | Undetermined
Z-0493-2007 | 36496 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1; CAPIOX Control Module; Catalog number 811113 (Not distributed within the United States). |  |  | October 16, 2006 | 2006 | Terumo Cardiovascular Systems Corp | The pump may lose communication with the central control monitor, causing the monitor to display a service message "System Computer Needs Service" and to become non-operational. | Other | Consignees were notified of the problem via letter dated 11/10/06 and informed that they would be contacted to schedule a service call to install software upgrades. Installation of software upgrades began at customer locations on 10/16/06. |  | 50 | 50 | Worldwide distribution --- including USA and countries of Australia, Belgium, Canada, China, Dominican Republic, Egypt, Germany, India, Indonesia, Japan, Malaysia, Mexico, Pakistan, Republic of China, Singapore, South Africa, South Korea, Thailand and United Arab Emirates. | Terminated | July 07, 2009 | 995 | 3743 | Undetermined
Z-0438-2007 | 36494 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1, 4 inch diameter Roller Pump; Model 801040. |  |  | October 16, 2006 | 2006 | Terumo Cardiovascular Systems Corp | The roller pump will pause in response to a pressure alarm condition, as designed, but then may not resume operation as expected after the alarm condition has been cleared. | Other | Consignees were notified of the problem via letter dated 11/10/06 and informed that they would be contacted to schedule a service call to install software upgrades. Installation of software upgrades began at customer locations on 10/16/06. |  | 1399 | 1,399 | Worldwide distribution ---- including USA and countries of Australia, Belgium, Canada, China, Dominican Republic, Egypt, Germany, India, Indonesia, Japan, Malaysia, Mexico, Pakistan, Republic of China, Singapore, South Africa, South Korea, Thailand and United Arab Emirates. | Terminated | July 07, 2009 | 995 | 3350 | Undetermined
Z-0436-2007 | 36491 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1; Central Control Monitor; Catalog number 802100. |  |  | October 16, 2006 | 2006 | Terumo Cardiovascular Systems Corp | Pump may stop without an audible tone as a result of quickly clearing the internal pump alarm condition or overpressure alert/alarm condition. | Other | Consignees were notified of the problem via letter dated 11/10/06 and informed that they would be contacted to schedule a service call to install software upgrades. Installation of software upgrades began at customer locations on 10/16/06. |  | 683 | 683 | Worldwide distribution ---- including USA and countries of Australia, Belgium, Canada, China, Dominican Republic, Egypt, Germany, India, Indonesia, Japan, Malaysia, Mexico, Pakistan, Republic of China, Singapore, South Africa, South Korea, Thailand and United Arab Emirates. | Terminated | July 07, 2009 | 995 | 1379 | Undetermined
Z-0432-2007 | 36490 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1; 100V -120V, 15 A (circuit breaker), 50/60 Hz (20A power source required) base; Model 801763. |  |  | October 16, 2006 | 2006 | Terumo Cardiovascular Systems Corp | A 'battery needs service' message may be displayed, or the remaining battery time may incorrectly be reported out as zero, even when the battery is fully functional. | Other | Consignees were sent a letter dated 11/10/06 informing them of the nature of the issues and that the firm would contact them to schedule a service call to install the software upgrade package. The software upgrade installation began at customers on 10/16/06, prior to issuance of the letter to consignees. |  | 489 | 489 | Worldwide distribution ---- including USA and countries Australia, Belgium, Canada, China, Dominican Republic, Egypt, Germany, India, Indonesia, Japan, Malaysia, Mexico, Pakistan, Republic of China, Singapore, South Africa, South Korea, Thailand and United Arab Emirates. | Terminated | July 07, 2009 | 995 | 696 | Undetermined
Z-0379-2007 | 37099 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Ventilator System Servo-i V3.1 |  |  | May 23, 2006 | 2006 | MAQUET Inc. | The Connection of certain nebulizers to the built in driver has lead to reduced or no nebulization effect on the Servo-i Ventilator | Other | Maquet Critical Care issued an ''Extended Warranty'' letter on 5/23/2006. On March 13, 2007, Maquet modified its recall strategy and issued a Recall Notification letter accompanied by a Safety Alert letter to all customers with Servo-i ventilators, serial numbers in the 14000-22772 range, with Servo nebulizer installed. The Safety Alert letter provides customers with information on how to identify if the unit is affected, and how to identify if the unit has shown any failures, as well as, any precautions. Firm initiated recall is on-going |  | 2436 | 2,436 units | Worldwide | Terminated | March 23, 2009 | 1035 | 2436 | Undetermined
Z-0412-2007 | 36154 | No | DSQ |  |  | Cardiovascular | PMA | Thoratec Dual Drive Console a drive component of the Thoratec Ventricular Assist Device, Model Numbers, 10025-2600-005, 10025-2601-007, 10025-2602-006, Thoratec Corporation, 6035 Stoneridge Drive, Pleasanton, CA 94588 |  |  | September 24, 2004 | 2004 | Thoratec Corp | The module power switch (at the back of console) may break while attempting to turn the Dual Drive Console OFF or ON. | Other | The firm issued on 1/25/07 an URGENT: MEDICAL DEVICE CORRECTION-notice informing its consignees about the 'Power Switch Replacement Instructions' and arrangements for their service department replacement of the switches. |  | 331 | 331 | Worldwide | Terminated | April 18, 2008 | 1302 | 331 | Undetermined
Z-0414-2007 | 37162 | No | NLM | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | AutoSuture Bladed Trocar, equipped with safety shield. Model number 179094. Available with smooth or threaded sleeve in sizes 5-15mm inner diameter and 70-100mm length. |  |  | October 19, 2006 | 2006 | Ascent Healthcare Solutions, Inc. | The tip of the trocars may separate before and while in use. In some instances the trocars have been found with exposed blades upon opening the packaging. | Process design | The recall was initiated on 10/19/2006 by direct distribution of the recall notice to direct accounts by the firm's sales force. The firm also posted the notification on their web site. All customers were instructed to examine their inventories for the product and to remove and quarantine all affected product. Instructions on returning the product to the firm were also included along with a response form. |  | 1566 | 1566 | Nationwide | Terminated | October 31, 2008 | 743 | 1566 | Undetermined
Z-0548-2007 | 37231 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Symbia S Series SPECT (Single-Photon Emission Computed Tomography) System with a Pinhole Collimator; Siemens Medical Solutions USA, Inc., Molecular Imaging, 2501 N. Barrington Road, Hoffman Estates, IL 60195-5203 USA; part number 08717741 |  |  | January 30, 2007 | 2007 | Siemens Medical Solutions USA, Inc. | Symbia systems, running on e.soft version 5.5 or earlier, could allow a fully-extended patient pallet to come into contact with the pinhole collimator, resulting in system damage. | Other | Siemens sent Customer Safety Advisory Notice/Urgent Field Correction Recall letters dated 1/30/07 to all users of Symbia systems with a pinhole collimator and e.soft software versions 5.5 or earlier. The accounts were informed of the potential for a fully-extended patient pallet to come into contact with the pinhole collimator, resulting in system damage. The customers were advised that a Siemens Customer Service Representative would be sent to their site to upgrade their system software to version syngeo MI Applications 2006A within 30 days. The upgrade will be done free of charge. |  | 16 | 16 units | Worldwide distribution --- including USA states of Alabama, California, Georgia, Indiana, Kentucky, Michigan, Ohio, Pennsylvania, Texas, Virginia and countries of Australia, Bolivia, Canada, China, France, Germany, Hong Kong and Spain. | Terminated | November 30, 2007 | 304 | 27 | Undetermined
Z-0536-2007 | 36864 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Microalbumin Multi-Calibrator Set, in vitro diagnostic. |  |  | September 06, 2006 | 2006 | Diagnostic Chemicals, Ltd. | Microalbumin Multi Calibrator mislabeled label on Level 5 and Level 6, cap label correct | Other | Diagnostic Chemicals Ltd, Canada notified customers by telephone and fax on 9/06/06 and advised to use the cap label to determine concentration of the calibrators. Accounts were asked to confirm and acknowledge receipt of the letter. |  | 38 | 38 kits | Nationwide. | Terminated | May 14, 2007 | 250 | 38 | Undetermined
Z-0617-2007 | 36962 | No | KFM | Nonroller-type blood pump. | Cardiovascular | Cardiovascular | 510(k) | Manual Drive Unit for Sarns (Terumo) Centrifugal Perfusion System; Catalog No. 164268. |  |  | December 12, 2006 | 2006 | Terumo Cardiovascular Systems Corp | Internal friction may cause the centrifugal manual drive to be inoperable when a disposable pump is attached. | Other | Consignees will be notified via letter dated 12/12/06, which provides instructions for identifying defective manual drives and to contact the firm if their unit is defective. |  | 2696 | 2696 | Woldwide, including USA, Canada, and Mexico | Terminated | June 19, 2008 | 555 | 2696 | Undetermined
Z-0566-2007 | 37287 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Coulter LH700 Series Hematology Analyzers, Part Number 6605632 with 2B Software (2B3, 2B4 and 2B5) |  |  | September 12, 2006 | 2006 | Beckman Coulter Inc | There is a potential for erroneous yet credible results for Body Fluids on the LH700 series hematology analyzers when a cassette label fails to read and the subsequent sample is cycled in Body Fluid mode. | Other | A Product Corrective Action (PCA) letter was mailed on Sep 12, 2006 to all LH 750 accounts to inform them that there is an issue with COULTER LH 700 Series analyzer software that could affect the Body Fluid analysis and that this this anomaly is present only in Software revisions of 2B. The letter instructs the customer on how to determine their software version.Immediate user preventive action was provided in the letter which is to reset the analyzer when a cassette label read error occurs. The letter advises that this information be shared with laboratory staff and be retained as part of their Quality System documentation. A customer response form was also included. |  | 1987 | 1893 units in the US; 94 units to Canada | Nationwide and Canada | Terminated | April 19, 2012 | 2046 | 1987 | Undetermined
Z-0616-2007 | 37479 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | The iSight PACS (picture archiving and communications system) a software package used with general purpose computer hardware to acquire, store, distribute, process, and display images/associated data throughout a clinical environment. |  |  | February 14, 2007 | 2007 | Stentor Inc A Phillips Medical Systems Co. | A defect may cause patient image orientation markers to be incorrectly labeled on Multi-Planar Reformation (MPR) images. Left and Right Anatomical orientation markers may be reversed. | Other | The firm has issued a notification letter to its consignees on 02/14/2007. This notification letter contains preventive measures. |  | 54 | 54 units | Worldwide Distribution ---- USA and countries of Australia, Netherlands, and Canada. | Terminated | December 03, 2007 | 292 | 54 | Undetermined
Z-0595-2007 | 37363 | No | MVU | Analyte specific reagents. | Pathology | Hematology | 510(K) Exempt | Vysis LSI ATM/p53: D13S319 /CEP 12/13q34 DNA Probe Set; fluorescence in situ hybridization (FISH) analyte specific reagents, in vitro diagnostic; Abbott Molecular Inc., Des Plaines, IL 60018 USA; list 05J83-001; |  |  | December 22, 2006 | 2006 | Abbott Molecular | The LSI D13S319 SO/13q34 SA/CEP 12 SG probe of the two mixture DNA-Probe was contaminated with LSI ATM SG/p53 SO probe. This contamination would result in 10 FISH signals in each normal cell rather than the expected 6 FISH signals in each normal cell. | Other | Abbott Molecular telephoned and/or sent recall letters to the Abbott FISH accounts on 12/22/06, advising them that an unexpected signal pattern with the Vysis LSI ATM/p53: D13S319 /CEP 12/13q34 Probe was observed, confirmed and identified as a result of contamination in a subset of vials of D13S319 /CEP 12/13q34 from lot number 79500. This contamination results in 10 FISH signals in each normal cell (4 SO, 4 SG, 2 SA) rather than the expected 6 FISH signals in each normal cell (2 SO, 2 SG, 2 SA). The customers were advised that lot 80178 was also being included in the recall, and were requested to review the information carefully and assess the risk to their patient population; discontinue and destroy any unused product immediately; complete and return the enclosed reply form; review the information with laboratory personnel; and communicate the information to health care providers served in cases where the affected product were used and reported. Abbott Molecular provided replacement product for that product destroyed, upon receipt of the reply form. Any questions were directed to Abbott Molecular Technical Service at 800-553-7042, option 2. |  | 109 | 109 probe sets | Worldwide including USA, Austria, Canada, Czech Republic, Denmark, Finland, France, Germany, Italy, Netherlands, Norway, Spain, and the United Kingdom. | Terminated | April 17, 2008 | 482 | 109 | Undetermined
Z-0594-2007 | 37356 | No | GHH | Fibrinogen/fibrin degradation products assay. | Hematology | Hematology | 510(k) | Clearview Simplify D-dimer, in vitro diagnostic for qualitative detection of D-dimer in human whole blood and plasma. |  |  | February 15, 2007 | 2007 | Inverness Medical Professional Diagnostics | Lot PX018A is exhibiting a decrease in sensitivity affecting the qualitative result around the cut off. | Other | Recall notification letters were sent to all customers on 2/14/07 by first class mail. |  | 382 | 382 kits | Worldwide, including USA, Puerto Rico, and Canada. | Terminated | May 08, 2007 | 82 | 382 | Undetermined
Z-0604-2007 | 37068 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Wallstent Endoscopic Biliary Endoprosthesis- 8mm x 40mm, UPN H965430100 |  |  | December 20, 2006 | 2006 | Boston Scientific Corporation | Guidewire restriction during use (excessive adhesive in the guidewire lumen) | Other | On December 20, 2006, Boston Scientific notified all US consignees by letter sent via Fed Ex Priority mail. Customers are instructed to remove affected product from inventory and return it to Boston Scientific. The notification also requested the completion of a Reply Verification Tracking Form as a method of documenting the presence or absence of affected product in inventory. Boston Scientific International Regulatory Affairs was notified and is responsible the monitoring foreign recall. |  | 8511 | 8,511 for Recalls Z-0604-0609-2007 | Worldwide distribution --- USA and countries of Austria, Belgium, France, Czech Republic, Germany, great Britain, Greece,Ireland, Italy, Luxemburg, Netherlands, Norway, Poland, Portugal, Spain, Switzerland, and Sweden. | Terminated | December 12, 2007 | 357 | 8511 | Undetermined
Z-0575-2007 | 37270 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | StarClose Vascular Closure System (Clip Applier &amp; Exchange System), |  |  | January 31, 2007 | 2007 | Abbott Vascular, Inc. | Premature release of the vessel locator wings, which stabilize the device prior to clip deployment, will result in no hemostasis. | Other | On 1/31/07, the firm initiated the recall and its notification was via letters explaining the reason for the recall and requesting the product be returned to its sales rep. |  | 97952 | 97,952 units | Worldwide distribution, including USA, Puerto Rico, and 410 other countries. There were no shipments to Canada and Mexico. | Terminated | February 26, 2008 | 391 | 97952 | Undetermined
Z-0597-2007 | 37393 | No | MCW | Intraluminal artery stripper. | Cardiovascular | Cardiovascular | 510(k) | ThromCat Thrombectomy Catheter System. 7F 2.5 -7.0 mm 150 cm. Catalogue number 63000-01. It is a single-use, disposable device that is used to perform percutaneous maceration and removal of thrombus and restoration of blood flow. |  |  | February 07, 2007 | 2007 | Kensey Nash Corp | Face seal may wear excessively and cause particulate matter. | Other | On 2/7/07, a voice mail message was sent to all KNC Sales Representatives instructing them to contact affected accounts immediately. On 2/9/07, recall notices were faxed to all affected accounts that had affected product. The recall notices instructed hospitals to quarantine the affected lots and to contact their KNC Sales Representative to arrange for retrieval of the product. |  | 165 | 165 units | Nationwide. | Terminated | June 12, 2007 | 125 | 165 | Undetermined
Z-0578-2007 | 37389 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | AXIOM Artis MP, Angiographic x-ray system, model number 5904466 |  |  | January 31, 2007 | 2007 | Siemens Medical Solutions USA, Inc | C-arm gantry could rapidly descend without command during use. | Other | In response to this issue, a Siemens Service Representative will visit the affected consignees beginning 02/01/2007, to check the spindle and adjust the spindle, if necessary. |  | 62 | 62 units | The products were shipped to medical facilities nationwide. | Terminated | May 05, 2009 | 825 | 70 | Undetermined
Z-0596-2007 | 37369 | No | KDC | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | JariTrak Table Clamp, 1 inch diameter, vertical post, Catalog Number 206-160. The JariTrak Table Clamp with vertical post (206-160) is part of the Table Mounted Hardware that is used with the JariTrak&#153; Retractor System. The JariTrak&#153; self-retaining retractor system consists of the Table Clamp with Vertical Post, Flexible or Horizontal Bar with Universal Joint, Oval or Segment Rings, Tilting Blade Clamps and various blades. The product is supplied non-sterile and may be packaged individually or as part of the JariTrak&#153; retractor set. |  |  | November 14, 2006 | 2006 | J. Jamner Surgical Instruments, Inc. | The JariTrak Table Clamp may not securely fasten to the bed rail of the OR table during set up of the retractor system. In the event that the clamp is not securely fastened to the bed rail, the system may move during surgery. The firm received several complaints from hospitals. | Other | On 11/14/06, an "URGENT-PRODUCT NOTIFICATION" letter (dated 11/14/06) and accompanying Fax Back Form were sent via Certified Mail Return Receipt Requested, to each customer that purchased or was supplied with the 206-160 Table Clamp. The letter instructs each surgical instrument coordinator (user) to review the inventory of 206-160 JariTrak" Table Clamps at the facility. The letter also provides instructions for the proper installation of the clamp to the bed rail. The customer is also asked to complete the Fax Back Form and fax it to back to Customer Service. J. Jamner Surgical Instruments, Inc. (J. Jamner) will replace any Table Clamps at the request of the customer. J. Jamner will account for all Fax Back replies and product returns. A second notification letter will be sent to any non-responding customers. Table Clamps returned from customers and from J. Jamner's inventory will be returned to the manufacturer, Koros USA Inc. or be destroyed by J. Jamner. |  | 30 | Domestic: 30 units | Nationwide | Terminated | January 10, 2008 | 422 | 30 | Undetermined
Z-0577-2007 | 37314 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | EKOS EndoWave Infusion System and EKOS Lysus Infusion Systems - drug delivery catheters. EKOS Corporation Bothell, WA |  |  | February 09, 2007 | 2007 | Ekos Corp | Ekos Corporation received three reports that distal radiopaque marker bands on the EndoWave Drug Delivery Catheter became detached from the catheter during use and remained in the patient. | Other | On 2/9/07 the firm faxed and sent via certified return receipt an URGENT: Device Recall letter to their consignees. The letters requested to immediately examine stock and segregate all recalled products. In addition, the firm''s representatives will contact, via phone or visit, each of the consignees. Product is to be returned to the firm. |  | 446 | 446 distributed | Nationwide. | Terminated | May 07, 2008 | 453 | 446 | Undetermined
Z-0583-2007 | 37079 | No | DWE | Cardiopulmonary bypass pump tubing. | Cardiovascular | Cardiovascular | 510(k) | Medtronic 2T10R1 Intersept Custom Tubing Pack. Pack 2T10R1 consists of two PVC lines; one of them is 132 inch in length and is capped. The second line is 102 inch in length and has a one way vacuum relief valve assembled in one end. Tubing is used in the vent roller heads of the heart lung machine. This tubing is connected to the cannula in the patient in order to divert blood to the circuit when the patient is on bypass. |  |  | January 02, 2007 | 2007 | Medtronic Perfusion Systems | One lot of Custom Perfusion Sets, Catalog 2T10R1 was incorrectly assembled. The one way vacuum relief valve on the 102 inch line is attached backwards, thus preventing flow. | Other | There was only one customer. The customer was sent a letter on December 29, 2006, requesting that the impacted product be disposed of or returned to Medtronic for disposal. |  | 17 | 17 | Illinois | Terminated | February 24, 2008 | 418 | 17 | Undetermined
Z-0413-2007 | 37070 | No | JJI | Enzyme analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ARCHITECT i System Assay CD-ROM, in vitro diagnostic analyzer, U.S. Version 21.0; for use on the ARCHITECT i System; Abbott Laboratories, Abbott Park, IL 60064; list number 06E58-21 |  |  | January 02, 2007 | 2007 | Abbott Laboratories MPG | The ARCHITECT AUSAB and ARCHITECT Anti-HCV assay parameter default interpretation screens when using ARCHITECT Assay CD-ROM US Version 06E58-21 does not align with the result interpretation options in the ARCHITECT AUSAB Reagent Package Insert (PI) 34-4162/R1 and ARCHITECT Anti-HCV PI 34-4152/R1. | Other | On 1/2/07, Abbott sent Product Correction, Immediate Action Required letters dated 12/29/06 to all ARCHITECT AUSAB (list number 1L82-25) and ARCHITECT Anti-HCV customers (list number 1L79-25) who have the affected assay CD-ROM. The end users were informed that for the ARCHITECT AUSAB and ARCHITECT Anti-HCV assays, the default interpretations in the Assay Parameters window, installed from ARCHITECT I System Assay CD-ROM US Version 06E58-21, are not align with the result interpretation options in the package inserts. The accounts were advised to reconfigure the names of the interpretations in Assay Parameters window, following the steps in the enclosed attachments A and B. The accounts were requested to keep a copy of the letter on file until the revised Assay CD-ROM is received. The accounts were requested to complete the customer reply form, indicating that the letter was received, understood and disseminated to the laboratory manager/supervisor responsible for Architect AUSAB and/or Anti-HCV testing, and fax it to Abbott by 1/12/07. |  | 51 | 51 affected analyzers | Nationwide | Terminated | December 09, 2009 | 1072 | 51 | Undetermined
Z-0576-2007 | 37298 | No | GZX | Microsurgical instrument. | Neurology | Neurology | 510(K) Exempt | AccuGuide Injection Needles (Surface electrodes) |  |  | January 22, 2007 | 2007 | Medtronic Xomed, Inc. | A labeling discrepancy was internally identified for the surface electrode component of REF 8263210 AccuGuide EMG Injection Needle, 30G X 25MM. The outside box label correctly indicates 'Content: 5 sterile Injection Needles and 10 non-sterile Surface electrodes'; however, the pouches for the 10 surface electrodes inside the box are incorrectly labeled as 'sterile'. | Other | Each consignee contacted via telephone 01/22/2007, and then sent a follow-up fax letter 02/02/2007, which explained the issue, and the options for dispositioning any product that they still had in stock.Since firm has determined that the risk of keeping and using these devices is negligible to no risk, as they are intended to be and typically used in a non-sterile manner, the firm is offering the consignees the option to keep their devices or have them replaced. |  | 48 | 48 boxes | Nationwide | Terminated | July 07, 2012 | 1993 | 48 | Undetermined
Z-0567-2007 | 37250 | No | KWY | Hip joint femoral (hemi-hip) metal/polymer cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Wright Medical Technology, GladiatorÂż Bipolar Hip Prosthesis, REF: GLBP-2241, Size 41mm O.D., use with 22 mm Head |  |  | January 03, 2007 | 2007 | Wright Medical Technology Inc | Product specification failure which could result in improper locking of the cup onto the femoral head. | Other | The firm initiated the recall by telephoning its distributors on January 3, 2007 to request the return of the products. They followed with letters (mailed 1/15/07) to hospitals, implanting surgeons, and distributors requesting return of the product and informing them of the revision surgery that was a result of the problem. |  | 10 | 10 units | Nationwide and Japan | Terminated | January 15, 2009 | 743 | 90 | Undetermined
Z-0574-2007 | 37170 | No | DSQ |  |  | Cardiovascular | PMA | AB5000 Console Circulatory Support System |  |  | January 08, 2007 | 2007 | Abiomed, Inc. | Unit may alarm ''Low Flow'' and ''Low Pressure'' due to manufacturing material in the pressure source. | Other | Abiomed notified customers by letter dated 1/08/07. Service reps will inspect units and replace the defective components. |  | 55 | 55 units | Worldwide, including USA, Japan, Germany, Australia, and Mexico. | Terminated | February 07, 2008 | 395 | 55 | Undetermined
Z-0582-2007 | 37248 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | BD Integra 1ml Insulin Syringe with Retracting Precision Guide Needle |  |  | January 27, 2007 | 2007 | Becton Dickinson &amp; Company | The needle becomes detached from the hub/syringe | Other | Recall notification letters were sent via UPS second day air on January 27, 2007. Distributors were asked to provide end user contact information to recalling firm. |  | 247824 | 247,824 units | Nationwide | Terminated | July 24, 2007 | 178 | 247824 | Undetermined
Z-0522-2007 | 37005 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Polyethylene tibial posterior stabilized components. Optetrak All Poly Tibial PS Component. Knee prosthesis. Catalog number : 204-11-13 |  |  | December 07, 2006 | 2006 | Exactech, Inc. | The outside profile on this manufactured lot of devices is oversized and does not meet specifications. | Other | Letter dated 12/08/06. The firm representative retrieved all four devices from the field. |  | 4 | 4 units | Distributed to one hospital in New York. | Terminated | July 11, 2012 | 2043 | 4 | Undetermined
Z-0547-2007 | 37307 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | One Touch Basic/Profile and One Touch Ultra-Blood Glucose Test strips- 50-count |  |  | February 02, 2007 | 2007 | Safety &amp; Supplies Co. Inc. | Counterfeit product (manufacturer unknown) | Other | A recall notification was sent on February 1, 2007 via first class mail to all customers. |  |  | Unknown | Nationwide to pharmacies. | Terminated | June 08, 2007 | 126 |  | Undetermined
Z-0495-2007 | 36980 | No | HQE | Vitreous aspiration and cutting instrument. | Ophthalmic | Ophthalmic | 510(k) | Bausch &amp; Lomb Lightning High Speed Vitrectomy Cutter, Catalog #CX4804, 1/20-gauge high speed vitrectomy cutter per package. The firm name on the label is Bausch &amp; Lomb incorporated, Rochester, NY. |  |  | December 14, 2006 | 2006 | Bausch &amp; Lomb Inc | The cutter tip could break during a procedure | Other | Recall letters were issued via e-mail to surgeons on 12/14/06 and via overnight mail on 12/14/06 to surgeons and operating room supervisors requesting immediate segregation and return of the product. |  | 8289 | Cat. CX4804 - 3,485 Cutter Tips | Worldwide distribution ---- including USA and countries of Canada, China, Japan, Australia, Hong Kong, India, Malaysia, New Zealand, Singapore, Korea, Taiwan, Thailand, Argentina, Brazil, Mexico, and The Netherlands. | Terminated | April 07, 2009 | 845 | 45288 | Undetermined
Z-0565-2007 | 37234 | No | GIF | Blood cell diluent. | Hematology | Hematology | 510(K) Exempt | CELL-DYN Diluent/Sheath Reagent, for use with CELL-DYN Ruby, Sapphire and 4000 systems. List #01H73-01. Packaged in 20 Liter cubitainers. |  |  | December 22, 2006 | 2006 | Abbott Laboratories | Results for tests of RBC, MCV, RDW, MPV and PLT may be falsely elevated or falsely decreased when the Cell Dyn Diluent/Sheath reagent is used with Cell-Dyn 4000, Ruby and Sapphire systems. | Other | Recall initated 12/22/2006. Product letters were sent to consignees, as well as a customer reply form. The letters were sent with customer reply forms. Responses will be tracked via the response forms included with letters. |  | 652 | 652 (20 L. each) cubitainers | Nationwide, including California and Nevada. An Abbott Laboratories facility in Hong Kong. | Terminated | November 28, 2007 | 341 | 652 | Undetermined
Z-0545-2007 | 37160 | No | CDD | Vitamin B[bdi1][bdi2] test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Sample Management System software for in vitro diagnostic testing |  |  | January 15, 2007 | 2007 | Siemens Medical Solutions Diagnostics | Under limited circumstances, an incorrect patient result could be printed on the optional chartable patient report. | Other | A recall communication was sent on 1/15/2007 to all customers by fax. |  | 87 | 87 units | Worldwide, including USA, Australia, Brazil, Canada, Chile, China, Columbia, Costa Rica, Cyprus, Czech Republic, Denmark, Ecuador, Finland, France, Germany, Greece, Ireland, Israel, Italy, Japan, Mexico, Netherlands, Norway, Poland, Portugal, Russia, South Korea, Spain, Sweden, Switzerland, Taiwan, Turkey, UAE and United Kingdom. | Terminated | January 18, 2008 | 368 | 87 | Undetermined
Z-0560-2007 | 37153 | No | HNN | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | Alcon Custom-PakÂż, part #10975-02, containing BD BeaverÂż Mini-Blade (part #BD6400 or part #BD6900); individual blades repackaged into custom ophthalmic surgical packs and resterilized by Alcon Laboratories, Inc., Houston, TX. |  |  | December 22, 2006 | 2006 | Alcon Laboratories, Inc | Individually packaged surgical blades, labeled as sterile, found with incomplete seals; blades purchased for use in custom surgical packs. | Other | Customers notified of recall via phone on 12/22/06, with follow-up fax and letter. A revised notification was made via phone and fax to each customer on 12/26/06. Customers instructed not to use blades and to return blades for a replacement. Recall to the user/hospital level. |  | 4 | 4 custom surgical packs. | Custom surgical packs sold to hospitals in AR, MN, NY, TX and WI. | Terminated | June 06, 2007 | 166 | 412 | Undetermined
Z-0535-2007 | 36605 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | FastCath Transseptal Catheter Introducers with Hemostasis Valve, DAIG division of St. Jude Medical, Minnetonka, MN, consists of a radiopaque sheath and dilator; Product # 406850. |  |  | October 05, 2006 | 2006 | St. Jude Medical / Daig Division | Limited number of FastCath Transseptal Catheter Introducers (REF #406850) are susceptible to the irrigation extension tube on the side port becoming disconnected when a pull force is applied. | Other | The affected customers were notified on 10/5/06 by letter. |  | 40 | 40 Units | Nationwide. | Terminated | May 25, 2007 | 232 | 40 | Undetermined
Z-0345-2007 | 36711 | No | LPN | Soft (hydrophilic) contact lens care products. | Ophthalmic | Ophthalmic | 510(k) | Krystal Klear Rinse (China only) |  |  | September 15, 2006 | 2006 | Advanced Medical Optics, Inc. | Due to a potential sterility issue with the product container. Three lots sold in Japan were found to have bacterial contamination, which compromised sterility.Because of this production-line issue at the firm's plant in China, AMO is recalling lots that were manufactured on the same production lines during the same production period. | Other | Beginning on September 15, 2006, a product notification letter was sent via facsimile or communicated via telephone with each of the customer accounts. Beginning on September 19, 2006, an AMO Japan sales representative visited each of the customer accounts obtaining COMPLETE Amino Moist Multi-Purpose Solution subject to recall. Letters dated November 21, 2006 were sent from AMO, Santa Ana, California beginning Novemer 27, 2006. |  | 3125885 | 182,678 units for all US products, 2,943,207 units worldwide | Worldwide including USA, Australia, China, Hong Kong, India, Indonesia, Japan, Korea, Malaysia, New Zealand, Philippines, Singapore, Sri Lanka, Taiwan and Thailand. | Terminated | February 17, 2010 | 1251 | 3125885 | Undetermined
Z-0553-2007 | 37289 | No | FLD | Sorbent hemoperfusion system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Clark Biocompatible Hemoperfusion Cartridge, sorbent: Heparinized polymer over acitivated carbon, 250mL, Clark Research and Development, Inc. Folsom, LA 70437 |  |  | January 19, 2007 | 2007 | Clark Research and Development Inc | Lack of sterility assurance based on lack of sterility validation for labeled sterilization directions | Other | The firm notified its consignees of the problem and the recall via letter (priority mail with delivery confirmation)dated 01/19/2007 beginning on 01/19/2007. This letter requested that the consignee cease the use of this product and contact the recalling firm for return instructions. Follow-up with non-reponders will be made by facsimile. |  | 276 | 276 | Nationwide | Terminated | July 17, 2007 | 179 | 292 | Undetermined
Z-0537-2007 | 36988 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Biosense Webster PREFACE Guiding Sheath, Multipurpose Short, for the introduction of intravascular electrophysiology catheters into any cardiac chamber. Catalog Number 301803MS |  |  | November 17, 2006 | 2006 | Biosense Webster, Inc. | Biosense Webster, Inc. has received a number of complaints regarding the radiopaque tip separating or partially detaching from the PREFACE Sheath during use in the left atrium. | Other | Initially a Medical Device Correction Advisory letter was mailed on November 20, 2006 to all domestic and international consignees recommending to customers that the Preface Sheath was not to be used with the NaviStar ThermoCool Catheter or any 8 French or larger catheter.On December 9, 2006, a Recall Customer letter and Certification form were sent out to all customers and Johnson &amp; Johnson affiliates who had purchased the product in the past three years. For customers in the US, two letters and forms were sent out -one addressed to the EP/Cardiology Lab and the second to the Risk Management Office.For the US, customers will return product directly to Biosense Webster afler receiving a return authorization number. Returned product will be segregated and labeled as quarantined product. Returned product will be used as necessary for the root cause investigation. Any product not used at the conclusion of the investigation will be destroyed.For international customers, product will be returned to the Biosense Webster warehouse in Roden, Netherlands or to the Biosense Webster US warehouse (depending on order placement) if export issues don''t present a problem. If product cannot be easily returned to Biosense Webster, the customers will be instructed to destroy the product and provide notification that the destruction was performed. |  | 4565 | 4565 | Worldwide distribution, including USA, Australia, Brazil, Canada, China, Colombia, Hong Kong, Indonesia, Israel, India, Japan, Korea, the Netherlands, New Zealand, Pakistan, Singapore, Thailand, and Taiwan. | Terminated | June 25, 2009 | 951 | 32317 | Undetermined
Z-0543-2007 | 37104 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Terumo L Series 1863 type arterial perfusion cannulae, straight open tip, 1/4'' connector, arterial cannula, 10 FR, 17.5 cm (7'') long; Catalog no. L7350. |  |  | December 07, 2006 | 2006 | Terumo Cardiovascular Systems Corp | The cannula connector is an incorrect size: It is actually 1/4" by 3/8" instead of the correct 1/4" by 1/4". | Other | Consignee was notified via telephone on 12/7/06 and returned the product. |  | 50 | 50 | Pennsylvania. | Terminated | December 05, 2007 | 363 | 60 | Undetermined
Z-0546-2007 | 37252 | No | DRG | Radiofrequency physiological signal transmitter and receiver. | Cardiovascular | Cardiovascular | 510(k) | LifeSync ECG System - Disposable leadwear, Ls-222, Ls-223 |  |  | January 10, 2007 | 2007 | GMP Companies/Lifesync Corporation | When the V-lead is separated from the rest of the lead set, the laminate may tear, exposing the dielectric layer causing a break in the silver ink. If the product is used during a defibrillation after such a de-lamination, it could burn the skin and may compromise the efficacy of the defibrillation. | Other | The recall was initiated on January 15, 2007. The firm''s representatives, hand carried, Device Recall Letter to Account Representative and Risk Manager. Distributors who received the Device Recall Letter, purged their stock of the affected devices and notified their customers.Recall Communication was hand carried to each account with 100% compliance by LifeSync Clinical/Sales Personnel. EFFECTIVENESS CHECKS of this recall will be achieved through Return Material Authorization via telephone with each affected account. |  | 272 | 272 cases | Nationwide | Terminated | July 27, 2009 | 929 | 272 | Undetermined
Z-0523-2007 | 37077 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Ultraview SL Command Module with Masimo Sp02 Option, a multiparameter patient monitor. |  |  | January 03, 2007 | 2007 | Spacelabs Healthcare, Incorporated | Potential for modules to reset, alarms returning to default settings, invasive pressure losing its labels and zero calibration and Sp02 alarm limits being frozen. This can result in delay in treatment. | Other | On 1/3/07 the firm issued a letter to their customers advising of the issues and provided guidance to the consignee to enable the Check Setup feature that will provide an audible prompt whenever the monitor resets. The letter advises that consignees will be contacted to schedule an update, once new software has been validated. |  | 2694 | 1347 devices; 1067 domestic and 280 international | Nationwide and worldwide. | Terminated | February 15, 2011 | 1504 | 2694 | Undetermined
Z-0542-2007 | 37014 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | LifeShield Latex-Free Primary I.V. Pump Set with Distal Microbore Patient Line and Gravity Flow Prevention Valve, Convertible Pin, 110 Inch with High-Pressure Filter, Orange Polyethylene-Lined/Light Resistant Tubing and Option-Lok for use with Omni-Flow Medication Management System; Hospira, Inc., Lake Forest, IL 60045 USA, Made in Costa Rica; list 12163-01 |  |  | December 20, 2006 | 2006 | Hospira Inc. | The tubing can separate from the set at the filter inlet post. | Other | Hospira sent product recall letters dated 12/20/06 to the accounts via Federal Express overnight letter on the same date. The letters informed the accounts that the tubing can separate from the set at the filter inlet post due to an inadequate solvent bond. The accounts were requested to examine their inventory for the affected lot, quarantine any found and call Hospira Customer Care at 1-877-946-7747 to arrange for return of the product for replacement. |  | 7656 | 7,656 sets | Nationwide. | Terminated | March 24, 2008 | 460 | 7656 | Undetermined
Z-0521-2007 | 37067 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker Tom Tessier Osseous Microtome: Plug for Bone Mill Systems (Parts 01-15400 and 01-15401), Stryker Leiblinger Inc., Kalamazoo, MI; Part 01-15407. |  |  | December 22, 2006 | 2006 | Stryker Leibinger USA | Metal shavings come off of the plug as a result of friction between the plug and the cutting cylinder, and may be mixed with the milled bone chips. | Other | Consignees were notified via recall letter dated 12/22/06 which instructed them to quarantine the product and to contact the firm for further instructions. |  | 26 | 26 | Nationwide. | Terminated | June 04, 2008 | 530 | 26 | Undetermined
Z-0418-2007 | 37083 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Cytomics FC 500 with CXP software, Versions 2.0 &amp; 2.1: Part #722645, Software Kit; CXP version 2.1 New User; Part #723113 Software Kit; CXP version 2.1 Upgrade; Part #629637 Software Kit; CXP version 2.0 New User; Part #629636 Software Kit; CXP version 2.0 Upgrade. The Cytomics FC 500 is a system for the qualitative and quantitative measurement of biological and physical properties of cells and other particles. |  |  | August 29, 2006 | 2006 | Beckman Coulter Inc | Two workflow scenarios associated with renaming regions may produce incorrect results. | Other | A Product Corrective Action (PCA) letter was sent on the week of August 29, 2006 to all accounts that have an FC500 or have ordered CXP or MXP software to inform them that a) incorrect statistics are generated when a polygonal region with vertical re-entrant vertices is renamed using the Region Properties dialog box. b)when a region is copied from one plot to another, and then the new region is renamed, the statistics are not updated with the new region name. (Letter was sent by US mail). Letter includes immediate user preventive action to remedy both issues. A response form is included for consignees to return to recalling firm. |  | 397 | 397 nationwide | Nationwide &amp; Canada | Terminated | April 19, 2012 | 2060 | 623 | Undetermined
Z-0411-2007 | 37066 | No | MNI | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Rocker Reducer, part of the SiLo Spinal System, REF 8880014, Rx only, Material: Stainless Steel, qty: 1 ea, non-sterile, Medtronic Sofamor Danek, 1800 Pyramid Place Memphis, TN 38132, USA, Manufactured at: Bartlett, Tennessee, USA |  |  | December 18, 2006 | 2006 | Medtronic Sofamor Danek Instrument Manufacturing | Due to a manufacturing error, the proximal end may deform causing the protrusions that engage the associated implant to break off of the instrument during use. | Other | The firm sent out the recall letter via Federal Express to direct consignees on 12/18/2006. The letter explained the problem, requested discontinuation of use and stated that sales representatives would be coordinating the retrieval and replacement of the device. |  | 58 | 58 units | Worldwide, including USA, the Netherlands, Canada, Australia, and the Bahamas. | Terminated | July 16, 2007 | 210 | 58 | Undetermined
Z-0410-2007 | 37057 | No | HNN | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | BD BeaverÂż Mini-Blade, Mini/Bulk 69, Catalog Number: 374823 |  |  | December 26, 2006 | 2006 | Becton Dickinson and Company | Product sterility may be compromised due to incomplete package seal | Other | On 12/26/2006 BD Medical, Ophthalmic Systems mailed a recall notification letter to US Distributors and US OEM''s. Customers are instructed to return any affected products to BD for a replacement. On 1/10/07, the firm expanded the recall to additional lots. Those lots are noted with an asterisk. |  | 43600 | 43,600 units | Worldwide, including USA, Asia, Australia, Canada, Mexico, Taiwan, Singapore, Korea, and Japan. | Terminated | March 07, 2016 | 3359 | 664050 | Undetermined
Z-0467-2007 | 36871 | No | MBB | Polymethylmethacrylate (PMMA) bone cement. | Orthopedic | Orthopedic | 510(k) | Simplex P with Tobramycin. Antibiotic PMMA Bone Cement. |  |  | November 07, 2006 | 2006 | Stryker Howmedica Osteonics Corp. | The expiration date on the external label is incorrect in that it does not identify the earliest expiration date of the subcomponents. The External label provides for an incorrect longer expiration date than the actual subcomponents will support. | Other | A recall notification was sent by Federal Express on 11/7/06. The firm will follow-up with telephone calls and follow-up letters as necessary. |  | 813 | 813 units | Nationwide | Terminated | February 28, 2007 | 113 | 813 | Undetermined
Z-0355-2007 | 36731 | No | MQP | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | TraXis Fixed Rasp in sizes: 7mm, 9mm, 11mm, 13mm and 15mm; for use with TraXis vertebral body replacement system |  |  | November 07, 2006 | 2006 | Abbott Spine | Tip of devices may dislodge while in use during surgery. | Other | All consignees that have the affected TraXis Fixed Rasps and TraXis Fixed Trials will be contacted via a documented phone script 11/07/2006. Affected inventory will be exchanged for re-worked instruments. |  | 66 | 66 units. | Worldwide, including USA, Spain, Mexico, United Kingdom, Germany, and Cyprus | Terminated | October 11, 2007 | 338 | 142 | Undetermined
Z-0475-2007 | 36973 | No | KWL | Hip joint femoral (hemi-hip) metallic cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Howmedica Hip Fracture Stem System |  |  | November 28, 2006 | 2006 | Stryker Howmedica Osteonics Corp. | Lack of assurance of Sterility (package integrity): Anomalies in seal width and strength that may affect the integrity of the sterile barrier. | Other | Recall notification letters and product accountability forms were sent on 11/29 via Fed Ex with return receipt. |  | 126 | 126 units | Nationwide | Terminated | May 08, 2007 | 161 | 1657 | Undetermined
Z-0470-2007 | 36956 | No | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | International Technidyne Corporation (ITC) ProTime Microcoagulation System Instrument-Catalogue Number-Protimeint &amp; Protimepro, Protimeintrf &amp; Protimeprorf, Protimeintl &amp; Protimeprol. |  |  | December 06, 2006 | 2006 | International Technidyne Corp. | Incorrect Test Results-ProTime ITC''s Microcoagulation System: Professional Instrument may report incorrect test results when more than 50 tests are in the memory and the technician uses the AUTO Send or PRINT RESULTS feature. | Other | Medical Device Correction letters were sent on 12/6/2006 by first class mail along with a replacement Operator's manual.On 3/19/07 the firm sent and additional letter to customers in order to ensure that self testers, as well as health care professionals and distributors were notified. |  | 1541 | 1,541 units | Nationwide to health care professionals and ITC distributors | Terminated | April 26, 2007 | 141 | 1541 | Undetermined
Z-0485-2007 | 37003 | No | FMT | Infant radiant warmer. | General Hospital | General Hospital | 510(k) | Fisher &amp; Paykel IW930 CosyCot Infant Radiant Warmer (with integrated bed) fitted with 9001W005J Elevator Module |  |  | October 09, 2006 | 2006 | Fisher &amp; Paykel Healthcare Ltd | This firm has received Product Complaint Reports concerning bending of the pivot plate. The elevator base of the CosyCot Infant Warmer incorporates a pivot plate which forms part of the mechanism allowing the CosyCot to raise and lower in height. A bent pivot plate generally results in the Warmer being unable to elevate properly. Other component failures may result. | Other | Fisher &amp; Paykel Healthcare commenced sending out Product Correction Notifications letter via U.S. Postal Service certified mail October 9, 2006. This Product Correction Notice explained the nature of the problem and include a reminder of the maximum recommended weight of accessories. Two FAX Response Forms were provided, Part A for initial receipt of the Product Correction Notice and Part B for confirmation of the completed upgrade. |  | 282 | 273 units in US &amp; 9 in Canada | Nationwide and Canada | Terminated | April 26, 2012 | 2026 | 282 | Undetermined
Z-0489-2007 | 37010 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | LCP Volar Distal Radius Plate Extra Articular 4H HD-RT-Long, LCP Volar Distal Radius Plate 2.4 mm, Catalog number 242.465 |  |  | December 05, 2006 | 2006 | Synthes | Hole is malformed, leading to reduced construct strength. | Other | All Sales Consultants were notified of the recall via teleconference on 12/5/06. The telephone script, part number and affected lot were communicated to the consultants during the teleconference. They were instructed to retrieve the affected product from the field. |  | 11 | 11 units | The products were shipped to medical facilities in NJ, TN, LA, and KS. The products were also shipped to sales reps in PA, SC, and CA. The product was also shipped to Canada. | Terminated | February 13, 2007 | 70 | 11 | Undetermined
Z-0488-2007 | 36889 | No | CEP | Luteinizing hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ARCHITECT LH Reagent, each kit contains bottles of Anti-Beta LH coated Microparticles and bottles of Anti-Alpha LH acridinium-labeled Conjugate; list numbers 6C25-22 (4 x 100 tests), 6C25-27(1 x 100 tests), and 6C25-30 (4 x 500 tests); Abbott Laboratories, Abbott Park, IL 60064-3500 |  |  | November 22, 2006 | 2006 | Abbott Laboratories MPG | The lots of ARCHITECT LH Reagent List 6C25-22, have the potential to generate patient specimen results that are elevated. Correlation studies have demonstrated an upward shift in slope for patient results of approximately 13% on average when compared to other currently available reagent lots and an average bias from patient specimens of 13% to 17% has been observed. | Other | Abbott initiated a field correction recall by letter dated 11/22/06 to all customers receiving the affected reagent. The accounts were informed of the elevated patient results with the affected lots of reagents. The accounts were requested to assess the inventory for the affected lots of reagents, discontinue use of the lots and destroy the remaining inventory of the lots, recording the number of kits on the customer reply form and faxing it to Abbott by 12/6/06, indicating that the letter was received and understood. Credit will be issued for the number of kits destroyed, as indicated on the reply sheet. Customers requiring additional information were directed to contact Abbott Customer Service at 1-877-422-2688. |  | 2960 | 2,960 kits | Worldwide distribution including U.S.A., Puerto Rico, Canada, Mexico, Venezuela, Colombia, Chile, Uruguay, Argentina, Brazil, Dominican Republic, Germany, Japan, Singapore, Hong Kong, Taiwan, Thailand, South Korea, Australia, New Zealand and Panama. | Terminated | December 09, 2009 | 1113 | 2960 | Undetermined
Z-0484-2007 | 36928 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Colleague Triple Channel Volumetric Infusion Pumps; Baxter Healthcare Corporation, Medication Delivery Division, Deerfield, IL 60015 U.S.A., Made in Singapore; product codes 2M8153, 2M8153R, 2M8163, 2M8163R |  |  | December 06, 2006 | 2006 | Baxter Healthcare Corp. | Baxter has identified the potential for a corruption of the memory chip in the pump to occur due to linear accelerator radiation exposure. If corruption occurs, it could result in an interruption of therapy with an audible and visual alarm notification to the user. | Other | Baxter sent the 12/6/06 Safety Alert letter to all Colleague Infusion Pump customers, to the attention of the Director of Radiation Oncology with copies to the Medical Physicist and Vice President of Nursing, via first class mail to advise them that the Colleague pump is not recommended for use in Linear Accelerator radiation therapy suites. The letters informed the users that there is a potential for a corruption of the memory chip to occur due to linear accelerator radiation exposure. If corruption occurs, it could result in an interruption of therapy with an audible and visual alarm notification to the user. The letter listed the error codes that may occur if the pump has been exposed to this type of radiation: 18:200, 18:416, 36:416, 19:113, 41:113, 314:432 and 302:435:XXXX:0006. If the pump displays any of the codes, the user was requested to submit the pump to their service facility. Failure code 199 may also occur and requires reconfiguration of the pump settings by the user's biomedical services organization. Any questions were directed to Baxter Global Technical Services at 1-800-843-7867, prompt 2. |  | 80212 | 80,212 pumps | Worldwide including USA, Puerto Rico, Aruba, Australia, Austria, Bahrain, Belgium, Bermuda, Brazil, Canada, Cayman, Chile, China, Colombia, Curacao, Denmark, Dominican Republic, El Salvador, Finland, France, Germany, Greece, Guatemala, Hong Kong, Honduras, Ireland, Israel, Italy, Jamaica, Korea, Kuwait, Lebanon, Macedonia, Malaysia, Malta, Mexico, Netherlands, New Zealand, Norway, Oman, Panama, Peru, Philippines, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Trinidad, Turkey, United Arab Emirates, United Kingdom and Venezuela. | Terminated | June 30, 2011 | 1667 | 347529 | Undetermined
Z-0482-2007 | 36802 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | BHM Reacher for Portable Ceiling Lifts; a portable ceiling lift accessory used to connect the lift to the BHM Track or KwikTrak Trolley; Manufactured by BMH Medical Inc., 2001 Tanguay Street , Magog, Quebec, Canada J1X 5Y5. The reacher comes in two models: part #700.08310 - 24'' handle length, and part #700.08320 - 36'' handle length |  |  | November 13, 2006 | 2006 | Arjo, Inc. | The carabineer may not be correctly attached to the reacher and the portable patient lift may drop. | Other | BHM sent recall letters to Arjo''s direct accounts on 11/13/06 via first class mail, informing them that fastening the carabineer to the exterior of the closed ring portion of the reacher is contrary to the manufacturer''s instructions and is an unsafe practice. The letters informed the users that Health Canada had received notification of an incident involving a reacher marketed by a competitor, and had conducted an investigation into all similar reachers available on the Canadian market. Enclosed with the letter was a copy of the BHM Reacher User''s Guide as well as a warning sticker that visually describes the safe method for connecting the carabineer to the reacher. The accounts were requested to place the warning stickers on their reachers, review the enclosed User''s Guide in detail, and ensure that all personnel using the reacher receive refresher training on its proper use and have access to the supplied User''s Guide. Once these measures have been completed, the account was requested to complete the enclosed response form and return it to BHM by fax, e-mail or mail. |  | 886 | 886 reachers | Nationwide | Terminated | January 19, 2010 | 1163 | 886 | Undetermined
Z-0465-2007 | 36849 | No | HRY | Knee joint femorotibial metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Depuy Preservation Uni-Compartmental Knee Fixed Tibial Tray, cemented, Right Medial/left Lateral, Size 5, sterile; Part number 1498-23-005. |  |  | November 15, 2006 | 2006 | Depuy Orthopaedics, Inc. | Chamfer Step Defect-The surgeon may have difficulty in assembling /seating the polyethylene insert into the tray due to a Chamfer step defect. | Other | The firm's sales force was notified of the recall via email sent 11/15/06. The sales force was instructed to deliver a recall letter to personnel in those facilities that had purchased the products, and the recall letters instruct the facilities to quarantine the product and respond to Depuy concerning product returns. The sales force was provided with a Dear Doctor letter for issuance to implanting surgeons at the sales forces' discretion. The recall was extended to Sizes 1-5 on Dec 5, 2006 with a second recall letter. |  | 558 | 558 for Recalls Z-0456-0465-2007 | Class II Recall - Worldwide Distribution ---USA and countries of Argentina and Colombia. | Terminated | April 30, 2008 | 532 | 558 | Undetermined
Z-0481-2007 | 36740 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Viasys V.I.P. Bird Sterling pediatric ventilator, Catalog Number 15654 |  |  | October 20, 2006 | 2006 | Viasys Respiratory Care, Inc. | The reason for this correction is due to the potential of a repeat ''failure to cycle'' event. | Other | On November 2, 2006, a product notification letter was mailed to all customers in the United States via overnight carrier (Federal Express). Consignees are informed in the notification letter that service calls are in process of being scheduled to update domestic consignees. On November 4, 2006, a product notification letter was mailed to all international customers via registed mail. Consignees are informed in the notification letter that replacement software, replacement instructions, and installation verification forms are included in their recall package. |  | 539 | 539 | Worldwide including USA, Canada, Estonia, Belgium, Saudi Arabia, Brazil, Thailand, United Arab Emirates, Syria, Portugal, Jordan, Italy, Latvia, Hong Kong China, Ecuador, Jamaica, Dominican Republic, New Zealand, Australia, Czech Republic, Bolivia, Austria, Greece, United Kingdom (England), Turkey, Chile, Japan, India, Mexico, Uruguay, Algeria, Vietnam, Costa Rica, Ireland, Switzerland, Bahamas, Singapore, South Africa, Philippines, Spain, France, Peru, Lebanon, Taipei, Malaysia, Netherlands, South Korea, Israel &amp; Slovak Republic | Terminated | September 03, 2008 | 684 | 2017 | Undetermined
Z-0479-2007 | 36636 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | BHM Kwiktrak Fixed Gate component of the Kwiktrak rail system used with the Maxi Sky 600 Ceiling Patient Lift; Manufactured by BMH Medical Inc., 2001 Tanguay Street , Magog, Quebec, Canada J1X 5Y5. |  |  | October 23, 2006 | 2006 | Arjo, Inc. | The Kwiktrak Gate stopper may fail and allow the ceiling lift to pass through to the end of the track and fall to the floor. | Other | Arjo Inc. USA sent recall letters to the direct accounts on 10/30/06, informing them that if the ceiling lift is moved too fast or too forcefully into a closed gate, the gate stopper can be damaged, causing its function to fail. One consequence can be that the gate may remain open, allowing the ceiling lift to pass through to the end of the track. The letters requested the customer to visually inspect their Kwiktrak Gate installation used with their Maxisky-600 Ceiling Lift to determine if there are any signs of impact or burr marks on the closed gate stopper. Illustrations of a damaged and normal gate stopper were included. The customers were requested to stop use of the Kwiktrak track system if any signs of impact or burr marks are found, and contact the Arjo Quality Department at 1-800-323-1245, ext. 6140 for special instructions. If no signs of damage were found, the customers were reminded to make sure that they manually drive the ceiling lift in a gentler manner to prevent damage to the gate stopper, making sure that all personnel using the ceiling lifts are properly trained. The accounts were informed that Arjo personnel will contact them in the next few weeks to install an improved system that will eliminate the risks of failure. |  | 3 | 3 gate systems | Michigan, Ohio and Illinois | Terminated | January 09, 2010 | 1174 | 3 | Undetermined
Z-0486-2007 | 36304 | No | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | GammaMed 12i radionuclide applicator systems, distributed by Varian Medical Systems, Palo Alto, CA |  |  | September 11, 2006 | 2006 | Varian Medical Systems | Medical device for patient treatment does not meet electrical safety standards. | Other | The recalling firm notified end users by e-mail and letter on 09/11/06 and advised that some devices did not comply with electrical safety standards. The firm stated that they plan to make modifications to improve electrical safety performance but users may continue to use the device to treat patients. The notification included a response form for users to fax back verification of receipt. |  | 31 | 31 total units | Worldwide distribution, including NC, GA, IN, IL and AZ in USA, Germany, South Korea, Switzerland, Belgium, and Brazil. | Terminated | May 18, 2007 | 249 | 32 | Undetermined
Z-0478-2007 | 36187 | No | FNG | AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Minstrel Patient Lift with SR weighing scale; a non-AC-powered patient lift; Arjo, Inc., 50 N. Gary Ave., Roselle, IL 60172; model number HMB002-US |  |  | August 31, 2006 | 2006 | Arjo, Inc. | The hanger bar assembly may detach from the lift during use due to fatigue failure of the bolt. | Other | Arjo sent an Urgent Device Field Correction Recall letter dated 8/31/06 to all affected end users to the attention of the Administrator/Risk Manager via FedEx on the same date. The letters informed the accounts of the potential for the hanger bar to detach from the scale due to fatigue failure of the pivot bolt. The accounts were advised that an Arjo Field Service Technician would contact them within five business days to schedule a visit to the account to perform an upgrade of their Minstrel lifts. Any questions were directed to Arjo Quality Department at 800-323-1245, ext. 6140. |  | 76 | 76 lifts | Nationwide | Terminated | October 16, 2008 | 777 | 76 | Undetermined
Z-0468-2007 | 36883 | component, external, limb, ankle/foot | ISH | External limb prosthetic component. | Physical Medicine | Physical Medicine | 510(K) Exempt | Re-Flex VSP and Low Profile Re-Flex VSP external prosthetic foot modules with integrated shock absorber, Model numbers RFXS0002, RFXS0003, RFLS0006 and RFL50007 |  |  | November 03, 2006 | 2006 | Ossur | Inner Tube Breakage: distributed low quality units can lead to breakage of the inner tube, resulting in an abrupt rotation of the shock module, compromising function and stability of the prosthetic foot | Other | By letter on Nov 9, 2006. An initial letter was sent to all involved customers stating the invoice number, the purchase order number and the original shipping date. The problem is explained and the detailed exchange process is outlined. Contact information is included (ref. enclosure 2.1). At the same time as the mailing, all customers are being contacted proactively by a customer service representative. |  | 138 | 138 worldwide (86 in US) | Nationwide in US plus Worldwide distribution in Canada, Austria, Switzerland, Czech Republic, Germany, Spain, France, United Kingdom, Greece, Israel, Italy, Netherlands and Norway. | Terminated | May 28, 2008 | 572 | 138 | Undetermined
Z-0447-2007 | 36649 | No | BTT | Respiratory gas humidifier. | Anesthesiology | Anesthesiology | 510(k) | REMstar Heated Humidifer Main PCA Repair Kit, Part number 1006265 |  |  | October 25, 2006 | 2006 | Respironics, Inc. | AC inlet connector failure that can result in thermal damage to the humidifier enclosure and/or ignition of materials external to the device | Other | The recalling firm issued a recall letter dated 10/25/06 to their domestic and Canadian dealers informing them of the problem and the need to contact the end users. These dealers were given two options: 1) contact the end user by forwarding a letter the recalling firm provided or 2) contacting the third party company with their distribution list. The recalling firm issued a recall letter dated 10/25/06 regarding the service repair kit and the need to contact end users that had the device repaired using this kit. The recalling firm issued a recall letter date 10/25/06 to the international customers informing them of the problem, contact the end users, and the need to destroy the recalled part. |  | 235 | 235 units | Worldwide Distribution ---- including USA and countries of Denmark, Austria, France, Canada, Netherlands, Japan, Hong Kong, Austrailia, Great Britain, Spain, Belgium, Germany, Poland, Ireland, Portugal, Saudia Arabia, Korea, Mexico, Finland, Taiwan, Philippines, Iceland, Brazil, Kuwait, New Zealand, Chile, India, Egypt, Norway, Turkey, Jordan, Thailand, Hungary, Russia, Peru, South Africia, Israel, Czech, Pakistan, Lebanon, Venezuela, Sweden, United Arab Emirates, Qatar, Slovakia, Bulgaria, Bermuda, and Colombia. | Terminated | December 15, 2009 | 1147 | 172235 | Undetermined
Z-0399-2007 | 37024 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Triathlon X3 CR Tibial Bearing Insert |  |  | December 04, 2006 | 2006 | Stryker Howmedica Osteonics Corp. | A product mix-up in which the labeling indicates a Triathlon X# CR Tibial Bearing Insert Size #2 - 11MM, however, the device inside the package is actually a Triathlon X3 CR Tibial Bearing insert Size # 7 - 11MM and vice versa. | Other | Notification letters and Product Accountability forms were sent on 12/8/2006 by Federal Express with return receipt. |  | 23 | 23 units | Nationwide and Canada. | Terminated | February 21, 2007 | 79 | 47 | Undetermined
Z-0445-2007 | 36626 | component, external, limb, ankle/foot | ISH | External limb prosthetic component. | Physical Medicine | Physical Medicine | 510(K) Exempt | SPS Blue Line Tube Clamp Adapters, Product No. BL120M |  |  | October 03, 2006 | 2006 | Southern Prosthetic Supply Co. | Device Fracturing: recent design and process changes leave the device susceptible to fracturing, which could cause the prosthesis to fail | Other | On October 6, 2006, SPS faxed a recall notification to the 187 customers who had purchased one or more BL120M Tube Clamp Adapters from SPS between August 29 and September 20, 2006. On October 9, SPS contacted those customers that had purchased five or more Adapters via telephone to confirm they had received the faxed recall notification. On October 10, SPS mailed another copy of the recall notification, via certified mail, return receipt requested to all of its customers that had purchased one or more BL120M Tube Clamp Adapters. The customers were told to send the BL120M Adapter back to SPS following replacement. |  | 943 | 943 units | Nationwide | Terminated | March 07, 2008 | 521 | 943 | Undetermined
Z-0299-2007 | 36624 | No | LGW |  |  | Neurology | PMA | Advanced Bionics Precision Linear Leads, part of the Precision Spinal Cord Stimulation System. Model Numbers: SC-2138-30, SC-2138-50, SC-2138-50T and SC-2138-70. |  |  | September 21, 2006 | 2006 | Advanced Bionics Corporation | A small number of unimplanted Precision Linear Leads may have been assembled with incorrect electrode material. At high stimulation levels, the metal may corrode and dissolved metals may enter the patient. | Other | In United States, the clinician recall notification letters were distributed by Fed-Ex or U.S. Postal Service certified mail on 09/21/2006. For the other regions, the in-country representatives started providing notification to the clinicians on 09/21/2006.The recall notification letters will instruct physicians to immediately return the unimplanted units. |  | 18 | 18 | Worldwide. | Terminated | January 25, 2012 | 1952 | 18 | Undetermined
Z-0454-2007 | 36821 | No | DLZ | Phenobarbital test system. | Toxicology | Toxicology | 510(k) | Phenobarbital- Reagent-Clinical Chemistry--List number: 1E08-20-Abbott Diagnostics Division |  |  | September 26, 2006 | 2006 | Abbott Laboratories Inc. | Erratic elevated results and or the inability to calibrate due to imprecision. | Other | Abbott sent out a product recall letter on 9/26, 2006, to all consignees that received lot numbers 39074HW00, 41008HW00 and 42063HW00. The letter states that the firm has identified a Phenobarbital performance issue that is demonstrated by erratic elevated results and/or the inability to calibrate due to imprecision. Until the cause of this issue is identified, they have discontinued shipment of Phenobarbital reagent. It will be necessary to identity an alternate means for Phenobarbital testing during this time. AXSYM and TDxITDxFLx Phenobarbital are available as alternate testing options (consignee may contact the local sales person or Customer Service). Consignees also advised if they forwarded any of the Clinical Chemistry Phenobarbital lots listed above to another laboratory to provide a copy of the letter to their customer. If there are questions, they are advised to call Customer Service (US consignee) or local sales person (foreign consignee). Letter advises if user currently using or have inventory of lot number 39074HW00, expiration 9/30/06, to run two levels of quality control (QC) every 6 hours and be alert to QC trends. For those using or have inventory of lot numbers 41008HW00 &amp; 42063HW00, they are advised to discontinue use of these lots and destroy them following laboratory procedures. Abbott will provide reimbursement for the destroyed kits. A customer reply form is provided. |  | 448 | Worldwide: Lot #41008HW00 - 162 kits; Lot #42063HW00 - 136 kits; Lot #39074HW00 - 150 Kits | Worldwide distribution ---- USA and countries of of Canada, Chile, Germany, Australia | Terminated | June 14, 2011 | 1722 | 448 | Undetermined
Z-0471-2007 | 36958 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Expo32 ADC Software Version 1.1C, Part Number 6418337-New Users Software Kit, and Part Number 175454 Upgrade Software Kit used with the Coulter EPICS XL Flow Cytometers |  |  | November 01, 2006 | 2006 | Beckman Coulter Inc | The firm has confirmed that the ratio parameter assignments (numerator and denominator) of a protocol may inadvertently change after the Parameter Selection Dialog box is accessed. This issue is limited to only protocols utilizing the Ratio parameter. | Other | A Product corrective Action (PCA) letter was mailed on 11/1/06 to Expo 32 software accounts. They are informed that the ratio parameter assignments (numerator and denominator) of a protocol may inadvertently change after the Parameter Selection Dialogue box is accessed. The issue is limited to only protocols utilizing the Ratio parameter. An immediate user preventive action to remedy this issue each time acquisition of samples that require a protocol utilizing the ratio parameter is desired is given within the letter: 1) Perform acquisition of samples in batches which use protocols containing the same ratio parameter assignments. 2) Before running samples, open the protocol intended for use and open the Parameter Selection dialog box to ensure the correct ratio parameter assignments for numerator and denominator are selected. Make any necessary changes and save the protocol. 3) Run samples. (Letter was sent by US mail). A response form is included. |  | 28 | 28 units | Nationwide including the states of MA, AL, CO, MI, IN, MD, MO NY, PA, TX, AR, CA, NJ, FL, and TN | Terminated | February 27, 2012 | 1944 | 28 | Undetermined
Z-0293-2007 | 36392 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1, 4 inch diameter Roller Pump; Model 801040. |  |  | October 06, 2005 | 2005 | Terumo Cardiovascular Systems Corp | An incorrect serial number was placed on one roller pump. | Other | The six inch roller pump labels were corrected during in-house servicing on 10/6/05 and 10/24/05. A corrected nameplate for the 4 inch roller pump with the incorrect serial number was printed and shipped to the foreign distributor on 10/11/06, and applied in the field by the international distributor on 10/18/06. |  | 1 | 1 | Worldwide-Michigan and Singapore. | Terminated | February 01, 2007 | 483 | 3 | Undetermined
Z-0466-2007 | 36868 | No | HTW | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | BOLD/INI-CLIP 2.2mm Drill Bit with AO Attachment. |  |  | November 10, 2006 | 2006 | Integra LifeSciences Corp. | Incorrect drill length-Cannulated drill bits were manufactured to incorrect specifications (shorter length) resulting in inadequate preparation of bone prior to insertion of the UNI-CLIP Staple. | Other | An Urgent Recall Notice was sent on 11/10/2006 to both hospitals/distributors and sales representatives by first class mail. |  | 107 | 107 in US. | Nationwide to distributors, hospitals and Integra Sales Reps. | Terminated | April 16, 2008 | 523 | 107 | Undetermined
Z-0448-2007 | 36706 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Bayer Advia 70 Hematology Analyzer, manufactured by Drew Scientific, Inc., Dallas, TX. |  |  | August 24, 2006 | 2006 | Drew Scientific Group PLC | Flags and alerts on the user interface will not automatically transfering to the LIS. | Other | Firm sent 'Product Notification #22' to all customers on 08/24/06. Users advised of potential risk to patient health and provided with instructions on how to eliminate risk by disabling the auto transfer feature and manually transfering data until a software patch is made available |  | 545 | 545 units. | Worldwide distribution --- including USA states of FL, MA, MN, ND, NM, NY, Puerto Rico and TX and countries of Austria, Bangladesh, Chile, China, Czech Republic, Egypt, France, Hong Kong, Italy, India, Iran, Korea, Mexico, Morocco, Phillipines, Portugal, Russia, Serbia and Montenegro, Spain, Syria, Thailand, Turkey, United Arab Emerates (Dubai), United Kingdom and Vietnam, West Indies and Yemen.- | Terminated | March 10, 2008 | 564 | 844 | Undetermined
Z-0444-2007 | 36306 | Yes | MNT | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Puritan Bennett KnightStar 330 Bi-Level Ventilator, Catalogue Numbers: Y-KS330-NA, Y-KS330-SD, DSY-KS330-NA, DSY-K5330-SD, DLY-KS330-NA &amp; DLY-KS330-SD |  |  | August 31, 2006 | 2006 | Puritan Bennett Corporation | Directions for Use-Puritan Bennett has determined that a service interval should be defined for the device. The firm has also determined that an exhalation filter should be employed when the device is in use. | Other | Domestic customers of record will be advised of this matter by letter sent via Certified US Mail, Return Receipt. The letter will provide recipients with simple instructions on how they may return their units for service, and provides scheduled dates for service for each of their devices, based upon Nellcor Puritan Bennett shipment records. |  | 7754 | 7,754 (in US) | Worldwide distribution ---- including USA and countries of Argentina, Australia, Belgium, Bolivia, Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Denmark, Ecuador, Egypt, France, Germany, Greece, Hungary, Ireland, Israel, Italy, Japan, Kuwait, Lebanon, Mexico, Netherlands, Norway, Poland, Portugal, Peru, Puerto Rico, Russia, Saudi Arabia, Scotland, Serbia &amp; Montenegro, Singapore, Slovak Republic, South Africa, Spain, Switzerland, Syria, Turkey, United Arab Emirates, United Kingdom and Venezuela. | Terminated | March 05, 2014 | 2743 | 7754 | Undetermined
Z-0338-2007 | 36934 | No | DYG | Flow-directed catheter. | Cardiovascular | Cardiovascular | 510(k) | Edwards Lifesciences-Swan-Ganz,Thermodilution Venous Infusion Port (VIP) Catheter with AMC Thromboshield (An Antimicrobial*Heparin Coating), REF:831HF75 |  |  | November 27, 2006 | 2006 | Edwards Lifesciences Technology SARL | Misbranding-One lot of Swan-Ganz Thermodilution VIP Catheters, Model 831HF75, lot 246HC197 was labeled with different model information (831F75) on the edge label . The edge label states the model is 831F75 which incorrectly indicates that product does NOT contain an antimicrobial heparin coating. (The tray lid label correctly states the model is 831HF75, containing the heparin coating) | Other | Edwards initiated the recall in the US by sending a Customer Letter by Federal Express on November 28, 2006, along with a Q&A; to their customers Service and Direct Customers to facilitate the uderstanding of the recall and the request for verification and return of all unused product. |  | 1110 | 1,110 units | Worldwide Distribution ----USA and countries of Columbia and Venezuela. | Terminated | January 23, 2012 | 1883 | 1110 | Undetermined
Z-0266-2007 | 36863 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | One Touch Basic/Profile and One Touch Ultra Blood Glucose Test Strips |  |  | October 14, 2006 | 2006 | Core Care Technologies, Inc | Counterfeit product (manufacturer unknown) | Other | The firm identified the customers who had been sent OneTouch test strips since August 2006 and contacted each by phone to identify the lot number of the test strips in the customer's possession. A letter was sent on 11/21/2006 by first class mail which identified the lot numbers identified in the letter from the law firm and providing instructions to follow if the affected lots were found. |  |  | unknown | Nationwide. | Terminated | June 06, 2007 | 235 |  | Undetermined
Z-0469-2007 | 36898 | No | LCX | Human chorionic gonadotropin (HCG) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Confirm Clearly Smart Pregnancy Test (Confirm Clearly Starter Kits, Product Code 30136 and Confirm Clearly Refills Kits, Product Code 30138) |  |  | November 17, 2006 | 2006 | Mizuho USA, Inc. | False positives-the firm has received a higher than expected rate of complaints regarding false-positive results. | Other | On November 14, 2006, Mizuho USA and SSL decided to cease further distribution of the product, and to contact SSL''s customers and ask them to return any unsold units. Mizuho USA informed SSL Americas by telephone.On November 21, 2006, SSL Americas sent Recall Notification Letters to each of its customers by US First Class, Certified, Return-receipt mail. The letter included a 'Fax Back' Notification to be use by the firm for recall effectiveness checks. |  | 474000 | 474,000 Kits | Worldwide-USA and Canada | Terminated | May 23, 2008 | 553 | 474000 | Undetermined
Z-0450-2007 | 36728 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Bayer Advia 70 Hematology Analyzer, manufactured by Drew Scientific, Inc., Dallas, TX. |  |  | September 20, 2006 | 2006 | Drew Scientific Group PLC | System has given inaccurate readings when running samples that trigger an extra rinse due to very high WBC, RBC or Platlet counts. | Other | Firm sent 'Technical Bulletin #39' to all customers on 09/20/06. Users advised of potential for inaccurate readings and provided with instructions on how to eliminate problem. |  | 545 | 545 units. | Worldwide distribution ---- including USA states of FL, MA, MN, ND, NM, NY, Puerto Rico and TX and countries of Austria, Bangladesh, Chile, China, Czech Republic, Egypt, France, Hong Kong, Italy, India, Iran, Korea, Mexico, Morocco, Phillipines, Portugal, Russia, Serbia and Montenegro, Spain, Syria, Thailand, Turkey, United Arab Emerates (Dubai), United Kingdom and Vietnam, West Indies and Yemen. | Terminated | March 10, 2008 | 537 | 844 | Undetermined
Z-0452-2007 | 36774 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | LifeScan brand OneTouch Sure Step- Blood Glucose Monitoring System Test Strips, 50 strips per package, |  |  | September 29, 2006 | 2006 | Lifescan Inc | Inaccurage test results: testing may lead to no test results, error message or inaccurately low test results. | Other | The firm has issued notification on 9/29/06 to consumers and hospital end users known to the firm and direct accounts that received the recalled product. The firm also posted the notice on the website, and customer service scripting was developed to notify callers. |  | 844920786 | 26,398 units (both recall numbers Z-0452/3-2007 | Nationwide to Al, AZ, CA, CN, DE, FL, GA, IL, IN, KT, KY, LA, MA, MD, MI, MN, MS, MO, NE, NJ, NM, and TX. | Terminated | November 27, 2007 | 424 | 844920786 | Undetermined
Z-0339-2007 | 36693 | Yes | JOH | Tracheostomy tube and tube cuff. | Anesthesiology | Anesthesiology | 510(k) | Bivona Adult Tracheostomy Tube, I.D. 8.5 mm, O.D 11.0 mm, Length 88 mm, TTS 11.8 mm, Mfrd by Smiths Medical Critical Care; Product code 670185. |  |  | August 31, 2006 | 2006 | Smiths Medical ASD, Inc. | Mislabeled- The outer diameter (OD) of the shaft is labeled as 11.0 m on the inner, outer and tray label in error. The OD is actually 11.8 mm. | Other | Distributors and hospital consignees were notified via recall letter dated 8/25/06 to return all unused product to the firm. Distributors were asked to notify their customers of the recall. |  | 197 | 197 | Worldwide Distribution ---- including USA states of Arkansas, California, Georgia, Illinois, Kansas, Minnesota, Missouri, Ohio, Pennsylvania, Texas and countries of Canada and Netherlands. | Terminated | May 04, 2007 | 246 | 197 | Undetermined
Z-0356-2007 | 36887 | No | CFN | Immunoglobulins A, G, M, D, and E immunological test system. | Immunology | Immunology | 510(k) | Roche IGGT COBAS INTEGRA Immunoglobulin G (Turbidimetric); Catalog no. 20766631322. in vitro diagnostic. |  |  | December 15, 2006 | 2006 | Roche Diagnostics Corp. | Falsely elevated IGGT results may be reported because the R2 reagent of the Tina-quant Gen.2 albumin assay has a carryover effect on the Immunoglobulin G (turbidimetric) assay on the COBAS INTEGRA 700 and 800 analyzers and the R1 reagent of the Tina-quant Hemoglobin A1c assay has a carryover effect on the Immunoglobulin G (turbidimetric) assay on the COBAS INTEGRA 800 analyzer. | Other | Device correction letters dated 12/15/06 instructing users to perform an extra wash cycle between applications when using both the ALBT2 and IGGT cassettes on the COBAS INTEGRA 700 and 800 analyzers. The recall was extended to via device correction letters dated 2/28/07 instructing users to perform an extra wash cycle between applications when using both the A1C2 and IGGT cassettes on the COBAS INTEGRA 800 analyzer. |  | 18905 | 18,905 | Nationwide. | Terminated | April 23, 2008 | 495 | 18905 | Undetermined
Z-0455-2007 | 36831 | No | NAE | Hydrogel wound dressing and burn dressing. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | HiZel, Trayless-sterile hydrogel wound dressing. Sizes 2.4 x 2.4'', 2.4 x 4.7'', &amp; 4.7 x 4.7''. Product is a transparent, flexible sheet. Packaged in a peelable pouch. |  |  | November 09, 2006 | 2006 | Embeem Inc. | Foreign material- this device, a sterile wound care product, was found to be contaminated with clumps of brown/black foreign material | Other | Recall letters were sent out on 11/9/2006 to the retail level |  | 1258 | 1,258 pieces | There are 47 customers which includes two large distributors/accounts and the rest are internet individual purchasers. These internet purchases amounts are 6 to 18 pieces at a time. There are no government accounts and no international accounts generating from EMBEEM. | Terminated | July 21, 2008 | 620 | 1258 | Undetermined
Z-0358-2007 | 36984 | No | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | I-Flow ON-Q PainBuster (100 ml, 2 mI/hr) Pump, PM012, Lot #662449 and 662742 |  |  | November 29, 2006 | 2006 | I-Flow Corporation | The pump flow rate labeling may not match the package labeling. The label on filter may not match the label on the top of the pump or the package labeling. | Other | A Recall Notification Letter was faxed to affected customers on November 29, 2006. Affected customers were notified of the identity of product subject to recall, the reason for recall, and instructed to quarantine immediately any remaining inventory.A Customer Response (FAX back) form was included with the Recall Notification Letter (dated November 29, 2006). Affected customers were requested to return the FAX back form to I-Flow within 48 hours of receipt. The fax back form indicates the quantity of remaining product from each lot the customer may have. |  | 3475 | 3475 | Nationwide | Terminated | December 11, 2009 | 1108 | 3475 | Undetermined
Z-0357-2007 | 36655 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Fibrin Analysis Catheter Testing System includes the component BD Vacutainer Luer-Lok Access Device; Catalogue #364902 as a finished sterile device, Non Pyrogenic in unopened undamaged package, Disposable/ Single use/ Do Not Reuse; for use with multi lumen; (4 Venous Catheter sizes) 20 cm. 12 fr. Central; 16 cm. 7-7.5-8 fr. Central; 20 cm 7.7.5-8 fr Central; 16cm 12 fr. Central. Manufactured in the USA for: CRS Medical Diagnostics, Inc., Pewaukee, WI. |  |  | October 18, 2006 | 2006 | Continental Medical Labs,Inc | Fibrin Analysis Catheter Testying System contains the ''Luer Lok Access Device'' that fails to lock securely to certain catheter devices because the male luer taper surface is not within the specification of the taper. | Process control | Consignee was notified by certified mail on 10/19/06. |  | 840 | 840 kits | Product was distributed to one consignee located in Wisconsin | Terminated | January 18, 2007 | 92 | 840 | Undetermined
Z-0396-2007 | 37139 | No | MQU | Eye sphere implant. | Ophthalmic | Ophthalmic | 510(k) | Perry-Kolberg Titanium Motility/Support System (Threaded Sleeve and Flat Peg), Model Number 100045 |  |  | December 20, 2006 | 2006 | Integrated Orbital Implants Inc | The threads on the threaded sleeve were manufactured with left-handed threads, requiring a modification to the surgical technique for proper insertion. Failure to properly identify the affected components and apply the modified insertion technique will result in an inability to insert the sleeve into the implant. | Other | All domestic customers were notified by phone call on 12//20/2006 which was followed up with a fax and the letter sent registered mail. Internationally, customers were notified by fax, email and mail.Customers were instructed to check their inventory of the product by lot number and visual inspection. They were instructed to isolate the affected inventory. They were provided with alternate instructions on how to use the product as it is, if needed. They were provided with a reply card to complete and fax back to the company. They were instructed that when replacement product becomes available, they will be contacted to arrange for the exchange of the affected product. |  | 234 | 234 | Worldwide, including USA, United Arab Emirates, Canada, Mexico, Costa Rica, Taiwan &amp; Romania. | Terminated | July 12, 2007 | 204 | 234 | Undetermined
Z-0417-2007 | 37065 | No | DDC | Thyroid autoantibody immunological test system. | Immunology | Immunology | 510(k) | AccessÂż Immunoassay Systems Thyroglobulin Antibody Reagent Kit, Part Number 33890 |  |  | October 26, 2006 | 2006 | Beckman Coulter Inc | Low end imprecision affecting both S0 calibrators and patient samples which can lead to inaccurate low level patient results when using the Access Thyroglobulin Antibody assay. | Other | A Product Corrective Action (PCA) letter was mailed on Oct 27, 2006 to all Thyroglobulin Antibody customers. They are informed that there has been an increase in customer feedback regarding low end imprecision affecting both S0 calibrators and patient samples, which can lead to inaccurate low level patient results when using the Access Thyroglobulin Antibody assay. The letter informs the customer that the firm recommends that they use an alternate method for testing TgAb. (Letter was sent by US mail). Customers are asked to complete and return the enclosed response form within 10 days. |  | 474 | 474 units | Nationwide and Canada | Terminated | April 30, 2012 | 2013 | 474 | Undetermined
Z-0388-2007 | 37001 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific Mach 1 Peripheral Guide Catheter, 6F, 55cm, LIMA Tip Shape, UPN # M001196620, Catalog # 19-662. Made in Mexico: Edificio G1 Local B, H, Parque Industrial La Mesa, Tijuana, Baja California 22650 MEX. |  |  | December 13, 2006 | 2006 | Boston Scientific | Excess strands of resin may exist in the lumen near the hub of certain batches/lots of the Mach 1 Guide Catheters. If this excess resin is present in the catheter, there is the potential of an embolization resulting from the excess resin detaching from the device during the procedure. | Other | An Urgent Medical Device Recall letter, dated 12/13/06, and a Clarification of Recall Notification letter, dated 12/14/06, were sent to all affected hospitals. The letter describes the issue, affected product, asks for customers to discontinue further distribution or use of any remaining product and asks for product to be returned. Replacement product will be available for any/all returned product. |  | 18 | 18 | Nationwide. | Terminated | June 28, 2007 | 197 | 51 | Undetermined
Z-0416-2007 | 36996 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ARCHITECT i2000SR Processing Module, for in vitro diagnostics, List Number 3M74-01; Temperature Controller Board Part #78560-109; distributed by Abbott Laboratories, Irving, TX. |  |  | December 12, 2006 | 2006 | Abbott Laboratories, Inc | The re-designed ARCHITECT Temperature Controller Board has been identified as being susceptible to electromagnetic interference (EMI) in the laboratory. This can cause a board reset condition, stopping temperature control function of the board. Situation does not stop analyzer operation and does not generate an error condition alerting operator when condition occurs. | Other | Firm contacted consignees by telephone and fax with instructions to verify proper functioning of the Temperature Controller Board starting December 13, 2006. |  | 648 | 648 units. | Worldwide, including USA, Argentina, Australia, Brazil, Canada, Chile, Columbia, Dominican Republic, Germany, Hong Kong, Jamaica, Japan, Korea, Mexico, New Zealand, Peru, Puerto Rico, Singapore, Taiwan, Thailand, Trinidad, Tobago, Uruguay, and Venezuela. | Terminated | September 24, 2007 | 286 | 705 | Undetermined
Z-0398-2007 | 36458 | No | FRO |  |  | General &amp; Plastic Surgery | 510(k) | Product is Mucotrol Concentrated Oral Gel Wafer contained in HDPE Bottle. |  |  | September 25, 2006 | 2006 | Belcher Pharmaceuticals Inc. | Processing and cleaning process validation were not completed before production. Product had some GMP failures related to the quality system. | Other | Consignees were notified of recall via letter dated 09/25/2006. |  | 1097 | 1097 bottles | Nationwide. | Terminated | July 22, 2008 | 666 | 1097 | Undetermined
Z-0397-2007 | 37143 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | PortraitÂż PSR3 System, an electro-surgical device used in dermatological applications |  |  | October 02, 2006 | 2006 | Rhytec Inc. | Inadequate Directions for Use: Software Upgrade to set maximum energy level at 2.5 Hz | Other | Rhytec notified users by letter dated 10/2/06 of the software upgrade. A service representative will perform the update. |  | 234 | 234 units | Canada, Australia, Brazil, Korea | Terminated | February 02, 2016 | 3410 | 234 | Undetermined
Z-0395-2007 | 37126 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific brand iLabÂż Ultrasound Imaging System, for intravascular use; |  |  | December 01, 2006 | 2006 | Boston Scientific Target | The product fails to comply with applicable electronic product performance standard, in that the electromagnetic energy is in excess of the labeled conformance standards for radiated emissions. | Other | A customer notification was sent December 1, 2006 indicating the need for the correction. |  | 24 | 24 units | Nationwide and Portugal. | Terminated | December 05, 2007 | 369 | 24 | Undetermined
Z-0393-2007 | 37021 | No | DTK | Cardiovascular intravascular filter. | Cardiovascular | Cardiovascular | 510(k) | G2 Filter System - Jugular/Subclavian Delivery Kit, for use in the prevention of recurrent pulmonary embolism. Catalog Number RF-320J |  |  | December 20, 2006 | 2006 | Bard Peripheral Vascular Inc | This action is being taken because BPV has recently received three complaints for the device that involve reports of introducer sheath tip damage. The damage may result in loss of tip integrity. | Other | Recall letters were sent to all consignees via Federal Express overnight delivery with signature receipt required on December 20, 2006. Each letter included a Recall and Effectiveness Check Form for each consignee to fill out and fax to Bard Peripheral Vascular, Inc (BPV). Customers will be asked to immediately remove product from inventory and isolated from use and contact BARD for return authorization. All returned product will be destroyed. |  | 5761 | 5761 units | Worldwide including USA, Australia, and Canada | Terminated | October 23, 2008 | 673 | 11770 | Undetermined
Z-0390-2007 | 37019 | No | DTL | Cardiopulmonary bypass adaptor, stopcock, manifold, or fitting. | Cardiovascular | Cardiovascular | 510(K) Exempt | ACS Rotating Hemostatic Valve, 0.096 inch, for use with a dilatation catheter - Reference Number 23242 |  |  | December 05, 2006 | 2006 | Abbott Vascular-Cardiac Therapies dba Guidant Corporation | Specific lots of the .096' and .115' RotatingHemostatic Valves (RHV) and the Duostat .096' Rotating Hemostatic Valve (RHV) potentially have an incomplete seal in the packaging pouch | Other | On December 5, 2006, the firm began notifying all affected customers, affiliates and distributors of the recall via letter. The letter advises that the firm has voluntarily recalled specific lots of the .096' and .115' Rotating Hemostatic Valves (RHV) and the Duostat .096' Rotating Hemostatic Valve (RHV) due to a potentially incomplete seal in the pouch. The letter advises Abbott Vascular will be replacing all returned product and that appropriate corrective actions will be implemented. |  | 24365766 | 6090751 - 791 devices, 6090752 - 656 devices, 6091151 - 381 devices, 6091152 - 132 devices | Worldwide, including USA, Canada, Germany, Italy, Greece, New Zealand, Netherlands, France, Japan, Spain, and Bulgaria. | Terminated | July 24, 2012 | 2058 | 85277830 | Undetermined
Z-0389-2007 | 37006 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | ev3 Protege GPS Self-Expanding Nitinol Stent Biliary System. SERB65-09-60-120. |  |  | December 06, 2006 | 2006 | Ev3, Inc | Two individual ev3 Protege GPS Biliary Stent System devices were incorrectly labeled as SERB65-09-60-120 and the actual device is SERB65-12-60-80. Length is shorter than expected. | Other | Letter 12/06/2006. An ev3 representative visited the affected hospital to determine the status of the recalled device. If the device is unused, the representative is to retrieve the device for return to ev3. The representative is to complete the Device Recall Field Action Form, along with a hospital representative signature, and fax the completed form to ev3. |  | 2 | 2 | IL, PA | Terminated | February 03, 2007 | 59 | 2 | Undetermined
Z-0381-2007 | 36963 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1 Integrated Centrifugal System Control Unit; Catalog number 801046. |  |  | November 08, 2006 | 2006 | Terumo Cardiovascular Systems Corp | Following installation of software version 1.30 in October 2006, the firm confirmed a complaint that the new software caused a discrepancy in the pump speed calculation and caused the pump to stop during clinical use and a ''connect motor'' message to be displayed. | Other | Customers were notified by phone, fax, email or by a Terumo representative of these problems between 11/08/06 and 11/13/06 and instructed to not use the unit until new software had been installed. All version 1.30 software was replaced and version 1.40 installed prior to 12/4/06. |  | 35 | 35 | Nationwide. | Terminated | June 11, 2007 | 215 | 35 | Undetermined
Z-0374-2007 | 36922 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Davol Salute Fixation System, 38cm Reusable Shaft (Laparoscopic) with Handle. To attach mesh to the abdominal wall. |  |  | December 27, 2006 | 2006 | Davol, Inc., Sub. C. R. Bard, Inc. | Inadequate directions for use in the Salute System. Labeling -Revised to provide updated instructions for use and proper servicing | Other | Davol notified users by a first letter on 12/27/06 providng a introductory letter explaining the Enhanced Servicing Program and changes made to the IFU (re: servicing, tip inspection, end of wire, and test fire protocol). A second letter will issue to accounts to identifying the serial number to be serviced and returned to Davol. |  | 4787 | 4787 (includes both types) | Worldwide, including USA, Austria, Italy, Israel, Finland, Switzerland, Portugal, Japan, France, Spain, Greece, Belgium, Germany, Sweden, Holland, Denmark, UK, and South Africa. | Open, Classified |  |  | 4787 | Undetermined
Z-0359-2007 | 36990 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Bard Composix Kugel Hernia Patch- |  |  | January 10, 2007 | 2007 | Davol, Inc., Sub. C. R. Bard, Inc. | PET recoil ring breakage-Memory recoil ring may break which could potentially lead to bowel perforation and or chronic enteric fistula | Other | Davol Inc. notified Hospital Administrators/Risk Health Managers,distributors, and surgeons of the expanded recall by letter beginning 1/10/07. Product under recall is to be returned to Davol for replacement. Davol is posting recall information to their website. |  | 32581 | 32,581 units | Nationwide | Open, Classified |  |  | 49184 | Undetermined
Z-0365-2007 | 37013 | No | KGO | Non-powdered surgeon's glove. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Sterile Neolon 2G Latex-Free Powder-Free Neoprene Surgical Gloves, Does Not Contain Natural Rubber Latex; 1 pair of gloves per package, 100 packages per case; Manufactured for Medline Industries, Inc., Mundelein, IL 60060 U.S.A., Made in Malaysia |  |  | December 18, 2006 | 2006 | Terang Nusa Sdn Bhd | Some of the packages labeled as latex-free actually contain latex gloves. | Other | Terang Nusa telephoned Medline Industries on 12/13/06 requesting them to retrieve the affected lot of size six gloves from their customers. Medline Industries notified their customers by letter dated 12/20/06, informing them of the possible latex gloves being labeled as the Neolon 2G Neoprene Surgical Gloves, and requesting them to inspect their inventories for the suspect size and lot of gloves. The accounts were instructed to complete the enclosed response questionnaire, indicating the amount of affected gloves found, and fax it back to Medline at 847-949-2643 to obtain a return goods authorization. Any questions were directed to Kathy Dunne at 847-359-1704. |  | 43 | 43 cases | Worldwide including USA, Canada, and Panama | Terminated | August 12, 2009 | 968 | 43 | Undetermined
Z-0368-2007 | 37012 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Allura FD 10 FC, X-Ray System, angiographic - a monoplane, digital Allura FD (flat detector) with a 10' Image Intensifier (II). |  |  | December 14, 2006 | 2006 | Philips Medical Systems North America Co. Phillips | Movement of the AD5 table or Lateral C-arm (LARC). The table and/or LARC will not move upon request, will move slowly when movement is requested or will not stop moving when requested. | Other | On 12/14/05 and 12/20/06 the firm issued a letter dated 12/12/06 to their customers. The letter explains the problem and states that the firm will be performing mandatory upgrades to correct the situation. Consignees will be contacted to arrange for service. |  | 140 | 140 devices total | Units were distributed to hospitals and medical centers nationwide. | Terminated | October 14, 2011 | 1765 | 140 | Undetermined
Z-0377-2007 | 37115 | No | FED | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Cook Flexor DL Dual Lumen Ureteral Access Sheath with AQ Coating, sterile, REF FUS-095013-DL. |  |  | January 05, 2007 | 2007 | Cook Urological, Inc. | Lack of assurance of sterility, as the packages may not have been sealed. | Other | Consignees were notified via recall letter dated 1/5/07 and requested to return the product. |  | 4 | 4 | Massachusetts, Maine, New Hampshire and Puerto Rico. | Terminated | May 01, 2008 | 482 | 4 | Undetermined
Z-0373-2007 | 37121 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Centricity PACS RA1000 Workstation; for diagnostic image analysis. GE Healthcare Integrated IT Solutions, Barrington, IL 60010 |  |  | January 12, 2007 | 2007 | GE Healthcare Integrated IT Solutions | The Centricity PACS RA1000 Workstation Software may manifest certain mammography image display problems only if the images were acquired through GE Senographe versions 2000D, DS or Essential. | Other | GE Healthcare IITS sent recall letters dated 1/12/07 to all customers who have the combination of GE Centricity PACS RA1000 Workstation software versions 2.1.X.X, or 3.0.X.X combined with the GE Senographe 2000D, the Senographe DS, or Senographe Essential mammography scanners, informing them that the entire breast image or a portion of the breast image may not be displayed in the monitor regions and may not be evident to the end-user. To mitigate this problem, the users were provided a Quick Reference Guide detailing how to use the Pan tool to move the image left, right, up and down in order to display any portion of the image that may not be displayed on the screen. The customers were requested to make sure that all necessary physicians and personnel at their location are made aware of the issues and understand the risk mitigation provided. A software patch to permanently resolve the issue will be installed at no charge when it becomes available. Any questions were directed to the GE Customer Care Center at 800-437-1171. |  | 198 | 198 units | Worldwide including USA, Australia, Belgium, Canada, China, Denmark, France, Germany, Hong Kong, Israel, Italy, Japan, Korea, Kuwait, Malaysia, Portugal, Saudi Arabia, Singapore, South Africa, Spain, Switzerland, Taiwan, Turkey and the United Kingdom. | Terminated | February 03, 2010 | 1118 | 198 | Undetermined
Z-0376-2007 | 37097 | No | JIT | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Bio-Rad D-10 Dual Program Calibrator Diluent, a component of the Bio-Rad D10 Dual Program diagnostic test. This is a specific part of the kit, and only the calibrator diluent is defective. It comes in two kits. |  |  | November 17, 2006 | 2006 | Bio-Rad Laboratories Inc | The Diluent in the D-10 Dual Program Calibrator reorder packs is defective, and can cause early calibration failure or elevated control and sample recovery results. | Other | The firm has issued phone calls and faxes to its US consignees starting 11/17/2006 and International notices began on 11/20/2006 explaining the reason for the recall and requesting the product be disposed. |  | 943 | 900 kits, 43 sets distributed. | Worldwide, including USA, India, Canada, France, Sweden, the UK, Hong Kong, Singapore, Germany, Brazil, Italy, China, Costa Rica and Panama. | Terminated | November 28, 2007 | 376 | 943 | Undetermined
Z-0378-2007 | 37051 | No | KAC | Ear, nose, and throat manual surgical instrument. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(K) Exempt | Bivona Laryngectomy Tube, Cuffless, Silicone, I.D. 9.5 mm, O.D 20.0 mm, Length 85 mm, Mfrd by Smiths Medical Critical Care; REF BOSL1L. |  |  | December 15, 2006 | 2006 | Smiths Medical ASD, Inc. | The label incorrectly states the outer diameter of the tube is 20.0 mm when it is actually 12.0 mm. | Other | Consignees were notified via recall letter dated 12/15/06. Distributors were requested to provide the firm with final user name and address information so that they could notify them of the recall. The firm is advising users that they can either return the product for replacement or continue to use it now that they are aware of the labeling error. |  | 117 | 117 | Worldwide, including USA, Canada, Netherlands and United Kingdom. | Terminated | December 04, 2007 | 354 | 117 | Undetermined
Z-0350-2007 | 37023 | Yes | DWC | Cardiopulmonary bypass temperature controller. | Cardiovascular | Cardiovascular | 510(k) | Jostra Heater Cooler Unit HCU30, system version 1.02-Low voltage machine |  |  | November 10, 2006 | 2006 | MAQUET Inc. | The HCU 30 system potentially will not be have the power capacity to run the compressor when the heaters are on during operation. This may lead to a situation where the user may not be able to cool the patient and/or cardioplegia sufficiently. | Other | Letters were mailed to customers on 11/10/2006 by Federal Express. |  | 13 | 13 units | Hospitals in CA, NC, FL, WY, NY and WV.S | Terminated | February 05, 2007 | 87 | 13 | Undetermined
Z-0372-2007 | 37089 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Ventilator System SERVO-i V3.1 PC1777 |  |  | May 23, 2006 | 2006 | MAQUET Inc. | Reported failures of the SERVO-i ventilators user interface system causing technical errors and increased fault frequencies. Switches may be activated unintentionally, or a function not activate when the switch is selected. Direct access knobs may not function or change the parameter in the wrong direction compared to how they are turned. | Other | Maquet sent an ''Extended Warranty'' letter on May 23, 2006. On March 13, 2007, Maquet issued a Recall Notification letter to all customers stating that they will replace all PC1777 boards in the serial number range 2400-3300, which have not previously been replaced. A safety alert letter was included outlining the failure indication and effect for the user of the failures, as well as, how to identify whether the Servo-i unit has shown any failure indications. |  | 1970 | 1, 970 units | Worldwide | Terminated | April 11, 2008 | 689 | 1970 | Undetermined
Z-0342-2007 | 36672 | No | LYJ |  |  | Neurology | PMA | Handhelds-Cyberonics VNS Therapy System, Model 250-''Handhelds'' preprogrammed with software v6.0, v6.1, v6.1.7, v7.0 and v7.1.3; manufactured by Cyberonics, Inc., Houston, TX. |  |  | October 17, 2006 | 2006 | Cyberonics, Inc | During programming, pulse generator may be inadvertently set to 8.0 mA output, regardless of the mA range selected by the clinician | Other | Firm sent safety alert to physicians on 11/14/06. Safety Alert informs physicians of the problem and provides instructions on how to avoid an event and correct an event should it occur. |  | 7404 | 7,404 Handheld PCs pre-programmed with software. | Worldwide Distribution ---- USA and countries of Argentina, Australia, Austria, Belgium, Canada, Cyprus, Czech Rep, Denmark, Finland, France, Germany, Greece, Hong Kong, Hungary, Iceland, Ireland, Israel, Italy, Japan, Jordan, Malaysia, Mexico, Netherlands, Norway, Poland, Portugal, S. Korea, Slovenia, South Africa, Spain, Sweden, Switzerland, Turkey, United Kingdom, and Venezuela. | Terminated | October 02, 2009 | 1081 | 10942 | Undetermined
Z-0363-2007 | 36256 | No | KFX | External assembled lower limb prosthesis. | Physical Medicine | Physical Medicine | 510(K) Exempt | Electronic Otto Bock Compact leg prosthesis System, 3C95/3C85, to be used exclusively for the exoprosthetic fitting of amputations of the lower limb. Microprocessor-controlled prosthetic knee joint used by lower extremity amputees. |  |  | August 10, 2006 | 2006 | Bock,Otto,Orthopedic Ind,Inc | A correction of the Otto Bock Compact prosthetic knee joint, material numbers 3C93 and 3C85, is being conducted. A limited number of Compacts have a defective solder connection between the contact wires and battery terminals. If exposed to strong vibrations these Compact knee joints have an increased probability of malfunction and a resultant risk to the Compact wearer. | Other | An Urgent Correction Notice, dated August 11, 2006, was sent to customers who purchased the Compact Prosthetic knee joint. The letter identifies affected units, asks for all in stock devices to be returned to Otto Bock, and for patients fitted with an affected device, a loaner knee will be provided as a temporary replacement while their device is serviced. |  | 176 | 176 | Nationwide. | Terminated | November 26, 2008 | 839 | 176 | Undetermined
Z-0340-2007 | 36333 | No | DWE | Cardiopulmonary bypass pump tubing. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Intersept Custom Tubing Pack. Do Not Reuse. The design and components of this product were specified by the user from an array of components and designs offered by the manufacturer. Any change in the cardiopulmonary bypass procedure that compromises the function of this Custom Tubing Pack is the responsibility of the user. The tubing is made of polyvinyl chloride (PVC). Portions of the tapered tubing are used in the vent, suction and arterial roller heads of the heart lung machine. Manufacturer: Medtronic, Inc., Minneapolis, MN 55432-5604 USA. Tube supplier: Natvar (Tekni-Plex Co.), 19555 E. Arenth Ave., City of Industry, CA 91748. |  |  | September 13, 2006 | 2006 | Medtronic Perfusion Systems | An out of specification condition exists with the tapered tubing found in some Custom Tubing Packs. The wall thickness is undersized and does not meet specifications. | Other | An Important Customer Notification, dated 09/06/06, was sent to affected customers. The letter identifies the affected model and lot numbers. Medtronic requests that affected product be returned. A suplemental pack can be built and shipped to the customer upon request. Medtronic will replace all returned product. A Recall Certificate is asked to be completed by customer and faxed to Medtronic. |  | 146 | 146 | Distributed in 6 states: CA, FL, IL, IN, LA and TN | Terminated | February 03, 2007 | 143 | 146 | Undetermined
Z-0361-2007 | 37011 | No | DKJ | Salicylate test system. | Toxicology | Toxicology | 510(k) | Salicylate-SL, in vitro diagnostic, Cat. No.511-20, Lot No. 29892 |  |  | December 05, 2006 | 2006 | Diagnostic Chemicals, Ltd. | Upon subsequent testing on different instruments (as required by other customers), DCL detected a non-linear response on certain analyzers. A decision was made to recall the product due to the fact that DCL customers sell the product to different end users, DCL cannot be certain which instruments the end users employ in their laboratories | Other | For product shipped by DCL to one customer notification letter dated 12/5/06 sent out to customer to return/discard product. Replacement will be sent to customer. For product shipped to DCL (USA) warehouse notification sent to discard kits remaining in stock. For product remaining at DCL Canada in-house notification sent out to discard in-house kits. |  | 55 | 55 kits shipped to one customer | PA | Terminated | January 27, 2016 | 3340 | 55 | Undetermined
Z-0362-2007 | 37071 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Suros ATEC Breast Biopsy System handpiece (9 gauge cannula 12 centimeters in length with a 20 millimeter aperture) for use with the ATEC Breast Biopsy and Excision System, disposable, sterile; Part no. 0912-20. |  |  | December 15, 2006 | 2006 | Suros Surgical Systems Inc. | The handpiece package may have an incorrect part number on the inner package. Error could result in over-penetration of the biopsy needle. | Other | Consignees were notified by phone beginning on 12/15/06 and by follow-up letter dated 12/19/06. |  | 1249 | 1249 handpieces | Nationwide. | Terminated | May 04, 2007 | 140 | 1249 | Undetermined
Z-0364-2007 | 36991 | No | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | i-STAT PT/INR Cartridge prothrombin time test in vitro diagnostic. |  |  | December 12, 2006 | 2006 | Abbott Point of Care Inc. | Patients treated with the antibiotic CUBICIN can cause a concentration dependent false prolongation of prothrombin time (PT) and elevation of INR (International Normalized Ratio) when using the i-STAT PT/INR cartridges. | Other | Customers were notified by letter sent by regular mail on 12/13/2006. |  | 337080 | 14,045 boxes of 24 | Nationwide to hospitals and health care facilities | Terminated | August 15, 2007 | 246 | 337080 | Undetermined
Z-0349-2007 | 36843 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | DRLock Volar Plate Screw System, Model #DVR-900. The product is a kit contained in an aluminum tray case that contains the following: IMPLANTS: |  |  | August 03, 2006 | 2006 | OrthoHelix Surgical Designs Inc | The 48 distal locking-screws (ranging from 10-24mm) contained in the kit are not within established specifications, specifically the minor diameter of the screws are undersized. | Other | The recalling firm contacted all consignees by phone on or about 8/3/06. |  | 180 | 180 | The product was distributed to 5 consignees located in the U.S.: Archway Medical, Columbia, SC; Gulf Coast Surgical, Stafford, TX; Vaugham Medical, Wheat Ridge, CO; Belemy Medical, Cleveland, OH; and Crystal Clinic, Akron, OH. | Terminated | September 10, 2012 | 2230 | 180 | Undetermined
Z-0343-2007 | 36785 | No | MJK | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | Welch Allyn PIC 50 Portable Intensive Care System, Multi-Parameter Monitor/Defibrillator; a portable 12V internal battery powered defibrillator; Manufactured by MRL Inc., a WelchAllyn Company, 1000 Asbury Drive, Buffalo Grove, Illinois 60089-Part Numbers-971081,971082,971083,971084. |  |  | November 15, 2006 | 2006 | MRL, Inc., A Welch Allyn Company | Shock Delay:The monitor-defibrillator may display an 'ECG Comm' error message on the display during use, which may prevent or cause an unacceptable delay the delivery of a defibrillation shock. This delay could result in failure to resuscitate the patient. | Other | MRL sent Urgent Medical Device Recall letters dated 11/8/06 via certified mail on 11/15/06 to the direct accounts informing them of an intermittent connection between the J5 cable and its mating connector on a circuit board inside of the device that could result in a malfunction that could prevent or unacceptably delay the delivery of a defibrillating shock, which possibly could result in failure to resuscitate the patient. The accounts were provided a list of the affected units'' parts/serial numbers and were requested to inspect their PIC 50 units for the affected serial numbers and return the identified units for servicing. Until their PIC 50 is serviced, the accounts were instructed to check the unit before each shift by using the self-test, which is automatically initiated by powering up the device. If the 'ECG Comm' error occurs during use, cycling the power off-on may clear the error. Even if power cycling clears the device, the accounts were instructed to contact Welch Allyn - MRL to obtain priority servicing of the affected unit. The accounts were also instructed to respond to the notification within 30 days by fax or mail, or by calling Welch Allyn - MRL Technical Support at 1-800-462-0777, to schedule their PIC 50 for preventative service. Welch Allyn will provide the accounts with a loaner PIC 50 after receiving the form with instructions to send back the PIC 50 in the same shipping box accompanying the loaner unit. Any questions were directed to Welch Allyn-MRL Technical Support at 1-800-462-0777. |  | 510 | 510 units | Worldwide-USA, Australia, Belgium, Canada, Egypt, Finland, France, Germany, Ireland, Italy, Jordan, Poland, Portugal, Saudi Arabia, Singapore, South Africa, Spain, Sultanate of Oman, United Arab Emirates and the United Kingdom. | Terminated | October 31, 2007 | 350 | 510 | Undetermined
Z-0610-2007 | 37147 | No | NGY | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | 19 Gauge Trocar Cannula Replacement Plug, Catalog Number 617.32; Manufactured by Alcon Grieshaber, Ltd., Schaffhauxen, Switzerland CH8203. |  |  | January 18, 2007 | 2007 | Alcon Research, Ltd | Product is misbranded; 19 gauge Schlera Plugs distributed in containers labeled as 19 gauge Trochar Canula Plugs. | Other | Recalled by letter on January 22, 2007. |  | 15 | 15 units. | Nationwide- Hospitals in MD and MN. | Terminated | September 24, 2010 | 1345 | 15 | Undetermined
Z-0377-2008 | 45805 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Roche Ammonia/Ethanol/CO2 Calibrator for use in conjunction with Roche/Hitachi, COBAS INTEGRA, cobas c111 and cobas c 501 clinical chemistry analyzers; Catalog No. 20751995190. Roche Diagnostics Corporation, Indianapolis, IN 46250 |  |  | November 14, 2007 | 2007 | Roche Diagnostics Corp. | High Quality Control Recovery: Undetected high quality control recovery in bicarbonate assay. | Process design | Consignees were notified of this recall by an Urgent Medical Device Recall letter sent on 11/14/07. The letter instructed the users to immediately discontinue use of Ammonia/ Ethanol/CO2 Control N and Ammonia/Ethanol/CO2 Control A (with listed lots) for quality control assessment of the bicarbonate assay. They were informed that controls may still be used to assess ammonia and ethanol, but alternative controls must be used for assessment of the bicarbonate assay. Consignees are instructed to destroy products at their facility. Consignees were notified of the availability of replacement product via letter dated . |  | 1904 | 1,904 | Nationwide Distibution | Terminated | September 22, 2009 | 678 | 4011 | Undetermined
Z-0707-2007 | 37632 | No | MRG | Vitamin D test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | 1,25-Dihydroxyvitamin D RIA Kit -For the quantitative determination of 1,25 Dihydroxyvitamin D in serum or EDTA plasma-DiaSorin, Stillwater, Minnesota 55082-0285, USA |  |  | June 05, 2006 | 2006 | Diasorin Inc. | Potential for the DiaSorin Kit 1,25-Dihydroxyvitamin D RIA Kit Control 1 (lot 548520) &amp; 2 (lot 548521) to recover out of the defined range, LOW. (If one or both of the kit controls recover outside the defined range, the run is considered invalid) | Other | Product recalled by letter dated 6/01/2006 |  | 715 | 715 | Worldwide Distribution-USA including states of UT, CO, IL, NY, NV, MI, NC, MN, VA, GA, CA, KS, VT, CT and IN, and Countries of Canada, Sweden, Germany, Belgium, Italy, France, Norway, Australia, South Korea, Taiwan, Israel, New Foundland, India, China, Brazil, Austria, Saudi Arabia, and Spain | Terminated | June 25, 2007 | 385 | 715 | Undetermined
Z-0784-2007 | 37105 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | GE Healthcare Engstrom Carestation Critical Care Ventilator with the High Performance Display Unit with SW version 3.x and 4.11; identification numbers 1505-9000-000 and M1057699 (upgrade). Datex-Ohmeda, Inc., a General Electric Company, doing business as GE Healthcare, P.O. Box 7750, Madison, WI 53707-7550, USA, www.gehealthcare.com. |  |  | February 02, 2007 | 2007 | Datex - Ohmeda, Inc | There are two unique conditions with the Engstrom Carestation that require corrective action. These include an anomaly in the control of the safety relief valve and a situation where the display unit may reset. | Other | Device Correction Letter to customers was mailed on 2/2/07 giving information about the correction and a short term recommendation until version 4.15 could be installed by a GE representative who would be contacting the customer. |  | 446 | 446 | GA, RI, WI, TX. Foreign - AUSTRALIA, BRUNEI DARUSSALAM, BRAZIL, COLOMBIA, COSTA RICA, SPAIN, FINLAND, FRANCE, UNITED KINGDOM, INDIA, ITALY, KOREA, MEXICO, MALAYSIA, PHILIPPINES, SAUDI ARABIA, SWEDEN, SINGAPORE, GERMANY. | Terminated | October 16, 2008 | 622 | 446 | Undetermined
Z-0727-2007 | 37641 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Sunrise Medical-Hoyer Advance Patient Lifts (Portable and Folding Patient Lift) Sunrise Medical, Inc. |  |  | February 21, 2007 | 2007 | Joerns Healthcare Inc. | Unapproved Design Change (by base supplier) consisting of a reduction in the number of mast-to-base welds (from four welds to two welds). | Other | Recall initiated by letter dated 2/21/2007. Sunrise Medical will schedule onsite appointments to complete the base replacement. |  | 581 | 581 units | Worldwide Distribution --- USA including states of AK, AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NH, NJ, NM, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV and country of Canada. | Terminated | October 07, 2010 | 1324 | 581 | Undetermined
Z-0724-2007 | 37582 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Smith &amp; Nephew Suture Passer 12 inch |  |  | March 01, 2007 | 2007 | Smith &amp; Nephew, Inc. Endoscopy Division | Product is labeled as reusable but there is a potential that the cannulated handle of the product cannot becleaned adequately following use. | Other | Smith &amp; Nephew notified accounts by letter on 3/1/07, users are advised to destroy product and return unused product in their inventory. A reply form is required to be completed. |  | 12292 | 12, 292 units for all recalls Z-0724-0726-07 | World wide- USA and Canada, Belgium. Germany, Denmark, Spain, Finland, France, Great Britain, Greece, Italy, Netherlands, Portugal, Austria, Ireland, Sweden, Argentina. Chile, Columbia, Brazil, China. Japan, Korea, Thailand, Hong Kong, Malaysia. Taiwan, Singapore, India, , Dubai, Mexico, Australia, Egypt. Switzerland, New Zealand and Norway | Terminated | July 14, 2008 | 501 | 12292 | Undetermined
Z-0791-2007 | 37780 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ARCHITECT i2000SR Processing Module, for in vitro diagnostics, List Number 3M74-01; manufactured by Abbot Laboratories, Irving, TX. |  |  | March 28, 2007 | 2007 | Abbott Laboratories, Inc | If the ARCHITECT i2000/i2000sr System becomes contaminated with microbes generating folate-like by-products, the ARCHITECT Folate assay may experience calibration failures and/or shifts in results. | Other | Notification Letters were sent to all customers who received the ARCHITECT Folate Reagent on 03/28/07. Customers were provided with instructions on how to conduct a Wash Buffer Unload procedure and were requested to perform this action daily. Customers were also asked to contact their local Abbott Customer Support Rep to schedule an Internal Decontamination procedure on the ARCHITECT System. Firm recalling to the user (laboratory) level. |  |  |  | Folate Reagent distributed to the following countries: Australia, Austria, Belgium, Bosnia and Herzegovina, Brazil, Canada, Chile, China, Colombia, Croatia, Czech Republic, Denmark, Dominican Republic, Finland, France, Greece, Hungary, Ireland, Israel, Italy, Jordan, Lebanon, Luxembourg, Netherlands, New Zealand, Norway, Oman, Peru, Poland, Portugal, Qatar, Russia, Saudi Arabia, Serbia and Montenegro, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Trinidad and Tobago, Turkey, UAE, United Kingdom. No US distribution. | Terminated | September 25, 2007 | 181 |  | Undetermined
Z-0723-2007 | 37383 | [6 empty] | Hamilton brand Disposable Precision Tips, |  |  | January 30, 2007 | 2007 | Hamilton Co | Some of the packaged tips have increased force during aspiration, which can lead to pipetting errors such as LLD error messages, and also has risk of causing a missed pipetting. | Other | The firm initiated its recall on January 30, 2007. The initial notification was sent via certified mail and e-mail. Subrecalls will be conducted by the distributor. |  | 37 | 37 cases | Nationwide (via a single distributor in New Jersey) | Terminated | November 28, 2007 | 302 | 37 | Undetermined
Z-0728-2007 | 37446 | No | CHP | Estradiol test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ARCHITECT Estradiol Reagent, each kit contains bottles of coated Microparticles and bottles of Conjugate; 400 Test Kit-list number 2K25-20 (4 x 100 tests) and 100 Test Kit-list number 2K25-25 (1 x 100 tests); Abbott Laboratories, Abbott Park, IL 60064-3500 |  |  | February 19, 2007 | 2007 | Abbott Laboratories | A small number of reagent kits were incorrectly configured. The affected kits contain either two bottles of conjugate or two bottles of assay diluent instead of one bottle of each. | Other | Abbott initiated a field correction recall by letter dated 2/19/07 to all customers receiving the affected reagent. The accounts were informed of the incorrectly configured reagent kits. The accounts were requested to assess the inventory for the affected lots of reagents, inspect each kit and destroy any kits incorrectly configured, recording the number of kits inspected and the number of kits discarded on the customer reply form and faxing it to Abbott by 3/2/07, indicating that the letter was received and understood. Replacement kits will be issued for the number of kits destroyed, as indicated on the reply sheet. Customers requiring additional information were directed to contact Abbott Customer Service at 1-877-422-2688. |  | 1944 | 1,944 kits | Worldwide distribution including U.S.A., Puerto Rico, Canada, Mexico, Venezuela, Chile, Uruguay, Argentina, Brazil, Germany, Japan, Singapore, Hong Kong, Taiwan, Thailand, South Korea, Australia, New Zealand and Panama. | Terminated | January 09, 2008 | 324 | 1944 | Undetermined
Z-0705-2007 | 37167 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ABL800 FLEX Series analyzer with FLEXQ module-equipped with software version 5.2 to 5.27 |  |  | January 16, 2007 | 2007 | Radiometer America Inc | In some situations pre-registered patient data is not reset when leaving the pre-registration mode in the software. This occurs when a Safe Pico blood sampler associated with a specific patient is pre-registered in the analyzer by scanning a barcode. | Other | Radiometer America issued a recall letter to their U.S. customers dated 1/16/2007. The letter informs the customers of this software anomaly and the conditions underwhich analyzer results can be stored with incorrect Patient Demographics. The letter goes on to inform the customers that a Radiometer Service Representative will be visiting each customer location shortly in order to upgrade the affected software. In the meantime, the letter recommends that the customer(s) not pre-register their samples. The letter is accompanied by an attached FAX FORM which the customer is asked to sign and complete in order to verify receipt of the recall letter and to fax back to the recalling firm. |  | 1345 | 64 in U.S. &amp; 1,281 to foreign accounts | Worldwide Distribution-USA including states of NY, KS, SD, OH, TX, KY, NC, IL, MO, AL, UT, CO, CA, NM, and MI and Country of Denmark | Terminated | May 08, 2012 | 1939 | 1345 | Undetermined
Z-0721-2007 | 37638 | No | MEQ |  |  | Gastroenterology/Urology | PMA | Urologix Microwave Delivery System 2.5 - 3.5 For use with prostatic urethra lengths of 2.5-3.5 cm only. Model TA1321D. |  |  | March 16, 2007 | 2007 | Urologix, Inc. | The catheter in the kit matches the description on the kit and pouch label; however, an incorrect label was placed onto the catheter during manufacturing. | Other | A letter dated 3/16/07 was FedEx overnight to consignees. The letter described the product, problem and gave instructions for returning and replacing the affected product. |  | 21 | 21 | Nationwide and the Netherlands. | Terminated | May 25, 2007 | 70 | 117 | Undetermined
Z-0722-2007 | 37334 | No | CEO | Phosphorus (inorganic) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Clinical Chemistry Phosphorus in vitro diagnostic, List Number: 7D71-30 and 7D71-20 |  |  | January 30, 2007 | 2007 | Abbott Laboratories Inc. | Patient results are falsely decreased by up to 15% at these levels: 1) Serum Plasma phosphorus greater than 8.0 mg/dL (2.60 mmol/L) 2) Urine phosphorus greater than 80.0 mg/dL (25.80 mmol/L) | Other | A Product Correction letter dated January 30, 2007 provided to all Clinical Chemistry Phosphorus customers that have received lot numbers 42020HW00 and 44037HW00 stating that patient results are falsely decreased by up to 15% at these levels: Serum/Plasma phosphorus greater than 8.0 mgldL (2.60 mmol/L) and Urine phosphorus greater than 80.0 mg/dl. (25.80 mmol/L). Modified linearity ranges are provided in the letter that are to be implemented when using the two lots. Instructions are also provided for changing the linearity range on the Architect cSystem as well as the Aeroset System. Customers are instructed that If they have forwarded any of Information the Clinical Chemistry Phosphorus lots listed above to another laboratory to provide a copy of this letter to them. A customer response form is provided. |  | 87722 | 42020HW00 - 1,351 units worldwide; 44037HW00 - 314 units worldwide | Worldwide, including USA, Canada, Trinidad &amp; Tobago, Bahamas, Panama, Cayman Islands, Jamaica | Terminated | August 10, 2011 | 1653 | 87722 | Undetermined
Z-0793-2007 | 37854 | No | LPL | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic | 510(k) | CIBASoftÂż Visitint Soft Contact Lenses, SKU #749591001385, Power -6.50, Diameter 13.8, Base Curve, 8.9, Expiration 2011-10. |  |  | April 30, 2007 | 2007 | Ciba Vision Corporation | Base curve of the lenses do not meet specification. | Other | Consignees were contacted via traceable mail on 4/30/2007. The initial notification was sent to Eye Care Professionals. The firm subsequently identified one distributor and one retailer. The recall letter was revised and sent to these consignees on May 2, 2007. Specific return instructions were attached along with a Recall Tracking Form that was to be completed and returned to CIBA Vision. They were instructed to return any of the affected product on hand to CIBA Vision for replacement. |  | 22 | 22 lenses | Nationwide. | Terminated | March 13, 2008 | 318 | 22 | Undetermined
Z-0719-2007 | 37598 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | BD 30ml Syringe Luer-Lok tip |  |  | February 15, 2007 | 2007 | Becton Dickinson &amp; Company | The unit label on a limited number of syringes is incorrect, citing a 20 ml size. All other levels of labeling are correct, identifying the product as 30 ml size. | Other | Distributors were sent recall notification letters on February 1, 2007Recall notification letters were sent to all direct accounts and potential end users via UPS second day air on February 15, 2007. |  | 332400 | 332,400 units | Nationwide. | Terminated | July 19, 2007 | 154 | 332400 | Undetermined
Z-0718-2007 | 37486 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Deltec branded CADD-Prizm PCS II ambulatory infusion pump and the CADD-Prizm VIP system ambulatory infusion pumps; (Both) Model 6101; Order #21-8861-01 (PCS II) and 21-8821-01 (VIP). Smiths Medical MD, Inc., St. Paul, MN 55112 USA; RX ONLY. |  |  | January 19, 2007 | 2007 | Smiths Medical MD, Inc. | The year of manufacture had been incorrectly entered as 1980 into the pumps software during the manufacture process. As a result the pump may go into an immediate 'Clock battery is Low/Service Immediately' or 'Clock Battery needs service soon' alarm message indicating that the pump requires servicing. | Other | All consignees who have received the affected CADD-Prizm PCS II ambulatory infusion pumps and the CADD-Prizm VIP System ambulatory infusion pumps were contacted via telephone by February 1, 2007. The problem will be explained to the consignee and they will be asked to return the pumps to Smiths Medical MD in order for the problem to be corrected. |  | 27 | 27 | Nationwide. | Terminated | May 25, 2007 | 126 | 27 | Undetermined
Z-0674-2007 | 37188 | No | BTI | Portable air compressor. | Anesthesiology | Anesthesiology | 510(K) Exempt | Medical Air Compressor |  |  | January 26, 2007 | 2007 | Draeger Medical, Inc. | Compressors not providing sufficient supply pressure to connected medical ventilator | Other | The recalling firm issued a Medical Device Recall Letter dated 1/22/07 to their customers via Certified Mail. An authorized service representative will inspect each potentially affected device. If a device is found to have an affected dryer assembly, it will be repaired free of charge. |  | 67 | 67 units | Nationwide. | Terminated | June 15, 2007 | 140 | 67 | Undetermined
Z-0716-2007 | 37642 | [6 empty] | QRS Card 12 Lead Resting &amp; Stress ECG Machine with Blue Tooth (wireless) interface. Cardiology Suite (CS 4.0) Software is used with the machine. |  |  | March 05, 2007 | 2007 | Pulse Biomedical Inc | No 510(k) marketing clearance for the product | Other | On 3/5/07, the recalling firm telephoned their customers to inform them of the problem and the need to return to the product. As a follow up to the telephone calls a letter dated 3/5/07 was sent via certified mail. |  | 8 | 8 units | The product was shipped to four direct customers in CA, NY, TN and US military (Cuba). The product was also shipped to Chile, Greece, and the Netherlands | Terminated | August 29, 2007 | 177 | 8 | Undetermined
Z-0715-2007 | 37507 | No | LYL | Soft (hydrophilic) contact lens care products. | Ophthalmic | Ophthalmic | 510(k) | Contact lens solution. The product is distributed under the following trade names: Bausch &amp; Lomb ReNuÂż Multi-Plus Multi-Purpose Solution, Equate Multi-Purpose Solution and Target Brand Multi-Purpose Solution. The product is distributed in up to three sizes: 4 fl.oz., 12 fl. oz. and 16 fl. oz |  |  | March 05, 2007 | 2007 | Bausch &amp; Lomb Inc | A higher than expected amount of trace iron in some bottles of finished product that over time affects the product's stability and color. | Other | The firm issued a press release on 03/05/2006. Consignees were notified on 03/06/2007 via a customized Recall Notification Package via overnight courier with tracking and delivery confirmation capability. Bausch &amp; Lomb Sales Representatives will visit the home offices of the major wholesale and distributors concurrent with the mailing of the notification packages. |  | 1500000 | 1.5 million units | Worldwide, including USA, Argentina, Brazil, Canada, Chile, Costa Rica, Korea, Mexico, Taiwan and Venezuela. | Terminated | January 09, 2008 | 310 | 1500000 | Undetermined
Z-0733-2007 | 37465 | No | DKZ | Amphetamine test system. | Toxicology | Toxicology | 510(k) | Bio-Rad-TOX/See Drug Screen Test (25 tests/box), Catalog number 194-5230; Product is distributed by Bio-Rad Laboratories, Inc., Hercules, CA |  |  | February 12, 2007 | 2007 | Bio-Rad Laboratories Inc | The product may produce a faint line which could be interpreted as a false negative test result . | Other | The firm intends to contact its subsidiary distributors and sales staff via Fax (Feb 12, 2007) and phone call. In addition, the recall notification will be posted on the firm's intranet site. Customers are instructed to destroy recalled product following standard laboratory practices. |  | 177 | 177 Kits | Worldwide distribution --- USA and countries of Australia and Italy. | Terminated | February 04, 2008 | 357 | 177 | Undetermined
Z-0710-2007 | 37597 | No | KMF | Liquid bandage. | General Hospital | General &amp; Plastic Surgery | 510(k) | Sea-Clens Wound Cleanser, 6 fl. oz., Manufactured by Coloplast Corp., Marietta, GA, #11701-159-36, #1063. |  |  | February 14, 2007 | 2007 | Coloplast Corp Skin Care Div | Coloplast Corp. is recalling Sea-Clens brands Sea-Clens wound cleaner which were found to have some particulates (foreign materials). | Other | The initial consignees of the product were contacted and notified of the recall February 14, 2007. The consignees were asked to contact their customers immediately and return all product to the manufacturer. |  | 7848 | 654 boxes, 12 bottles per box for a total of 7,848 bottles | Nationwide. | Terminated | June 25, 2007 | 131 | 7848 | Undetermined
Z-0714-2007 | 37358 | No | FRO |  |  | General &amp; Plastic Surgery | 510(k) | ProWick Shoulder Postoperative Dressing and Cold Therapy System, Model number AR-1625 |  |  | February 01, 2007 | 2007 | Integra Biotechnical LLC | During routine new product testing, it was discovered that the product may not be sterile, although sterilized in accordance with validated parameters. | Other | An email and/or fax recall notification was sent starting on 02/1/2007 by Arthrex, Inc. (the Primary Distributor) . A product recall notice in the form of a letter was mailed via US Postal Certified Mail, return receipt requested by Arthrex, Inc. to all of the Distributor''s customers starting on 02/1/2007, including those who have not ordered or received the Recalled Product. Customers were asked to 'segregate and protect the product from shipment, and return all unused products to Arthrex, Inc.' |  | 1769 | 1769 | Nationwide | Terminated | August 03, 2012 | 2010 | 1769 | Undetermined
Z-0777-2007 | 37710 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | BD Vacutainer Safety-Lok Blood Collection Set with Preattached Holder. Color green. 21 Gauge. Translucent, integrated protective shield that provides one-handed activation immediately after use, 0.75 in needle length, 12 in tubing, Luer adapter and holder, wings, sterile. |  |  | March 20, 2007 | 2007 | Becton Dickinson &amp; Company | The BD Vacutainer Safety-Lok Blood Collection Set with Preattached Holder has a problem where the non-patient needle (in the holder) may become separated from the hub and remain the stopper of an evacuated blood collection tube. | Other | Recall letters were sent to customers on 3/11/2007 via a third party, West Inc., Appleton, WI who will coordinate the mailing to all consignees. |  | 147200 | 147,200 units | Nationwide and Canada. | Terminated | July 25, 2007 | 127 | 147200 | Undetermined
Z-0713-2007 | 37018 | No | LDS | Medical washer. | General Hospital | General Hospital | 510(K) Exempt | ThermaSure Medical Device Drying Cabinet, series 1000 |  |  | December 21, 2006 | 2006 | Cenorin | Potential for user to receive an electrical shock. | Other | Letter dated 12/20/2006 notifying customers of potential shock hazard under certain fault conditions and that a corrective action upgrade is being developed. |  | 8 | 8 units | Worldwide. | Terminated | November 09, 2007 | 323 | 159 | Undetermined
Z-0708-2007 | 37454 | No | GIF | Blood cell diluent. | Hematology | Hematology | 510(K) Exempt | CELL-DYN Sapphire and CELL-DYN 4000 Diluent/Sheath Systems- List No. 01H73-01, Distributed and Manufactured by: Abbott Diagnostics Division, 5440 Patrick Henry Drive, Santa Clara, CA |  |  | April 17, 2006 | 2006 | Abbott Laboratories | High platelet background count when using the CELL-DYN 4000 System and/or CELL-DYN Sapphire System; with patient results unacceptable -out-of-range. | Other | On 4/17/06, the firm initiated the recall and its notification letters explained the reason for the recall and reported Abbott Field Service personnel will install and replace filters at all recall sites |  | 56075 | 56,075 units | Worldwide Distribution-USA including states of AL, AR, CA, CO, CT, DE, FL, GA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NV, NY, OH, OK, OR, PA, SC, TN, TX, VA, WA, WI, and WV and Countriers of Argentina, Australia, Brazil, Hong Kong, Panama, Peru, Puerto Rico, Singapore, South Korea, Taiwan, and Venezuela. | Terminated | November 28, 2007 | 590 | 56075 | Undetermined
Z-0795-2007 | 37805 | No | FTM | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Gynecare PROLIFT Total Pelvic Repair System |  |  | April 20, 2007 | 2007 | Ethicon, Inc. | The pre-printed outer box incorrectly identifies the product as GYNECARE PROLIFT Total Pelvic Floor System (Product Code PFRT01). The contents of the box actually contain GYNECARE PROLIFT Anterior Pelvic Floor System (Product Code PFRA01) | Other | Recall notifications were sent by 2nd day air UPS delivery on April 20, 2007 to all hospitals Director of Surgery, OR Managers and Materials Managers. |  | 160 | 160 units to US | Nationwide | Terminated | March 04, 2008 | 319 | 160 | Undetermined
Z-0794-2007 | 37837 | No | DTK | Cardiovascular intravascular filter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific Medi-Tech Stainless Steel Greenfield Vena Cava Filter, order #50-400, lot #7826572 |  |  | October 06, 2006 | 2006 | Boston Scientific | One lot of product, Stainless Steel Greenfield Vena Cava Filter, may have the incorrect expiration date of 2008-07 and the correct date is 2008-06. | Other | Consignees were notified by letter (10/06/2006) which was accompanied by a Reply Verification Letter. All consignees were asked to cease use of product upon receipt of the letter and to fax the Reply Verification Letter to recalling firm and return product to firm. |  | 24 | 24 units | Nationwide. | Terminated | July 03, 2012 | 2097 | 24 | Undetermined
Z-0703-2007 | 37615 | No | FAE | Penile rigidity implant. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | AMS DURA-II Universal Tips, 2cm, Product Number 72009361, Sterile, American Medical Systems, Inc. Minnetonka, MN |  |  | February 19, 2007 | 2007 | American Medical Systems | Mis-Labeling:The Dura-II 2 cm Universal Tips and the Dura-II 4 cm Distal Tips were incorrectly labeled with the opposite labels. | Other | The consignees were notified of the recall by letter starting Feb 7, 2007. Consignees were asked to return product to American Medical Systems. |  | 5 | 5 devices | Nationwide | Terminated | June 27, 2007 | 128 | 5 | Undetermined
Z-0701-2007 | 37596 | No | KDC | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | SCANLAN International Mobin-Uddin Vein Holder, Sterile, Disposable, single-use, catalog #1001-761, Lot AZ115806. |  |  | February 19, 2007 | 2007 | Scanlan International Inc | Incorrectly assembled by manufacturer | Other | The consignees were notified of the product recall by letter dated February 19, 2007. Consignees were asked to return product to the manufacturer and replacement product is being offered. |  | 35 | 35 units domestically ( 70 units to OUS) | Worldwide distribution --- USA including states of IN, MO and country of Japan | Terminated | June 27, 2007 | 128 | 35 | Undetermined
Z-0706-2007 | 37453 | No | GIF | Blood cell diluent. | Hematology | Hematology | 510(K) Exempt | CELL-DYN 3200 Diluent/Sheath Reagent, |  |  | September 29, 2004 | 2004 | Abbott Laboratories | Elevated platelet background count results-some products may show a higher than expected platelet background count when used on the CELL-DYN 3200 System and report patient results that are unacceptable-out-of-range | Other | On September 29, 2004, the firm initiated the recall and its notification was via letters explaining the reason for the recall and requesting the product be disposed. |  | 27411 | 27,411 units | Class II Recall-Worldwide Distribution- USA including states of IL, TX, GA, and PA. and countries of Argentina, Australia, Bahamas, Barbados, Brazil, British Virgin Islands, Colombia, Dominican Republic, Germany, Hong Kong, Japan, Korea, New Zealand, Puerto Rico, Singapore, Taiwan, Thailand, Trinidad &amp; Tobago, Uruguay, and Venezuela. | Terminated | May 01, 2007 | 944 | 27411 | Undetermined
Z-0603-2007 | 37049 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1; 220V -240V, 7A (circuit breaker), 50/60 Hz (10A power source required) base; Model 801764. (Not distributed within the United States). |  |  | August 01, 2005 | 2005 | Terumo Cardiovascular Systems Corp | The power supply may fail to function due to a defective capacitor. (Note; this device incorporates a redundant separate power supply not affected by this capacitor) | Other | The power supplies were replaced as complaints were received between August 2005 and February 2006. The firm issued a Device Correction Letter dated 6/6/07 notifying all users of the problem and advising them that their units will be corrected when replacement components become available. |  | 534 | 534 of both model numbers | Worldwide Distribution: USA, Australia, Belgium, Canada, Dominican Republic, Egypt, Germany, Hong Kong, India, Indonesia, Japan, Jordan, Malaysia, Mexico, Pakistan, Philippines, Russia, Saudi Arabia, Singapore and South Africa. | Terminated | August 04, 2009 | 1464 | 534 | Undetermined
Z-0600-2007 | 37017 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1, 4 inch diameter (small) Roller Pump; Model 801040. |  |  | September 01, 2004 | 2004 | Terumo Cardiovascular Systems Corp | The internal ribbon cable may loosen from the roller pump display board, resulting in loss of local control and/or local display. | Other | The firm corrected units in the field or brought them back for evaluation and repair as complaints were received between 9/04 and 4/06. The firm issued a Device Correction Letter dated 6/6/07 notifying all users of the problem and advising them that they will be visited and their units upgraded as replacement components become available. |  | 1484 | 1,484 of both models | Worldwide: USA, Australia, Belgium, Canada, Dominican Republic, Egypt, Germany, Hong Kong, India, Indonesia, Japan, Korea, Mexico, Malaysia, Singapore, South Africa, Taiwan, Thailand and United Arab Emirates. | Terminated | July 07, 2009 | 1770 | 1484 | Undetermined
Z-0599-2007 | 36399 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1; 220V -240V, 7A (circuit breaker), 50/60 Hz (10A power source required) base; Model 801764. (Not distributed within the United States). |  |  | November 01, 2005 | 2005 | Terumo Cardiovascular Systems Corp | The power manager subsystem may not initialize properly during start-up, which will cause the battery status monitoring and battery switching operations to not be fully functional. | Other | Firm repaired the affected units beginning in November 2005, upon receipt of complaints of failure. The firm issued a Device Correction Letter dated 6/6/07 notifying all users of the problem and advising them that they will be visited and their units upgraded as replacement components become available. |  | 485 | 485 of both model numbers | Worldwide distribution ---- Nationwide and countries of Australia, Belgium, Canada, Dominican Republic, Egypt, Germany, Hong Kong, India, Indonesia, Japan, Malaysia, Mexico, Pakistan, Philippines, Republic of Georgia, Russia and Saudi Arabia. | Terminated | June 23, 2009 | 1330 | 485 | Undetermined
Z-0786-2007 | 37437 | No | JCW | Penile inflatable implant. | Gastroenterology/Urology | Gastroenterology/Urology | PMA | AMS 700 CX MS Pump IZ Preconnected penile Prosthesis, product length 18 cm, product number 72404232. Penoscrotal Approach - with InhibiZone, 1 MS Pump, 2 Cylinders. |  |  | February 09, 2007 | 2007 | American Medical Systems | The carton labeling does not match the product packaged inside. | Other | AMS has contacted every account (via e-mail 2/9/2007) that either purchased or implanted the affected product. Unused pieces have been recovered and scrapped. |  | 5 | 5 | Worldwide, including USA, Canada and England. | Terminated | February 24, 2008 | 380 | 26 | Undetermined
Z-0735-2007 | 37585 | No | NTI | Drug metabolizing enzyme genotyping system. | Toxicology | Toxicology | 510(k) | AmpliChip CYP450 Test, CE-IVD, M/N: 4381866190, Roche Molecular Systems, Inc. |  |  | March 09, 2007 | 2007 | Roche Molecular Systems, Inc. | The current package insert for the AmpliChip CYP 450 test, lists an incorrect part number for DNase I. The DNase is listed in the "Other Materials Required" section as 'DNase I rec., RNase-free, P/N 04716728001 (Roche Applied Science)" . The Part Number listed for the DNase I listed is incorrect and of lower specific activity than the correct/validated DNase. | Other | A product advisory notice was issued on March 9, 2007 to all affiliates, indentifying the problem and listing corrective actions followed by a customer letter (Urgent Medical Device Correction) issued on 3/21/2007, requesting several actions be performed including discontinued use of incorrect DNAse I. |  | 89 | 89 | World wide Distribution --- USA including states of MI, IN, NC, CA KY, IN, VA, and FL.and countries of Netherlands, Germany, Sweden, Denmark, Italy, Spain, Belgium, Switzerland, and Hungary | Terminated | August 23, 2007 | 167 | 218 | Undetermined
Z-0700-2007 | 37462 | No | LGW |  |  | Neurology | PMA | The Precision Implantable Pulse Generator, Model Number SC-1110 |  |  | January 30, 2007 | 2007 | Advanced Bionics Corp | The Precision Implantable Pulse Generator with firmware Revision 3.01 incorrectly reports an error condition for a valid output from its self internal check. | Other | March 2007 by letter to physicians who had implanted the units notifying them of the recall and future upgrade by wireless RF telemetry. |  | 57 | 5 unimplanted devices; 52 implanted devices | Nationwide including Puerto Rico | Terminated | January 25, 2012 | 1821 | 57 | Undetermined
Z-0773-2007 | 36826 | No | MDB | Microbial growth monitor. | Microbiology | Microbiology | 510(k) | BacT/ALERT Classic System Software, Blood Culturing System |  |  | August 14, 2006 | 2006 | bioMerieux, Inc. | Plastic bottle samples assigned less sensitive default glass bottle algorithm instead of the plastic bottle algorithm on BacT/ALERT Microbial Detection Systems. | Other | Consignees were notified via A Field Corrective Action Notification on/about 10/18/2006. A second notification was sent on October 31, 2006. |  | 10790 | 10790 units, total for all models | Worldwide, including USA, Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Denmark, Dutch Antilles, Egypt, Ecuador, Estonia, Faeroe Islands, Finland, France, Germany, Greece, Hong Kong, Hungary, Iceland, India, Indonesia, Ireland, Israel, Italy, Japan, Kenya, Korea, Latvia, Lithuania, Luxembourg, Malaysia, Mexico, Morocco, Netherlands, New Zealand, Norway, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Romania , Russia, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Sri Lanka, Suriname, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, Uruguay, Venezuela, and Vietnam. | Terminated | December 08, 2009 | 1212 | 10790 | Undetermined
Z-0760-2007 | 37607 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Puritan Bennett 840 Ventilator System, Catalogue Numbers: |  |  | March 12, 2007 | 2007 | Nellcor Puritan Bennett | In rare instances, the safety valve may remain open after an occlusion is resolved. If the safety valve does not reset, the ventilator will detect a circuit disconnect condition. When this happens, the ventilator will annunciate a high-urgency audible alarm and the ventilator will not resume normal operation. | Other | The correction will extend to the end user. All users will be contacted by mail or through personal contact. In the U.S., all customers were mailed the 'Urgent Device Correction' letter along with a serial number listing of those units which records show were consigned to them beginning on March 12, 2007. The customer communication informs customers of the issue and provides instruction for continued safe use of the device pending the correction. |  | 3756 | 3756 | Worldwide, including USA, ARGENTINA, AUSTRALIA,BELGIUM, BRAZIL, CANADA, CHILE, COLOMBIA, EGYPT,FRANCE, GERMANY, GREECE, GUATEMALA, IRAN,IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KAZAKHSTAN,KOREA, LEBANON, MEXICO, MOROCCO, NICARAGUA, PANAMA, POLAND, PUERTO RICO, ROMANIA, RUSSIA,SINGAPORE, SOUTH AFRICA, SPAIN, TURKEY, UK,&amp; YUGOSLAVIA | Terminated | December 13, 2012 | 2103 | 3756 | Undetermined
Z-0759-2007 | 37604 | joint, knee, external limb component | ISY | External limb prosthetic component. | Physical Medicine | Physical Medicine | 510(K) Exempt | Medipro OP4 Prosthetic Knee Joint (external for prosthesis, not implant). |  |  | August 11, 2006 | 2006 | Medi Manufacturing Inc | Joints contain a component manufactured with an alloy that may not comply with manufacturer's specifications. | Other | Each consignee was contacted by phone on 8/11/2006. A follow up letter was sent to consignees on 8/11/2006 providing details on who to contact and how to return the affected devices for replacement. A self addressed, stamped response form was included to be completed and returned to Medi. |  | 8 | 8 units | Nationwide. | Terminated | April 12, 2011 | 1705 | 8 | Undetermined
Z-0755-2007 | 37691 | No | LLM | Epstein-Barr virus serological reagents. | Microbiology | Microbiology | 510(k) | DiaSorin ETI-EBNA-G kits, Catalog No. P001607A. In vitro diagnostic. |  |  | March 14, 2007 | 2007 | Diasorin Inc. | A lot to lot inconsistency was noted on kits of ETI EBNA G. | Other | All consignees will be notified by phone followed by a notification letter sent by fax March 14, 2007. The distributors are instructed to forward the notification to the customers who have received the affected product. |  | 203 | 203 | Nationwide and Canada. | Terminated | July 15, 2007 | 123 | 203 | Undetermined
Z-0767-2007 | 37759 | No | GES | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Miltex Stainless Steel Disposable Scalpel #22, Part Number 4-422. Each box contains 10 individually packaged scalpels. |  |  | April 09, 2007 | 2007 | Miltex, Inc. | Sterility may be compromised based on incomplete package seals (manufacturer notified Miltex) | Other | The recalling firm issued a recall letter dated 4/9/07 informing the customers of the problem and the need to return the product. The letter was sent via UPS with a return receipt required to verify delivery. Distributors were notified to sub-recall. |  | 400 | 400 boxes | Nationwide to distributors in CA, FL, IA, ID, NM, NJ, OH, PA, RI, and VA. | Terminated | August 07, 2007 | 120 | 400 | Undetermined
Z-0765-2007 | 37281 | No | HQE | Vitreous aspiration and cutting instrument. | Ophthalmic | Ophthalmic | 510(k) | Bausch &amp; Lomb Millennium Microsurgical System TSV 25 Ga. High Speed Vitrectomy Cutter, Catalog No. CX5825, packaged individually in a Tyvek pouch, 6/box, RX. The firm names on the label are Bausch &amp; Lomb Incorporated, Rochester, NY, and Bausch &amp; Lomb Incorporated, Kingston-upon-Thames, UK. |  |  | February 09, 2007 | 2007 | Bausch &amp; Lomb Inc | The cutter probe tip breaks during use. | Other | On 2/12/07, express mail letters were addressed to operating room supervisors to notify them of the recall. Additional E-mail notifications were issued on 2/14/07 to doctors and letters were addressed to them on 2/20/07 to make them aware of the recall, although they do not directly receive the product. The communications instructed them to not use the affected product and to return it, along with the business reply card. |  | 187506750 | 32,190 units of Catalog No: CX 5825 | Class II Recall - Worldwide Distribution --- USA including states of MO, NC, FL, CA, and TX and countries of Canada, China, Japan, Australia, Hong Kong, India, Singapore, Korea, Taiwan, Thailand, Argentina, Brazil, Mexico, and The Netherlands | Terminated | April 07, 2009 | 788 | 187520435 | Undetermined
Z-0731-2007 | 37189 | No | NFB | Noncontinuous ventilator (IPPB). | Anesthesiology | Anesthesiology | 510(k) | DeVilbiss PD1000A Pulse Dose Oxygen Conserving Cylinder |  |  | January 24, 2007 | 2007 | Sunrise Medical | Cylinder could lose oxygen at high pressure after the filling process | Other | The recalling firm began to notify customer via telephone on 1/24/07. An ''Urgent Medical Device Recall Letter'' was sent on 1/29/07. Customers were advised what model and serial numbers were affected and how to contact the firm and/or dealer to have the unit replaced. The recalling firm identified additional serial numbers that were involved in the recall therefore another Urgent Medical Device Recall Letter dated 6/12/07 was sent to the customers that received the addiitonal units. |  | 4169 | 4,169 units (Z-0731/0732-2007) | Worldwide-original shipments to distributors/dealers USA and Canada. | Terminated | August 27, 2009 | 946 | 4169 | Undetermined
Z-0769-2007 | 37760 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Magnetom Espree System with OR Table |  |  | March 30, 2007 | 2007 | Siemens Medical Solutions USA, Inc | If table top is not seated properly it can lock up/stick. | Other | The recalling firm issued a Safety Advisory Letter to their customers on March 30, 2007, informing them of the problem and steps to take to avoid the problem. |  | 3 | 3 units | The product was shipped to medical facilities in NY and TX. | Terminated | August 07, 2007 | 130 | 3 | Undetermined
Z-0761-2007 | 37518 | No | GWQ | Electroencephalograph. | Neurology | Neurology | 510(k) | Bio-logic-Ceegraph/Sleepscan Netlink Traveler, Bio-logic Systems Corp., a nebus company, Mundelein, IL 60060; Model 580-T2ASM2. (a Digital EEG/Sleep Recorder Electroencephalograph that includes a built-in pulse oximeter, body position sensor, snore monitor, chest, abdominal and air flow transducers and electrode array for EEG, EMG and EOG) |  |  | February 26, 2007 | 2007 | Bio-Logic Systems Corp | Battery and Battery Compartment; incorrect placement in the Netlink Traveler unit which caused a short in the wiring of the battery pack and overheating. | Other | On 2/26/07, Bio-logic Systems shipped replacement batteries, marked with part number 520-LIBAT4-3 and a label reading 'THIS SIDE OUT', and a new battery compartment cover with a label inside the cover to clarify the correct installation of the battery to each of the affected customers. The accounts were requested to remove and replace the old batteries, part 520-LIBTA4-2, and compartment covers, and return the old batteries and covers to Bio-logic Systems in the enclosed pre-paid shipping container, along with the completed accounting form. |  | 130 | 130 battery packs (112 domestic USA, 18 foreign) | Worldwide, USA , Canada, China, Hong Kong, India and Israel. | Terminated | August 26, 2008 | 547 | 130 | Undetermined
Z-0768-2007 | 37590 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Medfusion 3500 Syringe Infusion Pump with PharmGuard Medication safety Software versions 4.1.3 and 4.1.4. Medex, 4350 Green Parkway, Suite 200. Duluth, GA 30096 |  |  | February 03, 2007 | 2007 | Smiths Medical MD, Inc. | Software anomaly that causes the pump to experience a sudden and unexpected failure. This anomaly occurs during the pump's startup sequence when the pump's Preventative Maintenance Alarm goes off, preventing the pump from completing its power up routine. Lock-up occurs and the pump will not function. | Other | Consignees were initially notified of the recall via a Customer Information Bulletin which was sent via electronic mail 2/3/07 and confimed with a telephone conversation. A Product Recall Notification and a Product Recall Bulletin will be sent to all consignees via the mail on 2/16/07. |  | 23519 | 23,519 | Nationwide, Canada, and Italy. | Terminated | February 10, 2011 | 1468 | 23519 | Undetermined
Z-0758-2007 | 37591 | No | CCW | Rigid laryngoscope. | Anesthesiology | Anesthesiology | 510(K) Exempt | Conditional Fiber Optic Laryngoscope Kit, which contains laryngoscope blades, handles and 2 battery packs for use in Magnetic Resonance environments; Distributed by Teleflex Medical, Research Triangle Park, NC 27709; catalog number 005852300 |  |  | March 01, 2007 | 2007 | Teleflex Medical | The MRI Battery is out of specification for its magnetic content. | Nonconforming Material/Component | Teleflex telephoned the account on 3/1/07, and faxed them a follow-up recall letter on the same date. The account was informed that the batteries did not meet specification, and was requested to check the batteries in the laryngoscope kit shipped to them for the affected battery lot numbers, and return the affected batteries to Teleflex for replacement. Any questions were directed to Linda Todd at 1-800-334-9751, ext. 4951. |  | 1 | 1 kit | Georgia | Terminated | May 03, 2007 | 63 | 1 | Undetermined
Z-0754-2007 | 37742 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | RITAÂż StarburstÂż MRI SemiFlex (3-5 cm Diameter/25cm Length), Electro-surgical Device, Manufactured by: RITA Medical Systems, Inc., One Horizon Way, Manchester, GA 31816 |  |  | April 03, 2007 | 2007 | Rita Medical Systems, Inc. | The product may have a cracked tray which can compromise the sterility of the product. | Other | Notification by letter of the AngioDynamics, Inc. sales force and the domestic and international distributors began on April 2, 2007. They were asked to return all affected devices to the Manchester, GA division of AngioDynamics, Inc. A reply form was attached to be completed and returned via fax to AngioDynamics. |  | 2021 | 2021 devices | Worldwide, including USA, Austria, Canada, Croatia, Denmark, Egypt, Estonia, France, Finland, Germany, Greece, Hong Kong, Hungary, India, Israel, Italy, Japan, Latvia, Netherlands, Norway, Poland, Republic of Ireland, Romania, Russia, Saudi Arabia, Singapore, Solomon Islands, South Africa, Spain, Switzerland, UK, and United Arab Emirates. | Terminated | January 20, 2010 | 1023 | 5103 | Undetermined
Z-0749-2007 | 37698 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Mobilett XP, mobile X-ray system, Model number 1818454 |  |  | March 21, 2007 | 2007 | Siemens Medical Solutions USA, Inc | Mounting bolts for tank fork assembly and collimator flange may become loose. | Other | Siemens Service Personnel are visiting each consignee, beginning March 21, 2007, to replace the required mounting bolts and tighten them according to required torque specifications. |  | 186 | 186 units | The products were shipped to medical facilities nationwide. | Terminated | August 07, 2007 | 139 | 227 | Undetermined
Z-0747-2007 | 37528 | No | LOM |  |  | Microbiology | PMA | VITROSÂż Immunodiagnostics Products HBsAg Reagent Pack, Catalog # 6801322: 1 Reagent Pack box per sales unit (100 tests per box), and Catalog # 6802450: 5 Reagent Pack boxes per sales unit (100 tests per box), IVD. Ortho-Clinical Diagnostics Amersham UK. |  |  | February 01, 2007 | 2007 | Ortho-Clinical Diagnostics | Complaints of an increase in ''Reactive'' results with patient samples collected in sodium citrate or EDTA plasma collection tubes when using these lots of VITROS HBsAg Reagent Pack compared to samples collected in other tube types. | Other | On 2/01/07, all US end-user consignees who received the VITROS HBsAg Reagent Packs, were informed of this issue by Federal Express overnight letter, and informed that they may continue to use their current lot of VITROS HBsAg Reagent Pack providing that they followed the instructions in the letter. INSTRUCTIONS: Please do the following if you choose to continue using sodium citrate and/or EDTA plasma collection tubes: - Do not freeze samples collected in sodium citrate or EDTA plasma collection tubes prior to processing the assay. - Process samples collected into sodium citrate plasma tubes within 12 hrs. of collection. -Continue to perform supplemental testing to confirm all HBsAg results that repeat as 'Borderline' or 'Reactive', as directed in the Package Insert. - Complete and return the attached Confirmation of Receipt Form upon receipt of this notification. - Forward the information in this notification if you have distributed the product outside your facility |  | 48514125 | Cat. #6801322 (1 Reagent Pack box per sales unit): 11,761 units; Cat. #6802450 (5 Reagent Pack boxes per sales unit): 4125 units | Worldwide Distribution --- USA including Puerto Rico and countries of Tortola (British Virgin Islands) and Anguilla. | Terminated | August 03, 2012 | 2010 | 48514125 | Undetermined
Z-0748-2007 | 37606 | No | JNO | Trypsin test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | DiaSorin TRYPSIK Kits, Catalog Number P2573-(assay used as a procedure for the quantitative determination of trypsin-like immunoreactiviy (TLI) in human serum or plasma samples) |  |  | February 15, 2007 | 2007 | Diasorin Inc. | Assay Failure: Assay failures may occur due to the kit control running outside of the established range (invalidating the assay) | Other | The consignees &amp; distributors were notified by letter on 2/15/07 . The Distributors are instructed to forward the notification to the customers who have received the affected products. A response form will be requested from the customers. |  | 100 | 100 kits | Worldwide Distribution --- USA including states of UT, CO, IN, NC, CA, FL and countries of Canada, Spain, Germany, Belgium &amp; Italy. | Terminated | July 15, 2007 | 150 | 100 | Undetermined
Z-0737-2007 | 37602 | No | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | CoaguChek brand PT Test Strips; U.S. Catalog Number 3116247 (48 strip pack-professional use). |  |  | February 28, 2007 | 2007 | Roche Diagnostics Corp. | The product was erroneously distributed to home users in contrast to its use for healthcare professionals only. | Other | All consignees were notified by telephone on/or about 2/28/07 and some received a follow-up letter dated 3/2/07 or 3/5/07. |  | 10 | 10 | Nationwide-California, Georgia, Louisiana, Maryland, Oklahoma and Texas. | Terminated | June 13, 2007 | 105 | 10 | Undetermined
Z-0778-2007 | 37719 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | GEM Premier 4000 Laboratory Analyzer |  |  | March 27, 2007 | 2007 | Instrumentation Laboratory Co. | Instrument in Micro Sampling Mode failed to aspirate and no error code was generated and an incorrect patient report could result | Other | Instrumentation Labs on 3/2707 notified Service Reps to perform on site mandatory upgrade. The foreign counterparts were advised by email on 3/21/07 |  | 210 | 210 | Worldwide, including USA, Australia, Austria, Greece, Spain, Netherlands, Great Britain, France, Germany, Italy, and Sweden | Terminated | May 23, 2012 | 1884 | 210 | Undetermined
Z-0746-2007 | 37427 | No | LJT | Subcutaneous, implanted, intravascular infusion port and catheter. | General Hospital | General Hospital | 510(k) | Vaxcel Implantable Vascular Access System - Plastic Standard Port With 9F Polyurethane Catheter, Catalog/REF No. 45-360, UPN/Material No. M001453600, Single Use Only, STERILE, EO, Boston Scientific MEDI-TECH (617) 972-4000/(800) 225--3238, 480 Pleasant Street, Watertown, MA 02472 --- The port is provided as a kit, with various single use, disposable items necessary for achieving vascular access and implanting the port catheter. Each kit is provided in a plastic (PETG) tray with a heat-sealed tyvek lid. This inner tray is then placed in a second outer plastic (PETG) tray, along with the Directions for Use, and also sealed with a tyvek lid which is then labeled. The port kit is placed in a cardboard carton, with additional labels and sealed. |  |  | June 13, 2006 | 2006 | Boston Scientific Corp | Mispackaging: some kits of peelable sheaths/dilators may contain a 10F sheath instead of a 9F sheath--(Recalling Firm was notified by their supplier B. Braun Medical of the error) | Other | Boston Scientific sent Urgent Medical Device Recall Notification- Immediate Action Required letters- (dated 6/13/06) to all 11 end users via Federal Express. The letters informed the end users that some of the peelable sheaths/dilators in the Vaxcel Implantable Vascular Access kits may incorrectly contain 10F sheaths instead of 9F sheaths which had been supplied by B. Braun Medical Inc. Improper sizing of the sheath for the catheter may result in difficulty /inability to insert the catheters. The letters instructed the end users to immediately remove the affected lots from inventory, segregate the recalled product, contact Boston Scientific at 1-800-862-1284 to obtain a Return Authorization (RGA) number, pack the product in an appropriate shipping carton, affix the shipping label to the outside of the shipping carton, write the RGA number on the outside of the shipping carton, and to return the recalled product to Boston Scientific''s Distribution Center Attn: QA Returns, 500 Commander Shea Blvd., Quincy, MA 02171. Accompanying each letter was a Reply Verification Tracking Form to be completed by the end user and mailed or faxed to Boston Scientific even if the user did not have the recalled lots on hand. |  | 5 | 5 | NY, CA, MD, WA, TN. | Terminated | April 23, 2007 | 314 | 27 | Undetermined
Z-0744-2007 | 37254 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Mega Diagnostics Alcohol Standard (Cat.# A7504-STD) packaged in 5ml vials. |  |  | September 20, 2006 | 2006 | Mega Diagnostics | Potential for microorganism growth in the alcohol standard. | Other | The firm sent a recall notification to their customer in Indonesia and advised them to dispose of the product |  | 6 | 6 units | Country of Indonesia | Terminated | June 04, 2012 | 2084 | 18 | Undetermined
Z-0779-2007 | 37762 | No | GIR | Russell viper venom reagent. | Hematology | Hematology | 510(k) | American Diagnostica Inc. ACTICLOT dPT |  |  | February 28, 2007 | 2007 | American Diagnostica, Inc. | Stability of product not assured. An increase of clotting time for both normal and abnormal plasmas | Other | American Diagnostica notified accounts by letter on 2/28/07 and 3/02/07 and international distributors on 2/28/07. Users are requested to discontinue use and destroy. Kits will be replaced. |  | 249 | 249 kits | Worldwide, including USA, Canada, Germany, UK, Italy, Israel, Greece, and the Netherlands. | Terminated |  |  | 249 | Undetermined
Z-0694-2007 | 37687 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | BioHorizons 0.050 Hex Driver, Long, REF 300-351, BioHorizons Implant Systems, Inc., Birmingham, AL 35243, non-sterile, Rx only |  |  | March 12, 2007 | 2007 | BioHorizons Implant Systems Inc | Premature Wear; The material from which the hex driver was made was improperly hardened during manufacturing and could lead to premature wear. | Other | Letters and replacement drivers were sent to the consignees (both domestic and abroad) on 3/12/2007 explaining the problem and requesting that the consignee discard the instrument and use the replacement. They were delivered via United Parcel Service (UPS) 3rd day package delivery. |  | 436 | 436 units | Worldwide Distribution -- USA and countries of Canada, Norway, Australia, Spain, Brazil, Pakistan, Israel, Latvia, Japan, Columbia, Thailand, Turkey, Italy, South Korea, India, Taiwan, Ukraine, Russia, United Arab Emirates, Saudi Arabia, Hong Kong, Romania, UK, Kuwait, and Lithuania | Terminated | September 04, 2007 | 176 | 2732 | Undetermined
Z-0736-2007 | 37599 | No | KXT | Digoxin test system. | Toxicology | Toxicology | 510(k) | Roche COBAS Integra DIG, Digoxin Reagent for use on the COBAS INTEGRA models 400, 400 plus, 700 and 800 analyzers, as well as, Cobas c 501 analyzer; Catalog Number 20737836322. |  |  | March 16, 2007 | 2007 | Roche Diagnostics Corp. | The lower detection limit (LDL) may be a higher value than stated in the labeling. Note: To resolve this issue, the Digoxin LDL claim has been revised to 0.3 ng/mL for all reagent lots. | Other | A recall letter dated 3/16/07 was sent to each consignee advising them that the LDL claim in the label has been revised to 0.3 ng/ml for all reagent lots. |  | 4735 | 4,735 | Nationwide. | Terminated | June 12, 2008 | 454 | 4735 | Undetermined
Z-0780-2007 | 37510 | No | FWZ | Manual operating table and accessories and manual operating chair and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Doro Products Adapter, Product number 5979-300. The adaptor system is attached to OSI''s Spinal Frame, which is an operating room tabletop that is attached to a table base. The Frame is used to support patients during spinal surgery. |  |  | February 14, 2007 | 2007 | Orthopedic Systems Inc | Material used to manufacture accessory bracket does not meet product specification. The screw is normally composed of ''Torlon'', which is yellow. The part was made of "Tercite", which is blue. | Other | A notification letter was sent to all 8 consignees on 02/16/2007. The firm will send service representatives to review any devices on site and replace with the proper part if needed. |  | 8 | 8 units distributed. | The products were distributed to 8 consignees in the United States, to ID, IL, NY, and CA. | Terminated | December 06, 2007 | 295 | 8 | Undetermined
Z-0770-2007 | 37702 | No | FEB | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | 83 plus Mini-flex with accessory crate, Custom Ultrasonics, Inc. |  |  | December 12, 2005 | 2005 | Custom Ultrasonics, Inc. | The device was marketed without a 510k | PMA | On 12/12/05, the recalling firm telephoned their sales representatives to inform them of the problem. Letter were hand delivered to the hospitals by the sales representatives. |  | 4 | 4 units | Worldwide-The device was shipped to medical facilities in CA, GA, VA. and the UK. | Terminated | February 20, 2013 | 2627 | 4 | Undetermined
Z-0771-2007 | 37813 | No | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | Nemio XG, ultrasound system, SSA-590A with version 1.5 software, Toshiba American Medical Systems, Inc., Tusin, CA. |  |  | December 19, 2006 | 2006 | Toshiba American Med Sys Inc | Transducer Surface Temperature : It was found during testing that when using the system in the 2DCDI + PW mode, with the Scale Graph function active, that the transducer surface temperature could exceed the 27 degrees Celsius limit as specified in the IEC safety standard. | Other | Notification letters were sent to customers on December 19, 2206. Customers were instructed to avoid using Scale Graph function with live images, and instructed to call up the Scale Graph after images are frozen. Customers were informed that a Toshiba Customer Service Engineer will contact them to schedule a software upgrade when software becomes available. If customers have any questions they were instructed to contact their Toshiba Customer Service Engineer or call (800) 421-1968 and ask for the Regulatory Affairs Department. Customers were also instructed to distribute this information to all users of the product and also with the Bio-medical Department at their facility. |  | 37 | 37 | Nationwide Distribution including the states of CA, FL, GA, ID, MA, MN, NC, NY, PA, RI, TN, TX, WA, and WI. | Terminated | February 15, 2012 | 1884 | 37 | Undetermined
Z-0772-2007 | 37747 | walker, mechanical | ITJ | Mechanical walker. | Physical Medicine | Physical Medicine | 510(K) Exempt | LumexÂż Walkabout Lite Four-Wheel Rollator, Manufactured for: Graham-Field Health Products, Atlanta, Georgia 30360. |  |  | April 02, 2007 | 2007 | Graham Field Atlanta | The caster bolt stem/screw may loosen over time causing the wheel to come off. | Other | Graham-Field notified all of its Durable Medical Equipment (DME) dealers involved in the distribution of these products by certified US Mail on April 3, 2007. They were made aware of the need to adjust the product to lessen the risk of injury and were provided a re-work kit. The distribution of DME dealers includes the announcement of the field correction to the DME dealer, the letter to send to the end-users, rework supplies, rework instructions and proof of compliance cards. Graham-Field instructed the DME dealers to adjust any impacted products on site, as well as, any delivered by end-users, customers or other impacted third parties to the DME dealer for correction. DME dealers will be compensated for performing the corrective action. The installation of the rework kit involves the installation of lock washers and Locitite and is described in more detail in the attached rework instructions. Graham-Field does not have direct access to end user names, however it plans to reach end-users by financially supporting a certified mailing from DME dealers to their individual customers notifying them of this field correction. A follow-up 60 day certified mailing will be performed for non-responders. |  | 13300 | 13,300 units | U.S. Nationwide, Costa Rica, Ecuador, El Salvador, Guatemala, and Santo Domingo | Terminated | January 12, 2010 | 1016 | 13300 | Undetermined
Z-0730-2007 | 37628 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Flexxus&#153; Endoscopic Biliary Stent, 7.5 Fr |  |  | March 14, 2007 | 2007 | Conmed Endoscopic Technologies, Inc. | Blue safety clip incorrectly placed on the stent may prevent deployment | Other | ConMed Endoscopic Technologies notified consignes by letter dated 3/14/07 requesting them to return recalled devices. |  | 1 | 1 UNIT | Worldwide distribution, USA including states of DE, OH, MI, DC and countries of Belgium Japan. | Terminated | May 15, 2007 | 62 | 12 | Undetermined
Z-0781-2007 | 37819 | No | FWF | Surgical camera and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | SOPRO Camera Control Units, NTSC, SV420P; REF (Catalog) #72200505 v1.1; Made in France; Manufactured for Smith &amp; Nephew, Inc., Endoscopy, Andover, MA 01810. |  |  | January 19, 2007 | 2007 | Smith &amp; Nephew, Inc Endoscopy Division | Surgical camera control units were not designed and manufactured to pass the IEC 60101 testing for electrical fast transients (EFT) | Other | Recalling firm sent Recall notification letters to customers (Jan 2007) and distribution centers requesting device return. |  | 77 | 77 units. | Foreign Distribution Only-No U. S. A. distribution International distribution to: Argentina, Austria, Australia, Belgium, Brazil, China, Columbia, Denmark, Dubai, France, Germany, Greece, Italy, India, Japan, Korea, Malaysia, Netherlands, Norway, Portugal, South Africa, Spain, Sweden, Turkey and United Kingdom. | Terminated | September 02, 2008 | 592 | 102 | Undetermined
Z-0783-2007 | 37855 | No | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | B.Braun HyperFormer Pump Dispensing Set, Catalog Number V9900-02. The product is shipped in cases containing 10 sets. |  |  | April 11, 2007 | 2007 | B. Braun Medical, Inc. | The integrity of the sterile product barrier may be compromised | Other | An Urgent Product Removal Notice was issued to the distributors on 4/11/07. The notice informs the distributors of the problem, the need to conduct a sub recall and to return the product. |  | 180 | 180 units | Nationwide distribution --- including states of GA, MA, OH, and TX | Terminated | August 07, 2007 | 118 | 180 | Undetermined
Z-0698-2007 | 37403 | No | EQH | Ear, nose, and throat fiberoptic light source and carrier. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(K) Exempt | Boston Scientific SpyGlass Light Cable |  |  | February 07, 2007 | 2007 | Boston Scientific Corporation | Light Degradation in transmission-A film may deposit on the front face and limit transmission of light | Other | Boston Scientific notified accounts by letter dated 2/1/07 requesting return of product. |  | 13 | 13 units | Nationwide Distribution --- including states of CA, CO, FL, OH, MA, MN, MO NY, PA, SC,TX, and WA . | Terminated | November 15, 2007 | 281 | 152 | Undetermined
Z-0756-2007 | 37495 | No | MLN | Electrocardiograph electrode. | Cardiovascular | Cardiovascular | 510(K) Exempt | PADPRO ADULT Radiotranslucent Pad w/ZollÂż connector. Defibrillation/Pacing/Cardioversion/Monitoring Electrodes, REF/Catalog # 2516Z. RX ONLY. Made in USA for: CONMED CORPORATION Utica, New York, 13502. |  |  | March 08, 2007 | 2007 | Conmed Corporation | Firm was notified of three incidents over the past 15 months in which customers opened the pouch &amp; found an incorrect connector. This error makes the product unusable. The Zoll and Medtronic connectors are not interchangeable and are unique to their specific device. | Other | On 3/08/07, "URGENT DEVICE RECALL" letters (dated 3/05/07) will be sent via UPS to the US distributors and direct user accounts. The letters inform them of the problem and of the recall. The letters instruct them to immediately discontinue using the affected lots, inspect their stocks, and to return them as soon as possible. The letters instruct the consignees to notify their customers of the recall of the devices to the user level if they have been distributed, and request the return of the products. Additionally, when requested, ConMed will notify distributor's customers (sub-accounts) by the same means. Returned devices will be automatically replaced by ConMed Corporation. A separate letter is being sent to the foreign accounts today instructing them to return the devices back to ConMed Corporation for final disposition. The letters also instruct the consignees to complete the accompanying reply form and fax it to Nancy Crisino at (315) 624-3089. ConMed has contracted Stericycle in Indianapolis, IN, to handle the recall and returns from the domestic accounts. |  | 9712 | 9712 pairs | Worldwide, including USA, Canada, Romania, Chile, and Belgium | Terminated | July 16, 2012 | 1957 | 10632 | Undetermined
Z-0556-2007 | 37322 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | OEC 9800 Fluoroscopic X-ray System, Model Number A349855, GE Healthcare, Surgery. |  |  | February 09, 2007 | 2007 | GE OEC Medical Systems, Inc | During routine service, a cover may have been installed without the required proper lead shielding. | Other | Consignees were notified by telephone on 02/09/2007 and told to discontinue use of the units until they could be inspected by GE personnel. GE personnel visited all 8 sites on 02/12/2007. |  | 11 | 11 units of all models | Nationwide to AL, AZ, IA, OH, TX, VA, and WV. | Terminated | June 26, 2007 | 137 | 11 | Undetermined
Z-0497-2007 | 37150 | No | LFD |  |  | Dental | 510(k) | Salivart Oral Moisturizer, packaged in 2.5 fluid ounce (73.9 mL) spray cans, Part Number: 0386-0009-75; Manufactured by Gebauer Company, Cleveland, Ohio |  |  | January 19, 2007 | 2007 | Gebauer Company | Microbial Contamination-certain lots of product failed USP &lt;61&gt; Microbial Limits Testing for total aerobic count during 6 month stability testing. | Other | The recalling firm sent recall letters dated 1/23/2007 to their Distributor Customers and Non-distributor customers who received Salivart product from Lot Code: 06AA003. both letters include an attached recall response form which the customers are instructed to destroy the product, complete a form and fax back it to the recalling firm. The response forms provide verification that the form has been received by the customer(s). With the initiation of the RECALL EXPANSION for more lot codes to be recalled on 2/14/2007, the firm issued new recall letters dated, 2/13/2007, 2/15/2007 and 2/22/2007. These four letters provide the same basic information and also have attached recall response cards, however, they list the additional 6 lot codes subject to recall. The firm issued a Press Release concerning this recall through Cleveland (Business Wire) on 3/1/2007. |  | 34545 | 34,545 cans | The product was distributed throughout the U.S. to wholesalers, retail accounts (i.e. pharmacies, dental offices, and medical facilities) and full size samples of product were given away at trade shows to medical, and dental offices and some end users consumers). A total of 240 units of product derived from Lot Code: 06AA001, were distributed to a single Canadian distributor under the assigned Lot Code: 06BB002 | Terminated | December 17, 2007 | 332 | 34545 | Undetermined
Z-0641-2007 | 37477 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart FR2+ Defibrillator |  |  | February 28, 2007 | 2007 | Philips Medical Systems | Failure of tactile switches on one of two printed circuit board assemblies (PCBA) used in the FR2+ device. Potential delay or prevention of shock treatment due to PCBA board contamination, causing intermittent switch operation. | Nonconforming Material/Component | On 02/28/2007 the firm sent letters via return receipt to all domestic and foreign customers. The letter advises the customer of the issue and if the user encounters the problem, to press the button again with additional force. In addition, the firm sent with the letter a new Quick Reference Card. Customers are instructed to replace the old Quick Reference Card with the new one which instructs them to press 'On' and 'Shock' buttons firmly. |  | 9099 | 9099 units | Distribution is worldwide. Devices are used by first responders (EMT, EMS, Fire and Police personnel, etc.), and by trained personnel on commercial airlines and business offices. | Terminated | November 13, 2013 | 2450 | 9099 | Undetermined
Z-0645-2007 | 37390 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ADVIA Centaur CP System, Catalog/Part Number 086-A001. Automated Immunoassay Analyzer. --- Classification of the Device: Under 21 CFR, Part 862.2160, the ADVIA Centaur CP is a discrete photometric chemistry analyzer for clinical use and is classified as a Class I (general control) device. --- The ADVIA Centaur CP is part of Bayer's Immunochemistry analyzer family, which also includes the ADVIA Centaur and the ACS:180. Based on FDA guidance for deciding when regulatory filings are warranted for device modifications, as well as, the FDA Instrument Family and Replacement Reagent Policy Guidance, Bayer determined that the changes associated with the ADVIA Centaur CP did not affect the safety and effectiveness of the predicate device (ACS:180) and therefore no device modification 510(k) was warranted. The ADVIA Centaur CP instrument uses the same reagents as the ADVIA Centaur. The ACS:180 was cleared by FDA under 510(k) # K902336, and the ADVIA Centaur was cleared by FDA under 510(k) # K971418. The ADVIA Centaur CP system was released for sale in November 2005 and distribution of these systems began at that time. All ADVIA Centaur CP Systems are impacted. There is no expiration date for these systems. |  |  | November 09, 2006 | 2006 | Siemens Healthcare Diagnostics | Firm received complaints for signal 2 &amp; 4 errors, and shifts in Relative Light Units (RLU). Investigations showed that the On-Board Stability (OBS) of the acid &amp; base reagents may be compromised after 4 days after installation onto the Centaur CP system due to evaporation. | Other | A Support Bulletin was sent by e-mail to all affected Bayer HealthCare LLC Branches on 11/09/06 in order to inform them of the issue and to provide them with instructions for this field correction. Two Customer Bulletins were prepared and sent by e-mail to Bayer HealthCare LLC Branches on 11/09/06 both in the US and outside the US for communication with the affected customers. The first bulletin advised customers of the problem and the corrective action. A Confirmation (Fax-back) Form was included to ensure that all customers worldwide have been notified and have complied with the mandatory action. The second Customer Bulletin describes an optional alternative mitigation, which requires a Bayer representative to go to the customer site to implement. |  | 612 | 306 units: 101- Domestic, 205 - Foreign | Worldwide | Terminated | July 30, 2012 | 2090 | 612 | Undetermined
Z-0646-2007 | 37473 | No | JLW | Thyroid stimulating hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Thyrotest TSH POC Rapid membrane test. In vitro diagnostic. |  |  | January 24, 2007 | 2007 | ThyroTec Inc | Positive control showed little or no positive line. | Other | On 1/24/07 the recalling firm notified their only customer of the problem. On 2/5/07, the distributor issued recall letters to their customers informing of the problem and the need to destroy the kits. The kits from lot number 02106-06 were returned to the recalling firm. |  | 484 | 484 units | Nationwide. | Terminated | June 04, 2007 | 131 | 484 | Undetermined
Z-0799-2007 | 37064 | [6 empty] | Maxi Move Patient Lift with Scale; Manufactured by Arjo Hospital Equipment AB, Eslov, Sweden; Model KMBB4OSU2FUS |  |  | January 05, 2007 | 2007 | Arjo, Inc. | The hanger bar assembly may detach from the Maxi Move lifters with scale, due to an incorrect torque setting on the two bolts holding the scale in place. | Other | Arjo sent Urgent Device Field Correction Recall Notice letters dated 1/5/07 to the end user accounts who received the affected lifts, advising them of the potential for the hanger bar assembly to detach. They were informed that an Arjo Field Service Technician would call them within five business days to schedule the time for him to visit and perform the inspection and correction to the lifts involved. If they did not hear from the technician within that time, they were instructed to contact the Arjo Quality Department at 800-323-1245, ext. 6140. The accounts were also requested to complete and return the enclosed customer response form, acknowledging the receipt and understanding of the letter. |  | 121 | 121 lifts | Nationwide | Terminated | September 02, 2008 | 606 | 284 | Undetermined
Z-0812-2007 | 37868 | No | GKP | Coagulation instrument. | Hematology | Hematology | 510(k) | STA Compact and STA Compact CT |  |  | January 18, 2007 | 2007 | Diagnostica Stago, Inc. | Result of complaints of field failures of the silicon bridge rectifier, resulting in smoke billowing out of the unit and burnt circuit boards, wires and connectors. | Other | A recall notification has not been sent. The firm did issue a technical bulletin to their representatives.--Update: based on FDA input the firm issued a Recall Letter to all users dated May 25, 2007 |  | 693 | 693 units | Product was distributed to hospitals and laboratories nationwide. | Terminated | December 10, 2007 | 326 | 693 | Undetermined
Z-0643-2007 | 37133 | No | MDB | Microbial growth monitor. | Microbiology | Microbiology | 510(k) | BacT/ALERTÂż FA Culture Bottles, 30 ml, for in vitro diagnostic use, BioMerieux, Inc., Durham, NC 27704 |  |  | January 11, 2007 | 2007 | bioMerieux, Inc. | Bacillus sp. contamination was detected in inoculated and uninoculated bottles of BacT/ALERT FA Culture bottles. | Other | Consignees were notified by letter on/about 01/12/2007. |  | 2136 | 2136 cases (100 bottles per case) | Worldwide, including USA, Austria, Denmark, Finland, France, Greece, Guyana, Italy, Libya, Netherlands, Norway, New Caledonia, Poland, Portugal, Singapore, Slovenia, Spain, Tanzania, and Sweden. | Terminated | December 08, 2009 | 1062 | 2136 | Undetermined
Z-0621-2007 | 37423 | No | JSI | Selective culture medium. | Microbiology | Microbiology | 510(K) Exempt | Remel MacConkey Agar w/Sorbitol, growth medium, Catalog R01556, packaged 10/box. The firm name on the label is Remel, Lenexa, KS. |  |  | February 21, 2007 | 2007 | Remel, Inc. | False negative results for sorbitol fermenting bacteria such as E. Coli 0157. | Other | Recall letters dated 2/19/07 were issued on 2/21/07 via regular mail explaining the reason for recall and requesting all remaining inventory of the affected lot be discarded. A Product Inventory Checklist was enclosed to indicate the amount of inventory the customer discarded. This was to be returned via FAX or mail in the return envelope provided. |  | 694 | 694/10-pack boxes | Nationwide. | Terminated | March 06, 2008 | 379 | 694 | Undetermined
Z-0644-2007 | 37141 | No | MDB | Microbial growth monitor. | Microbiology | Microbiology | 510(k) | BacT/VIEW C.30a Software update (P/N 514515-1) diagnostic data management system, bioMerieux, Inc., P. O. Box 15969, Durham, NC 27704-0969 USA |  |  | December 19, 2006 | 2006 | bioMerieux, Inc. | Bottle data is not sent to BacT/ALERT 3D instrument. | Other | A customer notification advising customers of the issues and the precautions were sent on/about December 20, 2006. |  | 1233 | 1233 disks | Worldwide, including USA, Canada, UK, Italy, France, Spain, Germany, Poland, Chile, Japan and the Netherlands. | Terminated | December 08, 2009 | 1085 | 1233 | Undetermined
Z-0626-2007 | 37106 | No | IRT | Powered heating pad. | Physical Medicine | Physical Medicine | 510(K) Exempt | HoMedics Thera P Model King size Moist/Dry Heating Pad with Auto Shut-off; Model No. HP-500 (UPC 031262003652). |  |  | February 09, 2007 | 2007 | HoMedics, Inc. | Heating pads may have a loose electrical connection, which could result in excessive heat being generated that poses a risk of patient burns, fire, and property damage. | Other | Press release issued on 2/9/07 instructing consumers to return the product to store of purchase or Homedics. Retail stores notified via recall letter issued on 2/9/07. Retail stores are being instructed to accept returns and to post placards to notify customers of the recall. |  | 292108 | 292,108 of all models | Nationwide and Canada. | Terminated | January 04, 2010 | 1060 | 292108 | Undetermined
Z-0642-2007 | 36979 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific Medi-Tech(R) Katzen (TM) Infusion Wire, for the infusion of therapeutic agents (i.e. heparin, saline, thrombolytic agents, etc.) in the peripheral vasculature. The wire allows for the delivery of agents in either a 'pulse-spray' or 'slow weep' technique. Order No. REF: 46-193 (M001461930). |  |  | December 12, 2006 | 2006 | Boston Scientific | Boston Scientific is voluntarily recalling one lot/batch of Katzen Infusion Wires because they have identified that the label on the carton may indicate a different length device than what is actually in the carton. The affected batch was manufactured using the Katzen Core Assembly of 146 cm, instead of a 177cm assembly. | Other | An Urgent Medical Device Recall letter, dated December 12, 2006 was sent to affected hospitals Risk Manager beginning 12/12/2006. The letter describes the issue and product affected, requests that further distribution or use of any remaining product cease immediately, and provides information for returning product to Boston Scientific Replacement product will be issued for all recalled product returned from affected facilities. A Reply Verification Tracking Form is requested to be completed and returned to Boston Scientific by Wednesday, December 27, 2006. |  | 10 | 10 units | US only to: FL, GA, IN, NY. | Terminated | July 20, 2007 | 220 | 10 | Undetermined
Z-0638-2007 | 37318 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | VIASYS VELA Ventilator, Catalog Number-16186; Viasys Respiratory Care, Inc. |  |  | January 31, 2007 | 2007 | Viasys Respiratory Care, Inc.dba Bird Products | The graphical user interface may become slow in its response and in some cases non-responsive. Under some circumstances this situation may cause the ventilator to delay the annunciation of alarms. | Other | On January 31, 2007, a product notification letter was mailed to all customers in the United States via overnight carrier (Federal Express). Consignees are informed in the notification letter that service calls are in process of being scheduled to update domestic consignees. On January 31, 2007, a product notification kit (including a product notification letter, download software, download instruction sheets and installation verification fax-back sheets) were mailed to all international customers via register mail. Consignees are informed in the notification letter that replacement software, replacement instructions, and installation verification forms are included in their recall package. Domestic Consignees will have product corrected by VIASYS service personnel. Self address postage paid envelopes will be provided with notifications to aid in effectiveness checks |  | 878 | 878 | Worldwide Distribution ---USA and countries of Bangladesh, Bosnia, Brazil, China, Czech Republic, Ecuador, Estonia, Finland, France, Guatemala, Hong Kong, India, Israel, Italy, Japan, Latvia, Lithuania, Malaysia, Morocco, Netherlands, Portugal, Russian Federation, Taiwan, Thailand, Turkey &amp; Vietnam. | Terminated | September 18, 2008 | 596 | 878 | Undetermined
Z-0620-2007 | 37362 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Baxter Vented Paclitaxel Set; a sterile fluid pathway with polyethylene lined tubing, non-DEHP pump segment, 0.22 micron downstream high pressure filter, injection site, male Luer lock adapter; for use with rigid nonvented solution containers on Colleague and Flo-Gard 6201 and 6301 series pumps; Baxter Healthcare, Deerfield, IL 60015; product code 2C7553 |  |  | February 16, 2007 | 2007 | Baxter Healthcare Corp. | Baxter has received reports of disconnections/leaks during use of vented Paclitaxel administration sets. | Other | Baxter sent the 2/16/07 Urgent Product Recall letter to affected customers, to the attention of the Director of Nursing with copies to the Materials Manager, via first class mail to advise them of the potential disconnections/leaks during use of the Vented Paclitaxel Administration Sets. The accounts were instructed to immediately discontinue use and segregate the listed affected lot numbers, and contact Baxter Center for Service at 1-888-229-0001 to arrange for the return and credit for the affected product. Any questions were directed to the Center for One Baxter at 1-800-422-9837. |  | 200309 | 200,309 units | Worldwide, including Switzerland, Australia, Canada, Hong Kong, India, Korea, Mexico, Malaysia, New Zealand, Panama, Phillipines, Singapore, Thailand and Taiwan. | Terminated | March 17, 2008 | 395 | 215886 | Undetermined
Z-0632-2007 | 36872 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Flo-Gard 6201 Volumetric Infusion Pump, catalog #2M8063, 2M8063R (refurbished); Baxter Healthcare Corporation, Deerfield, IL 60015 |  |  | November 27, 2006 | 2006 | Baxter Healthcare Corp. | There is the potential for non-detection of, and no alarm for, upstream occlusion for Flo-Gard Volumetric Infusion Pumps, if an occlusion occurs above a flexible chamber such as a drip chamber, Buretrol or blood filter. | Other | Baxter sent the 11/27/06 Urgent Device Correction letter to all Flo-Gard Volumetric Infusion Pump customers, to the attention of the Vice President of Nursing and the Biomedical Engineer, via first class mail to alert them to the potential for non-detection of, and no alarm for, upstream occlusion under specific conditions for Flo-Gard Volumetric Infusion Pumps if the occlusion occurs above a flexible chamber (such as a drip chamber, Buretrol or blood filter), and to provide the addendum to the operator manual to address these specific conditions. The letters included the mitigations and risks associated with each of the conditions that could potentially cause this situation. This information was included in the letter, as well as the Operator''s Manual Addendum, which Baxter requested to attach a copy to all Flo-Gard Operator''s Manuals in their facility. Baxter instructed the customers to share this information with all staff members, services and facilities that use Flo-Gard pumps. Any questions were directed to Baxter''s Global Technical Services at 1-800-843-7867, prompt 2. Although Baxter previously notified customers on 3/5/99 that they would no longer support or service the Flo-Gard 6200 and 6300 devices as of 12/1/99, these product codes were included in this letter due to a similar potential for this issue. Should the customer still have these 6200 or 6300 devices in their possession, Baxter recommended that they be taken out of service due to obsolescence of the appropriate components. The accounts were requested complete the enclosed reply form acknowledging receipt of the letter and dissemination of the information to their staff. |  | 177468 | 177,468 pumps | Worldwide including USA, Puerto Rico, Argentina, Australia, Bahrain, Belgium, Bermuda, Brazil, Canada, Chile, China, Colombia, Costa Rica, Denmark, Ecuador, Egypt, El Salvador, Guatemala, Hong Kong, Honduras, India, Indonesia, Ireland, Italy, Korea, Lebanon, Malaysia, Mexico, Netherlands, New Zealand, Nicaragua, Norway, Oman, Panama, Peru, Philippines, Poland, Russia, Singapore, Slovakia, Slovenia, South Africa, Sweden, Taiwan, Thailand, Turkey, United Kingdom and Venezuela. | Terminated | January 19, 2010 | 1149 | 219566 | Undetermined
Z-0650-2007 | 37069 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | NexGen Complete Knee Solution Minimally Invasive Solutions MIS Technology Femoral Finishing Guide, instrumentation used in total knee arthroplasty procedures. Catalog No. 82015272521. |  |  | October 10, 2006 | 2006 | Zimmer Inc. | The device was cut to an incorrect angle, which may result in an incorrect bone cut. | Other | Product was removed from customer location during visit on 10/10/2006. |  | 1 | 1 | One medical facility in Michigan. | Terminated | December 04, 2007 | 420 | 1 | Undetermined
Z-0816-2007 | 37838 | [6 empty] | Cranex Tome panoramic x-ray unit with Spiral Tomography for cross-sectional imaging. manufactured by SOREDEX, PO Box 250, 00031 Helsinki, Finland. |  |  | April 28, 2006 | 2006 | Instrumentarium Dental, Inc. | There may be a material defect in the up/down movement axle of the unit. This may cause the upper part of the unit to suddenly come down. | Other | A recall letter, dated 04/28/2006, was sent to affected customers . The letter describes the issue, and requests that no patient operation is permitted until the unit is checked by authorized service technician. If the service technician noted wear of the screw, the assembly was replaced. |  | 9 | 9 | Worldwide, including USA, Canada, Bermuda, and Colombia. | Terminated | February 24, 2008 | 667 | 9 | Undetermined
Z-0640-2007 | 37444 | No | GIF | Blood cell diluent. | Hematology | Hematology | 510(K) Exempt | CELL-DYN 4000 Diluent/Sheath Reagent, List Number 01H73-01, Distributed and Manufactured by: |  |  | February 24, 2004 | 2004 | Abbott Laboratories | Incorrect Test Results-product may show a higher than expected platelet background count when used on the CELL-DYN 4000 System and report patient results that are unacceptable (out-of-range). | Other | Recall Letter-On 2/24/04, the firm initiated the recall and its notification explaining the reason for the recall and requesting the product be destroyed |  | 9800 | 9,800 units | Worldwide Distribution-USA including states of AL, AR, AZ, CA, CO, CT, DC, FL, GA, lA, IL, IN, KS, KY, LA, MA, MI, MN, MO, MS, NC, ND, NJ, NV, NY, OH, OK, OR, PA, SC, TN, TX, VA, WI, and WV. and countries of Australia, Argentina, Brazil, Hong Kong, Japan, Korea, Malaysia, Singapore. | Terminated | May 01, 2007 | 1162 | 9800 | Undetermined
Z-0815-2007 | 37653 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE Healthcare-Americas Signa 3.0T EXCITE Whole Body Magnetic Resonance System |  |  | June 01, 2006 | 2006 | GE Medical Systems LLC | To provide the 3.0T MRI systems users with the 8-channel brain coil the proper coil weighting software. The proper coil weighting factors for the eight-channel brain coil were not included in the software. | Other | The corrective action is already completed and all customers have been notified via a personal visit to each site, to install a software patch beginning June, 2006. |  | 30 | 30 units | Worldwide, including USA, China, Finland, Germany, Israel, Italy, Korea, Spain, and Taiwan. | Terminated | February 24, 2008 | 633 | 30 | Undetermined
Z-0636-2007 | 36685 | [6 empty] | Brilliance CT System, Big Bore Configuration, Model #728243. (version 2.2.1) |  |  | October 02, 2006 | 2006 | Philips Medical Systems (Cleveland) Inc | Two conditions have been identified with the Philips Medical System Brilliance Bore ,version 2.2.1 Tumor Localization, that may lead the user to misinterpret the displayed data. | Other | Customer advisory notification letters dated 10/4/06 and 10/25/06 were sent to consignees providing short term work around instructions. Software modifications installed by the recalling firm will follow as the long term correction. |  | 17 | 17 units | Class II Recall-Nationwide Distribution-USA including states of KY, MD, TX, RI, NJ, MA, LA, VA, PA, SC, CA, OH, and FL. | Terminated | March 26, 2012 | 2002 | 17 | Undetermined
Z-0639-2007 | 37450 | No | GIF | Blood cell diluent. | Hematology | Hematology | 510(K) Exempt | CELL-DYN 4000 Diluent/Sheath Reagent, |  |  | May 12, 2004 | 2004 | Abbott Laboratories | Incorrect Test Results-test results may show a higher than expected platelet background count when used on the CELL-DYN 4000 System and report patient results that are unacceptable (out-of-range). | Other | On May 12, 2004, the firm initiated the recall and its notification was via letters explaining the reason for the recall and requesting the product be disposed. |  | 6893 | 6,893 units | Class II- Worldwide Distribution-USA including the states of AL, CA, CO, CT, DE, FL, GA, IL, IN, KY, MA, MD, MI, MN, MO, MS, NC, NE, NH, NJ, NY, OR, PA, SC, TN, TX, VA, WA, WI, and WV and countries of Argentina, Brazil, Hong Kong, Japan, Malaysia, Singapore, Taiwan, and Thailand. | Terminated | May 01, 2007 | 1084 | 6893 | Undetermined
Z-0612-2007 | 37272 | No | MWS | Cardiovascular surgical instruments. | Cardiovascular | Cardiovascular | 510(K) Exempt | Guidant Acrobat Vacuum Stabilizer, Model number-OM-9100 |  |  | December 01, 2006 | 2006 | Guidant Cardiac Surgery | Vacuum Tubing Failure- vacuum tubing fails before the labeled two year shelf life. This may result in partial or complete vacuum loss during surgery. | Other | The firm has issued by FedEx letters dated 12/1/2006, to all consignees who received the product in the US, and has also informed their foreign distributors in the same manner on December 4, 2006. The firm''s field representatives were notified via teleconference on December 1, 2006. A hard copy of the notice was delivered to field reps on December 4 as well. |  | 57691 | 57,691 units distributed total | Product was distributed to a total of 529 consignees, 482 domestic and 47 international. Product was sent Nationwide, Europe and Asia. | Terminated | December 04, 2007 | 368 | 57691 | Undetermined
Z-0814-2007 | 37136 | No | MDB | Microbial growth monitor. | Microbiology | Microbiology | 510(k) | BacT/VIEWÂż Software, all versions. in vitro diagnostic database software. |  |  | December 19, 2006 | 2006 | bioMerieux, Inc. | BacT/VIEW will be unable to run scheduled macro applications unless the user touches the BacT/VIEW touch screen or clicks a function after resynchronizing the clocks. | Other | A customer notification was sent on/about December 19, 2006. Customers will be sent a new BacT/VIEW software version C.30b QTR 1 2007. |  | 1233 | 1233 units | Worldwide, including USA, Canada, UK, Italy, France, Spain, Germany, Poland, Chile, Japan and the Netherlands. | Terminated | December 08, 2009 | 1085 | 1233 | Undetermined
Z-0627-2007 | 37422 | No | KMK | Intravascular catheter securement device. | General Hospital | General Hospital | 510(K) Exempt | StayFix Fixation Device for Percutaneous catheters, REF 680ME, (small) 5F - 14F, Merit Medical, Unomedical, Ltd., Great Britain. |  |  | February 15, 2007 | 2007 | Merit Medical Systems, Inc | Sterility of some units may be compromised due to damaged packaging. | Other | Consignees were notified by letter on 02/23/2007. They were instructed to stop use of and return all affected product to Merit. |  | 195001 | 195,001 units | Worldwide, including USA, France, Germany, and United Kingdom. | Terminated | November 09, 2011 | 1728 | 250808 | Undetermined
Z-0817-2007 | 37852 | ph buffer | JCC | General purpose reagent. | Pathology | Pathology | 510(K) Exempt | Medtronic Buffer Solution pH 1.07 500ml |  |  | April 06, 2007 | 2007 | Medtronic Neuromodulation | Medtronic pH 1.07 buffer solution lot #2610596 was mislableled. | Other | Customers were sent a recall notice, a customer detail sheet for this product lot, and a reply form to be used when returning the suspect bottles. Distributor was provided with a formatted templates to be used in recalling the supsect bottles from their customers and instructions on how to manage the returns and ship the bottles back to Medtronic. |  | 238 | 238 | Nationwide. | Terminated | March 02, 2008 | 331 | 238 | Undetermined
Z-0637-2007 | 37102 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | COOK Triple Lumen Central Venous Catheter Tray with Cook SpectrumÂż GlideÂż Antimicrobial Catheter with EZ--PassÂż Hydrophilic Coating: Reorder Number C-UTLMY-701J-RSC-ABRM-HC-FST-A. |  |  | January 02, 2007 | 2007 | Cook, Inc. | These catheters do not include the proximal sidehole as required by the specification. | Other | Consignees were notified by phone on 1/2/07 and instructed to return the product. |  | 100 | 100 | Nationwide, including the states of Florida, Louisiana, Maine, Massachusetts, Minnesota, Missouri, New Jersey and North Carolina. | Terminated | December 05, 2007 | 337 | 100 | Undetermined
Z-0631-2007 | 37442 | No | LJT | Subcutaneous, implanted, intravascular infusion port and catheter. | General Hospital | General Hospital | 510(k) | Vaxcel Implantable Vascular Access System - Titanium Mini-Port with PASV Valve and 8F Silicone Catheter, REF/Catalog No. 45-220, UPN/Material No. M001452200, |  |  | September 27, 2006 | 2006 | Boston Scientific Corp | Sterile barrier may have been compromised. Firm received one complaint which demonstrated that the inner and outer seals of the port tray packaging had been compromised. | Other | On 9/27/06, Boston Scientific Corp. sent Urgent Medical Device Recall - Immediate Action Required letters (dated 9/27/06) via Federal Express to the 84 end users (medical facilities). The letters informed the users of the recall of the products due to a compromised sterile barrier. The letters instructed the users to immediately remove the affected lots from inventory, segregate them, contact Boston Scientific at 1-800-862-1284 to obtain a Return Authorization (RGA) number, return the product in an appropriate shipping carton, place the label on the outside of the shipping box, write the RGA # on the outside of the shipping carton, and return the recalled product to Boston Scientific Corp. Distribution Center, Attn: QA Returns, 500 Commander Shea Blvd., Quincy, MA 02171. Accompanying each letter was a Reply Verification Tracking Form to be completed by the user, and mailed or faxed back to Boston Scientific. |  | 42 | Domestic - 42 | Nationwide | Terminated | November 26, 2007 | 425 | 236 | Undetermined
Z-0821-2007 | 37848 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Definium AMX 700 Digital mobile x-ray system, GE Healthcare 3000 North Grandview Waukesha, WI |  |  | December 21, 2006 | 2006 | GE Medical Systems LLC | When using the Tech Switch on exposures &gt;750ms, &gt;100 kVp and &gt; 100 mAs the exposure terminates prematurely. | Other | Visit to each consignee beginning December 21, 2006. |  | 6 | 6 units | Product was distributed in the following states: MS, NE, NJ, and, TX | Terminated | February 24, 2008 | 430 | 6 | Undetermined
Z-0655-2007 | 37366 | No | DTW | Cardiopulmonary bypass level sensing monitor and/or control. | Cardiovascular | Cardiovascular | 510(K) Exempt | Terumo Perfusion System 8000 Base, 4 pump, 115v; Model 8000; Catalog No. 16400. |  |  | January 12, 2007 | 2007 | Terumo Cardiovascular Systems Corp | The hand crank may crack in the bearing/shaft area due to a manufacturing defect. (Note: Some of the hand cranks subject to this recall were shipped as replacements for hand cranks recalled in the Summer of 2006.) | Other | Consignees were notified by phone between 1/12/07 and 2/23/07 and informed replacement hand cranks would be shipped. |  | 7 | 7 | Worldwide, including USA, Belgium, Hong Kong, Japan, Taiwan and United Arab Emirates. | Terminated | June 19, 2008 | 524 | 83 | Undetermined
Z-0654-2007 | 37376 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Colorado 2 Spinal System, Anterior Plate, REF8619500, for export only, manufactured at: Warsaw, Indiana USA, Rx only |  |  | February 02, 2007 | 2007 | Medtronic Sofamor Danek USA Inc | Bone screws may pull though the plate. | Other | Medtronic Global Quality initiated the recall via letter on 02/02/2007. Sub-recalls will be managed by Medtronic EMEA quality (Europe, Middle East, and Africa) in the affected geographies. Distributors and surgical facilities will be contacted as well as regulatory bodies in each country. |  | 293 | 293 units | France, Germany, Italy, UK, and Tunisia. No US distribution. | Terminated | August 11, 2007 | 190 | 1700 | Undetermined
Z-0670-2007 | 37431 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Intuitive brand Da Vinci S Surgical System, an Endoscopic Instrument Control System, Model SS2000 Patient Side Cart, Model PS2000, Part Number: 380267-01, IS2000, Top Level, 4-Arm; Distributed and Manufactured by: Intuitive Surgical, Inc. 950 Kifer Road Sunnyvale, CA |  |  | April 12, 2006 | 2006 | Intuitive Surgical, Inc. | Da Vinci S Surgical System's spinal pin could limit mechanical motion of the arm and render the system unavailable for surgical use, and may result in system failure and the need to convert to an alternative surgical technique. | Other | The firm did not release notification to its consignees. The reworks were arranged as part of routine service visits initiated on 4/12/06. |  | 24 | 24 units | Nationwide. | Terminated | December 05, 2007 | 602 | 24 | Undetermined
Z-0807-2007 | 37756 | No | GAT | Nonabsorbable poly(ethylene terephthalate) surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Product is an EO Sterile, Polyester braided and coated nonabsorbable surgical suture with double needle, size 2/0 (3.0 Metric), NSH 26 mm TAPER Needle. Product is wrapped in a five-fold card with inserted flap ends (overwrap). A label is placed on the overwrap and the packet is then inserted into a tyvek/mylar pouch. One end of the overwrap card is printed with 'Open'. This end is positioned in the pouch so it is closest to the peel-open end. No suture is exposed or visible through the overwrap which must be torn to have access to the suture and needle. Product is boxed (12 to a box) and shrink wrapped. |  |  | April 03, 2007 | 2007 | C P Medical | Potential for sterility of product to be compromised. Suture package seals incomplete or missing. | Other | On 04/03/2007 the firm notified its Polyester product (Polybond) customer by telephone and followed up on 04/09/2007 with a letter via certified mail requesting the distributor to discontinue sales of the affected product lot, to notify their customers of the issue, and return stock to CP Medical. The firm provided the distributor with a letter dated 04/09/2007 to be provided to the distributor's customers requested those customers to discontinue use of the affected product lots and return stock to their distributor who in turn are to return product to CP Medical.On 04/20/2007 the firm notified its Polypropylene and Polydioxanone product customers by letter dated 04/17/2007 via certified mail requesting distributors to discontinue sales of the affected product lot, to notify their customers of the issue, and return stock to CP Medical. The firm provided the distributor with a letter dated 04/18/2007 to be provided to the distributor's customers requested those customers to discontinue use of the affected product lots and return stock to their distributor who in turn are to return product to CP Medical. |  | 14641 | 121 boxes (12/box) Lot I0625120, and 121 boxes (12/box) Lot I0627120-01 | Firm distributes products to human and veterinary medical device distribution centers worldwide. Polybond Polyester sutures, product CP523A, were distributed to one distribution center in South Africa.Polypropelene sutures, product 8665, were distributed to veterinary distribution centers in OH, OR, and NE for nationwide distribution.Polydioxanone sutures, product M398A, were distributed to human and veterinary distribution centers in AZ, IA, OH, LA, MA, MI, MO, NE, and TX for nationwide distribution. | Terminated | August 01, 2008 | 486 | 29569 | Undetermined
Z-0704-2007 | 31058 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | InstaTrak 3500 Plus system with cranial application (aid to the surgeon for locating anatomical structures in the human body) |  |  | January 26, 2005 | 2005 | GE OEC Medical Systems,Inc | Tracking inaccuracy of 5-10 mm may result when the shunt guide accessory is used with version 5.1 software. | Other | Consignees were visited by GE personnel to upgrade the software, beginning in April, 2004, to version 5.2 |  | 91 | 91 systems (version 5.1 software) | Worldwide distribution --- USA including states of AK, AZ, AL, CA, CT, IA, IL, MD, ME, MO, MT, MS, NC, NJ, NY, OH, PA, RI, SD, UT, TX, VA, WA and WI and countries of France, United Kingdom, Japan, Italy, India and Canada | Terminated | July 08, 2011 | 2354 | 91 | Undetermined
Z-0813-2007 | 37871 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | BD 5ml Syringe Luer-Lok Tip |  |  | April 09, 2007 | 2007 | Becton Dickinson &amp; Company | Small cracks in plastic trays of one lot of BD 5mL Syringe Luer Lok Tip, Bulk Sterile Convenience Paks that may adversely affect tray integrity and steriity. | Other | Recall communications were sent on April 2, 2007 and May 3, 2007 by UPS second day air to all direct consignees and potential customers. |  | 174300 | 174,300 units | Nationwide. | Terminated | August 09, 2007 | 122 | 174300 | Undetermined
Z-0800-2007 | 37635 | No | EBF | Tooth shade resin material. | Dental | Dental | 510(k) | 3M ESPE Rely X Veneer Cement Refill Syringes, A3 Opaque/Yellow Opaque Shade, manufactured by 3M. ESPE is a radiopaque , color stable light-cured resin cement indicated for bonding veneers fabricated of porcelainor composite. |  |  | February 12, 2007 | 2007 | 3M Espe Dental Products | Cement in 3M ESPE RelyX Veneer Cement Refill A3 Shade syringes, may not cure to the degree expected. | Other | All consignees including distributors were notified of the recall via telephone or mail on 2/27/2007. Communication with distributors included the product being recalled, the problem , recommended stop usage or distribution and directions for return shippment. Further, the consignees were also sent a letter which described the product and problem and what to do if the product was still availabe or had been used. |  | 99 | 99 | Worldwide distribution ---- USA including states of IN, PA, NY, CA, NV, AZ and country of Malaysia. | Terminated | February 24, 2008 | 377 | 99 | Undetermined
Z-0801-2007 | 37893 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Biohorizons 3.0mm Overdenture Implant System, REF 30120S2, (3 mm x12 mm-2mm collar), Biohorizons Implant Systems, Inc., Birmingham, AL 35243 |  |  | April 18, 2007 | 2007 | BioHorizons Implant Systems Inc | Dental implant system is missing ball attachment system. | Other | Letters explaining the problem and MBAS systems were sent via United Parcel Service (UPS) 3rd day package delivery to the consignees (both domestic and abroad) on 04/18/2007. |  | 34 | 34 units | Worldwide distribution ---- USA and country of Italy | Terminated | August 11, 2007 | 115 | 34 | Undetermined
Z-0665-2007 | 37568 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Cook ATB Advance PTA Dilatation Catheter, French size 5, rated burst pressure 15 atm, inflated diameter 7 mm, balloon length 4 cm, recommended introducer French size 6, sterile, REF (Global product no.) G26887, REF (order no.) ATB5-35-40-7-4.0. |  |  | February 28, 2007 | 2007 | Cook, Inc. | Mislabeled as to size: The outer label reads 7-4.0 balloon, when the balloon is actually 5-2.0. | Other | Consignees were notified via telephone on 2/28/07 and instructed to return the product. |  | 20 | 20 units | Illinois, North Carolina and South Carolina. | Terminated | May 29, 2007 | 90 | 20 | Undetermined
Z-0672-2007 | 37533 | No | KYG | Ocular surgery irrigation device. | Ophthalmic | Ophthalmic | 510(K) Exempt | BD Visitec Irrigating Cytosome, for use in cataract surgery. |  |  | March 02, 2007 | 2007 | Becton Dickinson and Company | Mislabeled: Carton label incorrectly labeled as Ref 581618 (.40 X 16mm) , may contain product Ref 581617 (.50 x 16mm) | Other | BD Opthalmic notifiied accounts by telephone on 3/2/07 and followed with letter. Accounts are requested to return product. |  | 1500 | 1500 units | Nationwide. | Terminated | March 05, 2008 | 369 | 1500 | Undetermined
Z-0669-2007 | 37405 | No | JTN | Antimicrobial susceptibility test disc. | Microbiology | Microbiology | 510(k) | BDÂż Sensi-DiscÂż Ticarcillin TIC-75, in vitro diagnostic, catalog # 231619, package of 10 cartridges, labeled in part ***BD Diagnostic Systems, 7 Loveton Circle, Sparks, MD 21152*** |  |  | February 19, 2007 | 2007 | Becton Dickinson &amp; Co. | Antibiotic susceptibility testing reagent was not manufactured according to specifications and may cause ineffective treatment regimen in patients with infections. The recalling firm received customer reports of false susceptibility results. | Other | The recalling firm ceased distribution of the product on 01/15/07 and subsequently notified US distributors by faxed letter and international consignees by e-mail on 02/19/07. The notification advised distributors to discontinue and return remaining inventory for replacement. The consignees were instructed to return the accompanying response form indicating inventory counts for credit. The notice also requested distributors to provide within three days the end-user lists of those receiving the product for subsequent direct notification by the recalling firm. The notification to consignees was initiated on 02/22/07 by overnight courier. The consignee notification advised that the product will give false susceptible results when testing with the identified QC organism using the Kirby Bauer method. Consignees were instructed to immediately discontinue use and return prouct identified by lot 6167435, EXP 06/30/08 and reply by returning the enclosed response form. The firm recommended to end users, primarily laboratory and testing facilities, that patient isolates yielding susceptible results with the identified lot of Ticarcillin Sensi-disc which had not been tested with the QC organism, be retested and the physicians notified. |  | 5280 | 528/10 packs | Worldwide, including USA, Australia, Canada, Hong Kong, Philippines, Belgium, Japan, Mexico, Singapore, Chile, India, and Venezuela. | Terminated | July 09, 2007 | 140 | 5280 | Undetermined
Z-0803-2007 | 37007 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | GE LightSpeed VCT Select / LightSpeed Pro32, Model 5126093, CT Scanner system. |  |  | December 21, 2006 | 2006 | GE Healthcare | Oil may leak onto the port window of the CT tube and may result in image artifacts that could lead to misdiagnosis. | Other | A Field Modification Instruction was issued to the field to inspect all systems and replace the tube window screws and the window filter, as well as clean any traces of oil in the tube window area. A 12/21/2006, Urgent Product Correction Notice/Customer letter has also been distributed informing the customers of the safety issue and the product correction. |  | 158 | 158 | Worldwide Distribution: USA in the following states and locations: AL, AK, AZ, AR, CA, CO, CT, DE, DC, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WV, WI, Guam and Puerto Rico. Foreign countries include ; Albania, Algeria, Argentina, Austria, Australia, Belgium, Bulgaria, Brazil, Belarus, Canada, Chile, China, Columbia, Czechoslovakia, Denmark, Finland, France, Germany, Great Britain, Greece, Honduras, Hong Kong, Iceland, India, Indonesia, Ireland, Israel, Italy, Japan, Korea, Kuwait, Latvia, Lebanon, Lithuania, Luxembourg, Malaysia, Martinique, Mexico, Netherlands, New Zealand, Norway, Panama, Philippines, Poland, Portugal, Reunion Island, Romania, Russia, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, United Arab Emirates, Venezuela, and Vietnam | Terminated | July 03, 2012 | 2021 | 1817 | Undetermined
Z-0684-2007 | 37168 | No | FFK | Electrohydraulic lithotriptor. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Swiss Lithocast Ultra Ultrasound Handpiece (EMS- PN:EL-236) used in combination with Swiss Lithoclast Ultra System . |  |  | March 16, 2005 | 2005 | E.M.S. - Electro Medical Systems SA | Handpiece may become dislodged causing cable damage and possible electric shock. | Other | EMS Lyon, France (manufacturer) notified the single US Distributor by letter to contact their accounts and return device for replacement. Boston Scientific, Natick, Ma conducted the sub-recall by letter dated 4/22/05 to all USA accounts. |  | 160 | 160 units | Nationwide | Terminated | February 02, 2016 | 3975 | 160 | Undetermined
Z-0661-2007 | 37452 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | MicroMaxx LAP/12-5 MHz Transducer, Laparoscopic Ultrasound Transducer |  |  | February 05, 2007 | 2007 | Sonosite, Inc. | Laparoscopic transducer exceeds the limits specified for radio-frequency emissions by up to 10db. | Other | On 2/5/2007 firm sent by FED-Ex notification letters to firm representatives. Representatives began hand delivery of recall letter to consignees and retrieval of LAP transducer for return delivery to firm. The letter advised customers that the LAP transducer exceeded the limits specified for radio-frequency emissions. The firm also advised that exceeding the limit may cause the LAP transducer to interfere with surrounding equipment that may be susceptible to radiated emissions and cause surrounding equipment to not perform as expected. |  | 6 | 6 Devices | Distribution to USA (Tennessee &amp; Ohio), France, Korea, and Germany. | Terminated | June 01, 2007 | 116 | 6 | Undetermined
Z-0668-2007 | 37384 | No | HCG | Neurovascular embolization device. | Neurology | Neurology | 510(k) | Matrix2 Detachable Coils, vascular embolization device. |  |  | February 07, 2007 | 2007 | Boston Scientific Target | The product may be missing the temperature indicator label on the inner pouch. | Other | The firm initiated the recall on 02/07/2007, and initially notified consignees via letter. |  | 57 | 57 units | Worldwide including USA, Argentina, Australia, Brazil, Czech Republic, Germany, Spain, Finland, France, Great Britain, Greece, Ireland, Netherlands, and New Zealand. | Terminated | November 27, 2007 | 293 | 57 | Undetermined
Z-0662-2007 | 37232 | No | GXI | Radiofrequency lesion probe. | Neurology | Neurology | 510(k) | RF Denervation Probe, 5cm, temperature sensing electrodes designed for use in radiofrequency lesion procedures for the treatment of pain. |  |  | January 19, 2007 | 2007 | Smith &amp; Nephew, Inc. Endoscopy Division | Product is non-sterile but labeled incorrectly as sterile. | Other | Smith &amp; Nephew notified accounts by letter sent via Federal Express January 19, 2007 |  | 56 | 56 units | Worldwide, including USA, Denmark, France, UK, Germany, Korea, Switzerland, Italy, Spain, and South Africa. | Terminated | January 18, 2008 | 364 | 497 | Undetermined
Z-0675-2007 | 37365 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Zimmer Trabecular Metal Reverse Shoulder System Instrumentation Conical Reamer, instrument for shoulder arthroplasty, REF no. 00-4309-021-00. |  |  | February 01, 2007 | 2007 | Zimmer Inc. | The reamer may lock up on the threads of the reamer body rather than spin freely as intended by design. | Other | Zimmer sales staff was notified by letter dated 2/1/07 to pull any of the product from stock and to provide the name and address of their hospital customer accounts for further notification by Zimmer. |  | 4 | 4 | California, Ohio and Texas. | Terminated | May 04, 2007 | 92 | 4 | Undetermined
Z-0667-2007 | 36947 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Disposable components of the superDimension Bronchus System: Locatable Guide Kit (AKI00017-01) and Extended Working Channel (AKI00041-01). |  |  | November 29, 2006 | 2006 | SuperDimension, Inc. | The action is being taken in response to certain incidents that have been reported to the company, one of which concerned an adverse event in which a patient incurred a pneumothorax during a procedure in which a superDimension Bronchus System was used. The biopsy results for the patient were taken from a different area than the confirmed navigation site that was indicated by the system. | Other | An Urgent Medical Device Field Action letter, dated 11/30/2006, was sent to medical facilities who have the superDimension Bronchus System installed. The letter infoms the facility to immediately stop using the Bronchus System, and move all inventory of the disposable system components to a controlled quarantine area. A superDimension representative will contact the consignees to schedule a time to retrieve the quarantined inventory, at which time a verification form will be completed. |  | 2000 | Up to 2,000 each Locatable Guide Kit and Extended Working Channel | Worldwide, including USA, Austria, Belgium, Canada, France, Germany, Hungary, Israel, Italy, Netherlands, South Africa, Spain, Switzerland, Turkey, and United Kingdom. | Terminated | November 20, 2008 | 722 | 2074 | Undetermined
Z-0811-2007 | 37865 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Certain MicroMiniplant Straight Healing Abutment. Dental implant. Model numbers: ISMHA32, ISMHA33, ISMHA34 and ISMHA36. |  |  | April 25, 2007 | 2007 | Biomet 3i | The Straight Healing Abutment did not fully seat to the restorative platform of the implant, leaving a gap between the abutment and the implant. | Other | Each consignee was notified on 4/25/2007 by telephone using a prepared script. A follow letter was sent by fax. Each consignee was instructed to check their respective inventory for the affected product and to return a response card and any affected product to the firm. |  | 362 | 362 units | Distributed worldwide, including USA, Spain and the Netherlands. | Terminated | November 18, 2008 | 573 | 362 | Undetermined
Z-0652-2007 | 37347 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Siemens syngo Dynamics 5.0 Workplace. Medical image report and archive system. |  |  | January 31, 2007 | 2007 | Siemens Medical Solutions, USA, Inc | The cardiac calculation feature may incorrectly calculate a derived V max value (reported out lower than it should be) and certain values imported into worksheets and reports may be incorrect due to a software defect. | Other | Distributors were issued a recall letter dated 1/29/07 and asked to notify their customers by issuing their subaccounts a similar letter. When available, new software will be provided. |  | 71 | 71 | Nationwide. | Terminated | April 30, 2008 | 455 | 71 | Undetermined
Z-0660-2007 | 37436 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | 8mm EndoWrist Bipolar Maryland Instrument, Part Number: 400172-04, |  |  | March 24, 2005 | 2005 | Intuitive Surgical, Inc. | The products were incorrectly programmed as training instruments, which allows for 30 uses (surgical procedures) instead of 10 uses. | Other | The firm initiated its recall in March 24, 2005, when the firm sent a letter via Federal Express mail to all 18 consignees, informing them of the issue, how to identify affected product, and instructions to return the affected instruments for a full credit. The form included a reply form that consignees were requested to fill out and return. |  | 8 | 8 | Product was distributed to 12 consignees in the US, who received 19 instruments, in FL, IN, LA, MN, NC, NH, NJ, NY, TN and TX. In addition, 5 consignees outside the US received product as well, in Italy, Sweden, France, Singapore and Germany. | Terminated | November 28, 2007 | 979 | 8 | Undetermined
Z-0810-2007 | 37761 | No | CEW | Parathyroid hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Roche PTH (parathyroid hormone test system) for use in conjunction with Roche cobas e 411, cobas e 601, Elecsys 1010, Elecsys 2010 and Modular E 170 analyzers; catalog no. 11972103122. |  |  | April 19, 2007 | 2007 | Roche Diagnostics Corp. | Negative bias caused by hemolysis interference at the low end of the reference range could cause erroneous diagnosis of hypoparathyroidism. | Other | The firm issued a letter dated 04/19/07 to all users informing them of the change in the hemoglobin interference claim. |  | 14503 | 14,503 | Nationwide. | Terminated | June 11, 2008 | 419 | 14503 | Undetermined
Z-0826-2007 | 37843 | No | GNX | Influenza virus serological reagents. | Microbiology | Microbiology | 510(K) Exempt | BinaxNOW A&B; Kit 10 Test Kit (Japan) |  |  | March 28, 2007 | 2007 | Binax, Inc. dba IMPD | Tip of the foam swab could detach from the shaft causing injury to the patient. | Other | Inverness Medical issued recall notification by letter on 3/28/07. Customers are instructed to discard the NP swabs provided in the kits. The firm will replace the NP swabs. In the meantime customers are to use the tests with nasal wash/aspirate samples (influenza) or nasal wash samples (RSV). Response letters are to be signed by customers and returned to Inverness Medical. |  | 3814 | 3814 kits | Nationwide and Japan. | Terminated | September 08, 2008 | 530 | 13934 | Undetermined
Z-0685-2007 | 37247 | [6 empty] | Plicator EPS Plication System |  |  | January 26, 2007 | 2007 | NDO Surgical, Inc. | Device arm may fail to open after being deployed and require surgical intervention to remove | Process design | NDO Surgical notified accounts by letter on 1/26/07 advising to quarantine units. Replacement units are being sent and recalled devices are requested to be returned. |  | 9 | 9 units | Worldwide distribution, USA including states of AZ, GA, NH, NY, TN and country of Germany. | Terminated | December 12, 2007 | 320 | 9 | Undetermined
Z-0709-2007 | 37385 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Abbott Precision Link Software (Version 2.1), (Version 2.2), (Version 2.3), (Version 2.4), (Version 2.5), and (Version 2.6); Serial Number(s): ALL; |  |  | December 18, 2006 | 2006 | Abbott Diabetes Care, Inc. | If the time and date are not set correctly on both the user's meter and computer, there is the potential for incorrect times and dates to be assigned to the results in the Precision Link Software. | Other | Firm initiated its recall on 12/21/2006. The firm provided customer letters with additional information (eg. flyer or insert) to customers that they had addresses for via fedex. In addition, the product notification was posted on the firm's internet website. Further, a flyer is to be included as a supplement to the user manual. |  | 67601 | 67,601 units | Worldwide Distribution: USA as well as abroad to Argentina, Aruba, Australia, Austria, Bahamas, Belgium, Canada, Colombia, Croatia, Curacao, Czech/Slovak Republics, Dominican Republic, Ecuador, El Salvador, Finland, France, Germany, Greece, Guatemala, Hong Kong, Ireland, Italy, Japan, Korea, Latvia, Lithuania/Estonia (Baltics), Mexico, Netherlands, New Zealand, Norway, Poland, Puerto Rico, Portugal, Saudi Arabia, Serbia, Bulgaria/Macedonia, Singapore, Slovenia, South Africa, Spain, St. Maarten, Sweden, Switzerland, Turkey, UK and Venezuela. | Terminated | December 20, 2010 | 1463 | 67601 | Undetermined
Z-0686-2007 | 37277 | [6 empty] | Aspyra, Inc. CyberMED Version 2.1 Pharmacy Software |  |  | January 31, 2007 | 2007 | Aspyra, Inc. | Medications prescribed for one patient were printed on the Medication Administration Record (MAR) of another patient. | Other | Current customers were mailed a Recall Letter January 31,2007 explaining the issue and informing the customers of an upcoming CyberMED Version 2.1 maintenance software release to correct the potential issue. A fax-back Acknowledgement of Receipt form was sent with the Recall Letter. |  | 10 | 10 | Nationwide to: CA, GA, IN, MI, MN, MS, NC, ND, &amp; OK | Terminated | July 22, 2011 | 1633 | 10 | Undetermined
Z-0688-2007 | 37291 | No | MNB |  |  | Obstetrics/Gynecology | PMA | HTA Console System, 230 Volt, English, Product Number: 560010 (HTA system is a software-controlled hysteroscopic thermal endometrial ablation system) |  |  | January 24, 2007 | 2007 | Boston Scientific Corporation | Defective integrated circuit board could result in loss of the system pump and patient injury (hot fluid 90 degree C into uterus) if in use during the recirculation/heating phase of treatment | Other | Boston Scientific notiifed accounts by letter dated January 24, 2007 via federal Express Mail. Accounts were requested to examine inventory, remove and return product. |  | 14 | 14 units | Worldwide-USA, UK, Denmark, Italy, and France | Terminated | December 12, 2007 | 322 | 83 | Undetermined
Z-0673-2007 | 37600 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Prostate Biopsy Needle, Model: NAC-1825BB, Soxe 18 GA x 25 CM, Single use, Ethylene sterilized, Non-pyrogenic, Remington Medical, Inc., Alpharetta, GA 30005. |  |  | January 19, 2007 | 2007 | Remington Medical Inc. | Pouches were not sealed prior to sterilization. | Other | Consignees were contacted by telephone on 01/19/2007. Each customer was informed the sterility of the product may be compromised as a result of unsealed pouches and requested to visually check their inventory. |  | 192 | 192 units | Nationwide. | Terminated | May 11, 2007 | 112 | 192 | Undetermined
Z-0217-2008 | 44962 | [6 empty] | Matrix MAX3 Actuator, Catalog Number: MAX30 A300515A254L000, SKF USA, Inc. Norristown, PA 19403 |  |  | July 24, 2007 | 2007 | Skf USA Inc | Sudden failure of the actuator during lifting operation | Labeling design | On or about 7/24/07, the firm sent an Important Safety Information letter and instructions to all known OEMs who purchase and sell SKF actuators for use in patient lifts. The letter reiterated and emphasized the importance of proper handling and maintenance for the actuator. The firm enclosed with this correspondence a proposed letter that the OEMs were asked to send to their customers, including the safety instructions and a new warning label. The firm also is making available for OEMs, dealers and their customers new warning labels to place on all actuators in service. The root cause of the problem is still under investigation by the recalling firm and additional instruction may follow later |  | 1298 | 1,298 units (Z-0215-0217-2008) | Worldwide Distribution-USA and Canada | Terminated | April 26, 2013 | 2103 | 1298 | Not_Computer
Z-0256-2008 | 45332 | No | LIP | Treponema pallidum treponemal test reagents. | Microbiology | Microbiology | 510(k) | BioPlex 2200 System, Cat. No. 665-1460A containing APF CD SW2_v1., Syphilis IgG IFU Packet contains the Syphilis IgG Assay Protocol File LAPF CD |  |  | September 25, 2007 | 2007 | Bio-Rad Labroatories Inc. | Mislabeling: Reporting terminology in the Syphilis IgG APF CD is not consistent with the distributed Instructions for Use. (Non-Reactive and Reactive rather than Negative and Positive) | Labeling Change Control | The recalling firm notified consignees by phone call on 9/25/2007, followed up with a fax notification. The firm plans to monitor the consignees that do not respond by tracking on a spreadsheet and following up with another phone call. Consignees are instructed to destroy the APF CD which will be replaced at a later date. |  | 36 | 36 | Nationwide | Terminated | April 16, 2008 | 204 | 36 | Not_Computer
Z-0259-2008 | 38445 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific 6F Guider Softip XF Guiding Catheter, 6 French, 40 degree angle, 90cm usable length, Material Number (UPN) H965100420, Catalog Number : 10042, Made in USA: 5905 Nathan Lane, Plymouth, MN 55442 USA |  |  | July 11, 2007 | 2007 | Boston Scientific Target | Mislabeling: Catheters were found to be mislabeled (90 cm labeled devices actually contains 100 cm length catheter) | Labeling mix-ups | Firm notified consignees by Urgent Medical Device Recall Letter dated 11 Jul 07 requesting return of product. |  | 47 | 47 units | Worldwide Distribution-USA, EU, Japan, South Korea, Mexico, Colombia, and Hungary. | Terminated | December 05, 2007 | 147 | 96 | Not_Computer
Z-0176-2008 | 38464 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | SmartSeal Wedge Femoral Pressuriser; Product Code/Catalog No: 3206002, DePuy Orthopaedics, Inc., 700 Orthopaedics Drive, Warsaw, Indiana 46581-0988-This product allows for effective and controlled pressurization of bone cement into the femur and the acetabulum during hip arthroplasty procedures |  |  | July 13, 2007 | 2007 | Depuy Orthopaedics, Inc. | Package Integrity: The outer peelable pouch seals may be compromised due to creasing, although the inner pouch seals are intact. | Packaging process control | Sales Force Managers were notified of this recall by an Urgent Information Recall Notice Letter, via email, on 7/13/07 and 7/23/07, to visit their Consignees. Letters for the Hospitals (Dear Materials Manager) and Dear Doctors Letters were provided as background by the Sales Force Managers during their visits. If any product remained in inventory, it should be quarantined immediately, and returned to DePuy. Hospitals should be notified immediately not to use the affected device. Replacement stocks will have a check mark on the bottom right front of package to indicate it has passed inspection. |  | 3530 | 3,530 | Nationwide Distribution -- including States of Arkansas, California, Colorado, Florida, Idaho, Illinois, Indiana, Kentucky, Maine, Michigan, Mississippi, New Mexico, North Carolina, Pennsylvania, South Dakota, Texas, Utah and Washington. | Terminated | June 23, 2009 | 711 | 3530 | Not_Computer
Z-0214-2008 | 45106 | [6 empty] | Stryker T2 Locking Screw, Fully threaded; 05 x 70mm, Catalog Number: 1896-5070S, Sterile, Qty: 1,Stryker Trauma GMbH, 24232 Schonkirchen, Germany |  |  | September 27, 2007 | 2007 | Stryker Howmedica Osteonics Corp. | Mislabeled: One lot of screws (05x70mm) is labeled as containing 70mm, but may actually contain 75mm length screws. | Packaging process control | Urgent Product Recall letters were sent on September 27, 2007, to all Stryker Branches/Agencies and hospitals who received the product requesting them to remove and quarantine the product for Stryker pick-up. |  | 50 | 50 | Nationwide | Terminated | May 08, 2008 | 224 | 50 | Not_Computer
Z-0523-2008 | 45814 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | life+cel, Replacement Battery for Cardiac Science Powerheart AED G3, Lithium Sulfur Dioxide, OEM P/N: 9146-001, voltage 12V, 7.5 amps, Mfg for Battery Zone, Inc ; Mfg by Nordix, Inc. |  |  | August 24, 2007 | 2007 | BatteryZone, Inc. | Marketed without 510 (K) approval | No Marketing Application | The original recall communication (important recall information) was sent out on 9/3/07 by Phoenix Solutions, a consultant to Battery Zone. A second recall communication, Urgent Device Recall, was issued by Battery Zone dated 10/25, and was sent out to all customers (even those who had returned product) on 11/5/07. Both communications were sent by first class mail. |  | 24 | 24 units | Nationwide; (primarily police/fire departments) | Terminated | July 01, 2013 | 2138 | 142 | Not_Computer
Z-0204-2008 | 44967 | No | NVR |  |  | Orthopedic | 510(k) | Medtronic Sofamor Danek, Satellite System Sphere PEEK, Size 12 mm, Material PEEK-Optima, quantity 1, Sterile, Rx only, REF 9000212, Manufactured at: Warsaw, Indiana, USA |  |  | September 13, 2007 | 2007 | Medtronic Sofamor Danek USA Inc | Labeling, including indications for use, lack 510 (K) approval | Labeling design | On September 12, 2007, the recalling firm's representatives were instructed by electronic mail to remove any manuals in the field and contact surgeons who have evaluated, implanted or been trained to implant the Satellite System and to collect any guides they may have on site. The sales representatives that have accounts where the Satellite product was shipped directly were mailed the information and instructed to hand deliver the 9/13/07, Urgent Device Recall notices to each consignees risk management office. The recall was then expanded on 10/31/2007, to include the surgical manual, the implants and associated instruments. The delivery of the second Urgent Device Recall Notices, dated 10/25/07, were handled in three ways: 1) The letter was delivered to all initial consignees that received satellite products along with a questionnaire (response card) to be completed by the Risk Manager or equivalent administrative function., 2) The letter without the questionnaire was sent directly to target surgeons who are known to have implanted or have been trained to implant Satellite Product 3) a field action confirmation form was sent to Sales Representative with surgeon customers that are receiving the recall communication. The form will be completed by the Sales Representative to confirm the surgeons have received the recall communication and have returned any of the surgical techniques and instruments. |  | 179 | 179 units | Nationwide | Terminated | August 18, 2009 | 705 | 6071 | Not_Computer
Z-0441-2008 | 45456 | No | LWJ | Hip joint femoral (hemi-hip) metallic cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Excia Cemented Hip Stems, Size 14mm, Part Number : NJ314K, Aesculap Implant Systems, Inc., Center Valley, PA 18034 |  |  | October 10, 2007 | 2007 | Aesculap Implant Systems | Sterility (package integrity) Compromised: Device package damaged during shipment | Package design/selection | The recalling firm issued an Important Recall Notice letter dated 10/10/07 to all their customers informing them of the problem and the need to return the product. |  | 167 | 167 units | Nationwide-USA including states of AZ, GA, MD, OK, PA, TX, and VA (distributors) medical facilities in FL, NY, and PA; and sales representatives in CA and FL. | Terminated | February 07, 2008 | 120 | 167 | Not_Computer
Z-0255-2008 | 45327 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Biomet Interlok 67 mm Fixed Cruciate Tibial Plate with Locking Bar; CO-CR-MO/TI 6AL 4V Alloy; Product Number/REF 141232, Biomet Orthopedics, Inc., Warsaw, IN 46582 |  |  | September 18, 2007 | 2007 | Biomet, Inc. | Mislabeled as to size: the package actaully contained a 71 mm tray. | Labeling mix-ups | Consignees were notified via letter dated 9/18/07 to discontinue use of the product and to return it to the firm. |  | 7 | 7 | Nationwide : Distributors in Georgia, Kentucky, Texas and Washington. | Terminated | June 04, 2008 | 260 | 7 | Not_Computer
Z-0263-2008 | 45489 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | Baxter Duploject Easy Prep, Fibrin Sealant Preparation and Application System ISO; Product Codes: 1500949, 1500950, 1500951, 1501250 &amp; 1501252, Baxter Healthcare Corporation, Deerfield, IL 60015 |  |  | October 25, 2007 | 2007 | Baxter Bioscience | Sterility (package integrity) compromised: This recall is being conducted due to incomplete seal of the product package at the corner of the outer tray and tyvek lid. | Packaging process control | An Urgent Product Recall Letter dated October 24, 2007, was issued with attached Product Recall Instructions and Customer Response forms to all consignees (Director of Materials Management, and Director of Risk Management). The recall letter requests sub recalls to the retail level and product quarantine and return. The recall notification letters were sent via commercial carrier (UPS Next Day Service). |  | 4782 | 4,782 units (6 units/case) | Worldwide-USA, Canada and European Union (France and UK) | Terminated | March 13, 2009 | 505 | 4782 | Not_Computer
Z-0264-2008 | 45539 | No | DTK | Cardiovascular intravascular filter. | Cardiovascular | Cardiovascular | 510(k) | Vena Tech LP Vena Cava Filter System, Item No/REF. 05010024, Model No: 40DLP, B Braun. Manufacturer: B. Braun Medical 204, avenue du Marechal Juin BP 331 92107 Boulogne Cedex , France. (The VenaTech LP vena cava filter is pre-loaded in a syringe (cartridge) and provided as a system with introducer accessories and instructions to accommodate delivery and implantation either via the femoral or jugular approach.) |  |  | October 30, 2007 | 2007 | B. Braun Interventional Systems | Sterility (package integrity) compromised: package sealing problem leaving packaging being partially open on one of the four sealed Tyvek pouch edges. | Packaging process control | B. Braun sent an Urgent Medical Device Removal letter (10/10/07) by express mail to each of the hospital accounts that received the recalled lot of VenaTech LP. B. Braun sales represenatatives will also contact each account directly to inquire if any inventory of the suspect lot remains and to provide guidance for return and replacement of any existing inventory. |  | 60 | 60 units | Nationwide: States of AL, GA, PA, FL, NE, WI, IL, LA, TX, AZ, MN, MA, CA, NY, OH, IN, and MI. | Terminated | March 02, 2008 | 124 | 60 | Not_Computer
Z-0275-2008 | 45326 | No | KXA | Hip joint femoral (hemi-hip) metallic resurfacing prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer Durom Hip Resurfacing Systems, Femoral Component 54 Code T; Catalog No.: 01.00211.154, Zimmer, Warsaw, IN, 46581-0708. |  |  | September 26, 2007 | 2007 | Zimmer Inc. | Mislabeled as to size; 54mm package contains 46 mm device and vice versa. | Packaging process control | Zimmer GmbH notified Distributors via a Product Recall Notification letter, dated 9/26/07, instructing them to return the products and to subrecall. |  | 1 | 1 | Nationwide | Terminated | June 04, 2008 | 252 | 6 | Not_Computer
Z-0132-2008 | 38381 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | VNUS ClosureFAST catheters, Model Number: CL7-7-100, VNUS Medical Technologies, Inc., San Jose, CA 95138 |  |  | March 06, 2007 | 2007 | VNUS Medical Technologies Inc | Sterility (package integrity) Compromised: Product sterility may be compromised due to a pinhole under the pouch foldover. | Package design/selection | The recalling firm has issued customer notification letters by mail on 03/06/2007. The firm is requesting that consignees discontinue their use of the device and return recalled product . The firm will provide replacements. |  | 92 | 92 units | Nationwide: Product was distributed to 18 health care providers and medical centers throughout the US in the following states: TX, UT, CA, CT, MI, NY, MI, OR, WA, VA, PA, and NJ. | Terminated | December 17, 2007 | 286 | 92 | Not_Computer
Z-0123-2008 | 44814 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Biomet Orthopedics, Vanguard PS Open Box Femoral, 67.5 mm Left-Interlok, For Cemented Use Only, CO-CR-MO Alloy, REF 183130, Knee Prosthesis Component, Biomet Orthopedics, Inc., Warsaw, IN 46581 |  |  | August 31, 2007 | 2007 | Biomet, Inc. | Mislabeled as to size; labeled as 67.5 mm, actually contains 62.5 mm size femoral component. | Packaging process control | Biomet notified consignees via letter dated 8/31/07, titled Urgent Medical Device Recall Notice to immediately discontinue use of the product and to return units to Biomet. |  | 16 | 16 | Nationwide. | Terminated | June 04, 2008 | 278 | 16 | Not_Computer
Z-0385-2008 | 44893 | No | NJL |  |  | Orthopedic | PMA | Depuy P.F.C. Âż E Knee System Stabilized Tibial Insert (STAB), 67 mm M/L, 45 mm A/P, 8.0 mm, 2.5, sterile; Product Code/REF 96-0430. |  |  | September 27, 2007 | 2007 | Depuy Orthopaedics, Inc. | An incomplete seal on the inner pouch allows exposure to oxygen, which may result in oxidation and would affect long-term performance of the implant. | Packaging process control | Consignee hospitals and implanting physicians were notified of the recall via an Urgent Information-Recall Notice letter and Dear Doctor letters dated 9/27/07. The hospitals were instructed to determine whether they had any of the recalled products and to contact their sales representative to pick up the product. Depuy sales representatives were notified to collect the products from consignee hospitals via email on 10/1/07. The decision to discuss this product defect with patients who had already been implanted was left to the implanting physicians. |  | 10 | 10 | Nationwide. | Terminated | April 30, 2008 | 216 | 1160 | Not_Computer
Z-0279-2008 | 45432 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific, Imager II Angiographic Catheter, (Outer Carton Label): UPN M001315081, 5 Fr, 100 cm, Catalog No. 31-508, Lot 48951, Description 5 Fr Flush TR, Sterile EO, Made in Ireland, Unit 7, 8 &amp; 9, Annacotty Business Park, Annacotty, C. Limerick, Ireland ****Inner Pouch Label: Boston Scientific, Imager II Angiographic Catheter,UPN M001315080, 5 Fr, 100 cm, Catalog No. 31-508, Lot 48951, Description 5 Fr Flush TR, Sterile EO, Made in Ireland, Unit 7, 8 &amp; 9, Annacotty Business Park, Annacotty, C. Limerick, Ireland |  |  | October 12, 2007 | 2007 | Boston Scientific Corporation | Mislabeled: Units within the identified lot/batch were incorrectly labeled as 5F on the outer carton and pouch when the product was actually a 4F device. | Packaging | Boston Scientific sent an Urgent Medical Device Recall -- Immediate Action Required letter to consignees on 10/12/07. The letter described the product and the mislabeling problem (labeled outer carton and pouch as 4F when they are actually 5F devices). The firm recommended segregating the affected product and to return it to Boston Scientific PLEASE NOTE: The firm's recall letter incorrectly addressed the mislabeling of this product recall. The mislabeling involves the labeled outer carton and pouch were labeled as 5F when the device enclosed in the pouch was a 4F device. |  | 70 | 14 boxes each containing 5 Units | USA only including states of NJ, CA,TX,AZ | Terminated | December 19, 2007 | 68 | 70 | Not_Computer
Z-0463-2008 | 45334 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Hipstar V40 Femoral Stem-127 degrees; Catalog No: 7850-5-001, N degrees 3 high offset 31 mm;Sterile; IMPLANT SIZE 1, Howmedica,Benoist Girard dt Cie, 203 bd de la Grande Della B.P.8;14291 Herouville Saint Clair |  |  | September 28, 2007 | 2007 | Stryker Howmedica Osteonics Corp. | Mis-labeled: The warning label "Do not use with heads more than +10mm offset" does not appear on 127 degree stem boxes. | Packaging process control | Urgent Product Correction/Recall notification letters were sent by Fed Ex return receipt on 9/28/07, requesting quarantine for Styker representatives to return or re-label. |  | 59 | 59 | Nationwide- USA including states of OR, AZ, NV, MI, AL, MI, FL, MN, and CA | Terminated | October 27, 2008 | 395 | 910 | Not_Computer
Z-0331-2008 | 44963 | [6 empty] | Axya Bone Anchor , 3.0 mm Titanium with Nylon Suture |  |  | September 27, 2007 | 2007 | Axya Medical | Sterility Compromised (Package Integrity): Sterility of the device may be compromised due to holes in the sterile packaging | Package design/selection | Axya Medical notified customers by letter, Urgent-Product Recall, dated 9/25/07 , via UPS . Users are requested to return product and complete the reply form. |  | 1008 | 1,008 units | Nationwide | Terminated | June 21, 2012 | 1729 | 17956 | Not_Computer
Z-0358-2008 | 45427 | Yes | CAC | Autotransfusion apparatus. | Anesthesiology | Anesthesiology | 510(k) | Haemonetics cardioPAT System, Product List Number 02050, Haemonetics Corporation, Braintree, MA 02184,(Cardiovascular Perioperative Autotransfusion) |  |  | October 15, 2007 | 2007 | Haemonetics Corporation | Labeling (additional) use instructions: to clarify proper position of cardioPAT device and post-op line to prevent fluid build up from patient's chest | Labeling design | Haemonetics issued a Labeling suppplement to users via a Field Notification letter on October 15, 2007. Users are provide additional Directions for use for the Operator's Guide and Quick Reference Guide. |  | 106 | 106 units | Nationwide: USA including states of CA, CT, FL, IN, KY, MI, NC,NE, NH, PA, RI, SC,TX, and WA | Terminated | March 19, 2012 | 1617 | 106 | Not_Computer
Z-0292-2008 | 45817 | No | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | Medrad Avanta Multi-Patient Disposable Hand Controller, Angiographic Injector and Syringe. Catalog number AVA 500 HC, For Use with MEDRAD Avant Fluid Management System, MEDRAD, Inc., Indianola, PA 15051 |  |  | October 25, 2007 | 2007 | Medrad Inc | Sterility (package integrity) compromised: Incomplete package seal; sterility may be compromised | Packaging process control | The recalling firm issued an Urgent Medical Device Field Correction Letter dated 10/25/07 via Federal Express to their customers informing them of the problem. Customers can either return the product or consistent with the indicated use and established pratice of sheathing, utilize sterile sheaths with the hand controllers for up to five patients. Each letter is being accompanied by a box of 25 sheaths. |  | 25977 | 25, 977 units | Worldwide-USA and countries of Germany, Israel, Finalnd, Italy, Netherland, Austria, France, Japan, Denmark, Ireland, South Africia, Australia, Saudi Arabia, Singapore, Canada, New Zealand, Spain, Sweden, Portugal, Korea, Great Britain, Turkey and The United Arab Emerites. | Terminated | April 22, 2008 | 180 | 25977 | Not_Computer
Z-0375-2008 | 44972 | No | JGS | Sodium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Abbott Point of Care, i-STAT G Cartridge; (Glucose/Glu); List No: 06F09-01; Part No: 120100; Abbott Point of Care, Inc., East Windsor, NJ 08520, USA; Product of Canada. |  |  | October 18, 2007 | 2007 | Abbott Point of Care Inc. | Difficult to Close -- Certain lots of i-STAT cartridges are difficult to close, will not close or do not stay closed. | Packaging change control | Abbott Point of Care sent and Urgent Notice to all i-STAT Customers (US and Canadian consignees) notifying them that certain lots of i-STAT cartridges have snap closures that are difficult to close, will not close or do not stay closed. The customers were asked to review the lot numbers on the enclosure and, if the facility has any cartridges from the lot numbers listed, they should remove from service immediately. |  |  |  | Worldwide Distribution- including USA and country of Canada. | Terminated | July 18, 2008 | 274 |  | Not_Computer
Z-0170-2008 | 45491 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Signa MFO/I Magnetic Resonance System, Model Numbers : 2210339, 2262700, 2276937, 2282853, 2282860, 2282880, 2352129, 2388781, 5105123, 5113339, 2138300-32, 2189500-3, 2198999-4, 2259620-4, 2262700-2, 2269512-2, 2286849-2, 2352129-2, 2377062-5, 2377062-8, Manufactured by GE Hangwei Medical Systems Co., LTD, Beijing PR, China |  |  | December 31, 2006 | 2006 | GE Healthcare | Incorrect patient positioning; when moving the patient into the bore may lead to wrist and hand injury. | Labeling design | Field Modification Instruction : FMI 67631 (DEC 2006) was communicated directly to the GE Healthcare service engineers who provide the modification at the customer site. Corrections were performed in DEC 2006 |  | 163 | 163 | Worldwide Distribution --- USA including states of AL, AR, AZ, CA, CO, CT, FL, ID, IL, IN, KS, KY, LA, MD, MO, MS, MT, NJ, NY, OH, OK, PA, SD, VA, and WA, and countries of United Arab Emirats, Argentina, Belgium, Brazil China, Ivery Cost, Germany, Spain, France, United Kingdom, Hungary, Italy, Japan, Korea, Mexico, Oman, Panama, Puerto Rico, Russia, and Turkey | Terminated | December 11, 2007 | 345 | 163 | Not_Computer
Z-0062-2008 | 44751 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | HeartStart XL Defibrillator/Monitor, Model: M4735A |  |  | September 04, 2007 | 2007 | Philips Medical Systems | Inadequate shipping package causing damage to control boards resulting in unit failure. Failure to deliver therapy (defibrillation, pacing or cardioversion) may occur as a result of undetected damage to the PCB assembly. | Package design/selection | Philips Medical notified accounts (Distributors/Healthcare Customers) by letter dated 9/4/07. Users are requested if possible to remove the unit from service. Firm will have Field representative conduct correction. |  | 147 | 147 boards | Puerto Rico, Canada, Argentina, Australia, Brazil, China, Domincan Republic, Czech Republic, Hong Kong, India, Japan, Korea, Macau, Singapore, Spain, South Africa, Slovenia, Sweden, Switzerland, Ireland, UK, France, Kuwait, Poland, and Germany | Open, Classified |  |  | 147 | Not_Computer
Z-1221-2007 | 38983 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Product is labeled as LTX XP Miniplant, Dental Implant, Sterile. |  |  | August 03, 2007 | 2007 | Biomet 3i | The LTX XP Miniplant, BIOMET 3i dental implant, was labeled incorrectly. The 13MM implants were labeled and packaged as 11.5MM implants and 11.5MM implants were labeled and packaged as 13MM implants. | Labeling mix-ups | International customers were contacted first by e-mail , then by fax letter 08/03/2007. |  | 240 | 240 (120 each model / lot) | Product was not distributed within the U.S. Product was distributed Internationally - Brazil, Argentina, Israel, Mexico, Colombia, Korea, Taiwan, Uruguay, Italy, Chile. | Terminated | September 11, 2009 | 770 | 240 | Not_Computer
Z-0109-2008 | 44939 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific-Renegade Hi-Flo Microcatheter Kit, |  |  | September 10, 2007 | 2007 | Boston Scientific Corporation | Incorrect Guidewire: Product may incorrectly contain a 135 cm length Transend-18 guidewire instead of the kit labeled 165 cm length Transend-18 guidewire. | Packaging process control | Boston Scientific issued a Urgent Medical Device Recall Letter dated September 10, 2007 to Domestic accounts via Federal Express. The notification instructed the account to immediately discontinue the useof the product under this recall, check their inventory, remove any affected product , and return it to Boston Scientific. The notification also requested the completion of a Reply Verification Tracking Form |  | 153 | 153 | Nationwide | Terminated | December 20, 2007 | 101 | 153 | Not_Computer
Z-0002-2008 | 38369 | equipment, laboratory, general purpose, labeled or promoted for a specific medical use | LXG | General purpose laboratory equipment labeled or promoted for a specific medical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | FP1000 Cell Preparation System Software versions 1.1 &amp; 1.2, Part Number 624922, for in vitro diagnostic use. |  |  | February 22, 2007 | 2007 | Beckman Coulter Inc | Pressure (from recapped tubes) or vacuum (from under-filled short-draw tubes) will cause inaccurate results if not properly vented. | Labeling design | A Product Corrective Action (PCA) letter was mailed on Feb 22, 2007, to all FP1000 customers to inform them of an issue associated with the FP1000 Cell Preparation System. Pressure (from recapped tubes) or vacuum (from under-filled short-draw tubes) will cause inaccurate results if not properly vented. Under these conditions the FP1000 will under-dispense specimen volume leading to lower than expected absolute counts. Specimen tubes that are recapped or underfilled must therefore be vented prior to running on the FP1000. If the initial state of the specimen tube is unknown, then it must be vented to ensure accurate results prior to running on the FP1000. Additional instructions are provided. A Fax Back Response Form was also provided. |  | 10 | 9 in the U.S.; 1 in Canada | Nationwide (MA, IN, IL, UT, RI, NJ), and Canada. | Terminated | April 26, 2012 | 1890 | 10 | Not_Computer
Z-0065-2008 | 44881 | No | LOM |  |  | Microbiology | PMA | Bio-Rad Laboratories-MONOLISA Anti-HBc IGM EIA kit, an in-vitro diagnostic test kit, Catalog # 26174, manufactured by Bio-Rad Laboratories, Redmond, WA. |  |  | September 05, 2007 | 2007 | Bio-Rad Laboratories | Incorrect Instruction: the package insert states to add 100L of conjugate; it should state to add 100uL. | Error in labeling | On 9/5/07 a letter titled "IMPORTANT PRODUCT INFORMATION" was sent to customers advising them of the issue and requesting customers make a note of the correction in their laboratory procedures and/or any existing copies of the package insert. On 9/19/07 the firm sent a second letter (dated 9/18/07) via DHL titled "URGENT PRODUCT CORRECTION" and supplied a corrected package insert to replace the original insert sent. |  | 46 | 46 | Nationwide: Kits were distributed to 27 hospitals, clinics and testing labs throughout the US. | Terminated | November 15, 2007 | 71 | 46 | Not_Computer
Z-0588-2007 | 37346 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | FRAZIER Surgical Suction Instrument, 8 Fr. with Control Vent and Obturator, REF 0033080. MANUFACTURED FOR CONMED CORPORATION 310 Broad Street, Utica, NY 13501. ASSEMBLED IN MEXICO. |  |  | November 10, 2006 | 2006 | Conmed Corporation | Firm was made aware of instances where the sterile barrier of the instruments had been compromised. | Packaging process control | On 11/10/06, ConMed Corp. sent "URGENT: DEVICE RECALL" letters (dated 11/10/06) to the US and foreign consignees by UPS overnight. The letters instructed the consignees to inspect their stocks of the instruments and that they return all devices. The letters also instructed the distributors to contact their customers that received the instruments and conduct sub-recalls. Additionally, when requested, ConMed notified the distributor's customers (conducted sub-recalls) by UPS overnight mail. This recall action is to the end user level. Accompanying each letter was an effectiveness check form to be completed by the consignee, and faxed back to ConMed Corp. The instruments are to be returned to ConMed Corp. in Utica, NY. For questions, the consignee can contact Ms. Nancy Crisino by phone at (315) 624-3078 or by fax at (315) 624-3089. |  | 711245 | 711,245 units | Worldwide | Terminated | July 31, 2012 | 2090 | 7401479 | Not_Computer
Z-0076-2008 | 44825 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Biomet --DCM AGC Molded Tibial Implant (12 mm x 80 mm)--Biomet Orthopedics , Interlok finish/posterior stabilized, 12 mm thick x 80 mm wide, UHMWPE/CO-CR-MO alloy, sterile; REF 155364, Biomet Orthopedic, Inc. |  |  | July 24, 2007 | 2007 | Biomet, Inc. | Incorrect Product -- Mislabeled as to size; a 18 mm x 65 mm tibula implant is enclosed in a package labeled as 12 mm x 80 mm. | Labeling mix-ups | The firm notified distributor by phone on or about 7/24/07 and followed up with an Urgent Medical Device Recall Notice/Letter dated 8/31/07. The letter requested that the product be discontinued and returned to Biomet. |  | 4 | 4 | Class II Recall - Worldwide Distribution --- USA including states of IA, SD, and countries of Brazil and Chile. | Terminated | July 24, 2008 | 366 | 5 | Not_Computer
Z-0055-2008 | 39383 | No | LGW |  |  | Neurology | PMA | Precision Implantable Pulse Generator (IPG), Model Number SC-110 |  |  | August 03, 2007 | 2007 | Advanced Bionics Corp | Mislabeled: Product labeled with the incorrect shelf-life information. | Error in labeling | There was one unit affected by this field action. An Advanced Bionics representative contacted and visited the facility on August 3, 2007. At this time, the affected device has been retrieved by the Company. |  | 1 | 1 unit affected | MD | Terminated | January 25, 2012 | 1636 | 1 | Not_Computer
Z-1092-2007 | 38019 | No | LPN | Soft (hydrophilic) contact lens care products. | Ophthalmic | Ophthalmic | 510(k) | COMPLETE MOISTUREPLUS MULTIPURPOSE SOLUTION, Formula 9451X; |  |  | May 25, 2007 | 2007 | Advanced Medical Optics, Inc. | In response to information received 05/25/2007 from the U.S. Centers for Disease Control and Prevention (CDC) regarding eye infections from Acanthamoeba, a naturally occurring water-borne organism which can contribute to serious corneal infections. CDC estimates a risk of at least seven times greater for those who use Complete MoisturePlus solution versus those who did not. | Labeling design | On Friday, May 25, 2007 at 11:16 ET, AMO issued a press release to the market related to the voluntary recall of AMO's COMPLETE MoisturePLUS TM Multi-Purpose Solution from the market. Similar press releases, specific to the geographic regions, were issued in the other markets beginning on Saturday, May 26, 2007.A copy of the United States press release is present on the AMO website (www.amo-inc.com) and was forwarded to the FDA for posting on the FDA web site. In addition, the FDA also completed a press release related to the AMO recall and it is available on the FDA web site.Regulatory agency notifications in all markets affected outside the United States were initiated beginning on Monday, May 28. Beginning on Thursday, May 31, AMO (via Stericycle) distributed RetaiI Customer notification packages via overnight UPS delivery to the 485 customer accounts in the United States. |  | 27807198 | 27,807,198 units in US (57,252,581 Worldwide) | Worldwide including USA, Canada, Europe, Japan, China and various Asia Pacific countries (to include Hong Kong, Australia, New Zealand, Korea, Vietnam, Singapore, India, Indonesia, Sri Lanka, Taiwan, Malaysia, Pakistan and the Philippines). | Terminated | January 10, 2012 | 1691 | 31602188 | Not_Computer
Z-1219-2007 | 39362 | [6 empty] | Biogel Skinsense Polyisoprene Surgical Gloves, Non-Latex, Sterile, (Size 8) Made in Malaysia |  |  | July 20, 2007 | 2007 | Molnlycke Health Care, Inc | Sterility compromised (package integrity): The primary packaging for a portion of the lot may contain gloves that were not sealed properly. | Packaging process control | Consignees were notified by registered mail on 07/23/2007. The notice contained a Recall Response Form to confirm receipt of the notice and to verify if the product is still in inventory. The notice instructed to consignee to inspect their inventory for Biogel Skinsense Polyisoprene Surgical Gove, lot #07B026, Size 8. If this lot and size are in their inventory, they were instructed to discontinue use, quarantine the product and return the product to Molnlycke Healthcare. |  | 6000 | 6,000 pair | Nationwide including states of AZ, CA, CT, KY, MD, MN, MO, TN, and WI | Terminated | March 17, 2010 | 971 | 6000 | Not_Computer
Z-0614-2007 | 37404 | No | KNS | Endoscopic electrosurgical unit and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Gyrus ACMI Tripolar Ultra Cutting Forceps: 10 mm X 32 cm-Reference Number: 006689-901 |  |  | February 19, 2007 | 2007 | Gyrus ACMI Corporation | Product sterility may be compromised due to failure of the tray seal packaging | Packaging process control | Gyrus ACMi issued recall letters on 2/19/07. Each account is asked to remove recalled product and return product . |  | 184 | 184 boxes (5 units/box) | Worldwide distribution ----USA and countries of Canada, The Netherlands, Japan, Portugal, Columbia, South Africa, Australia, and Spain. | Terminated | February 24, 2008 | 370 | 722 | Not_Computer
Z-0064-2008 | 39416 | No | KXT | Digoxin test system. | Toxicology | Toxicology | 510(k) | Abbott Diagnostic International, Ltd.-- AxSYM Digoxin III Reagent Pack, List (6L07-20), For In Vitro Diagnostic use. This reagent system is composed of one bottle of Digoxin-Alkaline Phosphatase Conjugate, one bottle of Anti-Digoxin (Rabbit) Coated microparticles, one bottle of MEIA buffer, and one bottle of Digoxin Probe wash solution. |  |  | July 20, 2007 | 2007 | Abbott Diagnostic International, Ltd. | Incorrect Results -- Error codes 1062, 1063, 1113, and 1118 observed when running patient samples on the AxSYM Digoxin III assay, rather than a concentration, is reported for the patient sample. | Labeling design | Abbott Diagnostics International Limited, Ltd., mailed a Product Correction Letter to customers on 7/20/02 advising an additional option(s) has been identified to resolve Error Codes: - Perform one of the two options to resolve these error codes 1062, 1063, 1113, and 1118 observed when running patient samples on the AxSYM Digoxin III assay. - Dilute the sample one to one (1 part sample plus 1 part Calibrator A) and repeat analysis. - The final concentration is equal to the reported concentration of the sample multiplied by the dilution factor - The minimum final concentration that may be reported out is 0.6 ng/mL, which corresponds to the sensitivity of the AxSYM Digoxin III assay (0.3 ng/mL) multiplied by the dilution factor (2). OR- Centrifuge the sample at 8,000 to 10,000 RCF for 10 minutes and repeat analysis. |  | 3104 | 3,104 | Worldwide Distribution -- USA including States of AL, AR, AZ, CA, CT, FL, GA, IA, ID, IN, IL, KS, KY, LA, MA, ME, MD, MI, MN, MS, MO, MT, NV, NJ, NY, NC, OH, OK, OR, PA, SD, SC, TN, TX, UT, VA, VT, WA, WI, WY, HI, Puerto Rico, and countries of Canada, Venezuela, Uruguay, Germany, Hong Kong, South Korea, Australia, and New Zealand | Terminated | September 09, 2011 | 1512 | 3104 | Not_Computer
Z-1191-2007 | 38351 | No | DRF | Electrode recording catheter or electrode recording probe. | Cardiovascular | Cardiovascular | 510(k) | Polaris Dx Diagnostic Catheter, for cardiac electrophysiology studies, Model M0049663S0 |  |  | June 26, 2007 | 2007 | Boston Scientific Corp | Product is mislabeled as hexapolar (6 electrodes), but is actually quadripolar (4 electrodes). | Labeling mix-ups | The firm issued notification letters to its consignees, sent 06/26/2007, via FedEx. The firm intends to remove the product and destroy it. |  | 30 | 30 units | Product was distributed to 3 consignees in Florida and Tennessee. | Terminated | November 27, 2007 | 154 | 30 | Not_Computer
Z-0095-2008 | 44788 | No | KWY | Hip joint femoral (hemi-hip) metal/polymer cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | FoundationÂż Hip System Bipolar Modular, Size 53mm OD; Cat. #412-01-053; Sterile; assembly required by physician prior to implant; manufactured by Encore Medical, L.P., Austin, TX 78758. |  |  | August 02, 2007 | 2007 | Encore Medical, Lp | Device mislabeled; hip prosthesis components validated for 5 years were labeled as having a 6 year expiration date. | Packaging change control | Firm notified consignees of recall via phone, e-mail and overnight letters beginning on 08/02/07. Consignees asked to return devices for a properly labeled replacement. |  | 137 | 137 units. | Worldwide, including USA, Greece, Japan, and Saudi Arabia. | Terminated | January 09, 2008 | 160 | 5174 | Not_Computer
Z-0114-2008 | 39409 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Medtronic Sofamor Danek METRx (TM) System Bayonet Electrosurgical Pencil, sterile, Manufactured at Warsaw, Indiana; REF (P/N) 9560575. |  |  | September 04, 2007 | 2007 | Warsaw Orthopedic, Inc. dba Medtronic SOFAMOR DANEK | Packaging weakness may compromise sterility and instructions for use are inadequate. | Package design/selection | Consignees were notified via letter dated September 4, 2007 to cease using the product and to contact the firm for instructions on how to return the product. |  | 22632 | 22,632 | Worldwide, including USA, Hong Kong, India, Mexico, Netherlands, Qatar, Saudi Arabia and South Korea. | Terminated | July 14, 2009 | 679 | 22632 | Not_Computer
Z-0008-2008 | 38894 | No | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | CVI OZ Palm Injector, used to perfuse contrast media into vessels for angiographic procedures. Ref : OZ 1035 (Cardiovascular Innovations LLC) |  |  | July 18, 2007 | 2007 | Lemaitre Vascular, Inc. | Sterility of device may be compromised due to breach in sterile barrier. | Package design/selection | The firm issued a "Adivsory Notice" to accounts 7/18/07, advising them to examine inventory and remove product with defective packaging. The firm issued an updated notification dated 9/24/07 advising users to return recalled lots instead of conducting an examination of the product at the user site. |  | 7860 | 7,860 units | Worldwide, including USA, Canada, Netherlands, and Germany. | Terminated | September 08, 2008 | 418 | 7860 | Not_Computer
Z-1262-2007 | 38950 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Merit Medical Surgical Convenience Kits, Custom and SPPT, Merit Medical, South Jordan, UT 84065. Part numbers: K05-01767A, K05-01769, K05-01769A, K09-03843M, K09-03951M, K09-05894F, K09-08610, K10-00033Q, K10-00382M, K10-00685K, K10-00984F, K10-01531F, K10-01826D, K10-01826F, K10-02902B, K10-03080D, K10-03378C, K10-03397B, K10-03590A, K10-03590B, K10-03659D, K10-03659F, K10-03659G, K10-03689, K10-03907B, K12-01342D, K12-01444B, K12-01468, K12-01484A, K12-01487, K12-01490A, K12-01593, K12-01638, K12-01661, K12-01667, K12-01667A, K12-01964, K12-01968, K12-02098, SPPT-100/C, SPPT-5F-10C, SPPT-5F-10L/A, SPPT-5F-7/C, SPPT-5F-7L/A. |  |  | August 15, 2007 | 2007 | Merit Medical Systems, Inc | Expiration Dating; Some Convenience Kits contain medications and/or skin preparations that may expire before the expiration date on the outer package label. | Error in labeling | An August 22, 2007, Product Correction/Removal letter was issued to all distributors requesting return of affected lots and sub recall notification to their consignees. |  | 16333 | 16,333 units | Nationwide | Terminated | May 13, 2011 | 1367 | 16333 | Not_Computer
Z-0003-2008 | 38430 | No | MSQ | Campylobacter fetus serological reagents. | Microbiology | Microbiology | 510(k) | BreathTek Urea Hydrolysis Rate Calculation CD, Version 1.3.5 - January 1,'03, in vitro diagnostic. Meretek Diagnostics, Inc., Lafayette, CO. |  |  | May 02, 2007 | 2007 | Meretek Diagnostics, Inc. | Diagnostic test kit for Helicobacter pylori was distributed for adult and pediatric use, but was not approved for pediatric use. | Labeling False and Misleading | Customers were notified by letter on 05/02/2007 and asked to remove the software program from any computers on which it had been installed and return the software disc and User's Guide to Meretek. |  | 92 | 92 units | Nationwide. | Terminated | November 26, 2007 | 208 | 92 | Not_Computer
Z-0033-2008 | 39402 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Smith + Nephew, Birmingham Hip Resurfacing, Acetabular Cup, 62 mm cup Ref.: 74120262. Hip implant component. |  |  | July 12, 2007 | 2007 | Smith &amp; Nephew Inc | The carton label may not reflect the correct size of the acetabular cup. | Packaging process control | The consignees were notified of the problem and the recall by initially by telephone on 07/12/2007 and followed by a letter dated 07/24/2007. |  | 1 | 1 unit | Worldwide, including USA, Australia, Belgium, Canada, Finland, France, Germany, Italy, The Netherlands, Norway, South Africa, Spain, Switzerland, and United Kingdom. | Terminated | February 02, 2010 | 936 | 177 | Not_Computer
Z-0012-2008 | 39369 | No | GES | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Shelf Pack: |  |  | August 10, 2007 | 2007 | Becton Dickinson &amp; Company | Complaints received of rust on the blades of BD Bard-Parker Surgical blades. | Packaging change control | Urgent recall letters were sent out 8/10/2007 via UPS second day air. |  | 4026 | 4026 units | Nationwide. | Terminated | March 07, 2008 | 210 | 121452 | Not_Computer
Z-0063-2008 | 44724 | No | DQA | Oximeter. | Cardiovascular | Anesthesiology | 510(k) | Medtronic Export XT Aspiration Catheter, 6F (US) |  |  | August 29, 2007 | 2007 | Medtronic, Inc. | Additional Warning in Instructions for Use for guide wire management. Complaints of wire lumen tearing during use or band marker becoming dislodged. | Labeling design | Medtronic has notifed sales representatives to visit Cath lab site with the letter dated August 29, 2007 and educate the user, provide the Safety Alert Letter and document delivery of the letter with the confirmation form. |  | 13940 | 13,940 units | Worldwide, including USA, Canada, India, Austria, Belgium, France, Germany, Italy, Netherlands, Spain, Sweden, Switzerland, and UK. | Terminated | September 08, 2008 | 376 | 13940 | Not_Computer
Z-1091-2007 | 38219 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Colleague Triple Channel Volumetric Infusion Pumps; Baxter Healthcare Corporation, Medication Delivery Division, Deerfield, IL 60015 U.S.A., Made in Singapore; product codes 2M8153, 2M8163, 2M9163 |  |  | June 20, 2007 | 2007 | Baxter Healthcare Corp. | A software anomaly is causing newly upgraded Colleague Triple Channel Infusion Pumps to alarm, display an error code (16:310:867:0002) and stop the infusion. This occurs during user programming with all three channels simultaneously infusing fluids. In reported cases, the pump stopped infusing and both an audible and a visual alarm notification were activated. | Software design | Baxter telephoned the direct accounts on 6/20-21/07, informing them that Baxter has received multiple reports of newly upgraded Colleague Triple Channel Infusion Pumps alarming, displaying an error code and stopping infusion. These reported events occurred during user programming with all three channels simultaneously infusing fluids. In all cases, the pump stopped the infusions and both an audible and visual alarm notification were activated. The accounts were requested to immediately remove the affected triple channel pumps, identified by two attached hangtags entitled "Battery" and "Power On", from service. The accounts were informed that their Baxter representative would be in contact with them to coordinate the replacement of the affected pumps with upgraded single channel pumps or non-upgraded triple channel pumps, if available. Any questions were directed to Baxter's Medication Delivery Services at 1-800-843-7867. Baxter sent follow-up Urgent Product Recall letters dated 6/22/07. |  | 4512 | 4,512 pumps | Worldwide, including USA, Canada, Australia, Belgium, New Zealand, South Africa, Thailand, and Singapore. | Terminated | December 20, 2011 | 1644 | 4512 | Computer
Z-0364-2008 | 38346 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Accuray, Multiplan Treatment Planning Software, Versions 2.0, 2.0.1 .6, 1.6.1, 1.6..1 or 1.6.3., Accuray, Sunnyvale, CA 94089. |  |  | March 21, 2007 | 2007 | Accuray Inc | Software problem - Sagittal, oblique or coronal sliced MR Images imported from a non-Accuray manufactured device into the product may result in inverted image orientation. | Software design | The firm issued an Urgent Advisory Notification on March 20, 2007, notifying users to only import axial sliced MR image series for treatment planning. Users should not import sagittal, oblique or coronal images series until further notice. The firm intends to issue a software patch to correct the problem and will contact user regarding resolution. Users can contact Accuray Customer Support at 1-877-668-8667 (USA) or 1-408-716-4700 (non-USA) or customersupport@accuray.com. |  | 25 | 25 units | Worldwide Distribution - USA including states of MD, CA, FL, WA, LA, IL, IN, MI, MN, NC, NJ, OK, WI, and OH, and countries of Korea, Greece, Germany, France, and Turkey. | Terminated | December 26, 2007 | 280 | 25 | Computer
Z-1095-2007 | 36812 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ABL 800 FLEX series blood gas analyzers containing software version 5.24. |  |  | November 13, 2006 | 2006 | Radiometer America Inc | Software anomaly. Protein removing program incorporated in the operating software for Versions: V3.833 and V5.24, of the ABL700 and ABL800 Blood Gas Analyzers, may allow for protein build up on the electrode membranes of blood gas analyzers. This condition can influence analyzer results. When the protein removal program is run too frequently, more than once a week, or when running the protein rem | Software design | The firm issued a 'Safety Alert Notice' [FAN 915-248] to all customers which received the recalled ABL700 and ABL800 Analyzers equipped with the referenced software. The 11/13/2006 notice describes the problem concerning the Protein Removal program to the user/customers and emphasizes that the frequency of running the program in these software versions must be no more than once a week in order to prevent the protein build-up problem. It also contains additional recommendations for the user/customers for operation of these analyzer and informs the customers that a Radiometer Service Rep. will be coming out and making a software upgrade to these systems within the next few weeks. The alert notification is accompanied by a 'FAX FORM' to be signed and completed by each customer in order to verify that the alert notification has been received and that the staff members responsible for operating and maintaining the analyzer(s) on-site have been informed of this situation and that the necessary changes in the Operator's Manual concerning the Cleaning, Protein Removal, and Decontamination programs have been made. |  | 464 | 464 | The recalled analyzers were distributed to customers nationwide throughout the U.S. and in some parts of Canada. | Terminated | May 09, 2012 | 2004 | 481 | Computer
Z-0451-2008 | 45843 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Philips Medical Systems, M3290A IntelliVue Information Center Software for M3170 Intellivue Patient Link, Catalog Number: 865007 and 865015; Running Software Versions J.00.23, J.00.24, J.00.25, and J.00.26, Manufactured and Distributed by; Philips Medical Systems, Andover, MA 01810 |  |  | November 13, 2007 | 2007 | Philips Medical Systems | Alarm failure : If changes are made to the Caregroup Editor configuration, Care Group alarm status, including alarm text and prompt tones, may fail to be displayed or annunciated at bedside monitors within the Care Group | Software change control | Philips Medical notified customers by letter dated November 2007 titled : Urgent Medical Device Correction Notice. Customers will be instructed to follow the Procedure to Mitigate Risk section * of the notice while they await the correction of their devices with a mandatory software upgrade to be performed by service representatives. * (no changes are to be made to the Caregroup Editor configuration on any system. If changes have already been made to the configuration prior to receipt of this notification, turn each bedside monitor "off " and then "on" in order to reset the system-Customer Care Contact telephone 1-800-722-9377) |  | 60 | 60 units | Worldwide Distribution -USA including states of CA, MA, TX, IN, MI, FL, MN, NC, MS, HI, RI, NH, and PA, and countries of Canada, Australia, Belgium, France, Germany, Ireland, Italy, Japan, Switzerland, United Arab Emirates, and The United Kingdom. | Open, Classified |  |  | 60 | Computer
Z-0147-2008 | 44961 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ARCHITECT c16000 Processing Module, for In Vitro Diagnostics; List Number 3L77-01; Affected software: v2.10, v2.11, v2.12, v2.20, v2.20DB, v2.30, v2.60, v3.00 and v3.10; Distributed by Abbott Laboratories, Irving, TX. |  |  | September 05, 2007 | 2007 | Abbott Laboratories, Inc | Incorrect Sample/Test Identification: Due to a software timing defect with certain versions of ARCHITECT cSystem Software, test results may be incorrectly associated to a Sample Identification during specific sample-processing conditions of the Robotic Sample Handler. | Software change control | Notification letters sent to consignees on 09/17/07. Consignees informed of the software defect and the conditions that would cause the software error to occur. Consignees told they would be contacted by an Abbott Field Service Rep who would install ARCHITECT System Software v3.11 and were asked to respond via fax using Customer Reply form acknowledging receipt and understanding of correction letter. Firm recalling to hospital/user level. |  | 89 | 89 units. | Worldwide Distribution: USA (including Puerto Rico), Argentina, Australia, Bahamas, Bermuda, Brazil, Cayman Islands, Canada, Chile, China, Colombia, Dominican Republic, Ecuador, El Salvador, Germany, Guatemala, Hong Kong, Jamaica, Mexico, New Zealand, Panama, Singapore, Thailand, Trinidad and Tobago, Uruguay, and Venezuela. | Terminated | October 04, 2010 | 1125 | 2159 | Computer
Z-0449-2008 | 45540 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Triathlon MIS Femoral Adjustment Block;Catalog Number: 6541-5-601;Stryker Orthopaedics, Mahwah, New Jersey |  |  | September 19, 2006 | 2006 | Stryker Howmedica Osteonics Corp. | Assembly problems: MIS instruments (Triathlon MIS AP Sizer Body - Left; Triathlon MIS Sizer Body - Right; Triathlon MIS Femoral Adjustment Block and Triathlon MIS Modular Capture) may not assemble properly. | Device Design | Important Product Correction letters and Product Accountability Forms were sent via Federal Express on 9/19/06 with return receipt. Foreign locations were notified by a Advisory Notice. Notification letters requested that all inventory of the affected items be examined to verify that it assemble correctly. Any product that does not assemble correctly was to be returned to Stryker, Mahwah, NJ. The Product Accountability form was to be completed by each location to notify Stryker that the letter ahd been received and inventory examined. Stryker representatives were to be contacted for product replacement. |  | 2868 | 2,868 units total for Recalls Z-0448-0450-2008 | Worldwide: USA including states of MI and IL. and countries of Australia, China, New Zealand, Singapore, France, Germany, Italy, Portugal, Spain, Sweden and UK. | Terminated | January 18, 2008 | 486 | 2868 | Computer
Z-0241-2008 | 38307 | No | LHN | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | OSI, Allegro 6800 Imaging Table (Operating), single pedestal AC or battery operated table, Orthopedic Systems, Inc., Union City, CA |  |  | February 28, 2007 | 2007 | Mizuho Orthopedic Systems Inc | Failure to anchor; due to a failure in the locking mechanism, the product may not anchor properly. | Software design | The firm issued a Product Correction Notification letter by mail on 02/28/2007 and follow up with phone calls. The letter notified users that a required no-cost upgrade of the removal and replacement of a microprocessor that controls the logic for the floor locking mechanism will be conducted by the recalling firm. |  | 117 | 117 units | Worldwide: USA product was distributed to 16 wholesale dealers, 212 medical facilities, 14 distributors throughout the US in the following states: AL, AZ, CA, FL, GA, HI, IL, IN, KS, KY, MA, MD, ME, MI, MO, NC, NE, NH, NJ, NM, NY, OH, OR, PA, SC, TX, VA, WA, WI, WV and WY and to the following countries Taiwan, Australia, German, Japan, Canada, Saudi Arabia, and The UK. | Terminated | December 18, 2008 | 659 | 343 | Computer
Z-0671-2007 | 37490 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | LIFEPAK CR Plus defibrillator. The device is a battery operated, portable automated external defibrillator (AED). The devices are designed for non-professional users and are primarily used in public access defibrillator environments such as airports, schools, or public libraries. |  |  | March 02, 2007 | 2007 | Medtronic Emergency Response Systems, Inc. | LIFEPAK CR Plus defibrillators with software version 1.16 or lower may miscalculate the HLC battery capacity and result in the CHARGE-PAK battery charger icon/indicator to display and falsely indicate battery needs replacement. | Software design | Firm sent letters on 03/02/2007 dated 03/2007 via certified return receipt to consignees with affected LIFEPAK CR Plus devices. The firm advised consignees to upgrade their devices' software, and indicated that each consignee will be provided a user-installable software upgrade for their device(s) within 45 days. |  | 13835 | 13,835 units | Distribution is worldwide. The device is designed for non-professional users and are primarily used in public access defibrillator environments such as airports, schools, or public libraries. | Terminated | April 26, 2012 | 1882 | 13835 | Computer
Z-0695-2007 | 37306 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Gambro Prismaflex Continuous Renal Replacement System, Catalog Number 602314700, Gambro Lundia AB, Sweden. |  |  | January 04, 2007 | 2007 | Gambro Renal Products, Inc. | Under certain conditions, * an excessive amount of anticoagulant may be infused into the patient. (* interruption of a self-test by an alarm and obstruction of the access lines) | Device Design | Customers were notified by letter on 12/19/2006. They were told that Gambro would schedule the required upgrades to their equipment and not to use the current syringe for delivering anticoagulant. |  | 153 | 153 units (Domestic) | Nationwide distribution -- including states of AL, CA, FL, GA, IN, MD, ME, MI, MS, NC, NH, NY, OH, PA, SC, TX, VA, and WI. | Terminated | January 03, 2012 | 1825 | 153 | Computer
Z-1045-2007 | 38072 | No | GEH | Cryosurgical unit and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Wallach Cryotip T-00337 Anorectal Tip used with the Cryosurgical LL100 or LL50X Cryosurgical Systems |  |  | June 06, 2007 | 2007 | Wallach Surgical Devices, Inc. | The tip may separate from the probe. | Device Design | Wallach Surgical notified accounts by fax on 6/6/07. A follw-up telephone call will be made to the 70 domestic consignees. Foreign consignees notiifed by fax or email with follow-up telephone calls. Accounts are requested to return product and replace with an alternate product. |  | 315 | 315 units | Worldwide, including USA, Canada, Australia, England, Greece, Jordan, Turkey, Hungary, Netherlands, Saudi Arabia, Spain, and Lithuania | Terminated | July 15, 2015 | 2961 | 329 | Computer
Z-0999-2007 | 37587 | Yes | NVZ |  |  | Cardiovascular | PMA | Vitatron C-Series C60DR DDDR Dual Chamber Rate Responsive Pacemaker includes the following models: C60A1 and C60A3. The Vitatron C60 DR is a dual chamber rate responsive pacemaker (activity sensing using an accelerometer) for permanent atrial and ventricular pacing. C-series software is VSF11/VSF12 Version 1.0. Vitatron is a division of Medtronic, Inc., Sales Office is Vitatron USA, 7000 Central Ave., NE, Minneapolis, MN 55423-3576. The head office is Vitatron B.V., P.O. Box 5227, 6802 EE Arnhem, The Netherlands. |  |  | February 05, 2007 | 2007 | Medtronic Inc. Cardiac Rhythm Managment | Pacing rate; Software anomaly affecting Vitatron dual chamber C-series and T-series pacemakers, related to automatic retrograde conduction testing, can inhibit the device from pacing if the patient's intrinsic rate falls below the programmed lower rate. | Software design | Advisory letters to physicians were mailed in February 2007. Sales respresentatives will follow-up with physicians by visit. A programmer software update will be released upon regulatory approval. The Advisory Letter provided short-term direction to reduce problems pending software upgrades. |  | 4294 | 4,294 | within the US to include: AL, AR, AZ, CA, CT, FL, GA, IL, IN, KY, MA, MD, MI, MN, NC, ND, NJ, NY, OH, OK, SC, TX. | Terminated | October 16, 2008 | 619 | 10553 | Computer
Z-1243-2007 | 44739 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Cerner Millennium RadNet Auto Launch Study and Auto Launch Report software functionalities. |  |  | August 23, 2007 | 2007 | Cerner Corp | Defects in the Auto Launch functionality make it possible for a mismatch of patient data. | Software design | The recalling firm initiated recall via telephone on 8/23/07 to their customers who had the 2007 version in their production environment, explained the reason for recall, and informed them they should turn off the affected functionality. The firm posted details of the situation on their website in two different documents (referred to as a Priority Review Flash), one dated 8/29/07 and the other dated 8/31/07. The 8/31/07 document was also mailed to their customers via certified mail on 9/4/07.The corrected software (EP28822 ver.1) was made available to the firm's clients on 9/7/07 via their website. Beginning 9/18/07, workarounds were discussed via telephone with the clients who had downloaded the software correction which contained a new bug that affected the dictation file. A new software correction to fix the dication file problem was posted on the recalling firm's website on 9/28/07 for clients to download and validate. A Priority review Flash #PR07-0176-2 client notification dated 10/4/07 was also posted on their website and mailed via certified mail return receipt. |  | 9 | Downloaded by 9 customers nationwide. | Distribution was made to health care facilities located in LA, MO, NJ, MS, MI, NE, ND, and AR. | Terminated | November 29, 2007 | 98 | 9 | Computer
Z-0996-2007 | 38130 | No | HNN | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | Premier Edge MVR 20g Microsurgical Knife, Item PE-3520 |  |  | June 08, 2007 | 2007 | Oasis Medical Inc | This voluntary recall has been initiated due to the product puncturing the packingduring rough handling. Once the packaging is damaged, the product is no longersterile as labeled. | Device Design | Notification of the recall sent to distributors and users. The 06/08/07 letter requests customers to notify down to the user level. Included in theRecall Notification Letter are instructions: "In addition, if you may have further distributed this product, please identify your customers and notify them at once ofthis product recall. Your notification to your customers may be enhanced by including a copy of this recall notification letter. This recall should be carried out to the user level. Your assistance is appreciated and necessary to prevent any inconvenience." |  |  |  | Worldwide, including USA, Switzerland, Germany, England, UK, Turkey, Czech Republic, and Lithuania. | Terminated | September 25, 2009 | 840 |  | Computer
Z-0074-2008 | 44787 | No | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Medtronic Sofamor Danek, TSRH SiLo 5.5 Spinal System, Sagittal Adjusting Screw Implants, Supplement Sets, REF: 8889245, Pedicle Screw Fixation System, Medtronic Sofamor Danek, Memphis, TN 38131 |  |  | August 24, 2007 | 2007 | Medtronic Sofamor Danek USA Inc | Disassembly: The movable cradle component could detach from the saddle under extreme manipulation of the associated rods prior to final tightening of the set screws. | Device Design | Medtronic sent letters, titled URGENT DEVICE RECALL, on 08/24/2007, to the Risk Managers of the hospital consignees explaining the potential for these sets to disassemble, and requesting the discontinuation of the use of and quarantine of the product. The product will be retrieved by the firm and replaced. |  | 20 | 20 units | Worldwide Distribution - USA, The Netherlands, Australia, and Canada. | Terminated | August 21, 2009 | 728 | 20 | Computer
Z-0905-2007 | 37811 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | FramelessArray software, Version 1.0. ; Medical Device firmware incorporated with the RadioCamera Extracranial System, distributed by Varian Medical System, Palo Alto, CA. |  |  | March 14, 2007 | 2007 | Varian Medical Systems Oncology Systems | Localization error; with planning data transferred to the Optical Guidance Platform via DICOM RT; The software incorrectly computes the center of the CT volume, resulting in a potential axial error ranging from 0.3 mm to 1.5 mm, affecting both Fractionated and Stereotactic Radiosurgery (SRS) Treatments | Software design | The recalling firm notified consignees by certified letter on 03/14/07. The notification was flagged as "Urgent Medical Device Correction" and advised end users that a localization error may occur when using the FramelessArray and affect its accuracy for Frameless treatments. Users were instructed to verify software in the devices and if running FramelessArray version 1.0, to discontinue use of the software when Stereotactic Radiosurgery plans are transfered via DICOM RT. The notification offered that firmware would be upgraded during scheduled service visit. |  | 178 | 178 devices | Worldwide : The software was distributed to direct accounts in the US and internationally distributed to accounts in Canada, Australia and Spain. | Terminated | July 03, 2008 | 477 | 178 | Computer
Z-1195-2007 | 38349 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian On-Board Imager Device versions 1.2 and 1.3, when used with Cone Beam Computed Tomography. Medical Linear Accelerator. |  |  | April 18, 2007 | 2007 | Varian Medical Systems Inc | Patient positioning errors may occur because of mechanical variations during product installation. | Device Design | The firm issued a notification letter to all customers April 18, 2007. The firm intends to arrange for service personnel to contact affected sites to verify the calibration of the OBI system, and make arrangements to correct any issues. The firm will also update installation manuals to clearly describe the OBI installation and calibration requirements. |  | 291 | 291 units | Worldwide, including USA, Puerto Rico, Australia, Belgium, Canada, China, Colombia, Czech Republic, Denmark, Finland, France, Germany, India, Ireland, Italy, Japan, Luxembourg, New Zealand, Norway, Poland, Portugal, Spain, Sweden, Switzerland, Netherlands, and United Kingdom. | Terminated | August 12, 2010 | 1212 | 291 | Computer
Z-0312-2008 | 44916 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | COULTER LH750 Analyzer, Part Number 6605632, Beckman Coulter, Brea, CA. (Beckman Coulter, Brea, CA. (Hematology analyzer) |  |  | April 18, 2007 | 2007 | Beckman Coulter Inc | Patient Mis-identification can occur:1)- When a positive identifier (Sample ID or Cassette/Position) is manually edited to a positive identifier that is already in the To Do list. The workstation will accept the entry and no error message will be generated, creating 2 samples with the identical positive identifier.2)- When manually entering a Patient ID, if a blank space is entered between th | Software design | Product Corrective Action (PCA) letters were mailedon April 18,2007, to all GENS, LH500, LH750 andLH780 accounts. They are informed that a potentialmisidentification can occur in 2 scenarios.1. When a positive identifier (Sample ID orCassetteIPosition) is manually edited to apositive identifier that is already in the ToDo list.The workstation will accept the entry and noerror message will be generated, creating 2samples with the identical positive identifier.2. When manually entering a Patient ID, if a blankspace is entered between the characters in thePatient ID field, the system will only accept thecharacters before the space.----Short term fix-----The letter explains the issues and provides thecustomer work around. (Letters were sent by USmail). A customer response form was provided. The root cause of the problem is still under investigation, software upgrades, when approved will be made to devices. |  | 2344 | 2,214 units in US; 130 units in Canada | World wide: USA and Canada | Terminated | April 16, 2012 | 1825 | 4294 | Computer
Z-1146-2007 | 38355 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ISOLOC 6.5 Software Part Number NT-NW-425-520. Software for radiotherapy treatment. |  |  | April 07, 2006 | 2006 | Northwest Medical Physics Equipment Inc | This correction was for several software bugs which caused image format problems; incorrect ruler display; software crashes; failure to update pixel spacing; display of unattainable treatment table positions; failure to update gantry angles; wrong shifts for two of three axes of treatment table movement. | Software design | Software to fix the problems was written and software users were notified of the availability by e-mail on 04/07/2006. They could download the fixed software from the firm's internet website. |  | 141 | 141 copies | Customers in USA, Korea, Spain and UK. | Terminated | April 19, 2011 | 1838 | 141 | Computer
Z-0366-2008 | 44790 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Philips Medical Systems, Cardiac Viewer or Pulmonary Viewer Application that can run on all Gemini TFs, GXLs and EBW Workstations. Philips Medical Systems, Cleveland, OH 44143. (The product is a software application that is installed in Philips Medical System CT scanners.) |  |  | September 10, 2007 | 2007 | Philips Medical Systems (Cleveland) Inc | Incorrect Measurement: Incorrect measurement will occur when the operator uses the distance or area measurement function, such as Graphic Tool line or ROI on zoomed and batch saved images from the Cardiac Viewer or Pulmonary Viewer and then views them in the CT Viewer or on a PACS System. | Software design | Consignees were notified of this recall by an Urgent Device Notification letter sent on 9/10/07. The letter instructed the users to not save any zoomed images to a batch file. Make all measurements of zoomed images in the CT viewer, which provides correct measurements. Do not perform measurements on any zoomed batched images that have been stored on a PACS System. Phillips in investigating the problem and will provide updates free of charge when completed. |  | 2544 | 2,544 | Worldwide Distribution - USA including states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, VT, WI, WV, and WY, and country of Canada. | Terminated | September 19, 2012 | 1836 | 2544 | Computer
Z-1153-2007 | 38361 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ISOLOC Software Version 6.0, Localization software for determining treatment machine position (to position patients) for high accuracy radiotherapy. |  |  | April 11, 2005 | 2005 | Northwest Medical Physics Equipment Inc | When ISOLOC version 6.0 was used in the centimeter mode, the display in the couch moves window had a rounding error relative to the Portal Image Marker window of up to 0.5 mm. | Software design | Software to fix the described problem was written and customers contacted by e-mail on 04/11/2005, that a software patch, titled 04112005, was available to their customers on the firm's internet website. |  | 84 | 84 | Worldwide, including USA, Korea, Spain, and UK. | Terminated | April 19, 2011 | 2199 | 84 | Computer
Z-1145-2007 | 38319 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Isoloc software Versions 6.5 and 4.5, Part Number NT-NW-425-520. Software for image guided radiation therapy. |  |  | January 02, 2006 | 2006 | Northwest Medical Physics Equipment Inc | Isoloc version 6.5 and 4.5 software issued a false warning message regarding Gantry angles when set to "Anatomic Landmark" setting. | Software design | Software to fix the described problem was written and customers contacted by e-mail on 01/06/2006 that a software patch, titled 01022006 was available to their customers on the firm's internet website. |  | 132 | 132 | Nationwide, Korea, Spain, and UK. | Terminated | April 19, 2011 | 1933 | 132 | Computer
Z-0148-2008 | 45398 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | Ellex, Eye Cubed, I 3 System-ABD Ultrasound Version 3.0 with Software Version 3.0 and Biometry A-scan modality, Ellex Innovative Imaging, Sacramento, CA 95827 |  |  | September 25, 2007 | 2007 | Ellex Inc. | Incorrect IOL calculations: software behavior with Version 3.0 Eye Cubed Ultrasound systems running Biometry A-scan enabled Version 3.0.0 software can cause incorrect IOL calculations. | Software design | A Global Customer Service Bulletin (GCSB)-Urgent Notification of Software Correction was issued on 25 September 2007, which provided short term workaround directions pending a free software undate promised in the near future. |  | 59 | 59 | Worldwide; including Canada and Mexico. U.S.A. distribution included states of AL, AZ, CA, CT, FL, GA, IL, MA, ME, MI, MN, NC, NE, NJ, NM, NY, OH, OR, PA, SC, TN, and TX. | Terminated | September 08, 2008 | 349 | 59 | Computer
Z-1154-2007 | 38363 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ISOLOC Software Version 6.0, Part Number NT-NW-425-520. Localization software for radiotherapy treatment system. |  |  | December 30, 2004 | 2004 | Northwest Medical Physics Equipment Inc | When ISOLOC version 6.0 software was used in the Anatomic Landmark mode, the couch moves display in the Couch Translator window was incorrect. | Software design | Software to fix the described problem was written and customers contacted by e-mail on 12/30/2004 that a software patch, titled 12302004 was available to their customers on the firm's internet website. |  | 79 | 79 | Worldwide, including USA, Korea, Spain, and UK. | Terminated | April 19, 2011 | 2301 | 79 | Computer
Z-1152-2007 | 38360 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ISOLOC Software Versions 6.0 and 4.2. Part Number NT-NW-425-520. Image guided radiotherapy treatment software. |  |  | April 28, 2005 | 2005 | Northwest Medical Physics Equipment Inc | When ISOLOC version 6.0 and 4.2 distance unit is set to centimeter, the localization coordinates were rounded to the nearest millimeter which could affect positioning by 0.5 mm. | Software design | Software to fix the described problems was written and customers contacted by e-mail on 04/28/2005, that a software patch, titled 04282005, was available on the firm's internet website. |  | 84 | 84 | Worldwide, including USA, Korea, Spain, and UK. | Terminated | April 20, 2011 | 2183 | 84 | Computer
Z-0319-2008 | 44892 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Biomet Optigun Cement Gun, Ratchet, non-sterile; REF 419500 and REF 4195 (in Europe), Biomet, Inc., Warsaw, IN |  |  | September 21, 2007 | 2007 | Biomet, Inc. | Loose Screw Component: A screw may come loose and fall from the device during use, and result in implantation with the bone cement. | Device Design | Consignees were notified via Urgent Medical Device Correction letter dated 9/21/07 to assure the screw components are fully seated and secure, to not use any guns with loose or missing screws, and to contact Biomet for instructions to return all instruments for evaluation and refurbishment. |  | 6590 | 6,590 worldwide (2,301 to U.S.) for both recalls Z-0318-0319-2008 | Worldwide: USA, Canada and Japan. | Terminated | June 30, 2009 | 648 | 6590 | Computer
Z-1155-2007 | 38366 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ISOLOC Software Version 4.2, Part Number NT-NW-425-520. Image guided radiation treatment software. |  |  | November 30, 2004 | 2004 | Northwest Medical Physics Equipment Inc | When ISOLOC Version 4.2 unit is set to centimeter and the user clicks the Reenter button in the digitizer window, the coordinates of the markers and target in the anatomic imaging system erroneously displays values ten times greater than the correct values. | Software design | Software to fix the described problem was written and customers contacted by e-mail on 11/30/2004, that a software patch, titled 11302004, was available to their customers on the firm's internet website. |  | 83 | 83 | Worldwide, including USA, Korea, Spain, and UK. | Terminated | April 19, 2011 | 2331 | 83 | Computer
Z-0240-2008 | 38256 | No | FIN | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Hill-Rom VersaCare Hospital Bed System; Models P3200 and P3201, Hill-Rom, Batesville, IN 47006 |  |  | November 12, 2006 | 2006 | Hill-Rom, Inc. | False Latch/Failure to Latch-The side rail may not latch, or the latch may not fully engage, resulting in the siderail being able to be lowered without pulling on the release handle. | Device Design | The firm notified consignees via an Urgent Device Correction letter dated 11/12/07 informing them of the problem and informed them they would be receiving upgrade kits which they could have their qualified personnel install themselves following video instructions also provided, or they could contact Hill-Rom for assistance. |  | 43731 | 43,731 beds total for both models were distributed; 3,282 have been corrected to date. | Worldwide, including USA, Argentina, Australia, Brazil, Canada, Dominican Republic, Emirate of Sharjah, England, Greece, Indonesia, Japan, Jordan, Kuwait, Lebanon, Malaysia, Mexico, Oman, Pakistan, Philippines, Qatar, Saudi Arabia, Selangor, Singapore, South Korea, Spain, Taiwan, Thailand, Turkey, United Arab Emirates and Venezuela. | Terminated | October 22, 2010 | 1440 | 43731 | Computer
Z-0287-2008 | 45499 | No | HQF | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | Ellex Integre Duo, Model No: LP1RG - Ellex Medical PTY, LTD 82 Gilbert Street, Adelaide, S.A. 5000, AUSTRALIA. (The Integre Duo is a dual-wavelength surgical photocagulator laser intended for use in a clinic, outpatient clinic, or a retinal specialist's office. ) |  |  | June 13, 2007 | 2007 | Ellex USA | Electric Shock from Power Supply: The power supply has a defect which can cause an electrical shock to user. | Device Design | Ellex, Inc. called each customer on 6/13/07, about the problem (defect which can cause electric shock) and advised that they would be performing modification of the power supply in the laser system and scheduling a field service visit to complete the work. Upon completion of the modification, the service technician completed the service record and left a copy with each client and informed them of the completed work. A copy of each completed service record was sent to Ellex Medical Pty. Ltd. On 9/27/07, Ellex sent a letter to all consignees, informing them that all modifications by service technicians' have been completed, and that their Solitaire LP4532 laser and/or Intregre Duo LP1RG is up to the manufacturers specifications and is in good operating order. |  | 1 | 1 | Worldwide Distribution - USA including states of WI, PA, CO, CA, NC, NY, SC, MD, and MS, and countries of Canada and Brazil. | Terminated | October 17, 2008 | 492 | 11 | Computer
Z-0304-2008 | 44913 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | c.cam Gamma Camera, Model No: 9VIR1200; Tomography System (Emission Computed); Siemens Medical Solutions USA, Inc., Molecular Imaging, 2501 N. Barrington Road, Hoffman Estates, IL 60195-5203; Made in Denmark |  |  | September 17, 2007 | 2007 | Siemens Medical Solutions USA, Inc. | Unintended Motion: A hardware fault in the c.cam system can cause the system to experience unintended motion, thereby causing a collision hazard. | Software design | Urgent Field Correction Recall letters dated 9/11/07, on Siemens letter head were sent to the end users via U.S. certified mail, informing them of the potential risk of unintended motion of the detector head resulting in a collision hazard. The end users were warned not to press any of the detection motion buttons while simultaneously holding the collision over ride button on the c.cam hand controller, and were provided steps to take if the system stops during an acquisition and displays the Collision Error message. They were told that a Siemens service representative will contact them within 90 days to schedule the repair of the c.cam at no cost to the end user. Any questions regarding the Important Safety Notice were directed to their service representative or Siemens office at: 1-800-888-7436 (USA). |  | 486 | 486 units | Worldwide: USA including Puerto Rico and Brazil, Canada, France, Germany, Great Britain, Israel, The Netherlands, and Portugal. | Terminated | January 09, 2009 | 480 | 486 | Computer
Z-1142-2007 | 38223 | No | LZH | Infusion pump. | General Hospital | General Hospital | 510(k) | The Companion ClearStar Pump System, for enteral feeding. |  |  | June 21, 2007 | 2007 | Ross Products Division Abbott Laboratories | The transducer adapter bracket on the pump could crack or break and potentially result in under delivery of the enteral product. | Equipment maintenance | Consignees were notified by letter dated 6/25/2007. A revised letter was sent 07/26/2007. |  | 27814 | 27,814 | Worldwide, including USA and Canada. | Terminated | September 26, 2012 | 1924 | 27814 | Computer
Z-0580-2007 | 37394 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Defibtech Sentry Semi-Automatic External Defibrilator (AED) sold as 'Lifeline AED' brand -Model DDU-100A (version 2.002 and earlier) |  |  | February 17, 2007 | 2007 | Defibtech, LLC | Device Malfunction-The self-test software may allow a self-test to clear a previously detected low battery condition. If this occurs, the operator may be unaware of the low battery and the device may be "unable" to deliver a defibrillation shock, which could result in failure to resuscitate a patient. | Software design | Defibtech LLC initiated notification to customers/distributors on 2/17/07 by email. Hard copy letters followed on 2/22/07 with acknowledgement card. Users provided with instructions for 'Immediate Recommendations' until the unit''s software is upgraded. Distributors, their customers, and direct accounts instructed to locate affected units and upgrade these units in the field. Distributors, theircustomers, and direct accounts shall return to Defibtech documentation confirming unitsoftware upgrade was completed..Press was issued by the firm on March 6, 2007 |  | 22197 | 22,197 units | Worldwide distribution ---- USA and countries of Canada, Uruguay, Jordon, Turkey United Kingdom, Saudi Arabia, Sweden France, South Africa, Cyprus, Spain, Germany India, Japan, Israel, Australia, Norway, Philppines Singapore, Poland, Germany, Italy, Switzerland Netherlands, Indonesia, United Arab Emirates Malaysia, Austria, Ireland, Kingdom of Bahrain, and Korea. | Open, Classified |  |  | 36219 | Computer
Z-1172-2007 | 38096 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Non-AC-powererd Patient Lifts for assisting patients when bathing and toileting-sold under names "Bath Helper" and "Shower Helper" , Sisus Corporation, Santa Cruz, CA 95062. |  |  | June 01, 2007 | 2007 | The Sisus Corporation | Poor welding may cause the device(s) to fail, which may result in the user, falling in the bathroom and/or a lack of assistance with bathing or toileting. | Device Design | On 6/8/2007, a letter was sent to all dealers and each direct account customer/user followed by a phone call to ensure the notification letter was received. Customers were instructed to discontinue use of the devices pending receipt of modification kits/ instructions to be shipped priority overnight. The firm plans to call customers/dealers to support them with the installations and ensure smooth implementation. Dealers were instructed to conduct recall to the user level. |  | 100 | approximately 100 units have been distributed. | Nationwide Distribution including states of PA, TX, NY, MI, AL, FL, CA, OH, CT, WA, MD, OR, NC, NH, VT, IL, CO, HI, LA, GA, TN, OK, MA, KS, SC, ME, GA, MN, IN, and DE. | Terminated | August 19, 2008 | 445 | 100 | Computer
Z-0138-2008 | 44953 | No | MJO |  |  | Orthopedic | PMA | Medtronic Sofamor Danek, Prestige Cervical Disc System, Implant Trial, REF/Catalog #: 6972226, SIZE 6 X 12, MATL: Stainless Steel/Titanium, Rx only, Manufactured at: Bartlett, Tennessee USA |  |  | September 04, 2007 | 2007 | Medtronic Sofamor Danek USA Inc | Implant mis-seating; Variance in size between the trial and the implant could cause the implant to be improperly seated. | Device Design | The firm notified their consignees of the problem and the recall by Urgent Medical Device Recall letter, telephone calls and electronic mail on 09/04/2007. Instructions include requests to immediately quarantine the product and hold for recalling firm collection. |  | 196 | 196 units | Nationwide | Terminated | November 03, 2008 | 426 | 196 | Computer
Z-1267-2007 | 38634 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Davol Salut Fixation Reusable 38cm Shaft (Laparoscopic) |  |  | July 16, 2007 | 2007 | Davol, Inc., Sub. C. R. Bard, Inc. | Malformed constructs (straight shots) may form causing injury to user/patient | Device Design | Davol issued recall letter notification on 7/16/07 via Federal Express. Sales of the disposable cartridges ceased July 31, 2007. Cartridges will not be requested to be returned to Davol. Customer requests to return cartridges will be managed on an account-by-account basis. A reminder Second Letter issued July 23,2007 as a follow-up to the letter of July 16, 2007.The second stage will be a series of three letters to customers following the cessation of disposable cartridge sales and to recall the re-usable handles. A Stock Status form will be provided to customers to inventory the product which they are returning.The Stock Status form will be used to document the handles being returned. In the event that the customer cannot locate one or more handles. All handles will be returned to Cranston and will be dispositioned. Account verification for the number of handles returned from the account will be performed. No verification testing will be performed on returned handles unless a formal complaint has been registered for a handle. |  | 4242 | 4,242 shafts | Worldwide-USA and Canada, Austria, Switzerland, Spain, Sweden, United Kingdom, Italy,Portugal, Greece,Holland, South Africa, Israel, Japan, Belgium, Denmark, Finland, France, and Germany | Open, Classified |  |  | 195104 | Computer
Z-1109-2007 | 38342 | No | DSQ |  |  | Cardiovascular | PMA | Thoratec Paracorporeal Ventricular Assist Device (PVAD) component of the Thoratec VAD system. Catalog number 14086-2550-000. |  |  | March 30, 2007 | 2007 | Thoratec Corp | Incorrect collect nut assembly (inadvertent use of the incorrect (Black) collect nut) during surgical implantation (wrong size collect nut) to secure the inflow cannula to the PVAD has resulted in patient deaths | Device Design | The firm has issued an Urgent Medical Device Correction notice dated March 30, 2007, to all VAD customers alerting them to the hazard. The letter instructs them that a Thoratec representative will be contacting them to make arrangements to apply a warning label to the affected devices. The letter also informs them to dispose of any expired original design caged tip atrial cannula and to 1) not use the black collet and nut with any current design VAD, 2) make sure that all current patients using PVADs are using the appropriate collet nut, and 3) provide a copy of the product safety bulletin to all outlying hospitals and/or physicians that may be following patients implanted with PVADs and provide contact information with the acknowledgement form. |  | 2484 | 2,484 units | Class I Recall - Worldwide Distribution --- USA, Canada, Australia, Mexico, Singapore, Taiwan, Thailand, Austria, France, Germany, Israel, and Switzerland. | Terminated | April 24, 2008 | 391 | 2484 | Computer
Z-1244-2007 | 38069 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Endoscopic Instrument Control System (da Vinci Surgical System) Model IS1200, A4.3 SW level, Manufactured by Intuitive Surgical, Inc, Sunnyvale, CA 94086 |  |  | May 03, 2007 | 2007 | Intuitive Surgical, Inc. | Under certain circumstances, the product's software may crash and require a manual override or restart before functioning again. In addition, there is an unrelated addendum to the User Manual to clarify an icon not described in the previous User Manual. | Software design | As of 9/7/2007, Intuitive Field Service Engineers have performed the upgrade via service visit on 292 of 293 U.S. systems listed on the affected sites list. A notification letter that the upgrade has been performed was left with each site after the upgrade has been completed. The firm's distributors received advance notice of the upgrades 9/13/2007. |  | 405 | 405 | Worldwide, including USA, Canada, Australia, Austria, Belgium, China, Czech Republic, Denmark, France, Germany, Greece, India, Italy, Japan, Korea, Malaysia, Netherlands, Norway, Romania, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Taiwan, Turkey, United Kingdom, and Venezuela. | Terminated | July 02, 2008 | 426 | 405 | Computer
Z-0015-2008 | 39389 | No | JIF | Ammonia test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Ammonia Liquid Stable Reagent, Enzymatic Method. Catalog number TR60101. in vitro diagnostic. |  |  | July 24, 2007 | 2007 | Thermo Fisher Scientific | Method is susceptible to positive interference from endogenous ALT in patient samples. | Device Design | The firm has updated the Limitations section of the Ammonia Reagent Package Insert and circulated a Technical Bulletin dated July 24, 2007 to all customers that have purchased Ammonia Reagent that is still within expiration date. A Notification of Change has been introduced and shall be placed within subsequent Ammonia Reagent Kits with a copy of the updated package insert; the notification of change informs customers that a change has been made to the package insert. Notifications of Change are placed within kits for 6 months from the date of issue. |  | 6882 | 6,882 units | Worldwide, including USA, Ecuador, and Canada. | Terminated | December 18, 2007 | 147 | 6882 | Computer
Z-0005-2008 | 38449 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | PathFinder Bone Drill-Awl-Tap, a part of the PathFinder spinal stabilization system, Part Number 1156-60. |  |  | July 18, 2007 | 2007 | Abbott Spine | Tip of PathFinder surgical instrument broke during procedure. | Device Design | Firm notified sales representatives with an approved phone script on 7/18/2007. |  | 11 | 11 units | Nationwide. | Terminated | April 02, 2008 | 259 | 11 | Computer
Z-0052-2008 | 44793 | Yes | LWP |  |  | Cardiovascular | PMA | Guidant CONTAK RENEWAL 3 (Models H170, H173, H175), Guidant CONTAK RENEWAL 3 HE (Models H177, H179). Guidant CONTAK RENEWAL 4 (Models H190,H195) *, Guidant CONTAK RENEWAL 4 HE (Models H197, H199)* Cardiac Resynchronization Therapy Defibrillator (CRT-D). Sterile EO. Guidant Corporation, 4100 Hamline Avenue North, St. Paul, MN 55112-5798 USA. (* models are not available in the USA). |  |  | March 16, 2007 | 2007 | Boston Scientific CRM Corp | Atypical Charge Time Behavior: 1)-End of Life (EOL) or ERI is displayed during mid-life (typically 24-48 months), even though battery capacity remains available. 2)- Extended ERI charge time limits: Charge times during mid-life may remain below a normal extension of the ERI charge time . ****** Prophylactic replacement of these devices prior to ERI is not recommended. As stated in the Product Up | Device Design | A Product Update titled "ERI Charge Time Limit Extended During Mid-Life and Mid-Life Display of Replacement Indicators", dated 03/10/2007 was posted to Company's website on 03/13/2007. A CRM Product News letter was sent to physicians beginning 03/16/2007. The CRM Product News Letter gave an overview of the content of the Product Update and stated the web location for finding the complete Product Update. The Product Update identified affected devices, gave an overview of normal charge time behavior at mid-life for specific ICD and CRT-D devices. The letter also described the atypical charge time behavior, projected rate of mid-life display of replacement indicators and Patient Management Considerations. ****** Prophylactic replacement of these devices prior to ERI is not recommended. As stated in the Product Update, "Devices that have triggered charge time-based ERI or EOL during mid-life have several months, and in most cases more than one year of remaining battery voltage and capacity, which allows the devices in this pattern to continue to provide brady and LV pacing and maximum energy shocks. However, if ERI or EOL is triggered, device replacement should be scheduled."*****Recalled units only cover devices implanted prior to July 2005**** |  | 29375 | 22,761 within USA, 4,609 Foreign-Recalled units only cover devices implanted prior to July 2005 | Worldwide | Terminated | November 23, 2007 | 252 | 29375 | Computer
Z-0110-2008 | 38264 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance CT Systems, 40 slice configuration. Computed tomography scanner. |  |  | May 25, 2007 | 2007 | Philips Medical Systems (Cleveland) Inc | Two potential artifacts may appear in some images from head scans on the Brilliance 40/64 systems. These artifacts may lead to misdiagnosis. | Software design | Firm notified its customers via a Product Safety Notification letter dated May 25, 2007, warning the user of a possible appearance of artifacts which could lead to a misdiagnosis and improper treatment. |  | 404 | 404 | Worldwide, including USA, Canada, Argentina, Austria, Australia, Bangladesh, Belgium, Brazil, Chile, China, Colombia, Denmark, Ecuador, Egypt, England, Equatorial Guinea, Finland, France, Germany, Greece, Indonesia, Ireland, Israel, Italy, Jamaica, Japan, Korea, Lebanon, Libya, Malaysia, Mexico, Morocco, Netherlands, New Zealand, Norway, Pakistan, Panama, Phillipines, Poland, Portugal, Puerto Rico, Russia, Saudi Arabia, Serbia, Singapore, South Africa, Spain, Switzerland, Sweden, Taiwan, Thailand, Tunisia, Turkey, Ukraine, United Arab Emirates, Uzbekistan, Venezuela, and West Indies. | Terminated | September 24, 2012 | 1949 | 1037 | Computer
Z-0262-2008 | 45464 | No | DXG | Single-function, preprogrammed diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Edwards Vigilance II Continuous Cardiac Output/Oximetry/Volumetric (CCO/SvO2/CEDV) Monitor Model: VIG2/VIG2E Software Versions: 00.51,00.60, 00.61, 00.62, 00.63, 00.64, 00.67, 00.69, Edwards Lifesciences LLC., Irvine, CA 92614 |  |  | May 17, 2007 | 2007 | Edwards Lifesciences, Llc | MAP offset errors: If users are "slaving" mean arterial pressure (MAP) from the bedside monitor into the Vigilance II Monitor for the purpose of the continuous calculation of systemic vascular resistance (SVR), at times, on some bedside monitors, the Vigilance II monitor may display a MAP offset of as much as 25mmHg. Since the MAP is used in calculations of various patient parameters such as SVR, | Software design | An Urgent Customer notification letter dated May 17, 2007 was mailed to all Vigilance II customers advising that If they are "slaving" mean arterial pressure (MAP) from their bedside monitor into the Vigilance II Monitor for the purpose of the continuous calculation of systemic vascular resistance (SVR), to be aware that at times, on some bedside monitors, the Vigilance II monitor may display a MAP offset of as much as 25 mmHg. Since the MAP is used in calculations of various patient parameters such as SVR, these may also not be accurate. A table is provided showing those bedside monitors with the problem: GE Solar, Phillips (HP) M1006A, M1006B, as well as, those bedside monitors verified to not have the problem: SpaceLabs Model 90402 and Nihon Kohden Model BSM 8000. The letter also serves to remind the customer of the Warning in the Vigilance II Directions for Use which is also provided in the letter that instructs user to verify slaved MAP values with those displayed on bedside monitor. A temporary work around for this offset problem is to manually enter the MAP based on the readings from the bedside monitor. All Vigilance II calculations from the MAP have been re-verified to be correct when the MAP is accurate. The problem is in the transfer of the MAP between the bedside monitor and the Vigilance II. Customers are also advised that if they have the GE Solar or the Phillips (HP) models listed in the letter or any monitor not listed as verified to be compatible with the Vigilance II (Space labs and Nihon Kohden models listed above) Edwards recommends verification of MAP accuracy prior to treating based on parameters from a Vigilance II in a "slave" configuration. If there is a difference, they should enter the MAP into the Vigilance II manually for a static calculation of SVR. Edwards will be supplying a software update to correct this problem. |  | 2680 | 2,680 units in North America distribution | Worldwide: USA and Canada | Terminated | October 09, 2009 | 876 | 2680 | Computer
Z-0257-2008 | 45429 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Abbott ACCELERATOR APS System Input/Output Module (IOM); List Number: 7L01-01; Abbott Laboratories, Abbott Park, IL 60064 |  |  | October 10, 2007 | 2007 | Abbott Laboratories, Inc | Sample problems; Five issues identified related to software utilized by the ACCELERATOR APS System Input/Output Module, to include: Under three specific conditions, the post-aspiration radio-frequency identification (RF ID) read of the tube carrier is not performed and the APS work cell may not appropriately generate Sample Presentation/Sample Queue errors with APS software version 1.1.1 or earlie | Software design | Firm notified consignee via letter on 11/10/07. Consignee notified of software issues and informed they would be contacted by an Abbott representative to schedule installation of software upgrade. |  | 1 | 1 unit. | Nationwide: including state of UT. | Terminated | March 10, 2008 | 152 | 1 | Computer
Z-0212-2008 | 45486 | No | DSQ |  |  | Cardiovascular | PMA | Thoratec Implantable Ventricular Assist Device (IVAD pump) |  |  | October 19, 2007 | 2007 | Thoratec Corp | Pneumatic Leaks: air leaks developing in the pneumatic driveline when subjected to excessive stress at the junction of the percutaneous line to the pump housing due to sharp bending when implanted (paracorporeal-external-position). | Device Design | An initial communication was sent to all affected IVAD customers on October 19, 2007, in the form of an "Urgent: Medical Device Correction" notice informing them of the problem. The correction notice requests that its customers/physicians take the following actions:1) Contact all ongoing patients with IVADs placed in the paracorporeal position toevaluate the percutaneous line adjacent to the pump housing to ensure that it is not subject to extreme flexing or bending at acute angles. 2) Remind implanting surgeons and support staff of proper percutaneous line management and provide the patient with a copy of the recall information provided. Complete and sign an Acknowledgement form and fax it to Thoratec Regulatory Affairs. This is a User Level Recall. The firm also issued a November 7, 2007, Press Statement- Urgent Thoratec Corporation Worldwide Recall of Implantable Ventricular Assist Device. |  | 766 | 766 | Worldwide: USA including states of AL,AR,AZ,CA,CO, FL,GA, IA,IL, IN,KS,KY,LA, MA, MD, ME,MI, MN, MO,NC, NJ, NY, OH, OK, OR, PA, TN, TX, UT, VA, WA, and WI and to the following countries: Austria, Canada, France, Netherlands, Singapore, Switzerland, and The United Kingdom | Terminated | May 26, 2009 | 585 | 766 | Computer
Z-0180-2008 | 38276 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Philips, Voxel Q workstations, untilizing AcQSim or AcQPlan software version 4.2, 4.2.1, 4.9, 4.9.1, 4.9.2, 4.9.3, 5.0, 5.0.1, 5.0.2 or 5.0.3 and any Mx8000, Mx8001DT, or Brilliance 6, 10, 16, 16P, and Big Bore CT systems connected to the VoxelQ, Philips Medical Systems, Cleveland, OH 44143 |  |  | March 01, 2007 | 2007 | Philips Medical Systems (Cleveland) Inc | Incorrect simulation and/or beam placement: An anomaly has been identified with the Philips VoxelQ workstation utlizing AcQPlan software version 4.2, 4.2.1, 4.9, 4.9.1, 4.9.2, 4.9.3, 5.0, 5.0.1, 5.0.2 or 5.0.3 when computing absolute patient marking coordinates with data containing reconstruction offsets from the Mx800, the Mx8001DT, and the Brilliance 6, 10, 16, 16P, and Big Bore CT systems. Whe | Software design | On November 1, 2007, after telcon with FDA, the firm revised its recall strategy to notify consignees via a November 9, 2007, Urgent Device Notification letter which indicates that warning labels will be affixed to the units that alerts the user of the reconstruction offset defect by Philips employees. |  | 520 | 520 Devices | Worldwide Distribution- USA including states of AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, WI, WV and WY and Canada | Terminated | March 01, 2012 | 1827 | 520 | Computer
Z-0078-2008 | 44915 | No | LKN |  |  | Gastroenterology/Urology | 510(k) | Spectra Optia apheresis system, a blood component separator. Part No. 61000, Gambro BCT Lakewood, CO USA 80215. |  |  | September 18, 2007 | 2007 | Gambro BCT, Inc. | Under specific conditions, apheresis machine could return up to 60mL of air to the patient. | Software design | All consignees were notified by telephone beginning 09/17/2007. They were given instructions on how to safely use the machine. Letters were also sent on 09/22/2007 restating the information. |  | 15 | 15 units | US, Canada, Australia, Europe. | Terminated | January 25, 2008 | 129 | 15 | Computer
Z-0254-2008 | 44811 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | NEMIO XG Ultrasound System, Model SSA-580A , Software versions: VI .5.0, Vl .5.000A, VI .5.000B, V2.0.0,Toshiba America Medical Systems, Inc., Tustin, CA. |  |  | August 29, 2007 | 2007 | Toshiba American Med Sys Inc | Multiple problems including : (1) The surface temperature of the transducer tip can rise and exceed the limit set by the international standard. (2) Incorrect display of the TI indices on the display monitor. The displayed value is greater or smaller than the true TI value, and (3) Manuals contain incorrect values for MI/TI values. | Software design | Urgent Medical Device Notification letters were sent to customers on August 29, 2007. Customers were notified of the several issues related to the Nemio SSA-550A and Nemio SSA-580A recall and asked to share this information with all users of the device. Customers were informed that Toshiba is preparing software and corrected manuals for distribution to address these issues. If customers have any questions they were instructed to contact their Toshiba Customer Service Engineer or call (800) 421-1968 and ask for the Regulatory Affairs Department. |  | 128 | 128 units | Nationwide | Terminated | April 05, 2012 | 1681 | 1406 | Computer
Z-0001-2008 | 38473 | No | GWQ | Electroencephalograph. | Neurology | Neurology | 510(k) | Bio-logic System Corp --Telescoping Camera Pole Assembly (part number 585-CYBPSM and 580-VIDPOL) and assemblies (580IPCART and 580 PNCART; Bio-Logic Systems Corporation, Mundelein, IL 60060; The Model 580-VIDPOL is used with the 580-PNCART that is wired for use with a Panasonic Camera. The pole is also a subassembly of the 585-CYBPSM with wires used in the 580-IPCART wired for the use with an Ipela Camera. Both carts are a component of the Ceegraph/Sleepscan Netlink with Video. |  |  | July 26, 2007 | 2007 | Bio-Logic Systems Corp | Bio-logic System Corp Camera Pole may weaken and detach at the base of the telescoping camera pole during transport causing pole to break at the welded base and fall. | Device Design | Bio-logic sent a Safety Alert letters dated 7/26/07 to the affected customers, warning them that using the pole to transport the cart may cause the weld at the base of the pole to weaken, causing it to break. As a precaution to ensure customer safety, the accounts were instructed to place the enclosed red sticker on the pole, below the tightening knob. An Instruction sheet is enclosed . The warning label reads: "Caution: Do not use as a handle to move cart. Always lower camera when transporting. Avoid sudden stops or collisions." The accounts were requested to complete and return an enclosed reply sheet acknowledging the receipt of the letter and the placement of the warning labels on the poles. Any questions were directed to Customer Support at 1-800-332-3114. |  | 311 | 311 units | Worldwide Distribution -- USA and China, Egypt, Hong Kong, Israel, Malaysia and Saudi Arabia | Terminated | August 12, 2009 | 748 | 311 | Computer
Z-0081-2008 | 39422 | No | DPS | Electrocardiograph. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Paceart System, 2006 Edition, Medtronic, Inc. 710 Medtronic Parkway, Minneapolis, MN 55432-5604 USA. The Paceart System can act as a Practice Management software application and is intended for use by facilities that provide services such as event monitoring, trans-telephonic pacemaker evaluation, and in-office evaluation of implanted cardiac devices. |  |  | August 27, 2007 | 2007 | Medtronic Inc. Cardiac Rhythm Managment | An issue with the Paceart System - 2006 First Edition (Get Connected Edition) exists. Under certain circumstances, Paceart System Generic ICD reports may not accurately reflect patient VT/VF, SVT/NST and Mode Switch/AT/AF episode detection data. When episode detection data is imported into the Paceart System from certain sources, the Paceart System Generic ICD reports display a zero, suggesting | Software design | A Medtronic letter was sent to consignees on 8/27/07. The letter described the issue and the product involved. It also notified customers that Medtronic is developing a software update that will correct the issue. An Acknowledgement form was attached to be filled by consignees as confirmation and understanding of the information received. |  | 721 | 721 | Nationwide. | Terminated | November 03, 2008 | 434 | 721 | Computer
Z-0182-2008 | 44944 | No | FLG | Ultrasonic cleaner for medical instruments. | General Hospital | General Hospital | 510(K) Exempt | Clarisonic Skin Care System - A hand held cleansing device for the skin. The Unit comes with a charger cradle for charging the device. Pacific Bioscience Laboratories, Inc., 13222 SE 30th St., Bellevue, WA 98005 |  |  | July 03, 2007 | 2007 | Pacific Bioscience Lab, Inc. | Overheating: The handle or charger/cradle can overheat, melting the plastic, which potentially can lead to smoke and fire and has caused mild user burns. | Device Design | On August 30, 2007, the firm issued a press statement via PR Newswire and The Associated Press. They also posted a Recall Notice on their website (www.Clarisonic.com) requesting consumers to return their affected products for an exchange. All retailers and professional resellers were notified by telephone, E mail and/or letter and asked to return unsold inventory, as well as, post the provided point of sales notices and send provided post cards to known customers requesting exchange of recalled product. All users can contact the firm via the website or by phone at 1-866-378-7360 (6 am-6pm, M-F, Pacific Time) to arrange for the device exchange. |  | 97500 | 97,500 | Worldwide Distribution - including USA and countries of Japan and Canada. Units were distributed to retailers and professional skin care clinics throughout the USA. In addition, units were distributed to one distributor in Japan and to 4 clinics in Canada. | Terminated | December 13, 2010 | 1259 | 97500 | Computer
Z-0113-2008 | 39400 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | GE Healthcare Voluson E8 ultrasound scanner with Software version 6.2.0 and 6.2.1 and DICOM stations with these software versions; GE Healthcare 9900 West Innovation Drive, Wauwatosa, WI 53226, USA. |  |  | July 06, 2007 | 2007 | Ge Healthcare | When a spectrum of a Pulsed Wave Doppler shows an angle correction unequal to 0 is reloaded from the archive or the image is sent to a DICOM station, this angle correction is mistakenly disregarded resulting in an underestimation of the flow velocities. | Software design | GE Healthcare sent customer letters on July 15, 2007 to affected customers. The letter included a short term mitigation recommendation and a long-term correction. The long-term correction will be software upgrade provided free of charge to all affected customers. |  | 521 | 521 Units | Nationwide. | Terminated | October 21, 2008 | 473 | 521 | Computer
Z-0059-2008 | 39392 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Philips Medical System - BV Pulsera Mobile C-arms with Remote Control Option, (Release 2) |  |  | August 23, 2007 | 2007 | Philips Medical Systems North America Co. Phillips | System Shut Down -- Unexpected Shut Down during use of the Optional Remote Control. | Software design | On 8/23/07, the firm sent out a Safety Notification letter, via certified mail. On 8/29/07, the firm sent out an URGENT - DEVICE CORRECTION letter via regular mail, dated 8/20/07. Both letters advised customers that if the device shuts down during use, it can be restarted within about 1 minute by switching the system off and then on again, using the system on/off buttons. The letter(s) states that software is being developed to correct the problem. |  | 596 | 596 units | Nationwide | Terminated | September 27, 2010 | 1131 | 633 | Computer
Z-0119-2008 | 44880 | No | JGS | Sodium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | OPTIÂż LION E-Plus Cassettes, Model BP7507, (25 per box), In Vitro Diagnostic, OPTI Medical Systems, Inc., Roswell, Georgia 30076 |  |  | July 18, 2007 | 2007 | Opti Medical Systems Inc | Incorrect results/stability failure; Actual device stability studies show sodium and chloride sensors age differently than anticipated when stored at the high end of labeled storage, therefore subsequent measurements give results outside the allowed error allowance before the 6 month expiration date shown on the device. | Device Design | An Urgent Medical Device Recall Letter dated 16 July 2007, was issued via E mail and Fax on/or about 7/18/2007 directing consignees to remove unused cassettes from inventory and destroy them. Replacement cassettes will be shipped upon receipt of information regarding destruction by calling customer service at 800-490-6784. Sub-recall by distributors was requested. |  | 7225 | 289 boxes of 25 cassettes | Worldwide Distribution-USA, Australia, Chile, China, Egypt, France, Great Britain, Guatemala, Indonesia, Italy, Korea, Lebanon, Mauritania, Morocco, Nepal, Russia, South Africa, Sweden, Tanzania, Thailand, and Venezuela. | Terminated | December 13, 2007 | 148 | 7225 | Computer
Z-0159-2008 | 38516 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MEVATRON M-2 6300,Medical charge-particle radiation therapy system (Medical Digital Linear Accelerator),MODEL NUMBER: 9401316,MANUFACTURED BY: Siemens Medical Solutions USA, Inc, 4040 Nelson Ave, Concord, CA 94520 |  |  | March 22, 2007 | 2007 | Siemens Medical Solutions USA, Inc | Mistreatment of radiation therapy due to inadvertent parameter changes or miss-calibration in Service Mode. The miss-calibration can affect the beam profile and Virtual Wedge angle. This leads to patient mistreatment and alteration of future treatment plans. A second reason is an unplanned Dosimetry calibration data set can be erroneously copied during "Service Mode." | Software design | Distribution of Update Instructions TH007/07/S was sent to all affected sites on March 26, 2007. This was intended to inform users that Service Mode is only for use by authorized personnel, and that unauthorized and/or incorrect softpot parameter changes performed in Service Mode can result in the risk of patient mistreatment.Current control system software versions do not allow unplanned Dosimetry calibration data sets to be erroneously copied during "Service Mode." |  | 6 | 6 units | Worldwide:USA including states of Alabama, Alaska, Arizona, Arkansas, California, Colorado, Delaware, District of Columbia, Florida, Georgia, Idaho, Illinois, Indiana, Kansas, Kentucky, Louisiana, Maine, Maryland, Massachusetts, Michigan, Minnesota, Mississippi, Missouri, Montana, Nebraska, Nevada, New Hampshire, New Jersey, New York, North Carolina, North Dakota, Ohio, Oklahoma, Oregon, Pennsylvania, Puerto Rico, Rhode Island, South Carolina, South Dakota, Tennessee, Texas, Vermont, Virginia, Washington, West Virgina, Wisconsin, and Wyoming and FOREIGN COUNTRIES: Canada, Mexico, Serbia Montenegro, Angola, Argentina, Australia, Austria, Belarus, Belgium, Bosnia, Herzeg-Bosnia, Brazil, Bulgaria, Chile, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, DPR of Korea, Ecuador, Egypt, El Salvador, Estonia, France, Germany, Greece, Hungary, India, Indonesia, Iran, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kuwait, Lebanon, Lithuania, Malaysia, Morocco, Nepal, Netherlands, New Zealand, Norway, P.R. China, Pakistan, Palestine, Peru, Philippines, Poland, Portugal, Qatar, Rep. of Yemen, Korea, Romania, Russian Fed., Saudi Arabia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, Ukraine, United Kingdom, Uzbekistan, Venezuela, Vietnam, and Zambia | Terminated | December 18, 2010 | 1367 | 2120 | Computer
Z-0144-2008 | 44957 | No | HIG | Hysteroscopic insufflator. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Storz Equipment, Fluid Management Stand, for Fluid Management Systems, Product Number 26601EQ-KT |  |  | August 17, 2007 | 2007 | Karl Storz Endoscopy America Inc | Outflow Difference: The accuracy of the Equimat displayed value could be influenced by external factors of the cart/stand and may result in a fluctuating or inaccurate reading. | Device Design | All affected customers were notified via a Customer Letter issued on September 10, 2007 via FedEx overnight mail. The letter describes the problem with the device and the health hazard risk, and provides instructions for users to return the old stand and make arrangements for replacement stands. |  | 287 | 287 | Nationwide | Terminated | February 29, 2012 | 1657 | 287 | Computer
Z-0213-2008 | 38376 | No | KPQ | Radiation therapy simulation system. | Radiology | Radiology | 510(k) | Ximatron Digital Imaging (Ximavision), V 7.5.51 with Service Pack (SP), Oncology Information System, Varian Medical Systems |  |  | June 25, 2007 | 2007 | Varian Medical Systems Inc | Unexpected loss of Treatement Field Add-Ons (Wedges, Blocks, etc) during Ximatron simulation. If an unapproved treatment field is created without an MLC, but with an add-on, the add on will be deleted without warning. If this is not detected and corrected, incorrect dose delivery will occur. | Software design | The firm issued an Urgent Medical Device Correction letter to its customers by mail on June 25, 2007. It informed consignees of the anomaly in the software and workarounds for the issue. In addition, the firm also stated that they are currently validating a new software modification to repair the problem. |  | 14 | 14 units | Worldwide Distribution-USA (MO and NY medical centers) and countries of Switzerland, Bahrain, Canada, France, Germany, Japan, Norway and Sweden.--- | Terminated | February 04, 2008 | 224 | 14 | Computer
Z-0075-2008 | 44813 | No | KXJ | Radiologic table. | Radiology | Radiology | 510(K) Exempt | Sunisa Flexi-DT Table, Model Numbers: 51335678, 2269647, 2242088-2, 5141609-2, 5115763, and 5128145, |  |  | August 01, 2007 | 2007 | GE Healthcare | Design Problems Including: 1) Hot Brake (Longitudinal Overheating) potential risk of burns 2) Brake Locks and Sudden Releases 3) Emergency Stop Circuit Button; open instead of closed 4) Longitudinal Brakes Release does not engage properly 5) Tilting Movement causing tabletop to slide; may cause risk of pinched fingers to patient and/or operator 6) Lodging Pins on floor Uncovered; may cause potenti | Device Design | Recalling firm initiated correction by visit on August 2007. Corrections include replacement of one or more of the following: charger board, pin lodging, rear bumper. Also, replace the interface board and remove resistor mounted on chassis, remove brake pedal and assembly of cosmetic brake or apply caution labels. |  | 96 | 96 | Class II Recall - Worldwide Distribution --- USA including States of -- AL, CA, FL, GA, IA, IL, IN, LA, ND, NY, OH, PA, TX, and WI, and Countries of -- BELGIUM, CZECH REPUBLIC, FINLAND, FRANCE, GERMANY, GUAYANA, HONG KONG, HUNGARY, IRELAND, ISRAEL, ITALY, KUWAIT, LATVIA, LEBANON, SAUDI ARABIA, SPAIN, SWEDEN, SWITZERLAND, and UNITED KINGDOM. | Terminated | March 02, 2008 | 214 | 96 | Computer
Z-0067-2008 | 45403 | Yes | LWS |  |  | Cardiovascular | PMA | Medtronic Sprint Fidelis 6930 Steroid eluting, tripolar, ventricular lead with tined tip and right ventricular (RV) defibrillation coil electrode. Sterile. This lead is designed for pacing, sensing, cardioversion and defibrillation therapies. Medtronic World Headquarters, Medtronic, Inc., 710 Medtronic Parkway, Minneapolis, MN 55432-5604 USA. |  |  | October 15, 2007 | 2007 | Medtronic Inc. Cardiac Rhythm Managment | Medtronic Cardiac Rhythm Disease Management has voluntarily suspended worldwide distribution of the Sprint Fidelis family of defibrillation leads because of the potential for lead fractures. In addition, the company recommends against new implants of the leads (Sprint Fidelis Models: 6930, 6931, 6948, 6949). Medtronic, its Independent Physician Quality Panel, and Bruce Lindsay, M.D., Professor of | Device Design | A press release was issued by Medtronic on 10/15/2007. An FDA Statement was also issued on 10/15/2007. An Urgent Medical Device Information letter, dated 10/15/2007, was sent to physicians. The letter states that Medtronic has suspended distribution of the leads, the leads should no longer be implanted and unused leads should be returned to Medtronic. This letter describes background information, performance and recommendations. Medtronic believes it is inappropriate to prophylactically remove Sprint Fidelis leads except in unusual individual patient circumstances. An Important Patient Information letter was addressed to Medtronic Heart Device Patients. Both letters can be found on Medtronic's website at www.medtronic.com/fidelis. Return of product will be achieved with the assistance of Medtronic Representatives. |  | 235000 | 235,000 for all models | Worldwide, including USA, Puerto Rico, Guam, and Northern Mariana Islands. | Terminated | June 21, 2011 | 1345 | 235000 | Computer
Z-1147-2007 | 38358 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ISOLOC 4.2 Software Part Number NT-NW-425-520. Image guided radiation treatment software. |  |  | May 31, 2005 | 2005 | Northwest Medical Physics Equipment Inc | When the ISOLOC (Version 4.2) profile distance unit is set to "centimeter" and the user clicked the "back" button in the digitizer window, the marker and target coordinates in the localization window erroneously display values ten times greater than the correct values. | Software design | Software to fix the described problems was written and customers contacted by e-mail on 05/31/2005, that a software patch, titled 05312005 was available on the firm's internet website. |  | 85 | 85 | Worldwide, including USA, Korea, Spain, and UK. | Terminated | April 19, 2011 | 2149 | 85 | Computer
Z-0330-2008 | 44959 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Allura Xper FD20/10, Release 1.2, Philips Medical Systems, Bothell, WA 98021 |  |  | September 24, 2007 | 2007 | Philips Medical Systems North America Co. Phillips | Inaccurate anatomical measurements when using the optional SmartMask feature | Device Design | The recalling firm sent customers an Urgent Device Correction notification letter dated 09/21/07 via certified mail - return receipt requested. This notification provided a short term work around for the device defect. For a permanent fix, the recalling firm is developing software update to check for mismatch at the start of a subtracted fluoroscopy run, and will contact customers to schedule an install. Installation is expected to begin 10/2007 and be completed 06/2008. |  | 53 | 53 units | Nationwide | Terminated | May 03, 2012 | 1683 | 521 | Computer
Z-0294-2008 | 45492 | [6 empty] | Medtronic 0-arm Imaging System, Product Catalog Number: B1-700-00027, Medtronic Navigation, Inc., Littleton, MA 01460 |  |  | October 17, 2007 | 2007 | Medtronic Navigation, Inc | Reversed Instrument Position : a defect in the software version 3.0 causes an error in the position of surgical instruments and relative motion are reversed 180' on the StealthStation image. (If the StealthStation is not used, or if the patient orientation is prone or supine, the problem does not occur) | Software change control | Medtronic Navigation notified customers by letter on October 17, 2007. Medtronic Field service will perform upgrade (software version 3.0.1) at user site. |  | 7 | 7 units | Worldwide: USA including states of CO, IN, NE, and MN and countries of Canada and Germany | Terminated | November 27, 2007 | 41 | 7 | Computer
Z-0421-2007 | 36967 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Desktop ProÂż R6.0 &amp; R6.1, Linear Medical Accelerator. |  |  | November 23, 2006 | 2006 | Elekta, Inc. | Unexpected Diaphragm movement when manual field size modifications are not saved. | Software design | Consignees were notified by letter on/about 11/23/2006 |  | 37 | 37 units | Worldwide distribution ---- including USA states of AZ, AL, CA, CO, FL, GA, IA, MI, NV, OH, OK, PA, WA, WV and Country of Canada . | Terminated | May 14, 2008 | 538 | 37 | Computer
Z-0121-2008 | 44950 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | The Brilliance CT Big Bore is a whole body computed tomography X-Ray system. The malfunctioning software application is the Tumor Localization (Tumor LOC) application. |  |  | September 27, 2007 | 2007 | Philips Medical Systems (Cleveland) Inc | The maximum intensity projection (MIP) images produced from respiratory gated images by the Tumor Localization (Tumor LOC) software application on Brilliance Big Bore CT scanners will be labeled incorrectly for certain patient orientations and view conventions when saved for later viewing. | Software design | Notification letters were sent to the installed customer base informing them of the hazard on 10/1/2007. A mandatory corrective software release due out mid October 2007, will be available at no cost to the customer. The software will be installed by Philips Field Service Engineers. |  | 206 | 206 systems | Worldwide, including USA, Canada, France, Finland, Korea, Australia, Italy, Switzerland, China, Austria, Thailand, Malaysia, Denmark, Norway, Jordan, Netherlands, England, South Africa, Spain, Germany, Bahrain, Belgium, New Zealand, Puerto Rico, South Korea, Saudi Arabia, Czech Republic, Sweden, Israel, and Lithuania | Terminated | March 02, 2012 | 1618 | 206 | Computer
Z-0009-2008 | 38951 | No | HRY | Knee joint femorotibial metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | EPIKÂż KNEE SYSTEM Angled Insert Impactor LM/RL; Catalog #801-03-051 Rev. A; non-sterile; knee prosthesis instrument. Distributed by Encore Medical, L.P., Austin, TX 78758. |  |  | August 09, 2007 | 2007 | Encore Medical, Lp | Angled Insert Impactor may scratch Tibial Insert on the articulating surface, causing increased wear at the damaged site. | Device Design | Firm notified consignees regarding recall via letter on 08/09/07. Consignees informed that Encore customer service representatives would be contacting them to exchange recalled units with replacements. |  | 5 | 5 units. | Products distributed to direct accounts in CA, FL and MD. | Terminated | March 10, 2008 | 214 | 5 | Computer
Z-0278-2008 | 45420 | No | OAT | Endosseous dental implant. | Dental | Dental | 510(k) | Straumann Palatal Implant , 4.8 L 4.2 mm, |  |  | September 28, 2007 | 2007 | Straumann USA LLC | Sticking Transfer Piece: The Transfer Piece which accompanies the Palatal Implant may get stuck during implant requiring removal by using mechanical force. | Device Design | Straumann USA, LLC sent Urgent Device Recall letters, via FedFx on 9/28/07, requesting the return of all Straumann Palatal Implant, articles 042.335 with lot numbers 1001, FE3776, F4662, F6563, E7083, and article 042.336S lot number 1001. |  | 32 | 32 units | Nationwide Distribution - USA including states of AZ, CA, CO, ID, IL, IN, GA, KY, NH , MA, MD, MI, NM, PA, TX, UT, VA, and WA. | Terminated | April 16, 2010 | 931 | 90 | Computer
Z-0037-2008 | 38371 | No | EZW | Implanted electrical urinary continence device. | Gastroenterology/Urology | Gastroenterology/Urology | PMA | Medtronic InterStim II INS Model 3058 with N'Vision InterStim-B Software supplied on the Model 8870 Version MMB_01/NNB_01 N'Vision with a Clinic an Programmer Model 8840. Implantable Neurostimulator for Urinary Control. |  |  | June 25, 2007 | 2007 | Medtronic Neuromodulation | A software issue may cause a lower than expected battery capacity estimate when evaluating an InterStim II Implantable Neurostimulator (model 3058) with an N'Vision Clinician Programmer(Model 8840) before implant . | Software design | Begining on July 11, 2007 a letter to the physicians was sent informing them of the issue and patient mangement of the two issues. The letter includes a Cover Letter, A Reply Form, and two Recall letters explaining the issue in detail. |  | 1885 | 1885 | Nationwide and the Netherlands. | Terminated | October 08, 2010 | 1201 | 1885 | Computer
Z-0041-2008 | 44791 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | PEMView Software used with Naviscan PEM Flex Solo I and PEM Flex Solo II PET Scanner, for nuclear medicine imaging. |  |  | August 01, 2007 | 2007 | Naviscan PET Systems | Naviscan has identified a potential safety issue with all versions of PEM View Software commonly used for analyzing patient images acquired on the PEMFIex product line.1. The Positron Emission Mammography Standardized Uptake Value (PEMSUV) and distance measurements for the product mentioned above may be inaccurate if a zoomed image has been opened, measurement tools are used, and the applet rema | Software design | Telephone notification regarding issues #1 and #2 were carried out on August 1 and 2, 2007 and included instructions for a validated work-around to avoid the software issues. Written notifications with these same validated work-around instructions were sent on August 2, 2007 via certified mail. Customers were notified that Naviscan has identified a potential safety issue with all versions of PEM View Software.Issue #3 was identified after the notification to customers for issues #1 and #2. Telephone notification regarding issues #3 was carried out via telephone and certified mailing beginning August 16, 2007. Customers were informed that Naviscan is preparing software to resolve these issues and that they would be contacted by Naviscan Customer Support to ensure they thoroughly understand the issues and the recommended procedure to prevent these issues from occurring. If customers have any questions they were instructed to contact Naviscan Customer Support at 1-888-628-4159 or service@navsicanpet.com. |  | 19 | 19 units | Nationwide. | Terminated | May 14, 2008 | 287 | 19 | Computer
Z-0261-2008 | 45402 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | ARJO,Carendo Multi-Purpose Hygiene Chair; Model BIB2003-01, a non-AC powered patient lift; Manufactured by Arjo Hospital Equipment AB, Eslov, Sweden |  |  | October 15, 2007 | 2007 | Arjo, Inc. | Male patient genitals may be pinched between the seat and the frame of the patient lift chair when the Care Raiser seat is returned to its lowest position. | Device Design | Arjo sent Customer Field Correction Notification letters dated 10/15/07 to the end user accounts who received the affected chairs, advising them of the potential for the pinching of male patient genitals between moving parts of the care raiser function of the Carendo Hygiene Chair. They were instructed to review the enclosed revised Operating and Product Care Instructions manual, highlighting the warnings on pages 18 and 20; ensure all users receive refresher training in the chair's proper use and have access to the revised manual; and complete and return the enclosed customer response form, acknowledging the receipt and understanding of the letter. Any questions were directed to the Arjo Quality Department at 800-323-1245, ext. 6140 or 6165. |  | 1342 | 1,342 hygiene chairs | Nationwide | Terminated | July 27, 2009 | 651 | 1342 | Computer
Z-0040-2008 | 44740 | No | EZW | Implanted electrical urinary continence device. | Gastroenterology/Urology | Gastroenterology/Urology | PMA | Medtronic InterStim INS (Model 3023) with N'Vision Inter Stim-B software supplied on the Model 8870 Version MMB_01/NNB_01 with N'Vision Clinician Programmer Model 8840 or InterStim iCon Patient Programmer Model 3037. Implantable Neurostimulator for Urinary Control. |  |  | June 25, 2007 | 2007 | Medtronic Neuromodulation | Software issue may cause false low battery alert to be displayed when interrogating an Interstim Implantable Neurostimulator (Model 3023) with an N'Vision Clinician Programmer (Model 8840), or with an Inter Stim iCon Patient Programmer (Model 3037). | Software design | Begining on July 11, 2007 a letter to the physicians was sent informing them of the issue and patient management for the two issues. The letter includes a Cover Letter, A Reply Form, and two Recall letters explaining the issue in detail. |  | 1885 | 1885 | Nationwide and the Netherlands. | Terminated | October 08, 2010 | 1201 | 1885 | Computer
Z-0384-2008 | 38345 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Siemens, Mevatron M2 Primus, Part No.: 19 40 035, Mid-Energy Medical Charge particle radiation therapy system, with pre-R2.0 Optivue Flat Panel Positioners, Siemens Medical, Concord, CA 94520 |  |  | March 23, 2007 | 2007 | Siemens Medical Solutions USA, Inc | Product's collision detection system may not be sensitive enough to trigger a stop action without a significant amount of force. | Device Design | The firm has issued a Customer Safety Advisory Notice dated 12/6/2006, to all sites that have the pre-R2.0 OPTIVUE Flat Panel Positioners installed. Users will have a new microswitch-based collision detection system, installed free of charge. |  | 9 | 9 units. | Worldwide Distribution- USA including MA, ME, NJ, NY, PA, MD, WV, NC, SC, GA, FL, AL, TN, KY, OH, IN, MI, IA, WI, SD, ND, MT, IL, NE, LA, OK, TX, WY, NV, CA, AK AND AR, and countries of Serbia, Australia, Austria, Belgium, Canada, Denmark, El Salvador, France, Germany, Hungary, India, Ireland, Italy, Japan, Kuwait, Malaysia, Mexico, Netherlands, New Zealand, Norway, China, Philippines, Poland, Portugal, Qatar, Republic of Korea, Russia, Saudi Arabia, Singapore, South Africa, Spain, Switzerland, Turkey, Ukraine, UK. | Terminated | February 11, 2011 | 1421 | 516 | Computer
Z-0056-2008 | 39408 | No | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | Angiomat Illumena Injector with Auxiliary Console, Mallinckrodt, Inc., Hazelwood, MO. |  |  | August 24, 2007 | 2007 | Mallinckrodt Inc | Synchronization problems: units that have both a main programmable console and an auxiliary programmable console, when simultaneously, injection parameters are programmed into the unit's main console and the unit's auxiliary consoles, mis-synchronization could occur resulting in a variance between the intended injection parameters (those programmed) and the actual parameters of the injection deliv | Device Design | On 9/20/2007, a notification was sent from Mallinckrodt to all direct accounts via certified mail. The notification details the potential for mis-synchronization of injection parameters if the main and auxiliary consoles are simultaneously programmed. The notification reminds end-users that simultaneous programming should not be performed and to ensure it is not performed, a notification label (provided in the notification) is to be affixed on both consoles of the injector system. Once the label is affixed to the system, end-users are instructed to complete the notification survey and return the survey to the Mallinckrodt. In addition, the notification includes an updated chapter insert (Chapter 11.1) for the injector system's product manual that emphasize the separation between programming of the main and auxiliary consoles. Questions regarding this matter may be directed to Mallinckrodt's Product Monitoring Department at 1-888-744-1414, menu option 1 and then menu option 2. |  | 176 | 176 | Worldwide to include:U.S.A. states of - AL, AZ, CA, CO, CT, FL, GA, HI, IA, IL, KS, KY, LA, MA, MD, MI, MN, MO, NE, NJ, NY, OH, OK, PA, TN, TX, UT, VA, WA, &amp; WI. and Foreign countries of Argentina, Belgium, Brazil, Canada, Germany, Japan, Singapore, &amp; Zambia. | Terminated | August 17, 2010 | 1089 | 176 | Computer
Z-0269-2008 | 45154 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Siemens ECAT ACCEL, Catalog No. 3554479, PET/CT Scanner, Manufactured by Siemens Medical Solutions USA, Inc., 810 Innovation Drive, Knoxville, TN 37932 USA, TEL: 865-218-2534 |  |  | July 23, 2007 | 2007 | Siemens Medical Solutions USA Inc | Incorrect Decay Correction Results - ECAT System's software versions 7.2.2 and 7.4 calculate the decay correction for emission data incorrectly, while performing whole-body studies with all transmissions followed by all emissions. | Software design | Siemans sent an Urgent Device Correction letter, dated June 15, 2007, to all affected consignees informing the user of the problem and requesting that they discontinue the practice of running all transmissions followed by all emissions until a corrective software patch is installed. The corrective software patch is expected to be installed by October 2007. |  | 105 | 105 units | Worldwide Distribution - USA and Countries of Argentina, Austria, Belgium, Brazil, Canada, Switzerland, Chile, China, Czech Republic, Germany, Denmark, Spain, Finland, France, Great Britain, Italy, Japan, South Korea, Netherlands, Norway, Saudi Arabia, Sweden, Slovak Republic, Turkey, and Taiwan. | Terminated | September 11, 2009 | 781 | 446 | Computer
Z-0265-2008 | 45514 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Healthcare Advantage Workstations, Advantage Worksation version 4.2 and 4.3 with Volume Viewer 2 (version 6.4-6.8) provided with Advance Vessel Analysis (AVA) and Advantage Worksation 4.4 with Volume Viewer 2 (version 6.8 or before) or Volume Viewer 3 (version 7.2 -7.3) provided with Advanced Vessel Analysis (AVA), Model 2378698-2,GE Healthcare, Waukesha, WI. 53188 |  |  | September 10, 2007 | 2007 | GE Healthcare | Aortic Length over-estimated: incorrect measurements when using the Aorta analysis protocol or customized protocols. Until a new software version is installed the aorta analysis protocol or the customized protocols for an aneurysm should NOT BE USED. | Software design | An Urgent Safety Notice to customers dated 9/10/07, was issued by the recalling firm informing them of the device defect, safety issues, warnings on future uses and promised new software updates to be installed by GE representatives, free of charge, beginning Nov. 2007. |  | 2518 | 2,518 | Worldwide Distribution - USA including states of AK, AL, AR, AZ, CA, CO, CT, CU, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, WA, WI, WY, and Guam, and countries of Canada, Mexico. Argentina, Brazil, Chile, Colombia, Dominican Republic, Ecuador, Australia, China, Honk Kong, Indonesia, Japan, Korea, New Zealand, Singapore, Thailand, Taiwan, United Arab Emirates, Albania, Austria, Azerbaijan, Bosnia, Belgium, Belarus, Switzerland, Czech Republic, Germany, Denmark, Egypt, Spain, Finland, France, Great Britain, Guadeloupe, Greece, Hungary , Ireland, Israel, Italy, Kuwait, Lebanon, Morocco, Netherlands, Norway, French Polynesia, Poland, Portugal, Russian Federation, Saudi Arabia, Sweden, Slovenia, Tunisia, Turkey, and South Africa. | Terminated | December 17, 2011 | 1559 | 2518 | Computer
Z-0061-2008 | 44942 | Yes | MVK |  |  | Cardiovascular | PMA | Battery Pack portion of the LifeVest device. The LifeVest is a wearable cardioverter defibrillator worn by a patient at risk for sudden cardiac arrest (SCA). |  |  | September 14, 2007 | 2007 | Zoll Lifecor Corporation | Possible malfunction of the battery pack. A significant electrostatic discharge (ESD) can cause the battery pack to malfunction. The malfunction causes the battery pack to discharge more rapidly and prevents the pack from recharging. | Device Design | Patients currently using the LifeVest should continue to do so while the product removal is underway. Advisory notices will be sent via first class mail 9/14/2007 to all active patients using the LifeVest device. Product replacements began on 8/31/07 by shipping out revised battery packs for new patient use. Shipments of new battery packs will continue based on production capacity and material availability. Any patients expected to be using the device longer than 5 months will be provided with 2 replacement Battery Packs. They will be instructed to return their existing packs in the same shipping box. |  | 2006 | 2006 packs | The product was shipped nationwide to patients. The product was also shipped to distributors in Germany, United Kingdom, Italy, Spain, France, and Denmark. | Terminated | June 13, 2008 | 273 | 2006 | Computer
Z-0978-2007 | 37803 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | LINAC Scapel Treatment Planning System, incorporating FastPlan software versions 4.0, 5.0, 5.0.1, 5.1, 5.1.1. |  |  | March 09, 2007 | 2007 | Varian Medical Systems Inc | A software anomaly in a medical device used for patient treatment could result in accidental or unintended exposure radiation exposure. The tilt rotation compensation algorithm does not adequately compensate for the tilt rotation. | Software design | The recalling firm notified medical device end users by letter dated 03/09/07 and advised of the software anomaly and provided details for user corrective action pending software upgrade by the manufacturer. |  | 80 | 80 devices | The medical devices were distributed to medical centers nationwide and worldwide. | Terminated | February 28, 2012 | 1817 | 80 | Computer
Z-0352-2007 | 36713 | No | LEG | Neuroleptic drugs radioreceptor assay test system. | Toxicology | Toxicology | 510(k) | ADVIA 1650 Chemistry System, Automated Clinical Chemistry Analyzer - human serum, plasma and urine tests. The following are the measured parameters that are within the scope of this ''field correction'': Affected assay: Valproic Acid. Affecting assays: Digoxin; Phenytoin; Gentamicin; Theophylline. |  |  | September 08, 2006 | 2006 | Bayer Healthcare, LLC (Diagnostics Division) | Bayer received a customer complaint where the Valproic Acid Quality control material recovered lower than target when it was run in a panel of other Therapeutic Drug (TDM) assays. The under recovery observed by the customer was approximately -25%. Investigation into this issue was initiated promptly following receipt of this complaint. | Software design | On 9/08/06, a Support Bulletin was sent to all affected Bayer HealthCare LLC Branches in order to inform them of the issue and to provide them with instructions for the 'field correction'. On 9/08/06, a Customer Bulletin was sent to all affected Bayer HealthCare LLC Branches both in the US and worldwide for communication with the affected end user customers. The bulletin advises of the problem and the corrective action. A Completion Confirmation (Fax-back) Form was included to ensure that all customers worldwide have been notified and have complied with the mandatory action. The Customer Bulletin is to be sent to all affected end user customers. Bayer Technical Service Specialists will follow-up with all of the ADVIA Chemistry System customers to make sure they have understood the bulletin and to ask that they return the FAX-back Completion Confirmation Form to Bayer HealthCare. If Bayer HealthCare does not receive the form in a timely manner, specialists in the Bayer Technical Care Center will contact the customer through whatever means possible in an effort to obtain 100% effectiveness on this communication. |  | 1184 | 1184 instruments total, all products | Worldwide, including USA, Puerto Rico, Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, China, Columbia, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, Hong Kong, Hungary, Ireland, Israel, Italy, Korea, Malaysia, Mexico, The Netherlands, New Zealand, Norway, Poland, Portugal, Russia, Serbia, Singapore, Slovakia, South Africa, Spain, Saudi Arabia, Sweden, Switzerland, Syria, Taiwan, Turkey, and United Kingdom. | Terminated | May 23, 2012 | 2084 | 1184 | Computer
Z-0293-2008 | 45567 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Toshiba-Interventional Angiography System, Infinix i, Model DFP-8000D, equiped with Digital Radiography System with Software version 3.4er000, Toshiba America Medical Systems, Inc., Tustine, CA 92780 |  |  | October 01, 2007 | 2007 | Toshiba American Med Sys Inc | System Lock-up: When performing DAS One Shot, the system may lockup due to a software timing error. | Software design | Notification letters (Urgent: Medical Device Correction-dated October 1, 2007) with an attached Customer Response form were sent to customers beginning on October 2, 2007. Customers were notified this action was being taken as a result of an investigation that resulted from a software issue that was found done by the manufacturer. Customers were informed that a Toshiba Customer Service Engineer would contact them to schedule a software upgrade when the software becomes available. If customers have any questions they were instructed to contact their or call (800) 421-1968 and ask for the Regulatory Affairs Department. |  | 23 | 23 | Nationwide-including states of AL, AR, FL, GA, MI, NY, OH, PA, TX &amp; WV | Terminated | April 05, 2012 | 1648 | 23 | Computer
Z-0290-2008 | 45462 | No | HWR | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Odyssey Femoral Holder/Driver, Part Number K001-7410, Non-sterile, 1 each, Wright Medical Technology, Inc., 5677 Airline Road, Arlington, TN 38002 |  |  | October 04, 2007 | 2007 | Wright Medical Technology Inc | Breakage: breaking of the connecting foot that is on the distal surface of the instrument | Device Design | The firm notified their distributors of the recall on October 4, 2007, with an Urgent Medical Device Recall letter explaining the problem and requesting return of the product from their inventory and from their consignees. The hospitals were notified of the recall on October 4, 2007 using the same Urgent Device Recall Letter. |  | 66 | 66 units | Worldwide-USA and Australia | Terminated | July 22, 2009 | 657 | 66 | Computer
Z-1141-2007 | 38222 | No | LZH | Infusion pump. | General Hospital | General Hospital | 510(k) | Embrace Enteral Pump |  |  | June 21, 2007 | 2007 | Ross Products Division Abbott Laboratories | Continued use of the pump may cause the door pivot point on the back of the pump to crack or break. This will cause the pump to shut down and stop the initiation of enteral feeding. | Equipment maintenance | Consignees were notified by letter dated 6/25/2007. The firm sent a revised letter on 07/26/2007. |  | 9600 | 9,600 units | The product was distributed throughout the US. | Terminated | September 20, 2012 | 1918 | 9600 | Computer