## Sheet1
Recall Number | Recall Event ID | Device Type | Product Code | Regulation Description | Medical Specialty | Review Panel | Submission Type | Trade Name/Product | Main Name | Recall Class | Date Posted | Year Posted | Recalling Manufacturer | Reason for Recall | FDA Determined Cause | Action | Consumer Instructions | Clean Quantity | Quantity in Commerce | Distribution | Recall Status | Termination Date | Time to Terminate | Merged Quantities | Fault_Class
Z-0407-2009 | 49522 | No | GAZ | Vacuum-powered body fluid suction apparatus. | General Hospital | General Hospital | 510(K) Exempt | Mentor Aseptic Transfer Set, Cat. #350-8400 (contains BD 60mL Luer-Lok Syringe). Aseptic transfer set manufactured by Mentor Texas, LP, Irving, TX 75038 (BD 60 mL Luer-Lok syringe manufactured by Becton Dickinson, Franklin Lakes, NJ 07417). |  |  | August 22, 2008 | 2008 | Mentor Texas, Inc | The Mentor Aseptic Transfer Set contains a component, the BD 60mL Luer-Lok Syringe, which is under recall by Becton Dickinson due to a package integrity issue. | Package design/selection | Firm sent notification, along with a copy of the Becton Dickinson recall letter, to consignees on 08/18/08. Consignees asked to inspect inventory of Mentor Aseptic Transfer Sets for affected BD 60mL syringes and to complete and return the included Device Accountability Worksheet. For additional information, please contact Mentor Customer Service at 800-235-5731. |  | 32480 | 32,480 Aseptic Transfer Sets (4 syringes/set) | Worldwide Distribution; product sent to all states (except Alaska), Washington DC and Puerto Rico. Product also sent to the following countries: Australia, Canada, Costa Rica, Korea, New Zealand and South Africa. | Terminated | August 26, 2010 | 734 | 32480 | Not_Computer
Z-0569-2008 | 45951 | No | ELZ | Preformed crown. | Dental | Dental | 510(K) Exempt | 3M ESPE Unitek Stainless Steel Permanent Molar Crowns , Product Code: 900321, Made in U.S.A by 3M ESPE Dental Products St. Paul, MN 55144-1000. |  |  | October 17, 2007 | 2007 | 3M Company / Medical Division | Mislabeled: Unitek Stainless Steel Permanent Molar Crowns were placed in packaging which incorrectly identified the product as Unitek Stainless Steel Primary Molar Crowns. | Packaging process control | 3M ESPE Dental Products initially contacted Distributors and Doctors on 10/22/07 through telephone calls, and sent a Device Recall letter via Fed Ex - Overnight on 10/22/07.The letter informed the users that Unitek Stainless Steel Permanent Crowns Lot # P060928, Product Code 900321 may be mislabeled as Unitek Stainless Steel Primary Molar Crowns. Users were asked to stop using and/or distributing the product and contact ESPE Dental Products at 1-800-237-1650 to initiate product return and replacement. |  | 975 | 975 | Worldwide Distribution USA including states of TX, IN, FL, NV, KY, TN, MN, WA, NJ, PA, IL, CA, and IA, and country of CANADA. | Terminated | October 17, 2008 | 366 | 975 | Not_Computer
Z-2166-2008 | 48608 | No | LOX | Percutaneous Transluminal Coronary Angioplasty (PTCA) Catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific Scimed Maverick 2 Monorail 12mm Balloon Catheter, 2.5 mm x 12 mm |  |  | November 01, 2005 | 2005 | Boston Scientific Corporation | Two lots may not be correctly labeled. Specifically , a 2.5 mm x 12 mm unit was incorrectly labeled as a 3.5 x 20 mm unit on the box and tyvek label. | Packaging process control | Consignees were sent a "Boston Scientific Urgent Medical Device Recall or Correction-Immediate Action Required"letter dated November 7, 2005. The letter addressed to Risk mangers, described the issue and product. The letter stated that immediate discontinued use of and segregation of the recalled product was required. A Reply Verification Tracking Form was requested to be completed and returned. |  | 67 | 67 Units | International Distribution --- including countries of Czech Republic, France, Germany, Great Britain, Italy, Spain, Sweden, India, and Iran. | Terminated | September 10, 2008 | 1044 | 67 | Not_Computer
Z-0151-2009 | 49439 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Ray TFC Device with End Caps: |  |  | August 18, 2008 | 2008 | Stryker Spine | The label on one side on the Ray TFC Device (Ray Cage) is incorrect. It reads 20 mm x 26 mm instead of 20 mm x 21 mm. | Labeling mix-ups | Urgent Product Recall notification letters were sent via Fed Express on August 18, 2008 to Stryker Branches and Hospital Risk Managers. The firm asked customers to examine their inventory and hospital locations to identify the product; reconcile all products on the Product Accountability Form and attached distribution sheet; and fax a copy of the Accountability Form and distribution sheet to Sarah Thompson within 5 days of receipt of the notice. All questions are to be directed to Susan Krasny at 201-760-8150. |  | 69 | 69 units | Nationwide Distribution | Terminated | November 04, 2008 | 78 | 69 | Not_Computer
Z-0476-2009 | 49550 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | VASOVIEW 4 Endoscopic Vessel Harvesting System, with VASOVIEW Flexible Endoscopic Tools in the following configurations: VH-10242: Uniport plus dissection cannula, short port blunt tip trocar, flexible endoscopic scissors; and VH-10243: Uniport plus dissection cannula, short port blunt tip trocar, flexible endoscopic bisector. |  |  | September 09, 2008 | 2008 | CTS, Inc. dba Guidant Cardiac Surgery | Packaging unsealed: Sterility of product may be compromised due to packaging issue. For further information, please contact the firm at 408-635-6874. | Packaging process control | The firm issued a letter entitled "URGENT FIELD ACTION - IMMEDIATE ACTION REQUIRED *** VASOVIEW(TM) 4 Endoscopic Vessel Harvesting System (Kit Only)" via federal express to its consignees on September 9, 2008. The firm has asked its consignees to return the product. Follow-up calls were made to ensure recall effectiveness. |  | 7935 | 7,935 units | Worldwide Distribution --- including USA and country of Australia. | Terminated | December 08, 2010 | 820 | 7935 | Not_Computer
Z-0542-2008 | 45510 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | VNUS 7F ClosureFAST Catheter, Order #: CF7-7-60 (60 cm working length), Manufactured by VNUS Medical Technologies, Inc., 5799 Fontanoso Way, San Jose, CA 95138 |  |  | August 08, 2007 | 2007 | VNUS Medical Technologies Inc | Instructions for Use change: Instructions for Use have been changed to include a modified placement recommendation. | Labeling design | The Recalling firm is updating the Instructions For Use, training materials, and procedural guides to incorporate the treatment recommendations. The recalling firm sent an Update to customers who received the product, Important Update Regarding the Use of the VNUS ClosureFAST Catheter, dated August 6, 2007 with supplemental device instructions. Future product distribution will include the updated labeling. |  | 16716 | 16,716 units | Worldwide | Terminated | May 22, 2008 | 288 | 22937 | Not_Computer
Z-1308-2008 | 46814 | No | LKK |  |  | General Hospital | PMA | Medtronic Intrathecal Catheter Pump Segment Revision Kit; Programmable implanted infusion pump; Model Number: 8596SC; Packaged with catheter model number 8731 or 8731SC. |  |  | January 09, 2008 | 2008 | Medtronic Neuromodulation | Incorrect Connector Pin: The Medtronic 8596SC Intrathecal Catheter Pump Segment Revision kits were packaged with the incorrect connector pin. | Packaging process control | Consignees were notified by an Urgent Medical Device Recall letter sent in January 2008. The letter informed users that the recalling firm would contact them regarding removing the affected products from their inventory. Users were requested to return a response form indicating the number of affected product. For additional information, contact 1-800-707-0933. |  | 18 | 18 | Nationwide Distribution including states of MI, OR, IN, NY, FL, PA, MN, TX, CO and UT. | Terminated | October 08, 2010 | 1003 | 18 | Not_Computer
Z-0718-2008 | 38335 | No | JTC | Microtiter diluting and dispensing device. | Microbiology | Microbiology | 510(K) Exempt | Hamilton Sample Carriers for the ML STAR line (SMP-CAR-24). Product Code: 866.2500; Part Number: 173400. Hamilton Company, Reno, NV 89520 (Accessory on the Microlab STAR series instruments.) |  |  | June 25, 2007 | 2007 | Hamilton Co | Incorrectly labeled-- Product has the wrong barcode labeling. The barcode error will not influence the results of testing. The instrument will discover the error and post an error notice. | Error in labeling | The firm sent an Urgent: Product Recall letter to consignees on June 25, 2007. The firm also issued an alert via their User news distribution. They instructed consignees to examine their stock and return any affected carriers. A fax back form was supplied with the notification. The firm will be providing replacements. |  | 179 | 179 units distributed. | Worldwide Distribution --- USA including states of OH, CA, NJ, NC, and countries of Japan, France, Switzerland, Spain, Italy, Germany, and the UK. | Terminated | April 15, 2008 | 295 | 179 | Not_Computer
Z-0545-2008 | 45520 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand Clinac, Accelerator, Linear, Medical charged-particle radiation therapy system; |  |  | September 10, 2007 | 2007 | Varian Medical Systems Inc | The collimator drive chain may break or slip off of its drive track, allowing the collimator to rotate freely without motor control; if undetected resulting in a treatment with the wrong collimator angle. | Labeling design | On 9/10/07, the firm issued an Urgent Medical Device Correction letter to all its direct consignees, informing them of the device defect and providing instructions on the recall. Short term user instructions were provided and the firm will schedule follow-up visits to make the needed permanent device corrections at each device site. |  | 418 | 418 (total for Recalls Z-0544-0545-2008) | Worldwide- Canada, Mexico, and the USA | Terminated | November 27, 2010 | 1174 | 418 | Not_Computer
Z-1697-2008 | 47871 | No | EMA | Dental cement. | Dental | Dental | 510(k) | Dental Cement-Bisco BISCEMÂż Dual-Cured Self-Adhesive Resin Cement, Opaque; The product is intended to serve as a temporary tooth filling or as a base cement to affix a temporary tooth filling, to affix dental devices such as crowns or bridges, or to be applied to a tooth to protect the pulp. |  |  | April 28, 2008 | 2008 | Bisco Inc | During product stability testing, Bisco determined that after 15 months of storage, BISCEMÂż Self-Adhesive Luting Cement used in self-cure mode only without light-curing may take longer than specified to set, resulting in reduced bond strength. Bisco also determined that product component(s) may thicken slightly over time. | Incorrect or no expiration date | Bisco sent Urgent Medical Device recall letters dated 4/25/08 to the U.S. consignees via 1st class mail on 4/28/08. The letters informed the users that after 15 months of storage, BISCEMÂż Self-Adhesive Luting Cement used in self-cure mode only without light-curing may take longer than specified to set, resulting in reduced bond strength. Future production lots will be labeled with a shelf life of 1 year with a requirement to refrigerate the product when not in use. The accounts were requested to destroy their remaining stocks of the product and record the number of used or partially used syringes destroyed, indicating the lot numbers of the units destroyed on the enclosed response form and fax it back to Bisco at 847-534-66101. Questions were directed to Bisco Dental Products Customer Service at 1-800-247-3368. |  |  |  | Nationwide including Puerto Rico and internationally to Argentina, Australia, Bulgaria, Brazil, Canada, Chile, China, Colombia, Costa Rica, Denmark, Ecuador, Finland, France, Germany, Greece, Guatemala, Hong Kong, Iceland, India, Indonesia, Israel, Italy, Jordan, Korea, Lebanon, Malaysia, Netherlands, Norway, New Zealand, Peru, Philippines, Portugal, Romania, Saudi Arabia, Serbia, Singapore, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Uruguay, Vietnam | Terminated | September 14, 2010 | 869 |  | Not_Computer
Z-0193-2009 | 49529 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | AXIOM Artis. Angiographic X-ray System. |  |  | August 22, 2008 | 2008 | Siemens Medical Solutions USA, Inc | Patients could possibly fall off the table if not properly secured when: 1) moving them onto the patient table, 2) moving them on the patient table or 3) removing them from the patient table. | Labeling design | The recalling firm issued a Customer Safety Advisory Notice dated 7/04/08 (including an Addendum to the Operators Manual) which alerts the operator of possible injury to patients if they are not properly secured to the table during examinations and describes the recommended fastening procedures required to properly secure the patient to the table.For questions or additional information, contact Siemens Medical Solutions, Inc. at 610-448-4634. |  | 1530 | 1,530 units | Nationwide Distribution. | Terminated | September 23, 2009 | 397 | 1638 | Not_Computer
Z-1641-2008 | 47471 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | CADD-MS"3 Ambulatory Infusion Pumps, Model 7400, Rx only, CE 0473, SN 499000. The device is designed for subcutaneous, intravenous, epidural and intrathecal infusion of medication. |  |  | February 22, 2008 | 2008 | Smiths Medical MD, Inc. | Smiths Medical has received reports that the device powers down without an alarm. Investigation concluded that if the battery cap is not fully tightened, the pump may power down and a brief "chirp" will sound. This may occur if the user has not sufficiently tightened the battery cap or if the battery cap is damaged. If a patient is asleep or does not regularly check their Pump, there is a potentia | Labeling design | A "Smith Medical Urgent Medical Device Correction" letter dated February 22, 2008 was distributed to consignees. The letter addressed to Risk/Safety Managers, Distributors, Clinicians, Patients and other Customers, instructs them to fully tighten the battery cap and to routinely monitor the pump for damage that may result in power down situations. Consignees are to return the "Urgent Medical Device Correction" form.If you have questions about the letter or did not received the letter contact Mike Herbert at 651-628-7049 or Jeff Esselman at 651-628-7920. |  | 2303 | 2303 | AZCAFLGAHIILKYMDMIMNNCNENJNMNYPATNTXUTCANADACHILECZECH REPUBLICGERMANYGREAT BRITAINITALYJAPANNETHERLANDSNORWAYSWEDEN | Terminated | January 11, 2011 | 1054 | 2303 | Not_Computer
Z-1118-2008 | 45911 | No | JHI | Human chorionic gonadotropin (HCG) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Fisher Healthcare Sure-Vue Serum/Urine hCG-STAT; in vitro diagnostic test; Catalog Number: FHC-A202; Product Codes: FHC-A202-KFI3O &amp; FHC-A202-KFI5O; Inverness Medical Professional Diagnostics, San Diego, CA 92121. |  |  | October 01, 2007 | 2007 | Innovacon Inc | Incorrect results: The Sure-Vue Serum/Urine hCG-STAT test, when interpreted at extended read times, may exhibit sensitivity to patient samples containing hCG at levels well below the cut-off, potentially resulting in a false positive interpretation by the user. | Labeling design | Inverness Medical sent an Urgent Medical Device Alert letter, dated October 1, 2007, to all Innovacon customers for the product via US Postal Service return receipt requested. The field correction notice removed the extended read time of 10 minutes. Customers were instructed to note that although the package insert currently allows interpretation up to 10 minutes, do not read at times greater than 3 minutes for urine and 5 minutes for serum. |  | 1825480 | 1,825,480 units | Nationwide | Terminated | August 12, 2011 | 1411 | 1825480 | Not_Computer
Z-1128-2008 | 46693 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | ev3 IntraStent LD Mega Large Diameter Biliary Stent, Model Number: S17-36, |  |  | January 18, 2008 | 2008 | Ev3, Inc | Incorrect information on label: Specific lots of ev3 IntraStent Unmounted Balloon Expandable Stents have the two symbols for length and diameter on the side and end flaps reversed. | Labeling Change Control | ev3 notified all consignees via "Medical Device Recall" letter on 1/18/08. The letter was addressed to the Hospital and it described the device affected and the problem. The letter provided recommendations to segregate the affected devices and requested the return of the "Device Recall Field Action Reconciliation Form." |  | 70 | 70 | Worldwide Distribution - USA, ARGENTINA, AUSTRIA, BELGIUM, BRAZIL, CANADA, CHILE, DENMARK, GERMANY, GREECE, HONG KONG, IRAN, IRELAND, ITALY, JORDAN, KUWAIT, LEBANON, NORWAY, POLAND, RUSSIA, SAUDI ARABIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TURKEY, UNITED ARAB EMIRATES, and UNITED KINGDOM. | Terminated | November 04, 2008 | 291 | 942 | Not_Computer
Z-0593-2009 | 50108 | No | MQE | Assisted reproduction needles. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Wallace Oocyte Recovery Needles,Oocyte Recovery sets, 750 mm tube with flushing connector, Product Number: KÂżLN-750, and KÂżLN-750. |  |  | October 23, 2008 | 2008 | Irvine Scientific Sales Co Inc | Smiths Medical International became aware that certain lots of Wallace oocyte Recovery Sets have a damaged needle tip. | Packaging | The recall was initiated by Smiths Medical on October 22, 2008 with the forwarding a recall notification to Irvine Scientific. Smiths Medical International sells the products to Irvine Scientific for distribution in the U.S. Irvine Scientific is the U.S. Agent. Irvine Scientific issued an Urgent Field Safety Notice Letter, dated 10/22/2008, with attached fax-back Customer Response form to their customers via US Postal service beginning on October 24, 2008. Irvine Scientific requested their customers to complete an inventory summary on the Customer Response form and contact the firm to arrange for the return of any remaining product in their inventory to the firm. |  | 71 | 71 units | Nationwide Distribution --- including states of CA, FL, PA, IL, OR &amp; UT. | Terminated | September 02, 2011 | 1044 | 71 | Not_Computer
Z-1054-2008 | 46743 | No | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | CIVCO Latex-Free Needle Guide, Sterile transrectal needle guide for use with Shimadzu UB1OR-065U transducer, Part #677-057, packaged 1 needle guide per tyvek pouch, 20 pouches per box. The label shows the product is Mfg by CIVCO, Kalona, IA. |  |  | January 16, 2008 | 2008 | Civco Medical Instruments Inc | Incorrect Needle Guide- The individual pouches contained V53W Endocavity Needle guide(p/n/034-054) instead of the Shimadzu T/R Needle guide (p/n134-054). | Packaging | On 1/16/08, the recalling firm contacted their customers and notified them of the problem. Pictures of the problems were provided for identification. The firm requested that all products be returned. |  | 207697 | Lot M102250 - 5 boxes; Lot M105440 - 2 boxes | Nationwide Distribution- including states of TX and GA | Terminated | April 03, 2008 | 78 | 207697 | Not_Computer
Z-0517-2008 | 46044 | No | FFT | Stomach pH electrode. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Medtronic Bravo pH Capsule with Delivery System, Stomach pH electrode, Product Numbers 9012B1011 and 9012B1001. Medtronic, Inc., Minneapolis, MN 55432 |  |  | December 03, 2007 | 2007 | Medtronic Neuromodulation | Unable to detach from source: The capsule may not detach from the delivery system following attachment to the esophageal wall. | Labeling Change Control | Consignees were notified by an Urgent Device Recall in December 2007. The letter emphasizes to users the current Instructions for Use regarding detachment of the capsule from the deliver system. It advises that users may continue to use the Bravo capsules and delivery systems. It instructs users to discard all copies of the Âż2005 Bravo pH Monitoring System Interactive Training CD and to use the enclosed Âż2007 Bravo pH Monitoring System Interactive Training. A "Bravo Delivery System Disassembly Procedure" was also provided with the letter. The letter requests consignees to complete and return the reply form included with the notification. For additional information, contact 1-800-707-0933. |  | 169365 | 169,365 individual Bravo capsule and delivery systems have been sold worldwide | Worldwide Distribution -- USA states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV and WY and countries of Australia, Austria, Belgium, Brazil, Chile, Czech Republic, Denmark, France, Germany, Greece, Hong Kong, India, Indonesia, Ireland, Israel, Italy, Japan , Korea, Kuwait, Lebanon, Liechtenstein, Luxembourg, Malaysia, Mexico, Netherlands, Norway, Panama, Portugal, Russian Federation, Saudi Arabia, Serbia and Montenegro, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates and United Kingdom. | Terminated | May 03, 2012 | 1613 | 169365 | Not_Computer
Z-0988-2008 | 46341 | No | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | Stryker Round Fluted Bur (6.0mm) Extra-Long for use with TPS Universal Drill Extra Long "R" Attachment, sterile, Stryker Instruments, Kalamazoo, MI; REF 5160-10-60. |  |  | January 10, 2008 | 2008 | Stryker Instruments Div. of Stryker Corporation | Lack of assurance of sterility. | Package design/selection | Stryker sent a Medical Device Recall Notification letter dated 1/10/08 to consignees. The letter instructed them to check their facility for product, cease using it, and to return the product using a mailing label provided. International subsidiaries were notified via email. |  | 10588 | 10,588 of all products | Worldwide Distribution - USA, Australia, Brazil, Canada, England, France, Germany, Hong Kong, India, Italy, Japan, Mexico, the Netherlands, New Zealand, Portugal, Saudi Arabia, Poland, Romania, South Africa, Spain, Sweden, Switzerland and United Arab Emirates. | Terminated | December 01, 2009 | 691 | 10588 | Not_Computer
Z-0892-2008 | 46151 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Peel-Away Introducer set, 7.0 FR; Catheter introducer, includes sheath: 7.0 FR., 15.5 CM long, Wire: .035" DIA., 50 CM long, 3 MM J needle: 18 ga., 7.0 cm long. Accepts devices up to 7.0 French,sterile, Catalog Number: C-PLI-7.0-35. Cook Inc., Bloomington, IN 47404 |  |  | December 07, 2007 | 2007 | Cook, Inc. | Packaging Problem: The peel-away sheath does not peel uniformly or completely. | Package design/selection | Cook Inc. initially notified consignees via FEDEX by an Urgent: Product Recall on 12/7/07. The recall was then expanded and consignees were notified on 12/18/08 by another Urgent: Product Recall. The letters instructed any distributors to inform their customers of recall, and all users to return the recalled products and enclosed Recall Response Form to the firm. For additional information, contact 1-800-346-2686. |  | 1881 | 1,881 of total products | Worldwide Distribution- including USA and countries of Argentina, Australia, Canada, Chile, Denmark, Japan, New Zealand and Singapore. | Terminated | April 15, 2008 | 130 | 1881 | Not_Computer
Z-0668-2008 | 46275 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker GMRS Proximal Femoral Surgical Protocol; |  |  | December 18, 2007 | 2007 | Stryker Howmedica Osteonics Corp. | The GMRS Proximal Femoral Surgical Protocol states " The GMRS Proximal Femoral components are fully compatabile with all Stryker V40 femoral heads." This statement is incorrect and conflicts with the Alumina Head label and packaging insert. | Labeling design | Stryker sent Urgent Product Recall notification letters and Product Accountability forms via FedEx on 12/21/2007, return receipt requested. An Urgent Product Recall letter dated 1/7/2008, was sent to surgeons who are GMRS users, advising that the GMRS Proximal femoral components are not compatible with Alumina ceramic heads. |  | 5869 | 5869 | Worldwide Distribution - USA, Canada, Mexico, China, Australia, Hong Kong, Singapore, Sweden, Germany, Brussels, the Netherlands, Italy, Poland, Dubai, United Kingdom, Spain, and South Africa. | Terminated | October 15, 2008 | 302 | 5869 | Not_Computer
Z-1820-2008 | 48116 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Allez Spine Del Mar Pedicle Screw System: IT-FN 1007, Del Mar Sterilization Tray (Part Number: 30013) |  |  | April 14, 2008 | 2008 | Allez Spine, LLC | This field action was initiated because the cleaning, decontamination and sterilization procedures specified in the "Instructions for Use" for the impacted systems are incorrect.The use of these parameters in performing cleaning, decontamination and sterilization procedures may not result in clean and sterile product. | Error in labeling | Allez Spine issued a safety notice with attached fax-back attachment on April 14, 2008 via FedEx to all its hospital users, distributors and field representatives informing them of the failure of the sterilization validation and requesting them to use the new recommended sterilization parameters. The notice included copies of the updated Instructions for Use for the Allez Spine Laguna Pedicle Screw System and the Allez Spine Del Mar Pedicle Screw System. The safety notice requested that the revised recommended instructions be used for all future cleaning, decontamination and sterilization. Copies of the revised instructions were also provided with the safety notice along with links to the company website where additional copies may be downloaded.Contact Mark De Baca at 415-357-9908 if you have any questions. |  | 2 | 2 systems | Nationwide to AZ, CA, CO, GA, IL, NV, TN, TX, UT | Terminated | January 25, 2010 | 651 | 66 | Not_Computer
Z-1248-2008 | 46679 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Edwards LifeStent FlexStar Biliary Stent System; Biliary catheter and accessories; Model Number: EX060803CD; Edwards Lifesciences LLC, Irvine, CA 92614. |  |  | December 05, 2007 | 2007 | Edwards Lifesciences, Llc | Incorrect Expiration Date: Products were labeled with an incorrect shelf-life expiration date. Therefore, the sterility of the device may be compromised. | Packaging process control | Consignees were notified by an Urgent Product Recall letter sent on 12/5/07. The letter instructed users to return any affected product to the recalling firm for replacement products. The letter also advised users that the likelihood of patient complications is extremely low for products that have already been implanted in patients. Users were requested to return a confirmation form after checking the inventory for any affected products. For additional information, contact 1-949-250-3779. |  | 678 | Lot #FR6G0122 - 3 UNITS, FR6G0123 - 2 UNITS, FR6H0193 - 4 UNITS, FR6H0201 - 6 UNITS | Nationwide Distribution | Terminated | January 11, 2012 | 1498 | 10320 | Not_Computer
Z-1314-2008 | 47210 | No | EXB | Ostomy pouch and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | GentleTouch System-Colostomy/Ileostomy Postoperative Kit; Ostomy pouch and accessories; UPC Code Number: 30003-020938; 2 3/4 in., 70mm; ConvaTec, Skillman, NJ 08558. |  |  | March 07, 2008 | 2008 | ConvaTec | Mislabeled: The carton label reads "For postsurgical use following a urostomy procedure." It should read "For postsurgical use following a colostomy or ileostomy procedure." | Labeling mix-ups | Consignees were notified by an Urgent Product Recall letter sent on 3/7/08. The letter instructed users to return any unused affected product to the recalling firm for a credit. For additional information, contact 1-800-422-8811. |  | 9299 | 9,299 units | Worldwide Distribution including USA and country of Canada. | Terminated | October 21, 2008 | 228 | 9299 | Not_Computer
Z-0032-2009 | 49465 | No | JDC | Elbow joint metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Biomet brand Discovery Elbow, 3 mm x 75 mm right ulna with bearing/bond coat, TI 6AL 4V Alloy/ARCOM UHMWPE, For cemented use only, sterile, Biomet Orthopedics, Inc., Warsaw, IN; REF 114813. The product intended use is for Orthopedic implant. |  |  | May 23, 2008 | 2008 | Biomet, Inc. | The component in the package is not the correct size. | Packaging process control | Consignees were notified by letter dated May 28, 2008 to locate and return the recalled lots. Contact Biomet Orthopedics at 1-800-348-9500, ext. 3983 for assistance. |  | 6 | 6 | USA--California, Illinois, Indiana, Massachusetts, Ohio Puerto Rico, and Canada, OUS: Germany and Great Britain. | Terminated | September 22, 2009 | 487 | 11 | Not_Computer
Z-0711-2008 | 38631 | No | GFF | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ConMed Linvatec, Ultrapower Burs, Linvatec Corporation, Largo, FL 33773 |  |  | August 09, 2007 | 2007 | Linvatec Corp. | Compromised sterility: ConMed Linvatec medical devices packaged within a certain Peel Pouch Sterile Package may have compromised sterility due to an improper seal. | Packaging process control | Conmed Linvatec sent their consignees Urgent New Recall Medical Device Recall Notification letters On August 9, 2007 by UPS overnight delivery. The letter requests the customer return the devices to Linvatec for credit. Domestic Distributors / Sales Representatives were notified by telephone and email. |  | 1181 | 1181 units | Worldwide Distribution. | Terminated | October 01, 2009 | 784 | 1181 | Not_Computer
Z-0282-2009 | 50113 | No | JEY | Bone plate. | Dental | Dental | 510(k) | HT X-DRIVE SCREW, Part # 91-2416, W.Lorenz Surgical. Screws for use in the stabilization and fixation of mandibular fractures and mandibular reconstructive surgical procedures. |  |  | July 11, 2008 | 2008 | Biomet Microfixation, Inc. | The packaging for the 91-2416, 2.4X16MM HT X-DRIVE SCREW, Lot 493200 actually contained the 99-6577, 2.0X7MM FOSSA X-DRIVE SCREW. | Labeling mix-ups | An "Urgent Medical Device Recall Notice" was sent to customers on 07/11/2008 by certified letter. The letter described the problem, the risk and asked that customers discontinue use of the product and return them to Biomet. Also, customers were provided with a Inventory Reconciliation sheet and asked to fax it back to Biomet. Please contact Biomet Microfixation at 1-800-874-771 or 904-741-4400, ext. 448 for assistance. |  | 10 | 10 | Nationwide and Canada | Terminated | September 11, 2009 | 427 | 10 | Not_Computer
Z-0021-2009 | 49435 | No | GAM | Absorbable poly(glycolide/l-lactide) surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | VICRYL RAPIDE (polyglactin 910) Braided Coated Synthetic Absorbable Suture, Undyed, Non-USP. |  |  | August 11, 2008 | 2008 | Ethicon, Inc. | Package defect compromised the integrity of the primary seal, which could lead to premature suture degradation and/or impair the sterile barrier of the product. | Packaging | Letters were sent to distributors, Director of Surgery, Director of Risk Management, Director of Materials Management and Director of Operating Room on August 11, 2008, requesting that all affected lot numbers of Vicryl Rapide be discontinued immediately from sale and/or use. A separate correspondence was sent to hospitals. Letters to Director of Surgery, Director of Risk Management, Director of Materials Management and Director of Operating Rooms were advised of the recall and asked to discontinue use of any of the affected lot codes of sutures. Return and credit instructions were provided to the Director of Material Management. Contact Ethicon, Inc. at 1-908-218-2553 for assistance. |  | 6912 | 6912 | Worldwide: US and Canada OUS: Australia, New Zealand and Indonesia. | Terminated | June 05, 2012 | 1394 | 6912 | Not_Computer
Z-0030-2009 | 49462 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Platinum Plus" LT , REF 46-601, UPN M001466010, LT/0.018/180, Sterilized with ethylene oxide gas, Rx only, Legal Manufacturer Boston Scientific Corporation, One Boston scientific Place, Natick, MA 01760-1537, Made in USA: 8600 NW 41st Street, Miami, FL 33166-6202. The Platinum Plus" Guidewire facilitates placement of a catheter during diagnostic or interventional intravascular procedures. The wire can be torqued to facilitate navigation through tortuous arteries and/or avoid unwanted side branches. |  |  | August 25, 2008 | 2008 | Boston Scientific Corporation | Boston Scientific is initiating a voluntary recall of the Platinum Plus Guidewire (peripheral). BSC has received reports that product may be labeled as a 260 CM long guidewire when the actual packaged device is a 180 CM long guidewire and the product may be labeled as a 180 CM long guidewire when the actual packaged device is a 260 CM long guidewire. There is no expected potential for injury to | Labeling mix-ups | A "Boston Scientific Urgent Voluntary Medical Device Recall or Correction - Immediate Action Required" letter dated 8/25/08 was sent to each affected account (To The Hospital Field Action Contact). The letter includes the customer steps for recall, a complete listing of all product batches/lots being recalled, the Reply Verification Tracking Form(s), and a return shipping label. Contact Boston Scientific Corporation at 1-763-494-1700 for assistance. |  | 47 | 47 | USA: CAFLGAILINLAMNMONJNVNYOHPATNTXUTWAWI andWVOUS: France, Germany, Great Britain, Italy, and the Netherlands. | Terminated | December 10, 2011 | 1202 | 82 | Not_Computer
Z-0156-2009 | 49452 | No | FZP | Implantable clip. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Weck Hemoclip Traditional Ligating Clips, Rx Only, Sterile, Teleflex Medical, RTP, NC 27709 USA. The use of the reusable Teleflex Medical metal ligating clips is transient, invasive and intended for the application of hemostatic clips during general open surgery. Ligation clips are non-absorbable to thermostatically restrict flow of fluids within vessels. Clips are removed from the cartridge and applied in such a manner as to engulf vessels. Teleflex Medical metal ligating clips are intended for use in procedures involving vessels or anatomic structures for which the surgeon determines ligating clips are the best choice. Surgeons should select the size type, and material of the clip based upon their experience, judgment and needs. |  |  | August 22, 2008 | 2008 | Telefelx Medical | A hole in the sterile unit blister pack was detected that would compromise sterility. | Packaging process control | Consignees were notified via UPS next day on/about August 22, 2008. They were instructed to cease use and distribution and quarantine all affected product immediately. If they have any of the affected product they were instructed to return all it to Teleflex Medical via provided UPS-Ground. An Acknowledgement form was provided to be completed and returned via fax to Teleflex Medical to confirm receipt of the notification and amount of product on hand for return. A second mailing to non-responding customers on the consignee list will be made approximately four to six weeks after the first mailing. A third mailing to non-responding customers on the consignee list will be made approximately four to six weeks from the second mailing. Consignees also may be contacted by phone, e-mail, or fax. A record of these contacts will be maintained. Upon return to Teleflex Medical, product will be quarantined, assessed for recall effectiveness, and either reworked or destroyed. Contact Telefex Medical at 1-800-334-9751 ext. 8066 for assistance. |  | 786036 | 786036 (total pieces) | Nationwide | Terminated | March 15, 2012 | 1301 | 786036 | Not_Computer
Z-0310-2008 | 45461 | No | JDK | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | VariLift Bone Plug w/ End Cap; 13mm Dia, 24mm long; Ref #: CIE1-13-24-BP; Sterile; Encore Orthopedics, Austin, TX 78758. |  |  | September 19, 2007 | 2007 | Encore Medical, Lp | Device marketed without label warning required in the device's 510(k) approval letter. (In part-Warning: This device is not intended for any spinal indications.....) | Labeling Change Control | The recalling firm notified consignees of its recall via an Urgent Device Recall letter on 09/19/07. Consignees notified of a labeling error regarding device usage and were told they would be notified by a customer service rep to facilitate return of the recalled devices. |  | 61 | 61 units. | Nationwide: including direct accounts in CO, KS and TX. | Terminated | March 10, 2008 | 173 | 175 | Not_Computer
Z-0689-2008 | 46353 | No | MIH |  |  | Cardiovascular | PMA | Medtronic brand AneuRx AAAdvantage Endovascular Stent Graft with the Xcelerant Delivery System; |  |  | December 10, 2007 | 2007 | Medtronic CardioVascular | Sterility compromised: Device sterility may be compromised as evidenced by a loss of outer package integrity. | Package design/selection | Medtronic sent Urgent Communication - Medical Device Correction Advisory Notice letters to all its direct consignees on 12/7/07, via Federal Express, informing them of the affected product and providing instructions on avoiding contamination of the sterile surgical field with these devices. |  | 7000 | 7,000 | Nationwide Distribution. | Terminated | July 28, 2008 | 231 | 7000 | Not_Computer
Z-0915-2008 | 46347 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | BD 10 ml Syringe Luer Lok Tip, Reorder #3039605, BD, Franklin Lakes, NJ 07417. |  |  | January 07, 2008 | 2008 | Becton Dickinson &amp; Company | Defective Seal: The convenience trays may have open seals, which can adversely impact tray sterility. | Packaging | BD notified distributors by an Urgent Product Recall Letter sent via UPS on January 7, 2008. In addition, the firm sent an Urgent Product Recall Letter on January 11, 2008 to customers identified by the distributors. The letter instructed the users to remove the product from its inventory and to return affected products to the firm using an enclosed customer packing list. For additional information, contact 1-888-237-2762. |  | 2926560 | 2,926,560 | Nationwide, including the states of CA, NY, MO, MA, IL, OH, TN, TX, MI, LA, MO, ID, GA, NE, SC and FL. | Terminated | May 19, 2008 | 133 | 4680660 | Not_Computer
Z-0447-2008 | 45460 | No | CBH | Tracheal tube fixation device. | Anesthesiology | Anesthesiology | 510(K) Exempt | Dale 242 Blue (Neonate, Infants and Small Children) Tracheostomy Tube Holder, Product Number: H84102421, Dale Medical Products, Inc., Plainville, MA 02762 |  |  | September 27, 2007 | 2007 | Dale Medical Products, Inc. | Revised Labeling : Instructions For use: "Do not trim fastener hook tabs. " Trimming can cause tabs to disengage at the trach plate | Labeling design | Dale Medical notified Distributors by letter dated 9/27/07, identified as an Urgent Medical Device Correction Letter. The letter requested distributors to notify all their customers by letter. The Correction Letter reminded users to take notice of new labeling " Do not trim fasterner hook tabs" |  | 65000 | 65,000 | Worldwide-USA and countries of Canada, Austria, Belgium, Denmark, Finland, France, Germany, Greece, Iceland, Italy, The Netherlands, Norway, The UK, Spain, Switzerland, Australia, Hong Kong,India, Japan, Korea, Malaysia, New Zealand, The Philippines,Singapore, Taiwan, ROC, Brazil, Bermuda, Chile, Costa Rica,Ecuador, Mexico, Panama, Venezuela, Thailand, Saudi Arabia, and The United Arab Emirates. | Terminated | May 09, 2012 | 1686 | 665000 | Not_Computer
Z-0585-2008 | 46251 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Boston Scientific SlimLine 200 Micron Holmium Fiber Delivery Device, Catalog No. 840-840, Sterile, Made In Israel, Manufactured By LUMENIS, P.O.Box 240, Yokneam, ISRAEL. |  |  | December 18, 2007 | 2007 | Lumenis, Inc. | Incorrect Label: The Units outer pouch is labeled 200 micro fiber however, the inner pouch's label affixed to the product indicated that it is a 365 micron fiber. The micron fiber contained in the inner pouch is 200 micro fiber. The inner product is mislabeled with the wrong size. | Labeling mix-ups | The firm notified Consignee by certified letter on 12/18/2007, and asked that all affected products be return. Any questions concerning this recall can be addressed by contacting the Manager, Regulatory Compliance, at 1-877-586-3646 ext. 3355. |  | 59 | 59 units. | MA | Terminated | April 09, 2008 | 113 | 59 | Not_Computer
Z-0447-2009 | 48470 | No | HSD | Shoulder joint humeral (hemi-shoulder) metallic uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Modular Shoulder Body Assembly, Left (SH-1540L-S) and Modular Shoulder Body Assembly, Right (SH-1540R-S). Product labeled in part, "Acumed 5885 NW Cornelius Pass Road Hillsboro, OR 97124-9432 " |  |  | May 13, 2008 | 2008 | Acumed LLC | The packages for Modular Shoulder Body Assembly Left (SH-1540L-S) and Modular Shoulder Body Assembly Right (SH-1540R-S) may contain assemblies for the opposite side as indicated. | Labeling mix-ups | The firm issued a Urgent Notice: Device Recall on 5/13/08. The firm notified consignees of the issue and asked them to return the product to Acumed with the enclosed product recall effectiveness form for customers. For further information contact Acumed customer service at 1-888-627-9957 |  | 20 | 20 units | World-wide distribution including states of GA, MD, CA, MO, and AL, countries of United Kingdom, South Korea, South Africa, and Spain. | Terminated | May 13, 2010 | 730 | 20 | Not_Computer
Z-1646-2008 | 47535 | No | MUH | Extraoral source x-ray system. | Dental | Radiology | 510(k) | i-CAT 17-19 Cone Beam 3D Dental Imaging System. This product is also sold under a private label as KaVo 3D eXam |  |  | March 19, 2008 | 2008 | Imaging Sciences Intl Inc | Incorrect scatter information in the manual. | Error in labeling | Firm issued Urgent Medical Device Recall letters dated 3/19/08 to their customers. The distributors were provided background information and specific instruction on the steps to perform to correct the problem.If you have questions or have not received this letter, contact Francine McCoonse at 215-997-5666 ext 1340. |  | 1203 | 1203 units | The products were shipped to distributors and dental offices nationwide. The products were also shipped to Canada, Austria, Belgium, Czech Republic, England, Finland, France, Germany, Hungary, Ireland, Italy, Latvia, Lithuania, Netherlands, Norway, Portugal, Lavenia, Spain, Sweden, Switzerland, Turkey, United Kingdom, Australia, New Zealand, Chile, Israel, South Korea, UAE, India, Argentina, Brazil, China, Columbia, Hong Kong, Jamaica, Korea, Lebanon, Malaysia, Morocco, Russia, Scotland, Taiwan, Thailand, and Singapore. | Terminated | May 05, 2009 | 412 | 1203 | Not_Computer
Z-0587-2008 | 46317 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | MyelotecÂż Steerable Video Guided Catheter, 3.0mm, Label Insert contained in LB0020-0, LB0020-1, LB0020-2, LB0020-3, and LB0020-4 (Product Code: 2010). Myelotech, Roswell, GA 30076. |  |  | December 19, 2007 | 2007 | Myelotec, Inc. | Incorrect Labeling information: Product may contain inappropriate information in its label insert that refers to the product's use with energy delivering instrumentation. | Error in labeling | Consignees were first notified by phone on 12/19/2007. A follow up Urgent Label Insert Recall Notification letter was sent via certified mail on/about 12/28/2007. Consignees were instructed to remove the subject label inserts from the affected product in inventory, replace with the enclosed revised label inserts and to complete the attached Recall and Accountability Form. The incorrect label inserts were to be returned, via package provided to Myelotec. |  | 1252 | 1,252 shipping containers | Worldwide Distribution --- USA including states of AL, CA, CT, FL, GA, ID, IN, ME, MN, NH, NJ, NY, OH, and OR, and countries of Aurba, Korea, Japan, Switzerland, South Africa, UK. | Terminated | December 03, 2008 | 350 | 1252 | Not_Computer
Z-1593-2008 | 47424 | No | MNI | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | PathFinderÂż Spinal Fixation System T-Handle Driver; Part 2151-7 (13" shaft); Non-sterile, Reusable Instrument; manufactured by Abbott Spine, Austin, TX 78727. |  |  | March 24, 2008 | 2008 | Abbott Spine | Reports received regarding disassembly of T-Handle Driver components during surgery dropping the closure top into the surgical site. | Error in labeling | Firm sent notification letters to sales agents via letter on 03/24/08. Enclosed with the notification letter, the firm provided updated product inserts for the InCompassÂż and PathFinderÂż Spinal Fixation System (Cat. #2199-0002-LB Rev. H and #1199-0002-LB Rev. H, respectively), as well as the "Disassembly and Assembly Instructions for PathFinderÂż &amp; InCompassÂż T-Handle Drivers" (Cat #1199-0040-MKC Rev. A).Sales Agents instructed to distribute updated inserts and instructions to all locations that handle the drivers for cleaning and to destroy all copies of old inserts.If you have questions or have not received the recall notice from the firm, contact Kirstin Widding at 512-533-1852. |  | 224 | 224 units. | Nationwide Distribution --- including states of AR, AZ, CA, CT, CO, FL, IL, LA, MD, MI, MO, MN, MT, NC, NM, NJ, NY, OH, OK, OR, PA, SD, TX, VA and WI. | Terminated | December 02, 2009 | 618 | 502 | Not_Computer
Z-1381-2008 | 47219 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Scorpio Total Knee Cruciate Retraining Femoral Component; Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis; Product Number: 70-300-7R; Stryker Corp, Howmedica Osteonics Corp, Mahwah, NJ 07430. |  |  | February 08, 2005 | 2005 | Stryker Howmedica Osteonics Corp. | Mislabeled: The labeling indicates a Scorpio CR Femoral Component, however, the device inside the package is actually a Scorpio PS Femoral Component. | Labeling mix-ups | Consignees were notified by letter on 2/8/05. For additional information contact 201-831-5825. |  | 7 | 7 units | Worldwide Distribution including USA and countries of Canada, Mexico, Italy, The Netherlands and the UK. | Terminated | July 16, 2008 | 1254 | 12 | Not_Computer
Z-0843-2008 | 45812 | No | DQA | Oximeter. | Cardiovascular | Anesthesiology | 510(k) | Nonin Model 2500 PalmSAT handheld pluse Oximeter , Nonin Medical, Inc., Plymouth, MN USA |  |  | October 11, 2007 | 2007 | Nonin Medical, Inc | Incorrect Label: The label on the back of the device may show Model 2500A rather than Model 2500. | Error in labeling | An Urgen Field Action letters was sent (10/11/07) to Nonin distributors who received the potentially mislabeled product. The letters provided labeling differences and included instructions for notifying their customers, with a response form to be returned to Nonin. For additional information please contact Kim Aves at 763-577-3196. |  | 240 | 240 | Worldwide Distribution ---- USA including states of ME, SC, SD, CA, MO, NY, MI, FL, IL, PA, NC, CT, and TX, and countries of Canada, Greece, Russia, Italy, Sweden, Australia, &amp; Japan. | Terminated | December 18, 2011 | 1529 | 240 | Not_Computer
Z-0012-2009 | 48638 | No | CGX | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Roche Diagnostics NAPA2, N-Acetyl-Procainamide, COBAS INTEGRA, Roche Diagnostics, Indianapolis, IN.; 20737852322. |  |  | June 11, 2008 | 2008 | Roche Diagnostics Corp. | Some of the labels are glossy, resulting in the bar code being difficult to read by the bar code reader. | Labeling Change Control | Consignees were notified via an Urgent Medical Device correction letter dated 6/11/08 informing users of the problem and providing them with work-around instructions. The instructions suggests 1) inserting the cassette rack slowly; 2) using mainly reagent slots A, B, or C; 3) tilting the rack slightly clockwise so the barcode scanner does not read the barcode at an exact 90 degree angle; and 4) confirming the rack is properly seated on the track and displayed on the Cassette Status screen. If the cassette barcode does not register after following the Work-Around instructions, then customers are required to contact Diagnostics Technical Support. |  | 2 | 2 | Nationwide Distribution | Terminated | September 22, 2009 | 468 | 3624 | Not_Computer
Z-1305-2008 | 46785 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific Mach 1 peripheral Guide Catheter 8F, RDC Renal double Curve, .091 in (internal dia), Shaft lenght 55cm, Catalog Number: 19-866, UPN M001198660, Sterile EO, Availmed S.A. de C.V. (Mexico), Boston Scientific Corp., Maple Grove, MN 55311 |  |  | February 07, 2008 | 2008 | Boston Scientific Corporation | Mislabeled: Seven lots/ batches of Boston Scientific Mach 1 Guide Catheter may be mislabeled with the incorrect device length. | Labeling mix-ups | Consignees were sent a Boston Scientific Urgent Voluntary Medical Device Recall letter on 2/7/08. The letter included the customer steps for recall, a complete listing of all product batches being recalled, the Reply Verification Tracking Form(s) and a return shipping label. |  | 4 | 4 | Worldwide Distribution including USA states of PA, AZ, NY, KY, CA, WV, OH, WA, MI, TN, IN, IL, NJ, KY, FL, AR, FL, NC, CO, TX, WV, WI, OR, AL, MN and ND and countries of BE, FR, GB, DE, IE, IT, JP, AU, SOUTH KOREA, NZ, TR, TW and ZA. | Terminated | November 03, 2008 | 270 | 94 | Not_Computer
Z-1567-2008 | 46921 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | FASTPAK batteries labeled Manufactured for Medtronic Physio-Control 11811 Willows Rd, Redmond, WA USA. Batteries are Nickel-Cadmium (NiCd) battery packs used to provide DC power to LIFEPAK 5, 10, 11, or 12 defibrillator/monitors. The FASTPAK battery has a capacity of 1 amp/hour (1.0AH). The FASTPAK battery can be charged in a Battery Support System, Battery Support System 2, Mobile Battery Service Station, or the LIFEPAK 10, 11 or 12 defibrillator/monitor when the device is powered by an external power adapter. The FASTPAK battery does not have a pushbutton fuel gauge like the LIFEPAK NiCd. The fuel gauge provides a visual indication of battery capacity. |  |  | February 27, 2008 | 2008 | Physio Control, Inc. | Batteries mislabeled - batteries labeled as 2.4 amp/hour batteries, however are 1.0 amp/hour batteries. | Labeling mix-ups | On February 27, 2008, the firm sent a "URGENT MEDICAL DEVICE - ACCESSORY RECALL" letter via certified return receipt to their consignees. The letter states the batteries are fully functional 1.0 amp/hour batteries mislabeled as 2.4 amp/hour batteries and that the batteries will provide half the monitoring time and half the number of defibrillation shock as of a 2.4 amp/hour LIFEPAK NiCd battery. Enclosed with the letter is a confirmation sheet for the consignee to indicate the number of recalled batteries they have and fax back to the firm. The firm will send out replacement batteries and the consignee is to then send back the mislabeled batteries.If there are any questions, contact the technical support group at 1-800-442-1142 - option 5 or visit www.physio-control-notices.come/fastpak. |  | 811 | 811 batteries | Batteries were distributed nationwide and to Canada, Netherlands, Asia Pacific and Latin America. | Terminated | March 07, 2011 | 1104 | 811 | Not_Computer
Z-1179-2008 | 46802 | No | EXB | Ostomy pouch and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Active Life Little Ones One-Piece Custom Urostomy Pouch 5/16 - 1 inch, 8-25mm; Ostomy pouch and accessories, Product Order Number 020917, Nypro Dominican Republic Inc., ConvaTech, Skillman, NJ 08558 |  |  | February 14, 2008 | 2008 | ConvaTec | Mislabeled: The market unit carton label reads 5/8 - 1 inch, and it should read 5/16 - 1 inch. The metric units (8-25mm) on the label are correct. | Labeling mix-ups | Consignees were notified by an Urgent Product Recall Letter on 2/14/08. Updated letters were issued on 2/21/08 and 4/10/08 with expanded lot numbers. The letters instruct users to return any affected products and a product recall response form. For additional information, contact (908) 904-2199. |  | 3986 | 3986 market units | Nationwide Distribution. | Terminated | October 21, 2008 | 250 | 3986 | Not_Computer
Z-0693-2008 | 45847 | No | MHY |  |  | Neurology | PMA | Medtronic Kinetra 7428, Dual Program Neurostimulator for Deep Brain Stimulation, Rx only, Medtronic, Inc., Minneapolis, MN 55432-5604 USA. (The Medtronic Kinetra Model 7428 Neurostimulator is a multiprogrammable device and a component of the Kinetra Neurostimulator System for deep brain stimulation.) |  |  | October 01, 2007 | 2007 | Medtronic Neuromodulation | Cessation of Therapy: separation of internal connections (lifted wirebonds) between the electronic circuit and battery may lead to sudden cessation of therapy. | Packaging change control | A letter "Medtronic Urgent: Device Recall" dated October 2007 was sent to Healthcare Professionals. The letter described the device defect, product involved, root cause, probability of occurrence, recommendations and also provided information for assistance. At the time of the recall all of the product remaining in date for use (not expired) had been implanted, therefore Medtronic provided Important Patient Management Information for the implanting physicians to use in patient care and decisions regarding patient explants. |  | 1427 | 1,427 | Worldwide: USA including the states of AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MT, NC, ND, NE, NH, NJ, NM, NY, NV, OH, OK, OR, PA, RI, SC, SD, TN, TX, WA, WI, WY, UT, VA, VT, and WV and countries of Australia , Austria , Belgium , Canada , Cyprus , Czech Republic , Denmark , Finland , France , Germany , Greece , Hong Kong , Hungary , Iceland , India , Iran , Israel , Italy , Japan , Latin America, Lebanon , The Netherlands , Norway , Poland , Portugal , Russian Federation , Saudi Arabia , South Africa , Spain , Sweden , Switzerland , Turkey , and The United Kingdom | Terminated | May 03, 2012 | 1676 | 3079 | Not_Computer
Z-0288-2009 | 49709 | No | JDQ | Bone fixation cerclage. | Orthopedic | Orthopedic | 510(k) | Biomet 2.0 MM Stainless Steel (SS) Crimp Sleeve, sterile, use w/2.0 MM S.S. cable # 350800, Biomet Orthopedics, Inc., Warsaw, IN; REF. 350805. |  |  | September 24, 2008 | 2008 | Biomet, Inc. | The outer package is properly labeled, but the inner package may be labeled as a femoral component. | Labeling mix-ups | The firm's sales force was notified of the problem via an Urgent Medical Device Removal Notice letter dated 9/24/08, and provided instructions to retrieve the product from customer accounts and to provide them with a copy of the removal letter. Distributors were to immediately locate and remove the identified implants from circulation, carefully follow the instructions, and fax a copy of the Response Form prior to the return of the product. If the product was further distributed, hospital personnel MUST be notified via the enclosed "Dear Biomet Customer" notice. The letter MUST be given to hospital personnel responsible for receiving recall notices but distributors are charged with the location and return of those products. |  | 48 | 48 | Worldwide Distribution --- USA including states of Arkansas, Florida, Kansas, New Jersey, and Ohio, and countries of Argentina, Canada, Chile Italy, Mexico and United Kingdom. | Terminated | June 29, 2009 | 278 | 48 | Not_Computer
Z-1614-2008 | 46905 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Fiducial Markers (0.9 x 3 mm) Soft tissue gold markers, Medical charged-particle radiation therapy system; Part Number: MT-NW-887-809, CIVCO Medical Solutions (MEDTEC), Orange City, IA 51041. |  |  | February 28, 2008 | 2008 | Med Tec Inc dba CIVCO Medical Solutions | Mislabeled: Fiducial markers were incorrectly labeled. They are marked (0.9x3mm cylindrical) Soft tissue gold markers, but contain (2mm spheres) Bone gold markers. | Labeling mix-ups | Consignees were notified via telephone beginning 2/28/08 and a Recall Notification letter sent on 2/29/08. Users were requested to return any affected products for a replacement. They were also requested to return a recall notification form indicating the amount of affected product in their inventory. For additional information, call 800-842-8688. |  | 33 | 33/3-marker pouches | Nationwide Distribution including states of WA, PA, DE, AL, CT, WI, FL, TX, MN, and IL. | Terminated | July 02, 2008 | 125 | 69 | Not_Computer
Z-0706-2008 | 46141 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Biomet CC Cruciate Tibial Tray, 67 mm.; Product No. 141232, Biomet, Inc., Warsaw, IN 46581 |  |  | December 12, 2007 | 2007 | Biomet, Inc. | Mislabeled as to size: Package labeled as containing a 67 mm tray actually contains a 75 mm tray and vice versa. | Packaging process control | Consignees were notified via an Urgent Medical Device Recall Notice dated 12/12/07 to cease use of the product and to return it. |  | 8 | 8 | Worldwide: USA including states of Georgia, Idaho, Indiana, Missouri, Ohio, and Virginia and country of Turkey. | Terminated | June 11, 2008 | 182 | 8 | Not_Computer
Z-1608-2008 | 46863 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MT-APSD-2.4 Type-S Thermoplastic Mask, Medical charged-particle radiation therapy system; Disposable non-sterile (43.2cm) standard perforated thermoplastic head, neck and shoulder mask (2.4mm), 1 device per package; CIVCO Medical Solutions (MEDTEC), Orange City, IA 51041. |  |  | February 25, 2008 | 2008 | Civco Medical Instruments Inc | Mislabeled: The mask is mislabeled with the incorrect part number but the outer packaging is correct. | Labeling mix-ups | Consignees were notified by a Recall Notification letter on 2/25/08. The letter informed users that the recalling firm will replace the affected product. It also requested that users return a recall notification acknowledging notification of recall. Users were instructed to contact their sales representative to arrange servicing or for additional information at 1-800-842-8688. |  | 160 | 160 | Worldwide Distribution including USA states of IL, PA, NC, TN, SD, CA, MI, WI, MA, and CO; and country of Japan. | Terminated | July 02, 2008 | 128 | 160 | Not_Computer
Z-1206-2008 | 45901 | No | MCY | Elastic bandage. | General Hospital | General &amp; Plastic Surgery | Enforcement Discretion | Smith &amp; Nephew, Inc., Exu-Dry Wound Burn Dressing, Product Number: 5999M36, Exu-Dry Pad 24"x36" Non-Permeable Heavy Absorbency, Smith &amp; Nephew, Inc.,Largo, Florida. |  |  | November 21, 2007 | 2007 | Smith And Nephew, Inc. Wound Management Division | Sterility (package integrity) compromised: Distribution of product that did not meet specifications; open seals. | Package design/selection | Consignees were notified of recall by certified mail, Urgent Device Recall letter dated 11/21/2007, which was sent to each consignee. The recalled product will be returned to the firm by pre-paid FedEX return labels. A second letter is to be issued to each distributor for recall to the retail level, as the firm's original recall strategy of distributor level recall, was deemed inappropriate by CDRH |  | 5100 | 5,100 Dressings | Worldwide-USA, England, Japan, South Africa, New Zealand, Australia and Canada. | Terminated | December 15, 2008 | 390 | 47554 | Not_Computer
Z-1549-2008 | 47474 | No | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Medex MX821L 72 inch (182.9cm) Administration Sets, Standard Drip, Sterile Single Use, Latex Free device packaged 25 units per case. |  |  | March 24, 2008 | 2008 | Smiths Medical Asd Inc | Misbranded/mispackaged device was distributed. Product labeled to contain 72 inch Standard Drip ( IV) Administration Sets, actually contained 60 inch (IV) administration sets with a' Y' site. | Packaging process control | On 3/24/2008, the firm issued separate URGENT: PRODUCT RECALL NOTIFICATIONs sent via certified mail to their distributor and medical customers. Both recall notifications inform the customers of this labeling error and ask that the customers: 1) check their inventories for the presence of the suspect product; 2) complete and return the attached 'Product Recall Notification Response Form'; and 3) make arrangements to return any unused affected product to the recalling firm in exchange for a credit or replacement product. In addition, the distributors are asked to either provide the recalling firm with a list of their sub-account customers so that Smith Medical can contact the enduser customers and notify them of the recall, or as an alternative measure, the distributor can notify their branches and customers directly, via their own certified letter. |  | 350 | 350 units | Worldwide Distribution --- USA including states of PA, NJ, IL, MN and ND and country of Canada. | Terminated | September 30, 2010 | 920 | 350 | Not_Computer
Z-1445-2008 | 47244 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | QWIX Screw 4.3mm diam; length 26 mm; The stabilization screw is indicated for fixation of bone fractures or for bone reconstruction. |  |  | February 27, 2008 | 2008 | Integra LifeSciences Corp. | Certain QWIX Fixation Screws (Reference numbers 111426S and 111428S) have been etched and labelled with an incorrect length. | Labeling mix-ups | Integra Reps were sent recall notification with a recall acknowledgement on 2/27/2008. The customers who were shipped affected product were also notified on that date. |  | 42 | 42 units | Nationwide Distribution --- including states of NC, IL, CA, NC and TX. | Terminated | August 21, 2008 | 176 | 62 | Not_Computer
Z-1193-2008 | 46726 | No | JDZ | Cement mixer for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker VertePort, 10 G Cement Cannula, REF 306-410, Qty 18; For use with the Stryker PCD Precision System and ACM Delivery System, Rx only, sterile, Stryker Instruments, Kalamazoo, MI. |  |  | February 13, 2008 | 2008 | Stryker Instruments Div. of Stryker Corporation | Mis-labeling; The instructions for use include a use for which the device is not intended. (Device not approved to capture a bone biopsy) | Labeling Change Control | Consignees were notified via Medical Device Recall Notification letter dated 2/13/08 to read the enclosed instructions for use, remove all old instructions for use at the consignee location and replace them with the new instructions for use, and to discard the old instructions for use. |  | 953 | 953 boxes Recalls Z-1192-1195-2008 | Nationwide. | Terminated | December 30, 2008 | 321 | 953 | Not_Computer
Z-0557-2008 | 45862 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | FRAZIER Surgical Suction Instrument, 7 French with Control Vent and Obturator (Bend with Obturator in Place), Product Code: 0031070, 50/C. STERILE, EO, Rx ONLY. CONMED CORPORATION 310 Broad Street, Utica, NY 13501 |  |  | September 12, 2007 | 2007 | Conmed Corporation | Sterility (package integrity) compromised -- Frazier and Poole suction instruments may have an inadequate seal. The product pouch seal made during assembly of this product, may be incomplete or not present. | Packaging process control | URGENT-Device Recall Letters (dated 9/10/07) and FIELD CORRECTIVE ACTION CERTIFICATION forms were sent via Certified Mail to affected customers the week of 9/09/07 (mailed on 9/12/07). The recall letters informed the consignees of the recall due to potentially inadequate seals of the instruments, and instructed them to review their inventory of the instruments checking them for inadequate seals. The letters also instructed the consignees to complete the form and fax it back by 9/21/07. For those consignees that have any of the recalled instruments or may have questions, please contact ConMed Quality Assurance at 1-800-552-0138 ext. 5225 between the hours of 7 am - 6 pm (MST). International customers were notified on 9/18/07. A second recall letter (dated 10/09/07) was sent to non-respondents of the initial mailing. ConMed plans to send a third recall letter to the non-responders during the week of 12/10/07 which addresses sub-recall notifications to the distributors' customers. |  | 53691 | 53,691 instruments | Nationwide and Worldwide Distribution | Terminated | August 30, 2012 | 1814 | 4700864 | Not_Computer
Z-0948-2008 | 39366 | No | BSP | Anesthesia conduction needle. | Anesthesiology | Anesthesiology | 510(k) | Functional Anaesthetic Discography (F.A.D.) Catheter |  |  | April 25, 2006 | 2006 | Kyphon Inc | Guidewire Breakages: during the Functional Anaesthetic Discography (F.A.D.) procedure. | Labeling design | A customer advisory letter was sent on 4/25/06 to consignees. In some cases some were issued in person. Follow up with field representatives was also conducted. Both contacts involved providing IMPORTANT CHANGES IN INSTRUCTIONS FOR USE that are to be attached to the on site users documents. The recalling firm is updating future distribution with the new instructions for use. |  | 23 | By Lot Number : J 6013102 = 147, Lot: J6033012 = 23 | Nationwide including states of AL, AZ, CA, FL, IN, NH, NY, DC, ME, MD, WI, and PA. | Terminated | June 10, 2008 | 777 | 66 | Not_Computer
Z-0227-2009 | 49484 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Hipstar V40 Femoral Stem |  |  | August 20, 2008 | 2008 | Stryker Howmedica Osteonics Corp. | The warning label "do not use with heads more than +10, offset" does not appear on 135 degree stem boxes. This warning is contained in the instructions for use. | Labeling mix-ups | Recall notification letters were sent by Fed Ex on August 20, 2008 to Stryker branches/agencies and Hospital Risk Management departments. Letters to Branches stated the reason for the recall and the potential hazard. Both letters state that it is recommended that physicians follow patients in a manner that is consistent with a physicians routine standard of care for total hip arthroplasty. Contact Stryker Howmedica Osteonics Corp. at 1-201-831-5000 for assistance. |  | 77 | 77 units | Nationwide | Terminated | June 01, 2009 | 285 | 77 | Not_Computer
Z-0695-2008 | 45889 | No | CBT | Arterial blood sampling kit. | Anesthesiology | Anesthesiology | Enforcement Discretion | Smiths Medical, Add-On Kids Kit , Blood Sampling System, Product Code: MX4037, Smiths Medical ASD, Inc, Dublin, Ohio 43016 |  |  | November 21, 2007 | 2007 | Smiths Medical Asd Inc | Misbranded; The product codes are labeled incorrectly with a green promotional label for "New split septum sampling site. Only use needle-free access devices .......", however this product code contains a sampling site which is accessed with a shrouded needle, NOT with a needle-free access device | Packaging process control | Smiths Medical mailed certified letters (Urgent Product Advisory Notifications) dated November 19, 2007, to all customers that received the affected lot. Firms are instructed to make the field correction (extra label removal) or mail the product back to Smiths for the correction. |  | 1440 | 1440 units | Nationwide including the states of CA, DC, IL, IN, MD, MN, NC, NE, SC, and TN and Puerto Rico. | Terminated | September 30, 2010 | 1044 | 1690 | Not_Computer
Z-1633-2008 | 47361 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | 6F Viking Guiding Catheter, Percutaneous catheter, Shape ALR 1.2, Part Number: 1001934-06, Abbott Vascular, Santa Clara, CA 95054. |  |  | March 17, 2008 | 2008 | Abbott Vascular-Cardiac Therapies dba Guidant Corporation | Mislabeling: The part number on the product label packaging may not match the product. | Labeling mix-ups | Consignees were notified via an Urgent Device Recall letter dated and hand delivered starting 3/18/08. The letter requests users to remove and quarantine the affected lots, which will be picked up and replaced by the recalling firm. Users were also requested to return the enclosed Efficiency Check Reconciliation Form. Abbott Vascular will exchange all returned devices. For additional information contact 1-800-227-9902. |  | 10 | 10 units | Nationwide Distribution. | Terminated | July 07, 2011 | 1207 | 17 | Not_Computer
Z-0698-2008 | 45949 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Spectranetics Quick-Cross Support Catheters, .035in x 90cm, REF#: 518-036, Sterile, Manufactured by: Spectranetics, Colorado Springs, Colorado 80907. |  |  | December 10, 2007 | 2007 | Spectranetics Corporation | Mislabeling: Catheters were placed into mis-labeled inner pouches as REF Number 518-037, 135 cm--the exterior box labeling is correct as 90 cm. | Packaging process control | Consignees were notified by Dear Healthcare provider letter on 12/10/2007 and were told that a sales rep would remove and replace all affected catheters. |  | 64 | 64 | Nationwide including states of FL, IL, IN, MO, NC, NY, TX, and UT. | Terminated | February 08, 2008 | 60 | 64 | Not_Computer
Z-1550-2008 | 47334 | No | FZW | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Zimmer Dermacarriers II skin graft carriers, 1.5 to 1 dermacarrier, Rx, latex free, sterile, Zimmer, Dover, Ohio; REF 00-2195-012-00. |  |  | April 03, 2008 | 2008 | Zimmer Inc. | Sterility compromised. Lack of assurance of sterility, as the packages may not have been sealed on the manufacturer's end. | Packaging process control | Consignees were notified via letter (Urgent: Medical Device Correction) dated 4/3/08 to identify their stocks on hand, verify package sealing, to discard any non-sealed packages, mail the Business Reply Card regardless if there are any remaining affected products. |  | 99345 | 99345 of all units | Worldwide Distribution --- including USA and countries of Algeria, Australia, Austria, Belgium, Brazil, Canada, Chile, China, Czech Republic, Denmark, Dubai, Ecuador, England, Finland, France, Germany, Ghana, Greece, Guatemala, Hong Kong, Hungary, Iceland, India, Iraq, Israel, Italy, Japan, Jordan, Kenya, Largos, Lebanon, Lithuania, Malta, Mexico, Morocco, Netherlands, Nigeria, Panama, Peru, Poland, Saudi Arabia, Slovakia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Tunis, Turkey and Venezuela. | Terminated | December 17, 2009 | 623 | 99345 | Not_Computer
Z-0483-2009 | 49586 | No | LTD |  |  | Toxicology | Contact ODE | RNAgentsÂż Total RNA Isolation System (Cat. # Z5110). Kit Box label/Box drawer label. For Laboratory Use. Promega Corporation, Madison, WI. USA. |  |  | June 04, 2008 | 2008 | Promega Corpopration | Promega Corporation received reports of leaking bottles of Phenol:Chloroform:Isoamyl Alcohol (Part # Z5112), a component of the RNAgentsÂż Total RNA Isolation System (Catalog # Z5110). Promega product Z5112 is a mixture of phenol, chloroform and isoamyl alcohol. If phenol contacts skin, it can cause burns and poisoning. The fumes of this product are also harmful. | Packaging | Consignees were sent a "Promega Urgent Product Alert-Risk of Injury" letter dated June 9, 2008. The letter described the product, problem and provided instructions on handling, inspection and disposition of product. In addition, E-mail notifications were sent to distributors and branch offices on May 29,2008. For questions and additional information, contact your local Promega representative or Promega Technical Service at 800-356-9526 and press "4", or e-mail at techserv@promega.com. |  | 836 | 836 | Worldwide Distribution: USA including states of IN, ME, NE, NY, WI, PA, NC, FL, OR, UT, MI, IL, AL, OR, GA, CA, MD, TX, MA, UT, MO, OH, and LA, and countries of Canada, Mexico, Brasil, China, Japan, Germany, Singapore, Great Britain, Australia and France. | Terminated | June 26, 2012 | 1483 | 836 | Not_Computer
Z-1182-2008 | 46877 | [6 empty] | System 83 Plus Washer-Disinfector, Endoscope use, DOS based system, Custom Ultrasonics Inc., Buckingham, PA |  |  | February 11, 2008 | 2008 | Custom Ultrasonics, Inc. | Inadvertent selection of the "wash" cycle | Labeling design | The recalling firm issued a Product Correction letter dated 2/11/08 to its customers informing them of the problem and the correction choices. The letter states that a key cover is available to reduce the error in the future and will be provided to each users upon return of the recall letter notification questionnaire. |  | 3190 | 3,190 units | Worldwide; USA medical facilities and physicians and Canada | Terminated | September 03, 2009 | 570 | 3190 | Not_Computer
Z-0578-2008 | 46158 | No | JEB | Manual operating table and accessories and manual operating chair and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | TENET Medical Engineering Wrist Stabilization Kit, Medium; REF (Part) #7692; Sterile; manufactured by TENET Medical Engineering, Calgary, Alberta, Canada T2Z 4M3. |  |  | November 23, 2007 | 2007 | Tenet Medical Engineering Inc | Sterility compromised: Sterility seals on some of the firm's wrist stabilization kit pouches may be compromised. | Packaging process control | TENET notified US distributor consignee via Urgent Product Recall Notification letter sent FedEx on 11/23/07. Consignee asked to return product and to contact their consignees to also return product. |  | 360 | 360 units. | Nationwide Distribution. | Terminated | May 01, 2009 | 525 | 820 | Not_Computer
Z-1620-2008 | 47521 | No | BTA | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | FLAT DRAIN -10MM, CHARIKER-JETER DRESSING KITS (LARGE AND SMALL) -Product numbers 6680217 and 6680221 - Wound Management Product. |  |  | March 31, 2008 | 2008 | Smith And Nephew, Inc. Wound Management Division | Incorrect labeling. Kit labels may incorrectly state that a flat drain is contained as a kit component while some kits may contain a round drain. | Labeling mix-ups | Consignees were notified of recall by letter (Correction and Removal Notification) sent to customers on 3/31/08. They were asked to check their stocks, identify, and return all affected units to Smith &amp; Nephew according to the instructions. Consignees are to contact the Customer Care Center at 1-800-876-1261, if they have additional questions. |  | 4796 | 4796 | Nationwide Distribution. | Terminated | November 18, 2008 | 232 | 4796 | Not_Computer
Z-0826-2008 | 45532 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Xoran MiniCAT System for ENT (server and work station), Xoran Technologies, Ann Arbor, MI |  |  | January 04, 2008 | 2008 | Xoran Technologies, Inc. | Misbranding; These devices do not bear the manufacturer's name or address. | Labeling design | Consignees were notified via letter, Urgent Medical Device Correction, dated 1/4/08 of the need to apply proper labeling to the products, and were sent instructions and labeling to affix to their units. |  | 71 | 71 | Nationwide. | Terminated | October 22, 2010 | 1022 | 71 | Not_Computer
Z-1538-2008 | 46916 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Stryker Wedge HS Suture Anchor II with Needles, 6.5 mm, Model number 3910-100-145 |  |  | February 07, 2008 | 2008 | Stryker Endoscopy | Product package may crack and cause loss of sterility. | Packaging | Stryker Endoscopy initiated the recall on 02/07/2008 by notifying its consignees via telephone, as well as sending out Urgent: Device Removal notices. Recipients of affected lots are to inspect their inventory for part/lot numbers and quarantine any affected product. Affected products are to be returned to Stryker Endoscopy and Customer Service should be contacted for replacement alternatives. Also, the "Acknowledgment of recipt" card needs to be signed and returned to Stryker Endoscopy. |  | 53 | 53 units | Product was distributed to 8 consignees in CA, OH, PA, CO, MA, and NJ. | Terminated | December 21, 2010 | 1048 | 927 | Not_Computer
Z-0849-2008 | 45912 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | NexframeÂż Stereotactic System Kits, Model Number(s): DB-1021-MR, DB-1031, DB-1040-BL, DB-1040-ST, DB-1041, |  |  | October 04, 2007 | 2007 | Medtronic Image Guided Neurologics, Inc. | Sterilty (package integrity) Compromised: Some failures were for damage to the outer pouch, while another set of failures were for the seals on this pouch. The seal between the inner tray and lid has not been compromised and the contents remain sterile, however, the sterility of the outer surfaces of the inner tray and lid cannot be assured. | Package design/selection | On 10/4/2007, consignees were notified by Medtronic field personnel, either by personal visits or telephone. All wholesalers of this product are OUS. Medtronic OUS Regulatory personnel communicated this issue directly to the wholesalers, who provided the letters to their customers. Medtronic field personnel are responsible for completing and returning reply cards confirming all notifications (100%). |  | 385 | 385 | Worldwide, USA, Taiwan, Thailand, Japan, Spain, Sweden and The Netherlands. | Terminated | July 13, 2012 | 1744 | 385 | Not_Computer
Z-1750-2008 | 47764 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Omnifit EON Surgical Protocol; |  |  | May 17, 2007 | 2007 | Stryker Howmedica Osteonics Corp. | The Table 2 Broach and Neck Trial Sizing (Literature # LSP48 04/04) on page 3 is incorrect, reading 5mm less than what it should be for certain sizes. | Labeling design | Styrker Branches/Agencies were notified of the incorrect information printed in the protocal on May 17, 2007 via an Important Product Correction letter. The letter requested that all incorrect copies of the protocol be located and destroyed immediately. They requested that the Acknowledgement Form be completed and faxed, and that a Stryker Orthopaedics Customer Service rep be contacted in order to reorder the destroyed product. |  | 2501 | 2501 units | Worldwide Distribution --- including USA and countries of Canada, Australia, Latin America and Greece. | Terminated | August 19, 2008 | 460 | 2501 | Not_Computer
Z-0934-2008 | 45407 | No | JPA | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(k) | Inotech brand Heparin Adsorbant, Catalog #: HA-1, in 10mm x 75mm polystyrene tubes, each containing 70mg of Heparin Adsorbent, 10 tubes per foil lined pouch, 10 foil pouches/box, Inotech Biosystems International, Inc., Rockville, MD 20855 |  |  | August 20, 2007 | 2007 | Inotech Biosystems Intl., Inc | Erroneous Coagulation Tests | Labeling design | The recalling firm notified consignees by phone on 08/20/07 and provided instructions for increasing the speed of centrifugation from 1,500 x g to 2,500 x g and decanting the supernatant aliquot carefully so as to not draw in resin fibers that may be stuck on the test tube walls during sample preparation for testing in the coagulation studies. The firm updated their recall notification on 10/01/07 as an Urgent Recall/Correction to alert healthcare practitioners to not use "lupus sensitive reagents" with the heparin absorbant product and that Platelet Factor 3 and any other negatively charged molecule in the plasma may interact with the resin causing false high readings due to the partial removal of the clotting initiating factors. |  | 37370 | 37,370 tubes | Worldwide; USA to Hospitals and Clinical Laboratories and to one domestic distributor for export to Canada. | Terminated | August 07, 2012 | 1814 | 37370 | Not_Computer
Z-0174-2009 | 49360 | No | FMK | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Premier Value Lancing Device. UPC Code: 840986016707, The Package Includes: 1 Lancing Device, 1 Clear Cap and 5 Premier Value Ultra Thin Lancets, Distributed by: Chain Drug Consortium, LLC, Boca Raton, FL 33431. |  |  | July 18, 2008 | 2008 | Facet Technologies | The Alternate Site Cap was not included in the package as indicated on the labeling. | Error in labeling | The consignees were notified by letter on/about 7/18/2008. The Return Material Authorization number was communicated to the customer. They were asked to place the items on hold and Can Am will obtain any quantities. The Returned Product Mechanics Process will: Inventory will be sent back to Facet in two stages, Inventory on hand at Cimetra will be sent back to Facet, Inventory received back from customers will be sent back to Facet in one shipment 30-60 days after notice. Inventory received back after this second shipment will be destroyed. |  | 2976 | 2976 units | Nationwide--GA and FL. | Terminated | August 24, 2009 | 402 | 2976 | Not_Computer
Z-1040-2008 | 37183 | No | ERL | Ear, nose, and throat electric or pneumatic surgical drill. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | Anspach MicroMax Pneumatic Motor System; ENT Surgical Drill, including handpiece and Accessories Attached. All Models; Black Motors; Anspach Effort, Inc., Palm Beach FL 33410. |  |  | December 18, 2006 | 2006 | The Anspach Effort, Inc. | Incorrect Sterilization Process --The Operating Manual for the product contains an incorrect statement regarding the sterilization of the Autolube foot control. The manual states that the product can withstand ETO sterilization, but there is no data to support this process. | Labeling False and Misleading | Anspach issued an Urgent Device Correction letter on 11/27/06, to all manual holders advising them of the error and requested that they update the manual with the corrected instructions. A reply receipt confirming notification was included. |  | 7399 | 7,399 total units | Worldwide Distribution | Terminated | July 06, 2012 | 2027 | 7399 | Not_Computer
Z-0216-2009 | 49427 | No | JJX | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Architect LH MasterCheck (List 6C25-05). An in-vitro diagnostic, consisting of 4 - 2 mL vials of LH MasterCheck, Levels 0, 1, 2 and 3. Manufactured for Abbott Laboratories, Abbott Park, IL 60064. |  |  | September 05, 2008 | 2008 | Abbott Laboratories | The values listed in the Architect LH MasterCheck, Lot 82520 data sheet are incorrect. When MasterChecks do not perform as intended, the integrity of the system cannot be verified and patient results would not be generated. Other commercially available materials are available to customers in order to meet CLIA requirements. | Error in labeling | A Product Correction, Immediate Action Required letter and Customer Reply Form dated 9/5/08 were sent to all Architect LH MasterChecks (List Number 6C25-05) customers who received Lot 82520M. Customers were informed that the values listed in the Architect LH MasterCheck Lot 82520 data sheet are incorrect. Customers were requested to check their inventory for the affected lot, and if any kits were found, to substitute the values printed on the data sheet with the values supplied in the letter. In addition, Customers were requested to keep a copy of the letter on file and to complete the Customer Reply Form, indicating that the letter was received, understood and disseminated to the laboratory manager/supervisor responsible for the Architect LH MasterCheck testing, and fax it to Abbott by 9/19/08 at 1-800-777-0051. Follow-up telephone calls will be placed to customers who do not provide a written reply via the Customer Reply Form. U.S. Customers please contact Abbott Laboratories Customer Service at 1-877-422-2688 for questions regarding this information. |  | 30 | 30 kits | Nationwide Distribution including Arizona, California, Florida, Illinois, Kansas, Louisiana, Massachusetts, Missouri, Montana, New York, Ohio, Pennsylvania, South Carolina, Texas, Utah and Virginia. | Terminated | March 18, 2010 | 559 | 30 | Not_Computer
Z-1539-2008 | 46960 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | 3DKNEE Tibial Insert Trial, Size 2, 13mm Right; Part #801-01-640; manufactured by Encore Medical, L.P., Austin, TX 78758. |  |  | February 20, 2008 | 2008 | Encore Medical, Lp | Tibial insert trial incorrectly marked for thickness; device thickness marked correctly as 13mm on the underside of the trial, but is incorrectly marked as 15mm on the side. | Error in labeling | Firm notified consignee by phone and asked them to return any outstanding product on 02/20/08. |  | 6 | 6 units. | Distribution --- country of Germany. | Terminated | September 02, 2008 | 195 | 6 | Not_Computer
Z-0285-2008 | 45495 | No | KPQ | Radiation therapy simulation system. | Radiology | Radiology | 510(k) | Si emans Evansi, with patient table, Part No: 44 81 586, and without patient table Part No: 44-81-669; Siemens Medical Solutions USA, Concord, CA 94520 |  |  | February 27, 2007 | 2007 | Siemens Medical Solutions USA, Inc | Loose Mounting Screws -- Cassette holder may become loose and result in the device dislodging and colliding with the patient. | Labeling design | Customers were sent a Customer Safety Advisory Notice on February 27, 2007, notifying them that the mounting screws on the cassette holders on the SIMVIEW 3000 and NT Simulators, as well as some MEVASIM Simulators, could loosen and can cause damage to the system, or personal injury. The firm requested that they check of tightness on the mounting hardware. A STOP Modification notice was issued on 06/01/2007, to all affected regional service organizations that are responsible for implementing updates due to new information. Updated letters and User instructions will be released in August 2007. |  | 287 | 287 units total, all recalls Z-0281-285-2008 | Worldwide Distribution - USA including states of IN, VA, MA, VT, NJ, NY, PA, DC, NC, SC, GA, AL, TN, MI, KY, OH, IN MI, MN, ND, MT, IL, MO, KS, NE, AR, OK, TX, AZ, CA, SD, LA, Puerto Rico, and countries of Argentina, Australia, Austria, Bangladesh, Belarus, Belgium, Brazil, Bulgaria, Canada, Colombia, Costa Rica, Croatia, Denmark, Ecuador, Egypt, El Salvador, Estonia, France, Germany, Greece, Hungary, India, Iran, Ireland, Italy, Kuwait, Lithuania, Malaysia, Mexico, Morocco, Netherlands, Norway, China, Israel/Palestine, Philippines, Poland, Portugal, Qatar, Republic of Korea, Russia, Singapore, South Africa, Spain, Thailand, Turkey, Ukraine, UK, and Vietnam. | Terminated | February 03, 2011 | 1437 | 287 | Not_Computer
Z-0215-2009 | 49371 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | VACUETTEÂż, 2 ml Lithium Heparin Venous Blood Collection Tubes. Reference #454237, 13x75 green cap-white ring, 24 racks of 50 pcs (1200 pcs in total, non ridged, Sterile). Manufactured by: Greiner bio-one NA, Inc., 4238 Capital Drive, Monroe, NC 28110. |  |  | July 29, 2008 | 2008 | Greiner Bio-One North America, Inc. | Labeling Error: Incorrect tube label reads "Z Serum Clot Activator" instead of "LH Lithium Heparin" tube.All other information on the tube label including color coding of tube was correct. Case and rack label information showed the correct information with 2 ml Lithium Heparin. | Labeling mix-ups | Consignees were notified by letter on 07/29/2008. Notification letters to distributors instructed them to stop distribution of the product and destroy all affected lots of the product. In addition, notification letters addressed to users instructed them to destroy and discard all products from the item and affected lots. The "Product Disposition Form" must be completed by the consignee and returned to Greiner Bio-One North America, Inc.For questions or additional information, contact Greiner Bio-One North America, Inc. at 704-261-7800. |  | 33 | 33 cases (39,600 pieces) | Worldwide including USA (states of AK, AL, AR, CA, CO, DE, GA, ID, IL, IN, KS, LA, MA, MI, MO, MS, NC, ND, NE, NH, NY, OH, OK, OR, PA, RI, SC, SD, TX, VA, VT, WA, and WI) and country of Canada. | Terminated | August 26, 2009 | 393 | 33 | Not_Computer
Z-1376-2008 | 46708 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Biomet Vanguard DCM CR Tibial Bearing, 14MM x 79/83MM, CR STD bearing, sterile, Biomet Orthopedics, Inc., Warsaw, IN; Part 183464. |  |  | February 12, 2008 | 2008 | Biomet, Inc. | Mislabeled: The packaging of the two products was mixed up, resulting in the products being labeled with an incorrect size. If implanted, the product may not fit properly, which could require repeat surgery. The problem could also result in a delay in the procedure. | Packaging process control | Distributors were notified via recall letter (Urgent Medical Device Removal Notice) dated 2/12/08 to locate the devices in their possession and at hospital sites, discontinue use, and return them to Biomet. Hospitals having received the devices will be provided with a recall letter dated 2/12/08 which describes the nature of the problem and the removal of these instruments. A third letter dated 2/12/08 was provided for implanting physicians. |  | 8 | 8 | Worldwide: USA and Australia | Terminated | December 22, 2009 | 679 | 12 | Not_Computer
Z-0306-2009 | 49397 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | CELL-DYN Sapphire Hematology Analyzer, List numbers 08H00-01, 08H00-03 |  |  | August 25, 2008 | 2008 | Abbott Laboratories | Exposed wire resulted in minor electrical shock and burn to the service technician | Device Design | Recall issued August 22, 2008. Technical service bulletin was issued for all field personnel via internal database. No customer communication is planned as customers are instructed *not* to remove the cover, which is necessary to access the area. |  | 958 | 958 units | Worldwide Distribution --- including USA and countries of Argentina, Australia, Brazil, Canada, Chile, China WFOTC, Colombia, Germany, Hong Kong, Japan, Mexico, New Zealand, Panama, Paraguay, Peru, Singapore, South Korea, Taiwan and Venezuela. | Terminated | January 06, 2011 | 864 | 958 | Computer
Z-0286-2009 | 49498 | No | FMQ | Protective restraint. | General Hospital | General Hospital | 510(K) Exempt | Posey Synthetic Leather and Biothane Waist and Wrist Restraints, Catalog Numbers: 2217SL and 2217B |  |  | July 25, 2008 | 2008 | J T Posey Company | Failure to restrain: This recall was initiated after a complaint was received by the firm regarding a Posey 2217SL unit indicating that a patient was able to pull the rivet or washer through the synthetic leather allowing the patient to have freedom of arm movement. A second complaint was received indicating that a patient was able to tear the strap at the adjustment notch on a Posey 221 7SL unit | Device Design | On July 29, 2008 Posey sent a Recall Letter to all affected customers. The recall letters were sent by certified mail with return receipt signature. The recall letter instructed the customers not to distribute or use the affected products. Any existing inventory subject to this recall was requested to be returned to Posey Company. Also, if they were a distributor, they were asked to please additionally forward a copy of the recall notice and reply form to each of their customers to whom you have further distributed the subject product. A Recall Fax Back form was provided with the recall letter for the affected facilities to fax back to inform Posey of their existing inventory.An RMA number was issured to each facility and a return UPS label was provided for any product that will be returned to Posey. Additionally, the sales representatives for each of the customers will be contacted by Posey Customer Service. |  | 145 | 145 units | Worldwide, USA, Australia, Canada, France &amp; The Netherlands. | Terminated | March 30, 2012 | 1344 | 145 | Computer
Z-0100-2009 | 49402 | No | DQA | Oximeter. | Cardiovascular | Anesthesiology | 510(k) | M3811B Philips Telemonitoring Clinical Review Software part of Philips Telemonitoring System, M3810A, software revision identified as Build 1.1.2.11d, B.02.07. Prescriptive medical device used to automatically collect and transmit medical information (weight, blood pressure, non-diagnostic ECG) over phone lines between provider and patient. |  |  | September 11, 2008 | 2008 | Philips Medical Systems | Multiprint report may contain incorrect vital data for patients | Software design | Philips contacted customers by sending an Urgent Medical Device Correction Notification on 9/11/08. Letters to customers sent via certified mail and receipt of delivery will be tracked.Customers are asked to follow the Action To Be Taken By Customer/User section of the Urgent Medical Device Correction Notification. This section instructs users to allow the upgrade of the software upon logging in to the system. Contact Philips at 1-866-246-7316 for assistance. |  | 52 | 52 units | Nationwide | Completed |  |  | 52 | Computer
Z-0331-2009 | 50096 | No | FCG | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Boston Scientific Easy Core" Biopsy System, 18 ga. x 15 cm, REF 43-453 , M001434531, Sterilized with ethylene oxide gas. |  |  | October 07, 2008 | 2008 | Boston Scientific Corporation | Boston Scientific Corporation initiated a recall of its easy Core" BIOPSY SYSTEM due to difficulty cocking or arming the cannula latch on the device. This difficulty may result in an inability to use the device. | Device Design | Consignees were sent a "Boston Scientific Urgent Medical Device Recall" letter dated October 21, 2008. The letter was addressed to The Hospital Field Action Contact. The letter described the problem and products being recalled. The letter stated that it was very important that the attached product table be used carefully and consider both inner and outer packaging product codes when searching for affected/recalled product as the UPN numbers on the inner and outer labeling are different. Requested to discontinue use of product, identify, and segregate and return affected products to BSC. Letter also requested to complete and return the Reply Verification Tracking Form. |  | 3808 | 3808 total for all products | Nationwide Distribution --- including states of AL, AZ, AR, CA, CT, FL, GA, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO,MT, NE , NV, NH, NH, NM, NY, NC, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WV, WI, and PR. | Terminated | December 10, 2011 | 1159 | 3808 | Computer
Z-0191-2009 | 49497 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker Osteosynthesis Reduction Spoon; |  |  | August 14, 2008 | 2008 | Stryker Howmedica Osteonics Corp. | Stryker has become aware of the potential for breakage of the Reduction Spoon designed prior to 2003. | Device Design | Recall notification letters were sent on August 14, 2008 to all branches and hospitals via Federal Express. Questions are to be directed to 201-831-5718 or a Trauma team member listed in the letter. |  | 1070 | 1070 units | Nationwide | Terminated | May 29, 2012 | 1384 | 1070 | Computer
Z-0324-2009 | 50085 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo MultiModality WorkPlace (MM WP). Software product(Model Number 10140720). |  |  | September 12, 2008 | 2008 | Siemens Medical Solutions USA, Inc | Values derived from Dynamic CT data sets may be incorrect. The corresponding images are not affected. | Software design | The recalling firm has issued a software update to affected customers via Update Instructions CT034/08/S. This software update corrects the values from the volume perfusion post processing application. A firm representative will visit each consignee to install the software update beginning 9/12/08. |  | 130 | 130 units | Nationwide Distribution. | Terminated | September 23, 2009 | 376 | 130 | Computer
Z-0522-2009 | 49065 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Philips Diagnostic Ultrasound System, Model HD3. Product labeled HD3 Ultrasound System, Distributed by Philips Ultrasound, 22100 Bothell-Everett Highway, Bothell, WA 98021. |  |  | August 12, 2008 | 2008 | Philips Medical Systems | The system has a software defect that affects the display of Estimated Fetal Weight (EFW) Growth Percentile result values. | Software design | On 8/12/08, the firm sent an Urgent - Medical Device Correction letter, dated July 2008, to all affected customers via UPS. The letter describes the reason for the recall, informs them that Philips Ultrasound is correcting the error, states that the software is under development, will be available in September 2008 and will contact the customer to upgrade their system. "In the interim, Philips recommends that you only use the Hadlock or Osaka growth tables to perform your EFW Growth Percentile calculations. You should make the appropriate selections in the Analysis Setup menu on your system." |  | 4279 | 4279 units | Worldwide Distribution --- including USA and countries of Argentina, Australia, Austria, Bahamas, Bahrain, Belgium, Bosnia, Brazil, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Dominican Rep, Ecuador, Egypt, El Salvador, France, Germany, Greece, Guam, Guatemala, Hong Kong, Hungary, India, Indonesia, Israel, Italy, Jordan, Kuwait, Malaysia, Martinique, Mexico, Moldova, Morocco, Netherlands, New Zealand, Peru, Philippines, Poland, Portugal, Puerto Rico, Romania, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Syrian Arab Rep, Taiwan, Thailand, Tunisia, Turkey, Ukraine, United Kingdom, Utd Arab Emirr., Venezuela, Viet Nam, Canada, Afghanistan, Albania, Argentina, Australia, Austria, Belgium, Bolivia, Bosnia, Brazil, Bulgaria, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican Rep, Ecuador, El Salvador, Estonia, Finland, France, Germany, Greece, Guatemala, Honduras, Hong Kong, Hungary, India, Indonesia, Iraq, Italy, Japan, Korea Rep, Latvia, Lebanon, Lithuania, Malaysia, Mauritius, Mexico, Moldova, Morocco, Namibia, Netherlands, New Zealand, Nicaragua, Norway, Oman, Panama, Paraguay, Peru, Poland, Portugal, Puerto Rico, Romania, Serbia, Singapore, SLOVAKIA, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, Uganda, Ukraine, United Kingdom, Uzbekistan, Venezuela, and Virgin Isl (Br). | Terminated | July 14, 2011 | 1066 | 6181 | Computer
Z-0188-2009 | 49448 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Maquet Alphastar 1132.01AX/BX, 1132.02AX/BX, and 1132.03AX/BX Universal Operating Table |  |  | September 16, 2008 | 2008 | MAQUET Inc. | Table tilts: The Alphastar Surgical Table piston rod head may break and cause a full inclination of the table top during Trendelenberg positioning if there is insufficient lubrication. | Equipment maintenance | Maquet sent Recall Notification letters to all customers on September 16, 2008 by Federal Express. Letters inform customers of the potential problem with the Alphastar tables, a description of the failure and hazards involved. Hospitals are asked to confirm whether yearly maintenance by an authorized technician have been performed. If preventive maintenance has not been performed, customers are requested to contact their Maquet Service rep. A Customer Service Response form is included and customers are to complete if to acknowledge whether of not the facility is still in possession of an affected table. If so, the model number, serial number and facility name &amp; address is to be completed so that Maquet reps can visit the facility and perform the corrective action. |  | 141 | 141 in US | Nationwide | Terminated | January 12, 2010 | 483 | 141 | Computer
Z-0195-2009 | 49546 | No | EWO | Audiometer. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(K) Exempt | GSI Audera (version 2.6 software) Viasys Healthcare, Inc. |  |  | September 16, 2008 | 2008 | Cardinal Health, Inc | GSI Audera systems requires a system software update due to issues with the Split-Screen and the Vestibular Evoked Myogenic Potential (VEMP) functions. A. Split Screen Field Correction:GSI Audera units with version 2.6 software may mislabel Auditory Evoked Potential (AEP) waveform responses with respect to the identification of the stimulus ear. When using 2.6 software in split-screen mod | Software design | Consignees were sent a Cardinal Health "Urgent: Medical Device Correction" letter dated September 16, 2008 on 9/22/08 and 9/26/08. The letter was addressed to the Customer/Distributor. The letter instructed consignees to perform installations if they have the following software: 1) For Version 2.6 software, you must install Version 2.6.5 software to correct a Split- Screen performance issue and to remove the VEMP function. 2) For Version 2.0 or 2.1 software, you must install the Protocol Update Software to remove the VEMP function. Versions 2.0 and 2.1 do not contain Split-Screen capability. "Attachment A-Verification Form" must be completed by the recipient/distributor.If you need assistance in making the software revision determination or have further questions, please contact the Cardinal Health (formerly VIASYS) Neurocare/GSI Technical Support Group at 1-800-356-0007, 1-800-700-2282, 608-273-5000, or by e-mail at gsiservice@viasyshc.com or gsiservice@cardinalhealth.com. |  | 297 | 297 | Nationwide including states of AL AZ CA CO CT DC FL GA HI IA IA IL IN KS LA MA MD ME MN MO MS NC NC NH NJ NM NY OH OK PA RI SC TN TX UT VA WA and WI. | Terminated | May 01, 2012 | 1323 | 297 | Computer
Z-0196-2009 | 49549 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | 26" Vision Elect HDTV Surgical Viewing Monitor, model number 0240030960, Manufactured by Stryker Endoscopy San Jose |  |  | September 22, 2008 | 2008 | Stryker Endoscopy | Knob parts stick together, resulting in poor functionality of the knob. | Device Design | Urgent: Device Correction letters dated 9/22/2008 were sent to Sales representatives and all accounts. Sales reps were also notified via email. Each Sales Rep was provided rework instructions, a list of affected accounts and serial numbers, and new knobs to install on affected units. Customers were asked to inspect their inventory for the part number, they were informed that their sales rep would be responsible for replacing the rotary knobs on all affected units, and that they should contact their sales rep to schedule the knob replacement. |  | 2265 | 2265 devices | Worldwide Distribution --- USA including states of Alabama, Arkansas, Arizona, California, Colorado, Connecticut, DC, Delaware, Florida, Georgia, Hawaii, Iowa, Idaho, Illinois, Indiana, Kansas, Kentucky, Louisiana, Massachusetts, Maryland, Michigan, Minnesota, Missouri, Montana, North Carolina, North Dakota, Nebraska, New Jersey, New Mexico, Nevada, New York, Ohio, Oklahoma, Oregon, Pennsylvania, Rhode Island, South Carolina, South Dakota, Tennessee, Texas, Utah, Virginia, Vermont, Washington, Wisconsin, West Virginia, and Wyoming, and countries of Australia, Canada, China, England, France, Germany, Greece, Iberia, India, Israel, Japan, Korea, Mexico, Netherlands, Poland, Portugal, Scandinavia, Singapore, South Africa, Switzerland, and Taiwan. | Terminated | December 21, 2010 | 820 | 2265 | Computer
Z-0181-2009 | 49464 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Biomet brand Modular Microplasty Cup Inserter, 3/8" thread, Biomet Orthopedics, Inc., Warsaw, IN; Model 31-400600. Non-powered, hand-held instrument intended for use as an impactor during orthopedic surgery. |  |  | July 28, 2008 | 2008 | Biomet, Inc. | The pin and clip may fracture during use. | Device Design | Consignees were notified by recall letter dated 7/28/08 to remove the recalled lots from use and to return them. Contact Biomet, Inc. at 1-800-348-9500 or 1- 574-372-3983 for assistance. |  | 85 | 85 | Nationwide, Australia, Belgium, Canada, Denmark, New Zealand and Sweden. | Terminated | September 22, 2009 | 421 | 85 | Computer
Z-0183-2009 | 49540 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Bard Vacora Biopsy Vacuum Assisted Biopsy Probe; Catalog numbers VB10118 and VB10140. |  |  | September 12, 2008 | 2008 | Bard Peripheral Vascular Inc | Some of the Vacora Biopsy Vacuum Assisted Biopsy Probe thumb wheels from these lots may fracture when fired using the prime/pierce option. | Device Design | Product Recall Notification Letter sent dated September 11, 2008 Attention: Risk Managers "All VB10118 and VB10140 catalog numbers with expiry dates prior to September 2008 are now beyond their expiration date. In accordance with accepted medical practice, they should have been discarded and not used. Do not use or further distribute any affected products. Once the product affected by the recall has been removed from your inventory; fill out the recall &amp; effectiveness check form, call Silvia Carrillo at 1-800-321-4254 ext 2727 to issue your with either a Return Authorization Number or Consignment Recall Number. fax the form to BVP at 1-800-440-5376 |  | 225 | 225 | Nationwide Distribution --- including states of AL, AR, CA, FL, KY, MA, NY, OH, SC, TN, TX, VA, and WA. | Terminated | September 19, 2012 | 1468 | 225 | Computer
Z-0122-2009 | 49584 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Axiom Artis dTA, Axiom Artis Modular Angiography X-Ray System, Model number 7008605. |  |  | September 02, 2008 | 2008 | Siemens Medical Solutions USA, Inc | Incorrect contour finding: Image may be calibrated to the wrong size catheter. | Software design | Siemens issued a Customer Safety Advisory Notice dated 8/29/08 to inform customers of the potential issue. The recalling firm issued Update Instructions AX040/08/S to the service engineers which provides information to fix the problem. |  | 133 | 133 units | Nationwide. | Terminated | October 05, 2009 | 398 | 577 | Computer
Z-0182-2009 | 49514 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Circlip component of the Stryker Flat Panel and Navigation Arm System. Stryker Flat Panel and Navigation Arm System are intended for use as a ceiling-mounted device to support a medical grade flat panel monitor or camera in the patient area. The Circlip within the flat panel and navigation arms is part of the mechanical stop designed to prevent the arm from rotating infinitely. It also prevents the cables routed through the arm from becoming twisted during routine movement and rotation of the arm. |  |  | September 05, 2008 | 2008 | Stryker Communications Corp | Circlip component used to suspend flat panel and navigation arm system may become dislodged. | Device Design | Notification letter was sent to consignees September 5, 2008. Field service will be performed at all affected sites. Old Circlips will be replaced with new clamps that do not allow the arm to fall. Contact Stryker Communications Corp. at 1-972-410-7310 for assistance. |  | 705 | 705 units | Nationwide and International Distribution | Terminated | October 02, 2009 | 392 | 705 | Computer
Z-0213-2009 | 49440 | Yes | MSX | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, Mobile Care Server used in conjunction with Web Viewer, Pocket Viewer, and Cellular Viewer WIT with L-WEB05 software. |  |  | July 16, 2008 | 2008 | GE Healthcare | Issue 1 -Possible failure of CIC Pro audible alarms: Requests for telemetry defaults from Mobile Care Server could cause excessive CIC Pro memory consumption. This excessive memory consumption may lead to potential loss of audio at the CIC or a system reset or reboot. Issue 2 - Mobile Care software/Mobile Viewers don not recognize ApexPro FH Telemetry transmitter: Mobile Care server does not ac | Software design | Consignees contacted via a GE Healthcare Urgent Medical device Correction letter dated 8/26/08 and supplemental instruction for installation of software" Mobile care server with software, L-WEB05 and iCentral with software, L-NET05 supplemental Instruction for installation". The letter was addressed to Healthcare Administrator/Risk manger, chief of Nursing and director of Biomedical engineering / IT. The letter also described the problem, affected Products,and Short Term Recommendations . |  | 292 | 292 | Worldwide Distribution --- USA including states of CALIFORNIA, IOWA, INDIANA, VERMONT, and WASHINGTON and countries of AUSTRALIA, AUSTRIA, BELGIUM, CANADA, CHINA, CROATIA, CYPRUS, DENMARK, FINLAND, FRANCE, GERMANY, GREECE, INDIA, INDONESIA, ITALY, JAPAN, REPUBLIC OF KOREA, KUWAIT, MALAYSIA, MEXICO, NETHERLANDS, NEW ZEALAND, NORWAY, QATAR, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, THAILAND, TURKEY, and UNITED KINGDOM. | Terminated | November 24, 2015 | 2687 | 292 | Computer
Z-0101-2009 | 49417 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | AXIOM Luminos TF, model number 10093902 |  |  | August 12, 2008 | 2008 | Siemens Medical Solutions USA, Inc | Liquids may enter the system and cause potential malfunction and possible hazard to patients, user or other persons. | Device Design | The recalling firm was made aware of a potential problem with the device from the manufacturer in Germany. The recalling firm sent a copy of the Customer Safety Advisory Notice, dated June 18, 2008, received from Germany to their customers. The recalling firm has representatives visiting each location to modify the units. |  | 211 | 211 units | Nationwide Distribution | Terminated | October 05, 2009 | 419 | 211 | Computer
Z-0456-2009 | 49460 | No | HGM | Perinatal monitoring system and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | GE Healthcare Centricity Perinatal (formerly Quantitative Sentinel) System - Alert and Reminder software; automatic patient data management providing clinical information at the bedside in Labor &amp; Delivery, Mother-Baby and the Neonatal Intensive Care Unit; GE Healthcare Integrated IT Solutions, Barrington, IL 60010 |  |  | September 25, 2008 | 2008 | GE Healthcare Integrated IT Solutions | Software anomalies in the Alert and Reminder feature could result in a delay of treatment. When attempting to select the last visible alert or reminder choice, the next choice on the list below the desired choice is selected, and an inconsistent color may be displayed for the same clinical element across a set of work stations. | Software design | The software was updated to version 6.80.1 on 4/22/08, however, the consignee was notified of the software defects relating to the last visible Alerts and Reminders choice and the inconsistent display for the same clinical element across a set of work stations, via Product Safety Notification letter dated 8/29/08, sent via first class mail on 9/24/08. If there were any questions, the customer was to contact Tech Support at 1-800-433-2009. |  | 1 | 1 unit | Nationwide Distribution --- state of Ohio. | Terminated | December 03, 2008 | 69 | 1 | Computer
Z-0164-2009 | 48661 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian RV Software |  |  | June 02, 2008 | 2008 | Varian Medical Systems Oncology Systems | A malfunction within the Varis software might result in a misadministration (underdose). The software does not correctly interpret the number, causing field parameters having decimal values to be incorrect. | Software design | Recall initiated on June 2, 2008. A product notification letter was distributed by certified mail to all affected consignees. |  | 18 | 18 | Worldwide Distribution --- including USA and countries of Canada, Austria, France, Germany, South Korea, Sweden, and the UK. | Terminated | May 03, 2011 | 1065 | 18 | Computer
Z-0103-2009 | 47839 | No | HNO | Keratome. | Ophthalmic | Ophthalmic | 510(k) | Zyoptix XP Microkeratome maxon motor, Ref #ZXP09183. The label shows the product is manufactured by Bausch &amp; Lomb Incorporated, Rochester, NY. |  |  | April 23, 2008 | 2008 | Bausch &amp; Lomb Inc | The plastic tip on the drive shaft may come off the shaft and cause the blade to stop oscillating. | Device Design | Phone calls were initiated on 4/28/08 and an Urgent Recall letter dated 4/23/08 was issued 4/28/08 via overnight mail to U.S. customers explaining the reason for recall. The letter also requested that customers stop using the motors identified and return of the motors for evaluation. A return form was enclosed. |  | 69 | 69 motors | Worldwide Distribution --- including USA and countries of Brazil, The Netherlands, Canada, India, Malaysia, Mexico, and Singapore. | Terminated | April 02, 2012 | 1440 | 113 | Computer
Z-0102-2009 | 49496 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo Imaging with versions VB20B, VB20D, VB20F, VB20G, VB25B, VB30A, VB30A_SP1, and VB30A_SP2. |  |  | August 22, 2008 | 2008 | Siemens Medical Solutions USA, Inc | Image may not be visible after merge. | Software design | The recalling firm issued Customer Safety Advisory Notice dated 8/22/08 to inform the consignees of the potential issue and provide instruction to avoid its occurrence. To avoid potential risk of this issue, users are asked to check whether the extended HL7 messages as stated in the letter are mandatory for their RIS-PACS workflow, and disable these messages if they are not. Siemens is preparing a modification under the UI IM086/08/S for the affected versions of syngo Imaging that will resolve the potential malfunction. The field modification will be immediately distributed and installed by the firm's Service as soon as it is available. |  | 59 | 59 units | Nationwide Distribution --- including states of AL, AZ, CA, CO, CT, FL, IA, KS, LA, MA, MD, MI, MN, MO, NC, NE, NJ, NY, OH, OK, OR, PA, TN, TX, and WI. | Terminated | November 19, 2008 | 89 | 59 | Computer
Z-0297-2009 | 48628 | No | KXT | Digoxin test system. | Toxicology | Toxicology | 510(k) | AxSYM Digoxin II Reagent Pack, Disk Version 1. (3D53-01 or 3D53-02) |  |  | May 22, 2008 | 2008 | Abbott Diagnostic International, Ltd. | AxSYM TDM/ Transplant Assay Disks 3D53-03 and higher do not contain the AxSYM Digoxin II Assay file.The instrument might require installation of the AxSYM Digoxin II assay file. | Software change control | A Product Information letter, dated May 22, 2008, was issued to customers. It requested that one of the following two actions be taken if the AxSYM Digoxin II assay file is not installed on your AxSYM instrument and you want to run the AxSYM Digoxin II assay. If you have one of the previously manufactured software disks listed, [AxSYM Digoxin II Assay Disk, LN 2C43-01, or higher; or AxSYM TDM/Transplant Assay Disk Version 1.0 (3D53-01) or Version 2.0 (3D53-02)], install the AxSYM Digoxin II assay file. Do not install any other assay file. If you do not have one of the assay disks listed, contact your local Customer Service representative to have this assay file installed or install the file from a previous version if you have it on hand. |  | 9069 | 9069 kits | Worldwide Distribution --- USA including states of AL, AZ, AR, CA, CT, FL, GA, IL, KS, MD, MA, MI, MN, MS, MO, NV, NJ, NY, ND, OK, OR, PA, RI, TX, VA, WA, Hawaii, and Puerto Rico, and countries of Canada, Mexico, Guatemala, Costa Rica, Venezuela, Colombia, Ecuador, Peru, Chile, Uruguay, Argentina, Puerto Rico, Germany, Japan, Singapore, Hong Kong, Taiwan, Thailand, South Korea, Australia, New Zealand, Curacao, and Panama. | Terminated | September 09, 2011 | 1205 | 9069 | Computer
Z-0305-2009 | 49097 | No | FMZ | Neonatal incubator. | General Hospital | General Hospital | 510(k) | Giraffe OmniBedÂż, labeled in part: ***GE Healthcare 8880 Gorman Road Laurel, MD 20723 U.S.A. www.gehealthcare.com***. |  |  | June 20, 2008 | 2008 | Datex Ohmeda, Inc. dba GE Healthcare | Manufacturer of neonatal incubators changed the device labeling to account for new use and care instructions. | Equipment maintenance | A notification letter (Urgent Medical Device Correction) was sent to consignees on 06/20/08. The notification letter addressed healthcare professionals, including neonatal, infection control and nursing departments at medical facilities. In add tion, the letter advised of revisions to the current Giraffe OmniBedÂżÂż cleaning instructions contained in the product Operators Manual. Instructions were enclosed with the notification letter and are intended to supplement those already present in the manual.Contact the GE Healthcare Technical Support center at 1-800-558-7044 for questions or additional information. If outside of the USA, please contact your local GE Healthcare Service Representative. |  | 7878 | 7,878 beds | Worldwide Distribution: The devices were widely distributed to national and international medical facilities. | Terminated | March 24, 2010 | 642 | 7878 | Computer
Z-0137-2009 | 49426 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Bicor X-Ray System, Model number 9023136. |  |  | August 11, 2008 | 2008 | Siemens Medical Solutions USA, Inc | Unintended movement of the system table and/or C arm. | Software design | Siemens Medical Solutions issued a letter to consignees on 8/11/2008 to inform them of the problem and a workaround should the issue occur prior to software revision. |  | 20 | 20 units | Nationwide. | Terminated | December 21, 2012 | 1593 | 975 | Computer
Z-0202-2009 | 49578 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Oxylog 3000 Emergency and Transport Ventilator Continuous Respirator with software version 1.10 pre-installed; Catalog number 2M86300. |  |  | October 03, 2008 | 2008 | Draeger Medical, Inc. | May experience an interruption of ventilation for approximately 5 seconds | Software design | The firm issued an Urgent - Medical Device Recall letter dated September 2008. Software version 1.10 may under specific conditions experience an interruption of ventilation for approximately 5 seconds. During that time period, the patient airway pressure may drop to zero. This phenomenon was identified by Draeger engineers and to date has not been reported by any user. The firm's investigations revealed that when ventilating in BIPAP mode with the pressure rise time (Ramp) set to the highest setting and the expiration time set to exactly 0.5 seconds, the inspiration phase can coincide with the internal auto zero function of the device. Under these circumstances a pressure spike will occur upon reaching Pmax +5mbar (cmH2O) which with software version 1.10 will lead to the inspiratory phase stopping. Ventilation will resume after approximately 5 seconds. The device will generate a Paw Measurement INOP alarm. Until the new software is available, customers are being instructed to avoid the use of the following combination of settings: 1) BIPAP mode; 2) Ramp = highest, and 3) Texp = 0.5 seconds. |  | 50 | 50 units | Nationwide Distribution --- including states of AZ, CA, FL, IL, IN, KY, LA, ME, MI, MO, NC, NE, NJ, NY, OH, PA, TX, UT, VT, WA, WV, and WY. | Terminated | October 16, 2009 | 378 | 50 | Computer
Z-0312-2009 | 49569 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | AXIOM Artis System dFC using software version VB31D with 30x40 cm2 flat detectors. |  |  | August 25, 2008 | 2008 | Siemens Medical Solutions USA, Inc | System may switch to emergency fluoroscopy mode. | Software design | A Customer Safety Advisory Letter dated 8/20/08 was issued with Update Instructions AX045/08/S. The letter informed the customer of the potential issue and provided instruction to avoid its occurrence. A software update (VB31E) was released to correct the issue. |  | 4 | 4 units | Nationwide Distribution including states of AR, CA, CO, FL, IA, IL, IN, KS, KY, MI, NY, TX, VA, and WY. | Terminated | January 26, 2010 | 519 | 21 | Computer
Z-0117-2009 | 49399 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Exact Couch with Exact Couch Top, a component of the C-Series Clinacs, Acuity and Ximatron treatment systems. Model numbers H14, H18, H27, H72, H77. A powered radiation therapy patient support assembly is an electrically powered adjustable couch intended to support a patient during radiation therapy. |  |  | August 11, 2008 | 2008 | Varian Medical Systems Oncology Systems | During couch movement, a patient's fingers may get pinched at certain points which could result in broken, disengaged or pinched fingers. | Device Design | Recall initiated on 08/11/2008. Notification will be sent to Varian Sales, Marketing and Service organizations as well as to all affected users. Contact Varian Medical Systems Oncology Systems, at 1-650-424-5731 for assistance. |  | 3296 | 3296 units | Worldwide including to the US and to Austria, Belgium, Czech Republic, Denmark, Finland, France, Germany, Greece, Hungary, Iceland, Ireland, Italy, Latvia, Lithuania, Luxembourg, Norway, Poland, Portugal, Romania, Slovakia, Slovenia, Spain, Sweden, Switzerland, The Netherlands, United Kingdom, Algeria, Australia, Bahrain, Bangladesh, Belorussia, Bolivia, Brazil, Canada, Chile, China, Colombia, Ecuador, Egypt, Guadeloupe, Hong Kong, India, Indonesia, Iran, Israel, Jamaica, Japan, Kazakhstan, Lebanon, Macau, Macedonia, Malaysia, Martinique, Mexico, Morocco, Nepal, New Zealand, Oman, Pakistan, Panama, Peru, Philippines, Russian Fed., Saudi Arabia, Serbia, Montenegro, Singapore, South Africa, South Korea, Sudan, Syria, Taiwan, R.O.C., Thailand, Turkey, U.A.E, Ukraine, Uruguay, Venezuela, and Vietnam. | Terminated | November 27, 2010 | 838 | 3296 | Computer
Z-0459-2009 | 50092 | No | CCW | Rigid laryngoscope. | Anesthesiology | Anesthesiology | 510(K) Exempt | Sun Med GreenLine/D Laryngoscope Blades Compatible with Fiber Optic Green System Handles, Sterile Disposable. Contents 20. |  |  | April 27, 2007 | 2007 | Sun Med, Inc. | A breakage problem of the acrylic bundles at both distal and proximal ends of the light tube for the MD Laryngoscope Blade GreenLine D Mac 3 has been detected. If breakage occurs during use, there is the potential for leaving approximately a one inch piece of acrylic tube in the patient's mouth. | Device Design | SunMed sent initial recall letter on April 27, 2007 to all distributors requesting that kits containing the MacIntosh size 3, consisting of lots number Gj and Ha be returned to facility. Replacement information was included in letter. On December 12, 2008, the firm issued a Device Recall Press Release with additional information on the potential for the tube to break during use leaving an approximate 1 inch piece of the acrylic tube in the patients mouth and subsequently requiring medical or surgical intervention to prevent serous injury to the patient. SunMed has notified its distributors and customers by certified mail and has arranged for return and replacement of all recalled products. Consumers with question may contact their distributor of the GreenLine/D or SunMed at 1-800-433-2797. |  | 2682 | 2,682 boxes (20 per box) | Nationwide Distribution | Terminated | July 06, 2010 | 1166 | 2682 | Computer
Z-0118-2009 | 49403 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | iSite PACS (Picture Archiving and Communication System), software versions 3.6.28.x and 4.1.x. |  |  | August 29, 2008 | 2008 | Philips Healthcare Informatics, Inc. | Two defects have been identified: 1) potential to display a patient on the canvas page that is different from the patient whose images are displayed on diagnostic monitors when using Conference Presentation States and when opening two or more studies from a folder or an exam worklist and 2) potential to miscalculate measurements when Pixel Spacing and Imager Pixel Spacing DICOM tag values are bot | Software design | A notification letter (Urgent Product Correction) dated 8/29/08 was sent via certified mail to all customers alerting them of the issue, requiring posting of notification. Philips Healthcare Informatics, Inc. strongly recommend all customers advise their iSite Radiology users not to open multiple patient exams simultaneously from a folder or worklist using Conference Presentation States. A software correction is being worked on, and should be available by end of Q3 2008 for version 4.1.45.2. A mandatory upgrade is required for all customers who are currently using iSite version 4.1.45.2 and former versions. For specific questions regarding these issues, contact the Customer Advocate (CARE) or Customer Care at Philips Healthcare Informatics, Inc. at 1-877-328-2808. In your communication with Customer Care at Philips Healthcare Informatics, Inc. , please refer to this notification and associated FCO 22945704-08292008-004. |  | 27 | 27 units. | Worldwide Distribution including USA (MN, CO and PA) and countries of Italy, Germany, Netherlands, Sweden, France, Portugal, Switzerland, Thailand, Japan and Australia. | Terminated | November 30, 2010 | 823 | 27 | Computer
Z-0301-2009 | 48712 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer Alumina Ceramic Femoral Head, 12/14 taper, 28 mm diameter, neck length +3.5 mm, sterile, Zimmer, Warsaw, IN; REF 00-6418-028-03 and 00-6428-028-03. |  |  | June 11, 2008 | 2008 | Zimmer Inc. | When used with cobalt/chromium hip stems, the average test values for ceramic head bursting falls below the guidance document limits if the taper is made by casting instead of by forging. | Device Design | The firm issued Product Compatibility Update # 13, dated 6/11/08, to consignees informing them which products are not compatible, and thus should not be used in combination. |  | 49474 | 49,474 of all products | Worldwide Distribution --- including USA and countries of Austria, Australia, Belgium, Brazil, Canada, Chile, China, Czech Republic, Denmark, Dubai, England, Finland, France, Germany, Greece, Hungary, India, Israel, Italy, Korea, Lebanon, Mexico, Netherlands, New Zealand, Poland, Romania, Singapore, Slovakia, Spain, Switzerland and Ukraine. | Terminated | November 10, 2009 | 517 | 49474 | Computer
Z-0496-2009 | 50091 | No | DSQ |  |  | Cardiovascular | PMA | Thoratec Heartmate II Left Ventricular Assist System (LVAS). Model/Catalog Number 1355 and Model/Catalog Number 102139. |  |  | October 24, 2008 | 2008 | Thoratec Corp | The firm has become aware that, over time, wear and fatigue of the percutaneous lead connecting the HeartMate II Left Ventricular Assist System (HM II LVAS) blood pump with external system controller may result in damage that has the potential to interrupt pump function and may require a reoperation to replace the pump. | Device Design | An initial "Urgent: Medical Device Correction" letter dated October 21, 2008 and a follow-up letter dated October 24, 2008 were sent to affected HeartMate II Left Ventricular Assist System (HM II LVAS) customers. The letters identified the problem, immediate action and preventive action to be taken by customers and recommended that the pump be replaced as soon as possible if damage to the percutaneous lead is confirmed. If customers suspect that a HM II LVAS patient may have a damaged percutaneous lead, please contact Thoratec Technical Services (800-456-1477) for assistance. A return response (Acknowledgement Form) was included with the letter for consignees to complete and fax it to Thoratec Regulatory Affairs at 925-847-8628. Direct questions to the Thoratec Regulatory Affairs Department at 925-730-4135 (Monday-Friday, 8 a.m.-5 p.m. PST). Clinicians and patients with questions may contact the firm at 800-528-2577, or if calling outside the USA, 925-847-8600 (7 days a week, 8 a.m.-5 p.m. PST). Copies of the press release may be found on the firm's website, www.thoratec.com, under Investor Relations/Press Releases. |  | 2335 | 2,335 units | Worldwide Distribution. | Terminated | January 26, 2011 | 824 | 2335 | Computer
Z-0163-2009 | 49567 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Axiom Aristos FX Multipurpose Radiography System, stationary X-Ray system, Model number: 7414803. |  |  | August 25, 2008 | 2008 | Siemens Medical Solutions USA, Inc | Tube support arm or the detecting mounting hardware may become loose at the joint with the telescope. | Equipment maintenance | Updated Instructions were issued in August 2008 (AX049/08/S) to Siemens Service Engineers to have them inspect and tighten any loose hardware on the affected systems. |  | 7 | 7 units | Nationwide distribution including states of CA, IN, KY, UT, VA and WI. | Terminated | October 20, 2008 | 56 | 7 | Computer
Z-0460-2009 | 49458 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Centricity PACS RA1000 Workstation (for diagnostic image analysis). GE Healthcare Integrated IT Solutions, Barrington, IL 60010. |  |  | September 25, 2008 | 2008 | GE Healthcare Integrated IT Solutions | Software anomalies result in patient safety issues involving patient jacket content intermittently becoming unintentionally out of synchronization with the images being displayed, and involving Default Display Protocols (DDPs) which are used to layout images when displaying a study. If the user does not check the Study Date Time of the exam on the image title bar, then they may interpret the curr | Software design | GE Healthcare Integrated IT Solutions notified their Centricity PACS RA1000 Workstation customers via letter (Product Safety Notification) dated 9/25/08 of the software anomalies resulting in patient safety issues involving the patient jacket content intermittently becoming unintentionally out of synchronization with the images being displayed, and involving Default Display Protocols (DDPs) which are used to layout images when displaying a study. If the user does not check the Study Date Time of the exam on the image title bar, then they may interpret the current exam as a historical study and vice versa. The letter provided the users with safety instructions to follow until the software is updated. Direct any questions to the Remote Online Center, Field Engineer or the GE Customer Care Center at 800-437-1171. The accounts were requested to complete the enclosed customer reply form, supplying the workstation software or build number, and fax the completed form to 847-277-5240. |  | 546 | 546 units | Worldwide Distribution including USA and countries of Australia, Belgium, Brazil, Canada, Cayman Islands, Chile, Denmark, Ecuador, France, Germany, Hong Kong, Italy, Japan, Korea, Malaysia, Malta, Mexico, New Zealand, Singapore, Spain, Switzerland, Taiwan, United Kingdom and Venezuela. | Terminated | September 06, 2010 | 711 | 546 | Computer
Z-0449-2009 | 49712 | No | KGI | Bone densitometer. | Radiology | Radiology | 510(k) | Hologic, Inc., 10 Year Fracture Risk Questionnaire Option for QDR X-Ray Bone Densitometers Software |  |  | September 24, 2008 | 2008 | Hologic, Inc. | Software error may lead to a high estimate of major fracture probability. | Software design | Hologic notified customers by telephone on 9/23 and 24/08 and instructed them to not use the software. A customer notification letter was sent via FedEx to customers on September 25, 2008.The letter requested de-installation of the software and return of the software to Hologic in a pre-paid envelope. |  | 15 | 15 software discs | Nationwide distribution in: CO, MA, MN, NM, NY, NC, OR, OH, PA, and TX. Also distributed in the UK. | Terminated | March 07, 2016 | 2721 | 15 | Computer
Z-0467-2009 | 49495 | No | KZD | Pressure infusor for an I.V. bag. | General Hospital | General Hospital | 510(K) Exempt | CASMED 500cc Unifusor Classic with piston gauge and stopcock valve |  |  | September 20, 2008 | 2008 | CAS Medical Systems, Inc. | Leaks: Infusion Cuffs may exhibit air leakage at the union of the tube and bag, and not maintain pressure. For further information, please contact the firm at 203-488-6056. | Equipment maintenance | CAS Medical Systems notified accounts by letter dated 9/19/08 requesting return and sub-recall be conducted to their accounts. The letter was entitled "UNIFUSOR(R) INFUSION CUFFS AND PRIVATE LABEL (CASMED-MANUFACTURED) INFUSION CUFFS *** URGENT PRODUCT CORRECTION NOTICE." For further information, please contact the firm at 203-488-6056. |  | 6384 | 6,384 pc | NationwideForeign: Malaysia, Austria, Belgium, Czech Rep., France, Germany, Greece, Italy, Norway, Spain, Sweden, Switzerland and the UK. | Terminated | August 05, 2010 | 684 | 80864 | Computer
Z-0092-2009 | 48432 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | PRIMEVIEW 3i System equipped with a Coherence Therapist R2.0 or R2.1, and PRIMEVIEW 3iR2.0 or R2.1, part number 7341428 |  |  | November 06, 2007 | 2007 | Siemens Medical Solutions USA, Inc | Flat panel positioning calibration could be off by as much as 4 mm without the machine discovering detail. | Software design | Service calibration inspection/correction Update instructions TH038/07/S was issued on 11/06/2007 to all customers. The instructions identified the affected products, stated the reason for the update, and listed the steps to update the software. Customer notification was also issued. Software patch planned for release in June 2008. |  | 13 | 13 units | Worldwide Distribution | Terminated | December 18, 2010 | 1138 | 251 | Computer
Z-0112-2009 | 49459 | No | HGM | Perinatal monitoring system and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Centricity Perinatal (formerly Quantitative Sentinel) System - Fluid Total Precision software; automatic patient data management providing clinical information at the bedside in Labor &amp; Delivery, Mother-Baby and the Neonatal Intensive Care Unit; GE Healthcare Integrated IT Solutions, Barrington, IL 60010 |  |  | September 25, 2008 | 2008 | GE Healthcare Integrated IT Solutions | On the I&O; chart, the IN, OUT and NET fluid totals values will not honor the numeric precision configuration, always displaying a whole number, losing decimal point accuracy. | Software design | GE Healthcare notified their Centricity Perinatal customers via letter dated 4/20/07 of the problem with the Fluid I &amp; O Module when fluid precision is required. The letter instructed those accounts that require fluid precision to contact GE Technical Support at 1-800-433-2009 to obtain a software patch to fix the problem. A follow-up letter dated 8/29/08 was sent on the same date to the accounts informing them that GE had re-classified the issue as a class II safety issue and that all customers need to have their software updated. The accounts were requested to contact GE Tech Support at 800-433-2009 to schedule a time to upgrade their system to 6.70.6 P02 as soon as possible. The accounts were requested to complete the enclosed customer reply form, supplying the workstation software or build number, and fax it back to 847-277-5240. |  | 470 | 470 units | Worldwide Distribution --- including USA and country of Canada. | Terminated | May 17, 2010 | 599 | 470 | Computer
Z-0214-2009 | 49136 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Sunquest Laboratory System. Software Version: Sunquest Laboratory v5.4.2, v6.1, v6.2 and v6.3. |  |  | August 01, 2008 | 2008 | Sunquest Information Systems, Inc. | While performing in-house testing for the Sunquest Laboratory System and under certain conditions, there have been instances in which an error that occurred at the database level or originated in the operating system (Cache)was not communicated or displayed in the Graphic UserInterface (GUI) application so the activity continued. While this error was registered in the error log in the SystemU | Software design | A notice (PSN-08-07) was issued to all customers on 08/01/2008 via e-mail. The notice advises customers of the problem and actions Sunquest Information Systems has taken to eliminate the Problem. Sunquest Client Support will contact their facility to schedule correcting the error trap setting. The notice also provides temporary procedural instructions to include monitoring error log using System Utilities (Function UTL) and suboption 2, System Error/Message Inquiry. Contact Client Support at 520-570-2252 for questions or additional information. |  | 139 | 139 | Worldwide Distribution including USA and countries of Canada, Denmark, Saudi Arabia and United Arab Emirates. | Terminated | November 28, 2011 | 1214 | 139 | Computer
Z-0454-2009 | 48660 | No | JQW | Pipetting and diluting system for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Vidiera NsP Nucleic Sample Preparation, Part Number: A22421, Version 1.0.41 |  |  | August 30, 2007 | 2007 | Beckman Coulter Inc | Incorrect results: under certain circumstances, when transfer volume is not within the set range, the Vidiera NsP software does not report sample exclusion and does not flag the excluded sample in the "Run Results" report. The software error associated with the excluded sample appears as "Liquid was not successfully transferred to well#..........the sample will be excluded from further proce | Software design | The customer notifications were initiated by Beckman Coulter on September 11, 2007 when all affected accounts were contacted via telephone, followed by a Product Corrective Action (PCA) was mailed via US Postal Service to all affected customers. The PCA informed the customers of the issue and instructed them to:(1) Review the "run Summary" event log for your samples. (2) Review the attached instructions describing the steps for identifying the samples affected by this error. (3) Complete and return the enclosed response form within 10 days so that we may maintain our records.The letter informed the customers that Beckman Coulter would provide a software update within 60 days. The letter also asked their customers to share the PCA with their staff and provide a copy of the PCA to any customers that they may have forwarded the affected products to.For questions regarding the PCA, please contact firm at 714-961-3634. |  | 6 | 6 | Nationwide distribution including states of PA, MI, UT, WA, IL, and CA. | Terminated | January 13, 2009 | 502 | 6 | Computer
Z-0617-2009 | 50220 | No | CAW | Portable oxygen generator. | Anesthesiology | Anesthesiology | 510(k) | DeVilbiss Model 515A Oxygen Concentrator |  |  | October 30, 2008 | 2008 | Sunrise Medical | potential for sieve beds to have been incorrectly assembled | Device Design | On 10/30/08, the recalling firm began to telephone customers of the problem. The recalling firm followed up with Urgent: Medical Device Recall letters, dated 11/7/2008. The letters requested that users cease the use of the affected device and immediately contact their provider to make arrangements to have the Concentrator replaced. The firm has set up a toll free line to assist customers with identification of affected units and the return process. |  | 428 | 428 units | Worldwide Distribution --- USA including states of AL, CA, FL ,GA, IN, IL, MA, MD, MI, MS, NC, NE, NH, NM, NY, PA, SD, TN, TX, VA, and WI, and countries of Canada, Columbia, Germany, Mexico, Netherlands, Singapore, and United Kingdom. | Terminated | October 27, 2012 | 1458 | 428 | Computer
Z-0283-2009 | 49475 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | One Touch Ping Glucose Monitoring System |  |  | August 04, 2008 | 2008 | Animas Corporation | display old, inaccurate values for bolus amount delivered and amount of planned bolus totals | Software design | On 8/6/08, the recalling firm telephoned all users to inform them of problem and the need to return the system. |  | 251 | 251 units | Nationwide Distribution | Terminated | November 12, 2008 | 100 | 251 | Computer
Z-0116-2009 | 49024 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Artiste MV system, Part number 8139789, medical linear accelerator equipped with COHERENCE therapist RTT4.1. Medical Linear accelerator for radiation therapy. The device utilizes a console system for selecting various treatment sequences and manages commands for all motorized movement. |  |  | July 04, 2008 | 2008 | Siemens Medical Solutions USA, Inc | Three issues: unexpected rotation of gantry, unexpected movement of table between beams during patient setup, and single exposure images will be overwritten by double exposure images. | Software change control | A letter marked "Urgent" dated July 4, 2008 was sent to customers. The letter provided a detailed description of the issues, the conditions under which these issues may occur, as well as recommendations to avoid these issues. Customer notification TH022/08/S was released on July 11, 2008. A software fix is currently being validated. All consignees notified by certified mail. Contact Siemens Medical Solutions USA, Inc. at 1-925-246-8407 for assistance. |  | 16 | 16 | Nationwide in LA, WI, DE, PA and OH, and internationally in Netherlands, Germany, Poland, Norway, Australia, Hungary, Belgium and France | Terminated | December 22, 2010 | 901 | 16 | Computer
Z-2138-2008 | 48974 | No | MVR | Implantable clip. | General &amp; Plastic Surgery | Cardiovascular | 510(k) | GEM 2753 Microvascular Anastomotic Coupler, 2.5 mm, REF GEM2753/I, P/N 5111-00250-010, Rx only, Irradiation Sterilized |  |  | May 14, 2008 | 2008 | Synovis Micro Companies Alliance Inc | Coupler rings may slip out of the delivery toolcaused by jaws larger than specification. | Device Design | Urgent Recall Notification was sent via Fed Ex on May 15, 2008 to 432 hospitals. The letters were addressed to the attention of the Risk Manager. The recall notification instructed the hospital to immediately discontinue use, examine their inventory and, if found, to remove the product and quarantine it for return. A second notification was sent on July 21, 2008 to non-responding consignees. |  | 3812 | 3,812 units | Worldwide Distribution --- including USA and countries of Canada, Japan, UAE, Korea, South Africa, Australia, Taiwan, Norway, Netherlands, Sweden, Greece, Norway, Turkey, Portugal, Taiwan, Denmark, Italy, Germany, Japan, Spain, Austria, and Great Britain. | Terminated | April 09, 2010 | 695 | 3812 | Computer
Z-0086-2009 | 49411 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Siemens Ultrasound, Acuson/Sonovista X300 ultrasound systems, software revisions Sonovista X300 software revision 1.0.06 and 1.0.07, Acuson X300 sofware revision 1.0.08 ultrasound system with onscreen display. |  |  | August 19, 2008 | 2008 | Siemens Medical Solutions USA, Inc. | Incorrect value calculations by the device may result in inaccurate aortic stenosis estimates. | Software design | Siemens initiated a Field Correction 08/19/2008. Service personnel will conduct a site visit and distribute a Customer Safety Advisory Notice, correct the problem and install a revised software version. |  | 169 | 169 units total | Worldwide Distribution: USA, Singapore, Germany, Denmark, Romania, Sweden, Poland, Portugal, Australia, Chile, India, Spain, Thailand, Malaysia, South Africa, Hungary, Saudi Arabia, Qatar, Peru, Italy, Azerbaijan, Greenland, Austria, Venezuela, Puerto Rico, Brazil and Japan. | Terminated | April 21, 2011 | 975 | 169 | Computer
Z-0570-2008 | 45447 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | BD Biosciences MultiSET Flow Cytometry Software Versions 2.1, 2.1.1, and 2.2 with OS X operating systems, BD Biosciences, San Jose, CA 95131. |  |  | September 06, 2007 | 2007 | BD Biosciences | Inaccurate results: Software error results in inaccurate display result statistics. If the user adjusts the lymph gate or attractors in the Lab Report view without using the Manual Gate function (as described in the system manual), the statistic results will not be updated to match the adjusted gate. | Software design | Becton Dickinson sent an Important Product Information letter by mail on 09/14/2007 informing customers of the need to follow the instructions for adjusting the lymph gate or attractors. The next software revision will alleviate this limitation. |  | 1136 | 1136 units | Worldwide Distribution - USA, Belgium, Brazil, Peru, Singapore, Canada, Jamaica, Chile, Ecuador, Hong Kong-China, Thailand, Mexico, Venezuela, Barbados, Philippines, Puerto Rico, Guatemala, Japan, Uruguay, Argentina, China, Trinidad and Tobago, Honduras, Australia, New Zealand, and India. | Terminated | April 15, 2008 | 222 | 1136 | Computer
Z-0933-2008 | 46293 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Linemaster Wireless Footswitch Accessory, Model SP-9970214-005112010 and Model SP-9970214-005112232. Sound Surgical Technologies, LLC, Louisville, CO 80027. |  |  | December 28, 2007 | 2007 | Sound Surgical Technologies, LLC | Failure to shut-off: Surgical equipment, activated by a wireless foot switch with VASER system, may remain powered on when switch is no longer depressed. | Device Design | Consignees were sent an Advisory Notice on 12/28/2007. The notification describes the defect. Sound Surgical will be contacting facilities to arrange a service call to install new cables between the footswitch receiver and the VASER system. Facilities may contact the Technical Service Center at 1-888-677-8765. |  | 127 | 127 switches | Worldwide Distribution including USA states of AL, CA, CT, FL, GA, IN, KY, LA, MI, NJ, NY, OH, OK, PA, TX, and VA. Foreign Distribution to Colombia, Costa Rica and Switzerland. | Terminated | April 01, 2008 | 95 | 127 | Computer
Z-0455-2008 | 38465 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Abbott - CELL-DYN 1800 System with CELL-DYN 22 Calibrators and the CELL-DYN Calibrators and CELL-DYN 16 Controls. Part/Model Number: 07H77-01. Abbott Laboratories, Santa Clara, CA 95054-1113 |  |  | October 27, 2006 | 2006 | Abbott Laboratories | Incorrect Results: Hemoglobin results in donor samples assayed on the CELL-DYN 1800 System vary more than expected across the range of claimed operating temperatures of 18 degrees Celcius to 30 degrees Celcius. | Software design | A Product Correction letter was sent to all customers with a reply form on October 27, 2006. Product information letters are being packed with the instrument's accessory kit and with calibrators and quality control products used with the CELL-DYN 1800 System. Correction letters with reply forms have been resent to nonresponding customers. |  | 4224 | 4,224 Systems | Worldwide Distribution, USA and countries of Argentina, Bolivia, Brazil, Canada, Chile, Columbia, Costa Rica, Dominican Republic, El Salvador, Germany, Guatemala, Honduras, Hong Kong, China, Japan, Mexico, Paraguay, Puerto Rico, Singapore, South Korea and Venezuela. | Terminated | January 06, 2011 | 1532 | 4224 | Computer
Z-0186-2008 | 44911 | No | MNI | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Size 8 Pedicle Screws of the Allez Spine Laguna Pedicle Screw System, Screw Size (diameter x length) 8 x 35, Catalog Number LG-PR0835, Allez Spine, LLC., Irvine, CA |  |  | July 30, 2007 | 2007 | Allez Spine, LLC | Screw Shank Failure; separation of the screw shank portion from the pedicle head portion of the Size 8 Pedicle Screw, when torquing down the single piece locking nut during implantation. | Device Design | On July 30, 2007 Allez Spine notified all its field representatives and distributors via email of a potential for failure of the Size 8 Pedicle Screws when used in combination with the single piece locking nuts and instructed them to return all unused inventory of Size 8 Screws and the two piece Locking Nuts with saddle. This action was to be carried out to the retail level-hospital, clinic, lab, and physicians sold the product. The email also had an attachment of a list of expected frequently asked questions (FAQ) from surgeons and answers to help them respond. A second Notice to all users was sent by the recalling firm on Feb 29, 2008, to request follow-up of implanted units. |  | 317 | 317 | Nationwide: including the states of AZ, CA, CO, GA, NV, TX and UT | Terminated | April 09, 2012 | 1715 | 1578 | Computer
Z-1196-2008 | 46784 | No | IMD | Hot or cold disposable pack. | Physical Medicine | Physical Medicine | 510(K) Exempt | Icy Hot Heat Therapy , Air Activated Heat Patch, Back and Large Areas, UPC Code 0-41167-08330-7, Product Code 08330, 4.1" x 8.4" (10.5 cm x 2.1 cm) each, 1 and 3 count patches, Distributed by Chattem, Inc. P.O. Box 2219, Chattanooga, TN 37409-0219 |  |  | February 12, 2008 | 2008 | Chattem Inc | Skin irritation and burns with product usage | Device Design | The firm initiated its recall on 02/08/2008 by issuing a press release warning the public of the hazard and directing them to stop its use and either return or discard the products. The firm initiated the recall at the retail level on 02/12/2008 by notifying its employee sales force by email, containing a Urgent Medical Device Recall Letter, with instructions to immediately email a copy of the recall letter to each Chattem consignee (wholesalers and direct retail accounts) with sub recall requests. |  | 666636 | 666,636 units (total for Recalls Z-1196-1197-2008) | Worldwide - USA, Canada, and Honduras | Terminated | October 05, 2009 | 601 | 712656 | Computer
Z-0947-2008 | 45861 | Yes | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | HomeChoice PRO Automated Peritoneal Dialysis Systems; Catalog Numbers; 5C8310, 5C8310R, R5C8320, R5C8320R, T5C8300, and T5C8300R, Manufactured for Baxter Healthcare Corporation, Deerfield, IL 60015 U.S.A., Made in U.S.A |  |  | November 20, 2007 | 2007 | Baxter Healthcare Renal Div | Unrecorded Infusion; If the system is powered down or a power failure occurs during a fill cycle, the HomeChoice system may not record the last 1-4 pump strokes infused into the patient when the power is restored. This situation may create a potential unrecorded delivery of approximately 15-60 mL of fluid being infused into the patient. | Device Design | Urgent Product Recall letters dated 11/20/07 were sent via first class mail to all of the affected customers, informing them of the potential for overfill if the system is powered down or a power failure occurs during a fill cycle. A Baxter representative will contact the customer to schedule the software update which corects this anomaly. The accounts were requested to complete and return the enclosed Customer Reply Form, indicating that they have disseminated the information to their staff and home patients. Any questions were directed to Baxter Global Technical Services at 1-800-553-6898, prompt 1.Baxter sent a follow-up letter dated 3/17/08 to the consignees, accompanied by a copy of the 11/20/07 letter. The letter reiterated the information in the 11/20/07 letter, emphasizing the serious consequences of an overfill for neonates and young pediatric patients, and providing information describing the symptoms of excess fluid in the abdomen and steps to be taken if the user suspects overfill of a neonate or young pediatric patient. Baxter requested that Baxter be contacted if an overfill should occur so Baxter can perform a complete investigation into the event. The accounts were requested to complete and return the enclosed Customer Reply Form, indicating that they have disseminated the information to their staff and home patients. Any questions were directed to Baxter Global Technical Services at 1-800-553-6898, prompt 1. |  | 32229 | 32,229 units | Worldwide; USA, Australia, New Zealand, China, Hong Kong, the Philippines, Thailand, Asia, Canada, Algeria, Austria, Bahrain, Belgium, Bulgaria, Croatia, Czech Republic, Denmark, Egypt, Estonia, Finland, France, Germany, Greece, Hungary, Iceland, Iraq, Ireland, Israel, Italy, Jordan, Kuwait, Latvia, Lebanon, Lithuania, Malta, Morocco, The Netherlands, Norway, Oman, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Senegal, Serbia, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Tunisia, Turkey, United Arab Emirates, United Kingdom, Ukraine, Latin America and Japan. | Terminated | October 26, 2009 | 706 | 103490 | Computer
Z-0574-2008 | 45881 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Stryker Disposable StrykeProbe 32cm Spatula Tip, Ref: 250-070-551, Electrosurgical Suction Tip, Stryker Endoscopy, San Jose, CA 95138. |  |  | November 16, 2007 | 2007 | Stryker Endoscopy | Cannot connect: Tips have electrodes that are too large to fit monopolar cables and are unusable. | Device Design | Stryker notified their sales representatives by phone. Stryker sent all consignees a Product Advisory Notice dated November 19, 2007, and verbally notified them of the recall via telephone. A response form with return postage was sent with the notice. The letter requested the hospitals to quarantine and return the product to Stryker. |  | 48 | 8 boxes of 6 units (48 total) | Nationwide Distribution. | Terminated | March 26, 2008 | 131 | 48 | Computer
Z-0576-2008 | 45984 | No | JJY | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Roche COBAS 6000 Clinical Chemistry analyzer; Catalog Number: 04745914001, Roche Diagnostics Corporation, Indianapolis, IN 46250. |  |  | December 04, 2007 | 2007 | Roche Diagnostics Corp. | Incorrect results: CRP assays may be calculated using incorrect calibration parameters which could result in falsely high or falsely low patient results being reported. | Software design | Roche Diagnostics contacted the consignee by phone on 12/4/07 and by follow-up letter dated 12/4/07. The user was instructed to discontinue use of two CRP assays on this analyzer. The letter described a workaround and that a new revision of software, due out first quarter 2008, will fix this. |  | 1 | 1 | USA Distribution in the state of Nebraska. | Terminated | August 04, 2008 | 244 | 1 | Computer
Z-0961-2008 | 45853 | No | LGW |  |  | Neurology | PMA | Advanced Bionics Precision Implantable Pulse Generator Kit, Model Number SC-1100 with firmware version prior to Revision 3.02, Advanced Bionics Corportation, Sylmar, CA 91342 |  |  | October 05, 2007 | 2007 | Advanced Bionics Corp | Incorrect Data -- Corruption of internal memory component results in an inability for the physician to reprogram the IPG with firmware version prior to Revision 3.02. When this occurs, the IPG will report an error code of "10h0" or "00h0" through the Remote Control. Under this condition, the IPG will cease to log in some data that could be used for informational purposes. ** It should be noted t | Device Design | Advanced Bionics sent an Important Notification letter on October 5, 2007, via FED-Ex or US Postal Service Certified mail. (International notification activity is in progress.) The recall letter advises that the firm has become aware of an issue with Precision Implantable Pulse Generators (IPG model numbers SC-1100, Sc1110, and ONS-1110) manufactured prior to July 2007. The letter describes the issue in detail. The firm is requesting that customer immediately locate the affected IPGs and set them aside so that they are not used. A Boston Scientific/Advanced Bionics sales representative will contact your facility to confirm these IPGs and arrange for the exchange of these devices.The firm enclosed Reply Verification Tracking Form listing the IPGs that are at your Clinic according to the firm's records. In addition, the firm recommends that if error code 10h0 or 00h0 is encountered in Implanted IPGs, it may be corrected by upgrading to a newer version of IPG firmware (software that runs in the IPG). The new version of firmware corrects this condition and helps prevent it from occurring in the future. This firmware upgrade is performed non-invasively, through wireless RF telemetry.If any of patients encounter error code "10h0" or "00h0" on their Remote Control, consignees are advised to contact their Advanced Bionics representative or Technical Services team at (866) 566-8913 to help identify whether the condition can be corrected through a firmware upgrade and, if so, they will help arrange for the upgrade. Please note that some devices may not be upgradeable. The firm request that all consignees return the notification acknowledgement form is included with the letter. |  | 12000 | Over 12,000 units | Class II Recall - Worldwide Distribution --- including USA and countries of EU, Canada, Australia and Israel. | Terminated | February 29, 2012 | 1608 | 12000 | Computer
Z-1462-2008 | 46740 | No | DKZ | Amphetamine test system. | Toxicology | Toxicology | 510(k) | AxSYM Drugs of Abuse/Toxicology Assay Disk Version 8.0; List Number: 3D54-08; Abbott La boratories, Diagnostics Division, Abbott Park, IL 60064 |  |  | January 21, 2008 | 2008 | Abbott Laboratories | Error Message: If the Positive and Negative Interpretation Cutoff parameters (116/117) are edited after installation of the revised assay file, VRTX error #0002 in Task 40 is generated and the instrument locks up when it attempts to report an AxSYM Amphetamine/Methamphetamine II patient result. | Software change control | Consignees were sent a Product Correction-Immediate Action Required letter on 1/21/08. The letter provided users with the steps to take to temporarily correct the problem until the recalling firm is able to send a corrected assay file. Users were requested to return a customer reply letter acknowledging receipt of the letter. For additional information, contact 1-877-4ABBOTT. |  | 1270 | 1,270 disks | Worldwide Distribution including USA territory of Puerto Rico, and countries of Australia, Canada, Chile, Colombia, Germany, Mexico, New Zealand, South Korea, Taiwan, Thailand and Venezuela. | Terminated | December 01, 2009 | 680 | 1270 | Computer
Z-1135-2008 | 46210 | No | FRW | Patient scale. | General Hospital | General Hospital | 510(K) Exempt | BHM Medical Scales, 800 lb. scale weight limit; Part Number 700.00510, Manufactured by BHM Medical Inc., Magog, QC, Canada J1X 5Y5; |  |  | December 19, 2007 | 2007 | Arjo, Inc. | Unintended Hanger Bar Detachment: The locknut on the scales may loosen during use and allow the hanger bar assembly of the patient lift to detach and fall. | Device Design | Arjo sent Urgent Device Field Correction Notification letters dated 12/19/07 to the end user accounts who received the affected scales, advising them of the potential for the hanger bar assembly to detach. The accounts were given the option immediately stopping use of the lifts with the affected scales or removing the scales from the lift so the lift can continue to be used. Instructions for removal of the scales were provided with the letter. The accounts were requested to complete and return the enclosed customer response form, acknowledging the receipt and understanding of the letter, and indicating whether they will complete the correction upon receipt of the correction parts or if they require the assistance of an Arjo Service Technician. |  | 772 | 772 scales | Nationwide including the states of Alabama, Arizona, California, Delaware, Florida, Hawaii, Iowa, Illinois, Indiana, Massachusetts, Michigan, Minnesota, Mississippi, North Carolina, North Dakota, New Hampshire, New Jersey, New York, Ohio, Pennsylvania, South Dakota, Tennessee, Washington and Wisconsin. | Terminated | November 18, 2009 | 700 | 798 | Computer
Z-0134-2008 | 44799 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Smiths Medical CADD-Sentry Pro Medication Safety Software - electronic pump protocol programming software - Administrator |  |  | August 14, 2007 | 2007 | Smiths Medical MD, Inc. | Incorrect dosage possible: A software anomaly in the CADD-Sentry Pro Administrator software (version B) where the values of a weight based protocol go to zero or round up to the next higher hundredths place, even if Administrator displayed the values correctly at data entry. | Software design | A Smiths Medical MD employee visited the consignees between August 14 and 31, 2007, to verbally notify them of the recall, collect the recall CD-ROM containing the affected software and replace with the updated software. |  | 3 | 3 CD-ROMs | UT, MD, CO | Terminated | October 31, 2008 | 444 | 3 | Computer
Z-0731-2008 | 46266 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Kodak DirectView Model DR 3000 System, Stationary X-ray System; Catalog #s 6551360, 6551386, 6552152, 6552160, 6551345, 6552319, 6552335. The units are labeled under the following firms: EASTMAN KODAK COMPANY (now known as Carestream Heath Inc.) Rochester, NY 14608, or CARESTREAM HEALTH, INC., Rochester, NY 14608. |  |  | December 21, 2007 | 2007 | Carestream Health, Inc. | Unintended Positioner Movement - The U-Arm Positioner starts to move 1) without command from the operator and/or 2) can unexpectedly start movement if the system is powered down and re-energized after a collision has occurred with the Quantum Table. | Software design | On 12/21/07, Service Bulletins (dated December 2007) were sent to consignees informing them of design modification, and providing modification instructions. The firm followed up with telephone calls to the direct user sites to schedule appointments to visit each site, in order to make the necessary modifications. The firm will be installing Safety Adaptation Board and software into all the affected DR3000 systems. |  | 19 | Domestic - 19 units | Nationwide Distribution including states of NY, PA, MA, MI, IL, NC, GA, FL, KS, MO, TX, IA, and WA. | Terminated | April 10, 2008 | 111 | 19 | Computer
Z-0732-2008 | 46318 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Siemens Syngo Imaging System with VB30A version of software. Model Number: 10014063, Siemans Medical Solutions, Malvern, PA 19355 |  |  | November 27, 2007 | 2007 | Siemens Medical Solutions USA, Inc | Inconsistent data - A buffer overflow may occur between information system and picture system if the system experiences a high load. (e.g. RIS/PACS communication is unusally high) | Software design | The recalling firm issued a Customer Safety Advisory, via certified mail. to the affected customers on 11/27/07. The letter informs the customer of the issue and also provides instruction to detect and avoid its occurrence. In addition, the firm is developing a software update to correct the issue, which will be installed on the affected systems when it is released. |  | 10 | 10 units | Nationwide Distribution, including states of AL, MO, NC, NY, OH, OR, TN, and TX. | Terminated | September 17, 2008 | 295 | 10 | Computer
Z-0583-2008 | 45807 | No | KNC | Obstetric table and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(K) Exempt | Hill-Rom Affinity Three Birthing Bed, Model P3700A. |  |  | November 21, 2007 | 2007 | Hill-Rom, Inc. | Hole/opening on the labor bar -- Amputation of a finger could result, if a finger is placed into the opening, as the hole/opening creates a shear point when the mechanism is activated. | Device Design | Consignees were notified of the problem via an Urgent Medical Device Correction letter dated 11/21/07. Users were informed that they would be receiving a free upgrade kit for all of their Affinity Birthing Beds. The Kit includes two plugs per bed, an insertion tool, and full written instructions. An instruction video is available at www.Hill-Rom.com. |  | 32379 | 32,379 for all model numbers | Worldwide Distribution-- including USA and countries of Argentina, Australia, Austria, Bahamas, Bahrain, Belgium, Bermuda, Brazil, Brunei, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Dominican Republic, Egypt, Finland, France, Germany, Greece, Guam, Honduras, Hong Kong, Hungary, Iceland, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Kuwait, Latvia, Lebanon, Malaysia, Mexico, Morocco, The Netherlands, New Calidonia, New Zealand, Nigeria, Northern Mariana Islands, Norway, Oman, Panama, Philippines, Poland, Portugal, Qatar, Republic of Korea, Saudi Arabia, Singapore, South Korea, Romania, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Uruguey, Venezuela, and Virgin Islands . | Terminated | September 22, 2009 | 671 | 32379 | Computer
Z-0079-2008 | 38311 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Intuitive Surgical, daVinci S Surgical System, Model Number: IS2000, A5.1 P1 and A5.0 P5 software level, Surgical Endoscopic Instrument Control System, Intuitive Surgical, Sunnyvale, CA 94086. |  |  | March 30, 2007 | 2007 | Intuitive Surgical, Inc. | System Lock-Up: Software anomalies could cause product failure during use, or on start-up. System transitions to a safe "soft-lock" state. | Software design | Intuitive Field Service Engineers installed a software upgrade beginning on 03/30/2007, and left a notification letter that the upgrade had been performed at each site. New software level is A5.1 P2. |  | 159 | 159 | Worldwide Distribution - USA, Belgium, China, France, India, Italy, Japan, Korea, Spain, Sweden, Switzerland, UK and Mexico. | Terminated | September 05, 2008 | 525 | 159 | Computer
Z-0728-2008 | 45953 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Invictus-- Bionime Rightest Blood Glucose Monitoring System |  |  | November 26, 2007 | 2007 | Invictus Scientific Inc | Abnormal Memory Recall: When the meter is fixed in 12 hour mode, the time frame display of all memory data will be shown incorrectly as AM or PM depending on when the user is recalling the memory stored in the meter. This will cause confusion as to when the last test was taken. | Software design | An Urgent: Device Recall Notification letter were sent to customers (direct accounts) beginning on November 20, 2007, via US Postal Service, with a return receipt requested. The distributors were instructed to inform consumers/users to fix the time format to the 24 hrs mode and to contact the distributors or the firm for replacement, if they prefer to use the time format in the 12 hrs mode. They were also instructed that if they have further distributed any of the recalled products, to please immediately contact those accounts and advise them of the recall situation, and have them return their outstanding recalled stocks. If customers have any questions, they were instructed to contact at 858-481-8485. |  | 21070 | 21,070 | Nationwide Distribution | Terminated | October 13, 2011 | 1417 | 21070 | Computer
Z-0571-2008 | 45798 | No | NQN | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(k) | Ventana Image Analysis System (VIAS) - PATHWAY anti-HER-2/neu (4B5) primary antibody with software version 2.3.0.2, Model Number: VIAS-799-10000-00, Pathology Image Analysis System, Ventana Medical Systems, Inc., Tucson, AZ 85755 |  |  | June 26, 2007 | 2007 | Ventana Medical Systems Inc | Incomplete upgrade: The software upgrade from February 1, 2007 may not have beencompleted. This software upgrade adds the algorithm (Slidetype) for thePATHWAY anti-HER-2/neu (4B5) antibody to the VIAS system. | Software change control | Beginning in 6/07, Ventana Medical Systems, Inc., made phonecalls to labs asking to confirm software version, date of upgrade if available and confirmation of whether site uses 4B5 and Imaging together. A phone script was provided. Customers with systems not ugraded were to upgrade before further use of system. A notification letter dated 8/17/2007 was later sent to VIAS customers advising that the software upgrade from February 1 , 2007 may not have beencompleted on their system. This software upgrade adds the algorithm (Slidetype) for thePATWAY anti-HER-2/neu (4B5) antibody to their VIAS system . This Slidetype is cleared for marketingand is required prior to imaging slides stained with PATHWAY anti-HER-2/neu (4B5).The letter also advises that their VIAS and NEXES systems are conneded to the internet which enables this upgrade to beperformed remotely with the support of the Ventana Medical Systems Customer Support Center(CSC). customers are instructed to call the CSC to carry out the upgrade or to schedule aTechnical Applications Specialist (TAS) to assist with the upgrade in their lab. Arrangements for the upgrade are to be made prior to imaging slides stained with PATHWAYanti-HER-2/neu (4B5). |  | 77 | 77 devices | Nationwide and Canada | Terminated | March 26, 2012 | 1735 | 77 | Computer
Z-1142-2008 | 46685 | No | LKK |  |  | General Hospital | PMA | Medtronic - SynchroMed EL Implantable Infusion Pump, Model Number: 8626-10. The contents of the inner package have been sterilized by ethylene oxide gas. 10 mL pump with reservoir valve. Medtronic, Inc., 710 Medtronic Parkway NE, Minneapolis, MN 55432-5604, USA. The implantable Medtronic SynchroMed EL Programmable Pump is part of the SynchroMed EL Infusion System designed to contain and administer parenteral drugs to a specific site. The implantable components of the SynchroMed EL Infusion System include the pump with or without a side catheter access port, catheters, and catheter accessories. |  |  | January 16, 2008 | 2008 | Medtronic Neuromodulation | Device/Drug Interaction - The company updated the labeling for the devices to include current patient management and treatment recommendations. The company received reports of inflammatory mass formations at or near the distal tip of intrathecal catheters which infuse opioids, baclofen, or chemotherapy drugs into patients. On January 16, 2008, Medtronic sent a letter to doctors who implant these d | Device Design | An Urgent Medial Device Correction letter was sent January 16, 2008, to Health Care Professionals. The letter describes the incidences, symptoms and recommendations for patient management. Excerpts from the approved Medtronic Professional Labeling are also included with the letter. The firm has requested unreported inflammatory mass in a patient with a Medtroinc device to the firm and to the FDA MedWatch Program by phone at 1-800-FDA-1088. Additional assistance may be obtained by contacting Medtronic Neuromodulation Technical Services at 1-800-707-0933. |  | 102792 | ~102,792 worldwide. 90,330 within US and 12,462 OUS | Worldwide Distribution -- USA including states of Washington D.C., Puerto Rico, and countries of Aruba, Australia, Austria, Belarus, Belgium, Brazil, Canada, China, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Egypt, Faroe Islands, Finland, France, French Polynesia, Germany, Greece, Guadeloupe, Hong Kong, Hungary, Iceland, India, Iran, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kuwait, Lebanon, Luxembourg, Malta, Mexico, Netherlands, Netherlands Antilles, New Caledonia, Norway, Pakistan, Poland, Portugal, Reunion, Romania, Russian Federation, San Marino, Saudi Arabia, Serbia and Montenegro, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Tunisia, Turkey, United Arab Emirates, United Kingdom, Vatican City State. | Terminated | June 06, 2011 | 1237 | 925128 | Computer
Z-1044-2008 | 46245 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Siemens Magnetom Harmony sygno. Model Number: 7104693, Siemens Medical Solutions USA, Malvern PA, 19355 |  |  | December 01, 2007 | 2007 | Siemens Medical Solutions USA, Inc | Magnet Quench, unintended- magnet of a mobile MRI system quenched into the exam room. | Device Design | The firm issued an Update Instruction package to consignees 12/07. As an immediate measure, the recalling firm is initiating field checks of mobile MRI systems per Update Instruction MR039/07/S. This update defines a safety inspection for mobile systems, including checks for potential ice formation and proper configuration of venting assemblies. |  | 10 | 10 units | Nationwide Distribution | Terminated | October 01, 2012 | 1766 | 210 | Computer
Z-1367-2008 | 45950 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Gambro cartridge blood tubing set used with Gambro Dasco Phoenix and Cobe Centrysystem 3 (and 3+) Hemodialysis System, Gambro, Lakewood, CO 80401 |  |  | September 10, 2007 | 2007 | Gambro Renal Products, Inc. | Kinked tubing on hemodialysis device may cause hemolysis. | Device Design | On September 10, 2007, the firm sent an Advisory Notice to all consignees alerting them that a kink in the tubing can result in physical hemolysis. The firm included with the Notice a Quick Reference Guide that provides drawing of the proper installation of the Phoenix and describes how to prevent bloodline kinking. A customer reply form was also included. The firm is in the process of making labeling changes that will be released with the updated software revisions. |  | 41000000 | 41,000,000 | Nationwide, including 17 VA facilities and 1 military facility. Foreign distribution to Canada, Korea, Mexico. | Terminated | January 03, 2012 | 1576 | 41000000 | Computer
Z-0987-2008 | 46294 | No | BZD | Noncontinuous ventilator (IPPB). | Anesthesiology | Anesthesiology | 510(k) | AirLife Infant nCPAP System Driver (Catalog No: 006900;) a nasal continuous positive airway pressure device, which is part of a larger system that includes an Infant Nasal CPAP generator, fixation device and patient interface (prongs and/or mask); Cardinal Health, McGaw Park, IL 60085 |  |  | January 04, 2008 | 2008 | Cardinal Health | Oxygen fluctuations - The AirLife Infant Nasal CPAP System exhibited O2 fluctuations when used with low O2 flow rates when the PTO/Auxillary port on the driver is used as a blender in administering O2 therapy via nasal cannula at flow rates of less than 4L/min. | Device Design | Cardinal Health sent an Urgent Medical Device Correction Notice dated 1/4/08 with a list of catalog number and affected serial numbers of the nCPAP Driver, and informed them that the nCPAP Driver may exhibited O2 fluctuations when used with low O2 flow rates when the PTO/Auxiliary port on the driver is used as a blender in administering O2 therapy via nasal cannula at flow rates of less than 4L/min. Customers were instructed to discontinue use of the PTO/Auxiliary port and were provided with a sticker reading "WARNINIG! DO NOT USE PTO/AUXILIARY PORT" that must be affixed to each unit in their possession. Sticker placement instructions and photo showing proper placement were provided with the letter. The accounts were informed that future upgrades to this device will eliminate the PTO/Auxiliary port and are expected to be completed in the field by June1, 2008. The accounts were requested to complete and return via fax at 847-689-9101, the enclosed acknowledgment form, indicating the number of affected drivers in their possession. Any questions were directed to Quality Systems-Professional Services at 1-800-292-9332 or 1-800-327-7503. |  | 314 | 314 units | Nationwide Distribution | Terminated | January 09, 2009 | 371 | 314 | Computer
Z-1141-2008 | 46414 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | iSite PACS image management system; Picture Archiving and Communications System; Radiology software package, Philips Healthcare Informatics, Foster City, CA 94404. |  |  | January 04, 2008 | 2008 | Philips Healthcare Informatics, Inc. | Incorrect display of one of the patient's images. This malfunction could cause an image to be incorrectly displayed, for example flipped left to right and anterior to posterior. The original image is still presented in the correct orientation next to the incorrect image when displayed on the monitor. | Software design | Philips Healthcare Informatics mailed customers notification: Urgent Product Correction regarding FCO#2954704-01/04/08-001-C. Firm released a corrected version of software (iSite PACS v.3.5.69) on January 21, 2008. Firm is scheduling customers to upgrade to the corrected version. |  | 197 | 197 units | Worldwide Distribution: USA, United Arab Emirates, Germany, the Netherlands, Denmark, Australia, France, Italy, Belgium, Ireland, Italy, and Canada. | Terminated | October 07, 2009 | 642 | 197 | Computer
Z-1104-2008 | 46306 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Centricity PACS RA1000 Workstation; for diagnostic image analysis. GE Healthcare Integrated IT Solutions, Barrington, IL 60010 |  |  | December 28, 2007 | 2007 | GE Healthcare Integrated IT Solutions | Report missing characters: Four special characters, when entered into the exam notes, are not transferred to the preview panel or hard copy printout (greater than [&gt;], less than [&lt;], quotation mark ["], and apostrophe [']). This may not be evident to the end user, as the symbols only appear when the Exam Note window is open. | Software design | GE Healthcare IITS sent Product Safety Notification letters dated 12/28/07 to all customers who have the GE Centricity PACS RA1000 Workstation software versions 2.1.X, or 3.0.X, informing them of the potential patient safety issue involving the exam notes window, which may be used to provide various patient information including lab results. To mitigate this problem, the users were instructed not to use the following four special characters (greater than sign &gt;, less than sign &lt;, double quotation mark " and apostrophe ') in the exam notes field in order to prevent any adverse patient events. For example, if a lab result was entered as GFR &lt;60, it should be typed as GFR less than 60. The customers were requested to make sure that all necessary physicians and personnel at their location are made aware of the problem and the recommended actions provided. A software patch to permanently resolve the issue will be installed at no charge when it becomes available. Any questions were directed to the GE Customer Care Center at 800-437-1171. |  | 557 | 557 units | Worldwide Distribution - USA, Australia, Belgium, Canada, Denmark, France, Germany, Hong Kong, Israel, Italy, Japan, Korea, Kuwait, Malaysia, New Zealand, Portugal, Saudi Arabia, Singapore, South Africa, Spain, Switzerland, Taiwan, Turkey, and United Kingdom. | Terminated | May 17, 2010 | 871 | 557 | Computer
Z-0905-2008 | 46330 | No | EZN | Ureteral dilator. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Pathway Balloon Expandable Ureteral Access Sheath- 11/14/28, Ureteral dilator, Catalog Number: 255100; Product Number: M0062551000, Onset Medical, Boston Scientific Corp., Natick, MA 01760 |  |  | November 30, 2007 | 2007 | Boston Scientific Corporation | Sheath Removal Difficulty: Users experience difficulty in removing the sheath. | Device Design | Consignees were sent an Urgent Medical Device Recall letter on November 29, 2007. The letter instructed users to check their inventory, remove any affected product and place them in quarantine. It also included instructions for returning the product. Users were asked to return a Reply Verification Tracking Form. Follow-up communication will be made to users who do not return the tracking form. |  | 89 | 89 | Nationwide Distribution | Terminated | October 22, 2008 | 327 | 3996 | Computer
Z-0973-2008 | 46200 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Siemens Coherence Impression Therapist system, Part No: 7341410, medical charged-particle radiation therapy system with software version 2.1. Siemens AG, Concord, CA. |  |  | October 02, 2007 | 2007 | Siemens Medical Solutions USA, Inc | Image Orientation Incorrect (Image Shifts) - 1) if the reference image has been calibrated for centering and the reference image is used for patient positioning, the image will shift when the positioning tools are used. 2) This problem occurs only when non-square reference images are used, such as when images are cropped. At this point, the images will refresh, and the center of the reference im | Software design | On 10/03/2007, the firm sent Consignees a Customer Safety Advisory Notice informing them of the two potential issues with the display of reference images used for position correction. Consignees will be informed when a solution to this issue has been validated. |  | 16 | 16 units | Worldwide Distribution-USA including 48 consignees in the US and 39 outside the US. | Terminated | December 24, 2008 | 449 | 94 | Computer
Z-0958-2008 | 45945 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Cytomics FC 500 MPL Flow Cytometry System with MXP Software, Software Version 2.1, Automated differential cell counter, Part Number: 626554, Beckman Coulter, Fullerton, CA 92834 |  |  | June 29, 2007 | 2007 | Beckman Coulter Inc | Data generation error: If the cytosettings is not refreshed or restarted during data importation, the output will include old and new data. | Software design | Consignees were notified by an Urgent Product Corrective Action letter sent on 6/29/07. The letter provided details on the data error and instruction on how to avoid the error. Users were asked to retain notification as part of the Quality Systems documentation and return a response form acknowledging receipt. For additional information, contact 1-800-526-7694. |  | 589 | 516 units in the US; 73 units in Canada | Worldwide Distribution- USA and country of Canada. | Terminated | December 13, 2012 | 1994 | 589 | Computer
Z-1038-2008 | 46360 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Leonardo Workstation Picture Archiving and Communication System, Model Number: 7129534, Software Version: VB30B, Siemens Medical Solutions, Inc., Malvern, PA 19355 |  |  | December 21, 2007 | 2007 | Siemens Medical Solutions USA, Inc | Incorrect display: The orientation labels will be incorrectly displayed on the reconstructed InSpace 3-D image if the orientation was not originally HFS on the acquisition system. | Software design | Siemens Medical Solutions issued a Customer Safety Advisory Notice dated 12/21/07 to its customers to inform them of the problem and instructions to avoid the problem until the firm's representative is available to upgrade the software. |  | 47 | 47 Units | Nationwide Distribution. | Terminated | September 30, 2008 | 284 | 47 | Computer
Z-1102-2008 | 46150 | No | LEH | Vancomycin test system. | Toxicology | Toxicology | 510(k) | Seradyn QMS Vancomycin reagents, Part Number: 0373589. Seradyn, Inc, Indianapolis, IN. |  |  | December 20, 2007 | 2007 | Seradyn, Inc. | False negative test results due to Interfering substances (heterophile antibodies) in the patient's blood may produce erroneously low results. | Device Design | Consignees were notified of the problem via Urgent Medical Device Correction letter dated 12/20/07. The letter provided information on the issues concerning the recall. The firm will revise the package insert to address the potential occurrence of false negative bias. If inconsistencies with Vancomycin level exist, confirm with an alternative test method. For further information please contact 1-800-428-4072. |  | 3682 | 1,819 kits plus bulk material for 1,863 kits. | Worldwide Distribution --- USA including states of California and New York, and countries of Australia, Belgium, Canada, Finland, Germany, Ireland, Japan, Spain and United Kingdom. | Terminated | December 30, 2008 | 376 | 3682 | Computer
Z-0661-2008 | 45410 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON Sequoia Diagnostic Ultrasound System, general purpose Diagnostic Ultrasound system. Siemens Medical Solutions USA, Inc., Mountain View, CA |  |  | September 25, 2007 | 2007 | Siemens Medical Solutions USA, Inc. | Measurement Error affecting all Sequoia systems using the PAL video standard configuration and have a perspective PC installed: 1) The Sequoia" system is operating at a software revision below 8.0. 2) The Sequoia" system is configured to PAL video mode (normal for systems running at 240V 50 Hz.). 3) While using any of the Perspective" advanced display options listed above, an image or clip is t | Software design | Consignees were notified via a Urgent Medical Device Notification letter/customer letter sent out on 9/25/2007 under Siemens update program US013/07/S. The Siemens Service organization coordinates the implementation of the field correction with the distributors for all countries. Consignees who do not respond to the recall communication will be visit ed and/or telephoned for follow up. |  | 280 | 280 | Worldwide Distribution including the countries of: Australia, Austria, Belgium, Brazil, Brunei, China, Denmark, Finland, France, Germany, Great Britain (UK), Greece, Hungary, India, Iran, Ireland, Italy, Latvia, Netherlands, New Zealand, Romania, Russian Federation, Singapore, Spain, Switzerland, Thailand. U.S customers were not affected. | Terminated | December 20, 2010 | 1182 | 280 | Computer
Z-0575-2008 | 45892 | No | NVI | Amphetamine test system. | Toxicology | Immunology | 510(k) | BioRad BioPlex 2200 Software, for BioPlex 2200 Instrument, an automated immunodiagnostic analyzer, Bio-Rad Laboratories, Inc., Hercules, CA 94547 |  |  | November 13, 2007 | 2007 | Bio-Rad Laboratories Inc | Incorrect values: An error was found in the BioPlex 2200 Software, Version 2.0 that results in assignment values being utilized for the BioPlex assay Calibrator Set that may be slightly different from those printed on the Value Assignment Data Sheet. Some values may be identical between the lot assignment combinations with the ANA Screen with MDSS kit. | Software design | Bio-Rad sent a Voluntary Field Correction notice dated November 1, 2007. Consignees were notified by phone call on Nov. 13, 2007, followed up by facsimile (domestic). Sub-recall to be by e-mail message for consignees outside the US. Distribution of the product ceased on Nov. 2, 2007. Bio-Rad will provide Calibrator Set CDs with only one lot of calibrator assignment values until a new version of software is available. |  | 7333800 | 85 units of calibrator lot 86280 and 16 units of calibrator lot 86290 were distributed. | Worldwide Distribution - USA, Canada, France, and Germany. | Terminated | July 23, 2008 | 253 | 7333800 | Computer
Z-0526-2008 | 45880 | [6 empty] | BD FACS Sample Prep Assistant II, Part Number: 337170; |  |  | October 02, 2007 | 2007 | BD Biosciences | Mixing Performance: SPA II devices manufactured with the affected product motor control board may exhibit suboptimal mixing performance. | Software change control | On 10/1/07, all consignees were notified via Important Product Information letter and also an on-site visit by the firm's sales representative. The letter informed users of the affected product and provided instructions on the recall. For additional concerns or questions contact BD Customer Support Center at 1-877-232-8995. |  |  |  | Worldwide Distribution --- USA including states of: NJ, IL, TX, PA, NY, MN, and AZ, and countries of Australia, Belgium, Hong Kong, Trinidad and Tobago.Australia, Belgium, Hong Kong, Trinidad and Tobago. | Terminated | December 13, 2010 | 1168 |  | Computer
Z-1317-2008 | 46398 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | LifeScan OneTouch Data Management Software Version 1.0, Part Number: 021-722-01(for OneTouch Select and UltraMini (Download) Blood Glucose Meters), LifeScan Inc, Milpitas, CA 95037. |  |  | December 18, 2007 | 2007 | Lifescan Inc | Meter Temporarily Freezes-- A software compatibility issue may cause the blood glucose meter to cease operations and freeze temporarily. | Software design | The firm sent an Urgent Recall letter to Health Care Professionals on 12/18/2007. The letter advised user to uninstall software drivers and provided options on how to uninstall. For further assistance, please contact LifeScan Customer Service at 1-866-679-9250. |  | 260 | 260 units | Nationwide Distribution including the states of AL, CA, IN, IO, MA, MD, MI, MS, NE, NY, NC, OR, PA, TN, TX, UT, VT, VA, WV, and WI. | Terminated | July 18, 2008 | 213 | 260 | Computer
Z-1400-2008 | 46940 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Oncor Expression Medical Linear Accelerator, Material Number: 07360717, with control console software versions 6.5/7.2, 8.0, 9.0 and 9.1, Manufactured by Siemens Medical Solutions USA, Inc, Concord, CA |  |  | January 11, 2008 | 2008 | Siemens Medical Solutions USA, Inc | Linear accelerator gantry may rotate in an unexpected direction and may cause patient injury on impact. | Software design | This is a mandatory software upgrade for consignees. The firm's representatives have contacted consignees via Urgent Device Notification letters dated February 14, 2008, notifying them of the new update and making arrangements to install it. |  | 5 | 5 units | Worldwide-USA including states of AL, AK, AZ, AR, CA, CO, DE, DC, FL, GA, IL, IN, KS, KY, LA, MD, ME, MI, MN, MS, MO, MN, NE, NV, NH, NJ, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, VT, VA, WA, WV, WI, and WY and countries of Angola, Australia, Austria,Belgium,Brazil,Canada,Chile,Colombia,Costa Rica, Croatia,Cyprus,The Czech Republic,Denmark,DPR of Korea,El Salvador,France,Germany,Greece,Hungary,India,Iran,Ireland,Italy,Japan,Jordan,Malaysia,Mexico, The Netherlands,New Zealand,Norway,P.R. China,Pakistan,Peru,Philippines,Poland,Portugal,Republic of Korea,Romania,Russian Fed.Saudi Arabia,Singapore,Slovakia,South Africa,Spain,Sweden,Taiwan,Thailand,Trinidad,Tobago,Turkey, U.A.E.,Ukraine, The United Kingdom, and Venezuela | Terminated | May 06, 2011 | 1211 | 3292 | Computer
Z-1349-2008 | 46889 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Agfa IMPAX CardioVascular Suite Results Management Nuclear Reporting Module, Outbound Report, Model/Catalog No.: L7W1P, Software Versions: 2.04.17, 2.04.19 and 2.04.20, AGFA HealthCare, Westerly, RI |  |  | February 22, 2008 | 2008 | AGFA Corp. | Reporting Differences: The Outbound Report software produces an Outbound Report (text version) which does not contain all the clinical content elements present in the PDF version of the report, also generated on the Nuclear Cardiology Reporting Module. | Software design | All five (5) sites utilizing this product were notified about the issue directly via telephone. An "Urgent Safety Notice" was sent via FED-EX to all the sites using the product on February 22, 2008. The letter describes the potential issue and mitigation. The recalling firm will also visit each site and make configuration changes to resolve this issue. |  | 5 | 5 units | Nationwide to five hospitals in DC, NY, OH, TN, and VA. | Terminated | July 27, 2010 | 886 | 5 | Computer
Z-1365-2008 | 47341 | No | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | BIOMET 3i Certain MicroMiniplant Straight Healing Abutment, Model/Item Number: ISMHA32, 3.4mm(D) X 0mm(P) X 2mm(H), BioMet 3i, Palm Gardens, FL. |  |  | February 06, 2008 | 2008 | Biomet 3i, Inc. | Seating Problems: The Micro-miniplant Straight Healing Abutments do not fully seat onto the restorative platform of the implant leaving a gap between the abutment and the implant. | Device Design | USA customers were contacted by telephone with a follow-up fax. International customers were first contacted by E-mail then by fax letter. Each customer was instructed to check their respective inventory for the affected part and corresponding lot number. They were then requested to fax back a response form to Biomet 3i Regulatory Services to provide feedback concerning their inventory. They were also instructed to return the affected product to Biomet 3i. |  | 270 | 270 | Worldwide | Terminated | July 07, 2008 | 152 | 270 | Computer
Z-0724-2008 | 45815 | No | CGL | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Nova - EZ CHEM Creatinine Meter: SRS Versions 5.1 and 5.2 |  |  | November 02, 2007 | 2007 | Nova Biomedical Corporation | Incorrect Results - Creatinine results were lower than the laboratory reference method. | Device Design | Nova Biomedical issued an Urgent: Medical Device Recall notification letter to consignees dated November 19, 2007. The letter addressed the issues with the product(s) and requested that if users had any remaining stock that it must be returned. The firm also provided a Customer Response form to be returned. Questions concerning this recall will be addressed by Gina Vogelsberg at 1-800-544-4624 or by email at gvogelsberg@ezem.com. |  | 133 | 133 units | Nationwide Distribution, including the state of NY | Terminated | December 09, 2008 | 403 | 719 | Computer
Z-1107-2004 | 45411 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Siemans- ACUSON Aspen Diagnostic Ultrasound System, general purpose Diagnostic Ultrasound system, Model Number: 8247887, Siemans, Mountain View, CA 94039 |  |  | September 18, 2007 | 2007 | Siemens Medical Solutions USA, Inc. | Software error: Error affects all Aspen customer systems using the PAL video standard configuration and have a Perspective PC installed. This can lead to a measurement error under this specific condition: On secondary capture images/clips stored from within any Perspective advanced display option (specifically: 3D Surface Rendering, 3D MultiPlanar Reconstruction, Color Doppler 3D MultiPlanar Reco | Software design | Siemens issued an Urgent Medical Device Notification to consignees 9/18/07, under Siemens update program US014/07/S. (Distribution on this product had ceased on 6/29/2005.) The firm has instructed the user to not take measurements or use measurements on these secondary captures under the conditions described in the notice, to place a copy of the instructions for use in the file, and to notify all users of the issue with the software. The Siemens Service organization has coordinated the implementation of the field correction with the distributors for all countries. Non-responding consignees will be visited or called by telephone for follow up. |  | 280 | 280 units | Worldwide Distribution --- including USA and countries of Argentina, Australia, Austria, Belgium, Brazil, China, Czech Republic, Denmark, Egypt, France, Germany, Great Britain (UK), Greece, Hong Kong, India, Italy, Jordan, Malaysia, Norway, Poland, Portugal, Romania, Russian Federation, Sweden, Switzerland, Thailand, and Turkey. | Terminated | April 21, 2011 | 1311 | 280 | Computer
Z-0720-2008 | 45527 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand VARiS 1.4g (Medical Charged Particle Radiation Therapy System, Record and Verify System); Linear Accelerator with RTP Exchange v6.2, v6.6, &amp; v8.0, All Models; Varian Medical Systems, Palo Alto, CA |  |  | October 02, 2007 | 2007 | Varian Medical Systems Inc | Incorrect Software Validation- The ARC treatment plans containing a Dose Dynamic MLC for IMRT delivery that will not be recognized properly by the system. Under these circumstances, an arc plan can be administered to a patient whereby the gantry will not rotate, resulting in delivery to a single location rather than through the gantry arc. | Software design | An Urgent Medical Device Correction letter was distributed October 2, 2007, to all direct consignees, informing them that the ARC treatment plans containing a Dose Dynamic MLC for IMRT delivery will not be recognized properly by the System. The firm provided corrective instructions and informed consignees that a product modification will be developed to resolve the issue. A Varian Service Representative will be scheduling a visit to upgrade the application. |  |  |  | Worldwide Distribution | Terminated | May 17, 2011 | 1323 |  | Computer
Z-0546-2008 | 45809 | No | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | ProTime Microcoagulation System Instrument, Catalogue Number: PROTIMEPRO, International Technidyne Corporation (ITC), Piscataway, NJ 08854 |  |  | November 19, 2007 | 2007 | International Technidyne Corporation | An increase in the frequency of non-conformance reports (NCMR) for certain displayed error messages. | Device Design | ITC sent to distributors, self test users and health care professionals on 11/19/07, an Urgent Medical Device Recall Letter requesting return of the devices. |  | 3482 | 3,482 units | Worldwide-USA and countries of China, Taiwan, Greece, Ireland, Chile, Finland, Italy, Spain, Norway, Germany, Switzerland., Singapore, Korea, Indonesia, Lichtenstein, Israel, Saudi Arabia, Latvia, Kuwait, and Slovenia. | Terminated | August 12, 2008 | 267 | 5910 | Computer
Z-0539-2008 | 45496 | No | KPA | Automated slide stainer. | Pathology | Pathology | 510(K) Exempt | Benchmark LT, Slide Staining System, Catalog Number: N750-BMKLT-FS, Ventana Medical Systems Inc, Tucson, AZ |  |  | September 29, 2006 | 2006 | Ventana Medical Systems Inc | Leaking Carboys/Safety Updates:The Ventana Staining Platform includes one or more carboys that may have the potential to leak if the spigot is either improperly installed or not fully closed. | Device Design | The firm has mailed three customer letters, a) the first letter dated 9/30/2005 describes the upgraded Voila CD that contains safety information; b) a second letter, dated 6/27/2006, identified as a "Customer Alert" was mailed describing an upgrade to the software system version 9.3/10.3 and the CD that includes safety improvements, c) the third letter dated 9/29/2006 and identified as a Safety Alert advises that as part of their effort to continuously improve laboratory safety, they are updating the labeling and instructions for use regarding installation and cleaning for the carboy as the Ventana staining platform includes one or more carboys that may have the potential to leak if the spigot is either improperly installed or not fully closed. Enclosed are pictogram labels that indicate "slip hazard" and "refer to instructions for use." The customer is instructed to apply these two labels to the carboys above the spigot as a reminder to laboratory personnel to keep the spigot in the full "off" position when not in use and to refer to the instructions for use for maintenance. In addition, the letter advises that there is an electrical shock warning for the power strip that is part of their system to indicate that it should be mounted off of the floor. This power strip is intended to connect the computer components of the staining system and should not be placed on the floor where fluid might contact it. Also included is a warning label for shock hazard to apply to their power strip. |  | 1712 | USA - 1,712 units; Foreign - 1,159 units (total for recalls-Z-0537-0541-2008) | Worldwide; USA, Australia, China, Canada, Germany, Great Britain, France, Italy, Hungary, Singapore, Korea, Taiwan, Spain, Poland, Greece, Austria, Sweden, Japan, Israel, Denmark, Turkey, Finland, Saudi Arabia, The Netherlands, Norway, Slovenia, Ireland, and Belgium | Terminated | June 06, 2013 | 2442 | 1712 | Computer
Z-0258-2008 | 38312 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Intuitive Surgical EndoWrist, 8 mm, PK Dissecting Forceps for use with the da Vinci and da Vinci S Surgical systems, Part Numbers: 400214-02, 400214-03, 420214-02 and 420214-03, Intuitive Surgical, Inc., Sunnyvale, CA |  |  | April 13, 2007 | 2007 | Intuitive Surgical, Inc. | Mislabeling-electrical isolation requirements: devices were incorrectly labeled with a CF symbol not their proper BF Symbol on the instrument housing. (All instruments meeting the CF isolation requirements are suitable for direct cardiac application, while instruments meeting BF isolation requirements are not suitable for direct cardiac application) | Device Design | On January 9, 2008, via an Urgent Device Recall Letter the recalling firm notifed users of the defect (mislabeling with heart symbol) and to Immediately Discontinue any direct cardiac applications. The recalling firm will arrange with each consignee a time to relabel their devices. Customers on April 17, 2007 had earlier been warned of the device defect via letter but no correction was indicated. |  | 1136 | 1,136 units | Worldwide Distribution --- USA including states of MN, PA, AK, NY, CA, MD, IN, AZ, CO, AR, TX, TN, FL, VA, MA, NE, NC, IL, OH, ME, WI, MD, NJ, MN, MI, VA, IA, OR, MO, NC, MS, AL, WA, GA, KY, LA, AK, AR, SD, SC, DE, NV, HI, NH, WV, and CT, and countries of Canada, Australia, Belgium, Spain, Hong Kong, Italy, Switzerland, and Sweden. ( 192 USA hospitals, 3 foreign hospitals and 5 foreign distributors) | Terminated | August 01, 2008 | 476 | 1136 | Computer
Z-0481-2008 | 45528 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | StelKast Proven Knee - Tibial Half Block Augmentation and Screw. Part Number SC2287-3-5, Size 3, Thickness 5 mm, StelKast, Murray, PA 15317 |  |  | October 05, 2007 | 2007 | Stelkast Co | The screws may fracture | Device Design | The recalling firm telephoned their customers on 10/5/07 to inform them of the problem and the need to return the product. As a follow up to the telephone calls, the recalling firm issued a recall letter (Field Notification) on 10/9/07 via certified mail. |  | 41 | 41 units | Nationwide: shipped to medical facilities in IL, NC, PA, TN, TX, VA, and WV. | Terminated | January 23, 2008 | 110 | 207 | Computer
Z-0183-2008 | 45441 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE OEC, Insta Trak 3500 Plus with the Cranial Application for Image Guided Surgery, GE Healthcare, Surgery, Salt Lake City, UT 84116 |  |  | October 18, 2007 | 2007 | GE OEC Medical Systems, Inc | Crack, Fracture or Detachment: The Transmitter Arm for the Mayfield-compatible Transmitter may become detached from the unit and enter into the surgical field. | Device Design | Consignees were sent an Urgent Recall Notice via mail dated 10/11/07, notifying the customer to examine the plastic end of the transmitter arm and to not use the transmitter arm, if any cracks or damage are observed. Also to stop using Steam Sterilization method and to utilize the other Sterilization methods indicated in the Operators Manual. The root cause is still under investigation GEHC Surgery will provide, at no charge, a solution to this issue when it becomes available. |  | 169 | 169 units | Worldwide Distribution --- USA including states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, IL, MA, MD, MI, MS, NC, NY, OK, OR, PA, SC, TX, UT, VA, WI, and WV, and countries of Canada, China, France, Hong Kong, Japan, Mexico, Spain, and Switzerland. | Terminated | August 10, 2012 | 1758 | 169 | Computer
Z-0365-2008 | 39353 | No | LZS |  |  | Ophthalmic | PMA | Advanced Medical Optics (AMO), VISX WaveScan WaveFront System (software). Part Number 0070-1478. Advanced Medical Optics, Santa Clara, CA 95051. |  |  | May 30, 2007 | 2007 | Visx Inc | Erroneous Treatment Calculations - Two software caused errors; in combination, in the WaveScan System Software, will result in an erroneous treatment calculation (overcorrection) in patients. (1) Installation of software for a certain brand of computer (Dedicated Computing LLC) can cause a software registry setting for an algorithm that identifies image reflection to be erroneously set in the "off | Software design | AMO sent an Action Required notification on May 30, 2007, to AMO Certified Physicians. The letter included a CD software update (0070-1721) to be installed by users that provided correct registry parameter settings. Upon completion of installation, a confirmation form will print out. Customers are requested to fax back confirmation to AMO (408-773-7055). |  | 26068 | 196 systems are affected in the U.S. and 133 outside of the U.S. | Worldwide Distribution --- USA including states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, ID, IL, IN, KS, KY, LA, MD, MI, MN, MO,MS, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK,OR,PA, SC, TN, TX, UT, VA, WA, WI, and WV and countries of Canada and Mexico. | Terminated | April 17, 2008 | 323 | 26068 | Computer
Z-0120-2008 | 44937 | No | JJX | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AmniSure ROM Test, Fetal Membrane Rupture Test for |  |  | October 03, 2007 | 2007 | AmniSure International LLC | Weak true postive/false negative results: weak true positive or false negatives in ruptured fetal membranes may occur due to a "hook effect" | Device Design | AmniSure notified users by Medical Device Correction (Recall) letter dated October 1, 2007 issued via US Mail. The letter provides additional labeling instructions and a checklist for trouble shooting, including a diluation procedure that allows users to determine and solve the problem of the hook effect with this device. |  | 319976 | 319,976 units | Worldwide, USA, Australia, Austria, Italy, Korea, Bulgaria, Russia, Spain, Kuwait, Switzerland, Luexemburg, Ireland, Denmark, Netherlands, Jordan, Israel, Chile, UK, Singapore, Basil, New Zealand, Turkey, France, Greece, and Equador | Terminated | February 06, 2008 | 126 | 319976 | Computer
Z-0733-2008 | 45906 | No | FPO | Wheeled stretcher. | General Hospital | General Hospital | 510(K) Exempt | Sechrist Hyperbaric Chamber Gurney, Standard Gurney &amp; Stretcher Model, P/N : 21502, Sechrist Industries, Inc., Anaheim, CA |  |  | October 31, 2007 | 2007 | Sechrist Industries Inc | Brake Problems: The stretcher brake cams may crack while attempting to lock the stretcher brake (Sechrist is supplied the stretchers by the manufacturer Stryker Medical) | Device Design | Urgent Medical Device Correction letters were sent to customers beginning on November 9, 2007, via FedEx &amp; DHL overnight. Customers were notified this action was being taken to providing notification to their customers as a continued implementation of the Stryker Medical field correction.The Notification letter instructed their customers to:1. Locate the gurneys with the serial numbers listed in this advisory.2. If any of the affected gurneys cannot be located, determine if they have been sold, transferred, or donated to another facility and notify us regarding their new location.3. Verify the brakes are working properly. The brake function may be checked by engaging the brake pedal and pushing on each end of the gurney to make sure that casters are locked securely.4. Remove the gurneys with non-functioning brakes from service until brake cams have been replaced and brakes are fully functional.5. Contract Sechrist for the replacement parts package.6. Complete the retrofit of the Gurney.7. Verify correct function of the brake assembly following retrofit.8. Complete and send the Correction Verification Form contained in the package to Sechrist. |  | 559 | 559 Units Total (483 in US) for all models | Worldwide-USA, Brazil, Bulgaria, Canada, Dubai, U.A.E., Estonia, India, Indonesia, Japan, Malaysia, Saudi Arabia, Sweden, Taiwan, Thailand, and Vietnam. | Terminated | April 19, 2012 | 1632 | 559 | Computer
Z-0519-2008 | 45913 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Medical Systems Information Technologies Clinical Information Center (CIC) Pro v5.0.3 &amp; v5.0.6 systems, GE Healthcare, Clinical Systems |  |  | November 08, 2007 | 2007 | General Electric Medical Systems Information Technology | Potential for Loss of Audible Alarming and ECG Tab Setting Changes; from Clinical Informaiton Center (CIC) Pro v5.0.3 &amp; 5.0.6 when used in combination with Telemetry and bedside monitors. | Software design | A "GE Urgent Medical Device Correction" letter was sent to consignees on 10/16/2007. The letter described the problem and product involved. The letter also gave short term recommendations. A long term solution is being worked on by GE and will be announced when determined. |  | 3116 | 3, 116 | Worldwide-USA including states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MO, MN, MS, MT, NC, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WI, WV, WY, PR, and District of Columbia and countries of The United Kingdom, Turkey, Thailand, Tsiwan, Syria , Switzerland, Sweden, Spain, Slovenia, Singapore, Saudi Arabia, Russia, Poland, Qatar, Pakistan, Philippines, Palestinian Territory, New Zealand, Norway, Netherlands, Nigerie, Malaysia, Lebanon, Jordan, Italy, Israel , Ireland, India, Honk Kong, Greece, Germany, France, Finland, Egypt, Denmark, Czech Republic, Cyprus, Costa Rica, Colombia,China, Chile, Canada, Delgium, Bahamas, Austria, and Australia. | Terminated | December 10, 2011 | 1493 | 3116 | Computer
Z-0295-2008 | 44941 | timer, general laboratory | JBS | General purpose laboratory equipment labeled or promoted for a specific medical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Nova 8 Analyzer Calibrator Pack (Catalog/REF # 15196), Nova Biomedical, Waltham, MA |  |  | September 17, 2007 | 2007 | Nova Biomedical Corporation | Elevated Normalized Ionized Calcium (nCa) and Normalized Ionized Magnesium (nMg) calculated values on patient samples | Device Design | Nova issued a Customer Advisory Notice notification by phoneon 9/17/07 to contact customers using a script . If the facility reports Normalized Ionized Calcium or Normalized Ionized Magnesium, Nova Biomedical will replace existing stock of Nova 8 calibrator packs. If the facility does not report Normalized Ionized Calcium or NormalizedIonized Magnesium, the account can continue to use the existing lot of calibrator packs.This contact will be documented within the Nova Biomedical complaint database. Intemational Customer Advisory Notice being sent to Nova Biomedical Subsidiariies and Distributors by Federal Express Envelope.The Federal Express envelopes tracking number will be used for delivery confirmation. |  | 2999 | 2,999 units | Worldwide, USA, Canada, Argentina, Bangladesh, Bolivia, China, Czech Republic, Denmark, Ecuador, Egypt, Finland, France,Germany,Greece, Hungary, India, Indonesia, Ireland,Italy, Japan, Korea, Mexico, Peru, Poland, Scotland, South Africa, Spain, Sri Lanka. Switzerland, Syria ,Taiwan, Turkey and The United Kingdom | Terminated | March 16, 2012 | 1642 | 2999 | Computer
Z-0150-2008 | 45405 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Datascope, Panorama Patient Monitoring Network; Panorma Telepack 608;Software Versions 8.1.X, 8.3, 8.3.1, 8.4.1, and 8.5.X, Part Number: 0998-00-0191-04, Datascope Corp., Mahwah, NJ |  |  | October 09, 2007 | 2007 | Datascope Corporation | Display problems: If an ECG cable, which has been damaged due to the ingress of liquid or by mechanical trauma, is utilized with the Telepack, it may cause the Telepack to switch between the 3-lead and 5 lead-input modes. During the switching process, the digital heart rate displayed at the Panorama Central Station will be frozen, and subsquent arrthymia alarms will not be announced. | Software design | Urgent Product Field Correction notices were sent to hospital administrators on October 9, 2007, by certified mail, return receipt requested. The letter describes the scenario that causes the problem and its remediation It also describes how the software upgrade will fix the problem and alerts hospitals that a Datascope Service Representative will be contacting them to arrange for the upgrade. |  | 338 | 338 sites | Nationwide | Terminated | November 04, 2008 | 392 | 338 | Computer
Z-0326-2008 | 38368 | [6 empty] | Actiwatch 64 AM, Bandless Part Number 198-0301-AM1, Respironics, Inc., Mini Mitter, a Respironics Company, Bend, Oregon |  |  | June 27, 2007 | 2007 | Mini-Mitter Company, Inc. | Communication Failure: the devices have a memory chip that has a limited number of read/write cycles which can trigger set up information and/or subject information anomalies or intermittent communication annomalies. If left unaddressed, the issues can lead to permanent communication failure. | Software design | A Voluntary Medical Device Recall letter dated June 26, 2007. The recall letter will instruct the customers to 1) fill out a business reply form to confirm contact information and devices serial numbers in their possession. Then FAX this info back immediately to Stericycle or Mini-Mitter. 2) Refer to the work scheduled rework date reserved for them and noted in their recall notification letter and return the device for service at that time; and 3) Package the device appropriately, insure it and return it via UPS using the provided UPS shipping label. Product returned to Mini-Mitter will be quarantined, decontaminated, and repaired. Min-Mitter will rework all the devices returned. Each will have a new memory chip installed, be re-calibrated, retested, and have a new battery installed if required. Min-Mitter will provide a new one year warranty. |  | 20 | 20 | Worldwide | Terminated | October 20, 2010 | 1211 | 1073 | Computer
Z-0488-2008 | 45870 | No | HNN | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | Premier Edge, Bulk Microsurgical Knife, 3.0mm Slit Knife, Item Code B-PE3830, Oasis, Glendora, CA |  |  | July 18, 2007 | 2007 | Oasis Medical Inc | Sterility (package integrity) compromised: the device is puncturing the packing during handling. Once the packaging is damaged, the product is no longer sterile as labeled. THIS IS AN EXPANSION OF A PREVIOUS RECALL | Device Design | A recall notification, Urgent Medical Device Recall Expansion Letter, dated 7.18.2007, was sent to customers advising the letter is to inform them of a voluntary product recall expansion involving the OASIS Medical Premier Edge Knives. A product label for ease in identifying the product is included. The letter further advises this voluntary recall has been expanded due to the product puncturing the packing during rough handling. Once the packaging is damaged, the product is no longer sterile as labeled. The letter instructs customers to immediately examine their inventory and quarantine product subject to recall. Once the product has been identified, product is to be returned by FedEx to OASISMedical. The letter further instructs that if the customer has further distributed this product, to identify their customers and notify them at once of this product recall. This recall should be carried out to the user level. |  | 28319 | 28,319 individual knives (Z-0485-0516-2008) | Worldwide: USA, Kuwait, Lebanon, Lithuania, Mexico, Netherlands, Norway, South Africa, Saudi Arabia, Slovakia, Spain, Sweden, Switzwerland, Turkey, United Kingdom, Ecuador, Finland, France, Germany, Greece, Italy, Denmark, Columbia, Chile, Brazil, Belgium, and Australia | Terminated | September 25, 2009 | 800 | 28319 | Computer
Z-0305-2008 | 45811 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Toshiba Aplio XG Diagnostic Ultrasound System, Model: SSA-790A with version 1.1 software, Toshiba America Medical Systems, Inc., Tustin, CA. 92780 |  |  | October 02, 2007 | 2007 | Toshiba American Med Sys Inc | Panel Lock-up: When using dynamic image acquisition modes the panel may lock up, if the operator attempts to perform a still image capture. If the panel locks up the system will have to be rebooted to continue the examination which may cause a delay in completion of the patient examination. | Software design | Notification letters, Urgent Medical Device Correction, with an attached Customer Response form were sent to customers beginning on October 2, 2007. Customers were notified this action was being taken as a result of an investigation that resulted from a software issue that was found by the manufacturer. Customers were informed that a Toshiba Customer Service Engineer would contact them to schedule a software upgrade when the software becomes available. If customers have any questions they were instructed to contact their or call (800) 421-1968 and ask for the Regulatory Affairs Department.--A short term instruction included "Avoid using the Freeze Key when performing dynamic image acquisitions" |  | 69 | 69 | Nationwide: USA including the states of AK, AZ, CA, FL, IA, IL, IN, KS, MI, MN, MO, MT, NE, NV, NY, OH, OR, PR, VA, WI &amp; WV | Terminated | April 20, 2012 | 1662 | 69 | Computer
Z-0297-2008 | 45511 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Siemens, ACUSON Antares, Ultrasound System (Standard and Premium Edition), 5.0 Systems with software version 200.0.054, P/N : 10032747 and 10032746 standard, 10032746, 10037591, and 10038202 for premium edition, Siemens, Ultrasound Division, Mountain View, CA 94039 |  |  | July 10, 2007 | 2007 | Siemens Medical Solutions USA, Inc. | Image errors: A software problem results in on-screen indications (graphics) that lead the user to believe that the patient's right and left and the transducer orientation are oriented the same. This is not the case and the actual patient's right and left are reversed from the on-screen indicators. | Software design | The firm sent a recall letter, Urgent Medical Device Notification, to consignees on 07/10/2007, to be followed by a service visit by Siemens personnel for a mandatory software update to be installed. Short term the users were instructed to avoid increasing the SieClear compounding levels to an Advanced SieClear compounding setting after an L/R invert action with Biopsy guidelines on-screen. |  | 418 | 418 units (Z-0297-0298-2008) | Worldwide: Product was distributed to medical facilities Nationwide including the states of OH, FL, IL, NY, CA, NH, TX, NJ, PA, MD, MO, GA, WA, WI, NE, WV, HI, MA, SC, MI, VA, AZ, UT, NH, RI, KY, SD, and IN, and internationally to Canada, Germany, Australia, Sweden, Saudi Arabia, Chile, South Korea, Mexico, France, UK, Taiwan, Singapore, Denmark, Brazil, New Zealand, Ireland, Belgium, Norway, Serbia, Vietnam, UK, Poland, China, Spain, South Africa, Austria, Croatia, Russia, Greece, Trinidad and Tobago, United Arab Emirates, India, Japan, Malaysia, Turkey, Thailand, Norway, Slovakia, Romania, Netherlands, Switzerland, Portugal, Italy, and Venezuela. | Terminated | December 20, 2010 | 1259 | 418 | Computer
Z-0299-2008 | 45397 | No | LOJ | Tumor-associated antigen immunological test system. | Immunology | Immunology | 510(k) | ARCHITECT i System Assay CD-ROM - WW (excluding USA accounts), Versions prior to Version 24.0; for use on the ARCHITECT i System; List Number (LN) 6E59,Abbott Laboratories, Diagnostics Division, Abbott Park, IL 60064 USA |  |  | September 18, 2007 | 2007 | Abbott Laboratories | The lower limit flag for auto-dilution protocol 3 (amniotic fluid) was set too low (15 ng/mL) for the ARCHITECT AFP assay file for the ARCHITECT AFP non-U.S. product. The assay labeling indicates that the limit flag should be set no lower than 20 ng/mL | Software design | Abbott affiliates were e-mailed copies of the Product Correction (Immediate Action Required) Recall Letter on 9/18/07 for translation where needed and dissemination to their ARCHITECT accounts who have the affected assay CD-ROMs. Each Abbott international organization will be responsible for contacting customers and determining the method of effectiveness checks based on their country requirements. Abbott sent Product Correction, Immediate Action Required letters dated 18 September 2007 to all active non-U.S. ARCHITECT customers who have the affected assay CD-ROMs. The end users were informed that current AFP file has been updated on a new ARCHITECT I Assay CD-ROM (LN 6E59, version 24). The Low-Linearity flag threshold for the Amino 1:50 dilution has changed from 15 ng/mL to 20 ng/mL for consistency with overall assay sensitivity. The accounts were advised to install the updated version of the current ARCHITECT AFP assay when they receive the version 24 ARCHITECT I Assay CD-ROM. The accounts were requested to keep a copy of the letter on file. The accounts were requested to complete the customer reply form, indicating that the letter was received, understood and disseminated to the laboratory manager/supervisor responsible for Architect testing, and fax it to Abbott by 9/28/07 at 1-800-777-0051. |  | 2490 | 2,490 | International Only: Australia, Canada, Chile, Costa Rica, Dominican Republic, Germany, Guatemala, Honduras, Hong Kong, Jamaica, Japan, Mexico, New Zealand, Peru, Singapore, South Korea, Thailand, Uruguay and Venezuela. There were no USA accounts. | Terminated | November 05, 2008 | 414 | 2490 | Computer
Z-0592-2008 | 46154 | No | KMI | Bed-patient monitor. | General Hospital | General Hospital | 510(K) Exempt | Hoana LifeBed Patient Vigilance System, (Intelligent Medical Vigilance System),Model Number: 010135-01, Product distributed by Hoana Medical, Inc., 828 Fort St. Mall, Suite 620, |  |  | November 07, 2007 | 2007 | Hoana Medical | Bed Exit Failure: Failure to recognize "Bed Exit" when the feature is active due to a software anomaly. | Software design | On 11/7/07, all consignees were notified via an untitled letter and on-site visit by the firm's sales representative, informing them of the affected product and providing instructions on the recall. The recalling firm will arrange for installation of a software fix at each consignee. |  | 216 | 216 units | Nationwide including states of HI, CA, TX and MA | Terminated | February 01, 2008 | 86 | 216 | Computer
Z-0551-2008 | 45873 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Abbott, ACCELERATOR Device Manager (ADM); List Number: 08H74-01; Abbott Laboratories, Abbott Park, IL |  |  | October 25, 2007 | 2007 | Abbott Laboratories, Inc | Sample/Patient Mis-identification: Two issues identified related to the ACCELERATOR Decision Manager (ADM), to include: 1) Sample Identification (SID) and/or Patient Identification (PID) numbers that contain more than 12 characters are truncated to the final 12 characters by the ADM; 2) for SIDs using alpha characters that are case sensitive, SIDs with lower case characters are transmitted by the | Software design | Firm notified consignees via Product Correction, Immediate Action Required letter(s) on 10/25/07. Consignees were notified of software issues and informed they would be contacted by an Abbott representative to schedule installation of software upgrade. |  | 19 | 19 units. | Nationwide:Distributed to hospitals/laboratories in IL, PA, TN, TX, UT, VA and WV. | Terminated | April 01, 2008 | 159 | 19 | Computer
Z-0677-2008 | 46157 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Siemens Mevatron KD2 Medical Charged-Particle Radiation Therapy System; (Interlock #51 Jaw Field Position) Model Number:08515520, Siemens Medical Solutions, Concord, CA 94520 |  |  | August 07, 2007 | 2007 | Siemens Medical Solutions USA, Inc | Incorrect field size -- Products manufactured prior to 1999 may experience jaw field size errors during interlock resulting in potential mistreatment if a light field check is not performed. This problem is present in machines with serial numbers less than 3094, as they do not contain redundant sensors for jaw position. | Device Design | The firm has sent consignees a Customer Safety Advisory Notice on August 7, 2007 informing them that a misplacement of the X and/or Y jaws can occur if interlock #51 Jaw Field Position is asserted and then cleared with the left-hand key rest. The firm recommends that if an interlock issue is encountered tha treatment of the patient be discontined until this issue is resolved. |  | 55 | 55 units | Worldwide Distribution | Terminated | December 01, 2010 | 1212 | 642 | Computer
Z-0552-2008 | 45888 | No | IWX | Personnel protective shield. | Radiology | Radiology | 510(K) Exempt | AADCO Rayshield Frame Mounted Overhead X-Ray Barrier; |  |  | November 01, 2007 | 2007 | Aadco Medical Inc. | Screws may become loose from the mounted shield compromising stability | Device Design | AADCO Medical Inc., issued a Product Safety Notification/recall letter to Hospital Administrators, Managers of Radiology/Cardiology and Radiologists and Cardiologists dated 11/1/07. AADCO will perform on site inspections with device tightening or replacement of parts as needed; at no cost the the users |  | 324 | 324 units | Worldwide; USA, Canada, Mexico, Brazil, Venezula, Chile, Bermuda, Spain, and Panama | Terminated | May 02, 2012 | 1644 | 324 | Computer
Z-0880-2008 | 46304 | No | GZT | Self-retaining retractor for neurosurgery. | Neurology | Neurology | 510(k) | Comfort T-Handle Hudson Connector with Impactor Cap; Part #1006-1; handle is component of the following consignment surgical kits: Harmony 2700-0001-PL, Harmony 2700-0009-PL, Harmony 2900-0001-PL, Harmony 2900-0002-PL, Cadence 1500-0002-PL and Cadence 1500-0008-PL; Distributed by Abbott Spine, Austin, TX 78727. |  |  | December 19, 2007 | 2007 | Abbott Spine | Dislodged Collar Spring; The collar spring binding on the shaft of generic handle used with multiple surgical instruments has the potential to release the device's internal bearings during use. | Device Design | All medical suppliers with consigned inventory from affected lots were contacted via phone beginning 12/19/07. Consignees notified of potential issue and were asked to inspect their inventory and contact Abbott Spine for return and replacement of the affected handles. |  | 230 | 230 handles. | Worldwide:Medical suppliers in the following states: AR, AZ, CA, CT, FL, IL, LA, MD, MI, MO, MN, NC, NJ, NM, NV, OH, OK, OR, PA, TX, VA and WI. and country of Spain. | Terminated | October 04, 2010 | 1020 | 230 | Computer
Z-0950-2008 | 45595 | No | LKK |  |  | General Hospital | PMA | Medtronic SynchroMed EL Programmable Pumps, Models 8626-10, Medtronic, Inc., 710 Medtronic Parkway, Minneapolis, MN 55432-5604 USA. (The implantable Medtronic SynchroMed EL Programmable Pump is part of the SynchroMed EL Infusion System designed to contain and administer parenteral drugs to a specific site) |  |  | August 03, 2007 | 2007 | Medtronic Neuromodulation | Pump Motor Stall; pumps can stall due to gear shaft wear. | Device Design | Begining August 3, 2007, consignees were notified by letter. The letter "Medtronic Urgent: Medical Device Correction Aug 2007" described the problem and the Affected Devices as well as Patient Risk, Patient Management Recommendations, Next Steps and Physician and Patient Support. |  | 81023 | 81,023 total for recalls Z-0950-0957-2008 | Worldwide-USA including states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, DS, TN, TX, UT, VA, VT, WA, WI, WV, and WY, and countries of HONG KONG, JAPAN, AUSTRALIA, CANADA, BELGIUM, CROATIA, AUSTRIA, CYPRUS, The CZECH REPUBLIC, DENMARK, EGYPT, FINLAND, FRANCE, GERMANY, GREECE, HUNGARY, INDIA, IRAN, ICELAND, IRELAND, ITALY, JORDAN, KWAIT, LEBANON, MALTA, The NETHERLANDS, NORWAY, POLAND, PORTUGAL, SAN MARINO, SAUDI ARABIA, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TUNISIA, TURKEY, UNITED ARAB EMIRATES, The UNITED KINGDOM, CHILE, VENEZUELA, BRAZIL, and MEXICO. | Terminated | December 27, 2011 | 1607 | 81023 | Computer
Z-0676-2008 | 45329 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Philips Medical Systems -- Cardiac Viewer or Pulmonary Viewer software application version 2.0 in use with Precedence SPECT/CT System and all CT EBW workstations using software version 3.5. Philips Nuclear Medicine, Milpitas, CA 95035 |  |  | September 28, 2007 | 2007 | Phillips Nuclear Medicine | Incorrect Measurement: The firm discovered a software defect in Cardiac and Pulmonary Viewer software application options used in the Precedence CT Sub-System. The defect causes incorrect measurements in derived images which are zoomed and saved in a batch file. | Software design | A Customer Advisory Notification was mailed to affected consignees on September 28, 2007. The firm informed consignees that incorrect measurement can occur when the Cardiac Viewer or Pulmonary Viewer is used to zoom and save derived images in batch files. The firm provided instructions for operators to prevent incorrect measurements. Phillips is currently investigating methods to correct this issue and will provide updates, free of charge, when they become available. If you have questions concerning this recall, please contact Phillips Medical at 1-800-722-9377, (option5: Diagnostic Imaging, option 5:NM). |  | 19 | 19 | Worldwide Distribution --- USA including states of FL, NY, VT, MD, TX, and MO and country of Canada. | Terminated | February 04, 2009 | 495 | 19 | Computer
Z-0518-2008 | 45986 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Magnetom Espree with Swiveling OR (Operating Room) Table, Magnetic Resonance Imaging System,Siemens Medical Solutions USA, Inc., Malvern, PA |  |  | November 05, 2007 | 2007 | Siemens Medical Solutions USA, Inc | Table may experience a deadlock situation | Software design | The recalling firm issued a Customer Safety Advisory to affected customers per Update Instructions MR031/07/S. The letter informs customers of the potential issue and provides instructions to avoid its occurrence. The recalling firm has also released the software update to correct the issue via Update Instructions MR034/07/S. Notifications were issued on or about November 5, 2007. |  | 4 | 4 units | The products were shipped to medical facilities in TX and NY. | Terminated | February 05, 2008 | 92 | 4 | Computer
Z-0739-2008 | 36595 | No | LKK |  |  | General Hospital | PMA | Medtronic SynchroMed EL Programmable Pump, Model 8626-10. Medtronic, Inc., 710 Medtronic Parkway, Minneapolis, MN 55432-5604 USA.-- (The implantable Medtronic SynchroMed EL Programmable Pump is part of the SynchroMed EL Infusion System designed to contain and administer parenteral drugs to a specific site) |  |  | August 03, 2007 | 2007 | Medtronic Neuromodulation | Pump motor stall due to gear shaft wear | Device Design | Beginning August 3, 2007, consignees were notified by letter. The letter "Medtronic Urgent: Medical Device Correction Aug 2007" described the problem and the Affected Devices as well as Patient Risk, Patient Management Recommendations, Next Steps and Physician and Patient Support. |  | 23 | 23, 895 total for recalls Z-0739-0746-2008 | Worldwide: USA including states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY and countries of HONG KONG, AUSTRALIA, CANADA, AUSTRIA, BELGIUM, CROATIA, The CZECH REPUBLIC, DENMARK, EGYPT, FINLAND, FRANCE, GERMANY, HUNGARY, IRAN, ICELAND, IRELAND, ITALY, THE NETHERLANDS, NORWAY, POTUGAL, SAN MARINO, SAUDI ARABIA, SPAIN, SWEDEN, SWITERLAND, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, IRAQ, BRAZIL, and MEXICO. | Terminated | June 06, 2011 | 1403 | 23 | Computer
Z-0280-2008 | 45454 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare CIC Pro Clinical Information Center, Central Station, Software Version 4.1. (CIC Pro v4.1). GE Medical Systems Information Technologies, Inc., Milwaukee, Wisconsin (System hardware consists of a processing unit and a monitor) |  |  | January 24, 2007 | 2007 | General Electric Medical Systems Information Technology | Two separate issues may occur with operating The CiC Pro v4.1 System with Telemetry. 1) Experience a loss of audible alarms when operating system memory resource reaches approximately 170MG. This will occur if CIC Pro continuously runs for average of 160 day, and 2) System may continuously reboot upon start up in an attempt to correct the issue. | Software design | GE Healthcare sent a certified letter titled "Urgent Medical Device Correction" to customers on 3/13/07. As a Short Term Solution, GE recommends that user reboot of system according to instructions and precautions in Service Manual at least once every three months. Also, take the necessary precautions to ensure that system is shut down properly. This will decrease the likelihood of a continuous reboot. If system continues to reboot, GE suggest that user contact Technical Support at 1-800-588-7044 (Press Option 2) and ask for Technical Support. As for a Long Term Solution, GE will provide a software update that will rectify the issue as soon as an upgrade is available. |  | 6217 | 6,217 Units (1,549 OUS, 4,668 within US) | Worldwide Distribution - USA (excluding DE, HI, ID, MT, ND, SD, VT), Peurto Rico, Washington DC and including the countries of Australia, Austria, Belgium, Canada, Czech Republic, Denmark, Finland, France, Germany, Greece, Hungaria, India, Indonesia, Ireland, Israel, Italy, Japan, Korea, Kuwait, Malaysia, Malta, Netherland, New Zealand, Norway, Pakistan, Philippines, Poland, Qatar, Russia, Saudi Arabia, Singapore, Slovakia, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, and United Kingdom | Terminated | May 16, 2012 | 1939 | 6217 | Computer
Z-0690-2008 | 45845 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Innova 2100IQ Versatile Cardiac and Vascular Imaging System, GE Healthcare, Waukesha, WI 53188 ( The system consists of an a monoplane positioner, a vascular or cardiac table, an X-ray system and a digital detector) |  |  | November 15, 2007 | 2007 | GE Healthcare | Frozen Image Problems: During an acquisition (fluoro and/or record), there have been cases reported where an image became "frozen" on the DL (digital leader acquisition system) live monitor screen. In such cases, the system continued to send out X-rays without reporting an error message. The result was that the live imaging screen displayed an older "frozen" image until the operator released the | Device Design | Customers were contacted by means of an Urgent Medical Device Correction Letter on November 9, 2007. The letter described the product and the safety issue. The letter also made short term recommendations and described a long term solution, which is pending as an expected hardware and/or software modification that will be made at no cost. Short Term Recommendation: Take extra care when working with motionless anatomy or devices where there is little evident motion to warn you of the frozen image occurrence. Always completely shut down and restart the system (not just reset) immediately after such an error occurs. Also, please immediately report this issue to your local GE Healthcare field service engineer. |  | 378 | 378 | Worldwide-USA including states of AR, CA, CO, FL, GA, IL, IN, LA, MA, MD, MO, NC, NJ, NM, NV, SC, TN, TX, WA, WI, AK, AL, AZ, ,CT, DE, HI, ID, KS, KY, ME, MI, MN, MS, MT, ND, NE, NH, NJ, NY, OH, OK, OR, PA, SD, UT, VA, WV, and Puerto Rico and countries of Hungary, Korea, Algeria, Argentina, Australia, Belarus, Belgium, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, Chile, China, Columbia, Czech Republic, Dominican Republic, Finland, France, Georgia, Germany, Great Britain, Greece, Hong Kong, India, Indonesia, Ireland, Italy, Japan, Kuwait, Latvia, Lebanon, Lithuania, Malaysia, Mexico, Morocco, New Zealand, Norway, Panama, Poland, Portugal, Russia, Saudi Arabia, Singapore, Spain, Sudan, Sweden, Switzerland, Syria, Taiwan, Turkey, and Venezuela. | Terminated | January 29, 2016 | 2997 | 869 | Computer
Z-0660-2008 | 44965 | No | CBR | Nitrous oxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Spacelabls Medical, Ultraview Multigas Analyzer, Model 91518,Spacelabs Healthcare, Issaquah, WA, 98027 |  |  | October 08, 2007 | 2007 | Spacelabs Healthcare, Incorporated | Inaccurate partial pressure readings; readings are not compensated for ambient pressure in mmHg or kPa. | Software design | The firm sent out their Urgent Medical Device correction/recall letters on 10/5/2007, and an updated versions on 11/30/07 and 12/7/2007, forwarding an attached set of work around directions for use and associated tables for calculations. This short term fix will be updated to a permanent fix when a new software revision is released and installed at no cost to each user as noted in the recall notification. Telephone contact was provided at 1-800-522-7025. |  | 722 | 722 devices | Worldwide | Terminated | April 12, 2011 | 1282 | 722 | Computer
Z-0462-2008 | 44934 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Philips Medical Systems, GEMINI GXL 6, 10, &amp; 16 slice PET/CT Systems, diagnostic imaging systems for fixed or mobile installations that combine Positron Emission Tomography (PET) and X-ray Computed Tomography (CT) systems, Philips Medical Systems, Cleveland, OH 44143 |  |  | September 21, 2007 | 2007 | Philips Medical Systems (Cleveland) Inc | Artifacts: Artifacts may appear in some images from head scans on the CT subsystem of the GEMINI PET/CT system. Reading images with these artifacts may lead to misdiagnosis or improper treatment. | Software design | Philips Medical notified domestic customers via Product Safety Notification dated May 30, 2007. The product safety notification includes a Technical Bulletin notifying customers of the issue and how to avoid it. This information will be used until the software correction becomes available. The software correction will be released in October 2008. Once software is available, the firm will begin correcting the field. |  | 126 | 126 | Worldwide Distribution - USA, Australia, Belgium, Canada, China, Denmark, England, Finland, France, Germany, Hong Kong, Hungary, India, Ireland, Italy, Japan, Jordan, Korea, Netherlands, Poland, Portugal, Saudi Arabia, Slovenia, Spain, Sweden, Taiwan, Turkey, and Venezuela. | Terminated | March 02, 2012 | 1624 | 187 | Computer
Z-0124-2008 | 44824 | No | DPS | Electrocardiograph. | Cardiovascular | Cardiovascular | 510(k) | Philips PageWriter Touch Cardiograph, |  |  | September 12, 2007 | 2007 | Philips Medical Systems | Incorrect display: When using the 16 lead Patient Interface Module (PIM), lead tracings of V3R and V4R are reversed in the extended lead modes of Pediatric and Balanced. | Software change control | Philips Medical Systems notified affected customers by Customer Letter and Urgent - Device Correction Notice dated 9/12/07. The correction will be a software upgrade. In the US, the Philips Response Center will contact US customers and arrange their upgrades, either via a software download, or by sending a Field Service Engineer to the customer site to perform the upgrade. Until new software is installed, customers are being advised to use the 12 lead module, and to disable or disconnect the 16 lead connection. |  | 266 | 266 units | Worldwide Distribution - USA, Canada, Australia, Hungary, New Zealand, Nigeria, Norway, Qatar, Taiwan, and UAE. | Open, Classified |  |  | 266 | Computer
Z-0708-2008 | 46142 | No | HCI | Aneurysm clip applier. | Neurology | Neurology | 510(k) | Caspar Titanium Aneurysm Clip Appliers, Catalog No/Part #: FT420T, Casper Applier F/Mini Ti-Clipsstr 146 mm, Aesculap, Inc., Center Valley, PA. 18034 |  |  | November 01, 2007 | 2007 | Aesculap Inc | Breakage: the inner rod of the applier may corrode and break | Device Design | The recalling firm issued an Import Recall Notice, on or about December 18, 2007, to all their customers to inform them of the problem and the need to return the devices . |  | 60 | 60 clips | Worldwide- USA hospitals in AZ, CA, CO, IL, KY, MI, MS, ND, NM, NY, OH, OK, PA, TX, VA, WA, and WI, to sales representatives in NC, OK, and MO and shipped to distributors in CA and Canada. | Terminated | March 28, 2008 | 148 | 144 | Computer
Z-0568-2008 | 45958 | No | LYJ |  |  | Neurology | PMA | Cyberonics, VNS Therapy System Model 250 Handheld programming system using Dell x5 handheld and preprogrammed with software v7.1 and higher; manufactured by Cyberonics, Inc., Houston, TX. |  |  | October 25, 2007 | 2007 | Cyberonics, Inc | Screen Freezes-- The Dell X5 Handheld PC screen will freeze caused due to incompatibility between the Microsoft 2002 OS and the model Dell X5 handheld computer. Once frozen, the handheld device becomes non-responsive to user input. | Software design | Cyberonics sent an initial notification (Important Safety Alert) to physicians via Certified Mail on 10/01/07. Letter notified physicians that the Dell X5 handheld computer screen will freeze following an interrogation. Once frozen the handheld device becomes non-responsive to user inputs. The firm informed the user that in the event that the computer becomes non-responsive, the user has two options, 1) Reinitialized the flash memory and, 2) Perform a software reset of the handheld computer. A second notification addressing the approval and addition of labeling was sent to physicians on 11/15/07. Physicians were asked to respond to both letters via an enclosed response card, acknowledging receipt of the notification. Additional information will be posted at www.VNStherapy.com. |  | 4201 | 4,201 (US) 199 (International) | Worldwide Distribution- USA and the countries of; Austria, Belgium, Denmark, France, Germany, Ile de la Reunion, Ireland, Israel, Italy, Japan, Marocco, Mexico, Netherlands, Ireland, Norway, Poland, Portugal, S. Korea, Slovenia, Sweden, Switzerland, Turkey, United Kingdom, and Venezuela. | Terminated | October 05, 2009 | 711 | 4201 | Computer
Z-0554-2008 | 45519 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Real-Time Position Management System (RPM), A Respiratory Gating System; Version 1.7.3, All Models; Varian Medical Systems, Palo Alto, CA 94304 |  |  | September 26, 2007 | 2007 | Varian Medical Systems Inc | Erroneously Resetting: RPM System version 1.7.3, software failure can occur that affects the gated radiotherapy treatment when phase-based gating is used. RPM 1.7 erroneously resets the phase-based gating thresholds to system default during the treatment session, if the user accesses the Session options during the treatment session. This can lead to difference between the planned and delivered | Software design | Varian sent an Urgent Medical Device Correction letter on September 26, 2007, to all its direct consignees, informing them that the RPM version 1.7 may reset the phase thresholds if Sessions options are accessed during treatment. A Service Representative will be scheduling a visit to user site to perform the modification requested to address the problem, when a corrected software version is available. |  | 59 | 59 units | Worldwide including countries of: Ireland, Italy, Spain, UK, Denmark, Finland, France, and Canada | Terminated | September 30, 2009 | 735 | 59 | Computer
Z-0478-2008 | 45513 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Hallmark Anterior Cervical Plate System: Primary Bone Screw, Semi Constrained, REF # 62-3140,14mm; Blackstone Medical, 90 Brookdale Drive, Springfield, MA 01104 (The Hallmark Anterior Cervical Plate Systems consists of an assortment of implantable titanium alloy plates and bone screws that are provided non-sterile) |  |  | May 08, 2006 | 2006 | Blackstone Medical, Inc. | Incomplete Design Validation: the laser marking on the head of the screw lacked complete process qualification | Device Design | On May 8, 2006, a communication to the field was sent via phone, e-mail and fax. Instruction was made to retrieve and return all product in the field. |  | 1991 | 1,991 semi-constrained primary and rescue screws. | Nationwide | Terminated | November 06, 2008 | 913 | 1991 | Computer
Z-0527-2008 | 45856 | Yes | KPF | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Baxter RenalSoft v.2.0 Patient Management Software Suite - HD Module; clinical data management software; Baxter Healthcare Corporation, McGaw Park, Illinois 60085 |  |  | November 14, 2007 | 2007 | Baxter Healthcare Renal Div | Pre-Treatment/Post-Treatment Report errors: 1) The Heparin Bolus Value is displayed on Pre and Post Treatment Reports for patients with "No Heparin" orders (RenalSoft v2.0, Renal Software Suite v.3.0 and v.3.1), and 2) Withheld Medications are incorrectly displayed as Administered on the Pre-Treatment report (RenalSoft v.2.0 only). | Software design | Baxter Healthcare Corporation sent an Urgent Device Correction letter, dated 11/14/07, to all of the affected customers, informing them of the Heparin Bolus and the Withheld Medication issues and provided them with interim mitigation instructions that they should implement while waiting for a software correction. A Baxter representative will contact the customer to schedule the software update when the upgraded software is released. Customers were requested to complete and return the enclosed Customer Reply Form, indicating that they have disseminated the information to their staff, and indicating if they still have RenalSoft software and require an upgrade. Any questions were directed to Baxter Global Technical Services at 1-800-553-6898, option 4.A follow-up letter dated 3/6/08 was sent to the consignees on the same date, informing them that the upgraded RenalSoft Software v.2.0.1 was now available for installation. The consignees were requested to complete and fax back the customer reply form indicating that they received the letter, and were informed that they would be contacted by a Baxter representative to schedule an installation of the software upgrade at no charge. |  | 25 | 25 units | Worldwide Distribution-USA including Puerto Rico, and countries of Australia, Austria, Canada, Chile, Denmark, Finland, Greece, Ireland, Italy, Mexico, The Netherlands, New Zealand, Norway, Panama, Portugal, Spain, Sweden and the United Kingdom | Terminated | February 03, 2010 | 812 | 34 | Computer
Z-1105-2008 | 46331 | No | HWA | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Triathlon Baseplate Impactor Extractor, Orthopedic manual surgical instrument, Catalog Number 6541-4-805, Non-Sterile, Stryker Howmedica Osteonics Corp, Mahwah, NJ 07430 |  |  | February 23, 2005 | 2005 | Stryker Howmedica Osteonics Corp. | Insertion/Removal Difficulties: The Baseplate Impactor/Extractor may not assemble/disassemble easily to the baseplate. | Device Design | Consignees were notified via an Important Market Withdrawal Letter sent 2/25/05. The letter instructed users to fax a Product Accountability Form to the recalling firm identifying any affected products and to return them to the firm. The letter informed users that the products will be reworked and returned at a later date. The letter also identified other products that could be used as replacements. For additional information, contact 201-831-5825 or 201-831-5110. |  | 293 | 293 total units | Worldwide Distribution | Terminated | May 30, 2008 | 1192 | 293 | Computer
Z-1461-2008 | 47401 | No | HNR | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | Alcon/Grieshaber Morris ILM Forceps, Direct Action; Manual ophthalmic surgical instrument, Catalog Number: 641.22; Non-sterile/Reusable; Alcon Laboratories, Inc., Fort Worth, TX 76134. |  |  | March 07, 2008 | 2008 | Alcon Research, Ltd | Device corrosion: Forceps are prone to corrosion and possibly premature fracture and/or malfunction. | Device Design | Consignees were notified by an Urgent Medical Device Recall that was sent via fax on 03/07/08. Users were instructed to cease using the affected products and to return a Recall Acknowledgement Form to the recalling firm. The sales representative will pick up the affected product and discuss replacement alternatives. For additional information, contact 1-800-757-9780. |  | 24 | 24 units | Nationwide Distribution including states of AL and CA. | Terminated | July 18, 2008 | 133 | 24 | Computer
Z-0088-2009 | 49490 | No | IOQ | Powered flotation therapy bed. | Physical Medicine | Physical Medicine | 510(K) Exempt | Envision E700 Low Airloss Therapy Surface, Made By Hill-Rom Manufacturing, Inc., Charleston, SC. |  |  | June 19, 2008 | 2008 | Hill-Rom Manufacturing, Inc. | A defect in the software of the device may not allow the patient bed exam alarm or the patient movement alarm to function correctly. | Software change control | An internal Modification Notice dated 9/1/08 was sent via email to Service Centers (because the device is owned by the manufacturer, Hill-Rom, each Service Center in possession of the device will be the consignee). Each Service Center will perform the correction on the units affected. Each device in the affected serial number range will be tested for the software defect. If the defect is found, the circuit board that contains the defective software will be replaced. The Field Action Plan is expected to be completed on July 24, 2008. Expected completion date of the Correction will be 12 months from July, 2008. |  | 779 | 779 units | Worldwide Distribution --- including USA and country of Canada. | Terminated | May 21, 2009 | 336 | 779 | Computer
Z-1186-2008 | 46345 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Biomet brand Modular Microplasty Cup Impactor, 3/8" - 24 Thread Insert; Model 31-400603. Instrument for hip prosthesis acetabular cup component. |  |  | January 16, 2008 | 2008 | Biomet, Inc. | The weld at the lock location may fracture during impaction. | Device Design | Biomet notified distributors and medical facilities by letter dated 1/16/08. Distributors with distributor-owned instruments at medical facilities were notified by email on 1/17/08, and provided with a letter to present to their associated medical facilities describing the recall. |  | 59 | 59 | Worldwide Distribution: USA, Germany, Mexico, Netherlands, Sweden and United Kingdom. | Terminated | September 22, 2009 | 615 | 1004 | Computer
Z-1615-2008 | 46965 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | TomoTherapy Hi-Art System with software versions 3.0 and 3.1 - Radiation Therapy System. The device is intended to be used as an integrated system for the planning and delivery of intensity modulated radiation therapy. |  |  | February 22, 2008 | 2008 | TomoTherapy Incorporated | Registration adjustment field numbers may appear on printed pages of the Register Tab as nonsense characters or values different than the numbers shown on the Operator Station (OS) display. Setup will move the couch to the correct offset position. If the couch is manually adjusted to the incorrect value, the patient may not be the correct position during treatment. | Software design | Consignees were sent a letter "TomoTherapy Incorporated Urgent Medical Device Correction Field Safety Notice" dated February 22,2008. The letter described the Issue, Product Affected, Cause, Recommended Customer Action and Resolution.If you have not received this letter or have questions, contact the firm as outlined below:North America: 1 866 368 4807Belgium: 0800 38783France: 0805 631 565Germany: 0800 000 7541Italy: 800 986 399Netherlands: 0800 0201364Spain: 800 300049Switzerland: 0800 001927United Kingdom: 0808 238 6035China/CNC (Northern): 10 800 712 1701China (Southern): 10 800 120 1701Hong Kong: 800 967912Japan: 0044 22 132374Singapore: 800 1204 683South Korea: 0079 81 4800 7204All other locations: +1 608 824 2900 or +32 2 400 44 44 |  | 32 | 32 | Worldwide Distribution --- USA including states of CA ,CT, DC, FL, IL, LA, MD, MI, MO, NE, NJ, NY, SC, TN, VA, WI , and countries of Canada, Germany, Taiwan, and Turkey. | Terminated | December 02, 2008 | 284 | 32 | Computer
Z-1564-2008 | 46807 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Precision RXi Digital Remote X-Ray Imaging R&F; System. It is a remote R&F; system consisting of a tilting table with integrated spot film device, 65 or 89 kW pulsing generator and high heat capacity X-Ray tube. The system can be configured to meet a wide variety of digital record and fluoroscopy exams for various clinical applications with either 32 cm or 40 cm image intensifier. |  |  | December 18, 2007 | 2007 | GE Healthcare | GE Healthcare Precision RXi tilt joystick became stuck in the Trendelenburg direction causing the patient to slide off the table. Also recall conducted to investigate the possibility of accidentally activation of the remote console joystick where a fatal accident occurred on a non-GE Remote R&F; System. . | Device Design | Two letters "GE Product Safety Notification" dated 12/18/07 and 3/7/08, addressed to Hospital Administrators, Manager of Radiology and Radiologists, were sent to each consignee. Each letter addressed a separate safety issue. Letters described the affected product, safety instructions, corrective action and contact information. The joystick can become stuck causing unintentional motion of the tube.Field Modification Instructions (FMI 10802/10803) have been issued to the field to install additional warning labels on the patient table, install a protector for the console joystick and inspect and correct the clearance between the console joystick and its mounting location.If your facility has a Precision RXi or Precision e system manufactured before March 2007, a GE Healthcare Field Engineer will schedule a field visit to install the added joystick protection. This upgrade will be performed at no charge.If you have any questions/concerns regarding this issue, please contact your local GE Healthcare sales or service representative. |  | 193 | 193 | (USA)AKCOILLAMAMIMSNCNJNYPASCTNUTVAWAWI(OUSA)BR CA MX AU C3 KR PH TW AE AT BE BF BG CM CY CZ DE ES FI FR GB GE GR GY HR HU IT LB LV MT NA NG PL PT RO RU SA SD SE SI SK SY TN TR UA YU ZA | Terminated | August 09, 2012 | 1696 | 265 | Computer
Z-1234-2008 | 46383 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Nautica (GK) Mattress Stretcher, manufactured by Georg Kramer GES.M.B.H. &amp; CO. KG A-4061 PASCH ING, Neuhauserweg 14 Austria. These mattress stretchers were sold for installation on or as a component of the following patient bath lifts: a) Manufactured by BHM Medical Inc., Magog, QC, Canada J1X 5Y5: Part number 88004051, Nautica Bath Lift System; b) Distributed by Whitehall Manufacturing, P.O. Box 3527, City of Industry, CA 91744, Manufactured by BHM Medical Inc., Canada: Part number 88004051.12, Nautica Bath Lift System |  |  | January 23, 2007 | 2007 | Arjo, Inc. | Repeated use of the ratchets that lock the position of the back and leg rest of the stretcher may damage the retention pin that prevents the latch from disengaging from the anchor system. This would allow the backrest to pivot freely down to the floor. | Device Design | BHM Medical initially notified the end users on 1/23/07 of the potential for the back and leg rest latch ratchets to become damaged, allowing the rests to pivot freely. Should a patient be in the stretcher at the time, the patient could fall, potentially sustaining cuts, bruises or even more severe injuries. The letters provided safety instructions to follow until a permanent fix is available, and requested the accounts to complete and return the enclosed response form acknowledging receipt of the letter and dissemination of the safety instructions to all employees operating the lifts with the affected GK stretchers.BHM sent Customer Field Correction Notification letters dated 1/23/08 to the accounts, along with the Technical Advisory Notice - TAN A dated 1/3/08, reiterating the safety instructions from the 1/23/07 letter and requesting the accounts to, if at all possible, remove the equipment from service until an authorized engineer can replace the ratchets in the backrest and leg rest latches.All consignees have been contacted. |  | 7 | 7 units | Nationwide Distribution --- including states of California, Connecticut, Florida, Mississippi and Texas. | Terminated | August 25, 2008 | 580 | 7 | Computer
Z-1532-2008 | 46908 | No | ITI | Powered wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Quickie Groove Power Wheelchair with QC (non-expandable) electronics, Model number EIPW13. |  |  | January 11, 2008 | 2008 | Sunrise Medical Inc | Product experiences a longer than expected stopping distance, which may cause injury to the user, objects, or the chair itself. | Software design | Initial phone contact began on January 11, 2008. Medical device correction letters were also sent to affected dealers. |  | 141 | 141 units total | Product was distributed to 126 retailers nationwide throughout the USA, as well as Canada and Australia. | Terminated | November 16, 2009 | 675 | 141 | Computer
Z-1541-2008 | 46961 | No | DXJ | Medical cathode-ray tube display. | Cardiovascular | Cardiovascular | 510(k) | IntelliVue Clinical Information Portfolio |  |  | March 04, 2008 | 2008 | Philips Medical Systems | Medications prescribed for one patient were printed on the record of another patient. | Software design | Philips Medical issued letter notification, dated 29 Feb 2008 to customers titled; "Urgent Medical Device Correction Notice".Customers were asked to not solely rely on the Worklist to remind them when to administer orders or interventions until the software has been upgraded. Customers should review the reminders and verify orders and interventions on the Medication Administration Record and Administration record. The correction will be in the form of a software upgrade and will be performed remotely by Philips service personnel. |  | 138 | 138 for products | NationwideCanada, Australia, Austria, Belgium, Denmark., France, Finland, Germany, Great Britain, Italy, Malaysia, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Taiwan, and UAE | Open, Classified |  |  | 138 | Computer
Z-2139-2008 | 47801 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | CELL-DYN Sapphire Hematology Analyzer, List number 08H00-01, LIS/ Middleware Interface Specification, Abbott Diagnostics Division, Santa Clara, CA |  |  | April 11, 2008 | 2008 | Abbott Laboratories | Table in LIS Specification depicting the association between Record ID and Numerical Result Label is incorrect. If a record ID is used to configure the system for mapping, results will come out nonsensical. | Software change control | Recall initiated on April 11, 2008. A Product Correction letter and customer reply form were sent to all currently active CELL-DYN Sapphire Customers. The letter requests that users be certain that the transmission of test results receipt by the host interface is mapped according to the Numerical Result Label, when configuring their CELL-DYN Sapphire Hematology Analyzer for results transmission to a Clinical Laboratory Information System or Middleware. Affected product was distributed to a total of 4 distributors. However, notification was made to all 154 customers that received the CELL-DYNE Sapphire product. If the customer or the health care providers that they serve have any questions regarding this information, U.S. customers should call Customer Support at 1-877-4ABBOTT. |  | 15 | 15 copies of the specification were sent worldwide | Nationwide Distribution. | Terminated | March 12, 2010 | 700 | 15 | Computer
Z-1514-2008 | 46297 | Yes | LWP |  |  | Cardiovascular | PMA | VITALITY EL, models T127, T167, T177 Implantable Cardioverter Defibrillator (ICD), pulse-generator, dual Chamber |  |  | January 07, 2008 | 2008 | Boston Scientific CRM Corp | Repetitive mechanical stress applied to the titanium case can induce component damage and device malfunction in certain ICD and CRT-D devices implanted subpectorally with the serial number facing the ribs. Devices implanted subcutaneously or in a subpectoral position with the serial number facing away from the ribs are not included in this advisory. This update identifies additional VITALITY ICD | Device Design | An Urgent Medical Device Information Update to May 2006 Product Advisory, dated January 04, 2008, was sent to implanting physicians and physicians following implanted devices beginning Monday, January 07, 2008. The letter described a malfunction that can occur when the ICD is implanted subpectorally with the serial number facing the ribs. Devices implanted subcutaneously or in a subpectoral position with the serial number facing away from the ribs are not included in this advisory. The physicians were asked to review records to determine if an affected device was implanted subpectorally. An AP radiograph should be done to determine specific device orientation for subpectoral implants. If a subpectoral implant is in a susceptible orientation, consider repositioning or replacement for physically active patients who regularly need device therapy. Follow patients with susceptible devices at 3-month intervals. |  | 25000 | BSC CRM estimates that the number of devices implanted in a susceptible orientation is likely less than 1% of the total active implants 25,000 (w.w.) | Nationwide and Worldwide: Bolivia, Panama, Iran, Saudi Arabia, Kuwait, United Arab Rep., Austria, Belgium, Cyrus, Czech Republic, Denmark, Finland, France, Germany, Greece, Iceland, Ireland, Italy, Netherlands, Norway, Portugal, Spain, Sweden, Switzerland, UK, India, Australia, New Zealand, Malaysia, Singapore, Hong Kong, China, Israel, Turkey, Thailand, and Canada. | Terminated | December 30, 2008 | 358 | 25000 | Computer
Z-1769-2008 | 47773 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Medical charged-particle radiation therapy system. Siemens Linear Medical Accelerators equipped with COHERENCE Therapist R2.0 and Oncologist R2.0. Product update, TH 030//07/S. The COHERENCE Therapist R2.0 software numbers are: 7358299, 7360857, and 8141835 and applies only to those parts used with the following devices: Coherence AG Therapist; Coherence Therapist System; Coherence Impression Therapist; PRIMEVIEW 3i; AG Therapist 3rd party V&R; Coherence Therapist 2.0 and PRIMEVIEW 3i 2.0. |  |  | October 03, 2007 | 2007 | Siemens Medical Solutions USA, Inc | Software. The firm issued an Advisory Letter Update instruction TH030/07/S as a result of a complaint that images viewed in the device may shift, which could result in mistreatment. The advisory alerts all its affected customers of the image issue to prevent potential mistreatment. | Software design | Customer Safety Advisory Notice was issued on 10/03/2007 to consignees, via certified mail. Consignees are strongly cautioned that they should not further correct the image alignment at this time as doing so will result in incorrect offset values which could lead to incorrect repositioning of the patient. Consignees are instructed to check the "Filter Images on Loading" and "Saving of Filtered Images" config settings, and apply any additional filters to the image BEFORE using the interactive shift or landmark registration tools. Do not correct any perceived issues. A software fix is being validated and will be available soon. |  |  |  | Class II Recall - Worldwide Distribution --- including USA and countries of Angola, Australia, Brazil, Canada, Croatia, Denmark, Egypt, France, Germany, Greece, Hungary, India, Iran, Italy, Japan, Malaysia, Netherlands, New Zealand, P.R. China, Poland, Portugal, Republic Korea, Russian, Fed. Saudi Arabia, Slovakia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Thailand, Trinidad, Tobago, Turkey, United States, Vietnam. | Terminated | April 29, 2011 | 1304 |  | Computer
Z-1138-2008 | 46373 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo US Workplace Picture Archiving and Communication System, software version 2.0, Siemens Medical Solutions, Inc., Mountain View, CA 94043 |  |  | December 06, 2007 | 2007 | Siemens Medical Solutions USA, Inc. | Inaccurate results: A software bug may cause inaccurate wall motion abnormality scoring results to be displayed. | Software design | Siemens Medical Solutions sent an Urgent Medical Device Notification letter, dated November 2007, to its consignees. The letter described action required by users, and that Siemens will contact the user with a new software version correcting this error. A field correction to update software is planned to be underway, designated to start on 12/21/2007, and the correction is expected to be completed within 45 days. |  | 17 | 17 units | Worldwide Distribution: South Korea, Malaysia, Lithuania and China, and 1 unit to New Jersey in the USA. | Terminated | December 20, 2010 | 1110 | 17 | Computer
Z-1695-2008 | 47824 | No | LZS |  |  | Ophthalmic | PMA | Excimer Laser System; WaveScan WaveFront System, Software Version 3.65, for use with the VISX STAR S4 and STAR S4 IR Excimer Laser Systems. System is a diagnostic instrument intended for the automated measurement, analysis and recording of refractive errors and wavefront aberrations to the target shape and then generates the commands for the VISX STAR laser to create the target shape on the cones. |  |  | November 17, 2006 | 2006 | VISX INCORPORATED, A SUBSIDIARY OF AMO INC | Firm became aware of a significant overcorrection in one eye of a patient who had received a wavefront-guided LASIK treatment at a facility in Germany. The probability of a recurrence of this error was determined to be extremely remote. To ensure that this type of incident will not happen again, a software update was developed. | Software design | VISX issued a notification letter on November 27, 2006, to all distributors and direct accounts. The letter informed the recipients that they were performing a field correction in the form of a mandatory upgrade on all VISX WaveScan WaveFront Systems and Offline Programming Modules to correct an issue. The field upgrade should be made by February 16, 2007. A Dear Doctor letter was also issued instructing surgeons to upgrade all Offline Programming Modules using an enclosed CD. |  | 1310 | 1310 units | Class 2 Recall - Worldwide Distribution: USA, Venezuela, Colombia, Italy, Saudi Arabia, Kuwait, Egypt, UAE, Jordan, Lebanon, Australia, Canada, France, Germany, India, Ireland, Japan, Sweden, Vietnam, Czech Republic, Slovak Republic, Bulgaria, Philippines, South Korea, Turkey, UK, Portugal, Spain, Argentina, Israel, Cyprus, India, Mexico, Greece, Singapore, Chile, Poland, Brazil, China, Thailand, Taiwan, Hong Kong and Russia. | Terminated | June 08, 2010 | 1299 | 1310 | Computer
Z-1478-2008 | 46781 | No | CCX | Patient position support. | Anesthesiology | Anesthesiology | 510(K) Exempt | Allen Medical C-Flex, Polar Head Positioner |  |  | November 27, 2007 | 2007 | Allen Medical Systems, Inc. | Handle jamming in the open position, preventing the device from adequately supporting the head during spinal surgery. Possible delay in surgery, or injury. | Device Design | Allen Medical notified accounts by telephone on 11/08/07 and the Sales Representative. The Hospitals confirmed that the device would be sent back for evaluation and upgraded. |  | 2 | 2 | MA, MS | Terminated | August 26, 2008 | 273 | 2 | Computer
Z-1184-2008 | 46336 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Beckman Coulter UniCel Dxl 800 Access Immunoassay System; |  |  | October 08, 2007 | 2007 | Beckman Coulter Inc | Incorrect results: There is an increased variance in assay performance when the UniCel Dxl 800 Modification (MOD ID 10027)and/or UniCel Dxl 800 Modification (MOD ID 6733) have been implemented and there are room fluctuations after assay calibration between the extremes of the 18ÂżC and 32ÂżC operating range. | Device Design | Beckman Coulter mailed a Product Corrective Action (PCA) letter the week ofOctober 8, 2007 to the customers who have the affected UniCel Dxl 800Access lmmunoassay systems advising that there is an increased variance in assay performance when the UniCel Dxl 800 Modification (MOD ID 10027)and/or UniCel Dxl 800 Modification (MOD ID 6733) have been implemented and there are room fluctuations after assay calibration between the extremes of the 18ÂżC and 32ÂżC operating range. The letter further advises that Beckman Coulter will contact the customer to arrange a date to remove the modificatlons but until the modification is completed, the system should be operated in aroom temperature of 20ÂżC to 30ÂżC. This can be acomplished by monitoringand recording room tamperature in the area of the instrument. Customers are also asked to review historical quality control records to ensure assay performance is within their laboratory specification and that patient results should be reviewed in context of other diagnostics tests and the clinical presentations of the patient. A response form is included. |  | 120 | 111 units in the US; 9 units in Canada | Nationwide and Canada | Terminated | January 11, 2012 | 1556 | 120 | Computer
Z-1378-2008 | 46938 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Radius Capspin is a non-sterile instrument for use with other devices in spinal surgery. |  |  | February 19, 2008 | 2008 | Stryker Spine | The Radius Capsin has separation of the tip from the inserter. | Device Design | Recall notification was sent via FedEx on February 19, 2008 to all branches/agencies. All questions are to be directed to Michelle Klein, Regulatory Compliance Specialist - 201-760-8256. |  | 52 | 52 | Product was distributed to a limited population during a product launch for the purposes of gathering feedback on the product. | Terminated | August 19, 2008 | 182 | 117 | Computer
Z-1322-2008 | 46870 | [6 empty] | Implantable Cardioverter Defibrillator |  |  | January 16, 2008 | 2008 | St Jude Medical CRMD | A condition that could lead to a ventricular sensing anomaly in Epic and Atlas family of lmplantable Cardioverter Defibrillators(ICDs) has been identified. A loss of ventricular sensing would prevent the ICD from being able to detect an arrhythmia. If an arrhythmia occurred and the device was unable to detect it, then a patient would not receive potentially life saving therapy. | Software design | A January 16, 2008, letter entitled, "Important Device Upgrade Information" was sent to physicians by FedEx. In the US, sales representatives will follow-up with the physician and will upgrade all programmers. For the rest of the world, it will be monitored by various St. Jude offices and local distributors with this information being fed back to St. Jude Medical CRMD in Sylmar. In addition, St. Jude Medical issued a press release on January 18, 2008 and information concerning this advisory can be found on St. Jude's website, www.sirn.com. |  |  |  | Nationwide and Internationally: Argentina, Australia, Austria, Barbados, Belgium,Bolivia, Brazil, Canada, Chile, China, Columbia, Costa Rica, Denmark, Dominican Republic,Ecuador, Finland, France, Germany, Hong Kong, Hungary, India, Italy, Japan, Korea, Malaysia,Mexico, Netherlands, New Zealand, Norway, Peru. Poland, Portugal, Republic of Panama,Singapore, Spain, Sweden, Switzerland, Taiwan, Thailand, Trinidad, United Kingdom, Uruguayand Venezuela. | Terminated | April 03, 2012 | 1539 |  | Computer
Z-1370-2008 | 46771 | No | NEP | Intravascular administration set. | General Hospital | General Hospital | 510(K) Exempt | Exacta-Mix 2400, Pharmacy Compounding System is used in the pharmacy to compound multiple source ingredients into on final solution based on user-specified sequences and volumes. |  |  | February 11, 2008 | 2008 | Baxa Corporation | Flush solution from TPN compounder may be added to patient's TPN bag. | Software design | Consignees were notified by letter (Exacta-Mix Compounder Software Safety Alert) on 02/11/2008 advising them of the necessity to install a waste bag prior to a machine flush. |  | 1000 | 1000 units | Nationwide, including military facilities in CA, FL, HI, MS, NC, TX, VA and VA facilities in NY, OR, SC, UT. Foreign distribution to Australia, Canada, China and New Zealand. | Terminated | March 16, 2012 | 1495 | 1221 | Computer
Z-1525-2008 | 38337 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | FACSDiva Software Versions 4.0.1 to 5.0.2, Flow Cytometry Acquisition &amp; Analysis software. Catalog number 641609 |  |  | April 05, 2007 | 2007 | BD Biosciences | When data file containing one or no fluoresence parameters is exported, the software will automatically apply compensation to this file and all subsequently exported files. | Software design | The firm has issued a customer Important Product Information letter via mail on 04/20/2007 to inform them of the software issues. The letter states that to avoid this problem, make sure each tube or well has at least two fluorescence parameters before exporting FCS 3.0 data files. Manually exporting tubes or wells one at a time or restarting the software after the export of an FCS data file containing one or no fluorescence parameters will prevent the problem from occurring. Customers are requested to fill out and return the response form and to contact BD Customer Support Center if there are any additional questions. A software fix is being developed in version 6.0 and a free maintenance release with no product enhancments (version 5.0.3). |  | 1074 | 1074 units | Worldwide Distribution --- USA including states of NY, NM, MI, GA, MD, MA, AL, CA, VA, NC, OH, TX, CO, NJ, PA, MI, CT, IA, MO, NV, KY, IL, TN, FL, MN, FL, UT, WA, OK, CT, WI, SD, AZ, and IN, and countries of Buenos Aires, BE, Brazil, Canada, Chile, China, Hong Kong, Japan, Singapore, Taiwan, Venezuela, Mexico, Trinidad and Tobago, and Columbia. | Terminated | September 30, 2009 | 909 | 1074 | Computer
Z-1640-2008 | 47441 | No | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Logix TPN Software Solution, catalog 2M8410, consists of the Logix Order Entry (OE) Installation CD, version 1.0.28 and Logix Compounder Manager (CM) Installation CD, version 3.1.8. The software is included in a number of Baxter's pharmacy compounders. The software assists the pharmacist in management of compounding, creating and authorizing nutritional solutions intended for IV administration. |  |  | March 28, 2008 | 2008 | Baxter Healthcare Corp. | Software anomaly results in inaccurate information being printed out on the Delivery Report, even though the compounder performed the compounding correctly. | Software design | Baxter sent the 3/28/08 Urgent Device Correction letter to affected customers, to the attention of the Director of Pharmacy, via first class mail to advise them of the potential inaccuracy of the Delivery Report due to a software anomaly. The accounts were instructed to review each delivery report for accuracy. If the report contains zero actual volumes the user must manually reprint the report. The letter outlined the process for reprinting the Delivery Report. Any questions were directed to the Baxter Nutritional Professional Services at 1-800-422-2751. The accounts were requested complete the enclosed reply form acknowledging receipt and understanding of the letter and dissemination of the information to their staff. A software upgrade will be provided.If you have questions about this letter or you did not receive this letter contact Baxter Nutrition Professional Services at 1-800-422-2751. |  | 5 | 5 units | Arizona, Minnesota, Wisconsin, Puerto Rico and Canada | Terminated | August 25, 2008 | 150 | 5 | Computer
Z-1643-2008 | 47524 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syn go Imaging XS is a picture archiving and communication system intended to display, process, read, report, communicate, distribute, store and archive digital radiological images, including digital mammography images. |  |  | March 10, 2008 | 2008 | Siemens Medical Solutions USA, Inc | The firm discovered a problem where the selected patient images on the device may display an additional image from another patient or study. | Software design | The recalling firm issued a Customer Safety Advisory via certified mail to affected customers per Update Instructions IM011/08/S. The letter informs customers of the potential issue and provides instructions to detect its occurrence. A software update to correct the issue is already available. A service representative will visit affected sites to install the software update.If you have questions concerning the advisory letter or you did not receive a letter contact Alexander Vogel at +49 (9131) 84-3840 or alexander.vogel@siemens,com and make reference to complaint 08-1KM-0073 date 2008-03-06. |  | 236 | 236 units | The product was shipped to medical facilities in AL, AZ, CA, CO, CT, FL, GA, IA, KS, KY, MD, MI, MN, MO, NC, ND, NE, NJ, NM, NV, NY, OH, OK, PA, TN, TX, WI, and WV. | Terminated | November 19, 2008 | 254 | 236 | Computer
Z-1703-2008 | 47522 | No | DRF | Electrode recording catheter or electrode recording probe. | Cardiovascular | Cardiovascular | 510(k) | Circular Mapping Catheter-Biosense Webster LASSO 2515 Variable Circular Mapping Catheter- The catheter is indicated for multiple electrode electrophysiological mapping of the cardiac structures of the heart, i.e., recording and stimulation only. The device is designed to obtain electrograms in the atrial regions of the heart. |  |  | March 25, 2008 | 2008 | Biosense Webster Inc | This recall was initiated after the firm received several complaints reporting issues with deflection. The catheter may become caught in the mitral valve, requiring surgical procedures, valve may be torn while trying to dislodge, unable to retract the catheter through the sheath and require surgical intervention to remove, pulling on catheter to remove may lead to large atrial septal tear, and sep | Device Design | A letter entitled, "Urgent Safety Field Notice" dated March 25, 2008, and certification form were sent to all customers and Johnson &amp; Johnson affiliates who purchased the product in the past year as of March 25, 2008, due to the expiration date. For customers in the US, two letters and forms were sent out - one addressed to the EP/Cardiology Lab and a second to the Risk Management Office.If you have questions or have not received a letter contact Customer Service at 1-800-729-9010. |  | 14295 | 14,295 units for both products | Nationwide and Worldwide to:Argentina,Australia,Austria,Belgium,Brazil,Canada,Chile,China,Columbia,Czechia,Denmark,Egypt,Finland,France,Germany,Greece,Hong Kong,Hungary ,Iran,Ireland,IsraÂżÂżl,Italy,Japan,Lithuania,Luxembourg,Mexico,Middle East,Netherlands,Norway,Panama,Poland,Portugal,Russia,Serbia,Slovenia, South Africa,Spain,Sweden,Switzerland,Taiwan,Turkey,United Kingdom | Terminated | May 03, 2012 | 1500 | 14295 | Computer
Z-1702-2008 | 47908 | No | HQC | Phacofragmentation system. | Ophthalmic | Ophthalmic | 510(k) | Opthalmic microsurgical system-AMO WHITESTAR SIGNATURE Phacoemulsification System, Product Number NGP680300. The device is a modular ophthalmic microsurgical system that facilitates both anterior segment (i.e., cataract) and posterior segment (i.e., vitreoretinal) ophthalmic surgery. |  |  | April 02, 2008 | 2008 | Advanced Medical Optics, Inc. | AMO initiated this field correction event after becoming aware of a trend in complaints associated with AMO Vitrectomy Cutter used in conjunction with the AMO WHITESTAR SIGNATURE Phacoemulsification System. The physician may be unable to make a cut during cataract surgery. | Device Design | On April 2, 2008, AMO began distribution of an Urgent Safety Notice to all customers in the United States via overnight carrier (Federal Express). The U.S safety notice was used as a template for all global AMO communications regarding the Urgent Safety Notice , many of which required translation into local language prior to distribution. Included in this correspondence is a listing of the serial numbers of the affected products.The Notice informed customers of the problem and gave the customers modifications to the Vitrectomy Cutter Priming Instructions. In addition, the notice informed the customers that an AMO Field Service Engineer will be scheduling a site visit to modify a pressure setting on the AMO Whitestar Signature system. Contact the AMO Phaco Technical support at 1-877-AMO-4LIFE for assistance. AMO will provide updates to the FDA, including all information regarding the recall effectiveness as part of the recall progress reports. |  | 191 | 191 units | Nationwide and to: AL, AR, CA, CO, FL, HI, IL, IN, MD, NC, ND, NJ, NY, OH, OR, PA, PR, SC, TX &amp; WA and worldwide to: Azerbaijan, Belgium, Germany, Denmark, Egypt, Spain, France, Great Britain, Croatia, Hungary, Israel, Italy, Jordan, Kuwait, Lebanon, Liechtenstein, Netherlands, Norway, Poland, Portugal, Saudia Arabia, Sweden, Turkey, South Africa, Australia, India, New Zealand and Singapore | Terminated | January 21, 2010 | 659 | 191 | Computer
Z-1752-2008 | 47245 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | PT-3B Control Unit - component of the EndoWave Infusion Catheter System. EndoWave Infusion System consists of three main components, a single use EndoWave Infusion Catheter System which is comprised of a Drug Delivery Catheter (CD) and removable Ultrasound Core (USC) and a reusable EKOS PT-3B Control System. Box is labeled EndoWave Infusion System Drug Delivery Catheter with Ultrasound Core |  |  | March 10, 2008 | 2008 | EKOS Corporation | Potential for unit to burn patient. | Software design | In March 2008, the firm planned to visit each consignee to install the new software version 01.05.02 and hand deliver letters, dated 02/20/2008, notifying consignees of the reason for the action. The old software version is 01.04.02. The software was upgraded because the Ultrasound Core heats very quickly and if removed from the DDC while operating it may result in burns to the patient or user. The software will turn off the Ultrasound Core automatically if removed from the DDC during operation, and will prompt to restart therapy if ultrasound is turned off during treatment.On 04/17/2008, the firm issued a new letter with the date of April 2008 informing customers of the issue. In addition the letter indicates that the corrective action, release of the new software is delayed to incorporate further changes, and that customers will be notified when the software is available for their unit. Customers are advised that the device is safe to use when operating according to the Instructions for Use. EKOS contact number 1-888-356-7435 |  | 218 | 218 units | Products were distributed nationwide to hospitals. | Terminated | May 13, 2011 | 1159 | 218 | Computer
Z-1825-2008 | 47840 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Centricity AW Suite 2.0 Software running on the Centricity RA 1000 or RA600/CA1000 Workstation; for diagnostic image analysis; GE Healthcare Integrated IT Solutions, Barrington, IL 60010 |  |  | May 09, 2008 | 2008 | GE Healthcare Integrated IT Solutions | Incorrect measurement of the aortic length in certain cases of aneurysm during the use of the aorta protocol or customized protocol of the Centricity AW Suite 2.0 running on the Centricity RA 1000 or RA600/CA1000 workstations. | Software design | GE Healthcare IITS sent Urgent Safety Notice letters dated 5/9/08 to all customers who have the AW Suite Software Version 2.0, informing them of the incorrect measurement of the aortic length in certain cases of aneurysm during the use of the aorta analysis protocol or customized protocols. The users were instructed not to use the aorta analysis protocol or the customized protocols for an aneurysm until the new software version can be installed. A software patch to permanently resolve the issue will be installed at no charge. Any questions were directed to the GE Customer Care Center at 800-437-1171. |  | 207 | 207 units | Nationwide and internationally to Australia, Belgium, Canada, Denmark, France, Germany, Great Britain, Hong Kong, Ireland, Israel, Italy, Japan, Malaysia, Malta, Portugal, Saudi Arabia, South Korea, Spain, Sweden, Switzerland and Turkey. | Terminated | September 24, 2010 | 868 | 207 | Computer
Z-1826-2008 | 48078 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Centricity PACS RA1000 Workstation; for diagnostic image analysis. GE Healthcare Integrated IT Solutions, Barrington, IL 60010 |  |  | May 06, 2008 | 2008 | GE Healthcare Integrated IT Solutions | A software anomaly results in incorrect study date and time information being displayed in the report screen and title, which may result in a potential patient misdiagnosis. | Software design | GE Healthcare IITS sent Product Safety Notification letters dated 5/5/08 to all customers who have the GE Centricity PACS Software Versions 3.0.0.X, 3.0.1.X, 3.0.2.X and 3.0.3.X informing them of the incorrect study date and time information being displayed in the report screen and title, which may lead to a potential patient misdiagnosis. The users were instructed to refer to the image overlay for the correct date and time until the new software version can be installed. A software patch to permanently resolve the issue will be installed at no charge. Any questions were directed to the GE Customer Care Center at 800-437-1171. |  | 177 | 177 units | Nationwide and internationally to Australia, Belgium, Canada, Denmark, France, Germany, Italy, Japan, Malta, South Korea, Spain, Switzerland and United Kingdom. | Terminated | February 03, 2010 | 638 | 177 | Computer
Z-1823-2008 | 47834 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Headset Model Number RTC-2K for use with the Serene Sound Digital MRI Compatible High Fidelity Stereo Sound System. |  |  | May 14, 2007 | 2007 | Resonance Technology Inc | Potential safety concerns - Use of this product with a MRI may result in the headset's cord overheating and patients experiencing burns to the skin. | Device Design | The strategy was notifying customers by written communication. The letter advises Resonance Technology, Inc. is voluntarily recalling the MRI Headset, Model No. RTC-2K and RTC- 551-070 because of potential safety concerns. Use of this product with a MRI may result in the headset's cord overheating and patients experiencing burns to skin. The letter also advises that Resonance Technology Inc. has designed and developed the significantly improved MRI Headset, Model No. RTC-661-060-09P-001 (upgrade) and began shipping this product in January 2007. Customers who received the Headset Model; RTC-2K and RTC- 551-070, are asked to upgrade to MRI Headset Model No. RTC-661-060-09P-001. Response is requested as to the number of headsets to be replaced. In addition, consignees are advised that if they further distributed this product, to identify their customers and notify them at once of this product recall. If you have questions, contact Olivia Shayesteh at 818-882-1997 or via email at Olivia@mrivideo.com |  | 20 | 20 units | Nationwide and Internationally: Canada, Japan, Taiwan, United Kingdom, Netherlands, Italy, Spain, Germany, Switzerland, France, | Terminated | July 07, 2011 | 1515 | 49 | Computer
Z-1766-2008 | 47529 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON Antares 5.0 ultrasound system at SW revision 200.0.054 or 200.0.059A, with one of the following options installed: 3-Scape Imaging, Factory, PN 05937011 or 3-Scape Real Time Imaging, UPG ACU, PN 10033688. Material number 10032746, 10037592, 10032746, 10037591 or 10038202. |  |  | March 21, 2008 | 2008 | Siemens Medical Solutions USA, Inc. | Software issue may result in distorted images and inaccurate measurements, which could lead to a misdiagnosis. | Software design | The firm initiated a field correction on 03/21/2008. Site visits to correct the problem and install the revised software version 200.0.063A are planned for all consignees. No other notification was made. The firm can be contacted at 650-694-5993. |  | 488 | 488 active units total | Product was distributed worldwide to 539 consignees, in the USA to MD, SC, AZ, FL, IL, WI, NJ, NC, CA, OH, VA TX, PA MO, NH, IA, NE, MN, NY, GA, MI, WA, HI, MA, SC, WV, VA, ND, IN, WY, UT, MS, TN, PR, SD, KY, OK, OR, and worldwide to Sweden, Saudi Arabia, Austria, New Zealand, Italy, Switzerland, Mexico, Germany, Australia, Canada, Great Britain, France, China, South Korea, Singapore, Denmark, Brazil, Poland, Norway, Faroe Islands, South Africa, Russian Federation, Romania, Trinidad and Tobago, Greece, India, Spain, Thailand, Ukraine, Nicaragua, Japan, India, Czech Republic, Egypt, Taiwan, Turkey, Malaysia, Qatar, Bosnia an dHerzogovina, Vietnam, Brunei Darussalam, Italy, Kuwait, Turkey, Portugal | Terminated | September 30, 2009 | 558 | 488 | Computer
Z-1591-2008 | 47423 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | Deltec CozmoÂżÂż Insulin Pump, model 21-1800. |  |  | February 22, 2008 | 2008 | Smiths Medical MD, Inc. | Smith Medical has received reports of Deltec CozmoÂżÂż Insulin Pumps powering down without an alarm. Investigation concluded that if the battery cap is not fully tightened, the Pump may power down and a brief "chirp" will sound. This may occur if the user has not sufficiently tightened the battery cap or if the battery cap is damaged. If a patient is asleep or does not regularly check their blood gl | Device Design | Consignees and the USA demonstration pump physicians were mailed an "Urgent Medical Device Correction" letter on 2/22/08. The letter stated Details on Affected devices, Description of the problem, Advice on action to be taken by the user, Additional Information and New Warning. These consignees will be required to return the Correction Confirmation Form. The 17 countries with international mailing addresses were sent the Urgent Medical Device Correction letter (Urgent Field Safety Notice) via the Smiths Medical International Regulatory department based in Watford UK. The Smiths Medical Distributors and Affiliates will be required to mail the Safety Notice to their direct users. The USA distributors were mailed the Urgent Medical Device correction letter with a cover letter. These consignees will be required to add the letter to in-house inventory. The USA prescribing physician and the the USA Managed Care Accounts (Insurance Companies) were mailed the Urgent Medical Device Correction letter with a cover letter. These consignees will not be required to respond and no follow-up will be done by Smiths Medical.Contact Mike Herbert at 651-628-7049 if you have any questions or did not receive the recall information from the firm. |  |  |  | Worldwide Distribution --- USA including states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY, and countries of AU, CA, CL, CN, DE, EG, FR, GB, GR, IL, KW, MX, NL, NZ, ON, PL, SA, SE, TR, UK. | Terminated | January 19, 2011 | 1062 |  | Computer
Z-1748-2008 | 47527 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Zimmer Spine Cyclone anterior cervical plate, 4 Level, 88MM (Ti-6Al-4V); REF 600-04-188. |  |  | March 26, 2008 | 2008 | Zimmer Inc. | The screw locking cap may fracture when the surgeon rotates it into the locked position. | Device Design | Consignees were notified via voicemail on 3/26/08 and a letter was sent to all distributors on 3/27/08 informing them that the plate should no longer be used and that all lots are being recalled. A letter dated 6/18/08 was issued to implant physicians recommending patient monitoring. Contact Zimmer Spine at 1-952-857-5636 for assistance. |  |  |  | Arizona, California, Florida, Georgia, Louisiana, Michigan, Minnesota, Montana, New Hampshire, North Carolina, Oklahoma, Pennsylvania, Tennessee, Texas, Utah, Washington and Wisconsin. | Terminated | November 12, 2009 | 596 |  | Computer
Z-1642-2008 | 47487 | No | IXI | Radiation therapy beam-shaping block. | Radiology | Radiology | 510(k) | AD 7 Patient table supplied with Allura Xper FD10 and Allura Xper FD10/10. Product is non-tilting but can be purchased with a tilt option. |  |  | March 14, 2008 | 2008 | Philips Medical Systems North America Co. Phillips | Potential for radiologic patient table to become immobile and unable to move again due to force sensor sensitivity to electromagnetic radiation. | Device Design | On 03/22/2004, firm sent letters via certified mail to notify customers of issue. Letter advises customers of potential risk for patients and users, and states to not use system for critical examinations or treatments if table movements suddenly stops or makes a freeing movement. The letter indicates a representative will contact customer for installation of a Hardware filter, free of charge, to prevent Electromagnetic interference (EMI) from affecting the force sensor. Please contact Philips at 1-800-722-9377 #5, #4, #1. |  | 39 | 39 units | Product is distributed to hospitals and medical centers nationwide to the following states: ND, FL, KY, MN, VA, ME, PA, NJ, CT, TX, CA, OH, TN, OR, GA, NY, SC, MO, WA, IL. | Terminated | July 18, 2012 | 1587 | 39 | Computer
Z-1705-2008 | 47774 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | OEC 9800 Image Intensified fluoroscopic x-ray system (includes OEC 9800, OEC FluoroTrak 9800 Plus, OEC 9800 Plus, and OEC 9800MD Motorized C-arm System), GE Healthcare, Surgery, Salt Lake City, UT. |  |  | April 15, 2008 | 2008 | OEC Medical Systems, Inc | Use of existing four-pedal footswitch on a different machine may cause various operational errors. | Device Design | Consignees were notified by letter on 03/21/2007 and advised that GE would be replacing the 4-pedal switches at some time in the future. Replacement began on 04/15/2008. Contact GE Healthcare at 1-800-874-7378 for assistance. |  | 78 | 78 units | Nationwide, including one VA facility in FL. No military or foreign distribution. | Terminated | December 21, 2011 | 1345 | 103 | Computer
Z-1650-2008 | 47885 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Ambulift, standard basic model manually powered patient lift. The Ambulift is used to move and lift patients from their beds, transport them to the bathroom, and lift them in and out of the bath. |  |  | May 05, 2008 | 2008 | Arjo, Inc. | The armrests of the lift chair may push up, allowing the patient to slip down and out of the seat. The resultant fall could result in serious or fatal patient injuries. | Device Design | Arjo sent Customer Field Correction Notification letters dated 5/2/08 on the Arjo letterhead to the end user accounts who received the affected Ambulift Patient Lifts, to the attention of the Administrator/Risk Manager, via UPS 2nd day delivery on 5/5/08. The letters informed the consignees of the incident with the patient slipping from the seat of the lift and sustaining fatal injuries and advised them to always ensure that a clinical assessment to verify that the resident has sufficient upper body strength to maintain a steady position on the chair, and that the resident has sufficient mental capacity to adhere to the caregiver's instructions, is carried out by a qualified nurse or therapist before lifting residents with the Ambulift. To further enhance safety, a seat belt is being made available, which must be fitted to the chair, enabling it to be used when required by clinical assessment. Instructions on how to apply, use, and maintain the seat belt will be provided with the belts. The belts will be provided at no cost. The accounts were requested to review, complete and return the enclosed customer response form, indicating the number seat belts needed. Any questions were directed to Arjo's Quality Department at 1-800-323-1245, ext. 6140 or 6165. |  | 11 | 11 lifts | Worldwide Distribution-USA includes the states of AL, AZ, DE, FL, KY, MO, NY, OK, and VA | Terminated | February 03, 2010 | 639 | 14 | Computer
Z-1309-2008 | 46834 | No | FTC | Ultraviolet lamp for dermatologic disorders. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | HOUVA III Phototherapy System with DermaSense Technology comes on one skid within 10 cardboard boxes. |  |  | February 19, 2008 | 2008 | National Biological Corp | Software allows operator to override "low line voltage" error warning and store light intensity value. | Software design | National Biological sent an "Important safety reminders for your Houva DermaSense phototherapy device" letters, dated 2/11/08 to users. It states that the user should collect a "Profile" of the light energy and record it on a daily/weekly flow chart. The chart should be in plain view so each operator can clearly determine the units output. If an error message appears in the display panel, treatment should not be continued unless the phototherapy booth operator was able to identify and correct the cause of the error. Make sure both breakers are in the "ON" position before starting up the unit each morning and before giving treatment. |  | 48 | 48 Units | Worldwide distribution --- USA including states of: CA, DE, FL, GA, ID, IL, IN, MA, MI, MN, NC, NH, NY, OH, OR, PA, TX, VA, WA and WI, and country of Canada. | Terminated | September 20, 2011 | 1309 | 48 | Computer
Z-2174-2008 | 48616 | No | LKK |  |  | General Hospital | PMA | Model 8731SC Intrathecal Catheter |  |  | April 14, 2008 | 2008 | Medtronic Neuromodulation | The SC catheters cannot completely engage with the Model 700-04M portal connector. The result of this condition is a possible leakage or disconnect of the SC catheter from the portal during the continuous therapy trial period. | Device Design | Medtronic informed all Belgium implanting and managing physicians of the SynchroMed II infusion system. Each customer was mailed a Dear Healthcare Professional letter (Medtronic Urgent: Field Safety Notice) on 4/3/08 providing important safety information and patient management recommendations regarding the device incompatibility. The firm recommends that healthcare professionals do not use SC catheters if a connection with an implantable access port is required. Also, if a continuous infusion trial is required with use of an implantable access port, they should use a model 8711 intrathecal catheter. The firm states that for patients who are currently undergoing a continuous trial with an implantable port, replacing the catheter connector before connecting to the SynchroMedd II using the Medtronic model 8578 revision kit is recommended. Medtronic also recommends discussing the risks of connection failure with patients. Local field reps will personally follow-up with affected customers to confirm and document notification. |  | 232 | 232 total from all models | International Distribution --- country of Belgium. | Terminated | September 10, 2008 | 149 | 232 | Computer
Z-1776-2008 | 47807 | No | ITI | Powered wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Q-Logic Controller Software used on the following power wheelchairs: Q6000, Q6000XL, Q6000Z, R4000, Q600, Q600XL, Q610, Q1107, Q600E, and Dynamo 1107 |  |  | April 11, 2008 | 2008 | Pride Mobility Products Corp | Software - Watch dog timer feature was disabled. | Software design | The recalling firm issued an Urgent Device Correction letter on 4/11/08. This letter instructs the dealers to contact the end users so that the units can be upgraded to enable the watch dog feature. |  | 4129 | 4129 | Class II Recall - Worldwide Distribution --- including USA and Canada. | Terminated | November 13, 2009 | 581 | 4129 | Computer
Z-1716-2008 | 47468 | No | BYJ | Portable liquid oxygen unit. | Anesthesiology | Anesthesiology | 510(k) | Generation 4 electronic liquid gauges used on the Caire Liberator, Stroller, Sprint, and Hi-Flow Portable oxygen unit. |  |  | October 23, 2007 | 2007 | Caire, Inc. | Moisture ingress into the electronic liquid level indicator may cause inaccurate liquid oxygen level readings. The oxygen could run out without the patient knowledge. | Device Design | A Technical Service Bulletin was sent to consignees on/about October 23, 2007 to notify customers of the potential issue. Technical support and information for ordering replacement Generation 4 liquid gauges were offered.For technical concerns or if you have not received the notification from the firm, contact U.S. Technical Service at 1-800-482-2473 or European Technical Service at +44 (0) 797 1260077. |  | 3470 | 3470 | U.S. Nationwide and the following foreign countries: Argentina, Austria, Brazil, Canada, China, Chile, Central America/Caribbean, Colombia, Costa Rica, Europe, Austria, Belgium, Denmark, France, Germany, Greece, Hungary, Irish Republic, Italy, Netherlands, Norway, Portugal, Spain, Sweden, Switzerland, United Kingdom, Mexico, South America, Peru and Uruguay. | Terminated | December 17, 2009 | 786 | 3470 | Computer
Z-2160-2008 | 48602 | No | IZJ | Cine or spot fluorographic x-ray camera. | Radiology | Radiology | 510(k) | General X-Ray System types with Digital Spot Imaging (DSI) software |  |  | May 23, 2008 | 2008 | Philips Medical Systems North America Co. Phillips | X-ray images may be stored in the wrong patient file or corrupted, images from two different patients may be combined in one patient's folder or may be combined on one image (corrupted). | Software design | On 5/20/08 the firm issued an URGENT - DEVICE CORRECTION letter to the consignees advising of the problem, requesting consignees to check the stored or printed images to ensure that it represents the data and images captured from the examination. The firm will issue a Field Change Order to their service representatives providing new software, when software is available. This should be issued by the end of June. |  | 41 | 41 units | Nationwide Distribution. | Terminated | September 28, 2010 | 858 | 41 | Computer
Z-1444-2008 | 46942 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Philips Medical Systems iSite PACS (Picture Archiving and Communications Systems) version 4.1.17.0 - Radiology Image management system. |  |  | January 29, 2008 | 2008 | Philips Healthcare Informatics, Inc. | When using the Freehand Region of Interest (ROI) tool, errors may occur in area calculation yielding incorrect results. The firm discovered that a descrepancy exists in the measurement provided by the tool. | Software design | The firm issued product notification letter February 20, 2008, to its consignees. The notification alerts customers to the behavior, requires them to post the notification and notify users not to use the Freehand ROI feature until they are upgraded to a corrected software version. Philips will deliver a software upgrade to correct the problem. |  | 2 | 2 units | Worldwide Distribution --- USA including states of MN, CO, and PA, and country of Germany. | Terminated | September 25, 2008 | 240 | 5 | Computer
Z-1515-2008 | 46941 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | SEQUENCER, MOSAIQ versions 1.3 and 1.4 with Elekta and Siemens Machines, Model number 40003-001000-RO, Manufactured by IMPAC Medical Systems, Sunnyvale, CA. |  |  | February 14, 2008 | 2008 | Impac Medical Systems Inc | Software issue may result in change to intended treatment field, potentially resulting in mistreatment. | Software design | Firm issued field advisory to its consignees on 02/14/2008. Phone calls will also be made to affected customers to ensure reception and understanding of the field notification. Firm advises users with affected product to review the jaw positions of IMRT fields to ensure that they match the RT Plan after an Apply Limits or Machine Change occurs, until software at their site can be updated to the corrected version. A software patch is under construction. |  | 36 | 36 units | Worldwide Distribution --- USA including states of OH, OK, NJ, MN, OR, TN, NY, CA, AZ, NV, IA, GA, AL, MD, MI, PA, IN, WA, FL, TX, CO, and countries to Canada, Australia, Germany, Greece and Sweden. | Terminated | June 24, 2009 | 496 | 36 | Computer
Z-1442-2008 | 46920 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Medical Systems On-Board Imager, Versions 1.0.15 and 1.2.05 Imaging Accessory to Medical Linear Accelerator. The On-Board Imager Device is used for verification of correct patient position in relation to isocenter and verification of the treatment fields in relation to anatomical and/or fiducial landmarks. |  |  | January 03, 2008 | 2008 | Varian Medical Systems Oncology Systems | If used with a third party radiation therapy treatment planning software system, mistreatment may occur because of a misalignment. | Software design | The firm has upgraded all user sites through normal course of service and installation activities. Notification letters were distributed starting January 03, 2008. |  | 291 | 291 units | Worldwide Distribution -- USA including states of GA, MI, NY, MN, TX, TN, CA, MD, KS, AL, KY, CT, WI, LA, NC, NJ, PA, IL, RI, FL, NE, AK, VA, WV, NV, MO, TN, OR, IN, MA, CT, AL, AZ, NH, ID, OH, SC, CO, CT, ND, WA, and UT, and countries of Sweden, Switzerland, Germany, Denmark, Japan, UK, France, Italy, Norway, Taiwan, New Zealand, Canada, Poland, Australia, Colombia, China, Spain, Belgium, Russian Federation, India, Iceland. | Terminated | November 27, 2010 | 1059 | 291 | Computer
Z-2019-2008 | 46287 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Cytomics FC 500 Flow Cytometry System with CXP Software Version 2.1 and 2.2 - CXP analysis SW V2.1 single user upgrade, Automated differential cell counter; Part Number 722395; Beckman Coulter; Miami, FL 33196. |  |  | May 04, 2007 | 2007 | Beckman Coulter Inc | Incorrect Display: In certain modes, the protocol will not initially display the correct data in some plots. In addition, when running EWL files, any edited SampleID1 is not automatically updated in all tubes in the panel. | Software design | Consignees were notified of this safety recall by an Urgent Product Corrective Action letter sent on 5/3/07. The letter provided instructions on how to ensure correct data acquisition. Users were requested to return a response form acknowledging receipt of the letter. For additional information contact 1-800-526-7694. |  | 501 | 445 units in the US; 56 units in Canada | Worldwide Distribution including USA and country of Canada. | Terminated | July 11, 2011 | 1529 | 501 | Computer
Z-2196-2008 | 48655 | No | KPQ | Radiation therapy simulation system. | Radiology | Radiology | 510(k) | Update Instructions "TH003/08/S" applies to Siemens brand SIMIVIEW 3000, SIMVIEW NT, and MEV ASIM S systems; |  |  | March 10, 2008 | 2008 | Siemens Medical Solutions USA, Inc | The firm sent an Update Instruction, "TH003/08/S" for Motion Enable Switch/holder assembly. The switch holder allows the motion enable switch to be activated if inserted partially; thereby enabling inadvertent motion and serious injury may result. | Software design | On June 25, 2008 all consignees were notified via written letter and on-site visit by the firm's sales representative, informing them that the switch holder allows the motion enable switch to be activated if inserted partially; thereby enabling inadvertent motion and serious injury may result and providing instructions on the recall. A new software version will be provided. For assistance contact Siemens at 1-925-246-8407. |  | 300 | 300 units | Worldwide and Nationwide | Terminated | December 22, 2010 | 1017 | 300 | Computer
Z-1606-2008 | 47518 | No | MTF |  |  | Immunology | PMA | Dimension Total Prostate Specific Antigen (TPSA) Flex Reagent Cartridge (RF451). Siemens, Newark, DE 19714 |  |  | March 24, 2008 | 2008 | Siemens Healthcare Diagnostics Inc. | False Results: The product may exhibit falsely elevated results. The affected lots demonstrate sporadic variability in recovery of calibrator and quality control; therefore this issue may not be readily detected upon calibration or during routine processing of Quality Control. | Device Design | The recalling firm issued a letter (Urgent Field Safety Notice) dated 3/25/08 to inform their customers to discontinue use of any remaining inventory, retest using an alternate lot for those samples from post-prostatectomy patients which recovered positive TPSA results less than 1.0 ng/nL. The firm has requested that you call the them for no-charge replacement product at 1-800-441-9250 and/or for technical assistance. |  | 1631 | 1631 cartons | The product was shipped to medical facilities nationwide. The product was also shipped to Germany, Singapore, Canada, Japan and AE | Terminated | October 15, 2008 | 205 | 1631 | Computer
Z-1167-2008 | 46408 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Zoll E Series Defibrillator with 12 Lead ECG, Automated external defibrillator, Software Version 3.00 and higher; Zoll Medical Corporation, Chelmsford, MA 01824. |  |  | January 15, 2008 | 2008 | ZOLL Medical Corporation, World Wide Headquarters | Incorrect Patient Records: Patient records that have been stored, then later transmitted or printed may not be the correct record. Records that are transmitted or printed immediately following acquisition are correct. | Software design | Consignees were notified by an Urgent Device Correction letter sent on 1/14/08. The letter instructed users to complete and return a response form. Upon receipt of the response form, the recalling firm will have sent a software update. The letter also provided instructions for users to follow in order to avoid the software error before the software update was received. For additional information, contact 1-800-348-9011. |  | 5018 | 5,018 units | Worldwide Distribution including USA and countries of Canada, AE, Australia, Belgium, Chile, Germany, Spain, France, UK, Hungary, India, Italy, Jordan, Kuwait, Netherlands, Russia, Saudi Arabia, Singapore and ZA. | Completed |  |  | 5018 | Computer
Z-1242-2008 | 46382 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Agfa Heartlab Cardiovascular Results Management Product: Versions 2.03 and 2.04. Heartlab, Inc., Westerly, Rhode Island 02891 |  |  | July 27, 2007 | 2007 | AGFA Corp. | Incorrect Results: Erroneous echocardiographic measurement values due to mathematical formula being mis-configured. | Software design | On 7/27/2007 to 8/16/2007, AGFA Healthcare notified, via telephone, consignees of the situation. On 7/27/2007, a formal customer notification letter was sent to affected consignees, providing them with a respective listing of patients that were potentially affected. AGFA informed the consignees that the firm would correct any incorrect patient ECG data. AGFA also informed their consignees that they would provide a configuration update to correct the problem. A notification confirmation form to be return to the firm was also attached to the notification letter, which also provided the firm's customer service number: (877) 777-2432. |  | 6 | 6 units | Nationwide Distribution -- USA including states of GA, NH, NY and TX | Terminated | October 10, 2008 | 441 | 6 | Computer
Z-1241-2008 | 46381 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Agfa HeartLab Cardiovascular DICOMstore Version 2.04.40.00 software, Heartlab, Inc., Westerly, Rhode Island 02891. The software is used in the Agfa HeartLab Cardiovascular System. |  |  | June 21, 2007 | 2007 | AGFA Corp. | MIsidentification: The electrocardiogram (ECG) data of one patient is misidentified as the data of another patient. | Software design | On 6/21-26/2007, AGFA Healthcare notified, via telephone, the affected consignees of the situation. On 7/2/2007, affected consignees were sent, via registered mail, an "Urgent Medical Device Correction" letter instructing them to: a) remove the affected software DICOMstore 2.04.40.00 and replace it with the prior software version, DICOMstore 2.04.34.02; and b) review patient ECG reports to determine if any patients were actually affected and to notify the firm. AGFA Healthcare also provided a customer service contact number (877) 777-2432. |  | 43 | 43 units | Worldwide Distributution - USA, including states of AZ, CA, DC, FL, IL, IN, KY, MA, ME, MN, MS, OH, OK, OR, NH, PA, SC, TN, TX, WI, and country of Canada. | Terminated | August 04, 2008 | 410 | 43 | Computer
Z-1220-2008 | 46281 | [6 empty] | VECTRA GENISYS, Model No. 2761, Chattanooga Group, Hixson, TN 37343, Made in the USA |  |  | December 17, 2007 | 2007 | Chattanooga Corp Inc | Device malfunction: causing electric shock and burn. | Device Design | Beginning on 12/17/2007, Chattanooga Corp. sent two recall notification letters, via UPS mail, to its consignees informing them of the situation. The letters stated that the recall was to the user level and requested that the user perform the software upgrade, which will eliminate the possibility of shock and burn. |  | 4047 | 4047 units | Worldwide Distribution - including USA | Terminated | February 19, 2010 | 795 | 22000 | Computer
Z-1162-2008 | 46280 | No | BTC | Reservoir bag. | Anesthesiology | Anesthesiology | 510(K) Exempt | 1 Liter Latex-Free Breathing Bags; Reservoir Bag; Vendor Part Number: 720059B, Medline Item Number: 40056; Available in various Anesthesia Circuit Kits and Complete Delivery System kits including: 1) Item Number CDS980230I, NICU-OR CDS; 2) Item Number DYNJ22691, Basin Supply Pack; 3) Item Number DYNJ24724, Spine-Crani Anesthesia Kit; 4) Item Number DYNJAAHY11, Hyperinflation System-UHS-1L; 5) Item Number DYNJAP0010, Anesthesia Circuit Ped-LF; 6) Item Number DYNJAP0011A, Pediatric Circuit-LF; 7) Item Number DYNJAP0012, Pediatric Anesthesia 40 In. CI; 8) Item Number DYNJAP0100, Pedi Breathing Circuit-LF; 9) Item Number DYNJAP0101, Anesthesia Circuit-Ped; 10) Item Number DYNJAP4012, Anesthesia Circuit Ped; 11) Item Number DYNJAP4013, Anesthesia Circuit; 12) Item Number DYNJAP4702A, Anesth Circuit Ped GS Line; 13) Item Number DYNJAP4707, Anesth Circuit 40in Ped-LF; 14) Item Number DYNJAP4722, Anes Cir 40 Ped GSEPW HYBABYH; 15) Item Number DYNJAP4729A, Circ Ped Anes Super-LF; 16) Item Number DYNJAP4730A, Circuit Pediatric Anesthesia; 17) Item Number DYNJAP4732, Anesth Circ Ped 40in 2B/V; 18) Item Number DYNJAP4740, Anesth Circ 40in Ped; 19) Item Number DYNJAP6101, Anesthesia Circuit; 20) Item Number DYNJAP6105, Anesth Circuit 60in Ped; 21) Item Number DYNJAP6106, Anes Circ Ped 60in BI GSE-LF; 22) Item Number DYNJAP6107, Anesth Circ Ped 60in ANG HME; 23) Item Number DYNJAP6108, Anesth Circuit 60in Ped; 24) Item Number DYNJAP6109, Anesth Circ Ped 60in B/V 1L-LF; 25) Item Number DYNJAP6111, Anes Cir Ped 60 GSEPW PALL FIL; 26) Item Number DYNJAP6112, Anesthesia-Peds Circuit; 27) Item Number DYNJAP6113, Circuit Pedi 60in; 28) Item Number DYNJAP6114, Anesth Circ 60in Ped; 29) Item Number DYNJAP6116, Anesth Circ 60in Ped; 30) Item Number DYNJAP6120A, Pediatric Circuit 40in LF; 31) Item Number DYNJAP6124, Anesth Circ Pedi 60in 1L GSL; 32) Item Number DYNJAP6127, Anes Circ Pedi 60in-LF; 33) Item Number DYNJAP6130, Anesth Circ Ped 60in 1L-LF; 34) Item Number DYNJAP6132A, Circuit Pediatric Anesthesia; 35) Item Number DYNJAP6133, Anesth Circuit Peds 60in-LF; 36) Item Number DYNJAP6136, Circuit Peds Anesthesia; 37) Item Number DYNJAP6137, Circuit Pediatric Anesthesia; 38) Item Number DYNJAP7217A, Anesth Circuit 72in Exp Ped BI; 39) Item Number DYNJAP7225A, Circuit Pediatric Anesthesia; 40) Item Number DYNJAP7227, Anes. 90in Exp Pedi -LF; 41) Item Number DYNJAP7233, Anes. Circuit 90in Exp Pedi-LF; 42) Item Number DYNJAP7234A, Anesthesia 72in Pediatric Exp; 43) Item Number DYNJAP7235A, Anesth Circ 72in Ped; 44) Item Number DYNJAP7236A, Anesthesia Pedi 72in; 45) Item Number DYNJAP7237, Circuit Pedi 90in EXP; 46) Item Number DYNJAP7241A, Circuit Pediatric Anesthesia; 47) Item Number DYNJAP7243B, Circ Anesth Ped 2 B/V LFree; 48) Item Number DYNJAP7246C, Circuit Pediatric Anesthesia; 49) Item Number DYNJAP7247, Anes Circ Peds 90in Exp; 50) Item Number DYNJAP7248, Anesth Circ Ped 72in Exp; 51) Item Number DYNJAP7261, Circuit Pediatric Anesthesia; 52) Item Number DYNJAP7262, Circuit Pediatric Anesthesia; 53) Item Number DYNJAP7263, Pediatric Anesthesia Circuits; 54) Item Number DYNJAP9005, Anesthesia Circuit; 55) Item Number DYNJAP9721A, Anesth Circuit Pedi HMEF LF; 56) Item Number DYNJAP9725, Pedi Expand Circuit; 57) Item Number DYNJAP9726, Anesth Circ Ped 90in; 58) Item Number DYNJAP9729, Anesth Circ 90in Exp Ped; 59) Item Number DYNJAP9730A, Circuit Pediatric 60in CORR-LF; 60) Item Number DYNJAP9732A, Circuit Pediatric Anesthesia; 61) Item Number DYNJAP9735, Peds Circuit; 62) Item Number DYNJAP9739, Circuit Pediatric Anesthesia; 63) Item Number DYNJAP9740, Circuit Pediatric Anesthesia; 64) Item Number DYNJAP9743, Circuit Pediatric Anesthesia; 65) Item Number DYNJAPF6003, Anesth Circ 60in Ped Unilimb; 66) Item Number DYNJAPF6005A, Anesth Circ 60in Uni Peds; 67) Item Number DYNJAPF6006, Anes Circ Peds 60in Uni; 68) Item Number DYNJAPF6009, Peds Anesthesia Circuit-LF; 69) Item Number DYNJAPF6011, Circuit Pediatric Anesthesia; 70) Item Number DYNJAPF601 |  |  | December 07, 2007 | 2007 | Medline Industries Inc | Separation from Bushing: The breathing bag may separate from the bushing during use. | Device Design | Consignees were notified by a Sub-Recall Notice letter sent on 12/5/07. The letter instructed users to check their inventory for any affected product and to return a response card indicating the quantity of affected product. The users were instructed to return the affected products to the recalling firm so that they can be reworked. Users who did not wish to return the affected products could request a replacement product and were instructed to destroy the affected products. The letter also informed users that affected products in the manufacturer's inventory were being replaced and these replaced products would contain the label "INSPECTED." For additional information, contact 1-866-359-1704. |  | 77750 | 77,750 bags | Nationwide Distribution. | Terminated | October 16, 2008 | 314 | 571500 | Computer
Z-1161-2008 | 46204 | Yes | MSX | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | S/5 iCentral, Cardiovascular Monitoring Devices, Versions L-NET03, L-NETC03, L-NET05 and L-NETC05, GE Healthcare, Wauwatosa, WI 53226. |  |  | December 13, 2007 | 2007 | GE Healthcare | Failure to alarm: An active monitor could become disconnected from the unit without any notification or alarm. | Device Design | Consignees were notified via a Medical Device Correction letter on 12/13/07. The letter explained the circumstances in which the "connection to monitor loss" alarm would not display, if a monitor was disconnected. The letter also included an attachment entitled "How to ensure a 'Connection to monitor lost' alarm on iCentral." For additional information, contact 1-800-558-7044. |  | 2146 | 2146 | Worldwide Distribution including USA states of AZ, CA, CO, CT, FL, GA, IA, IL, IN, KS, KY, LA, ME, MI, MN, MO, MS, NC, NH, NJ, NY, PA, SC, TN, TX, VA, VT, WA and WI; and countries of Algeria, Australia, Austria, Barbados, Belgium, Bolivia, Brazil, Bulgaria, Canada, Chile, China, Costa Rica, Croatia, Czech Republic, Denmark, Ecuador, Estonia, Finland, France, Georgia, Germany, Greece, Guatemala, Guans, Hong Kong, Hungary, Iceland, India, Iran, Ireland, Israel, Italy, Japan, Jordan, Korea, Kuwait, Libyan, Macedonia, Malaysia, Mexico, Netherlands, New Zealand, Nicaragua, Norway, Pakistan, Poland, Portugal, Puerto Rico, Qatar, Romania, Russian Federation, Saudi Arabia, Singapore, Slovakia, South Africa, South Korea, Spain, Sweden, Switzerland, Syrian Arab Republic, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, Uruguay, Venezuela and Yemen. | Terminated | August 13, 2013 | 2070 | 2146 | Computer
Z-1180-2008 | 46716 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | da Vinci S Surgical System, Model number IS2000 Revision number A51 P5; endoscopic surgical control system, Intuitive Surgical, Inc., Sunnyvale, CA 94086 |  |  | December 27, 2007 | 2007 | Intuitive Surgical, Inc. | Delay in responding: In certain circumstances, the device may not respond immediately to a user's command such as master clutch or camera control. | Software design | Beginning 12/28/2007, Intuitive Surgical Field Service Representatives visited each site to correct the issues. A notification letter was provided to all sites. |  | 9 | 9 units | USA Distribution: TX, NJ, WA, NV, HI, MI, NH and CA | Terminated | August 15, 2008 | 232 | 9 | Computer
Z-1155-2008 | 44884 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | GE LOGIQ 3 Expert ultrasound scanner with software versions 4.0.5, 4.1.1 and 4.1.2. GE Healthcare P.O. Box 414, Milwaukee, Wisconsin 53201 |  |  | September 04, 2007 | 2007 | GE Healthcare Clinical Systems | Inaccurate guidelines: The BE9C is an optional, small radius Bi-planar convex probe intended for use with the LOGIQ 3 Expert system in Urology application. An attempt by the user to activate the biopsy guidelines while this probe is in use will cause incorrect guidelines to be displayed on the image. The incorrect guidelines are recognized by their orientation that is left/right reversed from t | Software change control | GE sent consignees a "GE Urgent Medical Device Correction" letter dated 8/14/07. The letter informed the customers about the problem, offer a short-term workaround and informed them that a representative will schedule a visit to upgrade the software. |  | 62 | 62 | Worldwide Distribution - USA (NY &amp; AL), Korea, China, India, Brazil, Mexico, Germany, Greece, Italy, Poland, Russia, Saudi Arabia, Spain, and Czech Republic. | Terminated | May 29, 2012 | 1729 | 62 | Computer
Z-1711-2008 | 47335 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Grey Inter-Unit Interface Connectors ("IUI") on the following Alaris System modules built on or before March 26, 2004: (1) Alaris PC point-of-care unit ("PC Unit") (Model 8000)* (a.k.a. Medley PC Unit), |  |  | March 05, 2008 | 2008 | Cardinal Health 303 Inc DBA Alaris Products | A channel error on an affected Alaris System module will stop an active infusion (or monitoring) with an audible and visual alarm. Channel errors with the Alaris System due to the Inter-Unit Interface (IUI) connectors on all Alaris System modules (manufactured on or before March 26, 2004). Grey Inter-Unit Interface Connectors ("IUI") on the following Alaris System modules built on or before March | Device Design | An Urgent Medical Device Recall Notification letter was sent to consignees on April 10, 2008. Four customer letters were issued. The U19 IC issue affects only the Pump module and the U9 IC issue affects only the Syringe and PCA modules. Therefore there will be one letter for customers affected by all three issues Ul9 IC, U9 IC, and the IUI, one letter for affected Ul9 IC and U9 IC customers, one letter for U19 IC-only customers and one letter for U9 IC-only customers. All mailings were conducted through consignees (distributors as well as direct customers for the affected devices). Distributors will be sent multiple copies of the Recall Notification and Customer Response Cards. The Recall letters all include Required Actions sections which give their customers specific information about the issue and what the customers are to do and includes a web site link for each specific problem. Each recall letter also has chart of Cardinal Health contacts support. |  | 77515 | 77,515 Units | Class II Recall - Worldwide Distribution --- USA including states of AR, AZ, CA, CT, CO, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, ME, MD MI, MN, MO, MS, NC, NE, NH, NJ, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI and WY, Guam and country of Canada.****CDRH Comments: "Note for District Office to follow-up" The firm has not provided a grand total of devices affected by these problems. It is possible that devices suffering from U19 &amp; U9 problems may also experience IUI problems. This recall has 3 problems with 6 modules. One problem affects only 2 modules, one problem affects 6 modules. For simplicity and expediency, CDRH agrees to keep the file as is, rather than separate into 6 products and have complicated explanations of the problems.77,515 devices affected by IUI problem.85,286 devices affected by U19 problem.8,286 devices affected by U9 problem. **** | Terminated | September 13, 2011 | 1287 | 171087 | Computer
Z-1137-2008 | 45959 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | MultiDiagnost Eleva Flat Detector X-Ray System; Part Number: SW 9896 001 43361; Philips Medical Systems, |  |  | November 29, 2007 | 2007 | Philips Medical Systems North America Co. Phillips | Delay in Imaging: MultiDiagnost Eleva with Flat Detector x-ray system cannot perform fluoroscopy procedures while certain image file housekeeping operations are performed. Fluoroscopy cannot be performed and new exposures are not able to be acquired until the housekeeping operations are complete which may take several minutes. The loss of system availability could occur during an interventional | Software design | On 11/29/07 Philips Medical Systems sent a letter dated 11/28/07 to customers stating that if users perform certain image file housekeeping operations, fluoroscopy cannot be performed and new exposures cannot be acquired until those operations are complete. It also instructs that the unavailability of fluoroscopy and exposure can be avoided by not pressing the "Delete All Images" button until any pending Autopush operation is complete. FCO 70800067 will be issued to service representatives to install new software. |  | 18 | 9 units were distributed to hospitals in the US, 6 were installed, 3 were not installed or in use. | USA Distribution - CO, MN, NC, NJ and TX. | Terminated | September 17, 2010 | 1023 | 18 | Computer
Z-1546-2008 | 47386 | No | MUE | Full-field digital mammography system. | Radiology | Radiology | 510(k) | Selenia Full Field Digital Mammography System with Software v. 3.3.1.1; Full field digital, system, x-ray, mammographic; Model Number: SEL-00002; Hologic, Danbury CT 06810 |  |  | February 28, 2008 | 2008 | Lorad, A Hologic, Inc. | Marker Misalignment: The new software version contained a magnification factor that shows CAD markers misaligned with the identified indications. The marked area is misaligned up to 2cm from the indication. | Software design | Consignees were notified by an Urgent Recall Notice sent on 3/21/08. The letter informed users that representatives were dispatched to all locations and they have corrected the magnification factor problem. The letter instructed users to re-evaluate any cases read using CAD from the period of installation until the date of system correction. The letter also instructed users that only images read with CAD markers are affected. The letter requests users to return a confirmation of notification. For additional information, contact 1-877-371-4372. |  | 73 | 73 units | Worldwide Distribution including USA and countries of Canada and the Republic of Korea. | Terminated | February 03, 2009 | 341 | 73 | Computer
Z-1644-2008 | 47526 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | EVMS (Enterprise Visual Medical System); Picture archiving and communications system; Model Numbers: 5.30.4, 5.30.5, 5.30.6, 5.30.7, 5.40.0 and 7.0.0.1. Emageon, Inc., Hartland, WI 53029. |  |  | January 01, 2008 | 2008 | Emageon | Miscalculation: The calculation of the standard uptake value does not include the required calibration factor/decay correction factor. The resulting value is incorrectly calculated, which may lead to incorrect interpretation of tissue density. The calculation error exists in both the software and user guide. | Device Design | Consignees were notified by Urgent Device Correction letters that were sent on 1/9/08 and 1/18/08. The letters instructed users not to use the standard uptake value in EVMS for diagnostic interpretation when viewing PET studies unless they received service bulletin 445, 446 or 7.0.1. Users were also instructed to replace pages in the user's guide with the pages provided and to affix the enclosed warning sticker on all workstations. Users will be contacted by the recalling firm to schedule a software patch. |  | 99 | 99 | Nationwide Distribution including states of: MI, WI, CA, MN, AZ, CO, MN, MD, DC, MO, TX, IL, GA, FL, WI, SC, PA, NY, KS, TN, AL and OK. | Terminated | June 11, 2009 | 527 | 99 | Computer
Z-1094-2008 | 45960 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Synchron LX20 Clinical Chemistry System, Part Numbers: 466200,473032, Beckman Coulter, Inc. |  |  | April 05, 2007 | 2007 | Beckman Coulter Inc | Intermittent Failure of Stirrer Motor-Beckman Coulter has determined that incorrect Glucose results on LX20 orDxC instruments have been related to an intermittent failure of the stirrermotor.-Stirrer motors can stall without any flags or motion errors.-If a stall occurs during a GLUm test near the time of sample inject,results may be affected.-Affected GLUm results may be high or low. | Software design | GLUm stir motors greater than one year old were replaced. Replacements began April 5, 2007. Based on the Risk Assessment, BCI decided that Customer Notification is not needed. GLUm stir motors greater than one year old will be replaced on high volume LX/DxC systems by 3rd quarter of 2007. Replacements began April 5, 2007. A new stirrer design is being developed and is expected to be available before 4th quarter of 2007. Implementation Strategy of Stock Control: 1. A Stop Ship was not necessary based on risk assessment. The possibility of stirrer motor failure may increase with time and duty cycles. Systems running a high volume of tests may have a greater likelihood of exhibiting stirrer motor failure with motors greater than one year old. 2. New Customers were placed on a preventative maintenance plan to have their motor replaced at one year cycle. Completed 3. A Technical Update was issued to notify Hotline and Field Service Personnel of the issues. |  | 370 | 344 in the US; 26 in Canada | Worldwide Distribution - USA including Canada. | Terminated | January 15, 2013 | 2112 | 2599 | Computer
Z-1307-2008 | 46813 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Impax 4.5, Impax 5.0, Impax 5.2 and Impax 5.2x Picture Archiving Systems (PACs) with integrated Voxar 3D software; and Impax 5.0 and Impax 5.2 PACs with integrated MediCAD software. The softwares are used to provide digital X-ray images for pre-surgical orthopedic planning. AGFA Corp., Greenville, South Carolina 29601 |  |  | February 15, 2007 | 2007 | AGFA Corp. | Misidentification: The orthopedic planning X-ray images for one patient are misidentified as the images for another patient. | Software design | On 2/15/07, AGFA notified their consignees of the situation by sending them "Urgent Medical Device Correction" letters, instructing them on how to work around the problem until the firm was able to install the upgraded software that corrected the problem. On 12/11/2007, the firm reported that the upgraded software had been installed at all affected sites. |  | 879 | 879 units | Worldwide Distribution - USA including states of AR, AZ, CA, CO, DC, FL, GA, IL, IN, KY, KS, MA, MD, ME, MI, NC, NJ, NV, NY, NV, OH, OK, PA, SC, TN, TX, UT, VA and WA, and country of Canada. | Terminated | June 30, 2009 | 866 | 879 | Computer
Z-1233-2008 | 46800 | No | FNG | AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | BHM Medical Infrared (IR) Hand Control System, an optional wireless control component of the Arjo Maxi Sky 600 and BHM V4 Ceiling Patient Lifts; the hand control system is composed of two components, part 700.13640.33 - IR hand control, and part 492.00038 - IR receiver board |  |  | February 20, 2008 | 2008 | Arjo, Inc. | The ceiling lift may not stop lateral movement after releasing the left (or right) action button on a ceiling lift equipped with an infrared hand control. | Device Design | Arjo sent Urgent Device Field Correction Notification letters dated 2/15/08 on the BHM Medical letterhead to the end user accounts who received the affected IR remotes, advising them of the potential for the remote to become difficult to operate or to create uncontrolled movement of the ceiling lift. The accounts were instructed to remove the IR remote from service until the field correction is issued by BHM Medical. Instructions for removal of the IR equipment were provided in the Safety Advisory Notice included with the letter. The accounts were requested to complete and return the enclosed customer response form, acknowledging the receipt and understanding of the letter, and indicating the number of units requiring correction. |  | 25 | 25 units | Nationwide Distribution. | Terminated | September 11, 2010 | 934 | 25 | Computer
Z-1463-2008 | 46850 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Viasys Healthcare, Pulmonetic Systems LTV 1000 Ventilator |  |  | February 22, 2008 | 2008 | Cardinal Health 203, Inc | The failure of an electronic component could result in failure of the ventilator to breathe for the patient and/or failure of the ventilator to properly alarm to alert the caregiver. | Device Design | An Urgent Medical Device Recall letter dated February 22, 2008 was sent to ventilator consignees/owners and distributors. The letter describes the issue, identifies the affected devices, states that the devices will be corrected by Cardinal Health and provides direction to verify ventilator operation and audible alarm function daily. Service Centers were provided information to identify affected printed circuit board assembly replacement kits' lot numbers and asked to return them. An "Update" was mailed 03/04/2008 to identify additional ventilator serial numbers that also need correction. |  | 4809 | 4809 (2427 US, 2382 OUS) | Worldwide Distribution --- USA including Puerto Rico and countries of Algeria, Argentina, Australia, Austria, Bahamas, Belgium, Brazil, Canada, Chile, China Peoples Republic, Colombia, Denmark, Finland, France, Germany, Greece, Hong Kong, India, Indonesia, Israel, Italy, Japan, Malaysia, Mexico, New Zealand, Norway, Panama, Paraguay, Philippines, Portugal, Saudi Arabia, Serbia &amp; Montenegro, Singapore, South Africa, South Korea, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Vietnam, and West Indies. | Terminated | August 16, 2012 | 1637 | 20603 | Computer
Z-1479-2008 | 46888 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Sequoia Spinal System Surgical Kits; Pedicle screw spinal system; Non-sterile (kits intended to be sterilized prior to use); Rx Only; |  |  | February 13, 2008 | 2008 | Abbott Spine | Screw Head Dissociation: The polyaxial screw head may dissociate from screw shaft during surgery. Also, the polyaxial screw driver is experiencing difficulty (broken/bent driver shafts, difficulty attaching screws to the driver and jamming of the collet.) | Device Design | The recalling firm notified sales reps regarding the recall by phone on 02/13/08. The sales reps were instructed to discontinue kit use in surgeries and to return inventory. Surgeons who had used the Sequoia System in surgery were notified of the recall via letter on 02/20/08. Surgeons were asked to monitor for post-operative dissociation in implanted patients and to notify Abbott Spine sales reps if any incidences occur. |  | 38 | 38 surgical kits. | Nationwide Distribution including states of AR, AZ, CA, CT, CO, FL, IL, MD, MI, MN, NC, OH, OK, OR, PA, TX, VA and WI. | Terminated | July 28, 2008 | 166 | 38 | Computer
Z-1628-2008 | 47355 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | STERRAD 100S Sterilization System, Product Code 10101 |  |  | February 25, 2008 | 2008 | Advanced Sterilization Products | Oil mist filter fails allowing oil mist to be emitted into the vicinity of the sterilizer, which may result in the release of a mist, "haze," or "smoke." | Device Design | The firm sent an Urgent Device Correction letter dated 4/23/08, notifying customers of a situation that may occur with all models of STERRAD Sterilization Systems (100S, 50, 200, NX and 100NX). As a precaution, personnel should leave the room if mist is observed and avoid returning until the mist has cleared. ASP asks that customers circulate this letter to any STERRAD system users within their facility and also to the Chief Administrative Officer, the Director of Biomedical Engineering, Director of Materials Management and the Chief of Nursing. The letter instructs customers on what to do if they experience a filter failure - a) cancel the cycle if the sterilizer is running [and as always, do not use instruments from an incomplete cycle], b) leave the room, c) discontinue use of the STERRAD Sterilizer until the system is repaired. d) Personnel should avoid working in the room until the mist has cleared. The letter advises that ASP service will replace the sterilizer's filter and increase the frequency of filter changes by adjusting the Planned Maintenance (PM) schedule within the software on their STERRAD 100S and STERRAD NX Sterilizer. Per the notice, for each STERRAD Sterilizer model, ASP will do the following: a) STERRAD 200 Sterilizer - beginning April 28, 2008 through October 2008, ASP will install a new filter and gasket assembly at no charge. b) STERRAD NX Sterilizer - beginning April 28, 2008 through October 2008, ASP will install a new filter assembly at no charge and adjust the PM schedule to 650 cycles or 6 months (whichever occurs first). c) STERRAD 100S Sterilizer - beginning April 28, 2008 through October 2008, ASP will adjust the PM schedule to 750 cycles or 6 months (whichever occurs first). d) STERRAD 50 and 100NX Sterilizers - No adjustments to the oil mist filters or PM schedules are required at this time. PMs will be performed as currently scheduled. Information for MedWatch reporting is included in the letter. |  | 9950 | 9,950 | Worldwide Distribution -- including USA and countries of France, Germany, Italy, United Kingdom, Switzerland, Benelux, Uruguay, Venezuela, Argentina, Chile, Colombia, Ecuador, Belgium, Puerto Rico, Hong Kong, Indonesia, Mexico, Philippines, Australia, Austria, China, Greece, Ireland, Israel, Japan, Korea, Malaysia, Singapore, Taiwan, Thailand, Turkey, Peru, Canada, Czech Republic, Middle East, Portugal, S. Africa, Spain, Sweden, Hungary, Poland, Brazil, Russia, Egypt, India, and Slovenia. | Terminated | November 21, 2011 | 1365 | 9950 | Computer
Z-1530-2008 | 46372 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | BD FACSDiva Software v6.0 and 6.1 with the HP XW4400 Computer System |  |  | December 21, 2007 | 2007 | BD Biosciences | Values might not update in statistics views after certain functions are performed in a worksheet. | Software design | The firm issued an Important Product Information notification letter on 12/21/2007. The letter instructs users to reload data by closing and reopening an experiment before reporting results in order to determine if their statistics have been impacted. They are asked to fill out and return a form pertaining to this problem. BD Biosciences will release a version of BD FACSDiva software to correct this problem. The letter also stated that BD Biosciences rep will contact the users when the new software is available. User are asked to contact BD Customer Support Center at 1-877-232-8995 if they have any additional questions. |  | 74 | 74 units | Nationwide Distribution. | Terminated | September 24, 2008 | 278 | 74 | Computer
Z-1529-2008 | 46370 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Cytomics FC 500 Flow Cytometry System with Data Innovations Instrument Manager, Part Numbers 626553 and 6605628 |  |  | August 20, 2007 | 2007 | Beckman Coulter Inc | Labeling provided by two integrated software systems are not clear enough to avoid potential demographic and sample type mismatches under certain conditions. | Software design | A Product Corrective Action (PCA) letter, dated August 16, 2007, informed the customers that the User documentation and Online Help for the FC500 with Data Innovations lnstrument Manager are not clear enough to avoid potential demographic and sample type mismatches under certain conditions. The letter asks the customers to share the information with their laboratory staff and retain the notification as part of their laboratory Quality System documentation. They are also to review all sample identification, demographics, data and results before reporting results out. A response form is included. Questions and concerns regarding the notification should be directed to Beckman Coulter Customer Service at 1-800-526-7694. |  | 141 | 141 units | Worldwide Distribution --- including USA and country of Canada. | Terminated | October 15, 2012 | 1883 | 141 | Computer
Z-1181-2008 | 46846 | No | MEA | Infusion pump. | General Hospital | General Hospital | 510(k) | Hospira LifeCare PCA Label Utility, Version 1.0 CD. The installation/user guide is entitled "LifeCare PCA Barcode Label Utility Program." This software program is designed to allow entry of each drug concentration into a formulation library on a pharmacy computer for standardized printing of barcode labels for compounded drug product vials. |  |  | January 03, 2008 | 2008 | Hospira Inc. | Incorrect dosage labels created: Customers have experienced independent changes in drug concentration (mg/mL to mcg/mL and vice versa) and expiration dates, and the addition of trailing zeros to the barcodes. | Software design | Hospira sent letters dated 1/3/08 to the immediate recipients of the product via FedEx overnight letter on the same date. The letters informed the accounts of the reports of independent changes in drug concentration while editing a saved formulation using the Hospira LifeCare PCA Label Utility Software CD that was provided to some customers in a Hospira MedNet Drug Library Creation Team Introductory Kit-LifeCare PCA kit. The accounts were requested to remove the utility from any PC in their facility on which it is currently installed, destroy the CD and sign and return the enclosed removal verification form verifying that they have removed the software from their system. The letter listed commercially available software packages that the account could validate and implement in their facility. Any questions concerning the utility software withdrawal were directed to Hospira at 1-800-241-4002, option 4, then option 1. |  | 127 | 127 CDs | Nationwide Distribution and Canada | Terminated | March 09, 2010 | 796 | 127 | Computer
Z-1190-2008 | 46753 | No | MED | Liquid chemical sterilants/high level disinfectants. | General Hospital | General Hospital | 510(k) | Steris C1160 Universal Flexible Processing Tray, for use with STERIS SYSTEM 1 Processor, used for sterilization of flexible endoscopes. |  |  | February 19, 2008 | 2008 | Steris Corp | Failure to effectively sterilize: Under certain conditions, the design/operation of the C1160 Tray may compromise the ability of the SYSTEM 1 processor to correctly assess the state of the high pressure pump during the Sterile Cycle. The L6 sensor fails to alarm when the high pressure pump is not working properly. | Device Design | Steris sent an Urgent Safety Alert/Recall letter dated 2/8/2008, to all customers. The letter stated that any device sterilized in the C1160 Tray should not be used clinically prior to a completion of a successful diagnostic cycle. Until further notice, you must run a diagnostic cycle after each sterile cycle involving a C1160 tray or alternately you must discontinue use of these trays. Steris service technicians will schedule visits to correct the problem. |  | 28766 | 28,766 units | Worldwide Distribution, including USA, Puerto Rico, Canada, and Japan. | Terminated | February 24, 2012 | 1466 | 28766 | Computer
Z-1459-2008 | 46852 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | High-Definition 120 MultiLeaf Collimator, Medical charged-particle radiation therapy system; Software Version 7.0; Product Code: H56; Accessory to Medical Linear Accelerator. |  |  | January 30, 2008 | 2008 | Varian Medical Systems Oncology Systems | Dose Calculation Error: Software anomaly may result in failure of a MLC leaf to reach planned position, potentially resulting in misadministration of dose to a patient. | Software design | Consignees were notified by a Medical Device Notification letter sent on 01/29/2008. The letter advised users to cease use of the device until a corrected version of the software is available. A representative will contact users to upgrade the application once the revised software has been validated and released for distribution. |  | 3 | 3 units subject to correction | Nationwide Distribution including states of NC, FL and CA. | Terminated | August 18, 2010 | 931 | 3 | Computer
Z-0274-2008 | 47235 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Precision 500D Classical R&F; System with control room PC; Image-intensified fluoroscopic x-ray system; ; Model Number: 5117866-27. |  |  | August 01, 2007 | 2007 | GE Healthcare | Inaccurate cassette size display reading/ Inaccurate mAs reading: 1. When the device is in PBL mode and a larger film cassette is inserted into the table Bucky (replacing a smaller cassette) after a protocol is selected on the console, the collimator will not adjust to the larger cassette size. The collimator stays at the original smaller cassette size and allows the operator to take an exposure. | Software design | Consignees were notified by two "Product Safety Notification" letters on July 10, 2007. Each letter addressed one of the two safety issues with the affected product. The letters provided safety instructions on how avoid the software error and informed users that their representative would contact them to schedule a visit for a software correction installation. |  | 101 | 101 | Worldwide Distribution --USA including states of: AK, AL, AR, CA, CO, CT, FL, GA, IA, ID, IL ,IN, LA, MA, MD, MI, MO,MS, NC, ND, NE, NH, NJ, NM, NV, OH,OK, OR, PA, SC, SD, TN, TX, UT, VA, VT and WA; USA territory of: PR and countries: EGYPT, SAUDI ARABIA, MACEDONIA, FRANCE and GREAT BRITAIN. | Terminated | September 10, 2008 | 406 | 101 | Computer
Z-0246-2008 | 45153 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Biograph 16, Biograph PET/CT Scanner, Catalog No. 872858, Manufactured by Siemens Medical Solutions USA, Inc., 810 Innovation Drive, Knoxville, TN 37932 USA, TEL: 865-218-2534 |  |  | July 23, 2007 | 2007 | Siemens Medical Solutions USA Inc | Visual asymmetric appearance; An asymmetry can be introduced into the attenuation-corrected PET images when using iterative reconstruction in combination with a specific number of subsets and reconstruction matrix sizes. (This problem only occurs when using protocols other than the default Biograph protocols) | Software design | A certified letter-Urgent Device Correction- dated June 15, 2007 was sent to all affected consignees on July 23, 2007 informing them of the problem (PET iterative reconstruction algorithm) and requesting as a short term fix that they discontinue the practice of operating using protocols other than the default Biograph. Software upgrades as the long term fix were promised in the Urgent Device Correction Notice. Updated information received October 4, 2007: a software patch to fix PETsyngo 5.0 was released, software patch to fix PETsyngo 4.x was released, and target date for the software patch to fix PETsyngo 3.2 and 3.4 is the end of November 2007. |  | 162 | 162 units | Worldwide-USA and countries of Argentina, Austria, Australia, Belgium, Brazil, Canada, Switzerland, Chile, China, Czech Republic, Germany, Denmark, Spain, France, Great Britain, Hong Kong, Hungary, Ireland, India, Italy, Japan, South Korea, Kuwait, Lebanon, Mexico, Malaysia, Netherlands, Norway, Poland, Portugal, Sweden, Singapore, Slovak Republic, Thailand, Turkey, and Taiwan. | Terminated | September 11, 2009 | 781 | 415 | Computer
Z-1460-2008 | 46919 | No | KPQ | Radiation therapy simulation system. | Radiology | Radiology | 510(k) | Acuity with Conebeam Computed Tomography; Radiation therapy simulation system; Model Number: H77. |  |  | January 03, 2008 | 2008 | Varian Medical Systems Oncology Systems | Incorrect Image Orientation: The image orientation tag may not be set correctly if the patient has been scanned with an orientation other than Head First Spine. When such images are imported to a treatment planning system, they may appear with either a mirrored or a rotated view. | Software design | Consignees were notified by an Urgent Medical Device Correction letter issued on 02/05/2008. The letter provided recommended actions for users to follow in order to avoid the software error. The letter informed users that a service pack is being created to correct the issue and that the recalling firm will notify users when it becomes available. For additional information, contact 888-827-4265. |  | 16 | 16 UNITS | Worldwide Distribution--USA including states of CA, MA and PA, and countries of UK, Chile, Belgium, Canada and Germany. | Terminated | September 30, 2009 | 636 | 16 | Computer
Z-1557-2008 | 47456 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Medi-Lifter III Plus Total Lift, Model 7000-2000 and Summit Total Lift, Model 4800-2000; patient lifts with mast scale (600 lbs. capacity); |  |  | April 10, 2008 | 2008 | Arjo, Inc. | Scale could break and fall potentially causing injuries to someone who is on it. There is a potential for the mast scale load cell ssembly to fracture and allow the mast/boom to fall. Should a patient be in the lift at the time, the patient could fall and sustain injuries. | Device Design | Arjo sent Urgent Device Recall Customer Notification letters dated 4/1/08 on the BHM Medical letterhead to the end user accounts who received the affected patient lifts with mast scales, advising them of the potential of fracture at the load cell assembly which might cause the scale to fall. Should a patient be in the lift at the time, the patient could fall, potentially sustaining injuries. The accounts were instructed to immediately stop using the patient lifts. A BHM representative will contact the accounts within 5 days to arrange for the correction of the lifts. The accounts will be given the option of returning the lifts for repair (with shipping costs covered by BHM), or purchasing a replacement lift at a significantly reduced price. Any questions or concerns were directed to the BHM Service Department at 1-819-868-0441 or service@bhm-medical.com. |  | 37 | 37 lifts | Nationwide Distribution --- including states of California, Florida, Illinois, Michigan, Minnesota, New York, Ohio, Oklahoma, Oregon, Pennsylvania, Texas, Washington, and Wyoming. | Terminated | November 18, 2009 | 587 | 37 | Computer
Z-2037-2008 | 48755 | No | HGM | Perinatal monitoring system and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | GE Healthcare Centricity Perinatal (formerly QS) System - PFILS (Patient File Server) software; automatic patient data management providing clinical information at the bedside in Labor &amp; Delivery, Mother-Baby and the Neonatal Intensive Care Unit. |  |  | June 30, 2008 | 2008 | GE Healthcare Integrated IT Solutions | Patient information mix-up: When the PFILS application, configured with a backup server has experienced a network interruption, PFILS verify could result in the Centricity Perinatal application recording non-identifiable patient information to the incorrect patient file. | Software design | GE Healthcare notified their Centricity Perinatal customers via letter dated 7/20/07 of the problem with the PFILS application that could result in the Centricity Perinatal application recording non-identifiable patient information to the incorrect patient file. The letter provided the users with a workaround and provided them with a Frequently Asked Questions document which listed the availability of a software patch to fix the problem. A follow-up letter dated 6/27/08 was sent on 6/30-7/2/08 to those accounts who had not requested a PFILS Hotfix CD-ROM or had not had their software recently upgraded. The letter a copy of the 7/20/07 letter, the most recent FAQ, the PFILS Hotfix CD (PFIL67_66) and Installation Procedure for the PFILS Patch instructions for installation of the software patch. The accounts were advised that they can request an analysis of existing patient records to determine whether they have been affected by the PFILS issue, by submitting a non-critical Technical Support Request on the GE Centricity Perinatal Web site: http://www.gehealthcare.com/usen/security/cis/stupgrades.html, indicating "PFILS ANALYSIS REQUEST" in the Description of Issue field. Any questions were directed to Technical Support at 800-433-2009. |  | 1031 | 1,031 units | Worldwide: USA, Puerto Rico, Canada, and China. | Terminated | September 06, 2010 | 798 | 1031 | Computer
Z-1373-2008 | 46751 | Yes | MRM |  |  | Cardiovascular | PMA | Lumax Cardiac Resyncronization Therapy Defibrillator (CRT-D); Implantable, dual-chamber defibrillator; Model Number: HF-T; Catalog Number 355263; Software Version 603.U; Catalog Model Number 359251; ICS 3000 units are composed of three Catalog/Model Numbers: 349527, 349530, and 355547. The model 349530 is the component which is loaded/updated with software. |  |  | January 10, 2007 | 2007 | Biotronik Inc | Timing Feature Error: An implementation error allows the usage of the V-V timing feature, which is not approved for use in the U.S. | Software design | Consignees were notified by a letter sent 1/10/07. The letter advised users to stop usage of the affected software. Users were instructed not to implant additional devices until the software program has been corrected. The recalling firm requested that all affected software be scrapped or returned to them. Users were requested update the software with a corrected version as soon as it becomes available. |  | 53 | 53 of the 603.U discs distributed. | Nationwide Distribution including states of AL, AZ, CA, CO, DE, FL, GA, IL, LA, MI, MO, NC, NJ, NV, NY, OH, OR, PA, SC, TX and WA. | Terminated | May 05, 2011 | 1576 | 53 | Computer
Z-1571-2008 | 47253 | No | LHN | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | EUB-5500 VISION 5500 Diagnostic Ultrasound Scanner w/Step 5 (V05-7) Software. Imaging device intended to provide the physician with physiological and clinical information. |  |  | February 25, 2008 | 2008 | Hitachi Medical Systems America Inc | Miscalculation reading. A software error in the firm's EUB-5500/HI VISION diagnostic scanning system causes a miscalculation of the left ICA/CCA ratio when using the Carotid Calculation package for patient scans. | Software design | A Device Correction Letter was issue by the firm on 02/25/2008. The letter asks the affected customers not to use the value of ICA/CCA on the Measurement Report for diagnosis, but instead to follow the Immediate countermeasure detailed in the 'Note in Use' document which accompanies the Device Correction letter. The customers are asked to follow the Immediate Countermeasure until they receive a Temporary Countermeasure to correct the ICA/CCA ratio calculation. The Temporary Countermeasure will include: 1) a Disk to import a new carotid measure package; 2) an Installation booklet detailing how to import the new carotid measurement package and optional customization tools; and 3) an Operator manual addendum Upon release of the countermeasure, the firm's Applications or Servcie personnel shall perform the corrective action as identified in the Service Pack (A) Installation Instructions. Finally, upon release of the next software revision, service personnel will replace the software at each customer location. |  | 20 | 20 units | Nationwide Distribution --- including states of OH, FL, CA, TX, DE, VA, NY, MO, KS, NJ, MD and ID. | Terminated | February 24, 2012 | 1460 | 20 | Computer
Z-1902-2008 | 47409 | No | KFM | Nonroller-type blood pump. | Cardiovascular | Cardiovascular | 510(k) | Levitronix CentriMag Back-Up Console |  |  | March 17, 2008 | 2008 | Levitronix, Llc | Stops pumping: Interruption of CentriMag System support may occur when using a Valleylab Force FX-C Electrocautery Unit | Device Design | Levitronix issued an Urgent - Voluntary Device Correction letter on 3/17/08 mailed to the US distributor and requested to contact their accounts. The letter revises the warning in the CentriMag Console Manual. Levitronix is issuing revised manuals.On March 13, 2008 the firm emailed and mailed (FedEx) to the one Canadian Hospital and Health Canada was informed of the Voluntary Advisory Notice via email on Wednesday, March 19, 2008.Rest of the World: Wednesday, March 12, 2008 throughMonday, March 17, 2008. The Notice was issued as aVoluntary Advisory Notice after review and follow up with theUK MHRA. On 7/24/08 the firm issued a revised Recall Letter to advise users not to use Valley Lab Force FX-C or SSE2L Electrosurgery device with the firm's CentriMag Blood Pumping System and include the significant risk to health. The firm revised their letter and Black Box Warning and reissued the letters to the hospital, surgeons, etc. on 7/24/08.On 2/5/09 the firm issued a "Follow-Up to Urgent Device Correction of 3/17/08 " to advise users of the software upgrade to be installed in the CentrigMag Primary software Version 3.0 and CentriMag Back Up console , software version 2.0. Site visits will be made by Levitronix between Feb 2009 and April 2009 , software upgrade installed and new Operating Manuals provided. |  | 237 | 237 units | Worldwide Distribution: USA, Canada, Switzerland, UK, Italy, France, Greece, Spain, Slovenia, Czech, Turkey, South Africa, Argentina, Chile, Panama, Singapore, Taiwan, Israel, and Saudi Arabia. | Terminated | July 06, 2010 | 841 | 867 | Computer
Z-1474-2008 | 47225 | No | CAW | Portable oxygen generator. | Anesthesiology | Anesthesiology | 510(k) | Medline EndurO2 Oxygen Concentrator without Oxygen Monitor; Portable oxygen generator; Reorder Number: HCS021; Packaged in 1 unit per case; Product is an oxygen generator designed to provide oxygen to patient who has been prescribed oxygen therapy by a physician. |  |  | March 04, 2008 | 2008 | Medline Industries Inc | Overheating of Components: There is the potential for an electrical component in the oxygen concentrator to overheat, causing damage to the unit and making it nonfunctional. | Device Design | Consignees were notified by a letter sent on 3/4/08. The letter informed users that the recalling firm will replace the affected units with new units or issue a credit for the returned recalled units. Distributors were requested to notify their accounts of the recall. The letter requested users to complete and return an enclosed response sheet, indicating the number of units to be replaced or returned for credit. For additional information, contact 866-359-1704. |  |  |  | Nationwide Distribution | Terminated | June 17, 2010 | 835 |  | Computer
Z-0472-2009 | 49517 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | GE Centricity Ultra Laboratory System Software; for recording, reporting and distribution of lab results. GE Healthcare Integrated IT Solutions, Barrington, IL 60010. |  |  | September 22, 2008 | 2008 | GE Healthcare Integrated IT Solutions | Software computer, error: The abnormal result flag is not being sent in the HL7 messages on results that are lower than the normal range. For further information, please telephone GE Healthcare Integrated IT Solutions at 800-437-1171. | Software design | GE Healthcare notified their Centricity Ultra Laboratory Information System customers via letter dated 9/19/08. The letter was entitled "PRODUCT SAFETY NOTIFICATION." The letter states that the software correction was installed or will be installed to correct the abnormal results flag not being sent in the HL7 on results that are lower than the normal range. The letter summarized the safety concern and the action taken. Any questions were directed to GE Technical Support at 888-778-3373 or 905-479-1249. The accounts were requested to complete the enclosed customer reply form, supplying the software version number, and fax it back to 847-277-5240. |  | 3 | 3 units | Globally - United States including the states of New York and Texas; and the country of Canada. | Terminated | May 18, 2010 | 603 | 3 | Computer