## Sheet1
Recall Number | Recall Event ID | Device Type | Product Code | Regulation Description | Medical Specialty | Review Panel | Submission Type | Trade Name/Product | Main Name | Recall Class | Date Posted | Year Posted | Recalling Manufacturer | Reason for Recall | FDA Determined Cause | Action | Consumer Instructions | Clean Quantity | Quantity in Commerce | Distribution | Recall Status | Termination Date | Time to Terminate | Merged Quantities | Fault_Class
Z-2041-2009 | 52804 | No | HTT | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | ST ULTRAPOWER DIAMOND WHEEL (RED) 25.4mm, Catalog Number 00702109200, STERILE, Single-use. Made is USA. |  |  | July 31, 2009 | 2009 | Linvatec Corp. dba ConMed Linvatec | ConMed Linvatec is recalling Ultrapower Bur, Diamond Wheel for possibility of a packaging breach in the sterile barrier. | Package design/selection | Firm notified international and domestic consignees by an Urgent Medical Device Recall letter on July 31, 2009. Customers were asked to check their inventory; complete, sign, and return the enclosed Reply Form; and return the devices for credit. Questions should be directed to Customer Service at 1-800-237-0169. |  | 2502 | 2502 | Worldwide Distribution | Terminated | May 11, 2010 | 284 | 2502 | Not_Computer
Z-0619-2009 | 50235 | No | MEG | Piston syringe. | General Hospital | General Hospital | 510(k) | BD 1 ml Integra TB Syringe with Retracting Precision Glide Needle; 1 ml 27 g 1/2 (0.4mm x 13 mm); Sterile; REF # 305298 |  |  | November 24, 2008 | 2008 | Becton Dickinson &amp; Company | Expiry date on the shelf carton is incorrect. It reads 2013-13, it should read 2013-03. | Incorrect or no expiration date | Recall letters were sent on November 24, 2008 to all direct customers and potential end users of the product by UPS second day air. BD requests that inventory be examined and affected product be quarantined and returned as per recall letter instructions. Questions should be directed to BD customer service at 888-237-2762, M-F 8.OO AM to 5:00 PM EST. |  | 334400 | 334,400 units | GA, KS, OH, MO, FL, PA, KY and NY. | Terminated | March 27, 2009 | 123 | 334400 | Not_Computer
Z-0818-2009 | 50232 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Tornier AFFINITI Humeral Stem Standard 10mm, Catalog No. 0020010. Sterile. Manufactured for Tornier, 10750 Cash Rd., Stafford, TX 77477. Tornier's Affiniti Shoulder System is a non-constrained glenohumeral prosthesis intended for use as a total or hemi-shoulder replacement system. The total shoulder consists of a metal humeral stem, a metal humeral head and an ultrahigh molecular weight polyethylene glenoid. The hemi-shoulder consists of a metal humeral stem and a metal humeral head. |  |  | October 20, 2008 | 2008 | Tournier, Inc. | Tornier, Medical has received a report of a mislabeled sterile AFFINITI head product. As a result, Tornier has decided to voluntarily recall the potentially affected lots, including two (2) lots of AFFINITI Stems. | Labeling Change Control | An Urgent Product Recall letter, dated 10/17/08, was hand delivered to end users by distributors and sales representatives beginning 10/20/08. The recall letter identified the issue and requested end users to immediately discontinue use of the affected products and return any unused affected products to Tornier. Distributors were to notify any and all accounts that may be end users or to provide this information to Tornier, whom will contact the end users. A questionnaire and return instructions were provided with the letter. The questionnaire was asked to be completed and returned to Tornier as soon as possible. For additional information, contact Tornier Customer Service at 1-888-494-7950. |  | 401 | 401 total | AR, CA, FL, NC, NE, OH, PA, TN, TX, WA, and WI | Terminated | May 07, 2011 | 929 | 401 | Not_Computer
Z-0929-2009 | 50770 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Somatom Definition AS, Model number 8098027. |  |  | December 09, 2008 | 2008 | Siemens Medical Solutions USA, Inc | possible improper installation of tilt transport lock/tilt limiter | Labeling Change Control | The recalling firm received a letter dated 11/25/08 from the manufacturer informing them of the problem. As a result, the recalling firm has issued Update Instruction CT065/08/S to have the transport lock/tilt limiter on the device visually inspected. The tilted scan modes have been disabled until the modified lock/tilt limiter has been installed. |  | 42 | 42 units | Nationwide Distribution --- including states of AZ, CA, CO, FL, IL, KS, KY, MA, MI, MN, MO, NE, NY, OH, PA, SC, VA, and WY. | Terminated | September 23, 2009 | 288 | 42 | Not_Computer
Z-2000-2009 | 51896 | No | DKB | Clinical toxicology calibrator. | Toxicology | Toxicology | 510(K) Exempt | Roche/Hitachi Preciset TDM I, Roche Diagnostics, Indianapolis, IN; Catalog No. 03375790 190. |  |  | April 22, 2009 | 2009 | Roche Diagnostics Corp. | The bottle caps may be more difficult to remove than usual, and minor injuries may be caused by the application of the force required to open the bottles. | Packaging process control | Roche Diagnostics Corp. notified Consignees via letter dated April 22, 2009 and instructed to not use excessive force when opening bottles and contact the firm for product replacement instead. A Fax Back Form was included for customers to return via fax to 1-888-345-0480.For further information, contact your Roche Diagnostics representative or call 1-800-428-5074. |  | 3970 | 3,970 bottles of both products | Nationwide Distribution | Terminated | December 07, 2009 | 229 | 3970 | Not_Computer
Z-0616-2009 | 50198 | bandage, elastic | FQM | Elastic bandage. | General Hospital | General Hospital | 510(K) Exempt | Esmark Elastic Bandage, 4" x 3 yards, sterile, for single use only, latex free. Distributed by Cardinal Health, McGaw Park, IL 60085-6787. |  |  | November 18, 2008 | 2008 | Cardinal Health | The natural rubber latex bandages were mislabeled as latex free. | Labeling Change Control | A recall letter dated 11/17/08 was issued to the Directors of Materials Management and sent via DHL 2-day service on 11/18/08. Accounts were advised that the elastic bandages containing latex were mislabeled as latex free. Customers were requested to return their inventories of the affected product to Cardinal Health for credit or replacement by calling the Presource Customer Sales Operations Group at 800-766-0706 for return instructions. Customers were also requested to complete the enclosed Cardinal Customer Acknowledgement form by indicating the amount of affected product being returned and fax it to 847-689-9101. Please direct questions to Quality Systems, Professional Services at 800-292-9332. |  | 9521 | 9,521 units (of all devices) | Worldwide Distribution including USA, Guam and Korea. | Terminated | November 04, 2009 | 351 | 9521 | Not_Computer
Z-0906-2009 | 50737 | No | EMG | Dental hand instrument. | Dental | Dental | 510(K) Exempt | Obwegeser Ramus Clamp w/ wide shield, 20 cm, 8", Right. ORAL-MAXILLO-CRANIO-FACIAL SURGERY, Original OBWEGESER Instruments. Walter Lorenz Surgical, Inc., Subsidiary of Biomet Microfixation, Inc. NOTE: None distributed domestically. |  |  | May 21, 2008 | 2008 | Biomet Microfixation, Inc. | The Biomet Microfixation Obwegeser Ramus Clamp, Wide Right 01-7996 were incorrectly etched as the left part number 01-7997. | Labeling mix-ups | Customers were contacted by e-mail notification with read receipt. Customers were notified in the e-mail to immediately discontinue use of the recalled product and return the affected items to Biomet Microfixation, Inc. The e-mail also included an Inventory Reconciliation form for customers to complete and return via fax (904-741-1801) to Biomet Microfixation, Inc. For questions related to this notice, please call 1-800-874-7711 or 904-741-4400 extension 448, Monday through Friday, 8 a.m.-4 p.m. Eastern Standard Time (EST). |  | 4 | 4 | International Distribution including countries of Ireland and Japan. | Terminated | September 11, 2009 | 478 | 4 | Not_Computer
Z-0550-2009 | 49510 | No | GFF | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Burrs |  |  | August 08, 2008 | 2008 | Sterilmed Inc | SteriMed, Inc is conducting a product removal of a small group of devices individually packaged and sealed within a pouch. The removal is being conducted because the possibility exists that some of the pouches were not properly sealed on one end.A breach in packaging seal or a failure in packaging integrity has the possibility of risk to the patient in terms of transmitting organisms capable | Packaging process control | Consignees were notified of the recall via verbal communication in person by on-site technician and or their local sales representative on 8/8/08. Others were contacted via fax or e-mail followed by on-site visits 8/11/08. SteriMed also provided consignees (Materials Manager/Risk Manager) with a "Urgent Product Removal" recall letter (August 14, 2008) in person informing them of the potential of receiving affected product. Included in the Product Removal Notification was a Customer Reply Form to inventory quantity of product received and quantity returned. Contact SterilMed, Inc. at 1-888-541-0078 for assistance. |  | 125 | 125 | AL, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KS, KY, MA, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, WA, WI, and WV OUS: CANADA | Terminated | April 26, 2011 | 991 | 5070 | Not_Computer
Z-0560-2010 | 53878 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Toggleloc System, Biomet Ziptight Ankle Syndesmosis Fixation System - SS with Ziploop Technology, polyethylene/polypropylene/ polyester/ stainless steel, sterile Biomet Sports Medicine, Warsaw, IN.; REF 909857. |  |  | November 10, 2009 | 2009 | Biomet, Inc. | Packages labeled as containing a stainless steel device actually contain a titanium device, and vice versa. | Labeling mix-ups | Distributors, hospital consignees and an implanting surgeon were notified of the recall by an Urgent Medical Device Recall Notice, dated 11/10/09. Distributors were instructed to locate and remove the identified products from circulation. The instructions on the enclosed "FAX Back Response Form" should be followed and product from their accounts should be returned. Hospital consignees and the implanting surgeon were informed of the recall and advised of the potential health hazards presented by the mix-up. |  | 24 | 24 of both products. | Nationwide Distribution -- Including states of California, Florida, South Carolina, and Virginia. | Terminated | November 03, 2010 | 358 | 24 | Not_Computer
Z-0903-2009 | 50512 | No | OEX | Catheter introducer. | Cardiovascular | Cardiovascular | Enforcement Discretion | Spectranetics SLS II 14Fr. Laser Sheath Kit, REF 500-012, Sterile EO, individually packaged, Lot number 080314B, Spectranetics, 96 Talamine Court, Colorado Springs, Colorado 80907. |  |  | October 15, 2008 | 2008 | Spectranetics Corporation | Catheter pouch was mislabeled with the wrong size. | Error in labeling | Consignees were notified by sales representatives who visually inspected their inventory. The sales representatives were given a form to fill out to confirm whether the consignee had the catheters with the serial number still in inventory. The forms were faxed back to Spectranetics indicating if the catheters were used or not and if they were not used, were they labeled correctly. For assistance, contact Spectranetics Corporation at 1-719-447-2319. |  | 24 | 24 units | CA, FL, MI, NC, NY, OH, PA, TX, and VA. | Terminated | February 26, 2009 | 134 | 24 | Not_Computer
Z-1965-2009 | 52798 | No | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | iLink Bone Void Filler. Packaged in a double Tyvek pouch, then placed in a box. Model Number: 50-05-0010, Exp. 2010-04. |  |  | August 26, 2008 | 2008 | Theken Spine LLC | Failure to label the product properly. Additionally, 15 pieces were reworked and the expiration date was misprinted as 2011-04 instead of the actual date of 2010-04. | Error in labeling | Theken Spine, LLC contacted the Consignee August 26, 2008 via telephone and instructed them to inventory their stock and return affected product to the firm.For further questions, contact Theken Spine, LLC at 1-330-475-8662. |  | 22 | 22 units | Nationwide Distribution -- TX. | Terminated | March 09, 2012 | 1291 | 22 | Not_Computer
Z-2228-2009 | 53121 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | St. Jude Medical FAST-CATH Hemostasis Introducer 7F, REF 406702, Length 12 CM, Lot 6372 Max. Guidewire O.D. .038", 2010-01, Manufacturer, St. Jude Medical, 14901 DeVeau place, Minnetonka, MN 55345-2126, Sterile EO. |  |  | August 17, 2009 | 2009 | St. Jude Medical | St. Jude Medical has initiated a voluntary recall of two lots of Fast-Cath Hemostasis Introducers-reorder number 406709, lot number 4219 and reorder number 406702, lot number 6372. They discovered that the product lots could contain some items which have been mis-labeled with the incorrect French size. In one case, the pouch label does not correctly identify the content of the pouch. In another | Error in labeling | Field Representatives were sent a "St Jude Medical Urgent Product Information" letter to be hand delivered to their customers. The letter listed the products and problem and advised consignees to discontinue use of the product. A form was attached for the return of products to the firm.Contact your local St Jude Medical representative for further information. |  | 2000 | 2,000 | Worldwide Distribution -- US (CA, MI, PA and TN), China and Belgium. | Terminated | December 17, 2011 | 852 | 2490 | Not_Computer
Z-0838-2009 | 50136 | No | GXI | Radiofrequency lesion probe. | Neurology | Neurology | 510(k) | NeuroTherm Simplicity III RF Electrode |  |  | July 28, 2008 | 2008 | Neurotherm, Inc. | Distal tip may detach from the probe | Labeling design | NeuroTherrn sales representatives were notifiedtelephonically or through e-mail. End users were notified through visits by NeuroTherm sales representatives on 7/28/08 and 8/15/08. For questions, please contact NeuroTherm, Inc. at 1-978-777-3916. |  | 71 | 71 units | CO, FL, IL, OH, TN, TX , VA, and WYForeign: Australia, China, and UK, | Terminated | September 24, 2009 | 423 | 71 | Not_Computer
Z-0857-2009 | 50477 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Artiste MV Digital Linear Accelerator, when used in combination with Syngo RT Therapist Version 4.1 build 67 and automatic Motion of Flat Panel, part number 8139789, manufactured by Siemens Medical Solutions, Oncology Care System. |  |  | October 23, 2008 | 2008 | Siemens Medical Solutions USA, Inc | Under certain conditions, the flat panel may move unexpectedly and collide with the patient. | Labeling design | The firm issued an Urgent Medical Device Customer Safety Letter, dated 10/23/2008. With the letter included a detailed description of the issue, the condition under which this issue may occur, and recommendation to avoid them. Customers are to read the information carefully and distribute to the responsible persons. The Customer Safety Letter should be included their syngo RT Therapist Express manual and remain there until the next release of software is installed. In the interest of safety, the preventive measures should be performed and all affected personnel should be informed immediately. The firm is tracking responses via return cards and completion protocol. |  | 31 | 31 consignees have the configuration noted above. | Worldwide Distribution --- including USA and countries of Netherlands, Germany, Poland, Norway, Australia, Belgium, France, New Zealand, Malaysia, Spain and Canada. | Terminated | December 09, 2010 | 777 | 31 | Not_Computer
Z-1715-2009 | 52217 | No | JDL | Hip joint metal/metal semi-constrained, with a cemented acetabular component, prosthesis. | Orthopedic | Orthopedic | PMA | CONSERVE, TOTAL A-CLASS HEAD W/BFH TECH., REF 38AM-3604, O.D. Size: 36mm, Neck Length: -3.5mm, Taper: SLT, Sterile R, 1 each, Rx only, Wright Medical Technology, Inc., 5677 Airline, Road, Arlington, TN 38002. |  |  | May 14, 2009 | 2009 | Wright Medical Technology Inc | Incorrect packaging, swapped sizes | Packaging process control | Wright Medical Technology, Inc. notified hospitals, surgeons and their distributors of the recall situation via an "Urgent: Medical Device Voluntary Recall" notice dated May 14, 2009 explaining the problem, the hazard involved and requesting confirmation that the product was removed from their inventory. The letter included a "Fax Back" response page to ensure the receipt of the notice.For further questions, contact Wright Medical Technologies, Inc. at 1-800-874-5630. |  | 14 | 14 units | Worldwide Distribution -- Including states of NC, TN, IA, IN, MD, WI, WV, OK, WA, PA, TN, TX, FL and CA and country of Canada. | Terminated | December 09, 2009 | 209 | 29 | Not_Computer
Z-2326-2009 | 51156 | No | JFY | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | DCL (Diagnostic Chemicals Limited) Creatine Enzyme Reagent (R2) Catalog Number: 265-30. |  |  | February 12, 2009 | 2009 | Diagnostic Chemicals, Ltd. | Labeled with extended expiration date (24 months instead of 18 months) | Labeling Change Control | Genzyme Corp. notified consignees by Letter dated February 12, 2009 via Federal Express. Accounts are requested to discontinue use and discard units. Customers were also asked to complete and sign the Confirmation of Notification the fax to 1-902-628-6504.For further information, contact Genzyme Technical Support at 1-800-565-0265. |  | 82 | 82 kits | Worldwide Distribution -- United States (AL, CA, NJ, MO, PA, KY, FL, LA, MN, NY, NC, OH, TX and UT), Canada and Korea. | Terminated | May 23, 2012 | 1196 | 107 | Not_Computer
Z-0595-2009 | 50111 | No | LYT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Ossur 3" Ceramic Tip Pin, 4 pack, Model Number: 517DS |  |  | September 26, 2008 | 2008 | Ossur | Ossur has initiated this voluntary recall as a precautionary measure because the firm is unable to guarantee the sterility of the affected products. | Packaging process control | Ossur initiated contact with the customers to locate the pins on Sept. 26, 2008with the Product Manager for this device contacting each consignee by phone.The customers were informed the recall was initiated because the firm was unable to guarantee the sterility of the affected products. Customers were asked to check their inventory for the affected parts. If products were located, they were informed that an Ossur customer service representative will be emailing you with return information any they are to return these pins to the firm as soon as possible. They were also told Ossur does not currently have replacement product available, but they will contact them immediately when they are available.If the products were not located, they were asked to determine where these pins are currently located, if they have been applied to a patient? If so, are they able to track down that patient and bring them in for an evaluation? |  | 12 | 3 pouches (4 pins/pouch) = 12 pins | Nationwide Distribution --- including states of FL, KS, MI, MO, NY, &amp; WA. | Terminated | April 27, 2010 | 578 | 43 | Not_Computer
Z-2044-2009 | 52814 | No | OMP | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Engenex Cover Strip; Product Code 446471; |  |  | July 29, 2009 | 2009 | ConvaTec | Labeling: The expiration date does not appear on the product label. | Labeling design | Urgent Product Recall letters dated 7/29/09, were hand delivered to all direct customers by ConvaTec Sales Reps. Customers were asked to return any unused products to their supplier. The letter also stated that a credit would be issued for the product. Customers are to contact Karen Krenta at the firm's Customer Interaction Center, 1-800-422-8811. |  | 128 | 128 units | Nationwide Distribution | Terminated | October 02, 2009 | 65 | 456 | Not_Computer
Z-2016-2009 | 51966 | No | LIF | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | HACH SteriChek Residual Peroxide Reagent Strips, 5/100-strip bottles per kit, Kit Part #811905, Bottle part #1905LB, Hach Co., Ames, IA. |  |  | April 21, 2009 | 2009 | Hach Co | Some of the reagent strip bottles contain the incorrect strips | Labeling mix-ups | Hach Company issued an "Urgent Medical Device Recall" letter dated April 21, 2009 explaining the reason for recall and requested product remaining in inventory be returned. Subrecall to the end-user level was requested. A Return Form was also enclosed to be completed to indicate the amount of product being returned.Between April 22-23, 2009, the firm telephoned their customers to assure they received the recall letter.For further information, contact Hach Customer Service at 1-800-548-4381 extension 3560. |  | 275 | 275 kits | Nationwide Distribution -- NJ, CA, IN, PA, NV, CA, and ME. | Terminated | October 07, 2009 | 169 | 275 | Not_Computer
Z-0854-2009 | 50112 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Integra Ruggles Surgical Instrument ACF Distraction Screws |  |  | October 22, 2008 | 2008 | Integra LifeSciences Corp. | A packaging error had occurred where 16 mm screws were packaged and labeled as 12 mm screws and vice versa. | Packaging | Urgent: Product Recall Notification letters were sent via Fed Ex on October 22, 2008. The letter asked customers to quarantine any of the affected stock on hand and contact their local Integra Sales Representative for further assistance and for arrangement of product returns. Customers were to fill out the attached "Recall Acknowledgement and Return Form" indicating whether or not they possess any of the product identified in the letter. Additional questions can be directed to Christin Roach at 609-936-5413.A Urgent: Product Recall Second Notification dated 1/5/2009 was sent to customers providing them with additional information on the packaging error and included immediate and/or product used during surgery. This notification also requested that all surgeons be notified who were or may have been provided with any of the screws affected by this recall. |  | 351 | 351 units | Nationwide Distribution | Terminated | July 15, 2009 | 266 | 351 | Not_Computer
Z-0119-2010 | 53492 | No | MCM |  |  | Ear Nose &amp; Throat | PMA | Nucleus Non-Magnetic Plug for removable magnet implants, REF Z50100, STERILE EO, Cochlear Limited. |  |  | September 30, 2009 | 2009 | Cochlear Americas Inc. | Cochlear implant device component was mis-labeled. | Labeling mix-ups | On September 28, 2009, each surgeon who received a non-magnetic plug after December 5, 2008 was sent a notification letter via courier. The letter described the mislabeling issue and actions for consignees. Customers were requested to return affected product (and the quantity discarded) to the firm and review patient records to verify the implantation of the non-magnetic plug and not the mislabled product (sterile magnetic). Direct questions about the recall by calling at 1-800-523-5798. |  | 10 | 10 units | Worldwide Distribution -- United States (AL, AZ, CA, GA, IA, NC, MA, MI, MN, MO, NY, OH, OR, PA, TN, TX and WA), Brazil, and Argentina. | Terminated | December 22, 2009 | 83 | 10 | Not_Computer
Z-1814-2009 | 52680 | No | OMP | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ConvaTec, Engenex Tunnel Dressing, for use with Engenex Negative Pressure Wound Therapy, (1 x 15 cm) |  |  | July 13, 2009 | 2009 | ConvaTec | Labeling: the actual manufacturing date (2/2009) and actual expiration date (2/2011) are switched. | Labeling Change Control | Recall notification letters were hand delivered on July 13, 2009 to all direct customers (tier 1) by ConvaTec reps. Letters were provided to all direct customers so that they could notify all customers that received the product. |  | 764 | 764 | Nationwide distribution and ConvaTec sales rep in the UK and France. | Terminated | October 02, 2009 | 81 | 764 | Not_Computer
Z-2250-2009 | 53080 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | B. Braun Blistered Omnifix L.L. Syringe packaged in the Easypump LT 60-24, 60 ml x 2 ml/hr |  |  | August 01, 2008 | 2008 | I-Flow Corporation | I-Flow identified a B. Braun syringe package that had a tear/hole in the paper portion of the primary package.Used to provide an optional filling device for the Easypump. | Packaging process control | I-Flow sent a recall letter dated August 1, 2008 to consignees telling them of the above mentioned recall and asking them to Inspect their Easypump inventory to determine if they have any of the affected lots identified on this Recall Notification Letter. They were told not to distribute affected devices and quarantined remaining inventory immediately. They were asked to check the appropriate response on the attached I-Flow Fax sheet and return the form back to I-Flow within 48 hours of receipt.The I-Flow Customer Service number (949) 206-2700, and hours of operation (7:00 am - 5:00 pm PST) were provided to answer questions. |  | 1608 | 1,608 units | France | Terminated | October 07, 2009 | 432 | 2177 | Not_Computer
Z-0884-2009 | 50498 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Restoration Modular Conical Distal Stem; |  |  | November 21, 2008 | 2008 | Stryker Howmedica Osteonics Corp. | Stryker became aware of a lot for lot mix up. | Labeling mix-ups | Recall notification letters sent via Federal Express on Nov. 21, 2008 to Branches/Agencies, Hospital Risk Management, Chief of Orthopaedics and Surgeons. Letters explain the labeling mix up and the potential risks to patients if the wrong size stem is used. Stryker branches are asked to examine inventory and that of hospital locations to identify affected product. All affected product is to be reconciled on the Product Accountability form and the form faxed back to Stryker. Contact Stryker Howmedical Osteonics Corporation for additional information at 1-201-831-5718. |  | 4 | 4 | Class II Recall - Nationwide Distribution | Terminated | June 25, 2012 | 1312 | 4 | Not_Computer
Z-2024-2009 | 52642 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | ev3 The Endovascular Company, Protege GPS 6mm x 80mm Self Expanding Peripheral Stent System , Rx Only, and SERP65-06-80-120. |  |  | June 10, 2009 | 2009 | Ev3, Inc. | ev3 Inc. is conducting a voluntary recall of specific lots of the Protege GPS Self-Expanding Stent System because labeling did not match the actual stent size contained in the package. The affected product was labeled as a 6mm x 80mm Protege GPS stent but actually contained an 8mm x 40mm stent. This situation can lead to vessel damage or inadequate coverage of the target lesion. | Labeling mix-ups | Consignees were emailed a letter with a subject heading of "Medical Device Recall". The e-mail was dated June 10, 2009. The letter informed the consignee of the problem and product involved. They requested that any affected devices remaining in inventory be identified and returned and to fill out and return the "Device Recall Field Action Reconciliation Form". |  | 2 | 2 | International Distribution Only -- Brazil, Uruguay, &amp; Malaysia. | Terminated | December 11, 2011 | 914 | 13 | Not_Computer
Z-0888-2009 | 50485 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Solar Shoulder Humeral Stem |  |  | November 13, 2008 | 2008 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics became aware of a visual defect in certain blister packages of certain devices. | Packaging | Recall notification letters were sent via Federal Express on November 21, 2008 to Stryker Branches/Agencies, Hospital Risk Management, Chief of Orthopaedics and Surgeon. The letters explained the problem and indicated that there are no risk factors. Contact Stryker Orthopaedics for additional information at 1-201-831-5718. |  | 322 | 322 total units - all types | Nationwide. OUS: Australia, Canada, Japan, Argentina, Brazil, Germany, Spain, Columbia, Guatemala and Sweden. | Terminated | January 20, 2010 | 433 | 322 | Not_Computer
Z-2028-2009 | 52810 | No | MCW | Intraluminal artery stripper. | Cardiovascular | Cardiovascular | 510(k) | Diamondback 360 Orbital Atherectomy Device, 1.50 LT Classic, Model #DB-150L, CSI Part # 70025-02 |  |  | April 16, 2009 | 2009 | Cardiovascular Systems, Inc. | CSI is correcting (8) packages of Diamondback 360 due to a missing "Use By Date" on the device labeling. The labelling should have been a 12 month shelf life (2010-02) rather than no information for shelf life. The units affected by this correction are sterilized products, and CSI believes that there is no immediate risk to health associated with these distributed products recognizing the valid s | Incorrect or no expiration date | Cardiovascular System, Inc. issued an "URGENT: Medical Device Correction" letter dated April 16, 2009. The letter described the product and the problem. The letter informed the consignees that the "Device" re-labeling would be performed by company representatives. Devices will not be returned.For further information, contact Cardiovascular System, Inc. 1-877-274-0901. |  | 8 | 8 | Nationwide Distribution -- CA, GA, MI, NY, NY and TN. | Terminated | December 10, 2011 | 968 | 8 | Not_Computer
Z-1930-2009 | 52000 | No | DCN | C-reactive protein immunological test system. | Immunology | Immunology | 510(k) | Dimension Flex Reagent Cartridge C-Reactive Protein Extended Range (RCRP). |  |  | April 15, 2008 | 2008 | Siemens Healthcare Diagnostics, Inc. | falsely elevated results in EDTA plasma samples | Labeling False and Misleading | Siemens Healthcare Diagnostics issued an Urgent Field Safety Notice on April 15, 2008 informing their customers that the EDTA plasma specimen claim had been rescinded. Current lots in the firm's inventory have been put on hold and reworked with an Alert Card rescinding the EDTA plasma claim. Subsequent lots use a revised instruction for use stating "Do not use EDTA plasma" in the Specimen Collection and Handling Section.For further questions, contact Siemens Healthcare Diagnostics at 1-302-631-6564. |  | 25760 | 25,760 cartons | Worldwide Distribution -- US, Austria, Bosnia, Czech Republic, Hungary, Slovakia, Serbia, Denmark, Estonia, Latvia, Lithuania, Norway, Sweden, Belgium, Cyprus, Canary Islands, France, Great Britain, Germany, Ireland, Italy, Netherlands, Poland, Russia, Portugal, Saudi Arabia, Spain, South Africa, Switzerland, Turkey, Yemen, Bahamas, Brazil, Canada, Mexico, Australia, Bangladesh, China, Hong Kong, Japan, Malaysia, Pakistan, and Taiwan. | Terminated | September 09, 2009 | 512 | 25760 | Not_Computer
Z-0871-2009 | 50546 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer Natural-Knee II System Modular Cemented Tibial Baseplate, Right, Size 3, For cemented use only, sterile, Zimmer, Warsaw, IN; REF 6420-01-230. |  |  | November 25, 2008 | 2008 | Zimmer Inc. | Sterility compromised: Lack of assurance of sterility, as a corner flap of the sterile barrier packaging on the inner cavity may be trapped in the seal of the barrier packaging for the outer cavity and the seal integrity of the outer sterile barrier cannot be assured. | Packaging process control | Zimmer distributors were notified by email of letter dated 11/25/08 on 11/25/08. This letter is entitled "URGENT: DEVICE RECALL." Secondary consignees/user accounts were notified by a copy of the 11/25/08 letter, which was sent on or before 12/05/08. Consignees were instructed to stop using the product and to return it to Zimmer. For further information, please contact Zimmer, Inc. by telephone by dialing 800-613-6131. |  | 10 | 10 | Nationwide Distribution --- including states of Arkansas, Florida, Michigan, North Carolina, New York and Wisconsin. | Terminated | November 12, 2009 | 352 | 10 | Not_Computer
Z-0932-2009 | 50680 | No | FTM | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | AMS InteXen LP Collagen Dermal Matrix, REF 72404066. Product Size: 8 x 12 cm. AMS InteXen LP is a sterile non-perforated processed porcine collagen dermal matrix. Sterile EO. American Medical Systems, Inc., 10700 Bren Road West, Minnetonka, MN 55343 USA. |  |  | December 05, 2008 | 2008 | American Medical Systems | The product size shown on the carton labeling does not match the size on the product inside. | Labeling mix-ups | Customers were first notified by phone on 12/05/08. They were told of the situation and were requested to return the devices. Customers were also provided an Important Product Recall letter, dated 12/05/08. This letter identified affected product, told customers to discontinue use of product and to return it. An acknowledgement form was also asked to be returned. AMS Customer Service should be contacted at 1-800-328-3881 to receive a Return Authorization number and arrange for shipment of replacement product. |  | 5 | 5 | Nationwide Distribution --- including states of TN and TX. | Terminated | June 05, 2012 | 1278 | 5 | Not_Computer
Z-0068-2010 | 53435 | No | MAF |  |  | Cardiovascular | PMA | Boston Scientific LiberteÂżÂż MonorailÂż and Over-the-Wire Coronary Stent Systems to be relabeled as VeriFLEX". |  |  | September 11, 2009 | 2009 | Boston Scientific Corporation | Boston Scientific initiated a field correction for the LiberteÂż Bare-Metal coronary stent products. They have received reports from cardiac cath labs that TAXUS Liberte Paclitaxel-Eluting (TAXUSÂż LibertÂżÂż Drug-Eluting) coronary stents have been inadvertently selected when the interventional cardiologist intended to implant a LiberteÂż Bare-Metal stent, and LibertÂżÂż Bare-Metal stents have been ina | Labeling design | Boston Scientific issued an "Urgent Medical Device Field Correction" letter dated September 18, 2009. The letter was addressed to Risk Manger/Field Action Contact. The letter described the product, problem and actions which included renaming the LiberteÂżÂż Bare-Metal stent as VeriFlex" Bare-Metal stent. Enclosed with the letter included the Direction for Field Correction and a Reply Verification Tracking Form.For further information, contact you local Boston Scientific Sales Representative or call 1-800-811-3211. |  | 391769 | 391,769 | Nationwide Distribution -- United States including Puerto Rico. | Terminated | June 28, 2012 | 1021 | 391769 | Not_Computer
Z-2244-2009 | 53143 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific Renegade" Hi-Flo Microcatheter, Peripheral Vascular Use. Sterilized with Ethylene Oxide. Legal Manufacturer: Boston Scientific Cork Ltd., Business &amp; Technology Park, Model Farm Road, Cork, Ireland. |  |  | August 26, 2009 | 2009 | Boston Scientific Corporation | Boston Scientific has determined that the sterile barrier in the packaging of the affected product may be compromised. It has been reported that the pouch seal (top seal) may be breached. | Packaging | Consignees were sent a Boston Scientific Urgent Medical Device Recall Removal - Immediate Action Required letter Dated August 26,2009. The letter was addressed to Risk Manager / Field Action Contact. The letter described the problem and product involved. Advised consignees to segregate the affected product and return it to Boston Scientific. Requested the return and completion of The Account Reply Verification Tracking form enclosed. Replacements will be issued for all recalled product that is properly returned to Boston Scientific. |  | 80885 | 80,885 (51,491 US; 29,394 OUS) | Worldwide Distribution -- USA, excluding New Mexico and Wyoming, and countries of Europe, the Middle East and Africa, Japan, American Virgin Islands, Argentina, Aruba, Australia, Austria, Bahrain, Belgium, Brazil, Canada, Chile, China, Colombia, Czech Republic, Denmark, Dutch Antilles, Egypt, Finland, France, Georgia, Germany, Great Britain, Greece, Honduras, Hong Kong, Hungary, India, Iran, Ireland, Israel, Italy, Jordan, Kuwait, Latvia, Libya, Luxembourg, Macau, Malaysia, Mexico, Namibia, Netherlands, New Zealand, Norway, Oman, Pakistan, Panama, Poland, Portugal, Puerto Rico, Romania, Saudia Arabia, Singapore, Slovakia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Ukraine, United Arab Emirates, and Vietnam. | Terminated | August 24, 2012 | 1094 | 118920 | Not_Computer
Z-0664-2009 | 50128 | No | LWJ | Hip joint femoral (hemi-hip) metallic cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stelkast Dual Taper Wedge Pressed Fit Femoral Component. Hip Prosthesis, Part number SC1439 |  |  | October 07, 2008 | 2008 | Stelkast Co | Sterility may be compromised due to breach in packaging. | Package design/selection | On 10/7/08, Stelkast telephoned the two customers of the problem and the need to return the products. |  | 6 | 6 units | USA distribution to dealers in TX and WV. | Terminated | April 01, 2009 | 176 | 6 | Not_Computer
Z-0542-2009 | 49382 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Roche ISE Internal Standard Gen.2 for use on Roche cobas c systems, Roche Diagnostics, Inc., Indpls., IN; Catalog no. 04522320190. Product is an in vitro diagnostic. |  |  | September 15, 2008 | 2008 | Roche Diagnostics Corp. | Test results, low: Sodium, potassium and chloride results on the cobas 6000 ISE unit may exhibit a downward drift in controls and patient samples within the recommended calibration interval of 24 hours. For further information, please contact Roche Diagnostics Corporation by telephone at 800-428-2336. | Package design/selection | Consignees were notified by letter dated 9/15/08, which provided a work-around for them to use until the problem can be permanently resolved. The letter was entitled "URGENT MEDICAL DEVICE CORRECTION." For further information, please contact Roche Diagnostics Corporation by telephone at 800-428-2336. |  | 7909 | 7909 | Nationwide Distribution. | Terminated | December 01, 2009 | 442 | 7909 | Not_Computer
Z-2085-2009 | 52683 | No | CAZ | Anesthesia conduction kit. | Anesthesiology | Anesthesiology | 510(k) | Single Shot Epidural Tray, Disposable, Convenience Kit, Reorder Number CUS511-08, Integra LifeSciences Corporation, Salt Lake City, Utah 84104. |  |  | July 13, 2009 | 2009 | Integra Life Sciences Corporation | Lack of packaging seal integrity on various pain management kits may compromise sterility of contents. | Packaging process control | Consignees were notified by telephone on July 09, 2009 and by "Urgent Medical Device Recall" letter on July 13, 2009. Customers were instructed to immediately examine inventory, cease use and distribution of any affected product and to contact Integra Pain Management Customer Service by calling 1-800-241-2210 to arrange for return and replacement. Refer questions to Integra Pain Management by calling 1-800-241-2210. |  | 350 | 350 units | Worldwide Distribution -- US (states of AK, AR, AZ, CA, CO, CT, FL, GA, HI, ID, IN, KY, ME, MO, NC, NJ, NV, NY, OH, OK, PA, RI, SD, ST, TX and UT), Belgium and Japan. | Terminated | October 23, 2009 | 102 | 3594 | Not_Computer
Z-0837-2009 | 50140 | No | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | Atlantis PV Peripheral Imaging Catheter, model number H749364560, catalog number 36456, manufactured by Boston Scientific Corporation, Fremont, CA. |  |  | October 09, 2008 | 2008 | Boston Scientific Corporation | Product is mislabeled as having a shelf life of 3 years, where the actual shelf life is 1 year. | Labeling mix-ups | Recall initiated on October 6, 2008, and notification letters sent on October 8, 2008. Letters were sent via FedEx. Responses are being tracked along with delivery verification. For additional information contact Boston Scientific Corporation at 1-508-652-5594. |  | 1033 | 1033 units | Worldwide Distribution --- USA including Puerto Rico and country of Panama | Terminated | July 27, 2009 | 291 | 1033 | Not_Computer
Z-0587-2009 | 49603 | No | LKK |  |  | General Hospital | PMA | Medtronic SynchroMed EL Programmable Pump, Model 8627-10. 10 mL Reservoir. The contents [STERILE] of the inner package have been sterilized by ethylene oxide gas. The implantable Medtronic SynchroMed EL programmable Pump is part of the SynchroMed EL Infusion System designed to contain and administer parenteral drugs to a specific site. The implantable components of the SynchroMed EL Infusion System include the pump with or without a side catheter access port, catheters, and catheter accessories. Medtronic, Inc., 710 Medtronic Parkway, Minneapolis, MN 55432-5604, USA. |  |  | August 25, 2008 | 2008 | Medtronic Neuromodulation | Medtronic issued a letter to healthcare professionals providing important safety information regarding MRI (magnetic resonance imaging) effects on SynchroMed EL and SynchroMed II implantable infusion pumps. As stated in product labeling, the magnetic field of an MRI will temporarily stop the rotor of the pump motor and suspend drug infusion for the duration of MRI exposure for all SynchroMed pump | Labeling Change Control | An Urgent Medical Device Correction, Important Information on Potential MRI Effects letter, dated August 2008, was sent to health care professionals beginning August 25, 2008. The letter described the issue, provided patient management recommendations and included enclosures: Roller Study Procedure to confirm pump function, Post MRI Pump Interrogation Guidelines and New Information Regarding Potential MRI Effects. |  | 76253 | Approx 76,253 of all devices. | Worldwide Distribution --- USA including Washington D.C., and Puerto Rico, and countries of Algeria, Aruba, Austria, Belarus, Belgium, Croatia, Cyprus Czech Republic, Denmark, Egypt, Finland, France, French Polynesia, Germany, Greece, Hungary, Iceland, India, Iran, Iraq, Ireland, Israel, Italy , Jordan Kuwait, Lebanon, Luxembourg, Malta, Netherlands, Netherlands Antilles, Norway, Pakistan, Poland, Portugal, Reunion, Romania, Russian Federation, San Marino, Saudi Arabia, Serbia and Montenegro, Slovakia, Slovenia, Sough Africa, Spain, Sweden, Switzerland, Tunisia, Turkey, United Arab Emirates, United Kingdom, Vatican City State, New Caledonia, Brazil, Dominican Republic, Mexico, Hong Kong, Singapore, Canada, Australia and Japan. | Terminated | May 17, 2012 | 1361 | 76253 | Not_Computer
Z-2078-2009 | 52211 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | DYONICS 25 Fluid Management System Disposable Inflow/Outflow Tubeset With Forked Suction |  |  | May 27, 2009 | 2009 | Smith &amp; Nephew, Inc. Endoscopy Division | The sterility of the device cannot be assured | Package design/selection | Smith &amp; Nephew notified direct accounts by letter dated May 27, 2009 . International accounts notified by Smith &amp; Nephew Inc. Endoscopy Division to each affected Outside United States (OUS) country or distributor by electronic communication, with telephone follow-up as required. The distributor is expected to notify each of their customers that has been confirmed to have product included in this recall. |  | 67697 | 67,697 units | Worldwide Distribution -- USA, Korea, Australia, Hong Kong, India, Mexico, New Zealand, United Kingdom, Spain, Singapore, Austria, Poland, Russia, Finland, Denmark, Japan, Switzerland, China, Italy, Puerto Rico, France, Asia Pacific, Belgium, Canada, Norway, Thailand, The Netherlands, Malaysia, Brazil, Argentina, Chile, Colombia, Costa Rica, Venezuela, Turkey, Israel, Germany, UAE, Greece, and South Africa. | Terminated | May 08, 2012 | 1077 | 67697 | Not_Computer
Z-2111-2009 | 50802 | No | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | I-Flow Corporation Easy Pump LT 60-24 (60 ml x 2 ml/hr), REF: 04434366, assembled in Mexico. |  |  | February 11, 2008 | 2008 | I-Flow Corporation | I-Flow Corporation received two complaints involving Easypump LT 60-24 (Lot 612996) containing a 100 ml x 2 ml/hr pump in the sterile package instead of a 60 ml x 2 ml/hr pump (Illustration was included to show consignee what both the correct and incorrect part looks like). Lot 612996 was manufactured on January 6, 2006. As of February 7, 2008, there have been no further complaints received on t | Labeling mix-ups | I-Flow Corp. issued an "Important Recall Information" letter dated February 7, 2008 informing Consignees to quarantine and return any remaining inventory and notify any end users of the affected lot.For further information, contact I-Flow Corp. at 1-949-206-2653. |  | 4450 | 4,450 Distributed | Nationwide Distribution to B. Braun, Spain, who then further distributed to Italy, Germany, Netherland, Latvia, Spain, Portugal and Singapore. | Terminated | February 04, 2010 | 724 | 4450 | Not_Computer
Z-0869-2009 | 50535 | No | HRY | Knee joint femorotibial metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Triathlon PKR Insert X3 #1 LM/RL - 8mm, 5630-G-108, Lot code NXDMEE |  |  | November 25, 2008 | 2008 | Stryker Howmedica Osteonics Corp. | Label mix-up: Stryker Orthopaedics became aware of a lot for lot mix up between Triathlon PKR Insert X3 #1 LM/RL 8mm, 5630-G-108 lot code NXDMEE and Triathlon PKR Insert X3 #1 RM/LL- 12 mm 5630-G-122 lot code M8JMEE. | Labeling mix-ups | Recall notifications entitled "URGENT PRODUCT RECALL" were sent via Federal Express on 11/25/2008 to Stryker branches, hospital risk management, Hospital Chief of Orthopaedics. For further information, please contact Stryker Howmedica Osteonics Corporation by telephone at 201-831-5718. |  | 29 | 8 mm - 1 unit; 12 mm - 8 units. | Nationwide distribution: Product was distributed to 7 Stryker branches and one hospital. | Terminated | August 30, 2010 | 643 | 29 | Not_Computer
Z-2015-2009 | 52826 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific Expo, Angiographic Catheter, Sterilized with ethylene oxide gas, Made in Mexico, Parque Industrial La Mesa, Edificion G1, Suites B. H, Tijuana, Baja California 22650 MEX. |  |  | August 05, 2009 | 2009 | Boston Scientific Corporation | Boston Scientific is initiating a recall removal of numerous batches/lots of Wiseguide Guide Catheter, Impulse Angiographic (Diagnostic) Catheter, Expo Angiographic (Diagnostic) Catheter. Boston Scientific has determined that the sterile barrier in the packaging may be compromised. During internal inspections, they discovered gaps between the side seals and the top seal approximately 3/8 in len | Packaging | Boston Scientific issued an "Urgent Medical Device Recall Removal" letter dated August 5, 2009. The letter was addressed to Risk Manger / Field Action Contact. The letter described the problem and product involved. The Firm recommended to segregate affected product immediately and return to Boston Scientific and requested return of the "The Reply Verification Tracking Form" enclosed with the letter.For further questions, contact your local Boston Scientific Sales Representative or call 1-763-494-1634. |  | 367898 | 367,898 (US 167,496; OUS 224,292) | Worldwide Distribution -- US (including DC and Puerto Rico), Japan, United Arab Emirates, Albania, Austria, Belgium, Bulgaria, Bahrain, Switzerland, Czech Republic, Germany, Denmark, Algeria, Egypt, Spain, Finland, France, Great Britain, Greece, Hungary, Ireland, Israel, Iraq, Iceland, Italy, Jordan, Kuwait, Lebanon, Lithuania, Libya, Morocco, Netherlands, Norway, Pakistan, Poland, West Bank Gaza Strip/ Palestine, Portugal, Qatar, Romania, Russian Federation, Saudi Arabia, Sweden, Slovakia, Syria, Turkey, Ukraine, South Africa, Argentina, Australia, Brazil, Canada, Chile, Colombia, Dominican Republic, Ecuador, Guatemala, Indonesia, India, South Korea, Mexico, Malaysia, Nicaragua, New Zealand, Panama, Philippines, Puerto Rico, Paraguay, Singapore, El Salvador, Thailand, Taiwan, Uruguay, American Virgin Islands and Vietnam. | Terminated | January 07, 2012 | 885 | 787578 | Not_Computer
Z-1196-2009 | 50202 | No | MDS |  |  | Clinical Chemistry | PMA | Abbott brand FreeStyle Navigator, Receiver Warranty Kit, Blood Glucose Monitoring System, Part Number: 70789-01. |  |  | November 01, 2008 | 2008 | Abbott Diabetes Care, Inc. | The product has the Units of Measure set to mmol/L rather than mg/dL. The blood glucose results could be misinterpreted. This may lead to under treatment and potential for hyperglycemia. | Packaging process control | The Consignees of the affected FreeStyle Navigator Receiver Kit was contacted via telephone, Federal Express or an Abbott representative beginning on November 1, 2008. Contact Customer Service at 1-866-597-5520 with further questions. |  | 2 | 2 units | US Distribution to states of NY and AZ. | Terminated | June 03, 2009 | 214 | 2 | Not_Computer
Z-1706-2009 | 52561 | No | MMP | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General Hospital | 510(k) | GE Medical Systems, OEC, Vascular Drape, 00-902776-01, Sterile. 150" long x 87" wide. Distributed by GE OEC Medical Systems, Salt Lake City, UT 84116 USA. |  |  | August 15, 2008 | 2008 | GE Medical Systems, LLC | GE Healthcare has found that the sterile products manufactured by Contour Fabricators, Inc. (CFI) for GE healthcare (GEHC) may have seals that could be compromised during the shelf life of the products. This issue could result in the product not maintaining its sterile state. Eleven complaints and 3 MDR's have been filed. No illness or injuries reported. | Packaging process control | GE Healthcare issued an Urgent Recall Notice dated August 15, 2008. The letter described the problem, provided a list of products, instructed customers to stop using the products and to return or dispose. To receive credit, customers must complete the attached product return form and include it in the package with returned items. All customers were asked to complete and fax back the attached response form. |  | 2760 | 2760 | Nationwide Distribution -- AZ, CA, FL, MD, NY, OR, PA, SC and TX. | Terminated | May 15, 2012 | 1369 | 5820 | Not_Computer
Z-0192-2010 | 53543 | No | KFM | Nonroller-type blood pump. | Cardiovascular | Cardiovascular | 510(k) | Revolution Blood Pump, "revOlution Centrifugal pump Catalog# 050300000***Rx ONLY***SORIN GROUP ITALIA via Statale 12 Nord, 86 41037 MIRANDOLA (MO) (Italy) Distributed in the U.S. by: Sorin Group USA, Inc. 14401 W. 65th way Arvada (CO) 80004-3599" |  |  | September 25, 2009 | 2009 | Sorin Group USA, Inc. | A weakness in the plastic packaging tray could cause the tray to crack; compromising product sterility. | Packaging change control | Customers were notified by an Urgent: Field Safety Notice, dated 9/25/09, and instructed to: check their inventory for the affected product, complete the customer response form and return or visually inspect all affected product before use. For additional information contact Sorin Group Customer Service at 1-800-650-2623. |  | 808 | 808 units | Worldwide Distribution -- US, Canada, Japan, and Europe. | Terminated | December 14, 2009 | 80 | 1064 | Not_Computer
Z-1553-2009 | 51951 | No | GZL | Depth electrode. | Neurology | Neurology | 510(k) | Cyberkinetics Inserter Wand &amp; V2 Trigger Assembly, Expiration: July 14, 2009; Catalog Number: 4468. |  |  | February 04, 2009 | 2009 | Cyberkinetics Neurotechnology Systems, Inc. | Sterility of product cannot be assured | Packaging process control | Cyberkinetics Neurology Systems (CKI) notified the Distributor, Blackrock Microsystem, of the affected device. Blackrock notified the final user by email on February 4, 2009. |  | 2 | 2 units | Nationwide Distribution - State of MA. | Terminated | August 20, 2009 | 197 | 2 | Not_Computer
Z-1053-2009 | 50914 | No | JDC | Elbow joint metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer Coonrad/Morrey Total Elbow, Interchangeable Humeral Assembly, extra small, 4 inch length, sterile, Zimmer Inc., Warsaw, IN; REF 32-8105-027-04. |  |  | January 30, 2009 | 2009 | Zimmer Inc. | The surgical technique is inadequate for extra-small sized implants, resulting in the possibility of the provisional becoming stuck in the bone. | Labeling design | Consignees were notified by "Urgent Device Correction" letter dated January 30, 2009 and provided revised surgical techniques for these extra-small implants. Direct questions about this recall to Zimmer, Inc. at 574-267-6131. |  | 14613 | 14,613 of all products | Worldwide Distribution, including USA, Australia, Austria, Belgium, Brazil, Canada, China, Colombia, Czech Republic, Denmark, England, Finland, France, Germany, Hong Kong, Hungary, India, Italy, Japan, Jordan, Latin America, Netherlands, Russia, Singapore, South Africa, South Korea, Spain, Switzerland, Taiwan, Thailand, Turkey and Venezuela. | Terminated | December 22, 2009 | 326 | 14613 | Not_Computer
Z-0181-2010 | 53458 | No | MGR |  |  | General &amp; Plastic Surgery | PMA | Organogenesis, Inc., Apligraf (draftskin) packaging, Canton, MA 02021. |  |  | September 27, 2009 | 2009 | Organogenesis, Inc. | Unit contaminated with Staphylococcus epidermidis | Packaging process control | Organogenesis, Inc. initiated a recall notification dated September 28, 2009 via fax to Physicians asking for return of unused affected product. If the unit was used, Physicians are advised to use standard wound care precautions to assure the safety of the patient. Users were asked to acknowledge receipt by completing and returning the notification to the firm.For further information, contact Organogenesis, Inc. at 1-888-432-5232. |  | 60 | 60 units | Nationwide Distribution | Terminated | March 31, 2010 | 185 | 60 | Not_Computer
Z-1576-2009 | 52170 | No | MSH | Surgical apparel. | General &amp; Plastic Surgery | General Hospital | 510(k) | Inovel N95 HealthCare Particulate Respirator, Model/Catalog Number: FRN95-SEZ, Expiration Date: 07/13/2010. |  |  | July 23, 2008 | 2008 | Inovel LLC | Display boxes labeled as FRN95-MLEZ, Medium/Large Size Respirators were inadvertently mixed with Display Boxes labeled for FRN95-SEZ Small Size Respirators. | Packaging process control | The recalling process was initiated on July 23, 2007. The affected distributor was notified of the recall through a certified recall letter with an attached Customer Response card. The distributor was instructed to notify and obtain recall product from their customers and return all affected product to Inovel. Direct questions about the recall to Inovel at 1-866-546-6835. |  | 80 | 80 cases | US (Illinois only). | Terminated | July 17, 2009 | 359 | 80 | Not_Computer
Z-1552-2009 | 51785 | No | OEU | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | Enforcement Discretion | Medtronic Ideal MAXPAC Total System, Catalog Numbers: TS1262R6, TS1262R7, TS1420R, TS1420R1, TS1442R, TS1446R, TS1464T, TS1470R, TS1470T1, TS1480R3, TS1480R5, TS1480R6, TS1487R, TS1487R2, TS8040R28, TS8108R6, TS8163R10, TS8387R5, TS9412R3, TS9453R28, TS9764R9, TS9800R28, TS9800R29, TS9800R31, TS9987R33 and TS9987R34. |  |  | March 23, 2009 | 2009 | Medtronic Cardiovascular Revascularization &amp; Surgical Therap | Medtronic has identified a small number of One Source Kits and MAXPAC/Total System lots with incorrect Use Before Dates on the outer kit package label. The Use Before Date on the outer label of the kit/pack does not accurately reflect the Use Before Date of the components within the kit/pack. For affected product, the outer label lists a date which exceeds the earliest Use Before Date of the com | Packaging process control | An "Urgent Medical Device Recall Notice" dated March 23, 2009 was sent to customers describing the affected device. Consignees were asked to quarantine product, complete and sign a certificate indicating acknowledgement of the notification and the amount of affected product at the facility, and to return the affected product.For further questions, contact Medtronic at 1-763-514-9605. |  | 1363 | 1363 kits total (1215 US, 148 OUS) | Worldwide Distribution -- US including states of AZ, CA, CO, CT, DE, FL, HI, IL, IN, LA, MI, MN, NC, NJ, NM, NV, NY,OH, OK, PA, TN, TX, and WI and countries of Australia, Japan and Venezuela. | Terminated | May 22, 2012 | 1156 | 1363 | Not_Computer
Z-1020-2009 | 51177 | No | NQO |  |  | Orthopedic | PMA | Spacers of varying lengths used with the InSTANT Stranding System, Part numbers: TG3138, TG3139, TG3140, TG3141 and TG3142. Theragenics Corporation, 5203 Bristol Industrial Way, Buford GA 30518. |  |  | December 08, 2008 | 2008 | Theragenics Corporation | Some pouches were found to have weak seals potentially compromising sterility. | Packaging process control | Consignees were notified first by telephone on December 8, 2008 and sent a follow up letter dated January 13, 2009. They were instructed to return any of the affected lots remaining in their inventory. |  | 62 | 62 units | Nationwide Distribution -- states of VA, WI, and CO. | Terminated | April 06, 2009 | 119 | 62 | Not_Computer
Z-1073-2009 | 50520 | No | LWJ | Hip joint femoral (hemi-hip) metallic cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Dual Taper Wedge Pressed Fit Femoral Component. |  |  | October 31, 2008 | 2008 | Stelkast Co | A possible breach of the sterile barrier packaging (Tyvek pouch). | Packaging | The recalling firm telephoned the distributors to inform them of the problem and the need to return the product. |  | 187 | 187 units | Nationwide Distribution -- states of NC, TX, and WV. | Terminated | April 14, 2009 | 165 | 187 | Not_Computer
Z-1160-2009 | 50867 | No | LRS | Intravascular catheter. | General Hospital | General &amp; Plastic Surgery | Enforcement Discretion | Tri-State Centurion PICC Insertion Tray, sterile, 15 per case, Tri-State Hospital Supply Corp., Howell, MI.; Reorder CVI2415A. |  |  | January 14, 2009 | 2009 | Tri-State Hospital Supply Corporation | Lack of assurance of sterility, as the package seals may be inadequate. | Packaging process control | Tri-State Hospital Supply Corporation notified all consignees via Urgent Product Recall Notice letter dated January 14, 2009. Consignees were instructed to remove the product from inventory, to inform the firm of amount on hand, and to hold the product for retrieval. |  | 60 | 60 | Nationwide. | Terminated | December 31, 2009 | 351 | 1186 | Not_Computer
Z-1898-2009 | 52762 | No | LJT | Subcutaneous, implanted, intravascular infusion port and catheter. | General Hospital | General Hospital | 510(k) | SmartPort CT Port Access System, Catalog number CT80STPD, AngioDynamics Incorporated |  |  | July 16, 2009 | 2009 | Angiodynamics, INC | The product may contain an incorrect size catheter and introducer. | Packaging process control | Notification of the AngioDynamics, Inc. sales force began on July 16, 2009 and they were asked to assist in the return of all affected devices to the Manchester, GA division of AngioDynamics, Inc. Urgent Recall letters were sent to all customers on July 16, 2009, requesting the completion of a FAX BACK form. Those currently in possession of any remaining devices will be provided a Return Goods Authorization (RGA) number and pre-paid shipping boxes for return of the product to AngioDynamics, Inc. All returned devices will be physically and electronically quarantined until final disposition of the devices. All devices will be inspected for damage to the devices and any damaged devices destroyed (scrapped) via the AngioDynamics procedure for Disposition of Nonconformances, procedure. Non-damaged product may be reworked to remove the incorrect catheter and introducer and replaced with the correct catheter and introducer, repackaged and re-sterilized in accordance with the AngioDynamics procedure for Rework. If necessary, follow-up letters, phone calls or visits by the AngioDynamics sales representatives may be used to complete the effective checks. Level A (100%) of consignees to be contacted).Please contact your local representative or Recall Coordinator at 1-706-846-3126, for assistance. You may e-mail your questions to customerservice@angiodynamics.com. |  | 150 | 150 kits | AL, CA, CO, FL, GA, ID, IL, IN, KS, LA, MI, MO, MN, MS, NC, NY, OH, OR, PA, TN and TX. | Terminated | October 27, 2010 | 468 | 150 | Not_Computer
Z-0491-2010 | 50685 | No | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Zodiac Polyaxial Spinal Fixation System. A spinal interlaminal fixation orthosis is a device intended to be implanted made of an alloy, such as stainless steel, that consists of various hooks and a posteriorly placed compression or distraction rod. The device is implanted, usually across three adjacent vertebrae, to straighten and immobilize the spine to allow bone grafts to unite and fuse the vertebrae together. The device is used primarily in the treatment of scoliosis (a lateral curvature of the spine), but it also may be used in the treatment of fracture or dislocation of the spine, grades 3 and 4 of spondylolisthesis (a dislocation of the spinal column), and lower back syndrome. |  |  | November 03, 2008 | 2008 | Alphatec Spine, Inc. | Alphatec Spine recently discovered that this lot of screws has the incorrect length laser etched on one side of the screw body. As a result, they are conducting a voluntary recall of the Zodiac Titanium cannulated polyaxial screw 6.5 mm x 50 mm L. | Error in labeling | Alphatec Spine contacted consignees via phone on October 16, 2008 and requested that consignees immediately remove the Zodiac Titanium cannulated polyaxial screws (6.5 mm x 50 mm L) from the surgical set to help ensure that they will not be used in surgery. There is a script the company developed to assure consistent responses to customers, and a phone number for more information.The strategy calls for products to be returned to Alphatec Spine. The company will arrange to have Federal Express contact the consignees for pick-up of the screws. Follow up questions are directed to the company at 1-800-922-1356. |  | 56 | 56 | AZ, CA, FL, TX | Terminated | December 08, 2009 | 400 | 56 | Not_Computer
Z-1711-2009 | 52431 | No | EXQ | Urodynamics measurement system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Laborie Medical Technologies Urodynamic Catheters, Twin Lumen 6 Fr , 40cm For Cystometry. Catalog Number: CAT-206. |  |  | April 09, 2009 | 2009 | Laborie Medical Technologies Cor | Laborie Medical Technologies reports that there was a mix up in the labeling and packaging of their Urodynamic Catheters. There were 8Fr catheters mixed inside the 6Fr catheter packages. | Packaging process control | Laborie Medical Technologies notified customers by telephone/email beginning April 10, 2009 advising them of the mislabelled product.For further questions, contact Laborie Medical Technologies at 1-800-522-6743. |  | 28 | 28 boxes (10's) | Nationwide Distribution -- CA, GA, IL, KS, MD, PA, OH, OK, UT and WA. | Terminated | August 20, 2009 | 133 | 28 | Not_Computer
Z-1699-2009 | 52227 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer Allen medullary cement plugs, 1 - 20 dia. flange/10 mm dia. core, 1 - 24 mm dia. flange/12 mm dia. core (polyethylene with barium sulfate) with inserter, sterile, Zimmer, Warsaw, Indiana; Catatog Number: 00801102001. |  |  | June 08, 2009 | 2009 | Zimmer Inc. | The seal on the sterile barrier pouch may be inadequate, resulting in lack of assurance of sterility. | Packaging process control | Zimmer Orthopedic Surgical Products issued an "Urgent: Medical Device Recall" notice dated June 8, 2009 informing users of the affected devices and actions to be taken including the return of product and acknowledgement of receipt of notice by returning the Recall Certification Form by fax to 1-330-364-0974. |  | 133 | 133 | Worldwide Distribution -- United States, Canada and Japan. | Terminated | December 15, 2009 | 190 | 258 | Not_Computer
Z-1074-2009 | 51205 | No | KOD | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Magic3 (Intermittent Catheter), Sterile, Uncoated 6 French Pediatric Length REF # 52306 |  |  | January 14, 2009 | 2009 | Rochester Medical Corp | Rochester Medical observed some Magic3 and HydroSil(R) catheter packages in which the catheter was caught in the package seal during the packaging process. | Package design/selection | Consignees were contacted by telephone or in-person visit. The communication provided the reason for the recall and a list of affected products and their lot numbers, and instruction for the removal. Distributors were given instructions to notify their customers in turn. For the customers that were not reached via telephone call or in-person, a letter, "Voluntary recall of magic product samples," was used for the recipients of product samples and sent via e-mail. The letter was dated January 14, 2009 and asked customers to destroy all of their remaining sample inventory of the affected lot numbers. |  | 12 | 12 | Worldwide Distribution -- USA including states of CA, FL, GA, ID, IL, LA, MD, MA, MI, MN, NJ, NM, NY, OH, PA, and WA and countries of United Kingdom. | Terminated | April 20, 2011 | 826 | 1960 | Not_Computer
Z-1616-2009 | 52169 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer Alumina Ceramic Femoral Head, 12/14 taper, 32 mm diameter, neck length 0 mm, sterile, Zimmer, Warsaw, IN; Catalog Number: 6428-32-02. |  |  | May 29, 2009 | 2009 | Zimmer Inc. | The corner tab of the inner compartment lid was not folded under before the outer lid was sealed, resulting in lack of assurance of sterility. | Packaging process control | Zimmer, Inc. issued "Urgent: Device Recall" notices dated May 29, 2009 informing Consignees of the affected devices. Users were instructed to locate and return affected product and complete and return an Inventory Return Certification form via fax to the firm at 1-547-372-4265.For further questions, contact Zimmer, Inc. at 1-800-613-6131. |  | 147 | 147 of both products | Worldwide Distribution -- United States, Brazil, Canada, Chile, Czech Republic, Finland, France, Germany, Italy, Korea, Lebanon, Switzerland and United Kingdom. | Terminated | December 22, 2009 | 207 | 147 | Not_Computer
Z-1010-2009 | 50861 | No | LQR | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Boston Scientific Stonetome Stone Removal Device. Triple Lumen Sphincterotome, 5.5 Fr, 20 mm, 11.5 mm, sterile, Boston Scientific Corp., Natick, MA; REF 3511. |  |  | January 08, 2009 | 2009 | Boston Scientific Corp | The ultraviolet light barrier on the pouch may be discolored, which will allow ultraviolet light into the package and may cause the latex retrieval balloon to become brittle and to leak or burst when inflated. | Packaging | Consignees were notified via letter dated January 8, 2009 to discontinue use of the affected product and to return it to Boston Scientific. Customers were instructed to complete and fax the Reply Verification Tracking Form to Boston Scientific at 508-683-5578. Direct questions about this recall to your local Sales Representative. |  | 294 | 294 of all products. | Nationwide. | Terminated | October 22, 2009 | 287 | 294 | Not_Computer
Z-1555-2009 | 52007 | No | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Siemens Stratus CS Acute Care Troponin I CalPak. |  |  | April 28, 2008 | 2008 | Siemens Healthcare Diagnostics, Inc. | The firm received and confirmed that some lots of Cardiac Troponin I (cTnl) CalPaks (CCTNI-CR) may exhibit calibration issues that cause lower than typical signal during calibration with the current population of cTnl TestPaks (CCTNI). This issue may result in failed calibrations due to "Slope Error" messages or falsely elevated patient sample and quality control (QC) results on calibrations that | Labeling False and Misleading | Siemens Healthcare Diagnostics, Inc. issued an Urgent Field Safety Notice dated April 29, 2008 to all affected customers instructing them to discontinue use of 3 CalPak lots and reduce the shelf-life for the remaining lots.For further information, contact Siemens Healthcare Diagnostics, Inc. at 1-302-631-7672. |  | 4415 | 4415 cartons | Worldwide Distribution - including United States, United Arab Emirates, Austria, Bermuda, Belize, Canada, Germany, Egypt, India, Japan, Saudi Arabia and Bengal. | Terminated | July 14, 2009 | 442 | 4415 | Not_Computer
Z-0996-2009 | 51039 | No | KRO | Knee joint femorotibial metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer Most Hinged Tibia Baseplate Spacer, size 1, 6 mm, sterile, Zimmer, Warsaw, IN; REF 5020-06-001. |  |  | January 29, 2009 | 2009 | Zimmer Inc. | The package contains 11 mm screws instead of the specified 6 mm screws. | Packaging process control | The consignee in Germany was notified by an Urgent: Device Recall letter dated 1/29/09, and delivered by email on 2/19/09. The letter stated the reason for the recall and clinical implications. The firm asked the consignee to stop using the provisional and quarantine it immediately; physically count all affected product and record data on the Inventory Return Certification Form included with letter; fax a copy of the completed form to Zimmer, Inc at 574-371-8603; request an IRA from Lindsey Arnett; and return recalled product along with completed Inventory Return Certification Form and IRA to Simmer Distribution Center. The letter states that credit will be issued upon receipt of the returned product. Also, if product has been further distributed, a copy of the recall letter should be provided to the customers. Questions or assistance should be directed to Zimmer, Inc at 1-800-613-6131. |  | 3 | 3 | International Distribution only -- Country of Germany. | Terminated | November 12, 2009 | 287 | 3 | Not_Computer
Z-1703-2009 | 52289 | No | MRA |  |  | Orthopedic | PMA | Trident Acetabular System Surgical Protocol. Instructions for use of hip prosthesis. Literature Number LSP55. |  |  | June 09, 2009 | 2009 | Stryker Howmedica Osteonics Corp. | The Surgical Protocol has been modified to create separate and distinct surgical protocols, one for the TridentÂż PSL Shell (LSP68) and one for the TridentÂż Hemispherical Shell (LSP69) in order to clarify and highlight the differences in reaming technique. | Labeling design | Stryker sent Urgent - Product Correction - Surgical Protocol letters by Federal Express on June 9, 2009 to all Stryker branches/agencies, hospital Risk Management, hospital Chief of Orthopaedics, surgeons and all foreign consignees. |  | 3702 | 2376 in USA, 1326 internationally | Worldwide distribution. | Terminated | May 08, 2012 | 1064 | 3702 | Not_Computer
Z-1013-2009 | 50738 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Triathlon X3 Posterior Stabilized Tibial Insert |  |  | December 18, 2008 | 2008 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics became aware of a lot for lot mix-up between part numbers 5532-G-713, lot code LBD528 (Triathlon X3 Posterior Stabilized Tibial Insert, size 7-13 mm) and part 5532-G-813, lot code LBD512 (Triathlon X3 Posterior Stabilized Tibial Insert Size 8-13 mm) | Labeling mix-ups | Urgent Product Recall letters were sent via Federal Express on December 18, 2008 with return receipt to Stryker branches, Hospital Risk Management, Hospital Chief of Orthopaedics and Surgeons. The letter stated the issue, the potential hazards, and the risk mitigation factors. The firm asked customers to examine their inventory and hospital locations to identify affected product; retrieve all affected product and return it to their branch or agency warehouse for reconciliation; reconcile all products on the attached Product Accountability Form; and return affected product. Questions should be directed to Karen Ariemma at 201-831-5718. |  | 126 | 9 of lot code LBD528; 14 of lot code LBD512 | Nationwide Distribution: FL and IL. | Terminated | September 11, 2012 | 1363 | 126 | Not_Computer
Z-0486-2010 | 53555 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | E-Z CLEAN Non Stick Cautery Tip. Label reads in part: "MEGADYNE 11506 SOUTH STATE STREET DRAPER, UTAH 84020 USA***CAT/REF 0029M E-Z CLEAN NON STICK CAUTERY TIP***STERILE *** Expiration Date: 2104-07***" |  |  | October 13, 2009 | 2009 | Megadyne Medical Products, Inc. | Product labels reference an incorrect expiration date. Expiration date referenced on labels is 2104-07 instead of 2014-07. | Error in labeling | A notification letter dated October 13, 2009 was sent to the first level consignees. Consignees were instructed to examine their stock, remove affected product and contact Megadyne Customer Service at 1-800-747-6110 (United States) or 1-801-576-9669 for product return instructions. |  | 324 | 324 units | Worldwide Distribution -- United States (CA, FL, NE, LA, MT, OK, OR, PA, SC and UT) and New Zealand. | Terminated | February 17, 2010 | 127 | 324 | Not_Computer
Z-1601-2009 | 50093 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Duracon Total Knee, A-P Lipped Tibial Insert; 11MM |  |  | May 05, 2006 | 2006 | Stryker Howmedica Osteonics Corp. | A box packaged and labeled as a Duracon Total Knee Duration A-P Lipped Tibial Insert 11MM may actually contain a 9MM insert. | Labeling mix-ups | A notification letter and product accountability form were sent via Federal Express on May 5, 2006 with return receipt to branch accounts that had received the affected product. For questions about the recall contact Stryker Howmedica Osteonics Corp. by calling 1-201-831-5000. |  | 8 | 8 | Nationwide Distribution (AL, OH, MN, NM, and CA). | Terminated | July 07, 2009 | 1159 | 8 | Not_Computer
Z-1685-2009 | 50097 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Trident Hemispherical Acetabular Shell; 72mm, Arc Deposited, Hydroxylapatite Coated, Multi holed; Sterile; Stryker, France. |  |  | May 29, 2007 | 2007 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics became aware that Trident Hemispherical Shells, Catalog Number 508-11-74J, are actually marked and labeled as Catalog Number 598-11-72J, Lot Code 20282301 | Labeling mix-ups | Stryker Orthopaedics issued an "Important Market Withdrawal" notice via Federal Express dated May 29, 2007. Consignees were asked to locate and return affected product and fax the attached Product Accountability Form to 1-201-831-6069.For further questions, contact Stryker Orthopaedics at 1-201-831-5825. |  | 2 | 2 units | Distributed to 2 Stryker branches. | Terminated | August 17, 2009 | 811 | 2 | Not_Computer
Z-1061-2009 | 51136 | No | JDW | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Biomet Offset Tibial Tray Adaptor, 5.0 MM, sterile, Biomet Orthopedics, Inc., Warsaw, IN, REF 141491. |  |  | February 13, 2009 | 2009 | Biomet, Inc. | A 2.5mm locking insert was packaged with the 5.0mm Offset Tibial Tray Adaptor instead of the 5.0mm locking insert. | Packaging process control | Consignees were notified by "Field Safety Corrective Action" letter dated February 13, 2009 to conduct a recall to the hospital level and to locate and return the recalled product to Biomet Orthopedics. Customers were also instructed to complete the enclosed "Fax Back Response Form" and return it via fax to Biomet Orthopedics at 1-574-372-1683. Direct questions about this recall to 1-800-348-9500 or 1-574-372-3983, Monday through Friday, 8 a.m. to 5 p.m. |  | 17 | 17 | Worldwide Distribution including Germany, Italy, Netherlands, Spain and the United Kingdom. | Terminated | August 11, 2009 | 179 | 17 | Not_Computer
Z-1640-2009 | 51702 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | In Situ Bender Cutter Kit Wide Sterile, Catalog number 530.521S |  |  | April 03, 2009 | 2009 | Synthes Spine | Sterility may be compromised due to a ineffective package design. | Package design/selection | All affected Synthes DMF Sales Consultants were notified of the Medical Device Recall by an Urgent: Medical Device Recall letter notification which was sent via USPS certified mail on 4/6/09. The letter stated the problem and risk. The firm recommended that patients who have had a surgical procedure in which these products were utilized be monitored for signs/symptoms of infection for a period of six months following surgery. If customers have any of the identified devices, they are to call Synthes at 1-800-479-6329 to obtain a Return Authorization Number; complete the form; and return the letter with the recalled product to the address provided. Customers that do not have the identified product, are to complete the attached form and return the document to the address provided acknowledging their receipt of the recall information. Questions should be directed to their Synthes CMF Sales Consultant. |  | 2706 | 2706 units total | Worldwide Distribution -- USA including states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WV and the country of Canada. | Terminated | September 15, 2010 | 530 | 2706 | Not_Computer
Z-1396-2009 | 48692 | No | DDR | Myoglobin immunological test system. | Immunology | Immunology | 510(k) | Access Myoglobin Reagents Kits, Part Number: 973243. |  |  | September 27, 2007 | 2007 | Beckman Coulter Inc | Beckman Coulter initiated the field action as a result of recent internal testing following the CLSI EP-5A - Evaluation of Precision Performance of Clinical Chemistry Devices protocol, it was determined that the total imprecision for the Access Myoglobin Reagent is greater than or equal to 10% across the expected physiological range rather than the stated claim less than 8%. Based on internal tes | Labeling False and Misleading | An Urgent: Product Corrective Action (PCA) letter was mailed via US Postal Service on October 08, 2007 to the customers that received Access Myoglobin Reagent kits informing them about the issue. The letter also informed customers that the product insert in the Access Assay Manual will be revised with the new range to reflect the new imprecision claim. Customers were instructed to:(1) Share this information with their laboratory staff and retain the letter for theirQuality System documentation,(2) Complete and return the enclosed fax-back response form within 10 days so that Beckman Coulter, Inc. can be assured that the customer received the notification, and(3) Contact Customer Technical Support at 1-800-854-3633 or your local Beckman Coulter, Inc. representative for questions or additional information regarding the notification. |  | 20727 | 20,727 | Nationwide and Canada. | Terminated | May 27, 2009 | 608 | 20727 | Not_Computer
Z-0415-2010 | 53299 | No | NBH | Arthroscope. | Orthopedic | Orthopedic | 510(K) Exempt | AlMER ENDOFEMORAL 4MM OFFSET (PURPLE) |  |  | September 03, 2009 | 2009 | Smith &amp; Nephew, Inc., Endoscopy Div. | Device may have improper Laser Marking, 4mm device may be marked as 6mm and 6mm device may be marked as 4mm. | Labeling design | A notification letter was issued via Fedex on September 3, 2009 to domestic consignees and international consignees were notified via e-mail. The letter described the issue as the affected device may have improper laser marking. The 4mm devices are marked as 6mm and the 6mm devices are marked at 4mm. Customers were instructed to return the completed response form and affected product in inventory to Smith &amp; Nephew. Direct questions about the recall to your local Smith &amp; Nephew Sales Representatives or call 1-508-261-3655. |  | 134 | 134 individual units | Worldwide Distribution -- United States including Puerto Rico, Australia, Austria, Belgium, Canada, Costa Rica, Dubai, Finland, France, Germany, Great Britain, Netherlands, Norway, Portugal, Spain, Sweden, and Switzerland. | Terminated | May 08, 2012 | 978 | 134 | Not_Computer
Z-0455-2010 | 52886 | No | KWH | Toe joint polymer constrained prosthesis. | Orthopedic | Orthopedic | 510(k) | Tornier, LMP-50 (Lesser Metatarsal Phalangeal Implant), Size 50, Sterile EO, , Manufactured by Tornier, Inc, Edina, MN 55435. |  |  | April 06, 2009 | 2009 | Tornier, Inc. | Tornier, Inc. has received a report of a mislabeled product due to a product mix-up. The package label indicated that the product was a model LMP-50 however the product inside the package was Tornier FGT-50 implant (Primus Flexible Great Toe Implant With Grommetts). | Labeling mix-ups | Consignees were sent on April 6, 2009 a Tornier "Urgent-Product Recall" letter dated April 3, 2009. The letter was addressed to Healthcare professionals and distributors. The letter described the problem and product involved. Advised consignees to immediately discontinue use of the affected product and to return any unused affected product to their facility. Advised their distributors to notify all accounts. They also attached a questionnaire needed for inventory status. |  | 77 | 77 | Worldwide Distribution -- USA, including states of CA, FL, MD, MI, NJ, OH, and PA, and country of France. | Terminated | December 17, 2011 | 985 | 77 | Not_Computer
Z-1228-2009 | 51495 | No | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | MEDRAD Sterile Disposable Syringe Kit, catalog number DSK 130-Q, Angiographic Syringe for delivery of contrast media. |  |  | January 29, 2009 | 2009 | Medrad Inc | Possible compromised sterility due to holes in styrene tray | Packaging process control | On 1/29/09, Medrad issued an Urgent Medical Device Recall letter. Customers were requested to check their inventory for any of the affected batch. If the customer has affected product, they are instructed to contact the firm to make arrangement for the return of the product. Non responders were sent additional notifications via FedEx on 3/2/09 and 4/1/09. |  | 2958 | 2958 kits | Nationwide: AL, AZ, CA, FL, IL, LA, MA, MT, NC, NY, OH, TX, and WI. | Terminated | June 04, 2009 | 126 | 2958 | Not_Computer
Z-1351-2009 | 51409 | No | JDS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Stryker, Gamma3, The Compact Version of the Gamma Nail System, Operative Technique for Hip Fracture System, Trochanteric and Long Nails. |  |  | March 02, 2009 | 2009 | Stryker Howmedica Osteonics Corp. | Operative Technique: further clarification and guidance is needed to help reduce the potential for distal mis-drilling of the short Gamma3 Trochanteric Nails. | Labeling design | Stryker sent Urgent Product Recall letters by Federal Express on March 2, 2009 to Stryker Branches, Chief of Orthopaedics, Hospital Risk Management and Surgeons. Letters were sent to all who have used the Gamma3 products in the past. |  | 18232 | 18232 | Nationwide. | Terminated | August 17, 2012 | 1264 | 38964 | Not_Computer
Z-1319-2009 | 51439 | No | LRP | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Tri-State Centurion Endoscopy Tray, sterile, 10 per case, Tri-State Hospital Supply Corp., Howell, MI; Reorder MNS4635. |  |  | January 14, 2009 | 2009 | Tri-State Hospital Supply Corporation | Compromised sterility, as the package seals may be inadequate. | Packaging process control | A letter titled "Urgent Product Recall Notice" was sent by Tri-State Hospital Supply Corporation to all consignees via certified mail titled dated January 14, 2009. Consignees were instructed to remove the product from inventory, to inform the firm of amount on hand by sending the Accountability Information Form to the firm via fax (1-517-546-3356) or email and to hold the product for retrieval. Further questions should be addressed to Tri-State Hospital Supply Corp. at 1-517-546-5400 extension 1135. |  | 10 | 10 | Nationwide including states of NC, FL, TX, NJ, IL, MA, GA and IA. | Terminated | December 24, 2009 | 344 | 550 | Not_Computer
Z-1325-2009 | 51701 | No | NAW |  |  | Radiology | PMA | TheraSphere Administration Set (Part Number K125914-034). |  |  | February 24, 2009 | 2009 | MDS Canada Inc. DBA MDS Nordion | Instructions for the use of the pinch clamp that is an integral part of the TheraSphere Administration Set were not included with the distribution of the device. | Error in labeling | A User Bulletin (TS-2009-1) dated February 23, 2009 was issued to all TheraSphere user sites. The bulletin described the issue, affected units, and recommended actions for users. The bulletin instructed users on the proper use of the pinch clamp. MDS Nordion also updated the package insert with these instructions.For questions about the recall, customers in North America contact MDS Nordion Customer Service at 1-866-363-3330 and customers outside North America call 1-877-474-5999 and request Customer Service. Customers may also e-mail MDS Nordion at TheraSphere@mdsinc.com. |  | 26 | 26 | Worldwide Distribution -- US (states of AZ, CA, CO, FL, IL, IN, KS, MD, MI, NC, RI, OH, OR, PA, WI) and country of Canada. | Terminated | June 08, 2009 | 104 | 26 | Not_Computer
Z-1348-2009 | 51781 | No | ODP | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Synthes Zero-P StarDrive Screwdriver Shaft, T8, self-retaining, quick coupling, for use in spinal fusion surgery. Catalog # 03.617.902 Technique Guide. |  |  | February 10, 2009 | 2009 | Synthes Spine | Tip breakage, due to inadequate instructions for use. The Technique Guide provided incorrect information related to use of the Torque Limiting Attachment during screw insertion and final tightening. | Labeling design | All affected Synthes Spine Sales Consultants were notified of the Labeling Correction via email on February 10, 2009. They were instructed to review changes with their surgeons. They were supplied with an electronic version of the corrected Technique Guide on March 11, 2009. All affected Synthes Sales Consultants and Customers were notified of this labeling change via distribution of the updated Technique Guide and a notification letter. This packet was sent via USPS certified mail on March 26, 2009. To ensure a broad and accurate distribution of notification, affected consultants and customers are defined as having been shipped Zero-P 3.0 mm Titanium Cervical Spine Locking Screws. These devices are used only with the Zero-P system. |  | 1766 | 1766 units | Distributed Nationwide and Canada. | Terminated | October 05, 2009 | 237 | 1766 | Not_Computer
Z-0272-2010 | 53586 | No | HGB | Contraceptive tubal occlusion device (TOD) and introducer. | Obstetrics/Gynecology | Obstetrics/Gynecology | PMA | Richard Wolf Hulka Clip; for single patient use only; a sterile tubal occluding device; 2 clips per procedure tray/pouch, 10 pouches per shipping carton; Richard Wolf Medical Instruments Corp., Vernon Hills, IL 60061; item 4986.90. |  |  | October 16, 2009 | 2009 | Richard Wolf Medical Instruments Corp. | The sterile pouch seal may fail to remain intact during handling rendering the Hulka clip non-sterile. | Packaging | Richard Wolfe sent recall letters dated October 16, 2009 to the direct accounts via certified mail on the same date. The accounts were informed that the lot of Hulka clips did not meet their sterile seal burst test requirements, and that the sterile seal may fail to remain intact during handling, resulting in the clips becoming contaminated or non-sterile. They were instructed to inspect their inventory for the affected lot, quarantine any product found and call Richard Wolf at 800-323-9653, ext 225 or 380 to arrange for the return of the affected product for replacement. The accounts were also requested to complete and return the enclosed response form to Richard Wolf, even if they do not have any of the affected lot in stock. |  | 150 | 150 cartons | Nationwide | Terminated | March 16, 2011 | 516 | 150 | Not_Computer
Z-1197-2009 | 49019 | No | LTJ | Tumor-associated antigen immunological test system. | Immunology | Immunology | 510(k) | Beckman Coulter Access Immunoassay System Hybritech PSA Calibrators, Product Number: 37205. |  |  | January 08, 2008 | 2008 | Beckman Coulter Inc | Beckman Coulter incorrectly shipped Access Hybritech PSA calibrators containing dual calibration cards (Hybritech and WHO) to some Access Hybritech PSA Customers. The WHO calibration for PSA is not currently approved in the United States. | Packaging process control | An "Urgent: Product Corrective Action" letter dated January 10, 2008 was issued via fax or mail to customers. The letter described the issue and instructed consignees to discontinue use and discard the affected product. Consignees were also asked to complete and return a "Fax-Back Form" via fax (1-952-368-7610) or mail to Beckman Coulter, Inc. within 10 days of receiving the notification letter.Further questions should be addressed to Beckman Coulter Customer Technical Support at 1-800-854-3633. |  | 123 | 123 Units | Nationwide Distribution -- Including states of AL, AZ, CA, FL, GA, IA, IL, IN, KS, KY, MA, MD, MI, MN, MO, NC, NY, OK, OR, PA, SC, TN, TX, VA, WI and WV. | Terminated | March 15, 2012 | 1528 | 123 | Not_Computer
Z-1250-2009 | 51499 | No | MCW | Intraluminal artery stripper. | Cardiovascular | Cardiovascular | 510(k) | Diamondback 360, CSI ViperWire Orbital Atherectomy Guide Wire, REF: GW14335FRM. Sterilized with Ethylene Oxide. Manufactured for Cardiovascular Systems, Inc., St. Paul, MN 55112, USA. The ViperWire atherectomy guide wire has a shaft diameter of 0.014" and is constructed of stainless steel with a smooth finish and is silicone coated. The spring tip configuration is atraumatic, radiopaque, and can be bent to form a steerable system. |  |  | March 16, 2009 | 2009 | Cardiovascular Systems, Inc. | CSI is recalling 10 devices due to a mislabeling of the shelf life as shown on the device labeling. The labeling shown on the affected devices represents a 24 month shelf life rather than a 6 month shelf life that should have been indicated. | Incorrect or no expiration date | Customer letters were FedEx'd to CSI Sales Professionals on 03/16/2009, who delivered the letters and helped coordinate the removals of affected product. The letter identified the product, provided a description of the issue and provided instructions to immediately remove from use and return the affected guidewires to CSI. A Customer Acknowledgement Form was asked to be completed and returned with the devices. |  | 10 | 10 devices | Nationwide Distribution | Terminated | April 20, 2012 | 1131 | 10 | Not_Computer
Z-1251-2009 | 51212 | No | KRO | Knee joint femorotibial metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Modular Replacement System (MRS) Cemented Stems, Sterile, Catalog Numbers: 6485-3-000, 6485-3-008, 6485-3-009, 6485-3-010, 6485-3-018, 6485-3-019, 6485-3-300, 6485-3-308, 6485-3-309, 6485-3-310, 6485-3-318, and 6485-3-319. |  |  | January 30, 2009 | 2009 | Stryker Howmedica Osteonics Corp. | Instruction for use incorrect: There are no statements or warning on labels or Instructions for Use that 8, 9 and 10 mm diameter MRS Cemented Stems are indicated only for use in the GMRS Proximal Tibial and Distal Femoral Components. | Labeling False and Misleading | Urgent Product Recall letters, dated 1/30/2009, were sent to all Branches/agencies, Hospital Risk Management, Chief of Orthopaedics and Surgeons on January 30, 2009 via Federal Express. The letter asked consignees to examine their inventory and determine if they had any affected products, remove and quarantine affected product from their inventory, and complete the attached Product Recall Acknowledgment Form. The letter also stated that a Stryker Orthopaedics Sales representative would contact the consignee to reconcile and return the product . Consignees should contact their Stryker sales rep if they weren't not contacted by them. Questions should be directed to Colleen O'Meara at 201-831-5970. |  | 1161 | 1161 units | Nationwide Distribution | Terminated | April 08, 2010 | 433 | 1161 | Not_Computer
Z-1525-2009 | 52093 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Stryker Xia Dual Anterior Staple; |  |  | April 12, 2007 | 2007 | Stryker Spine | Labeling Issue: The large Xia Anterior Staple was laser marked with an "S", indicating small. | Error in labeling | Market Withdrawal letters were sent to Stryker branches who received product on April 13, 2007 by Federal Express. The letters asked that they examine their inventory and hospital locations to identify the implant; reconcile these implants on the attached Product Accountability Form and fax a copy to (201) 760-8456; retrieve and return the affected implants; and contact their Stryker Customer Service rep to re-order the product that they are returning as part of the regulatory action. |  | 38 | 38 in US | Nationwide Distribution. | Terminated | June 29, 2009 | 809 | 38 | Not_Computer
Z-1481-2009 | 52005 | No | HTY | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Zimmer Steinmann Pins, 7/64 in. (2.8 mm) diameter, 9 in. (229mm) length, style 6, nonsterile, qty 6, Zimmer, Warsaw, IN; Model Number: 00-0187-03-69. |  |  | May 01, 2009 | 2009 | Zimmer Inc. | The package contains pins with an incorrect diameter. | Packaging process control | Zimmer, Inc. issued an "Urgent: Device Recall" letter informing users of the affected product. Users were instructed to locate and quarantine all unused devices and ship them to the firm with a completed Inventory Return Certification form. Return Certification forms are also to be sent via fax to Zimmer, Inc. at 1-574-372-4265.For further questions, contact Zimmer, Inc. at 1-800-613-6131. |  | 37 | 37 | Worldwide Distribution -- US including states of CA, GA, IL, MA, MO, OR, TX and WI and Japan. | Terminated | November 12, 2009 | 195 | 37 | Not_Computer
Z-1559-2009 | 52041 | No | DQA | Oximeter. | Cardiovascular | Anesthesiology | 510(k) | Dixtal Novametrix Neonatal/Pediatric SpO2 Sensor, Foam Wrap Style |  |  | May 13, 2009 | 2009 | Dixtal Medical, Inc. | Neonatal/Pediatric SPU Sensors may be incorrectly packaged in pouches labeled for Adult/Pediatric SPU Sensors. | Labeling mix-ups | Urgent Product Recall Notification was issued to affected accounts by email/facsmile on May 8, 2009 requesting to cease distribution, quarantine and return recalled product. Contact Dixtal Medical with questions; to inquire about your on-hand inventory; or to arrange to receive replacement product at Customer Service: sales@dixtal.com, Phone Number 1-203-269-1112, Fax Number 1-203-269-1760. |  | 675 | 27 Boxes (Packaged 25 sensors/box) | Worldwide Distribution -- US (TX and CT) and countries of Italy, Japan, France and Turkey. | Terminated | March 23, 2012 | 1045 | 675 | Not_Computer
Z-1480-2009 | 52009 | No | JDZ | Cement mixer for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | Biomet Optivac Total Hip Kit, Maximum Capacity Cartridge, 1 kit, sterile, latex free, Biomet Orthopedics, Inc., Warsaw, IN; Product Number: 418000. |  |  | April 09, 2009 | 2009 | Biomet, Inc. | The package may contain two short mixing rods instead of one long and one short rod. | Packaging process control | Biomet Orthopedics issued an "Urgent Medical Device Recall Notice" dated April 9, 2009 instructing distributors to visit each hospital consignee, deliver a copy of the recall notification and replacement product and to pick up, and return the affected product.Further questions may be addressed to the Biomet Orthopedics at 1-800-348-9500. |  | 250 | 250 | Nationwide Distribution - Including AL, AR, FL, IL, LA, MN, NE, NC, NV, OH, OK, SC, SD, WA and WI. | Terminated | October 06, 2010 | 545 | 250 | Not_Computer
Z-0478-2010 | 53867 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | ProtÂżgÂż EverFlex" Self-Expanding Biliary Stent System (for United States distribution). 6 x 150mm, Model # PRB35-06-150-080. Sterile EO, ev3 Inc, 4600 Nathan lane North, Plymouth, MN 55442-2920, Rx only. |  |  | November 06, 2009 | 2009 | Ev3, Inc. | ev3 has determined that a lot of 150mm ProtÂżgÂż EverFlex Biliary Stent System contains a 100mm length Self-Expanding stent. The implantation of a stent 50mm shorter than expected may not cover the full length of the target lesion and may require subsequent intervention. | Labeling mix-ups | Two separate letters were sent to United States and international consignees on November 9, 2009. The letter to US consignees was addressed to Cath Lab Managers while the letter to foreign consignees was addressed to the Risk Manager or Cath Lab Manager. The letters described the affected product, gave an Issue Summary and Required Action, which advised consignees to locate and remove the recalled product. An Ev3, Inc. sales representative will contact consignees to arrange return of all unused product to Ev3, Inc. Direct questions regarding the recall to Ev3, Inc. by calling Customer Service at 1-800-716-6700. |  | 6 | 6 (United States) and 6 (Outside United States) | Worldwide Distribution -- United States (TX and FL), GERMANY, FRANCE and UNITED KINGDOM. | Terminated | March 08, 2012 | 853 | 6 | Not_Computer
Z-1656-2009 | 51923 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Welch Allyn AED 10 automatic external defibrillator and MRL Jumpstart (collectively "AED 10") |  |  | April 24, 2009 | 2009 | Welch Allyn Protocol, Inc | Possible for users to misunderstand signals. The original product labeling told users that a flashing low battery indicator signals that the battery is low and requires changing soon. If a user experiences a flashing low battery status indicator and does not understand that the device can continue to be used, the user might choose not to continue operating the device, thereby causing therapy to | Labeling design | The firm has developed expanded labeling for users and will revise the user manuals for software versins 2.03, 2.04, 2.05, and 2.06. Customers receiving the updated manuals will be sent a cover letter directing them to discard the original user manual and replace it with the updated manual. The April 24, 2009, cover letter will highlight the clarification of directions for use regarding what to do when a low battery indicator is displayed. |  | 46740 | 46740 units | Worldwide distribution: USA, Guam, Puerto Rico, Canada, Mexico, Andorra, UAE, Afghanistan, Australia, Azerbaijan, Belgium, Belarus, Bermuda, Bahrain, Brazil, British West Indies, Switzerland, China, Cyprus, Czech Republic, Germany, Denmark, Algeria, Egypt, Spain, France, French Guiana, French Lesser Antilles, Greece, Haiti, Hong Kong, Hungary, Ireland, Israel, India, Iraq, Italy, Japan, Kenya, Kyrgyz, Lebanon, Lithuania, Luxembourg, Latvia, Monaco, Malaysia, Netherlands, New Zealand, Phillipines, Pakistan, Poland, Portugal, Reunion Island, Russia, Saudi Arabia, Singapore, South Korea, Taiwan, United Kingdom, Uraguay, Yemen, and South America. | Terminated | August 11, 2011 | 839 | 46740 | Not_Computer
Z-1811-2009 | 52601 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | BD 60ml Syringe Luer-Lok tip, Sterile, BD, Franklin Lakes, NJ 07417, Made in USA. |  |  | July 07, 2009 | 2009 | Becton Dickinson &amp; Company | Insufficient Seal: Becton Dickinson received complaints of insufficient/open seals on one lot of 60 ml Luer-Lok Syringes. | Packaging process control | Becton Dickinson issued an "Urgent: Product Recall" notice via UPS tracking dated July 7, 2009 to all distributors informing them of the affected product. The firm asked all distributors to identify end customers in an attempt to notify all customers. The firms asks that all affected product be returned.Further questions should be directed to BD Customer Service representatives at 1-888-237-2762. |  | 133600 | 133,600 units | Nationwide Distribution -- Including states of IL, CA, VA, PA, and FL. | Terminated | October 02, 2009 | 87 | 133600 | Not_Computer
Z-1540-2009 | 51211 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | i-STAT EG7+ Cartridge; pH, PCO, PO, Na , K, iCa, Hct. |  |  | February 23, 2009 | 2009 | Abbott Point of Care Inc. | Equipment problem: i-STAT cartridges were not properly sealed as the packaging system jammed, causing compromised seals. | Packaging process control | Urgent Recall Notice letters pertaining to the EG7+ Cartridges were sent on February 23, 2009 by first class mail to all customers. The firm found one other affected lot of G3+ cartridges and expanded the recall to include one other lot. Letters for the expansion are dated May 2009. Customers were asked to use the attached form to record all of the box numbers if they have any cartridges from the affected lot. Cartridges from the box numbers listed will be replaced an any remaining unused cartridges from those specific boxes are to be returned as instructed on the form. To receive replacement cartridges, customers are to follow the instructions on the enclosed customer reply form. If any cartridges from the affected lot were forwarded to another facility, a copy of the recall letter needs to be provided to them and the information at the bottom of the customer reply form should be completed. Questions should be directed to 1-800-366-8020, Option 1. |  | 26200 | 1048 boxes of 25 cartridges each. | Worldwide Distribution -- USA, Austria, China, India, Saudi Arabia, South Africa, and Taiwan. | Terminated | July 29, 2009 | 156 | 54600 | Not_Computer
Z-0002-2010 | 52834 | [6 empty] | Edwards Lifesciences Fogaty Occlusion Catheter Ref # 62080822F. |  |  | June 22, 2009 | 2009 | Edwards Lifesciences, Llc | Several complaints were the two parts of the tubular package became loose under the shrink wrap that holds the DFU on the packaging tube. | Packaging change control | Edwards initiated a worldwide recall via customer letter dated June 29, 2009, sent via Federal express on same day, requesting that customers cease using Fogarty Occlusion Catheter, Model No. 62080822F products from the indicated lots and that they return all unused products. Edwards requested that the customers provide written verification of stock and identify any unused products to be returned to Edwards Lifesciences LLC, Att: Field Action Coordinator, ODW Logistics, Inc., 13725 Pipeline Ave., Chino, CA 91710. Additionally provided Customer Service Tel # 800.424.3278. |  | 9548 | 9548 | AL AR AZ CA CO CT DC DE FL GA HI IA ID IL IN KS KY LA MA MD MI MN MO MS MO NC ND NE NH NJ NM NV NK OH OR PA RI SC SD TN TX UT VA VT WA WI WV WY, Puerto Rico Australia, Belgium, Brazil, Bulgaria, Canada, Canary Island, China, Czech Republic, Egypt, Finland, France, Germany, Great Britain, Greece, Hungary, Iceland, India, Ireland, Italy, Japan, Luxembourg, Netherland, New Zealand, Poland, Portugal, Rumania, Russia, Singapore, Slovak Republic, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, and Turkey. | Terminated | February 23, 2012 | 976 | 9548 | Not_Computer
Z-1901-2009 | 52717 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Boston Scientific (1) Flexima Regular Biliary Catheter System with RO Marker, 8 FR (2.7 mm) with Glidex hydrogel Coating Locking Pigtail and Radiopaque Marker, REF /Catalog No. 27-260 UPN; M001272600, Use before 2011-11, Sterilized with Ethylene Oxide Gas, Single Use Only, Made in USA 780 Brookside Drive, Spencer, IN 47460-1080. |  |  | July 01, 2009 | 2009 | Boston Scientific Corporation | Boston Scientific is initiating a recall of a single batch/lot of this product. We received a report that the labeling on the device does not match the labeling on the product packaging. Specifically, the packaging pouch for the product and outer carton labeling reflect that the device is a Flexima Biliary Catheter, which is correct, but the heat shrink inscription directly below the hub on the | Labeling mix-ups | Boston Scientific issued an "Urgent Medical Device Recall -Immediate Action Required" letter dated July 1, 2009. The letter was addressed to Risk Manger / Field Action Contact describing the problem and units affected and that further distribution or use of affected product should cease immediately. The Reply Verification Tracking Form enclosed with the letter was requested to be completed and returned.For further information, contact Boston Scientific at 1-508-683-6036. |  | 1 | 79 units (1 box = 1 unit) | Worldwide Distribution -- United States (AZ, AR, CO, DE, FL, GA, MI, MT, NE, NC, OH and WI), Spain, Finland and Netherlands. | Terminated | December 02, 2011 | 884 | 1 | Not_Computer
Z-0979-2009 | 50697 | No | KNS | Endoscopic electrosurgical unit and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Gold Probe 7FR Working Length: 350 cm, Probe Tip O.D. 2.3 mm (7Fr), Rx Only. Electrical Connector: Standard. Single Use Only. Legal Manufacturer, Boston Scientific Corporation, One Boston Scientific Place Natick, MA 01760-1537 USA. Made in USA: 8600 N.W. 41st Street Miami, FL 33166-6202 USA. |  |  | December 24, 2008 | 2008 | Boston Scientific Corporation | Packaging tray for the Gold Probe Single-Use BiPolar Electrohemostasis Catheters may have a crack. Cracks in the tray or damage to the tray of the Gold Probe device may lead to a compromised sterile barrier. | Packaging process control | A notification letter (dated December 24, 2008) was issued to customers. The letter identified the issue. Consignees are instructed to: 1) Discontinue use and segregate affected product, 2) Complete a Reply Verification Tracking Form and fax to the Boston Scientific Corporation (BSC) at 508-683-5578 and 3) Package and ship recalled product as instructed in the letter. If affected product is found, customers are instructed to call BSC Customer Service at 1-866-868-4004 between 8:30 a.m. through 6:00 p.m. (Eastern Standard Time).Please direct questions about the recall to your local Sales Representative. |  | 10841 | 10,841 | Worldwide Distribution. | Terminated | May 11, 2010 | 503 | 109610 | Not_Computer
Z-0535-2010 | 52837 | No | LRK | Intraoral devices for snoring and intraoral devices for snoring and obstructive sleep apnea. | Dental | Dental | 510(k) | Repose G2 Bone Screw System, Sterile, REF 76353200; Lot # 62450300. |  |  | July 27, 2009 | 2009 | Medtronic Xomed, Inc. | Medtronic Xomed, Inc. is recalling the Repose Bone Screw System product and Bone Screw product due to improper labeling. The labeling for the Repose System was inadvertently placed on the Repose Screw and vice versa. | Labeling mix-ups | Firm notified customers by phone on 7/27/2009. Firm asked customers to check inventories for the affected products listed on the enclosed Checklist.1. Fill-in the "quantity on-hand "column on the checklist.2. If you have any this product lot, contact Medtronic ENT Customer Service at 800-874-5797 (select Option 1 for Customer Service) and ask for the REPOSEField Action Contact Person for help to arrange for returns and replacements or credit.3. Fax your Checklist to Medtronic ENT at the number provided.4. Return products clearly marked. |  | 7 | 7 | 9 consignees in the continental United States. | Terminated | March 15, 2010 | 231 | 20 | Not_Computer
Z-1900-2009 | 52722 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ArthroWand Saber 30 with Integrated Cable wand, Catalog number AC4330-01, Manufactured by ArthroCare Corporation, Sunnyvale, CA |  |  | August 05, 2008 | 2008 | ArthroCare Corporation | Potential Sterility Loss-- Due to wearing or puncture in the Tyvek lid, the product may lose sterility. | Package design/selection | Firm initiated this recall on August 5, 2008. Consignees were notified via letter asking them to contact consignees and delivered via courier. Distribution discontinued on July 9, 2008. |  | 23443 | 23,443 units (includes distribution of previously returned product) | Worldwide Distribution including USA and countries of Netherlands, Slovak Republick, Ireland, Spain, Switzerland, Singapore, Russian Federatin, Greece, Japan, Saudi Arabia, India, Slovenia, Israel, South Africa, South Korea, Poland, Thailand, Jordan, Belgium, Finland, Portugal, Cyprus, Italy, UAE, Kuwait, China, Bahrain, Denmark, Norway, Sweden, Australia, Germany, Austria, UK, New Zealand, Mexico and Canada. | Terminated | June 14, 2011 | 1043 | 23443 | Not_Computer
Z-0962-2009 | 50699 | No | NRA |  |  | Orthopedic | PMA | Biomet Oxford Unicompartmental Knee, Phase 3, Right Medical Tibial Tray, Standard Size B, CoCr, sterile, Biomet UK Ltd, Bridgend, UK; . |  |  | December 19, 2008 | 2008 | Biomet, Inc. | Packages labeled as "right" actually contain a "left" component. | Packaging process control | Biomet salesmen were notified by phone and instructed to retrieve the products. One customer was notified by letter dated 12/19/08 that use of the product should cease and that a salesman would pick up the product. A second customer had implanted the product and was hand delivered a letter dated 12/19/08 instructing that the patient should be monitored and symptoms treated as necessary. Contact Biomet at 574-362-3983 for assistance. |  | 2 | 2 | Massachusetts and West Virginia. | Terminated | July 07, 2009 | 200 | 2 | Not_Computer
Z-1897-2009 | 52640 | Yes | FIR | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Owner Manual Modifications for PlasmaPro Plasma Pump Model PP-05. Apheresis Technologies, Inc. Palm Harbor, FL. - No labeling or Packaging. |  |  | April 01, 2009 | 2009 | Apheresis Technologies, Inc. | Apheresis Technologies, Inc. is taking the following Corrective Action: Modification of the Owners Manual Only for the PlasmaPro Plasma Pump Model PP-05. Operator's Manual will modify tubeset loading instructions so as to reduce possible loading error which might result in the tube being pulled into the pump head. | Labeling design | PlasmaPro Plasma Pump Operators Manual was modified and distributed to firms consignees. Firms received written notification of re-issuance of Operator's Manual. |  | 7 | 7 Owner's Manuals | Nationwide Distribution -- Florida, Colorado and Utah. | Terminated | July 06, 2010 | 461 | 7 | Not_Computer
Z-0976-2009 | 50548 | [6 empty] | Turbo Elite 0.9mm Turbo Elite (TE) RX, Model #410-154, tip diameter 0.9mm/.038 in, 4 Fr, sterile EO. Manufactured by Spectranetics, 96 Talamine Court, Colorado Springs, CO 80907. |  |  | July 04, 2008 | 2008 | Spectranetics Corporation | The catheter pouch label was incorrectly labeled with the wrong size. | Error in labeling | The mislabeled catheter was returned to Spectranetics. The second possible mislabeled catheter was visually inspected by a Spectranetics representative and confirmed the catheter was labeled correctly. |  | 2 | 2 units | Nationwide Distribution including states of TX and FL. | Terminated | February 13, 2009 | 224 | 2 | Not_Computer
Z-1367-2011 | 53012 | Yes | NIQ |  |  | Cardiovascular | PMA | Endeavor Resolute Zotarolimus-Eluting Coronary Stent System with Rapid Exchange Delivery System, UPN 00613994400505, model #IDERES25024X. Manufacturer: Medtronic, Inc. Minneapolis, MN. Manufactured in: Medtronic Ireland, Galway, Ireland. Product is packaged inside a sterile, inner pouch, labeled in part "Use immediately". Inner pouch is supposed to be within a foil pouch containing desiccant, and moisture absorber. Foil pouch label reads in part "RX Endeavor Resolute Zotarolimus-Eluting Coronary Stent System...2.5 mm X 24 mm...REF IDERES25024X...Contents: One (1) Endeavor Resolute zotarolimus-eluting stent mounted on a rapid exchange stent delivery system...Guide Catheter/Minimum Inner Diameter &gt;=5F/0.056" (1.42 mm)...Maximum Guidewire Diameter 0.014" (0.36 mm)...Maximum Stent I.D. 3.00 mm (0.12"...Nominal Pressure 9 atm (912 kPa)...Rated Burst Pressure 15 atm (1520 kPa)...MR Conditional...Sterilized using ethylene oxide Consult instructions for use...Manufacturer: Medtronic, Inc. 710 Medtronic Pkwy NE Minneapolis, MN 55432 USA...Manufactured In: Medtronic Ireland Parkmore Business Park West, Galway, Ireland Tel:+353-91-708000 Fax:+353-91-757524 Investigational Devices / Returned Goods: Tel: (800) 556-4247 / Fax: (877) 518-2469...BioLinx Medtronic Polymer System". Foil pouch contents and Instructions for Use are placed within an outer box, which is labeled identically to the foil pouch. |  |  | August 06, 2009 | 2009 | Medtronic Vascular Galway Limited | The product was removed from investigational sites due to a packaging issue, in that the device units were packed without the foil pouch, oxygen absorber and Instructions For Use. | Packaging | The firm, Medtronic, sent their initial "Urgent Medical Device Recall Notice" to customers with affected product via e-mail on August 7, 2009. An additional "Urgent Medical Device Recall Notice" was sent to all customers on August 12-13, 2009 via FedEX and UPS. The notice described the product, problem and action to be taken by the customers. The customers were instructed to remove the affected units from their inventory and isolate the units in a secure location to prevent additional use. A Medtronic CardioVascular Clinical Customer Service (CCS) representative will contact their site to arrange for return of the units. This will be followed by a shipment of replacement units to their sites. If there were any questions regarding the return process, customers were to contact CCS at 800-556-4247.If you have any further questions or concerns regarding this notice, please contact the Sr. Clinical Research Manager at 707-543-2284 or ann.depaoli@medtronic.com. |  | 49 | 49 units | Nationwide distribution: AL, AR, CA, FL, IN, MD, MI, NY, OH, PA, RI, TX, VA, and VT. | Terminated | February 17, 2011 | 560 | 49 | Not_Computer
Z-1721-2009 | 51317 | No | IXQ | Radiologic table. | Radiology | Radiology | 510(K) Exempt | Foot Lok Cusion, Part Number MTAFS01, reusable nonsterile (20.3 x 42 x 25.4cm) cushion. |  |  | March 06, 2009 | 2009 | Med Tec Inc | The silkscreened indexing scale was reversed. | Error in labeling | Med Tec inc. (CIVCO Medical Solutions, Radiation Oncology) issued a "Recall Notification" notice dated March 6, 2009 via certified mail explaining the reason for recall and providing the customer a choice of returning the product for replacement or not take any action. A response form was enclosed to be returned via FAX to indicate the number of units in their possession and whether or not they are being returned. Telephone calls were initiated April 20, 2009 to nonresponders or for any consignee whose letter was returned as undeliverable.Foreign customers were notified by e-mail or phone beginning March 16, 2009.For further information, contact CIVCO Medical Solutions, Radiation Oncology at 1-800-842-8688. |  | 231 | 231 units | Worldwide Distribution -- Including the United States, Czech Republic, Colombia, Thailand, Morocco, Brazil, the United Kingdom, India, Costa Rica, The Netherlands, Russia, Spain, Israel, Taiwan, Mexico, Turkey, Korea, Australia, Malaysia, Italy and Japan. | Terminated | November 05, 2009 | 244 | 231 | Not_Computer
Z-0509-2010 | 53495 | No | LIT | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | IMPACT Catheters, Balloon Dilatation and Esophageal Dilatation Catheter, Model Number 307, STERILE, Rx only |  |  | September 23, 2009 | 2009 | Numed Inc | The firm received a complaint from the sole direct account that while performing an inventory review it was observed that the packaging of the IMPACT catheter was cracked and deteriorating. The product was still in its packaging and never used. The customer referred to it as looking like "paint crackle." The risk is that for continued sterility to be assured, the packaging must remain intact. | Labeling design | An informal e-mail was sent to B. Braun to give them notice not to distribute any Impact devices until further notice. This was followed up by a formal URGENT MEDICAL DEVICE RECALL letter (dated 9/24/09) to the sole consignee. The products are to be returned to the manufacturer, including any products further distributed. Customer questions should be directed to the recall contact (replace " AT " with @) nlaflesh ATnumedusa.com |  | 7367 | 7,367 units domestically | One distributor in Pennsylvania | Terminated | May 23, 2012 | 973 | 7367 | Not_Computer
Z-0980-2009 | 50223 | No | FAD | Ureteral stent. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Cook Medical Universa Soft Ureteral Stent and Positioner with Hydrophilic Coating, 5.0Fr/26cm, sterile, Cook Urological Inc., Spencer, IN; REF USH-526-R, GPN REF G49966. The product was placed in a shipping box which is preprinted with the word "firm" instead of "soft", but a product specific label stating "soft", as above, was also placed on the box. |  |  | November 10, 2008 | 2008 | Vance Products Incorporated | "Soft" stents were placed in "firm" stent outer boxes. The soft stent is only to be left in the patient for 6 months, but the firm stent can be left in the patient for 12 months, so this error could result in the stent being left in the patient for longer than indicated for the device. | Labeling mix-ups | Consignees were notified via recall letter, "Urgent: Recall Notice" dated November 10, 2008, and by follow-up phone call. Customers were instructed to isolate the recalled product, complete the enclosed form and fax to Cook Urological at 812-829-2353. After receipt of the completed form, Cook Urological will issue a Return Goods Authorization number for each customer to return the affected product to Cook Urological via FedEx or UPS. |  | 23 | 23 | Nationwide Distribution including states of Georgia, Oregon and Texas. | Terminated | November 25, 2009 | 380 | 23 | Not_Computer
Z-0975-2009 | 50161 | Yes | DTZ | Cardiopulmonary bypass oxygenator. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Surgical Membrane Oxygenator. |  |  | October 17, 2008 | 2008 | Medtronic Inc | Medtronic has identified an issue with the labeling of the I-4500 Silicone Oxygenators (P/N 61399402655). Lot numbers of this oxygenator manufactured between August 2007 and September 2008 have an incorrect "Recommended maximum blood flow range" on the label. The label on the device states a range of 5.0 - 8.5 L/min but should state a range of 5.0 - 6.5 L/min. These oxygenators are acceptable t | Error in labeling | A consignee letter " Urgent Medical Deive Notice" dated October 17, 2008 was sent to the impacted customers on October 17, 2008. The letter described the issue and product and confirmed the correct maximum internal flow of 6.5 L/min for the I-4500 Silicone Oxygenators. The letter states the product will not be removed from the field. Customers were instructed to complete the" Medtronic Product Notification Certificate (FCA 0903)" acknowledging receipt of the information and fax the completed certificate to 763-391-9603. Please direct questions about the recall to your Medtronic Sales Representative or call 763-391-9124. |  | 408 | 408 | Worldwide Distribution USA including states of CA, CO, FL, IL, IN, KY, LA, MD, MI, MN, MO, NV, NJ, NY, OH, OK PA, TN, TX, UT, VA, and WI, and countries of LATVIA, ITALY, ARMENIA, RUSSIAN FEDERATION, POLAND, CHILE and CHINA. | Terminated | April 19, 2011 | 914 | 408 | Not_Computer
Z-0974-2009 | 50887 | No | LYC | Bone grafting material. | Dental | Dental | 510(k) | Zimmer Dental Calcitite 4060-2 Dense Nonresorbable Hydroxylapatite Bone Grafting Material, 4 ampules, 0.5g each, 40-60 mesh size, sterile. Zimmer Dental, Carlsbad, CA; Item # 0003. |  |  | January 22, 2009 | 2009 | Zimmer Inc. | Fragments of the glass vial may be present in some of the packages from these lots. | Packaging process control | The consignee in Japan was notified by e-mail and "Urgent: Device Recall" letter dated January 22, 2009 to stop using the product and to return it to the firm. |  | 225 | 225 kits | International Distribution including the country of Japan. | Terminated | October 23, 2009 | 274 | 225 | Not_Computer
Z-0529-2010 | 53858 | No | HTS | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker Orthopaedics Sterile Knee Instrument, Patella Cutter; |  |  | November 11, 2009 | 2009 | Stryker Howmedica Osteonics Corp. | Disposable Patella Cutters - 30 mm (catalog number 8000-0030, lot V31257001) were packaged and labeled as 33 mm Disposable Patella Cutters (catalog number 8000-0033, lot V31257001 and vice versa. | Labeling mix-ups | An "Urgent Product Recall" letter dated November 11, 2009 was sent to all affected Stryker branches, hospital risk management, hospital Chief or Orthopaedics and surgeons by Federal Express. The letter described the issue, potential hazards and risk mitigation factors and requested that the Recall Acknowledgement form be completed and returned. The Recall Acknowledgement form has information regarding the return of affected product to the firm. Direct questions to Stryker Orthopaedics by calling 1-201-831-5832. |  | 12 | 12 | Nationwide Distribution (MA, NJ, FL, TN, OR and NY). | Terminated | August 30, 2010 | 292 | 12 | Not_Computer
Z-0973-2009 | 50848 | No | NGT | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Sterile Field Saline Flush; |  |  | January 13, 2009 | 2009 | Excelsior Medical Corp | Sterility Issue: There is the possibility that the affected lots have pouches that were not sealed properly and could potentially compromise the external sterility of the syringes. | Packaging process control | An Urgent Medical Device Recall letter (dated January 13, 2009) was sent to consignees via certified mail, with return receipt requested. The letter provided the reason for the recall and instructions to discontinue use and return (or destroy) the recalled product. The letter also instructed customers to complete the enclosed Customer Reply Form and fax (866-360-6211) to Excelsior Medical Corporation. If the affected product is destroyed on site, fax a copy of the certificate of destruction with the completed Customer Reply Form. Arrangements will be made by Excelsior Medical Corporation for the return of recalled product. Please direct questions to Excelsior Medical Customer Service at 800-487-4276. |  | 341580 | 341,580 units | Nationwide Distribution. | Terminated | August 06, 2009 | 205 | 341580 | Not_Computer
Z-1344-2009 | 51334 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Boston Scientific, FleximaÂż and PercuflexÂż Drainage Catheter Family (Pigtail Style) |  |  | March 18, 2009 | 2009 | Boston Scientific Corporation | Boston Scientific Corporation is initiating an urgent Field Correction related to its FleximaÂż and PercuflexÂż Drainage Catheter Family (Pigtail Style). During the attempted removal of the device, sutures were separated from the catheter and remained in the patient. | Labeling False and Misleading | Consignees were sent a Boston Scientific " Urgent Medical Device Field Correction Notification" letter dated March 18, 2009. The letter was addressed to "The Hospital Field Action Contact". The letter described the problem, product affected and the revised catheter removal instructions. They also requested the return of the Acknowledgement Form. |  | 586276 | 586,276 US; 276,013 OUS total for all products | Worldwide Distribution -- USA including DC, Puerto Rico and Guam and countries of United Arab Emirates Austria, Bosnia and Herzegovina, Bahrain, Switzerland, Germany, Egypt, Spain, Finland, France, Great Britain, Ireland, Italy, Kuwait, Netherland, Pakistan, Portugal, Saudi Arabia, Sweden, Slovak Republic, Tunisia, Turkey, South Africa, Netherlands, Costa Rica, Japan, Argentina, Australia, Brazil, Canada, Chile, China, Columbia, Ecuador, El Salvador, Guatemala, Honduras, Hong Kong, Macau, Malaysia, Mexico New Zealand, Panama, Peru, Philippines, Puerto Rico, Singapore, South Korea, Sri Lanka, Taiwan, Thailand, Uruguay, and Venezuela. | Terminated | February 23, 2012 | 1072 | 586276 | Not_Computer
Z-0950-2009 | 48172 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Architect / Aeroset Clinical Chemistry Iron/Magnesium Calibrator, List Number: 1E69-03, Lot Numbers: 54187M200 |  |  | April 24, 2008 | 2008 | Abbott Laboratories Inc. | An incorrect Magnesium calibrator 2 value (CAL 2) was provided for Clinical Chemistry Iron/Magnesium Calibrator lot number 54187M200. Use of the incorrect Magnesium CAL 2 value causes quality control and patient magnesium results to be reduced by as much as 17.24% above 0.8mEq/L (1.0 mg/dL or 0.4 mmol/L). | Error in labeling | On April 24, 2008, a Product Correction letter with attached Customer reply form was mailed via US Postal Service to all affected customers that have received the affected lots. Customer reply form were included in all US customers letter. The letter instructed customers to:1. Identify the lot number of Clinical Chemistry Iron/Magnesium Calibrator currently used in your laboratory2. If you have lot number 54187M200, cross out the incorrect Magnesium CAL 2 value in the value sheet, commodity 30-3535/R1, June 2005. Place a copy of this letter in all calibrator kits with this lot number.3. Enter the revised Magnesium CAL 2 value provided in this letter in your system. NOTE: Ensure that the value used corresponds to the units of concentration used in your laboratory. *ARCHITECT cSystems - Refer to Enter a calibrator concentration (cSystem) in Section 2, Installation procedures and special requirements of the ARCHITECT System Operations Manual. *AEROSET - Refer to Calibrator/Control Configuration in Section 2, Installation Procedures and Special Requirements of the AEROSET System Operations Manual.4. Calibrate the Magnesium assay. Verify the calibration with at least two levels of controls according to the established quality control requirement for your laboratory.Customers were also instructed to please retain this communication for future reference and if they have forwarded any Clinical Chemistry Iron/Magnesium Calibrator to another laboratory, please provide a copy of this letter to them. An 800 number was provided in the recall letter for any questions the affected customers or any of the health care providers they serve may have. |  | 1287 | 1287 | Worldwide Distribution --- USA including states of AL, AZ, CA, CT, FL, GA, IA, IL, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NH, NJ, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, and WI, and countries of Mexico, Chile, Puerto Rico, Canada, Thailand, Curacao, Guatemala, Argentina, Dominican Republic, Germany, New Zealand, Jamaica, Ecuador, Brazil, El Salvador, Hong Kong, and Bermuda. | Terminated | July 09, 2009 | 441 | 1287 | Not_Computer
Z-0990-2009 | 50552 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Zimmer Minimally Invasive Solutions Technology Offset Rasp Handle, Left, 30 degrees, Zimmer, Warsaw, IN; REF 00-7806-035-01. |  |  | December 19, 2008 | 2008 | Zimmer Inc. | The c-clip located near the connection end may fracture or fall off during use in surgical procedures. | Labeling design | U.S. consignees were notified by "Urgent: Device Recall" letter issued via e-mail on December 19, 2008. The letter described the problem and instructed customers to immediately quarantine and return the recalled product along with the completed Inventory Return Certification form to the Zimmer, Inc. Distribution Center. Consignees were also instructed to fax a copy of the completed Inventory Return Certification Form at 1-574-371-8603. International consignees were also notified, but Zimmer, Inc. will gather the instruments from those customers. Direct questions about the recall to Zimmer, Inc. at 1-800-613-6131. |  | 1461 | 1,461 of both products. | Worldwide Distribution --- including USA, Belgium, Czech Republic, Finland, Germany, Greece, Italy, Poland, Romania, Spain and United Kingdom. | Terminated | December 22, 2009 | 368 | 1461 | Not_Computer
Z-1818-2009 | 52379 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | T2 Ankle Arthrodesis Nail Surgical Protocol, Catalog Number LT2AA-OT, Stryker Orthopaedics |  |  | November 17, 2006 | 2006 | Stryker Howmedica Osteonics Corp. | Operative Surgical Technique was found to have a typographical error where hyphens were inadvertently omitted in incision lengths (i.e. 1520 cm should be 15-20 cm). | Labeling Change Control | Important Product Correction Letters, dated November 17, 2006, were sent to all Stryker Branches/agencies via Federal Express. The letter stated the issue and asked consignees to examine their inventory and contact their sales reps. It also told them to destroy the affected Operative Techniques by cutting them in half and discarding them, complete the attached Acknowledgment Form, and fax a signed copy of that form to 201-831-6069. The letter stated that Stryker Orthopaedics Customer Service reps would place consignees' re-orders for the product being destroyed. Also, the letter stated that in the interim a correct revision of the Operative Technique could be found online. Customers are to contact Dave O'Dell at 201-831-5277 or Rita Intorrella at 201-831-5825 with any questions. |  | 1270 | 1270 | Nationwide Distribution | Terminated | August 20, 2009 | 1007 | 1270 | Not_Computer
Z-0939-2009 | 50192 | No | LOX | Percutaneous Transluminal Coronary Angioplasty (PTCA) Catheter. | Cardiovascular | Cardiovascular | 510(k) | VOYAGER NC Coronary Dilatation Catheter, 4.5 X 12.0 mm. |  |  | November 05, 2008 | 2008 | Abbott Vascular-Cardiac Therapies dba Guidant Corporation | The recall was initiated after Abbott Vascular received reports that, in a small percentage of cases, the protective, yellow balloon sheath of the 4.5mm diameter VOYAGER NC Coronary Dilatation Catheter has been tight. If unnoticed by the user, there is a potential to cause damage to the balloon dilatation catheter during removal of the sheath. The balloon sheath covers the balloon, maintaining th | Package design/selection | An Urgent Device Recall letter (dated November 6, 2008) with an Efficacy Check Reconciliation form was issued and hand-carried to affected customers by their respective Sales Representative who assisted in locating and returning the recalled units. The letter described the problem and informed customers that Abbott Vascular will replace returned recalled product with like product or with VOYAGER NC sizes when they become available. Additionally, customers were instructed to contact their Abbott Vascular local account representative to review their inventory, complete and return the attached Efficacy Check Reconciliation Form and return affected products to Abbott Vascular. |  | 126 | 126 units | Worldwide Distribution. | Terminated | August 13, 2012 | 1377 | 224 | Not_Computer
Z-1714-2009 | 52457 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | SJM Peel Away Introducer, REF 405153, Lot 2694252, Length 14 cm, 7F (2.33 mm), Max Guidewire o.d .038" (.965 mm), contents 5, 2011-08, Manufacturer St Jude Medical 14901 DeVeau Place, Minnetonka, MN 55345-2126 USA, Sterile EO, Rx only. |  |  | June 24, 2009 | 2009 | St. Jude Medical | St. Jude Medical has initiated a voluntary recall of two lots of Peel Away Introducers . Isolated lots could contain some items that have been mis-labelled with the incorrect French size of the outer box label. Some 9F products have been labelled as 7F and vice versa. Only the outer box label is affected by this error, the labels on the individual product pouches correctly identify the product. | Labeling mix-ups | The Consignee was provided with a St. Jude Medical "Urgent Product Information" letter and a doctor Acknowledgement form, via e-mail on June 2009. The letter described the problem and the product involved and advised to isolate and discontinue use of the product involved.The Consignee will notify the end user and remove any remaining affected units and will notify St Jude Medical of the results. |  | 200 | 200 | OUS only: Belgium. | Terminated | March 29, 2012 | 1009 | 200 | Not_Computer
Z-1812-2009 | 52160 | No | LIO | Microbiological specimen collection and transport device. | Microbiology | Microbiology | 510(k) | Remel BactiSwab NPG Collection and Transport System, Sterile, 100 swab units/box, Catalog 12300. The firm name on the label is Remel, Lenexa, KS. |  |  | May 07, 2009 | 2009 | Remel, Inc | Incorrect expiration date on the product | Incorrect or no expiration date | The recalling firm issued Important Medical Device Product Recall Notice letters dated 5/7/09 via regular mail to their direct accounts explaining the reason for recall and requesting they discard all remaining units of the affected lot number in their inventory. A Product Inventory Checklist was enclosed to be returned in the enclosed self-addressed, stamped envelope to report the amount of product that was discarded. The Technical Services Department should be contacted at 800-447-3641 if there are any inquiries concerning the letter. |  | 103 | 103 packages | Nationwide Distribution | Terminated | January 15, 2010 | 253 | 103 | Not_Computer
Z-1733-2009 | 52111 | No | JSY | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | Remel Middlebrook 7H10 Agar, Ref R01600, packaged 10 units/package. The firm name on the label is Remel, Lenexa, KS. |  |  | May 06, 2009 | 2009 | Remel, Inc | The plate labeling is incorrect, but the outer box labeling is correct | Error in labeling | The recalling firm issued Important Medical Device Customer Recall Notice letters dated 5/6/08 via regular mail explaining the reason for recall and requesting all remaining inventory of the recalled lot be discarded. A Product Inventory Checklist was enclosed for completion to report the amount of inventory destroyed. A self-addressed postage-paid envelope was enclosed for return of the Product Inventory Checklist. Customers are to contact the firm's Technical Services Department at 800-447-3641 if they have inquiries concerning the letter. |  | 118 | 118 packages | Nationwide Distribution -- including states of FL, NJ, IN, FL, OH, KY, TX, and WI. | Terminated | October 15, 2009 | 162 | 118 | Not_Computer
Z-0166-2010 | 53353 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Koodinat M/MP Patient Table used with Angiostar Model 5294009. |  |  | July 01, 2009 | 2009 | Siemens Medical Solutions USA, Inc | The table swivel locking mechanism may not completely engage adn could result in uncontrolled patient movement during transfer resulting in a patient fall. | Labeling False and Misleading | The recalling firm issued a Customer Safety Advisory Notice dated July 1, 2009 to affected customers via Update Instructions AX025/09/S to inform customer of the potential issues and provide instructions to prevent occurrence. The firm also provided an addendum to User Manual. Direct questions about the recall to Siemens Medical Solutions USA, Inc. by calling 1-610-219-6300. |  | 107 | 107 units | Nationwide Distribution. | Terminated | January 12, 2010 | 195 | 469 | Not_Computer
Z-0627-2009 | 50139 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Excelsior SL-10 Microcatheter, placed in sterile pouch. Manufactured by Boston Scientific Corporation, Fremont, CA. |  |  | October 07, 2008 | 2008 | Boston Scientific Corporation | Packaging issue may result in failure to maintain sterile barrier. | Package design/selection | Consignees were notified by letter dated October 7, 2008. The letter identified the issue and affected product. Customers were instructed to: 1.) Discontinue use of and segregate recalled product, 2.) Complete and return the enclosed Reply Verification Tracking Form, and 3.) Package/ship the recalled product and send to the Boston Scientific Corporation (BSC).Return the completed Reply Verification Tracking Form via fax (800-836-4804) to the BSC Customer Service Call Center.A follow-up letter dated October 8, 2008 was also issued to customers. The letter notified customers of the correct date range of expiration for product within the scope of the recall which is September 2008 through October 2010. There were no other changes identified in the follow-up letter. |  | 21423 | 21,423 units | Worldwide Distribution. | Terminated | March 02, 2010 | 511 | 21423 | Not_Computer
Z-0501-2010 | 53860 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer Gender Solutions Nexgen Complete Knee Solution, Legacy Knee- Posterior Stabilized, Gender Solutions Female, Femoral Component, option, LPS-flex, size C, right, for cemented use only; REF 00-5764-013-52. |  |  | November 13, 2009 | 2009 | Zimmer Inc. | The inner package may interfere with the outer seal, resulting in lack of assurance of sterility. | Packaging process control | Zimmer sent a recall letter, dated 11/12/09, to each initial consignee on 11/13/09. Zimmer field staff is to remove the product from affected hospitals, and to provide them with a copy of the recall letter. |  | 22 | 22 | Nationwide, Korea and Switzerland. | Terminated | October 15, 2010 | 336 | 22 | Not_Computer
Z-0458-2010 | 53801 | No | MUD | Oximeter. | Cardiovascular | Cardiovascular | 510(k) | Nonin Model 7600B cerebral oximeter pod. Used with the Nonin Model 7600 Regional Oximeter with Equanox Technology and Bluetooth Wireless Technology. Nonin Medical, Inc., 13700 First Avenue North, Plymouth, Minnesota 55441-5443 USA. |  |  | October 23, 2009 | 2009 | Nonin Medical, Inc | A malfunction was detected with one of the components that, under rare circumstances, may result in the overheating of the POD. This overheating, while unlikely to occur, has the potential to cause a burn if the POD is in direct contact with the patient. | Device Design | An Urgent: Medical Device Recall letter was emailed to customers on 10/23/2009. The letter described the issue and hazard and asked that the use of the device be discontinued and to immediately contact Nonin for return and repair/replacement instructions. Questions or concerns should be directed to John Dalpee at 763-577-3166. |  | 62 | 62 | Worldwide Distribution -- USA, including states of LA, MN, NC, and NY and countries of Canada, Germany, Italy, Netherlands, and United Kingdom. | Terminated | March 29, 2012 | 888 | 62 | Computer
Z-0270-2010 | 53563 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo Imaging XS VA60 and VA70 in combination with JPEG Compression. |  |  | September 29, 2009 | 2009 | Siemens Medical Solutions USA, Inc | Distorted image. The "Distorted" image cannot be used for diagnostic findings, and, depending on the customer scenario, a repeated exam could be necessary if the image cannot be reported on any other workplace or recovered from another source. | Software design | The recalling firm issued a Customer Safety Advisory notice dated 9/14/09 to all affected customers via Update Instructions IM100/09/S. The notice informs the customer of the potential issue and provides instructions to avoid its occurrence. To avoid the potential risk of the issue, the local compression function can be disabled as described in the letter. Depending on the image data, the disabling of compression will result in increased storage usage.The issue will be resolved in the release of version VA70B.For all potentially affected customers with "JPEG Compression" license an update instruction will be released which defines the necessary corrective actions in case of "Distorted" images. The update instruction will be performed on request by the customer service group. Further questions regarding this issue are directed to the local SIEMENS Uptime Service Center or +49 (9131) 84-2056. |  | 28 | 28 units | The product was shipped to medical facilities in CA, FL, GA, IL, MD, MI, MO, NE, NC, NY, OH, PA, TN, TX, and WV. | Terminated | July 05, 2012 | 1010 | 58 | Computer
Z-0138-2010 | 52056 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Innova 2121-IQ/3131-IQ Biplane Cardiovascular Imaging System. |  |  | April 16, 2009 | 2009 | GE Medical Systems, LLC | Potential pump failure associated with the Coolix 4000 chiller of the Innova single plane and biplane systems that may stop the chiller from working. If a chiller failure occurs, the user will be informed with a message on the n-room monitor stating that three minutes are left before the X-ray is inhibited. This could lead to a delay in treatment and possible additional X-ray exposure and contra | Device Design | Consignees were sent on May 1, 2009 a GE Healthcare " Urgent Medical Device Correction" letter dated April 16, 2009. The letter was addressed to Hospital Administrator/ Risk Manger, Mangers of Radiology/ Cardiology, and Radiologists/ Cardiologists. The letter described the problem, Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information. |  | 61 | 15 AND 46 UNLOCATED | Worldwide distribution: USA, ARGENTINA, BELGIUM, BRAZIL, BULGARIA, CHINA, CROATIA, EGYPT, FRANCE, GERMANY, HONG KONG, HUNGARY, INDIA, ITALY, JAPAN, REPUBLIC OF KOREA, LATVIA, LIBYA, MEXICO, NORWAY, POLAND, RUSSIA, SINGAPORE, SPAIN, SWEDEN, TAIWAN, TURKEY, UNITED KINGDOM, ROMANIA, ISRAEL, UKRAINE, YEMEN, BELARUS, and AUSTRALIA. | Terminated | December 03, 2015 | 2422 | 234 | Computer
Z-0094-2010 | 53436 | No | GEH | Cryosurgical unit and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Wallach LL50x Cryosurgical Freezer CO2 Unpinned Yoke System |  |  | September 29, 2009 | 2009 | Cooper Surgical, Inc. | CO2 Cryosystem thermal load may exceed the exhaust gas phase, resulting in ice formation in the exhaust valve. Should this occur, the Cryo tip may cool at a less than optimal rate, resulting in ice formation in the exhaust valve. | Device Design | Cooper Surgical notified customers by letter titled "Notice of Recall - Urgent" dated September 29, 2009 advising them of the problem and requesting return of the device. International accounts were notified by letter Beginning October 1, 2009.For further information, contact Cooper Surgical by fax at 1-203-799-2002 or phone at 1-203-799-2000. |  |  |  | Worldwide Distribution -- United States, Canada, Australia, Belgium, England, France, Germany, Greece, Hungary, Iceland, Italy, Ireland, Latvia, Lithuania, Netherlands, Portugal, Spain, Solvenia, Sweden, Hong Kong, Republic of China, Taiwan, Angola, Cameroon, Congo, Croatia, Egypt, Ghana, Guam, Guineae Bissau, Indonesia, Kenya, Kuwait, Lebanon, Madagascar, Mali, Mauritania, Mozambique, New Zealand, Niger, Pakistan,Qatar, Russia, Sao Tome, Saudi Arabia, Senegal, Serbia, Singapore, South Africa, St Vincents &amp; Grenadines, Oman, Thailand, Turkey, UAE and Zimbabwe. | Terminated | April 17, 2012 | 931 |  | Computer
Z-0039-2010 | 53489 | Yes | BTM | Manual emergency ventilator. | Anesthesiology | Anesthesiology | 510(k) | Infant Manual Pulmonary Resuscitator with reservoir bag, face mask &amp; Gas sampling elbow. Product ID # 5801, 12-8713, 732-E, and 2206. |  |  | October 06, 2009 | 2009 | Unomedical | In some Manual Pulmonary Resuscitator units the duckbill valve and retention ring inside the patient valve housing may become dislodged and fall into the MPR bag/body. If this occurs, the MPR may lose its ability to generate the positive pressure necessary necessary to ventilate the patient. | Device Design | Unomedical issued Press on October 6, 2009. Letters were sent to customers on October 7, 2009 via UPS - return receipt. Letters explained the reason for recall and the potential hazard to patients. The letter contains a list of all affected lot numbers as well as a picture which shows that units with a clearly visible blue ring are NOT affected. Distributors are asked to examine inventory and quarantine all affected product and return the response form, even if no recalled product is in inventory. If the product was further distributed, the letter requests that a copy of the letter be sent to all customers and for returns, follow Distributor actions listed in the letter. Retailers and end users/Hospitals/EMT's are asked to perform a count of recalled product, segregate it and return the response form. All customers are asked to call 1-800-634-6003 where they will be assigned a RGA number and given instructions on how to return product.Contact information: 8:00AM - 5:30 PM Central Time Monday through Thursday and 8:00AM to 1:00 PM on Fridays, Central Time. |  | 200 | over 200 total lots | Worldwide distribution: USA, Argentina, Canada, Denmark, France, Greece, India, Italy, Japan, Mexico, Paraguay, Peru, Poland, Russia and Vietnam. | Terminated | August 10, 2010 | 308 | 200 | Computer
Z-0421-2010 | 52445 | No | IOR | Mechanical wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Kuschall K Junior Manual wheelchair |  |  | June 15, 2009 | 2009 | Invacare Corporation | The handle has been breaking at a calculated weight of 65 kg which may have serious consequences (broken bones) depending on the circumstances of the break (e.g., user on ramps, hills, etc.). | Device Design | The firm contacted their customers by telephone and then sent a recall notification letter to each customer on 6/15/2009. Questions are directed to email: Recalls@Invacare.com or phone: 800-333-6900. |  | 20 | 20 Wheelchairs | Product was distributed to the following states: CA, CO, MN, NJ, OR, PA, TX &amp; VA and one to Canada. | Terminated | November 21, 2011 | 889 | 20 | Computer
Z-0443-2010 | 53606 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Stryker Thor Anterior Plating System; |  |  | October 02, 2009 | 2009 | Stryker Spine | Surgeons may apply an accidental axial torque to and/or over-angulate the All-In One Guide while implanting the screw. This may cause the screw not to seat properly in the locking ring and may cause the Thor screwdrivers to twist or break at the tip. | Device Design | Recall notification letters were sent to all Stryker branches, Hospital Risk Management and Surgeons on October 19, 2009 by Federal Express. Questions are to be directed to Tiffani Rogers, Regulatory Compliance Manager at 201-760-8206. |  | 262 | 262 | Nationwide | Terminated | January 23, 2013 | 1209 | 9550 | Computer
Z-0562-2010 | 53820 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart Vehicle Wall Mount, Model M5528A |  |  | November 12, 2009 | 2009 | Philips Healthcare Inc. | HeartStart Vehicle Wall Mount (Model M5528A) with DC Power Module (Model M5529A)--Incorrect wiring during installation by user may result in a smoke or fire hazard. | Device Design | Philips Healthcare, Inc. notified customers by an "URGENT - Voluntary Medical Device Correction" letter, plus revised Instructions For Use (IFU) for the Vehicle Wall Mount accessory and the DC Power Module accessory of the HeartStart MRx which includes a revised description of correct installation and verification of performance. Customers were asked to follow the Actions in the customer/user section of the Urgent - Voluntary Medical Device Correction Notification. In this section, customers were given the instructions to correct their devices. Distribution of the Urgent - Voluntary Medical Device Correction Notification along with the IFU will be managed by Philips representatives in each affected geography. Direct questions or concerns about the recall to your local Philips representative. |  | 3265 | 3,265 units | Class II Recall -- Worldwide Distribution -- United States, Canada, Australia, Brazil, China, France, Germany, Hong Kong, Iceland, India, Indonesia, Italy, Japan, Jordan, Kenya, Libyan Arab Jamahiriya, Netherlands, New Zealand, Norway, Peru, Poland, Portugal, Romania, Saudi Arabia, Spain, Sri Lanka, Switzerland, Trinidad, Tobago, United Arab Emirates, and United Kingdom. | Terminated | January 25, 2012 | 804 | 3265 | Computer
Z-0418-2010 | 53593 | No | DXJ | Medical cathode-ray tube display. | Cardiovascular | Cardiovascular | 510(k) | IntelliVue Clinical Information Portfolio (ICIP) Critical Care Release D, Version D.02. |  |  | October 21, 2009 | 2009 | Philips Healthcare Inc. | A software defect exists in the ICIP Release 0.02 scheduling function for order management such that if a user edits either the Schedule Time or the Stop Time for an existing order, the corresponding intervention(s) for the order may not be created. | Software design | Philips Medical Systems notified accounts by letter on October 21, 2009, of the release ICIP version D.03, which contains a software correction . A Philips representative will schedule the implementation of this upgrade. Please contact the Philips Customer Care Center with questions or concerns about this correction at 1-800-722-9377. |  | 53 | 53 units | MD, RI, WV, and TXAustraliaBelgiumDenmarkFranceItalyMalaysiaSingaporeSpainSwitzerlandTaiwanUnited Arab EmiratesUnited KingdomandIreland | Completed |  |  | 53 | Computer
Z-0374-2010 | 53688 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | SureSigns Vital Signs Viewer (VSV):Spot Check Vital Signs Monitor with NBP (options to include Predictive Temperature only; SPo2 only; SPo2 and Recorder only; SPo2 and Predictive Temperature only; or SPo2, Predictive Temperature, and Recorder |  |  | October 30, 2009 | 2009 | Philips Healthcare Inc. | Use of these devices at their maximum volume setting may result in the premature failure of the internal speaker | Software design | Customer notifications dated October 2009, in the United States will be delivered via UPS. Customer notifications outside the US will be managed by Philips representatives in each affected geography. Customers will be instructed to lower the volume setting while using existing units during the first notification. As part of the second notification to the customers, the field action correction will be a customer installable software upgrade. Customers may contact Philips Response Center with questions or concerns at 1.800.722.9377.On December 14, the firm issued the follow-up notification letter implementing the action plan to provide the software upgrade kit for installation and instructions. |  | 221 | 221 units | Nationwide Foreign: CanadaAfghanistan, Albania, Algeria, Angola, Argentina, Austria, Australia, Bahrain, Bosnia &amp; Herzegowina, Bangladesh, Belgium, Bulgaria, Bahrain, Bolivia, Brazil, Switzerland, Chile, China, Colombia, Costa Rica, Croatia/Hrvatska, Cyprus, Czech Republic, Germany, Denmark, Dominican Republic, Algeria, Ecuador, Estonia, Egypt, Spain, Finland, France, Gabon, Ghana, Great Britain, French Guiana, Guadeloupe, Greece, Guatemala, Hong Kong, Hungary, Indonesia, Ireland, Israel, India, Iceland, Italy, Iraq, Jordan, Japan, Jamaica Republic of Korea, Kenya, Kosovo, Kuwait, Lebanon, Sri Lanka, Lithuania, Luxembourg, Latvia, Libyan Arab Jamahiriya, Morocco, Norway, New Zealand, Republic of Moldova, Macau, Malawi, Mali, Malta, Martinique, Mauritius, Mexico, Malaysia, Namibia, Nepal, Netherlands,Nicaragua, Nigeria, Norway, Philippines, Pakistan, Oman, Panama, Peru, Poland, Puerto Rico, Palestinian Territories, Portugal, Qatar, Romania, Republic of Serbia, Russian Federation, Saudi Arabia, Sweden, Singapore, Slovenia, Slovakia, Spain, Sri Lanka, Switzerland, El Salvador, Tanzania, Thailand, Tunisia, Turkey, Taiwan, Uganda, Venezuela, Vietnam, South Africa, and Zimbabwe United Arab Emirates | Terminated | February 20, 2013 | 1209 | 49648 | Computer
Z-0122-2010 | 53614 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Stryker Instruments Navigation System II- CART, Stryker Navigation, Kalamazoo, MI; Stryker REF 7700-100-000. |  |  | October 26, 2009 | 2009 | Stryker Instruments Div. of Stryker Corporation | Multiple software-related malfunctions may occur. The device may suddenly stop working, the screen may freeze, the screen may only be updated very slowly, the unit may not respond to user input, or the unit may only respond to user very slowly. | Software design | The firm issued a press release on 11/4/09. Consignees were notified by an Urgent: Medical Device Recall Notification letter dated 10/23/09, which advised of the affected products, the reason for the recall, the recall plan, and the actions to be taken. Customers were to immediately review this notification and distribute to all departments affected. Also, customers were to immediately quarantine any affected product and label as recalled product; complete and sign the enclosed Business Reply Form and fax a copy to 866-521-2762. Upon receipt of the form, a Stryker Navigation Field Service Rep would contact them to schedule the repair of their equipment. All U.S. units were corrected by 10/30/09. |  | 21 | 21 | Worldwide Distribution -- USA, including states of Florida, Kansas, Kentucky, Massachusetts, Michigan, New Jersey, Pennsylvania, South Carolina, Texas, and Washington and countries of Australia, China, Germany, Japan, Korea, Switzerland, and the United Kingdom. | Terminated | December 31, 2009 | 66 | 23 | Computer
Z-0543-2010 | 53771 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | i-Cat Classic, 3D Dental Imaging System manufactured by Imaging Sciences International, Hatfield, PA. |  |  | October 30, 2009 | 2009 | Imaging Sciences International, LLC | overhead carriage drops quickly | Device Design | Imaging Sciences International issued an "Urgent Medical Device Recall" notification dated October 30, 2009 to all customers informing them of the problem and that a field service representative would schedule a visit to install an upgrade kit.For further information, contact Imaging Sciences International at 1-800-205-3570, option 5. |  | 98 | 98 units | Worldwide Distribution -- Including the United States, Austria, Australia, Belgium, Brazil, Canada, France, Germany, India, Israel, Italy, Jamaica, Spain, South Korea and Taiwan. | Terminated | September 30, 2010 | 335 | 98 | Computer
Z-0544-2010 | 53577 | No | KFM | Nonroller-type blood pump. | Cardiovascular | Cardiovascular | 510(k) | Sarns Centrifugal Pump (with X-coating), catalog 164275X, packaged in 8 piece shelf boxes, and convenience kits with single pump units. |  |  | July 06, 2009 | 2009 | Terumo Cardiovascular Systems Corporation | The tubing- to- pump connection in a medical device used in surgical procedures may disengage resulting in blood loss. | Device Design | Terumo Cardiovascular Systems Corporation notified consignees of an "Urgent Medical Device Correction" by letter on July 06, 2009. The letter was addressed to hospital Perfusion, Cardiovascular Surgery, and Operating Room departments and contained specific measures to be taken to ensure proper use of the affected device. The firm recommended that users continue to use the devices and follow the corrective measures enclosed with the notification. Consignees were requested to route the notification to all users and confirm receipt of the notification by returning a customer response form.For further information, contact Terumo CVS Customer Service at 1-800-521-2818. |  | 20702 | 20,702 units | Nationwide Distribution | Terminated | November 09, 2011 | 856 | 34455 | Computer
Z-0420-2010 | 53681 | No | LBI | Neonatal phototherapy unit. | General Hospital | General Hospital | 510(k) | The Natus Blue Light Phototherapy is also called the "neoBLUE LED Phototherapy System". The "neoBLUE mini LED Phototherapy System" |  |  | October 01, 2009 | 2009 | Natus Medical Incorporated | The clamp connecting the neoBLUE mini phototherapy light to the accessory pole stand can sustain stress fractures or break when excessive force is used when tightening the clamp. No injuries have been reported. | Device Design | On 10/1/09 Natus notified domestic consignees by Urgent Product Advisory Notice letter that advised them of the issue, instructed them to inspect the clamps, and included pictures of the affected clamp. Consignees are instructed to return the recalled clamps for a new replacement to be shipped by Natus.International distribution partners were notified by letter on 10/1/2009 and instructed to identify and contact their consignees. Replacements will be sent to the distribution partner for their consignees. |  | 2127 | 2127 | Worldwide distribution: USA, Argentina, Australia, Austria, Belgium, Bermuda, Canada, Chile, Denmark, Finland, France, Germany, Greece, Hong Kong, Hungary, India, Indonesia, Iraq, Israel, Italy, Japan, Kenya, Kuwait, Lithuania, Malaysia, Mexico, Morocco, Netherlands, Norway, Philippines, Poland, Portugal, PR, Qatar, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, UAE, and UK | Terminated | September 22, 2011 | 721 | 2127 | Computer
Z-0541-2010 | 53519 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | GE Healthcare Voluson E6, GE Voluson E8, and GE Voluson E8 Expert Diagnostic Ultrasound Systems |  |  | October 22, 2009 | 2009 | GE Medical Systems, LLC | GE Healthcare has recently become aware of an incorrect operating procedure associated with software versions 9.0.0, 9.0.1, and 9.0.2 of the Voluson E6 or Voluson E8 Diagnostic Ultrasound Systems that may impact patient safety.When using the combination of 2D-CRI and PW with a steered Doppler angle, the displayed velocity scale is incorrect. The consequence can be an underestimation of flow v | Software design | Consignees were sent on 10/23/09 a GE healthcare "Urgent Medical Device Correction" letter dated 10/22/09. The letter was addressed to Hospital Administration/Risk manger, Primary Physician and Biomedical Dept. Manger. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information. Questions and concerns regarding this issue are directed to the local GE Healthcare service representative or to 1-800-437-1171.GE also plans to update all of the affected GE Voluson E6, GE Voluson E8, and GE Voluson E8 Expert Diagnostic Ultrasound Systems that were distributed. |  | 715 | 715 | AL, AR, CA, CT, FL, GA, IL, KS, KY, LA ME, MD, MA, MI, MN,MS, MO, NE NJ, NY, NC, OH, OK PA, SC, TN,TX, VA,WI.AUSTRALIA, AUSTRIA , BELGIUM, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, CYPRUS, CZECH REPUBLIC, DENMARK, ESTONIA, FINLAND, FRANCE, FRENCH POLYNESIA, GERMANY, GREECE, HUNGARY, INDIA, INDONESIA, IRAN, IRELAND, ITALY, JAMAICA, JAPAN, KOREA, KUWAIT, LITHUANIA, MALAYSIA. NETHERLANDS, NE ZEALAND, PARAGUAY, PERU, PHILIPPINES, POLAND, PORTUGAL, QATAR, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TUNISIA, UNITED ARAB, EMIRATES, UNITED KINGDOM, MEXICO, VENEZUELA. | Terminated | December 17, 2011 | 786 | 715 | Computer
Z-0149-2010 | 53251 | stretcher, hand-carried | FPP | Hand-carried stretcher. | General Hospital | General Hospital | 510(K) Exempt | Stryker Medical Stair-PRO Stair Chair with wheels and a track system, patient transport device, Model 6252. |  |  | September 02, 2009 | 2009 | Stryker Medical Div. of Stryker Corporation | The front of the handle may become detached from the chair, which could result in a patient fall or an operator injury. | Device Design | Consignees were sent a letter titled URGENT MEDICAL DEVICE CORRECTION dated 9/9/09, advising them of the problem and to remove the devices from service until they have been repaired. U.S. customers were notified that Stryker will complete the correction while international customers were sent a kit with instructions for completion of the correction. International consignees were told that Stryker will complete the correction if the consignee does not feel comfortable in doing so. Questions or comments regarding this issue should be addressed to Sarah Marie Puterbaugh of Stryker Medical at (800) 869-0770 ext. 6902. |  | 165 | 165 | Worldwide distribution - Colorado, Ohio, Wisconsin, and Germany, Italy, Netherlands, Poland and Switzerland. | Terminated | February 16, 2011 | 532 | 230 | Computer
Z-0414-2010 | 53475 | No | FTD | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | STERIS Harmony LL 500/700 Surgical Light, Vertical Spring-Loaded Arm. |  |  | September 11, 2009 | 2009 | Steris Corp | DrÂżger Medical AG &amp; Co., KGaA informed STERIS Corporation that they received (4) reported international cases where a welded seam at the joint of the spring-loaded arm of the DrÂżger SOLA 500/700 surgical light broke as a result of a crack that had slowly developed over time. The STERIS Harmony LL 500/700 surgical light is a Private Label version of the DrÂżger SOLA 500/700 surgical light. | Device Design | The firm sent Urgent Product Correction letters, dated Sept. 11, 2009, to their customers. The letter identified the problem and stated that Drager had supplied STERIS the replacement spring arms and that a STERIS Service Technician would perform the replacement. The letter also stated that the correction would be performed at no charge to the customers. According to the letter, customers would be contacted by their STERIS Service Representative to arrange a convenient time for replacement of potentially affected spring arms. In the meantime, the letter states that customers should use their surgical lights in accordance with the guidelines presented in the Operator's Manual. Questions or concerns regarding the letter should be directed to Barbara Rich at 440-392-7318 or your local STERIS account manager. |  | 388 | 388 units | Worldwide Distribution -- USA, including the states of AL, AR, CA, CO, CT, FL, GA, HI, IL, IN, KS, LA, MA, MN, MO, MS, NC, NH, NJ, NV, NY, OH, PA, RI, SC, TN, TX, UT, VA, WA, WI &amp; WY and the country of Canada. | Terminated | March 01, 2012 | 902 | 388 | Computer
Z-0579-2010 | 53871 | No | ITI | Powered wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Invacare Storm TDX SR Power Wheelchair with Stability Lock Model numbers: TDXSR, TDXSR-CG, TDXSR-CG-HD, TDXSR-HD, TDXSR-MCG, TDXSR-MCG-HD, TDXSRV-HD |  |  | September 15, 2009 | 2009 | Invacare Corporation | For some Invacare chairs, the stability lock feature may not be engaging properly or consistently. If the stability lock feature does not engage properly, the wheelchair may do one or more of the following: (1) veer to one side; (2) rock forward onto its front riggings; (3) drive in an unintended circular motion; and/or (4) fall forward or to one side. In these cases there is a risk to injury for | Device Design | On 9/15/2009 the firm sent recall notification letter to their customers instructing them to immediately contact the dealer to have the stability lock inspected and/or adjusted. If the dealer is unavailable or the repair will take longer than 60 days, they are to contact the firm and arrange for an inspection and/or adjustment. Invacare requested contact names to whom it can send additional information. Invacare plans to mail, via FEDEX, information regarding the recall to each consignee. Initial consignees received a response card and instructions to forward cards to any customer to whom they have distributed the product. The products are to be inspected at the cylinder label and if the label does not have a long alpha-numeric string demonstrated in the label, the cylinders on the chair have to be replaced. Questions are directed to Recalls@Invacare.com or to 800-333-6900. |  | 36 | 36 | Product was distributed to the following states: AZ, CA, CO, CT, DC, FL, GA, IA, IL, IN, KS, MA, ME, MI, MN, MO, MS, NC, ND, NM, NY, OH, SC &amp; TX. Product was also distributed in Canada. | Terminated | August 25, 2014 | 1805 | 36 | Computer
Z-0088-2010 | 53086 | No | LNQ |  |  | Gastroenterology/Urology | HDE - Humanitarian Device Exemption | Medtronic Enterra Therapy, Gastric Electrical Stimulation System, consisting of model 3116 Enterra Therapy Neurostimulator and model 4351 Unipolar Intramuscular Lead. Medtronic, Inc., 710 Medtronic Parkway, Minneapolis, MN 55432-5604 USA. |  |  | August 25, 2009 | 2009 | Medtronic Neuromodulation | Medtronic is in the process of updating the Enterra Therapy labeling to include information regarding the risks of bowel obstruction and/or perforation. Medtronic has received 15 reports since November 2002 of bowel obstruction and/or perforation of the bowel associated with Enterra Therapy System leads. Of the 15 reported events, 11 required surgical intervention (all patients recovered withou | Device Design | Medtronic Neuromodulation issued a "Medical Device Correction" letter beginning August 25, 2009 addressed to Healthcare Professionals. The letter described the Nature of the Issue, Scope, Potential Severity of the Issue, Enterra Therapy System labeling updating information warnings, Recommendations and other Additional Information.For further information, contact Medronic Neuromodulation Technical Services at 1-800-707-0933. |  | 1909 | 1909 (1095 US, 814 OUS) | Worldwide Distribution. | Terminated | May 22, 2012 | 1001 | 1909 | Computer
Z-0582-2010 | 53875 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Delta/Delta XL Monitors with Omega S Package |  |  | November 09, 2009 | 2009 | Draeger Medical, Inc. | monitor is rebooting/resetting | Software design | Draeger Medical Systems, Inc. issued and Urgent Medical Device Recall letter dated November 2009 to their affected customers informing them of the problem and a work around until software can be validated. A Technical Service Bulletin has been generated and distributed globally that instructs the affected International Drager subsidiaries to inform customers in their countries.For further information, contact Draeger Medical Systems, Inc. at 1-800-543-5047, extension 1-2349. |  | 4188 | 4,188 units | The product was shipped to medical facilities in CA, CO, CT, ID, IN, KY, MS, NY, OK, and WI. The product was also shipped to United Arab Emirates, Argentina, Austria, Australia, Bosnia Herz, Belgium, Benin, Brunei Darussalam, Bolivia, Brazil, Belarus, Canada, Switzerland, Chile, China, Columbia, Costa Rica, Cyprus, Czech Republic, Germany, Denmark, Egypt, Spain, Finland, France, Great Britain, Greece, Hong Kong, Croatia, Hungary, Ireland, India, Iran, Italy, Japan, Kenya, Korea, Kuwait, Kazakhstan, Lebanon, Lithuania, Latvia, Libya, Morocco, Mexico, Malaysia, Netherlands, Norway, Nepal, Oman, Poland, Portugal, Qatar, Reunion, Romania, Serbia, Russian Federation, Saudi Arabia, Sweden, Singapore, Slovak Republic, Thailand, Turkey, Taiwan, Ukraine, Venezuela, Vietnam, and South Africa. | Terminated | August 23, 2011 | 652 | 4188 | Computer
Z-0365-2010 | 51876 | No | NSX |  |  | General Hospital | Contact ODE | PICIS, CareSuite Family- OR Manager (4235), Preop Manager (4250), Anesthesia Manager (4220), PACU Manager (4245) &amp; Critical Care Manager (1110), patient information systems. |  |  | April 08, 2009 | 2009 | Picis Inc. | Demographic data, most notably allergy and precaution data, can be overwritten with incomplete data or blanks by the interface between non-Picis clinical systems (your "HIS" or "CIS") and certain Picis applications, if that interface is not configured properly. | Software design | On April 8th 2009, Picis initiated a customer notification letter to the installed base of potentially affected users, explaining the potential source of patient data conflict and to provide support in choosing the appropriate configuration option and implementation. The notification letters included with this report package contain additional details regarding configuration options. |  | 325 | 325 | Worldwide distribution: USA, Australia, Austria, Belgium, Denmark, Finland, France, Germany, Italy, Malaysia, Netherlands, Norway, Portugal, Singapore, Spain, Sweden, and United Kingdom | Terminated | December 22, 2014 | 2084 | 325 | Computer
Z-0456-2010 | 53462 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Centricity PACS (Picture Archiving and Communication System) RA1000 Workstation software; GE Healthcare Integrated IT Solutions, Barrington, IL 60010. |  |  | September 25, 2009 | 2009 | GE Healthcare Integrated IT Solutions | There are potential safety issues, including misdiagnosis or delay in treatment, associated with the use of GE Centricity PACS RA1000 Workstation software that may impact patient safety. | Software design | GE Healthcare Integrated IT Solutions sent Urgent Medical Device Correction letters dated September 23, 2009 to their Centricity PACS software versions 2.1.x and 3.0.x customers, to the attention of the Hospital Risk Manager, the Hospital Administrator, the Head of the Radiology Department and the PACS Administrator, on 9/25/09. The letters advised the users of the patient safety issues associated with the use of GE Centricity PACS related to unintentionally adding new images to an approved exam; deletion of images from PACS if Cancel Checked or Cancel Selected is chosen in the CA Tool - Archive Queue; inadvertently getting two or more exams out of sync with the corresponding reports resulting in the RIS-IC reports being associated with the wrong patients exam; and when rejecting a duplicate image(s) from an unspecified exam that is still in the Arrived or Verified status, the duplicate image(s) is also rejected from the Dictated, Transcribed, or Completed status exam. The letter provided the users with safety instructions to follow until the software is updated. Any questions were directed to the Remote Online Center, Field Engineer or the GE Customer Care Center at 800-437-1171. The accounts were requested to complete the enclosed customer reply form, supplying the workstation software or build number, and fax the completed form to 847-939-1479.Centricity PACS Software Versions 1.0.x and 2.0.x customers were sent separate letters explaining the issues and the safety instructions to follow, but there will be no software updates since those software versions are no longer being supported. Any questions should be directed to the Remote Online Center, Field Engineer or the GE Customer Care Center at 1-800-437-1171. |  | 1188 | 1,188 units | Worldwide Distribution -- United States including Puerto Rico and countries of Australia, Austria, Belgium, Brazil, Canada, Chile, China, Denmark, Ecuador, Egypt, France, Germany, Hong Kong, Ireland, Israel, Italy, Japan, Korea, Kuwait, Lebanon, Malaysia, Malta, Mexico, Netherlands, New Zealand, Portugal, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Taiwan, Turkey and United Kingdom. | Terminated | May 25, 2012 | 973 | 1188 | Computer
Z-0121-2010 | 53329 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Siemens ARTISTE syngo RT Therapist, Therapist Assist, Therapist Express Basic, and Therapist Express Assist Systems, manufactured by Siemens Medical Solutions, Concord, CA |  |  | September 10, 2009 | 2009 | Siemens Medical Solutions USA, Inc | A failed data transfer is possible that may result in data not being recorded in the interfaced system. Siemens Medical Solutions, Inc. USA became aware of this potential issue with its ARTISTE Linear Accelerator's Syngo RT Therapist v4.1 Workspace (containing the Siemens brand of Verify and Record application called PRIMEVIEW) and the LANTIS Oncology Information System product. | Software design | The firm issued a TH015/09/S Safety Modification syngo RT Therapist 4.1 Customer Advisory letter, dated September 10. 2009. the letter stated the reason for modification, the affected systems, working time, deadline, reimbursement, and reporting. The firm is preparing a software upgrade for the affected products to correct the problem. The letter was sent via certified mail. |  | 97 | 97 units | Worldwide Distribution -- USA, Spain, Germany, Poland, Australia, Malaysia, New Zealand, the UK, India, Ireland, Egypt, Italy, Belgium, Republic of Korea, and Canada. | Terminated | December 06, 2010 | 452 | 97 | Computer
Z-1991-2009 | 52861 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | EasyLink Informatics System included with the Dimension Vista System. |  |  | July 14, 2009 | 2009 | Siemens Healthcare Diagnostics, Inc. | incorrect results may be transmitted to LIS | Software design | Siemens issued an "Urgent Device Correction" letter dated July 2009. The letter informed the user of the affected product including instructions for avoiding the problem.For further information, contact Siemens Healthcare Diagnostics Technical Solution Center at 1-800-441-9250. |  | 359 | 359 units | Worldwide Distribution -- United States, Australia, Austria, Belgium, Canada, France, Germany, Italy, Japan, Malaysia, Netherlands, Portugal, Slovenia, Spain and Switzerland. | Terminated | January 05, 2010 | 175 | 359 | Computer
Z-0108-2010 | 53509 | No | MQP | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Synthes USA, TI Synex (TM) II Central Body 14mm - 19mm, Catalog number 04.808.001. |  |  | September 14, 2009 | 2009 | Synthes USA (HQ), Inc. | loss of device height | Device Design | Synthes USA issued an "Urgent: Medical Device Recall" letter to hospital accounts, Synthes personal and sales consultants dated September 23, 2009. Users were asked to take inventory, complete and return a Verification form and return all affected product to the firm.For further information regarding the recall, contact Synthes at 1-800-479-6329 or 1-800-620-7025 extension 5453. |  | 267 | 267 units | Worldwide Distribution -- Including the United States and Canada. | Terminated | July 09, 2015 | 2124 | 2275 | Computer
Z-2325-2009 | 51928 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ARCHITECT i System Assay CD ROM - US, list 6E58-22, Version 22.01, and ARCHITECT i System Assay CD-ROM - WW (excluding US), list 6E59-23, Version 23.0, containing the ARCHITECT LH assay file; Abbott Laboratories, Diagnostic Division, Abbott Park, IL 60064 USA; Product List Number: 6E58 (US) and 6E59 (WW) Versions 23 and lower. |  |  | April 24, 2009 | 2009 | Abbott Laboratories | There continues to be an elevated level of complaint activity for calibration error code 1227, when using ARCHITECT LH reagent lots (6C25-22 or 6C25-27) due to the use of assay disk versions that do not contain the updated ARCHITECT LH assay file contained on ARCHITECT i System Assay CD ROM versions 24 and higher (US LN 6E58, WW LN 6E59). | Software change control | Abbott issued a "Product Correction-Immediate Action Required" letter dated April 24, 2009. The letter provided users with the steps to take to determine what version of the ARCHITECT LH Assay is on their instrument and instructions for installing the assay file provided with the letter if needed. Users were requested to return a customer reply letter acknowledging receipt of the letter. For additional information, contact 1-877-4ABBOTT. |  | 5713 | 1,037CD-ROMs - US, 4,676 CD-ROMs - WW | Worldwide Distribution -- United States including Puerto Rico and Argentina, Australia, Brazil, Canada, Chile, Colombia, Costa Rica, Curacao, Germany, Guatemala, Hong Kong, Japan, Mexico, New Zealand, Peru, Singapore, South Korea, Taiwan, Thailand, Uruguay and Venezuela. | Terminated | November 30, 2009 | 220 | 5713 | Computer
Z-0064-2010 | 53033 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | 200996-Cable Assembly, Hybrid Extended Length, |  |  | July 17, 2009 | 2009 | Mako Surgical Corporation | Cables fail to provide adequate transmission. MAKO Surgical Corporation is recalling 200996 Cable Assembly, Hybrid Extended Length; 200997 Cable Assembly, Hybrid, Guidance Module; and 200998 Cable Assembly, Hybrid, UI Panel. These parts are sub components of the 20300 Rio Standard System. | Device Design | A Recall Notice, dated July 15, 2009, was sent to owners of the RIO system (hospital OR staff) with a signed acknowledgement returned to MAKO. Surgeons using the RIO system during the recall time period, regardless of whether their system has already been upgraded, are given the notification letter and asked to sign and return the acknowledgement memo. The letter stated that customers' RIO will be upgraded to replace the potentially faulty cable with a new cable. Their MAKOplasty Specialist (MPS) will notify them when their system was scheduled to have their fiber optic cable upgrade. The upgrade will be installed by a Field Service Engineer (FSE) free of charge and the RIO will be completely tested prior to being returned to service. Customers should direct questions or concerns to their MPS, FSE, or Jim Pomeroy, who can be reached at 954-927-2044. |  | 23 | 23 | Nationwide Distribution | Terminated | May 12, 2010 | 299 | 23 | Computer
Z-0164-2010 | 53453 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Siemens Biograph mCT-S 64, model number 10248672. |  |  | August 11, 2009 | 2009 | Siemens Medical Solutions USA Inc. | In the unlikely event that the device is modified to operate in the tilted position, the telescoping support mechanism will not engage properly. | Device Design | Siemens Medical Solutions USA, Inc. issued an "Urgent Field Correction - Recall" notice to consignees by letter dated September 3, 2009 informing them of the problem and that a Siemens service representative will be in contact shortly to service the affected product.For further information, contact Siemens at 1-800-888-7436. |  | 2 | 2 units | Worldwide Distribution -- United States (IL and OH), Germany, France and Hong Kong. | Terminated | December 22, 2009 | 133 | 3 | Computer
Z-0173-2010 | 53535 | No | HET | Gynecologic laparoscope and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Gynecare Morcellex Tissue Morcellator; |  |  | September 28, 2009 | 2009 | Ethicon, Inc. | There is the potential for the duckbill seal to become dislodged. | Device Design | Recall notification letters titled "URGENT VOLUNTARY PRODUCT RECALL" were sent on 10/8/2009 via UPS next day. The letters describe the lots within expiration that are recalled. The motor drive unit is not included in the recall. Letters have photos and illustrations of the product which clearly show the location of the product code and expiration date. Recalled product is to be returned to Stericyle using an enclosed pre-paid authorized shipping label. Ethicon will notify customers when the product is available.Questions are to be directed to Ethicon at 1-888-276-5901. |  | 60930 | 60,930 in US. | Worldwide distribution - US and United Arab Emirates, Argentina, Austria, Australia, Belgium, Brazil, Canada, Switzerland, China, Czech Republic, Germany, Denmark, Egypt, Spain, Finland, France, United Kingdom, Greece, Israel, Italy, Jordan, Republic of Korea, Mexico, Malaysia, Netherlands, Norway, New Zealand, Puerto Rico, Reunion, Russian Federation, Sweden, Slovina, Turkey, Saudi Arabia, Venezuela and S. Africa. | Terminated | August 02, 2010 | 308 | 60930 | Computer
Z-0174-2010 | 53564 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo Imaging XS VA60 and VA70 in combination with a NAS or HSM license |  |  | September 30, 2009 | 2009 | Siemens Medical Solutions USA, Inc | Incomplete amount of image data; specifically, The image count in the series table of the online and archive database will display the full amount of images received at the Syngo Imaging XS as archived. However, the zip file on the NAS or HSM will contain fewer images-those which have been archived until the time of the error/interruption. | Software design | The recalling firm issued a "Customer Safety Advisory Notice" dated 9/14/09 to all affected customers via Update Instructions IM101/09/S. The notice informs the customer of the potential issue and provides instructions to avoid its occurrence. Questions regarding this issue are directed to the local SIEMENS Uptime Service Center. |  | 2 | 2 units | The product was shipped to medical facilities in NY and OH | Terminated | July 05, 2012 | 1009 | 2 | Computer
Z-0176-2010 | 53676 | Yes | MNT | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | RESPIRONICS V60 VENTILATOR |  |  | October 26, 2009 | 2009 | Respironics California Inc | The recall was initiated to correct an issue with a component on the V60 Ventilator Power Management (PM) Printed Circuit Board (PCB). A component on the PCB has the potential to break off or become damaged if the PCB is removed from the ventilator. If this occurs, the ventilator may not properly switch between back up battery power and AC power. This can result in power loss during ventilation fo | Device Design | The recall notification was initiated October 26, 2009 with the firm mailing a Urgent Medical Device Field Correction letter to the affected customers. The letter informed the customers of the reason for the notice and how Respironics Field Service Technicians were going to correct this issue. The letter also gave the customers instructions to follow pending the completion of the field inspection/replacement.For questions or additional information contact Respironics US Customer Service at 1.800.345.6443, Option #5 then Option #1. |  | 524 | 524 | Nationwide and Internationally to :Libya, Hong Kong, France, Italy, India, Spain, United Kingdom,Germany, Netherlands and Saudi Arabia | Terminated | October 13, 2011 | 717 | 524 | Computer
Z-0557-2010 | 51180 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker Howmedica Osteonics, Primary Super SecurFit Stems, hip prosthesis, Sterile; |  |  | January 14, 2009 | 2009 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics became aware that there may be difficulty in mating stem to implantation/extraction instrument. | Device Design | Stryker sent Urgent product Correction Letters on January 14, 2009 by Federal Express to Stryker branches/agencies, hospital risk management, chief of orthopaedics and surgeon. |  | 786 | 786 total stems | Nationwide. Product was distributed to 3 Stryker branches/agencies and 2 hospitals in PA and CA. | Terminated | July 18, 2012 | 1281 | 786 | Computer
Z-0170-2010 | 53128 | No | LGW |  |  | Neurology | PMA | Eon Rechargeable Implantable Pulse Generator (IPG) System |  |  | December 22, 2006 | 2006 | Advanced Neuromodulation Systems, Inc | Product's external charging system on occasion was failing to communicate with the Implantable Pulse Generator preventing re-charging of the pain relief system. | Device Design | 12/22/2006, Marketing Update 5015 was issued to field sales representatives providing information on improving communication between the Internal Pulse Generator and its charger. May 15, 2007, Marketing Update 5022 was issued to field sales representatives to reiterate avoiding contact between the charger antenna and conductive materials and place the preferred side of the charging antenna face-down over the implant. Troubleshooting Guide was created in January 2008 for field sales representatives and technical services departments regarding isolation of the communication issue to the external devices instead of the implanted component. March 25, 2008, the Clinician's Manual was updated adding a recommendation for a minimum depth of pocket for device implantation. |  | 24750 | 24,750 units total | Worldwide | Terminated | July 23, 2010 | 1309 | 24750 | Computer
Z-0066-2010 | 52632 | No | FRC | Sterilization process indicator. | General Hospital | General Hospital | 510(k) | Sterrad CycleSure Biological Indicator, Product Code 14324. |  |  | June 23, 2009 | 2009 | Advanced Sterilization Products | Biological Indicator indicating a positive result does not maintain yellow color for the entire incubation period and some indicators experience media evaporation prior to the end of the labeled maximum incubation time. | Device Design | An Important Product Labeling Correction - New Instructions for Use dated, June 26, 2009, has been proposed to be sent by the firm for Cyclosure Biological Indicators to consignees. The revised instructions reduce the maximum read time for the indicators to 3 days (72 hours). The letter also states why ASP changed the Instructions. Questions are to be directed to ASP Customer Care Center. |  |  |  | Worldwide Distribution. | Terminated | November 29, 2011 | 889 |  | Computer
Z-0086-2010 | 52048 | No | KWY | Hip joint femoral (hemi-hip) metal/polymer cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Ortho Development Corp., Pivot BiPolar Head, 28mm, Hip Prosthesis, Femoral Head. |  |  | April 29, 2009 | 2009 | Ortho Development Corporation | A Pivot BiPolar Head, the femoral head in a hip prosthesis, might not move freely after assembly due a stacking up of manufacturing tolerances. | Device Design | Ortho sent a Field notification Memo dated 4/2/2009 to agents and distributors, giving them pre-implant examination guidelines. Decision to recall non-implanted units was made on 4/29/2009 and this was communicated by telephone to distributors. |  | 23 | 23 | AZ, CA, FL, MA, MI. No military, government or foreign distribution. | Terminated | May 25, 2010 | 391 | 23 | Computer
Z-0481-2010 | 53748 | No | MRZ | Infusion pump. | General Hospital | General Hospital | 510(k) | Animas ezManager Max Diabetes Management System version 2.0.9 |  |  | November 02, 2009 | 2009 | Animas Corporation | Printed hard copy of results is not the same as the system screens- carbohydrate to insulin ratio incorrect. | Software design | Animas issued an Urgent Medical Device Correction letter dated 11/2/09 to the pump user and another letter to the health care professional that prescribed the pump. The letters explain the issue. The patient letter is accompanied with a CD that contains the upgraded software and instructions to follow. |  | 20648 | 20,648 units | Nationwide distribution, and Canada, Cayman Islands, and Denmark. | Terminated | October 27, 2010 | 359 | 20648 | Computer
Z-0179-2010 | 53139 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | System 1000 family of Hemodialysis Instruments, including the System 1000, TINA, AURORA and AltraTouch; Baxter Healthcare Corporation, Deerfield, IL 60015 U.S.A. |  |  | August 31, 2009 | 2009 | Baxter Healthcare Renal Div | When the main power is lost temporarily, a software anomaly prevents the instrument from automatically changing the ultrafiltration rate (UFR) to the minimum value as described in the Operator's manual. | Software design | Baxter sent an Urgent Device Correction letter dated August 31, 2009 to all System 1000 Series of Hemodialysis Instruments customers via first class mail on the same date, to the attention of the Hemodialysis Administrator. The letters informed the accounts that a software anomaly in software versions 4.B.0.9, 4.B.1.1 and 4.B.1.3 was discovered when the main power was lost temporarily during a simulated treatment. The anomaly prevents the instrument from automatically changing the ultrafiltration rate (UFR) to the minimum value as described in the Operator's Manual. The accounts were informed that instruments with software versions 4.B.0.9, 4.B.1.1 and 4.B.1.3 will be upgraded to version 4.B.1.4 software, provided the instruments have a 486 Motherboard or Single Board Computer, and that instruments containing versions of 4.A software are not affected. The accounts were requested to complete the enclosed Customer Reply form, indicating the instrument serial number and software version for each machine at the site, and if they would be requesting Baxter Technical Services to upgrade their instruments or have the software sent to them to make the upgrades, and fax the sheet to 1-847-270-5457. Any questions were directed to Baxter Global Technical Services at 1-800-553-6898, prompt 4, option 1. |  | 8258 | 8,258 units | Worldwide Distribution -- United States (including Puerto Rico and the Virgin Islands), Algeria, Argentina, Austria, Bahrain, Bangladesh, Brazil, Bulgaria, Canada, Chile, China, Colombia, Croatia, Denmark, Dominican Republic, Ecuador, Egypt, Finland, France, Germany, Greece, Guatemala, Honduras, India, Indonesia, Ireland, Italy, Jordan, Korea, Latvia, Lebanon, Lithuania, Malaysia, Mauritania, Mexico, Morocco, Nigeria, Oman, Panama, Paraguay, Peru, Philippines, Poland, Romania, Russia, Saipan, Singapore, Slovakia, South Africa, Spain, Sweden, Taiwan, Thailand, Trinidad, Tunisia, Turkey, Ukraine, United Kingdom, Uruguay and Venezuela. | Terminated | November 15, 2010 | 441 | 8258 | Computer
Z-0065-2010 | 52964 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | CELL-DYN Emerald Hematology Analyzer, model number 09H39-01. Distributed by Abbott Diagnostics Division, Santa Clara, CA. |  |  | July 16, 2009 | 2009 | Abbott Laboratories | Software bug. The CELL-DYN Emerald WBC L1 flag is configured incorrectly. The flag is not generated according to system requirements. | Software design | Abbott Diagnostics issued a "Product Correction" notice dated July 16, 2009 describing the affected product. The notice was sent to all currently active users and included a customer reply form, a software update and installation instructions.For further information, contact Abbott Customer Support at 1-877-4ABBOTT. Customers outside the U.S. should contact your local hematology customer support representative. |  | 411 | 411 units | Worldwide Distribution -- United States (AL, AR, AZ, CA, CT, FL, GA, IL, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NJ, NV, NY, OK, OR, PA, RI, SC, TN, TX, VA, WA, WI and WV), Canada, Germany, Puerto Rico, Singapore, and South Korea. | Terminated | December 06, 2010 | 508 | 411 | Computer
Z-0180-2010 | 53083 | Yes | LWS |  |  | Cardiovascular | PMA | Guidant LATITUDEÂż Communicator, model 6481, Boston Scientific 4100 Hamline Avenue North, St. Paul, MN 55112-5798 USA. The LATITUDEÂż Patient Management system enables physicians to periodically monitor both patient and device status remotely. |  |  | August 07, 2009 | 2009 | Boston Scientific CRM Corp | The L ATITUDEÂż communicator (Model 6481) for a patient may not have been able to receive measurements from their weight scale or blood pressure monitor afer initial setup. The issue has been resolved. If any patient weight or blood pressure readings are stored in the patient's sensor, they will be uploaded to LATITUDE after the patient's communicator calls in to Boston Scientific's servers. It | Software design | Boston Scientific used phone calls as the delivery method for these communications. The communication began August 6, 2009 for the original 6 patients affected. Notifications via phone calls to additional patients affected began on August 17, 2009. Direct questions about the recall to the LATITUDE Customer Service by calling1-651-582-4000. |  | 97 | 97 | Nationwide Distribution -- AZ, AR, CA, CT, FL, GA, IL, IN, MA, MO, NV, NJ, NY, OH, OR, PA, SC, TX, UT, VA and WA. | Terminated | December 04, 2012 | 1215 | 97 | Computer
Z-0558-2010 | 53687 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Zimmer/CAS Power Cord, Sesamoid Plasty, NA, Zimmer/CAS, Montreal (Quebec), Canada; REF 20-8000-070-12. The device is the power cord component for the Sesamoid Plasty CAS workstation which connects the workstation to the power mains. |  |  | November 02, 2009 | 2009 | Zimmer Inc. | The power cord female receptacle is not recessed sufficiently and may present a shock hazard. | Device Design | Zimmer sales staff were notified by letter dated 11/2/09 and instructed to locate the units and to upgrade the cords and to notify consignees by copy of a letter addressed to risk managers and dated 11/2/09. |  | 71 | 71 | Nationwide. | Terminated | October 15, 2010 | 347 | 71 | Computer
Z-0189-2010 | 53494 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Siemens Medical Solutions USA, Inc., Siregraph CF X-Ray System, Model Number: 4466033. |  |  | September 23, 2009 | 2009 | Siemens Medical Solutions USA, Inc | possible injury to hand/fingers | Device Design | Siemens Medical Solutions USA, Inc. issued a "Customer Safety Advisory Notice" dated September 11, 2009 to inform customers of the potential hazard to patients when using the Siregraph CF system.For further information, contact Siemens at 1-610-448-4500. |  | 14 | 14 units | Nationwide Distribution -- FL, IL, MI, PA and TX. | Terminated | January 12, 2010 | 111 | 14 | Computer
Z-0459-2010 | 53161 | aid, transfer | IKX | Daily activity assist device. | Physical Medicine | Physical Medicine | 510(K) Exempt | Bath Safe Adjustable Shower Seat, Without Back and Arms, Catalog number 727142001 |  |  | July 09, 2009 | 2009 | Nutech Manufacturing Corp | Safety chair used in shower/bath may not meet specifications for weight load. | Device Design | The subsidiary distributor, Maddak, Inc. notified customers by A "Voluntary Product Recall Notice", phone and e-mail on 07/09/2009. The own label distributor instructed customers to return their inventory of affected products received between May 1, 2009 and July 7, 2009. Distributors were instructed to immediately quarantine then return all affected inventory to Maddak and notify their accounts. Distributors were requested to instruct dealers to display a "storefront" notice and provide consumers with a customer response form. Distributors were also requested to forward a request to dealers to post a website notice if they have one. Recovered products will be replaced. |  | 3627 | 3,627 total units | Worldwide Distribution -- USA and Canada. | Terminated | May 29, 2013 | 1420 | 3627 | Computer
Z-0083-2010 | 53073 | No | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | Foot Pedal, Foot Control for SC2000 and SC2100 Powered, simple cranial drills, burrs, trephines, and their accessories. Catalog Numbers: EMAX2-FP and FPLUS-FP. |  |  | August 13, 2009 | 2009 | The Anspach Effort, Inc. | The Anspach Effort, Inc. is recalling the Foot Control/ Foot Pedal for SC2000 and SC2100 Powered, simple cranial drills, burrs, trephines, and accessories. The Foot Control Cables lacks the shield connection from the cable to the connector body. The lack of shielding may result in radiated emissions which could possibly cause interference on other electrical equipment adjacent to the foot pedal | Device Design | Anspach issued an "Urgent Product Removal" Letter dated August 17, 2009 informing customers of the affected product. Consignees are asked to screen their inventory and return any affected product including a completed Product Replacement Form to the firm.For further questions, contact Anspach at 1-800-327-6887 or 1-561-627-1080. |  | 14 | 14 | Worldwide Distribution -- Canada, Indonesia, United Kingdom, Korea and United States (NC, OK, TN, CA, KS, MI and WA). | Terminated | January 05, 2010 | 145 | 14 | Computer
Z-0548-2010 | 53987 | No | HDD | Obstetric table and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(K) Exempt | Hill-Rom Affinity Four Birthing Bed, Hill-Rom Company, Inc., Batesville, Indiana; Model P3700B. |  |  | November 16, 2009 | 2009 | Hill-Rom, Inc. | The brakes may not hold or lock. | Device Design | Hill-Rom issued an "Urgent Safety Notice" notification dated November 16, 2009 to customers. The letter advised users of the problem and asking them to locate the beds and to remove them from use until corrected if possible and, if not possible to inspect them for the problem and remove them from use if it is present. The firm will visit each account to make corrections.For further information, contact Hill-Rom Technical Support at 1-800-445-3720. |  | 2551 | 2551 (US), 4171 (OUS) | Worldwide Distribution -- United States, Canada, Argentina, Australia, Bahamas, Brazil, Chile, China, Denmark, Dominican Republic, France, Great Britain, Hong Kong, India, Indonesia, Ireland, Israel, Italy, Japan, Korea, Kuwait, Libiya, Netherlands, Poland, Puerto Rico, Russia, Saudi Arabia, Serbia, Singapore, Spain, Sweden, Syria, Thailand, Turkey and United Arab Emirates. | Terminated | July 02, 2012 | 959 | 2551 | Computer
Z-0188-2010 | 53499 | No | FSY | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | XTEN LED DF cupola, P/N: 567815999. |  |  | October 07, 2009 | 2009 | MAQUET Inc. | Several reports have been received regarding one model of spring arm (Acrobat 200 series) in combination with surgical light from various Manufacturers. All incidents pertain to a crack on the front joint of the Acrobat arm. | Device Design | Urgent Device Correction letters, dated October 7, 2009, were sent out by FedEx. The letter identified the affected products and asked customers to call 1-888-627-8383 #1, if they have any questions. The letter also provided a Technical Description along with a Customer Fax Response form. The letter stated the potential hazard and recommended precaution. If customers need additional info or have any questions they should contact their local MAQUET representative. |  | 346 | 346 | Nationwide Distribution. | Terminated | June 04, 2012 | 971 | 940 | Computer
Z-0191-2010 | 53581 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Biomet 5.5 mm peek Allthread PEEK Knotless Anchor, knotless, sterile, Biomet sports Medicine, Inc., Ontario, CA; REF 904842P. |  |  | September 09, 2009 | 2009 | Biomet, Inc. | These anchors may be difficult to remove, or may not detach from the inserter shaft. | Device Design | Biomet issued an "Urgent Medical Device Recall Notice" dated September 9, 2009 informing consignees of the affected product and requesting return of the device. Distributors were instructed to notify their customers copy of the recall letter and to locate and return the products from user facilities.For further information, contact Biomet at 1-800-348-9500 or 1-574-372-3983. |  | 894 | 894 | Worldwide Distribution -- United States and Belgium. | Terminated | January 19, 2011 | 497 | 894 | Computer
Z-0480-2010 | 53612 | No | FWZ | Manual operating table and accessories and manual operating chair and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Maquet Accessories, Connection Bracket. Catalogue # 1180.36A0. |  |  | October 22, 2009 | 2009 | Maquet Inc. | Product monitoring has determined that the connection bracket shows potential mechanical play at the OR table interface for some procedures. This is due to the hole in the lock washer, which is responsible for locking the accessory to the OR table, being too large. | Device Design | Urgent Device Correction Letters were sent by Federal Express to affected US hospitals on October 22, 2009. The letter stated the technical description, potential hazard, and the recommended precaution. The letter asked customers to immediately verify whether their Connection Bracket is affected thru the FaxBack provided. Questions or need for additional information can be directed to customers' local MAQUET representative. |  | 9 | 9 units in US. | Nationwide Distribution -- USA, including states of SC and PA. | Terminated | May 18, 2010 | 208 | 9 | Computer
Z-0001-2010 | 53426 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Penumbra Neuron Delivery Catheter 070, percutaneous catheter, catalog numbers PND6F0701058, PND6F070956, PND6F0701058M, and PND6F070956M; manufactured by Penumbra, Inc., Alameda, CA. |  |  | September 22, 2009 | 2009 | Penumbra Inc. | Failure-- product can kink or ovalize, which may result in difficulty in catheter advancement or delivery of other devices and damage to catheter, which could lead to patient injury or death as a result of portions being retained in the patient. | Device Design | The firm issued an Urgent Voluntary Field Removal Notice letter on 09/22/2009 to its consignees, asking them to inspect their stocks; remove affected product and place in quarantine for return to Penumbra Inc.; and complete and return the attached product identification / return form. |  | 2246 | 2246 units | Worldwide Distribution -- USA, Canada, Hong Kong, Chile, the UK, Denmark, Norway, Sweden, Finland, Netherlands, Germany, Switzerland, Austria, and Spain. | Terminated | November 24, 2010 | 428 | 2246 | Computer
Z-0259-2010 | 53603 | No | ERL | Ear, nose, and throat electric or pneumatic surgical drill. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | Medtronic Integrated Power Console System, IP X1 |  |  | July 17, 2009 | 2009 | Medtronic Xomed, Inc. | An error code (Code 13) is displayed on some of the consoles when a Medtronic Stylus or EHS handpiece is connected to the IPC indicating that the handpiece has not been recognized by the console. | Software design | A Product Correction Notification was issued to each consignee on June 15, 2009. Medtronic sales representatives will visit each facility to advise consignees of the issue and to install a software upgrade that will correct the handpiece recognition issue on the console(s). |  | 493 | 493 | Worldwide Distribution | Terminated | January 20, 2010 | 187 | 493 | Computer
Z-0498-2010 | 51734 | No | GAD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ILLICO Posterior Thoracolumbar Access Retractor, Part Number: 73520 |  |  | January 28, 2009 | 2009 | Alphatec Spine, Inc. | The recall was initiated after Alphatec Spine received a complaint on the ILLICO Posterior Thoracolumbar access retractor. It was reported that one of the retractor arms in the retractor broke during surgery. | Device Design | On January 28, 2009 Alphatec Spine contacted their domestic consignees via telephone and the international consignee was contacted via telephone and electronic mail. Alphatec Spine's telephone call consisted of informing the consignees of the product affected, the problem, what actions needed to be taken and what action Alphatec Spine would take to arrange for pick-up of the affected products. |  | 14 | 14 | AZ, CA, OR, TX, and Tokyo, Japan. | Terminated | December 22, 2009 | 328 | 14 | Computer
Z-0133-2010 | 53461 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | VARiS 6.2, 6.1 and 6.0 when used with LINK, Model numbers H88 and H85. Manufactured by Varian Medical Systems, Palo Alto, CA. |  |  | September 15, 2009 | 2009 | Varian Medical Systems Oncology Systems | Software issue: Under certain circumstances, when a therapy treatment field is transferred to VARiS Vision 6.2, 6.1 and 6.0 via LINK procedures, the rotation direction will be reversed. | Software design | Recall initiated September 15, 2009. An "Urgent - Field Safety Notice" letter was distributed to all affected users, with a description of the problem and corrective action steps. The letter was also distributed to Varian sales, marketing, and service organizations, informing them of the issue. For further information in the USA and Canada, contact Varian at 1-650-424-5731 or 1-888-VARIAN5 (888- 827-4265). In Europe, contact +41 41 749 8844. |  | 100 | 100 units | Worldwide distribution to 100 consignees throughout the US in IL, NY, MI, NJ, AL, TN, MD, OH, FL, and TX, and to Brazil, Chile, China, Czech Republic, Egypt, Germany, Hungary, Israel, Italy, Japan, Malaysia, the Netherlands, Romania, Slovakia, Slovenia, South Korea, Spain, Taiwan, Thailand and the United Kingdom. | Terminated | July 29, 2011 | 682 | 100 | Computer
Z-0264-2010 | 53131 | No | LYJ |  |  | Neurology | PMA | Cyberonics VNS Therapy Programming M250 Programming System, model number 250, version 4.6, 6.1, 7.0 and 7.1. The Model 250 Handheld software is provided on flashcards packaged in software envelopes or pre-installed on handheld computers. |  |  | July 24, 2009 | 2009 | Cyberonics, Inc | Some VNS Therapy System replacement Demipulse generators reporting low lead impedance readings. In rare instances, a system diagnostic test using Model 250 Programming Software (versions 7.1 and earlier) may report "Lead Impedance: OK" when a short-circuit condition exists. | Device Design | Cyberonics sent an Important Safety Notice on July 24, 2009, to all physicians explaining the event, risks presented and recommended actions to take. Labeling updates are being developed concurrently with the notification. The firm will mail or hand-deliver the notifications via firm representatives. |  | 9500 | 9,500 programming devices | Worldwide: USA, Argentina, Australia, Austria, Belgium, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Cyprus, Czech Republic, Denmark, Finland, France, Germany, Greece, Hong Kong, Hungary, Iceland, India, Iran, Ireland, Israel, Italy, Japan, Jordan, Kuwait, Lebanon, Luxembourg, Malaysia, Mexico, Morocco, Netherlands, Norway, Poland, Portugal, Puerto Rico (US Territory), Russia, San Marino, Saudi Arabia, Singapore, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Turkey, United Arab Emirates, United Kingdom, and Venezuela. | Terminated | September 26, 2012 | 1160 | 9500 | Computer
Z-0561-2010 | 47370 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Xoran MiniCAT System for ENT (server and work station), Xoran Technologies, Ann Arbor, MI. |  |  | September 17, 2008 | 2008 | Xoran Technologies, Inc. | The system may overheat, causing poor image quality, interruption in image acquisition and/or damage to the system. Also, if the system is not allowed to cycle down prior to being restarted after an emergency stop, the x-ray and gantry motions will both resume. | Device Design | A letter dated 9/17/08 was sent to each account telling them that the PLC would be upgraded. The firm visited each account and made corrections. |  | 82 | 82 | Nationwide Distribution | Terminated | December 23, 2009 | 462 | 82 | Computer
Z-0530-2010 | 53584 | No | KDD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | BD Recalled component is labeled: |  |  | January 09, 2009 | 2009 | Biomet Microfixation, Inc. | Biomet Microfixation, Inc. is recalling their OnPoint TMJ Kit (Part Numbers 24-3050, 24-3055) and Arthrosimplicity TMJ Arthocentesis Kit (Part Numbers 922070, 922090, 922075). These two kits contain a BD sterile disposable 60 ml Luer-Lok syringe, Part number 309653 which is under recall by Becton Dickinson (Recall Number Z-2347-2008). Becton Dickinson has determined that unit package seal integ | Device Design | Consignees were notified by a combination of telephone, fax, e-mail and certified mail. Questions related to this notice should be directed to Christy Cain at 1-800-874-7711. |  | 121 | 121 Kits | Product was distributed Internationally to Argentina, Denmark, Norway, and UK. Product was distributed domestically to AL, CA, CT, FL, GA, HI, IA, IL, IN, KS, LA, MA, MD, MN, MO, MS, NC, OH, NY, PA, SC, TN, TX, UT, and VA. | Terminated | March 17, 2010 | 432 | 235 | Computer
Z-0266-2010 | 53497 | No | CCT | Laryngotracheal topical anesthesia applicator. | Anesthesiology | Anesthesiology | 510(k) | Nasal Mucosal Atomizer Device with syringe, LATEX FREE For Single Patient Use Only Packaged Clean -Non-Sterile- MADE IN CANADA. |  |  | March 05, 2009 | 2009 | Sharn, Inc. | Sharn Anesthesia Inc., Tampa, FL received reports of leaking containers following use of their Sharn Anesthesia Microblast nasal and pediatric laryngeal atomizer device. | Device Design | Recalling firm is contacting all direct accounts (distributors &amp; user facilities) by telephone and letter. They request completion and return of the of the Return Goods Form and of any affected product by UPS using the firm's account number. |  | 550 | 550 | Nationwide Distribution | Terminated | July 13, 2012 | 1226 | 3225 | Computer
Z-0554-2010 | 52805 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Zimmer / Abbott Spine Sequoia dorsal height &amp; revision tool, non-sterile; REF 3367-1. Surgical instrument for pedicle screw system. |  |  | August 20, 2009 | 2009 | Zimmer Inc. | The tip of the instrument may fracture during use. | Device Design | Zimmer sent a letter dated 8/17/09 to distributors and consignees on or about 8/20/09. Users were notified of the potential problem via the letter and advised that the instruments will be removed and replaced when a new version of the instrument becomes available (around 10/09). |  | 215 | 215 | Nationwide, Belgium and New Zealand. | Terminated | April 13, 2011 | 601 | 215 | Computer
Z-0523-2009 | 49065 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Philips Diagnostic Ultrasound System, Model HDI4000. Product labeled HDI4000 Ultrasound System, Distributed by Philips Ultrasound, 22100 Bothell-Everett Highway, Bothell, WA 98021. Product also marketed as Philips 4000 New Standard, Philips 4000 System and Multibeam System. According to the firm, these units have the same user manual as the HDI4000 system. The different names are used internally as a description for the different configurations of model HDI 4000. Over the product lifecycle, the product configuration changed as follows: MULTIBEAM SYSTEM : This was the original HDI 4000 system branded with the former company name 'ATL' and labeled, "Distributed by ATL Ultrasound Bothell, WA. USA 98041"; PHILIPS 4000 SYSTEM :This configuration changed the legacy 'ATL' brand / colors to the 'Philips' brand / colors; PHILIPS 4000 NEW STANDARD: This Philips brand configuration eliminated hardware that supported the 'CV'. The CV option was initially standard, but later became optional depending on customer preference. The product ID label is the same for both PHILIPS 4000 SYSTEM and PHILIPS 4000 NEW STANDARD. |  |  | August 12, 2008 | 2008 | Philips Medical Systems | The system has a software defect that affects the display of Estimated Fetal Weight (EFW) Growth Percentile result values. | Software design | On 8/12/08, the firm sent an Urgent - Medical Device Correction letter, dated July 2008, to all affected customers via UPS. The letter describes the reason for the recall, informs them that Philips Ultrasound is correcting the error, states that the software is under development, will be available in September 2008 and will contact the customer to upgrade their system. "In the interim, Philips recommends that you only use the Hadlock or Osaka growth tables to perform your EFW Growth Percentile calculations. You should make the appropriate selections in the Analysis Setup menu on your system." |  | 1902 | 1902 Units | Worldwide Distribution --- including USA and countries of Argentina, Australia, Austria, Bahamas, Bahrain, Belgium, Bosnia, Brazil, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Dominican Rep, Ecuador, Egypt, El Salvador, France, Germany, Greece, Guam, Guatemala, Hong Kong, Hungary, India, Indonesia, Israel, Italy, Jordan, Kuwait, Malaysia, Martinique, Mexico, Moldova, Morocco, Netherlands, New Zealand, Peru, Philippines, Poland, Portugal, Puerto Rico, Romania, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Syrian Arab Rep, Taiwan, Thailand, Tunisia, Turkey, Ukraine, United Kingdom, Utd Arab Emirr., Venezuela, Viet Nam, Canada, Afghanistan, Albania, Argentina, Australia, Austria, Belgium, Bolivia, Bosnia, Brazil, Bulgaria, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican Rep, Ecuador, El Salvador, Estonia, Finland, France, Germany, Greece, Guatemala, Honduras, Hong Kong, Hungary, India, Indonesia, Iraq, Italy, Japan, Korea Rep, Latvia, Lebanon, Lithuania, Malaysia, Mauritius, Mexico, Moldova, Morocco, Namibia, Netherlands, New Zealand, Nicaragua, Norway, Oman, Panama, Paraguay, Peru, Poland, Portugal, Puerto Rico, Romania, Serbia, Singapore, SLOVAKIA, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, Uganda, Ukraine, United Kingdom, Uzbekistan, Venezuela, and Virgin Isl (Br). | Terminated | July 14, 2011 | 1066 | 6181 | Computer
Z-0005-2009 | 53451 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Toshiba Infinix fluoroscopy DFP-8000D Digital Radiography System with version 3.40 or higher software. Controller for Infinix interventional angiography systems. |  |  | September 29, 2008 | 2008 | Toshiba American Medical Systems Inc | Audible signal is not available at all times when using the HLC fluoroscopy mode. Special means of activation are not designed in for selection of High Level Controls, as required by regulation. | Software design | Toshiba issued a letter entitled Important Electronic Product Radiation Warning, September 30, 2008, explaining the issue. The customers were requested to reply to the letter via fax-back form, and a Toshiba service representative would install the corrective action. |  | 129 | 129 | Nationwide | Terminated | October 04, 2012 | 1466 | 129 | Computer
Z-0986-2009 | 50753 | No | MNI | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Stryker Trio Small Offset Connector |  |  | September 05, 2006 | 2006 | Stryker Spine | Reports were received of the ball ring ejecting from the connector's body when other connectors have already been locked down. | Device Design | A notification letter dated September 12, 2006 and a Product Accountability Form were sent via Federal Express (with return receipt) to all affected branches/agencies. Further information can be found by calling 1-201-760-8256. |  | 2417 | 2,417 units | Nationwide Distribution including OR, CA, MA, VA, GA, FL, NY, NM, NJ, UT, KS, OH, MI, MN, IN, MT, TX, WI, AL, CO, and TN. | Terminated | April 14, 2009 | 952 | 2417 | Computer
Z-1235-2009 | 51423 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo Imaging, picture archiving and communications system (PACS), Model number 10 014 063 |  |  | January 12, 2009 | 2009 | Siemens Medical Solutions USA, Inc | Incorrect and/or incomplete data: Several potential safety issues while working with the syngo Image product including, but not limited to, data loss, errors in patient data due to different patient IDs, delays in image availability, inaccurate results during image measurements, incorrect review or diagnosis of displayed or printed images, measurement inaccuracies from the host modality and image | Software design | syngo Imaging has issued an Addendum to the User Manual, which was distributed to affected customers via Update Instructions IM118/08/S in November, 2008. Two safety related hotfixes (VB20G_HF4 and VB30A_HF26) will be released to those affected only after Update Instructions IM118/08/S has been performed. The addendum informs the user about safety relevant notes which pertain to syngo Imaging versions VB20G and VB30A. Update is a prerequisite for the following safety related hotfixes: VB20G_HF4 and VB30A_HF26. |  | 55 | 55 units | Nationwide Distribution. | Terminated | September 23, 2009 | 254 | 55 | Computer
Z-1266-2009 | 51563 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo Imaging model number 10014063. Affected systems are V30A or earlier with an NFS connected archive and V30A systems where image data was stored to LTS using software version older than V30A. |  |  | February 09, 2009 | 2009 | Siemens Medical Solutions USA, Inc | An issue concerning archived image data that is sporadically saved with corrupted data may occur. This issue may occur when syngo Imaging is configured to archive image data on a network file system (NFS) connected via a connected archive (Mount Point Solution). | Software design | A Customer Safety Advisory Notice, dated 02/09/2009 was sent to affected customers via Update Instruction IM013/09S. The letter informs customers of when the malfunction occurs and the potential issue, provides instructions to avoid its occurrence, and how the issue will be resolved. Customers are to ensure that the Safety Advisory is placed in the system's instructions for use and contact their SIEMENS Uptime Service Senter if they have any questions. |  | 20 | 20 units | Nationwide Distribution -- including states of AL, CA, CT, FL, LA, MI, MO, NC, NJ, OH, PA, TN, and WI. | Terminated | September 23, 2009 | 226 | 20 | Computer
Z-1123-2009 | 51276 | No | GDC | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Operon Electrical Operating Room Table. |  |  | February 02, 2009 | 2009 | Berchtold Corp. | Difficulty in pressing the button needed to release a bed part resulted in an injury to a caregiver's finger which became stuck in metal latch under the gas cylinder of the bed piece. | Device Design | An official notification letter dated February 3, 2009 was hand-carried by a local Berchtold Corporation Technical Service representative to each account. The letter indicated that the Berchtold Corporation Service Technician will complete recommended upgrades. Direct questions concerning this recall to your BERCHTOLD Technician Service at 1-800-243-5135 and press the Number 2. |  | 453 | 453 units | Worldwide Distribution -- US and Canada. | Terminated | June 04, 2009 | 122 | 453 | Computer
Z-0995-2009 | 47442 | No | LYJ |  |  | Neurology | PMA | Cyberonics VNS Therapy Demipulse Duo Generator, Model 104; an implantable, multiprogrammable, vagus nerve stimulator. |  |  | February 07, 2008 | 2008 | Cyberonics, Inc | Reset/disabling of the VNS Therapy Demipulse Generator and Demipulse Duo Generator due to magnet interference, resulting in the loss of stimulation. | Software design | Cyberonics notified consignees of recall via letter on 02/07/08. Consignees were notified of reset issue in which during a strong magnetic field, the software may report an error even though no error occurred. This error causes the generator to reset, which disables the generator (i.e., no stimulation delivered).Physicians should inform any potential and currently implanted patients of the possibility of this event. Patients should continue with their regularly scheduled visits as determined by their physician.If this event is suspected, physicians should interrogate the device as described in the Physicians Manual to confirm the generator has been reset and Cyberonics should be contacted for any additional instruction. |  | 108 | 108 units. | Worldwide Distribution - USA, Belgium, Germany, Norway, Spain, Sweden, the Netherlands, and United Kingdom. | Terminated | February 28, 2011 | 1117 | 795 | Computer
Z-1003-2009 | 50918 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | GE Healthcare Infinia HawkeyeÂż 4, Model Numbers: 5178728, ASM500060, H3000WC and H3000WD. |  |  | June 02, 2008 | 2008 | GE Medical Systems, LLC | GE Healthcare has recently became aware of excessive pressure applied by the collimator pressure sensitive cover of Infinia (Infinia, Infinia HawkeyeÂż, Infinia HawkeyeÂż 4) that may impact patient safety. It was reported at a customer site that the pressure used to activate the Pressure Sensitive Device (PSD) should be lowered and that the body contouring software requires correction to prevent po | Device Design | Consignees were sent from June 2008 to February 2009 a GE "Product Safety Notification" letter dated April 2008. The letter was addressed to Hospital Administrators and Managers of Radiology/Cardiology, Radiologists/Cardiologists. The letter described the Safety Issue, Affected Product, Safety Instructions, Product Correction and Contact Information.A GE Healthcare service representative will contact each consignee to schedule the installation of the Pressure Sensitive Device (PSD). Please contact your local sales or service representative for further information regarding this recall. Further information can be found by calling 1-800-437-1171 in the US or Canada or 1-262-896-2890 if outside the US. |  | 288 | 122 US, 166 OUS | Worldwide Distribution US (and Puerto Rico) including states of AL, AK, AZ, AR, CA, CO, CT, DE, DC, FL, GA, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, SC, SD, TN, TX, UT, VT, VA, WA, WV and WI and countries of Australia, Austria, Belgium, Brazil, Canada, China, Colombia, Czech, Denmark, Ecuador, Egypt, Finland, France, Germany, Greece, Hungry, India, Indonesia, Ireland, Israel, Italy, Japan, Korea, Malaysia, Martinique, Mexico, Netherlands, New Zealand, Norway, Pakistan, Panama, Philippines, Poland, Portugal, Republic of Georgia, Russia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, Ukraine and the United Kingdom. | Terminated | May 24, 2012 | 1452 | 1256 | Computer
Z-1239-2009 | 51492 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | Edwards Lifesciences Swan-Ganz Intellicath CCO/VIP Thermodilution Catheter with AMC Thromboshield (an Antimicrobial *Heparin Coating) For use with Vigilance Monitor. REF: 139HF75 |  |  | February 09, 2009 | 2009 | Edwards Lifesciences, Llc | Cardiac output measurements were unavailable and the following message is displayed in the Vigilance monitor: "Fault. CCO Catheter Verification. Use Bolus Mode". | Software design | Edwards initiated the worldwide recall via Urgent Customer Notification letters dated February 9, 2009 and are being sent by Federal Express starting on February 10, 2009. These letters are requesting that customers cease using Swan-Ganz" Continuous Cardiac Output (CCO) Thermodilution Catheter products from the indicated lots and that they return all unused product. Edwards Lifesciences requested that the customer provide written verification of stock and identify any unused product to be returned. Questions should be directed to Technical Service at 800-822-9837. |  | 107 | 107 units | Worldwide Distribution -- AL, AR, AZ, CA, CT, FL, GA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MO, MS, MT, NC, NE, NJ, NM, NY, OH, OK, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, WY, Hawaii, and Puerto Rico and countries of Canada, Austria, Belgium, Switzerland, Germany, Denmark, Spain, Finland, France, UK, Greece, Ireland, Israel, Italy, Kuwait, Netherlands, Norway, Portugal, Sweden, Colombia, Australia, and Singapore. | Terminated | April 11, 2011 | 791 | 10754 | Computer
Z-0989-2009 | 50872 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | Deltec CozmoÂż Insulin Pump, Model 1700 containing 1800 software. |  |  | November 20, 2008 | 2008 | Smiths Medical MD, Inc. | Smiths Medical has become aware of a display irregularity with the Deltec CozmoÂż Insulin Pump, Model 1800 and Model 21-1700 containing Model 1800 software. There have been adverse events reported that the amount of Extended Bolus delivered was not accurately displayed by the Pump. The amount of extended Bolus displayed on the Pump Home Screen 2 and in the Bolus Summary Report is less than what w | Software design | Consignees were sent a Smiths Medical "Urgent Device Recall" letter dated November 20, 2008. The letter was addressed to Distributors, Clinicians, Patients and other Customers. The letter included a Description of the Problem, Advice on Action to be Taken by the User and requested the return of the Confirmation Form.In particular, users were given the option to check the "Replacement Pump" option on the Confirmation Form initiating the process for pump replacement with a recertified pump followed by a return of the defective product. The Confirmation Form can be faxed to 1-800-628-6322 in the U.S. or 1-651-628-7485 outside the U.S., completed at www.cozmore.com/notice or calling 1-800-501-5748 in the U.S. |  | 586 | 586 | Worldwide Distribution --- Including states of AL, AK, AZ, AR, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MI, MA, MD, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, WY and DC and countries of AUSTRALIA, NEW ZEALAND, UNITED KINGDOM, IRELAND, CANADA, CHINA, GREECE, POLAND, SLOVENIA, SWEDEN, ISRAEL and BRAZIL. | Terminated | June 05, 2012 | 1293 | 17047 | Computer
Z-1154-2009 | 51297 | No | FTD | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Alm Prismatic Surgical Light 5004 SF; |  |  | February 23, 2009 | 2009 | MAQUET Inc. | The screws holding the Prismatic triple surgical lights may break due to dynamic multidirectional bending stress. The firm has received one complaint report. | Device Design | An "Urgent Device Correction" letter dated February 23, 2009 was mailed by FedEx Tracking to all affected consignees. The letter described the issue, potential hazard and recommended precaution. Customers were instructed to complete the enclosed "Customer Response Fax Form" and return it via fax (1-732-667-1903) to Maquet, Inc. Maquet, Inc. will schedule repair service of the affected units once the "Customer Response Fax Form" is received. Direct questions to your local Maquet, Inc. Representative. |  | 54 | 54 of all units, both article numbers | Nationwide Distribution. | Terminated | September 29, 2009 | 218 | 54 | Computer
Z-1094-2009 | 50990 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | GE Healthcare Infinia Hawkeye(R), Model numbers: |  |  | September 08, 2008 | 2008 | GE Medical Systems, LLC | There is a possibility for misregistration when brain scans are performed with Infinia with the Hawkeye option. This misregistration could present the potential for misdiagnosis when the patient headrest sags under the weight of a head. | Software design | Consignees were sent from September 2008 - Feb 2009 a GE "Product Safety Notification" letter dated 4/2008. The letter was addressed to Hospital Administrators, and Managers of Radiology/Cardiology, Radiologists/Cardiologists. The letter described the Safety Issue, Affected Product, Safety Instructions, Product Correction and Contact Information. It instructed consignees to continue to use the Infinia Hawkeye system normally for all types of Hybrid scans and stated that as a general practice tomograms should be reviewed carefully using the Xeleris Volumetrix processing protocol to ensure proper registration. Local sales or service reps should be contacted for further information regarding the notification. |  | 417 | 153 US, 264 OUS | Worldwide Distribution -- USA including states of AL, AK, AZ, AR, CA, CO, CT, DE, FL, DC, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK ,OR, PA, PR, SD, TN, TX, UT, VT, VA, WA, WV, and WI and countries of AUSTRALIA, BELGIUM , BRASIL, CANADA, CHINA, COLUMBIA, CZECH REPUBLIC, DENMARK, ECUADOR, FINLAND, FRANCE, GEORGIA, GERMANY, INDIA, INDONESIA, ISRAEL, ITALY, JAPAN, KOREA, MEXICO, NETHERLANDS, NEW Zealand, NORWAY, POLAND, PUERTO RICO, RUSSIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TURKEY, UKRAINE, UNITED KINGDOM. | Terminated | May 24, 2012 | 1354 | 639 | Computer
Z-1065-2009 | 50171 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Roche/Hitachi Modular Analytics System, clinical chemistry analyzer Modular D2400 Module, Roche Diagnostics Corp., Indpls., IN; GMMI No. 04998618001. |  |  | November 12, 2008 | 2008 | Roche Diagnostics Corp. | If a masked module is unmasked during the reagent registration process, the calibration parameters may be assigned to the wrong reagents, resulting in the reporting of incorrect control and sample results. | Software change control | Consignees were notified via an Urgent Medical Device Correction letter dated 11/12/08 and instructed to only execute "module unmasking" when the unit is in "standby" mode and to consult with their physicians at the facility to determine the clinical implications and whether any clinical intervention is needed for patients treated based upon results obtained with this software. Questions should be directed to Roche Diagnostics Technical Support at 1-800-428-2336. |  | 22 | 22 systems in all | Nationwide Distribution -- including states of California, Florida, Georgia, Hawaii, Indiana, Maryland, New York, North Carolina, Pennsylvania, and Virginia. | Terminated | September 22, 2009 | 314 | 22 | Computer
Z-0002-2009 | 53157 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Revolution XR/d system control room PC model 5117866-2 with Software collector 5115434 |  |  | February 14, 2008 | 2008 | GE Medical Systems, LLC | GE Healthcare has recently become aware of a potential risk of observing a blank image after exposure onRevolution XR/d systems equipped with HP 8200 PC user consoles that may impact patient safety.It was reported that images acquired with off-center collimation for extremities with theposition selected as Digits to front or Digits to back are displayed as blank. Investigation intoth | Software design | Consignees were sent a GE Healthcare "Product Safety Notification" letter dated March 7, 2008. The letter was addressed to Hospital Administrators, Mangers of Radiology and Radiologists. The letter listed the Safety Issue, Affected Product, Safety Instructions, Corrective Action and Contact Information. |  | 129 | 129 | Nationwide Distribution -- AL, AZ, CA, CO, FL, GA, HI, IL , IN, KY, LA, MD, MA, MI,MN,MS, MO, MT, NE, NH, NJ, NM, NY, NC, OH, OK OR, PA, SC, TN, TX, VA, WA, and WV. | Terminated | January 07, 2012 | 1423 | 129 | Computer
Z-1172-2009 | 51135 | Yes | LWS |  |  | Cardiovascular | PMA | Medtronic RV Lead Integrity Alert , Cat. # SW012, 1.0 software in EnTrust(R) (D153ATG, D153VRC, D154ATG, D154VRC) and EnTrust(R) Escudo (D144DRG, D144VRC) defibrillators. |  |  | November 03, 2008 | 2008 | Medtronic Inc. Cardiac Rhythm Managment | Medtronic has identified an issue related to installation or removal of the Lead Integrity Alert software in EnTrust(R) and EnTrust(R) Escudo" defibrillators. In those devices only, installation or removal of LIA will inadvertently turn off two audible patient alerts. | Software design | Between November 24 and December 9, 2008 US CRDM are responsible to hand deliver the consignee letter, a Medtronic "Urgent Patient Management Information," dated November 2008, to impacted physicians to make them aware of the issue and continue to promote and install LIA (by using the new LIA tipcard). For any physician listed that a confirmation sheet was not received by end of day, December 9, 2008, the letter was mailed to the physician on Dec 10, 2008. The letter described the issue and the product involved. The letter also gave recommendations and informed consignees in regard to a future update to the LIA software for EnTrust devices. |  | 55755 | 55,755 | Worldwide Distribution -- USA including states of AL, AK, AZ, AR, CA, CO, CT, DE, FL, GA, HI, DC, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MI, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY and countries of Argentina, Aruba, Australia, Austria, Bahamas, Barbados, Bolivia, Brazil, Canada, Cayman Islands, Chile, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Finland, France, Germany, Greece, Hong Kong, Hungary, Iceland, India, Ireland, Israel, Italy, Jamaica, Jordan, Kuwait, Lebanon, Macedonia, Martinique, Netherlands, Netherlands Antilles, Norway, Oman, Peru, Poland, Portugal, Puerto Rico, Qatar, Reunion, Saudi Arabia, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Turkey, United Arab Emirates, United Kingdom Vatican City State, Venezuela, and Virgin Islands. | Terminated | May 10, 2012 | 1284 | 55755 | Computer
Z-0999-2009 | 50875 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Viasys Healthcare, Sonora Transcranial Doppler (TCD) System. Model Sonara. Part Number of Units: 982A0470 Sonara Bilateral. |  |  | January 06, 2009 | 2009 | Cardinal Health, Inc | Internal testing has identified several several scenarios where unexpected System lockups, shutdowns or system errors would render the Sonara unavailable for use until the Sonara can be restarted. These issues can be corrected by installing the Version 6.0 b045 software updating the Audio Driver and updating to XP Service Pack 3. Additionally, improper wiring could cause the cooling fan to not op | Device Design | Consignees were sent a Cardinal Health "Urgent: Medical Device Correction" letter dated January 12, 2009 beginning the week of February 2, 2009 and continuing into week of February 9, 2009. The letter was addressed to customers describing the problems and systems affected and solutions. Cardinal Health requested customers to complete the actions listed in the letter to install the Field Correction and return the completed "Verification Form" via mail, e-mail or fax to Cardinal Health-NeuroCare. Further information can be obtained by calling 1-800-356-0007 ext 2191 or fax to 1-877-244-3805. |  | 64 | 64 systems | Worldwide Distribution -- USA including states of FL, DC, LA, NJ, NY, PA, TN, TX, WI, CA, MO, and LA and countries including THAILAND, SINGAPORE, SERBIA-MONTENEGRO, PORTUGAL, PHILIPPINES, POLAND, NORWAY, NETHERLANDS, KOREA, ITALY, ISRAEL, HUNGARY, HONG KONG, GREAT BRITAIN, SPAIN, GERMANY, CZECH REPUBLIC, SWITZERLAND, CANADA, BULGARIA, and BOSNIA-HERZEGOVINA. | Terminated | July 03, 2012 | 1274 | 100 | Computer
Z-1039-2009 | 51065 | Yes | MSX | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, S/5 iCentral and iCentral Client, |  |  | January 26, 2009 | 2009 | GE Healthcare | GE Healthcare is providing information about two issues associated with iCentral L-NETO5 and L-NETC05 v4.3, 4.4, 4.5 and 5.0 when used in combination with Dash 2000, 3000, 4000 and 5000; Solar 7000, 8000, 8000M and 8000i; Eagle 3000 and 4000 and ApexPro. The iCentral does not provide audio or visual alarming for the following: A FIB, VT&gt; and PVC(1) alarms when used with ApexPro and VT&gt;2 and PVC | Software design | Consignees were sent a GE Healthcare "Urgent Medical Device Correction" letter dated January 26, 2009 addressed to Healthcare Administrator/Risk Manager, Chief of Nursing and Director of Biomedical Engineering. The letter described the safety issue, affected product details, safety instructions, product correction and contact information. Direct questions to GE Healthcare at 1-800-558-7044, Option 2 and ask for Technical Support or contact your local GE Healthcare Distributor. |  | 2146 | 2,146 | Worldwide Distribution, USA (and PR) including states of AZ, CA, CO, CT, FL, GA, IL, IN, IA, KS, KY,LA, ME, MI, MN, MS, MO, NH, NJ, NY, NC, OH, PA, SC, TX, VT, VA, WA, and WI and countries of AUSTRALIA, AUSTRIA, BARBADOS, BELGIUM, BOLIVIA, BRASIL, BULGARIA, CANADA, CHILE, CHINA, COSTA RICA, CROATIA, CZECH REPUBLIC, DENMARK, ECUADOR, ESTONIA, FINLAND, FRANCE, GEORGIA, GERMANY, GREECE, GUATEMALA, GAYANA, HONG KONG, HUNGARY, ICELAND, INDIA, IRAN, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KOREA, KUWAIT, LIBYAN ARAB JAMAHIRIYA, MACEDONIA, MALAYSIA, MEXICO, NETHERLANDS, NEW ZEALAND, NICARAGUA, NORWAY, PAKISTAN, POLAND, PORTUGAL, QATAR, ROMANIA, RUSSIA FEDERATION, SAUDI ARABIA, SINGAPORE, SLOVAK IA, SOUTH AFRICA, SLOVEN IA, SOUTH KOREA, SPAIN, SWEDEN, SWITZERLAND, SYRIA ARAB REPUBLIC, THAILAND, TUNAS IA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY, VENEZUELA and YEMEN. | Terminated | August 13, 2013 | 1660 | 2146 | Computer
Z-1231-2009 | 51223 | No | FPO | Wheeled stretcher. | General Hospital | General Hospital | 510(K) Exempt | Stryker Electric Stretcher with the Quick Drop CPR option, Model 1550. |  |  | February 27, 2009 | 2009 | Stryker Medical Div. of Stryker Corporation | The head section backrest (fowler) may not raise up, or, if up, may drop suddenly due to wearing out of the actuator threads on the fowler. | Device Design | Consignees were notified by an Urgent - Medical Device Correction letter dated 2/27/09, advised of the problem and instructed to locate stretchers and verify that there is no grinding noise when activating the fowler electrically. Consignees were also instructed to remove any stretchers with a fowler that was non-operational, or which makes a grinding sound, from service until the firm makes repairs and contact the firm. Stretchers will be corrected within the next four months by a Stryker field service representative. The enclosed post card should be returned to confirm receipt of the notification. The notice should be forwarded to any users that have borrowed or purchased the stretchers and the firm should be notified of their new location. Also, if any of the affected stretchers have been disposed of or are no longer in use, the firm needs to be provided with their serial numbers. Questions or concerns should be directed to Andrea Kuthe at 269-324-6581. |  | 3607 | 3607 | Worldwide Distribution | Terminated | April 25, 2011 | 787 | 3645 | Computer
Z-1042-2009 | 50704 | No | JJQ | Colorimeter, photometer, or spectrophotometer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Diasorin LIAISON (R) Chemiluminescence Analyzer Model # 15970. |  |  | December 05, 2008 | 2008 | Diasorin Inc. | DiaSorin determined that there was a potential for a short to occur in a a residual starter pump electrical cable in the LIAISON (R) Analyzer. | Device Design | DiaSorin Inc authorized Field Service Engineer has visited the sites and performed the necessary field correction. Upon completion of their work, the field service engineer generated a Field Service Report and a copy was left with the customer. |  | 9 | 9 | Worldwide Distribution -- USA including states of NC, NJ, PA, CA, WA, MO, UT, and OR and country of Canada. | Terminated | May 22, 2012 | 1264 | 9 | Computer
Z-1133-2009 | 51287 | No | JXG | Central nervous system fluid shunt and components. | Neurology | Neurology | 510(k) | Medtronic Innervision Snap Shunt Ventricular Catheter, BioGlide, Catalog Number 27782B. Sizes: 5 CM in length, REF (240) 27782B. Each catheter is provided in a sealed Tyvek plastic tray placed within a sealed Tyvek/Mylar pouch. The sealed pouch is then placed within a chipboard carton along with product IFU and patient wallet cards. The product is ETO sterilized. |  |  | February 12, 2009 | 2009 | Medtronic Neurosurgery | The ventricular catheters may become detached from the Snap Base assembly after implantation and this condition may increase the need for emergency revision surgery. | Device Design | An Urgent - Medical Device Recall notification was sent on February 13,2009 to the attention of risk managers, advising them to cease use of the Innervision Snap Shunt Ventricular Catheter, BioGlide, Catalog Number 27782 ( all versions and all lot numbers), the Snap Shunt Ventricular Catheter, BioGlide, Catalog Number 27802 ( all versions and all lot numbers), and the Snap Shunt Ventricular Catheter, BioGlide, Catalog Number 27708 ( all version and all lot numbers). Each customer is to cease the use of the listed product and any unused unit needs to be returned to Medtronic. Further, each account is required to account for product used/implanted and to advise surgeons of the recall information. Each account will recieved an individualized Customer Product Reconciliation Form provided by their Medtronic Sales Representative for determining the status of any product shipped to them. |  | 3048 | 3048 of all units | Worldwide Distribution -- USA, Ireland, and Australia. | Terminated | November 08, 2010 | 634 | 3048 | Computer
Z-1040-2009 | 50163 | No | NQN | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(k) | QCA and iQCA software, a video microscopy software system for quantitative estrogen receptor immunohistochemistry; Cell Analysis Inc., Evanston, IL 60201 or Highland Park, IL 60035. |  |  | November 06, 2008 | 2008 | Leica Microsystems, Inc. | A change in the software to open up the software architecture to include other analytes other than Estrogen Receptor was made without FDA review and approval. | Software change control | Leica Microsystems sent Medical Device Correction letters dated 11/6/08 to the direct accounts, informing them that all versions 3.7 and above of QCA and iQCA software are not in compliance with the original 510(k) cleared software for scoring of immunohistochemical stains in clinical applications. The QCA/iQCA software is only cleared for Estrogen Receptor staining under clinical applications. The accounts were provided with red warning stickers that state: "WARNING The clinical application of QCA software has been cleared only for the measurement of images from microscopic slides of breast cancer specimens stained for the presence of estrogen (ER) nuclear receptor protein only. Performance characteristics of the product, when used outside of this intended use, have not been established." They were instructed to apply the stickers to any existing manuals and or at each point of use, and to instruct all users about the notification. The accounts were also requested to complete and return the enclosed reply form confirming that the changes have been made. |  | 5 | 5 units | Nationwide Distribution -- including states of CA, IL, KS, and NY. | Terminated | March 16, 2011 | 860 | 5 | Computer
Z-1090-2009 | 49548 | No | MKQ |  |  | Pathology | PMA | BD PrepStain System, Model Number: 799-14000-00; Catalog Number: 05CR000021, 05CR00021R, 799-13000-00, 799-13000-00R, 799-14000-00 and 799-14000-00R; Catalog Number: 05CR000021, 05CR000021R, 799-13000-00, 799-13000-00R, 799-14000-00 and 7914000-00R. |  |  | September 26, 2008 | 2008 | Bd Diagnostic Systems Tripath | Software Error in the non-gyn program of some devices may cause non-gyn samples to be incorrectly transferred from centrifuge tubes to slides potentially result in mismatched patient results. | Software change control | BD Diagnostics began contacting known non-gyn customers by phone on 09/26/2008. Customers were notified by Urgent - Important Product Information letter on 09/29/2008. They were advised to stop using the system until they complete a list of actions provided in the letter. A telephone contact number for Technical Support was provided in the letter. A permanent software based solution will be provided to customers when completed. They anticipate the software will be available in October 2008. |  | 662 | 662 units | Worldwide Distribution | Terminated | January 15, 2013 | 1572 | 662 | Computer
Z-1092-2009 | 51159 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Innova 3100 / 3100 IQ, Cardiovascular Imaging System, GE Healthcare, Waukesha, WI 53188 (The system consists of an a monoplane positioner, a vascular or cardiac table, an X-ray system and a digital detector) |  |  | March 06, 2008 | 2008 | GE Medical Systems, LLC | GE Healthcare reported an issue with the Exposure Switch Redundancy system with Innova Systems 3100 &amp; 4100, in which this error stays on, even after the pedal is released, inhibiting further x-rays, and a system reset is required to recover. | Software design | On site visits were performed from 3/6/08 - 12/19/08 to update software on all 5 units. |  | 4 | 4 | Worldwide Distribution -- USA including states of MD, TX, and VA and country of China. | Terminated | September 04, 2014 | 2373 | 5 | Computer
Z-1236-2009 | 51319 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Cerner Bridge Medical Transfusion Administration and Specimen Collections, Versions 3.4.0, 3.4.1 and 3.4.2. |  |  | January 19, 2009 | 2009 | Cerner Corp | Specimen labels printing incorrect patient demographics. | Software design | Cerner Corporation issued a "Priority Review Flash" dated January 15, 2009 via e-mail to consignees using the software. The letter described the violative nature of the software and explained the proper navigation controls needed to prevent further issues.Further questions should be addressed to Bridge Client Support at 1-800-786-0239. |  | 4 | 4 | Nationwide distribution to states of OH, NC, MO and MI. | Terminated | March 22, 2012 | 1158 | 4 | Computer
Z-1237-2009 | 51432 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris Patient-Controlled Analgesia (PCA) Module (formerly Medley PCA Module) Model 8120 used with the Alaris PC Unit with software versions 8 through 9.1. |  |  | March 04, 2009 | 2009 | Cardinal Health 303 dba Cardinal Health | The Alaris PCA module can potentially infuse above or below the intended infusion dose. | Software design | The field action was initiated on March 11, 2009 with Cardinal Health sending by registered return receipt mail a Safety Alert to all customers using the Alaris PC Unit with software versions 8 through 9.1, which encompasses the subset of those customers that use the Alaris PCA Module. The Safety Alert was addressed to the Directors of Nursing, Pharmacy, Risk Management, and Biomedical Engineering at each hospital facility. If a customer did not respond to the Safety Alert, then Cardinal Heath notified that customer by telephone to verify receipt of the Safety Alert.This Safety Alert Informed the customers about a potential risk, and required them to: (1) Determine whether you use version 8 through 9.1 of the Alaris PC Unit. (2)Forward this Safety Alert to all Data Set owners (3) To immediately identify the potential risks review enclosed FAQs and Instructions for Data Set Review. (4)Continue to follow the Institute for Safe Medication Practices (ISMP) and the Joint Commission recommendation for safe medication administration - two nurses checking to ensure that the correct prescription is being entered into the PCA pump when first ordered and when there are any changes in the prescription (paying close attention to dosing units and infusion modes). The customers were also instructed to call our Support Center at 1-888-562-6018 with any questions. An updated Customer Recall Notification was forwarded by registered return receipt mail by Cardinal Health on April 15,2009. This notice reinforced the information in the first notice in addition to informing the customers of the hazard of an over-infusion or under-infusion of medication. Also, customers were given instruction for reporting any adverse reactions experienced with the use of the affected product and/or quality problems via the FDA's MedWatch Program. |  | 164000 | Approximately 164,000 units | Worldwide Distribution -- USA, Canada, Germany, Malaysia, and Singapore. | Terminated | June 26, 2012 | 1210 | 164000 | Computer
Z-1232-2009 | 51253 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare CIC Pro" Software Versions 5.0.3, 5.0.6 ,5.0.7, 5.0.8 and 4.1. |  |  | February 27, 2009 | 2009 | GE Medical Systems Information Tecnology | GE Healthcare is aware of potential network time issues associated with the time change aspects of certain CARESCAPE CIC Pro" software versions. These issues may impact patient safety. No reported patient injuries or illnesses for the use of this device in conjunction with this device correction. | Software design | Consignees were sent a GE Healthcare "Urgent Medical Device Correction" letter dated February 27, 2009. The letter was addressed to the Chief of Nursing/Chief of Biomedical Engineering and Risk Manager/Healthcare Administrator. The letter describes the Safety Issue, Affected Product Details, Safety instructions, Product Correction, and Contact Information.Direct questions to your local GE Healthcare Field Service Representative or the GE Healthcare Technologies Support Center by calling 1-800-558-7044 (Option 2, Technical Support). |  | 12737 | 12,737 | Worldwide Distribution including US (states of AL, AK, AZ, AR, CA, CO, CT, DE, FL, GA, DC, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, OH, PR, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WY, WI, and WV) and countries of UNITED ARAB EMIRATES, AUSTRIA, AUSTRALIA, BOSNIA AND HERZEGOVINA, BANGLADESH, BELGIUM, BRAZIL, BAHAMAS, CANADA, SWITZERLAND, CHILE, CHINA, COLOMBIA, COSTA RICA, CUBA, CYPRUS, CZECH REPUBLIC, GERMANY, ESTONIA, EGYPT, SPAIN, FINLAND, FRANCE, GHANA, GREECE, HONG KONG, HUNGARY, HONDURAS, INDONESIA, IRELAND, ISRAEL, INDIA, IRAQ, BRITISH INDIAN OCEAN TERRITORY, ITALY, JORDAN, JAPAN, SOUTH KOREA, KUWAIT, LEBANON, LITHUANIA, MALTA, MEXICO, MALAYSIA, NIGERIA, NETHERLANDS, NORWAY, NEPAL, NEW ZEALAND, PANAMA, PERU, PHILIPPINES, PAKISTAN, POLAND, PORTUGAL, QATAR, RUSSIAN FEDERATION, SAUDI ARABIA, SUDAN, SW DEN, SINGAPORE, SLOVENIA, SLOVAK REPUBLIC, SYRIA, THAILAND, TURKEY, TRINIAD, TOBAGO, TAIWAN, UKRAINE, UNITED KINGDOM, VENEZUELA, YEMEN, and SOUTH AFRICA. | Terminated | September 15, 2014 | 2026 | 12737 | Computer
Z-1254-2009 | 51560 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | AXIOM Sireskop SD Model numbers 8890415 and 8890407. The product is intended for use as radiographic and fluoroscopic imaging device. |  |  | January 08, 2009 | 2009 | Siemens Medical Solutions USA, Inc | During interventional use and applications, the possibility exists for liquids to enter the systems. | Equipment maintenance | A Customer Safety Advisory Notice dated January 8, 2009 was issued to affected customers via Update Instructions AX069/08/S. This letter informed customers of the potential issues and provided instructions to avoid its occurrence. The firm has released an Addendum to the User Manual warning of the potential hazards if liquids penetrate the systems during interventional use and applications.Direct questions about this recall to Siemens Medical Solutions USA, Inc. by calling 1-610-219-6300. |  | 240 | 240 units | Nationwide Distribution. | Terminated | September 01, 2009 | 236 | 909 | Computer
Z-1424-2009 | 51952 | No | DWE | Cardiopulmonary bypass pump tubing. | Cardiovascular | Cardiovascular | 510(k) | Medtronic, Transfer Spike, Blood Processing Accessory, REF BT945, Sequestration Kits |  |  | February 25, 2009 | 2009 | Medtronic Cardiovascular Revascularization &amp; Surgical Therap | Sterility may be compromised. Non-vented caps that connect to the 3-way stopcock attached to the transfer spike could prevent Ethylene Oxide from contacting the inside surfaces of the stopcock during sterilization, leading to a potentially non-sterile product. Use of the product may cause serious health consequences from potential sepsis/infection and the associated remote risk of death. | Device Design | Consignees were sent a Medtronic "Urgent Medical Device Recall Notice" letter dated March 5, 2009. The letter was addressed to Risk Manager. The letter described the product and problem and requested the consignee to complete, sign and return the certificate acknowledgement. |  | 20 | 20 units | Distribution to one medical facility in MN. | Terminated | May 07, 2011 | 801 | 20 | Computer
Z-1452-2009 | 52027 | No | KGO | Non-powdered surgeon's glove. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Esteem Blue with Neu Thera Surgeons Glove; a sterile, single use powder-free synthetic polyisoprene surgical glove, latex-free, used as an underglove when double-gloving; 1 pair of gloves per sterile pouch, 200 pairs per case; REF2D73PB90 - size 9; Cardinal Health, McGaw Park, IL 60085 USA, Made in Thailand. |  |  | April 24, 2009 | 2009 | Cardinal Health | Increase in the number of customer complaints due to cuff tears. | Device Design | Urgent: Product Recall letters dated 4/24/09 were sent to the direct accounts on 4/24/09 via UPS 2-day mail, to the attention of the Director of Materials Management. The accounts were informed of the increased complaints of cuff tears, and were instructed to stop using them and destroy them. The letter included a list of gloves that could be substituted for the recalled gloves. The accounts were requested to call the appropriate customer service group to arrange for credit for the destroyed product, and to complete and return by fax the enclosed acknowledgement form, indicating the number of gloves destroyed. |  | 29697 | 29,697 cases total | Worldwide Distribution -- USA, Australia, Austria, Belgium, France, Germany, Japan, New Zealand, Switzerland, Spain, UK, Netherlands, Norway, Sweden, and Turkey. | Terminated | May 11, 2010 | 382 | 29697 | Computer
Z-1422-2009 | 51681 | Yes | JOQ | External programmable pacemaker pulse generator. | Cardiovascular | Cardiovascular | 510(k) | EP-4 Computerized Cardiac Stimulator; SJM Part Numbers: 09-1527-0000 / 09-1527-0002; |  |  | March 23, 2009 | 2009 | St. Jude Medical | Non-compliance with IEC Standard 60601-1 Clause 57.6 relative to fusing of AC input line. The neutral line is not fused. The hot line in fused. Clause 57.6 calls for both lines to be fused. | Device Design | All foreign customers were sent Urgent Product Safety Information letters on March 23, 2009. Letters described units affected, the reason for action, and the actions to take to eliminate risk. It is recommended that the information in the letter be circulated to all users, risk managers and the biomedical engineering department responsible for maintaining and inspecting electrical equipment located within their institution. The letter also stated that a SJM representative will be in touch to perform the EP-4 Stimulator fuse modification. Customers are to complete and return the attachment. |  | 694 | 694 total (US, Canada, and OUS) | Worldwide distribution. | Terminated | January 20, 2010 | 303 | 694 | Computer
Z-1510-2009 | 52076 | No | DXY | Implantable pacemaker pulse generator. | Cardiovascular | Cardiovascular | PMA | Medtronic SIGMA 100/200/300 Series IPGs, Single Chamber Pacemaker; models SS103, SS106, SS203, and SS303. |  |  | May 18, 2009 | 2009 | Medtronic Inc. Cardiac Rhythm Disease Managment | One or more bond wire pairs will lift or separate from the bonding terminals on the device hybrid. This may present clinically as loss of rate response, premature battery depletion, loss of telemetry, or no output. | Device Design | Medtronic Sales representative hand delivered an Important Patient Safety Information letter, dated May 2009, to physicians beginning 05/18/09. The letter describes the issues, provides a predicted failure rate for the 3 populations of devices, and provides Patient Management Recommendations. The letter recommended that physicians should advise their patients to seek medical attention immediately if they experience symptoms (e.g., fainting or lightheadedness). It was also recommended that physicians should consider device replacement for patients who are both pacemaker dependent and who have been implanted with a device in the affected subsets. Medtronic will offer supplemental device warranty if the device is not already at elective replacement time. The last recommendation was that physicians should continue routine follow up in accordance with standard practice for those patients who are not pacemaker dependent. The letter also provides Physician and Patient Support. Letters will be mailed to patients with registered devices beginning 05/27/09. Questions should be directed to a local Medtronic Representative or Medtronic Technical Services at 800-505-4636. |  | 9434 | 9434 | Worldwide Distribution -- USA, including Washington D.C., Guam, Puerto Rico, and Virgin Islands; Albania, Algeria, Argentina, Armenia, Aruba, Australia, Austria, Bahamas, Bangladesh, Belarus, Belgium, Bolivia, Bosnia and Herzegovina, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, Estonia, Finland, France, Georgia, Germany, Greece, Guadeloupe, Honduras, Hong Kong, Hungary, India, Iran, Ireland, Israel, Italy, Japan, Jordan, Kenya, Kuwait, Latin America, Latvia, Lebanon, Luxembourg, Martinique, Mauritius, Mexico, Morocco, Nepal, Netherlands, Netherlands Antilles, Norway, Oman, Pakistan, Palestinian Territory Occupied, Peru, Poland, Portugal, Qatar, Reunion, Romania, Saudi Arabia, Senegal, Slovakia (Slovak Republic), Slovenia, South Africa, Spain, Sri Lanka, Sudan, Suriname, Sweden, Switzerland, Trinidad and Tobago, Tunisia, Turkey, United Arab Emirates, United Kingdom, Uruguay, Vatican City State (Holy See), and Venezuela. | Terminated | March 28, 2012 | 1045 | 85378 | Computer
Z-1352-2009 | 51693 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Biomet Dual Offset Broach Handle, Clamping Style, Left, REF 31-555403. For use in primary hip arthroplasty. |  |  | March 31, 2009 | 2009 | Biomet, Inc. | The handle body may fracture or crack at the slot where the locking lever is inserted to clamp down onto the broach. | Device Design | Biomet notified their distributors by recall letter dated 3/31/09, instructed them to pick up the insturments at the hospitals and to notify the hospital personnel of this recall action in writing. |  | 178 | 178 of all products. | Nationwide, Canada and Finland. | Terminated | December 23, 2009 | 267 | 178 | Computer
Z-1224-2009 | 44848 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1 Central Control Monitor. Terumo Cardiovascular Systems, Ann Arbor, MI; Catalog number 802100. |  |  | January 12, 2009 | 2009 | Terumo Cardiovascular Systems Corporation | The monitor may exhibit start-up failure, video/display issues, calibration issues, and partial or complete loss of functionality during use. | Device Design | Consignees were notified via an Urgent Medical Device Correction: Safety Advisory letter dated 1/12/09 advised of the reason for the correction, potential problems, and provided clinical information should the problem(s) arise. Customers are to review the notice, assure that all users are aware of the notice, and confirm receipt of the communication by completing the attached Customer Response Form and faxing it to the number indicated on the form. Questions or concerns should be directed to Terumo CVS Customer Service at 1-800-521-2818. |  | 951 | 951 | Worldwide Distribution -- USA, Australia, Belgium, Canada, Dominican Republic, Egypt, Germany, Honduras, Hong Kong, India, Indonesia, Japan, Jordan, Kuwait, Malaysia, Mexico, Pakistan, Philippines, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Taiwan, Thailand, Turkey, United Arab Emirates, and Viet Nam. | Terminated | July 14, 2009 | 183 | 951 | Computer
Z-1436-2009 | 51982 | No | DXJ | Medical cathode-ray tube display. | Cardiovascular | Cardiovascular | 510(k) | CareVue Chart Release C.O |  |  | May 05, 2009 | 2009 | Philips Healthcare Inc. | Software: A software defect under certain circumstances could lead to the unintended administration of a duplicate dose of a scheduled medication within a single 24 hour period | Software design | Philips Medical notified accounts by letter on 5/6/09 the Urgent - Medical Device Correction notification. While customers await the correction of their devices, customers are given the following instructions:Clinicians should not rely solely on the worklist within CareVue Chart to determine actionable orders for the patient until their software is upgraded. Clinical users should always use the medication administration record (MAR) and the administration record in determining the care their patients receive in addition to the critical care worklist. Philips is currently developing a correction that will be available at no additional charge. A Philips representative will notify customers when the software correction becomes available. If further assistance is needed, customers are to contact the Philips Customer Care Center at 1-800-722-9377. |  | 103 | 103 units | Worldwide Distribution -- USA including states of CA, FL, MA, OH, OK, RI, VT, and WV and countries of Canada, Australia, Austria, Belgium, Denmark, Finland, France, Germany, Italy, Malaysia, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Taiwan, United Arab Emirates, and the United Kingdom. | Terminated | March 06, 2012 | 1036 | 103 | Computer
Z-1401-2009 | 51863 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Panorama Patient Monitoring Network with ambulatory telepacks, and software version 8.5.5, 8.5.6, 8.6.1 or 8.7. |  |  | April 10, 2009 | 2009 | Mindray DS USA, Inc., dba Datascope Patient Monitoring | Software Anomaly: Software anomaly can occur when the Panorama is attempting to play an audio file, either a physiological or technical alarm sound, at which time the system software may cause the system to freeze due to an interaction in the alarm system. | Software design | Urgent Product Field Correction letters, dated March 27, 2009, were sent by certified mail, return requested, and signature required. The letters identified the affected product and the problem. It also stated the adverse effect on patients and that customers' Panorama software will be upgraded to a current software revision. A Datascope Service Representative will contact customers to arrange for the software upgrade of the Panorama(s) in their facility. This will be done at no cost to the customer. |  | 294 | 294 sites | Worldwide Distribution -- USA and Australia. | Terminated | August 03, 2010 | 480 | 294 | Computer
Z-1323-2009 | 51668 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Lactosorb Pectus Stabilizer, model number: 01-3805. Distributed and Manufactured by : Biomet Microfixation, Jacksonville, FL. |  |  | March 20, 2009 | 2009 | Biomet Microfixation, Inc. | This recall is based on the occurrence of known adverse effects identified in the device instructions for use. Bending, fracture, loosening, rubbing and migration of the devices can occur as a result of excessive activity, trauma or load bearing.Implantation of foreign materials can result in an inflammatory response or allergic reaction. | Device Design | An "Urgent Medical Device Recall Notice" was issued via certified letter or email on March 20, 2009 to consignees. Customers were instructed to immediately remove and return the affected in-stock product to Biomet Microfixation, Inc. Customers were also asked to complete and return the enclosed Inventory Reconciliation sheet via fax (904-741-9425) to the firm, and to notify recipients of any recalled products that had been further distributed.For questions about the recall, contact Biomet Microfixation at 1-800-874-7711 or 904-741-4400, extension 468, Monday through Friday, 8am-4pm EST. |  | 2117 | 2,117 | Worldwide Distribution -- US (including DC and states of AL, AZ, CA, CO, FL, GA, HI, IL, IN, KS, KY, LA, MI, MO, MN, NC, NE, NH, NJ, NM, NY, OH, OK, PA, TX, VA and WI) and countries of Argentina, Austria, Belgium, Chile, Colombia, Czech Republic, Denmark, Estonia, Finland, France, Germany, Guatemala, Israel, Italy, Lativa, Netherlands, South Africa, Spain, Sweden, Turkey, United Arab Emirates, United Kingdom and Venezuela. | Terminated | September 11, 2009 | 175 | 2117 | Computer
Z-1418-2009 | 51994 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Surgical Table Headrest, Catalog/Part #0788-300-012, component of VERTIER Surgical Table. Distributed as part of the VERTIER Surgical Tables or individually by Stryker Communications, Flower Mound, TX 75028. |  |  | May 08, 2009 | 2009 | Stryker Communications Corp | Potential for health care provider/hospital staff member to sustain injury to fingers or hands when removing/handling the headrest component of the VERTIER Surgical Table. | Device Design | An "URGENT: Device Recall" letter dated May 11, 2009 was issued to consignees. Consignees were informed that Stryker service representatives will visit each location to affix a caution label identifying the pinch point on each headrest. Additionally, consignees will be provided with an updated surgical table user manual. Direct questions to Stryker Communications Corporation by calling 972-410-7100. |  | 230 | 230 headrests. | Nationwide Distribution and Canada. | Terminated | April 26, 2010 | 353 | 230 | Computer
Z-1646-2009 | 51695 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | Roche ACCU-CHEK Spirit insulin pump. Part numbers 4540492001, 4759702001 and 4759729001. |  |  | April 30, 2009 | 2009 | Disetronic Medical Systems, Inc. | The "up" and "down" buttons, which are used for changing the program in the menu or to administer additional insulin through a bolus delivery, may experience intermittent failure or completely cease to operate. | Device Design | The firm issued a press release on 4/30/09 and sent letters with attached Urgent Medical Device Recall (Field Corrective Action) notification dated 5/4/09 to distributors, health care providers and users. Disetronic will replace pumps within the affected serial numbers (SN02119552 to SN10006093) that have intermittent or complete loss of function of the up and/or down button promptly upon contact from the customer with a new pump that is not affected by the potential button failure. For pumps within the affected serial numbers (SN02119552 to SN10006093) which are functioning properly, no replacement is available/possible at this time. Disetronic will gradually phase in replacement pumps for the remaining pumps within the affected range, as they become available. Customers are to complete the enclosed Customer Reply Card to acknowledge receipt of the notification.On 5/17/11, after being notified of the recall's ineffectiveness (in that only 40% of the pumps had been replaced), the firm re-issued recall notification letters including the aforementioned information to their consignees. |  | 14929 | 14,929 | Worldwide Distribution -- USA and Canada. | Terminated | June 05, 2012 | 1132 | 14929 | Computer
Z-1398-2009 | 51207 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Cell-Dyn Ruby Hematology Analyzer with System Software Version 2.0ML, List Number 08H67-01, Manufactured by Abbott Diagnostics Division, Santa Clara, CA. |  |  | February 12, 2009 | 2009 | Abbott Laboratories | Burn Risk-- when power is cycled to the instrument, there is a possibility that the Mixer Motor Assembly may reach 100 degrees Celsius (212 degrees Fahrenheit ). If the customer touches the motor, burns may occur. | Device Design | A "Product Correction" letter dated February 12, 2009 and "Customer Reply" form were sent to all currently active CELL-DYN Ruby customers using Software Version 2.0 ML. The letter described the affected product, necessary actions and contact information for customers. Customers were instructed to acknowledge receipt of the "Product Correction" letter by completing the enclosed Customer Reply form and return via fax to 1-800-777-0051. A product information letter is being added to the accessory kits shipped with new instruments and version 2.0 ML System Software upgrade kits. For questions regarding this information, contact Abbott Laboratories Customer Support by calling 1-877-422-2688. Customers outside the US, please contact your local hematology customer support representative. |  | 324 | 324 units distributed | Worldwide Distribution -- US, Argentina, Canada, Columbia, Germany, Jamaica, Japan, Mexico, Puerto Rico, Singapore, South Korea and Venezuela. | Terminated | December 27, 2010 | 683 | 324 | Computer
Z-1195-2009 | 51193 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Innova 2121-IQ/3131-IQ Biplane Cardiovascular Imaging System. The Innova Biplane systems are the first and only digital flat panel biplane systems with a full sized lateral plane to cover lateral anatomy without requiring multiple contrast injections and radiation exposures. The product is indicated for use in generating fluoroscopic images of human anatomy for diagnostic and interventional cardiac angiography procedures. These systems can be operated in a mobile or fixed site environment. |  |  | November 21, 2008 | 2008 | GE Medical Systems, LLC | GE Healthcare has become aware of a potential image storage failure on Innova 2121 IQ/ 3131 IQ cardiovascular imaging systems that may impact patient safety. This intermittent communication failure on the image storage chain may result in the following situations: Record abort, Record not available, Fluorostore function not available, Review not available, Boot/Reset failure. Fluoroscopy imag | Device Design | A GE Healthcare "Urgent Medical Device Correction" letter dated February 2, 2009 was sent to consignees via mail. The letter was directed to Hospital Administrator/Risk Manger, Managers of Radiology/Cardiology and Radiologists/Cardiologists addressing the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information.Further questions should be addressed to GE Healthcare at 1-800-437-1171. |  | 85 | 85 | Worldwide Distribution -- USA including states of CA, CO, FL, HI, IA, IL, IN, KS, MD, MA, MN, MO, NY, OH, OR, PA, TN, TX, WA, WV, WI and VA and countries of the United Kingdom, Romania, Japan, Taiwan, Bahrain, France, Malaysia, India, Canada, Chile, Indonesia, Ireland and Italy. | Terminated | November 21, 2013 | 1826 | 85 | Computer
Z-1255-2009 | 51071 | equipment, laboratory, general purpose, labeled or promoted for a specific medical use | LXG | General purpose laboratory equipment labeled or promoted for a specific medical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Roche MagNA Pure LC Workstation; Product Numbers: 12236931001, 12236931444, 12236931692 and 03670325001. Robotic workstation for automated isolation of nucleic acids from different kinds of crude sample material such as whole blood, serum, blood cells, culture cells, tissue, bacteria and fungi. |  |  | February 09, 2009 | 2009 | Roche Diagnostics Corp. | Protruding screw threads may cut the user when cleaning the interior of the instrument. | Device Design | In a letter titled "Urgent Medical Device Correction" dated February 9, 2009, consignees were notified of the protruding screws and instructed to use caution when cleaning the instrument until a protective plate is installed by the firm's field representative.Further questions should be addressed to Roche Molecular Diagnostics Technical Support at 1-800-526-1247. |  | 371 | 371 | Nationwide Distribution including Puerto Rico and states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, HA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NV, NY, OH, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI and WV. | Terminated | December 07, 2009 | 301 | 371 | Computer
Z-1397-2009 | 50201 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Centricity TriWin Laboratory Information System; GE Healthcare Integrated IT Solutions, Barrington, IL 60010 |  |  | November 17, 2008 | 2008 | GE Healthcare Integrated IT Solutions | A software defect in the Centricity Lab (TriWin) product may incorrectly link patient information to a wrong report. Also, when viewing results in Patient Viewer, and a coagulation result is higher or lower than the linearity limit, then all of the results following the higher or lower result will show the same result as the first flagged result. | Software design | GE Healthcare notified their Centricity TriWin Laboratory customer via letter dated 11/17/08 of the software anomalies resulting in patient safety issues involving the patient information incorrectly being linked to a wrong report, and when viewing results in Patient Viewer, and a coagulation result is outside of a linearity limit, then all of the results following the out of linearity result will show the same result as the first flagged result. The user was informed that a software patch has been prepared to resolve the issues and will be installed on their system. Any questions were directed to GE Technical Support at 888-778-3375 or 905-479-1249, select option 1. The accounts was requested to complete the enclosed customer reply form, supplying the workstation software or build number, and fax it back to 905-305-0046 to schedule the software upgrade. |  | 26 | 26 units | Worldwide Distribution -- USA and country of Canada. | Terminated | September 11, 2010 | 663 | 26 | Computer
Z-1326-2009 | 50740 | No | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Smiths Medical BCI(R) Advisor(R) Vital Signs Monitor, model 9200, catalog numbers: 920654225, 920674225, 920674235, 925454220, 925454235, 925454320, 925454325, 925454330, 925454335, 925457325, 925474325, 925654225, 925654325, 925754325, 925754335, and 92M774325. |  |  | December 22, 2008 | 2008 | Smiths Medical PM, Inc. | Recent improvements were made to the Advisor case plastic resulting in the inserts being deeply embedded into the plastic. This affects the ability of the monitor to be adequately secured on the wall mount bracket or roll stand bracket, resulting in the monitor falling off. There have been no adverse events or injuries reported. | Device Design | An Urgent Product Safety Information, Voluntary Recall Notification, dated 12/22/08, was sent to affected consignees. The notification contained a Safety Action Bulletin which provided actions to be taken. The Safety Action Bulletin described the issue, requested customer to contact Smiths Medical to ensure proper mounting hardware is utilized for those consignees using a roll stand or wall mount. If the consignee further distributed the product, they were asked to forward on the Safety Action Bulletin to the end user. Page 3 of the Safety Action Bulletin was asked to be completed and returned to Smiths Medical PM, Inc.A second "Urgent Product safety Information Voluntary Recall Notification" letter dated January 16, 2009 was sent to consignees. The second letter included a risk statement associated with the recalled product. |  | 143 | 143 monitors (17 US, 126 OUS) | Worldwide Distribution -- USA including states of FL, GA, NJ, NY, and TN and countries of Argentina, Brazil, Ecuador, Greece, Peru, Romania, Russia, Singapore, Slovenia, Trinidad and Tobago, United Kingdom, and Uruguay. | Terminated | December 17, 2011 | 1090 | 143 | Computer
Z-1263-2009 | 51447 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Carina Home Ventilator. |  |  | January 21, 2009 | 2009 | Draeger Medical, Inc. | The device may not generate a pilot line disconnect alarm if the pilot line becomes disconnected, which may potentially result in a substantial reduction in tidal volume delivery. | Software design | An "Urgent - Medical Device Recall" Letter dated January 2009 was issued to consignees via certified mail. The letter described the issue, specific conditions, background information, and precautions for customers. An additional letter dated January 22, 2009 was included with the "Urgent - Medical Device Recall" letter which instructed consignees to inform an enduser of the contents of the notification letter because some of the consignees are home care device/service providers. When the new hardware/software solution is available, affected devices will be updated free of charge. Direct questions about the notification letter to Drager Medical, Inc. by calling 1-800-543-5047 (press "1" at the prompt and then press "2349#"). Direct questions regarding updating your device when the new software is available, contact Drager Service Technical Support at 1-800-543-5047 (press "4" at the prompt). |  | 40 | 40 units | Nationwide Distribution -- states of CA, FL, GA, HI, IL, LA, MA, MN, OH, and TX. | Terminated | January 05, 2010 | 349 | 40 | Computer
Z-1252-2009 | 51473 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Servo Ventilator 300 (Model SV300) and |  |  | March 23, 2009 | 2009 | MAQUET Inc. | The expiratory valve in the SV 300 ventilator may stay in the open position if the printed circuit board fails. | Device Design | An "Urgent: Device Correction" letter dated March 23, 2009 was sent by Federal Express to customers. The letter described the affected product, problem, indications, potential hazard and actions for users. Customers were instructed to complete the "Device Correction -- Customer Fax Response" form and return via fax (732-667-1903) to Maquet Inc. RA/QA. Direct questions to your local Maquet Inc. representative. |  | 7847 | 7,847 units | Nationwide Distribution. | Terminated | November 16, 2009 | 238 | 7847 | Computer
Z-1233-2009 | 51494 | No | MUE | Full-field digital mammography system. | Radiology | Radiology | 510(k) | Siemens Mammomat Novation DR, full field digital mammography system. Model Number 06646900. |  |  | January 23, 2009 | 2009 | Siemens Medical Solutions USA, Inc | Image may appear dark, requiring repeat image acquisition. | Software design | A Customer Safety Advisory Notice dated December 16, 2008 was issued to affected customers via Update Instructions SP040/08/S. The letter described the potential issue and provided instructions to avoid its occurrence. A software update to correct the issue is currently being developed and will be installed on affected systems when it becomes available. |  | 497 | 497 units | Nationwide Distribution. | Terminated | March 19, 2013 | 1516 | 497 | Computer
Z-1060-2009 | 50856 | No | KNS | Endoscopic electrosurgical unit and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | PROSTIVA RF Model 8929 Hand Piece is the delivery system component of the PROSTIVA RF Therapy System. |  |  | June 05, 2008 | 2008 | Medtronic Neuromodulation | Reported incidents of needle retraction failure when using the PROSTIVA RF Model 8929 Hand Piece. The Medtronic Prostiva RF Model 8929 Hand Piece is the delivery system component of the PROSTIVA RF Therapy System. | Device Design | A Medtronic "Urgent Medical Device Correction" letter addressed to "Healthcare Provider" was sent August 6, 2008. The letter described the issue, product involved and the Prostiva hand piece Disassembly Procedure that should be followed in the event the user is unable to retract the needles during a procedure. The letter also requested the return of the "Medtronic Reply Form/receipt Confirmation" form. |  | 40168 | 40,168 | Worldwide Distribution -- AL, AZ, AR, CA, CO, CT, DE, DC, FL, GA, HI, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, NE, NV, NJ, NM, NY, NC, ND, OH, OK, OR, PA, PR, RI, SC, SD, IN, TX, UT, VA, VT, WA, WI, WY, and WV and countries of Australia, New Zealand, China, Austria, Denmark, Finland, France, Germany, Italy, Kuwait, Lebanon, Norway, Qatar, Russian Federation, Spain, Sweden, Switzerland, and United Kingdom. | Terminated | May 17, 2012 | 1442 | 40168 | Computer
Z-1005-2009 | 50142 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Welch Allyn AED 10 Automated External Defibrillator |  |  | December 15, 2008 | 2008 | Welch Allyn Protocol, Inc | Potential for device to shut down prematurely under certain circumstances. | Device Design | On 4/13/07 the firm released Service Bulletin part no. 810-2614-00, SB - AED 10 S/W V020600 DOWNGRADE to V020500. That mandatory bulletin, for units with software version 2.06.00 and up indicates that all returning AED 10 units shall have their software version downgraded to 2.05.00. |  | 89 | 89 units | Worldwide distribution: USA, Belgium, Canada, France, Ireland, Israel, Mexico, and United Kingdom. | Terminated | February 15, 2012 | 1157 | 89 | Computer
Z-0955-2009 | 50796 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | Meter Remote for OneTouch Ping Glucose Management System |  |  | October 27, 2008 | 2008 | Animas Corporation | incorrect dosing recommendation | Software design | The recalling firm issued Urgent: Medical Device Correction letters dated 10/27/08 to their customers/user to inform them of the problem. The letter stated that the firm was in the process of correcting the issue and that a new OneTouch Ping Meter-Remote would be provided as soon as one became available. In the interim, customers could still safely use their OneTouch Ping System provided that they use the ezCarb calculator on the meter-remote when it is paired and communicating with the pump. The letter listed guidelines to ensure that the device is properly communicating with the pump. The letter stated that customers should follow the directions in the Owner's Booklet to troubleshoot any communication issues. Questions regarding this letter were to be directed to Animas at 866-423-4087. A patient acknowledgment letter dated November 18, 2008 was also issued. Another Urgent: Medical Device Correction letter, dated November 2008, was issued along with the replacement device. This letter asked customers to download their current meter-remote data, set up their replacement meter-remote, and then return their current meter-remote. |  | 5093 | 5093 units | Nationwide Distribution --- including states of KY, UT, FL, OH, MA, and TX. | Terminated | December 15, 2009 | 414 | 5093 | Computer
Z-0836-2009 | 50538 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Centricity PACS RA1000 Workstation (for diagnostic image analysis); GE Healthcare Integrated IT Solutions, Barrington, IL 60010. |  |  | December 12, 2008 | 2008 | GE Healthcare Integrated IT Solutions | There is a potential patient safety issue involving the use of the F12 keyboard Approve &amp; Complete function. The issue may cause erroneous information to be added to a report when trying to approve and complete it. | Software design | GE Healthcare Integrated IT Solutions notified their Centricity PACS RA1000 Workstation software customers via letter (Urgent Medical Device Correction) dated 12/12/08 of the software anomaly resulting in a patient safety issue involving the use of the F12 keyboard Approve &amp; Complete function. The issue may cause erroneous information to be added to a report when trying to approve and complete it. The letter provided the users with safety instructions to follow until the software is updated. Any questions were directed to the Remote Online Center, Field Engineer or the GE Customer Care Center at 800-437-1171. The accounts were requested to complete the enclosed customer reply form, supplying the workstation software or build number, and fax the completed form to 414-918-4653. |  | 138 | 138 units | Worldwide Distribution including USA and countries of Canada, China, Denmark, Germany, Hong Kong, India, Ireland, Italy, Japan, Korea, Lebanon, Mexico, Netherlands, Saudi Arabia, Spain, Taiwan and United Kingdom | Terminated | May 18, 2010 | 522 | 138 | Computer
Z-0473-2009 | 49539 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Innova 2100IQ Versatile Cardiac and Vascular Imaging System, GE Healthcare, Waukesha, WI 53188 ( The system consists of an a monoplane positioner, a vascular or cardiac table, an X-ray system and a digital detector) The product is indicated for use in generating fluoroscopic images of human anatomy for vascular angiography diagnostic and interventional procedures and optionally, rotational imaging procedures. It is also intended for generating fluoroscopic images of human anatomy for cardiology, diagnostic and interventional procedures. |  |  | January 30, 2008 | 2008 | GE Medical Systems, LLC | Software error, computer: GE Healthcare has became aware of an issue with Error Management affecting the Innova 2100, 3100/3100 IQ &amp; 4100/4100IQ. It was reported that when the foot switch was pressed, the live monitor screen was blank and it looked like there were no X-rays being emitted. But the message "Acquisition in progress" was displayed on the DL monitor, even after the foot switch was r | Software design | Consignees sent a "GE Urgent Medical Device Correction" letter dated February 28, 2008. The letter was addressed to Hospital Administrators/Hospital Risk Management Department/Managers of Radiology-Cardiology/ Radiologists-Cardiologists. The letter described the affected products, safety Issue, short term recommendations, long term solution and contact Information. The first batch of letters was received by customers on 5/28/08. For further information, please contact GE Healthcare at 262-544-3894 |  | 427 | 427 | USA including states of AL, AK, AZ, AR, CA, CO, CT, DE, FL, GA, HI, ID, IL IN, KS, KY, LA, ME, MD, MA, MI,MN, MS,MO, MT, NE, NV, NH, NJ, NM, NY,NC, ND, OH, OK OR, PA, RI, SC, SD, TN, TX UT,VT, VA, WA,WV, and WI; the United States territory of Puerto Rico; and countries of ALGERIA, AUSTRALIA, BELARUS, BELGIUM, BOSNIA, BRAZIL, BULGARIA, CANADA, CHILE, CHINA, COLOMBIA, CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, ,FINLAND, FRANCE, GEORGIA, GERMANY, GREECE, HONG KONG, HUNGARY, INDIA, INDONESIA, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN,KOREA, KUWAIT, LATVIA, LEBANON, LITHUANIA, MALAYSIA, MEXICO, MOROCCO, NETHERLANDS, NEW ZEALAND,NORWAY, PANAMA, POLAND, PORTUGAL RUSSIAN FEDERATION SAUDI AR ARABIA, SINGAPORE, SLOVAKIA, SPAIN, SUDAN, SWEDEN SWITZERLAND, SYRIA, TAIWAN, TURKEY UNITED KINGDOM, and VENEZUELA. | Terminated | March 07, 2016 | 2959 | 1127 | Computer
Z-0629-2009 | 50207 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Cyberknife Robotic Radiosurgery System. A radiation therapy device, MultiPlan (MP) Treatment Planning Software and Iris Variable Aperture Collimator, Software version 3.0. |  |  | November 14, 2008 | 2008 | Accuray Inc | System may use random incorrect data to calculate dose. Resulting dose calculation error can exceed 100% of correct dose which may lead to serious patient injury. | Software design | An Urgent Advisory Notification letter dated November 14, 2008 was issued to all sites via Fed Ex. The letter identifies the issue and required actions. The user was instructed not to modify current extrapolation depth settings and to contact Accuray Incorporated Customer Support prior to making any changes.Please direct any questions or concerns regarding this issue to Accuray Incorporated Customer Support at 1-877-668-8667 (USA) or 1-408-716-4700 (outside of USA) or customersupport@accuray.com. |  | 4 | 4 units | Nationwide Distribution. Product was distributed to 4 medical facilities in the US. | Terminated | December 02, 2010 | 748 | 4 | Computer
Z-0855-2009 | 50143 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Eclipse 8.5 (build 8.2.2X) systems with distributed Calculation Framework (DCF) 8.2.22, model number H48, manufactured by Varian Medical Systems, Inc., Palo Alto, CA. |  |  | October 02, 2008 | 2008 | Varian Medical Systems Oncology Systems | When using the multiple static segment option, the device may alter dose delivery. | Software design | An Urgent Medical Device Correction Letter was distributed to all affected users, with a description of the problem and user corrective action steps on October 2, 2008. The Letter will also be distributed to the Varian Sales, Marketing, and Service organizations, informing them of the issue. The firm recommends that customers DO NOT USE the Multiple Static Segment (MSS) option in Varian Leaf Motion Calculator version 8.2.22. The firm also recommends setting the default delivery method for Varian Leaf Motion Calculator to Sliding Window. When using Eclipse 8.5, always follow the Safety Instructions found in chapter 2 of "Eclipse External Beam Planning Instruction for Use", P/N B501357R01A, March 2008. If customers believe a patient treatment has been affected by this issue, they should notify a Varian Service or Support person as soon as possible. |  | 64 | 64 units | Worldwide Distribution --- including USA and countries of Switzerland and the Netherlands. | Terminated | November 27, 2010 | 786 | 64 | Computer
Z-0620-2009 | 50472 | No | FWZ | Manual operating table and accessories and manual operating chair and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Foot switch, Part Number 1009.81. Maquet GmbH &amp; Co. KG, 76437 Rastatt, Germany. |  |  | December 08, 2008 | 2008 | MAQUET Inc. | A bearing point of the pedal of the foot switch may come loose. With continued use of the foot switch, one or both springs beneath the pedal can fall out of their mounting, and it is possible the pedal comes to rest on the switch causing an unintended motion of the operating table during patient use. | Device Design | An Urgent Device Correction letter (with an enclosed Customer Response form) was sent via Federal Express on Monday, December 8, 2008 to all customers who had received tables. Customers were notified in the letter of the problem and actions to be taken including not to use any affected foot pedal for safety reasons. In addition, customers were instructed to complete and return the Customer Response Form via fax (732-667-1903) to Maquet Inc. Please call USA Maquet, Inc. at toll free number 1-888-627-8383 #1 (then select the option for Surgical Workplace products) for questions. |  | 587 | 587 units total (54 units in US) | Worldwide Distribution. | Terminated | June 18, 2012 | 1288 | 587 | Computer
Z-0538-2009 | 49112 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Depuy DELTA Xtend, Standard Humeral PE Cup, Dia 42/+9 STD, sterile, Depuy France; REF 130742209. |  |  | July 10, 2008 | 2008 | Depuy Orthopaedics, Inc. | Compressed air may become trapped between the cup and the stem, preventing the tapered walls from engaging and causing the cup to feel unstable. | Device Design | A letter, "Urgent Device Correction" dated July 8, 2008 was sent via certified mail to implanting surgeons advising them of the issues and guidance on how to mitigate the problem. A communication was also sent to all sales representatives and territory offices on July 15, 2008. For questions, please contact Depuy Orthopaedics, Inc. Customer Service at 1-800-473-3789. |  | 1521 | 1,521 of all devices. | Nationwide Distribution. | Terminated | October 28, 2009 | 475 | 1521 | Computer
Z-0646-2009 | 50480 | No | KPQ | Radiation therapy simulation system. | Radiology | Radiology | 510(k) | Drive Belt for Radiation Therapy Simulation systems SIMVIEW 3000 and SIMVIEW NT with CT option or IDI. |  |  | November 06, 2008 | 2008 | Siemens Medical Solutions USA, Inc | Excessive wear on drive belts used for Image Intensifier and/or X-ray head drive mechanisms may result in uncontrolled movement, which may cause injury. | Device Design | Firm has reissued the Urgent Medical Device Customer Letter regarding the belt replacement and a reminder message to replace the belts yearly during the annual Preventative Maintenance. The letter requests that customers follow the guidelines, which consist of making any required SAD adjustments utilizing the Image Intensifier drive, the X-Ray Head Drive, or both; being vigilant of any change in the type or level of noise from either of the drive belt assemblies while they are in use as this may be an indication of wear; discontinuing clinical operation if there is any doubt, and contacting their local Siemens Service organization to inspect the belts. Currently, the firm is distributing the current belt kit to replace the entire installed base of active SIMVIEW Simulators (3000 and NT models). |  | 223 | 223 units | Worldwide Distribution | Terminated | December 09, 2010 | 763 | 223 | Computer
Z-0581-2009 | 49608 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension EXL Clinical Chemistry System used with software versions 8.5.1 and 8.5.1SP3. |  |  | September 30, 2008 | 2008 | Dade Behring, Inc. | Error messages, "Wash Aspirator Failure" and "Wash station unable to recover" are not generated when one or both HM wash probes are totally occluded. | Software design | The recalling firm notified consignees by telephone to inform them of the problem. A follow-up letter, "Urgent Field Safety Notice" dated 9/30/08 was issued with information describing the problem, the possible impact of an occluded wash probe on assay results, and emphasizing regular performance of routine wash probe maintenance as detailed in the operator guide to prevent wash probe occlusion. For questions, please contact the Siemens Technical Solution Center at 800-441-9250. |  | 57 | 57 units | Nationwide Distribution including states of AR, CA, CO, IA, ID, IL, IN, LA, MD, MI, MO, MS, NC, NE, OH, OK, PA, SD, TN, TX, VA, and WI. | Terminated | May 05, 2009 | 217 | 57 | Computer
Z-0129-2009 | 48188 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Cytomics FC 500 Flow Cytometry System with CXP Software Version 2.2, Part Number: 626553. This product is used for the qualitative and quantitative measurement of biological and physical properties of cells and other particles. These properties are measured when the cells pass through one or two laser beams in single-file. |  |  | October 11, 2007 | 2007 | Beckman Coulter Inc | Software error: This field action was initiated after Beckman Coulter determined that the CXP software does not refresh the dataset automatically under certain circumstances. The User documentation does not clearly state that the user should refresh the dataset after changes are made. | Software design | Software error: A Product corrective Action (PCA) letter was mailed via US Postal Service on October 17, 2007 to all Cytomics FC 500 Flow Cytometry System with CXP Software accounts. This letter was entitled "URGENT: PRODUCT CORRECTIVE ACTION." Customers were informed that the user documentation does not clearly state that the user should refresh the dataset after changes are made. The letters provide the customers with an explanation and action to be taken whenever a protocol containing data is modified by making changes to cytosettings, regions or gates.Customer were instructed to review all data plots and results before reporting results out. any uncharacteristic patterns in the data or unexpected results must be verified by running Quality Control and if necessary, preparing new samples.In addition, customer were instructed to please share this information with their laboratory staff and retain this notification as part of their laboratory Quality System documentation. Also they were asked to complete and return the an attached Customer Response Form within ten days.Customers with questions or concerns regarding the notification, could call Beckman Coulter Customer Service at 1-800-526-7694 in the United States or Canada. Outside the US and Canada, they were to contact their local Beckman Coulter Representative. |  | 1188 | 594 Worldwide: 518 in the United States and 76 in Canada | Globally - United States and Canada | Terminated | January 13, 2009 | 460 | 1188 | Computer
Z-0870-2009 | 50540 | No | LMD | Medical image communications device. | Radiology | Radiology | 510(K) Exempt | Centricity Enterprise Web 3.0 software; the software allows physicians convenient, quick access to medical images and related data they need via a web browser, password and a standard PC; GE Healthcare Integrated IT Solutions, Barrington, IL 60010. |  |  | December 12, 2008 | 2008 | GE Healthcare Integrated IT Solutions | Software error: There are two potential safety situations with Centricity Enterprise Web software - a forced log off may occur while using the Centricity Enterprise Web during an open session, and Centricity Enterprise Web does not utilize calibration performed on image measurements in Centricity RA1000 Workstation. | Software design | GE Healthcare Integrated IT Solutions notified their Centricity Enterprise Web software customers via letter (Urgent Medical Device Correction) dated 12/12/08 of the software anomalies resulting in patient safety issues involving a forced log off during an open session while using the Centricity Enterprise Web and measurements displayed in Centricity Enterprise Web will use the calibration provided by the original scanning modality. Centricity Enterprise Web does not utilize calibration performed on image measurements in Centricity RA1000 Workstation. The letter provided the users with safety instructions to follow until the software is updated. Any questions were directed to the Remote Online Center, Field Engineer or the GE Customer Care Center at 800-437-1171. The accounts were requested to complete the enclosed customer reply form, supplying the workstation software or build number, and fax the completed form to 414-918-4653. For further information, please contact GE Customer Care Center by telephone at 800-437-1171. |  | 742 | 742 units | Worldwide Distribution including USA and countries of Australia, Austria, Bahrain, Belgium, Canada, China, Denmark, Egypt, France, Germany, Hong Kong, India, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Korea, Kuwait, Malaysia, Malta, Martinique, Netherlands, New Zealand, Nigeria, Norway, Portugal, Reunion, Russia, Saudi Arabia, Slovenia, Singapore, South Africa, Spain, Sudan, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom and Venezuela. | Terminated | May 18, 2010 | 522 | 742 | Computer
Z-0643-2009 | 50501 | No | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | ACIST CVi Angiographic Injection System with Interface Cable for Toshiba Infinix-i cardiovascular X-ray system, Part Number 700318-002. The ACIST injection system is an angiographic injection system used in interventional cardiology, radiology, and vascular surgical procedures. The ACIST injection system supplies radiopaque contrast media to a catheter at a user-determined variable flow rate and volume which can be instantaneously and continuously varied. Manufactured by: ACIST Medical Systems, Inc., 7905 Fuller Road, Eden Prairie, MN 55344, USA. The ACIST CVi injection system is able to synchronize with certain X-ray imaging systems from Siemens, Toshiba, GE and Philips. |  |  | November 10, 2008 | 2008 | Acist Medical Systems | Multiple Injections may result. It has been determined that when using the ACIST CVi Angiographic Injection system to synchronize with the Toshiba Infinix -i cardiovascular x-ray system there is an injection scenario that may result in multiple injections of contract media when one injection was intended. | Software design | ACIST is distributing an Urgent Medical Device Correction letter, dated November 10, 2008, to affected customers. The letter describes the issue and the scenario when the issue to occur, instructs them to discontinue use of the synchronized mode between the ACIST Angiographic Injection System model CVi and the Toshiba Infinix-i X-ray fluoroscopic system by disconnecting the interface cable immediately. The letter states that ACIST is working with Toshiba to resolve this issue. A response form is asked to be returned either by fax, e-mail or mail. If there are any questions, they should be directed to either ACIST at 952-995-9300 or local Sales Reps or Clinical Applications Specialist. |  | 18 | 18 cables | Worldwide Distribution --- USA including states of AR, AZ, FL, and NY and countries of France, Germany, Italy, Japan, and Spain. | Terminated | May 07, 2011 | 908 | 18 | Computer
Z-0647-2009 | 50500 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | CT syngo Dual Energy Viewer on the syngo Multimodality Work Place (MMWP) and on the CT workplace, Model number 10094833. |  |  | November 13, 2008 | 2008 | Siemens Medical Solutions USA, Inc | Indicated orientation does not match actual orientation of the patient. | Software design | The recalling firm issued a Customer Safety Advisory dated 10/8/08 to affected customers via Update Instructions CT052/08/S. The letter informs customers of the potential issue and provides instructions to avoid its occurrence. Customers should consider "no saving of calculated Dual Energy data sets from non-supine and mirror-option scans via "Save Sets CT"-button." The firm also highly recommends for customers "to check the correct display of "left", "anterior" and "posterior"markings, e.g. by checking the position of the spine with respect to the posterior marking in the image." The error should be fixed within certain upgrade and installation activities on the affected software. |  | 374 | 374 units | Nationwide Distribution | Terminated | March 19, 2013 | 1587 | 374 | Computer
Z-0608-2009 | 50167 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Hitachi CXR4 Computed Tomography Scanner. This product is an x-ray imaging device that produces cross-sectional images of he body at different angles. The system reconstructs, processes, displays, and stores the collected images. The device output can provide an aid to diagnosis when used by a qualified physician and is intended for general purpose CT applications. |  |  | November 05, 2008 | 2008 | Hitachi Medical Systems America Inc | Software error: The device has a software error that can result in a condition where the displayed image orientation does not match the display orientation markers (right-left reversal). | Software design | On 11/5/2008 Hitachi Medical Systems sent recall notification letters entitled "URGENT - DEVICE RECALL CORRECTION" to its customers informing them of the problem with the software. They stated that a Hitachi service representative will contact each customer to schedule a software upgrade. For further information, please contact Hitachi Medical Systems America Inc. by telephone at 330-425-1313. |  | 42 | 42 | Nationwide Distribution --- including states of AZ, CA, DE, FL, ID, IL, IN, KS, KY, LA, MA, MO, NJ, NY, OH, OK, PA, PR, SC, TX, VA, and WI. | Terminated | September 21, 2012 | 1416 | 42 | Computer
Z-0601-2009 | 50144 | No | FNM | Alternating pressure air flotation mattress. | General Hospital | General Hospital | 510(K) Exempt | Synergy Dynamic, Powered Patient Rotation Bed, Model number 5607. |  |  | October 09, 2008 | 2008 | Hill-Rom Manufacturing, Inc. | These devices have the potential to overheat. | Device Design | All end users (those who either purchased the device and any customer that has rented) were notified by an Urgent Field Safety Notice on/about 10/31/2008. The letter informed the customers of the safety risk and list steps they can take to mitigate the risk until a Hill-Rom representative is able to modify the device. For the devices that were sold, the consignee of each device was notified by letter for the safety risk and what they can do to mitigate the risk until a Hill-Rom representative is able to modify the device. For the devices that are rented, an internal notification was sent via e-mail to Service Centers (Because the device is owned by the manufacturer, Hill-Rom, each Service Center in possession of the device will be the Consignee). Each Service Center will perform the correction on the units affected. |  | 2083 | 2,083 units (Total all four units) | Nationwide Distribution. | Terminated | December 10, 2009 | 427 | 2083 | Computer
Z-0455-2009 | 49010 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Cytomics FC 500 Flow Cytometry System with CXP Software, Part Numbers: 626553, Version 2.2 |  |  | August 01, 2007 | 2007 | Beckman Coulter Inc | Software error: The recall was initiated after Beckman Coulter identified that the CXP User documentation does not adequately characterize the use of the Live Gate feature.The primary use of a Live Gate is for the international removal of unwanted events, such as debris. The data outside of the Live Gate will not be stored to the final listmode file, this reducing the data set size.The softwa | Software design | The customer notifications were initiated by Beckman Coulter on July 23, 2007 when a Product Corrective Action (PCA) was mailed via US Postal Service to all affected customers. The PCA informed the customers that the user documentation does not adequately characterize the use of the Live Gate feature. The letter provides the customer with explanation of Live Gate and actions to be taken. For further information, please contact Beckman Coulter, Inc. by telephone at 714-993-5321. |  | 565 | 565 | Globally to the United States and Canada. | Terminated | January 13, 2009 | 531 | 565 | Computer
Z-0448-2009 | 48734 | No | LPS | Urinary homocystine (nonquantitative) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Homocysteine Reagent Test Kit. The Homocysteine test kits were manufactured under 2 brand names. The first brand name Catch, the label reads "CATCH Catch Inc., Bothell, WA 98011 HOMOCYSTEINE REAGENT KIT, Catalog Number A11167, 100 Test Kit, Manufactured at Intersect Systems, Inc., Longview, WA 98632". The kit is made up of 5 individual portions. These include: Homocysteine Reagent R1 HCY A (28mL), Homocysteine Reagent R2 HCY R2 (13.5 mL), Homocysteine HCY C (2.8 mL), Homocysteine 0.0 umol/L (2mL) and Homocysteine 27.0 umol/L (2mL). The second brand name Equal the label reads "EQual DIAGNOSTICS FOR IN VITRO DIAGNOSTIC USE ONLY HOMOCYSTEINE 3-Part Kit w/calibrators, Catalog Number E02-058-03, Equal Diagnostic, Exton, PA 19341". The kit is made up of 5 individual portions. These include: Homocysteine Reagent R1, Homocysteine Reagent R2, Homocysteine Reagent R3, Homocysteine Calibrator 0.0 umol/L and Homocysteine Calibrator 27.0 umol/L. All individual kits are packaged into individual boxes and then packaged into 1 box. Homocysteine Reagent is intended for in vitro quantitative determination of total homocysteine in serum and plasma. |  |  | January 18, 2007 | 2007 | Catch Incorporated | A high homocysteine level can be reported as falsely low. | Device Design | On 01/18/07 an "URGENT: PRODUCT CORRECTIVE ACTION" letter was sent to the consignees. The letter states the circumstances in which a high homocysteine rate can result in a false low homocysteine rate, a work around for the problem which a flag must be programmed into the user defined homocysteine test parameters and instructions to advise their customers about the situation was addressed in the letter so the consignees can take the recommended corrective action. A confirmation sheet was included with the letter to be signed by a representative of each distributor and returned to Catch Inc. Contact Catch Inc. at 1.425.402.8960 for assistance. |  | 2699 | 2,699 kits | CA and PA. | Terminated | September 28, 2010 | 1349 | 2699 | Computer
Z-0321-2009 | 50077 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Somatom Sensation 64 X-Ray System, Model number 8377520. This is a computed tomography x-ray system which is intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data from the same axial plane taken at different angles. |  |  | September 03, 2008 | 2008 | Siemens Medical Solutions USA, Inc | software error - does not allow protocol to properly synchronize with ECG triggers | Software design | The recalling firm issued a Customer Safety Advisory Notice dated 9/3/08 to affected customers via Update Instructions CT045/08/S. This letter informs customers of the potential issued and provides instruction to avoid its occurrence. For additional information, contact Siemens Medical Solutions USA, Inc. at 1-610-448-4500. |  | 9 | 9 units | Nationwide | Terminated | September 23, 2009 | 385 | 177 | Computer
Z-0342-2009 | 50083 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE OEC 9900 Elite GSP fluoroscopic x-ray system GE Healthcare, Surgery, Salt Lake City, UT. The system is intended to provide an x-ray display of human anatomy. |  |  | August 29, 2008 | 2008 | OEC Medical Systems, Inc | Under certain conditions, x-ray system may experience operational impairment. | Device Design | Consignees were notified by letter on 08/29/2008. Interim solutions were provided for proper start-up procedures and for avoidance of mechanical impact to the C-arm. A start Up caution label was included with the notification. Further information available at 800-874-7378. |  | 448 | 448 units | Nationwide, including AK, KY, NC, VA WA. and WI | Terminated | April 24, 2012 | 1334 | 1054 | Computer
Z-0339-2009 | 49607 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | TomoTherapy HI-ART Systems with 3.1.2, 3.1.3 or 3.2.1 software. The affected applications include Planning Station, Planned Adaptive, Data Management System, and TomoPortal. The product is intended for use in radiation therapy. |  |  | September 18, 2008 | 2008 | TomoTherapy Incorporated | TomoTherapy Inc. discovered an issue on the Hi-Art system during the course of ongoing testing. It has been identified that treatment plans involving very small structures (volume &lt;2cc), reading in a high dose gradient region, may exhibit inaccuracies with the DVH curve and the Dose Statistics for the small structures. This error can only happen if planning was performed using very small voxels | Software design | Consignees were sent on 9/18/08 a "TomoTherapy Incorporated Urgent Medical Device Correction Field Safety Notice" letter dated September 17, 2008. The letter described the issue, product affected, recommended actions and resolution. Questions may be directed to TomoTherapy Inc. Customer Service Center at 1-868-368-4807. |  | 165 | 165 | AKALARAZCACOCTDCFLHIIAILINKYLAMAMDMIMNMONENJNMNVNYOHOKORPAPRSCSDTNTXVAWAWI andWYBelgiumCanadaSwitzerlandGermanySpainFranceGreat BritainIndiaItalyJapanKoreaNetherlandsTurkey andTaiwan | Terminated | April 19, 2011 | 943 | 165 | Computer
Z-0285-2009 | 49135 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare Solar 8000M System, a multiparameter physiological patient monitoring system intended for use on adult, pediatric and neonatal patients, within a hospital or facility environment |  |  | April 02, 2008 | 2008 | GE Healthcare | A software error has been discovered when using Combination Monitoring with Solar 8000M/i - PDM - Telemetry in combo mode. This error has only been noted with this particular configuration and mode of monitoring. It has not been observed in other monitoring applications. The software error will result in a reset of the Solar 8000M/i, causing a loss of monitoring for approximately 1-2 minutes a | Software design | A GE Healthcare consignee letter was sent out 9/9/08. The letter dated August 26, 2008 was addressed to Healthcare Administrator/Risk Manger. The letter described the problem, product and provided recommendation. |  | 2782 | 2782 | Nationwide to: AL, AR, AZ, CA, FL, GA, IA, ID, IL, IN, MA, MI, MN, MO, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, VA, and WI. OUS to include: Australia, Austria, Belgium, Brazil, Czech Republic, Denmark, Egypt, France, Germany, Indonesia, Ireland, Italy, Jordan, Korea, Lebanon, Lithuania, Malaysia, Netherlands, Norway, Pakistan, Saudi Arabia, Spain, Sweden, United Kingdom, and Venezuela. | Terminated | May 17, 2012 | 1506 | 2782 | Computer
Z-0628-2009 | 50147 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | TomoTherapy HI-ART Systems with versions 2.2.4, 3.1.2, 3.1.3 or 3.2.1 software. The affected applications include Planning Station, Planned Adaptive, Data Management System, and TomoPortal. |  |  | October 14, 2008 | 2008 | TomoTherapy Incorporated | TomoTherapy Incorporated identified a potential issue with the Hi-Art system during the course of ongoing testing. The Operator Station Calibration panel provides access to view and modify machine specific configuration settings. Access to these settings has always been restricted to individuals with appropriate security rights, being limited to only the "Superuser" and "Field Service engine | Software design | A TomoTherapy Incorporated "Urgent Medical Device Correction Field Safety Notice" letter dated October 17, 2008 was sent to all consignees. The letter addressed the Issue, Product Affected, Recommended Actions and Resolution. Please direct questions to the TomoTherapy, Inc. Customer Care Center by email (www.support@tomotherapy.com) or by telephone, North America: 1-866-368- 4807 and all other locations: 1-608-824-2900 or 32-2-400-44-44. |  | 215 | 215 | Worldwide Distribution --- USA including states of AL, AK, AZ, AR, CA, CO, CT, FL, GA, HI, IL, IN, IA, KY, LA, MD, MA, MI, MN, MO, NE, NV, NJ, NM, NY, OH,OK , OR, PR, PA, SC, SD, TN, TX, VA, WA, WI, WY, and DC, and countries of BELGIUM, CANADA, SWITZERLAND, CHINA, GERMANY, SPAIN, FRANCE, GREAT BRITAIN, HONG KONG, ITALY, JAPAN, KOREA, NETHERLANDS, SINGAPORE, TAIWAN and TURKEY. | Terminated | February 10, 2011 | 849 | 215 | Computer
Z-0001-2009 | 50138 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | NeuViz Dual Multi-Slice Computed Tomography (CT) Scanner System (software). |  |  | April 10, 2008 | 2008 | Philips Medical Systems North America Co. Phillips | An issue associated with the "Timed Scan" application was found. The system does not wait for the programmed time interval and starts the x-ray after selection of the scan button when certain criteria is met. | Software design | A letter was issued to customers prior to the release of the field correction during June 2008. The system software will be updated to correct the non-compliances. For information or support concerning this issue, please contact the Neusoft Medical Systems Service Support Department (helpdesk@pnms.neusoft.com) or contact Philips Medical Systems North America Co. Phillips at 425-487-7602. |  | 19 | 19 units in the US. | Nationwide Distribution --- USA including states of CO, FL, GA, IL, KY, NC, TN and TX. | Terminated | October 29, 2010 | 932 | 19 | Computer
Z-1038-2009 | 50839 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Colleague Triple Channel Volumetric Infusion Pumps. Baxter Healthcare Corporation, Medication Delivery Division, Deerfield, IL 60015 U.S.A., Made in Singapore. Product codes: 2M8153, 2M8163, 2M9163, DNM8153, DNM9163, BRM8153, BRM8153T, BRM8153RT, PNM8153, PNM8153T, PNM8153RT, CNM8153, GNM8153, HNM8153, TRM8153 and WNM8153. |  |  | January 23, 2009 | 2009 | Baxter Healthcare Corp. | Baxter has identified failure codes that lead to a delay or an interruption of therapy, hazards associated with improper cleaning, and damaged battery messages for the Colleague Infusion Pumps. | Device Design | Baxter Healthcare Corporation sent an "Urgent Device Correction" letter dated January 23, 2009 to all Colleague Infusion Pump customers, to the attention of the Vice President of Nursing with copies to the Chief Operating Officer, Director of Biomedical Engineering and Director of Risk Management to advise them of failure codes that lead to an interruption of therapy, hazards associated with improper cleaning, and damaged battery messages associated with the Colleague pumps. The letters listed the failure codes involved that lead to an interruption of therapy which could lead to serious injury and/or death; advised institutions to have contingency plans to verify that back-up pumps are available; and provided steps for addressing an interruption of therapy with any failure code. The letters also provided information about the hazards associated with improper cleaning of the pumps and instructions for addressing damaged battery messages. Customers were instructed to complete the attached Customer Reply Form acknowledging receipt of the letter and fax it to 888-569-1585. Direct questions to your local Baxter Representative or to Baxter Global Technical Services by calling 1-800-843-7867, Prompt 2, between 7AM and 5 PM Central Standard Time (CST). |  | 81775 | 81,775 pumps | Worldwide including USA, Puerto Rico, Aruba, Australia, Austria, Bahamas, Bahrain, Belgium, Bermuda, Brazil, Canada, Cayman, Chile, China, Colombia, Curacao, Denmark, Dominican Republic, El Salvador, Finland, France, Germany, Greece, Guatemala, Hong Kong, Honduras, India, Ireland, Israel, Italy, Jamaica, Korea, Kuwait, Lebanon, Macedonia, Malaysia, Malta, Mexico, Netherlands, New Zealand, Nicaragua, Norway, Oman, Panama, Peru, Philippines, Qatar, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Trinidad, Turkey, United Arab Emirates, United Kingdom and Venezuela. | Terminated | October 04, 2011 | 984 | 302872 | Computer
Z-0928-2009 | 50539 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Centricity PACS RA1000 Workstation (for diagnostic image analysis). GE Healthcare Integrated IT Solutions, Barrington, IL 60010. |  |  | November 21, 2008 | 2008 | GE Healthcare Integrated IT Solutions | There is a potential patient safety issue involving patient jacket content, which may result in unintended historical image(s) being displayed. | Software design | GE Healthcare Integrated IT Solutions notified their Centricity PACS RA1000 Workstation software customers via letter (Urgent Medical Device Correction) dated 11/21/08 of the software anomaly resulting in a patient safety issue involving patient jacket content, which may result in unintended historical image(s) being displayed. The letter provided the users with safety instructions to follow until the software is updated. Please direct any questions to the Remote Online Center, Field Engineer or the GE Customer Care Center at 800-437-1171. The accounts were requested to complete the enclosed customer reply form, supplying the workstation software or build number, and fax the completed form to 414-918-4653. |  | 476 | 476 units | Worldwide Distribution including USA and country of Canada. | Terminated | September 06, 2010 | 654 | 476 | Computer
Z-0985-2009 | 50746 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Orthosoft Navitrack, OS Unicondylar Knee 1.0 - Universal; software application for orthopedic surgical stereotaxic instrument. |  |  | August 19, 2008 | 2008 | Zimmer Inc. | Incorrect parameters: If the user switches from one surgeon profile to another surgeon profile after completing anatomical landmark acquisition tasks, the application will continue to use the initial surgeon profile parameter values in the calculation of the predicted hip-knee-ankle alignment estimates, which may have different resection or component thickness parameters than are appropriate for | Software design | Zimmer notified consignees via letter dated 8/19/08, which advised them not to change the selected surgeon profile after the acquisition of landmarks has begun, and the issued will be fixed with the next software release. |  | 75 | 75 (34 U.S. and 41 international) | Worldwide distribution: USA, Australia, Austria, Belgium, Canada, France, Germany, Italy, Singapore, South Africa, Switzerland and United Kingdom. | Terminated | April 28, 2010 | 617 | 75 | Computer
Z-0944-2009 | 50541 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Symbiq One-Channel list number 16026-04 and Two-Channel list number 16027-04 Infusion System, a general purpose infuser, Distributed and Manufactured by Hospira Inc, Morgan Hill, CA. |  |  | December 03, 2008 | 2008 | Hospira Inc | Hypersensitivity to normal motor friction/drag may cause malfunction errors and alarms that are designed to stop pump operations. | Software design | Recall initiated 12/03/2008. Consignees were notified of the issue via an Urgent: Device Field Correction letter, dated 12/3/2008, and were requested to respond appropriately to any alarm or warnings displayed by the device. The firm is making arrangements to screen each Symbiq pump in customer facilities and implement the appropriate enhancements. |  | 11840 | combined 11,840 units | Worldwide Distribution --- including USA and countries of Canada and Australia. | Terminated | April 03, 2012 | 1217 | 11840 | Computer
Z-0919-2009 | 50064 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Cardiac Science Powerheart Automated External Defibrillator G3 9390A |  |  | October 13, 2008 | 2008 | Cardiac Science Corporation | Potential for AED to not deliver defibrillation therapy. | Software design | On October 13, 2008 Cardiac Science began sending via Fed-Ex letters and software upgrade kits. Mailing of letters and software upgrade kits is anticipated to be completed by November 21, 2008. The letter instructs consignees to install the software update, assure that the AED information test file generated by the update software is returned to the firm, and complete the postage paid response card and return it to the firm. If the affected devices are no longer in the possession of the consignee, the consignee is to contact the firm and forward the update kit and letter to the person who has the AED in their possession. A separate letter is being sent to distributors. The letter advised to install the software upgrades for any units in stock or to forward the software upgrade along with the end user letter enclosed and that the AED information text file generated by the update software must be returned to the firm. The firm will visit the locations of individual national account customers and install the update. OEM partners will be notified and provided the information and update kits to provide to their customers. |  | 2572 | 2,572 devices | Worldwide Distribution: USA, Puerto Rico, Virgin Islands, Australia, Belgium, Bermuda, Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Czechoslovakia, Denmark, Dominican Republic, France, Germany, Greece, Guam, Hong Kong, Hungary, Ireland, Israel, Italy, Jamaica, Japan, Kenya, Kuwait, Libya, Malaysia, Mexico, Netherlands, New Zealand, Norway, Oman, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, UK, and Ukraine. | Terminated | December 08, 2011 | 1151 | 24183 | Computer
Z-0940-2009 | 49610 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Stryker InTouch 1.0 Critical Care Bed; Model 2130. |  |  | January 05, 2009 | 2009 | Stryker Medical Div. of Stryker Corporation | The touch screen on the footboard may display a "tilt over range" error, resulting in a precautionary stop of the bed motions which result in the bed appearing to have locked out some of its bed positioning features. | Device Design | Consignees were notified of the problems via letter dated January 5, 2009, advised to check their beds and to take them out of service if the foot board touchscreen displayed a "tilt over range" error, and that a Stryker Medical field service representative would contact them to perform software upgrades on the affected beds. Please direct questions to Stryker Medical at 1-269-324-6644. |  | 1288 | 1,288 of all models | Worldwide Distribution including USA and countries of Canada, Japan and Switzerland. | Terminated | November 05, 2010 | 669 | 1288 | Computer
Z-0861-2009 | 50496 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Medical Systems Eclipse treatment planning system, Model Number H48, versions 6.5/7.x and Inspiration 8.0. |  |  | November 21, 2008 | 2008 | Varian Medical Systems Oncology Systems | Possible incorrect wedge orientation in treatment: Software labeling was issued without warning message relating to preservation of user selected values for Siemens wedge orientation, wedge limits and SAD values for Elekta machines, and collimator operating limits for Saturne Machines when the software version is upgraded. | Software design | An Urgent Device Correction product notification letter/field safetly letter dated November 21, 2008, was distributed to all affected users, with a description of the problem and user corrective action steps. The letter was also distributed to Varian Sales, Marketing and Service organizations, informing them of the issue. |  | 516 | 516 units | Worldwide Distribution: USA, Austria, Belgium, Canada, Czech Republic, Finland, France, Germany, Greece, Hungary, Italy, Ireland, Lithuania, Norway, Portugal, Romania, Slovakia, Spain, Sweden, Switzerland, Netherlands, Turkey and the United Kingdom. | Terminated | March 10, 2011 | 839 | 516 | Computer
Z-0982-2009 | 50700 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | DePuy DNP Anatomic Fracture Repair System, JIG DNP Anatomic, Depuy, Warsaw, IN; REF DNPAJIG. |  |  | December 23, 2008 | 2008 | Depuy Orthopaedics, Inc. | The DNP nail plate may not attach to the alignment guide (jig) instrument, making it difficult to insert the nail plate into position or to drill needed holes accurately. | Device Design | The firm's sales force was instructed via "Urgent Device Correction" letter issued December 29, 2008 to inspect all of these instruments in their possession, and those instruments at their hospital customers, and to leave a copy of the letter with each hospital customer who has the instruments on site.The letter describes the issue and corrective actions for customers. Please direct questions about this recall to Depuy Orthopaedics, Inc. at 888-783-7156. |  | 3250 | Approx. 3,250 of all products | Nationwide Distribution. | Terminated | October 18, 2010 | 664 | 3250 | Computer
Z-0863-2009 | 50522 | No | JCX | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | DeVilbiss Model 7310PR-D VacuAide Compact Suction Unit. A portable, AC-powered or compressed air-powered suction pump. The device may include a microbial filter. |  |  | October 27, 2008 | 2008 | Sunrise Medical | Improper suction: Regulator knob does not turn properly which can affect the suction level. | Device Design | DeVilbiss notified dealers and users (through the dealers) to recall the units, by a telephone call October 27, 2008 followed by an Urgent Medical Device Recall mailing November 3, 2008. The letter advised customers what models are affected and how to contact the firm of dealer to have the affected unit replaced. Corrective and preventative action on current manufacturing has been implemented. |  | 78 | 78 units | Worldwide Distribution: USA, Mexico, Guatemala, Colombia, and Costa Rica. | Terminated | September 25, 2012 | 1429 | 78 | Computer
Z-0635-2009 | 50227 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Terumo TenderFlow Pediatric Arterial Cannulae, 3.3mm (10 Fr) OD, with 1/4 inch non-vented connector, 9 inch (22.9 cm) long; Catalog no. 813572. |  |  | October 16, 2008 | 2008 | Terumo Cardiovascular Systems Corporation | The introducer or stylet is difficult to remove from the pediatric arterial cannula body, which may result in the need to remove the cannula and replace it. | Device Design | An Urgent Medical Device Recall letter dated 10/16/08 was sent to consignees advising them of the problem; requesting that they immediately verify whether or not they had unused cannulae with the affected lot numbers; requesting that they isolate any affected product to prevent use; complete the Customer Response Form, indicate on form if it is medically necessary to use product without introducer; fax Customer Response Form to Terumo CVS at the number listed; and Terumo CVS would contact customers to arrange for the return of unused product. The letter stated that credit would be issued for any unused TenderFlow cannulae that were returned. The quantities to be returned should be indicated on the Customer Response Form. Consignees were advised that, if it was medically necessary to use the affected cannulae, they should be used without the introducer and in accordance with instructions for such use. |  | 288 | 288 (170 plus 78 samples to US and 40 international) | Worldwide Distribution --- including USA and countries of Australia, Belgium and Canada. | Terminated | August 04, 2009 | 292 | 2892 | Computer
Z-0951-2009 | 49474 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ACCELERATOR APS Centrifuge Module (Hettich Robotic centrifuge); List Number: 07L02-01; manufactured by Inpecco SpA, Segrate (MI), Italy. |  |  | September 09, 2008 | 2008 | Abbott Laboratories, Inc | ACCELERATOR APS Centrifuge Module being recalled due to potential for micro-cracks to form over time in the metal centrifuge rotor buckets that may cause the centrifuge buckets to disconnect from the rotor in the centrifuge during centrifugation. | Device Design | Notification letters sent to consignees on 09/09/08. Consignees provided instructions for inspecting the buckets for cracks and were asked to contact customer service reps for replacement buckets. Consignees also asked to respond via fax using Customer Reply form acknowledging receipt, understanding and completion of correction letter. |  | 6 | 6 | Nationwide Distribution --- including states of TX, UT, and VA. | Terminated | October 04, 2010 | 755 | 6 | Computer
Z-0624-2009 | 45967 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Medtronic O-Arm Intraoperative Imaging System |  |  | November 27, 2007 | 2007 | Medtronic Navigation, Inc | Navigation accuracy problem: Accuracy may fall outside of specification when used in combination with the Medtronic Synergy Experience StealthStation System Spine Software with the O-Arm Intraoperative Imaging System and tracker. | Software design | Medtronic Navigation notified Field representatives on 11/27/07 to perform site visits to implement FCA 2007-005 - O-Arm Intraoperative Imaging System Accuracy, and deliver the Customer Letter notification. Confirmation of the visit was documented. |  | 54 | 54 units | Worldwide Distribution: USA, Canada, Czech Republic, Denmark, Germany, Netherlands, S. Korea, Poland, South Africa, and Singapore | Terminated | January 22, 2016 | 2978 | 54 | Computer
Z-0970-2009 | 50759 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Centricity PACS RA1000 Workstation (for diagnostic image analysis); GE Healthcare Integrated IT Solutions, Barrington, IL 60010. |  |  | November 20, 2008 | 2008 | GE Healthcare Integrated IT Solutions | Unread exams may inadvertently get marked as dictated in PACS workstation when the user signs off the report in Talkstation and opens the next exam. | Software design | GE Healthcare Integrated IT Solutions notified their Centricity PACS RA1000 Workstation with Talk Extend software customers via letter (Urgent Medical Device Correction) dated 11/10/08 of the software anomaly resulting in a patient safety issue involving Talk Extend, which could contribute to a delay of care or misdiagnosis due to delay in communicating pertinent exam results or failure to treat due to lack of an attached dictated report by the radiologist for physician review. The letter provided the users with safety instructions to follow until the software is updated. Any questions were directed to the Remote Online Center, Field Engineer or the GE Customer Care Center at 800-437-1171. The accounts were requested to complete the enclosed customer reply form, supplying the workstation software or build number, and fax the completed form to 847-939-1479. |  | 22 | 22 units | California, Florida, Illinois, Iowa, Louisiana, Oregon, Pennsylvania, Texas, Utah, Virginia and Wisconsin. | Terminated | February 03, 2010 | 440 | 22 | Computer
Z-0945-2009 | 50755 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX Cardiovascular Polling Procedure version 1.00.09, used in IMPAX 6.2.1 and 6.3.1.; Catalog numbers: E18R4000, L5NQJ000, L57OP000, LBCDB000, and E2ZCK000. |  |  | December 24, 2008 | 2008 | AGFA Corp. | Polling procedure may fail when multiple merge/split is carried out. | Software design | All sites potentially affected by this product were notified of the issue directly via telephone. Each site was directed to observe a ten minute polling window as an interim mitigation until the normal mitigation by software Service Update is performed at the site by Agfa Support. An Urgent Safety Notice was sent via FED-EX on/about 12/24/2008 to all sites. The letter describes the potential issue and both interim and formal mitigation. Acknowledgment, via FAX-Back, that the information was received and understood was requested from the sites. |  | 5 | 5 units | Nationwide Distribution --- including states of FL, IL, OR, SC, and SD. | Terminated | April 01, 2010 | 463 | 5 | Computer
Z-0872-2009 | 50676 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | SERVO Ventilator 300A (SV300A); |  |  | December 19, 2008 | 2008 | MAQUET Inc. | SERVO Ventilator 300A could cause hypoventilation if the use of the PC or PRVC mode with the Automode function set to "On" in Neonatal and Pediatric patient ranges with small tidal volumes and short set inspiratory rise time. | Device Design | URGENT: Device Correction letters were sent out on December 22, 2008. Letters alert users to the potential problem and the situation in which it may occur. Letter indicate the set inspiratory rise time should be set in the range of 7-10% and that the humidifier chamber water level be maintained according to manufacturers instructions. The letter is enclosed as a supplement to the SV300A Operating Manual. Consignees were directed to complete a "Device Correction-Customer Fax Response" to indicate number of units of the affected product and that they followed the instructions for including the attached supplemental instructions to the operating manual. For questions about the recall, contact your local Maquet, Inc. representative. |  | 3688 | 3,688 units in US | Nationwide | Terminated | July 08, 2009 | 201 | 3688 | Computer
Z-0936-2009 | 50774 | No | MDM | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Stryker AVS AL Spacer Trial Inserter; |  |  | September 08, 2006 | 2006 | Stryker Spine | DESIGN FLAW: Issues were reported concerning fracture of handle, difficulty in removing inserter and inserter difficult to lock. | Device Design | Market Withdrawal letters were sent out on September 8, 2006 to all branches who received device. The letter asked customers to examine their inventory and hospital consignment locations to identify the product. Also, customers needed to reconcile all products on the attached Product Accountability Form and fax a copy back. The affected products were to be retrieved and returned using the "Product Returns Procedure". |  | 37 | 37 | Nationwide Distribution --- including state of NJ. | Terminated | February 05, 2009 | 881 | 37 | Computer
Z-0969-2009 | 50758 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Centricity PACS RA1000 Workstation (for diagnostic image analysis); GE Healthcare Integrated IT Solutions, Barrington, IL 60010. |  |  | January 12, 2009 | 2009 | GE Healthcare Integrated IT Solutions | If a user site created a custom default display protocol (DDP) to display the current exam versus the historical exam and the user-defined DDP fails, the system displays the exams using a GE default hanging protocol. This may create a potential for misdiagnosis if the user does not review the date and time stamps in the title bars. | Software design | GE Healthcare Integrated IT Solutions notified their Centricity PACS RA1000 Workstation software customers via letter (Urgent Medical Device Correction) dated 1/12/09 of the software anomaly resulting in a patient safety issue involving Default Display Protocols. The issue may create a potential for misdiagnosis if the user does not review the date and time stamps in the title bars. The letter provided the users with safety instructions to follow until the software is updated. Any questions were directed to the Remote Online Center, Field Engineer or the GE Customer Care Center at 800-437-1171. The accounts were requested to complete the enclosed customer reply form, supplying the workstation software or build number, and fax the completed form to 847-939-1479. |  | 778 | 778 units | Nationwide, Canada, Bermuda, Mexico, Ecuador, Chile, and Brazil | Terminated | September 06, 2010 | 602 | 778 | Computer
Z-0948-2009 | 50751 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | iSite PACS (Picture Archiving and Communications System) Versions 4.1.x up to and including 4.1.51.3, manufactured by Philips Healthcare Informatics, Foster City, CA. |  |  | December 29, 2008 | 2008 | Philips Healthcare Informatics, Inc. | Product may display clinical reports with the study date and time in GMT format, which may not be obvious to the end user, who may expect to see the local time in the report. | Software design | Product notification was issued on 12/29/2008, instructing users to identify affected products, and instructing users to refer to image overlays for local study date and time information. The firm is planning a field correction expected to be available in "Q1 2009". For additional information, contact Philips at 650-293-2624. |  | 40 | 40 units | International - Italy, Denmark, Germany, Netherlands, Sweden, Spain, Portugal, Switzerland, Spain, Turkey, France, Saudi Arabia, Australia, Thailand, New Zealand, Japan, and Brazil. | Terminated | November 29, 2010 | 700 | 40 | Computer
Z-0949-2009 | 50787 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Artiste MV Digital Linear Accelerator System Equipped with 160 MCL Accessory Holder, Manufactured by Siemens Medical Solutions, Oncology Care System, Concord, CA |  |  | December 20, 2008 | 2008 | Siemens Medical Solutions USA, Inc | The range of tolerance is to high for the accessory holder, therefore, Wedges and other accessories do not latch to the locking mechanism correctly. | Device Design | On 12/20/08, the firm issued letters via certified mail to direct consignees, informing them of the affected product and providing instructions on the recall.Contact Siemens Medical Solutions USA, Inc. at 1-925-246-8407 for additional information. |  | 43 | 43 units | Nationwide and Worldwide | Terminated | December 09, 2010 | 719 | 43 | Computer
Z-0931-2009 | 50523 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Low Profile U-Joint Driver for the SynFix Mini-Open System, Catalog number 03.802.231. |  |  | November 20, 2008 | 2008 | Synthes Spine | screwdriver tip may break off and lodge in the screw head | Device Design | The recalling firm emailed all Spine Sales Consultants on 11/20/08. They were instructed to immediately examine their stock and discuss appropriate technique with surgeon users. In addition, a Medical Device Correction letter dated 11/19/08 was issued via USPS certified mail to the medical facilities. This initial letter stated the issue and requested the users take the following actions: discard all SynFix Technique Guides and replace with the updated literature contained in the mailing packet with the letter, and is identified by code J7022-E (Rev. E); to ensure receipt and understanding of the information, the Verification Section of the letter needs to be completed and returned to Synthes by either fax: 610-719-5120 or scan/e-mail: lewis.lynne@synthes.com. An Urgent: Medical Device Recall letter, dated 1/16/2009, was also sent out. This latter letter requested that users examine their stock immediately and remove any of the affected devices from their inventory. Customers are also to call Synthes at 1-800-479-6329 to obtain a Return Authorization Number; complete the Verification Section; and return the letter with the recalled product to the address listed. There are also actions for those who do not have the affected product. Questions should be directed to 800-620-7025 x 5453 or 610-719-5453. |  | 381 | 381 units | Class II Recall - Worldwide Distribution --- USA including states of AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NE, NJ, NM, NV, NY, OH, OK, PA ,SC, TN, TX, UT, VA, WA, and WI, and countries of Canada and Switzerland. | Terminated | October 27, 2009 | 341 | 381 | Computer
Z-0927-2009 | 50515 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Artiste MV SA Linear Accelerator with syngo RT Therapist Version 4.1 build 67 and build 122. Manufactured by Siemens Medical Solutions Oncology Care Systems, Concord, CA. |  |  | October 23, 2008 | 2008 | Siemens Medical Solutions USA, Inc | If a plan contains a sequence of beams, including at least 1 ARC treatment beam and if during the setup stage of a beam following an ARC beam in the delivery sequence a termination occurs, and the user selects the immediate resumption feature, syngo RT Therapist Express systems will display and record a second delivery of a prior beam and not the delivery of the current beam in certain circumstanc | Software design | An Urgent Medical Device Customer Safety Letter dated October 23, 2008 was issued to consignees via certified mail. The letter identified the issue and preventive measures for consignees. Monitoring will be done via return cards. Follow-up to be conducted via telephone. |  | 36 | 36 units | Worldwide Distribution including USA and countries of Netherlands, Germany, Poland, Norway, Australia, Hungary, Belgium, New Zealand, France, Malaysia, Spain, and Canada. | Terminated | December 09, 2010 | 777 | 36 | Computer
Z-1500-2009 | 51887 | Yes | LDF | Cardiovascular permanent or temporary pacemaker electrode. | Cardiovascular | Cardiovascular | 510(k) | Arrow Balloon and stylet temporary pacing catheter |  |  | February 11, 2009 | 2009 | Arrow International Inc | Leakage: Material not lipid resistant, may crack and/or leak. Use of this device with lipid containing solutions or 70% isopropyl alcohol may cause leakage or air embolism. | Device Design | Arrow International issued a letter dated 2/18/09 to all customers informing them of the problem and if a lipid resistant stopcock is required to contact the firm for replacements. New labels state: Warning - do not use with lipid containing solutions or 70% isopropyl alcohol, as these solutions may cause leakage or air embolism. |  | 900000 | 900,000 total, all products | Worldwide distribution: USA, Argentina, Australia, Bahamas, Belarus, Belgium, Brazil, Chile, China, Colombia, Costa Rica, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Finland, France, Germany, Greece, Guatemala, Hungary, Iceland, India, Indonesia, Ireland, Israel, Italy, Japan, Latvia, Lithuania, Malaysia, Mexico, Netherlands, New Zealand, Norway, Pakistan, Peru, Philippines, Poland, Singapore, Slovakia, South Africa, South Korea, Spain, Sri lanka, Sweden, Taiwan, Thailand, Tunisia, Turkey, UK, Ukraine, and Venezuela. | Terminated | July 22, 2010 | 526 | 900000 | Computer
Z-1505-2009 | 50773 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Siemens Medical Solutions E.CAM Patient Handling System (PHS) Field of View (FOV) indicator bars; a component of the E.CAM Emission Computed Tomography System; part number 43 66 618 (right) and 43 81 088 (left). |  |  | December 18, 2008 | 2008 | Siemens Medical Solutions USA, Inc. | Patient fingers can get pinched if they attempt to pull themselves up from the patient handling system without any assistance and grabbing onto the flip-up FOV indicator bars. | Device Design | Siemens sent Urgent Field Correction Recall letters dated 12/18/08 to all users of the E.cam Systems. The accounts were informed of the potential risk to patients when they attempt to pull themselves up from the patient handling system (PHS) without any assistance and grabbing the field of view (FOV) indicator bar (aka flip-up armrests). The customers were advised to always assist the patient when getting on/off of the PHS. They also were advised that a Siemens Customer Service Representative would contact them within 90 days to schedule a corrective upgrade to the PHS with a modified FOV guide to mitigate the risk for patient injury free of charge. Any questions were directed to contact Siemens Medical Solutions USA, Inc. at 1-800-767-2313 (USA) for assistance. |  | 1107 | 1,107 units | Worldwide distribution: USA, Guam, Argentina, Australia, Austria, Bangladesh, Belgium, Bosnia and Herzegovina, Brazil, Canada, Chile, China, Czech Republic, Ecuador, Egypt, Finland, France, Germany, Great Britain, Hong Kong, India, Ireland, Israel, Italy, Japan, Korea, Latvia, Lebanon, Mauritius, Mexico, Myanmar, Netherlands, New Zealand, Norway, Pakistan, Peru, Philippines, Poland, Portugal, Russia, Saudi Arabia, Spain, Sweden, Switzerland, Taiwan, Turkey and United Arab Emirates. | Terminated | May 03, 2011 | 866 | 1107 | Computer
Z-1478-2009 | 52026 | No | HGM | Perinatal monitoring system and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Philips Avalon Fetal Monitor FM20; Model Number: M2702A. |  |  | May 12, 2009 | 2009 | Philips Healthcare Inc. | Interruption of power-up/boot sequence may cause the paper scale and recorder speed settings of the monitor to revert to factory default may lead to misinterpretation of the fetal trace by the user and to potential incorrect clinical decisions | Software change control | Philips Healthcare issued an "Urgent-Medical Device Correction" notification dated May 15, 2009 to Customers in the United States via UPS and notifications outside the US will be managed by Philips representatives in each affected geography. Customers were asked to follow the guidelines in the Action to be Taken by Customer/User section of the Correction Notice until they receive their software upgrade.Questions can be made to Philips response center at 1-800-722-9377. |  | 5663 | 5,663 units | Worldwide Distribution | Open, Classified |  |  | 10451 | Computer
Z-1987-2009 | 52662 | No | KWY | Hip joint femoral (hemi-hip) metal/polymer cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Apex Hip System Bipolar Head Bipolar Head, 28 mm ID, 59 mm OD |  |  | July 10, 2009 | 2009 | Omni Life Science | The APEX Hip System Bipolar Head implant may not properly capture the femoral head assembly of the bipolar head onto the femoral head. | Device Design | Omni Life Science notified distributors and implanting surgeons by an Urgent, APEX Hip System Bipolar Head Voluntary Recall letter dated June 29, 2009. The method of notification is both e-mail and FedEx overnight. The letter requires a signature acknowledging receiving the letter and the letter is to be faxed back to Omni Life Science. The Distributors/Sale Representatives are instructed to fill in the "Quantity Returned" on the recall letter and fax back to Omni. They are to ship all existing Apex Hip System Bipolar Head inventory by 07/10/2009 to the directed address. |  | 17 | 17 units | Nationwide Distribution -- Including states of AZ, CA, CO, FL, IL, MI, NJ, OK, UT, VA, and WI. | Terminated | June 28, 2012 | 1084 | 295 | Computer
Z-2054-2009 | 52434 | No | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | EUB-6500/HI VISION 6500 Diagnostic Ultrasound Scanners, DICOM Option EZU-FC5, Software Versions: V01-03, V01-05, V02-04, V03-02, V04-02, V04-03, V04-04, V05-03, V05-05A, V05-07A and V05-30A. |  |  | May 12, 2008 | 2008 | Hitachi Medical Systems America Inc | The firm's parent company discovered a DICOM UID (Universal Identifier) code mix-up which can occur due to a software error. This error can cause the images from two (2) different patient scans to become intermingled and consequently, labeled incorrectly. | Software design | HMSA issued a 'DEVICE CORRECTION' notification to their customers on 5/12/2008. The notification describes the potential data transmission problem (potential to archive studies done on different patients into the same patient folder) and the conditions under which this problem can occur. The notification cautions the user/customer to verify all images stored on the PACS (Picture Archiving and Communications System) in order to ensure that they do not contain images from a different patient. The DEVICE CORRECTION letter also states that the firm will be installing updated software on all affected systems beginning in June, 2008, in order to address this issue. It also states that a Hitachi Service Rep. will be contacting each customer and making arrangements to set up a convenient time to come out and perform a software upgrade. Lastly, the letter provides a toll-free telephone # to the firm's 'Customer Service Center: 1-800-800-4925, which the customer can use to obtain additional information, if desired. |  | 10 | 10 units | The recalled device system units were distributed to customers located in the following states within the U.S ony.: AR, CA, and TX. | Terminated | July 20, 2011 | 1164 | 16 | Computer
Z-1912-2009 | 52694 | No | GAA | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Broselow/Hinkle Pediatric Emergency System; |  |  | June 30, 2009 | 2009 | Vital Signs Inc | Not compatible: The luer connector on the Intraosseus needle provided does not fit securely on the luer connection of the IV Extension Set in the module. | Device Design | Vital Signs, Inc. issued an "Urgent Medical Device Correction" letter dated June 30, 2009 to all direct and indirect accounts by first class mail. Consignees were informed of the affected device, asked to complete a Confirmation form for return and to return unused product for replacement.For further questions, contact Vital Signs, Inc. at 1-973-790-1330 extension 404. |  | 23429 | 23,429 kits | Worldwide Distribution -- US, Canada, United Arab Emirate, Venezuela, Guam, Korea and United Kingdom. | Terminated | August 27, 2012 | 1154 | 23429 | Computer
Z-1887-2009 | 52757 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Agfa Impax R6.3 DB Server. |  |  | July 16, 2009 | 2009 | AGFA Corp. | Measurements are inaccurate on multiple workstations. | Software design | Agfa Healthcare issued an "URGENT FIELD SAFETY NOTICE" dated July 16, 2009. The notice informed of the problem and provided details to identify potentially compromised data that could result in measurement inaccuracies. A feedback form was provided to be completed and returned to the firm.For further information, contact Agfa Healthcare at 1-773-592-4769. |  | 1 | One unit | Nationwide Distribution -- IL | Terminated | April 28, 2011 | 651 | 1 | Computer
Z-1946-2009 | 52232 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | ARJO Carendo Multi-Purpose Hygiene Chair; Model BIB2003-01, a non-AC powered patient lift; Manufactured by Arjo Hospital Equipment AB, Eslov, Sweden. |  |  | June 11, 2009 | 2009 | ArjoHuntleigh | Male patient genitals may be pinched between the seat and the frame of the patient lift chair when the Care Raiser seat is returned to its lowest position. | Device Design | ArjoHuntleigh sent Urgent Device Field Correction letters dated June 11, 2009, along with the Urgent Field Safety Notice ID 090526AHE dated May 26, 2009, via UPS Ground delivery. Customers were informed of the affected device and were requested to complete the response form with new contact information and fax it back to ArjoHuntleigh at 630-567-5012 or 630-925-7969. For further questions, contact ArjoHuntleigh at 1-800-323-1245. |  | 1987 | 1,987 hygiene chairs | Nationwide Distribution | Terminated | August 19, 2010 | 434 | 1987 | Computer
Z-1883-2009 | 52767 | No | BRY | Anesthetic cabinet, table, or tray. | Anesthesiology | Anesthesiology | 510(K) Exempt | Pyxis Anesthesia System 3500, Cardinal Health, Clinical Technologies and Services, San Diego, CA. |  |  | December 24, 2008 | 2008 | Cardinal Health Manufacturing LLC | Cardinal Health has sent out a Safety Alert Notification to inform its customers that the SR2b Software for the Pyxis Anesthesia System 3500 (PAS 3500) has received several customer reports of system "non-responsiveness" (lock ups) during the use of the product since its release in September 2008. | Software design | The "Safety Alert Notification" letter was mailed to all Pyxis Anesthesia System 3500 customers on December 17, 2008 by registered return receipt mail. The letter was addressed to the Director of Pharmacy at each facility. The mailing packages contain the "Safety Alert Notification" letter and the revised pages of the Pyxis Anesthesia System 3500 User Guide and Safety Manual. The "Safety Alert Notification" letter reinforced to customers how to maintain back-up processes to allow quick access if a system lock-up condition is experienced.The letter informed customers the situation will be corrected and a solution will be implemented. If you experience any system unresponsiveness during the use of the product, please report the issue to the Cardinal Health Technical Support Center at 1-800-727-4102. If you have any questions regarding this issue, please contact your local field support representative or the Customer Advocacy Department at Customer.Advocacy@cardinalhealth.com. |  | 250 | 250 | Nationwide and Germany. | Terminated | March 28, 2012 | 1190 | 250 | Computer
Z-1808-2009 | 52697 | No | HQC | Phacofragmentation system. | Ophthalmic | Ophthalmic | 510(k) | American Optisurgical, Horizon Phacoemulsification System, Model: HZN-2030, for ophthalmic surgery. |  |  | March 24, 2009 | 2009 | American Optisurgical Inc | Device operation interrupted: A software bug causes Phaco to stop when vacuum goes above 300 mmHg while in Multi Burst mode. Use of this product at the parameters specified above could lead to a delay in surgery that might present a health risk. | Software change control | American Optisurgery sent an E-mail to customers on April 2, 2009 stating the reason for recall and informed them to immediately examine their inventory and quarantine product subject to recall. In addition, if they have further distributed this product, please identify customers and identify them at once of this product recall. Their notification to customers may be enhanced by including a copy of this recall notification letter. They were then told to refer to the enclosed Recall Questionnaire for instructions on what to do with the recalled product: The recall should be carried out to the user level and return the enclosed response form as soon as possible. If there are any questions, they are advised to contact David Salzberg at (800) 576-1266. |  | 12 | 12 units | Worldwide distribution: USA (California), Brazil, Algeria, France, South Korea, India, and Peru. | Terminated | September 04, 2009 | 164 | 12 | Computer
Z-2008-2009 | 52209 | No | LGZ | Infusion pump. | General Hospital | General Hospital | 510(k) | Model 100 enFlow Fluid/Blood Warmer |  |  | March 23, 2009 | 2009 | Enginivity LLC | 1. Fluid contamination of the electronic circuitry can cause short circuits resulting in permanent fuse failure or localized circuit board overheating with possible smoldering and discoloration. 2. Electromagnetic (EMI) Interference associated with the Warmer may produce artifact in ECG, EEG, or EMG records (cardiac or neuro monitoring). | Device Design | Vital Signs, Inc. issued Product Safety Notification on March 23, 2009 informing users of potential EMI and fluid ingress issues. All warmers in the field will be replaced beginning in May 2009. |  | 3052 | 3052 units | Worldwide Distribution -- USA, Philippines, Saudi Arabia, India, Israel, Australia, Mexico, Germany, Republic of Singapore, China, Guatemala, Switzerland, Dubai United Arab Emirates, New Zealand, Malaysia, India, Ecuador, &amp; United Kingdom. | Terminated | May 02, 2012 | 1136 | 3052 | Computer
Z-1558-2009 | 52019 | Yes | MSX | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Datex-Ohmeda S/5 Device Interfacing Solution, N-DISVENT 02. |  |  | May 08, 2009 | 2009 | GE Medical Systems, LLC | GE Healthcare has recently become aware of an issue in the calculations associated with Datex-Ohmeda S/5 Device Interfacing Solution N-DISVENT-02 interfaced between certain Datex-Ohmeda patient monitors (Datex-Ohmeda S/5 Anesthesia Monitor (AM), Datex-Ohmeda S/5 Critical Care Monitor (CCM), Datex-Ohmeda S/5 Compact Anesthesia Monitor (CAM), Datex-Ohmeda S/5 Compact Critical Care Monitor (CCCM)) an | Software design | An "Urgent Medical Device Correction" letter was issued to consignees on or about April 9, 2009. The letter was addressed to Healthcare Administrator/Risk Manager, Chief of Nursing, Director of Biomedical Engineering. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information. A representative from GE Healthcare will contact customers to schedule implementation of the correction when it is available for installation at the customer's facility. Direct questions regarding this notice to GE Healthcare Technical Support or your local GE Distributor and refer to FMI 36042. |  | 2638 | 2,638 | Worldwide Distribution -- US (including DC and states of AL, AK, AZ, AR, CA, CO, FL, IL, IN, IA, KS, KY, LA, ME, MD, MI,MN, MO,MT, NE, NH, NJ, NY, NC, OH, OK OR, PA, SC, TX, UT,VT, VA, WA, WV and WI) and countries of AUSTRALIA, AUSTRIA, BAHRAIN, BELGIUM, CANADA, CHINA, CYPRUS, DENMARK, ESTONIA, FRANCE, GERMANY, INDIA, ITALY, JAPAN, JORDAN, KENYA, REPUBLIC OF KOREA, KUWAIT, LIBYAN ARAB JAMAHIRIYA, LITHUANIA, NETHERLANDS, NEW ZEALAND, NORWAY, PAKISTAN, PHILIPPINES, PORTUGAL SAUDI ARABIA, SERBIA, SOUTH AFRICA, SWEDEN SWITZERLAND, THAILAND, TURKEY, UKRAINE, UNITED ARAB EMIRATES and UNITED KINGDOM. | Terminated | February 04, 2012 | 1002 | 2638 | Computer
Z-2029-2009 | 50785 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Newport e360 Ventilator, Rx only. |  |  | December 15, 2008 | 2008 | Newport Medical Instruments Inc | In December 2008, Newport initiated a field correction to upgrade e360 ventilators with software versions 3.0 or higher to software version 3.3. As a result of this upgrade, some customers experienced unexpected device alert alarms. Newport's investigation into these complaints confirmed that the 3.3 software responds too quickly to certain situations, giving a false device alert. Newport condu | Software design | Newport Medical initally sent out a Field Correction Notice to a limited number consignees dated December 17, 2008 and then the notice was expanded to include all e360 consignees dated June 15, 2009. The firm notified consignees that they have revised the e360 software to version 3.8. A customer verification form was attached to verify upgrade completion. For further questions, contact Newport Customer Technical Support Department at 1-800-451-3111 or +1-714-427-5811 Extension 500. |  | 4117 | 4,117 units (214 Domestically; 3,903 Internationally) | Worldwide Distribution -- US (including AL, AZ, CA, CO, FL, GA, IL, IN, KY, MA, MD, ME, MI, NJ, NM, NY, OH, OK, OR, PA, UT and VA), Albania, Algeria, Argentina, Armenia, Australia, Austria, Bangladesh, Belgium, Bolivia, Brazil, Brunei, Bulgaria, Canada, Chile, Colombia, Czech Republic, Dominican Republic, Ecuador, Egypt, England, Finland, Germany, Greece, Guatemala, Hong Kong, India, Indonesia, Japan, Jordan, Kazahkstan, Kenya, Korea, Kuwait, Lebanon, Liechtenstein, Lithuania, Malaysia, Mauritius, Mexico, Morocco, Nambia, Nepal, New Zealand, Nicaragua, Nigeria, Pakistan, Palestine, Panama, Paraguay, Peru, Philippines, Poland, Puerto Rico, Russia, Saudi Arabia, Singapore, Slovenia, South Africa, Sri Lanka, Taiwan, Thailand, Tunisia, Turkey, AUE, Ukraine, United Kingdom, Yemen Republic and Zimbabwe. | Terminated | November 04, 2011 | 1054 | 4117 | Computer
Z-2048-2009 | 52993 | No | HLI | Ophthalmoscope. | Ophthalmic | Ophthalmic | 510(K) Exempt | Stratus OCT Model 3000 instruments distributed with version 6.0.0 software, product number 2660021119710 (new instrument) or 2660022119710 (refurbished instrument), and Stratus 6.0.0 Software Kit, product number 2660021133164. |  |  | July 31, 2009 | 2009 | Carl Zeiss Meditec Inc | Mis-labeled image. A bug in version 6.0.0 software and 6.0.0 software kit used with Stratus OCT Model 3000 can produce a report that incorrectly labels OD or OS. | Software design | Firm initiated notification to customers on July 2, 2009, by sending a letter via FedEx to US customers, and via e-mail and FedEx to outside US consignees. Customers were informed of the issue and steps needed to correct it. A revised software version will be provided to the customers. Any questions should be directed to Carl Zeiss Meditec, Inc. Customer Care at 1-800-341-6968. |  | 244 | 244 units | Worldwide - US, Austria, UK, Japan, Singapore, Argentina, France, Hong Kong, New Zealand, Canada, Colombia, Germany, Korea, Venezuela, India and China. | Terminated | April 19, 2010 | 262 | 244 | Computer
Z-2004-2009 | 52724 | No | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | Roche CoaguChek XS System, CoaguChek XS Plus Care Kit, Roche Diagnostics, Indianapolis, IN; Model Number 05021537001. |  |  | July 22, 2009 | 2009 | Roche Diagnostics Operations, inc. | The recommended cleaning procedures for the meter may cause a malfunction and falsely elevated test result reports. | Device Design | Roche Diagnostics Operations, Inc. notified Consignees by letter dated July 22, 2009 informing them of the problem and instructing them to use revised cleaning procedures.For further questions, contact Roche Diagnostics Operations, Inc. at 1-317-521-3194. |  | 1327 | 1,327 in US | Worldwide Distribution | Terminated | April 27, 2011 | 644 | 66326 | Computer
Z-1444-2009 | 51704 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Sybaritic, NannoLight MP50. |  |  | April 12, 2008 | 2008 | Sybaritic, Inc | Sybaritic Inc. had an issue with the 585nm filter used with the NannoLight MP50. The filter may cause unwanted burning and should not be used for vascular lesion treatments. | Device Design | A "Sybaritic" notification letter dated April 15, 2008 was issued to consignees. The letter described the modifications to the protocols for the NannoLight MP50 and the temporary replacement of 585 Filter for the 560 Filter for Vascular Lesions until further notice. They requested the immediate use of the protocols replacing all other versions.A follow-up letter dated June 3, 2008 was issued to customers. The letter included an attachment " Product Notice Nannolight MP50 585 nm "special" filter from Sybaritic, Inc. The letter described the new 585 nm "bandwidth filter". Also the notification advised consignees to discontinue the use of the old 585 filter and to use 560nm filter until they receive the new 585 nm bandwidth filter.Please direct any questions to your Account Manager or to Sybaritic, Inc. by calling 1-800-445-8418 or 1-952-888-8282. |  | 59 | 59 | Worldwide Distribution -- US (states of WA, CA, MN, FL and AR), POLAND, THAILAND, CHILE, CANADA, KOREA, AZERBAIJAN, INDIA, SPAIN, ARGENTINA, RUSSIA, GREECE, PHILIPPINES, POLAND, INDIA, JORDAN, TURKEY, MEXICO, URUGUAY,BELGIUM, LEBANON, BULGARIA and SINGAPORE. | Terminated | July 05, 2012 | 1545 | 59 | Computer
Z-1889-2009 | 52679 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | NX 2008 Central Monitoring System (CMS). Product code: E2FB6 |  |  | June 22, 2009 | 2009 | AGFA Corp. | After opening and closing a study on a Central Monitoring System while a study is still open on an in-room NX system, a problem of image mix-up can occur on the in-room NX system. | Software design | An "Urgent Safety Notice" was sent to all potentially affected sites on/about June 22, 2009 to make them aware of the issue. A FAX Back form is requested with each letter to ensure someone at the account acknowledges the information was received and understood. If a FAX Back form is not received, the account will be contacted by telephone. A software service update to solve this issue is expected to be released by August 15, 2009. A Mandatory Service Bulletin will be deployed to correct the problem in the software. If you have questions, please contact Agfa Healthcare at 1-800-581-2432 prompts 1, 3 and then 2 for CR Support. |  | 97 | 97 units | Worldwide - IL, PA, TX, Canada, Europe, Asia Pacific and Latin America | Terminated | July 22, 2010 | 395 | 97 | Computer
Z-1951-2009 | 52803 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | PowerProMax PRO5100M Battery Single Trigger, |  |  | July 31, 2009 | 2009 | Linvatec Corp. dba ConMed Linvatec | There is a potential the devices may self activate for their PowerPro, PowerProMax, and MPower I Handpieces manufactured prior to June 1, 2008. | Device Design | Consignees were notified via letter by FedEx delivery or equivalent method. If the product has been further distributed the consignee is requested to notify their consignees. The letter indicated the following risk mitigation actions until affected handpieces are returned to the manufacturer for preventive maintenance: 1) Apply supplied label to the identified handpieces by specific serial number, 2) If self-activation is experienced, discontinue use immediately and return to ConMed Linvtec and 3) Advise/remind the consignee to follow the manufacturers preventive maintenance intervals specified in the instructions for use manuals. Direct questions to ConMed Linvatec at 1-800-237-0169.Recall Expansion: New amended customer and distributor letters will be sent to those identified customer/distributor accounts. The letter will address the additional units. |  | 5553 | 5,553 (4,235 INITIAL AND 1,318 EXPANDED) | Worldwide Distribution including United States. | Terminated | December 19, 2011 | 871 | 23426 | Computer
Z-1995-2009 | 51278 | No | MLD | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Datex-Ohmeda S/5 Compact Critical Care Monitor with L-CICU01, L-CICU02A, L-ICU05 &amp; L-ICUO5A software. |  |  | February 06, 2006 | 2006 | GE Medical Systems Information Tecnology | GE Healthcare (Datex-Ohmeda) S/5 Critical Care Monitor and S/5 Compact Critical Care Monitor with L-CICU01, L-CICU02A, L-ICU05 &amp; L-ICUO5A software had an issue where the Heart Rate (HR) was displaying as ---(3 dashes) when the monitor was used continuously for 49 days without a cold start. A missed alarm or missed HR value could cause a delay in recognition of a cardiac event that may result i | Software design | GE Healthcare Technologies issued a Marketing Bulletin "Ref: MB MS HKI 06007" dated February 6, 2006 to explain a software update. |  | 507 | 507 | ITALY, AUSTRALIA, AUSTRIA, BELGIUM, BOLIVIA, BRASIL, BULGARIA, CHILE, CHINA, COLOMBIA, COSTA RICA, CYPRUS, CZECH REPUBLIC, DENMARK, EGYPT, ESTONIA, FINLAND, FRANCE, GERMANY, GREECE, HONG KONG, HUNGARY, INDIA, IRAN, ISRAEL JORDAN, LITHUANIA, ITALY, MACEDONIA, MALAYSIA, MEXICO, NETHERLAND, NICARAGUA, NORWAY, POLAND, PORTUGAL PUERTO RICO, RUSSIA, SINGAPORE, SLOVAK IA, SOUTH AFRICA, SPAIN, SWEDEN, SYRIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, VENEZUELA. | Terminated | May 17, 2012 | 2292 | 2760 | Computer
Z-1990-2009 | 51999 | No | FNG | AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | BHM Combi Sling, 100% Polyester Shell/100% Polyester Strap; soft polyester net fabric sling for ceiling and floor patient lifts; BHM, Magog, Canada, made in China. |  |  | April 07, 2009 | 2009 | B.H.M. Medical, Inc. | Premature failure of the stitching at the junction of the shoulder strap and sling body of the BHM/Medi-Man Combi Sling used with patient lifts. | Device Design | Arjo sent "Urgent Device Recall Customer Notification" cover letters dated April 15, 2009 to the end user accounts who received the affected BHM/Medi-Man Combi Slings, to the attention of the Administrator/Risk Manager, advising them of the potential for possible stitching failure that could lead to a hazardous situation. Enclosed with the letter was a copy of the BHM Medical Field Safety Notice, FSN 10/2009/ca, dated April 3, 2009 which was sent to their BHM customers in the U.S. on April 7, 2009. The FSN identified the part numbers of the BHM/Medi-Man Combi Slings, described the potential health hazard, and listed the actions to take. The accounts were instructed to inspect and assess their slings' suitability for use per the Slings Maintenance Instructions. If the sling is showing this kind of breakage, the account was instructed to remove it from use and discard it immediately. If the inspection reveals no deterioration of the stitching, the sling is safe to use. The accounts were requested to complete and return the enclosed FSN receipt form within 10 days, acknowledging the receipt and understanding of the letter, and indicating the number of slings in use prior to inspection, the number of slings in use after the inspection, and the number of slings rejected. The accounts will be contacted for the replacement of their eligible old slings (in use up to two years). |  | 1396 | 1,396 slings | Nationwide | Terminated | January 07, 2011 | 640 | 1396 | Computer
Z-1929-2009 | 52055 | No | KEY | Automated slide stainer. | Pathology | Pathology | 510(K) Exempt | VP 2000 Processor; an automated tissue stainer |  |  | September 15, 2008 | 2008 | Abbott Molecular | Thermal runaway conditions, in that the target temperature was exceeded and the temperature display could not track the actual temperature within the heated slide basins of the VP 2000 Slide Processor. | Device Design | Abbott sent Device Recall/Urgent Field Safety Notice letters dated 9/15/08 to the VP2000 customers on 9/12/08, informing them that thermal runaway conditions may occur within the heated slide basins of the VP2000 Slide Processor. The customers were given actions to take to prevent this condition: a) VP2000 should be powered down when not in use and when the controlling computer is not in use; b) Liquid level in all heated basins should be checked periodically to ensure that there is adequate reagent present for slide processing and to keep the sensor probe immersed for accurate temperature sensing; and c) If large, rapid and random fluctuations in indicated basin temperature occurs the unit should be shut down and a Field Service Representative be contacted. Any questions were directed to their local Abbott Technical Service Representative. |  | 261 | 261 units | Worldwide Distribution -- USA, Australia, Austria, Belgium, Canada, Columbia, Czech Republic, Egypt, France, Germany, Hong Kong, India, Ireland, Italy, Japan, Korea, Malaysia, Norway, Russia, Saudi Arabia, Singapore, Switzerland and the United Kingdom. | Terminated | January 07, 2010 | 479 | 261 | Computer
Z-1945-2009 | 52126 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Maxilift Patient Lift, Non-AC Powered; a patient lift with a four point tilting frame, adjustable chassis, and optional electronic scale, to move patients in a hospital/nursing home situation; Arjo Inc., Roselle, IL 60172; Model numbers 21102406, 21200001, 21200006, 21000-01, 210024-06, 211010-06, 211024-06, 211026, 212000-06, 212000-06R, 212000-6, 212024-06, 212026, 212624, 218100-01, KGA0200, KMA0520, KMA1003-04, KMA1004, KMA1005, KMA1006, KMA1040, KMA1501-06, KMA1513-06, KMA211024, KMA82124-24, KMB2351, KMB4800-06, KMB4900-06, KMB5501-12US, KRA0300, KRA0310, KRA1000, KRA1500-12, KRA2100, KRA3000-12, MA0500, MA0510, MA1010, MA1020, MA1100, MA1500, Maxi Pillar, Maxilift, Maxipillar, MAXIPILLER, MB0500R, MB0510, and MB0610. |  |  | April 27, 2009 | 2009 | ArjoHuntleigh | The roll pin on certain configurations of the Maxilift hanger bar could break due to a combination of severe corrosion and fatigue. Breakage of the roll pin would lead to the detachment of the hanger bar from the jib of the lifter. | Equipment maintenance | ArjoHuntleigh sent a Field Safety Notice DCHU-UK/0209/001FSN to all affected end users via UPS Ground on 4/27/09. The notice informed the accounts of the potential for the roll pin to break due to a combination of severe corrosion and fatigue, leading to the detachment of the hanger bar from the jib of the lift. The accounts were advised to identify if they have any Arjo Maxilift patient lifters in their facility, identify which version of Arjo Maxilift hanger bar they have using the photographs in Figures 1,2 and 3 in the letter. Any lifters found with a hanger bar attachment as shown in Figure 1 must be immediately and permanently taken out of use, as the replacement hanger bar assembly is no longer available. Any lifters with a hanger bar attachment as shown in Figure 2 must be taken out of use until a correction has been performed by ArjoHuntleigh. Any lifters with a hanger bar attachment design as shown in Figure 3 are not affected and can continue being used. The accounts were provided with a product upgrade offer and a MaxiLift Preventative Maintenance Schedule. The accounts were requested to complete and fax, e-mail or mail the enclosed Customer Response Form to Arjo at 630-925-7974 or quality@arjousa.com, indicating the number of Maxilifts in each category that they have and that the products have been taken out of service for replacement or correction, listing the model numbers and serial numbers of their lifts. Questions were directed to the ArjoHuntleigh Customer Support or Quality Department at 1-800-323-1245. |  | 8015 | 8,015 lifts | Worldwide Distribution -- Countries of USA and Mexico. | Terminated | February 25, 2011 | 669 | 8015 | Computer
Z-1729-2009 | 52636 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Clinac with Version 7.x Software, manufactured by Varian Medical Systems, Palo Alto, CA. |  |  | May 29, 2009 | 2009 | Varian Medical Systems Oncology Systems | Unexpected Movement: if the stereotactic motion disable function is turned on, the couch can be moved via float mode unexpectedly. | Device Design | Varian Medical Systems issued an "Urgent Medical Device Correction" dated May 29, 2009 distributed to all affected users, as well as Varian Sales, Marketing and Service organizations, informing them of the issue. Consignees were informed that a Varian Service representative will be in contact to schedule the proper correction.For further questions, contact your local Varian Customer Support District Manager. |  | 1933 | 1,933 units | Worldwide Distribution. | Terminated | May 18, 2011 | 719 | 1933 | Computer
Z-2051-2009 | 52811 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | REF IA-2000-S, LightWave Suction Ablator, STERILE, Rx ONLY, MADE IN USA. |  |  | July 31, 2009 | 2009 | Linvatec Corp. dba ConMed Linvatec | ConMed Linvatec has determined there is a possibility that the device may not shut off if the ablate button is released, which may result in users/patients suffering burns. | Device Design | ConMed Linvatec is notifying their direct consignees by an Urgent Medical Device Recall letter, requesting that customers check their inventory for the affected products; complete, sign, and return the Reply Form; and return of any affected lot numbers of the product. The product is to be returned to ConMed Linvatec. Questions should be directed to Customer Service at 1-800-237-0169. |  | 28074 | 28,074 units | Worldwide Distribution | Terminated | February 25, 2011 | 574 | 28074 | Computer
Z-2073-2009 | 53129 | No | LKK |  |  | General Hospital | PMA | Medtronic Indura 1P Intrathecal Catheter, model 8709SC, Method of Sterilization: Ethylene Oxide. Medtronic, Inc., Minneapolis, MN 55432-5604 USA. The implanted infusion system components consist of a Medtronic pump and a model 8709SC Catheter. The catheter connects to the pump at the catheter port. Contents: 81.4-cm Catheter with length markers and guide wire, 15 T-guage introducer needle, Sutureless pump connector with attached 7.6-cm catheter and connector pin, Transparent strain-relief sleeves, Anchors. |  |  | August 27, 2009 | 2009 | Medtronic Neuromodulation | Current labeling for Medtronic Sutureless Connector (SC) catheters and revision kits (hereafter referred to as SC catheters ) incorrectly states that SC catheters are compatible with Medtronic IsoMedÂż constant-flow infusion pumps. SC catheters are compatible with Medtronic SynchroMedÂż II and SynchroMed EL pumps. Medtronic has determined that SC catheters are not compatible with IsoMed pumps. | Device Design | A "Medical Device Correction Notification" letter was issued on August 14, 2009 to implanting and managing physicians. The letter described the nature of the issue, scope of affected product, potential severity of the issue and provided recommendations. Physicians were instructed that Sutureless Connector (SC) catheters must not be used with IsoMed pumps for future revisions and implants. A copy of the letter was also sent to the Risk Manager at each hospital associated with physicians using SC catheters. Healthcare professionals were asked to complete and return the enclosed reply card.Direct questions to your Medtronic field representative or contact Medtronic Neuromodulation Technical Services at 1-800-707-0933. Patient Management information is also available at http://www.professional.medtronic.com under the Heading, "Advisories". |  | 20000 | 20,000 of all products | Worldwide Distribution -- US (including Washington D.C. and Puerto Rico), Algeria, Aruba, Australia, Austria, Belgium, Canada, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Egypt, Finland, France, Germany, Greece, Hong Kong, Hungary, Iceland, India, Ireland, Israel, Italy, Jordan, Kuwait, Lebanon, Luxembourg, Malta, Netherlands, Netherlands Antilles, New Caledonia, New Zealand, Norway, Poland, Portugal, Reunion, Romania, San Marino, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom and Vatican City. | Terminated | December 12, 2011 | 837 | 20000 | Computer
Z-1928-2009 | 52118 | No | JJH | Clinical sample concentrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AMPLILINK Software, version 3.1.0 to 3.2.2 with COBAS AmpliPrep. Catalog number 03563383001, catalog number 04855094001, catalog number 04807197001, catalog number 04862392001, catalog number 05093236001, and catalog number 05201357001. |  |  | June 02, 2009 | 2009 | Roche Molecular Systems, Inc. | Software bug: A unique series of events involving the COBAS AmpliPrep instrument running AMPLILINK software version 3.2.2 that results in specimens being amplified and detected with the wrong parameters. The error can occur on either COBAS AmpliPrep/COBAS TaqMan 48, COBAS AmpliPrep/COBAS TaqMan (docked) or cobas s 201 system (docked) running Amplilink Software versions 3.1 or 3.2 series up to and | Software design | Urgent Medical Device Correction Letters were sent on June 2, 2009 by first class mail. The letter described the issue; impact on results; clinical implications; the workaround; and the actions required, which include following the workaround until a new version of AMPLILINK software is available and filing the letter for future reference. Questions or concerns should be directed to Roche Diagnostics Technical Support at 1-800-526-1247. |  | 282 | 282 | Nationwide Distribution | Terminated | August 13, 2010 | 437 | 282 | Computer
Z-1938-2009 | 51281 | No | MOS | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Mirror assembly to magnetic resonance device (1.5T Atlas SPEEDER Head-Neck Coil) Coil Part #Q7000021, Quality Electrodynamics, Mayfield Village, OH 44143. |  |  | December 29, 2008 | 2008 | Quality Electrodynamics Llc | The firm received a complaint indicating that the mirror is high risk and could drop on a patient. | Device Design | Quality Electrodynamics, LLC notified Customers via email beginning December 29, 2008 requesting the affected device be returned to the firm for improvement.For further information, contact Quality Electrodynamics, LLC at 1-440-484-2228. |  | 5 | 5 units | Nationwide Distribution: IL and OH. | Terminated | April 25, 2012 | 1213 | 5 | Computer
Z-2316-2009 | 52858 | No | LMB | Medical image storage device. | Radiology | Radiology | 510(K) Exempt | Centricity Enterprise Archive (EA) versions 3.0 to 3.0.7 software; GE Healthcare Integrated IT Solutions, Barrington, IL 60010. |  |  | August 14, 2009 | 2009 | GE Healthcare Integrated IT Solutions | Software error: There is potential safety issue with Centricity Enterprise Archive (EA) 3.0.x software where study split operations are not correctly replicated to a secondary 'shadow' archive. | Software design | GE Healthcare Integrated IT Solutions sent Urgent Medical Device Correction letters dated August 7, 2009 to the attention to the Director of Radiology, the Hospital Administrator, the Head of the Radiology Department and the PACS Administrator. The letters advised users of the potential safety issue and provided safety instructions to follow until the software is updated. Users were asked to complete the enclosed Customer Reply form and fax the completed form to 1-847-939-1479. Any questions should be directed to the Remote Online Center or the GE Customer Care Center at 1-800-437-1171. |  | 212 | 212 units | Worldwide Distribution -- United States, Belgium, Canada, Cayman Islands, China, Denmark, France, Germany, Great Britain, Hong Kong, India, Israel, Italy, Korea, Malaysia, Malta, Mexico, New Zealand, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Taiwan and Turkey. | Terminated | February 25, 2011 | 560 | 212 | Computer
Z-1888-2009 | 52723 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARTISTE MV Systems (Part Number 8139789), equipped with a syngo RT Therapist (Part Number 8162815) or RT Therapist Assist system (Part Number 8162807), with Software version NET VA10A for RTT4.1.122 or RTT 4.1.125, manufactured by Siemens Medical Solutions, Concord, CA. |  |  | April 09, 2009 | 2009 | Siemens Medical Solutions USA, Inc | Unexpected treatment plan alteration-- Software issue may result in unintended modifications to treatment | Software design | Siemens Medical Solutions USA, Inc. notified consignees of the affected product by advisory letter via Certified Mail on April 09, 2009.For further information, contact Siemens Medical Solutions at 1-925-602-8157. |  | 53 | 53 units | Worldwide Distribution -- US (MA, UT, OH, NJ, NE, FL, OH, MI and PA), Germany, Poland, the UK, Ireland, Italy, France, South Korea, and Belgium. | Terminated | April 12, 2011 | 733 | 53 | Computer
Z-2231-2009 | 53045 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Medical Systems, Eclipse Proton Convolution Superposition DCF, version 8.6.15, Model H48, radiation therapy treatment planning system. |  |  | July 17, 2009 | 2009 | Varian Medical Systems Oncology Systems | Software Issue/Mistreatment Potential-- Due to software issues, dose distribution may be incorrect and unintended radiation doses may be delivered. | Software design | Varian Medical Systems Recall initiated on July 17, 2009. A product notification letter was distributed to the affected user, with a description of the problem and user corrective action steps. The letter will be distributed to the sales, marketing and service organizations via certified mail. A software correction is planned. |  | 1 | 1 | Product was distributed to one user in Texas | Terminated | November 27, 2010 | 498 | 1 | Computer
Z-2239-2009 | 53132 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | ILUMAVision Software, version 2.1.0 Picture Archiving and Communication system |  |  | August 03, 2009 | 2009 | IMTEC IMAGING L.L.C. | Under certain circumstances, software Version 2.1.0 has potential for incorrect data imaging and patient identification. | Software design | Firm sent notifications to all ILUMAVision customers on 8/3/2009 and offered upgraded software. The notification letter indicated the the reason for the recall and identified the potential issues associated with the device. Contact 3M IMTEC Technical Support at 1.800.879.9799, if there are any questions. |  | 223 | 223 | Worldwide - US - Foreign: Australia, Austria, Belgium, Bulgaria, Canada, Croatia, Egypt, France, Germany, Hong Kong, Italy, Korea, Malaysia, Mexico, The Netherlands, Poland, Saudi Arabia, Spain, Turkey, and United Kingdom | Terminated | January 27, 2010 | 177 | 223 | Computer
Z-0003-2009 | 53158 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Revolution XR/d systems with control room PC model 5117866-2 with Software collector 5194381. GE Healthcare, 3000 North Grandview, Waukesha, WI 53188. |  |  | February 14, 2008 | 2008 | GE Medical Systems, LLC | GE Healthcare has recently become aware of an error in the generator software on the Revolution XR/dconfigured with HP 8200 PC consoles that reports a mAs reading that is higher than is actually exposed to patients. | Software design | Consignees were sent a GE Healthcare "Product Safety Notification" letter dated March 7, 2008. The letter was addressed to Hospital Administrators, Mangers of Radiology, and Radiologists. The letter listed the Safety Issue, Affected Product, Safety Instructions, and Contact Information. |  | 83 | 83 (63 USA, 20 OUS) | Worldwide Distribution -- USA, including states of AL, AZ, CA, CO, FL, ID, IL, IN, LA, MA, MI, MO, NY, NC, OH, OR, PA, SC, SD, TN, TX, WA, and WI and countries of Canada, China, Japan, Korea, and Taiwan. | Terminated | December 18, 2011 | 1403 | 83 | Computer
Z-2243-2009 | 52729 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | TomoTherapy Hi-Art SystemÂż, Version 3.X |  |  | August 10, 2009 | 2009 | TomoTherapy Incorporated | TomoTherapy received one customer report indicating that after starting a patient procedure from the StatusConsole, the radiation on light and the audible indicator on the Status Console came on as expected,however the Operator Station (OS) did not display the dose rate and the procedure time, or the couch position updates. Within a few minutes, the customer pressed the Stop button on the Sta | Software design | Consignees were sent a TomoTherapy "Urgent Medical Device Correction Field Safety Notice" letter dated August 10, 2009. The letter addressed the Issue, Product Affected, Investigation, Recommended Actions and Resolution.Customers were recommended to perform the following actions (in the event that the dose rate, procedure timer, and couch position updates are not present on the OS, after pressing Start on the status console and the radiation on light and audible indicators are active):1) Immediately press the Stop button on the Status Console and 2) contact your authorized TomoTherapy Service Representative for further direction. The affected issue will be corrected in the next software version, which is scheduled for release in early 2010. The TomoTherapy Customer Contact Center can be reached 24 hours a day, 7 days a week, by e-mail atsupport@tomotherapy.com or by phone at 1-866-368-4807 (North America). |  | 259 | 259 (169 US; 91 OUS) | Worldwide Distribution -- US (states of AL, AK, AZ, AR, CA, CO, CT, DC, FL, GA, HI, IL IN, IA, KS, KY, LA, MD, MA, MI, MN,MO, NE,NV,NJ, NM, NY, OH, OK OR, PA, PR, SC, SC, TN, TX, UT, VA, WA, WI and WY) andBelgiumCanadaSwitzerlandChinaGermany Spain FranceEnglandHong KongIndiaItaly Japan KoreaNetherlandsPolandSaudi ArabiaSwedenSingaporeTurkey andTaiwan. | Terminated | December 17, 2011 | 859 | 259 | Computer
Z-2252-2009 | 53207 | No | LZH | Infusion pump. | General Hospital | General Hospital | 510(k) | Flexiflo Quantum Enteral Pump Model # 50596, 50598 and 59322. |  |  | September 04, 2009 | 2009 | Abbott Nutrition, Abbott Laboratories | The firm initiated this voluntary recall because some of its Flexiflo Quantum Enteral Pumps, have a 110-volt AC power cord plug that may crack or fail. This potential power cord plug failure may cause electrical shock and/or a delay in feeding that could represent a potential health and safety risk. | Device Design | An "Urgent Device Recall" letter dated September 4, 2009 was issued via FedEx overnight delivery to all direct consignees. The letter described the affected product, issue and immediate actions for customers. Direct consignees were further instructed to immediately contact affected customers and notify them about the Urgent Device Recall. Abbott Nutrition (AN) Device Call Center (DCC) will telephone those customers who do not responded promptly to the "Urgent Device Recall" letter with the completed AN Account Summary of the Flexiflo Quantum Enteral Pump Shipments. AN Device Call Center will make arrangements for the service or replacement of the recalled pumps. Please contact the AN Device Call Center at 1-877-457-0249, Monday through Friday between the hours of 8:30 am and 5:00 pm Eastern Standard Time (EST). An AN representative will work with you to arrange service or replacement of the recalled enteral pumps. |  | 5173 | 5,173 pumps | Nationwide Distribution (states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV and WY). | Terminated | September 19, 2012 | 1111 | 5173 | Computer
Z-2232-2009 | 52959 | No | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Oridion Medical Capnostream20 capnograph/pulse oximeter monitor and monitor with printer. Part numbers CS 08651, CS 08653, CS08657, and CS08659. |  |  | June 30, 2009 | 2009 | Oridion Medical 1987 Ltd. | 1) The monitor's ability to withstand electrostatic discharge (ESD) to exposed metal connectors on the back of the monitor which may cause the monitor's display to freeze.2) The unit generates electromagnetic radiation above the limits set by applicable standards for medical devices | Device Design | Oridian sent a letter to all its distributors June 25, 2009, instructing them to gather all the units to be exchanged for factory reconditioned units that have the hardware and software enhancements and meet all performance requirements in full. |  | 326 | 326 units | Nationwide. | Terminated | October 01, 2010 | 458 | 326 | Computer
Z-2112-2009 | 50687 | No | LZH | Infusion pump. | General Hospital | General Hospital | 510(k) | EnteraLite Infinity Enteral Feeding Pump, REF INFP01, Zevex Enteral Nutrition Delivery Systems, Salt Lake City, UT 84123. |  |  | December 18, 2008 | 2008 | Zevex International, Inc. | Enteral feeding pumps may not accurately deliver fluid, potentially under-infusing or over-infusing. | Software design (manufacturing process) | All of Zevex' customers and distributors were notified by an Urgent Medical Device Correction letter, dated 12/18/2008. The letters explained the issue, the patient impact, and what necessary actions should be taken. Customers were instructed to contact Zevex to schedule the upgrade of their installed base of pumps at 800-970-2337 or sales@zevex.com. Distributors were to contact their customers and notify them of the letter. |  | 32545 | 32,545 | Nationwide, including 1 VA facility in CA. Foreign distribution to 1 consignee in Canada. | Terminated | October 27, 2011 | 1043 | 32863 | Computer
Z-2110-2009 | 53087 | No | DXN | Noninvasive blood pressure measurement system. | Cardiovascular | Cardiovascular | 510(k) | Biphasic LIFEPAK 15 Monitor/Defibrillator. |  |  | August 12, 2009 | 2009 | Physio Control, Inc. | Some LIFEPAK 15 Monitor/Defibrillators have a Non-Invasive Blood Pressure (NIBP) module that was supplied by CAS Medical Systems, Branford, Connecticut who is recalling the modules. The NIBP modules may be susceptible to sudden changes in cuff pressure. Sudden changes in cuff pressure may possibly be caused by bumping or squeezing thecuff. In some instances this may result in a screen display | Software design | The firm notified customers by letter dated August 2009. Customers will have the Noninvasive Blood Pressure (NIBP) module replaced. If additional questions about the notification, contact Technical Support by calling 1-800-442-1142, option 5, Monday-Friday 6:00 a.m. to 4:00 p.m. (Pacific Standard Time).. |  | 445 | 445 devices worldwide | Worldwide Distribution -- US, AUSTRALIA, GERMANY and NETHERLANDS. | Terminated | August 10, 2012 | 1094 | 445 | Computer
Z-2213-2009 | 52230 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Stryker Epic II critical care bed, 115v, Stryker Medical, Portage, MI; Model 2030. |  |  | August 10, 2009 | 2009 | Stryker Medical Div. of Stryker Corporation | The backrest of the bed may not stay in an upright position (may drift down when left raised). | Device Design | Stryker Medical issued an "Urgent - Medical Device Correction" letter dated August 10, 2009 informing customers to locate the affected product listed and verify whether it is functioning properly. In the case the device is not properly functioning, remove from service and contact the firm as soon as possible at 1-800-STRYKER, option 3. Customers are also asked to return an enclosed return receipt post card and also forward the notification to users who have loaned or purchased the device. A representative from Stryker field service will be contacting you to perform device corrections shortly.For further information, contact Stryker Medical at 1-269-324-6527. |  | 12599 | 12,599 total products | Worldwide Distribution -- United States, Brazil, Canada, Chile, China, India, Korea, Latin America, Mexico, Singapore and Switzerland. | Terminated | April 25, 2011 | 623 | 12599 | Computer
Z-2221-2009 | 50488 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ACCELERATOR APS System Input/Output Module (IOM); List Number: 7L01-01; manufactured by Inpecco SpA, Segrate (MI), Italy. |  |  | October 27, 2008 | 2008 | Abbott Laboratories, Inc | When an ARCHITECT c16000 System is connected to an ACCELERATOR APS System, the software on the ACCELERATOR APS System may not generate an error warning when a Sample Presentation/Sample Queue error occurs, which could result in contamination of subsequent samples. | Software design | Abbott Laboratories notified Customers of the affected device via a "Product Correction" letter dated October 27, 2008. Consignees provided with a temporary work around until a service representative upgrades the system software.For further information, contact your local Abbott Laboratories, Inc. service representative. |  | 15 | 15 units | International Distribution -- Spain, Italy, the United Kingdom, Sweden, Switzerland, the Netherlands, France, Canada and Scotland. | Terminated | December 02, 2009 | 401 | 15 | Computer
Z-2222-2009 | 51275 | No | LOJ | Tumor-associated antigen immunological test system. | Immunology | Immunology | 510(k) | Abbott AxSYM Cancer Assay Disk Version 5.0, for in vitro diagnostic use with the AxSYM System; Abbott Laboratories, Abbott Park, IL 60064; List Number 3D50-05. |  |  | February 20, 2009 | 2009 | Abbott Laboratories | The AxSYM Cancer Assay Disk (LN 3D50-05 version 5.0) contains an AxSYM AFP assay file where the Low Concentration flag for the AFP 1:101 automated dilution protocol the instrument to report results between 0 ng/mL and 0.4 ng/mL (40.4 ng/mL accounting for the 1:101 dilution). The sensitivity of the assay is 0.4 ng/mL. | Device Design | Abbott sent "Product Correction Immediate Action Required" letters dated February 2, 2009 to all AxSYM AFP customers. Customers were requested to install the updated version of the AxSYM AFP assay using version 6.0 of the AxSYM Cancer Assay Disk provided with the letter and to retain a copy of this communication for their records. Any questions regarding the letter were directed to Customer Service at 1-877-4ABBOTT. |  | 8205 | 8,205 disks | Worldwide Distribution -- United States (including Puerto Rico), Australia, Bahamas, Barbados, Bermuda, Brazil, Canada, Cayman Islands, Chile, Columbia, Costa Rica, Curacao (Neth Ant), Dominican Republic, Germany, Guatemala, Hong Kong, Jamaica, Japan, Mexico, New Zealand, Peru, Singapore, South Korea, Taiwan, Thailand, Trinidad and Tobago, Turks and Caicos and Venezuela. | Terminated | August 23, 2010 | 549 | 8205 | Computer
Z-2223-2009 | 52218 | No | JJQ | Colorimeter, photometer, or spectrophotometer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | BioTek Gen5CL Microplate Data Collection and Analysis Software. |  |  | March 12, 2009 | 2009 | BioTek Instruments, Inc. | Software design allowed false negatives to be read when wells are dark. | Software design | BioTek Instruments, Inc. issued a "Product Recall Notification" dated March 13, 2009 information customers of the affected product. Follow-up letters with software installation instructions were sent March 25, 2009.For further information, contact BioTek Instruments, Inc. at 1-888-451-5171 extension 405. |  | 13 | 13 | Worldwide Distribution -- United States (MN, CA, IL, MA, MI, CA, GA, FL, AZ and TX) and Peru. | Terminated | May 18, 2015 | 2258 | 13 | Computer
Z-2019-2009 | 52862 | No | KGC | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Bard * Button Device Continuous Feeding Tube with 90 degree Adaptor, 18F x 1.7cm, NON STERILE, REF 000282, Bard Access Systems, Inc., Salt Lake City, UT 84116 USA. |  |  | January 20, 2009 | 2009 | C R Bard Inc | Possible restriction/occlusion in adaptors/tube accessories causing difficulty to infuse during routine feeding treatment. | Equipment maintenance | Customers were notified by letter and instructed to not use or further distribute any of the affected lot numbers. Customers were further instructed to return the recalled products to Bard Access Systems (BAS) through a return goods authorization number. Questions and requests for assistance should be directed to BAS Customer Service at 800-290-1689. |  | 4381 | 13, 337 total of all products | Worldwide Distribution -- US, Canada, Austria, Australia, Belgium, Finland, France, Germany, Great Britain, Holland, Italy, Norway, Spain, Sweden and Switzerland, Japan, and Latin America. | Terminated | December 09, 2009 | 323 | 4381 | Computer
Z-1939-2009 | 52769 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | The BladderScan BVI 9400 is a portable ultrasound instrument that provides a noninvasive measurement of urinary bladder volume. The device consists of an ultrasound probe that scans the patient's bladder, and a battery operated console that provides measurement related information. The BVI 9400 measures ultrasonic reflections on multiple planes inside the body and produces a three dimensional image. Based on the image, the BVI 9400 calculates and displays bladder volume. |  |  | May 18, 2009 | 2009 | Verathon, Inc. | Three specific issues are being addressed:1. Calculation and use of an incorrect year in dates after 12/31/2009.2. Double scans, continuous scans, and double printing.3. Calibration sensitivity. | Software design | U.S. customers (including Government) were sent a post card on or about 5/18/2009 and were called about the upgrade. Units are to be upgraded by the customers using an online tool called ScanPoint. Customers may elect to return their units to the firm for upgrading.Canadian customers were phoned by Verathon.For foreign units sold and distributed through one of the recalling firm's related firms, Service Bulletin SB-0014 was sent to the distributors located in The Netherlands, United Kingdom, and Australia.Contact Verathon Customer Care at 1-800-331-2313 if you have any questions. |  | 1978 | 1,978 | Worldwide - US, Canada, Malaysia, Hong Kong, Singapore, Argentina, New Zealand, Barbados, Dominican Republic, and Costa Rica. | Terminated | October 07, 2010 | 507 | 1978 | Computer
Z-2009-2009 | 52766 | No | CAW | Portable oxygen generator. | Anesthesiology | Anesthesiology | 510(k) | Cobham AMOGS Advanced Medical Oxygen Generating System, 70309-03900-101. Concentrator, Oxygen, Molecular Sieve, S/A P/N: 70309-03900-101, Mfg P/N 3261165-0101. The firm name on the label is Carleton Life Support Systems Inc., Davenport, IA. |  |  | July 24, 2009 | 2009 | Carleton Life Support Systems Inc | The backup oxygen supply control valve on the oxygen concentrator could fail resulting in it sticking in the open or closed positions. | Device Design | Cobham Life Support Systems, Inc. issued letters to their two customers via e-mail. One customer letter was dated July 24, 2009 and the other was dated July 28, 2009. Both letters notified the customer of the reason for the recall and that the usage of the concentrator must be restricted to exclude actual use with a patient until corrected. The letter informs customers they are redesigning the backup supply control valve and will coordinate the retrofit in the future. For further information, contact Cobham Life Support at 1-563-383-6421. |  | 23 | 23 units | Nationwide Distribution -- CT and AL. | Terminated | March 22, 2012 | 972 | 23 | Computer
Z-1931-2009 | 52200 | No | DSQ |  |  | Cardiovascular | PMA | iPulse Circulatory Support System |  |  | April 23, 2009 | 2009 | Abiomed, Inc. | iPulse Console SC1035 displayed a "Low Pressure / Low Flow" alarm and stopped pumping during patient transport in battery operation. | Device Design | An "Urgent Voluntary Medical Device Field Correction Notice" was issued April 8, 2009 to affected customers. The customers were notified that the iPulse Console SC1032 displayed a "Low Pressure/ Low Flow" alarm and stopped pumping during patient transport in battery operation. Instructions were provided to users in case they encounter this failure mode. ABIOMED Field Service Engineers will contact customers to schedule a time to visit site and correct affected consoles. Please contact ABIOMED's Field Service at 1.800.554.8666, if you have any questions. |  | 79 | 79 | Domestic: MA, MI, SC, NY, NJ, OH, CA, PA, AR, MI, MD, CT, IL, WI, TX, NJ, TN, IN, MI, DE, FL, IA, WA, KY, and WVForeign: Germany, Greece, Italy, Latvia, Spain, and Turkey | Terminated | June 05, 2014 | 1869 | 79 | Computer
Z-1810-2009 | 52643 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | CyberKnife Treatment Delivery System, a subsystem of the CyberKnife Robotic Radiosurgery System, manufactured by Accuray Inc., Sunnyvale, CA. |  |  | June 22, 2009 | 2009 | Accuray Inc | Targeting accuracy out of specification, Error alert does not render the system down, which may result in mistreatment in the wrong area. | Software design | Accuray, Inc. issued an "Urgent Device Correction" notice dated June 22, 2009. The consignees were informed of the defect including methods for detecting the error. The firm is in the process of developing a software patch to address the issue.For further information, contact Accuray, Inc. Customer Support at +1-877-668-8667 (US) or +1-408-716-4700 (outside US) or customersupport@accuray.com. |  | 61 | 61 units | Worldwide Distribution -- US including states of MA, FL, AL, TN, DE, CA, NJ, OK, CO, NC, IL, PA, MD, DC, MS, CN, NV, RI, SC, UT, VA, MI, NE, MO, TX and AL and countries of Canada, Russia, UK, Ukraine, Saudi Arabia, Switzerland, Turkey, Taiwan, India, China, Korea and Thailand. | Terminated | January 05, 2011 | 562 | 61 | Computer
Z-1885-2009 | 52761 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IntelePACS Picture Archiving Communications System InteleViewer module, Model #3.X.X. |  |  | October 27, 2008 | 2008 | Intelerad Medical Systems Inc. | Product software versions 3.5.2.P1 to 3.5.2.p56 inclusive and 3.6.1.P1 to 3.6.1.P82 inclusive were defective. | Software design | Intelerad Medical Systems issued an Urgent Medical Device Filed Safety Notice" dated October 24, 2008 describing the affected product and steps required for system upgrade. Users are requested to complete and return an acknowledgement form via email to quality@intelerad.com.For further information, contact Intelerad Medical Systems at 1-514-931-6222. |  | 17 | 17 (including 11 free trials) | Worldwide Distribution -- United States (NM, AZ, CT, MN, TX, NC, TN, FL and CO), Canada, Australia, and New Zealand. | Terminated | September 24, 2010 | 697 | 18 | Computer
Z-1809-2009 | 52654 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varis, Aria Radiation Oncology - Version 8.1.15, 8.5.11, and 8.6.07, Model number HIT-INFSYS, Manufactured by Varian Medical Systems Inc., Palo Alto, CA |  |  | May 08, 2009 | 2009 | Varian Medical Systems Oncology Systems | Potential mistreatment-- due to software issues, retired plans may be used and treated repeatedly. | Software design | An Urgent Medical Device Correction Letter/ Field Safety Letter, dated May 8, 2009, were distributed to all affected users, with a description of the problem and user corrective action steps. The Letter will also be distributed to the Varian Sales, Marketing, and Service organizations, informing them of the issue. The firm intends to produce a software fix for the problem. If further clarification is needed, users should contact their Varian Customer Support District or Regional Manager. |  | 1587 | 1,587 units | Worldwide Distribution | Terminated | December 23, 2011 | 959 | 1587 | Computer
Z-1582-2009 | 52214 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance iCT software v2.5.0, Model 728306. The Brilliance iCT is packaged INA carton and the software is contained within the Brilliance CT scanner. |  |  | December 22, 2008 | 2008 | Philips Medical Systems (Cleveland) Inc | Software anomaly. An artifact may appear in brain scans and brain perfusion studies, when using 80kV and a UA, UB, or UC filter with this CT scanning device. | Software design | An "URGENT - Field Safety Notice" dated December 22, 2009 was issued via certified mail to affected customers. The letter described the affected product, problem, hazards involved, and actions for customers to take. Philips will provide customers with a software update for all affected systems within next 6 months. The firm also sent a second notification in the form of an "URGENT -Medical Device Correction" report dated January 29, 2009 which contained information noted in the "URGENT - Field Safety Notice". Direct questions to your local Philips representative or Philips Customer Care Center (1-800-722-9377), option 5: Other Imaging Systems, option 1: CT or to your local Philips Healthcare office. |  | 44 | 44 | Worldwide Distribution -- US including DC (states of CA, GA, IL, IN, KY, NY, MD, MN, OH, OR, and MD), Germany, Greece, Israel, Italy, Japan, Netherlands, Norway, Russia, South Korea, Switzerland, Taiwan, Canada, and Thailand. | Terminated | September 24, 2010 | 641 | 44 | Computer
Z-1674-2009 | 52129 | walker, mechanical | ITJ | Mechanical walker. | Physical Medicine | Physical Medicine | 510(K) Exempt | Essential Medical Supply brand Essential Rollator with Loop Hand Brakes, models W1700B (Blue) and W1700R (Red), Four Wheel Aluminum Walker/Rollator |  |  | May 08, 2009 | 2009 | Essential Medical Supply, Inc. | The caster may separate from the frame completely causing injury to the user. | Device Design | Essential Medical Supply Inc issued a Notice of Quarantine Letter to all customers on 04/30/09, identifying affected products as W1700B &amp; W1700R Essential Rollators, Serial #'s KLD-E07090001 through KLD-E08020570, received after 11/05/2007. Firm issued a Device Recall Letter to all direct customers on 05/08/09 instructing them to do the following: (1) segregate and quarantine any affected product; (2) identify and notify any customers that may received the affected products, instructing them to return the product to their dealer; and (3) either return the affected products to Essential Medical Supply or destroy the products and send proof of destruction. |  | 2265 | 2265 | Nationwide distribution. | Terminated | September 15, 2010 | 495 | 2265 | Computer
Z-1245-2009 | 51934 | No | IMB | Moist steam cabinet. | Physical Medicine | Physical Medicine | 510(K) Exempt | Hydration Station with Pedestal (Part # E30003-01), Hydration Station with Pedestal &amp; DVD (Part # E30003-01D), Sybaritic, Inc., 9229 James Avenue South, Minneapolis, MN 55431. |  |  | May 15, 2007 | 2007 | Sybaritic, Inc | Sybaritic conducted a recall on the Hydration Station 806 because of the potential for hot water to spurt through steam ports into unit and cause burns. Also, potential for customers to inadvertently place their legs over the side wall of the unit directly onto the steam ports and cause burns.There were 18 injuries and 8 malfunctions reported. | Device Design | Sybaritic, Inc. issued a letter dated May 15, 2007 addressed to "Hydration Station Customer". The letter informed the customer of the new 2007 Hydration Station Upgrade Improvement Kits including instructions and advised the customer to coordinate the installation of the components with their Technical Service Department.Further questions can be addressed to Sybaritic, Inc. Technical Administration at 1-800-445-8418 extension 373. |  | 370 | 370 | Worldwide Distribution - US including states of AL, AZ, CA, CO, CT, FL, GA, ID, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MS, MO, MI, NE, NV, NH, NJ, NM, NY, NC, OH, OR, PA, SC, TN, TX, UT, VA, WA, WV and WI and CANADA. | Terminated | July 05, 2012 | 1878 | 370 | Computer
Z-1679-2009 | 51713 | Yes | LWP |  |  | Cardiovascular | PMA | Boston Scientific TELIGEN, Implantable Cardiac Defibrillator models E102, E110, the following models are not available in the US: models F102, F110. |  |  | March 23, 2009 | 2009 | Boston Scientific CRM Corp | Lead connection issues or fracture increases probability of inappropriate therapy: Inappropriate shocks, inappropriate anti-tachycardia pacing therapy, or inappropriate inhibition of needed bradycardia pacing therapy. | Device Design | Boston Scientific sent an Urgent Medical Device Information letter, dated March 23, 2009, to physicians following implanted devices. The letter described the issue, the rate of occurance and provided Patient Management Recommendations. |  | 49148 | 49,148 total | Worldwide distribution: USA, Guam, Puerto Rico, Aruba, Austria, Belgium, Brazil, Canada, Cayman Islands, Colombia, Cyprus, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, Hong Kong, Hungary, Ireland, Israel, Italy, Kuwait, Latvia, Lebanon, Luxembourg, Mexico, Netherlands, New Caledonia, Norway, Philippines, Poland, Polynesia, Portugal, Reunion, Russia, Saudi Arabia, Slovakia, South Africa, Spain, Sweden, Switzerland, United Kingdom, and Venezuela. | Terminated | October 18, 2012 | 1305 | 49148 | Computer
Z-1569-2009 | 52099 | No | BRY | Anesthetic cabinet, table, or tray. | Anesthesiology | Anesthesiology | 510(K) Exempt | Amico E-Cylinder Canister Holder, distributed by Stryker Communications, Flower Mound, TX 75028. |  |  | May 19, 2009 | 2009 | Stryker Communications Corp | Design flaw on the E-Cylinder canister holder may result in early fatigue and failure of the bracket screws, which could cause the cage to become detached from its mount and fall. | Device Design | An "Urgent: Device Recall" letter dated May 19, 2009 was issued to customers. Customers were advised to discontinue use of all affected E-cylinder holders. Consignees were asked to return the "Product Accountability Form" included with the recall letter, after which consignees will be provided with an improved cylinder holder at no cost. Direct questions to Stryker Communications at 1-972-410-7100. |  | 66 | 66 units. | Nationwide and Canada. | Terminated | April 26, 2010 | 342 | 66 | Computer
Z-1605-2009 | 52036 | No | IMB | Moist steam cabinet. | Physical Medicine | Physical Medicine | 510(K) Exempt | DERMALIFE Takara. Part #E30006-02 - Takara; Renew 2G, w/Pedestal. |  |  | June 22, 2007 | 2007 | Sybaritic, Inc | A correction or removal, conducted to reduce a risk to health posed by a device, was not reported in writing to FDA.Sybaritic, Inc. conducted a recall on DERMALIFE 2G equipment (SpaJet, Spa Oceana, Spa Feng Shui) because of IR Fan failure which resulted in the overheating and failure of the IR emitter and surrounding assembly. | Device Design | Sybaritic, Inc. issued a "Safety Alert" dated April 15, 2008 informing "All Owners/Users of DermaLife 2G Equipment" of the affected device. The letter advised consignees on device use until the mandatory Filed Correction Kit is delivered. For further questions, contact Sybaritic, Inc. at 1-800-445-8418. |  | 10 | 10 | Worldwide Distribution -- US including states of CA, IA, MN, NE, NJ, NY, TX and WA and countries of Macedonia, Hungary, Greece, South Africa, Italy, Hong Kong, United Arab Emirates, Greece, Croatia, Australia, Kuwait, Turkey, Mexico, Belgium, England, Barbados, France, Japan, Lebanon, Moldova, Poland, Portugal, Romania, Russia, Spain and Thailand. | Terminated | July 05, 2012 | 1840 | 363 | Computer
Z-1543-2009 | 51252 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Centricity PACS RA1000 Workstation software (for diagnostic image analysis); GE Healthcare Integrated IT Solutions, Barrington, IL 60010. |  |  | March 24, 2009 | 2009 | GE Healthcare Integrated IT Solutions | Software anomalies may allow the image orientation to be misinterpreted, posing a potential patient safety issue. | Software design | GE Healthcare Integrated IT Solutions notified their Centricity PACS RA1000 Workstation Software Versions 2.1.x and 3.0.x customers via an Urgent Medical Device Correction notice dated March 2, 2009 of the affected device resulting in patient safety issues. The letter provided users with safety instructions until the software is updated. Consignees are requested to complete the enclosed customer reply form, supplying the workstation software or build number, and fax the completed form to 1-847-939-1479.Centricity PACS RA1000 Workstation Software Versions PathSpeed 7.x, 8.x, Centricity 1.X and 2.0.X customers were sent separate letters explaining the recall including instructions. There will be no software updates since those software versions are no longer being supported.For questions, call the Remote Online Center, Field Engineer or the GE Customer Care Center at 1-800-437-1171. |  | 984 | 984 units | Worldwide Distribution - US including Puerto Rico and countries of Brazil, Canada, Chile, Ecuador, Mexico, Australia, China, Hong Kong, Japan, Korea, Malaysia, New Zealand, Singapore, Taiwan, Austria, Belgium, Switzerland, Germany, Denmark, Egypt, Spain, France, Great Britain, Ireland, Israel, Italy, Kuwait, Lebanon, Malta, Netherlands, Portugal, Saudi Arabia, Sweden and Turkey. | Terminated | September 06, 2010 | 531 | 984 | Computer
Z-1542-2009 | 50824 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | TumorLoc software for the Brilliance Big Bore (BB) and the CT Extended Brilliance Workstation (EBW), Model #4550 116 02431 (TumorLoc). |  |  | January 28, 2009 | 2009 | Philips Medical Systems (Cleveland) Inc | A software error in the TumorLoc software was discovered upon receipt of a customer complaint. There is the potential for data to be lost during reconstruction of images in which slices can either be discarded, or interleaved based upon their series ID. | Software design | Philips Medical issued an 'URGENT -Medical Device Correction' document dated January 28, 2009. This notification informs consignees of the TumorLoc imaging issue, the inherent hazards involved should this problem occur, the circumstances under which this problem may occur and corrective actions which the user/customer should perform in the event that this problem is encountered. It also informs that user/customer that Philips will have their Service Engineers contact the customers within 6 months to install a TumorLoc upgrade. For further questions, contact the Phillips Customer Care Center at 1-800-722-9377. |  | 356 | 356 units | Worldwide Distribution - US, Australia, Austria, Bahrain, Bangladesh, Belgium, Brazil, Canada, China, Czech Republic, Denmark, Finland, France, Germany, Greece, India, Israel, Italy, Jordan, Libyan Arab Jamahiriya, Lithuania, Malaysia, Netherlands Antilles, Netherlands, New Zealand, Norway, Philippines, Portugal, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Spain, Sweden, Switzerland, Tahiti, Taiwan, Thailand, United Kingdom and Venezuela. | Terminated | September 19, 2012 | 1330 | 356 | Computer
Z-1577-2009 | 52171 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | XoranConnect, software version 4.1.0.6, Xoran Technologies, Ann Arbor, MI. |  |  | February 03, 2009 | 2009 | Xoran Technologies, Inc. | An incorrect patient name label may appear on the screen displaying the selected patient's scan. | Software design | The firm released software version 4.1.0.7 by uploading it through a remote software upgrade on February 3, 2009. User accounts were notified by letter dated February 9, 2009. Direct questions to Xoran Technologies at 1-800-709-6726. |  | 14 | 14 | Nationwide Distribution (Alabama, Arizona, Colorado, Georgia, Illinois, Louisiana, New York, North Carolina, Oklahoma, Rhode Island, Tennessee, Texas and Utah). | Terminated | November 19, 2009 | 289 | 14 | Computer
Z-1650-2009 | 52436 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo Imaging Version V31 Picture Archiving and Communication System, model number 10014063. |  |  | June 01, 2009 | 2009 | Siemens Medical Solutions USA, Inc | Risk of misinterpretation of patient response to treatment: Date displayed may be incorrect. Filtering based on examination date may lead to an incorrect list of examinations. | Software design | syngo Imaging has issued a Customer Safety Advisory Notice, dated May 28, 2009, to affected customers via Update Instruction IM019/09/S. The notice informs the customers of the potential issue and provides instructions to avoid its occurrence. A software update to correct this issue is currently being developed and will be installed on affected systems by sygo's service representatives when it becomes available. |  | 16 | 16 units | Nationwide. | Terminated | March 19, 2013 | 1387 | 16 | Computer
Z-1652-2009 | 52430 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Hitachi ECLOS Computed Tomography (CT) Scanners, Software Version: 2.11 or less. |  |  | April 14, 2009 | 2009 | Hitachi Medical Systems America Inc | Incorrect scale on image: A software error can occur if two (2) different Field-of-View (FOV) settings are registered in a protocol using Multi-Reconstruction. This could cause the image generated to be displayed with the wrong scale factor rendering anatomical measurements incorrect. | Software design | HMSA issued an URGENT - DEVICE CORRECTION notification to their customers on 4/14/2009. The notification describes the potential software error and the condition under which these image distortions can occur. The notification goes on to provide the customer with a work-around procedure which they can use to avoid encountering this issue with their current software configuration(s). The letter states that the firm will be installing updated software on all affected systems beginning in June, 2009, and that a Hitachi Service Rep. will be contacting each customer and making arrangements to come out and install a software upgrade on the affected system(s) at no cost to the customer. Lastly, the letter provides a telephone # to the firm's Applications Helpline: 1-877-228-2777, which the customer can use to obtain additional information, if desired. |  | 4 | 4 units | Nationwide distribution, including Puerto Rico (PR). | Terminated | September 21, 2012 | 1256 | 51 | Computer
Z-1582-2009 | 52223 | No | DRN | External counter-pulsating device. | Cardiovascular | Cardiovascular | 510(k) | Cardiomedics CardiAssist External Counterpulsation System, circulatory assist device, Series 4000. |  |  | April 09, 2009 | 2009 | Cardiomedics Incorporated | Cuff may underinflate, or device may stop working during use. Cardioassist software version 4.3.1 utilized a signal from the plethysmograph that was found not to be sensitive enough due to the shape of the waveform in certain circumstances where there was excessive motion artifact. | Software design | Service visit to customer in May 2009 to upgrade and validate CardioAssist 4000 software to version 4.3.5. |  | 1 | 1 | Florida | Terminated | August 27, 2012 | 1236 | 1 | Computer
Z-1654-2009 | 51979 | No | DXJ | Medical cathode-ray tube display. | Cardiovascular | Cardiovascular | 510(k) | Philips IntelliVue Clinical Information Portfolio |  |  | May 05, 2009 | 2009 | Philips Healthcare Inc. | Error in therapy scheduling: If scheduled orders are documented and edited prior to the orders' original schedule time, ICIP can create a second instance of the same order marked as pending. A second instance of an order may result in a patient receiving duplicate therapy or patients may not receive their scheduled therapies. | Software design | Philips notified customers by Urgent Medical Device Correction Notification letter on May 6, 2009. The Customer notifications in the United States will be delivered via UPS. Customer notifications outside the US will be managed by Philips representatives in each affected geography.Customers will be asked to follow the guidelines in the Action to be taken by Customer/User section of the Urgent Medical Device Correction Notice until they receive their software corrections. The corrections will be in the form of software upgrades and will be performed by Philips service personnel.On May 27, 2009, Philips issued the second Urgent Medical Device Correction notification. While customers await the correction of their devices, customers are given the following instructions:Clinicians should not rely solely on the worklist within lCIP to determine actionable orders for the patient until their software is upgraded. Clinical users should always use the medication administration record (MAR) and the administration record in determining the care their patients receive in addition to the critical care worklist. |  | 105 | 105 units | Worldwide distribution, including USA, Australia, Belgium, China, Denmark, France, Germany, Ireland, Italy, Kenya, Malaysia, Netherlands, Singapore, Spain, Switzerland, Taiwan, UAE, and UK. | Completed |  |  | 105 | Computer
Z-1611-2009 | 52090 | No | MNH | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Stryker Mantis Rod Inserter; Non Sterile, Manufactured by Stryker Spine SAS, Cestas, France. Distributed by Howmedica Osteonics Corp; 325 Corporate Drive, Mahwah, NJ 07430. |  |  | April 30, 2009 | 2009 | Stryker Spine | The Mantis Rod Inserter Shaft was found to have deformation of the stainless stell split sphere when subjected to high forces, prohibiting the user from being able to grip the rod implantation and the split sphere may separate and fall into the patient. | Device Design | Stryker Spine issued an "Urgent Product Recall" notification dated April 29, 2009 via Federal Express to all Consignees. Users were instructed to quarantine all affected product and fax a copy of the Product Accountability Form to 201-760-8406. A Stryker sales representative will retrieve the quarantined product. For further information, contact Stryker Spine at 201-760-8206. |  | 175 | 175 in US | Nationwide Distribution. | Terminated | January 23, 2013 | 1364 | 562 | Computer
Z-1246-2009 | 51675 | No | HTQ | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Zimmer TM-400 Broach Handle; Part No. 96-210-10001. The product is a surgical instrument for use in conjunction with Zimmer TM-400 Trabecular Metal implants during spinal surgery. |  |  | April 29, 2009 | 2009 | Zimmer Inc. | The broach handle may fracture at the threaded tip during use, which would result in surgical intervention to retrieve the broach or the trial as the threaded tip fragment prevents attachment of another handle to retrieve the device. | Device Design | Distributors and user accounts were notified by "Urgent: Medical Device Recall" letter dated April 29, 2009. The letter described the affected product, reason for recall, health risks, and action for consignees. Distributors were instructed to retrieve the devices and to return them to Zimmer no later than May 13, 2009. This instrument recall will result in an inability to implant the associated devices until a new instrument becomes available.For questions or assistance about the recall please contact Zimmer Spine by calling 800-777-7505. |  | 435 | 435 | Nationwide, Austria, Australia, Belgium, Canada, Finland, Germany, Greece, Korea, South Africa, Sweden, Switzerland, Taiwan and the United Kingdom. | Terminated | October 08, 2010 | 527 | 435 | Computer
Z-1597-2009 | 51953 | No | GDC | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ASC2000 Ambulatory Surgical Table. Model #MTS110003. |  |  | April 17, 2009 | 2009 | Steris Corp | Upon investigation, the firm determined that repetitive rubbing of the hydraulic hoses against the interior of the stainless steel shroud may result in the rupture of the hydraulic hoses causing loss of control of the table surface. | Device Design | A Field Correction Notice dated April 17, 2009 was issued to customers. The letter described the potential problem of the unit. The letter also notified customers of the corrective actions to be taken by the firm. A Steris technician will visit affected accounts to inspect the hydraulic hoses and replace any that show abrasion and will make other preventive corrections. For further information or if you have questions regarding the servicing of your unit, please contact the STERIS Field Service Dispatch at 1-800-288-2805. |  | 412 | 412 units | Worldwide Distribution -- US (states of AL, AR, AZ, CA, CO, CT, FL, GA, ID, IL, IN, KY, LA, MD, MI, MN, MO, NC, NH, NJ, NY, OH, OK, OR, PA, PR, SC, TN, TX, VA and WV), AUSTRALIA, BANGLADESH, BELGIUM, BRAZIL, CANADA, CHINA, COLOMBIA, ECUADOR, GERMANY, HONG KONG, INDIA, INDONESIA, ITALY, KOREA, MEXICO, PHILIPPINES, PORTUGAL, SINGAPORE, SPAIN, SRI LANKA and THAILAND. | Terminated | March 01, 2012 | 1049 | 412 | Computer
Z-1649-2009 | 51483 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Siemens Medical Solutions USA, Inc., Symbia T Series SPECT-CT (Single-Photon Emission Computed Tomography and Computed Tomography) System; |  |  | March 11, 2009 | 2009 | Siemens Medical Solutions USA, Inc. | Patient injury can occur when the patient is lying on the patient bed and the automatic collimator exchanger is undergoing a collimator change operation at the launch of an acquisition workflow. Detector motions required for this change can cause collision with the patient or pallet if it is in brain scan position, i.e. the patient or pallet is into the field of view. | Software design | Siemens sent Urgent Field Correction Recall letters dated 1/20/09 to all users of the Symbia S and Symbia T Systems. The accounts were informed of the potential injury to patients while lying on the patient bed and the automatic collimator exchanger is undergoing a collimator change operation at the launch of an acquisition workflow. The customers were advised to perform all necessary collimator changes before the patient is positioned on the patient bed, and to examine each of their workflows that include a Tomo Acquisition activity and make sure the Auto Collimator Change check box on the Camera Parameters Tab is not checked. They also were advised that a Siemens Customer Service Representative would contact them within 60 days to schedule a corrective software upgrade free of charge. Any questions were directed to contact Siemens Medical Solutions USA, Inc. at 1-800-767-2313 (USA) for assistance. The accounts were requested to complete the Mandatory Fax-Back Form acknowledging receipt and understanding of the letter. |  | 148 | 148 units | Worldwide distribution, including USA, Australia, Austria, Canada, China, Czech Republic, Denmark, Finland, France, Germany, Great Britain, Ireland, Italy, Japan, Korea, Kuwait, Netherlands, Norway, Poland, Portugal, Qatar, Saudi Arabia, South Africa, Spain, Sweden and Switzerland. | Terminated | September 11, 2010 | 549 | 300 | Computer
Z-1568-2009 | 52047 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Dental Application on the Extended Brilliance Workspace v4.0 and v4.0.1. Model #455011202691 for Computed Tomography (CT); Model #455011002031 for Nuclear Medicine (NM). |  |  | April 27, 2009 | 2009 | Philips Medical Systems (Cleveland) Inc | Philips Healthcare has decided to recall the Dental Application on the Extended Brilliance Workspace v4.0 and v4.0.1 due to the fact that when images are saved or filmed that contain the Curved Multi-planar Reformation Section (cMPRS) views that were created with a short curve the Right/Left annotations will be reversed. | Software design | An "URGENT - Field Safety Notice" dated April 27, 2009 was sent to all consignees via certified mail. The letter described the affected products, issue, and actions for the customer/user. The notification letter provides a work around until the new software is available. For further information or support concerning this issue, please contact your local Philips representative: Customer Care Center (1-800-722-9377, option 5: Diagnostic Imaging, option 5: Radiation Therapy, option 3: Oncology) or to your local Philips Healthcare office. |  | 417 | 417 units | Worldwide Distribution -- US including DC (states of AL, AR, AZ, CO, DC, FL, IL, KY, LA, MA. ME, MI, MN, MO, NJ, NY, OH, OR, PA, SC, TN, TX, VT, and WI) and countries of Africa, Albania, Austria, Australia, Bangladesh, Belgium, Brazil, Canada, China, Columbia, Cyprus, Czech Republic, Denmark, Ecuador, Finland, France, Germany, Holland, India, Iraq, Israel, Italy, Japan, Korea, Lithuania, Mexico, Netherlands, New Zealand, Oman, Pakistan, Peru, Philippines, Poland, Puerto Rico, Russia, Singapore, Spain, South Africa, Sweden, Switzerland, Syria, Tahiti, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Uzbekistan, and Venezuela. | Terminated | September 21, 2012 | 1243 | 417 | Computer
Z-1580-2009 | 52176 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | TumorLoc Software Application for the EBW (Extended Brilliance Workstation, Model # 455011002031, sold as an accessory for the GEMINI PET/CT Systems running CT software versions 3.5.x. , CT host 2.2. |  |  | February 19, 2009 | 2009 | Philips Medical Systems (Cleveland) Inc | A software anomaly in the TumorLoc software for the firm's Brilliance CT scanners was discovered. The same software anomaly was subsequently determined to be present in the firm's GEMINI PET/CT units. The software defect results in a the potential for data to be lost during reconstruction of images in which slices can either be discarded, or interleaved based upon their series ID. | Software design | An "URGENT - Field Safety Notice" dated February 12, 2009 was sent via FedEx to domestic customers. The notice described the affected products, issue, involved hazard should problem occur and corrective actions for the customer/user. The letter also notified users/customers that a Philips Healthcare Field Service Engineer will contact and visit their site within 6 months to install a TumorLoc upgrade. Notification will also be sent to the international consignees by Philips personnel. For further information or support concerning this issue, please contact your local Philips representative: Customer Care Center (1-800-722-9377, option 5: Diagnostic Imaging, option 5: Radiation Therapy, option 3: Oncology) or to your local Philips Healthcare office. |  | 11 | 11 units | Worldwide Distribution -- US, Columbia, Germany, France, Netherlands, Canada, Belgium and Slovenia. | Terminated | September 21, 2012 | 1310 | 27 | Computer
Z-1549-2009 | 51501 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Centricity PACS (Picture Archiving and Communication System) software; GE Healthcare Integrated IT Solutions, Barrington, IL 60010. |  |  | April 03, 2009 | 2009 | GE Healthcare Integrated IT Solutions | There is a potential safety issue associated with the use of GE Centricity PACS software related to configuration of DICOM query retrieval with rejected images. | Software design | GE Healthcare Integrated IT Solutions sent an "Urgent Medical Device Correction" letter dated April 2, 2009 to their Centricity PACS software versions 2.1.x and 3.0.x customers with attention to the Director of Radiology, the Hospital Administrator, the Head of the Radiology Department and the PACS Administrator. The letters advised users of the patient safety issue associated with the affected device and provided safety instructions to follow until the software is updated. User were asked to complete the enclosed Customer Reply form and fax the completed form to 1-847-939-1479.Centricity PACS Software Versions 1.0.x and 2.0.x customers were sent separate letters explaining the issue and the safety instructions to follow, but there will be no software updates since those software versions are no longer being supported.Any questions should be directed to the Remote Online Center, Field Engineer or the GE Customer Care Center at 1-800-437-1171. |  | 1188 | 1,188 units | Worldwide Distribution - United States including Puerto Rico and countries of Australia, Austria, Belgium, Brazil, Canada, Chile, China, Denmark, Ecuador, Egypt, France, Germany, Hong Kong, Ireland, Israel, Italy, Japan, Korea, Kuwait, Lebanon, Malaysia, Malta, Mexico, Netherlands, New Zealand, Portugal, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Taiwan, Turkey and United Kingdom. | Terminated | May 25, 2012 | 1148 | 1188 | Computer
Z-1556-2009 | 52015 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE Signa Ovation 0.35 T MR System. |  |  | April 16, 2009 | 2009 | GE Medical Systems, LLC | GE Healthcare has identified a potential safety issue with the Signa OpenSpeed and Signa Ovation MR Scanner tables that may impact patient safety. The table may begin lowering on its own while a patient is on the table. If this occurs while the table is in the magnet bore, the table may tilt and could result in patient injury. The patient's head or feet depending on patient orientation, may hit | Device Design | GE Healthcare issued an "Urgent Medical Device Correction" letter dated April 16, 2009 addressed to Hospital Administrators, Mangers of Radiology/Cardiology and Radiologists/Cardiologists. The letter describes the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information. A GE Healthcare Field Engineer will schedule a field visit to correct the device.Further questions or concerns may be addressed to GE Healthcare at 1-262-521-6681. |  | 203 | 203 | Worldwide Distribution -- US including PR and states of AL, AZ, AR, CA, CO,CT, DE, FL, GA, ID, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MN, MS, MO,MT, NV, NH, NE, NJ, NY, NC, OH, OK,OR, PA, SC, SD, TN, TX, VA, WA, WV, WI and countries of BRAZIL CHINA, DENMARK, EGYPT, ENGLAND, GERMANY, GUATEMALA, HUNGARY, ITALY, IVORY COAST, JAPAN, KOREA, KUWAIT, OMAN, MEXICO, OMAN, PHILIPPINES, RUSSIAN FEDERATION, SPAIN, SWITZERLAND, TURKEY, UNITED ARAB EMIRATES and VENEZUELA. | Terminated | April 23, 2013 | 1468 | 374 | Computer
Z-1548-2009 | 51485 | No | MUH | Extraoral source x-ray system. | Dental | Radiology | 510(k) | Gendex VixWin Platinum Imaging Software, Versions 1.0, 1.1 and 1.2; Gendex Dental Systems, 901 West Oakton Street, Des Plaines, IL 60018. |  |  | March 06, 2009 | 2009 | Gendex Dental Systems | There is a compatibility issue when the VixWin Platinum versions 1.0/1.1/1.2 software is used in conjunction with Easy Dental, PracticeWorks, EagleSoft and SoftDent practice management software, resulting in patient files that may contain images that are mixed between patients. | Software design | Gendex issued an "Urgent Medical Device Report" dated March 6, 2009 to the direct account dealers of the affected device via first class mail. The noticed described the issue and provided a software update directed to fixing the problem. In addition, customers were asked to complete and return an Acknowledgement Form by fax to 1-847-718-0601.Further questions should be directed to the Gendex Technical Support team at 1-888-275-5286. |  | 1256 | 1,256 units | Nationwide Distribution. | Terminated | January 07, 2011 | 672 | 1256 | Computer
Z-1802-2009 | 52125 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Medtronic INTREPID Spinal System Coverplate, size: Small, Part Number: REF 7961330; medium, REF 7962330; large, REF 7963330; STERILE R. |  |  | February 06, 2009 | 2009 | Medtronic Sofamor Danek USA Inc | The firm received complaints regarding the use of the product in patients with poor bone quality, translational instability, difficulty implanting the cover plate and case of cover plate detachment. | Device Design | Medtronic Sofamor Danek USA, Inc. instructed Sales Representative on starting February 6, 2009 by phone to immediately return product from the field.The firm further issued a "Subject: Medtronic Intrepid Intervertebral Body Fusion Device Recall" notice dated April 2009. The recall notices were mailed directly to Risk Managers and Surgeons, with Recall Questionnaires and Surgeon Confirmation forms included respectively.For further questions, contact your Medtronic Sales Representative or the Medtronic Global Quality Department at 1-800-876-3133 extension 6333. |  | 1059 | 1059 units | Worldwide Distribution -- United States, Australia, Germany, The Netherlands, and South Africa. | Terminated | April 19, 2010 | 437 | 9087 | Computer
Z-1915-2009 | 52676 | device, prosthesis alignment | IQO | Prosthetic and orthotic accessory. | Physical Medicine | Physical Medicine | 510(K) Exempt | Stryker Orthopaedics PAL Pelvic Alignment Level, Ref: PAL 400; Howmedica Osteonics Corp, 325 Corporate Drive, Mahwah, NJ 07430. |  |  | September 11, 2007 | 2007 | Stryker Howmedica Osteonics Corp. | The bubble in a vial of the Pelvic Alignment Level (PAL) was wider than the level marking. Upon opening the PAL and removing the vial, it was discovered that the vial was leaking | Device Design | Stryker Orthopaedics issued an "Important Product Removal" letter dated September 11, 2007 via Federal Express informing consignees of the affected devices. The firm requested that users identify and return any affected product and account for product by returning a Product Accountability Form via fax.For further information, contact Stryker Orthopaedics at 1-201-831-5825. |  | 615 | 615 | Nationwide Distribution | Terminated | October 15, 2009 | 765 | 615 | Computer
Z-1795-2009 | 52575 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | HeartStart MRx/MRxE Defibrillator/Monitor Models: M3536A (HeartStart MRx) and M3536J (HeartStart MRxE). Note: M3536JIMRxE was sold to Japan only. |  |  | July 07, 2009 | 2009 | Philips Healthcare Inc. | External transport use environments, the mechanical/electrical connection between the padstherapy cable including pads/CPR cable and the MRx/MRxE may wear at an increased rate ultimately prevent the device from sensing that the padstherapy cable is connected, or, cause the device to inappropriately identify external paddles or internalpaddles when the pads therapy cable is connected | Device Design | Philips Healthcare issued an "Urgent - Voluntary Medical Device Correction" notice dated July 2009 providing information on the affected device, namely, how to identify the affected devices and instructions on actions to be taken. Consignees are asked to contact their local Philips representative to confirm receipt of the recall letter and to arrange for service.Contact Philips Healthcare for further questions at 1-800-722-9377. |  | 11579 | 11,579 units | Worldwide Distribution -- United States, Afghanistan, Albania, Australia, Austria, Bahrain, Brazil, China, Columbia, Croatia, Czech Republic, Denmark, Finland, France, Germany, Greece, Hong Kong, Iceland, India, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Kenya, Latvia, Lebanon, Libyan Arab Jamaica, Lithuania, Netherlands, New Caldonia, New Zealand, Norway, Panama, Peru, Poland, Portugal, Reunion, Saudi Arabia, Singapore, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Tunisia, United Arab Emirates, United Kingdom and Venezuela. | Terminated | February 19, 2013 | 1323 | 11579 | Computer
Z-1934-2009 | 52802 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | Edwards Lifesciences "VAMP Jr" (Venous Arterial blood Management Protection system), Lot numbers (all product manufactured between 8/01/06 and 2/29/08.) |  |  | March 14, 2008 | 2008 | Edwards Lifesciences AG | Edwards Lifescience initiated a recall of on Edwards "VAMP Jr" Pressure Monitoring Products manufactured between 8/01/06 and 2/29/08 due to the potential for the cap on the reservoir of the "VAMP Jr" to become disconnected. | Equipment maintenance | Edwards initiated the recall in the U.S. via a Urgent Customer lNotification letter with attached response confirmation fax-back attachment dated March 10, 2008 and was sent by Federal Express on March 12, 2008 requesting that customers cease using all "VAMP Jr" products and that they return all unused product.Customers were instructed to fax-back the confirmation attachment to 800-422-9329 and contact Edwards Customer service at 888-570-4012 to arrange for return of any unused product and obtain information about replacement product.Edwards Lifesciences requested that the customer provide written verification of stock and identify and unused product to be returned. This will enable a 100% effectiveness check for this recall. For product that was shipped to distributors, contact was made directly from Edwards Lifesciences Corporate Office in Irvine, CA. All product that was distributed internationally will be handled regionally. (i.e. US, Canada, Europe/EMEA, Japan, Intercontinental-including Latin America and Asia Pacific) |  | 61199 | 61,199 units | Worldwide Distribution including USA and the countries of Australia, Austria, Belgium, Brazil, Canada, Canaric Island (Spain), China, Columbia, Czech Republic, Denmark, France, Germany, Hungary, Italy, Japan, Kuwait, Netherlands, New Zealand, Norway, Puerto Rico, Spain, Sweden, Turkey, and United Kingdom. | Terminated | July 03, 2012 | 1572 | 61199 | Computer
Z-1955-2009 | 52718 | No | FMT | Infant radiant warmer. | General Hospital | General Hospital | 510(k) | Stabilet Infant Warming System, model numbers 200, 300, 1250, 1500, 200/3000, 2000, 2200/3200, 3000, and 3200. |  |  | July 17, 2009 | 2009 | Draeger Medical Systems, Inc. | Hot particles fell into the infant bassinet beneath a Radiant Infant Warmer. The resulting fire injured an infant. These devices are obsolete. | Equipment maintenance | On 7/20/09, Draeger Medical, the recalling firm, issued a press release to medical media outreach organizations and an Urgent Medical Device letter to all known customers. This information explained the incident and the need to remove and replace the units in question as soon as feasibly possible. |  | 6489 | 6489 units | Worldwide distribution: USA, Brazil, Canada, China, Dominican Republic, France, Korea, Lebanon, Mexico, Nigeria, Panama, Poland, Saudi Arabia, Taiwan, Turkey, United Kingdom, Uruguay, and Venezuela. | Terminated | June 01, 2011 | 684 | 6489 | Computer
Z-1906-2009 | 51336 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Specimen Gate Lab - MSMS Data Suite, Product Code: 5002-0310. |  |  | January 15, 2009 | 2009 | Perkin Elmer | If Result Codes are used and if they are manually edited, they are correct unless the laboratory specifically uses the View Assay button in theSpecimen History screen to open another assay, and then returns back to the original assay | Software design | Perkin Elmer began notifying US customers January 15, 2009. Notification included instructions for how to avoid the situation and instructions for how to obtain a correction. The firm requested consignees to send in a Response Form by fax (358-2-2678-357) or by email (TurkuQMresponse@perkinelmer.com).For further information, please contact your closest Specimen Gate support center at specimen.gate.support@perkinelmer.com for North America and specimengatesupportfi@perkinelmer.com for outside of North America. |  | 13 | 13 units | Worldwide Distribution -- United States (AL, AZ, CA, CO, CT, FL, GA, IN, KS, KY, MI, MO, OH, OK, OR, TN and TX) and Canada. | Terminated | September 09, 2011 | 967 | 39 | Computer
Z-1902-2009 | 52524 | No | INX | Air-fluidized bed. | Physical Medicine | Physical Medicine | 510(K) Exempt | CLINITRONÂż CII Air Fluidized Therapy Units. |  |  | June 24, 2009 | 2009 | Hill-Rom Manufacturing, Inc. | Incorrectly repaired power cord wires can potentially overheat resulting in smoldering sound foam in the base of the therapy bed. | Equipment maintenance | Hill-Rom issued an "Urgent Field Safety Notice" letter dated February 2, 2009 addressed to Facility Risk Managers and Facility Administrators. The notice detailed the issue and instructed users on actions to be taken.For further information, please contact Hill-Rom Customer Service at 1-800-638-2546. |  | 625 | 625 units (total both units) | Worldwide Distribution -- United States, Abu Dhabi, Brazil, Canada, Chile, Columbia, Mexico, Newfoundland, Puerto Rico, Saudi Arabia, Singapore, Taiwan, Turkey and Venezuela. | Terminated | July 11, 2012 | 1113 | 625 | Computer
Z-1722-2009 | 51314 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE Healthcare Signa Ovation 0.35T, Model Numbers: |  |  | November 03, 2008 | 2008 | GE Medical Systems, LLC | GE Healthcare has identified a potential pinch point hazard on Ovation MR Scanner when using CTL Array XL Body Flex and Opened Body coil. A patient's hand may be pinched between the magnet bore ceiling cover and the CTL Array, XL Body Flex, or Open Body coil if the patient puts his/her hand on the top of the coil. 2/20/09: Recall being expanded to include the OpenSpeed Systems. | Device Design | GE Healthcare issued a "Product Safety Notification" letter dated October 23, 2008. The letter was addressed to Hospital Administrators, Mangers of Radiology/Cardiology &amp; Radiologists/Cardiologists. The letter described the Safety Issue, Affected Product, Safety Instructions, Product Correction and Contact Information for the Signa Ovation 0.35T Systems.A second GE Healthcare "Urgent Medical Device Correction" letter dated March 5, 2009 was sent 3/9/09 to consignees which addressed the expanded recall to the Signa OpenSpeed0.7T MR Systems.For further questions, contact your local GE Healthcare sales or service representative. |  | 43736 | 308 (143 US / 142 OUS / 23 TBD) | Worldwide Distribution-- US (including states of AL, ALZ, AR, CA, CO,CT, DE, FL, FA, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MS, MT, MO, NE,NH, NV, NJ, NY, NC, OH, OK, OR, PA, PR, SC SD, TN, TX, VA, WA, WV and WI), CHINA, CYPRUS, ARGENTINA, BRAZIL EGYPT, DENMARK, FEDERAL REPUBLIC OF GERMANY, FEDERAL REPUBLIC OF FRANCE, GET 7 D CHINA, GREECE, HUNGARY, ITALY, IVORY COAST, JAPAN, KOREA, KUWAIT, MALAWI, MEXICO, OMAN, PANAMA, PHILIPPINES, RUSSIAN FEDERATION, SAUDI ARABIA, SPAIN , SWITZERLAND, TURKEY, UKRAINE, UNITED ARAB EMERITUS, UNITED KINGDOM and VENEZUELA. | Terminated | May 24, 2012 | 1298 | 43750 | Computer
Z-1027-2011 | 52903 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | SEQUENCER component of MOSAIQ versions 1.30 (all builds), 1.40 (all builds), 1.50 (all builds), 1.60 (all builds), and 2.0 (all builds). Manufactured by IMPAC Medical Systems, Inc. Sunnyvale, CA. |  |  | July 17, 2009 | 2009 | Impac Medical Systems Inc | Potential for improper treatment. A software error may result in potential major underdosing to targeted areas from jaw blockage of the MLC aperture (jaw intrusion), and dosing to non-targeted areas from interstitial leaf leakage (jaw extrusion). | Software design | Impac Medical Systems Inc sent an Important Safety Notice letter to all affected consginees. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to contact Impact Software support as soon as possible if they have the affected configurations identified in the letter. Customers were asked to complete the form included with the letter and and fax to:European customers: +44 1293 654401 American &amp; rest of world customers 702-992-5002 or e-mail to support@impac.com Customers were asked to distribute the Safety Notice to any and all users of IMPAC software at their organization who are potentially affected by this issue. For questions e-mail support@impac.com or callEuropean customers: 44 1293 654320American &amp; rest of world customers: 800-488-4672. |  |  |  | Worldwide Distribution - USA including AL, CA, CT, FL, GA, IA, IL, KY, MI, MN, MS, ND, NJ, OH, OR, PA,TN, TX, VA, WA, WI, and WV and the country of Canada. | Terminated | November 25, 2011 | 861 |  | Computer
Z-1958-2009 | 50783 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Stryker Reflex Hybrid Level Anterior Cervical Plate; |  |  | June 14, 2006 | 2006 | Stryker Spine | Design Flaw: The design did not prevent over angulation of the screws when the recommended instrument, All in One Guide, was not used. | Device Design | Stryker Spine issued an "Important Market Withdrawal" notice dated June 14, 2006. Consignees were advised to complete and return a Product Accountability Form via fax to 1-201-831-6069 prior to returning the affected product. Product return was to be performed during certain phase periods when upgrades were being shipped.For further questions, contact Stryker Spine at 1-201-831-5825. |  | 38987 | 38987 plates | Nationwide Distribution -- Including states of NC, OR, AZ, TN, CA, MA, MD, IL, GA, FL, NJ, NM, KY, LA, PA, MI, UT, MO, OK, KS, OH, MN, NE, IN, MT, TX, AL, CO and WI. | Terminated | August 20, 2009 | 1163 | 38987 | Computer
Z-1933-2009 | 52809 | Yes | LWP |  |  | Cardiovascular | PMA | Boston Scientific TELIGEN, Implantable Cardiac Defibrillator, Models E102 and E110. Sterilized using ethylene oxide. Boston Scientific, 4100 Hamline Avenue North, St. Paul, MN 66112-5798, USA. |  |  | July 20, 2009 | 2009 | Boston Scientific CRM Corp | Boston Scientific Cardiac Rhythm Management retrieved devices that could be subject to a potential for acute non-secure lead connections when implanted. Nonsecure lead connections can manifest as high impedance, or lead electrogram artifacts resulting in oversensing and/or inappropriate delivery of therapy. | Device Design | Boston Scientific Corporation field personnel began retrieving devices containing the original set screw system on Friday, July 17, 2009. If needed a letter, dated 07/17/2009 was left with the hospital administrator letting them know that product has been collected for return to Boston Scientific Corporation. Replacement devices will be provided at no charge. If the field personnel was unable to contact the hospital in the time provided, a notification letter dated July 22, 2009 was sent via Federal Express to the hospital. This letter was directed to the EP Catheter Lab Manager with a cc to the Hospital Risk Manager, Hospital Administrator and Purchasing and Central Supply. The letter asked that the affected devices not be implanted and to contact their local representative for pick-up and return. The firm will replace the affected devices. Please direct questions to your local Sales representative or Boston Scientific Corporation CRM Technical Services by calling 1-800-227-3422. |  | 104 | 104 | U.S. Distribution only (states of AZ, CA, CT, FL, GA, IA, IL, KY, MD, MO, NH, NJ, NY, OH, PA, SC, TN, TX, WI, and Washington D.C.) and Puerto Rico. | Terminated | June 26, 2012 | 1072 | 149 | Computer
Z-1718-2009 | 51322 | No | KPQ | Radiation therapy simulation system. | Radiology | Radiology | 510(k) | AdvantageSim MD versions 7.4 through 7.6, Model 5160092. |  |  | February 02, 2009 | 2009 | GE Medical Systems, LLC | It was reported by a customer site that when using Advantage Sim MD on Advantage Workstation, the series could be incorrectly labeled in image view when multiple series of an exam are loaded simultaneously in Advantage Sim MD and if their series dates are different. A mismatch of series label for structure sets may lead to under-treatment of a tumor due to too small coverage of volume treated. | Software design | GE Healthcare issued an "Urgent Medical Device Correction" letter dated March 11, 2009. The letter was addressed to AdvantageSim Customers, Hospital Administrators/ Risk Mangers and Oncology Department Mangers. The letter addressed the Safety Issue, Affected Product, Safety Instructions, Product Correction and Contact Information.For further questions, contact GE Healthcare at 1-800-437-1171 (US), 0120-055-919 (Japan) and for other countries, contact your local GE Healthcare sales or field service representative. |  | 230832 | 687 units (351 US / 336 OUS) | Worldwide Distribution -- Including states of AL, AZ, CA, CO, CT, FL, GA, DC, ID, IL, IN, IA, KS, LA, ME, MD, MA, MI, MN, MS,MO, MT,NE, NH, NJ, NM, NY, NC, ND, OH, OK PR, PA, SC, SD, TN, TX, UT, VA, WA and WI and AUSTRIA, BELGIUM, BRAZIL, CANADA, CHILE, CHINA, CZECH REPUBLIC, DENMARK, EGYPT, FINLAND, FRANCE, GERMANY, HONG KONG, INDIA, IRELAND, ISRAEL, ITALY, JAPAN, KOREA, MALAYSIA, NETHERLANDS, POLAND, PORTUGAL, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, TOKELAU, TURKEY, UNITED ARAB EMIRATES and the UNITED KINGDOM. | Terminated | December 10, 2011 | 1041 | 230832 | Computer
Z-1719-2009 | 51120 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare MUSE Cardiology Information Systems vs. 7.0, 7.0.1, 7.0.2, 7.1, 7.1.1. |  |  | February 02, 2009 | 2009 | GE Healthcare, LLC | GE Healthcare has become aware of an issue resulting from a workflow situation related to the automated labeling of patient data in the MUSE Cardiology Information System. When a prior record has been confirmed with an invalid (default) patient identifier (PID), MUSE will accept future records with the same default PID as correct. If subsequent records don not contain a name, MUSE will pre-popu | Software design | GE Healthcare issued an "Urgent Medical Device Correction" letter dated October 13, 2008, addressed to Healthcare Administrator/Risk Manager, Cardiology Administrators and Head of Cardiology. The letter listed the Safety Issue, Affected Product Details, Safety Instructions, Product Correction, Background Information and Contact Information.For further questions, contact GE Healthcare at 1-800-558-7044. Outside the US, please contact your local GE Healthcare Service Representative. |  | 327 | 327 | Worldwide Distribution -- Including states of AL, AK, AZ, AR, CA, CO, CT, FL, GA, ID, IL, IN, KY, LA, MD, MA, MI, MN, MS, MO, NE, NV, NH, NJ, NM, NY, NC, OH, OK, OR, PA, SC, TN, TX, UT, VT, VA, WA, WI, WY and DC and countries of UNITED KINGDOM , TAIWAN, SWEDEN, SPAIN, SOUTH AFRICA, SINGAPORE, SAUDI ARABIA, QATAR, NEW ZEALAND, NETHERLANDS, KUWAIT, REPUBLIC OF KOREA, JORDAN, JAPAN, ITALY, IRELAND, GERMANY, FRANCE, FINLAND, ESTONIA, DENMARK, CANADA, BELGIUM, AUSTRIA and AUSTRALIA. | Terminated | January 16, 2014 | 1809 | 327 | Computer
Z-1720-2009 | 52022 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, CT Perfusion 4 on Advantage Workstation, CT and PET/CT consoles. |  |  | May 04, 2009 | 2009 | GE Medical Systems, LLC | GE Healthcare has recently become aware of the need to adjust default settings associated with the use of CT Perfusion 4 on the Advantage Workstation, ST and PET/CT consoles that may impact patient safety.The default parameter settings in CT Perfusion 4 may not be optimal in the case of data acquired with longer time sampling intervals such as axial scans with inter scan delay (ISD) greater th | Software design | GE Healthcare issued an "Urgent Medical Device Correction" letter dated April 23, 2009. The letter was addressed to Hospital Administrators / Risk managers and Radiology department mangers. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information instructing the user to follow the enclosed User Guide Addendum to mitigate further problems.For further questions, contact GE Healthcare at 1-800-437-1171 (US), 0120-055-919 (Japan); for other countries, contact your local GE Healthcare field service representative. |  | 444 | 444 | Worldwide Distribution -- Including states of AZ, CA, CT, DC, FL, HI, IL, IA, MA, MI, NJ, NY, NC, OK, PA, PR, TX, WA, WV and WI and countries of ARGENTINA, BELGIUM, BOLIVIA, BRAZIL, CANADA, CHILE, CHINA, CZECH REPUBLIC, DENMARK, FINLAND, FRANCE, GERMANY, GREECE, GUATEMALA, HONDURAS, HONG KONG, INDIA, INDONESIA, ISRAEL, ITALY, JAPAN, REPUBLIC OF KOREA, MALAYSIA, MEXICO, NETHERLANDS, NEW ZEALAND, PAKISTAN, POLAND, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SLOVAKIA, SOUTH AFRICA, SPAIN, SWEDEN, TAIWAN, THAILAND, TURKEY, UNITED KINGDOM and VENEZUELA. | Terminated | May 16, 2012 | 1108 | 444 | Computer
Z-1734-2009 | 52079 | No | GKH | Automated blood cell diluting apparatus. | Hematology | Hematology | 510(K) Exempt | CellTracks AutoPrep System for in vitro diagnostic use, software version 2.2 or greater. Distributed by Veridex, 33 Technology Drive, Warren, NJ 07059. Manufactured by Immunicon, Huntingdon Valley, PA. Immunicon was acquired by Veridex in August of 2008. |  |  | May 06, 2009 | 2009 | Veridex, LLC | Bottle degradation: Investigation has determined that degradation may occur in the plastic waste bottle of the CellTracks AutoPrep System. | Device Design | Veridex, LLC issued a recall notice to consignees dated March 12, 2009 detailing the affected product. The firm will provide replacement bottles to all customers in the near future.For further questions, contact Veridex Customer Technical Services at 1-877-837-4339. |  | 196 | 196 instruments | Worldwide Distribution | Terminated | October 05, 2010 | 517 | 196 | Computer
Z-1815-2009 | 52523 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Hospira Phoenix Infusion System with MedNet Software, Symbiq One-Channel Infusion System, Software versions v2.1 &amp; v3.0, Manufactured by Hospira Inc., Morgan Hill, CA. |  |  | March 31, 2009 | 2009 | Hospira Inc | Potential delay/underinfusion of critical therapy-- Devices experience increased frequency of alarms that result in temporary pump stoppage. | Software design | Hospira, Inc. issued an "Urgent: Device Field Correction" dated March 31, 2009 via Federal Express to Healthcare Professionals and Valued Hospira Customers. Consignees were informed of the affected devices and will be contacted by a Hospira representative to arrange necessary device upgrades.For further questions, contact Hospira Global Product Safety and Complaints at 1-800-441-4100. |  |  |  | Product was distributed to 59 consignees to all 50 States. | Terminated | December 18, 2010 | 627 |  | Computer
Z-1813-2009 | 52439 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Respironics Esprit Ventilator System, Model Number V1000 with Optional External Battery; Product Number: 1001456. |  |  | April 22, 2009 | 2009 | Respironics California Inc | The recall was initiated because through routine product monitoring Respironics discovered that under certain conditions, when operating on AC line power, a portion of the current from the External Battery charger is erroneously diverted to the ventilator. When this occurs, the External Battery charger may not transition correctly. Prolonged operation in this condition results in overcharging and | Device Design | Respironics California, Inc. issued an "Urgent - Medical Device Field Correction" notice dated April 22, 2009. Customers were informed of the affected device including corrective actions taken by the firm. A Respironics representative will contact customers to schedule service of the device.Further questions can be addressed by contacting Respironics US Customer Service at 1-800-345-6443. |  | 508 | 508 | Worldwide Distribution: United States, Hong Kong, India, Mexico, South Africa and Vietnam. | Terminated | July 13, 2012 | 1178 | 508 | Computer
Z-1697-2009 | 49148 | No | ILY | Infrared lamp. | Physical Medicine | Physical Medicine | 510(K) Exempt | Foot Choice Therapeutic Foot Massager, Model Number YS-322. |  |  | August 07, 2006 | 2006 | Suarez Corporation Industries | Complaints of product causing burns to consumers. | Device Design | On August 9- August 14, 2006, the firm sent a letter to all customers which included an "Important Customer Reminder." The Reminder now listed additional warnings about serious burns occurring regardless of the control settings. Direct questions to Suarez Corporation Industries by calling 1-330-494-5504. |  | 71512 | 71,512 units | Nationwide and Canada. | Terminated | October 13, 2010 | 1528 | 71512 | Computer
Z-1875-2009 | 52032 | No | IOQ | Powered flotation therapy bed. | Physical Medicine | Physical Medicine | 510(K) Exempt | FlexicairÂż MC3 (model M4000) Low Airloss Therapy Units, for the prevention of pressure sores. |  |  | May 12, 2009 | 2009 | Hill-Rom Manufacturing, Inc. | The devices may emit smoke from the blower box assembly. | Equipment maintenance | Urgent Field Safety Notice letter dated May 11, 2009. Hill-Rom service centers were notified by letter on 05/13/2009. Action to be taken are as follows: Rental Customers; In the next 60 days, Hill Rom will be removing the FlexicairÂż II and FlexicairÂż MC3 low Airloss Therapy Units from service to correct the devices. During that time, Hill-Rom may provide an equivalent substitute therapy to accommodate the patients who continue to need a pressure reducing therapy device. If t his malfunction occurs before a substitution or a corrected device is provided, unplug the device, remove the patient from the surface and notify Hill-Rom immediately. Facility-Owned Customers: In the next 60 days, for all facility owned Flexicair II and FlexicairÂż MC3 Low Airloss Therapy Units, a Hill-Rom service technician will either replace the affected components in the blower assembly or deliver an equivalent substitute unit from Hill-Rom's rental fleet. If this malfunction occurs before the device is corrected or a substitute provided, do as follows: unplug the device, remove the patient from the surface, and immediately notify Hill-Rom Technical Support using the attached listing their geographical region. |  | 889 | 889 units | Nationwide and Canada | Terminated | April 03, 2012 | 1057 | 946 | Computer
Z-0544-2009 | 49457 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Siemens syngo Dynamics Workplace, version 5.0, Siemens Medical Solutions, Ann Arbor, MI; 10090310. |  |  | July 23, 2008 | 2008 | Siemens Medical Solutions, USA, Inc | If the text is manually modified and additional phrases added as a result of worksheet interactions, some of the phrases may not populate the report. | Software design | Consignees were notified of the problem via a Customer Safety Advisory Notice dated 7/23/08, which provided guidance on how to minimize the risk of the problem occurring, and told customers that a service patch will be installed when it becomes available. |  | 306 | 306 of all products. | Nationwide Distribution. | Terminated | October 18, 2010 | 817 | 306 | Computer
Z-1675-2009 | 52435 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Somatom Definition Model AS/AS+, Model number 8098027. |  |  | May 29, 2009 | 2009 | Siemens Medical Solutions USA, Inc | Support mechanism may not engage properly due to improper stowage after opening which may result in internal component damage or potential injury. | Equipment maintenance | The firm has issued a modification of the telescope support arm parking bracket via Update Instruction CT028/09/S. A firm service representative will install the modified support arm parking brackets on the affected systems. |  | 93 | 93 units | Nationwide Distribution -- including states of AR, AZ, CA, CO, CT, FL, IA, IL, IN, KS, KY, MA, MD, MN, MO, MT, NC, ND, NE, NH, NY, OH, PA, SC, SD, TN, VA, WA, WI, and WY. | Terminated | September 01, 2009 | 95 | 93 | Computer
Z-1873-2009 | 52721 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Welch Allyn Vital Sign Monitor 5300 Series/Class II. (VSM 300). |  |  | July 28, 2009 | 2009 | Welch Allyn Protocol, Inc | The Welch Allyn VSM300 Vital Signs Monitor may experience a loss of audio which includes a loss of audio alarms. | Device Design | An "Urgent: Correction" letter dated July 28, 2009 was sent to customers describing the risk associated with loss of audible alarms and recommend that the users mitigate the risk by checking for proper audible alarm function each time they power on the device. (Current labeling describes the variety of beeps and tones that should be present). Users were instructed: 1) if the audio fails, do not use the device, 2) Remove it from service and 3) contact Welch Allyn Protocol, Inc. at the following telephone number (1-800-289-2501) to arrange the return and repair of your device to Technical Services. Direct questions about this recall to Welch Allyn Protocol, Inc. by calling 1-503-530-7935. |  | 68123 | 68,123 | Worldwide Distribution -- US, United Arab Emirates, Afghanistan, Austria, Australia, Belgium, Bulgaria, Bahrain, Botswana, Canada, Switzerland, China, Cyprus, Germany, Denmark, Egypt, Spain, Figi, France, Great Britain, Greece, Hong Kong, Croatia, Indonesia, Ireland, Israel, India, Iceland, Italy, Jordan, Kenya, Kyrgyzstan, South Korea, Kuwait, Lebanon, Sri Lanka, Malaysia, Nigeria, The Netherlands, Norway, New Zealand, Oman, Phillipines, Pakistan, Poland, Puerto Rico, Panama, Portugal, Qatar, Russia, Saudi Arabia, Sweden, Singapore, Thailand, Taiwan and South Africa. | Terminated | August 10, 2011 | 743 | 68123 | Computer
Z-1790-2009 | 52119 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | The Alaris PC Unit, Model Numbers 8000 and 8015 (formerly Medley PC Unit) with software version 4 and above. |  |  | June 12, 2009 | 2009 | Cardinal Health 303 dba Cardinal Health | There is a potential risk for medication not being delivered when IV tubing sets are occluded in the Alaris Pump module. This can occur if IV sets are set up in the infusion pumps but not used for an extended period of time. If users fail to clear the occlusion, then the occluded tubing may cause medication not to be delivered. A delay of medication may result in serious injury and/or death, an | Device Design | The recall was initiated on June 12, 2009 with the firm forwarding via UPS an Urgent: Medical Device Recall Notification with attached Customer Response Card to the Chief Administrative Officer, Director of Pharmacy, Director of Biomedical Engineering, Director of Nursing &amp; Director of Risk Management of all of the affected consignees. Customers were also asked to please promptly complete and return the enclosed Customer Response Card to expedite the correction process.The Recall Notification informed the consignees of:(A) PROBLEM #1: Occlusion Warning Message for Alaris Pump Module; (B) PROBLEM #2 Syringe volume Warning Message for Alaris PCA module; (C) PROBLEM #3 Electrostatic Discharge (ESD) protection circuitry; (D) PROBLEM #4 Fluid Ingress and (E) PROBLEM #5 Alaris System Inter-Unit Interface (IUI) Connectors Labeling. The recall notice included "Required Action for Users &amp; Follow-up actions by Cardinal Health for all five problems. Customers were also asked to report any adverse reactions experienced with the use of this product, and/or quality problems to the FDA's MedWatch Program. In addition, customer were provided the following contact information for the recall: (1) Cardinal Health Recall Center for general questions: By phone at 1-888-562-6018 7am to 5pm (Pacific) ; (2) Customer Advocacy for adverse reports: By phone at 1-800-854-7120, Option 1, Option 1, Option 3 or by email at customerfeedback@cardinalhealth.com 24 hours a day, Sunday-Saturday; (3) Technical Support for Technical Questions regarding the Alaris System: By phone at 1-888-562-6018, 6am-5pm (Pacific).The firm followed-up with the release of a nationwide press statement regarding this issue on July 29, 2009. The press release indicated serial numbers of affected devices, as well as CareFusions short term instructions to customers, and the firms strategy to fix the affected devices can be found at: www.cardinalhealth.com/alaris/medical-device-recall/ (This informatio |  | 186000 | 186,000 units | Worldwide Distribution -- US (including Puerto Rico and Guam), Germany, England, Canada, Netherlands, Malaysia, Singapore, Australia, and New Zealand. | Terminated | March 19, 2012 | 1011 | 1016000 | Computer