## Sheet1
Recall Number | Recall Event ID | Device Type | Product Code | Regulation Description | Medical Specialty | Review Panel | Submission Type | Trade Name/Product | Main Name | Recall Class | Date Posted | Year Posted | Recalling Manufacturer | Reason for Recall | FDA Determined Cause | Action | Consumer Instructions | Clean Quantity | Quantity in Commerce | Distribution | Recall Status | Termination Date | Time to Terminate | Merged Quantities | Fault_Class
Z-3176-2011 | 59455 | No | KRI | Cardiopulmonary bypass accessory equipment. | Cardiovascular | Cardiovascular | 510(K) Exempt | Terumo Sarns Level Sensor II Pads. |  |  | August 12, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | Level Sensor II mounting pads may not stick to the venous reservoir and may detach either at case setup or during a case. | Incorrect or no expiration date | Terumo Cardiovascular Systems Corp. sent an "Urgent Medical Device Removal" notice dated August 10, 2011 to all affected customers.The notice instructs customers to discard all remaining level sensor pads that do not have an expiration date on the package, and notify Terumo of the number of affected pads via a response form. Terumo will replace recalled product.For questions on this recall call Terumo Customer Service at (800) 521-2818, 8 am-6 pm EST. |  | 1531440 | 1,531,440 devices | Worldwide Distribution -- Nationwide Distribution, and the island of Puerto Rico. | Terminated | June 20, 2012 | 313 | 1531440 | Not_Computer
Z-0001-2012 | 57858 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Probe, Tactile, Left, ASM, Long, |  |  | January 20, 2011 | 2011 | Medtronic Navigation, Inc | Distal end of tactile probe may break off during use. | Labeling design | Medtronic sent an "Urgent Product Safety Notice" to all affected customers dated Jan 20, 2011.The letter included affected product and updated instructions for use. On-site training was given for the use of the probes by Medtronic Representatives.For information on this recall call Medtronics at (720) 890-3409. |  | 6 | 6 probes | Worldwide Distribution: Nationwide distribution including the states of CO and WI; and the country of South Korea | Terminated | October 03, 2011 | 256 | 6 | Not_Computer
Z-3104-2011 | 59567 | No | DIC | Alcohol test system. | Toxicology | Toxicology | 510(k) | Drug Check, Distributed by Express Diagnostics Intl Inc., 1550 Industrial Drive, Blue Earth, MN 56013. |  |  | November 01, 2010 | 2010 | Express Diagnostics Int'l., Inc. | Express Diagnostics Intl'l Inc is recalling DrugCheck devices because they were informed via letter from the alcohol test strip manufacturer on October 27, 2010 that the product was not 510 K cleared for in vitro diagnostic use and could not be labeled as such in the US. | Error in labeling | Express Diagnostics Intl., Inc. sent an Updated Recall Notice dated November 1, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed if they have open orders of the affected product to contact customer service about their order. Customers were asked to complete and fax the form on page 3 of the document to 507-526-2252. Further shipment of products could not continue until Express Diagnostics received the signed form.For questions regarding this recall call 507-526-3951. |  | 56650 | 56,650 | Nationwide Distribution including AZ, CA,K CT, FL, GA, IL, IA, LA, MD, MA, MN, MS, MO,K MT, NE, NJ, NY, ND, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WV, WI, and WY. | Terminated | July 06, 2012 | 613 | 56650 | Not_Computer
Z-3172-2011 | 59453 | No | MCY | Elastic bandage. | General Hospital | General &amp; Plastic Surgery | Enforcement Discretion | Wound Dressing Kit - Tier I, 10 MM Flat Drain; product code 0877061. |  |  | July 27, 2011 | 2011 | Medela Inc | Medela determined that some of the individual kits or bags of lot #370118 of "0877061 Wound Dressing Set, Tier I, 10 MM Flat Drain", were labeled incorrectly as "0877067 Wound Dressing Set, Tier III 15 Fr Round Channel Drain". The outside label on the case is correctly labeled as "0877061 Wound Dressing Set, Tier 1,10 MM Flat Drain". | Error in labeling | Medela Inc. sent an "Urgent Device Recall" letter dated July 26, 2011 to all affected customers.The letter described the product, problem and actions to be taken. The customers were instructed to do the following: 1) not to use the affected lot and place them in quarantine 2) stop sale and distribution and set aside their existing inventory of the affected product 3) conduct a physical count and record the data on the Product Inventory Form included with this letter 4) complete and return the Product Inventory Form via fax to 800-995-7867 or e-mail suction@medela.com, even if they have no affected product on hand 5)if they have the affected product in stock, Medela Customer Service will ship substitute product to the customer at no cost. 6) and to record the number of units on hand for return and replacement on the enclosed Product Inventory Form. The completed form should be returned via fax or e-mail to Medela even if the account has no inventory of the affected lots. If you have any questions about this recall please call (877) 735-1626 or e-mail at suction@medela.com. |  | 3000 | 3,000 kits | Nationwide Distribution including the states of California, Florida, Georgia, Illinois, Kansas, Kentucky, Massachusetts, Michigan, Mississippi, New Jersey, New Mexico, North Carolina, Ohio, Pennsylvania, South Carolina, Tennessee and Texas. | Terminated | February 16, 2012 | 204 | 3000 | Not_Computer
Z-2959-2011 | 59119 | No | FJS | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Bard Aspira Drainage Kit, Product Code 4991502. |  |  | May 21, 2011 | 2011 | Bard Access Systems | Labeling for drainage kit contained incomplete information. | Error in labeling | The firm, BARD, sent a letter entitled "Labeling Discrepancy on Aspira Drainage Accessory Kits" dated June 12, 2011 to its customers. The letter provides the customers with the information missing from the labeling and/or IFU stating that the product can be safely used for peritoneal and pleural applications.If you have any questions please do not hesitate to contact BARD at 1-800-290-1689. If there are any clinical questions please contact Medical Services and Support at 1-866-893-2691. |  | 39290 | 39,290 units | Nationwide distribution. | Terminated | November 02, 2011 | 165 | 68530 | Not_Computer
Z-0104-2012 | 60004 | No | ODP | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Synthes Spine Zero-P Instruments and Implants. |  |  | September 23, 2011 | 2011 | Synthes USA (HQ), Inc. | LABELING CORRECTION - Following review of labeling and promotional materials associated with the Zero-P Implant system, it was found that the materials reference a study where PEEK-CR is utilized as a cervical interbody spacer. PEEK-CR is not currently indicated for use as a cervical interbody spacer. | Labeling False and Misleading | Synthes (USA) sent an "URGENT: MEDICAL DEVICE LABELING CORRECTION" letter dated September 23, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter instructs customers to review the updated supplemental labeling page included with the letter and replace/attach the technique guide page in any Zero-P technique guide that customers may have. The Verification Section of the letter is to be completed and returned to Synthes via fax at 610-251-9005. For questions regarding this notice call 1-800-620-7025 ext. 5403. |  |  |  | USA (nationwide) including Puerto Rico, Washington, DC and the states of AL, AK, AZ, AR, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, and WY. | Terminated | December 18, 2014 | 1182 |  | Not_Computer
Z-2955-2011 | 59392 | No | IKD | Electrode cable. | Physical Medicine | Physical Medicine | 510(K) Exempt | Supreme Electrophysiology Extension Cable, REF 401984, For use with Hexapolar Catheters, Sterile EO, Rx only, St Jude Medical, 14901 DeVeau Place, Minnetonka, MN 55345. |  |  | June 01, 2011 | 2011 | St. Jude Medical | St. Jude Medical is recalling a single lot (3334664) of Supreme Electrophysiology Extension Cable which were packaged with a Quadripolar cable instead of a hexapolar Cable. St. Jude Medical has determined that the incorrectly labeled product poses no safety risk since the extension cable can be safely used with a corresponding Quadripolar catheter but cannot be incorrectly connected to a Hexap | Labeling mix-ups | St. Jude Medical sent an "IMPORTANT MEDICAL DEVICE FIELD ADVISORY NOTICE" letter dated June 8, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter instructs customers to locate the affected product and to complete and return an attached SJM AFD Reconciliation Form. The form will be used to replace product and as documentation that the customers received the notice. Fax the completed form to 877-497-031.Contact the Director of Product Marketing at 512-329-1054 for questions regarding this notice. |  | 32 | 32 | Nationwide Distribution-USA (nationwide) including the states of FL, NC, NY, and OR. | Terminated | August 16, 2012 | 442 | 32 | Not_Computer
Z-2952-2011 | 59382 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Spectranetics Turbo-Booster Laser Guide Catheter, 8 Fr, REF 518-063, Sterile EO, Spectranetics Corporation, 96 Talamine Court, Colorado Springs, CO 80907. |  |  | July 18, 2011 | 2011 | Spectranetics Corporation | Packaging integrity may be compromised, thus sterility can not be assured. | Package design/selection | Spectranetics sent an "PRODUCT RECALL NOTICE: EU SHIPMENTS OF TURBO-BOOSTER LASER GUIDE CATHETERS" letter dated July 15, 2011 to all affected customers. The letter identifies the product, problem, and the actions to be taken by the customers.The letter instructs customers to check their inventory and set aside affected product. Customers are to contact their sales representative for information on how to return the affected product and receive replacement product.Further information is available at +011.49.719.447.2539. |  | 123 | 123 units | Foreign Distribution-including the countries of Belgium, Cyprus, Germany, Greece, Italy, Spain, Switzerland, and United Kingdom. | Terminated | August 12, 2011 | 25 | 180 | Not_Computer
Z-0068-2012 | 60000 | No | KGT | Drape adhesive. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | STRIP T's STERILE ADHESIVE ORGANIZER, Model #SC5000-4R, Lot #090710, 510(k) exempt. One (1) adhesive organizer is packed in a Tyvek pouch and 60 pouches are placed into a 'Shelf Box'. The shelf box is placed into a shipping container. |  |  | March 05, 2008 | 2008 | Kapp Surgical Instrument Inc | A portion of the labeled expiration date is missing. The expiration date on approximately half of the pouch labels distributed are missing the last digit of the year of expiration. However, the shelf box containing the pouches is labeled correctly. | Incorrect or no expiration date | Surgical Concepts, Inc ( A subsidiary of Kapp Surgical Instrument Inc) sent a notification letter dated March 5, 2008, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The firm ceased distribution of the suspect product on December 5, 2007. The letter informs the firm's customers (both distributors and medical user/customers) of the expiration date labeling omission of the last digit of the year of expiration. It goes on to explain that the incomplete expiration date is present on the individual packages only and that the shelf boxes are labeled correctly with the complete/correct expiration date. The letter provides the customers with a contact person and a toll free telephone number 1-(800) 282-5277, to call in order to make arrangements for receive replacement labels for any misbranded packages received. |  | 59 | 59 boxes (3540 units) | Nationwide Distribution-- including the states of AK, CA, CO, FL, GA, ID, IL, IN, LA, MA, MD, MI, MO, MT, NC, NM, OH, PA, TN, TX, UT, and VA | Terminated | February 19, 2013 | 1812 | 59 | Not_Computer
Z-2754-2011 | 59111 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Medtronic 6F Zuma Guide Catheter with AL 1.5 Curve with side holes |  |  | June 08, 2011 | 2011 | Medtronic Vascular | Mislabeled; Product labeled as 6F Zuma Guide with AL 1.5 curve with side holes may contain a 5F Zuma Catheter with SAL 1.0 curve | Labeling Change Control | The firm, Medtronic, sent an "Urgent Medical Device Recall Notice(for US)" letter dated June 7, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to segregate the product so that it will not be used. A Medtronic Sales Representative will visit the customers to account for all units they received and initiate a credit as appropriate.If you have any questions, please call (978) 739-0371 or (978) 739-3259. |  | 5 | 5 units | Nationwide distribution: USA including states of: CA, FL, and MA. | Terminated | April 17, 2012 | 314 | 5 | Not_Computer
Z-0070-2012 | 60007 | No | MJN | Vascular clamp. | Cardiovascular | Cardiovascular | 510(k) | LeMaitre 9F Pruitt F3 Outlying Carotid Shunt with T-port |  |  | October 04, 2011 | 2011 | Lemaitre Vascular, Inc. | Extended Expiration on the outer box label 2016-05, the correct expiration date is on the tray label: 2014-05 | Incorrect or no expiration date | The firm, LeMaitre Vascular, Inc., sent an "URGENT: 9F PRUITT F3 Carotid Shunt Device Field Safety Notice" dated October 4, 2011 to its customers. The notice describes the product, problem and action to be taken. The customers were instructed to identify their inventory; return the affected products for exchange with a properly labeled device; contact Le Matire customer service for replacement at 800-628-9470, and complete and return the RECALL NOTICE RECONCILIATION FORM via fax to 781-425-5049, even if no product is on hand. If you have any questions concerning this safety notice, please contact the Quality Assurance Engineer at 781-221-2266 ext. 183. |  | 144 | 144 units | Nationwide distribution: USA including states of: CA, FL, IA, IL, MA, MI, NE, NJ, NY, OH, SC, TX, and WI. | Terminated | March 14, 2012 | 162 | 144 | Not_Computer
Z-3119-2011 | 59691 | No | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | Operations and Technical Manuals for the firm's WarmAir Model 135 Hyperthermia System (115V, 100V, and 230/240V), 510 (k) #K101148, Serial Numbers: 994-1350001 through 104-1356894. The WarmAir 135 Hyperthermia System is a patient temperature management device which provides forced air warmed by the controller to ... a blanket that is placed over or around adult, pediatric or neonatal patients in order to warm them. The heated air is blown through a connecting flexible hose to a disposable blanket to provide patient therapy by the means of warmed air. The system can be used in long-term care facilities, surgical facilities, hospitals including the Post-anesthesia Care Unit (PACU), Intensive Care Unit (ICU), Surgical Intensive Care Unit (SICU), Emergency Room (ER), Operating Room (OR), medical and surgical floors, or any other department or hospital facility requiring patient temperature management. |  |  | July 20, 2011 | 2011 | Cincinnati Sub-Zero Products Inc | Cincinnati Sub-Zero is conducting a field correction due to changes that have been made to the WarmAir Model 135 devices Operation and Technical Manuals in order to comply with FDA device labeling regulations. | Labeling design | Cincinnati Sub-Zero Products, Inc. sent an 'URGENT Medical Device Recall Field Action Notice" dated July 27, 2011, to all affected user/customers. The letter includes: description of device; the problem with the suspect manuals, and provides contact information for the customers to use to request new versions of the manuals. Upon receipt of the new manuals, the customers are asked to immediately collect and destroy/discard all of the old manuals at their facility and to replace them with the new versions. After the manuals have been discarded and replaced, the customers are asked to complete and return the attached "Recall Response Form" via fax to (513) 772-9119 or scan and email to WA135manual@cszinc.com or mail to Cincinnati Sub-Zero, 12011 Mosteller Road, Cincinnati, OH 45241.For additional information the customers are provided with an e-mail address: WA135-2011@cszinc.com as well as a toll-free telephone number and fax number to reach the firm and obtain additional information as necessary. The contact numbers are: 1 (800) 989-7373, or (513) 772-8810, and/or fax # (513) 772-9119.. |  | 6561 | 6561 units were distributed. | Nationwide distribution: AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MT, NC, NE, NJ, NM, NY, OH, OK, PA, RI, SC, TN, TX, UT, VA, WA, and WI. | Terminated | July 05, 2013 | 716 | 6561 | Not_Computer
Z-0187-2012 | 59461 | No | LLI | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Invacare 600 lb. Bariatric Bed, Head motor extension cable used on component numbers BAR5490 and BAR5490IVC. These components are part of the following Bed Model Number(s): BAR600, BARPKGCA, BAR600IVC, BARPKGIVC-1633, and BARPKG. |  |  | June 23, 2011 | 2011 | Invacare Corporation | Invacare Corporation decided to recall the products because risk analysis reveals that miskeying or disconnection of the head motor extension cable during bed set-up may result in product damage, or in rare instances, the potential for fire. | Labeling design | On 9/19/11, Invacare sent Recall Notificaiton Letters to their customers. They instructed their customers to quarantine the product, complete and return a response form. |  | 35747 | 35,747 | Distribution Nationwide and to United Arab Emirates, Dominican Republic, Egypt, Guam, Mexico, &amp; Virgin Islands. | Terminated | July 02, 2013 | 740 | 35747 | Not_Computer
Z-3110-2011 | 59436 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | Cell Marque antibody CD99 (H036-1.1) |  |  | July 12, 2011 | 2011 | Cell Marque Corporation | Incorrect label on product bottle, although labeling on outer box is correct. | Labeling mix-ups | The firm, Cell Marque, notified customers by telephone on July 12, 2011. The customers were instructed to return the product. Cell Marque immediately issued replacement of the product.If you have any further question or require additional documentation, contact the VP of Regulatory and Quality Affairs at (916) 746-8900. |  | 2 | 2 units | Worldwide distribution: USA (nationwide) include state of : CA and country of: Great Britain. | Terminated | August 31, 2011 | 50 | 2 | Not_Computer
Z-1373-2011 | 57502 | No | EHD | Extraoral source x-ray system. | Dental | Radiology | 510(k) | Orthoralix 8500 DDE Panoramic Dental X-Ray System; Extraoral x-ray unit with timer; Gendex Dental Systems, Des Plaines, IL 60018; part number 110-0234G1 (domestic) and 110-0234G3 (international) |  |  | December 21, 2010 | 2010 | Gendex Dental Systems | The Automatic Exposure Control (AEC) feature located on the control panel of the Orthoralix 8500 Panoramic Dental X-Ray System was functional, but was subsequently disabled in units manufactured starting 5/17/10. North American units manufactured between 12/4/09 and 5/16/10 had functioning AEC buttons, but their manuals stated they were disabled. Units intended for International distribution bet | Labeling Change Control | Gendex sent a recall letter to their accounts on December 21, 2010, stating that the AEC button on the control panel was a non-functional feature that becomes illuminated upon activation and that Gendex would send them a kit to correct this issue and inactivate the button. A second letter was sent to the accounts on January 31, 2011, explaining that a corrective kit was no longer necessary, letting them know if the AEC feature was active or inactive for their system, and providing them with the appropriate user manual addendum for their system.Customers were asked to complete and return the required Acknowledgement Form by fax to 215-997-5665 or e-mail: regulatoryaffairs@gendex.com.For questions regarding this recall call 800-323-8029. |  | 574 | 574 systems | Worldwide Distribution - USA including AL, AZ, CA, CO, CT, FL, GA, CO, CT, FL, GA, ID, IL, KS, KY, LA, MA, MD, MI, MN, MO, NC, NE, NJ, NM, NY, OH, OK, PA, TN, TX, VA, WA, and WI and the countries of Australia, Canada, Germany, New Zealand, Taiwan, and Vietnam. | Terminated | November 09, 2011 | 323 | 574 | Not_Computer
Z-0012-2012 | 59943 | No | ECI | Orthodontic appliance and accessories. | Dental | Dental | 510(K) Exempt | Biomet Microfixation outer packaging labeled in part:"REF: SP-2411***RUBBER BAND, LATEX FREE SMALL, HEAVY, 300/PACKAGE 4MM, (3/16" INNER DIAMETER) NON LATEX ELASTICS LOT 307850***DISTRIBUTED BY: BIOMET MICROFIXATION 1520 Tradeport Drive, Jacksonville, FL 32218 USA www.biometmicrofixation.com***." |  |  | August 18, 2011 | 2011 | Biomet Microfixation, Inc. | On 09/18/2011, Biomet Microfixation, Jacksonville, FL, recalled their SP-2411, Rubber Band, Lot 307850 because it was labeled and sold as "LATEX FREE" although it contains "LATEX". The firm's HHE identified this as a potential high health risk to the patients or end users with latex allergies. | Error in labeling | Biomet Microfixation, Jacksonville, FL sent a Urgent Medical Device letter dated August 19, 2011, to all affected customers. They were also notified by phone, and followed-up by e-mails starting on 08/18/2011. The letter identified the product the problem and the action needed to be taken by the customer.The customer was instructed If you have any of the affected products in your facility, please remove it and return it immediately to Biomet Microfixation for credit. In order to return these products, please utilize the Return Goods Authorization (RGA) number provided on the reconciliation form.Please follow the instructions on the form and fill out the attached Inventory Reconciliation sheet and FAX to (904) 741-9425. If you have any further questions, please call 1-800-874-7711 or (904) -741-4400, extension 9468, Monday through Friday, 8am-4pm ET. |  | 22 | 22 packs | Worldwide Distribution-- USA (nationwide) including the states of FL, KY, MN, NJ, SD, and VA., and the country of Australia. | Terminated | February 06, 2012 | 172 | 22 | Not_Computer
Z-0013-2012 | 59952 | No | EXB | Ostomy pouch and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Natura Durahesive ConvaTec Moldable Technology Post-Operative /Surgical System (57 mm) |  |  | June 22, 2011 | 2011 | Convatec Inc. | Post-operative kits 57 mm may contain 70 mm pouches. | Labeling mix-ups | The firm, ConvaTec, sent a "MEDICAL DEVICE RECALL" letter dated June 22, 2011 via UPS 2nd day air to all customers and sales representatives that received product. The letter describes the product, problem and actions to be taken. The customers were instructed to immediately stop distributing and quarantine the recalled lot; perform a count of recalled product currently in inventory; complete and return the Recall Response Form via fax to (908) 904-2120 or mail to 200 Headquarters Park Drive, Skillman, New Jersey 08558; return the Recall Response Form even if no recalled product is in inventory; call the toll free number 1-800-582-6514 for return instructions and RGA number-(recall product must be returned in a separate container from any other product returns); mark all return shipments "Recalled Product" and provide supplied RGA number, and forward this letter to customers, if they distributed the product.Should you have any questions, call 1-800-582-6514. |  | 283 | 283 units | Nationwide distribution: AL, AZ, CA, FL, GA, IL, IN, KY, LA, MA, MD, MI, MO, NC, NJ, NV, NY, OH, OK, OR, PA, RI, TN, TX and UT. | Terminated | January 11, 2012 | 203 | 283 | Not_Computer
Z-3173-2011 | 59830 | [6 empty] | BardPort Titanium Implanted Port. |  |  | August 30, 2011 | 2011 | Bard Access Systems | Implanted port has discrepant product information on package labeling. | Labeling Change Control | Bard Access System, Inc. sent a "Label Discrepancy Notification dated August 30, 2011 to all affected customers.The letter included affected product and informed them of the labeling discrepancy. Customers could either use the device or return it to Bard for replacement. For information on this recall call (801) 595-5568. |  | 1375 | 1375 units | Worldwide Distribution: Nationwide Distribution including the states of; AL, AR, AZ, CA, CT, DC, FL, GA, HI, IL, KY, LA, MA, MD, MI, MO, NJ, NV, NY, OK, PA, RI, SC, TN, TX, UT, VA, WI, and WV; and the countries of Canada and Belgium. | Terminated | March 30, 2012 | 213 | 1375 | Not_Computer
Z-3109-2011 | 59366 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | REF 91200, Juggerknot system, Soft Anchors, -Two implants size 1, Biomet Sports Medicine, STERILE. |  |  | July 21, 2011 | 2011 | Biomet, Inc. | The device packaging was punctured by the device, resulting in loss of implant sterility. | Package design/selection | Biomet Sports Medicine notified distributor accounts by telephone on July 08, 2011 with instructions to quarantine all lots of the recalled product. Urgent Medical Device Recall Notice letters dated July 21, 2011 were sent to all affected customers. The letter identified the recalled product, problem, and actions to be taken. The customers were instructed to immediately locate and discontinue use of the recalled product. The distributors are advised to communicate this notice to their customers if the recalled products were further distributed. The letter also states that distributors are responsible for the return of the recalled product. Customers were asked to complete a response form and fax to 574-372-1683, within 3 business days. Customers were instructed to provide phone confirmation at 800-348-9500, extension 3755 or 3756, when the recall notice is received. Further questions should be directed to 574-371-3755 or 574-371-3756, Monday through Friday, 8 AM to 5 PM. |  | 4925 | 4925 | Worldwide Distribution - USA (nationwide) and the countries of: Columbia, Europe and Korea. | Terminated | April 26, 2013 | 645 | 4925 | Not_Computer
Z-3038-2011 | 59465 | No | DRF | Electrode recording catheter or electrode recording probe. | Cardiovascular | Cardiovascular | 510(k) | Duo-Decapolar Electrophysiology Catheter, Manufacturing Part Number D-1171-35-S, Catalog Number D728260RT |  |  | April 29, 2011 | 2011 | Biosense Webster, Inc. | The recall was expanded because between February and April 2011, Biosense Webster, Inc. notified customers that a voluntary Customer Notification had been Initiated due to the inaccessibility of the electronic version of the Instructions for Use (e-IFU) for the "10 ft. Interface Cable, 10 Pin - Shielded Tip Pins to 12 Pin - Blue", Catalog Number 0128709,through the J&J; Gateway (i.e., e-IFU websi | Labeling Change Control | Biosense Webster sent an "URGENT FIELD SAFETY NOTICE" letter dated April 29, 2011 to all affected customers. The letter describes the product, problem, and the action taken by the firm.The letter notifies and advises customers that the updated IFUs for ISMUS Catheter with Auto ID Technology and the Duo-Decapolar catheters have been released and posted on the e-IFU website as of April 22, 2011 and April 21, 2011. Customers were instructed to complete the Acknowledgement Form and return the form to Biosense. Customers with questions or require additionial information were instructed to contact their Biosense Webster representative. |  | 1530 | 1530 units | Nationwide Distribution | Terminated | October 12, 2012 | 532 | 1855 | Not_Computer
Z-1372-2011 | 57736 | No | MLN | Electrocardiograph electrode. | Cardiovascular | Cardiovascular | 510(K) Exempt | Medstorm Adult Radiolucent Zoll Medical (Part # 16381, HeartSync, Ann Arbor, MI. |  |  | December 17, 2010 | 2010 | Heart Sync, Inc | The seal on the electrode pouch may be open, exposing the electrode to the outside environment. The package does have a caution statement that states, " Do not open package until immediately prior to using electrodes. and Electrodes may dry out when removed from packaging and exposed to air. | Packaging | Heart Sync, Inc sent an URGENT MEDICAL DEVICE RECALL letter dated December 23, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to:Examine their inventory and quarantine product that was subject to the recall that was not sealed properly. Contact any customers if they further distributed the product and notify them at once of this product recall. Contact the firm if they had any unused product. The firm would issue a UPS Call Tag to have the product returned to them and the firm would send replacement product as soon as they were notified.Complete and return the Recall Return Response Form to the firm in the attached prepaid envelope.For any questions regarding this recall call 734-213-5530. |  | 300 | 300 | Worldwide Distribution - USA including AK, AZ, CA, FL, IL, KS, LA, MA, ND, NM, OH, TX, WA, WI and WV; and Canada | Terminated | July 13, 2011 | 208 | 300 | Not_Computer
Z-3024-2011 | 56839 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Spacelabs Medical elance Vital Signs Monitor, Model 93300. |  |  | September 23, 2010 | 2010 | Spacelabs Healthcare, Llc | The monitor has a 3 second delay in the waveforms and respiration alarm; the delay is not disclosed in the IFU. | Error in labeling | Spacelabs Healthcare sent an "URGENT-MEDICAL DEVICE CORRECTION" letter dated September 24, 2010 to all affected customers. The letter describes the product, problem, and the action to be taken.The letter includes the revised Operator's Manual as a replacement for customers current manual.For Technical Support call Spacelabs Medical at 1-800-522-7025 and select 2. |  | 2517 | 2517 | Worldwide Distribution-USA (Puerto Rico) and countries of ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA, BAHRAIN, BANGLADESH, BELGIUM, BOLIVIA, CHILE, CHINA, COLOMBIA, CYPRUS, CZECH REPUBLIC, DOMINICAN REPUBLIC, ECUADOR, EGYPT, EL SALVADOR, FRANCE, GERMANY, GREECE, GUATEMALA, HONDURAS, HONG KONG, INDIA, INDONESIA, IRAQ, ISRAEL, ITALY, KENYA, KOREA, REPUBLIC OF, KUWAIT, LEBANON, MACEDONIA, MALAYSIA, NEPAL, NETHERLANDS, NICARAGUA, OMAN, PALESTINE, PANAMA, PARAGUAY, PERU, PUERTO RICO, PHILIPPINES, QATAR, ROMANIA, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SPAIN, SRI LANKA, SWITZERLAND, SYRIAN ARAB REPUBLIC, THAILAND, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, VENEZUELA, and VIET NAM. | Terminated | August 17, 2011 | 328 | 2517 | Not_Computer
Z-1251-2011 | 57604 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Elekta Leksell Stereotactic System, used for localization (spatial reference) for cranial surgery using X-ray or CT and MRI Image data. |  |  | October 02, 2009 | 2009 | Elekta, Inc. | New protocols for MR sequences may result in higher RF energies deposited during MR scanning, generating heat in uninsulated fixation posts. | Labeling design | Elekta sent Important Notice Affected Product: Fixation Posts, A331, dated June 22, 2009. The Important Notice will be added to the Customer User manual upon receipt. |  | 379 | 379 units | Nationwide | Terminated | February 17, 2011 | 503 | 379 | Not_Computer
Z-0147-2012 | 60240 | No | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | Medrad(R) Sterile Disposable Stellant Syringe Kit, Catalog # SDS-CTP-SPK, Lot # 112865. |  |  | October 24, 2011 | 2011 | Medrad Inc | Packaging may exhibit open seals - sterility compromised | Packaging | Medrad sent a " Urgent Medical Device Recall letter" dated October 24, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The customer was instructed if they have any of the affected product lot, please contact MEDRAD customer support to make arrangements for the return of the product. Call 1 -800-633-7231, select option 5, then select option 1. Indicate that you are calling about the affected syringe kit recall.Complete and sign the attached response form and fax it back to MEDRAD at (412) 406-0942.Please complete and fax this form even if you do not have any of the referenced product so we can maintain proper accountability for all syring kits that have been shipped. |  | 17380 | 17,380 | Nationwide Distribution -- including the States of : AL, AZ, AR, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, KS, KY, LA, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, VT, VS, WA, WV, and WI. | Terminated |  |  | 17380 | Not_Computer
Z-3053-2011 | 59693 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Katzen" Infusion Wire, 0.035 in x 180 cm, |  |  | August 01, 2011 | 2011 | Boston Scientific Corporation | Boston Scientific is voluntarily recalling one lot/batch of Katzen" Infusion Wires because we have identified that the label on the carton and the inner pouch indicated an incorrect device length. The labels indicate a device length of 180cm while the actual device length is 145cm. | Labeling mix-ups | Boston Scientific sent an "Urgent Medical Device Recall Removal" letter dated August 1, 2011 to all affected customers.The letter described the problem and product being recalled and advised customers to segregate the affected product immediately and return it to Boston Scientific. Boston Scientific requested the completion and return of the Reply Verification Tracking Form.For additional information on this recall call Boston Scientific Quality Systems at (763) 494-1133. |  | 33 | 33 | (USA) Nationwide distribution including the states of AZ, AR, CA, CO, FL, GA, IL, IN, NV, NY, NC, OH, PA, RI, and TX. | Terminated | August 10, 2012 | 375 | 33 | Not_Computer
Z-1236-2011 | 57505 | No | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | Stryker 5190 TPS Series Burs. Powered simple cranial drills, burrs, trephines, and their accessories. |  |  | November 30, 2010 | 2010 | Stryker Instruments Division of Stryker Corporation | Product was packaged with an incorrect packaging material and therefore has a potential to be non-sterile which, if contaminated, could lead to a patient infection. | Packaging | The firm, Stryker, sent a "MEDICAL DEVICE RECALL NOTIFICATION" letter dated December 2, 2010 to all customers. The letter described the product, problem and the action to be taken by the customers. The customer were asked to complete and fax in the business reply form indicating they have received the notification and to indicate whether or not they have affected product on hand, and send back any affected product to Regulatory Affairs at Stryker Instruments. Stryker will then send a replacement product to the customer.Upon receipt of the completed business reply form, the customer will be contacted to arrange return of any affected product they have in their inventory. Once the affected product is received, a replacement 5190 TPS Series Bur will be sent to the customer.For questions regarding this recall, please contact Stryker Instruments at 1-800-800-4236 ext. 3584 or kathleen.kelly.smith@stryker.com. |  |  |  | Worldwide distribution: USA including states of: CA, FL, LA, MI, MT, NE, NH, OH, PA, SD, TN, TX, VT, and WA; and countries including: Sweden and Australia. | Terminated | April 13, 2011 | 134 |  | Not_Computer
Z-0049-2012 | 59978 | [6 empty] | ArthroCare brand MiniMagnum Knotless Fixation Implant Device with Inserter Handle Plus SpeedStitch Magnum Wire Suture Cartridge; Catalog Number: OM-6007; |  |  | May 11, 2011 | 2011 | ArthroCare Corporation | ArthroCare initiated the recall of one lot of the MiniMagnum Knotless Implant Fixation Device where the incorrect diver block was packed with the implant. | Packaging process control | The firm, ArthroCare Corporation, sent a letter dated May 10, 2011, to its customers. The letter described the product, problem and actions to be taken. The customers were asked to immediately return the product; check their inventory, refer to the copy of the Mini Magnum outer label to help identify the product, and if the affected product lot is in their inventory, notify ArthoCare Product Support at (800) 797-6520, press 2 and return the product immediately. The letter also stated, "In addition, if you do not have any product from this lot remaining in your inventory, we request that you contact ArthroCare at the number above, or by e-mail at ra@arthrocare.com, to inform us of the lack of remaining product. This will ensure that you will not be contacted again and that we have accounted for product shipments from this lot." If you have any questions, please call the Chief Regulatory Officer and Senior Vice President, Quality Systems and Assurance at (408) 735-6229. |  | 32 | 32 units | Nationwide distribution: USA including states of: CA, FL, NY, OH and OK. | Terminated | October 24, 2011 | 166 | 32 | Not_Computer
Z-1095-2011 | 57660 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Triathlon Femoral Distal Fixation Peg; Catalog Number 5575-X-00; Manufactured by Stryker Howmedica Osteonics Corp., Howmedica Osteonics Corp., 325 Corporate Drive, Mahwah, NJ 07430 USA A Subsidiary of Stryker Corp. |  |  | December 22, 2010 | 2010 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics has become aware that the labels on the Triathlon Distal Femoral Fixation Pegs contain an incorrect expiration date of January 2015 (2015-1) when the actual expiration date should read June 2014 (2014-6). | Incorrect or no expiration date | The firm, Stryker, sent a "UPDATED (Corrected Lot Code) URGENT PRODUCT RECALL" letter dated January 6, 2011, via Fed Ex ,with return receipt to all customers. The letter described the product, problem and action to be taken by the customers. The customers were instructed to fax the attached Product Recall Acknowledgement form to (201) 831-6069; contact the hospitals in their territory that have the affected product to arrange return of the product, and return the affected product to the attention of Regulatory Compliance, Stryker Orthopaedics, 325 Corporate Drive Mahwah, NJ 07430.If you have any questions, feel free to contact the Manager, Divisional Regulatory Compliance at (201) 972-2100 or the Associate Product Manager, Knee Marketing Team at (201) 831-5028. |  | 15 | 15 units | Nationwide distribution. | Terminated | October 03, 2012 | 651 | 15 | Not_Computer
Z-1607-2011 | 57759 | No | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | CIVCO Latex Free NeoGuard Transducer Cover, Ref 676-117, Sterile (2.6 x 30cm) cover, packaged 24/box. The firm name on the label is CIVCO, Kalona, IA. |  |  | January 21, 2011 | 2011 | Civco Medical Instruments Inc | Non-sterile transducer covers were incorrectly labeled as sterile and contained incorrect temperature storage information. | Labeling design | The recalling firm notified their customers via phone and e-mail letter entitled Urgent - Medical Device Correction, dated 1/21/11, explaining the reason for recall. The letter stated that the firm will replace the affected product and requested that customers dispose of the product from the affected lots. A response form was enclosed to report final disposition of the product and should be returned to the firm. If there are any questions, customers can call their customer service representative at 800-445-6741 between 8 am and 5 pm CST. |  | 31 | 31/24-cover boxes | Nationwide Distribution -- WA, CA, and MO. | Terminated | July 09, 2013 | 900 | 31 | Not_Computer
Z-1231-2011 | 57550 | No | GYK | Shunt system implantation instrument. | Neurology | Neurology | 510(K) Exempt | Integra Peritoneal Introducer Sheaths, 61 cm;Integra NeuroSciences/REF 901-124; Manufacturer Integra NeuroSciences Implants S.A., 2905 Route Des Dolines, 06921 SOPHIA ANTIPOLIS Cedex, France; Distributed by Integra NeuroSciences, 311 Enterprise Drive, Plainsboro, NJ 08536, USA |  |  | December 22, 2010 | 2010 | Integra LifeSciences Corp. | Peritoneal Introducer Sheaths were incorrectly labeled as 61 cm instead of 46 cm. | Labeling mix-ups | Integra Life Sciences sent an Urgent Product Recall letter dated December 22, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to inspect their inventory and quarantine any of the affected product. Contact Integra Customer Service at (800) 654-2873, select option 2 to arrange for product returns. Customers were asked to complete the attached "Recall Acknowledgement and Return Form" indicating the current status of quantities of affected product in their inventory. For questions call (609) 936-2237. |  | 95 | 95 units | Nationwide Distribution including CO, FL, GA, MO, NC, NJ, PA, TX, UT, and WI. | Terminated | May 07, 2012 | 502 | 95 | Not_Computer
Z-1598-2011 | 57697 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Hospira Plum XL Micro/Macro Infusion Pump, Single Channel, with a 0.1-999 mL/hr flow rate range; list number 11846 |  |  | February 14, 2011 | 2011 | Hospira Inc. | Hospira has received reports of the Plum XL, XLM and XLD Infusion Pumps with no audible alarm at the low audio level setting. | Labeling design | Hospira sent Urgent Device Recall letters dated February 14, 2011 to their customers on the same date, informing them that Hospira had received complaints of the Plum XL, XLM and XLD infusion pumps in which the audible alarm failed. If the audible alarm fails and the user does not notice the visual alert, the user may not be aware of the change in pump status such as air-in-line or occlusion. This may result in a delay or interruption of therapy which may result in serious injury and/or death. Clinicians were instructed to weigh the risk associated with continued use of the device versus removal from service. If they elect to continue the devices, they were provided with instructions to perform alarm tests prior to each clinical use (during their pump cleaning process). If the alarm is not audible, they were instructed to discontinue use of the pump and contact Hospira Global Product Safety and Complaints at 1-800-441-4100. They were also instructed to ensure that all personnel in their facility are following Hospira's Technical Service Manual for routine decontamination of the pumps and cleaning of the buzzers during preventative maintenance, including inspection of the buzzers for leg lifting. The customers were requested to contact Hospira Advanced Knowledge Center at 1-800-241-4002 for technical assistance.The accounts were also requested to complete the attached reply form and return it to Hospira via fax at 1-866-382-4228. |  | 23385 | 23,385 pumps | Worldwide distribution, including United States, the U.S. Virgin Islands, Puerto Rico, Argentina, Australia, Bahamas, Bahrain, Barbados, Belgium, Bosnia, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Croatia, Ecuador, EI Salvador, Finland, France, Greece, Guatemala, Hong Kong, Hungary, Ireland, Israel, Italy, Jordan, Korea, Kuwait, Luxembourg, Mexico, New Zealand, Oman, Panama, Pakistan, Peru, Philippines, Poland, Portugal, Qatar, San Marino, Saudi Arabia, Singapore, Slovakia, South Africa, Spain, Sweden, Taiwan, Turkey, United Kingdom, Uruguay, United Arab Emirates and Venezuela. | Open, Classified |  |  | 205689 | Not_Computer
Z-3045-2011 | 59713 | No | JOJ | Sterilization process indicator. | General Hospital | General Hospital | 510(k) | VerifyÂż SixCess Indicator Strips, Model #: PCC005, PCC006, PCC007, PCC008, PCC031, PCC033, PCC034, PCC037 &amp; PCC038, Mfr. By: Albert Browne, LTD, Chancery House 190 Waterside Rd. Hamilton Industrial Park, Leicester, Leicester, LE5QZ, United Kingdom. |  |  | March 16, 2011 | 2011 | Steris Corporation | During an inspection the firm was notified that language contained in labeling and promotional literature directs Customers that the products can be used to replace biological indicators and/or release loads containing implants. | Labeling False and Misleading | STERIS sent a Customer Notification letter dated March 16, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review and follow the Challenge Packs' revised instructions for use enclosed as Attachment A (270-4) and Attachment B (275F 3).For any questions regarding this recall customers were instructed to contact their STERIS Account Manager or call 440-392-7421. |  | 2500 | 2,500 (one per customer) | Worldwide Distribution - USA (nationwide) and the countries of Lebanon and Qatar. | Terminated | March 19, 2013 | 734 | 2500 | Not_Computer
Z-0144-2012 | 60057 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | BardÂż Ventralex" ST Medium Circle with Strap Hernia Patch Diameter 6.4cm /2.5", Catalog Number: 5950008. |  |  | October 04, 2011 | 2011 | Davol, Inc., Sub. C. R. Bard, Inc. | Mislabeled: BardÂż Ventralex" ST Medium Circle was packaged in a pre-printed carton identified for a BardÂżVentralex" ST Large Circle. | Packaging process control | Davol Inc. issued an Urgent; Medical Device Recall letter on 10/4/11 for the mislabeled Bard Ventralex ST Medium Circle. The letter identified the affected product and explained the reason for recall. Customers are to immediately examine their inventory and quarantine any affected product. In addition, if any product was further distributed then those customers should be notified of the recall as well. Accounts are to contact Davol Customer Service at 1-800-556-6275 for instructions on how to return product to Davol. The attached recall effectiveness check form should be completed and returned. Questions should be directed towards Davol Customer Service Department at 1-800-556-6275 or C. R. Bard's Medical Services &amp; Support Department at 1-800-562-0027. |  | 96 | 96 units | Nationwide Distribution | Terminated | September 20, 2013 | 717 | 96 | Not_Computer
Z-3044-2011 | 59375 | No | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | TSRH(R) Spinal System SCREW, REF 837-645, Size: 6.5mm x 45mm, Mat'l: TITANIUM ALLOY, Rx only, QTY: 1 EA, Medtronic Sofamor Danek USA. |  |  | July 12, 2011 | 2011 | Medtronic Sofamor Danek USA Inc | Screws are 5 mm longer than the size etched on the screw. | Error in labeling | On 07/12/2011, the recalling firm, Medtronic notified the Sales Managers and Sales Force by conference call and email regarding the imminent notification and follow-up activities. A "VOLUNTARY RECALL" notification letter was sent to Risk Managers and Sales Representatives on 07/13/2011 with a delivery date of 07/14/2011. The letter states to immediately remove the recalled product from use and return to Medtronic. The letter includes a questionnaires form to be completed and returned as per the instructions. If appropriate, personal follow up with Sales Representatives will be made to assure that Risk Managers understand the recall communication and to facilitate return of product and the associated questionnaire documentation. For questions, contact your Medtronic Sales Representative or the Global Quality Department at 800-876-3133, extension 3173. |  | 17 | 17 units | Nationwide Distribution (USA) including the states of: OH, NY, and WI | Terminated | July 20, 2012 | 374 | 17 | Not_Computer
Z-3207-2011 | 59290 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Synthes T1 Matrix Locking Cap of the Synthes Matrix System, Part Number 04.632.000. |  |  | March 11, 2011 | 2011 | Synthes USA (HQ), Inc. | Labeling Correction -- The locking cap was not able to be loosened, resulting in the breakage of the screwdriver or the rod to be cut and the assembly to be removed (screw, cap, and rod). This poses a potential risk of significant prolongation of OR time. | Labeling Change Control | The firm initiated this action on March 11, 2011 by issuing an Urgent: Medical Device Labeling Correction letter to all consignees. The letter identified the affected product and stated the reason for the correction. It also instructed the consignee to review and attach the supplementary technique page to any MATRIX Spine System Technique Guides that they may have had in their facility. Customers are to complete the Verification Section of the letter and return it to the firm via mail, fax, or e-mail. If customers have questions, they can call 610-719-5063 or their Synthes Spine Sales Consultant. |  |  |  | Nationwide Distribution -- AL, AZ, AR, CA, CO, CT, DE, FL, GA, ID, IL, IN, IA, KS, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WV, and WI. | Terminated | December 18, 2014 | 1378 |  | Not_Computer
Z-0893-2011 | 54768 | No | RFZ |  |  |  |  | Hemochron Signature Elite Whole Blood Microcoagulation (IVD) System. |  |  | March 31, 2010 | 2010 | International Technidyne Corp. | There is a mismatch between the class of the laser barcode scanner and the labeling of the Hemochron Signature Elite: some of the devices were inadvertently equipped with class II laser barcode scanner instead of the intended class I barcode scanner. | Labeling mix-ups | Urgent - Medical Device Recall - Hemochron Signature Elite letters, dated February 24, 2010, by Federal Express to all direct customers. The letter informed customers of the issue and stated that the problem does not negatively affect the performance of the product and to continue using the product as intended. In addition, customers are to notify the appropriate personnel about the notice, complete and return the attached form, and an ITC representative will contact customers to discuss arrangements to re-label their devices. Customers can contact ITC Technical Support at 800-579-2255 (US) or 732-548-5700 ext 4700 (International) if they have any questions regarding this issue and notice. Customers can also email sigelitescan@itcmed.com with questions. |  | 1640 | 1640 units (1303 US, 337 foreign) | Worldwide Distribution -- United Arab Emirates, Belgium, Canada, Switzerland, Germany, Denmark, France, United Kingdom, Ireland, Iceland, Italy, Kuwait, Liechtenstein, Netherlands, Oman, Qatar, Saudi Arabia, Sweden, and Singapore. | Terminated | September 25, 2012 | 909 | 1640 | Not_Computer
Z-3074-2011 | 59680 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Habib, 4x Kit, Bipolar Resection Device RITA (Radiofrequency Ablation Devices). |  |  | August 30, 2010 | 2010 | Angiodynamics, INC | The labeling contains errors which result in the products being misbranded | Labeling False and Misleading | AngioDynamics, Inc. sent out a "Medical Device Recall" letter dated August 30, 2010 to all affected customers.The letter included; products affected, problem and asked them to return the product back to AngioDynamics. Customers were asked to fill out and return the included Recall Reply Form. For question on this recall contact AngioDynamics at (800) 772-6446. |  | 132 | 132 units | Worldwide Distribution: (USA) Nationwide distribution and the countries of Austria, China, Denmark, Egypt, France Germany, India, Italy, Latvia, Netherlands, Poland, Russian Federation, South Africa, Switzerland and the United Kingdom | Terminated | March 12, 2012 | 560 | 769 | Not_Computer
Z-3268-2011 | 59467 | No | LFR | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Contour, Contour TS Blood Glucose Test Strips and Bonus Packs |  |  | August 12, 2011 | 2011 | Bayer Healthcare, LLC | All 10- or 25- count glucose test strip vials, either loose vials or found in patient starter kits of Contour, Contour TS and Contour" USB blood glucose meters are being recalled a packaging issue which, under certain conditions, can cause some test strips packaged in some small-count vials to operate outside of US FDA-cleared performance specifications and produce a low-bias reading during blood | Packaging | Bayer sent an "URGENT DIABETES TEST STRIP RECALL" letter dated August 12, 2011, to its customers with the name of the test strips, sizes recalled, reason for the recall, potential injury, and instructions to discontinue use and return the test strips to Bayer. The recall letter also included a recall response form that was to be faxed back to Bayer at 1-800-876-2243. For questions call Customer Logistics at 1-800-348-2637. |  | 197342756237020 | 3,617,287 vials/ 54,555,460 test strips | Nationwide Distribution and Puerto Rico | Terminated | July 02, 2012 | 325 | 197342756237020 | Not_Computer
Z-0174-2012 | 60255 | No | KZH | Syringe needle introducer. | General Hospital | General Hospital | 510(k) | 1.0mL 28 G x 1/2" blister packaged insulin syringe Catalog Number 329424; Lot Number 0039920 |  |  | June 15, 2011 | 2011 | Becton Dickinson &amp; Company | Experiment labels may be on shelf cartons of 1.0ml 28G x 1/2 blister packaged insulin syringes. | Labeling mix-ups | The firm, BD, sent an "PRODUCT RECALL NOTIFICATION" letter and response form dated June 15, 2011 to their consignee/customers. The letter describes the product, problem and actions to be taken. The customers were instructed to immediately review their inventory for the affected product, remove the product from inventory, quarantine the product and return it to BD; if the product was further distributed, they should identify their customers and notify them at once, and complete and return the attached Tracking/Verification Form via fax to: 201-847-4853.If you have any questions, contact Customer Service at 1-800-237-2762, Monday -Thursday, 8:00am-6:30pm and on Friday, 8:00-5:30pm Eastern Standard Time. |  | 10230 | 10,230 shelf cartons (1,0230,000 syringes) | Nationwide Distribution. | Terminated | July 27, 2015 | 1503 | 10230 | Not_Computer
Z-3042-2011 | 59525 | No | LYL | Soft (hydrophilic) contact lens care products. | Ophthalmic | Ophthalmic | 510(k) | Equate Muti-Purpose Solution, For Soft Contact Lenses, Sterile, 12 Fl oz. (355 mL), Distributed by Wal-Mart Stores, Inc., Bentonville, AR 72716 |  |  | June 30, 2011 | 2011 | Bausch &amp; Lomb Inc | Muti-purpose solution bottles, lot GC1107 were packaged in a carton that was marked with an incorrect lot code and expiration date. | Packaging | Bausch &amp; Lomb Inc. notified the affected customer via teleconference and e-mail on June 30, 2011. The e-mail describes the product, problem, and actions to be taken by the customer.The customer was instructed to hold and stop shipments of any inventory of the affected product. An official recall was initiated with the customer on July 8, 2011 via e-mail and teleconference.Customer was asked to review their inventory and store returns, and send any of the affected product to Bausch &amp; Lomb Returns Dept., 130 Commerce Drive, Greenville, SC 29615. |  | 1200 | 1200 bottles | Distribution-in the state of AK only | Terminated | November 16, 2011 | 139 | 1200 | Not_Computer
Z-3113-2011 | 59632 | No | JXN | Propoxyphene test system. | Toxicology | Toxicology | 510(k) | DrugCheckÂż, Distributed by Express Diagnostics Int'l Inc., 1550 Industrial Drive, Blue Earth, MN 56013. Following cup Model numbers involved in the recall: 60670, 60700, 60713, 60717, 60870, 60923, 61012, 61019, 61023, 61024, 61040, 61042, 61057, 61102, 61105, 61203, 61204, 61206, 61208, 61301, 61303, 61401, 70603, 70809, 71032, 71043, 71055, 71065, 71202, 71204, 71205, 71401, 71402, 60718-3, 60720-6, 60925-A, 60927-3, 60930-A, 60980-3A, 61002-3, 61010-4, 61023-A, 61028-3A, 61039-4, 61040-A, 61055-A, 61056-3, 61057-6, 61100-3, 61102-3,61104-3,61105-5 CRL,61120-3,61202-4,61204-4,61204-A,61206-3A,61207-4,61302-3, 61303-A, 61401-A, 70660-3, 70716-3, 71045-A, 71055-6, 71102-3A, 71204-4, 71206-3A. |  |  | October 04, 2010 | 2010 | Express Diagnostics Int'l., Inc. | Express Diagnostics Int'l, Inc (EDI), is recalling DrugCheckÂż Drug Screen Cups containing buprenorphine (BUP) and propoxyphene (PPX) because they must be labeled as "forensic use only", meaning they cannot be distributed or used for treatment or diagnostic purposed in the United States. | Labeling False and Misleading | The firm, Express Diagnostics Int'l, Inc., sent a "RECALL NOTICE" letter dated October 4, 2010 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to notify their customers of the labeling change, and complete and return the DISTRIBUTOR/CUSTOMER CONFIRMATION FORM via fax to Express Diagnostics Int'l, Inc., at 1-507-526-2252. The letter also noted that "Further shipment of product containing BUP and/or PPX cannot continue until Express Diagnostics Int'l, Inc. receives the signed form."If you have any questions, call 507-526-3951. |  | 196750 | 196,750 | Nationwide distribution:AZ, CA, CT, GA, MD, MA, MN, MO, NH, NJ, OR, SC, and WA. | Terminated | July 06, 2012 | 641 | 196750 | Not_Computer
Z-0164-2012 | 60143 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | BioPro Femoral Head 28 mm +18, Item 19010. |  |  | September 21, 2011 | 2011 | Biopro, Inc. | BioPro Femoral Head Outer Packaging was found lacking proper seal ,compromising sterility. | Packaging | Biopro, Inc notified all affected customers via telephone on September 21, 2011 with instructions to quarantine and return affected product to the manufacturer. The firm sent a follow up "RECALL NOTICE" dated September 29, 2011 to all affected customers. The letter described the product, problem, and actions to be taken by the customers. The letter gives instructions on how to return affected product. Contact Customer Service at 1-800-252-7707 for questions regarding this recall. |  | 8 | 8 units | USA (nationwide) including the states of FL, MI, and TX. | Terminated | February 15, 2013 | 513 | 12 | Not_Computer
Z-3102-2011 | 57590 | No | EZO | Urethrotome. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | OLYMPUS Pediatric Knives and Resection Electrodes |  |  | December 22, 2010 | 2010 | Olympus America Inc. | The device sterile package seal shows evidence of being breached or was breached | Packaging process control | Olympus America Inc. sent an "URGENT PRODUCT RECALL" letter dated December 22, 2010 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter instructs customers to check their inventory for the affected product and to destroy or discard, if found. A questionnaire was attached to the letter for customers to complete and return via fax to Olympus at 484-896-7128.Contact the firm at 484-896-5688 for questions regarding this recall. |  | 1305 | 1305 devices | USA (nationwide) including the states of AL, AR, CA, CO, CT, FL, IA, ID, IL, IN, KS, KY, LA, MA, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NV, NY, OH, OR, TX, VA, VT, and WI. | Terminated | September 28, 2012 | 646 | 1305 | Not_Computer
Z-0060-2012 | 59998 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | ANCHORLOK(R) LEADING EDGE(R) SOFT TISSUE ANCHOR SYSTEM, 2.5mm, 2.0 SUTURE, REF 5100-006, 1 each, STERILE R, Wright Medical Technology, Inc., 5677 Airline, Road, Arlington, TN 38002 |  |  | September 20, 2011 | 2011 | Wright Medical Technology Inc | The product is labeled with an expiration date without adequate justification. | Incorrect or no expiration date | The firm, Wright Medical Technology, Inc., sent an "URGENT: PRODUCT RECALL NOTIFICATION" letter dated September 20, 2011 to all affected consignees. The letters identified the products, the problem, and the actions to be taken. The letter instructed consignees to identify affected inventory. If recalled products are located consignees are instructed to contact customer service at 1-800-238-7117, for specific instructions on how to return the recalled products. The letter asked that this communication be forward to any associated sites if the recalled products were further distributed. Consignees are to complete and return the enclosed notification form by fax to 901-867-7401 or by email to: cathy.park@wmt.com. For questions regarding this recall contact 901-867-4324. |  | 596 | 596 units | Worldwide Distribution - USA (nationwide) and countries of: Argentina, Canada, Denmark, France, Great Britain, Ireland, Japan, and Spain. | Terminated | July 17, 2012 | 301 | 11960 | Not_Computer
Z-1228-2011 | 57679 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Thomas Medical Products, Acuity Bleed back Control Valve, Acuity Rotating Hemostasis Valve, Acuity Universal Cutter, and Contamination Sleeve for use during vascular interventional procedures. Contamination Sleeve part no. CON-006-02; Acuity Bleedback Control Valve part no. CSP-001-00; Acuity Rotating Hemostasis Valve part no. CSP-002-01; Acuity Universal Cutter part no. CSP-003-00; Acuity Universal Cutter part no. CSP-003-01. |  |  | December 16, 2010 | 2010 | Thomas Medical Products Inc | Compromised sterility/weak or open seals. | Packaging process control | GE Healthcare Thomas medical products Recall notification dated 16 December 2010 (GEHC Ref # 35003) was sent by letter via UPS Next Day Air and/or USPS certified mail. |  | 28679 | 28,679 for all products | Distribution Nationwide USA and Germany. | Terminated | August 08, 2012 | 601 | 28679 | Not_Computer
Z-0042-2012 | 59891 | No | MSQ | Campylobacter fetus serological reagents. | Microbiology | Microbiology | 510(k) | IDkit:HP" Breath Test for Detection of H. pylori, Manufactured by: Exalenz Bioscience Ltd 4 Ha' maayan St. Modiin, Israel 71700, Distributed by: Exalenz Bioscience Inc., 1313 N. Market St., Sute 5100, Wilmington, DE 19808 USA |  |  | June 07, 2011 | 2011 | Exalenz Bioscience Ltd | The product is in violation of the FD&C; Act in that it was illegally marketed without proper clearance (no 510(k) submitted). | No Marketing Application | The firm, Exalenz Bioscience Ltd. notified consignees located in the United States of the recall by telephone, electronic mail, or written correspondence between June 9, 2011 and June 30, 2011. The letter informed consignees that Exalenz will immediately cease marketing the IDkit:HP TWO test in the United States. Consignees were instructed to return shipment of the IDkit:HP TWO test kits that they have in their possession to the companys US warehouse. For question call the VP of North American Sales at 704-770-6162 or VP of Marketing at 617-794-6936. |  | 1240 | 1240 units | Nationwide Distribution (USA) - including the states of: FL, IL, LA, MA, MS, NY, TN, TX and VA. | Terminated | July 06, 2012 | 395 | 1240 | Not_Computer
Z-0160-2012 | 57234 | No | EKM | Gutta percha. | Dental | Dental | 510(K) Exempt | Microseal Master Cone Part Number: 815-9033 |  |  | September 09, 2010 | 2010 | Ormco Corporation | The recall was initiated because SybronEndo has confirmed that the Microseal Master Cones has been mislabeled. | Error in labeling | Sybron Dental Specialties, Inc. sent an URGENT MEDICAL DEVICE RECALL letter dated September 20, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were requested to return the affected lots of product.. Products, would be replaced at no charge or credited to their account. Customers were instructed to contact SybronEndo Customer Care at 1-800-346-3636 directly to handle the arrangements of a quick return and replacement. The affected product should be returned to SybronEndo at the following address. Please label returned product "RECALLED PRODUCT Attention: Customer Returns"SybronEndo 1332 South Lone Hik AvenueGlendora, CA 91740ADDITIONALLY, THEY KINDLY REQUEST CONSIGNEE COOPERATION IN COMPLETING AND FAXING BACK THE ENCLOSED ACKNOWLEDGEMENT/RETURN FORM. THIS FORM MUST BE COMPLETED AND RETURNED WHETHER OR NOT YOU HAVE ANY PRODUCT TO RETURN.Customers were also asked to identify and recover the affected product lots listed above that may have been shipped to their customers.For questions regarding this recall call 909-962-5600. |  | 200 | 200 units | Worldwide Distribution - USA including AR, CA, FL, PA, GA, IL, TX, WA, NY, IN, AL &amp; MT; Internationally to Canada, CIS (Commonwealth of Independent States, Israel, United Kingdom, Sweden, France, Germany, and Italy. | Terminated | November 21, 2011 | 438 | 200 | Not_Computer
Z-3009-2011 | 59275 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Siemens Axiom Artis U mobile x-ray system |  |  | June 29, 2011 | 2011 | Siemens Medical Solutions USA, Inc | When the heat unit of the angiography system reaches or exceeds 65%, the cool down time display will appear with an acoustic signal. If users ignore the warning and continue the radiation, the tube may be autoprotected when it reaches its temperature limitation causing the X-Ray to lock until the tube cools down. Following a tube auto-protection, an error message "Door is Open" or "Err 421/240: | Labeling False and Misleading | Siemens sent a "Customer Safety Advisory Notice" letters to all affected users dated June 13, 2011. The letter identified the product, the problem, and the actions to be taken. Siemens included in the notice an addendum to the instruction for use to each system and request customers to take special attention. Customers were instructed to place the Safety Notice in the System's instruction for use. All appropriate personnel should be notified of the correction notice. If product was distributed further, customers were to forward the safety notice to the new owner and inform the firm of the new owner. |  | 18 | 18 | Nationwide Distribution (USA) including the states of: CA, CT, FL, KS, MN, MO, NJ, NY, NC, PA, and UT. | Terminated | January 24, 2014 | 940 | 18 | Not_Computer
Z-3080-2011 | 59400 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Model 754/754M Disposable MRI Ventilator Circuit |  |  | May 27, 2011 | 2011 | Impact Instrumentation, Inc. | Model 754/754 M Disposable MRI Ventilator Circuits are being distributed without market clearence (510k). | Labeling False and Misleading | Impact Instrumentation Inc. sent an "Urgent Product Recall" dated May 27, 2011 to all affected customers.The letter informed customers the product is not a cleared product and to return the product back to Impact Inst. for credit. A notification form is included in the letter to be return to the firm.For information on this recall contact the firms Customer Service Department at (800) 969-0750, 8:30 am to 5:00 pm EST. |  | 492 | 492 units | Nationwide Distribution; including the states of CA, FL, IL, NJ, OH, TX, UT, MN, MO, MS, OH, OR, TX, AR, and PA. | Terminated | May 23, 2013 | 727 | 492 | Not_Computer
Z-3187-2011 | 59764 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ENDOPATHÂż ETS Compact Flex (240 mm shaft length) Linear Cutters, Model Number(s): SCW45; CTS45; SCB45; CTS45NK; 6CB45 and SCG45. Mfr. By: Ethicon Endo-Surgery, LLC, Guaynabo, PR 00969 |  |  | July 29, 2011 | 2011 | Ethicon Endo-Surgery Inc | Ethicon Endo-Surgery initiated a voluntary global recall for specific production lots of ENDOPATHÂż ETS Compact Flex 45mm articulating linear cutters because the possibility exists that damage to the blister pack may have compromised the sterility of the device. | Packaging | Ethicon Endo-Surgery Inc sent an URGENT DEVICE RECALL letter dated August 11, 2011 to all affected customers. The letter identified the product, the problem, the action to be taken by the customer. Customers were instructed to examine their inventory immediately to determine if they have any of the affected product. Fill out the Business Reply Form and fax it to 1-866-853-2802 if they have any of the affected product. To ensure adequate identification of the return product at their facility customers should complete the Customer Purchase Order Field on the Business Reply Form. For any questions regarding this recall call 1-866-708-2330 and reference Event 2471. For clinical or product support call 1-800-USE-ENDO, option 6, or contact your local sales representative. |  | 27582 | 27,582 instruments | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, VA, WA, WI and WV and the countries of ARGENTINA, AUSTRALIA, BELGIUM, BRAZIL, CANADA, CHINA, CZECH REPUBLIC, GREECE, INDIA, ITALY, JAPAN, REPUBLIC OF KOREA, LIBYAN ARAB JAMAHIRIYA, MALAYSIA, MEXICO, POLAND, PUERTO RICO, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SPAIN, THAILAND and UNITED ARAB EMIRATES | Terminated | February 15, 2013 | 567 | 27582 | Not_Computer
Z-0152-2012 | 60100 | No | DKZ | Amphetamine test system. | Toxicology | Toxicology | 510(k) | ProScreen PSCup-5M Waived, Catalog Number: PSCup-5M Waived |  |  | June 07, 2011 | 2011 | Ameditech Inc | The recall was initiated because Ameditech, Inc. received a complaint from US Diagnostics (USD) that First Check Drug Screen Cups were packaged in the ProScreen Cup boxes. Some of US Diagnostics' customers received ProScreen 5-panel cups that contain First Check THC cups. The recalled products are mislabeled. If the USD customer receives the FirstCheck tes, there will be no instructions for use | Packaging process control | The recall communication was initiated on June 9, 2011 with Ameditech Inc. forwarding an Urgent Medical Device Recall notification with an attached Verification Form (via phone and written notification) to the affected distributor. The letter provides the distributor with an explanation of the problem identified and actions to be taken. The distributor was required to: cease distribution of the affected lot and provide remaining inventory count to Ameditech; destroy any remaining inventory of these lots and document destruction per site requirements; and ensure that all user of the device have received a copy of the Urgent Medical Device Recall. In addition, the distributor was instructed to complete and return the enclosed Verification Form with in 10 days.For any questions about the information contained in this notice, contact Ameditech at amd.complaint@alere.com or call 858-535-1968. |  | 2000 | 2000 test cups (800 kits) | Nationwide Distribution (USA) including the state of Alabama. | Terminated | January 18, 2012 | 225 | 2000 | Not_Computer
Z-1391-2011 | 57805 | Yes | CAC | Autotransfusion apparatus. | Anesthesiology | Anesthesiology | 510(k) | Haemonetics cardioPAT Cardiovascular - autotransfusion System |  |  | January 25, 2011 | 2011 | Haemonetics Corporation | Update Instructions for Use for the CardioPAT&reg; Manual for Leak Detection. User may mis-interpret the air-leak rate displayed by the cardioPAT device when making a determination about chest-drain removal, which could result in the patient experiencing a pneumothorax. | Labeling design | Haemonetics issued a Field Correction Notice on 1/25/11 . with the Air Leak Detection addendum and index be inserted into the appropriate document, applied to all cardioPAT@ devices on site. The serial number for each cardioPAT device on site should be recorded on the attached acknowledgementand the completed acknowledgement should be returned to your Haemonetics Sales Consultant. A Haemonetics Sales Consultant will train on the content of these items and insert them into the Operator Manuals |  | 309 | 309 units | Worldwide distribution, including USA, Australia, Belgium, Canada, Switzerland, China, Germany, Finland, France, United Kingdom, Hong Kong, Israel, Italy, Japan, Netherlands, Russian Federation, and Sweden. | Terminated | January 14, 2013 | 720 | 309 | Not_Computer
Z-3085-2011 | 59469 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Microline Surgical ReNew Laproscopic Long Fenestrated Grasper Forceps Tip |  |  | July 18, 2011 | 2011 | Microline Surgical, Inc. | Mislabeled: ReNew product number 3262 (Long Fenestrated Grasper) was packaged and labeled asproduct number 3642 (Super-Atrau Fenestrated Grasper). ReNew product number 3642 (Super-Atraumatic Fenestrated Grasper) was packaged andlabeled as 3262 (Long Fenestrated Grasper | Packaging process control | Microline Surgical sent a RECALL NOTICE letter dated July 18, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately examine their inventory and quarantine product subject to the recall. If customers further distributed the product, customers should identify their accounts and notify them immediately of the product recall. Customers were instructed to contact Microline Customer Service to obtain a Return Authorization Number for return of the recalled item. Customers were asked to complete and return the enclosed Recall Return Response Form as soon as possible. For any questions customers were instructed to call the Regulatory Affairs Manager at 978-867-1726. |  | 170 | 170 units | Worldwide Distribution - USA including FL, MI, MN, TN and WI and the countried of Canada, The Netherlands, Thailand, UK, Norway, and Italy. | Terminated | September 09, 2011 | 53 | 470 | Not_Computer
Z-3188-2011 | 58478 | No | KIH | Cement dispenser. | Orthopedic | Orthopedic | 510(K) Exempt | Orthovita(R) Cortoss(TM) Delivery Gun |  |  | March 31, 2011 | 2011 | Orthovita, Inc. | Recalling firm has discovered through internal testing that the Cortoss(TM) Delivery Gun may abrade its packaging material under certain transportation and handling conditions. This abrasion may cause small holes or tears in the packaging material, which may not be immediately visible to the user. Although t the Cortoss(TM) Delivery Gun is provided in double pouches and there have been no report | Packaging | Orthovita sent an "URGENT FIELD SAFETY NOTICE" letter dated March 31, 2011 to all customers. The letter described the product, problem and actions to be taken. The customers were instructed to Identify, quarantine and discontinue use of the affected product. The letter included an acknowledgement form to be completed and returned as per instructions.. The letter states that the recalling firm will contact customers to arrange for the return or destruction of any affected product. A pre-paid shipping label is provided for product return. The letter also states that the recalling firm will provide replacement product at no cost. For any questions regarding this notice call 484-323-8860. |  | 3174 | 3174 | Product was distributed within the US to medical facilities in AL, AZ, CA, CO, CT, DE, FL, GA, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, NE, NH, NJ, NY, NC, OH, OK, PA, RI, SC, SD, TN, TX, UT, VA, WA, WV, and WI. Government accounts include Bethesda Naval Center, 8901 Rockville Pike, Bethesda, MD; Lackland Air Force Base, 2200 Berquist Drive, Suite 1, Lackland AFB, TX; Hunter Holmes McGuire VA Hospital 1201 Broad Rock Blvd., Richmond, VA; Madigan Army Medical Center, Madigan Army Medical Center, Tacoma, WA; Ralph H. Johnson VA Medical Center, 109 Bee Street, Charleston, SC; VA Medical Center, 1481 W. 10th St., Indianapolis, IN; VA Hospital of Palo Alto, 3801 Miranda Ave., Palo Alto, CA; VA Loma Linda Healthcare System, 11201 Benton Street, Loma Linda, CA; and VA Medical Center/NY Harbor Healthcare, NY, NY. | Terminated | September 25, 2012 | 544 | 3174 | Not_Computer
Z-0161-2012 | 60003 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Synthes Spine T-PAL Spacer System. Transforaminal posterior atraumatic lumbar spacer system. |  |  | September 23, 2011 | 2011 | Synthes USA (HQ), Inc. | LABELING CORRECTION for Medical Device - previous labeling and promotional materials associated with the T-PAL Spacer system reference the use of T-PAL spacer with allograft material. T-PAL Spacer is currently indicated for use with autograft material. All accounts with record of receiving T-PAL Implant will receive notification of the labeling correction and a copy of the updated brochure. All | Error in labeling | SYNTHES sent an URGENT MEDICAL DEVICE LABELING CORRECTION letter dated September 23, 2011, to all affected customers. The labeling has been revised. Updated information packets have been sent to all accounts which included the final labeling and communications to accounts and sales force. Customers were instructed to review the updated supplemental labeling provided in the packageCustomers were instructed to attach the supplementary technique page to an T-PAL Spacer Brochure that may be in their facility. They can access an electronic version of the Brochure by accesssng Product Literature at the following link:http://us.synthes.com. The new version can be identified by code J10120 located at the bottom right corner of the back cover.To confirm receipt and understanding of the information, customers were requested to complete the Verification Section of the letter and return it to Synthes by either of the following methods:Mail provided postage paid envelopeFax 610-251-9005Scan/e-mail to FieldAction@synthes.comFor any questions call 1-800-420-7025 x5403 or contact their Synthes Spine Sales Consultant. |  |  | N/A - Labeling Correction | Worldwide Distribution - USA including AL, AK, AZ, AR, CA, CO, CT, FL, GA, ID, IL, IN, IA, HI, KS, KY, LA, MD, MA, MI, MN, MS, MT, NE, NV, NH, NJ, NM, NY, NC, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, and WI. and Puerto Rico and Washington, D.C; Internationally to Canada and Japan | Terminated | December 18, 2014 | 1182 |  | Not_Computer
Z-2046-2011 | 58240 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Clinac High Energy Medical Linear Accelerator, a Trilogy Radiotherapy Delivery System; Model #s: H14, H18, H27, H29, HCX; Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA |  |  | February 24, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | The wedge angle labeling on the wedge tray may not correctly match the wedge body. However, the wedge body is labeled with the correct wedge angle. | Labeling mix-ups | The firm, Varian Medical, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated February 24, 2011, to all affected users. The letter described the product, problem and actions to be taken. The customers were instructed to IMMEDIATELY inspect all wedges to verify that the wedge angle label on the wedge body matches the wedge angle on the wedge tray; if any wedge is identified with mismatched labeling follow the written instructions provided in the letter; always measure wedge data at installation of new or replacement wedges; verify all treatment accessories of "matched" linear accelerators, and please advise the appropriate personnel working in your Radiotherapy Department of the content of this letter.Varian has enhanced their manufacturing quality control process to further ensure correct labeling of wedges. Also, Varian will replace any wedge that is identified as mislabeled, at the request of the customer. If you require further clarification, please feel free to contact your local Varian Customer Support Service Manager at: 1-888-VARIAN5 (888-827-4265) for USA and Canada; +14 41 749 8844 Europe; North American: support-americas@varian.com; Australia/New Zealand:support-anz@varian.com; Europe: support-emea@varian.com; South East Asia: seasia.apps.helpdesk@varian.com; China/Asia: china.apps.helpdesk@varian.com; Japan: Japan.apps.helpdesk@varian.com; Latin America:soporte.al@varian.com. |  | 5541 | 5541 units | Worldwide distribution. | Terminated | November 05, 2012 | 620 | 5541 | Not_Computer
Z-1009-2011 | 57576 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | AutoSuture SURGIPRO Polyproylene Mesh Clear 3" x 5" Sterile |  |  | December 15, 2010 | 2010 | Covidien LP | Sterility compromised due to breach in sterile barrier | Packaging process control | Covidien notified customers via letter dated November 15, 2010. The letter identified the product, the problem, and action to be taken by the customer. Customers were instructed to examine inventory and return the affected lot to Field Returns Department, 195 McDermott Road, North Haven, CT 06473. For questions regarding this recall call 203-492-8091. |  | 132 | 132 units | Worldwide Distribution - USA including CA, CT, FL, HI, IL, IN, KS, MA, MD, MI, MN, NC, NJ, NY, OH, PA, SC, TX, and WI and the countries of Canada and Italy | Terminated | January 10, 2013 | 757 | 132 | Not_Computer
Z-2239-2011 | 58685 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | BardÂż Soft Mesh 4"x6" (10cm x 15cm) Flat, Product Code: 0117010. |  |  | April 25, 2011 | 2011 | Davol, Inc., Sub. C. R. Bard, Inc. | Mislabeling: complaints of product labeled as BardÂż Soft Mesh 4" x 6" containing BardÂż Soft Mesh 6" x 6" | Labeling mix-ups | Davol issued an URGENT: MEDICAL DEVICE RECALL notification letter to customers dated April 25, 2011, informing them the recall involving a single lot of the BARD Soft Mesh - Lot# HUUF0530. This recall has been initiated because this lot number may contain a different size patch than is indicated. Customers were instructed to immediately examine inventory and quarantine product subject to recall. Customers are asked to identify and notify their customers to whom recalled product may have been further distributed. A Recall Effectiveness Check reply form should be completed and returned to Davol via fax (401-825-8753). For directions on returning product or for questions regarding the recall, contact Davol Customer Service Department at 1-800-556-6275 or C.R. Bard's Medical Services &amp; Support Department at 1-800-562-0027. |  | 198 | 198 units | Nationwide Distribution | Terminated | October 28, 2011 | 186 | 198 | Not_Computer
Z-0307-2012 | 60125 | No | IWB | Radionuclide radiation therapy system. | Radiology | Radiology | 510(k) | Elekta Leksell Gamma Plan |  |  | June 23, 2011 | 2011 | Elekta, Inc. | There is a risk that the user may interpret the instructions in the manual too literally, and adjust the TMR 10 dose description to obtain an ETCPD that matches previous clinical prescriptions, which in turn could lead to a prescription dose that deviates a few percent from the intended dose. | Labeling False and Misleading | Elekta sent an IMPORTANT NOTICE letter dated June 23, 2011 to affected customers. The notice identified the product, problem, and corrective actions to be taken by the user. The letter advise all customers to follow any instructions or recommendations covered in the Notice. The letter states that Elekta is working on a new software version and customers will receive the future version when released. For questions contact your local Elekta representative. |  | 2 | 2 | Nationwide Distribution (USA) - including the states of: NM and WA | Terminated | January 31, 2012 | 222 | 2 | Not_Computer
Z-0968-2011 | 57231 | No | GFA | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Sternal Saw II Sterile Blades for the Sarns Sternal Saw System, Catalog number 5589, Terumo Cardiovascular System Corp, Ann Arbor, MI. |  |  | October 20, 2010 | 2010 | Terumo Cardiovascular Systems Corporation | The firm discovered that some outer box labels contain incorrect information regarding the method of sterilization. The label incorrectly indicates the product was ethylene oxide sterlized but the product actually was sterilzed with gamma radiation. | Labeling design | Terumo Cardiovascular Systems Corporation firm sent out an 'URGENT MEDICAL DEVICE CORRECTION: SAFETY ADVISORY" letter dated October 20, 2010, to the Director of Cardiac Operating Room Services or Operating Room Nurse Manager.. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the Safety Advisory, assure that all users were aware of the notice, and confirm receipt of the Safety Advisory by faxing the attached Customer Response Form to the fax number on the form. For questions regarding this recall call 1-800-521-2818. |  | 181902250 | 42,650 units, 4265 boxes of 10 each | US, Indonesia, UAE, Hong Kong, Japan, Singapore, Malaysia, Taiwan, Belgium, Uraguay, Venezuela, Columbia, Mexico, Guam, Vietnam and Thailand. | Terminated | December 14, 2011 | 420 | 181902250 | Not_Computer
Z-2676-2011 | 58807 | Yes | BTO | Tracheostomy tube and tube cuff. | Anesthesiology | Anesthesiology | 510(k) | Jet-Ventilator Catheter for Adults, 14G, STERILE, NO LATEX, REF 30-02-918-1, VBM Medizintechnik GmbH, Einsteinstrasse 1 7212 Sul z a.N. Germany |  |  | May 09, 2011 | 2011 | VBM Medical Inc. | Product labeled with incorrectly as being appropriate for use in adults when the intended population for use of the product is children and infants. | Error in labeling | VBM Medical Inc sent an "URGENT FIELD SAFETY NOTICE" dated May 9, 2011, to all affected customers. The letter identified the product, the problem, and the action needed to be taken by the customer.The reference and lot numbers, details of the error, along with the incorrect and correct box label, and instructions to the customers to discard the box with the incorrect label or send the complete product back to VBM Medical Inc. for re-packaging. For any further or detailed questions, please contact VBM Medical Inc. at (317) 776-1800.On 7/12/11 upon receiving the recall classification letter, the firm sent out an additional letter to its direct accounts informing them of the recall's classification and reiterating the instructions listed above. |  | 29 | 29 | Worldwide Distribution -- USA ( nationwide ) including the country Montreal QC | Terminated | March 14, 2012 | 310 | 417 | Not_Computer
Z-1740-2011 | 57705 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | UC Tibial Insert, Size 4, 13mm, REF 163-3413, Sterile R, Ortho Development, Draper, Utah 84020. |  |  | January 07, 2011 | 2011 | Ortho Development Corporation | Components of a replacement knee system could be mislabeled. | Error in labeling | The firm, Ortho Development, sent a "Field Notification Memo" date January 7, 2011 to its customers. The memo described the product, problem and action to be taken. The customers were instructed to promptly return any remaining parts of the listed product; locate all the parts in their inventory; identify any of the listed parts that may have been used in surgery; verify the quantity they have on hand and the quantity used in surgery and reconcile them with the quantity sent; if discrepancy exists, notify customer service resolution; prepare items for return and contact customer service to obtain RMA number. Once reconciled, arrangements would be made for sending new inventory and the return of unused parts.If you have any questions regarding this issue, please notify Regulatory Affairs Manager at (801) 619-3450. |  | 25 | 25 units | Worldwide distribution: USA including states of: AZ, CA, FL, MI, TX, and UT; and country of: Japan. | Terminated | December 27, 2011 | 354 | 72 | Not_Computer
Z-2688-2011 | 58444 | No | KNF | Unipolar endoscopic coagulator-cutter and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | LiNA Loop Endoscopic Monopolar Loop for LSH (Laparoscopic Supracervical Hysterectomy). Large loop, 4mm connector, 1 X REF: EL-200-4 and 10 X REF: EL-200-4, Manufactured by LiNA Medical ApS, Surgery Division. |  |  | December 01, 2009 | 2009 | LINA Medical ApS | Potential for the LiNA Loop to break during operating procedure. | Labeling design | Lina Medical Aps notified customers by newsletter in December 2009 followed by a letter to distributors including warnings added to instructions for use in January 2010. |  | 200 | 200 pieces | GA, MN, PA, WA, TN, MA, IN, Ontario Canada | Terminated | February 23, 2012 | 814 | 2720 | Not_Computer
Z-0986-2011 | 54350 | No | CAE | Oropharyngeal airway. | Anesthesiology | Anesthesiology | 510(K) Exempt | KING LT -D Laryngeal Tube, Size 4, KLTD204, King Systems Corporation 15011 Herriman Boulevard, Noblesville, IN 46060. |  |  | December 15, 2009 | 2009 | King Systems Corp. | King Systems was issued a warning letter from the Center for Devices and Radiological Health, Office of Compliance on October 28, 2009. The warning letter asserts that Kings Systems KLT(S)-D airway exceeds cleared indications in that it is promoted for "airway management in patients over 4 feet in height, for controlled or spontaneous ventilation, difficult and emergent airway cases, and ambulat | Labeling False and Misleading | On October 29, 2009, the firm, Kings System, sent a "Urgent - Product Correction" letter to the customers. The letter described the product, problem and actions to be taken. The customers were instructed to remove the promotional material from the King System website which includes literature, and Instructions for Use. The firm instructed all Sales force and distributors to cease dissemination of promotional materials, literature and communications.Note: The firm revised indications, instructions for use, product labels and labeling so it is aligned with the 510Ks. The customers were also instructed to replace current instructions for use with revised instructions for use and destroy the current IFU; return Kings LT(S)-D oropharyngeal airway product inventory to King Systems; and complete and return the attached Instructions for Use Destruction Signature Form via fax to King Systems at 317-776-5197.For further information, please call 317-776-6823 ext 257 or email: LTSD@kingsystem.com. |  |  |  | Worldwide distribution: USA and countries including: Canada, Costa Rica, and Panama. | Terminated | August 16, 2012 | 975 |  | Not_Computer
Z-0309-2012 | 60314 | No | CAZ | Anesthesia conduction kit. | Anesthesiology | Anesthesiology | 510(k) | Arrow International StimuCath(R) Continuous Nerve Block Kit, ASK-19608-SFH, Lot # MF0127039 |  |  | June 20, 2011 | 2011 | Arrow International Inc | Some of the outer trays used in the packaging of the ASK-19608-SFH Continuous Nerve Block Kits, lot # MF0127039 were found to have an insufficient seal. Therefore, the packaging may have been compromised, and sterility of the product cannot be guaranteed. | Packaging | The firm, Arrow International, sent an "URGENT MEDICAL DEVICE RECALL" letter dated June 20, 2011, to its customer. The letter included a stock response form to be completed and returned to the firm. The letter also described the product, problem and action to be taken. The customer was instructed to check their stock for the products included in the scope of the recall; cease use and distribution; quarantine all affected product immediately; complete and return the enclosed Recall Acknowledgement &amp; Stock Status Form to their sales representative, and return any affected product freight collect to: Arrow International, Inc., 312 Commerce Place, Asheboro, NC 27203, attn: Asheboro Quality Assurance Manager. Arrow International can provide replacement with a similar product or credit their account.If you have questions or concerns, please call this toll free number 800-233-3187. |  | 17 | 17 kits | Nationwide distribution: CT | Terminated | October 01, 2012 | 469 | 17 | Not_Computer
Z-2237-2011 | 58563 | No | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Nasal Cannula (Adult) Salter Style with modified nasal prongs, 7" (2.1M) Q supply tube and CO2 sample tube with male luer connector 22mm (I.D) x 6mm (O.D) Adaptor. Adult Divided Cannula. |  |  | April 19, 2011 | 2011 | Salter Laboratories, Division of Regulatory Affairs | Complaint received that the product in the box was different from labeled product. | Labeling mix-ups | The firm, Salter Labs, sent an "URGENT PRODUCT RECALL" subtitled Adult Divided Cannula Product Code 4707 Lot 012411, letter dated April 19, 2011 via certified mail to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to examine their inventory; return any of the lot in their possession to Salter; contact Salter Customer Service at 1-800-235-4203 to arrange for the return of the product, and complete and return the Customer Reply Form via fax to 661-854-6816 or Toll Free 1-800-628-4690.Should you have any questions, please contact Salter Labs at 1-800-235-4203. |  | 22150 | 22,150 pieces | Nationwide distribution. | Terminated | August 23, 2011 | 126 | 22150 | Not_Computer
Z-2698-2011 | 58702 | No | HRY | Knee joint femorotibial metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Oxford Femoral Knee System REF 154600, LOT 2272736, WARNING DO NOT RESTERILIZE, STERILE UNLESS PACK DAMAGED, BIOMET UK Ltd, (510K# K011138) |  |  | April 22, 2011 | 2011 | Biomet U.K., Ltd. | This recall was initiated due to a customer complaint which reported that while peeling off the Tyvek lid on the blister pack, very little effort was required to remove it and this could lead to contamination of the device. | Packaging | Biomet sent a "FIELD SAFETY NOTICE" dated April 21, 2011, to all affected accounts. The notice identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use of the affected product, locate and remove from inventory and quarantine pending return to their Bioment distributor. Customers were also instructed to complete and return the attached "Fax Back Response Form" to their local Biomet distributor.For questions customers should contact their local Biomet Contact. |  | 12 | 12 devices | Outside the US to Austria, Belgium, and Germany | Terminated | August 03, 2011 | 103 | 12 | Not_Computer
Z-2701-2011 | 58145 | No | NJE |  |  | Neurology | HDE - Humanitarian Device Exemption | CODMAN ENTERPRISEÂż Vascular Reconstruction Device and Delivery System 4.5mm x 28mm |  |  | March 21, 2011 | 2011 | Codman &amp; Shurtleff, Inc. | Mislabeled: One of the stick on stickers on the sterile pouch inside of the product box indicates the stent size to be 22mm in length, when in fact the product is 28mm in length. | Labeling design | Codman &amp; Shurtleff, Inc.sent an "Urgent Medical Device Correction Notice" on March 21, 2011.The letter recommend the incorrect label be removed and discarded. There are 3 additional peel-off labels, 2 on the outer box and one on the pouch. All 3 of these contain the correct printed size information (28 mm) that can be used for applying to the inventory sheet, patient record or patient card, as needed. It is NOT required to return any of this product to Codman.Customer Service at 1-800-225-0460 or contact your local Codman Neurovascular Representative |  | 380 | 380 units | Worldwide - Nationwide (USA), Canada, Korea. | Terminated | February 17, 2016 | 1794 | 380 | Not_Computer
Z-2711-2011 | 58809 | No | ETA | Total ossicular replacement prosthesis. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | Kurz, Heinz Kurz Gmb, Medizintecnik, Tuebinger Strasse 3, 7214 Dusslingen, Germany, Surgical Implant, Clip - Piston MVP Titanium, REF: 1006 711, Lot 2 09 13 06. |  |  | April 27, 2011 | 2011 | Heinz Kurz GmbH Medizintechnik | Outer package labels product .75 mm longer than actual size. | Labeling False and Misleading | Kurz Medical, Inc. sent an "URGENT VOLUNTARY MEDICAL DEVICE RECALL" letter dated May 9, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers.The letter instructs customers to return unused product for replacement by using an enclosed FedEx label. If the product has been implanted , the letter requests the customers send the date of surgery and size of the implant. |  | 40 | 32, only 8 consigned with U. S. hospitals | Nationwide Distribution--USA (nationwide) including the states of MN, NC, and WA. | Terminated | November 23, 2011 | 210 | 40 | Not_Computer
Z-2709-2011 | 59043 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Product Brand Name: Versafitcup Double Mobility Liner Inserter |  |  | May 23, 2011 | 2011 | Medacta Usa Inc | Medacta International has recently realized that the base of the Liner Inserter Ref 01.26.10.0018Lot 105549 of the Versafitcup Double Mobility system could become rusty after the washingprocedures. This is due to an incorrect raw material used by one of our suppliers during theproduction phase of this lot. For this reason, Medacta International has made the decision toremove all the pieces | Labeling False and Misleading | Medacta USA sent an "URGENT SAFETY NOTIFICATION-VERSAFITCUP DM LINER INSERTER" letter dated May 23, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter instructs customers to check their inventory for the affected device, immediately sequester it and not use the device in surgery. The letter instructs customers to pack the affected device for shipment to Medacta USA under Return Material Authorization 1590. Customer will receive a replacement within a few days.Questions regarding this recall are directed to the Director of Regulatory and Quality at 805-437-7085 ext. 26. |  | 27 | Worldwide: 18 pieces. In the USA: 9 pieces | Worldwide Distribution--USA (nationwide) including the states of CA, ID, IL, and OH, and 9 products to foreign countries. | Terminated | July 12, 2011 | 50 | 27 | Not_Computer
Z-0418-2012 | 59739 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | FOUNDATION HIP SYSTEM Series 460 Hip Stem STERILE R CEMENTED Sz 16.5mm |  |  | August 17, 2011 | 2011 | Encore Medical, Lp | The sterility has the potential to be compromised. | Packaging | Encore Medical Lp sent a Urgent Field Safety Notice dated August 17, 2011. to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to immediately quarantine these devices if you have not already done so. Contact Customer Service to obtain replacement devices as well as an RMA. Return all affected devices by October 31, 2011, using the RMA Customer Service provided in step #3.If you have any further questions you may call (512 ) 834-6330. For International Customer Service please call (512) 834-6275. |  | 169 | 169 units | Worldwide Distribution -- USA ( nationwide ) including the states of WI, NY, MO, TX, FL, CA, UT, AZ, Puerto Rico, and the country Japan. | Terminated | May 11, 2012 | 268 | 169 | Not_Computer
Z-0506-2012 | 60581 | No | FLL | Clinical electronic thermometer. | General Hospital | General Hospital | 510(k) | Walgreens Quick Read Digital Thermometer; an electronic digital thermometer; Model WIC551226 |  |  | November 23, 2011 | 2011 | Walgreen Company | The thermometers have been reported to have inaccurate readings | Labeling design | On November 23, 2011, Walgreens issued a recall notice electronically to their stores, instructing them to pull and quarantine the Thermometers and placing a Stop Scan on the products in the computer cash register system on the same date. Walgreens also removed the thermometer from their internet sales site. |  | 92000 | estimated 92,000 units | Nationwide through Walgreens retail stores and their internet site. | Terminated | August 14, 2017 | 2091 | 92000 | Not_Computer
Z-1852-2011 | 58028 | container, specimen, non-sterile | NNI | Specimen transport and storage container. | Pathology | General Hospital | 510(K) Exempt | Large Port Specimen Sock, used to convert the pour/accessory port of a Cardinal Suction canister into a specimen retainer. The sock is used as an accessory in a suction collection system. Its placed in a suction canister to remove solids from the liquid aspirated during surgery. The solids may consist of patient tissue that are then sent to pathology or retained for other purposes. There is no direct contact between this device and the patient. |  |  | February 18, 2011 | 2011 | Cardinal Health Inc | In one lot of product (lot# 1091745), the label for the small version of the sock (catalog number 65652-123) was improperly used for the larger version of the product (65652-124) | Labeling mix-ups | On February 18, 2011 Cardinal Health sent a consignee letter with the following instruction: 1) Quarantine either Catalog Number with the inticated Lot Number.2) Contact the appropiate Customer Service group for return instructions:Hospital - 800-964-5227Federal Government - 800-444-1166Distributirs - 800-635-6021All other customers - 888-444-54403) Return the enclosed acknowledge form via fax to 847-689-9101, even if you do not have affected product, as Cardinal Health is required to confirm receipt of this notification from you, and to prevent further notices.4) Notify any customer to whom you may have distributed product affected by this recall. |  | 3360 | 3,360 each | Worldwide Distribution - USA including AL, AR, AZ. CA. CO,, CT, DC, DE, FL, GA, HI, IA, ID,IL,IM,KS, KY, LA, MA, MD, ME. MI MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, WY, &amp; Puerto Rico - and the country of Canada. | Terminated | October 04, 2011 | 228 | 3360 | Not_Computer
Z-0988-2011 | 56963 | No | GDX | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Southmedic/Personna Safety Scalpel with Personna Plus Blades, Reorder No. 73-0610, Size #10, General Surgical Use, Sterile, unless package is damaged or opened., Southmedic Inc., 50 Alliance Blvd., Barrie, ON., Canada L4M 5K3 www.southmedic.com, Made in Canada. |  |  | September 03, 2010 | 2010 | Southmedic, Inc. | Scalpels labeled as sterile have the potential to be unsealed and therefore not sterile. | Packaging process control | On 9/3/2010, the firm notified consignees about the recall by facsimile letter. The letter instructed consignees to return scalpels for credit or replacement. The firm enlarged the scope of the recall on 11/15/2010. An additional 927 cases were recalled. This includes 11 additional lots. The firm sent out URGENT MEDICAL DEVICE RECALL letters dated 11/15/2010 to its direct accounts. The letter identified that affected product and the reason for the recall. The consignees were instructed to check their inventory, complete and return the attached Voluntary Recall Notification Confirmation Form, return the affected product to the manufacturer. Product return question should be directed towards Marc Pauze at 800-463-7146 ext 299 and regulatory questions should be directed towards Tish Anger at 800-463-7146 ext 342. |  | 1330 | 1330 all products | Worldwide Distribution -- USA, including states of TX, MO, NJ, LA, IL, NY, KY, PA, VA, WA, CA, MA, MD, IN, NV, TN, NC, AZ, IA, WI, UT, NE, GA, FL, MN, ME, OH, SD, and KS and countries of Australia and Canada. | Terminated | May 18, 2011 | 257 | 1330 | Not_Computer
Z-0382-2012 | 60437 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | CyberKnife Robotic Radiosurgery System: Lung Optimized Treatment option. |  |  | November 07, 2011 | 2011 | Accuray Incorporated | User Facility reported an anomaly where inhale and exhale CT pairs used for treatment planning did not represent the same magnitude of respiration that was displayed during treatment delivery. | Labeling design | Accuray sent a Urgent Medical Device Correction letter dated November 7, 2011, via Federal Express to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The information provided in this letter will be incorporated in user manuals shipped with the CyberKnife system.If you have any questions or concerns regarding this issue, please do not hesitate to contact Accuray Customer Support at 1. (877) 668-8667 ( USA) 1. (408) 716-4700 ( non USA ) |  | 30 | 30 | Worldwide Distribution -- USA ( nationwide ) including the states of FL, TN, IL, NJ, CO, PA, MD, CA, MS, MN, SC, WA, OK, GA, AZ, and TX., and the countries of France and United Kingdom. | Terminated | February 27, 2012 | 112 | 30 | Not_Computer
Z-0274-2012 | 60328 | No | FTM | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Peri-Strips Dry with Veritas Collagen Matrix Staple Line Reinforcement (PSDV), sterile EO, Rx Only, Synovis Surgical Innovations, A Division of Synovis Life Technologies, Inc, 2575 University Ave W., St Paul, MN 55114-1024. |  |  | October 26, 2011 | 2011 | Synovis Surgical Inovation Div. of | Synovis Surgical Innovations is voluntarily recalling one lot of Peri-Strips Dry with Veritas Collagen Matrix Staple Line Reinforcement (PSD-V) product because some of the devices were mislabeled. | Labeling mix-ups | Synovis Surgical Inovation sent a "Recall Notification" letter dated October 26, 2011. The letter described the problem and product affected. Customers were instructed to identify and quarantine the affected product and to complete and fax the Product Return/Replacement Form. |  | 132 | 132 units | Nationwide Distribution including CT, GA, NJ, NY, OH, TX, VA. | Terminated | April 24, 2012 | 181 | 132 | Not_Computer
Z-2760-2011 | 58335 | [6 empty] | J-Vac Drain Adaptor Manufactured for: Ethicon, Inc. Somerville, New Jersey, 08876-0151. |  |  | March 24, 2011 | 2011 | Ethicon, Inc. | There is a potential for the sterile barrier to be compromised in multiple lots of Blake silicone drains, Blake silicone drain kits, Blake cardio connectors, J-VAC reservoirs, J-VAC drain adaptors. | Packaging process control | Ethicon, Inc. sent an "URGENT VOLUNTARY PRODUCT RECALL" letter dated March 24&amp; 25, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter instructs customers to discontinue use of the products and to return the products to Stericycle. The letter includes shipping labels for product return. Ethicon will send replacement product. Customers are asked to return a completed Business Reply Card to Stericycle regardless of whether or not they have the products subject to this recall.Customers with questions regarding this recall are directed to call 1-866-664-1401. |  | 846990 | 846, 990 | Worldwide Distribution-- USA (nationwide) and Puerto Rico and the countries of Armenia, Australia, Austria, Azerbaijan, Columbia, Cyprus, Czech Republic, Denmark, French Guinea, French Polynesian, Germany,Israel, Italy, Japan, Kuwait , Netherlands, New Caledonia, New Zealand, Norway, Singapore, Slovenia, South Africa, Spain, Belarus, Belgium, Brazil, Canada, Chile, Dominica Republic, Egypt, Estonia, Finland, France, Greece, Guadeloupe, Hong Kong, Hungary, Ireland, Latvia, Lithuania, Malaysia, Martinique, Mozambique, Portugal, Reunion, Russia, Saudia Arabia, Sweden, Switzerland, Ukraine, and United Kingdom. | Terminated | January 10, 2014 | 1023 | 846990 | Not_Computer
Z-2753-2011 | 59024 | No | NEU | Implantable clip. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Bariatric Target Ring (Nitinol) (a thin, circular nitinol ring) Model # K-9536-N, Lot #08-001, Lot #08-002 &amp; Lot #08-003. One ring per plastic pouch, three rings are then placed into a plastic box with lid. This box is then placed into a solid white box. Item is shipped and sold non-sterile. |  |  | May 24, 2011 | 2011 | Kapp Surgical Instrument Inc | During an FDA Inspection on 3/21/11 - 3/22/11 at Kapp Surgical an uncontrolled instruction for use was discovered which contained sterilization parameters which differed from the instructions contained in the device master record. Kapp Surgical decided to conduct the recall because we cannot be sure that incorrect instructions were not shipped. | Labeling mix-ups | Kapp Surgical Instrument, Inc. sent an "URGENT MEDICAL DEVICE RECALL" letter dated May 24, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers.The letter instructs customers to check their inventory and quarantine any product affected by this recall. The firm will make arrangements for returning the products. Replacements will be made by Kapp Surgical.Questions about this recall are directed to the QA Manager at 800-282-5277. |  | 222 | 222 units (rings) | Product was distributed to the following states: OH, PA &amp; SC. Product was also shipped to Saudia Arabia. | Terminated | February 19, 2013 | 637 | 222 | Not_Computer
Z-1003-2011 | 56349 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand C-Series Clinacs and Acuity, Exact Couch Top, Model Numbers: |  |  | April 19, 2010 | 2010 | Varian Medical Systems Oncology Systems | The product has been shipped since 2005 with the wrong label regarding max weight to be used on accessory rails. | Labeling mix-ups | Varian issued via customer technical bulletin on April 19, 2010, to Varian Sales, Marketing and Service organization. Notification. The bulletin identified the product, the problem, and the action to be taken by the customer. Customers were asked to fill out the attached form and send it to their local Varial Medical Systems Field Service Office.For questions regarding this recall call 650-483-3153. |  | 13 | 13 | Worldwide Distribution - Japan only | Terminated | December 23, 2011 | 613 | 13 | Not_Computer
Z-0967-2011 | 57188 | [6 empty] | StaXxÂż XD Expandable Cartridge Convex Expandable Implant 25mm (L) x 9mm (W) x 7mm (H) |  |  | October 22, 2010 | 2010 | Spine Wave Inc | Package Insert Incorrect | Packaging process control | Spine Wave notified each affected customer on October 22,2010, via telephone of the product, the issue and action they should take. Customers were asked to segregate the product from other products and immediately return to Spine Wave.For questions regarding this recall call 203-944-9494. |  | 12 | 12 units | Nationwide Distribution including Delaware, Massachusetts, Maryland, and New York | Terminated | March 15, 2011 | 144 | 12 | Not_Computer
Z-2646-2011 | 58894 | No | KCT | Sterilization wrap. | General Hospital | General Hospital | 510(k) | Smith &amp; Nephew Hip Arthroscopy Repair Instrument Tray |  |  | May 19, 2011 | 2011 | Smith &amp; Nephew, Inc. Endoscopy Division | Hip Arthroscopy Repair Instrument Tray shipped with an incorrect Instructions For Use (IFU) | Labeling mix-ups | Smith &amp; Nephew sent an "URGENT FIELD CORRECTION NOTIFICATION" letter dated May 19, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.Customers were instructed to confirm receipt of the corrected IFU document (10600729 rev B.) and communicate to appropriate facility individuals. Once appropriate individuals are informed, customers were instructed to complete and return information requested at the bottom of the letter. The completed information should be returned to Smith&amp; Nephew via fax at 978-749-1654 or by mail at Smith &amp; Nephew Endoscopy, 150 Minuteman Road, Andover, MA 01810.Questions may be answered by contacting the Manager, QA and Compliance at 978-749-1397. |  | 11 | 11 units | Worldwide Distribution--No USA Distribution; countries of Italy, Portugal, United Kingdom, and Venezuela. | Terminated | May 09, 2012 | 356 | 11 | Not_Computer
Z-0282-2012 | 59913 | No | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | Varian Medical Systems, Flexible Probe with Blocking Water, GM11002420, (Part of the Segmented Cylinder Set, GM11004150) |  |  | May 20, 2010 | 2010 | Varian Medical Systems, Inc. | The Instructions For Use ( IFU) distributed with product prior to 2004 lacks the correct limit to the sterilization cycles and checks regarding the firm seating of the blocking washer. | Labeling Change Control | Varian Medical Systems sent a Medical Device Recall (PNL-GM-42697-2) in May 2010 for the blocking washer on the flexible probe, Varian part number GM11002420, a component of the Segmented Cylinder Applicator Set. The letter was to inform users that a flexible probe with an unstable or shifted blocking washer must not be used due to the potential for treatment length error. The blocking washer may move from rough handling or sterilization. Users were told to discard any instances of the obsolete Instructions for Use (IFU) in use and replace it with the IFU (PGM11004150 4 March 24, 2010.pdf). A Receipt Verification Card was included requesting consignees to fill out, fax or e-mail immediately to Varian BrachyTherapy. |  | 81 | 81 parts | Worldwide Distribution: USA, Germany, Belgium, Canada, Mexico City, Spain, Brazil, Switzerland, Uzbekistan, India | Terminated | December 05, 2011 | 564 | 81 | Not_Computer
Z-2660-2011 | 58864 | No | JDC | Elbow joint metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Condyle Kit with Hexalobular, REF 114700 DISCOVERY ELBOW SYSTEM, HUMERAL CONDYLE SET - HEXALOBULAR, CO-CR-MO / YI-GAL-4V ALLOY, IF USED FOR REVISION SURGERY, INSTRUMENTS 414926 &amp; 414923 MAY BE REQUIRED LOT 183300, 1 SET, BIOMET ORTHOPEDICS 56 EAST BELL DRIVE, P.O. BOX 587 WARSAW, IN 46581 USA EXPIRY DATE: 2021-03 |  |  | April 29, 2011 | 2011 | Biomet, Inc. | The firm initiated this recall after becoming aware that a Discovery Condyle Kit with Hexalobula contained two male condyles instead of a male and female condyle. | Packaging process control | Biomet sent an "URGENT MEDICAL DEVICE RECALL NOTICE" dated May 25, 2011 to all affected customers. The notice describes the product, problem, and actions to be taken by the customers.The notice instructed customers to locate and remove the product. Customers were asked to complete a response form and fax to 574-372-1683, within 3 business days. Additionally, customers were instructed to ship the package priority mail with a copy of the response form.Questions related to this notice should be directed to 574-371-3755 or 574-371-3756, Monday through Friday 8 AM to 5 PM. |  | 10 | 10 KITS | Worldwide Distribution--USA (nationwide) including the states of AR and SD and the country of Japan. | Terminated | August 29, 2011 | 122 | 10 | Not_Computer
Z-0487-2012 | 56540 | No | GWQ | Electroencephalograph. | Neurology | Neurology | 510(k) | Harmonie software in use with: |  |  | April 16, 2009 | 2009 | Stellate Systems | The following implied or stated indications in Stellate labeling with Harmoniae software have not yet been cleared by the Food and Drug Administration (FDA) for sale in the United States of America. | No Marketing Application | Stellate Systems sent a "DEVICE CORRECTION" letter dated April 6, 2009 to all affected customers. The letter described the product, problem, and the actions to be taken by the customers. The firm provided information on how to verify if customers had the affected product. Software updates will be available to all customers in possession of the affected product. Customers were instructed to call the Customer Support line at 1-888-742-1306 with any questions or concerns they may have regarding the notification. |  | 1041 | 1,041 units | Worldwide Distribution-USA (nationwide) and the countries of Australia, Canada, Denmark, Greece, Hong Kong, Hungary, Italy, Japan, Mexico, Norway, Rep. DE Panama, Spain, Switzerland, Taiwan ROC, and Turkey. | Terminated | December 27, 2011 | 985 | 1041 | Not_Computer
Z-0468-2012 | 51146 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | STERRAD NX Sterilizer Hydrogen Peroxide Gas Plasma Sterilizer Product Code 10133. |  |  | December 10, 2008 | 2008 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) has discovered a defect in the barcode labeling of the STERRAD NX System Cassettes 10133, Lot #08I032. This is the only lot number affected by this action. The barcodes for Lot #08I032 were printed improperly and are not capable of being read by the STERRAD unit. | Error in labeling | Advance Sterilization Products sent an "URGENT: PRODUCT RECALL" letter dated December 10, 2008 to all affected customers. The letter described the product, problem, and the actions to be taken by the customers. The letter instructs customers to examine inventory and return all sealed cassettes, using an enclosed prepaid UPS return label to: Stericycle, 2670 Executive Drive Suite A, Indianapolis, IN 46241. Additionally, a Business Reply Card and packing slip were included with the letter for customers to complete and return. Questions or further assistance was directed to ASP Customer Care Center at 888-783-7723, option 2. |  | 776 | 776 cases | Nationwide Distribution. | Terminated | December 16, 2011 | 1101 | 776 | Not_Computer
Z-2602-2011 | 58763 | No | MBL | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Encompass Hip Stem, Plasma Spray, 12/14 Taper, Standard, REF 108-XXXX, Sterile R, Ortho Development, Draper, Utah 84020. |  |  | May 10, 2011 | 2011 | Ortho Development Corporation | Tests on PVC packaging material resulted in alert levels for cytotoxicity. | Package design/selection | Ortho Development Corporation notified consignees by letter on May 11, 2011. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to return all remaining affected units to Ortho. Replacement product was to be shipped by May 31. |  | 113 | 113 units | Nationwide Distribution - USA including CA, FL, and TX. | Terminated | November 09, 2011 | 183 | 113 | Not_Computer
Z-2553-2011 | 58730 | No | HFS | Nipple shield. | General Hospital | General Hospital | 510(K) Exempt | Contact Nipple Shield (Sterile) - 24mm (Standard); Individually wrapped silicone contact nipple shields for immediate use for latch-on difficulties while breastfeeding the baby under professional care only; Manufactured for Medela, Inc., P.O. Box 660, McHenry, IL 60051-0660, Made in Switzerland; article #67203S |  |  | May 18, 2011 | 2011 | Medela Inc | The product sterility may be compromised due to possible pinholes in the rigid tray of the package. | Packaging | The firm, Medela, sent an "Important Device Recall" letter dated May 9, 2011 to all of their customers/direct accounts. Medala also began telephoning or visiting the direct accounts/customers on May 18, 2011. The letter described the product, problem and actions to be taken. The customers were instructed to do the following:1) not to use the affected lots and place them in quarantine2) stop sale and distribution and set aside their existing inventory of the affected product 3) conduct a physical count and record the data on the Product Inventory Form inlcude with this letter4) complete and return the Product Inventory Form via fax to 800-995-7867 or email jim.hughes@medela.com, even if they have no affected product on hand5)If you have the affected product in stock, Medela Customer Service will ship substitute product to the customer at no cost. Each affected product will be replaced with (2) non-sterile nipple shields.5) and to record the number of units on hand for return and replacement on the enclosed Product Inventory Form. The completed form should be returned via fax or e-mail to Medela even if the account has no inventory of the affected lots. Questions about the recall process were directed to Medela Customer Service at 1-888-835-5968 or e-mail at medelareturns@medela.com. Distributor/retailers were requested to notify their customers of the recall, using the enclosed customer letter. |  | 4320 | 4,320 units | Nationwide distribution: USA including states of: AR, CA, CO, CT, DE, FL, GA, IA, IL, IN, MA, MD, ME,MI, MN, MS, MT, NC, ND, NH, NJ, NY, OH, OK, PA, SC, TX, UT, VA, VT, WA, WI, and WV. | Terminated | February 16, 2012 | 274 | 4320 | Not_Computer
Z-0351-2012 | 60347 | No | FCK | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | The EnCorÂż biopsy probe, designed for use with EnCorÂż and EnCorÂż Enspire Vacuum Assisted Biopsy (VAB) driver systems. |  |  | November 03, 2011 | 2011 | Bard Peripheral Vascular Inc | This recall was voluntarily initiated by the firm due to the product 's damaged packaging trays; cracked EnCor (Trademark) packaging trays, plastic PETG tray the probe is packaged in was cracked. | Packaging | The firm sent the customer notification letter on 11/07/11, via FedEx with proof of delivery notice. The letter notified customers of the reason for recall that units from the affected lot may have packaging damage resulting from transportation. The letter instructed customers to check all inventory locations , remove any identified product, to not use the affected product, and return the affected product to Bard Peripheral Vascular, Inc., 1415 W. 3rd Street, Tempe, Arizona 85281. The letter also instructed customers to fill out the recall and effectiveness check form and fax it to 1-800-894-6772. Customers were instructed to call 1-800-321-4254, Option #2 Ext 2727 for a return authorization number to facilitate the expedient return of the product.For questions regarding this recall call 480-894-9515 |  | 64570 | 64570 units | Worldwide Distribution - USA (nationwide) and the countries of Canada, Europe, LAPAC, and Australia | Terminated | March 29, 2012 | 147 | 64570 | Not_Computer
Z-0949-2011 | 57470 | No | FQH | Jet lavage. | General Hospital | General Hospital | 510(K) Exempt | The Stryker InterPulse system is a single patient use, sterile, disposable, pulsed lavage device. The system consists of an irrigation handpiece, irrigation tubing, irrigation tips, suction tips, suction tubing, and splash shields. |  |  | October 04, 2010 | 2010 | Stryker Instruments Division of Stryker Corporation | There is potential for a breach in sterility in the blister packaging. This can lead to the use of non-sterile product which, if contaminated, could lead to a patient infection. | Packaging process control | Stryker sent a Medical Device Recall Notification letter dated October 8, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to destroy and dispose of all affected product and to complete the business reply form and return to Stryker to receive replacement interpulse products at no charge.For questions regarding this recall call (269) 323-7700. |  | 954 | 954 | Worldwide Distribution - USA (including all 50 states) and the countries of Brazil, Canada, France, Germany, Japan, Netherlands, South Africa, South Korea, Spain, Sweden, Switzerland and the UK. | Terminated | June 06, 2011 | 245 | 3187 | Not_Computer
Z-2409-2011 | 58681 | No | EZD | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | SureCath Set with Accessories, Coloplast Corp.-MN 06093, Sterile EO., Manufacturer Coloplast -A/S 3050 Denmark. |  |  | April 21, 2011 | 2011 | Coloplast Manufacturing US, LLC | Coloplast has initiated a recall on SpeediCath with Insertion Supplies and SureCath Set with Insertion Supplies because they contain Iodine prep pads which may expire before the overall product expiration date. SpeediCath with Insertion Supplies and SureCath Set with Insertion Supplies are both packaged with a set of accessories which includes Iodine prep pads. The Iodine prep pads included i | Incorrect or no expiration date | Coloplast Manufacturing US, LLC sent a Coloplast "Product Field Correction Notice" dated April 21, 2011. The letter was addressed to "Dear Valued Customer". The letter described the problem and the product involved in the recall. Customers were advised to provide the current inventory levels of the lots affected and the firm will be providing them with new iodine prep pads with sufficient shelf life. This product will include customer instructions, informing the customer to discard the original prep pads and use the new prep pads instead. If they prefer, they could themselves send the affected product back to Coloplast and they in return will attach new prep pads to the product for them.If they have further distributed the product, they advised consignees to contact their customers and to include the letter provided to them to inform them of the recall. |  | 438412 | 438,412 | Nationwide Distribution including AL, AZ, CA, CO, CT, DE, FL,GA, ID, IL, IA, MD, MA, MI, MN, MS, NJ, NY, NC, OH, OK, OR, PA,PR, SC, TN, TX VA, WA, WI. | Terminated | April 24, 2012 | 369 | 564759 | Not_Computer
Z-0347-2012 | 60335 | No | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | Product is labeled in part - Pouch label: "***NanoFUSE***DBM***Manufactured and Distributed By: Nanotherapeutics, Inc. 13859 Progress Blvd, Suite 300 Alachua, FL 32615 (386) 462-9663***Description: nanoFUSE DBM***Size***Order No.:***ID***Expiration***STERILE R***Donated Human Tissue Allograft***SINGLE PATIENT USE ONLY***Rx Only***nanotherapeutics***" |  |  | March 10, 2011 | 2011 | Nanotherapeutics, Inc | On 03/10/2011 Nanotherapeutics Inc, Alachua, FL initiated a recall of their Origen DBM with Bioactive Glass Catalog # 22-2002 (2cc), 22-2005 (5cc) and 22-2010 (10cc). The same product is also distributed as NanoFUSE DBM Catalog # NAN109-02 (2cc) NAN109-05 (5cc) and NAN109-10(10cc). The products were manufactured without an approved 510k. Direct accounts were notified of the recall on 04/08/2011 | No Marketing Application | Nanotherapeutics sent an Urgent: Medical Device Recall letter dated April 12, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to remove any unimplanted/unsold product shipped to them. Customers were asked to notify their customers and instruct them to make notification until the end user (hospital/physician) is reached. Recall Instructions:Immediately examine inventory and quarantine product subject to recall.Return all quarantined product subject to recall to:Nanotherapeutics, Inc.Attn: Receiving Department13859 Progress Blvd, Suite 300Alachua, Florida 32615For questions call 386-462-9663 |  | 131 | 131 units | Nationwide Distribuiton including TX, FL, AZ, CA, NY, and MA. | Terminated | April 03, 2012 | 390 | 12814 | Not_Computer
Z-2606-2011 | 58833 | No | JJX | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access Immunoassay Systems |  |  | May 10, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed that Access Ostase Calibrator and QC vials from the affected lots have shown an increased rate of fracturing and may leak the contents in the box during shipment to customers.Operator injury may occur due to contact with broken glass, exposure to infectious agents, and/or exposure by skin contact to agents that may cause sensitizatio | Package design/selection | Beckman Coulter sent an PRODUCT CORRECTIVE ACTION (PCA) letter on May 10, 2011 to all affected customers who purchased Access Ostase Calibrators and Access Ostase QC. The letter identified the customers with an explanation of the problem and the potential impact should the problem occur and the action needed to be taken. Customer were instructed to:(1) Inspect all Access Ostase Calibrator and QC vials and packaging for signs of failure prior to use, including:-Cracking, particularly along the bottom of the vial,-Absence of reagent and/or leaking of contents into the box(2) Human source material used in the preparation of Access Ostase Calibrator and QC material has been tested and found non-reactive for Hepatitis B (HBV), Hepatitis C (HCV), and Human Immunodeficiency Virus (HIV-1 and HIV-2). No known test method can offer complete assurance that infectious agents are absent. Access Ostase Calibrator and QC vials should be handled in accordance with Good Laboratory Practices, and should be handled as if capable of transmitting infectious disease.(3) Skin contact with Access Ostase Calibrator and QC material may cause sensitization. Wear suitable gloves when handling vials. If contact with skin occurs, wash immediately with plenty of soap and water.Customers were told to share this information with your laboratory staff and retain this notification as part of your laboratory Quality System documentation. If you have forwarded any of the affected products listed above to another laboratory, please provide a copy of the letter to them.In addition, customers were requested to complete and return the enclosed Response Form within 10 days so we are assured you have received this important communication.For replacement of defective vials:(1) In the United States, please contact Client Services at 1-800-526-3821, option 1.(2) In Canada, please contact Customer Service at 1-800-463-7828.(3) Outside of the United States and Canada, contac |  | 702 | 702 kits | Nationwide in US and worldwide to: Algeria,Argentina,Australia,Austria,Belgium,Canada,China,Colombia,Czech Republic,France,Georgia,Germany,Greece,Hong Kong,Hungary,India,Italy,Japan,Jordan,Lebanon,Liechtenstein,Mexico,Morocco,New Zealand,Panama,Reunion,Romania,Russian Federation,Saudi Arabia,Serbia,Singapore,Slovakia,South Africa,Spain,Sweden,Switzerland,Taiwan,Turkey,United Kingdom | Terminated | July 11, 2012 | 428 | 891 | Not_Computer
Z-0966-2011 | 57708 | No | KOD | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Cook Urological, Inc. Koyle Diaper Stent 6.0 Fr, Sterile, GPN G17667, REF KDSCF-060050. |  |  | December 13, 2010 | 2010 | Vance Products Inc. dba Cook Urological Inc. &amp; Cook OB/GYN | The firm received a complaint that 15 units of the Koyle Diaper Stent were received in packages that had not been sealed. Each stent should be sealed in a poly-Tyvek pouch. | Packaging process control | The firm sent an URGENT RECALL NOTICE to one consignee on 12/13/2010 (the other consignee returned all their product as a complaint). The letter informed the consignee of the recall and asked the customer to return the recalled product. |  | 25 | 25 stents | CA, CT | Terminated | January 31, 2011 | 49 | 25 | Not_Computer
Z-2395-2011 | 58518 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ES0101, AEM Disposable Curved Scissors, 1/2", 35cm, Sterile R, Encision Inc., Boulder, CO 80301 |  |  | April 07, 2011 | 2011 | Encision, Inc. | Sterility may be compromised due to holes in packaging. | Package design/selection | Encision, Inc. sent an Important Safety Notice letter dated April 7, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to either carefully inspect any remaining pouches for readily visible holes in the Tyvek pouch, or return the affected lots to Encision for inspection.Customers were also asked to fill in and fax back the attached Confirmation to 303-444-2693. |  | 598 | 574 Ten-Count Boxes, 24 Kits | Worldwide Distribution - USA (nationwide) and the countries of Australia and New Zealand | Terminated | August 09, 2011 | 124 | 2052 | Not_Computer
Z-2609-2011 | 58767 | Yes | MRM |  |  | Cardiovascular | PMA | Paradym DR Model 8550 and Paradym VR Model 8250, Sterile EO, Manufactured in Italy by Sorin CRM S. r.l., Via Crescentino s.n., 13040 Saluggia(VC) Italy |  |  | May 02, 2011 | 2011 | ela Medical Llc | On February 28, 2011, the company discovered that since the beginning of the US commercial distribution of Paradym DR 8550 and Paradym VR 8250 ICDs in October 2010, the devices included an incorrect version of the Patient Booklet.The company determined that there is no risk of adverse event associated with this labeling error. | Error in labeling | Sorin CRM USA instructed its field representatives on April 27, 2011 via a conference call to attach the correct Patient Booklet (ref. N582,) as an addendum to the device's external packaging, along with a cover letter "Quality Note" dated May 2, 2011 explaining the discrepancy. The field representatives were fedexed the information on May 2, 2011.The letter instructed the Field Representative to attach the correct patient booklet and an addendum, to the outside of the device box.Field Representatives were asked to complete the acknowledgement form by identifying S/N's in which they attached the pouch and their signature of completion and fax the form to 866-500-6096 or e-mail a scanned copy to sorin.ptusa.sorin.com.For questions regarding this recall call 877-663-7674." For affected units that were already implanted as of April 27, 2011, the patient booklet given to the patient may include information written by the physician or hospital staff that is related to the implanted system (ICD and leads), physician references and patient medications. For this reason and given the fact that the company determined that there is no risk for the patient associated with this labeling error, the company decided to not replace the patient booklet for this population. |  | 275 | 275 | Nationwide Distribution including AL, AZ, AR, CA, CO, CT, FL, GA, IA, MA, MI, MO, MS, NJ, NY, NC, OH, OK, PA, SC,A TN,T X, and WA. | Terminated | April 26, 2012 | 360 | 275 | Not_Computer
Z-2623-2011 | 58834 | No | FMG | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Baxter Interlink System Large Bore 4-Way Stopcock, Volume 0.65 mL, Injection Site, Rotating Male Luer Lock; a sterile, nonpyrogenic fluid pathway; product code 2N5601. Intended use: For the administration of drugs and solutions to a patient's vascular system through a vascular access device. |  |  | May 17, 2011 | 2011 | Baxter Healthcare Corp. | There is the potential for pinholes and tears in the individual packaging of the stopcocks, compromising the sterility of the product. | Package design/selection | Baxter Urgent Product Recall letters dated May 17, 2011 were mailed via first class mail to the direct accounts, to the attention of the Director of Nursing and the Director of Materials Management, on the same date, informing them that potential pinholes and tears in the individual packaging could compromise the sterility of the stopcock. The accounts were requested to locate and remove all affected inventory from their facility and return it to Baxter after calling Baxter Healthcare Center for Service at 1-888-229-0001 to obtain the returned goods authorization The accounts were also requested to complete the attached customer reply form confirming their receipt of the letter and fax it to Baxter at the number provided on the form. Dealers, wholesalers and distributors were requested to notify their customers of this action so they can comply with the procedures listed above. Any questions regarding the communication were directed to The Center for One Baxter at 1-800-422-9837 during the hours of 8:00 am to 5:00 pm CST. The Baxter Field Corrective Action Gatekeepers were also notified of this action via e-mail on 5/17/11, and instructed to notify their customers and ministries of health of the recall. |  | 4800 | 4,800 units | California and internationally to Japan and Australia | Terminated | June 18, 2012 | 398 | 4800 | Not_Computer
Z-0193-2012 | 60097 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker Restoration ADM System X3 Acetabular Insert Howmedica Osteonics Corp., 325 Corporate Drive, Mahwah, NJ 07430; Stryker France ZAC Satolos Green Pusignan Av de Satolas Green 69881 Mey Sieu Cedex France |  |  | June 03, 2011 | 2011 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics has become ware that there is the potential for interpretation of the product labeling which may lead to an incorrect implant being used. | Labeling False and Misleading | The firm, Stryker, sent an "URGENT PRODUCT CORRECTION" letter and Product Correction Bulletin dated June 3, 2011 via Fed Ex to its customers. The letter and bulletin describes the issue, potential hazards, risk mitigation and product correction instructions. The customers were instructed to complete and return the enclosed PRODUCT CORRECTION ACKNOWLEDGMENT FORM within 5 days via fax to: 201-831-6069.If you have any questions, feel free to call (201) 972-2100 or Hip Marketing Team at 201-831-5280. |  | 1677 | 1677 units | Nationwide distribution. | Terminated | November 13, 2013 | 894 | 16891 | Not_Computer
Z-0322-2012 | 59814 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | ZMR Porous Revision Hip Prosthesis Surgical Technique, utilized for the following ZMR devices: Porous Stem,Nitrited Porous Revision, Slotted Porous Bowed Stems with and without Compression Nut, Splined Straight Stem, Taper Stem and Bodies, Calcar Bodies, and Spout Bodies. |  |  | October 21, 2011 | 2011 | Zimmer Inc. | This recall was initiated because it may not be clear to the surgeon whether the device can be used based on the instructions for use. | Error in labeling | The firm Zimmer, sent an "URGENT DEVICE CORRECTION" letter dated August 26, 2011 to Marketing Personnel, Distributors, Sales Representatives, Distribution Operation Managers, and Surgeons. The letters identified the reason for the recall, the product, and required actions to be taken. The letter also included Medwatch reporting information and a Certificate of Acknowledgement form to fax back to the recalling firm to confirm receipt of the recall notice. For questions contact Zimmer at 1-877-946-2761. |  | 119284 | 119,284 total devices | Worldwide Distribution - USA (nationwide) including Puerto Rico and the countries of: Algeria, Argentina, Australia, Austria, Belgium, Brazil, Canada, China, Czech Republic, Denmark, Dubai/UAE, Eschbach, Finland, France, Germany, Greece, Hong Kong, Hungary, Italy, Japan, Jordan, Israel, Korea, Lebanon, Mexico, Netherlands, New Zealand, Peru, Poland, Portugal, Russia, S. Africa, Saudi Arabia, Singapore, Slovakia, Spain, Switzerland, Taiwan, Thailand, Turkey, UK, Ukraine, and Venezuela, | Terminated | August 29, 2013 | 678 | 119284 | Not_Computer
Z-1947-2011 | 57385 | No | HBX | Cranioplasty material forming instrument. | Neurology | Neurology | 510(K) Exempt | DePuy Mitek FMS Irrigation Intermediary Tube Without One-way Valve Arthroscopy tubing |  |  | December 03, 2010 | 2010 | DePuy Mitek, Inc., a Johnson &amp; Johnson Co. | Product sterility may be compromised due to breach in sterile packaging | Packaging process control | Depuy Mitek sent an Urgent Voluntary Product Recall letter dated December 2, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to isolate all inventory of the affected product and call 1-800-382-4682, PROMPT 3, for return good authorization and replacement product. Customers were instructed to return all identified affected product to:DePuy MitekATTN: Recall Coordinator50 Scotland BoulevardBridgewater, MA 02324Record their assigned return material authorization (RMA) number on the outer shipping box.Use the Prepaid Federal Express shipping label for the return of the product.Include a photo copy of the completed business reply form.For questions or concerns call 1-800-382-4682, PROMPT 3. |  | 7344 | 7344 units (372 boxes x 24 pieces) | Worldwide Distribution - USA (nationwide) and the coutnries of United Arab Emirates, Australia, Belgium, Switzerland, Czech, Denmark, France, UK, Italy, Netherlands, Poland, and Slovakia. | Terminated | April 02, 2013 | 851 | 149720 | Not_Computer
Z-2634-2011 | 58887 | No | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | ***Processed at: Regeneration TECHNOLOGIES, INC. Alachua, FL, USA 32615***Distributed by: SpinalGraft Technologies, LLC 4340 Swinnea Road STE 39 Memphis, TN 38118***OSTEOFIL DBM Paste, 5cc Tissue ID: 321933 Reorder: 002405 Lot #(Donor ID): 101050353 Expiration: 02/21/2013***STERILE***" |  |  | April 09, 2011 | 2011 | RTI Biologics, Inc. | RTI Biologics issued a recall of Osteofil DBM Paste, 5cc that was distributed on 03/04/2008 to one customer. The product was labeled as 5.0cc but should have been labeled as 0.5cc. All customers were notified of this recall on 04/09/2011. | Labeling mix-ups | RTI Biologics sent a "VOLUNTARY RECALL" letter dated April 9, 2008 to the one distributor, Spinal Graft Technologies. LLC. The letter describes the product, problem, and action to be taken by the distributor.The letter states that some of the product was issued a return authorization and requests that the distributor recover the remaining product from distribution. The recovered product will be issued a return authorization. RTI states the grafts can be returned in a non frozen state.Recall effectiveness will be determined by return of the product. |  | 26 | 26 units | Domestic Distribution--TN only. | Terminated | July 24, 2012 | 472 | 26 | Not_Computer
Z-2385-2011 | 58199 | No | JEY | Bone plate. | Dental | Dental | 510(k) | ***BIOMET MICROFIXATION BIOMET MICROFIXATION 1520 TRADEPORT DRIVE JACKSONVILLE, FL 32218 USA www.biometmicrofixation.com***6 HOLE REGULAR DOUBLE Y PLATE***"1.5MM SYSTEM PLATE DOUBLE Y-SHAPE, REGULAR 0.6 MM TITANIUM***NON-STERILE PRODUCT*** |  |  | July 28, 2010 | 2010 | Biomet Microfixation, Inc. | Biomet Microfixation, Jacksonville, FL is recalling their 1.5 6 Hole Reg Double Y, Lots 785710, 753120, 693500 and 733600. This product is labeled as 01-7110-K (1.5 6 Hole Reg Double Y), however the package contains a 01-7112-K (1.5.6 Hold Long Double Y). | Labeling mix-ups | The firm, Biomet Microfixation, sent an "URGENT MEDICAL SAFETY ALERT NOTICE" letter dated July 28, 2010 to its customers/distributors. The letter described the product, problem and actions to be taken. The customers were instructed to inspect their stock and ensure that the plates were properly distributed into their trays and that they have the sufficient quantity of plates that they need. In addition, they were instructed to complete and return the attached Inventory Reconciliation form via fax to: Attn: Customer Relations Specialist at 904-741-9425 and notify the recipient of this notice, if they have further distributed the product.If you have any questions, please contact the Customer Relations Specialist at 800-874-7711 or 904-741-4400, extension 468, Monday through Friday, 8am-4pm ET. |  | 180 | 180 | Worldwide distribution: USA (nationwide) including states of: FL, IN, KY, MT, NC, NE, NY, OH, TX, VA, and WI; and countries including: Argentina, Canada, Colombia, Czech Republic, Japan, Korea, Mexico, Norway, South Africa, Spain and Turkey. | Terminated | June 24, 2011 | 331 | 180 | Not_Computer
Z-2387-2011 | 58652 | No | DJG | Opiate test system. | Toxicology | Toxicology | 510(k) | DrugCheck, Manufactured by: Express Diagnostics Int;l Inc. 1550 Industrial Drive, Blue Earth MN 56013 USA, Model numbers : |  |  | March 18, 2011 | 2011 | Express Diagnostics Int'l., Inc. | Express Diagnostics Int'l, Inc. has initiated a medical device recall on various model numbers of DrugCheck devices,This recall was due to incorrect product expiration dating after findings from an FDA inspection. Use of expired product may not work properly, and may give incorrect screening results. | Incorrect or no expiration date | The firm, Express Diagnostic Int'l, Inc., sent an "Urgent Recall of DrugCheck Product" letter dated April 19, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately quarantine the product; notify Express Diagnostic of the quantity even if no product is remaining; complete and return the DISTRIBUTOR/CUSTOMER RECALL CONFIRMATION FROM via fax to 1-507-526-2252 as soon as possible. In addition, if they have further distributed the product to identify their customers and notify them of the recall. Provided procedures to be followed for them to destroy the quarantined product. Express Diagnostics will replace the DrugCheck Drug Screen devices.If you have any questions, please contact the Quality Assurance Manager at 507-526-3951, ext. 117. |  | 6925 | 6925 devices | Worldwide distribution: USA (nationwide) including states of: FL, IA, MD, OK, and WA; and countries including: AUSTRALIA and ITALY. | Terminated | July 06, 2012 | 476 | 6925 | Not_Computer
Z-0904-2011 | 52716 | No | EPT | Surgical microscope and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | VISU 200 on S8 Floor Stand, Manufactured by Carl Zeiss Surgical GmbH, Germany |  |  | May 18, 2009 | 2009 | Carl Zeiss Meditec Inc | Potential Retina Damage-- An error in the user manual may result in misapplied filter and injury during treatment. | Labeling mix-ups | Carl Zeiss Meditec Inc sent a SAFETY ALERT letter dated May 18, 2009, via Federal Express to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were reminded to use the correct retinal protection filter (GG475) when using the OPMI VISU 200 system during ophthalmic surgery.Verify the position of the retinal protection filter and inform their staff of the customizable feature. Verify that the correct filter was in place. Customers were asked to complete the acknowledgement form indicating receipt of the alert. Sign and return the form to:Carl Zeiss Meditec in Dublin, California via the enclosed self-addressed, pre-paid Federal Express envelope.For any questions customers should call (925) 557-4422. |  | 2540 | 2540 units | Worldwide Distribution including USA (nationwide) and the UK | Terminated | February 09, 2011 | 632 | 2540 | Not_Computer
Z-0604-2012 | 60479 | No | DRY | Cardiopulmonary bypass on-line blood gas monitor. | Cardiovascular | Cardiovascular | 510(k) | CDI 101 Hematocrit/Oxygen Saturation Monitoring System. The CDI 101 Hematocrit/Oxygen Saturation Monitoring System is intended for use during cardiopulmonary bypass procedures where continuous monitoring of the blood's oxygen saturation, hematocrit and hemoglobin levels is needed. |  |  | December 07, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | Terumo is recalling the CDI 101 Hematocrit/Oxygen Saturation Monitoring System because they have not yet submitted a 510(k) to FDA to market the device. | No Marketing Application | 12/7/11 &amp; 12/8/11 All consignees received a "Transition and Notification Guide" dated 12/6/11 and a "URGENT MEDICAL DEVICE REMOVAL" letter dated 12/4/11, which will inform them of the removal of the CD l 101 and provide the user with various options. Terumo will remove CD l 101 units from the US market over the next 18 months. |  | 366 | 366 | Nationwide Distribution - Including the states of: NV, MO, SC, WA, MI, FL, GA, NY, CT, IL, VA, NY, TN, IN, HI, TX, AL, OH, MD, AZ, CA, KS, MD, PA, and NM. | Terminated | August 30, 2013 | 632 | 366 | Not_Computer
Z-2194-2011 | 58490 | No | HBG | Manual cranial drills, burrs, trephines, and their accessories | Neurology | Neurology | 510(k) | Integra NeuroSciences, Combo Kit, Lot 171237, Integra LifeSciences 3498 West 2400 South, West Valley City, UT 84119. |  |  | February 22, 2011 | 2011 | Integra LifeSciences Corp. | Outer packaging of Combo Kit was mislabeled with an incorrect lot number. | Labeling mix-ups | Integra LifeSciences Corp. sent a " Urgent Product Recall Notification" letter dated February 22, 2011, to the only affected consignee. It was reported the Combo Kits of SPOO98 / INS- 4500 ( Hermetic Large Style Ventricular Catheter set and drill set for NY) - Catalog # 31046 Lot Number 171237, Outer packages were labeled incorrectly. The customer was advised to return this, and they would be replaced with correctly labeled product. The issue is easily recognizable, and does not pose a safety risk to patients.An acknowledgement form was included to be returned to Integra. Further information is available at 609-936-2485. |  | 7 | 7 Kits | Nationwide Distribution -- including the state of New York. | Terminated | May 13, 2011 | 80 | 7 | Not_Computer
Z-2715-2011 | 59174 | No | FPB | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Pall Total Nutrient Admixture Filter, 1.2 micron filter |  |  | June 15, 2011 | 2011 | Pall Life Sciences Puerto Rico Llc | Incorrect filter media used (0.2 um filter used instead of the required 1.2 um) within the TNA1 filter used with total parenteral nutrition mixtures. | Labeling mix-ups | Pall Life Sciences LLC, sent an "URGENT MEDICAL DEVICE RECALL" letter dated June 15, 2011 to all affected customers. The letter described the product, problem, and actions to be taken by the customers.The letter instructed customers to maintain the product in quarantine. If distributed, the letter request that customers notify the user to discontinue use. Additionally, customers were asked to complete and return an attached Recall Response Form. The letter also included instructions on how to return the product.Please contact the Customer Service Department at 516-801-9923, if you have questions. |  | 15575 | 15,575 units | Worldwide Distribution--USA (nationwide) including the states of CA, CO, CT, IL, KS, LA, MA, MO, MS, NC, NY, OK, SD, TX, WA, and WY, and Puerto Rico and the countries of Colombia and Costa Rica. | Terminated | January 10, 2013 | 575 | 15575 | Not_Computer
Z-2798-2011 | 58863 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | Terumo(tm) Hypodermic Needle, sterile, individually packaged into a "blister-pack" flexible plastic bottom-web thermo-formed to create a pocket, heat-sealed with paper top-web,1,000 needles per case. |  |  | June 08, 2010 | 2010 | Terumo Medical Corporation | Hypodermic needle packaging defect could compromise sterility. Affected lot numbers are: MC1725, MC2025, MC2025 | Packaging | The firm, Terumo Medical Corporation sent an "URGENT PRODUCT BULLETIN: RECALL OF TERUMO PRODUCT" letter dated June 8, 2010. Consignees were contacted via telephone on June 7 thru 9, 2010 to confirm receipt of the Product Recall Bulletin. The letter describes the product, problem and the actions to be taken. The letter instructed customers to immediately discontinue use and quarantine any available stock of the listed product codes. The letter stated that distributors must not distribute any of the listed product codes, quarantine any available stock and send the recall bulletin immediately to each customer that received the recalled products to inform them of this issue. Additionally, Terumo Medical Corporation sent a "Terumo Customer Instructions for Returning Recalled Product and Customer Confirmation of Product Recall Notification" letters dated June 8, 2010 to all affected facilities. Customers were instructed to fill-out the requested information and return to Terumo Medical Corporation (TMC) Product Recall Coordinator. Consignees were contacted via telephone on June 8 and 9, 2010 to confirm receipt of the recall information and compliance with the requested action. For questions or need for further information, call (800) 283-7866 and ask to speak with the Product Recall Coordinator. |  | 453000 | 453,000 needles | Nationwide Distribution (USA) -- AZ, CA, CO, FL, GA, IL, KS, KY, MA, MO, NC, NJ, NY, OH, PA, TN, TX, WA, UT, and VA | Terminated | July 14, 2011 | 401 | 453000 | Not_Computer
Z-0273-2012 | 60123 | No | DTS | Cardiopulmonary bypass cardiotomy return sucker. | Cardiovascular | Cardiovascular | 510(K) Exempt | IBC INTERNATIONAL BIOPHYSICS CORP. 2101-2 EAST ST. ELMO, SUITE 275, AUSTIN, TX 78744 SUCTION HANDLE (STERILE) PART NUMBER: 1990S LOT NUMBER 041811-2065 EXPIRATION DATE: 18/APR/2015 |  |  | October 05, 2011 | 2011 | International Biophysics Corp | Potential for the sterility to be compromised. | Packaging | International Biophysics Corp. sent a" Medical Device Recall Notice" dated October 5, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Use of the product in the referenced lot number should cease immediately. Customers were asked to follow the instructions in the letter to return the affected product.If you have any further questions please call (512) 814-0070. |  | 1200 | 1200 units | Nationwide Distribution including the states of AZ, CA, CO, DE, MS, OR, PA, TN, and TX | Terminated | March 23, 2012 | 170 | 1200 | Not_Computer
Z-2884-2011 | 58921 | No | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | LPCCC2: Low Profile Abutment Non-Hexed Castable Cylinder by Biomet 3i |  |  | March 14, 2011 | 2011 | Biomet 3i, LLC | Biomet 3i, Palm Beach Gardens, FL is recalling the Model LPCCC2 Low Profile Abutment, Non-Hexed Castable Cylinder, Lots 2010101036 and 2010101037  Exp. 10/2015. Package labeling may not correctly describe the package contents. Product was distributed in Europe only. | Labeling mix-ups | Biomet 3i sent a "Medical Device Advisory Notice" letter dated February 17, 2011 to all affected customers.The letter described the problem, instructed customers to returned any unused product and to confirm receipt of the notification.For additional information, please contact Customer Service at 800-342-5454. |  | 232 | 232 unit | Worldwide distribution: http://webapps.cdrh.fda.gov/division-tracking/findTrackable.do?docNum=RCL110656&reviewType;=RCL | Terminated | July 27, 2011 | 135 | 232 | Not_Computer
Z-2837-2011 | 59263 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Baxter Interlink System Non-DEHP I.V. Catheter Extension Set, 6.6 (17 cm), Volume 0.8 mL, Injection Site, Male Luer Lock Adapter; an Rx sterile, single use nonpyrogenic fluid pathway; 200 sets per case; Baxter Healthcare Corporation, Deerfield, IL 60015 USA; product code 2N3375 |  |  | July 06, 2011 | 2011 | Baxter Healthcare Corp. | Some of the individual extension set packages were improperly sealed, compromising the sterility of the product. | Packaging | The firm, Baxter, sent an "Urgent Product Recall" letter dated July 6, 2011 via first class mail to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to locate and remove all affected product from their facility; call Baxter Healthcare's Center for Service at 1-888-229-0001 between 7:00 am and 6:00 pm Central Time, Monday through Friday to return affected product and receive credit; complete and return the attached CUSTOMER REPLY FORM to Baxter via fax at 1-847-270-5457 or scan and email to fca@baxter.com and if you are a dealer, wholesaler or distributor/reseller, notify your customers of this recall. If you have questions regarding the communication, call The Center for One Baxter at 1-800-422-9837 during the hours of 8:00 am to 5:00 pm CST. |  | 9800 | 9,800 sets | Worldwide distribution: USA (nationwide) including: AL, LA, MO, SD, TX and WV; and country of: Canada. | Terminated | April 18, 2012 | 287 | 9800 | Not_Computer
Z-2171-2011 | 58216 | No | HBG | Manual cranial drills, burrs, trephines, and their accessories | Neurology | Neurology | 510(k) | Ventricular Bolt Monitoring Kit W/Cranial Access, Drain, 1104HMC, Integra LifeSciences, 3498 West 2400 South, West Valley City, Utah 84119. |  |  | February 22, 2011 | 2011 | Integra LifeSciences Corp. | Outer box of Monitoring Kit was mislabeled with the incorrect part number. | Error in labeling | Integra LifeSciences Corp. sent an Urgent Product Recall Notification letter dated February 22, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to quarantine any affected product and arrange for return to Integra by contacting Integra Customer Service Returns at 1-800-654-2873 (choose option 2) to obtain a Return Merchandise Authorization number and place an order for replacement product. Customers were asked to complete the attached "Recall Acknowledgement and Return Form" indicating whether or not they possess any of the product identified. For any questions regarding this recall call 609-936-2485. |  | 45 | 45 kits | Nationwide Distribution including CA, CT, GA, HI, IL, KY, MA, NC, NJ, NM. NV, OH, TX, WA, and WI. | Terminated | May 10, 2011 | 77 | 45 | Not_Computer
Z-0460-2012 | 60051 | No | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | SRS FAST SET PUTTY 3 CC BONE VOID FILLER-STERILE, Part Number SRS-005-FSP, Lot number N001101. |  |  | September 29, 2011 | 2011 | Synthes USA | A recent shipment carrying Norian SRS was lost by the carrier. Since the product was lost before final product release testing, Synthes is unable to confirm that Part Number SRS-005-FSP- Lot Number N001101 meets final product release requirements. The product contains all final labeling and cannot be visually distinguished from the product that has met final product release requirements. | Labeling Change Control | All Synthes accounts with a history of purchasing the product in the past two years will be notified by letter of the incident. All firm sales consultants will also receive written notification of the incident. |  | 131 | 131 units | Synthes USA sent a "NOTICE: MISSING SHIPMENT-NORIAN SRS" letter dated September 29, 2011 to all accounts with a history of purchasing the product in the past two years. The letter describes the product, problem, and actions to be taken by the customers. The letter instructs customers to check their inventory to verify they have not received the affected product. Call 1-800-620-7025 ext. 5450 for questions concerning this notice. | Terminated | December 18, 2014 | 1176 | 131 | Not_Computer
Z-2833-2011 | 59096 | No | JCT | Tracheal prosthesis. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Boston Scientific, WALLGRAFTÂżÂż OVER-THE-WIRE, Tracheobronchial Endoprosthesis, Made in IRELAND, Ballybrit Business Park, Galway, IRELAND, STERILE, EO. |  |  | May 25, 2011 | 2011 | Boston Scientific Corporation | Boston Scientific is initiating an important field correction related to WALLSTENT", WALLSTENT" RP, and WALLGRAFTÂż Tracheobronchial Endoprosthesis products. This letter is intended to clarify and re-emphasize that U.S. WALLSTENT", WALLSTENT" RP, and WALLGRAFTÂż Tracheobronchial Endoprosthesis stents are NOT currently cleared by the FDA for use in benign tracheobronchial strictures. | Labeling False and Misleading | The firm, Boston Scientific, sent an "Urgent Medical Device Field Correction" letter dated May 25, 2011 to its customers. The letter described the product, the problem and actions to be taken. The letter states that no product is being recalled and asks consignees not to return the product. The customers were instructed to read the enclosed information and post it in a way to ensure it is easily accessible to all users of this device; complete and return the Reply Verification Tracking Form via fax to: CV Field Action Center, 1-866-213-1806, and if they are a distributors to forward this notification to their customers. If you have any questions or would like assistance with this field correction, please contact your local Sales Representative and/or call 763-494-1133. |  | 23218 | 23,218 FOR ALL PRODUCTS | Worldwide distribution. | Terminated | August 10, 2012 | 443 | 23218 | Not_Computer
Z-2146-2011 | 58181 | No | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Single Day Infusor 2 mL/h Portable Elastomeric Infusion System; an Rx sterile disposable drug delivery system; Baxter Healthcare Corporation, Deerfield, IL 60015; REF 2C1071KJP |  |  | March 15, 2011 | 2011 | Baxter Healthcare Corp. | Some of the Single Day Infusor System 2 mL/h have a foil over-pouch that was incorrectly labeled as the Half Day Infusor System 5 mL/h, product code 2C1073KJP, lot 10H054. | Labeling mix-ups | Baxter Healthcare Corp. sent Urgent Product Recall letters dated March 15, 2011, via first class mail to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The accounts were requested to examine all of their inventory locations for both product code 2C1071KJP and 2C1073KJP, lot 10H054, and return the affected lot for replacement or credit. The accounts were also requested to complete the attached customer reply form listing the number of held for return, and fax it to Baxter at the number provided on the form. Any questions regarding the communication were directed to Baxter's Center for One Baxter at 1-800-422-9837. |  | 3696 | 3,696 units | Worldwide Distribution - USA including AL and the countries of Mexico and Spain | Terminated | March 12, 2012 | 363 | 3696 | Not_Computer
Z-2799-2011 | 58768 | No | NSQ |  |  | Dental | Enforcement Discretion | "***REF IWTH56***Certain EP HEALING ABUTMENT 5mm(D) X 5mm(P) X 6mm(H).*** Healing abutment*** STERILE.***Rx Only.*** Made in USA*** Implant Innovations, Inc. Global Headquarters 4555 Riverside Drive, Palm Beach Gardens, FL 33410 USA Tel: (800) 342-5454 www.3i.online.com. Outside US 561-776-6700***Implant Innovations Siberica S.L., WTC Almeda Park, Ed. 1 Planta 1a Pl. De la Pau s/n 08940-Cornella de Llobregat (Barcelona) Spain. Phone: +34 934 705500. Fax +34 933 717 849. Lot Number: 602691.***" |  |  | February 23, 2010 | 2010 | Biomet 3i, LLC | Biomet 3i of Palm Beach Gardens, FL is recalling their Certain EP Healing Abutment (IWTH56). The Healing Abutment dimensions did not match outside label and laser etching on part. The product was actually 8mm but the package indicated 6mm. | Error in labeling | The firm, BIOMET 3i, initially contacted each US customer by telephone using a prepared phone script. Following telephone contact, the firm followed up each call with a faxed "Urgent: Medical Device Recall" letter dated March 1, 2010. International customers were notified first by email then by fax. The letter and email described the product, problem and actions to be taken. The customers were instructed to return the healing abutment(s) for replacement if they have not already opened and used the product; respond with the attached response fax and return the product(s) to the: Regulatory Affairs Department, BIOMET 3i, 4555 Riverside Drive, Palm Beach Gardens, FL 33410, Incident No. IC95863. BIOMET will send a replacement healing abutment(s) once the returned product is received.If you have any questions or concerns, please contact the Recall Coordinator at 561-776-6906. |  | 49 | 49 | Worldwide distribution: USA (nationwide) and countries of: Australia, Europe and Japan. | Terminated | November 09, 2011 | 624 | 49 | Not_Computer
Z-0852-2011 | 57352 | No | FAD | Ureteral stent. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Cook Medical, Universa Soft Ureteral Stent and positioner, 6 Fr x 24 cm, REF USH-624-R, sterile, Cook Urological, Inc., Spencer IN. |  |  | October 18, 2010 | 2010 | Vance Products Incorporated | The firm received a customer complaint that a box labeled as Universa Firm Stents but the box actually contained Universa Soft Stents. The inner product label is correct but the outer box is wrong. | Packaging | The firm sent URGENT: RECALL NOTICES dated 10/18/2010 to its direct accounts. The notice identified the affected product and instructed firms to pull the recalled product from stock and isolate it. Customers are to complete the form and e-mail or fax the form back to Cook. Product should be returned to the recalling firm using the Return Goods Authorization number. Cook Urological will credit customer accounts upon receipt of the returned product. |  | 25 | 25 units | Worldwide Distribution -- USA, including states of WI, TX, NY, SC, AZ, TN, FL, AR, CA, AL, and WA and country of Canada. | Terminated | July 27, 2011 | 282 | 25 | Not_Computer
Z-0205-2012 | 60392 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | SUPER TORQUE MB Angiographic Catheter product is labeled in part: "***5F (1.65 mm)***110 cm .038" (0.97 mm)***MULTIPURPOSE A SUPER TORQUE MB***Max. 1200 psi (8274 kPa)***Markers***OPEN end***Sideholes***SPECIAL***Angiographic Catheter*** REF Cat No.***Lot***Use By***Sterile EO***Assembled in Mexico Cordis Corporation, 14201 NW 60th Ave. Miami Lakes, Florida 33014, USA***EC***REP*** Cordis Cashel, Cahir Road Cashel, Co Tipperary, Ireland***Cordis***a Johnson&Johnson; company***" |  |  | November 21, 2011 | 2011 | Cordis Corporation | Cordis Corporation is issuing a Correction for their SUPER TORQUE MB Angiographic Catheter after receiving a report of marker bands dislodged during a procedure. Model Numbers 532598A, 532598B, 532598C, 532598D, SRD5724MB and SRD5727MB are affected by the Cordis Correction. | Error in labeling | Cordis Corporation sent a Correction letter dated November 21, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. Customers are instructed to read the "Description" and "Recommendations" sections, sign and return an enclosed Acknowledgement Form, and to maintain awareness of the communication until the information has been incorporated into the Cordis Super Torque MB Catheter labeling. Contact Cordis at 800-781-0282 for questions regarding this notice. |  | 310 | 310 units | Worldwide Distribution-USA (nationwide) including Puerto Rico. | Terminated | July 23, 2013 | 610 | 119000 | Not_Computer
Z-0514-2012 | 60425 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo Imaging XS, all versions |  |  | September 20, 2011 | 2011 | Siemens Medical Solutions USA, Inc | Unintended behavior when using syngo Imaging XS, all versions. The "Rename" functionality can cause an unintended patient merge if the renamed patient is sent to another system. | Labeling False and Misleading | Siemens Medical Solutions USA, Inc sent a Customer Safety Advisory Notice letter dated August 3, 2011 to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.If you detect patient data inconsistencies in your system caused by this problem, please contact Siemens Service for support.For further questions please call (610) 219-4834. |  | 38 | 38 | Nationwide Distribution including the states of CA, FL, GA, IN, IA, KS, MN, MO, NE, NV, NJ, NY, NC, OH, PA, TN, TX, and WI. | Terminated | March 05, 2014 | 897 | 38 | Not_Computer
Z-1668-2011 | 56968 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Abbott brand CELL-DYN Emerald Diluent Reagent, 10 mL, List Number: 09H48-02, |  |  | September 30, 2010 | 2010 | Abbott Laboratories | The product bar code label, when scanned, indicates that the volume is 10 mL rather than the correct volume of 10,000 mL. Analytical results are not impacted, although, there is a potential for delay in generating results. | Error in labeling | Product Recall notification letters were sent on 9/30/2010. The letter identified the affected product, an explanation of the problem, patient impact, and necessary actions. Consignees are to contact their local hematology customer support representative to arrange for replacement of the CELL-DYN Emerald Diluent Lot. Once the replacement lot has been received, customers should destroy the affected lot according to their laboratory procedures. A copy of the letter should be retained for their records and the Customer Reply Form should be completed and returned via fax or e-mail. US customers should contact Customer Support at 1-877-4ABBOTT, if they have questions. |  | 322 | 322 units | Worldwide Distribution -- USA, including states of PA, NY, CA, NC, IL, TX, KS, UT, VA, FL, AR, TN, MN, NE, MI, and DE and the country of Germany. | Terminated | April 07, 2011 | 189 | 322 | Not_Computer
Z-0264-2012 | 60365 | No | DRE | Vessel dilator for percutaneous catheterization. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific Super Sheath and Super Sheath R/O Introducer Sheaths, Sterile EO, Manufactured by Togo Medikit Co., Ltd. 17148-6 Aza Kamekawa, Oaza Hichiya, Hyuga City, Miyazaki Prefecture 883-0062, Japan, Distributed by Boston scientific Corporation One Boston Scientific Place, Natick, MA 01760-1537 USA. |  |  | November 03, 2011 | 2011 | Boston Scientific Corporation | Boston Scientific is initiating a medical device recall removal regarding nine lots/batches of Super Sheath and Super Sheath R/O Introducer Sheaths. Boston Scientific was informed through product complaints that the 0.035" ID dilator may be labeled incorrectly as a 0.038" ID dilator. | Packaging process control | Boston Scientific sent an URGENT MEDICAL DEVICE RECALL REMOVAL -IMMEDIATE ACTION REQUIRED letter dated November 21, 2011. The letter described the product and the problem. Customers were instructed to segregate the affected product immediately and return it to Boston Scientific. Distributors were advised that the recall removal depth is to the hospital level and the recall removal notification should be forwarded to their customers.For questions regarding this recall call 763-494-1133. |  | 2660 | 2660 | Worldwide Distribution - USA including AZ, AR, CA, CO, FL, GA, IL, KY, LA, MD, MA, MN, MO, NE, NJ, NM, NY, OH, OK PA, RI, TN, TX, UT, VA, WY, WI, and WV and the countries of BELGIUM, BRAZIL, CZECH REPUBLIC, FRANCE, GERMANY, GREAT BRITAIN, IRAQ, ITALY, NETHERLANDS, RUSSIAN FEDERATION, SOUTH KOREA, SPAIN, SWEDEN, and UKRAINE. | Terminated | January 15, 2013 | 439 | 2660 | Not_Computer
Z-2790-2011 | 59035 | No | BTT | Respiratory gas humidifier. | Anesthesiology | Anesthesiology | 510(k) | HUDSON RCI, Humidifier Adaptor, Rx only, Teleflex Medical, Research Triangle Park, NC 27709. |  |  | June 06, 2011 | 2011 | Teleflex Medical | The packaging may not have been properly sealed, and therefore the sterility of the product cannot be guaranteed. | Packaging | Teleflex Medical sent an "URGENT MEDICAL DEVICE RECALL NOTIFICATION" letter dated June 6, 2011 to all affected customers. The letter described the product, problem, and the action to be taken.The letter instructs customers to discontinue use and quarantine the affected product. Customers must complete and return a Recall Acknowledgement Form via fax to 1-866-804-9881, Attn: Customer Service. A representative will contact customers with instructions on how to return the affected product. Customers with no affected product are also instructed to complete and return the form to verify receipt of the notification.Contact Customer Service at 1-866-246-6990 if you have questions regarding this recall. and a record of returns. A third letter was addressed to the Risk Manager with the above instructions. |  | 218880 | 228, 960 eaches | Worldwide Distribution--USA (nationwide) including the states of AL, AZ, CA, CO, FL, GA, IL, IN, KS, KY, MI, MN, MO, MS, NM, NE, NY, OH, OK, PA, TX, VA, WI and WM and the countries of Japan, Philippines, and Thailand. | Terminated | May 01, 2014 | 1060 | 218880 | Not_Computer
Z-2140-2011 | 58342 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Biomet Interlok 67 MM fixed Cruciate Tibial Plate with locking bar CO-CR-MO/TI 6AL 4V ALLOY, REF 141232, Sterile, BIOMET ORTHOPEDICS, WARSAW. IN I-Beam Tibia Plates are knee joint replacement prostheses intended for application with or without bone cement. |  |  | March 03, 2011 | 2011 | Biomet, Inc. | The firm is recalling the product due to the product's packaging being labeled "I-Beam 75mm. Tibial Plate", but the package actually contains the "Interlok 67mm. CR Tibial Plate". | Packaging process control | Biomet Orthopedics Inc initiated an electronic communication with Direct Accounts on March 3, 2011 and instructed them to quarantine and return all remaining/unused devices to the firm. For questions regarding this recall call 574-372-3983. |  | 8 | 8 | Worldwide Distribution -USA including FL, IN, and NC and the countries of Japan and the Netherlands | Terminated | July 27, 2011 | 146 | 8 | Not_Computer
Z-2794-2011 | 59057 | No | GFA | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Stryker 4.0 mm. Angled Aggressive 12 Degree Blade, Precision Cutters, Stryker Instruments, Kalamazoo, MI 49001 USA, (269) 323-7700 (800) 253-3210, Sterile, Single Use, Rx ONLY |  |  | June 13, 2011 | 2011 | Stryker Instruments Division of Stryker Corporation | Recalled device was packaged with the incorrect angled blade. The labeling indicates a 12 degree angle but the product has a 40 degree angle. | Labeling mix-ups | Stryker Instruments Division of Stryker CorporationOn 6/13/11 sent an 'URGENT MEDICAL DEVICE RECALL NOTIIFICATION" letter to all affected customers that included the device name, lot number, dates of distribution, reason for the recall, risks associated with use of the device, a photo of the device, instructions to quarantine and return the devices, and a form to fax back to the firm to confirm receipt and adherence to the recall letter. |  | 95 | 95 Units | Worldwide Distribution - USA including CA, LA,, MI, NJ, OH, OR, AR, IN, and Internationallly to Australia, Duisburg, Newbury and Hong Kong. | Terminated | August 03, 2011 | 51 | 95 | Not_Computer
Z-2103-2011 | 58271 | No | OCJ | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | ConfidenceÂż Introducer Needle Side Hole 13G x 4-Inch |  |  | March 22, 2011 | 2011 | DePuy Spine, Inc. | Mislabeled: Package labeled (11g x 6 inch) contains a 13g x 4-inch side hole needle | Packaging process control | The firm, DePuy Spine, sent an "URGENT-VOLUNTARY PRODUCT RECALL NOTIFICATION" letter dated March 22, 2011 to its consignees/customers ( hospital sites and domestic sales distributors). A product reconciliation form was provided to assist in the retrieval of affected product. The letter described the product, problem and actions to be taken. The customers were instructed to return any needles that they may have to their local Distributor for DePuy Spine products and complete and return the Product Reconciliation Form via fax to 508-828-3762; call Customer Service at 877-379-4871 to receive a return goods authorization number and/or mail the form with the return product to DePuy Spine Inc., 50 Scotland Boulevard, Bridgewater, MA 02324, Attn: CONFIDENCE NEEDLE RECALL COORDINATOR, RGA #.If you have any questions or concerns with regard to this notice, please contact the DePuy Spine Regulatory Compliance Department at 508-828-2790. |  | 67 | 67 units | Worldwide distribution: USA including states of: AZ, IA, LA, TX, NE, NJ ,NY, PA, VA, and WA; and countries of: Canada, Australia, Belgium, France,Hungary, Italy, Malaysia, Solvenia, UK, South Africa, Switzerland, and UAE. | Terminated | February 20, 2013 | 701 | 282 | Not_Computer
Z-1747-2011 | 58138 | No | MGR |  |  | General &amp; Plastic Surgery | PMA | Organogenesis Apligraf - Interactive burn and wound dressing, supplied as a living, bilayered skin substitute |  |  | February 21, 2011 | 2011 | Organogenesis, Inc. | Product sterility may be compromised | Packaging process control | Organogenesis contacted the accounts by telephone and fax on 2/24/11 to advise of the contamination and identification fo the organism for appropriate patient follow-up. |  | 7 | 7 units | Nationwide. | Terminated | July 28, 2011 | 157 | 7 | Not_Computer
Z-2028-2011 | 57340 | No | GWM | Intracranial pressure monitoring device. | Neurology | Neurology | 510(k) | Codman Disposable Twist Drill |  |  | November 30, 2010 | 2010 | Codman &amp; Shurtleff, Inc. | Packaging defect may compromise sterility of device | Package design/selection | The firm, Codman, sent an "Important Notice Urgent- Voluntary Product Recall Notification" letter dated November 30, 2010 to its customers via FedX. On January 7, 2011, Codman expanded the recall to four additional product codes.On January 31, 2011, the firm issued a supplemental recall letter and identified one additional product code: 80-11197. The letters described the product, problem and actions to be taken. The customers were instructed to identify the unused affected products in their inventory and follow the recall instructions provided: 1) If you have NO product to return, please advise Codman as soon as possible using the FACSIMILE form attached and fax to 508-828-6118 and 2) if you DO have product to return, please call 1-800-225-0460, Option 4, for instructions for return.Codman is presently working toward supplying replacement products as soon as possible. Please call Customer Service to check on the status of product availability at 1-800-225-0460.Should you have any questions regarding this recall, please call Customer Service at 1-800-225-0460 or your local Codman Representative. |  | 1549 | 1549 units | Worldwide distribution: USA including states of :AL, FL,GA, KY, MA, MI, MN, MS, OR, PA, RI, UT,WI and WV; and countries including: Argentina, Australia, Canada, Bahamas , Belgium, Dubai, Japan, Madrid, Mexico, New Zealand, Saudia Arabia, UAE, Uruguay, and United Kingdom. | Terminated | March 14, 2013 | 835 | 14792 | Not_Computer
Z-2192-2011 | 58451 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | BD Nokor Admix Needle |  |  | March 01, 2011 | 2011 | Becton Dickinson &amp; Company | Incorrect unit package labels which reflect Reorder #305216 instead of the correct item, Reorder #305213, were used. | Packaging | BD sent an Urgent: Product Recall letter with return response forms to direct consignees on 03/01/11 and 03/02/11 via UPS. The letter identified the affected product and lot along with the reason for the recall. The letter asked customers to complete the enclosed Distributor Packing List and to follow the instructions provided on the Packing List for returning the affected product. The firm will issue a credit upon receipt of the affected product. Customers can contact BD Customer Service if they need assistance or have questions regarding the recall. |  | 161000 | 161,000 units | Worldwide Distribution -- USA, including the states of NJ, IL, LA, NM, OH, NV, and VA, and country of Canada. | Terminated | November 30, 2012 | 640 | 161000 | Not_Computer
Z-0229-2012 | 57585 | No | OJV | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Angiotech***Custom Bone Marrow Tray for University of CA, Davis*** 1 - 10cc Syringe L/L; 1 - 20cc Syringe L/L; 1 - 19ga x 1-1/2" Filter Needle; 1 - 21ga x 1-1/2" Needle; 1 - 25ga x 5/8" Needle; 1 - 8ga x 4" Bone Marrow Needed "J" Type; 1 15ga x 4" Bone Marrow Needle "I" Type; 1 - T- Lok Extraction Cannula; 1 - 10% Providone Iodine Swab Sticks (3/Pkg.); 1 - Probe Guide; 1 - Obturator; 2 - 1% Lidocaine (5mL); 1 - #11 Mini Scalpel; 2 - 4" x 4" Sponges in stacks of 5; 1 - 2" x 3" ADH Dressing; 1 - Towel; 1 - Fenestrated Drape; 1 - Male Luer Cap; 1 - CSR Wrap***Catalog Number: CB0140TL***Sterile - For single patient use. STERILE EO. Caution: Federal (USA) Law restricts this device to sale by or on the order of a physician |  |  | December 22, 2010 | 2010 | Medical Device Technologies, Inc. | Medical Device Technologies, Inc. dba Angiotech, is recalling HSG Procedure Trays and Bone Biopsy Trays distributed between December 2007 and December 14, 2010 for package integrity. | Packaging | The firm, Angiotech, sent an "URGENT PRODUCT RECALL NOTIFICATION" letter dated December 22, 2010 to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to review their current inventory, segregate the affected lot numbers and complete and return the attached Customer Acknowledgement Form via fax to our Quality Assurance department at 1.352.338.0662 or 1.800.333.0440 to arrange for return of the identified product.If you have any questions concerning this notification, call (352) 338.0440 ext. 355 or ext 358. |  | 225 | 225 boxes | Worldwide distribution: US (nationwide) and countries of: Canada, Columbia, Costa Rica, Japan, Mexico, Peru and Saudi Arabia. | Terminated | November 30, 2011 | 343 | 32491 | Not_Computer
Z-2873-2011 | 59002 | [6 empty] | "***REF: 91-6704***THINFLAP SYSTEM CROSS-DRIVE, SELF-DRILLING SCREW 1.5X4MM (.059X.159") (WXL)***TITANIUM***NON-STERILE PRODUCT***BIOMET MICROFIXATION 1520 TRADEPORT DRIVE JACKSONVILLE, FL 32218 USA*** www.biometmicrofixation.com***QTY 00001***". |  |  | May 17, 2010 | 2010 | Biomet Microfixation, Inc. | On May 17, 2010, Biomet Microfixation issued a Safety Alert on their Thinflap System Cross Drive, Self Drilling Screw, 1.5x4mm, Model #: 91-6704, Lot#: 173770. The product is labeled incorrectly. Some packages of 91-6704 lot 173770 are labeled as cross drive screws, but contain center drive screws. All affected customers were notified of the recall on May 17, 2010. | Labeling mix-ups | Biomet Microfixation notified distributors and customers by letter, email and telephone calls. Distributors and customers were instructed to fill out and return the Inventory Reconciliation sheet, via Fax to 904-741-9425. Distributors were also instructed to notify their customers if the product was distributed further.For questions regarding this recall customers were instructed to call 1-800-874-7711 or 904-741-4400, ext. 468. |  | 927 | 927 | Worldwide Distribution - USA including AL, CA, GA, HI, IL, MO, NC, NY, TX and the country of Japan. | Terminated | October 12, 2011 | 513 | 927 | Not_Computer
Z-2124-2011 | 58470 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | BardÂżVentralex" Hernia Patch Medium Circle with Strap Diameter 6.4cm /2.5",\ |  |  | April 11, 2011 | 2011 | Davol, Inc., Sub. C. R. Bard, Inc. | Mislabeled: BardÂż Medium Ventralex Hernia Patch 6.4cm /2.5" packaged in a pre-printedcarton identified for a BardÂż Ventralex Small Circle 4.3cm /1.7" | Labeling mix-ups | Davol sent an Urgent: Medical Device Recall letter on April 11, 2011, via FedeX Priority Overnight to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their inventory and remove and return the recalled lot. Distributors were requested to notify their accounts.For questions contact the Davol Customer Service Department at 1-800-556-6275 or C. R. Bard's Medical Services &amp; Support Department at 1-800-562-0027. |  | 146 | 146 units | Nationwide Distribution - USA including AZ, CA, CO, CT, FL, GA, IA, IL, IN, KS, MA, MD, MI, MO, MT, NC, NJ, NV, NY, OH, OR, PA, RI, SC, TN, TX, VA, WA, and WV. | Terminated | October 27, 2011 | 199 | 146 | Not_Computer
Z-0467-2012 | 60590 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | MORPHEUS CT PICC with STYLET, 5F Dual Lumen 65 cm Insertion Kit, Catalog No./REF 1200923, STERILE, AngioDynamics Queensbury, NY 12804 USA --- COMMON/USUAL NAME: Percutaneous, implanted, long-term intravascular catheter; CLASSIFICATION NAME: Percutaneous, implanted, long-term intravascular catheter; Device Listing # D095274. |  |  | December 05, 2011 | 2011 | Angiodynamics Worldwide Headquarters | AngioDynamics distributed the affected PICC catheters that were beyond their expiration date. | Incorrect or no expiration date | AngioDynamics sent an "URGENT-MEDICAL DEVICE RECALL" letter dated December 2, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. A Recall Reply Form was attached to the letter for customers to complete and return to the firm via fax at 518-798-1360. Contact your local sales representative or AngioDynamics Customer Service at 800-772-6446 for questions regarding this notice. |  | 16 | 16 units domestically | Nationwide Distribution-including the states of Florida, Georgia, Montana, and Tennessee. | Terminated | September 27, 2012 | 297 | 16 | Not_Computer
Z-2894-2011 | 59113 | No | CAZ | Anesthesia conduction kit. | Anesthesiology | Anesthesiology | 510(k) | Arrow ASK-20608-CMC StimuCath Continuous Nerve Block Kit. |  |  | June 17, 2011 | 2011 | Arrow International Inc | Some of the outer trays used in the packaging were found to have an insufficient seal. Therefore, packaging may have been compromised, and sterility of the product cannot be guaranteed. | Packaging | Arrow International, Inc. sent an "Urgent Medical Device Recall" letter dated June 17, 2011 to all affected customers.The letter describes the problem and provides recommendations including; quarantine affected product, return of product, and send an acknowledgement form back to manufacturer.For additional information call 800 233-3187. |  | 95 | 95 kits | Nationwide Distribution including the state of North Carolina. | Terminated | October 01, 2012 | 472 | 95 | Not_Computer
Z-2783-2011 | 59115 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | Bard PowerLine 6 Fr. Dual-Lumen Polyurethane Catheter with SureCuff Ingrowth Cuff, REF 0700610, Sterile, Contents...1 Each - PTFE Introducer and Dilator, 6.5 FR. x 10 cm..., Assembled in Mexico. |  |  | June 20, 2011 | 2011 | Bard Access Systems | Kit label indicates the wrong dimension for an enclosed introducer. | Labeling design | Bard Access Systems, Inc. notified customers by letter dated June 16, 2011. The letter notified customers of the mislabeling issue and offered a replacement product if desired. For additional information contact Bard Access Systems at 1-800-443-3393. |  | 6485 | 6,485 units | Worldwide distribution - USA (nationwide) including the states of: HI, KY, TX, VT and the country of Canada. | Terminated | March 01, 2012 | 255 | 6485 | Not_Computer
Z-1727-2011 | 56962 | No | OJZ |  |  | Orthopedic | HDE - Humanitarian Device Exemption | Stryker Biotech OP-1 Putty (OP-1, BMP-7, bone morphogenetic protein) |  |  | October 25, 2010 | 2010 | Stryker Biotech | A change to the Preparation for Use sectionwas identified where 2.5cc was replaced with 2cc-3cc volume and was not approved | Labeling Change Control | The firm, Stryker Biotech, sent a "Medical Device Correction" letter on October 15, 2010 via return receipt letter to all customers. The letter contains labeling with the approved Preparation for Use Section of 2.5 cc volume and included the revised copy of the Package Insert. Additionally, the letter described the product, problem and actions to be taken. The customers were instructed to follow the new instructions provided and to keep the new instructions in mind when using the OP-1 Putty.If you have any questions concerning this communication, please contact 508-416-5326 or Stryker Biotech's Pharmacovigilance department at PVG@stryker.com. |  | 2861 | 2861 units | Nationwide distribution. | Terminated | March 20, 2012 | 512 | 2861 | Not_Computer
Z-2770-2011 | 59003 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Langston Dual Lumen Catheters, REF 5515, 7F, Sterile EO, Rx ONLY, 1 x Catheter and 5 x Catheter, Vascular Solutions, Inc., 6464 Sycamore Court, Minneapolis, MN 55369. |  |  | June 09, 2011 | 2011 | Vascular Solutions, Inc. | Vascular Solutions, Inc investigation of a recent Device Experience Report has made them aware of a potential problem with their 7F Langston dual lumen pigtail catheter (Model 5515) with lot number 548264.Recently it was reported that pouches containing Langston catheters (Model 5515) were missing a seal,posing a risk of product contamination. It is possible that a compromised sterile barrier | Packaging process control | Vascular Solutions, Inc. sent an "URGENT MEDICAL DEVICE FIELD ACTION- LOT SPECIFIC" letter dated June 7, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter instructed customers to check their inventory for the affected product and provided information on how to examine the product. Customers are to complete and return a Field Action Customer Inventory Form via fax at 763-656-4250 or e-mail at www.vascularsolutions.com. VSI Customer Service will contact customers to set up return shipment of affected product. All affected product will be replaced upon receipt of the shipped product.Customers should contact their local Sales Representative for questions concerning this recall. |  | 23 | 23 (5 pack) boxes | Worldwide Distribution--USA (nationwide) including the states of CA, FL, IL, MA, MI, MO, NM, NC, OK, and WA and the countries of Austria and Canada. | Terminated | April 05, 2012 | 301 | 23 | Not_Computer
Z-1367-2011 | 53012 | Yes | NIQ |  |  | Cardiovascular | PMA | Endeavor Resolute Zotarolimus-Eluting Coronary Stent System with Rapid Exchange Delivery System, UPN 00613994400505, model #IDERES25024X. Manufacturer: Medtronic, Inc. Minneapolis, MN. Manufactured in: Medtronic Ireland, Galway, Ireland. Product is packaged inside a sterile, inner pouch, labeled in part "Use immediately". Inner pouch is supposed to be within a foil pouch containing desiccant, and moisture absorber. Foil pouch label reads in part "RX Endeavor Resolute Zotarolimus-Eluting Coronary Stent System...2.5 mm X 24 mm...REF IDERES25024X...Contents: One (1) Endeavor Resolute zotarolimus-eluting stent mounted on a rapid exchange stent delivery system...Guide Catheter/Minimum Inner Diameter &gt;=5F/0.056" (1.42 mm)...Maximum Guidewire Diameter 0.014" (0.36 mm)...Maximum Stent I.D. 3.00 mm (0.12"...Nominal Pressure 9 atm (912 kPa)...Rated Burst Pressure 15 atm (1520 kPa)...MR Conditional...Sterilized using ethylene oxide Consult instructions for use...Manufacturer: Medtronic, Inc. 710 Medtronic Pkwy NE Minneapolis, MN 55432 USA...Manufactured In: Medtronic Ireland Parkmore Business Park West, Galway, Ireland Tel:+353-91-708000 Fax:+353-91-757524 Investigational Devices / Returned Goods: Tel: (800) 556-4247 / Fax: (877) 518-2469...BioLinx Medtronic Polymer System". Foil pouch contents and Instructions for Use are placed within an outer box, which is labeled identically to the foil pouch. |  |  | August 06, 2009 | 2009 | Medtronic Vascular Galway Limited | The product was removed from investigational sites due to a packaging issue, in that the device units were packed without the foil pouch, oxygen absorber and Instructions For Use. | Packaging | The firm, Medtronic, sent their initial "Urgent Medical Device Recall Notice" to customers with affected product via e-mail on August 7, 2009. An additional "Urgent Medical Device Recall Notice" was sent to all customers on August 12-13, 2009 via FedEX and UPS. The notice described the product, problem and action to be taken by the customers. The customers were instructed to remove the affected units from their inventory and isolate the units in a secure location to prevent additional use. A Medtronic CardioVascular Clinical Customer Service (CCS) representative will contact their site to arrange for return of the units. This will be followed by a shipment of replacement units to their sites. If there were any questions regarding the return process, customers were to contact CCS at 800-556-4247.If you have any further questions or concerns regarding this notice, please contact the Sr. Clinical Research Manager at 707-543-2284 or ann.depaoli@medtronic.com. |  | 49 | 49 units | Nationwide distribution: AL, AR, CA, FL, IN, MD, MI, NY, OH, PA, RI, TX, VA, and VT. | Terminated | February 17, 2011 | 560 | 49 | Not_Computer
Z-0911-2011 | 56772 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Medical Systems, Eclipse Client software versions with build numbers from 7.3.10 through 8.6.15 (i.e. excluding 8.6.17); Model Number: H48; Radiation Treatment Planning System |  |  | August 17, 0010 | 10 | Varian Medical Systems, Palo Alto, CA | Possible incorrect dose calculation: An anomaly in certain Eclipse versions where, under certain specific sequences of user actions, the dose distribution may fail to be re-calculated even though changes have been made which affect dose. | Software design | Varian Medical Systems Urgent Medical Device Correction letters were sent to all affected users on August 6, 2010. Letters were also sent to Varian Sales, Marketing and Service organizations to inform them of the recall. |  | 7019 | 7019 | Worldwide Distribution. | Terminated | July 21, 2011 | 730823 | 7019 | Computer
Z-0909-2011 | 51550 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | CyberKnife System Robotic Radiosurgery System, (Robocouch Patient positioning System) Synchrony Respiratory Tracking System, a Radiation therapy device, Model Numbers: 026430, 028186, Product is manufactured and distributed by Accuray, Inc., Sunnyvale, CA |  |  | March 16, 2009 | 2009 | Accuray Inc | The product has the potential for the extension coupling of the Synchrony extension mounting hardware to become loose over time. The assembly weighs approximately 17 pounds; an unexpected contact with users or patients may result in serious injury if the boom descends unexpectedly. | Device Design | Accuracy Inc., sent an URGENT ADVISORY NOTIFICATION letter, dated March 16, 2009, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to inspect their system to see if it was affected. Customers were instructed to contact Accuracy Customer Support at 1-877-668-8667 (USA) or 1-408-716-4700 or e-mail customersupport@accuracy.com to schedule service. |  |  |  | Worldwide Distribution - USA (nationwide) and the countries of Turkey, Russia, France, Italy, the Netherlands, the UK, Spain, Greece, Switzerland, Germany, India, Taiwan, China, Hong Kong, Korea, Thailand, Malaysia and Japan | Terminated | January 19, 2011 | 674 |  | Computer
Z-1362-2011 | 57544 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ConMed System 2450 Electrosurgical Generator, REF 60-2450-100, 100 V, ConMed Electrosurgery, Centennial, CO. |  |  | December 27, 2010 | 2010 | ConMed Electrosurgery | Electrosurgical generator may allow possibility of electrical shock and/or burns to user or patient. | Device Design | ConMed Electrosurgery sent an Urgent Device Recall letter dted December 27, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to notify anyone who may have received the affected product.. They were to review their inventory and return a Business Reply Form via fax to 303-699-9854 and to contact ConMed Electrosurgery Technical Services at 800-552-0138 to schedule a time to have their equipment upgraded. |  | 119 | 119 units | Worldwide Distribution, including USA (nationwide) and the countries of Australia, Belgium, Bolivia, Canada, Chile, China, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, El Salvador, Finland, France, Honduras, Hong Kong, Hungary, India, Iraq, Israel, Italy, Jamaica, Japan, Jordan, Lebanon, Lithuania, Malaysia, Mexico, Netherlands, New Zealand, Norway, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Puerto Rico, Qatar, Romania, Saudi Arabia, Singapore, South Africa, South Korea, Spain, Switzerland, Taiwan, Thailand, Turkey, Uruguay, Venezuela, Vietnam. | Terminated | August 01, 2011 | 217 | 2323 | Computer
Z-0057-2012 | 59442 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | MicroMaxx ultrasound system and M-Turbo ultrasound system is Medical Diagnostic Ultrasound. |  |  | July 18, 2011 | 2011 | Sonosite, Inc. | This is a software anomaly. In some circumstances, if the Systolic/Diastolic (S/D) ratio value is interpreted directly from the display, the clinician could potentially use an incorrect value of the S/D ratio in his/her assessment. | Software design | The firm SonoSite, Inc., sent an "Urgent - Medical Device Correction" letter dated September 29, 2011 to its foreign and domestic customers. The letter described the product, problem, and actions to be taken. The customers were instructed to follow the three (3) steps provided in the letter to verify if the extended OB/GYN Calculations (EOBC) is enabled on their system. Additionally, SonoSite will perform a software upgrade to M-Turbo and/or MicroMax ultrasound system in order resolve the issue. The customers were instructed to contact Sonosite Technical Support via website: www.sonosite.com/sdratio; email: regulatory@sonosite.com (include contact information and address), or call in U.S. and Canada: 1-877-590-4982. If you have any questions or concerns regarding this matter, do not hesitate to contact in US and Canada at 1-877-657-8118, Monday-Friday, 6 am-5 pm PST. Customer can also send an email to Technical Support at service@sonosite.com |  | 4236 | 4236 units (US) and 4089 units (internationally) | Worldwide distribution: USA (nationwide) including Puerto Rico; and countries of: Afghanistan, Algeria, American Samoa, Argentina, Australia, Bolivia, Botswana, Brazil, Canada, China, Costa Rica, Cyprus, Dominican Republic, Estonia, France, Gambia, Germany, Great Britain, Guam, Guatemala, Honduras, Hong Kong, Hungary, India, Iran, Ireland, Italy, Japan, Jordan, Kenya, Korea, Kuwait, Mauritius, Mexico, Netherlands, New Zealand, Nigeria, Northern Mariana Islands, Oman, Panama, Paraguay, Peru, Philippines, Portugal, Qatar, Russian Federation, Saudi Arabia, Singapore, Slovakia, Slovenia, Tanzania, Uganda, United Arab Emirates, Viet Nam, and Zambia. | Terminated | September 14, 2012 | 424 | 4236 | Computer
Z-1245-2011 | 57379 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | GE Centricity Laboratory software; GE Healthcare IT, Barrington, IL 60010. The intended use: the Centricity Laboratory is used to facilitate the general clinical, anatomic pathology and cytology laboratory workflow such as order entry, results entry, instrument interfacing, results reporting and patient record retrieval. |  |  | December 03, 2010 | 2010 | Ge Healthcare It | The GE Centricity Laboratory System Software has safety issues with the "Telephone List" function, the "Anatomic Pathology" application, and the Instrument Interface IO module that may impact patient safety. | Software design | GE Healthcare, sent "Urgent Medical Device Correction" letters dated December 12, 2010 to GE Centricity Laboratory software customers on the same date. The letters advised the users of the product, the problem, and the action to be taken by the customer. The letters provided the users with specific safety instructions to follow until the software was updated. The customers were requested to complete the enclosed customer reply form and fax the completed form to 847-620-2780. Any questions should be directed to their GE Service Representative or the GE Help desk at 888-778-3375. |  | 26 | 26 units | Worldwide Distribution -- USA, including AR, AZ, CA, CO, DC, DE, FL, GA, HI, IA, IL, KY, MA, MI, MN, MO, NC, NH, NJ, NY, OR, TN, VA, and WI, , and internationally to the countries of Australia, England, Hong Kong, India, Malaysia, New Zealand, Qatar, Scotland and Singapore. | Terminated | July 15, 2011 | 224 | 26 | Computer
Z-0519-2012 | 60518 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Smith &amp; Nephew, SPIDER2 Limb Positioner, including accessories Battery, foot pedal and charger |  |  | July 26, 2011 | 2011 | Smith &amp; Nephew, Inc. Endoscopy Division | SPIDER2 may unlock (loss of traction) resulting inpotential movement of the patient when the device is within close proximity to an RF source. Lossof traction can potentially result in the unintended repositioning of the extremity and potential for tissue or cartilage damage | Device Design | Smith &amp; Nephew Advanced Surgical Devices Division sent an "URGENT- PRODUCT RECALL NOTIFICATION-URGENT" letter dated July 26, 2011 to all affected customers. The letter described the product, problem, and the actions to be taken by the customers. Customers were instructed to complete and return a form via fax at 405-917-8650. Contact the firm at 405-917-8556 for questions regarding this recall. |  | 52 | 52 units | Worldwide Distribution-USA (nationwide) and the countries of Canada and France. | Completed |  |  | 52 | Computer
Z-0520-2012 | 60592 | No | MNR | Breathing frequency monitor. | Anesthesiology | Anesthesiology | 510(k) | Lifescreen Apnea is Lifescreen ECG Screening System, Version 3.31. |  |  | November 11, 2011 | 2011 | Del Mar Reynolds Medical, Ltd. | When Lifescreen is left open and a new recording is loaded from Sentinel Cardiology InformationManagement System, customers can get corruption in the data or a wrong ECG after the first32 hours. Both recordings have to be longer than 32 hours for this to happen. | Software design | Spacelabs Healthcare sent an "URGENT-MEDICAL DEVICE CORRECTION" letter dated November 11, 2011 to all affected customers. The letter described the product, problem, and actions to be taken by the customers. Customers are advised to close Lifescreen every time before loading another recording from the Sentinel database. Spacelabs is in the process of upgrading software and will contact customers to arrange for an update to affected systems. If you have any questions about this corrective action program, please contact Spacelabs Healthcare at 1-800-522-7025 and select 2 for Technical Support. |  | 523 | total 523 units and 171 kits (34 units and 7 kits in US; 489 units and 164 kits outside US) | Worldwide Distribution-USA (nationwide) including the states of AR, CA, CO, CT, FL, GA, KY, LA, MA, MD, MI, MN, NC, NY, OH, TN, TX, UT, WV, and PUERTO RICO and the countries of AUSTRIA, BELGIUM, CHILE, CHINA, COLOMBIA, COSTA RICA, DENMARK, FINLAND, FRANCE, GERMANY, GREECE, HONDURAS, HONG KONG, HUNGARY, INDIA, ISRAEL, ITALY, KUWAIT, MALTA, NETHERLANDS, NEW ZEALAND, NORWAY, OMAN, POLAND, PORTUGAL, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, Slovakia, SPAIN, SWEDEN, SWITZERLAND, TURKEY, UNITED ARAB EMIRATES, and UNITED KINGDOM | Terminated | June 05, 2012 | 207 | 523 | Computer
Z-0457-2012 | 60571 | Yes | LWS |  |  | Cardiovascular | PMA | St. Jude Medical Riata (8Fr), Models: 1560, 1561, 1562, 1570, 1571, 1572, 1580, 1581, 1582, 1590, 1591, 1592 |  |  | November 28, 2011 | 2011 | St Jude Medical CRMD | St. Jude Medical has confirmed the failures associated with all cause insulation failure on the St. Jude Medical Riata and Riata ST Silicone Endocardial Defibrillation Leads with specific emphasis on externalized conductors. | Device Design | St. Jude Medical sent a Medical Device Advisory Important Product Information Update letter dated November 28, 2011, Physician Communication dated December 15, 2010, Product Performance Report Extracts, and a Physician Device Advisory Notification to physicians by FedEx priority overnight service for receipt on November 29, 2011. The letters provides the customers with an explanation of the problem, the product, and the actions to be taken. The issuance of the recall letter is in conjunction with St. Jude's recent release November 2011 Product Performance Report (available online at http://sjmprofessional.com). The letter also explains about the Lead Peformance Summary, Root Cause, Clinical Implications, Rate of Occurrence from Complaints and Returns, New Peer Reviewed Publications, Data on Optim Insulated Leads (Durata and Riata ST Optim), and Recommendations and Mitigations. Physicians with any questions or concerns can contact their local St. Jude Medical representative or the Technical Services Department at (800) 722-3774. |  | 128000 | 128,000 units total | USA (Nationwide) and worldwide to: Austria, Belgium, Bulgaria, Canada, Czech Republic, Denmark, Estonia, Finland, France, Germany, Hungary, Iceland, Italy, Romania, Russia, Netherlands, Norway, Poland, Portugal, Russia, Serbia, Slovakia, Spain, Sweden, Switzerland, The Netherlands, UK, Australia, China, Hong Kong, Indonesia, Japan, Korea, Malaysia, Philippines, Singapore, Taiwan, Thailand, Vietnam, New Zealand, Argentina, Brazil, Colombia, Costa Rica, Ecuador, Mexico, Puerto Rico, Uruguay, Egypt, Greece, India, Iran, Israel, Joran, Kuwait, Lebanon, Oman, Qatar, Saudi Arabia, South Africa, Turkey, Malta, Cyprus, Iraq, Egypt, Lybia, Yemen, and U.A.E. | Terminated | May 09, 2012 | 163 | 128000 | Computer
Z-0607-2012 | 60660 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | SKYLight Imaging Systems, Model 882050; |  |  | December 08, 2011 | 2011 | Philips Medical Systems | During preventative maintenance, a crack was discovered in the lower portion of the old style, non-common arm for the detector assembly, which could result in the detector dropping in an uncontrolled manner. | Device Design | Urgent Medical Device Correction/Field Safety Notices, dated Dec 08, 2011, were mailed to affected customer sites on December 9th and 12th, 2011 by Return Receipt certified mail. The letter identified the affected product, described the issue and when it could occur, along with the hazards involved. Customers were advised to stop using the system and contact a Philips service representative for an immediate inspection. Customers were also to follow the Instructions For Use to avoid the issue. Philips service will contact customers to arrange a time for Philips to inspect the system. The letter should be placed in the Instructions for Use. The Field Change Order will be released on December 16, 2011. If further information or support is needed, customers should contact Customer Care Solutions Center at 1-800-722-9377. |  | 129 | 129 | Worldwide Distribution -- USA, Canada, Italy, Spain, France, United Kingdom, Sweden, Switzerland, Germany, Japan, Denmark, and the Netherlands. | Terminated | August 13, 2012 | 249 | 129 | Computer
Z-0466-2012 | 60459 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | SoftMic GUI versions: 4.0.3.0-4.0.3.15, 4.0.4.1-4.0.4.7. |  |  | August 25, 2011 | 2011 | SCC Soft Computer | On 08/25/2011, SCC Soft Computer initiated a correction on the following SoftMic GUI versions: 4.0.3.0-4.0.3.15, 4.0.4.1-4.0.4.7. For clients using SoftMic Result Entry, when a certain series of keystrokes on the test grid are performed while entering results and statusing a test, the test status may not be maintained upon saving. The status may be removed or copied from another test on the order | Software design | SCC Soft Computer initiated a Field Correction dated August 25, 2011. to all affected customers.Customers with the affected software have been notified and are scheduling receipt of the hot fix..Further questions call (727) 789-0100. |  | 144 | 144 | Worldwide Distribution -- USA (nationwide) and Canada. | Terminated | October 03, 2014 | 1135 | 144 | Computer
Z-0838-2011 | 57271 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | AngioScore, Inc. AngioSculpt PTA Scoring Balloon Catheter; |  |  | November 16, 2010 | 2010 | AngioScore Inc. | Reports of distal bond failure which has resulted in peeling of the bond and/or detachment of the distal end of the scoring element. May result in vessel injury or embolism. | Device Design | Medical Device Recall notifications were sent to all domestic consignees on Nov 16, 2010 by overnight mail. Foreign distributors, other than Japan, were sent recall letters by e-mail or international overnight delivery on November 16, 2010. The Japanese distributor received notification and notified their customers on November 10, 2010. |  | 17682 | 17682 total units, all sizes and lots | Worldwide Distribution: USA, China, Finland, Italy, Japan, Norway, Saudi Arabia, Spain, Sweden, Switzerland, UAE, United Kingdom | Terminated | October 04, 2011 | 322 | 17682 | Computer
Z-0922-2011 | 56356 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Stryker brand Cart with Isolation Transformer, |  |  | April 16, 2010 | 2010 | Stryker Endoscopy | The electrical cord power plug design may allow the prongs to develop a fracture inside or outside the plug, if the plug is misused. | Device Design | Stryker Endoscopy sent an URGENT: Device Recall letter dated April 16, 2010, to Stryker Sales reps and all customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to:Locate the affected units.Visually inspect each cord for damage.If the cord has bent or cracked prongs the unit should be taken out of service immediately. The customer was instructed to call a Regulatory Affairs Recall Representative at (408) 754-2738 for further instruction. If the cord did not show signs of damage, it could remain in service.An acknowledgement form was included to document confirmation of receipt of the letter. Upon completion of the power cord inspection, customers were to sign the acknowledgement form indicating that all of their units were inspected. Customers were asked to mail the form to Stryker Endoscopy or fax to (408) 754-8378 or scan the card and email to PowerCord@stryker.com.For questions regarding this recall call the Stryker RA/QA Department at (408) 754-2738 or (800) 435-0220, Reference RA 2010-034. |  | 7371 | 7371 | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, ID, IL, IN, IA, KS, KY, LA, MA, ME, MI, MN, MS, MO, MT, NE, NH, NJ, NM, NV, MY, NC, ND, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WI, WV, and WY and the countries of Canada, Chile, China, India, Korea, Mexico, Singapore, Taiwan and Latin America. | Terminated | January 19, 2011 | 278 | 7371 | Computer
Z-0462-2012 | 60461 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | syngo RT Therapist and syngo RT Therapist Connect/MOSAIQ OIS. |  |  | November 02, 2011 | 2011 | Siemens Medical Solutions USA, Inc | The software release will bring current Installed Base system to the forward production version. It is a result of the final system testing addressing several safety issues that are not critical for the safe use of the device and multiple performance defects. | Software design | Siemens sent Customer Information letters and Update Instructions TH024/11/S for the installation of software release 4.2.108 beginning on November 2, 2011 to all affected consignees. Sites with active devices for the Oncology system were sent the current Update Instructions TH014/11/S for the installation of software release 4.2.108 and Customer Information letter beginning November 22, 2011.Consignees were asked to include the Customer Information in their syngo RT Oncologist System Owner Manual chapter "Safety Advisory Letters" where it should remain. For questions regarding this recall call 925-602-8157. |  | 147 | 85 affected customers RT Therapa[ist, 62 customers Oncology | Worlwide Distribution - USA (nationwide) including UT, OH, TN, FL, WI, IN, KY, CA, PA, SD, NY and the countries of India, Germany, United Kingdom, Australia, Poland, Japan, Croatia, India, Columbia, Belgium, Puerto Rico, China, Sweden, South Africa and Czech Republic. | Terminated | March 21, 2012 | 140 | 147 | Computer
Z-0029-2012 | 59787 | No | FCG | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | OLYMPUS ViziShot EBUS Aspiration Needle |  |  | August 25, 2011 | 2011 | Olympus America Inc. | Firm has received complaints from users of the ability to wipe material from the needle's stylet and the presence of material in patient cytology samples obtained by the needle following endobronchial ultrasound fine needle aspiration of a lesion. Olympus Medical Systems Corporation (OMSC) in Japan has confirmed that metal particles can occur on the stylet due to friction between the stylet and n | Device Design | OLYMPUS sent an URGENT MEDICAL DEVICE CORRECTION letter dated August 26, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. A revised Instruction Manual with the new WARNINGS will be sent to the customer's attention.Customers were instructed to examine their inventory to determine if they have the affected product.To share the information on the new WARNINGS with all of the appropriate staff at their facility and any other organization where the product has been transferred.Complete and return the enclosed questionnaire even if they do not have affected product. Customers were instructed to fax the completed questionnaire to 484-896-7128.For any question call 484-896-5688. |  | 22498 | 22,498 | Nationwide Distribution including the District of Columbia and the following States: AL, AK, AZ, AR, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, NE, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, and WY. | Terminated | September 28, 2012 | 400 | 22498 | Computer
Z-0488-2012 | 60452 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | SoftLab GUI versions: 3.18.1.0-3.18.1.21; 3.19.1.0-3.19.1.25; 3.19.2.0-3.19.2.21; 4.0.0.2; 4.0.1.0-4.0.1.16; 4.0.2.0-4.0.2.10; 4.0.3.0-4.0.3.16; 4.0.4.0-4.0.4.6; and 4.5.0.0-4.5.0.8. |  |  | May 03, 2011 | 2011 | SCC Soft Computer | On 05/03/2011, SCC Soft Computer, Clearwater, FL, initiated a correction on the following SoftLab GUI versions: 3.18.1.0-3.18.1.21; 3.19.1.0-3.19.1.25; 3.19.2.0-3.19.2.21; 4.0.0.2; 4.0.1.0-4.0.1.16; 4.0.2.0-4.0.2.10; 4.0.3.0-4.0.3.16; 4.0.4.0-4.0.4.6; and 4.5.0.0-4.5.0.8. A client reported that results were not flagged correctly. This can occur when a test is ordered in Order Entry, then collected | Software design | SCC Soft Computer sent a "CORRECTION COMMUNICATION" notification letter to all affected customers. The letter describes the product, Risk-to-Health notification task, potential hazard, and corrective action. The customers were asked to acknowledge receipt of the task and grant permission for SCC to make the required and recommended changes. For any questions, call (727) 789-0100. |  | 220 | 220 | Worldwide Distribution - USA (nationwide) and the country of: Canada. | Terminated | August 07, 2013 | 827 | 220 | Computer
Z-0928-2011 | 56931 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Scorpio Fixed Femoral A/P Sizing Guide; non-sterile instruments; Howmedica Osteonics Corp., 325 Corporate Drive, Mahwah, NJ 07430; A subsidiary of Stryker Corporation |  |  | September 27, 2010 | 2010 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics has become aware that there is a potential for the shoulder screw on the Scorpio Sizing Guide (the "Sizing Guide") to disassociate from the main body of the instrument. | Device Design | Stryker Orthopaedics sent an 'Urgent Update Product Recall' letter dated December 30, 2010, along with Product Acknowledgement forms to all customers (US and international). The letter described the product, problem and action to be taken by the customers. The customers were asked to complete and return the PRODUCT RECALL ACKNOWLEDGMENT FORM within five (5) days via fax at 201-831-6069; contact the hospitals in their territory that have the affected product to arrange return of the product; and return the affected product to the attention of Regulatory Compliance, Stryker Orthopaedics, 325 Corporate Drive, Mahwah, NJ 07430.If you have any questions, feel free to contact the Manager, Divisional Regulatory Compliance at (201) 972-2100 or the Associate Product Manager, Knee Marketing Team, at (201) 831-5028. |  | 6 | six units | Worldwide distribution: USA and countries including: Australia, Brazil, China, Columbia, France, Greece, India, Italy, Korea, Poland, Singapore, South Africa, and United Kingdom. | Terminated | July 31, 2013 | 1038 | 243 | Computer
Z-0370-2012 | 60441 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand Eclipse Release 10 with RT Chart Release 10 (Client Build 10.0.28 or Build 10.0.34), An |  |  | October 31, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | An anomaly has been identified with Varian Eclipse and RT Chart where adding a reference pointin RT Chart and editing the field dose contributions may change the Prescribed Dose Per Fractionin Eclipse. | Device Design | Varian Medical Systems sent an "URGENT MEDICAL DEVICE CORRECTION/URGENT FIELD SAFETY NOTICE" dated October 31, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. Varian has corrected the issue and will contact customers to schedule installation of the corrected version of the system. Contact the Varian Oncology Help Desk at 1-888-827-4265 for questions regarding this notice.. |  |  |  | Worldwide Distribution. | Terminated | June 06, 2014 | 949 |  | Computer
Z-0924-2011 | 56859 | No | JEA | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Mizuho ProFx 6850 Operating Room Table, with orthopedic accessories, AC Powered, manufactured by Mizuho OSI, Union City, CA. |  |  | March 30, 2010 | 2010 | Mizuho Orthopedic Systems Inc | Device component spar gear box locks up and may result in failure of push pull tube, which may result in release of breaking and impact. | Device Design | All customers will be contacted to arrange for removal and shipping arrangements. Verification will be done by the Repair and Verification instructions for the leg spars, via registered letters.*******UPDATED INFO---The firm sent another notification letter---dated 12/22/10---which states in part: "****The upgrade involves the correction of your 6850 PROfxÂż spars (6850-2100) at your facility by a Qualified Technician. The spars shall be upgraded to remove the 6850-2123, Drive Bracket, Spar Push Rod and replace it with the 6850-2170, Drive Bracket, Spar Push Rod. Upon completion of the upgrade, the 6850-2123, Drive Bracket, Spar Push rod shall be returned to Mizuho OSI for processing. ******* |  | 17 | 17 units | Worldwide Distribution -- USA &amp; Germany. | Terminated | May 17, 2011 | 413 | 17 | Computer
Z-0394-2012 | 60440 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | OR table TITAN, Surgical Table. |  |  | January 21, 2011 | 2011 | Trumpf Medical Systems, Inc. | Product is being recalled due to unintended movement on the Titan Tables. | Device Design | Trumpf Medical Systems sent an initial notification to customers dated January 24, 2011. A subsequent "URGENT SAFETY NOTICE-UPDATE" letter dated May 26, 2011 was sent to all affected customers. The letter described the product, problem, and actions to be taken by the customers. A Corrective Action Plan includes a software update to be installed by a Trumpf Technical Service Representative. Contact Technical Service at 888-474-9359 for questions regarding this notice. |  | 94 | 94 beds total for all types. | Worldwide Distribution. | Terminated | August 13, 2012 | 570 | 98 | Computer
Z-0051-2012 | 59848 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Outlook ES pump is intended for use with B. Braun Medical Inc. Horizon Pump IV Sets. Used to regulate the flow of primary and secondary fluids when positive pressure is required. The infusion system is capable of delivering fluid from a negative head height (when the IV fluid container is lower than the pump), and provides clinically accepted volumetric accuracy for all standard IV fluids, including blood, lipid, and Total Parenteral Nutrition (TPN). Model number: 621-300ES, 621-400ES, 621-400ESR. |  |  | September 07, 2011 | 2011 | B Braun Medical, Inc | B Braun Outlook ES Safety Infusion System "Run" Light Emitting Diodes (LEDs) pumps stop infusing and a backup alarm sounds, but the "Run" LEDs advance as if the pumps were infusing. | Software design | B. Braun Medical, Inc. sent an "URGENT: FIELD CORRECTION" letter dated September 7, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. A firm representative will contact all affected customers in order to schedule the field correction. The correction consists of revising the Door and Main Processor software, inspecting the main board, and replacing the board, if warranted. A Field Correction Information Form was included in the letter for customers to complete and return. Contact Technical Product Support at 1-800-627-7867 for questions regarding this notice. |  | 6321 | 6321 | Nationwide Distribution | Terminated | May 04, 2015 | 1335 | 6321 | Computer
Z-0867-2011 | 54651 | No | LMD | Medical image communications device. | Radiology | Radiology | 510(K) Exempt | Stryker Digital Capture Pro 2 device (240-050-800, CD version) and (240-050-810, DVD version) designed for digital documentation of medical procedures. Manufactured by Stryker Endoscopy, San Jose, CA. |  |  | December 11, 2009 | 2009 | Stryker Endoscopy | Device will not able to assign the correct time/date stamp for saved filed after December 31, 2009. On Jan 1, 2010, the device system date for all SDC Pro 2 units will re-start from 1 Oct 2001 and begin incrementing from that point onwards. The system date will be incorrect and users will not be able to correct this issue. Problem will not affect how pictures or videos are captured using SDC Pro | Software design | Stryker Endocoscopy sent an URGENT: Device Correction letter dated December 11, 2009, to the OR Supervisor/Risk Manager and all Stryker Sales Representatives. The letter identified the product, the problem, and the action to be taken by the Sales Representatives. Sales Representatives were to identify affected accounts. Go to customer accounts and install the upgrade for each affected unit. Upon completion of the upgrade, the sales rep was to sign the acknowledgment form indicating that all of their affected accounts received the appropriate software fix and mail to Stryker Endoscopy or fax both sides of the postcard to 408-754-8378 or scan in both sides of the postcard and email to sdcpro2recall@stryker.com.For questions regarding this recall call 408-754-2738 or 408-754-2000. |  | 3070 | 3070 units | Worldwide Distribution - USA including AL, AK, AR, AZ, CA, CO, CT, DC, DE, FL, GA, ID, IL IN, IA, KS, KY, LA, MA, MD, MA, MI, MN, MS, MO, MT, NE, NH, NJ, NM, NV, NY, NC, ND, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WI, WV, and WY and the countries of Africa, Australia, Brazil, Cnanda, China, France, Germany, Hong Kong, Iberia, Italy, Japan, Korea, Latin America, Mexico, Netherlands, Poland, Puerto Rico, Scandinavia, Singapore, Switzerland, Thailand, and the UK | Terminated | January 08, 2011 | 393 | 3070 | Computer
Z-1257-2011 | 57499 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Toshiba TSX-301A AQUILION ONE (MPDCT0297EA) - Whole Body X-Ray CT Scanner |  |  | September 07, 2010 | 2010 | Toshiba American Medical Systems Inc | Incorrect Dose Calculation. The software that performs the pre-scan dose calculation uses an incorrect parameter for an exam Plan using a tube voltage of 80kV or 100kV. This results in a pre-scan expected dose that is approximately 10% lower than what the displayed value should be when the scanner executes the exam Plan. | Software design | Toshiba America Medical Systems (TAMS) issued notification to consignees 9/7/2010, via USPS return receipt mail. The letter contains a return reply form that is to be faxed to TAMS as proof of notification. |  | 396 | 44 units distributed between 09/2007 - 07/2010 | Nationwide distribution. | Terminated | April 20, 2012 | 591 | 396 | Computer
Z-0033-2012 | 59879 | No | FJP | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Sodium Bicarbonate Mixing and Delivery System Software v 7.1. Model # MD-380, Serial Numbers: 5103031102, 5101101101, 5102161007, 5103221007, 5104221007, 5105101008, 5109131010, 5111231011, 5114301012, 5102131101 &amp; 5113091012. Manufactured by: Isopure Corp, Simpsonville, KY |  |  | May 19, 2011 | 2011 | Isopure Corp | Software problem causes transfer line between mixing and storage tank to not be rinsed after disinfection, leaving residual chlorine in the line. | Software design | Isopure sent a letter dated May 19, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Isopure devised an immediate solution which must be implemented at their faclity in conjunction with the disinfection procedure. Customers were asked to inform all appropriate personnel, clinical and technical, who may be involved in operating the system during the normal disinfection schedule.On 5/19/2011, the firm sent letters to their customers via US mail and email. They also telephoned their customers to notify them of the problem. |  | 11 | 11 units | Nationwide Distribution including AR, CT, GA, MA, ME, NJ, PA, TN and WV. | Terminated | February 19, 2013 | 642 | 11 | Computer
Z-1250-2011 | 57159 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Xeleris Processing and Review Workstation 1.0, 1.1, 2.0, 3.0 |  |  | January 12, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware of safety issues relevant for the Xeleris Workstation 1.0, 1.1, 2.0, 2.1, 3.0:1. Measurements taken on the CT Images on the CT system are different from the Xeleris Workstation2. Inaccurate PET SUV values may be calculated when data is retrieved from PACS after being processed in the Xeleris.3. Wrong marking of Early and Late result series on display in Bra | Software design | The firm, GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated January 12, 2011 to all consignees/customers. The letter described the product, problem and actions to be taken by the customers. The customers were instructed to please ensure that all potential users in their facility are made aware of this safety notification and the recommended actions. A GE Healthcare Service representative will perform the required software update on each affected system.If you have any questions or concerns regarding this notification, contact your local GE Healthcare Service Representative. You may also contact GE Healthcare call center in US or Canada at 800-437-1171 and other countries at 262-896-2890. |  | 6046 | 6046 | Worldwide distribution: USA including DC, PR, and Guam; and countries including: AUSTRIA, AUSTRALIA, ARGENTINA, ALGERIA, AZERBAIJAN, BELGIUM, BOSNIA &amp; HERZEGOVIA, BULGARIA, BRAZIL, CANADA, CHILE, CHINA, COLOMBIA, CYPRUS, CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, ECUADOR, FINLAND, FRANCE, GERMANY, GREECE, GUADELOUPE, HONG KONG, HUNGARY, INDIA, INDONESIA, IRELAND, ISRAEL, ITALY, JAMAICA, JAPAN, KENYA, KUWAIT, LUXEMBOURG, LEBANON, MALY ASIA, MARTINIQUE, MEXICO, MOROCCO, NETHERLANDS, NEW ZEALAND, NORWAY, PAKISTAN, PANAMA, PHILIPPINES, PORTUGAL POLAND, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SLOVAK IA, SPAIN, SOUTH AFRICA, SOUTH KOREA, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TURKEY, UKRAINE, UNITED KINGDOM, URUGUAY, VENEZUELA, and VIET NAM. | Terminated | October 16, 2012 | 643 | 6046 | Computer
Z-0387-2012 | 59812 | No | JOW | Compressible limb sleeve. | Cardiovascular | Cardiovascular | 510(k) | ActiveCare DVT |  |  | August 28, 2011 | 2011 | Medical Compression Systems | This voluntary field action is being undertaken because MCS has received a limited number of reports involving the use of a non-MCS AC/DC adaptor with the ActiveCare DVT and ActiveCare+S.F.T. devices. The firm found that an ActiveCare DVT device may work improperly if connected to a non-MCS AC/DC adaptor with a higher nominal voltage. As specified in each product's User Manual, the ActiveCare DVT | Device Design | Medical Compression Systems (DBN) Ltd, sent out "VOLUNTARY FIELD ACTION" letters dated September 15, 2011 to all affected customers. The letter included description of the device and problem. Customers were instructed to check their inventory and make a list of serial numbers of all applicable ActiveCare+S.F.T. and ActiveCare DVT devices in their inventory, include the list on the attached excel file (Field Action Verification Form) and email to maya@mcsmed.com; even if they do not have any such devices. The letter also states that within 15 days from the email, they will receive instructions for the next steps. In addition, the letter stated -YOU MAY CONTINUE TO USE THE ActiveCare+S.F.T.Âż DEVICES WHILE AWAITINGTHE INSTALLATION OF THE OVP.In case you have questions or you need any assistance, you can contact the firm at (317) 485-7887 or e-mail to: maya@mcsmed.com or mahlet@mcsmed.com. |  | 2217 | 2217 units | Nationwide distribution. | Terminated | July 11, 2012 | 318 | 2768 | Computer
Z-0843-2011 | 57293 | Yes | FJK | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius Medical CombiSet BVM Hemodialysis Blood Tubing Set |  |  | November 19, 2010 | 2010 | Fresenius Medical Care Holdings, Inc. | Hemodialysis Blood tubing Sets may develop kinking of the arterial line. | Device Design | Fresenius Medical issued recall notification dated 11/19/10, "URGENT RECALL FMCNA CombisetÂż BVM Hemodialysis Blood Tubing Sets (Part Numbers 03-2695-9 and 03-27957)", by certified mail with return receipt. Customers have been instructed to check their stock immediately to determine if they have any affected lots on hand. If customers have the affected lots, they have been instructed to discontinue use immediately and place the product in a secure segregated area for return to Fresenius Medical Care North America. Customers have been instructed to contact Fresenius Medical Care Customer Service Team Representative for instructions on how to return the recalled product. |  | 190080 | 190,080 units (7920 cases x 24 units) | Distribution Nationwide and Canada | Terminated | November 03, 2012 | 715 | 190090 | Computer
Z-0890-2011 | 56326 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | CyberKnife Treatment Planning System, a subsystem of the CyberKnife Robotic Radiosurgery System radiation therapy device, manufactured by Accuray Inc., Sunnyvale, CA. |  |  | July 06, 2010 | 2010 | Accuray Inc | Latches responsible for securing the cover to the device may come lose if not properly secured and the cover may unexpectedly drop off. A design change is planned. | Device Design | Accuray sent an "URGENT DEVICE CORRECTION" letter dated July 6, 2010, to all customers. The letter described the product, problem and actions to be taken by customers. The customers were instructed to assure the LINAC cover is secure following reattachment and adhere to and follow the instructions included in the letter. Note: A new design is being developed, and will be distributed to user base once available.If you have any questions or concerns regarding this issue, please do not hesitate to contact Accuray Customer Support at +1-877-668-8667 (USA) or +1-408-716-4700 (non-USA) or customersupport@accuray.com. |  | 207 | 207 units | Worldwide distribution: US including: AK, AL, AR, AZ, CA, CO, CT, FL, GA, IL, IN, KS, LA, MA, MD, MI, MN, MO, NC, ND, NH, NJ, NM, NV, NY, OH, OR, PA, SC, TN, TX, UT, VA, WA, WI and WV, and countries including: Japan, China, Vietnam, Taiwan/China, South Korea, Hong Kong/China, Thailand and Malaysia. | Terminated | October 04, 2011 | 455 | 207 | Computer
Z-0891-2011 | 54146 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | 4D Integrated Treatment Console, versions: 8.0.12, 8.0.23 SP2, 8.1 SP2, 8.1.1.3, 8.1.1.4, 8.1.2, 8.1.3, 8.1.4, 8.3, 8.6, 8.6.1, 8.6.2, 8.8.0; Model Number: H51; Distributed and manufactured by: Varian Medical Systems Inc., Palo Alto, CA |  |  | October 16, 2009 | 2009 | Varian Medical Systems Oncology Systems | Software anomaly may result in a design deficiency causing Plan Edit Settings to malfunction under specific circumstances, where the user selects to use a Cone Beam CT set up field. | Device Design | The firm, Varian Medical Systems, sent a Product Notification Letter/Field Safety Letter entitled "URGENT MEDICAL DEVICE CORRECTION" dated October 16, 2009, to all affected customers. The letter described the product, problem, actions to be taken by the customers and Varian. The customers were instructed to use caution when creating CBCT imaging field and limit use of the 'select all' button when acquiring treatment parameters; review departmental policies and procedures regarding the acquisition of any gantry angle at the 4D ITC, and advise the appropriate personnel of this letter. The firm, Varian, recommended that only vertical, longitudinal and lateral couch parameters are acquired when creating a CBCT field at the 4D ITC; notified all affected customers of this document, and will correct this issue in a future release of the 4D ITC software. If owners of $D ITC, version 8.0.12 and higher, have any questions regarding this product notification letter, please contact Varian Medical Systems Help Desk at 1-888-VARIANS (888-827-4265) for USA and Canada, +41 41 749 8844 or support-emea@varian.com for Europe;Emails for North American: support-americas@varian.com; Australia/New Zealand: support-anz@varian.com; South East Asia: seasia.apps.helpdesk@varian.com; China/Asia: china.apps.helpdesk@varian.com; Japan: Japan.Apps.Helpdesk@varian.com and Latin American: soporte.al@varian.com, and Internets for Oncology Systems- www.myvarian.com and Varian Medical Systems- www.varian.com. |  | 3629 | 3629 | Worldwide distribution: USA and countries including: Australia, Asia, China, Japan, Latin American, New Zealand , North America, South East Asia. | Terminated | July 20, 2011 | 642 | 3629 | Computer
Z-0591-2012 | 60625 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Agfa Computed Radiography Systems with DX-G Digitizer. |  |  | December 06, 2011 | 2011 | AGFA Corp. | Occasionally images can be lost or sent to an incorrect folder or the digitizer may stop either during startup or during a cassette scan cycle. | Software design | Agfa HealthCare sent a 'URGENT SAFETY NOTICE" letter dated December 6, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Agfa Service will contact your facility and schedule a software upgrade if this has not already been performed. Please complete the attached feedback form as soon as possible and fax it to 864-421-1662.If you have any questions about this matter, please contact your local Agfa HealthCare Service at 877-777-2432. |  | 125 | 125 | Worldwide Distribution-- USA (nationwide) including the states of AZ, CA, CO, FL, GA, HI, IL, IN, KS, LA, ME, MI, MN, MO, NE, NH, NJ, NY, OH, OK, PA, SD, TX, VA, WA, WI and WV, and Canada. | Terminated | April 11, 2012 | 127 | 125 | Computer
Z-0039-2012 | 59316 | microtome, rotary | IDO | Tissue processing equipment. | Pathology | Pathology | 510(K) Exempt | Thermo Scientific Microm: Rotary Microtome: REF HM355S-3 |  |  | August 12, 2011 | 2011 | Richard-Allan Scientific Company | The firm has received two complaints of finger lacerations in operators of this device. | Device Design | ThermoFisher Scientific sent "URGENT MEDICAL DEVICE CORRECTION, RE: Thermo Scientific HM355S-3 User Safety Advisory Notice" to all affected customers.The letter states the manufacturer identified a need to enhance the safety instructions and the description of the safety devices in the user manual. It mentions the HM3555S-3 User Manual has been modified to include more specific safe work practices and a new Quick -Reference Card summarizing the manual's safety enhancements and that the changes have been included on a CD that is enclosed with the letter. The firm requests confirmation of receipt of the new User Manual and Quick-Reference Card by October 31, 2011 to be sent by FAX to (269) 372-2674For questions on this recall call the firms Technical Support representatives at 1-800-522-7270 08:00 AM- 4:30 PM. |  | 786 | 786 | Worldwide Distribution: Nationwide distribution; including the country of Canada. | Terminated | July 25, 2013 | 713 | 786 | Computer
Z-0603-2012 | 60390 | No | JLW | Thyroid stimulating hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | The GSP 2021-0010 (Genetic Screening Processor) instrument is a fully automated, high through put batch analyzer for time- resolved and prompt fluorescence analysis of samples in microtitration plates. It is intended for in vitro quantitative / qualitative determination of analytes in body fluids. The chemistries running on the GSP instrument are used to test blood specimens dried on filter paper as an aid in screening newborns for congenital disorders. |  |  | October 24, 2011 | 2011 | PerkinElmer LAS, Inc | Malfunction of the GSP disk remover unit may cause a punched disc to remain in the well during the measurement of the plate. This may cause a false negative screening result being reported for 17-OHP, resulting in failure to identify cases of congenital adrenal hyperplasia (CAH). | Device Design | Recall (correction) letters were sent to all customers on October 24, 2011. An improved disk remover design is under development to eliminate the reported problem. Until the final field correction is available, customers should follow the precautions provided. The needed protocol update will be performed by the local service engineer. If customers observe an increase in the occurrence of the problem, the laboratory should contact their local PerkinElmer representative. |  | 52 | 52 | Worldwide Distribution - USA, including the states of: Missouri, Florida, Texas, Oregon, Alabama, Louisiana, New Jersey, Georgia, Virginia, Kentucky and the countries of: Australia, Belgium, Brazil, Egypt, France, Germany, Italy, Netherlands, Norway, Qatar, Saudi MOH, Switzerland, and the UK. | Terminated | July 20, 2012 | 270 | 52 | Computer
Z-0853-2011 | 53047 | No | JAI | Powered radiation therapy patient support assembly. | Radiology | Radiology | 510(k) | RoboCouch Patient Support System, model number 025007, a subsystem of the Cyberknife Robotic Radiosurgery System, manufactured by Accuray Incorporated, Sunnyvale, CA. |  |  | July 21, 2009 | 2009 | Accuray Inc | Mechanical Failure-- Patient Support System's table top may unexpectedly descend several inches. | Device Design | Consignees were notified verbally about the issue when an inspection was scheduled. |  | 26 | 26 | Worldwide Distribution -- USA, including states of CA, FL, VA, AL, CT, MA, NC, TN, PA, RI, and NE and countries of Turkey, France, UK, Saudi Arabia, and Switzerland. | Terminated | January 12, 2011 | 540 | 26 | Computer
Z-0444-2012 | 60455 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | SoftLab used with SA INST 4.0.2.0 to 4.0.2.51, 4.0.3.0 to 4.0.3.13, 4.5.0.0 to 4.5.0.9, 4.5.2.0 to 4.5.2.8 and SA INST 3.1.3.0 to 3.1.3.29, 3.1.6.0 to 3.1.6.17, 4.0.0.0, 4.0.1.0 to 4.0.1.32. |  |  | July 26, 2011 | 2011 | SCC Soft Computer | SCC Soft Computer, Clearwater, FL, initiated a correction for the SoftLab used with SA INST. A test performed on the Roche Cobas Integra series of instruments when used with the GenInst/Dispatch listening program could cause incorrect flagging for a test result posted to the lab database. | Software design | The firm, SCC Soft Computer, sent a "Correction Communication" letter on July 26, 2011 to customers/consignees. The letter describes the product problem and actions to be taken. The firm will correct the software to recognize and support all known formats of the resulting date and time sent by an instrument. The software is available via a recommended patch or hot fix to the client's current release. The customers were asked to acknowledge receipt of the task and grant permission for SCC to correct the affected configuration files. If you have any questions, call (727) 789-0100. |  | 265 | 265 | Worldwide Distribution - USA (Nationwide) and Country of: Canada. | Terminated | May 21, 2013 | 665 | 265 | Computer
Z-0486-2012 | 60357 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Healthcare AW Server version 2.0 with release 2.0 - 4.0 |  |  | November 17, 2011 | 2011 | GE Healthcare, LLC | 2 Unrelated Software Issues: An incorrect patient name maybe displayed when switching between two Volume Viewer tabs with different patients. Also, if a second patient exam is selected while the application is starting, the second patient's data will be displayed with the first patient's name display on the applications tab. | Software design | Consignees were sent on 11/17/11 a GE "Urgent Medical Device Correction" letters dated November, 16 2011 The letter was addressed to Hospital Administrators, Radiology Mangers and Risk Managers. The letter described the Safety Issue, Safety Instructions Affected Product Details, Product Correction and Contact Information. For the first issue, customers were asked to contact their field engineer to install the correct configuration file. Also, customers should not apply the recommended changes from the Customer Release Note as those changes cause the issue to occur. With the second issue, customers are to wait until the application has fully initialized before making another patient exam selection. GE Healthcare will correct all affected AW server installations via a software upgrade at no cost to the customer. A representative will contact customers to arrange for the correction. Questions should be directed to customers' local GE Healthcare Service Representative or at the numbers provided in the letter. |  | 308 | 308 | Worldwide Distribution -- USA, including states of AZ, CA, CO, CT, DE, DC, FL, HI, IL, IN, LA, MD, MI, MN, MS,MO, NJ, NM, NY, NC, OH, OK, PA, SC, SD, TN, TX, UT, VA, WA, and WI and countries of UNITED ARAB EMIRATES, TURKEY, THAILAND, SWITZERLAND, SWEDEN, SPAIN, SOUTH AFRICA, SINGAPORE, SAUDI ARABIA, REUNION, QATAR, POLAND, NORWAY, NEW ZEALAND, LEBANON, SOUTH KOREA, JAPAN, ITALY, ISRAEL, HONG KONG, GUADELOUPE, GREAT BRITAIN, GERMANY, FRANCE, FINLAND, ESTONIA, EGYPT, DENMARK, CANADA, BRAZIL, BOSNIA AND HERZEGOVINA, BELGIUM, and AUSTRALIA. | Terminated | July 03, 2012 | 229 | 308 | Computer
Z-0381-2012 | 60214 | No | FST | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | UltraLite Headlight module-grey |  |  | October 10, 2011 | 2011 | Integra LifeSciences Corp. | Headlight assembly (module) may become loose and separate from the headband assembly | Device Design | Integra Life Sciences notified consignees by URGENT DEVICE RECALL letter dated October 10, 2011, via FedEx, email or telephone conference calls. The letter identified the affected product, the problem, and the actions to be taken. Customers were instructed to identify and return any affected product. Customers were requested to complete and return the attached Field Corrective Action Acknowledgement Form to Integra Surgical indicating receipt and review of this notification. For assistance and additional information contact your local and Integra Sales Representative or call Integra Surgical Customer Service at 1-800-431-1123. |  | 60 | 60 units | Worldwide Distribution - USA (nationwide) including Puerto Rico, and the countries of: Africa, Algeria, Australia, Austria, Brazil, Canada, Check Republic, China, Chile, Costa Rica, Denmark, Dominican Republic, Ecuador, Ethiopia, Finland, France, Georgia, Germany, Greece Hungary, Guatemala, India, Ireland, Israel, Italy, Iraq, Japan, Jordan, Kenya, Kuwait, Lebanon, Lithuania, Macedonia, Mexico, Netherlands, New Zealand! Norway, Oman, Pakistan, Peru Philippines, Poland, Portugal, Romania, Russian Federation, Saudi Arabia, Singapore, Slovakia, Spain, Sweden, Switzerland, Taiwan, Thailand, Trinidad, Turkey, Qatar, United Kingdom, Yemeni United Arab Emirates, Uruguay, and Venezuela. | Terminated | October 01, 2012 | 357 | 7216 | Computer
Z-0055-2012 | 59526 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | 8F Safe Sheath Introducer, Catalog number: V18; |  |  | August 01, 2011 | 2011 | Angiodynamics, INC | The product may contain an incorrect size catheter and introducer. | Device Design | AngioDynamics, Inc contacted affected customers initially via telephone and a follow up "Notification of Medical Device Correction" letter was sent dated August 1, 2011. The letter informed them of the problem and provided instructions. Customers were instructed to utilize the 8F Introducer Sheath provided within the 8F Smart Port CT Implantable Port during port implantation. The 8F Valved Sheath (V18) may continue to be utilized during the placement of the 7.5F Smart Port Implantable Port (for example: CT75STSD; CT75STSD-NF). For question on this recall contact AngioDynamics, Inc at (800) 772-6446. |  | 1095 | 1095 total units | (USA) Nationwide distribution including the states of CA, FL, GA, IA, IL, IN, KY, MD, MI, MO, NC, NE, NJ, NY, OH, OK, OR, TN, UT, WA and WI. | Terminated | March 07, 2012 | 219 | 1095 | Computer
Z-0048-2012 | 59768 | No | OOI | Instrumentation for clinical multiplex test systems. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Abbott m2000rt System software; an automated system for performing fluorescence-based PCR that results in quantitative and qualitative detection of nucleic acid sequences; Abbott Molecular Inc., Des Plaines, IL 60018 USA; catalog #09K15-01 |  |  | September 08, 2011 | 2011 | Abbott Molecular | The Abbott m2000rt System software has software bug which can affect customers who manually enter a Test Order onto the m2000rt and sorts a column on the calibration or control grids of the Test Order screen by double clicking the column header. This causes a mismatch between specific concentration values and the designated calibrators and controls in the final results, which could cause an inval | Software design | Abbott Molecular Inc. sent a Field Correction Recall letter dated September 8, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The customer was instructed to follow the instructions identified until an upgrade to your m2000rt software can be installed. Abbott Molecular estimates that this software version will be available the first quarter of 2012.Please complete the enclosed reply sheet and fax it to (847) 775-6728 at your earliest convenience to acknowledge receipt. We appreciate your assistance in this matter. |  | 926 | 926 units | Worldwide Distribution -- USA (nationwide) including the countries of Algeria, Argentina, Armenia, Australia, Austria, Azerbaijan, Bahrain, Belgium, Benin, Bolivia, Bosnia and Herzegovia, Brazil, Bulgaria, Burkina Faso, Cameroon, Canada, Chad, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Dem. Rep. of the Congo, Denmark, Egypt, Estonia, Ethiopia, Finland, France, Gabon, Gaza &amp; Jericho, Germany, Greece, Guatemala, Guinea, Haiti, Honduras, Hong Kong, India, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Korea, Kuwait, Libya, Lithuania, Malawi, Malaysia, Mali, Mauritius, Mexico, Moldova, Morocco, Mozambique, Netherlands, New Zealand, Nigeria, Oman, Pakistan, Panama, Philippines, Poland, Portugal, Republic of Yemen, Romania, Russia, Saudi Arabia, Senegal, Serbia, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Tajikistan, Tanzania, Thailand, Togo, Tunisia, Turkey, United Arab Emirates, Uganda, Ukraine, United Kingdom, Venezuela, Vietnam, Zambia and Zimbabwe | Terminated | March 08, 2016 | 1643 | 926 | Computer
Z-0428-2012 | 60516 | No | LGZ | Infusion pump. | General Hospital | General Hospital | 510(k) | Level 1Âż Normothermic I.V. Fluid Administration Sets, |  |  | November 21, 2011 | 2011 | Smiths Medical ASD, Inc. | Disconnected/ loose Luer connectors at the patient end of the triple lumen tubing of the HOTLlNE Disposable Administration Sets and Level 1 D/DI-60HL Normothermic IV Administration Sets | Device Design | Smiths Medical notified consignees by letter dated November 28, 2011 titled: Urgent Field Safety Notice. The letter identified the affected product, the problem, and the actions to be taken. Customers were instructed to inspect inventory and segregate unused affected product. Distributors were instructed to immediately stop distributing and quarantine all inventory. The letters included a Confirmation Form to be complete and return, by fax to 781-610-9859 or by email to hotlineconnectorrecall@smiths-medical.com. For questions consignees were instructed to contact Smiths Medical at 1781-763-9330. |  | 7631 | 7631 units | Worldwide Distribution  (USA) (nationwide) - Argentina, Austria, Australia, Belgium, Bahrain, Bermuda, Canada, Chile, China, Columbia, Czech Republic, Germany, Denmark, Estonia, Egypt, Finland, France, Great Britain, Guernsey, Greece, Guatemala, Hong Kong, Hungary, Indonesia, Ireland, Israel India, Iceland, Italy, Japan, South Korea, Kazakhstan, Lebanon, Luxembourg, Mexico, Malaysia, New Caledonia, Netherlands, Norway, New Zealand, Panama, Poland, Portugal, Qatar, Singapore, Sweden, Slovak Republic, South Africa, Spain, Sri Lanka, Switzerland, Taiwan, and Vatican City State. | Terminated | August 07, 2014 | 990 | 805506 | Computer
Z-0518-2012 | 60381 | No | BSF | Carbon dioxide absorber. | Anesthesiology | Anesthesiology | 510(K) Exempt | The IS Can Disposable CO2 absorber x 6 --- |  |  | February 09, 2011 | 2011 | Intersurgical Inc | Intersurgical Incorporated has issued a voluntary recall of the IS Cans Carbon Dioxide Absorber due to a possibility that loose granules may enter the inspiratory path of the breathing system. | Device Design | Intersurgical Inc. sent an "URGENT MEDICAL DEVICE RECALL" letter dated February 8, 2011 to all affected customers. The letter described the product, problem, and actions to be taken by the customers. The letter instructs customers to immediately discontinue use and quarantine any affected product found. A Recall Response Form was attached to the letter for customers to complete and return via fax at 315-451-3696, Attn: Customer Service. Contact Customer Service at 1-800-828-9633 if you have questions concerning this recall. |  | 110 | 110 units | Worldwide Distribution- USA (nationwide) including the states of Florida, North Carolina, New York, Ohio, and the country of Canada. | Terminated | January 15, 2013 | 706 | 110 | Computer
Z-0888-2011 | 51217 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Medical Systems Floorstand. An accessory of the Optical Guidance Platform, Model number HZ5 for use with a charged particle accelerator (such as the Clinac 600C). |  |  | January 26, 2009 | 2009 | Varian Medical Systems Oncology Systems | Potential patient impact -- limit switches for Floorstand may be inadequate for some models and allow for potential impact. | Device Design | Varian sent an urgent device correction notification letter 01/26/2009 to all affected users, with a description of the problem and user corrective action steps via certified mail. The letter will also be distributed to Varian Sales, Marketing and Service organizations. |  | 3 | 3 units | Product was distributed to 3 consignees in Florida and Wisconsin. | Terminated | February 24, 2012 | 1124 | 3 | Computer
Z-0896-2011 | 54718 | No | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | Acuson Antares Ultrasound System |  |  | September 30, 2009 | 2009 | Siemens Medical Solutions USA, Inc. | Software safety guard measures on the device malfunctioned to prevent the temperature from going over its limits when the device is used. | Software design | Field notification in the form of a Customer Safety Advisory Notice was initiated on 09/30/2009 by the firm, and the notice was mailed to all affected customersSiemens sent a CUSTOMER SAFETY ADVISORY NOTIFICATION letter on September 30, 2009. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to disconnect and reconnect the V5Ms transducer between uses until their local Customer Service Engineer installs the software update to their system. For any questions customers were to contact their local service support person.For questions regarding this recall call 650-694-5398. |  | 452 | 452 | Worldwide Distribution - United States (nationwide) and the countries of Australia, Austria, Azerbaijan, Brazil, Belarus, Canada, Chile, China, Czech Republic, Denmark, Egypt, France, Germany, Greece, Guatemala, Croatia, India, Indonesia, Italy, Jordan, Japan, Latvia, Malaysia, Mexico, Moldova, Netherlands, New Zealand, Norway, Poland, Portugal, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, South Korea, Spain, Sweden, Taiwan, Turkey, Ukraine, United Kingdom, Uruguay, Uzbekistan, Venezuela, and Vietnam. | Terminated | June 18, 2012 | 992 | 452 | Computer
Z-0037-2012 | 59898 | Yes | LWS |  |  | Cardiovascular | PMA | LATITUDEÂż Paceart Integration Software, version 1.00 Model 6472. |  |  | August 12, 2011 | 2011 | Boston Scientific CRM Corp | Boston Scientific has determined that under certain conditions, LATITUDE Paceart Integration (LPI) softwaremay not properly transfer data resulting in the incorrect summary of programmed parameter values in the ICD Summary tab display and Current Programming Tachy summary section of the Paceart report. | Software design | Consignees were sent on 8/12/11 a Boston Scientific letter dated August 2011. The letter was addressed to "Dear Doctor". The letter described the product and problem. It also gave Background Information, Description of Potential Data Translation Error, and Additional Information and Action information.A letter, including an HRS form, was also sent to all US LATITUDE Clinics that have active LPI software to notify them of the need to install the corrected LPI Software. |  | 398 | 398 | Nationwide Distribution -- AL, AK, AZ, AR, CA, CO, CT DE, DC, FL, GA, IL,IN, IA, KS, KY, LA,ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OJ, OK, OR, PA, RI, SC, SD, TN, TX, UT, VT, VA, WA,WV, WI, and WY. | Terminated | November 06, 2012 | 452 | 398 | Computer
Z-0866-2011 | 57418 | No | GBX | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Merit One Step Centesis Drainage Catheter |  |  | November 17, 2010 | 2010 | Merit Medical Systems, Inc. | The catheter's valved hub may fail resulting in an inability to drain fluid. | Device Design | Merit Medical Systems, Inc. sent a Product Safety Notice letter dated November 18, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer.Customers were instructed to provide a copy of the Safety Notice form to their applicable customers. Advise them to immediately stop using the affected product and to discard it.Customers were asked to complete and sign the Safety Notice forms and fax them to 1-801-208-3378 and send the original Safety Notice Form with both signatures via Federal Express 2nd day air to:Customer ServiceMerit Medical Systems, Inc.1600 West Merit ParkwaySouth Jordan, UT 84095-2416entirety and faxed them For questions regarding this recall call 801-316-4929 or Customer Service at 801-208-4365. |  | 20863 | 20,863 | Worldwide Distribution - USA including AL, AR, AZ, CA, CO, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NH, NJ, NM, NV, NY, OH, OR, PA, RI, SC, TN, TX, UT, VA, VT, WA, WI, AND WV and the countries of Canada, Finland, Germany, Netherlands, Singapore, Sweden, Switzerland, and the UK | Terminated | March 23, 2012 | 492 | 20863 | Computer
Z-0884-2011 | 56986 | No | DSA | Patient transducer and electrode cable (including connector). | Cardiovascular | Cardiovascular | 510(k) | CARTO 3 System Interface Cable, 34-pin to 10-pin, Red, 10' - Part Number: D-1286-18; Catalog Number CR3410CT |  |  | October 05, 2010 | 2010 | Biosense Webster, Inc. | When connecting this cable to the mapping system, loss of Body Surface ECG signals may occur. This may result in inaccurate locations appearing for catheters displayed 3D on the CARTO monitor. | Device Design | Biosense Webster issued an Urgent Field Safety Notice letter dated October 5, 2010. The letter identified the product, the problem,and the action to be taken by the customer. Customers were instructed to return the affected cable, accompanied by the attached Return Authorization form.For questions regarding this recall call Biosense Webster at 909 839-8874. |  | 1692 | 1692 | Worldwide Distribution: USA, including the states of: AK, AL, AR, AZ, CA, CO, CT, FL, GA, IA, IL, KS, KY, MA, MD, MI, MN, NC, NE, NJ, NV, NY, OH, OR, PA, SC, TN, TX, UT, WA, WI, and WV, and the countries of Belgium, China, Israel, Japan, Singapore, and Thailand. | Terminated | May 07, 2012 | 580 | 1692 | Computer
Z-0507-2012 | 56545 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Microsoft Amalga Unified Intelligence System (UIS), This is a software-only device, WHIE Implementation v1.1i. A calculator/data processing module, for clinical use. |  |  | July 27, 2010 | 2010 | Microsoft Corporation Health Solutions Group | The database may incorrectly match patient data resulting in the incorrect patient data being displayed. This could potentially present a risk to patient safety. | Software design | Microsoft notified the sole consignee by FAX on 7/22/2010 with an Urgent Medical Device Safety Information. The letter identified the affected product, the description of the problem, and actions to be taken to correct the problem, along with some additional information for the consignee. In the interim, the firm proposes to block access to patient records identified as mismatched. The firm also stated that the issue has been addressed in version 2.0. The customer should report any incidents to Michael Stokes, Director of Policy Compliance at 208-938-5397 or via e-mail at Michael.Stokes@microsoft.com. |  | 1 | 1 | Nationwide Distribution -- Wisconsin. | Terminated | December 20, 2011 | 511 | 1 | Computer
Z-1341-2011 | 57454 | No | JBP | Activated whole blood clotting time tests. | Hematology | Hematology | 510(k) | Celite ACT Cartridge |  |  | December 03, 2010 | 2010 | Abbott Point Of Care Inc. | Celite ACT cartridges have increased numbers of values outside the accepted allowable error. | Device Design | On behalf of Abbott, Genco Pharmaceutical Services sent an "URGENT RECALL NOTICE" dated November 2010, on December 3, 2010 via Fedex to all consignees/customers. Abbott sent out notification packages to foreign consignees/customers on December 3, 2010. The notice described the product, problem and action to be taken by customers. The firm recommended that customers use cartridges within 12 hours of removal from refrigerated storage based on the instructions provided with the product. The customers were instructed to complete and return the Business Reply cards to acknowledge receipt of the communication and provide the letter to any customers that have recieved the product. If you have any additional questions, please contact Abbott Point of Care Technical Support at 800-366-8020, option 1 or your Abbott Point of Care Representative.On February 27, 2012 Abbott sent an Important Product Information letter via Fed Ex to inform customers that i-Stat Celite ACT cartridges manufactured from the beginning of 2012 can now be stored for two weeks at room temperature as indicated on the cartridge box. Foreign consignees were notified on 2/28/2012. Users are instructed to complete business reply cards to acknowledge receipt of the communication and provide the letter to any customers that have received the product. Any additional questions can be answered by contacting 8003668020 option 1. |  | 1164275 | 1164275 cartridges | Worldwide distribution. | Terminated | July 16, 2013 | 956 | 1164275 | Computer
Z-0594-2012 | 60549 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Pinnacle3 Radiation Therapy Planning System, software version 9.0. |  |  | November 18, 2011 | 2011 | Philips Medical Systems (Cleveland), Inc. | Phillips Medical Systems have recently determined that a software (Pinnacle3 Software version 9.0) nonconformance can cause incorrect beam geometry. | Software design | Consignees were sent on 11/18/11 a Philips "Urgent Medical Device Correction" letter dated 2011 Nov 16. The letter identified the affected product and described the Problem, Hazard Involved, How to Identify Affected Products, Action to be Taken by Customer/User, Actions Planned by Philips and Further information and Support. If customers observed the issue, they were to manually make the suggested change. In addition, customers were to place the letter in the Instructions for Use until otherwise notified. The firm will deliver 9.2 software release free of charge. Customers in North America are to contact Customer Care Solutions Center at 1-800-722-9377, if they have questions or need support concerning the recall. |  | 715 | 715 (463 US) | Worldwide Distribution -- USA, including the states of AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WV, and countries of AUSTRALIA, AUSTRIA, BELGIUM, CHINA, DENMARK, FRANCE, GERMANY, HONG KONG, ITALY, JAPAN, MALAYSIA, MONTENEGRO, NETHERLANDS, NEW ZEALAND, OMAN, SOUTH AFRICA, SOUTH KOREA, SPAIN, SWITZERLAND, TAIWAN, THAILAND, TURKEY, UNITED KINGDOM, and CANADA. | Terminated | February 28, 2014 | 833 | 715 | Computer
Z-0515-2012 | 60424 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo Imaging XS, versions VA70A and VA70B Client/Server Installations |  |  | September 19, 2011 | 2011 | Siemens Medical Solutions USA, Inc | Unintended behavior when using syngo Imaging XS, versions VA70A and VA70B Client/Serve installations. Loading additional images/series from Patient Browser using the thumbnails can cause an unintended deletion of images/series. | Software design | Siemens Medical Solutions USA, Inc sent a Customer Safety Advisory letter dated August 24, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.If you detect patient data inconsistenceis in your system caused by this problem, please contact Siemens Service for support.For further questions please call (610) 219-4834. |  | 12 | 12 | Nationwide Distribution including the states of CA, IN, IA, NV, NE, NY, OH, TN and TX. | Terminated | January 24, 2014 | 858 | 12 | Computer
Z-0516-2012 | 60684 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Portal Vision, R-Arm with 4D Integrated Treatment Console; |  |  | November 23, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | An anomaly was discovered when using the 4DITC in combination with an R-Arm to acquire double exposure MV images. | Software design | The firm, Varian Medical Systems, sent an "URGENT MEDICAL DEVICE CORRECTION/URGENT FIELD SAFETY NOTICE" dated December 1, 2011 to all affected sites via Certified mail, return receipt. Canadian accounts were sent by Fed Ex and internationally to Varian International Customer Support Managers for translation and notification of customers. The notice described the product, problem and actions to be taken. The customers were instructed to always acquire both the treatment port and the open port images BEFORE displaying the digital graticule, when using a standard double exposure template; and to advise the appropriate personnel working in their radiotherapy department of the content of this letter.Special instructions for customers outside the USA and Canada: these customers were instructed to complete and return the attached Proof of Notification or Receipt Verification Card to Varian Medical Systems, once they have read this document. Varian is developing a correction for this issue and a Customer Service rep will contact the customers when the correction is available to schedule its installation on their systems.If you have any questions, contact USA and Canada at 1-888-827-4265; Europe: +41 41 749 8844; Email: North America: support-americas@varian.com; Australia/New Zealand: support-anz@varian.com; Europe: support-emea@varian.com; South East Asia: seasia.apps.helpdesk@varian.com; China/Asia: China.apps.helpdesk@varian.com; Japan: Japan-Apps.Helpdesk@varian.com and Latin America: soporte.al@varian.com or Internet: Oncology systems customer site: www.myvarian.com and Varian Medical Systems public site: www.varian.com. |  | 1486 | 1486 devices | Worldwide distribution: USA (nationwide) and countries of: Algeria, Argentina, Australia, Austria, Azerbaijan, Bahrain, Belarus, Belgium, Brazil, Canada, Chile, China, Colombia, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, Finland, France,, Georgia, Germany,Greece, Hong Kong, Hungary, Iceland, India, Indonesia, Ireland, Israel, Italy, Japan, Kazakhstan, Lativia, Macedonia, Malaysia, Martinique, Mexico, Montenegro, Morocco, Nepal, Netherlands, New Zealand, Norway, Oman, Panama, Philippines, Poland, Portugal, Puerto Rico, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom and Venezuela. | Terminated | August 06, 2014 | 987 | 1486 | Computer
Z-0900-2011 | 57436 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Horizon Cardiology ECG Management |  |  | June 16, 2010 | 2010 | McKesson Provider Technologies - Medical Imaging Group | When importing ECG files from Mortara ECG carts using DICOM interface, the Horizon Cardiology ECG Management System does not display the amplitudes of the waveform properly. | Device Design | McKesson Medical sent an advisory notice letter dated June 16, 2010, to all affected customers. Advisory notice ADV01646 was distributed to 13 affected clinical sites from August 30, 2010 to September 28, 2010. The letters identified the product that is affected, the circumstances under which patient safety issue may occur, the risk, and the immediate actions that need to be taken along with a temporary fix. Customers were instructed to use the ECG cart printout to verify results of procedures acquired by Mortara ECG carts. Users should be aware of the problem when referrring to past ECG procedures and when performing a visual comparison with ECG procedures acquired by Mortara ECG carts.Customers should disable the Glasgow Serial Comparison for ECG procedures.McKesson Support would contact the customer to coordinate the deployment of the software update on all the affected applications and Horizon Cardiology client applications. For any questions regarding this recall, customers should call the Customer Support department at 1-866-777-0202 and speak to their Support Manager. |  | 16 | 16 units | Nationwide Distribution including AZ, CA, FL, GA, MA, PA, UT, and WI. | Terminated | August 17, 2012 | 793 | 16 | Computer
Z-0585-2012 | 60132 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Discovery XR650, Digital Radiographic Systems. |  |  | November 22, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware of a potential issue associated with the patient barrier rotating arm that is a part of the Definium 8000 and Discovery XR650 image pasting systems that may impact patient and user safety. The patient barrier rotating arms that are used to position and stabilize the patient during image pasting procedures may fail to retain residual friction at the arm pivot. Thi | Device Design | The firm, GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated November 17, 2011 to its customers. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. The customers were instructed to ensure that all potential users in their facility are made aware of this safety notification and the recommended actions; discontinue use of the patient barrier and remove from service (away from patients) if either one the rotating arm locking levers is damaged (does not hold the rotating arm in place when engaged); take extra precautions to hold the patients barrier rotating arm during release of the locking lever and repositioning of the arm and store the rotating arms in the lowest pivot position (down) when not in use. The letter also stated that the patient barrier may continue to be used if the rotating arm locking lever securely holds the rotating arm in place when it is engaged.Additionally, GE Healthcare will modify the rotating arms to restore the residual friction so that the arm will retain position (rather than swing freely) while the locking levers are disengaged. Also, GE will update the product labeling to improve awareness and identify the hazard relative to proper storage.If you have any questions, call Center in US: 800-437-1171; Japan: 0120-055-919 and for other countries, please contact your local GE Healthcare Service Representative if you have any questions or concerns regarding this notification. |  | 229 | 229 | Worldwide distribution: USA (nationwide) including states of: AL, AK, AZ, AR, CA, CO, CT, DE, DC, FL, GA, HI, ID,IL, IN, IA, KS, KY, LA,ME, MD, MA, MI, MN, MS, MO, MT, NH, NJ, NM, NY,NC, OH, OR, PA, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, and WY; and countries of: VENEZUELA, UNITED KINGDOM, UNITED ARAB EMIRATES, TURKEY, THAILAND, SWITZERLAND, SWEDEN, SPAIN, SOUTH AFRICA, SINGAPORE, SAUDI ARABIA, RUSSIAN FEDERATION, REPUBLIC OF KOREA, PORTUGAL, NORWAY, NEW ZEALAND, NETHERLANDS, MARTINIQUE, JORDAN, JAPAN, ITALY, ISRAEL, IRELAND, HONG KONG, GREECE, GERMANY, FRENCH POLYNESIA, FRANCE, FINLAND, DENMARK, CZECH REPUBLIC, CHINA, CHILE, CANADA, BOSNIA AND Herzegovina, BELGIUM, AUSTRIA, AUSTRALIA, and ARGENTINA. | Terminated | October 16, 2012 | 329 | 860 | Computer
Z-0901-2011 | 57377 | No | LSZ |  |  | Anesthesiology | PMA | One Cat. Life Pulse High Frequency Ventilator, Catalog Number: 203A, Manufactured by: Bunnell Incorporated Salt Lake City, Utah, Rx Only. |  |  | October 25, 2010 | 2010 | Bunnell, Inc. | Product software coding error may compromise patient safety. | Software design | Customers were notified by an "Urgent: Medical Device Recall or Correction" letter starting 10/22/10. The letter identified the product and described the issues. Customers were given the option of either returning their ventilator(s) for rework or to receive installation/re-test instructions with an upgrade kit to rework themselves. The letter states that all Life Pulse ventilators in clinical use must be upgraded and there is no charge for the software upgrade. |  | 660 | 660 total | Worldwide Distribution -- USA, Australia, Canada, Czech Republic, Korea, Malaysia, Trinidad &amp; Tobago, and Uruguay. | Terminated | April 17, 2012 | 540 | 660 | Computer
Z-0389-2012 | 60344 | [6 empty] | Magnus Hybrid operating table columns 1180.01A1 and 1180.01B1 |  |  | September 19, 2011 | 2011 | Maquet Inc. | A component of the Magnus hybrid OR system fails to respond to a position command. | Device Design | Maquet sent a PRODUCT RECALL URGENT MEDICAL DEVICE CORRECTIVE ACTION letter dated September 19, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. To avoid the potential of a malfunction occurring prior to their scheduled Service procedure, customers were instructed to ensure that the MAGNUS table-top is positioned at the "0" pivot position and that the MAGNUS column pivot brake is secured. If it is necessary to pivot the MAGNUS table-top prior to procedure start, ensure that the column pivot brake is completely released and then completely secured when returned to "0" pivot position. Customers were instructed to ensure that all user of the affected product and other necessary persons are informed of the product correction notice. For technical support or questions call 1-888-627-8383 |  | 45 | 45 UNITS | Worldwide Distribution - USA including VA, LA, MO, NC, MD, and AZ and the countries of AE, AU, CA, CH, CN, DE, DK, ES, IE, IL, SP, PL, RU, and TH. | Terminated | December 08, 2011 | 80 | 45 | Computer
Z-1031-2011 | 57571 | No | GEH | Cryosurgical unit and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Presice Cryosurgical System, model FPRCH2047, and the following system is not approved in the US: CryoHit-P Cryoblation System, model FPRCH2058. The system is a software controlled cryoblation system designed to destroy targeted tissue using a minimally invasive percutaneous approach. Freezing and thawing of the tissue is achieved through the use of the Joule-Thompson effect, using a high-pressure gas distributed through closed-end needles placed into the patient;s target organ. |  |  | December 20, 2010 | 2010 | Galil Medical, Inc. | This voluntary correction has been initiated due to the discovery of a damaged optocoupler on one of our Presice Systems during periodic preventive maintenance. The optocoupler sends a signal across the locking mechanism in the Presice and Seednet MRI Systems, enabling the software to recognize if the channel on the system manifold is locked or unlocked. If an optocoupler is damaged or malfunction | Device Design | Galil Medical sent an "Urgent Medical Device Correction" letter dated December, 30, 2010, to all affected consignees. The letter described the product, the problem involved, and the action to be taken by the consignee. Consignees were instructed to make sure the locking mechanism on the Precise and Seednet MRI Systems were in the "locked" position when a needle is inserted into the manifold. Gently pull on the black plug to ensure needle is secure in locking mechanism. Apply the enclosed alert sticker on the bottom of the distribution panel and the bottom of the system monitor of the Precise System and on the side of the mobile unit itnerface panel and the top of the system monitor of the Seednet MRI System. Consignees were asked to take a picture of the alert sticker after it has been applied to the appropriate location on their Presice or Seednet MRI System(s) and e-mail to yoni.kadosh@galilmedical.com. Consignees were instructed to complete and return the enclosed form as soon as possible.For questions regarding this recall call Customer Service at (877) 639-2796.problem involved. |  | 63 | 63 | Worldwide Distribution - USA including CA, CT, FL, GA, MA, MD, MN, MO, NC, NY, OH, PA, SC, TN, TX, VA, and WA and the countries of CANADA, DENMARK, FINLAND, FRANCE, GERMANY, HONG KONG, ISRAEL, ITALY, KOREA, NETHERLANDS, SPAIN, SWEDEN, TAIWAN, TURKEY, UNITED KINGDOM, CHINA, and JAPAN.CANADA, DENMARK, FINLAND, FRANCE, GERMANY, HONG KONG, ISRAEL, ITALY, KOREA, NETHERLANDS, SPAIN, SWEDEN, TAIWAN, TURKEY, UNITED KINGDOM, CHINA, JAPAN. | Terminated | March 27, 2012 | 463 | 85 | Computer
Z-0522-2012 | 60426 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo.plaza |  |  | September 21, 2011 | 2011 | Siemens Medical Solutions USA, Inc | Potential malfunction when using syngo.plaza. On syngo.plaza workplaces large studies can be loaded into the viewer if transferred from a DICOM node or imported even while image transfer is not yet completed. In these situations more images of that study may exist which are not shown in the syngo.plaza viewer. | Software design | Siemens Medical Solutions USA sent a Customer Safety Advisory Notice dated March 10, 2011 to the affected customer. The letter identified the product the problem and the action needed to be taken by the customer.The letter informed the customer of the potential issues and provides additional information on the implemented resolution.For further questions please call (610) 219-4834. |  | 1 | 1 | Nationwide Distribution New York | Terminated | January 31, 2014 | 863 | 1 | Computer
Z-1004-2011 | 56759 | No | NQQ | Ultrasonic pulsed echo imaging system. | Radiology | General &amp; Plastic Surgery | 510(k) | Optovue iVue optical coherence tomography system, computer controlled opthalmic imaging system, manufactured by Optovue, Fremont, CA |  |  | July 20, 2010 | 2010 | Optovue Inc. | Use of the CAM attachment may result in a positioning anomaly and mistreatment in the eyes of the patient. | Software design (manufacturing process) | Optovue Inc. sent a Field Modification/Correction Notice dated July 20, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were informed that they could continue to use the iVue until the correction has been made. Customers were instructed to review any existing iVue patient data to ensure that this situation and an improper designation of eye assignment has not occurred. Upon completion of the updates, customers were to sign the attached form indicating they acknowledge that their iVue has been updated and they understand that they must review existing data. For any questions regarding this recall call 1-866-344-8948 or by mail at 45531 Northport Loop W, Fremont, CA 94538. |  | 22 | 22 devices in the US. | Nationwide | Terminated | February 08, 2011 | 203 | 22 | Computer
Z-1027-2011 | 52903 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | SEQUENCER component of MOSAIQ versions 1.30 (all builds), 1.40 (all builds), 1.50 (all builds), 1.60 (all builds), and 2.0 (all builds). Manufactured by IMPAC Medical Systems, Inc. Sunnyvale, CA. |  |  | July 17, 2009 | 2009 | Impac Medical Systems Inc | Potential for improper treatment. A software error may result in potential major underdosing to targeted areas from jaw blockage of the MLC aperture (jaw intrusion), and dosing to non-targeted areas from interstitial leaf leakage (jaw extrusion). | Software design | Impac Medical Systems Inc sent an Important Safety Notice letter to all affected consginees. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to contact Impact Software support as soon as possible if they have the affected configurations identified in the letter. Customers were asked to complete the form included with the letter and and fax to:European customers: +44 1293 654401 American &amp; rest of world customers 702-992-5002 or e-mail to support@impac.com Customers were asked to distribute the Safety Notice to any and all users of IMPAC software at their organization who are potentially affected by this issue. For questions e-mail support@impac.com or callEuropean customers: 44 1293 654320American &amp; rest of world customers: 800-488-4672. |  |  |  | Worldwide Distribution - USA including AL, CA, CT, FL, GA, IA, IL, KY, MI, MN, MS, ND, NJ, OH, OR, PA,TN, TX, VA, WA, WI, and WV and the country of Canada. | Terminated | November 25, 2011 | 861 |  | Computer
Z-0261-2012 | 60405 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Software: SA BASE 3.1.4.x, 3.1.6.x and SA LabMic 4.0.0.x, 4.0.1.x, 4.0.2.x, 4.0.3.x, 4.0.4.x. |  |  | August 08, 2011 | 2011 | SCC Soft Computer | On 08/08/2011 SCC Soft Computer, Clearwater, FL initiated a correction on following SoftLab used with: SA BASE 3.1.4.x, 3.1.6.x and SA LabMic 4.0.0.x, 4.0.1.x, 4.0.2.x, 4.0.3.x, 4.0.4.x. A client reported a flagging error where the result was flagged as critical in the Instrument Menu but the printed report did not show the corresponding flag. All affected clients were notified of the issue on 0 | Software design | SCC Soft Computer notified all affected customers on August 8, 2011, of the recall via the SCC Soft Computer proprietary Task management system to manage client communication. Clients use this tool to log in requests for services and receive and provide updates to any issue. Each client user has a unique password and every action is date and time-stamped. Each Task has a unique identifier and a full audit trail.SCC recommended to the customer until Live on the correction avoid making edits in the instrument menu.For further questions please call (727) 789-0100. |  | 255 | 255 | Worldwide Distribution - USA ( nationwide) and Canada. | Terminated | November 22, 2013 | 837 | 255 | Computer
Z-1011-2011 | 57275 | No | CEM | Potassium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ABL90 FLEX Blood Analyzer Part Number: 393-090. |  |  | November 01, 2010 | 2010 | Radiometer America Inc | Due to a software error, a patient mix up can occur on the ABL90. If a sample is pre-registered and the sample "fails" the next result will inherit the patient data from the previous result. | Software design | On 11/16/10 the firm sent Medical Device Field Correction letters to their customers. The letter states that all systems will be upgraded and until each system is upgraded, they instructed their customer to "always query for patient ID prior to introducing the sample." They also instruct each customer to complete the "Fax Return Form" acknowledging receipt of the letter. If customers have questions regarding the letter or the upgrade, they should contact Technical Support at 1-800-736-0600 Opt 4. |  | 432 | 432 units | Worldwide Distribution -- USA, including states of MI, NJ, &amp; PA and countries Australia, Austria, Belgium, Denmark, Finland, France, Germany, Indonesia, Italy, Japan, Kuwait, Malaysia, Netherlands, Norway, Poland, Russia, Spain, Sweden, Switzerland, United Arab Emirates &amp; United Kingdom. | Terminated | May 09, 2012 | 555 | 432 | Computer
Z-1010-2011 | 57595 | No | ONG | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Fraxel re:storeÂż Dual Laser System----a laser surgical instrument for use in general and plastic surgery and in dermatology. |  |  | December 15, 2010 | 2010 | Solta Medical, Inc. | The product has the potential to cause inadvertent laser firing which may lead to patient/operator injury if the problem recurs. | Software design | The firm, Solta Medical, issued three "DEVICE CORRECTION" letters, one dated December 15, 2010, the second one date December 17, 2010 and the third one dated December 20, 2010, to all affected customers. The letters described the product, problem and action to be taken by the customers. The customers were instructed to follow the instructions carefully, sign and date the bottom of the instruction form upon completion of the update and promptly return the signed form and USB stick in the enclosed, return-addressed envelope (Product Support, Solta Medical, 25881 Industrial Blvd., Hayward, CA 94545) Note: Solta Medical sales personnel will assist with the software update process so, if your local representative has already updated your system, you do not need to do anything further.If you have any questions, please call Solta Product Support at (510) 259-7291. |  | 352 | 352 units | Worldwide Distribution: USA and countries including: Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, China, Egypt, Germany, Greece, Hong Kong, India, Italy, Japan, Korea, Kuwait, Mexico, Netherlands, New Zealand, Pakistan, Philippines, Poland, Russia, Singapore, South Africa, Spain, Thailand, Turkey, Ukraine, United Kingdom, and Vietnam | Terminated | November 14, 2011 | 334 | 352 | Computer
Z-0150-2012 | 59991 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | RSP Humeral Socket Reamer (Size 32 SML): |  |  | September 28, 2011 | 2011 | Encore Medical, Lp | DJO Surgical has determined that the alignment pins cannot be consistently cleaned using the technique provided in the cleaning/sterilization section of the instrument IFU. | Device Design | DJO Surgical sent an URGENT FIELD SAFETY NOTICE letter dated September 28, 2011, to all affected customers. The letter identified the product. the problem, and the action to be taken by the customer. Customers were instructed to return the RSP Humeral Socket Reamers and RSP Baseplate Rim Planers. The RSP Humeral Socket Reamers will be reworked and returned to the customer. The Baseplast Rim Planer will be replaced when a suitable design is available. During the interim until instruments are returned, the use of FMP acetabular reamers size 42mm, 43mm, 45mm (DJO PIN 803-05-226, 227, 200) for the humeral preparation was validated to be effective to prepare the humerus for the implant.Customers were instructed to immediately notify the hospitals and surgeons using these devices and document it on the Field Safety Notice Response Form. For questions regarding this recall call 512-832-9500. |  | 399 | 399 Reamers | Worldwide Distribution - USA including AR, AZ, CA, CO, FL,GA, ID, IN, IL, LA, MA, MD, ME, MI, MO, NJ, NM, NY, OH, OK, PA, PR, SC, TN, TX, UT, VA, WA, and WI and Internationally toGermany, South America, Italy, Saudi Arabia, United Kingdom, and Canada. | Terminated | March 29, 2013 | 548 | 532 | Computer
Z-0032-2012 | 59853 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | MDC PACS - release R2.3 SP1 . |  |  | September 01, 2011 | 2011 | Philips Healthcare Inc. | Clinical measurements on true size printouts may be inaccurate | Software design | Philips Medical notified all affected customers with a "Field Safety Notice" dated August 2011.The letter provides instructions to the customer for how to avoid the problem. Customers are informed that until the software upgrade is recieved from Philip, do not use the true-size printing feature if the system is configured for more than one DICOM printer. Customers will be instructed that they will receive a software upgrade when it is available. For questions on this recall contact Philips representative at (800) 722-9377. |  | 20 | 20 units | Worldwide Distribution: Nationwide distribution including Puerto Rico; and the countries of Argentina, Brazil, China, France, Greece, India, Israel Mexico, and Turkey. | Terminated | March 13, 2017 | 2020 | 20 | Computer
Z-0263-2012 | 60155 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Infinity(R) Acute Care System (IACS) Monitoring Solution (M540) - Physiological Patient Monitoring Solution |  |  | October 18, 2011 | 2011 | Draeger Medical, Inc. | During product training on the Infinity(R) Acute Care System (IACS) Monitoring Solution (M540) it was discovered that the weight-based drug dosage calculation may indicate incorrect recommended values. In addition, a misalignment of ECG and ART waveforms delivered from an Infinity(RO Acute Care System (IACS) Monitoring Solution (M540) was observed on the Infinity Central Station. | Software design | The firm initiated their recall of this device on October 17, 2011 by issuing a letter to the consignee instructing customers to follow the Instructions for Use and use the Infinity Central Station only for remote assessment of a patient's status. They recommend that you do not consider the waveforms on Infinity Central Station for primary diagnosis purposes. In case of a high priority alarm situation, "Check the patient and treat if necessary".If the misalignment issue is observed, it can be cleared by undocking and redocking the M540 or by changing the order of waveforms displayed on the Cockpit screen (or any other change to the auto display screen layout). |  | 476 | 476 | Nationwide distribution - IL | Terminated | April 01, 2013 | 531 | 476 | Computer
Z-0272-2012 | 60164 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX CardioVascular (CV) Admin Tool |  |  | June 06, 2011 | 2011 | AGFA Corp. | Potential for misattribution of data and for a report to become unavailable when the user requests for it to be moved to another study that already has a report associated with it. | Software design | An 'URGENT FIELD SAFETY NOTICE" was sent via FED-EX by June 6, 2011 to all the sites using the affected product. The letter described the issue and mitigation. The letter recommends that users not move reports to studies where a report already exists. If a report is being moved because the existing report is irrelevant to the study, then the existing report should be deleted prior to moving the new report to that study. If the report is being moved because the user wants to combine the content of the reports, then the "merge" function should be utilized. The letter also states that a software enhancement is currently available. Acknowledgment via FAX-Back was requested from the sites. Questions regarding the issue should be directed to customers' local Agfa HealthCare Service at 877-777-2432. The firm has identified 185 additional sites and an "URGENT FIELD SAFETY NOTICE" was sent on October 12, 2012. |  | 323 | 323 units | Worldwide Distribution -- USA, Virgin Islands and Canada | Terminated | June 26, 2012 | 386 | 323 | Computer
Z-1064-2011 | 53054 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Linear Medical Accelerator ARTISTE MV systems, when equipped with a syngo RT Therapist, Therapist Express Basic, syngo RT Therapist Assist or Therapist Express Assist workspace with software version NET VA10A for RTT 4.1.122 or NET V10B for RTT 4.1.125, manufactured by Siemens Medical Solutions, Concord, CA. |  |  | July 30, 2009 | 2009 | Siemens Medical Solutions USA, Inc | Potential for Mistreatment-- when treatment plans intended for non-ARTISTE machinese are loaded, if the plans contain unspecified beam parameters, attributes or unknown block codes, the therapy system may remove those parameters. | Software design | The firm,Siemens, sent two Customer Advisory letters (TH004/09/S and TH008/09/S) contained in the Update Instructions to all consignees/customers with active devices. The letters describe the product, problem and action to be taken by the customers. The customers were instructed to follow the updated instructions. Siemens service representative will sign a document noted as CORRECTED upon completion of the software installation.If you have any questions, call 1-925-602-8157. |  | 151 | 151 units | Worldwide distribution: USA and countries including: Germany, France, Poland, Norway, Australia, Malaysia, Spain, New Zealand, the UK, India, Republic of Korea, Canada, Ireland, Italy, Egypt, Belgium, Philippines, Saudi Arabia, Netherlands, Hungary, Norway and South Africa. | Terminated | April 12, 2011 | 621 | 151 | Computer
Z-0278-2012 | 60430 | No | LNX |  |  | General Hospital | 510(k) | SCC Soft Computer, Softlab Laboratory information System with DORA versions: 2.0.0.4 with Hot Fix 1.2808, 1.2809, 1.2956, or 1.3476; 2.0.0.5; 2.0.0.6; 2.0.0.7 with Hot Fix 1.6595; and 2.1.0.3. |  |  | March 22, 2011 | 2011 | SCC Soft Computer | Potential for patient data to be sent to the wrong patient record. A software change affected how the Softlab handles recovery from certain types of table access conflicts. When such conflicts occur, this function has the potential to cause faulty writing to the database. | Software change control | Clients were notified through SCC Soft Computer's proprietary Task Management System On 03/22/201. All clients have received a hot fix that corrected the issue or made the hosparam change that prevents the issue. |  | 12 | 12 | Sold throughout US and Canada. | Terminated | October 29, 2012 | 587 | 12 | Computer
Z-0281-2012 | 60269 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Draeger Infinity(R) Delta, Delta XL, Gamma XXL and Kappa Monitors are used to monitor, physiologically, the vital signs of patients (with Arrhythmia Detection or Alarms). |  |  | October 31, 2011 | 2011 | Draeger Medical, Inc. | Firm received 2 complaints reporting Ventricular Fibrillation (Vfib) arrhythmias were not recognized by the monitor and did not alarm when the Masimo SET pod is used for SpO2 measurements. | Software design | Firm initiated their recall of this device on October 25, 2011 by sending an "Urgent Medical Device Recall" letter to consignees. A Technical Service Bulletin will be generated and distributed globally that instructs the affected International Draeger subsidiaries to inform customers in their countries. The software anomaly has been corrected in software version 7.3. Draeger Service Representatives will schedule time with customers to upgrade the software in affected monitors. The software upgrade will be provided free of charge. |  | 14894 | 14,894 | Worldwide distribution- USA including the states of AL, AR, AZ, CA, CO, CT, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NH, NJ, NM, NV, NY, OH, OK, PA, RI, SD, TN, TX, UT, VA, WI and WV and International distribution including the countries of Algeria, Angola, Argentina, Australia, Austria, Belarus, Belgium, Bolivia, Botswana, Brasilien, Brazil, Brunei Darussalam, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Deutschland, Dominican Republic, Ecuador, Egypt, El Salvador, Equatorial Guinea, Estonia, Finland, France, Frankreich, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Indien, Indonesia, Iran, Iraq, Ireland, Israel, Italy, Japan, Jordan, Jordanien, Kazakhstan, Kenya, Kuwait, Latvia, Lebanon, Libya, Lithuania, Madagascar, Malaysia, Malta, Mauritius, Mayotte, Mexico, Moldavia, Moldova, Morocco, Netherlands, Norway, Norway, Peru, Poland, Portugal, Puerto Rico, Qatar, Reunion, Romania, Rumania, Russian Fed., Saudi Arabia, Serbia, Seychelles, Singapore, Slovenia, Slowenien, South Africa, South Korea, Spain, Sri Lanka, SÂżdafrika, Sweden, Switzerland, Syria, Taiwan, Thailand, Tschechische Republic, Tunisia, Turkey, Turkmenistan, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Venezuela, Vietnam, and White Russia. | Terminated | July 01, 2013 | 609 | 14894 | Computer
Z-0014-2012 | 59949 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Dynamic Multileaf Collimator (DMLC) |  |  | May 25, 2011 | 2011 | Elekta, Inc. | It is possible to load different patients in the record and verify (R&V;) system and in the multilead control software (MCS). | Device Design | Elekta, Inc. sent an "IMPORTANT NOTICE" letter dated May 25, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. The letter recommends several precautions customers should use until Elekta finds a solution. Customers will be contacted when a solution is found. |  | 27 | 27 | Nationwide Distribution-USA including the state of CT, FL, GA, IA, MN, MO, NJ, NY, ND, OH, OK, PA, RI, TX, VA , and WA. | Terminated | January 14, 2014 | 965 | 27 | Computer
Z-1002-2011 | 56357 | No | FAT | Esophageal dilator. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Stryker Endoscopy, Infravision Esophageal Kit, |  |  | May 11, 2010 | 2010 | Stryker Endoscopy | The silicone sheath component of the product detaches from the catheter upon removal of the e-kit from the body. | Device Design | On 5/11/10, Stryker Endoscopy initiated the recall notification letters explaining the reason for the recall and requesting the product be returned. |  | 4125 | 4125 | Nationwide and 3 foreign consignees (Canada, Singapore and Switzerland). | Terminated | February 01, 2011 | 266 | 4125 | Computer
Z-0277-2012 | 60436 | No | LNX |  |  | General Hospital | 510(k) | SoftLab GUI versions 4.0.2.8-4.0.2.10; 4.0.3.12-4.0.3.13; and 4.0.4.1-4.0.4.4. |  |  | May 03, 2011 | 2011 | SCC Soft Computer | On 05/03/2011 SCC Soft Computer, Clearwater, FL initiated a correction on the following SoftLab GUI versions 4.0.2.8-4.0.2.10; 4.0.3.12-4.0.3.13; and 4.0.4.1-4.0.4.4. A client reported the PT and INR results did not get flagged. The PROT should have been flagged as HIGH and the INR should have been flagged as PANIC. Neither test was flagged. All affected clients were notified of the issue on 05/03 | Software design | SCC Soft Computer initiated a Field correction on May 3, 2011. All customers with affected software have been notified of this issue. Clients were instructed to have a "valcheck" utility run to determine if there are any tests in their system affected by the issue. They were supposed to indicate whether they wanted to schedule to receive the patch or hot fix.If you have any further questions please call ( 727 ) 789-0100. |  | 69 | 69 | Worldwide Distribution-- USA ( nationwide) and Canada. | Terminated | August 07, 2013 | 827 | 69 | Computer
Z-0146-2012 | 58162 | No | OLO | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | Labeling for RIO Robot Unit is comprised of three main components: 20399 RIO SURGICAL ARM, 201251 RIO GUIDANCE MODULE, 200294 RIO CAMERA STAND ASSEMBLY***PN 203999 SN ROB 125 2010-12 V 100/120/230 A 9.6/8.0/4.2 Hz 50/60 Class I Equipment. Conforms to IEC 60601-1/A2: 1995, EN 60601-1/A2: 1995 UL 60601-1: 2003, CAN/CSA-C22.2 No. 601.1-M90.***Manufactured in USA***MAKO SURGICAL CORP. 2555 DAVIE ROAD, FT. LAUDERDALE, FL 33317. |  |  | January 20, 2011 | 2011 | Mako Surgical Corporation | MAKO Surgical Corp. is recalling their RIO Robotic Arm Interactive Orthopedic System (RIO) due to software issue that exist that could potentially result in a bone resection. No adverse events reported. | Software change control | The firm, MAKO Surgical Corp., sent an "RECALL NOTICE 2.4. SITES" letter dated February 1, 2011 to its customers. The letter describes the product, problem, actions to be taken which includes a workaround and an updated user guide identifying the appropriate workflow. The customers were instructed to follow the user guide and workflow and to complete and return the enclosed confirmation memo to their MAKOplasty Specialist or return it in the pre-addressed, pre-paid envelope provided. MAKOplasty included with the letter new steps in the surgical workflow in the use of the RIO system to ensure error value accuracy; worked on a software update that will not allow the user to proceed with error values that do not support the required accuracy-this software was to be supplied to customers in February 2011, and trained MAKOplasty Specialist will make available written updates to the RIO use instructions as well.If you have any questions or concerns, please do not hesitate to speak with your MAKOplasty Specialist or contact the Manager of Customer Relations at 954.927.4022 x447. |  | 70 | 70 | Worldwide distribution: USA including states of: AZ, CA, CO, FL, GA, IA, IL, LA, MI, MO, MS, NC, NJ, NV, NY, OK, OH, OR, PA, RI, TX, UT, VA, WA, WI, and WV; and countries of: Scotland, Italy and Korea. | Terminated | December 13, 2011 | 327 | 70 | Computer
Z-1001-2011 | 56354 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand C Series ClinicÂżs (Includes Trilogy, and Novalis Tx), Model Numbers: H14, H26, H27, H29, Distributed by and/or Manufactured by: Varian Medical Systems Inc., Palo Alto, CA |  |  | December 01, 2009 | 2009 | Varian Medical Systems Oncology Systems | A circuit design anomaly whereby a faulty balance potentiometer can lead to an undetected electron beam asymmetry. | Device Design | The firm, Varian Medical Systems, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated December 1, 2009 via Federal Express to consignees/customers. The letter described the product, problem and actions to be taken by the customers and the firm. The customers were instructed to perform a verification of field symmetry on electron beam energies to determine the threshold and report any field symmetry issues to their local Varian service office for attention. The firm, Varian, is developing new functionality to detect the symptoms of this component failure condition under all conditions of electron beam operations and asserting a beam interlock if it is detected. Note: Varian indicated until the solution can be developed, they will conduct an audit of installed Clinacs to identify those the suspect "blue" electron beam balance potentiometers installed, and will execute a program of prophylactic replacement of these potentiometers.If you require further clarification, please contact your local Varian Customer Support District or Regional Manager at phone: USA and Canada-1-888-VARIANS (888-827-4265); Europe- +41 41 749-8844 or email support-emea@varian.com; North America: support-americas@varian.com; Australia/New Zealand: support-anz@varian.com; South East Asia: seasia.apps.helpdes@varian.com; China/Asia: china.apps.helpdesk@varian.com; Japan: Japan.apps.helpdesk@varian.com; Latin America: soporte.al@varian.com; Internet: Oncology Systems - www.myvarian.com; Varian Medical Systems -www.varian.com |  | 4411 | 4411 | USA and countries including: Algeria, Angola, Argentina, Australia, Austria, Azerbaijan, Bangladesh, Belgium, Belorussia, Bolivia, Brazil, Canada, Chile,, China, Columbia, Costa Rica, Czech Republic, Denmark, Ecuador, Egypt, Estonia, Finland, France,French Polynesia, Germany, Guadeloupe, Guatemala, Hong Kong, Hungary, Iceland, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Kazakhstan, Kuwait, Latvia, Lebanan, Luxembourg, Maca, Macedonia, Malaysia, Martinique, Mexico, Moldova, Morocco, Nepal, Netherlands, New Zealand, Norway, Oman, Pakistan,, Philippines, Poland, Portugal, Qatar, Reunion, Romania, Russia, Saudi Arabia, Singapore, Slovakia, South Africa, South Korea,, Spain, Sweden, Switzerland, Syria, Taiwan, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom, and Uruguay.. | Terminated | March 14, 2012 | 834 | 4411 | Computer
Z-0258-2012 | 60021 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Precedence System; |  |  | September 27, 2011 | 2011 | Philips Medical Systems | A mechanical malfunction of the button on the hand controller on the Philips Precedence and SKYLight SPECT systems may get stuck and cause unintended motion of a detector. | Device Design | Philips Medical Systems sent a "Customer Information" letter dated September 27, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. The letter lists additional precautions the customers should use while operating the device. Philips will notify customers for replacing the affected product through an Action for Performance-Proactive Field Change Order to resolve the issue. Contact Philips Customer Care Solutions at 1-800-722-9377, option 5, for questions concerning this notice. |  | 658 | 658 - all units in recall | Worldwide Distribution-USA (nationwide) including Puerto Rico. | Terminated | December 21, 2012 | 451 | 658 | Computer
Z-0211-2012 | 60154 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX Cardiovascular Results Management (RM) Software Versions RM 2.04.37.04 to RM 7.8 SU2 |  |  | June 06, 2011 | 2011 | AGFA Corp. | Improper merging of IMPAX Cardiovascular Results Management (RM) reports causes the signed report to become unavailable. | Software design | AGFA Healthcare sent an "URGENT FIELD SAFETY NOTICE" letter dated May 27, 2011, via Fed Ex to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed not to carry out a merge post finalization of an RM reports. If it is necessary to modify a finalized report, AGFA recommends that they make the modification through one of the following actions:Amend the report through the RM interface, orCall AGFA Service to change the status of the report to "preliminary" and then proceed with the finalization of the report.AGFA issued an Acknowledgment, via FAX- Back that the information was received and understood to affected customers. For any questions customers were instructed to contact their AGFA Healthcare Service at 877-777-2432. |  | 323 | 323 | Worldwide Distribution - USA (nationwide) and Canada | Terminated | July 02, 2013 | 757 | 323 | Computer
Z-0932-2011 | 57527 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | 4D Integrated Treatment Console (4DITC); |  |  | December 14, 2010 | 2010 | Varian Medical Systems, Inc. Oncology Systems | Multiple isocenter plans will load on the 4DITC and are available for imaging and treatment, but the 4DITC software does not recognize that mulitplex isocenter is present. | Software design | Urgent Medical Device Correction/Field Safety notification letters were distributed to all affected users, and to Varian Sales, Marketing and Service Organizations on December 14, 2010 by certified mail. ***12-20-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.*** The letters described the issue and recommended user actions along with Varian's actions. Users are to use caution when acquiring actual couch positions for any plan. In addition, each plan should be reviewed for quality assurance purposes in accordance to set departmental or organizational policies. The firm updated the procedure and enclosed a copy with the letter. It was also posted to the customer support site. If further clarification is needed, customers are to contact their local Varian Customer Support District or Regional Manager. |  | 4188 | 4188 | Worldwide Distribution. | Terminated | July 22, 2011 | 220 | 4188 | Computer
Z-1023-2011 | 54658 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Radiation therapy device. Medical charged particle radiation therapy system with software version. Model Numbers: Multi Plan Treatment Planning System, Software version 3.0, 3.1, 3.5.0, and 3.5.1. |  |  | January 15, 2010 | 2010 | Accuray Inc | When a user modified the align center of an existing plan during the treatment planning process the effective depths of each beam in the plan were not updated by the system. These depths were incorrect and led to an incorrect dose calculation. Software patch /update to address this issue. Implementation at affected sites will initiate, once the patch is available. | Software design | The firm, Accuray, sent an "URGENT DEVICE CORRECTION" letter dated January 18, 2010, to all customers. The letter described the product, problem and action to be taken by the customers. The customers were instructed to load the plan and recalculate the dose to reflect the current state of beam, compare the reloaded plan and the recalculated plan to verify maximum dose and prescription values, and re-prescribe and save to overwrite to existing plan if necessary. The firm is developing a patch/software update to address the issue and a Accuray Customer Support person will contact the customer regarding further resolutions.If you have any questions or concerns regarding this issue, contact Accuray Customer Support at +1-877-668-8667(USA) or +1-408-716-4700 (non USA) or email: customersupport@accuray.com. |  | 99 | 99 units | Worldwide Distribution: USA and countries including: Canada, Turkey, India, Russia, Italy, France, Belgium, Netherlands the UK, Spain, Athens, Saudi Arabia, Switzerland, the Ukraine, Germany, Korea and Hong Kong. | Terminated | February 11, 2011 | 392 | 99 | Computer
Z-0198-2012 | 60145 | No | DZJ | Bone cutting instrument and accessories. | Dental | Dental | 510(k) | Right IBO Blade , for use with Stryker TPS Reciprocating Saw, REF 5100-37-901S2, Stryker Instruments, Kalamazoo, MI. |  |  | October 18, 2011 | 2011 | Stryker Instruments Div. of Stryker Corporation | There is a potential for these blades to break where the blade is welded on to the arbor. This hazard can result in the following harms: tissue damage due to the broken blade being aspirated or swallowed, additional surgery to remove a broken blade, infection due to an unretrieved device component, complications associated with increased time under anesthesia to retrieve a broken blade, osteomyeli | Device Design | The firm, Stryker, sent an "URGENT MEDICAL DEVICE RECALL NOTIFICATION" letter dated October 18, 2011 to Risk Managers, Office Managers, and Oral Surgeons. The letter described the products under recall, the problem and a series of actions to be taken. Customers were asked to immediately locate recalled products and quarantine. Products were to be shipped back to Stryker. The letter included a Business Reply Form to be faxed back to Stryker Instruments Regulatory Department, 866-521-2762. The letter instructs to forward this notification to all affected locations if the recalled products were further distributed. Customers with questions can call Stryker at 1-800-253-3210. |  | 113 | 113 | Worldwide Distribution - USA (nationwide) and the countries of Brazil, Canada, China, Germany, Greece, Hong Kong, India, Jordan, Miramar, Saudi Arabia, Singapore, Spain, Switzerland, United Kingdom. | Terminated | July 12, 2013 | 633 | 12631 | Computer
Z-0191-2012 | 60174 | No | LKN |  |  | Gastroenterology/Urology | 510(k) | Spectra Optia Apheresis System, REF: 61000, Caridian BCT, Lakewood, CO 80215. |  |  | September 21, 2011 | 2011 | Caridian BCT, Incorporated | Possible higher than expected platelet loss following therapeutic plasma exchange or plasmapheresis. | Software design | CaridianBCT sent a "Safety Alert" letter dated September 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The firm will be sending a recall letter to consignees in November 2011, stating they are voluntarily recalling the software and will provide an upgrade once the software has been approved.If you have any further questions please contact your CaridianBCT Representative, the CaridianBCT Support Center at 1-877-3-FYU-BCT ( US Toll free 1-877-299-4228 ) or (303) 231-HELP ) ( 1 (303) 231-4357, or your local CaridianBCT Customer Service office. |  | 852 | 852 units (US: 333; Foreign: 519) | Worldwide Distribution -- USA ( nationwide ) and the countries of Australia, Belgium, Canada, Chile, China, Colombia, Hong Kong, India, Japan, Malaysia, Mexico, Pakistan, Republic of Korea, Saudi Arabia, Singapore, South Africa, Taiwan, Thailand and Vietnam. | Terminated | February 01, 2013 | 499 | 852 | Computer
Z-0182-2012 | 60036 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Extended Brilliance Workstation-NM (EBW-NM) |  |  | September 15, 2011 | 2011 | Philips Medical Systems | A problem related to a software error has been detected in the Philips Multiple Gated Radionuclide Angiocardiography (MUGA) application. A false negative interpretation due to a falsely elevated ejection fraction or a false positive interpretation due to a false low ejection fraction could occur. | Software design | Philips sent all affected sites an Urgent Medical Device Correction Letter on September 27, 2011 for the MUGA issue. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the information contained in the letter with all members of their staff who need to be aware of the contents of the communication. Customers were asked to retain a copy of the communication with the equipment instruction for use. Additionally, a field change order was released on September 26, 2011, customers will be receiving release notes as part of their correction informing them of the changes and corrections to the software. The new software will be installed on affected units by Philips Field Service Engineers.For customers in North America contact Customer Care Solutions Center at 1-800-722-9377. Select option 5. In all other countries contact the local Philips Healthcare office. |  | 1181 | 1181 systems | Worldwide Distribution - USA (nationwide) and Canada. | Terminated | January 30, 2013 | 503 | 1181 | Computer
Z-1024-2011 | 57464 | No | CCW | Rigid laryngoscope. | Anesthesiology | Anesthesiology | 510(K) Exempt | Laryngoscope Handle, Reusable, Rx, Product Code 4558GSP or 4559LED, Vital Signs Colorado, Inc., 11039 E. Lansing Circle, Englewood, CO 80112. |  |  | December 20, 2010 | 2010 | Vital Signs Colorado Inc. | Battery which powers the laryngeal scope light may dislodge resulting in loss of illumination. | Device Design | Healthcare Systems sent an URGENT MEDICAL DEVICE CORRECTION letter dated December 20, 2010, to all affected consignees advising them of the possible problem and giving them interim safety instructions until new battery cartridges were available. Customers were instructed to complete and return the Confirmation Form which was included with the letter.For information regarding this recall call 973-956-5300. |  | 2202 | 2202 units | Worldwide Distribution - USA including AL, AR, AZ, CA, CO, FL, GA, HI, IA, IL, IN, KY, MA, MN, MO, NH, NJ, NY, OH, OK, SC, TN, TX, WA, and WV and the countries of Australia, Belgium, Canada, Dominican Republic, Ecuador, Germany, Hong Kong, Latvia, New Zealand, Taiwan, Turkey, and the UK | Terminated | July 01, 2011 | 193 | 2202 | Computer
Z-0201-2012 | 60118 | No | DYW | Orthodontic plastic bracket. | Dental | Dental | 510(k) | Spirit MB Brackets, part #494-1211. |  |  | May 18, 2011 | 2011 | Ormco Corporation | Ormco Corporation is voluntarily recalling three lots of Spirit MB Brackets due to a manufacturing error. | Device Design | Ormco Corporation sent an "URGENT MEDICAL DEVICE RECALL" letter dated June 13, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. The letter instructs customers to check their inventory for the affected product. If any of the affected product is found, customers are directed to contact Ormco Customer Care at 1-800-854-1741 to receive a RMA number. The RMA number will allow a quick return and replacement or credit. Additionally, a Recall Return/ Acknowledgement Form was attached for customers to complete and return via fax to 909-962-5605. |  | 135 | 135 packs | Worldwide Distribution-USA (nationwide) including the states of FL, MN, TX, and VA, and the countries of Australia, France, Germany, Italy, Japan, Mexico, Morocco, Netherlands, New Zealand, South Africa, Turkey, and Ukraine. | Terminated | April 24, 2012 | 342 | 145 | Computer
Z-0993-2011 | 57507 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Arstasls One Access System; Models: AAD100, AAD101, Part Numbers: FG-02279, FG-03010; The system consists of the following components: (1) Device Sheath/Anchor, (2) Shaft |  |  | October 19, 2010 | 2010 | Arstasis, Inc | Device has the potential to fracture and sheath separation during use. | Device Design | The firm, Arstasis, initiated a market withdrawal and sent notification via letters dated October 19, 2010, to customers informing the customers that Arstasis is implementing a product improvement and will replace all existing inventory that the customer has on their shelf with the improved product. The customers were instructed to work closely with their local Arstasis territory manager to ensure that their product replacement is handled satisfactorily, and complete and return the PRODUCT REPLACEMENT and PRODUCT RETURN forms to Arstasis Customer Service via fax at 1-877-688-8417 or email to customerservice@arstasis.com.If you have any questions or concerns, contact Customer Service toll free at 1-877-594-4545. |  | 330 | 330 units | Worldwide Distribution | Terminated | February 16, 2011 | 120 | 330 | Computer
Z-0209-2012 | 60420 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | SoftPath GUI versions: 3.17.5 to 3.17.9, 4.2.1 to 4.2.3, and 4.3.0 to 4.3.0.4. |  |  | July 13, 2011 | 2011 | SCC Soft Computer | Soft Computer, Clearwater, FL, initiated a correction on the following SoftPath GUI versions: 3.17.5 to 3.17.9, 4.2.1 to 4.2.3, and 4.3.0 to 4.3.0.4. SoftPath architecture does not correctly handle AP macros and data fields inside report sections that are stored in the database. Improper information can be displayed in all places where the diagnosis from the database is read and displayed. | Software design | The firm, SCC Soft Computer, sent a "CORRECTION COMMUNICATION" notice on July 13, 2011 to its customers. The notice describes the product, problem and actions to be taken. The customers were instructed to restrict security access to the template set-up files and to acknowledge receipt of the notice; grant SCC personnel permission to check their report templates and correct any that are affected; indicate if they wish to schedule the utility to be run and if they would like to have the SoftWeb parameter PathPDFLinkOnly set to "Y". SCC will review the templates on the customers systems; discuss any irregularities requiring change and move or remove AP macros and data elements as needed.If you have any questions, contact the Director of Quality Management at (727)789-0100. |  | 142 | 142 | Worldwide distribution: USA (nationwide) including states of: AR, AL, AZ, CA, CT, FL, GA, IL, IN, KS, KY, LA, MA, MD, MI, MO, MN, NH, NJ, NY, OH, OK, PA, SD, TX, VA, WA, WI, WV including Puerto Rico; country of Canada. | Terminated | July 17, 2012 | 370 | 142 | Computer
Z-1029-2011 | 55087 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | SEQUENCER for Mosaiq interfaced to Electa SL75-5 machines. |  |  | March 10, 2010 | 2010 | Impac Medical Systems Inc | Computer interface is defected. Computer based verification system for recording patient treatment on a radiotherapy treatment machine fails to record treated field. | Software design | A User Notice was sent by the firm to its consignees. It identified the affected product and the reason for the recall. It also discussed the clinical impact and interim mitigations. The firm states in the letter that a fix for the problem and been developed and is available. Customers are to contact IMPAC Support for the patch to the SL75-5 interface. If customers have questions, concerns, or requests, they can email support@impac.com or one of the numbers listed in the letter for the applicable country. The letter should be distributed to all users affected by the issue at the customer's organization. |  | 40 | 40 | Worldwide Distribution -- US, Canada, Germany, and France. | Terminated | April 06, 2011 | 392 | 40 | Computer
Z-1046-2011 | 56448 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Omnifit M/S PSL Porous Coated Acetabular Shell 40MM; |  |  | August 02, 2010 | 2010 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics determined that specific lots of the Omnifit Acetabular Shells may have a screw hole wall thickness that falls above the design specification. | Device Design | The firm, Stryker Orthopaedics, sent an "URGENT PRODUCT RECALL" letter dated August 2, 2010, via Federal Express to all customers. The letter described the product, problem and action to be taken by the customers. The customers were instructed to examine their inventory to identify the affected product; retrieve all affected product and return it their branch or agency warehouse for reconciliation; reconcile all product(s) on the attached Product Recall Acknowledgement Form and fax to (201) 831-6069, and return the affected product via Inter-Org Transfer and ship to the attention of Regulatory Compliance, Stryker Orthopaedics, 325 Corporate Drive, Mahwah, NJ 07430.If you have any questions, feel free to contact the Manager, Divisional Regulatory Compliance at (201) 972-2100. |  |  |  | Worldwide distribution: USA including states of: AL, AR, CA, FL, IN, ME, MI, MS, NC, NJ, OK, OR, WA , and WI; and countries including: China, Hong Kong and Korea. | Terminated | November 09, 2012 | 830 |  | Computer
Z-1021-2011 | 56923 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Vital Connect 4.0 &amp; 4.1, radiological image processing software |  |  | August 23, 2010 | 2010 | Vital Images, Inc. | There is a potential alignment error of fused PET/CT images in the Vitrea Core component of Vitrea Enterprise Suite 1.2 and 1.3. The error occurs during panning of the fused PET/CT images. Because the PET data has a larger voxel size than the CT data, it is possible to pan the CT image in more frequent, smaller steps than the PET image, so that the fused PET image does not move until the CT data | Software design | The firm, Vital Images, Inc,., sent a "Urgent Software Alert" letter dated August 23, 2010 to all customers. The letter described the product, problem and actions to be taken by the customers. The customers were instructed to not use the PET/CT fusion functionality until the software updates are provided and to re-examine any past results. If you have any questions or comments, contact Customer Support at support@vitalimages.com or at 1-800-208-3005. |  | 178 | 178 | Worldwide Distribution: USA including states of: AK, AZ, CA, CO, CT, FL, GA, IL, IN, IA, KY, LA, ME, MD, MA, MI, MN, MS, MO, NH, NJ, NM, NY, NC, ND, OH, PA, SC, TN, TX VA, WV, WI, and WY; and countries of: AUSTRALIA, HONG KONG, UNITED KINGDOM, NETHERLANDS, FRANCE, BELGIUM, CANADA, CHINA, AUSTRALIA, SWITZERLAND, and ITALY.AUSTRALIA, HONG KONG, UNITED KINGDOM, NETHERLANDS, FRANCE, BELGIUM, CANADA, CHINA, AUSTRALIA, SWITZERLAND, ITALY. | Terminated | November 06, 2012 | 806 | 178 | Computer
Z-1019-2011 | 57558 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | PCR Eleva. The marketing brochure is labeled in parts: "Philips Medical Systems is part of Royal Philips Electronics... www.medical.philips.com... Philips Medical Systems Nederland B.V". |  |  | December 23, 2008 | 2008 | Philips Medical Systems North America Co. Phillips | There is a risk of selecting the wrong patient, when patient is selected in the patient list and the examination tab pressed with background query are running at the same time. | Software design | On 1/30/2009, Philips began sending out URGENT-Device Correction PCR Eleva notifications to their consignees.The consignees were informed of the hazards as when a patient is selected in patient list and examination tab pressed with background query running at the same time, there would be a risk that the wrong patient is selected.The firm issued the Field Corrective Order (FCO) 73200029 to their representative to correct the problem. The corrective action would be a software upgrade that resolve the issue.All consignees can contact Philips Call Center at 800-722-9377, #5, #2 and reference "FCO 73200029" with any questions. |  | 686 | 686 units | Nationwide Distribution -- USA, including states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WI, WV, and WY and to LANDSTUHL ARMY BASE, APO in United Arab Emirates. | Terminated | August 25, 2011 | 975 | 686 | Computer
Z-0181-2012 | 60022 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Clinac, Trilogy, Trilogy Tx and Novalis linear accelerators. |  |  | September 09, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | Varian has received reports in which a user has remotely rotated the gantry into contact with the couch or with the patient, in both the manual mode and the automate mode. | Device Design | The firm, Varian Medical Systems, sent an "URGENT MEDICAL DEVICE CORRECTION/URGENT FIELD SAFETY NOTICE" letter dated September 9, 2011 to all affected customers via Federal Express with attached instructions. The letter provides the description of problem; references, details, recommended user actions, and varian actions.***12-20-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.***In addition, the letter also states "THIS LETTER IS NOT A SUBSTITUTE FOR THE CAREFUL REVIEW OF ALL USER DOCUMENTATION ASSOCIATED WITH YOUR PARTICULAR EQUIPMENT AND SOFTWARE."If you have any questions, contact Varian Oncology Help Desk at USA and Canada: 1.888.VARIANS (888.827.4265); Europe +41 41 749 8844 and/or Internet: Oncology Systems customer site-www.myvarian.com and Varian Medical Systems public site-www.varian.com. |  | 5823 | 5823 devices | Worldwide Distribution. | Terminated | July 24, 2012 | 319 | 5823 | Computer
Z-0213-2012 | 60101 | No | KGI | Bone densitometer. | Radiology | Radiology | 510(k) | DICOM Option for the Norland Illuminatus Software Revs 4.2.0 though 4.3.1. Models Norland XR-600, XR-800, XR-46, EXCELL, XR-36 and Eclipse DXA Bone Densitometers. |  |  | September 23, 2011 | 2011 | Cooper Surgical, Inc. | The DICOM Software, Norland Illuminatus DXA Software with the DICOM Option (Part Number 434A133) version 4.2.0 through 4.3.1, which is installed on the XR-600, XR-800, XR-46 and EXCELL DXA Bone Densitometers, could lead to improper patient information getting into a previous patient's record. | Software design | The firm, CooperSurgical, sent a "Notice of RECALL-URGENT" letter dated September 23, 1011 to its customers. The letter described the product, "Problem/Situation", and "When can this problem occur?". CooperSurgical will send a new revision of Illuminatus software to the customers once available.If you have any concerns about your existing database of patients, please contact our Norland Technical Support team at (203) 601-5200 Extension 9888 or 9889. |  | 184 | 184 | Worldwide distribution: USA (nationwide) including states of: AZ, AR, CA, FL, GA, IN, KS, KY, LA, MD, MA, MI, MO, NJ, NY, NC, OK, PA, TN, TX, UT, and WI ; and countries of: ARGENTINA, AUSTRALIA, BELGIUM SPAIN, CANADA, CHINA, DENMARK, EMIRATES, ENGLAND, FRANCE, GERMANY, GREECE, IRELAND, MALTA, POLAND, SAUDI ARABIA, SOUTH KOREA, and TAIWAN. | Terminated | April 09, 2013 | 564 | 184 | Computer
Z-0177-2012 | 59425 | No | KPT | Radionuclide dose calibrator. | Radiology | Radiology | 510(k) | Capintec CRC 25 PET, item number 5130-3217 with Software Revisions 2.01 and 2.02 |  |  | July 08, 2011 | 2011 | Capintec Inc | Software error in Rev. 2.01 and Rev 2.02 for Capintec CRC 25PET | Software design | Capintec, Inc. sent a Software Error Alert letter dated July 8, 2011, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to contac their Capintec Customer Support Representative to receive a software upgrade. For questions call Capintec Customer Support at 1-800-631-3826. |  | 69 | 69 | Worldwide distribution - United States including CA, FL, IL, MA, NJ, OK, PA, and VA; and internationally to Brazil, China, India, Indonesia, Italy, Japan, Korea, South Korea, and Turkey. | Terminated | December 10, 2013 | 886 | 69 | Computer
Z-0175-2012 | 59996 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Stryker Mantis Rod Inserter Assembly Manufactured by: Stryker Spine SAS Z. I. Marticot - 33610 - Cestas France; Distributed in the USA by: Stryker Spine 2 Pearl Court. Allendale, NJ 07401-1677 USA |  |  | June 24, 2011 | 2011 | Stryker Spine | There have been reports of difficulty in rod rotation using the newly designed Mantis Rod Inserter. | Device Design | Stryker notified branches/agencies, surgeons and hospitals beginning June 24, 2011 with an Urgent Product Recall letter, dated June 24, 2011, via Fed Ex. The letter identified the affected product, the issue, potential hazards, risk mitigations, and the return process. Customers were asked to examine their inventory and reconcile any inserters and shafts with the catalog numbers provided using the Customer Response Form. Customers are to fax the form to Regulatory Compliance at 201-760-8370 within 5 days of receipt of the letter. The affected product should be retrieved and returned. Questions should be directed to 201-760-8298. |  | 1138 | 471 units US, 667 units international | Nationwide Distribution | Terminated | November 13, 2012 | 508 | 1138 | Computer
Z-0176-2012 | 59684 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | BD Microtainer MAP Microtube for Automated Process with K2EDTA. |  |  | February 11, 2011 | 2011 | Becton Dickinson &amp; Company | BD Microtainer MAP Microtube is being recalled due to the potential clotting of blood specimens resulting from a reduced amount of EDTA additive. | Device Design | A Becton Dickinson Sales Consultant conducted a "on-site" visit to the single Customer account, Florida Hospital, on February 11, 2011. All units were successfully recovered. All product recovered was quarantined and destroyed. Becton Dickinson believes that all reasonable efforts to communicate the issue to the affected consignee have been completed. If you have any questions or request additional information please do not hesitate to call (201) 847-5636. |  | 10800 | 10,800 units | Distributed only to Florida. | Terminated | November 10, 2011 | 272 | 10800 | Computer
Z-0294-2012 | 60031 | scale, stand-on, patient | FRI | Stand-on patient scale. | General Hospital | General Hospital | 510(K) Exempt | Health o Meter Professional Remote Display Digital Physician Scale, Model 498KL; digital patient weighing scale; Pelstar LLC, 11800 South Austin Avenue, Unit B, Alsip, IL 60803, U.S.A., Made n China; Model 498KL |  |  | October 07, 2011 | 2011 | Pelstar, Llc | There is a potential for an electrical short when installing or replacing the batteries in the remote display head of the scales. This short may cause the batteries to smoke, burn or burst. | Device Design | Pelstar LLC. sent a Urgent Device Recall letter dated October 7, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letter informed them of the potential of an electrical short caused when installing or replacing the batteries which may result in minor injuries to the person installing the batteries. The customers were requested to immediately examine their inventory and quarantine all product subject to the recall. In addition, if they distributed this product to end users, they were requested to identify all customers and notify them of this product recall at once in writing, using the enclosed recall notification letter. The customers were also requested to complete and return the enclosed mandatory response form as soon as possible. Any questions were directed to Health O Meter Professional at 800-815-6615 or 708-377-0600. |  | 2498 | 2,498 units | Worldwide Distribution -- USA (nationwide) including the countries of Canada, New Zealand, Israel and Lebanon. | Open, Classified |  |  | 2498 | Computer
Z-0293-2012 | 60304 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Ingenia 1.5T and 3.0T MRI Systems, Model Numbers: 781396 and 781377, 510(k) #K110151. System units equipped with software version: 4.1.1 are subject to recall/field correction. |  |  | August 03, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | The firm determined that during scanning procedures, for some larger patients, some reddening of the skin of can occur in certain instances due to RF heating when the torso of the patient is positioned within the body coil. | Software design | Philips Medical Systems ( Cleveland) Inc. sent a Urgent Field Safety Notice dated August 3, 2011, to all affected customers. Philips Field Service Engineers were dispatched to the affected customer sites to deliver the recall notification and to perform the necessary software corrections. The letter lists the affected MRI systems; describes the problem; the actions to be taken by the customer/user; it also informs the customer of the upcoming software correction which will be applied to each unit by the Philips Field Service Engineers in order to address this issue. If any additional information is required, the customer is instructed to contact their local Philips Field Service Engineer. For further questions please call Customer Care Solutions Center (CCSC) at 800-722-9377. |  | 54 | 26 units - 12 units distributed in U.S. &amp; 16 distributed to foreign customers | Worldwide Distribution -- USA (nationwide) including the states of AL, AZ, CA, GA, OR, MI, NJ and TX., and the countries of AU, BE, CH, DE, FR, JP, NL and SE. | Terminated | February 22, 2013 | 569 | 54 | Computer
Z-0991-2011 | 57522 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | EndoPlege Sinus Catheter Kit, REF EP, 9 Fr x 18.9", Contains Phthalates, Sterile EO, Sterile R, Edwards Lifesciences LLC, Irvine, CA 92614. |  |  | December 03, 2010 | 2010 | Edwards Lifesciences Llc | Defective product may result in the silicone valve becoming dislodged from the introducer which can result in injury. | Device Design | The firm, Edwards Lifesciences, sent an "URGENT PRODUCT RECALL - ACTION REQUIRED" letter dated December 3, 2010 to customers. The letter describes the product, problem and actions to be taken by the customers. The customers were instructed to check their inventory and return the unused devices; complete and return the INTROCSC Introducer Recall Response form to the attention of Customer Service, Edwards LifeSciences, One Edwards Way, Irvine, CA at 949-250-3489. If you have any questions, please contact Edwards Customer Service at 800-424-3278 from 8:00am - 4:30pm Pacific Time. |  | 3427 | 3,427 kits | Nationwide Distribution | Terminated | October 27, 2011 | 328 | 3517 | Computer
Z-0359-2012 | 60409 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | RayStation Version 2.0.0.15 Consult instruction for use. RaySearch Laboratories AB Sveavagen 25, SE-111 34 Stockholm Sweden. support@raysearchlabs.com |  |  | October 19, 2011 | 2011 | RAYSEARCH LABORATORIES AB | The firm has found an issue with DICOM RT Image representation of DRR images from RayStation 2.0 which needs to be clarified in the user manual. The symptoms involve the Pixel Spacing and Image Position attributes in the RT images being incorrect, thus enabling the user to make measurements that are erroneous and possibly misalign the patient. | Device Design | The firm, RaySearch Laboratories, sent an email with an "Urgent Field Safety Notice Medical Device Correction" notice dated October 19, 2011 to its customers. The notice describes the product, problem and actions to be taken. The customers were instructed to pass on the information in the field safety notice to everyone involved and create routines to compensate for the incorrect DRR images; if the DICOM DRR images are not used to make any form of absolute measurements, no action needs to be taken. Raysearch laboratories will provide a script to update these DICOM files to have the correct values upon request.For further information please contact RaySearch Laboratories AB Quality and Regulatory Affairs Manager at +46 722 366 110 or david.hedfors@raysearchlabs.com; RaySearch Americas Director of Customer Support at +1 877 778 3489 or freddie.cardell@raysearchlabs.com. |  | 19 | 19 units | Worldwide distribution: USA (nationwide) including states of: MA, NJ, and NY; and country of: Netherlands | Terminated | May 11, 2012 | 205 | 19 | Computer
Z-0106-2012 | 58984 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | "***Assembled for Invivo Corporation 12501 Research Parkway Orlando FL 32826***Made in U.S.A.***MRI Patient Monitoring System Model 865214***Use only with Model 989803169201 Power Adapter***REF 453564155341***SN US11401652***". |  |  | August 30, 2010 | 2010 | Invivo Corporation | Philips Invivo Corporation of Orlando, FL is recalling their Expression MRI Patient Monitoring System. During transport of the system's cart with a docked Display Controller Unit (DCU), a front caster (wheel) may become loose and fall off causing the cart to become unbalanced and fall. | Device Design | Philips Invivo Corporation issued a Field Safety Notice "URGENT - MEDICAL DEVICE RECALL" letter to all affected customers notifying them that the Expression MRI Patient Monitoring System could pose a risk for patients and users. The notice identifies the following: what the problem is and under what circumstances it can occur; the actions that should be taken by the user in order to prevent risks for patients or users and the actions planned by Philips Invivo to correct the problem. This notice contains important information for the continued safe and proper use of the equipment and customers were advised to retain a copy with the equipment Instruction for Use. Customers were instructed to contact Philips Invivo Customer Support at (877) 468-4861 (option 1, then option 3) for questions or support concerning this issue. |  | 285 | 285 units | Worldwide Distribution - USA (nationwide) including the states of: AK, AL, AR, AZ, CA, CO, CT, FL, GA, HI, ID, IL, IN, KS, KY, LA, MA, ME, MD, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NM, NY, OH, OK, OR, PA, RI., SC, SD, TN, TX, VA, WA, and WI and the countries of: Australia, Austria, Canada, Denmark, Germany, Israel, Malaysia, Netherlands, Norway, Singapore, Spain, Switzerland and, United Kingdom. | Terminated | November 15, 2012 | 808 | 285 | Computer
Z-0320-2012 | 60250 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Elekta Synergy XVI R.4.5 and R4.6 |  |  | October 14, 2011 | 2011 | Elekta, Inc. | The software did not operate as expected. | Software design | Elekta sent an IMPORTANT NOTICE letter dated October 14, 2011 to all affected consignees. The letter identified the affected product, the problem, and the actions to be taken. This notice contained precautionary measures that are required for the operation of the equipment. The notice advise all customers to follow any instructions or recommendations covered in the Notice. The notice instructs Users to file this document in the Important Notice section of the appropriate User Manual. For questions contact your local Elekta representative. |  | 77 | 77 | Nationwide Distribution (USA) - including the states of: AZ, CA, CT, GA, ID, IL, LA, MD, MA, MI, MN, MO, NJ, NY, NC, ND, OH, OK, OR, PA, RI, TN, VA, WA, WI and Puerto Rico | Terminated | November 03, 2014 | 1116 | 77 | Computer
Z-0332-2012 | 60238 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Head Ring Posts with part number 970.280 - re-usable components of the Frame Array Module (of the Optical Guidance Platform and Floorstand devices. |  |  | October 04, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | An anomaly has been identified with the Head Ring posts used by both the Optical Guidance Platform FrameArray module and the Floorstand where the head ring posts may be damaged due to excessive mechanical stress resulting in possible failure during usage. | Device Design | Varian Medical Systems sent an "URGENT MEDICAL DEVICE CORRECTION/URGENT FIELD SAFETY NOTICE" letter dated October 4, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. Customers are advised to verify that they are using the current design of the head ring posts (part # 970.284) and to contact Varian if replacements are needed. Customers may contact the Varian Help Desk for questions regarding this notice at 1-888-827-4265. |  | 820 | 205 kits ( Each kit contains 4 Head Post Rings) 820 HPR's | Worldwide Distribution. | Terminated | July 25, 2012 | 295 | 820 | Computer
Z-0335-2012 | 60228 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance 64 Computed Tomography X-Ray System, 510(k) #K033326, Model Number: 728231, Serial Numbers: 9745, 90169, and 95414. The recalled device systems are equipped with Version 3.5 Beta 1 software. |  |  | September 22, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | Software errors. A number of non-conformances were found for the Brilliance Big Bore and Brilliance 64 Computed Tomography X-ray Systems equipped with software version; 3.5 Beta1. The non-conformances stem from software errors identified in the operating software. | Software design | Philips Healthcare released a 'Field Change Order' (FCO #72800531) on September 12, 2011. Philips sent an - URGENT Medical Device Correction letter dated October 17, 2011, to all affected customers. The letter listed the affected device systems; described the nature of the software problem; and any possible hazard inherent to this situation. The letter informed customers on how to identify the affected products and what action to take. In this instance, the customer was asked to review the Customer Release Notes which were provided to each customer location once the software update is installed. The letter notified the customer that a Philips representative would be installing a software upgrade in order to address the referenced software issues free-of-charge, through a Field Change Order. Customers were instructed to contact their Philips representative or local Philips Healthcare Office for further information or support for this issue. For North America and Canada customers were instructed to contact the Customer Care Solutions Center (1-800-722-9377, option 5: Enter Site ID or follow the prompts). |  | 3 | three (3) units were distributed. | Worldwide Distribution - USA including Texas and the countries of Italy, India and China. | Terminated | February 22, 2013 | 519 | 10 | Computer
Z-0363-2012 | 60458 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | SoftPath GUI versions: 4.2.3.3-4.2.3.6 and 4.3.0.5-4.3.0.12. |  |  | July 20, 2011 | 2011 | SCC Soft Computer | On 07/20/2011 SCC Soft Computer, Clearwater, FL, initiated a correction on the following SoftPath GUI versions: 4.2.3.3-4.2.3.6 and 4.3.0.5-4.3.0.12. A client reported that the HPV results on a patient report were incorrect. The HPV results were from the previous order on the patient and not the current order. All affected clients were notified about the issue on 07/20/2011. | Software design | SCC Soft Computer sent a "CORRECTION COMMUNICATION" dated July 20, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. SCC personnel will review the customized Lab Results AP macro and/or the Micro Results AP macro on customer systems and will either discuss with them the irregularities requiring change or reconfigure macros as necessary. |  | 97 | 97 | Worldwide Distribution-USA (nationwide) and the country of Canada. | Terminated | May 21, 2013 | 671 | 97 | Computer
Z-0355-2012 | 60451 | No | LNX |  |  | General Hospital | 510(k) | SoftPath GUI Versions: 4.3.0.10, 4.3.0.11, and 4.3.0.12. |  |  | June 13, 2011 | 2011 | SCC Soft Computer | On 06/13/2011, SCC Soft Computer initiated a correction on the following SoftPath GUI Versions: 4.3.0.10, 4.3.0.11, and 4.3.0.12. A client reported when they verified the contents of the fax reports, they found one of the specimen diagnoses to be missing. Parts A through D were present as well as Part F, however, Part E diagnosis was missing for one of the patient reports. All affected clients wer | Device Design | SCC Soft Computer initiated a Field Correction on June 13, 2011. All customers with affected software have been notified of this issue.SCC recommends until the corrected patch is loaded to temporarily use one of the following alternative working solutions: 1. Use MS Word Editor functionality within SoftPath application for creating patient reports. In order to ensure Text Control editor is not inadvertently activated, the templates for Text Control editor will be deleted and reinstalled at the clients request by their SCC representative after receipt of the correction. 2. In Report Template set-up, the bottom margin should be made slightly larger than the size of the footer plus its distance from the bottom page border. For further explanation, please consult your SCC representative for further details, or call (727) 789-0100. |  | 14 | 14 | Class II Recall - Worldwide Distribution -- USA (nationwide) and the country of Canada. | Terminated | May 31, 2012 | 353 | 14 | Computer
Z-1249-2011 | 56182 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Radiotherapy delivery system (C-Series Clinac, Trilogy and Novalis Tx), model numbers H14, H17, H18, H23, H26, H29, HCX |  |  | June 23, 2010 | 2010 | Varian Medical Systems Oncology Systems | Position errors may result in misdelivery of radiation. | Device Design | Varian sent Urgent Medical Device Correction letters on June 23, 2010 by Federal Express to all customers. ***12-20-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.***Field Safety notification including customer distribution lists, was e-mailed to Varian international Customer Support Managers. Managers will translate to local language if necessary, notify sites by appropriate methods. |  | 5741 | 5741 units | Worldwide distribution. | Terminated | May 03, 2011 | 314 | 5741 | Computer
Z-0348-2012 | 60237 | No | HWE | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | Orthopedic | 510(K) Exempt | Synthes Oscillating Saw Attachment/Large w/Key for Small Battery Drive. |  |  | October 17, 2011 | 2011 | Synthes USA (HQ), Inc. | The potential exists for the Large Oscillating Saw Attachment to disengage from the Small Battery Drive handpiece. If the attachment disengages from the handpiece, there is a potential for injury to the patient and/or user. | Device Design | Synthes sent a "NOTICE: MEDICAL DEVICE RECALL" letter dated October 17, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. The firm requests that customers return their affected device(s) for servicing. Instructions are provided for customers to follow. Customers having questions regarding this notice should contact the firm at 610-719-5450. |  | 5 | 5 | Nationwide Distribution including the states of CO, LA, NY, and OK. | Terminated | December 18, 2014 | 1158 | 5 | Computer
Z-0305-2012 | 60121 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Elekta Synergy XVI 4.5 |  |  | March 09, 2011 | 2011 | Elekta, Inc. | If "Confirm" or "Unconfirm" settings are clicked during the "Terminated Checking" phase of treatment delivery, Desktop Pro cancels the loaded beam and sets the Beam MUs to zero. Desktop Pro then sends the delivered MUs to MOSAIQ a a zero value. | Software design | Elekta sent an IMPORTANT NOTICE letter dated March 9, 2011 to affected customers. The notice identified the product, the problem, and the corrective action to be taken by the customer. The notice advise all customers to follow any instructions or recommendations covered in the Notice. For questions call +44(0)1293 654200. |  | 56 | 56 | Nationwide Distribution (USA) including the states of: AZ, CA, GA, IL, LA, MD, MA, MI, MN, NJ, NY, NC, OH, PA, TN, VA, WI and Puerto Rico | Terminated | August 09, 2012 | 519 | 56 | Computer
Z-0362-2012 | 60416 | No | LNX |  |  | General Hospital | 510(k) | SoftReports version 1.1.8.2.4 with Hot Fixes 1.7199.7 or 1.7200.7. |  |  | October 25, 2011 | 2011 | SCC Soft Computer | On 10/25/2011 SCC Soft Computer, Clearwater, FL initiated a correction on the SoftReports version 1.1.8.2.4 with Hot Fixes 1.7199.7 or 1.7200.7. When SoftReports Designer opened a layout, the system did not load the layout with the previously saved value in the NewPage property. This caused multiple patient records to be printed on the same page. All affected clients were notified on the issue on | Software design | SCC Soft Computer sent a "Correction Notification" letter dated October 25, 2011 to all affected customers. The letter described the product, problem, and actions to be taken by the customers. The hot fixes that were previously loaded which introduced this issue were corrected and delivered to all affected customers as a mandatory correction in hot fixes 1.7199.8 and 1.7200.8 and higher. |  | 5 | 5 | Nationwide Distribution. | Terminated | March 16, 2012 | 143 | 5 | Computer
Z-0364-2012 | 60411 | [6 empty] | TrueBeam and TrueBeam STx, Model Number: H19, Ref/FSCA identifier: CP-06381 are intended to provide stereotactic radiosurgery and precision radiotherapy for lesions, tumors, and conditions anywhere in the body where radiation treatment is indicated. |  |  | October 31, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | An anomaly has been identified with the TrueBeam and TrueBeam STx systems where, under certain tuning conditions, the electron beam emerging from the bend magnet may have an elongated spot shape. | Device Design | On 11/04/2011, an Urgent Medical Device Correction Letter was distributed to all affected users, with a description of the problem and user corrective action steps. Varian is developing a corrective action for this issue. A customer service representative will contact you when the correction is available to schedule installation on your system. Questions regarding the letter should be directed to the Varian Oncology Help Desk Contact Information at USA and Canada: 1-888-827- 4265 or Europe +41-41-749-8844. |  |  |  | Worldwide distribution including the US and Internationally. | Terminated | July 25, 2012 | 268 |  | Computer
Z-0360-2012 | 59886 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance 64 Computed Tomography System, Model #728231, 510(k) #K 033326, Serial Numbers: 95504, 91003, 300030, 95130, and 95157. |  |  | September 06, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | Software errors. A number of non-conformances were logged during in-house testing and external validation testing of Brilliance 64 and Ingenuity CT software version 4.0 Beta 2. The non-conformances are due to software errors identified in these device systems. | Software design | Philips Healthcare released a 'Field Change Order' (FCO #72800542) on September 15, 2011. Philips sent an URGENT - Medical Device Correction letter dated October 10, 2011, to all affected customers. The letter listed the affected device systems; described the nature of the software problem; and any possible hazard inherent to the situation. The letter was mailed by Philips to their customers on October 27, 2011. The letter provided information telling the customer how to identify the affected products and what action to take. In this instance, the customer was asked to review the Customer Release Notes which were provided to each customer location once the software update was installed. The letter notified the customer that a Philips representative would be installing a software upgrade in order to address the referenced software issues free-of-charge, through a Field Change Order. The letter closes by providing information which all customers in Canada and North America can use to contact the firm through a toll-free telephone number 1-800-722-9377, option 5; enter side ID or follow the prompts.For questions regarding this recall call 440-483-4918. |  | 5 | 5 units | Worldwide Distribution - USA including Ohio and the countries of Australia, China, Belgium, Israel, and Sweden | Terminated | March 15, 2013 | 556 | 7 | Computer
Z-0083-2012 | 52105 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Instrument Manager -Software Versions 8.05, 8.06, 8.07, 8.08, and 8.09. |  |  | May 11, 2009 | 2009 | Data Innovations, Inc. | The "previous result" retrieved from the Specimen Management Database may not be the most recent "previous result". | Software design | Data Innovations sent an "IMPORTANT PRODUCT NOTICE" dated May 11, 2009 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter offers temporary workarounds for customers to use until corrected versions of the software is complete. Versions of the corrections will be provided on a staggered schedule. Customers will be notified by e-mail. Questions regarding this notice are directed to Support at 802-658-1955. |  | 188 | 188 | USA (nationwide) and the country of Canada. | Terminated | October 25, 2011 | 897 | 188 | Computer
Z-0383-2012 | 34843 | No | JEY | Bone plate. | Dental | Dental | 510(k) | Product is 40mm Flexible Tube. ***NON-STERILE PRODUCT***CAUTION: FEDERAL LAW RESTRICTS THIS DEVICE TO SALE BY OR ON THE ORDER OF A PHYSICIAN." |  |  | February 17, 2006 | 2006 | Walter Lorenz Surgical, Inc. | Lorenz Engineer discovered flex tube could break if bent too much. | Device Design | W. Lorenz Surgical notified consignees of recall by an Urgent Advisory Notice letter. Consignees were requested to acknowledge receipt of recall and provide product disposition. Initial letter was sent on March 2, 2006, March 15, 2006, follow-up letters sent on April 4, 2006 and May 15, 2006. For additional information call (904) 741-4400 or fax (904) 741-4500. |  | 76 | 76 pieces | Worldwide Distribution - USA (nationwide) including the states of: GA and TX and the country of Japan. | Terminated | December 12, 2011 | 2124 | 182 | Computer
Z-0079-2012 | 59948 | No | DXS | Cardiopulmonary bypass coronary pressure gauge. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Intersept Custom Tubing Pack, which are designed as specified by the user. The following models contain a Pressure Display Set. Sterilized by Ethylene Oxide. Non-pyrogenic. Manufacturer: Medtronic, Inc. 710 Medtronic Parkway, Minneapolis, MN 55432-5604. |  |  | September 09, 2011 | 2011 | Medtronic Perfusion Systems | Pressure Display Set (models 61000, 61006, 62000, 62006) DFU will be updated to clarify that the intended use of the disposables is only for cardiopulmonary bypass surgery (procedures up to 6 hours). Use beyond 6 hours may result in air leakage in the system resulting in a pressure reading that is not representative of the actual line pressure. | Device Design | Medtronic Perfusion Systems, sent an "Urgent Medical Device Notification letter dated September 9, 2011, via UPS 2-day delivery to all affected customers. The letter identified the affected devices, described the issues, and the actions needed to be taken. The letter indicates when approved, the updated Directions for Use for the Pressure Display Box and Pressure Display Sets will be sent to customers. Customers were asked to complete the Customer Confirmation Certificate and return to Medtronic per the instructions on the form. The letter states that if the recalled products have been further distributed, forward this letter immediately to all customers who received the affected products, and gather proof of notification. For further questions concerning this notification, contact Medtronic Cardiovascular Technical Support at 877-526-7890 or Medtronic Sales Representative. |  | 74804 | 74,804 | Worldwide Distribution - USA (nationwide) including Puerto Rico, and Virgin Islands (U.S.) and the countries of: Argentina, Australia, Austria, Belarus, Belgium, Bulgaria, Canada, Chile, China, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, El Salvador, France, Georgia, Germany, Greece, Hong Kong, Hungary, Iceland, India, Indonesia, Israel, Italy, Jamaica, Japan, Korea, Kuwait, Latvia, Lebanon, Libya, Malaysia, Netherlands, New Zealand, Nicaragua, Norway, Pakistan, Poland, Portugal, Romania, Saudi Arabia, Singapore, Slovenia, Spain, Sweden, Switzerland, Taiwan, United Arab Emirates, United Kingdom, Venezuela, and Yemen. | Terminated | October 11, 2012 | 398 | 270088 | Computer
Z-0388-2012 | 60418 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Vitrea (available in various configuration as Vitrea Enterprise Suite). Vitrea Enterprise Suite 1.2, Vitrea Enterprise Suite 1.3, Vitrea enterprise Suite 6.0 and Vitrea Enterprise Suite 6.1. |  |  | November 11, 2011 | 2011 | Vital Images, Inc. | Vital Images has found two issues that may display incorrect Standard Update Values (SUV) when viewing certain PET/CT images. This issue affects the VitreaCore PET/CT feature for Vitrea Enterprise Suite 1.2, 1.3, 6.0, 6.1, including all updates to those versions. | Software design | Vital Images, Inc. sent an URGENT DEVICE RECALL letter dated November 11, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to inform all potential users of the issues. Customers were to be contacted by their Service Account Manager to discuss and schedule resolutions.For questions customers were to contact Vital Images Customer Support with any questions toll free at 1-800-208-3005, in Europe +31 (0) 70 4135801, in Asia t +86 10 8588 1800, or online at support@vitalimages.com. |  | 338 | 338 (233 US) | Worldwide Distribution - USA including AL, AK, AZ, AR, CA, CO, CT, DC, FL, GA, IL, IN, IA, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NJ, NM, NY, NC, ND, OH, OK OR, PA, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, and WY and the countries of IRELAND, FRANCE, UNITED KINGDOM, SWITZERLAND, AUSTRALIA, SLOVENIA, GERMANY, NETHERLANDS, BELGIUM, PAKISTAN, BRAZIL, ARGENTINA, SWEDEN, QATAR, CHINA, FINLAND, ITALY, DENMARK, PORTUGAL, REPUBLIC OF KOREA, SPAIN, and CANADA | Terminated | December 03, 2013 | 753 | 338 | Computer
Z-0344-2012 | 60302 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Philips GEMINI TF 64 Diagnostic CT X-ray/PET Imaging System, Catalog/Model #882471, 510(k) #K052640. System units equipped with software version: 3.5.2 are subject to recall/field correction. |  |  | April 25, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | In response to complaints received by Philips Medical concerning certain identified software errors reported by their customers, the firm is conducting a recall/software correction to several of their GEMINI PET/CT X-ray system units in the field. The firm believes that if these software defects were to recur, it could impact the performance of reliability of the these PET/CT system units. | Software design | Philips Healthcare sent a "CPE CUSTOMER LETTER" dated October 28, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. Additionally, device operation information is provided to customers. A Philips Field Service Engineer will be visiting each customer site to install the necessary software upgrades. Contact the Customer Care Solutions Center at 1-800-722-9377 for questions concerning this notice. |  | 10 | 10 system units were distributed. | Nationwide Distribution including the states of CO, LA, MD, MI, MN, PA, TX, VA, VT, and WI. | Terminated | February 22, 2013 | 669 | 38 | Computer
Z-0330-2012 | 60231 | No | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | Varian brand GammaMed 12i Accessory: Clamping adapters for bronchial catheters, 1.8mm diameter, GammaMed 12i; Ref: BT-00264, Model 2-H63; Product is distributed by Varian Medical Systems Inc., Palo Alto, CA and Manufactured by Varian Medical Systems Haan GmbH, |  |  | October 06, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | The Product Models 1 &amp; 2 have a design flaw problem where the clamping adapter canobstruct movement of the source cable in the event the bronchial catheter is dislodgedfrom the connector. | Device Design | Varian Medical Systems sent an "URGENT MEDICAL DEVICE CORRECTION LETTER/FIELD SAFETY NOTICE" letter dated October 6, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. Contact the Varian Help Desk at 1-888-827-4265 for questions regarding this notice. |  |  |  | Worldwide Distribution--USA (nationwide). | Terminated | October 22, 2012 | 382 |  | Computer
Z-0954-2011 | 57498 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Covidien, LigaSure, Blunt Tip Laparoscopic Sealer/Divider, 5 mm-37 cm, REF LF1537, Sterile EO, Covidien, Valleylab, a division of Tyco Healthcare Group LP, 5920 Longbow Drive, Boulder, CO 80301 |  |  | December 08, 2010 | 2010 | Covidien LP | During surgical procedures, the device may energize when clamping without pushing the activation button, which could result in unintended energy delivery to tissue. | Device Design | The firm, Covidien, sent a "URGENT: PRODUCT RECALL" letter dated December 6, 2010, to the customers. The letter described the product, problem and the action to be taken by the customers. The customers were instructed to remove product from inventory, immediately quarantine the product and return any remaining stock of the lots listed to: Covidien Energy-based devices, Attention: R82, Building 4/Receiving, 5920 Longbow Drive, Boulder, CO 80301. The customers were also instructed to complete and return the Customer Response Form immediately via mail (in the envelope provided) or fax to (303) 516-8307.To return any identified product for credit or replacement, contact Customer Service at 1-800-722-8772 Option 1.Should you have any medical questions concerning this recall, please contact your local sales representative or the Covidien Clinical information Hotline at (800) 255-8522 (option 1). |  | 25009 | 25,009 units | Worldwide distribution: USA and countries including: Australia, Canada, France, Germany, Japan, Puerto Rico, Singapore, and Uruguay. | Terminated | March 27, 2012 | 475 | 25009 | Computer
Z-0112-2012 | 58454 | No | BYF | Venturi mask. | Anesthesiology | Anesthesiology | 510(K) Exempt | Mask adult, elastic strap style, Percent-O2 Lock air entrainment system, w/7' tube (Linde) - 50/cs; |  |  | March 25, 2011 | 2011 | Salter Laboratories, Division of Regulatory Affairs | A small number of oxygen mask components were found to have 100% obstructed airways rendering them incapable of delivering oxygen to patients undergoing respiration therapy. | Equipment maintenance | The firm, Salter Labs, called all their customers on March 25, 2011 and asked that they hold all affected product. In addition, the firm sent an "URGENT PRODUCT RECALL" letter dated April 4, 2011 and a follow-up letter dated April 7, 2011 via e-mail and certified mail to all customers. The letter described the product, problem and actions to be taken. The customers were instructed to examine their inventory, return any affected lots listed, contact Salter Customer Service at 1-800-235-4203 to arrange for return of affected product and complete and return the Customer Reply Form via 661-854-6816 or toll free 1-800-628-4690.Should you have any questions, please contact Salter Labs at 1-800-235-4203. |  | 800 | 800 all varieties | Worldwide distribution: USA (nationwide) and countries of: Belgium, Canada, Colombia and Iceland. | Terminated | October 26, 2011 | 215 | 800 | Computer
Z-0139-2012 | 60006 | [6 empty] | Picis Anesthesia Manager, PACU Manager,Critical Care Manager-8.2 Service Pack 4 or higher |  |  | August 12, 2011 | 2011 | Picis Inc. | Periop/Critical Care Manager EHR Software Application displayed one patient's demographic and health information in place of another patient that was in current view. | Software design | Picis Inc. sent a "Safety Notification letter" dated August 12, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The customers were instructed until this service pack is applied, we recommend: * For patients who have been transferred from other facilities within the same VISN, you should discontinue the use of these Picis applications. Manually check the patient's name with that displayed in the Picis patient blue banner to ensure you catch all of these cases. * For all other patients, you may safely continue the use of these Picis applications.For further questions please call (781) 557-3000. |  | 10 | 10 | Nationwide Distribution -- AZ, DE, FL, IL, NY, OH, PA, WI and WV. | Terminated | March 23, 2012 | 224 | 10 | Computer
Z-0987-2011 | 57501 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Physiological Monitoring, Patient Monitor |  |  | December 17, 2010 | 2010 | Philips Healthcare Inc. | Some PAS-201 active speaker assemblies in use with the Philips IntelliVue Information Center (PIIC) have experienced intermittent audio or loss of audio. Speaker failures may delay recognition of an alarm condition, if operators do not recognize that there is no alarm sound from the monitor and the monitor is not connected to another audible alarm annunciation system. | Device Design | Philips notified all affected customers with an Urgent Medical Device Correction notification letter dated December 17, 2010. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to observe their system for any of the described symptoms. If they experience intermittent audio or there is no sound from their PAS-210 speaker while it is not connected to another audible alarm annunciation system, customers were instructed to remove the speaker from use and contact their local Philips service representative. Customers were advised not to rely exclusively on the audible alarm system for patient monitoring, the most reliable method of patient monitoring combines close personal surveillance with correct operation of monitoring equipment.ACTIONS PLANNED BY PHILIPSPhilips will replace affected speaker assemblies free of charge.A Philips Healthcare representative would contact the customer when the replacement was available. Field Service Engineers will remove and replace defective speakers to correct the problem. The defective speakers will be returned to vendor (Beach Wire) to be discarded. This correction will be done free of charge for customers. |  | 1071 | 1071 | Worldwide Distribution - USA including AL, AR, AZ, CA, CO, CT, DC, FL, GA, IA, ID, IL, IN, KS, KY, MA, MD, ME, MI, MN, MO, MS, MT, MS, NC, ND, NE, NH, NJ, NM, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA WI, WV, WY, and Canada | Terminated | April 18, 2017 | 2314 | 1071 | Computer
Z-0304-2012 | 60204 | No | DXN | Noninvasive blood pressure measurement system. | Cardiovascular | Cardiovascular | 510(k) | Ambulatory Blood Pressure (ABP) Monitor, Model 90217A. |  |  | October 12, 2011 | 2011 | Del Mar Reynolds Medical, Ltd. | If external pressure is applied to the cuff of the Ambulatory Blood Pressure, Model 90217A, during deflation, the monitor will stop working and display an EC75 error code. | Device Design | On October 21, 2011, a Medial Device Correction customer letter was mailed to all US customers. A return receipt was requested.The letter identified the affected devices and described the problem. If the device displays an EC75 error code, customers can cycle the power off and then on to restore normal functionality. Please advise patients of this mitigation to avoid a delay in obtaining results. The firm is in the process of upgrading their software to avoid the issue and will contact customers to arrange for an update to their monitors.Customers can call Spacelabs Healthcare at 1-800-522-7025, select 2 for Technical support for any question about this recall. |  | 785 | 785 units worldwide (83 in US and 702 outside US) | Worldwide Distribution | Terminated | May 30, 2012 | 231 | 785 | Computer
Z-0311-2012 | 60146 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Roche Cobas c 111 Analyzer |  |  | October 18, 2011 | 2011 | Roche Diagnostics Operations, Inc. | Internal investigations have revealed a potential printer malfunction on the Cobas c 111 analyzer, where one or more numbers or characters can be missing on the hard copy result printout, or the printer may skip leaving blanks which may result in an incorrect print layout. The missing number of characters can be in any position, therefore it could be in the middle of a result, (e.g. 48.53 mmol in | Software design | Roche Diagnostics Operations, Inc sent a Urgent Medical Device Correction dated October 18, 2011, via UPS Ground (receipt required) to all affected customers and one distribution center, with instructions to disable the Auto Print Result feature until new software is available. All printed results will need to be confirmed against the results indicated on the display screen. Questions are to be directed to Roche Diagnostics Technical Support at 1-800 428-2336, |  | 135 | 135 | Worldwide Distribution - USA including West Virginia and Texas. | Terminated | February 21, 2013 | 492 | 135 | Computer
Z-1221-2011 | 57733 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Hill-Rom TotalCare Bariatric Bed, P1840, Hill-Rom, Batesville, IN. |  |  | February 03, 2011 | 2011 | Hill-Rom, Inc. | The firm received one report that a patient was weighed on the bed at 53 Kg when the patient actually weighed 73 Kg. The patient expired the next day and it is inconclusive as to whether the scale was a contributing factor. The firm investigated and found that the software for the scale system did not clear the last patient weight from the history of the scale system. | Software design | The firm, Hill-Rom, sent a "URGENT MEDICAL DEVICE CORRECTION" letter dated January 19, 2011, along with a software upgrade (USB stick/flash drive) with instructions, to all affected customers. The letter described the product, problem and action to be taken by the customers. The customers were instructed to requests that facility service personnel immediately upgrade their software; complete and return the Customer Response Form via fax to 812-931-2410 or mail to: Hill-Rom Company, 1069 State Route 46 E, Batesville, IN 47006, Attn: MOD Coordinator J12 and forward a copy of this notice to any other facility personnel. Note: The customer can request that Hill-Rom technicians complete the software upgrade.If you have any questions concerning this Urgent Medical Device Correction, please contact Hill-Rom Technical Support at 800-445-3720. |  | 40 | 40 beds. | Worldwide distribution: USA and countries including: AUSTRALIA, BRAZIL, CANADA, CHINA, EGYPT, GERMANY, GREECE, ITALY, JAPAN, KUWAIT, MALAYSIA, MEXICO, NETHERLANDS, OMAN, POLAND, QATAR, SAUDI ARABIA, SPAIN, SWEDEN, THAILAND, UNITED KINGDOM, VENEZUELA, and YEMEN. | Terminated | September 26, 2012 | 601 | 945 | Computer
Z-0310-2012 | 60282 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | cobas c 311 Analyzer, Part Number: 04826876001 |  |  | October 26, 2011 | 2011 | Roche Diagnostics Operations, Inc. | The cobas c 311 analyzer software may incorrectly calculate the volume of the ISE Reference Electrolyte solution remaining on the analyzer. In some cases the incorrect results obtained will not be flagged with any data alarms. | Software design | October 26, 2011 Urgent Medical Device Corrections were sent to all direct accounts in the USA that have installed the cobas c 311 analyzer and included a faxback form to be completed and faxed to 1-877-766-7452. Consignees were instructed to temporarily replace the ISE Reference Electrolyte bottle based upon the frequency calculated from the maximum number of ISE samples run per day and activate the yellow threshold alarm on the system. Customers with question and concerns are directed to call Roche Diagnostics Technical Support at 1-800-428-2366, available 24-hours, 7 days a week. |  | 940 | 940 worldwide | Worldwide Distribution -- USA including the states of Maryland and Texas. | Terminated | July 09, 2012 | 257 | 940 | Computer
Z-0113-2012 | 60025 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Radio used in the Precess MRI Patient Monitoring System. 3160 (Precess) and HE97B, 94044, 94045 (radio and radio kits). |  |  | September 23, 2011 | 2011 | Invivo Corporation | Philips Invivo has determined that it is possible for the system's alarm limits and other device settings to spontaneously change without user intervention. This problem is associated with intermittent and frequent loss of communication between the wireless Processing Unit and the Display Controller Unit. | Device Design | A Field Safety Notice entitled Urgent - Medical Device Correction letter, dated September 2011, with attached Field Safety Notice for customers were sent to distributors. The letter described the reason for the recall and the intention of the Field Safety Notice. Also, the hazards involved and instructions on how to identify units that are affected by the recall were provided . The actions to be taken by the user and the actions planned by the Philips Invivo were explained. Users were to follow the instructions provided on how to safely use the device until the unit can be serviced by Philips Invivo. For further information or support concerning this issue, customers are to contact Philips Invivo Customer Support at (877)-468-4861 or their local Philips Invivo representative. |  | 875 | 875 | Worldwide Distribution -- Australia, Austria, Brazil, Canada, China, France, Germany, Malaysia, Norway, Portugal, Spain, Sweden, The Netherlands, United Arab Emirates, United Kingdom, and US. | Terminated | August 07, 2013 | 684 | 875 | Computer
Z-0963-2011 | 57609 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | SRD S5 Mast Roller Pump System 85, S5 Component, Part No.: 50-80-70, Made in Germany, Distributed by: Sorin Group USA, Inc., Arvada, CO 80004. |  |  | November 12, 2010 | 2010 | Sorin Group USA, Inc. | Touch screen may become unresponsive, inhibiting user input. | Device Design | All affected US customers and distributors were notified by certified mail on Nov 12, 2010 via a Field Correction letter. They were told that they would be contacted by Sorin to arrange an appointment to inspect and replace affected touch screens on site. They were told that they could safely continue using their S5 systems in accordance with the Operator's Manual and previously supplied instructions until the inspection and necessary replacement had been completed. A customer response form was also included for completion and return to the firm. Customers can contact the firm if they have any questions. |  | 16 | 16 units | Worldwide Distribution. | Terminated | July 13, 2012 | 609 | 6671 | Computer
Z-0131-2012 | 59147 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | SYNCHRON CX7 Delta Clinical System Part #: 468205 |  |  | August 23, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed software limitations which affect only manually-loaded reagent cartridges on the CX4/5/7 CE, CX4/5/7 Delta and CX9 ALX Clinical Systems. | Software design | Beckman Coulter sent an "URGENT: PRODUCT CORRECTIVE ACTION" letter dated September 20, 2010 with an attached Fax Back Response Form (via US Postal Service for US Customers) to all customers who purchased the SynchronÂż CX4/5/7 CE, CX4/5/7 Delta and CX9 ALX Clinical Systems. The letter provides the customers wtih an explanation of the problem identified and actions to be taken. Customers were instructed to share this information with laboratory staff and retain this notification as part of their laboratory Quality System documentation. If they forwarded the affected product listed above to another laboratory, they were asked to provide a copy of this letter to them. Questions regarding this Product Corrective Action letter, were directed to Customer Support Center (Call Center) on the Beckman Coulter website athttp://www.beckmancoulter.com/customersupport/support or call (800) 854-3633 in the United States and Canada. Outside the United States and Canada, they were directed to their local Beckman Coulter Representative. |  | 1070 | 1070 systems total in US | Worldwide Distribution - USA (nationwide) and the country of Canada | Terminated | June 28, 2012 | 675 | 1070 | Computer
Z-0074-2012 | 59933 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | CareFusion EnVe Ventilator, Model 19250-001. Manufactured by: CareFusion, 22745 Savi Ranch, Yorba Linda, CA 92887 USA. EnVe is a portable critical care ventilator. It is suitable for a wide range of life support and critical care applications. The ventilator normally operates from external DC power. When the ventilator is portable, a Removable Battery Pack powers the unit. |  |  | September 12, 2011 | 2011 | CareFusion 203, Inc. | CareFusion has identified potential risks associated with the EnVe ventilator manufactured between December 2010 and May 2011. There are 3 issues that can result in ventilation delivery to the patient being interrupted. CareFusion will be contacting facilities to coordinate implementation of the necessary corrective actions. | Device Design | The firm, CareFusion, sent an "URGENT PRODUCT RECALL" letter dated September 12, 2011 to its customers.The letter identified affected product, described the 3 issues providing the description of the event, potential risk, and provided required action for Users. The letter also stated that CareFusion would be contacting the facility to coordinate the corrective action.If you have any questions and/or need support, please contact CareFusion Technical Support at 1-800-554-8933 from Mon. - Fri. 8:00am -5:00pm Pacific Time. |  | 128 | 128 units | Nationwide distribution: USA including states of: CA, CO, FL, IL, MI, MN, MT, ND, NJ, NM, NY, OK, PA, TX, UT, WI, and WV. | Terminated | November 22, 2012 | 437 | 128 | Computer
Z-0331-2012 | 60229 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Model CAT-850B Catheterization Table, a component of the Toshiba INFX-8000 series X-Ray Interventional System. |  |  | October 03, 2011 | 2011 | Toshiba American Medical Systems Inc | The recall was initiated because Toshiba has confirmed that during patient transfer from a gurney to the catherization table, the tabletop may rotate unexpectedly when sufficient force is applied to the tabletop. The result of the tabletop rotating unexpectedly is the patient being tranferred may fall from the tabletop. | Device Design | Toshiba America Medical Systems sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated October 3, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. Customers were instructed to follow the user instructions by positioning the tabletop at the foot-end position, and ensuringsufficient assistance is available, when transferring a patient from the gurney to the CAT-850B catheterization table. Customers were also instructed to follow the instructions for transferring a patient to the CAT-850B that are included in the operation manual. A Customer Reply Form was included with the letter for customers to complete and return via fax to (877) 349-3054. The customer reply form can also be sent via e-mail to raffairs@tams.com. Questions regarding the recall letter can contact the Director of Regulatory Affairs at (800) 421-1968 or contact their local Toshiba Representative at (800) 521-1968. |  | 300 | 300 units | Nationwide Distribution. | Terminated | March 18, 2013 | 532 | 300 | Computer
Z-0337-2012 | 59767 | No | FMZ | Neonatal incubator. | General Hospital | General Hospital | 510(k) | Dual Incu i (Atom Infant Incubator Model 100) |  |  | August 23, 2011 | 2011 | Atom Medical Corporation | The caster with lock on the incubator or infant warmer may become loose or break. | Device Design | The firm, Atom Medical, sent an "Urgent: Field Safety Notice" letter dated August 18, 2011, to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately identify all devices in their inventory and quarantine them until the rework is complete; identify any customers to whom they have further distributed the product; contact their customers at once and arrange a time in which the field service technician can perform the rework (the parts needed to complete the rework will be sent from Atom Medical's export department ); complete rework of all affected devices and document rework on the Distributor FSCA/Recall Response Form and return via Fax to: +81 3 3812-3199 or email to kira-utc@atomed.co.jp; t-sugino@atomed.co.jp or kenichi.shimada@atomed.co.jp (all forms should be completed and returned by January 31, 2012.) If you have any questions, contact the Manager Quality Control or Manager, Regulatory Affairs at email addresses above, phone: +81-3-3815-3632 or fax +81-3-3812-3199. |  | 105 | 105 units | Worldwide distribution: USA (nationwide) including state of: PA and countries of: Algeria, Austria, Australia, China, Egypt, Germany, Italy, Japan, Korea, Kuwait, Malaysia, Portugal, Russia, Saudi Arabia, Spain, Switzerland, Thailand, The Netherlands, and UAE. | Terminated | September 09, 2013 | 748 | 540 | Computer
Z-0345-2012 | 59934 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Odyssey |  |  | November 04, 2010 | 2010 | PerMedics, Inc. | Anomaly 1 - Invalid Export of Beam Table and or Collimator Angle for DICOM RT Ion, DICOM Image and RTP Link export files:During PerMedics' internal code reviews for lEC 61217 compliance verification, an issue was identified by PerMedics engineers in three export files: D1COM RT Ion Plan, D1COM RT Image and RTP Link. In some cases, this issue can result in table and/or collimator angle values t | Device Design | PerMedics sent an Odyssey TPS URGENT NOTICE OF CORRECTION letter dated November 4, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to:Anomaly 1 User Corrective Action:Install the mandatory new version correcting the software as soon as it is available (see PerMedics Corrective Action below). In the meantime, the following recommendations were provided if they planned to use Odyssey RT Ion Plan, DICOM RT Image, and/or RTP Link export:Workaround for the invalid exported table and/or collimator angle errors: If, prior to the software correction, any of these affected file types were exported from Odyssey, manually verify that the table and collimator angles are correct. Additionally, apply appropriate changes to any angles deemed incorrect that were obtained from Odyssey's DICOM RT Ion Plan, DICOM RT Image, and/or RTP Link files.Only those machines with table and/or collimator angles that increase in the clockwise direction will have the incorrect angles.Anomaly 2 User Corrective Action:Install the mandatory new version correcting the software as soon as it is available (see PerMedics Corrective Action below). In the meantime, the following recommendations are provided if the user- Changes the patients scheduled start date prior to the software correction:Workaround for Incorrect Monitor Unit Values when Changing Treatment Start Date: If, prior to the software correction, the user changes the patient's treatment start date on any fully-computed IMRT plan, IMAT plan or plan containing a "Region DVH" prescription, the user should force a full recommendation oft hose patient plans prior to fixing and exporting the plan.The following procedure can be used to force a full re-computation of the plan:- Open the patient's plan worksheet- Change the algorithm type of the plan- Change the algorithm type back to its original value- Close t |  | 15 | 15 | Worldwide Distribution - USA including NY, GA, CA and the countries of Hungary, Brazil, China, India, Italy, and Mexico | Terminated | March 02, 2012 | 484 | 15 | Computer
Z-0115-2012 | 59326 | No | MUE | Full-field digital mammography system. | Radiology | Radiology | 510(k) | GE SENOGRAPHE DS FULL FIELD DIGITAL ESSENTIAL MAMMOGRAPHY SYSTEM. |  |  | September 22, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware of inaccurate measurement on GE magnified images provided by Senographe DS and Essential Mammography systems when reviewed on IHE compliant workstations. This issue may impact patient safety. | Software design | GE Healthcare sent an " URGENT MEDICAL DEVICE CORRECTION" letter dated September 21, 2011 to all affected customers. The letter describes the products, problem, and actions to be taken by the customers. The letter provides safety instructions and advises customers to complete and return a Customer Reply Form via fax to +33(0)1 30 70 39 30. A GE Healthcare service representative will contact customers if the updated software version is required. Contact the Call Center at 800-437-1171. |  | 1344 | 1344 | Worldwide Distribution-USA (nationwide) including Puerto Rico and Washington, D.C. and the countries of ALGERIA, AUSTRALIA, AUSTRIA, BELGIUM BOTSWANA, BRAZIL, BULGARIA, CANADA, CHILE, CHINA, COLOMBIA, CROATIA, CZECH REPUBLIC, DENMARK, EGYPT, FINLAND, FRANCE, GEORGIA,GERMANY, GREECE, GUADELOUPE, GUYANA, HONDURAS, HONG KONG, HUNGARY, INDIA, IRAQ, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KUWAIT, LATVIA, LEBANON, LIBYAN ARAB JAMAHIRIYA, LITHUANIA, LUXEMBOURG, MALAYSIA, MALTA, MARTINIQUE, MEXICO, MOROCCO, NETHERLANDS, NEW ZEALAND, NIGERIA, NORWAY, PAKISTAN, PHILIPPINES, POLAND, PORTUGAL, REPUBLIC OF KOREA, RUSSIAN FEDERATION SAINT PIERRE AND MIQUELON, SAUDI ARABIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SUDAN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, THE FORMER YUGOSLAV REPUBLIC OF MECEDONIA, TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, AND VENEZUELA. | Terminated | October 16, 2012 | 390 | 3457 | Computer
Z-0325-2012 | 59885 | No | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | VariSource iX, Software Version 1.1, Part Number AL20020000 VariSource iX Afterloader. |  |  | March 01, 2010 | 2010 | Varian Medical Systems, Inc. | VariSource iX console software may cause incorrect dwell times to be delivered to patient and may result in a mistreatment. | Software design | Varian Medical Systems sent a "MEDICAL DEVICE RECALL" letter dated April 1, 2010 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. Customers were asked to conduct a simple test to establish if their machine is affected by the problem. Device users were requested to complete an attached receipt verification form and return it to Varian BrachyTherapy via fax or e-mail. Call the Contact Service Dispatch number at 1-800-864-1672 for questions regarding this detail. |  | 3600 | 72 worldwide/50 in U.S. | Worldwide Distribution-USA (nationwide) | Terminated | May 30, 2013 | 1186 | 3600 | Computer
Z-0333-2012 | 60244 | Yes | LWS |  |  | Cardiovascular | PMA | LATITUDEÂż Patient Management System, Models 6441*, 6442*, 6465*, 6488 (USA). |  |  | October 12, 2011 | 2011 | Boston Scientific CRM Corp | Current Boston Scientific defibrillator systems measure shock lead impedances up to 200 ohms, and export results to the programmer and LATITUDE monitoring system via the Daily Measurements feature. A measured shock lead impedance that is greater than 125 ohms or less than 20 ohms will generate an alert message on programmer screens and a Red Alert within LATITUDE (if activated). However, the mea | Software design | Boston Scientific sent an "IMPORTANT DIAGNOSTIC INFORMATION" letter dated October 2011 to all affected customers. The letter described the product, problem and actions to be taken by the customers. |  | 25300 | 25,300 (14,700 US; 10,600 OUS) | Worldwide Distribution-USA (nationwide) including DC and Puerto Rico and the countries of AUSTRIA, BELGIUM, DENMARK, FINLAND, FRANCE, GERMANY, IRELAND, ITALY, NETHERLANDS, NORWAY, PORTUGAL, SPAIN, SWEDEN, SWITZERLAND, UNITED KINGDOM, CANADA, AUSTRALIA, NEW ZEALAND, CYPRUS, CZECH REPUBLIC, GREECE, GUADELOUPE, HUNGARY, BANGLADESH, CHINA, COCOS ISLAND, HONG KONG, INDIA, INDONESIA, KOREA, MALAYSIA, PHILIPPINES, SINGAPORE, SOUTH KOREA, SRI LANKA, THAILAND, RUSSIA, SAUDI ARABIA, SLOVAKIA, SLOVENIA, SOUTH AFRICA, NEW CALEDONIA, POLAND, PORTUGAL, QATAR, REUNION,IRAN, ISRAEL, JORDAN, KUWAIT, LATVIA, LEBANON, LUXEMBOURG, MONACO, MOROCCO, JAPAN, ARGENTINA, ARUBA, BAHAMAS, BARBADOS, BELIZE, BERMUDA, BRAZIL, CAYMAN ISLANDS, CHILE, COLOMBIA, COSTA RICA, DOMINICAN REPUBLIC, ECUADOR, EL SALVADOR, FALKLAND ISLANDS, HAITI, HONDURAS, JAMAICA, MEXICO, PANAMA, PERU, TRINIDAD &amp; Tobago, URUGUAY, VENEZUELA, AND NETHERLANDS ANTILLES. | Terminated | December 04, 2012 | 419 | 25300 | Computer
Z-1605-2011 | 57611 | No | IXL | Spot-film device. | Radiology | Radiology | 510(K) Exempt | Philips BV Pulsera, Rel. 2.3, 718095. |  |  | December 14, 2010 | 2010 | Philips Medical Systems North America Co. Phillips | The fixing strap of the belt holding the CBX Laser to the Image Intensifier may be detached during a procedure when used with the BV Libra and the BV Pulsera/Endura Systems. | Device Design | On 01/06/11, Philips sent an URGENT Device Correction Notice, dated 01/05/2011, to their consignees. The consignees were informed that the fixing strap of the Laser Aiming Device Image Intensifier may inadvertently detach when used with BV Libra, BV Endura Rel. 2, and BV Pulsera Rel 2.3 Mobile X-Ray systems. Specifically the draw latch of the fixing strap holding the CBX Laser to the Image Intensified can be accidentally released during a procedure when used with the above Mobile X-Ray systems. While waiting for Philips to implement their corrective action, the consignees are advised to fix the closed latch body of the draw latch to the metal band using a Tie Wrap. Without securing the latch with a Tie Wrap, the device should not be used.The consignees can contact Philips Call Center at 1-800-722-9377, select 5, and reference FCO 71800035 with any questions about this recall. |  | 214 | 214 units | Nationwide Distribution -- AK, AL, AR, AZ, CA, CO, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NH, NJ, NM, NV, NY, OH, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and Washington, DC. | Terminated | February 28, 2012 | 441 | 227 | Computer
Z-1374-2011 | 56394 | No | LZS |  |  | Ophthalmic | PMA | AMO brand WaveScan WaveFront System, OPM software versions 3.67 and 3.67, |  |  | July 13, 2010 | 2010 | AMO Manufacturing USA, LLC | The failure to comply with Minimum System Requirements, increased system errors could result, and a miscreated treatment table (incorrect LASIK refractive treatment) could result. | Software design | Abbott Medical Optics, Inc. sent an Urgent Medical Device Field Correction letter to all affected customers starting July 15, 2010. The letter identified the product, the problem, and the action to be taken by the customer.Customers were asked to take the following actions:Surgeons were instructed to complete the facsimile attachment to acknowledge receipt of the notification letter and fax to 714-247-4530.For those using the affected product a 3.68 OPM disc must be requested by e-mailing a request to OPM@pacful.com.For any questions please contact AMO Technical Service Representatiave at 1-877-266-4543. |  |  |  | Worldwide Distribution - USA (nationwide) and the countries of Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, China, Columbia, Czech Republic, Denmark, Egypt, Finland, France, Germany, Great Britian, Greece, Hungary, India, Indonesia, Iraq, Ireland, Israel, Italy, Jordan, Kuwait, Lebanon, Malaysia, Mexico, Netherland, New Zealand, Peru, Philippines, Portugal, Russia, Russia Federation, Saudi Arabia, Singapore, South Korea, Spain, Sweden, Taiwan, ThailandChina, India, South Korea, Singapore, Taiwan, Thailand, Turkey, United Arab Emirates, Venezuela, and Vietnam. | Terminated | June 22, 2012 | 710 |  | Computer
Z-2578-2011 | 58200 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation treatment Planning System, XiO Release 4.2.0 and above |  |  | September 17, 2010 | 2010 | Computerized Medical Systems Inc | When a dialog box is displayed and the dose status is changed to "Dose not current", beam edits can be made in the graphics area using the mouse (rotate collimator, reposition collimator, move isocenter, move weight point, etc.). Then if CANCEL is selected on the dialog box, any edits made on the dialog box are undone, the dose is left unchanged, but the beam edits made in the graphics area with t | Software design | The firm, Elekta CMS Software, sent an "IMPORTANT SAFETY NOTICE" dated September 6, 2010 to its affected customers. The notice describes the product, problem and actions to be taken (workaround). The customers were instructed to distribute this notice to any and all users of CMS software at their organization who are potentially affected by this issue. The notice also states that this problem has existed since XiO Release 4.2.0 and will be resolved in XiO Release 4.70.00. Elekta will notify the sites when this solution is available. If you have any questions, please call 314-993-0003 or Toll Free: 800-878-4267. |  | 1293 | 1293 | Worldwide distribution: USA (nationwide) including Puerto Rico, and countries of: Albania, Algeria, Argentina, Australia, Austria, Azerbaijan, Bahamas, Belgium, Bolivia, Bosnia and Herzegovina, Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Ecuador, Egypt, El Salvador, Estonia, Finland, France, Georgia, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Kuwait, Latvia, Libya, Lithuania, Malaysia, Malta, Mexico, Mongolia, Morocco, .Netherlands, New Zealand, Nicaragua, North Korea, Philippines, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Suriname, Switzerland, Taiwan, Tajikistan, Thailand, Turkey, Turkmenistan, Ukraine, United Kingdom, Uruguay, Venezuela, Yemen, and Zimbabwe. | Terminated | April 08, 2014 | 1299 | 1293 | Computer
Z-2573-2011 | 58811 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | The Ultraview SL 91517 Capnography module is a mainstream and sidestream gas analyzer designed to measure the concentration of carbon dioxide in a gas mixture, to aid in determining the patient's ventilatory, circulatory, and metabolic status. |  |  | May 12, 2011 | 2011 | Spacelabs Healthcare, Llc | Capnography Module, Model 91517, resets 20 seconds after power failure, instead of after three minutes as specified. This reset clears the RAM and the module returns to default settings. | Software design | Spacelabs Medical sent a "URGENT-MEDICAL DEVICE CORRECTION" letter dated May 12, 2011 to all affected customers. The letter identifies the product, problem, and actions to be taken. Customers were recommended to set all their Capnography Module Configuration Manager alarm settings to "ON". The letter also recommended that customers brief their staff about the possibility of losing changes to the default settings and to restore them after a power interruption.A field service engineer will be contacting customers to set up a date and time when affected modules can be upgraded at no cost to the customer. For questions about this corrective action program, call Spacelabs Medical at 1-800-522-7025, Select 2 for Technical Support. For outside the US, customers can call 1-425-657-7200, x5089. |  | 5883 | 5883 units | Worldwide Distribution--USA (nationwide) and countries of ALGERIA, ARGENTINA, AUSTRALIA, BAHRAIN, BANGLADESH, BELGIUM, BOLIVIA, BRAZIL, CANADA, CHILE, CHINA,COLOMBIA, CYPRUS, CZECH REPUBLIC, DOMINICAN REPUBLIC, ECUADOR, EGYPT, ESTONIA, FRANCE,GERMANY, GREECE, GUATEMALA, HONG KONG, INDIA, ISRAEL, ITALY,JORDAN, KUWAIT, LEBANON, MALAYSIA, MEXICO, MOROCCO, NEPAL,NETHERLANDS, NEW ZEALAND, NORWAY, PAKISTAN, PALESTINE, PANAMA,PERU, POLAND, PORTUGAL, ROMANIA, SAUDI ARABIA, SINGAPORE, SLOVAKIA, SOUTH AFRICA, SPAIN, SRI LANKA, SWEDEN, SWITZERLAND,SYRIAN ARAB REPUBLIC, TAIWAN, REPUBLIC OF CHINA, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, VENEZUELA, and VIET NAM. | Terminated | December 02, 2011 | 204 | 5883 | Computer
Z-2604-2011 | 58270 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System, XiO Release 4.50.00 and above |  |  | May 19, 2010 | 2010 | Computerized Medical Systems Inc | XiO Software: when a treatment machine is defined in Source File Maintenance (SFM), the MLC Model is set to Elekta, and beams using this treatment machine are added to a new plan in Teletherapy, XiO automatically adds an MLC to the beam. This is because the Elekta linac will always have an MLC in place when the plan is actually delivered. | Software design | A Safety Notice, dated 5/20/10, was sent to Elekta / Computerized Medical Systems Inc clients in May 2010. The notice indicates the problem, the clinical impact. The notice states a future XiO Release will resolve the issue. A return postcard is included with the mailing to verify consignee received the notice. For sites not returning the postcard by 5/28/10, the recalling firm will contact them via phone or email to confirm receipt of the notice. The notice also informed the consignees they would be contacted when a solution was available. |  | 57 | 57 | Nationwide distribution to medical facilities including Puerto Rico. Distribution was also made to two Military and Government consignees. Foreign distribution was made to Australia, Belgium, Brazil, Canada, Colombia, France, India, Ireland, Israel, Italy, Malaysia, Netherlands, Philippines, Poland, Spain, and United Kingdom. | Terminated | April 02, 2014 | 1414 | 57 | Computer
Z-2592-2011 | 58269 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Focal Radiation Treatment Planning System, Focal Release 4.50.00 and above |  |  | May 10, 2010 | 2010 | Computerized Medical Systems Inc | Focal: for XiO plans calculated in relative dose mode, the Focal Plan Review application will display the DVH and maximum dose correctly. However, the isodose display is not scaled correctly. If the user evaluated the plan based only on the isodose display, they might misinterpret the actual dose delivered to the patient. | Software design | Elekta / Computerized Medical Systems Inc sent an undated Safety Notice to consignees beginning on May 21, 2010. The notice identified the product, the problem, and a workaround. |  | 308 | 308 | Worldwide Distribution - USA (Nationwide including Puerto Rico) Foreign distribution was made to Albania, Australia, Austria, Belgium, Canada, Chile, China, Cyprus, Denmark, Egypt, Estonia, Finland, France, Germany, Greece, Hong Kong, Hungary, India, Israel, Italy, Japan, Lithuania, Malaysia, Malta, Mexico, Netherlands, New Zealand, Portugal, Russia, Serbia, Singapore, Slovenia, South Africa, South Korea, Spain, Switzerland, Taiwan, Thailand, and Turkmenistan. | Terminated | April 02, 2014 | 1423 | 308 | Computer
Z-2595-2011 | 58260 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System, XiO Release 4.0.2 and above |  |  | May 28, 2010 | 2010 | Computerized Medical Systems Inc | XiO Software: when Digitally Reconstructed Radiographs (DRRs) are generated in XiO, and DICOM is used to export the DRRs (File-Print-DRR), the exported images are associated with the Beams, but NOT the Plan. The Plan association had always been done manually by the user. This is functionality that has never been provided on XiO. The problem occurs because XiO is not storing a persistent DICOM | Software design | Elekta / Computerized Medical Systems Inc sent an undated Safety Notice to all affected customers. The notice indicated the problem, the clinical impact, and the workaround. The notice stated the sites will be notified when a solution is available.No date was given for the mailing or the solution. |  | 1292 | 1292 | Worldwide Distribution - USA (Nationwide distribution, including Washington, DC and Puerto Rico) and the countries of Albania, Algeria, Argentina, Australia, Austria, Azerbaijan, Bahamas, Belgium, Bolivia, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, El Salvador, Estonia, France, Georgia, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Kuwait, Latvia, Libya, Lithuania, Malaysia, Malta, Mexico, Mongolia, Morocco, Netherlands, New Zealand, Nicaragua, North Korea, Philippines, Poland, Portugal, Qatar, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Tajikistan, Thailand, Turkey, Turkmenistan, Ukraine, United Kingdom, Uruguay, Venezuela, Yemen, and Zimbabwe. | Terminated | April 10, 2014 | 1413 | 1292 | Computer
Z-2599-2011 | 58250 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Focal Radiation Treatment Planning System, Focal Release 4.1.1 through 4.40.00. |  |  | February 06, 2009 | 2009 | Computerized Medical Systems Inc | Focal Software: if individual MLC leaf position edits are made, the original MLC aperture is retained. A subsequent aperture edit should not affect the MLC leaf positions however it was found that instead the leaves will then snap (unintentionally) back to the original aperture. Also, if the patient is sent to XiO after the MLC is edited, an isocenter shift made in XiO will also cause the MLCs to | Software design | Elekta CMS Software sent an "ADVISORY NOTICE" dated February 6, 2009 to its affected customers. The notice identified the product, the problem, and the actions to be taken. Customers were informed that the workaround method available is to (edit the MLC pattern using the aperture first and the individual MLC leaf edits second). The firm states that this issue will be resolved in Focal version 4.50.00 and above. A "User Notice" will be sent to all users still using the affected software notifying them of the updated software available to correct the problem. If you have any questions, please call 314-993-0003 or toll free 800-878-4267. |  | 760 | 760 | Worldwide Distribution -- USA (nationwide) including the state of: Puerto Rico and countries of: Albania, Algeria, Argentina, Australia, Austria, Azerbaijan, Belarus, Belgium, Bosnia and Herzegovina, Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Egypt, Estonia, France, Germany, Hungary, India, Israel, Italy, Japan, Malaysia, Mexico, Mongolia, Morocco, Netherlands, New Zealand, Nicaragua, Peru, Philippines, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovenia, South Africa, South Korea, Spain, Sudan, Switzerland, Taiwan, Thailand, Turkey, Turkmenistan, United Kingdom, Uruguay and Venezuela. | Terminated | April 08, 2014 | 1887 | 760 | Computer
Z-2598-2011 | 57801 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco Radiation Treatment Planning Workstation, Monaco Release 2.03.00 and above |  |  | January 13, 2011 | 2011 | Computerized Medical Systems Inc | Monaco: When the isocenter location is edited in the IMRT Activity screen, the shift coordinates are not updated in the IMRT Plan report. | Software design | Computerized Medical Systems Inc drafted a User Notice which explains the reason for correction, the clinical impact, and the workaround method. The User Notice was sent January 13, 2011. A return postcard is included for the customer to confirm receipt of the User Notice. The firm intends for this issue to be corrected in Monaco Release 3.00 in March 2011. |  | 137 | 137 | Worldwide Distribution - USA (Nationwide distribution to AK, AZ, CA, CO, FL, IL, IN, MA, MI, MN, MO, NC, NM, NY, OH, OR, PA, SC, TX, VA, WA, and WI and the countries of Algeria, Australia, Austria, Belarus, Belgium, Brazil, Canada, Chile, China, Cyprus, Czech Republic, Denmark, Estonia, Finland, France, Germany, Greece, India, Israel, Italy, Japan, Mexico, Netherlands, New Zealand, Phillipines, Poland, Portugal, Singapore, Slovenia, South Korea, Spain, Switzerland, Taiwan, Thailand, Turkey, and the United Kindom. | Terminated | April 02, 2014 | 1175 | 137 | Computer
Z-2585-2011 | 58837 | No | LFL |  |  | General &amp; Plastic Surgery | 510(k) | Integra CUSA NXT Ultrasonic Tissue Ablation System Service Module |  |  | May 02, 2011 | 2011 | Integra Limited | The insulated covering of a wire cable that powers the vacuum pump in the CUSA NXT Service Module can become abraded over time resulting in the exposure of bare wire. If this bare wire were to come in contact with the metal surface of the pump, it could result in a short circuit and a loss of aspiration. | Device Design | Integra sent a "FIELD ACTION NOTIFICATION" dated May 2, 2011 to all affected customers. The letter identified the product, the problem and the actions to be taken. Customers were informed of a Field Corrective Action that is being initiated. All devices (customer owned, Integra Loaner Pool, and Sales demos) will undergo rework of the vacuum pump assembly. The rework will be scheduled, at customer sites by Integra Field Service Technicians or returned to a designated Integra Service Center for rework. A loaner will be presented to customer accounts upon request. The letter asks that each customer complete and return the enclosed Recall Acknowledgement and Return Form to Integra, maintaining a copy of the letter and form for your records. For any additional information please call 609-936-2485. |  | 105 | 105 units | Worldwide -- USA (nationwide) and countries of France, Germany, Hungary, Ireland, Italy, Netherlands, Spain, Sweden and Taiwan | Terminated | March 26, 2014 | 1059 | 105 | Computer
Z-2591-2011 | 58853 | No | OUB | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | MiraDry System consisting of the DTS3000 Console, MiraDry Handpiece |  |  | April 24, 2011 | 2011 | Miramar Labs, Inc. | Reports of skin burns on axilla after treatment with the miraDry System. | Device Design | The firm,Miramar, has a service provider in Japan that assists in dealing with customers. The service providers were provided a copy of the investigation report and communicated directly with the customers to set up appointments for the field correction. The customers were personally visited by a representative from Miramars service provider in Japan. All devices had field correction performed on April 26-28, 2011. |  | 8 | 8 units | International only: Japan | Terminated | June 17, 2011 | 54 | 8 | Computer
Z-2584-2011 | 59037 | No | DPS | Electrocardiograph. | Cardiovascular | Cardiovascular | 510(k) | Philips SureSigns VS3 Vital Signs Monitor Models: |  |  | June 10, 2011 | 2011 | Philips Healthcare Inc. | Speakers on the SureSigns Patient Monitors and Vital Signs, failure may be identified by the display of a Speaker Malfunc message in the alarm message area or an AUDIO FAILED message in the main screen of the device or by the absence of audible sound | Device Design | Philips Healthcare sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated May 2011 to all affected customers. The letter identifies the products, problem, and actions to be taken.The letter states that customers may continue to use the affected products while awaiting replacement of the speaker, provided that the user routinely checks for the display of the "Speaker Malfunct." INOP. Additionally, the letter states that if the customer experience this INOP or there is no sound from the SureSigns Monitor, the customer is to remove the device from use and contact their local Philips service representative.The letter explains that Philips is completing the design for a replacement speaker assembly, which is expected to be available within 8 weeks. When available, Philips will replace the affected speaker assemblies free of charge.Customers are instructed to contact their local Philips Customer Care Center at 1-800-722-9377 if further information or support is required. |  | 5225 | 5225 units | Worldwide Distribution--USA (nationwide) and countries of ARAB JAMAHIRI, AUSTRIA, AUSTRALIA, BAHRAIN, BANGLADESH, BELGIUM, CANADA ,CHILE, COLOMBIA, CROATIA, CYPRUS, FRANCE, GERMANY, HUNGARY, INDIA, IRELAND, ITALY, KUWAIT, LEBANON, LIBYAN, MALAYSIA, MEXICO, NAMIBIA, NETHERLANDS, NEW ZEALAND, OMAN, PHILIPPINES, POLAND, PORTUGAL, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA ,SOUTH KOREA, SPAIN, SWITZERLAND, TAIWAN, TANZANIA, THAILAND, UNITED ARAB EMIRATES, and UNITED KINGDOM. | Terminated | August 15, 2013 | 797 | 5740 | Computer
Z-1929-2011 | 57288 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE |  |  | March 08, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware of potential safety issues associated with the CARESCAPE" Monitor B850 when Patient Data Module (PDM), Tram Acquisition Module (Tram), Patient Side Module (E-PSM or E-PSMP), E- MASIMO or E-(P)RE(S)TN modules or Bed-to-bed alarming feature is used.1. When CARESCAPE Monitor B850 is used with the above listed modules, some SpO2 alarms may be delayed or missed.2. Whe | Software design | The firm, GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated February 25, 2011 to its consignees/customers. The letter was addressed to Healthcare Administrator / Risk Manager, Chief of Nursing, and Director of Biomedical Engineering. The letter described the recall product the problem, Safety Issues, Affected Product details, Product Correction, Safety Instructions and Contact Information. The customers were instructed to follow the safety instructions provided in the letter. GE Healthcare is working on a software correction that will be provided to their customers once it is available.If you have any questions regarding this medical device correction or the identification or affected items please contact Technical Support at 1-800-558-7044, or your local Service Representative. |  | 3817 | 3817 | Worldwide distribution: USA including states of:AL, AR, AZ, CA, CO, CT, FL, GA, IA, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NE, NJ, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA,WA, WI, and WV; and countries including: JERUSALEM, UNITED KINGDOM, TURKEY, THAILAND, SWITZERLAND, SWEDEN, SPAIN, SLOVENIA, SLOVAKIA, SINGAPORE, SAUDI ARABIA, QATAR, PORTUGAL, POLAND, NORWAY, NEW ZEALAND, NETHERLAND, MALTA, MALAYSIA, LEBANON, KUWAIT, ITALY, ISRAEL, INDIA, GERMANY, FRANCE, FINLAND, ESTONIA, DENMARK, CZECH REPUBLIC, CHILE , CANADA, BELGIUM, BAHRAIN, AUSTRIA, AUSTRALIA. | Terminated | June 05, 2012 | 455 | 3817 | Computer
Z-2580-2011 | 58787 | No | ITH | Power traction equipment. | Physical Medicine | Physical Medicine | 510(k) | MTD 4000 - Mettler Traction Decompression System, Model: ME 4000 |  |  | April 14, 2011 | 2011 | Mettler Electronics Corp | The recall was initiated because Mettler has confirmed the possible failure of internal mechanical component and software detection for potential failure regarding Electronics Traction Device - MTD4000. The firm are initiating the recall because component failure might result in patient injury. Use of the device should cease immediately. | Device Design | Mettler Electronics sent an "URGENT MEDICAL DEVICE RECALL" letter dated April 14, 2011 to all affected consignees. The letter identified the product, the problem and the actions to be taken. Customers were instructed to immediately cease use of the recalled device, examine their inventory and quarantine product subject to recall. In addition, consignees were told if they have further distributed the affected product to please identify their customers and notify them at once of this product recall. The letter asks that that each includes a "Recall Return Response Form" that consignees are required to complete and return. If you have any questions please contact Mettler Electronics Corp 714-533-2221 x331, Monday through Friday, between the hours of 8:30 A.M. and 4:00 P.M. Pacific Time. |  | 221 | 221 | Worldwide Distribution --USA (nationwide) and countries of Taiwan, Ireland and Haiti. | Terminated | July 10, 2012 | 453 | 221 | Computer
Z-1907-2011 | 58006 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | 8V3c transducer used with the ACUSON SC2000 ultrasound system. 8V3c transducer model number 10433823; ACUSON SC2000 model number 10433816. |  |  | February 04, 2011 | 2011 | Siemens Medical Solutions USA, Inc. | When using the 8V3c transducer in color mode, the derated Ispta 3 (Spatial Peak-Temporal Average Intensity) may exceed the limit stated in the instructions for use. | Software design | A Customer Safety Advisory Notification was initiated on 2/4/2011 and sent via certified mail. The letter identified the affected product and described when the problem occurs along with the associated potential risk. The firm recommends that users not use the 8V3c until the software issue is fixed. The firm is working to release a software update and customers' local Customer Service Engineer will install new software on their system when it is available. Customers should direct their questions to their local service support person. The information in the letter should be shared with all personnel within the organization who need to be aware. |  | 386 | 386 units | Worldwide Distribution. | Terminated | October 04, 2011 | 242 | 386 | Computer
Z-2529-2011 | 58910 | Yes | NIP |  |  | Cardiovascular | PMA | Boston Scientific Innova Over-the-Wire Self-Expanding Stent System, Catalog No. 39181-07403. 7 mm x 40 mm x 130 cm (stent diameter x stent length x delivery system length). Made in USA, Two Scimed Place, Maple Grove, MN 55311 USA. The Innova Self-Expanding Stent System is comprised of two components: the implantable endoprosthesis and the stent delivery system. The stent is a laser cut self-expanding stent composed of a nickel titanium alloy (Nitinol). On both the proximal and distal ends of the stent, radiopaque markers made of tantalum increase visibility of the stent to aid in placement. The stent is constrained within a 6F (2.1 mm maximum OD) delivery system. The delivery system is a triaxial design with an outer shaft to stabilize the stent delivery system, a middle shaft to protect and constrain the stent, and an inner shaft to provide a guidewire lumen. The delivery system is compatible with 0.035 in (0.89 mm) guidewires. When ready to be implanted, the stent is deployed by retracting the middle shaft of the delivery system. A radiopaque marker at the distal end of the delivery system aids in visibility during deployment. As the stent is exposed to body temperature, it expands to appose the vessel wall. |  |  | May 19, 2011 | 2011 | Boston Scientific Corporation | Boston Scientific Corporation is conducting a recall of the Innova Self-Expanding Stent System because they have received 6 complaints involving no deployment/partial deployment of the Innova Stent. Potential health hazard events resulting from this type of failure include increased procedure time, vessel wall injury, and/or emergency surgery to remove the partially deployed stent. This recall d | Device Design | Boston Scientific sent an Urgent Medical Device Recall Removal - Immediate Action Required letter, dated May 13, 2011, to all affected customers via registered mail. The letter explained the issue, identified affected product, stated that further distribution and or use of the devices should cease immediately, and asked that devices in inventory to be segregated and immediately returned to Boston Scientific. An Account Reply Verification Tracking Form was also asked to be returned. Boston Scientific Sales Representative were assist in the return process by providing a Returned Goods Authorization (RGA) number.Customers were instructed to contact their local Sales Representative for any questions. |  |  |  | Worldwide Distribution - USA (nationwide) and the countries of Product removal in US is being managed through SuperNOVA Clinical Trial suspension in accordance with IDE #G100291. OUS distribution to include: Austria, Finland, France, Germany, Italy, Netherlands, Spain, and Switzerland. | Terminated | February 04, 2014 | 992 |  | Computer
Z-2577-2011 | 58210 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System, XiO Release 4.0.0 through 4.40.00 |  |  | April 08, 2008 | 2008 | Computerized Medical Systems Inc | XiO: When the patient is positioned with their feet toward the CT gantry, the coordinates exported to the LAP IsoMark laser positioning system are not correct. | Software design | The firm, CMS, Inc., sent a "XiO Customer Advisory" letter dated April 8, 2008 to all affected customers. The letter describes the product, the problem, and the actions to be taken ( workaround). The letter also states the problem will be resolved in a future release. No release number or date was given.If you have any questions, please call 408-380-8023. |  | 75 | 75 | Worldwide distribution: USA (nationwide) and countries of: China, France, Germany, Japan, Mongolia, Netherlands, Nicaragua, Poland, Russia, South Korea, Switzerland, and Turkey. | Terminated | April 08, 2014 | 2191 | 75 | Computer
Z-2579-2011 | 58197 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System, XiO Release 4.3.0 and above |  |  | October 03, 2010 | 2010 | Computerized Medical Systems Inc | XiO: When the beam template is recalled, XiO prematurely initiates a dose calculation, re-snaps the MLC or Port to the structure of the current patient , but then fails to re-calculate dose for the new MLC or Port position. If no subsequent change is made to force a dose recalculation, XiO will display the dose from the original MLC/Port shape dose instead of the dose from the re-snapped MLC/Port | Software design | The firm, Elekta CMS Software, sent an "IMPORTANT SAFETY NOTICE" dated September 14, 2010 to its affected customers. The notice describes the product, problem and actions to be taken. The customers were instructed to distribute this notice to any all users of CMS software at their organization who are potentially affected by this issue, and complete and return the attached postcard to confirm receipt of the Safety Notice. The notice also states that this problem has existed since the XiO Release 4.3.0 and is resolved in the XiO release 4.26.00. Elekta will notify the customers when this solution is available.If you have any questions, please call 314-993-0003 or toll free 800-878-4267. |  | 1208 | 1208 | Worldwide distribution: USA (nationwide) including Puerto Rico and countries of: Albania, Algeria, Argentina, Australia, Austria, Azerbaijan, Bahamas, Belarus, Belgium, Bolivia, Bosnia and Herzegovina, Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Ecuador, Egypt, El Salvador, Estonia, Finland, France, Georgia, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Kuwait, Latvia, Libya, Lithuania, Malaysia, Malta, Mexico, Mongolia, Morocco, .Netherlands, New Zealand, Nicaragua, Philippines, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Suriname, Switzerland, Taiwan, Tajikistan, Thailand, Turkey, Turkmenistan, Ukraine, United Kingdom, Uruguay, Venezuela, Yemen, and Zimbabwe. | Terminated | April 03, 2014 | 1278 | 1208 | Computer
Z-2593-2011 | 58262 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System, XiO Release 2.1.0 and above |  |  | May 27, 2010 | 2010 | Computerized Medical Systems Inc | XiO Software: The Patient Scan position is being confused with the treatment orientation of the patient. The Patient Scan position is determined from the data entered in PFM. Subsequently, when a New Teletherapy Plan is created for a patient in Teletherapy, the user is prompted with Head/Foot Gantry on the text page. This prompt is asking for the orientation of the patient when they are treated. | Software design | Computerized Medical Systems Inc sent an undated Safety Notice to all affected customers. The notice indicated the problem, the clinical impact, and the workaround. The notice stated the sites will be notified when a solution is available. |  | 1400 | 1400 | Worldwide Distribution - USA (Nationwide distribution, including Washington, DC and Puerto Rico) Foreign distribution was made to Albania, Algeria, Argentina, Australia, Austria, Azerbaijan, Bahamas, Belgium, Bolivia, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, El Salvador, Estonia, Finland, France, Georgia, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Kuwait, Latvia, Libya, Lithuania, Malaysia, Malta, Mexico, Mongolia, Morocco, Netherlands, New Zealand, Nicaragua, North Korea, Philippines, Poland, Portugal, Qatar, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Tajikistan, Tanzania, Thailand, Turkey, Turkmenistan, Ukraine, United Kingdom, Uruguay, Venezuela, Yemen, and Zimbabwe. | Terminated | April 09, 2014 | 1413 | 1400 | Computer
Z-2568-2011 | 58215 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System, XiO Release 4.3.0 through 4.50.00. |  |  | May 13, 2009 | 2009 | Computerized Medical Systems Inc | In XiO, when the beams are renumbered, it is possible to get into a state where the dose displayed for a beam does not match the beam placement on the patient. This can occur when dose is calculated, beams are renumbered, and the dose is re-summed only. | Software design | Computerized Medical Systems, Inc. sent an CUSTOMER ADVISORY letter dated May 2009, to all affected customers. The letter identified the product, the problem, and the action needed to be taken by the customer.Computerized Medical Systems, Inc have created "patches" to resolve the problem. A User Notice was sent in March 2011, to all users still using the affected software notifying them of the updated software available to correct the problem. This letter was sent by US Mail but no return receipt was included.For further questions, please call (408) 380-8023. |  | 1116 | 1116 | Worldwide Distribution -- (USA) nationwide Distribution, including Puerto Rico, and countries of Albania, Algeria, Argentina, Australia, Austria, Azerbaijan, Bahamas, Belgium, Bolivia, Bosnia, Herzegovina, Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Ecuador, Egypt, El Salvador, Estonia, France, Georgia, Germany, Greece, Guatemala, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Kuwait, Libya, Lithuania, Malaysia, Malta, Mexico, Mongolia, Netherlands, New Zealand, Nicaragua, Philippines, Poland, Portugal, Qatar, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Tajikistan, Thailand, Turkey, Turkmenistan, Ukraine, United Kingdom, Uruguay, Venezuela, Yemen, and Zimbabwe. | Terminated | April 09, 2014 | 1792 | 1116 | Computer
Z-2574-2011 | 58214 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System, XiO Release 1.5.0 and above. |  |  | May 14, 2008 | 2008 | Computerized Medical Systems Inc | In XiO, when Telepathy is exited, re-entered, and the temporary plan re-called, the OLD beam weight will be displayed but the dose distribution will reflect the modified beam weight. | Software design | Computerized Medical Systems, Inc. sent an CUSTOMER ADVISORY letter to all affected customers in May 2008. The letter identified the product, the problem and the action needed to be taken by the customer.The issue was resolved in XiO Release 4.50.00 which has been available since June 2009.A User Notice was sent in March 2011, to all the customers reminding them of the availability of the correction. This notice was sent hardcopy via US Mail but there was no return receipt.For further questions, please call ( 408) 380-8023 |  | 936 | 936 | Worldwide Distribution -- USA (nationwide) including Puerto Rico, and countries of Albania, Algeria, Argentina, Australia, Austria, Azerbaijan, Bahamas, Belgium, Bolivia, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Ecuador, Egypt, El Salvador, Estonia, France, Georgia, Germany, Greece, Guatemala, Hungary, India, Indonesia, Israel, Italy, Japan, Libya, Lithuania, Malaysia, Mexico, Morocco, .Netherlands, New Zealand, Nicaragua, North Korea, Philippines, Poland, Portugal, Qatar, Russia, Saudi Arabia, Serbia, Singapore, South Africa, South Korea, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Tajikistan, Tanzania, Thailand, Turkey, Ukraine, United Kingdom, Uruguay, Venezuela, Yemen, and Zimbabwe. | Terminated | April 08, 2014 | 2155 | 936 | Computer
Z-2572-2011 | 58233 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Focal Radiation Treatment Planning System, Focal Release 4.3.1 through 4.34.02. |  |  | March 31, 2008 | 2008 | Computerized Medical Systems Inc | Focal: In absolute dose mode a beam with a port may be calculated with open normalization, with the result that the displayed dose reflects open normalization while the monitor units and QA plan reflect blocked normalization. This is particularly important when these beams also include compensators. The problem is detectable by looking at the delivered dose to the weight point in the patient as | Software design | Computerized Medical Systems Inc sent a "CUSTOMER ADVISORY" letter dated March 2008 to all affected customers. The letter describes the product, the problem and the actions to be taken. The firm provided a workaround for customers that states "Alternatively the plan can be imported directly into XIO". A copy of the Advisory was shipped with software within the affected range to those requesting a software upgrade. The issue was resolved in Focal Release 4.40.00 introduced in June 2008.The firm will send out a User Notice to all sites still running affected software to remind them that the issue resolution is available.For further questions, please call (403) 830-8023. |  | 307 | 307 | Worldwide Distribution -- USA (nationwide) and the countries of Albania, Argentina, Australia, Belgium, Bosnia and Herzegovina, Brazil, Canada, Chile, China, Costa Rica, Croatia, Czech Republic, France, Germany, Greece, India, Italy, Japan, Netherlands, Peru, Portugal, Qatar, Romania, Russia, Serbia, South Africa, South Korea, Spain, Thailand, Turkey, United Kingdom, Uruguay, and Venezuela. | Terminated | April 09, 2014 | 2200 | 307 | Computer
Z-2567-2011 | 58227 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System, XiO Release 4.3.1 through 4.40.00. |  |  | December 09, 2008 | 2008 | Computerized Medical Systems Inc | XiO Software: Under certain conditions, when the port shape is altered, dose is not recalculated. The port-change flag is not being set correctly when the port shape is regenerated. The dose distribution will not match the blocking displayed and the mis-match might not be obvious. Monitor Unit calculations will also be based on an incorrect dose distribution. This problem has existed since Re | Software design | Computerized Medical Systems, Inc. sent a "XiO CUSTOMER ADVISORY" letter dated December 8, 2008 to all affected customers. The advisory describes the product, problem, and the actions to be taken. The customers were instructed to set the prompt to "Yes" to avoid the problem from occurring according to written instructions provided in the advisory. This problem has existed since Release 4.3.1 and is expected to be resolved with XiO Release 4.50.00. Questions should be directed to the manufacturer's representative at 408-830-8023. |  | 784 | 784 | Worldwide Distribution-- USA (nationwide) including Puerto Rico and countries of Albania, Algeria, Argentina, Australia, Austria, Azerbaijan, Bahamas, Belgium, Bolivia, Bosnia and Herzegovina, Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Ecuador, Egypt, El Salvador, Estonia, France, Georgia, Germany, Greece, Guatemala, Hungary, India, Indonesia, Israel, Italy, Japan, Libya, Lithuania, Malaysia, Mexico, Mongolia, Netherlands, New Zealand, Nicaragua, Philippines, Poland, Portugal, Qatar, Russia, Saudi Arabia, Serbia, Singapore, South Africa, South Korea, Spain, Sri Lanka, Sweden, Switzerland, Tajikistan, Thailand, Turkey, Ukraine, United Kingdom, Uruguay, Venezuela, Yemen, and Zimbabwe. | Terminated | April 09, 2014 | 1947 | 784 | Computer
Z-2569-2011 | 57843 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco Radiation Treatment Planning Workstation. Monaco Release 2.04.00 and above. |  |  | October 27, 2010 | 2010 | Computerized Medical Systems Inc | Monaco: In XiO IMRT (Dynamic MLC Delivery) software, the radiation beam angle is set to zero. If different beam angles are selected via Monaco software, the beam angles will remain at zero. | Software design | Computerized Medical Systems, Inc. sent a User Notice to the affected consignees on October 27, 2010, explaining the reason for correction, the clinical impact, and the workaround method. (VERIFY USE NOTICE DATE) A return postcard was included for the customer to confirm receipt of the User Notice. The firm intends for this issue to be corrected in Monaco Release 3.00 in March 2011. |  | 11 | 11 | Worldwide Distribution - USA including CA, IL, MN, NM, OH, NY, TX, and WA and the Netherlands. | Terminated | April 02, 2014 | 1253 | 11 | Computer
Z-2564-2011 | 58258 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco Radiation Treatment Planning System, Monaco Release 2.0.1 and above |  |  | May 31, 2010 | 2010 | Computerized Medical Systems Inc | Monaco: due to a software defect, the Radiological Depth (cm) and SSD to DRP (cm) parameters are reported incorrectly on the Beam Summary Report. This occurs for all Monaco IMRT plans. | Software design | Computerized Medical System sent a "USER NOTICE" letter dated June 24, 2010, to all affected users. This notice identified the product, the problem, and the actions to be taken. There is no current workaround. All customers will be mailed a postcard to be returned as proof of receipt and understanding of the notification. Consignees will be contact via telephone or e-mail after June 30, 2010 if the postcard is not returned This notice is to be distributed to any and all users of CMS software at your organization who are potentially affected. The notice states the issue will be resolved in Monaco Release 2.04, but no date for this release was given.For further questions, please call (403) 830-8023 |  | 62 | 62 | Worldwide Distribution -- USA (nationwide) and the countries of Australia, Austria, Belgium, China, Finland, France, Germany, India, Israel, Italy, Japan, Mexico, Netherlands, New Zealand, South Korea, Spain, Switzerland, Taiwan, Thailand and United Kingdom. | Terminated | April 10, 2014 | 1410 | 62 | Computer
Z-1928-2011 | 58014 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Hamilton-G5 Ventilator with software version V2.00x and Neonatal Option or V2.01x and Neonatal software. Hamilton-G5: catalog number 159001; |  |  | February 16, 2011 | 2011 | Hamilton Medical Inc | The adaptive volume/pressure regulation can be automatically and permanently reduced to PEEP + 3 mbar when a closed suctioning maneuver is performed on the patient and may be undetected by the operator. | Software design | A Medical device Safety Alert and Corrective Action along with a cover letter, was sent to all locations on February 16, 2011. The letter identified the affected product along with the reason for recall. The letter states that operators deactivate the "high tidal volume" alarm as described in appendix K.4 of the operator's manual until the ventilator is upgraded with a corrected software. Distributors will upgrade the affected devices with the corrected software as soon as possible. Questions can be directed to 800-426-6331, ext 215. |  | 97 | 97 units | Nationwide Distribution -- MA, FL, GA, NC, MI, NC, MO, TX, and WV. | Terminated | March 21, 2012 | 399 | 97 | Computer
Z-2563-2011 | 58676 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | GE Centricity Laboratory Instrument Interface to Data Innovations Instrument Manager software; GE Healthcare IT Solutions, Barrington, IL 60010. |  |  | May 03, 2011 | 2011 | Ge Healthcare It | The GE Centricity Laboratory Instrument Interface to Data Innovations is processing a preliminary result instead of the final result from the Process System Manager when there is a test rerun conducted. | Software design | GE Healthcare sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated May 3, 2011 to all affected customers. The letter describes the product, problem, and action to be taken by the customer. The letter provided the users with specific safety instructions to follow until the software was updated. The customers were requested to complete the enclosed customer reply form and fax the completed form to 847-620-2780. Any questions should be directed to their GE Service Representative or the GE Help desk at 888-778-3375. |  | 14 | 14 units | Worldwide Distribution-- USA (nationwide) including states of Georgia, New York, Ohio, Oregon, and Tennessee and countries of Canada, England, and Qatar. | Terminated | October 04, 2011 | 154 | 14 | Computer
Z-1930-2011 | 58275 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, CARESCAPE Monitor B650 |  |  | March 23, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware of a potential safety issue associated with the CARESCAPE Monitor B650 &amp; B850 Cardiac Output measurement using Patient Data Module (PDM), and activation of ECG Leads Off alarm in specific cases.1. When the CARESCAPE Monitor is used to measure Cardiac Output values with the Patient Data Module (PDM) or TRAM Module, Cardiac Output values may be calculated incorrectl | Software design | The firm, GE Healthcare, sent "Urgent Medical Device Correction" letters dated March 23, 2011 to its consignees/customers. The letters were addressed to Healthcare Administrator / Risk Manager, Chief of Nursing, and Director of Biomedical Engineering. The letters described the Safety Issues, Safety Instructions, Affected Product Details, Product Correction, and Contact Information. The customers were instructed to follow the safety instructions documented in the letter. GE Healthcare is working on a software correction that will be provided to their customers once it is available and will contact the customer to schedule the software update.If you have any questions regarding this medical device correction or the identification of affected items, please contact Technical Support at 1-800-558-7044, or your local Service Representative. |  | 1443 | 1443 | AL, AZ, AR, CA, CO, CT, FL,GA, ID, IL, IN, IA, KY, LA, MD, MA, MI, MN, MS, MO, MT, NE, NV, NJ, NM, NY, NC, OH, OK, OR, PA,PR, SC, TN,TX, UT, VA, WA, WV, WI.UNITED KINGDOM, UNITED ARAB EMIRATES, TURKEY, TUNISIA, THAILAND, SWITZERLAND, SWEDEN SLOVENIA, SLOVAKIA, SPAIN, SINGAPORE, SAUDI ARABIA, QATAR, PORTUGAL POLAND, NORWAY, NORTH KOREA, NEW ZEALAND, NETHERLANDS, MALTA, MALAYSIA, LITHUANIA, LEBANON, KUWAIT, ITALY, ISRAEL, IRELAND, INDIA, ICELAND, HONG KONG, GERMANY, FRANCE, FINLAND, DENMARK, ESTONIA, CZECH REPUBLIC, BELGIUM BAHRAIN, COLOMBIA, CHILE CANADA, BRAZIL, AUSTRIA, AUSTRALIA. | Terminated | February 04, 2012 | 318 | 8345 | Computer
Z-2560-2011 | 58229 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco Radiation Treatment Planning System, Monaco Release 1.0.0 through 1.0.2. |  |  | June 30, 2008 | 2008 | Computerized Medical Systems Inc | Monaco: If the user selects Monte Carlo Photon as the Secondary Algorithm (to be used for final dose calculation through segments), and selects SKIP Forward button at any time during stage 2 optimization, the system may only use the pencil beam algorithm for final computation of dose even though Monte Carlo was selected. Other than specific output to the console window, there will be no other in | Software design | Computerized Medical Systems, Inc. sent a "MONACO CUSTOMER ADVISORY" letter dated June 2008 to all affected customers. The letter describes the product, problem, and the action to be taken. The customer was instructed to use caution when using the Skip Forward button in stage 2 according to written instructions provided in the letter.This defect exists in Monaco Release 1.0.0 through 1.0.2 and will be resolved as soon as possible. Note: The issue was resolved in Monaco Release 2.0.1 and above.Questions should be directed to the manufacturer's representative telephone number at 408-830-8023. |  | 47 | 47 | Worldwide Distribution-- USA (nationwide) including the states of IL, MI, MO, OR, SC, TX, VA, and Puerto Rico and the countries of Canada, China, Cyprus, Estonia, France, Germany, Greece, Italy, Japan, Netherlands, Philippines, Poland, Portugal, Slovenia, South Korea, Spain, and United Kingdom. | Terminated | April 04, 2013 | 1739 | 47 | Computer
Z-2559-2011 | 58226 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Focal Radiation Treatment Planning System, Focal Release 4.34.00 through 4.50.00. |  |  | June 12, 2009 | 2009 | Computerized Medical Systems Inc | Focal Software: When a patient is defined as prone and the shifts are manually entered and the LOCK SHIFT box is checked the X coordinate (left/right) and the Z coordinate (anterior/posterior) are reversed (left to right, anterior to posterior). This could potentially lead to incorrect isocenter definitions from prone patients and thus a mislocation of dose. | Software design | Computerized Medical System sent a CUSTOMER ADVISORY dated June 12, 2009, to all affected users. This advisory identified the product, the problem, and the action needed to be taken by the customer. There were no patients mistreated as a result of this issueThe firm will send out a User Notice to all sites still runningaffected software to remind them that the issue resolution is available.For further questions, please call (403 ) 830-8023 |  | 462 | 462 | Worldwide Distribution -- USA (nationwide) including Puerto Rico, and the countries of Algeria, Australia, Austria, Azerbaijan, Belarus, Belgium, Bosnia and Herzegovina, Brazil, Canada, Chile, China, Colombia, Croatia, Egypt, Estonia, France, Germany, Greece, Hungary, India, Israel, Italy, Japan, Malaysia, Mexico, Mongolia, Morocco, Netherlands, Nicaragua, Peru, Philippines, Poland, Portugal, Russia, Saudi Arabia, Serbia, Singapore, South Africa, South Korea, Spain, Sudan, Switzerland, Taiwan, Thailand, Turkey, Turkmenistan, United Kingdom, and Venezuela. | Terminated | April 09, 2014 | 1762 | 462 | Computer
Z-2562-2011 | 57845 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Focal Radiation Treatment Planning Workstation. Focal Release 4.51.00 and above. |  |  | November 09, 2010 | 2010 | Computerized Medical Systems Inc | Focal: If a plan involving Dynamic MLC Delivery is sent from XiO to Focal and called up there in Plan Review, the angle of the gantry is set to zero for all beams. If this plan is then sent from Focal to the R&V; System, the gantry angles all remain at zero which could lead to patient mistreatment. | Software design | Elekta CMS Software sent a USER NOTICE dated November 5, 2010, to all affected customers. The notice identified the product, the problem, and the action to be taken by the customer. A return postcard was included for the customer to confirm receipt of the User Notice. The firm intends to correct the software so it does not modify anything about the plan sent from XiO.Customers were instructed to contact their local customer support office with any questions. |  | 155 | 155 | Worldwide Distribution - USA (nationwide) and the countries of Algeria, Australia, Belarus, Belgium, Brazil, Canada, Chile, China, Egypt, Estonia, Finland, France, Germany, Hong Kong, India, Italy, Latvia, Libya, Lithuania, Mexico, Morocco, Netherlands, New Zealand, Philippines, Poland, Romania, Singapore, Slovenia, South Africa, Spain, Suriname, Switzerland, Taiwan, Turkey, Turkmenistan, and the United Kingdom. | Terminated | April 02, 2014 | 1240 | 155 | Computer
Z-2561-2011 | 57760 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco Radiation Treatment Planning Workstation. Monaco Release 2.03.00 and above. |  |  | January 13, 2011 | 2011 | Computerized Medical Systems Inc | Monaco: For VMAT plans, the measured dose of radiation delivered is lower than the software calculated dose by about 5% in Monaco Release 2.03.00 and 2.04.00. | Software design | Computerized Medical Systems, Inc. sent a "USER NOTICE" on January 13, 2011, to all affected users. This notice identified the product, the problem, and the action needed to be taken by the customer. A return postcard is included for the customer to confirm receipt of the user notification. The firm intends for correction in the dose calulation engine to appear in the Monaco Release 3.00 in March 2011. Any further questions please call (403) 830-8023. |  | 94 | 94 | Worldwide Distribution-- USA (nationwide) and countries of Austria, Australia, Belarus, Canada, China, Cyprus, Czech Republic, Denmark, Estonia, Finland, France, Germany, Greece, India, Israel, Italy, Japan, Mexico, Netherlands, Poland, Singapore, Slovenia, South Korea, Spain, Switzerland, Taiwan, Thailand and UK. | Terminated | April 03, 2014 | 1176 | 94 | Computer
Z-2414-2011 | 58661 | No | FMZ | Neonatal incubator. | General Hospital | General Hospital | 510(k) | Datascope/ Maquet IABPs. System 98/98XT Intra-Aortic Balloon Pump; CS100/CS100i Intra-Aortic Balloon Pump; CS300 Intra-Aortic Balloon Pump; |  |  | March 16, 2011 | 2011 | Maquet Datascope Corp - Cardiac Assist Division | Several models of the Intra-Aortic Balloon Pumps are being recalled due to a potential problem with the fan's assembly which may cause the power supply to detect an overheating event and shut down. | Device Design | Maquet sent an Urgent Product Field Correction letter dated March 16, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to note the WARNING in the IABP Operating Instruction. Customers were advised to refer to the IAB Instructions for Use, Section D (Manually inflating and Deflating a Catheter). Section D of the IAB instruction for Use reiterates that an IAB should not remain inactive for more than 30 minutes and also provides detailed instruction on how to keep an IAB active during IABP failure. Customers were advised that their Service Representatiave would arrange to replace their supply which would contain a new fan assembly. Upon completion of the power supply replacement, customers were asked to sign a service repair order to verify satisfactory completion of the work. For technical questions contact Technical Support Department at 1-800-777-4222 and press 4.For scheduling a service visit by a Field Service Representatiave, contact the firm's US Call Center at 1-800-777-4222 and press 3.Maquet issued a press release on June 6, 2011 advising consumers of the issue and that their recall was considered a class one. |  | 885 | 885 units | Worldwide Distribution | Terminated | October 25, 2012 | 589 | 885 | Computer
Z-2594-2011 | 58253 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Focal Radiation Treatment Planning System, Focal Release 4.40.00 and above. |  |  | January 19, 2009 | 2009 | Computerized Medical Systems Inc | Focal: the couch position can be moved between scans of the same CT study. When the images are imported into Focal, the software is not checking the DICOM Image Position (Patient) tag and therefore is not aligning the images sets if there are different couch positions indicated. A shift between image sets will exist of the same amount that the CT couch was moved. | Software design | Elekta CMS Software sent an "ADVISORY NOTICE" dated January 19, 2009 to its affected customers. The notice identified the product, the problem, and the acton to be taken. Customers were informed that the workaround method available is to (use the image fusion feature already in Focal to relign the scans). The firm states that this issue is resolved in Focal version 4.62.00, which was released in March 2011. If you have any questions, please call 314-993-0003 or toll free 800-878-4267. |  | 31 | 31 | Worldwide Distribution -- USA (nationwide) including states of: CA, CO, FL, IL, IN, KS, MO, NC, OH, OK, PA, TN, and WA. and countries of Australia, Austria, Canada, India, Netherlands, Portugal, Singapore, Spain, and United Kingdom. | Terminated | April 10, 2014 | 1907 | 31 | Computer
Z-2596-2011 | 58256 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Focal Radiation Treatment Planning System, Focal Release 4.3.0 through 4.34.01 |  |  | November 30, 2007 | 2007 | Computerized Medical Systems Inc | Focal: when a plan is sent to Focal from XiO and the structure properties are edited in Focal CT Sim activity, these changes result in changes to the MLC or field size for the plan when they should not. XiO will recalculate dose when modified plan is returned to XiO however the user may not notice that it has changed form the plant they had approved earlier. Editing the structure properties (col | Software design | Computerized Medical Systems Inc sent a Customer Advisory, dated 11/30/07, was sent to consignees December 2007 to notify them of the issue and how to avoid the workaround. The advisory listed the problem, clinical impact, and a workaround. The firm will be sending each site a reminder that the issues is resolved. |  | 233 | 233 | Worldwide Distribution - USA (Nationwide AL, AR, CA, DE, FL, GA, IL, IN, MA, MD, MI, MOP, NE, NY, OH, PA, RI, TX, VA, WA, WI, and WV and the countries of Albania, Argentina, Belgium, Bosnia and Herzegovina, Brazil, China, Costa Rica, Croatia, Czech Republic, France, Germany, Greece, India, Italy, Japan, Netherlands, Peru, Portugal, Qatar, Romania, Russia, Serbia, South Africa, South Korea, Spain, Turkey, United Kingdom, Uruguay, and Venezuela. | Terminated | April 10, 2014 | 2323 | 233 | Computer
Z-1943-2011 | 57189 | No | DTY | Cardiopulmonary bypass in-line blood gas sensor. | Cardiovascular | Cardiovascular | 510(k) | Terumo Cardiovascular Systems Corporations, CDI 101 Hematocrit/Oxygen Saturation Monitoring System. In-line blood-gas sensor used in cardiopulmonary bypass. |  |  | November 01, 2010 | 2010 | Terumo Cardiovascular Systems Corporation | Inaccurate result. Possibility for displayed values for potassium, CO2, hematocrit, O2 saturation, pH, hemoglobin and other measured blood parameters to be inaccurate. | Device Design | On 11/1/2010, Terumo sent "URGENT MEDICAL DEVICE CORRECTION" notices to Chief Perfusionist, Department of Cardiovascular Surgery or Director of Operating Room Services. The letter informed the consignees about the incidents of the CDI 100 and 101 displaying inaccurate values. The letter states that when a displayed value is significantly different from expectations based on the clinical situation, the user should verify the accuracy by independent means before treating. The letter included "Caution" stickers that were to be applied to the CDI system operators manual. |  | 468 | 468 monitors | US, India, Colombia, Taiwan, Venezuela, Vietnam, Puerto Rico, Malaysia, Sri Lanka, Brazil, Indonesia, Pakistan, Chile, Germany, Hong Kong, Japan, AE, Singapore, Austria, Belgium and Thailand. | Terminated | April 17, 2012 | 533 | 1500 | Computer
Z-1884-2011 | 52820 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | TEE Transducer V5M, when used with the Acuson S2000 Ultrasound System, software versions: VA15, VA15A or VA15B; Model Numbers: 10041461 (Acuson S2000) and 8263703 (V5M Transducer); |  |  | April 07, 2009 | 2009 | Siemens Medical Solutions USA, Inc. | Device function failure--When performing a TEE exam with the transducer, CW (Continuous Wave) mode is not functional. Due to work flow, this may not be apparent until the transducer is inserted into the patient's esophagus. | Software design | Siemens sent a Customer Safety Advisory Notification letter on April 7, 2009, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer.Customers were instructed to pass the notice on to all those within their organization who need to be aware of the issue until the corrective action is completed.For questions regarding this recall call 650-969-9112. |  | 61 | 61 units | Worldwide Distribution - USA (nationwide) and the countries of Australia, Canada, Switzerland, China, Czech Republic, Germany, Denmark, Great Britain, Hong Kong, Indonesia, India, Italy, South Korea, Norway, Turkmenistan, and Taiwan | Terminated | April 05, 2011 | 728 | 61 | Computer
Z-1853-2011 | 54844 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Acuson Antares Ultrasound Imaging System with software versions 4.0 and 5.0. |  |  | March 01, 2010 | 2010 | Siemens Medical Solutions USA, Inc. | Image captured from the device may be erroneous. | Software design | Acuson Customer Safety Advisory notification letters were sent March 1, 2010 by certified mail. |  | 3539 | 3539 | Worldwide distribution to USA, Andorra, United Arab Emirates, Netherlands Antilles, Angola, Austria, Australia, Azerbaijan, Bosnia, Bangladesh, Belgium, Bahrain, Brunei Darussalam, Brazil, Belarus, Canada, Switzerland, Chile, China, Colombia, Costa Rica, Cyprus, Czech Republic, Germany, Denmark, Ecuador, Egypt, Spain, Finland, Faroe Islands, France, Great Britain, Georgia, Greece, Guatemala, Hong Kong, Croatia, Hungary, Indonesia, Ireland, Israel, India, Iran, Italy, Jordan, Japan, South Korea, Kuwait, Lebanon, Sri Lanka, Lithuania, Latvia, Morocco, Moldova, Mexico, Nicaragua, Netherlands, Norway, New Zealand, Oman, Peru, Philippines, Pakistan, Poland, Portugal, Qatar, Romania, Serbia, Russian Federation, Saudi Arabia, Sweden, Singapore, Slovak Republic, Suriname, Thailand, Turkey, Trinidad and Tobago, Taiwan, Ukraine, UA, Uruguay, Venezuela, Vietnam, Yemen, and South Africa. | Terminated | June 18, 2012 | 840 | 3539 | Computer
Z-1862-2011 | 57852 | No | HTJ | Depth gauge for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | Xcelerate Specialty Patella Milling System Per File K2399; Catalog Number I-K2399PR00, I-K2399PR10; Stryker Howmedica Osteonics, Howmedica Osteonics Corp., 325 Corporate Drive, Mahwah, NJ 07430 |  |  | January 27, 2011 | 2011 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics has become aware that the 10 mm on the bone remaining depth gage is incorrect thus allowing for the reamer to leave as little as 8.9mm (in a nominal condition) patella bone. | Device Design | Stryker sent an Urgent Product Recall letter dated January 27, 2011, via FedEX for January 28, 2011, delivery with return receipt. The letter identified the product, the problem, and the action to be taken by the customer. Customers were insructed to fax back the attached Product Recall Acknowledgement Form within 5 days to 201-831-6069. Customers were also instructed to contact the hospitals in their territory that have the affected product to arrange return of the product. Product should be returned to the attention of: Regulatory ComplianceStryker Orthopaedics325 Corporate DriveMahwah, New Jersey, 07430. The orange Product Remediation sticker should be attached to the return, indicating the Product Remediation # RA 2009-296. Mark the outer box with words "Product Recall."For questions regarding this recall call (201) 972-2100. |  | 52 | 52 units | Nationwide Distribution | Terminated | October 15, 2012 | 627 | 52 | Computer
Z-2702-2011 | 58836 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Eclipse Radiation Treatment Planning System, model H48; |  |  | May 03, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | In affected Analytical Anisotropy Algorithm versions when modeling the second source, the collimator jaw positions are read from the first control point of the field. This can cause the delivered dose to be higher than planned locally by up to 10%. | Software design | Varian Medical Systems, Inc. Oncology Systems sent a "Urgent Medical Device Correction" letters on May 3, 2011***12-12-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.***The letter contained recommended actions. For additional information contact Varian Medical Systems, Inc. at 888- 827-4265. |  | 10649 | 5201 units ***12-12-12 AMENDED TO: 5448 units*** | Worldwide distribution; Nationwide (USA) and the countries of Vietnam, Venezuela, United Kingdom, United Arab Emirates, Ukraine, Turkey, Tunisia, Trinidad and Tobago, Thailand, Taiwan, Syria, Switzerland, Sweden, Sri Lanka, Spain, South Korea, South Africa, Slovenia, Singapore, Saudi Arabia, Russia, Romania, Qatar, Portugal. Poland, Philippines, Panama, Pakistan, Norway, New Zealand, Netherlands, Morocco, Mexico, Malaysia, Macau, Lithuania, Libya, Latvia, Kuwait, Kazakjstan, Japan, Italy, Israel, Ireland, Iraq, Indonesia, India, Hungary, Hong Kong,, Guam, Greece, Germany, France, Finland, Egypt, Ecuador, Denmark, Czech Republic, Colombia, China, Canada, Brazil, Belgium, Belarus, Bangladesh, Bahrain, Austria, Australia, Argentenia, Algeria and Canada. | Terminated | July 09, 2014 | 1163 | 10649 | Computer
Z-2707-2011 | 58576 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | MaxPlus bifuse extension set (minibore) needleless connectors, |  |  | April 20, 2011 | 2011 | Carefusion 303, Inc. | Carefusion has initiated a product recall for the MaxPlus bifuse extension set (minibore) needleless connectors Model MP9232-C as a result of a potential for occlusion. The product may be occluded at the bifused component. | Device Design | CareFusion initiated on May 3, 2011, two distinctive "URGENT MEDICAL DEVICE RECALL NOTIFICATION" via mail delivery, to all affected customers and distributors. The notices identified the product, the problem, and the actions to be taken. Customers were instructed to 1.) Immediately complete and return the enclosed pre-addressed and postage paid Recall Response Card to CareFusion. Note on the card your distributor, the MaxPlus bifuse extension set lot numbers and quantities that will be returned.2.) All recalled product should be returned directly to the distributor from whom it was purchased. Customers will receive their return goods number through their distributor or CareFusion. 3.) Also, enclosed is a Frequently Asked Question (FAQ) Letter to address any questions or concerns related to the use of the MaxPlus" bifuse extension sets.The second notification was specifically addressed to distributors. The distributors were provided specific instructions for product returned. For customers who may have questions, please contact your distributor directly, or call the CareFusion Support Center at 1-800- 562-6018. The Support Center will be in operation from 7 a.m. - 4 p.m. PST. For distributors who may have any questions or require additional materials, please contact the CareFusion Support Center at 1800- 562-6018. Additionally, instructions were given to report adverse reactions experienced via the FDA's Med Watch Program. |  | 26150 | 26,150 units | Nationwide Distribution -- USA including states of: CA, IA, IN, and TX | Terminated | November 28, 2011 | 222 | 26150 | Computer
Z-1865-2011 | 57890 | No | IXI | Radiation therapy beam-shaping block. | Radiology | Radiology | 510(k) | Axxent Flexishield Mini |  |  | February 03, 2011 | 2011 | Xoft Inc. | Xoft Axxent Flexishield Mini product may may shed particles identified as tungsten, and may look like suspicious calcifications on follow up scans if the product was used during Intraoperative radiation therapy. | Device Design | Xoft, Inc. sent an Urgent Field Removal Action letter dated February 3, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to stop using and return all units and lots of Flexishield Mini Catalog Number: F5300 in their inventory. Customers were instructed to fill out the attached form and fax to 866-222-3404, then call Customers Service toll free directly at 877-963-8327 to make arrangements to return the product and receive a credit if applicable. Customers were asked to complete and return the form even if they did not have any units to return.For questions regarding this recall call 408-419-2300. |  | 165 | 165 units | Nationwide Distribution including AL, CA, CO, FL, IL, MI, MO, MS, NJ, NY, OK, PA, TN, and WI. | Terminated | June 08, 2011 | 125 | 165 | Computer
Z-2705-2011 | 58376 | No | HAO | Nonpowered neurosurgical instrument. | Neurology | Neurology | 510(K) Exempt | Sequoia final driver closure top retention, non-sterile, REF 3384-2, Zimmer Spine, Austin, Texas, The device is used to apply final torque to closure tops of the implant. |  |  | April 04, 2011 | 2011 | Zimmer Inc. | A rounded or partially rounded tip on the screwdriver may not provide the required locking torque to lock the closure top to the polyaxial screw. When the hex tip of the screwdriver is not fully engaged and/or the screwdriver is used off axis, the hex tip of the screwdriver can become rounded or partially rounded. | Device Design | The firm, ZIMMER, sent an "URGENT MEDICAL DEVICE RECALL" letter dated April 4, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure that each instrument tray had both final drivers (with and without retention feature); contact their customer service representative if they do not have both versions; ensure proper seating of the drive by moving the driver side to side, and complete and return the Distributor Certification Form via fax to: Zimmer Spine at 512-258-0995. The letter states the firm will be distributing the version of the screwdriver with no retention feature beginning in March 2011 and distributing the new surgical technique in April 2011. The firm is including an additional final driver in the instrument tray for final tightening of the closure tops. Please keep Zimmer Spine informed of any adverse events associated with the device or any other Zimmer spine product by calling 800-326-0635. For any further questions or concerns contact the Vice President Quality Assurance &amp; Regulatory Affairs at 800-655-2614. |  | 505 | 505 devices | Worldwide distribution: USA (nationwide) including the states of: AL, AZ, CA, CT, CO, FL, GA, LA, MI, MN, MO, NC, NY, OH, OK, PA, TN, TX, WA and WI; and countries of: Czech Republic, Denmark, France, Singapore, Spain, Switzerland, and United Kingdom. | Terminated | August 28, 2013 | 877 | 505 | Computer
Z-1863-2011 | 57742 | No | JJX | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Cardiac Marker Control Level 1, List Number 136601-06F12-02; Cardiac Marker Control Level 2, List number 136602/06F13-02; and Cardiac Marker Control Level 3, List number 136603/06F14-02. |  |  | December 17, 2010 | 2010 | Abbott Point Of Care Inc. | Presence of CO2 in the head space of the vial can result in the generation of high results in Cardiac Marker Control Level 1, 2, and 3 which is used along with the Cardiac Marker Control Calibration Verification Set. | Device Design | Urgent Recall Notice letters with business reply cards were sent Fed ex on December 17, 2010 to all accounts. International notifications were sent on December 21, 2010. The letter identified the affected product and lot numbers along with a background explanation of the issue. The firm recommends that customers remove the cap from the vial; squeeze the vial 10 times, prior to thawing; replace cap on the vial; and then follow the remaining steps for use as outlines in section 14-12 of the i-STAT 1 System Manual. Customers are to complete and return the Business Reply card to acknowledge receipt of the letter. If the affected product had been further distributed then a copy of the letter should be provided to those customers as well. If additional information is needed, customers should contact Abbott Point of Care Technical Support at 800-366-8020, Option 1 or their Abbott Point of Care representative. |  | 2360 | 2360 units total | Worldwide Distribution | Terminated | October 16, 2012 | 669 | 2360 | Computer
Z-2674-2011 | 58909 | No | LKN |  |  | Gastroenterology/Urology | 510(k) | Spectra Optia Apheresis System, REF: 61000, CaridianBCT, Lakewood, CO 80215. |  |  | December 09, 2010 | 2010 | Caridian BCT, Incorporated | An additional air detection system is being added to all existing systems to detect air in the system's return line. | Software design | CaridianBCT sent a "MEDICAL DEVICE RECALL" letter dated December 2010 to all customers. The letter describes the product, problem, and actions to be taken by the customers.Customers were advised to continue using their systems as long as they followed the instructions in the Operator's Manual. An Acknowledgement of Receipt form was included with the letter to be completed and faxed to 1-303-876-9277 or e-mailed to Regulatory.Affairs@caridianbct.com. A CaridianBCT representative will schedule a visit to facilities to upgrade their systems within the next 6 months.Further information is available by contacting the CaridianBCT Support Center at 877-339-4228 or 303-231-4357. |  | 594 | 594 units | Worldwide Distribution--USA (nationwide) and countries of Australia, Belgium, Canada, China, Hong Kong, Japan, Malaysia, Mexico, Pakistan, Saudi Arabia, Singapore, South Korea, Taiwan, Thailand, and Vietnam. | Terminated | September 12, 2011 | 277 | 594 | Computer
Z-2691-2011 | 59039 | No | MND | Hemorrhoidal ligator. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Disposable Hemostasis Clip, DCH-7-230, Rx only, Single Use only, Cook Medical |  |  | June 07, 2011 | 2011 | Cook Endoscopy | The disposable clip may not deploy after being positioned inside the patient. | Device Design | Wilson-Cook Medical, Inc. (Cook Endoscopy) sent an "URGENT: PRODUCT RECALL" notice dated June 7, 2011 and a "PRODUCT RECALL EFFECTIVENESS SURVEY' dated June 21, 2011 to all affected consignees. The notice described the product, problem, and actions to be taken. Medical users were instructed to quarantine any affected products and return to Cook Medical. The distributors were instructed to contact their customers and coordinate the return of the affected products to Cook Medical. The recalled notice and survey should be completed and returned to Cook Medical via fax to 812-339-7316 or scanned and emailed to CustomerRelationsNA@cookmedical.com. If you have any questions regarding this matter, please contact the Customer Relations Department at toll free-1-800-457-4500, press 4 then enter extension 2146 or by email at CustomerRelationsNA@cookmedical.com . |  | 307 | 307 units | (USA) Nationwide Distribution including the states of: CA, HI, MA, ME, NC and NY. | Terminated | May 15, 2012 | 343 | 307 | Computer
Z-2678-2011 | 58532 | No | DRY | Cardiopulmonary bypass on-line blood gas monitor. | Cardiovascular | Cardiovascular | 510(k) | Terumo Cardiovascular Systems, CDI blood parameter monitoring system 500 with Arterial and Venous blood parameter modules and Hematocrit/Saturation probe, Catalog AVHCT. |  |  | April 21, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | Inaccurate readings may result after Methylene Blue, an intravascular dye, is administered to the patient. The presence of Methylene Blue may interfere with the readings from the shunt sensor for the CDI 500 system in a way that would cause inaccurate pH values. Because the pH value is used to calculate other values, these parameters could also be affected by the presence of Methylene Blue: K | Device Design | Terumo sent URGENT MEDICAL DEVICE CORRECTION letters to Chief of Perfusion, Department of Cardiothoracic Surgery or Director of Operating Room Services, dated 4/21/2011. The letter explained the reason for the correction and the potential hazard. The letter stated that Terumo was updating the Operators Manual for all CDI systems and would provide the manuals to all users when available. Consignees with questions should call 1-800-521-2818. Monday- Friday 8 am- 6 pm. |  | 3767 | 3767 units (all CDI 500 products) | Worldwide distribution, including USA, Australia, Barbados, BELGIUM, BRAZIL, CANADA, Chile, Colombia, Costa Rica, Dominican Republic, ECUADOR, El Salvador, France, Germany, Guatemala, Holland, Honduras, Hong Kong, India, Indonesia, Japan, Jordan, Korea, Libya, MALAYSIA, MEXICO, PAKISTAN, PANAMA, Paraguay, Saudi Arabia, SINGAPORE, South Africa, Spain, TAIWAN, THAILAND, The Phillipines, TRINIDAD, UNITED ARAB EMIRATES, Uruguay, Venezuela and Vietnam. | Terminated | August 06, 2012 | 473 | 5582 | Computer
Z-2692-2011 | 58560 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Software versions |  |  | April 13, 2011 | 2011 | Eos Imaging | Print true size, print requests from the EOS acquisition station or the sterEOS image review station may, in certain cases, be affected by a different enlargement factorduring printing. T | Software design | EOS imaging sent a "FIELD SAFETY NOTICE" letter dated April 13, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter instructs customers to wait for the printing queue to empty before sending a second "print true queue" request with different print settings. Customers should then click the button shown on the instruction sheet to view the print queue.EOS imaging will provide a software update to correct all affected systems at no cost to the customers. An EOS representative will contact customers to schedule a time to upgrade the software.Questions can be answered by calling 33 1 55 25 63 90. |  | 15 | 5 EOS Systems, 10 sterEOS workstations | Worldwide Distribution--USA (nationwide) including the states of CA, IA, MO, VA, and WI, and countries of Canada and Europe. | Terminated | January 18, 2012 | 280 | 15 | Computer
Z-2685-2011 | 56408 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Beckman Coulter UniCel DxI 800 Access Immunoassay System with Unicel DXI system software version 4.3 |  |  | July 06, 2010 | 2010 | Beckman Coulter Inc. | Test results may have incorrect ranges. If a default sample type setting on the tests screen is changed while running Unicel DXI system software version 4.3, the new setting is not saved to the system database. The changed sample type setting will reset to the previously saved setting when power is interrupted to the system, or if an operator reboots the system. If the default sample type i | Software design | Beckman Coulter forwarded on 7/6/10 a Product Corrective Action letter with attached PCA Response Form (via US Postal Service for US Customers) to all customers who purchased the UniCel Dxl 600 and Dxl 800 Access Immunoassay systems. The letter provides the customers with an explanation of the problem identified and an action to be taken. Customers were instructed to:A revised version of UniCel DxI System Software that corrects this issue is currently in development. Until the new software version is available, Beckman Coulter has provided an alternate method for changing the default sample type setting. See the appendix included with this letter for additional information.Refer to the attached Questions and Answers document for additional details and instructions, and to determine if their laboratory is running UniCel DxI Sytem Software version 4.3. Customers may wish to confirm that proper ranges and units of measure are being reported for any Access assays with sample type settings that were changed while running software version 4.3.Customers were asked to share this information with their laboratory staff, and retain the notification as part of their laboratory Quality Ststem documentation. Complete and return the enclosed response form within 10 days so Beckman Coulter can be assured that consignees received the notification.If they need assistance or have any questions regarding this notification, customers were told to contact Technical Support at 1-800-854-3633 in the United States or Canada. Outside the United States and Canada, contact your local Beckman Coulter Representative. |  | 626 | 626 units | Nationwide in the US and to Canada | Terminated | May 07, 2012 | 671 | 794 | Computer
Z-1728-2011 | 57814 | No | JJQ | Colorimeter, photometer, or spectrophotometer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Hc2 System Software Suite 4.0 version 3.0, available as component in Qiagen user guide, catalog #5050-00162, and as a component within a Qiagen PC, available as catalog # 5050-11140A. |  |  | January 12, 2011 | 2011 | QIAGEN Gaithersburg, Inc. | The software component of a diagnostic medical device may be defective and could cause incorrect patient values. | Software design | The firm, Qiagen, sent a "Notice of Product Field Action" letter dated January 12, 2011 to all customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately discontinue the process of running multiple assay protocols on a single microplate to eliminate the potential for inaccurate patient result reporting, and to complete and return the acknowledgment section via fax to QIAGEN 661-702-3854 or scan the document and email to techservice-na@qiagen.com. The firm informed users that they will provide further detailed instructions for a workaround and software fix. A QIAGEN representative will contact the customers who indicated that they run multiple assays on a plate and are currently using the HC2 System Software Version 3.0 (Suite 4.0), to assist in the identification of those patients that may have been affected by the software defect.If you have any questions or concerns, please contact Qiagen Technical Services at 800-344-3631, Option 2. |  | 163 | 163 | The recalled product was sold to medical laboratories nationwide and in Canada. | Terminated | November 14, 2011 | 306 | 163 | Computer
Z-2673-2011 | 58878 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Instrumentation Laboratory GEM Premier 3500 Blood Gas Analyzer |  |  | May 11, 2011 | 2011 | Instrumentation Laboratory Co. | Software upgrade: misuse of the barcode gun by repeatedly clicking can cause the instrument to attach test results to an incorrect patient ID. | Software design | Instrumentation Laboratories sent a "FIELD SAFETY NOTIFICATION" letter dated April 2011 to all affected customers. The letter described the product, problem, and the actions to be taken by the customers.Included with the letter was an upgrade package containing Software V7.2.2 disk along with installation instructions. Also included with the upgrade kit was a Upgrade Tracking Form to be completed and returned via fax at 781-861-4207 or via e-mail at ra-usa@ilww.com.Customers with questions were directed to contact their local representative. |  | 840 | 840 units | Worldwide Distribution-- USA (nationwide) and the countries of AUSTRALIA, AUSTRIA, BELGIUM, CANADA, CHINA, CZECH REPUBLIC, DENMARK, EGYPT, FINLAND, FRANCE, GERMANY, GREECE , INDONESIA, ISRAEL, ITALY, JAPAN , KOREA, REPUBLIC OF CHINA, MALAYSIA, NETHERLANDS , PHILIPPINES, POLAND, PORTUGAL, SPAIN, SWEDEN, SWITZERLAND, THAILAND ,TURKEY, and UNITED KINGDOM. | Terminated | March 15, 2013 | 674 | 840 | Computer
Z-2662-2011 | 58186 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Elekta Impac Software Sequencer Verify &amp; Record System, for MOSAIQ |  |  | March 10, 2011 | 2011 | Impac Medical Systems Inc | There's a problem in the database conversion program that is used when upgrading from the Multi-ACCÂżS to MOSAIQ or from one version of MOSAIQ to another. If the gantry start angle is incorrect prior to pressing beam-on, the patient will be mistreated. | Software design (manufacturing process) | The firm, Elekta, sent a "USER NOTICE" to its customers with an attached confirmation of receipt form. The notice describes the product, problem and actions to be taken. The customers were instructed to follow the Interim mitigation instructions noted in the notice -verify the Arc start angle of the Simple ARC with MLC treatment field definition prior to delivery and if the customer identifes a field with an errant gantry start angle, to cease using that plan immediately and consult with the original plan for that field and make the correction to the gantry start angle. In addition, the customers were instructed to complete and return the Confirmation of Receipt via fax to: North America and Asia at 702-992-5002 or Europe and Africa +44 (0)1293 654401, Attn: Elekta OIS Su[pport and/or scan and email to: support@impact.com, europe.support@impact.com or by mail; and distribute this notice to any all users of IMPAC software at their organization who are potentially affected by this issue.For questions or requests for upgrade, email support@impac.com (North America/Asia) or europe.support@impac.com (Europe) or call USA 1-800-488-4672. |  | 840 | 840 | Worldwide distribution: USA (nationwide) and country of Canada. | Terminated | September 16, 2011 | 190 | 840 | Computer
Z-2601-2011 | 59000 | No | LDR | Infusion pump. | General Hospital | General Hospital | 510(k) | "***DOSI-FLOW 1 I.V. EXTENSION SET WITH I.V. FLOW REGULATOR REGULADOR DE CAUDAL PARA SETS DE INFUSION I.V.***REF 259-0250***50 UNITS***STERILE OE***LATEX-FREE***LEVENTON LEVENTON, S.A.U. C/Newton, 18-24 08635 Sant Esteve Sesrovires BARCELONA-SPAIN***Distributed by: Victus www.victus.com 4918 S.W. 74th Court Miami, FL 33155***" |  |  | March 11, 2011 | 2011 | Leventon S. A. U. | Leventon, S.A.U. Barcelona, Spain is recalling DOSI-FLOW REGULATOR, Part Number 259-0250, Lot Numbers 101444L and 101494L. Product has potential to leak. | Device Design | The firm, Leventon, S.A.U. sent an "URGENT: MEDICAL DEVICE RECALL" letter dated March 11, 2011 to the affected customer. The letter identifies the product, problem, and actions to be taken by the customer.The letter instructs the customer to examine their inventory and quarantine any product subject to the recall. They were also instructed to identify their customers who have received the product, notify them of the recall, send a copy of the recall notification letter and request that the customers return any unused product. The customers were also instructed to complete and return the attached RECALL RESPONSE FORM within 5 working days via fax to 00 34 93 6532556 or e-mail to dsalvatierra@leventon.es.Questions concerning this recall should be directed to the QA Manager at 00 34 93 8176316 or e- mail to dsalvatierra@leventon.es. |  | 1000 | 1000 cases (50,000 units) | Nationwide Distribution: FL only. | Terminated | June 28, 2011 | 109 | 1000 | Computer
Z-2663-2011 | 58442 | No | KZD | Pressure infusor for an I.V. bag. | General Hospital | General Hospital | 510(K) Exempt | Infusable and Infusa Scan Single Patient Use Pressure Infusors; Vital Signs Inc. |  |  | March 04, 2011 | 2011 | GE Healthcare, LLC | Infusable and InfusaScan Pressure Infusors may leak and not apply desired pressure to their IV solution bag and subsequently the rate of infusion may be inadequate. | Device Design | The firm, GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated February 25, 2011 with an attached return response form to its customers via Fed Ex. The letter describes the product, problem and actions to be taken. The customers were instructed to do the following:1) DO NOT use infusable or InfusaScan pressure infusors (listed lot numbers).2) Isolate all affected product and contact Vital Signs Customer Service at +1-973-956-5300 to arrange for the return and replacement of the infusor.3) If they have forwarded any of the affected product to any other healthcare institution, forward a copy of this letter to those institutions.4)Return only unused product.In addition, the customers were instructed to ensure that all potential users in their facility were made aware of this safety notification and the recommended actions, and to complete and return the attached confirmation form via fax to +1-973-956-5440.Please contact your local GE Healthcare Service Representative if you have any questions or concerns regarding this notification. You may also contact GE Healthcare Call Center: United States or Canada: 800-437-1171 and Other countries: 262-896-2890. |  | 954888 | 954,888 units | Worldwide distribution: USA (nationwide) states including: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, VA, VT, WA, WI, WV, WY; and countries of: Argentina, Australia, Chile, Colombia, Canada, Denmark, Dominican Republic, Finland, Germany, Greece, Hong Kong/Macao, India, Israel, Italy, Korea, Lebanon, Lativia, Mexico, Malaysia, Nicaragua, Netherlands, New Zealand, Paraguay, Philippines, Portugal Panama, Puerto Rico, Peru, Russia, Saudi Arabia, Singapore, Slovenia, Spain, Sweden, Taiwan, Turkey, United Kingdom, Uruguay, Venezuela, and Vietnam. | Terminated | October 31, 2012 | 607 | 954888 | Computer
Z-1882-2011 | 58044 | No | LLC |  |  | General Hospital | 510(k) | iSite PACS with iSite ViewForum Applications, R7.4V6 L1 L2, L3, |  |  | February 23, 2011 | 2011 | Philips Medical Systems | A software anomaly that occurs when clinicians use Volume Vision to view images in 3D and uses the measurement tool to measure regions of interest. This could lead to an incorrect interpretation of the measurements which could cause an incorrect diagnosis and incorrect treatment plan. | Software design | Philips sent an Urgent Field Safety Notice letter dated February 23, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to assure that all colleagues were aware of the problem.When using the Volume Vision software, check the correctness of measurements (graphics and numerical values) prior to saving a captured image containing measurements.Save captured images containing measurements immediately after creating them and before altering the measurements in the original image view.For any subsequent measurements, create a new capture and save it immediately.For further information or support concerning this issues, customers should contact their local Philips representative at +1-877-328-2808. |  |  |  | Worldwide Distribution. | Terminated | August 18, 2011 | 176 |  | Computer
Z-1893-2011 | 58009 | No | KQM | Surgical camera and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ***REF VP6528***Qty 1***Articulating Arm, Secondary Flat Panel***RX Only***See Instructions for Use***CONMED LINVATEC, CONMED LINVATEC, 11311 CONCEPT BLVD, LARGO, FL 33773-4908 USA*** |  |  | September 17, 2010 | 2010 | Linvatec Corp. dba ConMed Linvatec | If the VHM Arm is left in the locked position, the gas spring can lose pressure and releasing the lock can result in a sudden downward movement. | Device Design | An Urgent Service Bulletin Notification was mailed to each ConMed Linvatec customer who received the affected product. The letter identified the affected product and stated the reason for notification. The letter asks customers to review the GCX Corporation Service Bulletins and updated Installation/User Manual. ConMed Linvatec recommends that customers periodically check the VHM Arm to see that it is functioning properly. In addition "Sudden Movement" stickers, provided by GCX, were also included to all consignees as part of the mailing. If customers in the US have questions or need technical support, they can contact ConMed Linvatec Endoscopy Customer Service at 888-292-0100 or e-mail Customerserv1@linvatec.com. |  | 373 | 373 | Worldwide Distribution -- Australia, Canada, Morocco, Saudi Arabia, Taiwan, Province of China, and throughout the U.S. | Terminated | May 17, 2011 | 242 | 373 | Computer
Z-2645-2011 | 58206 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ERGO++ Stereotactic Radiation Treatment Planning System, ERGO++ Release 1.6.3 and 1.6.3.1 |  |  | December 11, 2009 | 2009 | Computerized Medical Systems Inc | ERGO Release 1.6.3 is overestimating the MU values. | Software design | Computerized Medical Systems, Inc. sent an "USER NOTIFICATION" dated December 11, 2009, to all affected customers. The notice identified the product, problem and suggested a workaround to minimize the risk. The notice stated that this workaround does NOT substitute the need for regular Quality Assurance. Customers are instructed to review page 39 of user manual for further explanation of warnings. Customers were asked to complete the Notice for Customer section and return to 3D Line Medical System. This issue has been resolved in Release 1.7.0 and above. No follow-up notice will be distributed to customers. For further information please contact: CMS Software (314) 812-4460 (St Louis Office) |  | 45 | 45 | Nationwide distribution to medical facilities in CA, MO, NC, ND, NV, NY, OH, OK, TX, and WA. Foreign distribution was made to Brazil, Cuba, Czech Republic, France, Germany, India, Italy, Japan, Malaysia, Poland, Portugal, Serbia, and Spain. | Terminated | April 01, 2014 | 1572 | 45 | Computer
Z-1868-2011 | 57939 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Curlin 6000 IOD Ambulatory Infusion Pump |  |  | February 28, 2011 | 2011 | Zevex International, Inc. | Infusion pumps may exhibit false error codes which render the device inoperable until a recovery process is performed. | Software design | The firm, MOOG, sent an "URGENT DEVICE RECALL NOTIFICATION" letter dated February 28, 2011 to its customers. The letter described the product, problem and actions to be taken. The Customers were instructed to do the following:1) Identify the affected Pumps:a) If they have an affected pump, fill out and return the enclosed Business Reply Card within five (5) working days.b) If their pump is not affected by this recall, the pump can be used for its intended purposes.2) Removal of affected Product:a) remove affected products from service and use non-affected pumpb) If you do not have an alternate pump available for use, MOOG recommended that they weigh the risks and benefits for delivering medications using the affected pump versus the risk of not using any pump for each patientWARNING: Affected pumps should not be used to deliver critical medications where a delay in therapy could result in serious adverse health effects, injury or death; and affected pumps should not be used with the neonatal population. 3)Should they Have an Affected pump which exhibits Error Code 45 or is used in therapeutic situations specified in the warning, please contact MOOG Customer Service at (800) 970-2337, Monday thru Friday, 7:00am to 5:00pm MST to arrange for a loaner pump and the return of your affected pump.If you have any additional questions regarding this recall, please contact MOOG's Customers Service Department at (800) 970-2337. |  | 3407 | 3,407 units | Worldwide distribution: USA and countries including: Australia, Canada, Costa Rica, Croatia, Dubai, France, Germany, Ireland, Italy, Japan, Jordan, Kuwait, Malaysia, New Zealand, Qatar, Saudi Arabia, Singapore, South Africa, Taiwan, Turkey, UAE, United Kingdom. | Terminated | August 06, 2012 | 525 | 18306 | Computer
Z-2627-2011 | 58873 | No | MDB | Microbial growth monitor. | Microbiology | Microbiology | 510(k) | BACTEC" FX Instrument-Top, catalog #441385. |  |  | August 10, 2010 | 2010 | Becton Dickinson &amp; Co. | In vitro diagnostic instrumentation may exhibit random operational errors that could delay effective patient treatment. | Software design | BD Diagnostics Systems sent an "URGENT PRODUCT RECALL" letter dated August 2010 to all affected customers. The letter describes the product, problem and action to be taken.The letter described the type of error messages the customers may experience. All affected units will be scheduled for on site software update as well as a disk drive replacement. A BD field service engineer will contact the customers to schedule the on site visit.Customers experiencing issues prior to the scheduled visit or have questions should contact the BD Diagnostics Service Department at 800-638-8663. |  | 907 | 907units | Worldwide Distribution--USA (nationwide) including countries of Australia, Belgium, Canada, China, Colombia, New Zealand, Hong Kong, Japan, Korea, Philippines, Singapore, Taiwan, and Thailand. | Terminated | July 05, 2011 | 329 | 907 | Computer
Z-2622-2011 | 58228 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Interplant Radiation Treatment Planning System, Interplant Release 3.4.0 |  |  | December 01, 2008 | 2008 | Computerized Medical Systems Inc | Interplant Software: If the trajectory needle with a non-zero retraction is adjusted in any of the planning views (Transverse, 3D) and the program is closed and reopened, the Needle Definitions page in the Dose Plan Summary (Plan View) displays the needle with a 0.00 retraction even when the needle is displayed correctly in the planning view. If the Eliminate Leading Spacers function is applied t | Software design | Computerized Medical Systems, Inc. sent a "CUSTOMER ADVISORY" letter dated December 1, 2008, to all affected customers. The letter identified the product, the problem, and the actions to be taken. Customers were instructed to discontinue use of the product. There is no workaround at this time. This issue was resolved in Interplant Release 3.4.1 software. All possibly affected users have been shipped the Interplant Release 3.4.1 software in May 2010. If you have any questions, please call 314-993-0003 or toll free 800-878-4267. |  | 41 | 41 | Worldwide Distribution -- USA (nationwide) including states of: CA, KS, KY, MA, MD, ME, MO, NC, NE, NY, OH, PA, TX and VA and countries of: Australia, Brazil, Chile, China, Germany and Japan. | Terminated | April 04, 2013 | 1585 | 41 | Computer
Z-1900-2011 | 51497 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Siemens brand ACUSON S2000 Ultrasound System with software versions: |  |  | January 14, 2009 | 2009 | Siemens Medical Solutions USA, Inc. | The product has a software problem in which previous patient measurement data gets associated with another patient's image. | Software change control | The firm, SIEMENS, issued a "CUSTOMER SAFETY ADVISORY NOTIFICATION" letter to all customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure that the first image in Exam Review is from the current patient; ensure that the Heart Cycle measurement tool is not selected on the Cale menu when entering Exam Review; be aware of this issue as described in the letter and perform four (4) or less measurements of the same measurement label. Should you have any questions, please call (650) 694-5398. |  |  |  | Worldwide distribution. | Terminated | April 06, 2011 | 812 |  | Computer
Z-2581-2011 | 58854 | No | DXQ | Blood pressure cuff. | Cardiovascular | Cardiovascular | 510(k) | Cuffable Blood Pressure Cuffs; Manufactured for Vital Signs, Inc., a GE Healthcare Company 20 Campus Road, Totowa, NJ 07512 Made in Mexico. |  |  | April 26, 2011 | 2011 | GE Healthcare, LLC | BP cuffs may not properly inflate due to a leak | Device Design | GE Healthcare sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated April 26, 2011 to all affected customers. The letter identified the product, the problem and the action to be taken. Customers were instructed to discontinue use of the product, isolate all affected products, return only unused products and distribute this letter to any other healthcare institutions that are potentially affected by this issue. The letter asks that each customer complete and return the enclosed confirmation form via fax to 800-535-7923, in order to confirm receipt of the notification letter. For any questions, contact your local Vital Signs Account Manager or the International Customer Service at 1-800-932-0760. |  | 2684996 | 2,684,996 units | Worldwide Distribution -- USA (nationwide) including Puerto Rico and countries of: Argentina, Australia, Belgium, Canada, Chile, Ecuador, Finland, Germany, Hong Kong, Israel, Italy, Latvia, Lebanon, Nicaragua, Panama, Peru, Saudi Arabia, Slovenia, UK, Uruguay and Venezuela | Terminated | May 28, 2013 | 763 | 2684996 | Computer
Z-2611-2011 | 59004 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Pathfinder Holter Analyzer is a Programmable Diagnostic Computer. |  |  | May 24, 2011 | 2011 | Spacelabs Healthcare, Llc | A potential hazard associated with the use of Spacelabs Healthcare Pathfinder Holter Analyzer software. The second 24 hours of a 48 hours EVO recording will not be analyzed. | Software design | Spacelabs Healthcare sent a "URGENT-MEDICAL DEVICE CORRECTION" letter dated June 3, 2011 to all customers. The letter describes the product, problem, and the actions to be taken by the customers. Spacelabs Healthcare recommended that until software upgrades can be done, the customers should weigh the benefits versus the risks when deciding whether or not to continue using the product. Spacelabs Healthcare also recommended that customers advise their staff that the second 24 hours of a 48 hour EVO recording would not be analyzed.Field service engineers will be contacting customers to set up a time and date to install the revised version of the Pathfinder software. Spacelabs will upgrade Pathfinder software regardless of whether the consignees are using EVO recorders to ensure this hazard does not happen in the future.Consignees can call Spacelabs Healthcare at 1-800-522-7025 and select 2 for Technical Support, or outside the US, 44 1992 507740 for any questions about this recall. |  | 58 | 58 units total (5 units in US and 53 units internationally) | Worldwide Distribution--USA (nationwide) including states of CA and NC and countries of Czech Republic, Denmark, Hong Kong, Israel, The Netherlands, and United Kingdom. | Terminated | October 26, 2011 | 155 | 58 | Computer
Z-2625-2011 | 58822 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | 12-lead ECG Cable Assembly contains the following cable and attachments: |  |  | May 11, 2011 | 2011 | Physio Control, Inc. | The 12-lead ECG cables have experienced premature failures which cause the shielding and conductors to break internally creating ECG noise or "lead off" indications. | Device Design | Physio-Control, Inc. sent an "URGENT-MEDICAL DEVICE ACCESSORY INFORMATION" letter dated May 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers.The customers were advised to follow the instructions:-Use standard hard paddles or therapy electrodes for ECG monitoring if monitoring cannot be performed using the ECG cable. Refer to Troubleshooting Tips for ECG Monitoring in the LIFEPAK 12or LIFEPAK 15 Operating Instructions-Monitoring chapter.-Disconnect the ECG cable from the monitor after use and coil it loosely for storage. Please refer to the Storage Instructions for 12-lead ECG Cable.-Inspect cables daily for bulges which may indicate damaged internal wires. If bulges are visible, replace the ECG cable.-Perform periodic 6 month function testing: Refer to Maintenance and Testing Schedule in the LIFEPAK 12 or LIFEPAK 15 Operating Instructions- Maintaining the Equipment chapters for instructions on how to perform a patient ECG cable function check to confirm proper operation ofthe cable.Customers can visit Physio-Control's website at www.physio-control-notices.com/12LeadECG for additional information or todownload copies of current LIFEPAK 12 defibrillator/monitor or LIFEPAK 15 monitor/defibrillator Operating Instructions and Operator's checklists.Customers can also call Physio-Control's Technical Support Team at 1-800-442-1142, option 5, from 6:00 A.M. to 4:00 P.M. (Pacific), Monday-Friday. |  | 21825 | 21,825 units distributed worldwide | Worldwide Distribution--USA (nationwide) including Puerto Rico and countries of Australia, Brazil, Chile, Germany, Hong Kong, Mexico, The Netherlands, and New Caledonia. | Terminated | August 12, 2011 | 93 | 21825 | Computer
Z-2610-2011 | 58779 | No | DRY | Cardiopulmonary bypass on-line blood gas monitor. | Cardiovascular | Cardiovascular | 510(k) | CDI TM 101 Monitor, CDI TM 101 Hematocrit/Oxygen Saturation Monitoring System, |  |  | April 28, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | Terumo discovered through its own testing that the presence of electromagnetic interference (EMI) in the operating room may affect the accuracy of the valuesdisplayed by the CDI"101 Hematocrit/Oxygen Saturation Monitoring System: hematocrit (HCT), hemoglobin (Hgb) and Oxygen Saturation. | Device Design | Terumo Cardiovascular Systems, Corp issued an "URGENT MEDICAL DEVICE CORRECTION" letter dated April 28, 2011, to all affected customers. The letter identifed the product, the problem and the action needed to be taken by the customer. After a solution has been verified, Terumo CVS will implement a correction for all affected units. Customers were instructed to review this Medical Device Correction notice, Assure that all users are aware of this notice. Confirm receipt of this and to fax back the Customer Response form.For further question please contact Terumo CVS Customer Service at 1-800-521-2818, Monday - Friday, 8 AM - 6 PM EST. |  | 622 | 622 | Worldwide Distribution -- USA ( nationwide) and the countries of Colombia, Vietnam, India, Indonesia, Pakistan, Canada, Chile, Vietnam, Chile, Hong Kong, United Arab Emirates, Singapore, Australia, Malaysia, Singapore, Taiwan, Belgium, Thailand, Vietnam and India | Terminated | November 06, 2012 | 558 | 622 | Computer
Z-1899-2011 | 57855 | No | KSD | Heat-sealing device. | Hematology | Hematology | 510(K) Exempt | 2380 MINI Hand-Held Tube Sealer Power Source |  |  | January 31, 2011 | 2011 | Haemonetics Corporation | Battery pack may have exposed wire becoming hot | Equipment maintenance | The firm, Haemonetics, issued a "Field Notification: 2380 Sealer battery pack" letter dated January 28, 2011 to its customers. An international English version of this letter (CL101020-IE) has been released to foreign Distribution countries. The letter described the product, problem and actions to be taken. The customers were instructed to do the following:1) Record the manufacturing date of battery, equipment model, serial number, and photo of the affected battery.2) Report the defect through the Haemonetics PIR system by emailing: sebra2380@haemonetics.com with all of the information listed including the appropriate contact information for their company.3) Recycle or dispose of the nickel-cadmium (NiCad) battery pack in accordance with the local, state, federal, or country specific regulations. Under no circumstances should the battery pack be incinerated as this can cause explosion and personal injury. The operation manual will be updated to provide additional information about the proper installation and removal of the battery. This updated manual will be distributed with new devices upon release.Should you have any questions or concerns regarding this matter, please contact the firm at sebra2380@haemonetics.com. |  | 50031 | 50,031 units | Worldwide distribution: USA and countries including: Australia, Austria, AEU, Belgium, Bangladesh, Canada, China, France, Germany, Hong Kong, India, Indonesia, Italy, Japan, Korea, Malaysia, Mexico, New Zealand, Pakistan, Philippines, Singapore, Sri Lanka, Taiwan and Turkey. | Terminated | March 06, 2012 | 400 | 50031 | Computer
Z-1901-2011 | 58208 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Prismaflex System, Catalog Number 107493, |  |  | March 14, 2011 | 2011 | Gambro Renal Products, Incorporated | Prismaflex touch screen may become unresponsive during Continuous Renal Replacement Therapy procedure. | Software change control | The firm, GAMBRO Renal Products, sent an "URGENT - MEDICAL DEVICE RECALL" letter dated March 14, 2011 to all customers. The letter described the product, problem and actions to be taken. The customers were provided with response procedures to use if the frozen screen problem occurs while using the product. The customers were instructed to closely monitor the return line for potential air and/or clots in the circuit; review the instructions provided in the attached copy of Chapter 5: Troubleshooting, "Softkeys won't work." of the Prismaflex Operators Manual in conjunction with the attached Troubleshooting Checklist; be aware if performing Manual Termination with Blood Return, air detection is not provided, and complete and return the Customer Reply Form via fax to: Regulatory Affairs at 1-866-447-4471 or 303-222-6916, or by email to: GRPRegulatory.Affairs@us.gambro.com. If you have any questions or concerns, please contact for Clinical questions: ICON 800-554-1312 (available 24/7) or your local Gambro Representative; for Technical questions: Gambro Customer Support 800-525-2623, option 2 (Available 6:00am to 5:30pm MT, Monday thru Friday). |  | 1300 | 1300 units | Nationwide distribution. | Terminated | September 01, 2011 | 171 | 1300 | Computer
Z-1932-2011 | 54483 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | Bio Rad brand VARIANT II TURBO Hemoglobin Alc Program, 1600 and 2000 tests Cartridge Sets. Catalog Numbers: 270-2417 and 270-2415; Distributed and Manufactured by: Bio-Rad Laboratories, Inc. Hercules, CA |  |  | December 01, 2009 | 2009 | Bio-Rad Laboratories Inc | Some customers have reported the occurrence of some ramping baselines on the chromatograms. The ramping baseline can affect hemoglobin test results. | Software change control | Bio Rad Laboratories, Inc. sent an Urgent Medical Device Correction letter dated December 1, 2009, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediatey examine their records to identify all customers who received one of the affected lots.Collect a Customer Medical Device Correction Response Form from every customer.Once forms were received from all applicable customers, complete the following Medical Device Correction Response Form and fax it to Bio-Rad CSD Regulatory Affairs at 510-741-3954 or pdf to CSD Recall Coordinator (excluding USSD).For questions regarding this recall call 510-724-7000. |  | 1099 | 1099 | Worldwide Distribution - USA (nationwide) and the countries of Australia, New Zealand, Taiwan, Singapore, India, China, S. Korea, Hong Kong, France, Sweden, United Kingdom, Italy, Hungary, Germany, Brazil, Mexico and Colombia. | Terminated | August 30, 2011 | 637 | 1099 | Computer
Z-1961-2011 | 52928 | No | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | Auxiliary 2 MHz Continuous Wave (CW) transducer, model number 8251046, when used with ACUSON Sequoia ultrasound systems at SW revision 12.X. Manufactured by Siemens Medical Solutions USA, Inc., Ultrasound Business Unit, Mountain View, CA. |  |  | March 21, 2009 | 2009 | Siemens Medical Solutions USA, Inc. | Overheating. A software bug may cause the transducer to overheat when used with certain transducer models. The transducer can become noticeably hot within a few minutes, which may cause patient discomfort or burns. | Software design | The firm, SIEMENS, sent a "Customer Safety Advisory Notification" to its customers on 3/21/2009. The notification described the product, problem and actions to be taken. The customers were instructed to disconnect any transducers that may be in an active transducer port and chose another imaging transducer prior to use of the Aux CW transducer. If they do not have another transducer to select other than those listed in the notification, they should close the study in progress and start a new study before using the Aux CW transducer-this procedure should be followed for each new patient and /or exam type. In addition, the customers were instructed to pass this notice on to all those within their organization who need to be aware of this issue until the corrective action is completed. SIEMENS will release a software update that will correct the issue. The customers local service engineer will install the updates on their system as soon as it becomes available.Should you have any questions regarding this notification, you may contact your local service support person and/or call (650) 969-9112. |  | 2591 | 2591 units | Worldwide distribution. | Terminated | March 09, 2012 | 1084 | 2591 | Computer
Z-0011-2012 | 59956 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Plum A+ Hyperbaric Single Channel Infusion Pumps; Hospira, Inc., Lake Forest, IL 60045; list 11005 |  |  | September 16, 2011 | 2011 | Hospira Inc. | Hospira has received reports of continuous recycling and/or rebooting of Plum A+ infusion pumps when the backlight intensity and/or display contrast settings for the LED display has been adjusted from the original default setting. This condition could cause a delay in therapy if the clinician is unable to start the device due to the continuous recycling and/or rebooting. | Software design | The firm, Hospira sent "URGENT DEVICE RECALL" letters dated September 16, 2011 to all affected customers. The letters identified the products, problem and action(s) to be taken. The customers were instructed to read the information provided carefully, complete the attached Reply Form and return it to Hospira via fax at 1-888-345-5358. The letter states that a software upgrade to address the issue is currently being developed. Customers were instructed to notify their consignees if they have further distributed the affected product. The customers were requested to contact Hospira Advanced Knowledge Center at 1-800-241-4002, option 4, for technical assistance. *** Follow-up letters dated June 26, 2012 were sent to the accounts listing the device modules with software versions v13.2, v13.3, v13.40 and v13.5 that will require the software upgrade. *** |  | 120 | 120 pumps | Worldwide Distribution - USA (nationwide) including Puerto Rico and the US Virgin Islands, and countries of: Argentina, Australia, Bahamas, Barbados, Belgium, Brazil, Canada, Chile, Colombia, Costa Rica, Denmark, Dominican Republic, Ecuador, France, Greece, Hong Kong, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Korea, Kuwait, Lebanon, Libya, Malaysia, Mexico, Netherlands, New Zealand, Oman, Pakistan, Peru, Philippines, Poland, Portugal, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, United kingdom, United Arab Emirates, Uruguay and Vietnam. | Open, Classified |  |  | 290217 | Computer
Z-2182-2011 | 58466 | No | IKZ | Powered patient rotation bed. | Physical Medicine | Physical Medicine | 510(K) Exempt | KCI RotoProne Critical Care Therapy System buckles for securing patient to the surface during proning therapy, Model 209500. |  |  | February 18, 2011 | 2011 | KCI USA, Inc. | Buckles used to secure patient to surface during proning therapy can come open if not properly engaged by user. | Device Design | Firm sent a letter with the "Warning: Improper Insertion of the RotoProne Therapy System Buckle Can Lead To Unintentional Bed Exit" to all consignees on 2/23/2011. The letter identified the affected product and the reason for the recall. Customers were instructed to read the attached product warning regarding the proper method to attach the RotoProne buckles and to post it with the device to ensure continued proper use of the buckles. If customers had questions regarding the proper use of the buckles, they were to contact KCI Customer Technical Service at 1-800-275-4524. |  | 229 | 229 units | Worldwide Distribution - USA and Kuwait. | Terminated | October 24, 2011 | 248 | 229 | Computer
Z-2047-2011 | 55154 | No | CFT | Prolactin (lactogen) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Active Prolactin ELISA, 96 Wells, DSL-10-4500 |  |  | August 03, 2009 | 2009 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed that the Active Prolactin ElLISA kit, lot number 890291 over estimates the expected dose values for patient serum samples due to a shift in the assigned calibrator values.Beckman Coulter recommends that patient results generated with the affected lot be evaluated in the context of other diagnostic tests and the clinical presentatio | Equipment maintenance | The firm, Beckman Coulter, sent an "URGENT: PRODUCT CORRECTIVE ACTION" Letter dated August 03, 2009 to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to do the following:Action/Resolution:Discontinue use and discard remaining product with lot number 890291 of Active Prolactin reagents. Contact Beckman Coulter at 1-800-231-7970 for replacement product.Newer lots (lot number 991210 or higher) of the product have been realibrated to restore the performance as described in the Instructions for Use (IFU). For method correlation information, they were told to contact Technical support at the phone number listed below.Consignees were asked to share this information with their laboratory staff and retain this letter for their Quality System documentation and asked to complete and return the enclosed response form within 10 days via fax to: 786-639-4000 so that Beckman can be assured that they received the notification.If you have any other questions regarding this notification, please contact DSL Technical Support Center at 1-800-231-7970 in the United States and Canada , or contact their local Beckman Coulter Representative. |  | 30 | 30 Kits | Nationwide distribution: USA including states of: FL, GA and NC. | Terminated | April 27, 2012 | 998 | 30 | Computer
Z-2200-2011 | 58505 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Optical Guidance Platform, Version 2.6 and 2.6.1, Model Number: HZl, Manufactured and Distributed by: Varian Medical Systems Inc., Palo Alto, CA. |  |  | March 30, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | A software anomaly has been identified with the Optical Guidance Platform (OGP) Software v2.6 and v2.6.1 where the transfer of datasets from treatment planning systems other than the FastPlan system result in a lateral offset error. | Software design | On 3/30/2011, Varian distributed an Urgent Medical Device Correction Urgent Field Safety Notice to all affected users, with a description of the problem and identification of the affected product. The letter also informed customers of how they can determine if the problem affects them. The letter states that there is no work around for the issue unless they possess a FastPlan system. Varian is preparing a software patch to correct the issue and it will be applied to all OGP v2.6 and v2.6.1 systems. Customers are to advise the appropriate personnel of the content of the letter. If further clarification is needed, customers are to contact their local Varian Customer Support District or Regional Manager. |  | 598 | 598 units | Worldwide Distribution | Terminated | June 19, 2012 | 447 | 598 | Computer
Z-2064-2011 | 58160 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Defibtech Lifeline AED and ReviveR AED, |  |  | April 01, 2011 | 2011 | Defibtech, LLC | Software with version 2.004 software or earlier may cause an affected AED to cancel shock during the charging process | Software design (manufacturing process) | The firm, Defibtech, sent an "URGENT MEDICAL DEVICE SAFETY INFORMATION AND CORRECTION" letter dated March 11, 2011 to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to do the following:1) Remove the pads from the storage pouch on the back of their DDU-100 series AED and inspect the 9-digit serial number label.2)Carefully note the 9-digit serial number on the back of the AED, then visit the Defibtech website at www.defibtech.com/fa11 and enter their serial number to determine if their AED is affected; email their serial number and contact information to fa@defibtech.com or call Defibtech directly at 1-877-453-4507 or 1-203-453-4501 and Defibtech will inform them if their unit is affected.3) If their AED is affected, confirm receipt of this notification with Defibtech as soon as possible by website and/or confirmation card. Complete and return the confirmation form via fax to 1-203-738-1072. Note: "it is very important for you to contact us either through the Defibtech website or return this card to confirm that you have received this notice. This information will expedite the process of upgrading your AEDs."If your DDU-100 series AED serial number DOES NOT fall within the listed range, your AED is not affected by this recall.A letter was also sent to the distributors and direct accounts that have received the affected product by first class mail with a prepaid, return written acknowledgement card. The field action consists of a software change to the AED. All affected AEDs operating with software version 2.004 or earlier are to remain in the field and in service throughout the field action process.Upon receipt of contact information from affected customers, Defibtech will send an upgrade kit to customers consisting of a software data card, instructions on how to perform the upgrade, and a pre-paid upgrade confirmation postcard that must be returned to Defibtech upon completion |  | 91472 | 91, 472 units | Worldwide distribution: USA (nationwide) and countries including: AUSTRALIA, AUSTRIA , BAHRAIN, BRAZIL, CANADA, CHILE, CHINA , CYPRUS, CZECH REPUBLIC, EGYPT, EL SALVADOR, FRANCE, GERMANY, GREECE, HONG KONG, HUNGARY, INDIA, INDONESIA, IRELAND, ISRAEL, ITALY, JORDAN, KOREA, REPUBLIC OF LITHUANIA, MALAYSIA, MEXICO, NETHERLANDS, NORWAY, PHILIPPINES, POLAND, PORTUGAL, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SWITZERLAND, THAILAND, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM AND URUGUAY. | Open, Classified |  |  | 91472 | Computer
Z-2051-2011 | 58231 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON S2000 ultrasound system |  |  | March 15, 2011 | 2011 | Siemens Medical Solutions USA, Inc. | The Firm discovered a software error on the Accuson S2000 ultrasound systems affecting the measurement results in the patient report. | Software design | Siemens issued a Customer Safety Advisory Notification letter to all customers on March 15, 2011, regarding a software error on the ACUSON S2000 ultrasound systems affected measurement . They are in the process of releasing a software update that will correct this issue. Siemens recommends reviewing the patient report at the end of each study to confirm that the report contains only measurements relevant to the exam just performed. No patient injury has been reported. Further questions regarding this recall please call (650) 694-5398. |  | 1533 | 1533 units | Worldwide Distribution including USA, states of AZ, VA, CA, NE, IL, MN, TN, FL, NM, and DC | Terminated | June 18, 2012 | 461 | 1533 | Computer
Z-2061-2011 | 53053 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | Siemens ACUSON X300 Premium Edition Systems with software versions 5.0.00 and 5.0.01 (model numbers 10348531, 10348532 and 10348533), when using the V5Ms transducer (model number 08263703) and the Plus option (10427795) |  |  | July 24, 2009 | 2009 | Siemens Medical Solutions USA, Inc. | Product Failure-- When the 64-Channel System V5M is used in the 128-Channel mode, an over current fault occurs which shuts the system down. | Software design | Siemens sent a Customer Safety Advisory Notification letter on July 24, 2009. The letter identified the product, the problem, and the action to be taken by the customer. A software update will be conducted in the field. Until that update occurs, customers were instructed to image without the CW doppler to avoid a system shut down.Customers were instructed to contact their service support person with any questions.For questions regarding this recall call 650-694-5993. |  | 70 | 70 units | Worldwide Distribution - USA including CA, CT, FL, MA, and TX and the countries of United Arab Emirates, Germany, Ecuador, Spain, France, Great Britian, Saudi Arabia, Sweden, Thailand, and South Africa. | Terminated | April 25, 2011 | 640 | 70 | Computer
Z-2188-2011 | 57194 | No | DRY | Cardiopulmonary bypass on-line blood gas monitor. | Cardiovascular | Cardiovascular | 510(k) | CDI Blood Parameter Monitoring System 500 with Arterial Blood parameter module Catalog number 500A , Terumo Cardiovascular Systems Corporations, Ann Arbor, MI. |  |  | November 01, 2010 | 2010 | Terumo Cardiovascular Systems Corporation | The CDI 500 monitors may display inaccurate values. These errors include inaccurate potassium, CO2, hematocrit, O2 saturation, pH, hemoglobin and other miscellaneous inaccurate values. | Device Design | The firm, TERUMO, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated October 27, 2010 to its customers. The letter described the product, problem and actionst to be taken. The customers were instructed to reveiw this notice; assure that customers are ware of this notice; confirm receipt of this notification by faxing the attached Customer Response Form, and place the attached stickers in a visible place in the Operators' Manuals, such as the inside of the front cover, if additional stickers are needed, the customers should contact Terumo CVS Customer Service at 1-800-521-2818. The letter states that when a displayed value is significantly different from expectations based on the clinical situatioin, the user should verify the accuracy by independent means before treating. The letter included "Caution" stickers that were to be applied to the CDI system operatros manual.The letter also states "Please note that we are not recommending that you stop using your CDI monitor system."If you have any questions or concerns, please contact Terumo CVS Customer Service at 1-800-521-2818. Customer Service hours are from Monday-Friday, 8AM - 6PM EST. |  | 3260 | 3260 all products | Worldwide distribution: USA (nationwide) and countries of : Australia, Barbados, Belgium, Brazil, Columbia, Costa Rica, Dominican Republic, Ecuador, El Salvador, Guatamala, Hong Kong, Honduras, Jordan, Japan, Libya, Mexico, Malaysia, Panama, PH, Saudi Arabia, Singapore, South Africa, Thailand, Trinidad, Taiwan and Uruguay, Germany, Saudia Arabia, Paraguay, South Korea, Dominican Republic, Phillipines, South Africa and UAE. | Terminated | July 31, 2014 | 1368 | 3260 | Computer
Z-2065-2011 | 53209 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Acuson X300 systems at software versions 1.00.00 to 1.0.10 and 2.0.00 to 2.0.02, model numbers 10037409, 10132987, and 10038837, manufactured by Siemens Medical Solutions ultrasound business unit, Mountain View, CA |  |  | July 30, 2009 | 2009 | Siemens Medical Solutions USA, Inc. | Software Error with potential for misinterpretation of data-- The previous patient's last capture clip image can appear in the next patient's study. | Software design | The firm, Siemens, sent a "CUSTOMER SAFETY ADVISORY NOTIFICATION" letter to its customers on July 30, 2009. The letter described the product, problem and actions to be taken. The customers were instructed they need to update any of the patient registration settings to reset the cine buffer so that no data is transferred between studies, if the system software is not being updated. Siemens informed the customers that a mandatory update software is being installed on their system free of charge.If you have any questions, contact (650) 969-9112. |  | 2312 | 2312 units | Worldwide distribution: USA (nationwide) and countries including: Albania, algeria, Argentina, Australia, Austria, Azerbaijan, Belgium, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, Chile,China, Colombia, Costa Rica, Croatia, Czech Republic, Denmark, Ecuador, Egypt, El Salvador, Estonia, Finland, France, Georgia, Germany, Greece, Hungary, India, Indonesia, Iran, Ireland, Italy, Japan, Kazakhstan, Kenya, Kosovo, Lativia, Lithuania, Luxembourg, New, Zealand, Nicaragua, Norway, Paraguay, Peru, Philippines, Poland, Portugal, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uzbekistan, Venezuela, Viet Nam. | Terminated | April 25, 2011 | 634 | 2312 | Computer
Z-2173-2011 | 58180 | No | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | NEXGEN COMPLETE KNEE SOLUTION FEMORAL AND PROVISIONAL IMPACTOR/EXTRACTOR, NONSTERILE, REF 00-5901-026-00, Zimmer UK, Ltd., Zimmer Warsaw, IN. |  |  | March 14, 2011 | 2011 | Zimmer Inc. | The firm is initiating a recall due to the potiential for the spring clip to break off during use and fall into the surgical site. The firm has received 24 complaints of the spring breaking, four of which were reported that the patient had to undergo additional x-rays to ensure the fractured clip had not been left in-vivo. | Device Design | The firm, Zimmer, sent an "URGENT MEDICAL DEVICE RECALL" notice dated March 14, 2011 to its Consignees/customers. The letter described the product, problem, and the actions to be taken. The customers were instructed to stop using the device and quarantine immediately; carry out a physical count of all affected product in their territory and record this data on the Inventory Return Certification Form and fax to (574) 372-4265; ensure they have provided acknowledgement via "email return receipt" upon receiving notification; return the recalled product along with the completed Inventory Return Certification Form and Sterilization Certification Form to: Zimmer Distribution Center, Attn: Product Service, 1777 West Center Street, Warsaw, IN 46580-Note: The instrument must be returned even if the spring clip is not missing from the instrument; and if they distributed these affected lots further, provide a copy of this letter to the those customers when they implement these recall instructions. The letter stated that a new or refurbished instrument with a redesigned spring clip will be provided as replacement for the recalled device. For shipping assistance, questions or assistance in notifying your accounts about the recall, please contact Zimmer, Inc. at 1-800-613-6131. |  | 802 | 802 devices | Worldwide distribution: USA (nationwide)including states of: HI, CA, NY, MI, IN, GA, FL, WI, MN, NJ, IL, VA, TX, NC, PA, MS, WA, OR, OH, KS, TN, NH, AZ, NV, UT and AK and countries of : Australia, Canada, China, Germany, India, Japan, Singapore and Sweden. | Terminated | August 28, 2013 | 898 | 829 | Computer
Z-2177-2011 | 58380 | No | JFY | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Siemens Dimension(R) EZCR Flex(R) Reagent Cartridge. Catalog number DF270B. In vitro diagnostic test for the quantitative measurement of creatinine in human serum, plasma and urine on the Dimension Vista(R) System. |  |  | June 10, 2010 | 2010 | Siemens Healthcare Diagnostics, Inc. | Patient results may be falsely depressed or falsely elevated across the assay range. | Device Design | Siemens issued an Urgent Field Safety Notice dated June 2010 to all Dimension(R) EZCR customers; to notify them that Siemens confirmed an issue which many cause EZCR results to be falsely depressed or falsely elevated across the assay range. The letter asks users to discontinue use of this lab test kit. |  | 860 | 860 | Nationwide distribution. | Terminated | March 19, 2013 | 1013 | 860 | Computer
Z-2063-2011 | 58246 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Optical Guidance Platform, version 2.6 and 2.6.1; Reference/FSCA Identifier: CP-03899; Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA |  |  | March 03, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | The optical guidance platform may not be properly enforcing a 24 hour time limit between optical camera recalibrations. | Software change control | The firm, Varian, sent an "URGENT MEDICAL DEVICE CORRECTION-URGENT FIELD SAFETY NOTICE" letter dated March 3, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately check the Calibration timeout on their optical guidance platform. If this exceeds 24 hour, please set the factory default of 16 hours to ensure that the camera will be calibrated daily. The instructions are located in the document. The customers were also instructed to download the most recent version of the Optical Guidance Platform Instructions for Use at www.MyVarian.com or request a copy from their local Varian Customer Support District or Regional Manager. In addition, the customers were instructed to advise the appropriate personnel working in their radiotherapy department of the content of this letter. Varian is preparing a software patch which will be applied to affected units to correct this issue.If you require further clarification, please feel free to contact your local Varian Customer Support District or Regional Manager; USA and Canada -1888-VARIAN5 (888-827-4265) or Europe- +41 41 749 8844. |  | 599 | 599 units | Worldwide distribution. | Terminated | June 19, 2012 | 474 | 599 | Computer
Z-2166-2011 | 58400 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | DigiView software release 3.4.4.4 through 3.5.5.2. , 2008 Digisonics, Inc, 800-940-3240. |  |  | November 04, 2008 | 2008 | Digisonics, Inc | Failure of software to update some measurements when tracings or points are amended after measurement calculations have been performed. | Device Design | Digisonics, Inc., notified consignees by telephone and e-mail on 10/17/2008, informing them that a problem had been found in the Save Linked Images functionality of DigiView Release 3.6 (and all releases which start with 3.6 such as 3.6.1.). The problem is the failure of the software to update some measurements when tracings or points are amended after measurement calculations have been performed. A utility program, called Linked Image Measurement Verify was being supplied to all sites as a patch to the DigiView program. Release 3.6.6 was to be provided to all customers having previous 3.6 version. Problem was fixed in version 3.6.6 and 3.7.1.If you have any questions, please contact Digisonics Customer Support at (800) 940-3240. |  | 70 | 70 units total | Worldwide Distribution -- USA including States of : TX, OH, NC, MA, NY, IL, FL, OK, MO, ID, AZ, AR, MI CA, MD, and country of Canada.*** Due to the age of this recall, district is recommending no audit checks*** | Terminated | May 31, 2011 | 938 | 70 | Computer
Z-2180-2011 | 58395 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Digisonics, DigiView and ERS software release 3.6.4.4 through 3.6.5.2+, 2008 Digisonics, Inc., 3701 Kirby Drive, Houston TX 77098, 800-940-3240. |  |  | January 20, 2009 | 2009 | Digisonics, Inc | Data from one patient's study had the potential for being put into another patient's report. | Device Design | Digisonics notified end users via e-mail January 20, 2009 and January 21, 2009, that Data from one patient's study was being put into another patient's report. The problem occurred when a System bug was found in DigiView Release 3.6.4.4, and higher versions in which under certain circumstances (specific workflow) a failure to switch to a patient's images occurs. Two software patch versions will be made available to sites on or after February 5, 2009, to correct the reported problem.If you have any further questions, please contact Digisonics Customer Support at (800 ) 940-3240. |  | 70 | 70 units | Worldwide Distribution -- USA including States of : OH, NC, LA, TX, CA, MO, NY, MD, MA, DC, IA, AZ, PA, MI, NC, FL, IN, SC, ID, MN, AR, OK, IL, CT, and countries of Canada, Lebanon, and New Zealand. | Terminated | May 31, 2011 | 861 | 70 | Computer
Z-2175-2011 | 58182 | No | NSX |  |  | General Hospital | Contact ODE | Philips IntelliVue Clinical Information Portfolio software as follows: |  |  | March 16, 2011 | 2011 | Philips Healthcare Inc. | Edits to an order can result in unintended scheduled interventions and a patient may receive additional interventions or medication doses that are not intended. | Software design | Philips Healthcare notified accounts on 3/16/11 via the Urgent Medical Device Correction Notification. Customer notifications in the United States will be delivered via UPS. Customer notifications outside the US will be managed by Philips representatives in each affected geography. Customers will be asked to follow the guidelines in the Action to be taken by Customer/User section of the Urgent Medical Device Correction Notice until they receive their software correction. The correction will be in the form of a software upgrade and will be performed by Philips service personnel.All customers potentially affected by this issue will receive the software correction as part of the field action. |  | 387 | 387 units | Worldwide distribution: USA, Canada, Australia, Austria, Belgium, China, Czech Republic, Denmark, Finland, France, Germany, Hong Kong, Ireland, Italy, Kenya, Latvia, Malaysia, Netherlands, Norway, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Taiwan, Thailand, United Arab Emirates, and United Kingdom. | Open, Classified |  |  | 387 | Computer
Z-2176-2011 | 58252 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Optical Guidance Platform, version 2.6 and 2.6.1; Reference/FSCA Identifier: CP-03976; Model Number: HZ1; Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA |  |  | March 03, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | The anomaly that has been identified with the Optical Guidance Platform (OGP) software may not be always displaying the correct transfer date on the patient file. | Software design | The firm, Varian Medical Systems, sent an "URGENT MEDICAL DEVICE CORRECTION/URGENT FIELD SAFETY NOTICE" letter dated March 3, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to review the Optical Guidance Platform (OGP) Instructions For Use on Loading a treatment plan; before positioning the patient, compare the plan details loaded into the OGP with the plan details printed from the treatment planning systems; verify that the patient information, isocenter, gantry and couch angles are identical as requested in the Instructions For Use, and advise the appropriate personnel working in their radiotherapy department of the content of this recall letter.Varian has made changes to the OGP Instructions For Use -Varian strongly recommends you download the most recent version at www.MyVarian.com or request a copy from your local Varian Customer Support District or Regional Manager. Varian also is preparing a software correction which will be applied to affected unites to correct this issue.If you require further clarification, please feel free to contact your local Varian Customer Support District or Regional Manager at the follow: USA and Canada - 1-888-827-4265; Europe +41 41 749 8844; E-mail-North America: support-americas@varian.com; Australia/New Zealand: support.anz@varian.com; Europe: support-emca@varian.com; South East Asia: seasia.apps.helpdesk@varian.com; China/Asia: china.apps.helpdesk@varian.com; Japan: Japan.Apps.Helpdesk@varian.com, and Latin America: soporte.al@varian.com; Internet: Oncology Systems customer site - www.myvarian.com and Varian Medical Systems public site-www.varian.com. |  | 599 | 599 units | Worldwide distribution. | Terminated | June 19, 2012 | 474 | 599 | Computer
Z-2165-2011 | 58060 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Instrumentation Laboratory GEM Premier 4000 |  |  | February 14, 2011 | 2011 | Instrumentation Laboratory Co. | Measurement (amperometric) spike can occur early in cartridge life on the glucose and lactate sensors during patient blood analysis, leading to erroneously highresults | Device Design | Instrumentation Laboratory notified accounts by a Field Safety NotifIcation Letter dated Feb 2011. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to use whichever of the following mitigation options works best for their institution. Both options will have the same effect to eliminate the potential impact of a rare spike on the glucose and lactate sensors. Customers were instructed to complete the enclosed Response Tracking Form and return immediately fax to 781-861-4207 or e-mail to ra-usa@ilww.com.Customers were advised to contact their local representative with any questions. |  | 4379 | 4379 units | Worldwide Distribution - USA (nationwide) and the countries of CANADA ISRAEL ITALY APAN KUWAIT MALAYSIA MEXICO NETHERLANDS NORWAY PANAMA POLAND PORTUGAL PUERTO RICO SAUDI ARABIA SLOVENIA SPAIN SWEDEN SWITZERLAND TAIWAN TURKEY UNITED ARAB EMIRATES UNITED KINGDOM and URUGUAY. | Terminated | May 23, 2012 | 464 | 4379 | Computer
Z-2203-2011 | 55543 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Software for SSH-880CV Aplio Artida; System Diagnostic Ultrasound |  |  | January 06, 2010 | 2010 | Toshiba American Medical Systems Inc | Toshiba America Medical Systems (TAMS) has become aware that the FLEX M software has reversed left and right annotation and post FLEX M measurement error, which may result in incorrect values. | Software design | The recall communication was initiated on 1/6/10 with Toshiba forwarding an Urgent: Medical Device Correction letter with attached Customer Notification Reply Form (via US Postal Service for US Customers) to all customers who purchased the FLEX-M mode on Aplio Artida Diagnostic Ultrasound System. The letter provides the customers with an explanation of the problem identified and an action to be taken. Customers were instructed to avoid performing M image measurement using images saved in FLEX-M mode until the corrective action is completed. If customers need to perform measurement for images saved in FLEX-M mode, they were instructed to use manual calibration. Customers were instructed to note that even though the corrective action is taken, the existing saved images (saved images before the corrective action is taken) are not modified. Therefore, when referring to the measured values on the images which are saved in FLEX-M mode before the corrective action is taken, customers were instructed to carefully handle the measurements results because they may be incorrect.Customers were instructed to complete the attached form and fax it to the toll free number at the top of the form. The form can also be sent via e-mail to raffairs@tams.com.If customers have any questions regarding the medical device correction letter, they can contact the Director of Regulatory Affairs at (800) 421-1968 or their local Toshiba representative. |  | 28 | 28 | Nationwide Distribution -- Including states of IA, MT, PA, CA, OR, GA, LA, OH, TX, MA, OH, WI, and Puerto Rico. | Terminated | April 20, 2012 | 835 | 28 | Computer
Z-2052-2011 | 57744 | Yes | LWS |  |  | Cardiovascular | PMA | Boston Scientific's LATITUDE Patient Management System: |  |  | January 10, 2011 | 2011 | Boston Scientific CRM Corp | Boston Scientific has determined that data within certain printed reports obtained from the LATITUDE Patient Management System may not be accurate, in that, the Most Recent Presenting Electrogram (MRP-EGM) for some of the selected patients may be either missing or incorrect. | Software design | Consignees were sent a Boston Scientific "LATITUDEÂż Medical Records Information" letter dated January 2011. The letter was addressed to "Dear Doctor". The letter described the problem and the product involved in the recall.The letter also included the Rate of Occurrence, Action and Further Information sections. |  | 25836 | 25,836 | Worldwide Distribution -- USA, including DC and PR, and countries of AUSTRALIA, CANADA, AUSTRIA, BELGIUM, DENMARK, FINLAND, FRANCE, GERMANY, IRELAND, ITALY, NETHERLANDS, NORWAY, PORTUGAL, SPAIN, SWEDEN, SWITZERLAND, and UNITED KINGDOM. | Terminated | November 15, 2012 | 675 | 25836 | Computer
Z-2168-2011 | 57594 | [6 empty] | Philips FloTrak Elite modules: |  |  | January 04, 2011 | 2011 | Respironics Novametrix, LLC. | Software: two displayed respiratory parameters are incorrect due to errors in the system software: MValv (alveolar minute ventilation) and Vt/kg (tidal volume/patient weight). In both instances, the displayed parameters are higher than actual. | Software design | Respironics issued Field Action Notification Letter via Fed'x on 1/4/11, notifying customers of the issue and to instruct them on conditions for continued use and return of the device. The device will be updated with revised software to correct the calculation errors. Any questions or need additional information about this recall action, Contact Respironics US Customer Service at 1-800-345-6443 for customers residing in the USA or +011 (203)697-6488 for International customers. |  | 38 | 38 units | NationwideForeign: Belgium, China, Columbia, Egypt, England, India, Italy, Saudi Arabia, and Spain | Terminated | March 07, 2013 | 793 | 220 | Computer
Z-2169-2011 | 57446 | centrifuges (micro, ultra, refrigerated) for clinical use | JQC | General purpose laboratory equipment labeled or promoted for a specific medical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Statspin Express 4 Centrifuge |  |  | August 18, 2010 | 2010 | StatSpin, Inc., dba Iris Sample Processing | Rotors may fail due to wire spokes breaking | Device Design | Iris notified customers by letter in August 2010. The letter identified the product, the problem, and the action to be taken by the customer. Customers were requested to install the new rotors and discard the older rotors with the wires. There was no need for customers to return the rotors, but the firm requested customers fax or e-mail the enclosed form to verify that they have installed the new rotors. For any questions customers should call 1-800-782-8774. |  | 2090 | 2090 units | Worldwide Distribution - USA (nationwide) and the countries of Canada, Australia Belgium, Czech Republic, Finland, France, Greece, Kuwait, Norway, Turkey,UK, and Norway. | Open, Classified |  |  | 2090 | Computer
Z-2179-2011 | 55466 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Toshiba America Medical Systems Inc. SSH-880CV Aplio Artida; Diagnostic Ultrasound System |  |  | April 20, 2010 | 2010 | Toshiba American Medical Systems Inc | There is a possibility of the system shutting down owing to the power supply unit causing the protection function to activate thus turning off the system prematurely. | Device Design | The recall was initiated on 4/20/10 with Toshiba America Medical Systems (TAMS) forwarding a Medical Device Correction with a Customer Reply Form (via US Postal Service for US Customers) to all customers who purchased the SSH-880CV Aplio Artida Diagnostic Ultrasound System. The letter provides the customers with an explanation of the problem identified and an action to be taken. Customers were instructed to avoid this problem until corrective action is taken, in case this problem occurs; please contact their Toshiba Representative. If emergency use is necessary, use the following procedure which allows the system to be used temporarily: 1. Turn OFF the breaker at the back of the system and confirm that the error code near the breaker disappears. 2. Turn ON the breaker again to supply power to the system. Customers were instructed to complete and return the attached form and fax it to the toll free number at the top of the form. This form can also be sent via email to raffairs@tams.com. If customers had any questions they can call Paul Biggins at (800) 421-1968 or contact their local Toshiba Representative at (800) 521-1968. |  | 14 | 14 | USA, including Puerto Rico. | Terminated | May 25, 2011 | 400 | 14 | Computer
Z-2156-2011 | 58323 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Brightview XCT, 3/4" |  |  | March 28, 2011 | 2011 | Philips Medical Systems | The current Headholder to support Bright View XCT brain imaging allowed for imaging the brain but doe allow for imaging of the associated shoulder/neck area of the body. | Device Design | Customer information sheets will be provided to all customers when the field change order commences in April 2011.Philips will implement a full change order for the installation of the new headholder to replace the old headholder, free of charge, to all affected Philips BrightView XCT SPECT/CT units.For questions regarding this recall call 408-463-3000. |  | 1 | 1 | Worldwide Distribution - USA including US distribution to OH, MA, FL, VT and KY and the countries of Norway, Saudi Arabia, Belgium, Brazil, Japan, China, Poland, Austria, United Kingdom, Spain, Indonesia, Malaysia, Taiwan, Germany, United Arab Emirates, Netherland, Switzerland, Ireland, Czech Republic and Canada. | Terminated | June 22, 2011 | 86 | 50 | Computer
Z-2154-2011 | 58379 | No | JFY | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Enzymatic Creatinine (ECREA) Flex(R) reagent cartridges, K1270 for the Dimension Vista(R) 500 System |  |  | June 10, 2010 | 2010 | Siemens Healthcare Diagnostics, Inc. | Results may be falsely depressed up to 0.4 mg/dL or falsely elevated up to 0.6 mg/dL across the assay range. | Device Design | Siemens sent an Urgent Field Safety Notice letter dated June 2010 to all Dimension Vista(R) ECREA customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to:Discontinue use of ECREA method and transition to the Dimension(R) Jaffe creatinine (CREA) method.To notify anyone to whom they may have distributed the product.Complete and return the Effectiveness Check Questionnaire attached to the letter by fax to (302) 631-8467.For questions regarding this recall call (302) 631-6299. |  | 1399 | 1,399 | Nationwide Distribution - including AL, AZ, CA, CO, DE, FL, GA, ID, IL, IA, MD, MA, MI, MT, NH, NJ, NY, NC, OH, OR, PA, SC, TX, VA, WA, and WI. | Terminated | December 19, 2013 | 1288 | 1399 | Computer
Z-2132-2011 | 58567 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Synchron LXi 725, Part Number: 476501. |  |  | April 05, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has received reports of incorrect electrolyte results on Synchron LX systems due to a variety of maintenance-related hardware issues. | Equipment maintenance | The recall communication was initiated on 04/05/2011 with Beckman Coulter forwarding a Product Correction letter with attached customer response form to the affected customers via US Postal Service. The letter informed customers of the products affected, the issue, the impact, and the actions to be taken.Customers were instructed to follow the steps provided in order to address the issues related to microbial contamination; and to complete and return the enclosed response form within 10 days so that receipt of this important information could be confirmed. In addition, the letter requested customers share the information with their laboratory staff and retain the notification as part of your laboratory Quality System documentation.For customers who have any questions regarding this update, the letter indicated to contact Customer Support Center Call Center via the Beckman Coulter website at http://www.beckmancoulter.com/customersupport/support/. They may also contact Beckman Coulter at (800) 854-3633 in the United States and Canada or through your local Beckman Coulter representative.Beckman Coulter will issue two additional recall notification ("UPDATE TO URGENT PRODUCT CORRECTIONSynchron LX20, LX20 PRO, LXi 725 Clinical Systems") beginning on 05/20/2011, one to Laboratory Directors and one to Physicians. The notices will be mailed via US Postal Service on Friday, May 20th. Customers who have subscribed to email notifications will receive an email containing the letters on Monday, May 23th. Both letter informs the Laboratory Directors and Physicians that Beckman Coulter has received reports of erroneous/aberrant sodium, potassium, chloride and calcium results that have been released from the laboratory. Electrolyte shifts can be low or high and may occur above, below or within the normal range.Laboratory Directors were instructed to immediately disseminate the Attachment to this letter (via notification/alert procedure appropriate to your |  | 1284 | 1284 units total for all products | Worldwide Distribution -- Australia, Belgium, Canada, China, Colombia, Czech Republic, Ecuador, Egypt, Finland, France, French Polynesia, Germany, Hong Kong, India, Ireland, Italy, Japan, Republic of Korea, Kuwait, Mexico, Netherlands, Oman, Panama, Philippines, Puerto Rico, Saudi Arabia, Singapore, Slovakia, South Africa, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, and the United States. | Terminated | February 06, 2013 | 673 | 1284 | Computer
Z-2147-2011 | 53051 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | Siemens ACUSON X300 Systems with software versions 3.0.00 and 3.0.02, model numbers 10348531, 10348532 and 10348533, manufactured by Siemens Medical Solutions, Mountain View, CA. |  |  | July 24, 2009 | 2009 | Siemens Medical Solutions USA, Inc. | Mistreatment risk-- when a user defined formula is used to calculate a value in a vascular exam, data is switched laterally when it is internally transferred and displayed on the summary report within the system. | Software design | A field correction was initiated and a Customer Safety Advisory Notification was sent out on 7/29/09. The letter identified the affected product and stated that there was a potential for user defined calculations to be incorrect within the vascular package when using the listed software versions of the product. The letter also discussed the steps that customers could take to avoid the potential risk of the issue. A mandatory update software was developed and will be installed free of charge. Customers should have been contacted to schedule a time for the installation. Questions were to be directed to the local service support person. |  | 154 | 154 units | Worldwide Distribution -- Austria, Brazil, China, Columbia, Germany, Denmark, Egypt, Great Britain, Honduras, Ireland, India, Italy Japan, South Korea, Mexico, Malaysia, Peru, Poland, Portugal, Romania, Russia, Sweden, Singapore, Thailand, Turkey, Taiwan, United States, Uruguay, Venezuela, and South Africa. | Terminated | May 09, 2011 | 654 | 154 | Computer
Z-2149-2011 | 53211 | [6 empty] | 4Z1c Matrix Array Transducer, model number 10033682, used with ACUSON SC2000 ultrasound systems, model number 10433816, manufactured by Siemens Medical Solutions, Ultrasound Business Unit, Mountain View, CA. |  |  | July 10, 2009 | 2009 | Siemens Medical Solutions USA, Inc. | Burn risk-- due to coolant leaks, patients may suffer burns or blisters due to increased temperature. | Equipment maintenance | A customer safety advisory notification was sent out to customers on 7/10/09, and a mandatory software update will also be conducted. |  | 42 | 42 units | Nationwide. Foreign: Japan, Italy, South Korea, Singapore, Yemen, Germany, Australia, and Greece. | Terminated | May 05, 2011 | 664 | 42 | Computer
Z-2112-2011 | 58273 | No | CEM | Potassium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ABL90 FLEX Portable Blood Gas Analyzer System, Part Number: 393-090, 510 (k) #K092686. All units equipped with software versions: 2.4.1680.25 and all previous versions are subject to recall. |  |  | January 12, 2011 | 2011 | Radiometer America Inc | Following the receipt of two (2) reported customer complaints, the firm determined that the current operating software for their ABL 90 Analyzer does not prevent the further use of the sensor cassette when there is the possibility of iron rich plaque build-up in the reference electrode. If the device operator/user ignores the error message and continues to obtain biased measuring results using the | Software design | The firm, Radiometer America, sent two Urgent "Medical Device Field Correction" letters issued on February 4, and on February 24th, 2011. The letters described the product, problem and actions to be taken. The customers were instructed to replace the sensor cassette. In order to prevent this situation from occurring, the firm informed the customers in the recall/correction notification that they intend to upgrade all ABL 90 Flex analyzers to a new software version. The new software will perform a check during a normal system FLUSH in order to allow the analyzer to detect the bias and will notify the customer immediately of this condition and will automatically disable the cassette thus preventing further use. The letter goes on to remind the device customers that until such time as the software upgrade is installed, they should make sure to adhere to the procedures stated in the Reference Manual which call for the replacement of the sensor cassette immediately upon receiving message #1310. The letter further instructs the customer to complete and return the attached FAX RETURN FORM via fax to 440-871-0463. The letter also states that the expected release date for the new software should be sometime in the second half of February 2011. If you have any questions regarding this letter, or the upgrade, please contact Technical Support staff at 1-800-736-0600 Opt 4. |  | 565 | 32 units were distributed in the U. S. and 533 units were distributed world-wide | Worldwide distribution: USA including states of CA, DC, PA, MN, NY, NJ, TX, OH, MO, WI, UT, and MI; and country of Canada. | Terminated | May 09, 2012 | 483 | 565 | Computer
Z-2119-2011 | 58230 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | The 4D Integrated Treatment Console; Model Number: H51; Reference: 4DITC versions 8.6 and below; File Reference: CP-2007-0288; Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA |  |  | February 28, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | A software anomaly, in that the Multileaf Collimator (MLC) message displayed on the 4DITC on versions 8.6 and below, does NOT interlock beam delivery. | Software design | Varian Medical Systems sent an Urgent Medical Device Correction Letter on February 28, 2011, to all affected customers.Varian has created the procedure described in CTB-GE-579 to inform users on proper interfacing of the demountable BrainLAB m3 on the Varian Clinac. The procedure ( CTB-GE-579) is enclosed and posted to the customer site: http://www.MyVarian.com.This design fix in 4DITC v10.1 or later will be provided to all affected sites.If you require further clarification, please feel free to contact your local Varian Customer Support District or Regional Manager.USA and Canada - 1-888-VARIANS (888-827-4265)Europe - + 41 41 749 8844 |  | 17 | 17 | Class 2 Recall - Worldwide Distribution -- USA including states of ; CA, NV, NJ, GA, IL, and TX and countries of Australasia, Belgium, Brazil, China P.R., Denmark, France, Germany, Germany ( BrachyTherapy) Hong Kong, India Pvt. Ltd, Italy, Japan, Middle East, Netherlands, Russia, Spain, Switzerland and United Kingdom. | Terminated | March 30, 2012 | 396 | 17 | Computer
Z-2125-2011 | 58424 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARIA, RT Chart Application v 8.0 through 8.9 |  |  | March 09, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | An issue was identified with the gantry rotation direction when creating opposing field within the RT Chart function of ARIA, or possibly with third party planning systems and delivering the plan at 4DITC. | Software design | Varian Medical Systems, Inc. Oncology Systems sent Urgent Medical Device Correction Field Safety Notice letters on March 9, 2011, to all affected customers. ***12-20-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.***The letters identifeid the product, the problem, and the action to be taken by the customer. Customers were instructed to:Manually enter the gantry angel 0.0 when the plan indicates angel 0.0. Manual entry is only allowed to one decimal place, so the range that triggers the issue cannot be created if the keyboard is ued to enter the gantry angel.Users with third party plan sources should verify whether gantry angles with two or more decimal places are supported. If so, customers may be affected by this issue, and should take steps suited to the device.In all cases, the display at the Treatment console is indicative of what direction the gantry will move - users are advised to be especially attentive to these values, particulary when non-zero couch rotation values are used.Users are advised to always run a simulation of the plan, including all plan motions, at the treatment unit prior to first patient treatment. For further clarification customers were insturcted to contact their local Varian Customer Support District Manager. For questions regarding this recall call 650-424-6471. |  | 612 | 612 units | Worldwide Distribution - USA including CA, DC, FL, GA, MD, MN, MS, NY, OH, and TN and the countries of Canada, Australia, Austria, Bahrain, Bagladesh, Belarus, Belgium, Brazil, Columbia, Costa Rica, Cypress, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, Guadeloupe, Guatemala, Hong Kong, India, Indonesia, Ireland, Israel, Italy, Japan, Kuwait, Latvia, Lithuania, Luxemborg, Malaysia, Mexico, Morocco, Netherlands, New Zealand, Norway, Panama, Philippines, Poland, Portugal, Reunion, Romania, Russia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Tobago, Tunisia, Turkey, Ukraine, Emirates, and the United Kingdom. | Terminated | March 13, 2012 | 370 | 612 | Computer
Z-2127-2011 | 58313 | No | OBO | Ophthalmoscope. | Ophthalmic | Ophthalmic | 510(k) | Cirrus HD-OCT Model 4000 instruments, new and refurbished, distributed with, or upgraded to, version v 1.0 - 4.5, 5.0.0 or 5.1.0 |  |  | March 10, 2011 | 2011 | Carl Zeiss Meditec Inc | A software defect exists in versions 5.0.0 and 5.1.0 where when following a specific workflow, exam data may be saved under the name of the first patient | Software design | Carl Zeiss Meditec, Inc. sent an Field Correction Action letter dated March 9, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The letters requested acknowledgement of the letter and a response and the fact that a follow up packet with revised software will be sent to all customers with V 5.0.0 or 5.1.0 software.For questions, please call 1-855-853-9461. |  | 6888 | 3465 US; 3423 OUS | Worldwide Distribution - USA including AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, LA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, SC, TN, TX, UT, VA, VT, WA, WI, and WV, and the countries of Algeria, Argentina, Australia, Bangladesh, Belgium, Boliva, Brazil, Canada, France, Chili, China, Columbia, Denmark, Ecuador,Germany, Great Britain, Greece, Guatemala, Hong Kong, India, Japan, Lebanon, Malaysia, Mexico, New Zealand, Pakistan, Peru, Singapore,South Korea, Spain, Sri Lanka, Taiwan, Uruguay and Venezuela | Terminated | March 29, 2012 | 385 | 6888 | Computer
Z-0836-2011 | 56024 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Medical Systems FastPlan software, Versions 5.5 and 5.5.1, Model Number HZ3, manufactured by Varian Medical Systems, Palo Alto, CA |  |  | March 23, 2010 | 2010 | Varian Medical Systems Oncology Systems | Due to software Issue, an error occurs when performing cone planning using MRI, which may result in the patient receiving a higher radiation dose than expected, potentially to healthy tissue. | Device Design | Varian sent Urgent Medical Device Correction letters on March 23, 2010, by Federal Express to all customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to run test cases or otherwise determine whether the MR images that they receive have the 'Series Date' parameter populated.Customers should contact their MR system manufacturer to get a confirmation that the 'Series Date' field is populated.Customers were instructed to confirm by visual inspection that the fusion results are loaded appropriately in Cone Planning before starting to create or edit a plan.For questions regarding this recall contact your local Varian Customer Support Distirct or Regional Manager.Phone: USA and Canada - 1-888-VARIANS (888-827-4265) Europe -+41 41 749 8844E-mail: North America: support-america@varian.com Australia/New Zealand: support-anz@varian.com Europe: support-emea@varian.com South East Asia: seasia.apps.helpdesk@varian.com China/Asia: china.apps.helpdesk@varian.com Japan: Japan.Apps.Helpdesk@varian.com Latin America: soporte.al@varian.comInternet: Oncology Systems customer site - www.myvarian.com Varian Medical System public site - www.varian.com |  | 218 | 218 units | Worldwide distribution - USA nationwide and the countries of Australia, Brazil, Canada, Chile, Colombia, Georgia, Germany, Hong Kong, India, Ireland, Japan, Macau, Malaysia, Russia, Singapore, Spain, Sri Lanka, and Taiwan | Terminated | February 24, 2012 | 703 | 218 | Computer
Z-2204-2011 | 55542 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | CHVH-001A/1B Variable Helical Pitch (vHP) Software for TSX-101A: Aquilion 64; System, X-Ray, Tomography Computed |  |  | April 09, 2010 | 2010 | Toshiba American Medical Systems Inc | The recall was initiated by Toshiba America Medical Systems (TAMS) because the software associated with TSX-101A: Aquilion 64 can prevent scanning from being performed at the specified helical pitch (couch movement speed). | Software design | Toshiba America Medical Systems (TAMS) forwarded an Urgent: Medical Device Correction letter, dated 4/6/10, with attached Customer Reply Form, via US Postal Service to all customers who purchased the CHVH-001A/1B variable helical pitch (vHP) software for TSX-I0IA: Aquilion 64. The letter provides the customers with an explanation of the problem identified and actions to be taken. Customers were informed that a Toshiba service representative would contact them to schedule an appointment in order to install a corrective measure that will prevent the occurrence of the problem. Toshiba would like customers to set a scan range different from that for the previous scan for the second and subsequent vHP scans, if multiple vHP scans with the same scan conditions are performed in Auto Mode.Customers were instructed to complete and return the attached form and fax it to the toll free number at the top of the form. The form can also be sent via e-mail to raffairs@tams.com. Customers with any questions can contact the Director of Regulatory Affairs at (800) 421-1968 or contact their local Representative at (800) 521-1968. |  | 35 | 35 | Nationwide Distribution -- MN, OH, MO, MI, NY, KS, MD, TX, NJ, IL, ID, TN, AZ, PA, HI, GA, MT, WA, FL, NC, WI, ME, &amp; CA. | Terminated | August 17, 2012 | 861 | 35 | Computer
Z-2201-2011 | 56389 | No | LFX | Rubella virus serological reagents. | Microbiology | Microbiology | 510(k) | Software CD is labeled in part: |  |  | July 19, 2010 | 2010 | Bio-Rad Laboratories | The EVOLIS Assay Protocol Files Software is configured in such a way that may result in a miscalculation of quantitative test results and/or an error in reporting positive/negative sample values. | Software design | Bio-Rad Laboratories sent an Urgent Product Correction letter dated July 19, 2010, to all customer sites via FedEx. There are two issues related to The EVOLIS Assay Protocol Files (APF) Software that could potentially affect patient results. Issue 1 involves quantitative sample test/calculations/results interpretation for assays with linear regression, point to point or 4PL data reduction. Issue 2 involves sample test results with " Result Report Flags''. A software patch to correct these two issues will be sent to each customer site. Bio-Rad Laboratories will work with each customer to complete documentation confirming that the software patch has been successfully installed on each instrument. If you have any questions or concerns regarding these two issues, please contact Technical Support at 1-800-224-6723, option 2 and then option 3. |  | 272 | 272 | Nationwide Distribution -- Including Puerto Rico. | Terminated | August 25, 2011 | 402 | 272 | Computer
Z-2467-2011 | 58703 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Hitachi HIVISION Preirus/Avius Ultrasound equipped with V03.02 Software - Doppler Trace Measurement program. 510 (k) Numbers: Preirus - K093466 (6/1/10), Avius - K102901 (10/27/10). |  |  | April 18, 2011 | 2011 | Hitachi Medical Systems America Inc | Software anomaly. The Doppler Trace Measurement feature on certain models of the firm's diagnostic ultrasound systems, may, under certain conditions whereby multiple measurements are done sequentially, record incorrect values. The recorded value is displayed to the user on a report; however depending upon the measurement area, displayed velocity values could be lower than at the true velocity lead | Software design | The firm, Hitachi, sent a "DEVICE CORRECTION NOTICE" letter dated April 15, 2011 to its customers. The letter describes the product, problem and actions to be taken. The customers were inform of the nature of the software error in the Doppler Trace Measurement feature; the circumstances under which this error can occur; and the safety issue involved should incorrect values generated lead to a false diagnosis of stenosis in a patient. The letter provides the customer with a workaround solution in order to avoid this issue. The customer were also instructed to call the HMSA Applications Helpline at 877-871-2777 to reprogram their system to temporarily eliminate the noted safted issue and to immediately have all selections unchecked on the Doppler trace measurement result store setting set up screen. The notification letter also states that Hitachi Medical Systems America (HMSA) will be having their service personnel schedule a visit at each customer location and install a software upgrade on each affected system unit in order to directly address, and permanently resolve this issue. If you need any further information regarding this notice, please contact the Manager of Regulatory Affairs at 800-800-3106 x3720 or email to: thistlewaited@hitachimed.com. |  | 24 | 24 system units | Nationwide distribution: USA including states of: CA, MD, KY, FL, MN, NJ, IA, OH, TX, IL, and IA. | Terminated | February 19, 2013 | 673 | 24 | Computer
Z-2291-2011 | 58420 | No | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | LOGIX Order Entry (OE) Software, all versions, catalog 2M8410; Baxter Healthcare, Deerfield, IL 60015; |  |  | March 29, 2011 | 2011 | Baxter Healthcare Corp. | The LOGIX OE automated backup does not retain the three most recent backup files and instead will only retain the last three backups of the previous year. The software stores backup files in numeric order rather than chronological date order. | Software design | Baxter Healthcare Corp. sent an URGENT DEVICE CORRECTION letter dated March 29, 2011, via first class mail to all affected customers. The letter identified the product, the problem and the action to be taken by the customer. It was discovered there is an issue with the LOGIX OE automated backup not retaining the three most recent backup files and instead will only retain the last three backups of the previous year. The accounts were requested to turn off the automated backup function on the LOGIX OE Software, and perform manual backups following the instructions included in the Urgent Device Correction letter, and retain a copy of the letter at each computer on which LOGIX OE Server is installed. The customer were also requested to complete the attached customer reply form and fax it to Baxter at the number provided on the form. Any questions regarding the communication were directed to Baxter's Medical Information Services at 1-800-422-2751. |  | 84 | 84 units | Worldwide Distribution -- USA including Puerto Rico, and the country Canada. | Terminated | January 05, 2017 | 2109 | 84 | Computer
Z-2420-2011 | 58658 | Yes | CAC | Autotransfusion apparatus. | Anesthesiology | Anesthesiology | 510(k) | Haemonetics OrthoPAT -Orthopedic Perioperative Autotransfusion System |  |  | April 29, 2011 | 2011 | Haemonetics Corporation | 2 Models Model 1:OrthoPAT devices with no centrifuge drain spill containment system may have potential for fluid contamination leading to potential spark or smoke may occur are recalledModel 2:OrthoPat equipped with a spill collection drainage system updated with revised Instructions for Use and Quick reference Guide Addendum | Device Design | The firm, Haemonetics, sent an "Urgent Field Recall" letter dated April 29, 2011 to all customers. The letter describes the product, problem and actions to be taken. This recall covers the 2 models. Model 1: OrthoPAT devices without drain ports (Serial numbers 00001 through 02279-T). All OrthoPAT devices without drain ports worldwide, will be removed from distribution and scrapped. Model 2: OrthoPAT devices with drain ports (serial numbers beginning with 02280-T and higher). All OrthoPAT devices with drain ports will remain in clinical use and will receive updated labeling.The customers were instructed to follow the enclosed directions to confirm their possession of any units of this model and arrange for their return to Stericycle; immediately cease use of devices not equipped with a spill collection drainage system ( to determine whether their OrthoPAT device is equipped with the spill collection drainage system, please refer to the attached illustrations/instructions); complete and return the IMPORTANT-CUSTOMER CONFIRMATION sheet via fax to: Stericycle at 1-866-792-5450, and RETURN the OrthoPAT devices NOT equipped with spill collection drainage system (for instructions on how to arrange return of OrthoPAT Devices, call a Stericycle Representative at 1-866-918-8736). A Haemonetics representative will be in contact with the customers facility within 14 working days to assist with the recall actions.The firm, Haemonetics, will update the OrthoPAT labeling. The labeling will have specific instructions to permanently deploy the centrifuge spill collection bag at all times. Device cleaning instructions in the event of a blood spill are also clarified and include specific instructions to unplug the device from the AC outlet in the event of a fluid spill on the device. If a device does experience a fluid spill, new instructions direct the user to return the device to Haemonetics for evaluation and repair. Operator Manuals will be replaced with a new revisi |  | 2976 | 2976 units | Worldwide distribution: USA (nationwide) and countries including: Austria, Australia, Belgium, Switzerland, Canada, China, Czech Republic, Germany, Spain, Finland, France, Great Britain, Greece, Guatemala, Ireland, Israel, India, Italy, Japan, Lithuania, Luxembourg, Latvia, Monaco, Netherlands, Norway, Panama, Poland, Puerto Rico, Russian Federation, Saudi Arabia, Singapore, Slovak Republic, EI Salvador, Turkey and Taiwan., | Terminated | April 21, 2014 | 1088 | 2976 | Computer
Z-2415-2011 | 58677 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Pinnacle Destination Guiding Sheath, 6 French, 90 cm, Straight Tip, Tuohy-Borst Valve, Lot MD24, MD31, ME07, &amp; ME14. 510(k) K012812 |  |  | June 24, 2010 | 2010 | Terumo Medical Corporation | Molding defect in destination product that render the device unusable. | Equipment maintenance | Terumo Medical Corporation sent an Urgent Product Bulletin Recall letter dated June 24, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Terumo requested customers to immediately discontinue use, and quarantine any available stock of the listed product codes. In order to facilitate the exchange of these devices, a Terumo Territory Manager contacted all customers by July 9th to confirm receipt of the Terumo Destination Bulletin and scheduled an appointment to visit each customer in person by July 31st to personally manage the exchange. For questions or for further information, call (800) 283-7866 and ask to speak with the Product Recall Coordinator. |  | 214 | 214 units | Worldwide Distribution - USA including AL, AK, AZ, TX, LA, MS, MN, MO, CA, GA, FL, HI, KS, IA, ID, IL, TN, NM, OR, OK, WA, and WI and the countries of Belgium, Brazil, and Panama | Terminated | July 07, 2011 | 378 | 258 | Computer
Z-2408-2011 | 58205 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System, XiO Release 1.0.2 and above. |  |  | September 10, 2010 | 2010 | Computerized Medical Systems Inc | When RT Plans are exported from XiO and RT Structure Sets and Images are exported from Monaco, the DICOM iso-center coordinates may not be accurate. | Software design | Elekta / Computerized Medical Systems Inc began mailing an Important Safety Notice to all affected customers in September 2010. The notice identified the product, the problem, and the action to be taken by the customer. Customers were advised not to mix treatment plan information from multiple planning systems. Customers were asked to return the postcard as proof of receipt and their understanding of the Safety Notice. For questions regarding this recall call 408-830-8000. |  | 1377 | 1377 | Worldwide Distribution - USA (nationwide and Puerto Rico) and the countries of Albania, Algeria, Argentina, Australia, Austria, Azerbaijan, Bahamas, Belgium, Bolivia, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Ecuador, Egypt, El Salvador, Estonia, Finland, France, Georgia, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Kuwait, Latvia, Libya, Lithuania, Malaysia, Malta, Mexico, Morocco, .Netherlands, New Zealand, Nicaragua, North Korea, Philippines, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Suriname, Switzerland, Taiwan, Tajikistan, Tanzania, Thailand, Turkey, Turkmenistan, Ukraine, United Kingdom, Uruguay, Venezuela, Yemen, and Zimbabwe. | Terminated | April 08, 2014 | 1306 | 1377 | Computer
Z-2411-2011 | 58785 | No | FNK | Hydraulic adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Eleganza 3 Hospital Bed |  |  | May 02, 2011 | 2011 | Linet Americas Inc | Linet has been made aware that there has been breakage of the Mobi-lift handle on a few beds which may cause failure during patient assist use. | Device Design | The firm, Linet America, notified US customers only via phone on May 5, 2011. A follow up "service advisory information" letter dated May 11, 2011 was sent to all customers. The letter described the product, problem and action to be taken. The customers were instructed to remind/retrain staff to ensure that the Mobi-lift handles are always in the stowed position during patient transport and after use, and complete and return the attached Inspection procedure for each affected Eleganza 3 bed via mail to: Linet Americas, 10420-R Harris Oaks Blvd., Charlotte, NC 28269 or fax to: 704-248-5655. Linet Americas will be receiving a supply of new Mobi-lift design and will begin the process of executing service calls to replace the Morbi-lift handles with the new design.If you have any questions, contact the Director of Quality and Service Operations at 704-248-5650 or 980-406-6446. |  | 1601 | 1601 units | Worldwide distribution: USA (nationwide) including states of: KY, MA, NC, and NY; and country of: Canada. | Terminated | July 19, 2012 | 444 | 1601 | Computer
Z-1957-2011 | 58072 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Medical Systems, Varis Vision - RT Chart version 7.3.10. For radiation therapy. |  |  | February 07, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | An issue with the Varis Vision software where treatment can occur without the dynamic Multileaf Collimator component and may deliver a higher than expected dose of radiation. | Software design | Varian Medical Systems Urgent Medical Device Correction letters were sent to all affected customers on February 7, 2011. |  | 421 | 421 | Worldwide Distribution. | Terminated | December 23, 2011 | 319 | 421 | Computer
Z-1956-2011 | 57990 | No | GKL | Automated cell counter. | Hematology | Hematology | 510(k) | Coulter GEN's |  |  | January 31, 2011 | 2011 | Beckman Coulter Inc. | This recall was initiated because Beckman Coulter identified a problem with some date fields that are transmitted to the Laboratory Information System (LIS). Starting Jan 1, 2011, the date fields will transmit as a blank field to the LIS if the year is greaterthan 2010. The following dates are affected: Run date, Control expiration date, Date of birth(For individuals born in 2011 or later), a | Software design | The firm, Beckman Coulter, sent an "URGENT: PRODUCTION CORRECTION" letter dated January 31, 2011 with attached PCA Response Form (via US Postal Service for US Customers) to all customers who purchased the Coulter LH 750 Analyzer, Coulter LH 780 Analyzer, Coulter LH 500 Series System, and Coulter GEN S System. The letter describes the product, problem and actions to be taken. The customers were instructed to work with their US vendor to assess how this affects their laboratory. The actions they choose are dependent on the configuration of their LIS.ACTION FOR ISSUE 1:Use Sample Drawn Date and Date of Birth information available at your LIS in lieu of the analyzer transmitted fields BIRTH and UDATE.Use the current date when the results are received at the LIS in lieu of the transmitted Run Date [DATE]. Print results from the Workstation to keep a copy of the analysis date on the analyzer, if necessary.Use the control expiration date from the analyzer display or printout.ACTION FOR ISSUE 2:If you use the LIS 1G1 transmission format, use the General Information group field names in place of the Retic Parameter group field names to identify each reticulocyte run (see Table in Issue 2 section).Check with your LIS vendor to determine if it is possible to switch to the LH 750 Workstation transmission format. Format requirements are LIS driver specific.For patients with Dates of Birth in 2011 for both Issue 1 and Issue 2, results should be reviewed to determine if decision rules or age-related limits sets were triggered inappropriately.RESOLUTIONThese issues will be corrected in a future software release for the LH Series products. The GENÂżÂżS product has been discontinued; no software revisions will be forthcoming for the GENÂżÂżS product line. Customers were also instructed to complete and return the PRODUCT CORRECTION RESPONSE FORM via fax to 786-639-7500/7501/7502/7503/7504; mail to Beckman Coulter, Inc., Attn: Regulatory Affairs Mail Code 310B06, 11 |  | 9419 | 9419 (total units distributed in U.S. &amp; foreign countries) | Worldwide distribution: USA (nationwide) and countries including: Algeria, Andorra, Angola, Antigua and Barbuda, Australia, Austria, Azerbaijan, Bahrain, Belgium, Bermuda, Bulgaria, Burundi, China, Colombia, Croatia, Czech Republic, Denmark, Ecuador, Egypt, Ethiopia, France, Germany, Greece, Guadeloupe, Guam, Guatemala, Hong Kong, Iceland, India, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kenya, Korea, Kuwait, Latvia, Lebanon, Libyan Arab Jamahiriya, Lithuania, Macao, Malaysia, Mayotte, Mexico, Monaco, Morrocco, Myanmar, Namibia, Netherlands, Netherlands, New Zealand, Norway, Oman, Pakistan, Panama, Philippines, Poland, Puerto Rico, Qatar, Reunion, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Tunisia, Turkey, United Arab Emirates, United Kingdom, Uruguay, Venezuela, Vietnam, and Virgin Islands (US). | Terminated | September 25, 2012 | 603 | 9419 | Computer
Z-2407-2011 | 58267 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System, XiO Release 4.50.00 and above |  |  | May 27, 2010 | 2010 | Computerized Medical Systems Inc | XiO Software: For a specific clinical setup, the MLC leaf positions on the exported DRR images might not match the MLC leaf positions displayed in XiO. | Software design | The firm, Elekta, issued an "important Safety Notice" undated, to its customers. The notice indicates the problem and the clinical impact. There is no reasonable workaround for this issue. The notice states a future XiO Release will resolve the issue and the customers will be notified when this solution is available. A return postcard was included with the mailing to verify consignee received the notice. For sites not returning the postcard within one month, the recalling firm will contact them via phone or email to confirm receipt of the notice. The customers were instructed to distribute this notice to any and all users of the CMS software at their organization who are potentially affected by this issue.If you have any questions, please call 314-993-0003 or 800-878-4267. |  | 415 | 415 | Nationwide distribution, including Puerto Rico, to medical facilities. Distribution was also made to Military and Government consignees. Foreign distribution was made to Albania, Algeria, Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, China, Cyprus, Czech Republic, Ecuador, Egypt, Estonia, Finland, France, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Kuwait, Latvia, Lithuania, Malaysia, Malta, Mexico, Netherlands, New Zealand, Poland, Portugal, Russia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Switzerland, Taiwan, Thailand, Turkey, Turkmenistan, and United Kingdom. | Terminated | April 10, 2014 | 1414 | 415 | Computer
Z-2403-2011 | 58636 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Capsule Neuron UMPC is a component in the DataCaptor" Connectivity System. |  |  | August 25, 2010 | 2010 | Capsule Tech Inc. | Software: Data cached on the Neuron due to server unavailability is not being delivered to the server after server availability is restored.If a compact flash failure occurs, the file system may be corrupted and result in system hangs, lock-ups or crashes.Batteries can discharge completely in a Capsule Neuron" UMPC even though it is connected to power source | Software design | Capsule Tech initiated three corrective action notifications for the Neuron" UMPC units three failure modes as follows:a) Data cached on the Capsule Neuron" UMPC due to server unavailability is not being delivered to the server after server availability is restored:i) A customer advisory was provided to a subset of consignees, along with DataCaptor software version 6.6.5 on December 1, 2010.ii) A customer advisory was provided to the remaining consignees, along with DataCaptor software version 6.6.7 on January 27, 2011.If a compact flash failure occurs, the file system may be corrupted and result in system hangs, lock-ups or crashes:i) A customer advisory was provided to all consignees on March 23, 2011, via phone call, email and FedEx;.The correction will be provided as a software update. A formal recall notification letter will be provided to all consignees alerting them to the availability of the correction once it is released, via phone call, email and FedEx. c) Batteries can discharge completely in a Capsule Neuron" UMPC even though it is connected to power. Note: the Capsule Neuron" UMPC operates normally when connected to power.i) A customer advisory was provided to all consignees on April 23, 2010, via FedEx.For questions regarding this recall contact Capsule technical support team at (800) 260-9537 or e-mail support@capsuletech.com.. |  | 3144 | 3,144 Capsule Neuron UMPC units | Worldwide Distribution - USA (nationwide) and the country of Canada | Terminated | March 20, 2012 | 573 | 3144 | Computer
Z-2401-2011 | 58568 | No | NJR | Streptococcus spp. serological reagents. | Microbiology | Microbiology | 510(K) Exempt | Xpert Infinity Software, V 4.1a. |  |  | April 01, 2011 | 2011 | Cepheid | A mismatch of Patient ID or Sample ID and the Host Order from the hosted LIS and its requested test results could occur when running in batch mode and cartridges are introduced in rapid succession. | Software design | Cepheid sent an Corrective Action Notification letter dated April 13, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Cepheid recommended customers using 4.1a software on the INFINITY 48 System that is interfaced to a LIS with Host Ordering to make these alterations in operationof the system.After scanning the Patient ID (optional). Sample ID and Cartridge barcode, the host test order information should apear highlighted in the Order Test workspace.Once the highlighted host order information appears select SUBMIT.Warning. If the host test order informaiton is not highlighted, cancel the order by selecting CLOSE. Re scan Patient ID (optional). Sample ID and cartridge barcode and verify highlighted host test order before SUBMIT. Contact Cepheid Technical Support at 888-838-3222 in the US or 33 5 63 82 53 19 outside of the US for any questions regarding this Corrective Action Notice. |  | 19 | 16 US, 3 ROW | Worldwide Distribution - USA including CA, UT, TX, NJ, FL, NC, CO, IN, VT, MN and the countries of the UK, Denmark and Saudi Arabia. | Terminated | June 20, 2011 | 80 | 19 | Computer
Z-1950-2011 | 57815 | No | MLN | Electrocardiograph electrode. | Cardiovascular | Cardiovascular | 510(K) Exempt | Medstorm Adult Radiolucent Leads-Out Electrode - Physio-Control, Part 16242. |  |  | December 05, 2008 | 2008 | Heart Sync, Inc | The "leads out" connector had bent pins which would not allow connection to the Medtronic Physio Control Defibrillator cable during a product demonstration. | Equipment maintenance | The firm sent a recall notice to their distributors on 12/3/2008 asking them to quarantine product and return it to HeartSync. The letter stated that customers' current inventory would be exchanged for new product. Customers were to call 1-800-828-4681 to provide the necessary information. |  | 21000 | 2100 (210 cases of 10 each) | Nationwide Distribution | Terminated | April 13, 2011 | 859 | 21500 | Computer
Z-2386-2011 | 58431 | No | OHH | Powered breast pump. | Obstetrics/Gynecology | Obstetrics/Gynecology | Enforcement Discretion | AC Power Adaptor (transformer), TENPAO Model U090100Q31; used with the Pump In Style Advanced family of breast pumps; Medela Inc., McHenry, IL 60051-0650; Made in China; Article #9207010. |  |  | April 04, 2011 | 2011 | Medela Inc | The power adaptor (transformer) used with the Pump In Style Advanced family of breast pumps have power cords that may easily become damaged if not properly handled. | Device Design | Medela Inc., sent an " IMPORTANT FIELD ACTION " letter dated April 4, 2011, to all affected customers via The United States Postal Service ( USPS) United Parcel Service ( UPS) and E-mail .The letter identified the product, the problem, and the action to be taken by the customer. It was discovered the power adaptor/transformer used with the Pump in Style Advanced family of breast pumps can develop damage in the form of twists, kinks, bends, bulges or exposed wires in the power cord. Use of the power adaptor with a damaged cord could result in a short in the cord resulting in sparks, smoke, burning and melting of the transformers. If the users have a damaged power adaptor or power adaptor cord, they were requested to discontinue use of the adaptor and visit Medela's website at www.poweradaptorhelp.com to request a replacement adaptor. Questions about the recall process were directed to Medela Customer Service at 1-855-367-0104 or e-mail at poweradaptor@medela.com. Distributor/retailers were requested to post the recall notice within their establishment visible to their customer base, and to notify their customers of the recall. Information on the recall was posted on Medela's website www.medelapoweradaptorhelp.com on April 4, 2011. |  | 478823 | 17,681 units alone, 461,142 units with products | Worldwide Distribution - Nationwide Distribution and Puerto Rico and countries of Canada, Bermuda, Cayman Islands, Hong Kong and Haiti.CHI-DO proposes phone call audit checks of 2 consumers, 2 retailers and 1 hospital account by each district. | Terminated | October 16, 2012 | 561 | 478823 | Computer
Z-2247-2011 | 58471 | No | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Philips MicrostreamÂż Model M1923A FILTERLINE H SET INF/NEO 25UN PHILIPS |  |  | April 07, 2011 | 2011 | Philips Healthcare Inc. | Fine plastic strands on inner surface of the infant/ neonatal airway adapter may become dislodged and inhaled by the patient. | Equipment maintenance | The firm, Philips, sent a Field Safety Notice titled "URGENT MEDICAL DEVICE RECALL" letter dated April 2011 to its customers. Customer notifications outside the US will be managed by Philips representatives in each affected geography. The letter described the product, problem and actions to be taken. The customers were instructed to immediately identify all products from affected lots; remove from inventory and dispose in accordance with local regulations (new products will be sent proactively to each affected customer). Direct accounts were instructed to notify additional customers to whom the product has been distributed. These customers are instructed to follow the instructions in the "Action to be taken by Customer/User" section of the Notice. Questions regarding the recall can be directed towards the customer's local Philips representative at 1-800-722-9377. |  | 25400 | 25,400 units (1016 boxes) | Worldwide distribution. | Terminated | March 14, 2013 | 707 | 25475 | Computer
Z-2294-2011 | 58600 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Zoll E Series BLS Defibrillator/Pacemaker/Monitor |  |  | April 21, 2011 | 2011 | ZOLL Medical Corporation | Zoll E Series BLS Defibrillator may fail to auto charge and deliver energy | Software design | The firm, Zoll, sent an "Urgent Medical Device Correction" letter dated April 22, 2011 via UPS to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately alert all E Series BLS users to this problem; instruct users to not touch patient and keep the patient still during ECG analysis, Instruct users that if they experience this problem, to press the Analyze button again, and contact ZOLL Technical Service for the software upgrade kit. The letter also states that affected units can remain in service until the software has been updated. If you have any questions, please contact 24/7 technical support at 1 (800) 348-9011 or +1 (978) 421-9460. |  | 874 | 874 units | Worldwide distribution: USA (nationwide) and countries of: Australia, Canada, Finland, Germany, Malaysia and United Kingdom. | Terminated | January 23, 2014 | 1008 | 874 | Computer
Z-1970-2011 | 53195 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | LANTIS oncology Information System, an accessory to the Linear Accelerator, as part of the COHERENCE RT Therapist, but also cleared under the ONCOR Avante-Garde and Expression, as well as the syngo RT Therapist system, as part of the ARTISTE And PRIMUS system. Manufactured by IMPAC Medical Inc., Sunnyvale, CA, device distributed by Siemens Medical Solutions Oncology Care System, Concord, CA |  |  | August 21, 2009 | 2009 | Siemens Medical Solutions USA, Inc | When editing the note type of a previously entered note, then reverting to the original type, data may be lost. | Software design | Siemens sent an Urgent Medical Device Correction letter to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer.Customers were instructed to configure their system to lock all new notes automatically.To automatically lock new notes, open eChart Navigator, right mouse click in the notes section, select "configuration" and select "Automatically Lock New Notes."If the customer chooses not to configure their system to automatically lock new notes, follow at least one of the suggestions listed below:If the customer must edit a note type, use the append feature and document that they are making an edit to the note.Lock all new notes that are entered via eChart Notes. To lock a note manually, after creating the note, right mouse click and select "Lock,"Access and enter patient notes via eChart Notes until a LANTIS upgrade has been performed at your site.When entering patient notes via eChart Navigator, make sure that you are entering the notes in the correct section.\When entering patient notes via eChart Notes, make sure that the correct note type has been selected.Customers are to keep the Customer Safety Advisory Notice together with their LANTIS 8.30 User Manual where it should remain until the measures defined above are performed. For questions call 1-888-826-9702. |  | 30 | 30 affected units | Worldwide Distribution - USA including DE, WI, TX, TN, NJ, PA, NE, MA, and AL and the countries of South Africa, Australia, Ireland, Germany, Brazil, South Korea, India, and Japan. | Terminated | April 13, 2011 | 600 | 30 | Computer
Z-2233-2011 | 58373 | No | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | MicroStream Filterline Infant/Neonatal products: Part # 006285 - FILTERLINE H SET INF/NEO 25 UN PHYSIO (Inf/Neo Airway Adapter). The adult/pediatric airway adapter and the adult/pediatric FilterLine and VitaLine sets are not affected by this recall. |  |  | March 25, 2011 | 2011 | Oridion Medical 1987 Ltd. | fine plastic strands on inner surface of the infant neonatal airway adapter may become dislodged and inhaled by the patient | Equipment maintenance | The firm, Oridion, sent an "Microstream Filterline Field Correction Notice" letter dated March 25, 2011 to its US and foreign customers via telephone, email with the documentation discussed in the call, and letter. The letter described the product, problem and actions to be taken. The customers were instructed to place all affected product held in their inventory on ship hold; gather these units in their inventory and at the customer sites and exchange with the new product; inspect filter line starter (sample) kits shipped with Capnostream 20 and Microcap patient monitors for affected product, if affected Infant/Neonatal filter lines are found, remove them from the starter (sample) kit; send a purchase order in order to obtain replacement product for units in their stock and at customers sites, and complete and return the Inventory Reporting Form Excel Spread Sheet within two weeks of receiving the email even if they or their customers do not have the affected inventory. Oridion shall conduct a removal of the products from the related production lots from the field and an immediate replacement with product from subsequent production lots that meet all requirements. This action will be conducted by Oridion's distributors or by Oridion directly in some cases. Oridion anticipates that new product will be available for shipment during the week of April 4th.Questions regarding the recall can be directed towards the Customer Service Coordinator at: 1-888-674-3466 x1234 or NA-FCA@oridion.com. |  | 350 | 350 units in US (14 boxes with 25 units ea) | Worldwide distribution. | Terminated | March 19, 2012 | 360 | 10525 | Computer
Z-2214-2011 | 57591 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | BrightSpeed Excel Select, BrightSpeed Edge Select ,BrightSpeed Elite Select and BrightSpeed Elite Computed Tomography X-ray systems with software version 09BW35.11 or 09HW30.4. |  |  | January 19, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware of a potential set of circumstances that could cause X-ray continuation during an unexpected table stop on certain BrightSpeed systems (BrightSpeed Excel/Edge/Elite Select and BrightSpeed Elite Computed Tomography X-ray Systems ) if this event were to occur on your system, patient safety may be impacted. | Software design | The firm, GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION GE Ref: 22932" letter dated January 19, 2011 to its consignees/customers. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction, Contact Information, and Additional Dosimetric Information.A second letter dated March 11, 2011 was sent to the consignee/customer letter consisting of a cover letter and an attachment. "This letter is to notify you of an error in the January 19, 2011, Urgent Medical Device Correction Letter, that you have already received (GE Ref: 22932). In the Safety Issue section of that letter it is written: "If this issue were to occur, it would result in the remainder of the exposure being delivered at the patient position where the table stopped up to the prescribed scan time or 30 seconds, whichever is shorter." The stated 30 seconds limitation is incorrect for your particular model. This information has been corrected on the attached Urgent Medical Device Correction letter, so please discard the original letter if it is still in your possession. "GE Healthcare will provide a software update for all affected systems to address the issue. A GE Healthcare service representative will contact you to arrange for this correction.The customers were also instructed as follows: prior to GE's installation of a software update, continue to monitor patient scanning closely. Should the table stop unexpectedly and X-ray remains on, manually stop the scan, and ensure that all potential users in your facility are made aware of this safety notification and the recommended actions.If you have any questions, please contact GE immediately at the following: United States: 800-437-1171; Japan: 0120-055-919; Korea: 1544-6119; Australia/New Zealand: 800-659-465 and China: 800-810-8188. For other countries please contact your local GE Healthcare Service Representative if you have any questions or concerns regarding this notifi |  | 479 | 479 | Worldwide distribution: USA (nationwide) including states: AL, CA, CO, CT, FL, GA, ID, IL, IN, IA, KS, KY, LA, MI, MN, MS, MO, NE, NH, NJ, NM, NY, NC, OH, PR, RI, SC, SD, TN, TX, and WV; and countries including: URUGUAY, TURKEY, SWITZERLAND, SWEDEN, SPAIN, SOUTH AFRICA, SAUDI ARABIA, RUSSIAN FEDERATION, ROMANIA, REUNION, PUERTO RICO, PORTUGAL, POLAND, PARAGUAY, NEW ZEALAND, NETHERLANDS, MEXICO, MARTINIQUE, LEBANON, JAPAN, ITALY, INDONESIA, INDIA, HUNGARY, GREECE, GERMANY, FRANCE, EGYPT, COLOMBIA, CHINA, CHILE, BRAZIL, BELGIUM, BELARUS, BAHRAIN, AUSTRALIA, ANGOLA, ALGERIA. TUNISIA, THAILAND, SYRIAN ARAB REPUBLIC, PHILIPPINES, PARAGUAY, OMAN, NEW ZEALAND, MOROCCO, KUWAIT, KAZAKHSTAN, ISLAMIC REPUBLIC OF IRAN, ARGENTINA. | Terminated | July 10, 2012 | 538 | 479 | Computer
Z-2293-2011 | 58545 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo Dynamics version 9.0, Siemens Medical Solutions, USA, Inc., Ann Arbor, MI. |  |  | March 24, 2011 | 2011 | Siemens Medical Solutions, USA, Inc | If one or more graphs are in a table on the report, and if any of those graphs are deleted, it may lead to reports failing to save and/or marking the study as read might leave the study in an unread state or might mark the study as read but leave the report with a "Preliminary" watermark displayed. | Software design | The firm, SIEMENS, sent a "Customer Safety Advisory Notice" letter dated March 24, 2011 to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to place graphs in sections and not tables in the report; avoid deleting graphs from the reports in run-time and include this safety advisory notice in their operator's manual where it should remain until the service patch is applied. SIEMENS will be issuing a service patch to fix the defect. The letter states the patch will be available in the next few weeks.If you have any questions, call +1 (734) 205-2400. |  | 29 | 29 devices | Nationwide distribution: USA including states of: CA, CT, FL, ID, MD, MI, MN, NJ, NY and PA. | Terminated | July 10, 2012 | 474 | 29 | Computer
Z-2238-2011 | 58789 | No | KXJ | Radiologic table. | Radiology | Radiology | 510(K) Exempt | Maquet Magnus Carbon-fibre table top; Device Part Number 1180.16xx (x represents a digit from 0-5 and F-for the US market) |  |  | April 15, 2011 | 2011 | Maquet Inc. | Maquet GmbH's quality monitoring detected that incorrect cover sheets were installed with some Magnus carbon fibre table tops. | Device Design | Maquet sent an Urgent Medical Device Corrective Action letter dated 4/15/2011, along with a Action Response Form, via Federal Express to their three customers informing them the incorrect cover sheets were installed with some of the affected "MAGNUS 1180.16F0 carbon-fibre table tops." The firm will have a Service Technician contact you to service the OR table, and will be exchanging the cover sheet free of charge. To the customers, please follow the instructions on the response form. At this time there have been no reports that injury has occurred.If you have additional questions, please contact yur local MAQUET Customer Service at 1-888-627-8383 |  | 3 | 3 | LA, MO, VA | Terminated | August 21, 2013 | 859 | 3 | Computer
Z-2243-2011 | 58533 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARTISTE , ONCOR and PRIMUS digital accelerators with syngo RT Therapist v4.1 and 4.2 system; |  |  | March 31, 2011 | 2011 | Siemens Medical Solutions USA, Inc | Siemens found a potential safety issue related to patient treatment with table offsets on syngo RT Therapist (RTT) 4.1 and 4.2 which could result in dose being delivered to wrong location. | Software design | Siemens initiated an Urgent Medical Device Correction letter on March 31, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to:use only relataive table setup with all planned table values '0' if applicable.In cases where the treatment plan contains at least one table parameter other than '0' to consider the following:Use only one PVG plus resulting offset per fraction.Once an offset has been applied do not use 'apply offset' at RTT for any further offset.Select the PVG and treatment beams in the same auto-sequence. In casethis is not possible, the table position for each treatment beam, in the fraction, needs to be verified separately (especially for treatments with beams having different table posiitons. If an in-session resumption or individual download of beams has to be performed, make sure that the treatment table is in the correct position.In order to prevent any unintended table movement select manual table movement for all directions in the syngo RT Therapist service UI.Scenario 2 - Manual change of table posiiton followed by a PVG calculated offset.Always deliver the PVG's at the planned table position. Avoid manual change of table posiiton prior to applicaiton of PVG.Select the PVG and treatment beams in the same auto-sequence. In case this is not possible, the table position for each treatment beam, in the fraction, need to be verified separately (especially for treatments with beams having different table positions).If an in-session resumption or individual download of beams has to be performed, make sure that the treatment table is in the correct position.In order to prevent any unintended table movement select manual table movement for all directions in the syngo RT Therapist service UI.Customers were informed that a software update would be provided as soon as possible.Customers were asked to include the Customer Sa |  | 164 | 164 active devices | Worldwide Distribution - USA including CA, DE, FL, KY, LA, MA, MI, MO, NC, NE, NJ, NY, OH, PA, SD, TN, UT, WI, and WV and the countries of Australia, Belgium, Canada, Croatia, Egypt, France, Germany, Hungary, India, Ireland, Italy, Japan, Malaysia, Netherlands, New Zealand, Norway, P.R. China, Philippines, Poland, Republic of Korea, Russian Federation, Saudi Arabia, South Africa, Spain, Turkey, and the United Kingdom. | Terminated | March 30, 2012 | 365 | 164 | Computer
Z-1990-2011 | 58225 | No | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Posterior Referencing Instrumentation for the NEXGENÂż COMPLETE KNEE SOLUTION MODULAR T-HANDLE, REF 00-5901-061-00, NONSTERILE, Zimmer UK, Ltd.,, Zimmer Warsaw,IN |  |  | March 14, 2011 | 2011 | Zimmer Inc. | The firm is initiating a recall due to the potential for breakdown of the AlTiN PVD (black) coating during surgery and the coating could fall into the surgical site. The firm has received 22 reports of the black coating coming off the instruments after cleaning and sterilization cycles and 5 resulted in a delay of surgery. | Device Design | The firm, Zimmer, Inc., sent an "URGENT MEDICAL DEVICE RECALL" letter dated March 14, 2011 to its consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed to do the following:Instructions for Handling of Affected Product:1) Stop using the devices and to quarantine immediately2) Carry out a physical count of all affect product in their territory and record data on Inventory Identification Form included with this letter-Note: The instruments must be returned even if the instrument is not exhibiting the black powder3)Ensure they have provided acknowledgement via "email return receipt" upon receiving this notification4) Fax copy of completed form to (574) 372-42655) Return the recalled product along with the completed Inventory Identification Form and the Sterilization Certification Form to: Zimmer Distribution Center, Attn: Product Service, 1777 West Center Street, Warsaw, IN 46580A new or refurbished instrument will be provided as replacement for their returned product once the inventory is available.Instructions for Health Care Professional Notification1) Provide a list of all hospital that currently use or have used the affected product listed in notice by utilizing the User Facility Documentation Form and the Health Care Professional Documentation Form2) Fax a copy of the completed form(s) to: Zimmer, Inc. at (574) 372-4265Zimmer will provide a similar notification via FedEx to these hospitals making them aware of this recall and provide a "Dear Doctor" letter for their surgeons.Important: If you have distributed these affected lots further, please provide a copy of this letter to these customers when you implement these recall instructions.For shipping assistance, questions or assistance in notifying your accounts about the recall, please contact Zimmer, Inc., at 1-800-613-6131.ADDENDUM INFOThe information listed above was updated to include the additional 5 |  | 552 | 552 | Worldwide distribution: USA (nationwide) and countries including: Australia, Canada, Germany, Japan, New Delphi and Singapore. Addendum: Finland, Switzerland, UK, Denmark, Spain, Portugal, France, Netherlands, Italy, South Africa, Norway, Spain, and Sweden. | Terminated | March 19, 2013 | 736 | 9462 | Computer
Z-1983-2011 | 58048 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Maxi Sky 1000/BHM Medical V10 Bariatric Ceiling Lift System; a patient transfer lift; ArjoHuntleigh, Addison, IL 60101; Model numbers LF20010, LF20011, LF20013, LF20014, LF20015, LF21015, LF21019, 9120001, 9120004, 9120005, 9120008, 9120009 and 9121009 |  |  | March 21, 2011 | 2011 | Arjo, Inc. dba ArjoHuntleigh | The battery-powered control board in the ceiling lift motor may overheat, resulting in smoke and possibly flames and the lift becoming inoperable. | Device Design | ArjoHuntleigh sent an Urgent Field Safety Notice dated March 4, 2011, to all affected customers via UPS Ground on March 21, 2011. The notice informed the accounts of the potential for smoke and flames coming out of the Maxi Sky 1000/V10 bariatric ceiling lifts. The accounts were requested to immediately cease use of the affected lifts, powering them down by activating the emergency stop, and tag the lifts to prevent further use. If there is a clinical need to continue use of the lift before it is retrofitted or replaced, specific instructions were given to reduce the risk of a device malfunction. The accounts were requested to complete the Customer Response Form online, using the enclosed card's instructions. Questions were directed to the ArjoHuntleigh at 1-800-323-1245, ext 57985. |  | 1246 | 1,246 units | Nationwide Distribution | Terminated | February 16, 2012 | 332 | 1246 | Computer
Z-2227-2011 | 58584 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | UnicelÂż DxH 800 CoulterÂż Cellular Analysis System. |  |  | April 14, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed the auto prune will not remove patient results if the patient record includes demographics where the results are less than 30 days old. | Software design | The recall communication was initiated on 4/12/11 with Beckman Coulter forwarding an Urgent Product Correction (PC) letter with attached PC Response Form (via US Postal Service for US customers) to all customers who purchased the UniCelÂż DxH" 800 CoulterÂż Cellular Analysis System PN 629029. The letter provides customers with an explanation of the problem identified and an action to be taken. For Issue 1 &amp; 2, customers were instructed to perform manual database maintenance procedures to locate and delete unnecessary patient result and demographic information that extends beyond the life (in days) of the control lots. Ensure an alternative method of analysis is available during database maintenance.For Issue 3, customers were instructed to not attempt to transmit the files again after receiving an incomplete report. Instead, submit the missing file(s) by one of the following methods:- Submit your data at the electronic IQAP website (eIQAP) at(http://www.beckmancoulter.com/qap/index.jsp)- Fax printouts of the missing files to 786-639-3999- Email an electronic copy (.CSV) of the control file to IQAP.Reports@beckman.com.- Include the IQAP number on any communication. A corrected report will be returned to the customers' laboratories.For issue 4, customers were instructed to submit their control data to IQAP by one of the following methods:-Submit their data at the electronic IQAP website (eIQAP) at(http://www.beckmancoulter.com/qap/index.jsp)-Fax printouts of the missing files to 786-639-3999-Email an electronic copy (.CSV) of the control file to IQAP.Reports@beckman.com.-Include your IQAP number on any communication.Customers were instructed to complete and return the enclosed Response Form within ten days. Customers were instructed to contact Beckman Coulter Customer Service at 800-526-7694 in the United States or Canada, or contact their local Beckman Coulter representative. |  | 673 | 673 units | Worldwide Distribution -- US, Canada, Australia, Belgium, Colombia, Croatia, Czech Republic, France, Germany, Greece, Hong Kong, Hungary, India, Israel, Italy, Japan, Republic of Korea, Macao, Malaysia, Myanmar, Netherlands, Norway, Philippines, Puerto Rico, Qatar, Russian Federation, Saudi Arabia, Singapore, Slovakia, Spain, Sweden, Switzerland, Taiwan, Turkey, and United Kingdom. | Terminated | May 23, 2012 | 405 | 673 | Computer
Z-1984-2011 | 55910 | No | GKA | Abnormal hemoglobin assay. | Hematology | Hematology | 510(k) | VARIANT nbs Sickle Program Resin Update CD-ROM; Model Number:250-3020, included in 250-3000; Distributed and Manufactured by Bio-Rad Laboratories, Inc. The VARIANT nbs Sickle Cell Program is intended as a qualitative screen for the presence of hemoglobins F, A, S, D, C and E in eluates of neonatal blood collected on filter paper by high-performance liquid chromatography(HPLC). The VARIANT nbs Sickle Cell Program is intended for Professional use only. |  |  | April 14, 2010 | 2010 | Bio-Rad Laboratories Inc | The affected resin Update CD-ROMs cause all user-defined settings in the Setup/Test/Data Setup sub screen and Set/Test/Pattern Setting sub screen to be overwritten when the Update Kit procedure is performed. | Software design (manufacturing process) | Bio-Rad Laboratories sent out the Urgent Medical Device Correction notice on 4/14/2011, by Federal Express. An Update was sent on May 13th, advising customers of a minor discrepancy in the parameters included on the affected Resin Update CD-ROMs. Customers were instructed to examine your inventory for the affected reorder pack lots. Please complete and return the Medical Device Correction Response FormQuestions regarding this recall please call (510) 724-6114. |  | 658 | 658 units | Worldwide Distribution -- USA , Brazil, Canada, Caribbean, France, Italy, and United Kingdom. | Terminated | August 08, 2011 | 481 | 658 | Computer
Z-2000-2011 | 51504 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Siemens LANTIS 6.1 Oncology Information System with MEDOne Module, Linear Medical Accelerators; Product Numbers: 4503178, 5835991, 5851055; Product is manufactured by IMPAC Medical, Inc., Sunnyvale, CA. |  |  | March 06, 2009 | 2009 | Siemens Medical Solutions USA, Inc | The product has a software problem; it does not calculate the Creatinine Clearance correctly for patients less than two years old. | Software change control | On 3/6/2009, a Safety Advisory Letter was distributed to all affected users, with a description of the problem and identifying the affected product. Customers were instructed not to use the LANTIS to calculate Creatine Clearance for any patient under two years of age. Customers should direct their questions with regards to the letter to their local Siemens Service organization. A Siemens Representative will update the systems. |  | 17 | 17 | Worldwide Distribution | Terminated | April 15, 2011 | 770 | 17 | Computer
Z-2218-2011 | 58549 | No | DQA | Oximeter. | Cardiovascular | Anesthesiology | 510(k) | AVOXimeter 1000, AVOXimeter 4000 |  |  | March 23, 2011 | 2011 | International Technidyne Corp. | An incorrect calibration code that affects the total hemoglobin measurement was applied to affected Avoximeter instruments. | Device Design | International Technidyne Corp. sent a letter dated March 23, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer.Customers were instructed to check their inventory to determine if they have any of the affected product. If they have any of the affected product, customers were instructed to stop using it and remove it from their inventory. Customers were also instructed to complete the attached Customer Account Tracking Form and return it by fax to (732-635-0144) or e-mail (techsupport@itcmed.com) or mail to:ITC Technical Support20 Corporate Place SouthPiscataway, New Jersey 08854For any questions call the Distributor or ITC Nexus Dx Technical Support at 800-631-5945 (US) or 732-548-5700 (International), ext 4707. |  | 17 | 17 | Worldwide Distribution - USA including TX, KS, OH, OR, FL, GA, NY, MA, and DC and the countries of Taiwan and Thailand | Terminated | October 25, 2012 | 582 | 17 | Computer
Z-1999-2011 | 58202 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | GemStar SP Infusion Suite Database, Version 1.0; an Rx PC software application; Hospira, Inc., Lake Forest, IL 60045; list 13092-04-01 |  |  | March 23, 2011 | 2011 | Hospira Inc. | The software enables programming of the GemStar device for a Bolus Lower Lockout Limit outside of the hard limits entered into the GemStar SP Infusion Suite Database PC application. | Software design | The firm, Hospira, sent "URGENT DEVICE FIELD CORRECTION" letters dated March 23, 2011 to their customers on the same date. The letter described the product, problem and preventative action to be taken by the customers until the software can be upgraded. The customers were requested to contact Hospira Global Product Safety and Complaints at 1-800-441-4100 to report adverse events, and Hospira Advanced Knowledge Center at 1-800-241-4002 for technical assistance. The accounts were also requested to complete the attached Reply Form and return it to Hospira via fax at 1-262-577-6921.For additional information and technical assistance, contact Hospira Advanced knowledge Center at 1-800-241-4002, option 4 (available 24 hours a day/7 days per week). |  | 139 | 139 units | Worldwide distribution: USA (nationwide) and countries including: Canada, Chile, Colombia and Singapore | Open, Classified |  |  | 139 | Computer
Z-2216-2011 | 58534 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARTISTE , ONCOR and PRIMUS accelerators that utilize the RT Therapist v4.2 system in combination with syngo RT Oncologist 4.2. |  |  | March 31, 2011 | 2011 | Siemens Medical Solutions USA, Inc | Siemens has become aware of a potential safety issue that may result in mistreatment when customers are configured with the ARTISTE Linac RT Therapist v4.2 in combination with RT Oncologist v.4.2 and use the remote "Online 2D Image Review. | Software design | The firm, Siemens, sent an "Urgent: Medical Device Correction Customer Safety Advisory Notice" on March 31, 2011 via Siemens Customer Service Group or certified mail to its customers. The notice describes the product, problem and actions to be taken. The customers were instructed to not use the "Online 2D Image Review" on the Syngo RT Oncologist 4.2; to always calculate the offset directly on the RTT 4.2 before it is applied for the treatment with the patient on the treatment table; include this notice in their syngo RT Therapist 4.2 System Owner Manual chapter "Safety Advisory Letter" where it should remain until the measures defined in the notice are performed, and in the interest of safety, perform the preventative measures noted in this notice and inform all affected personnel immediately.Siemens is currently working on a solution for the incremental loading workflow in syngo RT Therapist 4.2 and will provide a software update as soon as possible. Siemens will notify the customers as soon as an update for their system will be available. UPDATE: The release of Safety Update UI TH010/11/S to distribute the software update for RTTv4.2.94O04 began on April 15, 2011.If you have any questions, contact Regulatory Affairs at 925-602-8157. |  | 45 | 45 active devices | Worldwide distribution: USA (nationwide) including states of: CA, FL, OH, PA and UT; and countries including: Australia, Belgium, Croatia, Germany, Japan, South Africa, Sweden and United Kingdom. | Terminated | March 30, 2012 | 365 | 45 | Computer
Z-2220-2011 | 58581 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Eclipse Treatment Planning System, Client software versions 8.0 through 8.6; Model Number: H48; Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA |  |  | April 04, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | An anomaly in Eclipse Treatment Planning System, versions 8.0 to 8.6, whereby Actual SSD (source surface distance) value may not correspond to the Planned SSD value after Planning Approval. | Software design | The firm, Varian Medical, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated April 4, 2011, to its customers. ***12-20-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.***The letter describes the product, problem and actions to be taken. The customers were instructed to do the following:1) Set the Planning Approved status in Eclipse (Eclipse displays both the Actual and Planned SSD values during plan approval) and verify both SSD values prior to approving the plan.2)Verify the SSD values in the Eclipse Treatment Plan Report prior to approving the plan for treatment or proceeding with treatment delivery (this report lists both the Actual and Planned SSD values).3) Prior to approving the plan in RT Chart, verify both the Actual and Planned SSD values in the Parameters Workspace by selecting Field Parameters&gt;Geometry tab.4) Prior to treatment, they should verify intended SSD values during independent MU calculation.5)Re-familiarize themselves with the concepts of Actual SSD and Planned SSD by reviewing the listed information.Please advise the appropriate personnel working in your radiotherapy department of the content of this letter.Varian mitigated the product behavior in version 8.9.If you require further clarification, please feel free to contact the Varian Oncology Help Desk ; phone: USA and Canada - 1888-827-4265; Europe: +41 41 749 8844; email: North America: support-americas@varian.com; Australia/New Zealand: support-anz@varian.com; Europe: support-emea@varian.com; South East Asia: seasia.apps.helpdesk@varian.com; China/Asia: china.apps.helpdesk@varian.com; Japan: Japan.Apps.Helpdesk@varian.com; Latin America: soporte.al@varian.com; Internet: Oncology Systems-www.myvarian.com; Varian Medical Systems -www.varian.com; or your local Varian Customer Support District or Regional Manager. |  | 5288 | 5288 units | Worldwide distribution. | Terminated | June 27, 2012 | 450 | 5288 | Computer
Z-2002-2011 | 58125 | No | GCY | Nonpowered, single patient, portable suction apparatus. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Sarns Flexible Weighted Sucker in K9 Pack (custom pack), catalog number 63635-01, Sterile and Non-pyrogenic. Sterilized by ethylene oxide. Terumo Cardiovascular Systems, Ashland, MA . |  |  | February 25, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | The firm has received two reports of the Terumo Weighted Flexible Sucker becoming ensnared in cardiac tissue after it was inserted into a ventricle through a valve. The firm has determined that the products's instructions for use are inadequate; the product should not be used to drain a cardiac chamber or inserted through a valve. However this is not stated in the IFU. The firm is conducting | Device Design | SEE UPDATE BELOW IN ALL CAPS.The firm, Terumo Cardiovascular Systems (Terumo CVS), sent an "URGENT MEDICAL DEVICE RECALL" letter dated February 25, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were informed that the product's Indications For Use (IFU) are inadequate; the product should not be used to drain a cardiac chamber or inserted through a valve. However, this is not stated in the IFU. In addition, the customers were instructed to read and understand the changes and clarifications to the product's IFU; review the Medical Device Correction notice; assure that all users are aware of this notice, and confirm receipt of this communication by faxing the attached Customer Response Form via fax to: Terumo CVS Customer Service at 800-292-6551 (US only).If you have any questions or concerns, contact Terumo CVS Customer Service at 1-800-521-2818. Customers Service hours: Monday-Friday, 8AM-6PM.ON 8/30/2011 THE FIRM ISSUE A LETTER ENTITLED "URGENT MEDICAL DEVICE RECALL" THAT STATED "AFTER FURTHER CONSIDERATION, TERUMO CVS DECIDED TO DISCONTINUE SALES OF THE WEIGHTED FLEXIBLE SUCKER AND IS ASKING USERS TO RETURN ALL UNUSED PRODUCT" TO THE FIRM. |  | 11690 | 11,690 | Worldwide distribution: USA (Nationwide) and countries including: BRAZIL, BELGIUM, BRAZIL, DOMINICAN REPUBLIC, UNITED ARAB EMIRATES, and MALAYSIA. | Terminated | December 10, 2012 | 654 | 11690 | Computer
Z-2219-2011 | 58429 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | PercuNav Field Generator Stand is a Computed tomography, X-Ray System. Part Number 453561431911. |  |  | April 01, 2011 | 2011 | Philips Medical Systems North America Co. Phillips | The PercuNav Field Generator Stand could tip over during set up, during use, or during transport. | Device Design | On 4/4/11, Philips Healthcare began sending out the URGENT-Medical Device Correction notification letter to their consignees.Philips informed their customers of a potential tipping hazard of the PercuNav Field Generator Stand. The notification letter also provides instructions to customers to ensure the articulating Mounting Arm is positioned so the Field Generator is not further than the specified distance from the back of the clamp on the stand's central pole during set up or during use. Customers are instructed to apply a provided label on stand to warn of a potential for tipping due to excessive force. For transport, customers are instructed to remove the Field Generator from the Mounting Arm and place the Field Generator in the storage compartment on the system cart. The Mounting Arm should then be collapsed and folded into a 90 degree angle.Customers are informed that a Philips field service representative will contact them to schedule installation of a service kit on the stand, free of charge, once service kits are available.Customers can call Philips Customer Service at 1-866-767-7822 and reference to FCO 79500197 for any questions about this recall. |  | 65 | 65 units total (33 units distributed to US and 32 units to foreign consignees) | Worldwide Distribution -- USA, Australia, Austria, Belgium, China, France, Germany, Hong Kong, India, Korea, Netherlands, Singapore, Sweden, Switzerland, Taiwan, Thailand, and United Kingdom. | Terminated | March 15, 2012 | 349 | 65 | Computer
Z-2714-2011 | 56848 | No | LNX |  |  | General Hospital | 510(k) | SoftReports versions 1.1.6.x and 1.1.7.x |  |  | June 23, 2010 | 2010 | SCC Soft Computer | SCC Soft Computer is issuing a correction for SoftReports Versions 1.1.6.x released 02/2008 and 1.1.7.x released 06/2008. Potential hazard - when using SoftReports Layout Designer, results displayed in the report may not match the patient information in the report header. | Software design | SCC Soft Computers sent a "RISK TO HEALTH NOTIFICATION TASK" letter dated June 9, 2010 to affected customers. The letter identified the product, problem, and actions to be taken by the customers.Customers were offered an upgrade or hot fix. The letter requested acknowledgement of receipt of the task notification and permission to install the hot fix. |  | 49 | 49 | Worldwide Distribution--USA (nationwide) and the country of Canada. | Terminated | July 26, 2012 | 764 | 49 | Computer
Z-2720-2011 | 58971 | No | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | ECMO Heater Model 333, Cardiopulmonary bypass temperature controller. |  |  | May 02, 2011 | 2011 | Cincinnati Sub-Zero Products Inc | The probe jack contacts (terminals) on certain of the firm's ECMO Heater devices are located too close to the solid state relay contacts (terminals) which can lead to the terminals coming into contact with each other. If this situation occurs, the service technician, user, and/or patient could be exposed to electric shock during servicing and/or use. | Device Design | Cincinnati Sub-Zero Products Inc issued an "Urgeant Medical Device Field Action" letter on May 2, 2011 to all affected customers.The letter includes; description of product and problem, affected lots and instructs customers to discontinue its use. The product is to be returned to the Service Department for an update and the customer is instructed to complete and mail Field Action Form.For additional contact the firm at 1-800-989-7373 or (513) 772-8810. |  | 116 | 116 units | Worldwide Distribution and countries of The Netherlands, Kuwait, India, Saudi Arabia, Taiwan and Australia. (USA) Nationwide including the states of AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IL, IN, LA, MA, MD, MI, MIN, MO, NC, NE, NJ, NM, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, WA, and WI. | Terminated | March 07, 2013 | 675 | 744 | Computer
Z-1845-2011 | 57613 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | CellQuest Pro SW, version 4.0.2. and later; |  |  | September 07, 2007 | 2007 | BD Biosciences, Systems &amp; Reagents | Results and gates could be mismatched if you utilize the undo function to undo a gate that was deleted. | Software design | Important Product Information letters, dated September 2007, were sent the week of 9/14/2001 via Federal Express to all affected customers. The letters identified the affected product along with the reason for recall. The letter states that "as long as customers do not use the Undo function to undo a gate deletion, the software can be used with full functionality." The letter also states that BD will change the issue in the next software version. If customers believe that they may have analyzed sample files by utilizing the Undo function, then it is recommended that the sample files be re-analyzed. If customers need more information they should contact the Customer Support Center at 877-232-8995. |  | 4552 | 4552 | Worldwide Distribution | Terminated | March 31, 2011 | 1301 | 4552 | Computer
Z-3055-2011 | 59402 | No | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | Flexible Probe with Blocking Washer (part of Segmented Cylinder Set) |  |  | July 13, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | Blocking washers on the Flexible Probe with Blocking washers may shift and cause potential for treatment length error. | Device Design | Varian Medical Systems sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated July 13, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter instructs customers to locate old design applicator probes and guiding tubes and to cease use of the affected products. Information in the letter included instructions on how to return the products to Varian BrachyTherapy.Customers should contact their local Varian Customer Support District or Regional Manager for questions regarding this notice. |  | 1019 | 1019 total | Worldwide Distribution | Terminated | June 20, 2012 | 343 | 1019 | Computer
Z-3108-2011 | 59439 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | HiÂżArt versions 4.0.x, and HD versions 1.0.x, H-0000-0003. |  |  | July 13, 2011 | 2011 | Accuray Incorporated | As a result of an internal review, TomoTherapy has identified an issue with HiÂżÂżArt versions 4.0.x, and HD versions 1.0.x that we would like to bring to your attention.During DICOM export of plan level images with a non square exported Field of View (FOV), an anomaly in the process of squaring the plan level image may cause the image to shift with respect to ROIs and dose. When the anomaly occu | Software design | TomoTherapy sent a "Urgent Medical Device Correction" letter dated July 13, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.This issue will be remedied with HiÂżÂżArt software version 4.1 and HD software version 1.1 and later.If you have any questions, please contact the TomoTherapy Technical Solutions Center by e-mail "support@tomotherapy.com" . Please refer to the letter for what phone number you should call depending on where you are. |  | 346 | 346 | Worldwide Distribution-- USA (nationwide) including the states of AL, AK, AZ, AR, CA, CO, CT, DC, FL, GA, HI, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MO, MT, NE, NJ, NM, NY, NC, ND, OH, OK, OR, PA, PR, SC, SD, TN, TX, UT, VA, WA, WV, WI and WY and the countries United Arab Emirates United Kingdom Turkey Switzerland Sweden Spain Saudi Arabia Poland Netherlands Italy Germany France Belgium Taiwan Singapore Malaysia Korea Japan India Hong Kong China Australia Mexico El Salvador and Canada | Terminated | March 18, 2014 | 979 | 346 | Computer
Z-3103-2011 | 59217 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | BrachyVision, versions 8.2, 8.5, 8.6 and 8.9. |  |  | August 11, 2010 | 2010 | Varian Medical Systems, Inc. | Medical device for brachytherapy treatment may cause users to enter incorrect syntax and lead to potential misinterpretation of patient treatment plan. | Software design | Varian Medical Systems, Inc. sent a "Medical Device Recall" letter dated August 05, 2010 to all affected customers.The letter included description of product affected and problem. Customers were advised to review template files to ensure that there are no characters immediately after the string in the customized dose quality parameters. Device users were requested to return a response form.For information on this recall contact Varian Technical Support at (800) 360-7909. |  | 1252 | 1252 units | Worldwide and Nationwide Distribution. | Terminated | November 25, 2011 | 471 | 1252 | Computer
Z-3093-2011 | 59781 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Encision AEM Suction Irrigation Electrode, L-Diamond, |  |  | August 24, 2011 | 2011 | Encision, Inc. | Electrode tips have been reported breaking off during use or cleaning. | Device Design | Encision Inc. sent out an "Urgent Product Recall" letter dated August 24, 2011 to all affected customers.The letter included: affected product, problem and instructed customers to discontinue use and return all affected electrodes to Encision. An enclosed reply form was to be faxed back to Encision indicating status of firm's inventory.For information on this recall call Encision at (303) 339-6917. |  | 169 | 169 units | Worldwide Distribution: Nationwide distribution, and the country of New Zealand. | Terminated | September 04, 2012 | 377 | 1078 | Computer
Z-3100-2011 | 59351 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand 4D Integrated Treatment Console (4DITC) Radiation Therapy, Interpretation of a RapidArc Treatment Plan as a 3D Conformal Arc Treatment Plan; version v8.6 or 8.8 in combination with treatment planning systems; Model Number: H51; Reference/PSCA Identifier: CP-03063; Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA |  |  | June 27, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | In certain configurations, the 4DITC can interpret a RapidArc or VMAT treatment field as 3D Conformal Arc Treatment Plan. If a RapidArc or VMAT treatment field is incorrectly interpreted as Conformal Arc, the field will be delivered with a constant dose rate and gantry speed for the entire duration of the delivery, which can lead to delivered dose being different than planned. | Software design | Varian sent an "Urgent Medical Device Correction/Urgent Field Safety Notice" letter dated June 27, 2011 to all affected customers. The letter identified the product, problem and recommended user actions to follow. Customers were instructed to notify all appropriate personnel of the correction notice. The letter states that Varian Service will contact customers to arrange to perform update to 4D Integrated Treatment Console (4DITC) software. For further clarification, contact your local Varian Customer Support District or Regional Manager. The letter also provides Varian Oncology Help desk contact information 1-888-827-4265 for USA and Canada and +41417498844 for Europe. |  | 645 | 645 | Worldwide Distribution - USA (nationwide) | Terminated | December 05, 2013 | 892 | 645 | Computer
Z-1526-2011 | 57851 | [6 empty] | Fresenius 2008T Hemodialysis Machine |  |  | January 19, 2011 | 2011 | Fresenius Medical Care Holdings, Inc. | Cursor on the dialysis screen may disappear and not respond to user input | Device Design | Fresenius Medical Care contacted customers by mailing a Customer Notification letter "Urgent correction FMNCA Model 2008T Hemodialysis Machine" via Certified-Mail, Return Receipt requested on January 19, 2011. The Technical Services department began contacting consignees by telephone on January 13, 2011, to schedule an appointment for installing the electronic board on the 2008T machines at the customers' facilities. For questions call Technical services 1-800-227-2572. |  | 1268 | 1268 units | Worldwide Distribution - USA and Canada | Terminated | July 08, 2013 | 901 | 1268 | Computer
Z-3091-2011 | 57806 | No | LGW |  |  | Neurology | PMA | Software version 3.3 for Rapid Programmer devices, Model 3832. |  |  | June 02, 2009 | 2009 | Advanced Neuromodulation Systems | Product programmed with unapproved software was distributed. | Software change control | St. Jude Medical Neuromodulation contacted all affected customers by telephone between June 2, 2009 and June 5, 2009. Customers were given information about the product, problem and action to be taken.All affected devices were returned to the firm as instructed. |  | 14 | 14 units | Distributed in New Hampshire and Puerto Rico. | Terminated | September 27, 2012 | 1213 | 14 | Computer
Z-3083-2011 | 59483 | No | ONW | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius 2008K@home |  |  | May 04, 2011 | 2011 | Fresenius Medical Care Holdings, Inc. | Upgrade software to implement new on screen graphics and instructions, the addition of a bracket on the side of the machine to temporarily hold the bloodlines during set up in order to conform to approved 510(k) | Device Design | Fresenius Medical NA issued a formal field correction notification dated May 4 ,2011, sent by certified mail with return receipt requested. The letter identified the product the problem and the action needed to be taken by the customer. The customer was notified of the retrofit that would be done to their machines. The patients' facilities were made aware of the changes that would be made to each of their patient's 2008K@home machines on May 4, 2011. On May 13, 2011, all facilities with active 2008K@home machines were sent a notification letter of the retrofit.Please return the receipt for confirmation that you have received this notice.Please call (781) 699-2053, to schedule a certified technician to conduct the machine retrofit. |  | 705 | 705 units | Nationwide Distribution.Canada - FMCNA is not conducting the recall to Canada, as Health Canada did not require the same modifications to the machines | Terminated | August 14, 2014 | 1198 | 705 | Computer
Z-3084-2011 | 59577 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Maquet HL Console for Heart Lung Machine HL20 |  |  | June 20, 2011 | 2011 | Maquet Inc. | Maquet HL 20 Twin Pump Modules may unintentionally increase speeds for a 1.5-2 second period of time before returning to its original speed. | Device Design | Maquet sent a PRODUCT RECALL NOTIFICATION LETTER/return response forms, dated June 17, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use of the affected product until a software update was performed by their trained technicians. Customers were also asked to forward the information contained in the notification letter to all users and to all appropriate Biomedical Staff. Customers were instructed to acknowledge receipt of the notification letter by completing and faxing back the attached response form to 973-807-9210 whether or not they have any of the affected product. For questions customers were instructed to contact their local MAQUET representative. For questions regarding this recall call 973-709-7994. |  | 299 | 299 units | Nationwide Distribution | Terminated | October 16, 2012 | 484 | 299 | Computer
Z-3075-2011 | 58489 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, DASH 3000, 4000, 5000 Patient Monitors |  |  | July 05, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has recently become aware of an issue with the Dash 3000/4000/5000 patient monitors configured with the 802.11b wireless LAN option that may impact patient safety. The Dash monitor may exhibit reboots on some networks when the 802.11b wireless LAN option is used to connect to the hospital wireless LAN infrastructure. This reboot results in a loss of patient monitoring for approxima | Software design | GE Healthcare sent an URGENT MEDICAL DEVICE CORRECTION letter dated June 15, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. GE recommended actions that should be taken by a qualified service representative. For any questions regarding the medical device correction or identification of affected items contact Technical Support at 1-800-558-7044 or customers should contact their local Service Representative.UPDATE: A second letter dated 12/13/11 was sent on 1/20/12 to consignees. The revised letter provides an update to the safety instructions and product correction. |  | 8029 | 8029 devices | Worldwide Distribution - USA including AL, AK, AZ, AR, CA, CO, CT, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MI, MN, MS, MO, MT, NE, NV, NJ, NM, NY, NC, OH, OK, OR, PA, SC, TN, TX, VT, VA, WA, WV, WI and the countries of UNITED KINGDOM, TURKEY, THAILAND, TAIWAN, SWEDEN, SPAIN, SLOVENIA, SINGAPORE, SAUDI A ARABIA, RUSSIAN FEDERATION, REPUBLIC OF KOREA, PORTUGAL, POLAND, NORWAY, NEW ZEALAND, NETHERLANDS, MEXICO, KUWAIT, JAPAN, ITALY, ISRAEL, INDIA, GERMANY, FRANCE, FINLAND, ESTONIA, ECUADOR, DENMARK, CANDA, BELGIUM AUSTRIA, AUSTRALIA, and ALGERIA. | Terminated | September 11, 2012 | 434 | 8029 | Computer
Z-1592-2011 | 56994 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Straumann coDiagnostiX, dental implant planning system. |  |  | September 30, 2010 | 2010 | Straumann Manufacturing Inc. | Software: Surgical Error may occur when the implant length is changed after positioning the surgical sleeve. | Software design | Straumann notified accounts of this correction by letter datedSeptember 30, 2010 and delivered via Federal Express Delivery Service. The upgrade is available on-line by logging into the coDiagnostiX system and enabling the automatic upgrade to occur. |  | 20 | 20 systems | Nationwide | Terminated | September 19, 2011 | 354 | 20 | Computer
Z-3070-2011 | 59588 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Entroy Pool Lift; a patient lift used in conjunction with a chair or stretcher to lift them in/out of a swimming/therapy pool; Arjo Hospital Equipment AB, Eslov, Sweden; Models GAB1000-01-xx, GAB1000-11-xx and GAC100x |  |  | August 04, 2011 | 2011 | Arjo, Inc. dba ArjoHuntleigh | The Entroy Pool Lift boom may rapidly lower un-commanded to the bottom of its stroke, with a potential for injury to the resident or caregiver. | Device Design | ArjoHuntleigh sent an Urgent Field Correction Notice I.D. 110701AHE to the affected end user via UPS 2nd Day on August 4, 2011. The notice informed the account of the potential for lift boom to rapidly lower un-commanded to the bottom of its stroke. The account was requested to immediately cease use of the affected lift, and disconnect the power to the lift to prevent further use until corrected. The account was requested to complete and sign the Customer Response Form, and return it to ArjoHuntleigh to schedule a service call to correct the lift. Questions were directed to the ArjoHuntleigh at 1-800-323-1245, ext 54485. |  | 311 | 1 unit in the U.S., 311 units in the rest of the world | Worldwide Distribution - USA including CT | Terminated | November 28, 2011 | 116 | 311 | Computer
Z-1594-2011 | 57915 | No | KWA | Hip joint metal/metal semi-constrained, with an uncemented acetabular component, prosthesis. | Orthopedic | Orthopedic | PMA | ARCOS Modular Revision System Torque Limiting T-Handle 55 inch- pounds, REF 31-301850, Biomet Orthopedics Warsaw, IN. |  |  | January 28, 2011 | 2011 | Biomet, Inc. | The firm received two complaints of the tabs on the Torque Limiting T-Handle fracturing during surgery. The firm performed internal testing and found that the tabs may fracture when used in the "OFF" or "locked" setting for hand-reaming the femoral canal. | Device Design | Biomet, Inc. made telephone calls to distributor accounts on January 28, 2011, with specific instructions that the Arcos Torque Limiting T-Handle was not to be used in the "OFF" position. Distributors were required to use the Arcos Torque Limiting T-Handle on the "55" setting for insertion of locking screws and troch bolt. Biomet, Inc. sent an URGENT MEDICAL DEVICE RECALL NOTICE letter dated February 21, 2011, to their distributors. The letter stated that the firm was currently manufacturing replacement handles which would be distributed to the field shortly. Consignees were advised to immediately discontinue use of the T-Handle in the "OFF" position.Send back the current T-Handle as soon as they receive a replacement.Carefully follow the instructions on the enclosed "FAX Back Response Form."Fax a copy of the Response Form to 574-372-1683 prior to return of product.Use priority carrier for shipment.If the product has been further distributed, consignees must notify hospital personnel of this action, via the enclosed "Dear Bioment Customer" notice. This letter must be give to hospital personnel responsible for receiving recall notices. Consignees are charged with the location and return of these products. Consignees were instructed to confirm receipt of this notice by calling, 800-348-9500, ext. 3755 or 3983. For any questions call 574-371-3755 or 574-372-3983. |  | 40 | 40 units | Worldwide Distribution - USA including the states of CA, NY, FL, VA, UT, IL, MI, KY, OH, MA, PA, SD, MO, TX, NV, Puerto Rico, and the countries of Australia, New Zealand, Canada. | Terminated | July 27, 2011 | 180 | 40 | Computer
Z-1572-2011 | 57732 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MOSAIQ Oncology Information System |  |  | January 11, 2011 | 2011 | Siemens Medical Solutions USA, Inc | The potential for injury to a patient should the treatment data be incorrectly uploaded to patient records. | Software design (manufacturing process) | Sent an Urgent Medical Device Correction Safety Advisory Notice to customers on 1/11/11. Letters will be delivered by Siemens Customer Service Group or sent by certified mail. Letter identified the affected product and stated there is a potential risk due to a problem identified by the manufacturer. Customers are asked to refer to the attached Safety Notice from Elekta IMPAC Software. A Siemens Service Engineer will contact customers site to arrange for a patient matching criteria audit to determine if you have a problem. Customers should include the Safety Advisory Notice including the Safety Notice from Elekta IMPAC Software in their product related documentation until further notice. If customers have any questions, they should contact their responsible SIEMENS Service Engineer or Application Specialist. |  | 10 | 10 units | International Only Distribution -- Germany and Japan | Terminated | April 19, 2012 | 464 | 10 | Computer
Z-3052-2011 | 59226 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | High-Definition 120 Multileaf Collimator |  |  | June 06, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | When using the HD120 Multi Leaf Collimator for beam shaping, and under specific treatment conditions, dose may be delivered outside of the intended treatment field that exceeds the expected transmission dose. | Device Design | Varian Medical Systems sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated June 6, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.***12-20-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.***The letter describes various jaw position recommendations to be followed when planning and treating dynamic fields on the HD120 MLC. Contact the Varian Oncology Help Desk at 1-888-827-4265 for questions regarding this notice. |  | 106 | 173***12-20-12 AMENDED TO: 106 units*** | Worldwide Distribution. | Terminated | December 23, 2011 | 200 | 106 | Computer
Z-1591-2011 | 57432 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Anchor Tissue Retrieval System; ETO sterilized, single use disposable device; port size 15 mm, 1550 mL capacity; 3 units per box; Manufacturer: Anchor Products Co., 52 Official Rd, Addison, IL 60101 U.S.A.; Ref TRS175SB |  |  | December 09, 2010 | 2010 | Anchor Products Company, Inc. | The metal component within the unit can break and cause a malfunction. Use of the device may represent a potential health hazard if the metal component is not removed after breakage. | Device Design | The firm, Anchor, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated December 8, 2010 to all customers. The letter described the product, problem and actions to be taken. The customers were instructed to examine their inventory and return any remaining units to their distributor for replacement; and complete and return the attached ANCHOR RECALL RETURN RESPONSE FORM via fax to (314) 961-4535 or mail to: Progressive Medical, Inc., Attn: Marketing Manager, 11085 Gravois Industrial Court, St. Louis, MO 63128, even if they have no inventory. A representative of Anchor Products then visited each of the direct accounts to help them implement their subrecall of the devices from their customers.For additional information please contact: Vice President, Anchor Products Company at 1-800-543-9124 or 1-630-543-9124; and the distributor at 1-800-969-6331 or 1-314-961-4535. |  | 2256 | 2,256 units | Worldwide distribution: USA including state of: Missouri; and countries including: Australia, Belgium, Canada, France, Greece, Ireland, Japan, Spain and the United Kingdom. | Terminated | January 20, 2012 | 407 | 91041 | Computer
Z-1548-2011 | 57563 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | The LIFEPAK12 Defibrillator/Monitor. |  |  | October 01, 2010 | 2010 | Physio Control, Inc. | L1FEPAK 12 Operating Instructions that describethe process for disconnecting the device from AC (wall) power are not always being followed. When theinstructions are not followed and the Power ON button is pushed within 2 seconds after disconnectingthe device from AC power, the service indicator will illuminate and the device may fail to charge the defibrillator and/or initiate pacing. | Device Design | Physio Control, Inc. began mailing the URGENT-MEDICAL DEVICE INFORMATION letter to their consignees on October 1, 2010. The letter identified the product, the problem, and the action to be taken by the consignee. The consignees were advised to follow the current LIFEPAK 12 Operating Instructions-Section 7.5-AC Power Adapter Operation and wait at least 2 seconds to push the Power ON button after disconnecting the device from AC power. If the instructions are not followed and the Power ON button is pushed within 2 seconds after disconnecting the device from AC power and the service indicator illuminates, it's important to be aware that turning the device off, then back on again, will reset the device and restore full functionality.The complete copy of the current LIFEPAK 12 Operating Instructions can be downloaded from the company's website at www.physio-control.com/products/ defibrillators, Select LIFEPAK 12, Select Documents.The consignees can call the firm's Technical Support Team at 1-800-442-1142, option 5, between 6:00 A.M. to 4:00 P.M. (Pacific), Monday - Friday. |  | 39 | 39, 180 units (17168 units internationally and 22012 units in US) | Worldwide Distribution - USA and the countries of ARGENTINA, AUSTRALIA, BAHRAIN, BERMUDA, BOLIVIA, BRAZIL, BRUNEI, CANADA, CAYMAN ISLANDS, CHILE, CHINA, COLOMBIA, COSTA RICA, EGYPT, GERMANY, GREECE, HONG KONG, INDIA, ISRAEL, JAPAN, JORDAN, KENYA, KOREA, KUWAIT, LEBANON, MALAYSIA, MALTA, MEXICO, MOROCCO, NETHERLANDS, NEW CALEDONIA, NEW ZEALAND, OMAN, PAKISTAN, PERU, PHILIPPINES, POLAND, PUERTO RICO, QATAR, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SRI LANKA, TAIWAN, THAILAND, TRINIDAD, TRINIDAD AND TOBAGO, TURKEY, UAE, UNITED KINGDOM, URUGUAY, VENEZUELA, YEMEN, and ZIMBABWE.Devices were distributed to the following countries:ARGENTINA, AUSTRALIA, BAHRAIN, BERMUDA, BOLIVIA, BRAZIL, BRUNEI, CANADA, CAYMAN ISLANDS, CHILE, CHINA, COLOMBIA, COSTA RICA, EGYPT, GERMANY, GREECE, HONG KONG, INDIA, ISRAEL, JAPAN, JORDAN, KENYA, KOREA, KUWAIT, LEBANON, MALAYSIA, MALTA, MEXICO, MOROCCO, NETHERLANDS, NEW CALEDONIA, NEW ZEALAND, OMAN, PAKISTAN, PERU, PHILIPPINES, POLAND, PUERTO RICO, QATAR, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SRI LANKA, TAIWAN, THAILAND, TRINIDAD, TRINIDAD AND TOBAGO, TURKEY, UAE, UNITED KINGDOM, URUGUAY, VENEZUELA, YEMEN, and ZIMBABWE. | Terminated | April 05, 2012 | 552 | 39 | Computer
Z-3048-2011 | 59568 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Infinity Gamma X XL Portable Patient Monitor. |  |  | July 29, 2011 | 2011 | Draeger Medical Systems, Inc. | Monitor keys may spontaneously become inoperative or active. This may cause a membrane switch panel malfunction, which could result in the monitor discharging a patient automatically. | Device Design | Draeger Medical sent an "Urgent Medical Device Recall Notification" letters dated July 2011 to all affected customers.The letter included product and problem affected. The letter informed the customer a Draeger representative will contact them to upgrade the system. Until the solution is implemented, Draeger recommends the Delta and/or Gamma XXL may be used. If the affected monitors become inoperative or activate without user interaction, Draeger advises removing the monitor from service and contacting a Draeger representative.For questions or for information on this recall please call Draeger Service at (800) 543-5047 (press 1 at the prompt, then 2, then 32349). |  | 29 | 29 | Worldwide Distribution - (USA) Nationwide distribution including the states of AK, FL, LA, MN, MS, MO, NJ, NC, OH, OK, PA, TN, TX, NC, NH, NJ, NY, WV, and WI; and the countries of Australia, Austria, Argentina, Bahrain, Brazil, Bulgaria, Canada, China, Colombia, Czech Republic, France, Germany, Georgia, Greece, Hungary, Italy, Ireland, India, Kazakhstan, Latvia, Macedonia, Malaysia, Mexico, Netherlands, Norway, Russian Fed., Romania, Saudi Arabia, Slovakia, Spain, South Africa, Switzerland, United Arab Emirates, United Kingdom, Uruguay, Vietnam, and Yemen. | Terminated | August 16, 2012 | 384 | 61 | Computer
Z-3057-2011 | 59576 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Maquet HL Console for Heart Lung Machine HL20: |  |  | June 20, 2011 | 2011 | Maquet Inc. | HL 20 Consoles for 4-pump and 5-pump configuration can malfunction if an excessive amount of liquid is used causing circuit board malfunction. | Device Design | Maquet sent an URGENT PRODUCT RECALL NOTIFICATION letter dated June 17, 2011. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to acknowledge receipt of the notification by completing and sending back the attached response form whether or not they have any affected product, fax to 973-807-9210 no later than June 30, 2011. For any questions customers were instructed to contact thier local MAQUET representative. |  | 229 | 229 units | Nationwide Distribution | Terminated | October 16, 2012 | 484 | 229 | Computer
Z-3040-2011 | 59688 | No | LGZ | Infusion pump. | General Hospital | General Hospital | 510(k) | enFlow IV Fluid Warmer Part number 980105VS; |  |  | June 30, 2011 | 2011 | Vital Signs Devices, a GE Healthcare Company | enFlow IV Fluid Warmer may have an issue which results in excessive external surface temperature on the underside of the warmer. | Device Design | GE Healthcare sent UMDC Letters/UMDC Confirmation Letters to customers on June 30, 2011. GE Healthcare sent UMDC Letters/UMDC Confirmation Letters on October 24, 2011 reflecting a new recall strategy. |  | 1307 | 1307 units | Worldwide Distribution - USA (nationwide) and the country of Vietnam | Terminated | June 14, 2013 | 715 | 1307 | Computer
Z-1566-2011 | 57969 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1, 100/120V System 1 Base, REF 801763, Terumo Cardiovascular Systems Corp., Ann Arbor, MI. |  |  | February 14, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (CVS) has received 19 reports of a circuit breaker for the Terumo Advanced Perfusion System 1 tripping and causing the system to switch to battery back-up power. In two cases, the incident was caused when the circuit breaker came into incidental contact with the base of an IV pole; in others there was no known cause. Terumo CVS' investigation has determined that | Device Design | Terumo sent an URGENT MEDICAL DEVICE CORRECTION letter dated February 14, 2011, addressed to Chief of Perfusion, Department of Cardiovascular Surgery or Director of Operationd Room Services. The letter explains the problem with the circuit breakers and hazards of the problem. The letter states that a service technician will visit the facility to install a guard. The service will be performed at the next scheduled preventative maintenance visit for those consignees with a service contract. For customers without a service contract, Terumo CVS will contact the consignees to schedule an appointment for the correction.Customers were instructed to review the Medical Device Correction notice.Assure that all users are aware of the notice.Confirm receipt of the communication by faxing the attached Customer Response Form to the fax number indicated on the form. For questions or concerns customers were instructed to contact Terumo CVS Customer Service at 1-800-521-2818. |  | 728 | 728 devices. | Worldwide Distribution - USA including AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IN, KY, LA, MA, MD, ME, MI, MO, MT, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WI and WV and the countries ofCANADA, COLUMBIA, JAPAN, RUSSIA, EGYPT, TURKEY, PAKISTAN, KUWAIT, COSTA RICA, TAIWAN, SAUDI ARABIA, INDIA, SOUTH KOREA, JORDAN, THE PHILLIPINES, INDONESIA, HONDURAS, DOMINICAN REPUBLIC, REPUBLIC OF GEORGIA, HONG KONG, UNITED ARAB EMIRATES, SINGAPORE, AUSTRALIA, MALAYSIA, TAIWAN, GERMANY, BELGIUM, MEXICO, THAILAND, VEITNAM, SOUTH AFRICA | Terminated | March 20, 2012 | 400 | 1113 | Computer
Z-1577-2011 | 57885 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Automatic Cutting Needle, Soft Tissue Biopsy Needle, NAC-1825M, 18ga (1.3mm) x 25cm (10), Remington Medical, Inc. |  |  | February 07, 2011 | 2011 | Remington Medical Inc. | There is a potential for the needle hubs to break during use. | Device Design | Remington Medical consignees were notified via facsimile, email and certified mail starting 02/07/2011. Distributors are requested to communicate this recall to their customers. If a customer does not respond to the recall notice within one week of the recall notification, the customer will be contacted by telephone. If this is unsuccessful, the customer will be sent a second Recall Notice letter and Customer Inventory Reconcile Fax by certified mail by 28 February 2011. If there is no response at that point, RMI will consider the customer's inventory depleted. All return product for this recall will follow RMIs Returns Goods Authorization procedure (SOP7-200). All product returns and customer responses will be documented. Upon completion of the recall, the status will be updated and submitted for the FDA's recall closure process. |  | 8161 | 8161 units | U.S. Nationwide | Terminated | August 12, 2011 | 186 | 46878 | Computer
Z-3036-2011 | 56806 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Product is the Ultraview Spacelabs Compact monitor, models 91367, 91369, and 91370, with touch screen display and wireless option. |  |  | September 17, 2010 | 2010 | Spacelabs Healthcare, Llc | Monitor may reset repeatedly which can result in a loss of patient monitoring for several minutes. Cause is a software problem. | Software design | Spacelabs Medical sent an "URGENT - MEDICAL DEVICE CORRECTION" notification letter to U.S. consignees by mail on September 17, 2010 and to international consignees by e-mail on September 24, 2010. The letter identified the recalled product, warns customers of the potential hazard and identifies some ways to mitigate the risk if they decide to continue use. The letter states to correct this situation, the recalling firm Field Service Engineers will be contact customers to schedule a time to install a free software update. If customers have any questions about this corrective action program, please contact Spacelabs Medical at 1-800-522-7025 and select 2 for Technical Support, or outside the United States 1-425-657-7200 x5089. |  | 579 | 579 total (517 to US acc'ts; 62 to international acct's. | Worldwide Distribution - USA (nationwide) and the countries of: International distribution to: Canada, China, France, Germany, Italy, and the United Kingdom. | Terminated | August 22, 2011 | 339 | 579 | Computer
Z-3032-2011 | 59456 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Spacelabs Medical Patient Monitor is a Physiological, Patient monitor With Arrhythmia Detection or Alarms. The Models are 91367, 91369, 91370, 91387, &amp; 91393 are sold with with Perioperative Mode (option D), which includes Start case/End case functionality. |  |  | July 08, 2011 | 2011 | Spacelabs Healthcare, Llc | The clinician did not realize the monitor was in END CASE mode and was expecting audible alarms. Visual alarms continue to work normally. | Device Design | Spacelabs sent an "URGENT - MEDICAL DEVICE CORRECTION" letter dated July 25, 2011 along with the Perioperative section of the Operations Manual to all consignees in the US. This letter includes recalled model 91369, 91370, 91387, &amp; 91393 (option D). On July 27, 2011, Spacelabs sent out an updated recall notification letter to include one additional model 91367 to all consignees in the US. The letter identifies the products, problem, and actions to be taken. On August 2, 2011, a recall notification letter (translated as necessary) and the Perioperative section of the Operations Manual was emailed to all international subsidiaries and distributors of record. The notification letter stated that the Operating Room staff may not be aware that alarm tones are off if the device has been left in the END CASE state. If the customers make use of the Perioperative START CASE / END CASE feature, they should advise the clinicians to familiarize themselves with this feature as described in the Perioperative section of the Operations Manual. In particular, they should know how to confirm the state of the alarm tones prior to starting a case. The letter states that the firm is in the process of upgrading the software and will be contacting customers to update monitors at no cost. If customers have any questions about this corrective action program, please contact Spacelabs Healthcare at 1-800-522-7025 and select 2 for Technical Support, or outside the United States 1-425-657-7200 x5089. |  | 2812 | total 2812 units distributed worldwide | Worldwide Distribution - USA (nationwide) and the countries of: Argentina, Australia, Bahrain, Bolivia, Brazil, Canada, Caymen Island, Chile, China, Colombia, Cyprus, Ecuador, Finland, France, Germany, Guam, India, Indonesia, Israel, Italy, Jordan, Mexico, Morocco, New Zealand, Nicaragua, Pakistan, Paraguay, Peru, Philippines, Poland, Saudi Arabia, Singapore, Slovakia, South Africa, Spain, Switzerland, Syria, Taiwan, Thailand, Trinidad and Tobago, Turkey, United Arab Emirates, Uruguay, and Venezuela. | Terminated | May 23, 2012 | 320 | 2812 | Computer
Z-3026-2011 | 59408 | No | BXB | Powered exercise equipment. | Physical Medicine | Physical Medicine | 510(K) Exempt | AlterG Anti-Gravity Treadmill; |  |  | July 15, 2011 | 2011 | Alter-G, Incorporated | Potential installation error on AlterG Anti-Gravity Treadmill that could cause greater than expected unweighting. | Device Design | Alter-G, Inc. notified all affected customers by telephone, e-mail and followed by a "Medical Device Customer Advisory Notification" letter dated July 15, 2011. Customers were provided information on the installation error, the potential harm to patients, and the actions to be taken. The letter instructs customers to discontinue use of the Alter G until the recalling firm can evaluate the machine and correct any installation error. The letter states that the recalling firm is planning to visit affected customer's facility as quickly as possible and will contact customers to schedule this visit. Questions or concerns should be directed to David Tjon at 510-270-5369. |  | 16 | 16 potentially affected units | Nationwide Distribution (USA ) - including the states of: OK and TX. | Terminated | August 16, 2011 | 32 | 16 | Computer
Z-3028-2011 | 59218 | No | KST | Vacuum-assisted blood collection system. | Hematology | Hematology | 510(K) Exempt | BD Vacutainer Push Button Blood Collection Set; REF 367324 Rx only, 23 G x 3/4" x 12 " (0.6 x 19 mm x 305 mm); |  |  | April 22, 2011 | 2011 | Becton Dickinson &amp; Company | BD Vacutainer Push Button Blood Collection Set can potentially undergo cannula separation from the hub. | Device Design | BD Diagnostics sent a "Lot Specific Product Recall" letter dated April 22, 2011 to all affected customers.The letter asked customers return any affected product and provides instructions for its return. The letter ask customers return the attached Packaging List with the shipment.For question on this recall contact BD Diagnostics at (201) 847-5117. |  | 1744000 | 1,744,000 | Nationwide (USA) Distribution. | Terminated | June 06, 2012 | 411 | 1778960 | Computer
Z-3022-2011 | 59394 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Eclipse Treatment Planning System, 4D Structure Merge Tool Option in Eclipse Release 8.5 -10 (Client Build 8.2.2x through 10.0.34) with 4D Planning; Model Number: H48; |  |  | July 07, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | An anomaly has been identified in the 4D Structure Merge option of certain versions of Varian's Eclipse Treatment Planning System. The treatment field energy may change from the current selection to a previous selection without user intervention and without invalidating the MU or calculated dose distribution. | Software design | Varian Medical Systems sent an "URGENT MEDICAL DEVICE CORRECTION/FIELD SAFETY NOTICE" letter dated July 7, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter instructs customers to cease using the 4D Structure Merge tool option in the Eclipse until further notice. Additionally, it is recommended that customers ensure that a qualified physician reviews the accuracy and placement of all patient structures used for plan development and evaluation and to conduct standard quality assurance procedures prior to patient treatment. Customers will be contacted by a Customer Service representative when the technical correction for this issue is available.Contact your local Varian Customer Support District or Regional Manager if you have questions regarding this notice. |  |  |  | Worldwide Distribution | Terminated | April 04, 2014 | 1002 |  | Computer
Z-1609-2011 | 57846 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart MRx Defibrillator Monitor (M3535A/M3536A) with software (SW) versions F.01.00 or F.01.01 and with all of the following hardware options EtCO2, IP and Temp. |  |  | February 07, 2011 | 2011 | Philips Healthcare Inc. | Software: EtCO2 and Ventilation Rate values may be labeled incorrectly. | Software design | Philips Healthcare representatives issued an Urgent Medical Device letter on 2/7/11. Customers are asked to follow the Action to be taken by Customer/User section of the Urgent Medical Device Correction Notification. In this section, customers are given the following instructions: Philips software (SW) version F.01.03 (for M3535A and M3536A) and R.01.00 (for M3536A, English only) corrects this problem.If you have affected product a Philips Service Engineer will contact you to install software version F.01.03 or R.01.00. Until your software is updated your HeartStart MRx monitor/defibrillator can remain in service. If you are using the device in one of the use modes listed under problem description, make sure that operators are aware that even though the EtCO2 and ventilation rate values can display interchangeably (unexpectedly or randomly switch) under either the ventilation rate label or the EtCO2 label, the values themselves can be identified by their color (for example if the ventilation rate label is white, the ventilation value is displayed as white). Customers are to contact their local Philips representative at 1-800-722-9377, if they need further information or support regarding this issue. |  | 2382 | 2382 units | Worldwide -- USA, Canada, AUSTRALIA, AUSTRIA, BAHRAIN, BELGIUM, BRAZIL, CHINA, COLOMBIA, CZECH REPUBLIC, DENMARK, EGYPT, ESTONIA,FRANCE, GERMANY, HONG KONG, INDIA, IRELAND, ISRAEL, ITALY, KENYA, LATVIA, MALAYSIA, MEXICO, NEW ZEALAND, NORWAY, PHILIPPINES, REUNION, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SPAIN, SWITZERLAND, THAILAND, UNITED ARAB EMIRATES, and UNITED KINGDOM. | Terminated | March 16, 2017 | 2229 | 2382 | Computer
Z-3012-2011 | 59059 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Healthcare, Seno Advantage 2.1 and 2.2 Workstations with RIS CCOW integration, GE Healthcare 3000 North Grandview, Waukesha, WI 53188 |  |  | May 31, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware that patient mismatch could occur on some Seno Advantage 2.1 and 2.2 Workstations with RIS CCOW integration. This issue may impact patient safety.When Seno Advantage workstation is synchronized with RIS using CCOW (an HL7 standard protocol designed to enable disparate applications to synchronize in real-time), the following issues could occur: 1.Two different | Software design | GE Healthcare sent a URGENT MEDICAL DEVICE CORRECTION letter to all affected customers dated May 31, 2011. The letter identified the product , the problem and the action needed to be taken by the customer.GE Healthcare will upgrade all affected workstations at no cost. A GE Healthcare service representative will contact you to arrange for this correction.If you have any questions or concerns regarding this notification, please call one of the following phone numbers:United States: 800 437 1171Japan: 0120 - 055 - 919For other countries, please contact your local GE Healthcare Service Representative |  | 391 | 391 | Worldwide Distribution--USA (nationwide) including the states of: AZ, CA, CO, CT, FL, GA, HI, IL, IN, IA, LA, MD, MA, MI, MN, MO, MT, NE, NJ, NY, NC, OK, OR, PA, SC, TN,TX, and WI, and the countries of UNITED KINGDOM, TAIWAN, SWITZERLAND, SWEDEN, SPAIN, PORTUGAL NORWAY, NEW ZEALAND, LUXEMBOURG, LEBANON, KOREA, JAPAN, ITALY, IRELAND, GERMANY, FRANCE, DENMARK, CANADA, BELGIUM AUSTRIA, AUSTRALIA FINLAND. | Terminated | June 29, 2012 | 395 | 391 | Computer
Z-1492-2011 | 58015 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | ACCU-CHEK FlexLink Plus, subcutaneous infusion set. Part Number 05511089001. |  |  | February 24, 2011 | 2011 | Roche Insulin Delivery Systems Inc. | There is the potential to kink or bend the cannula when inserting the ACCU-CHEK FlexLink Plus infusion set. This can lead to under delivery and elevation of blood glucose levels. | Device Design | Roche Insulin Delivery Systems sent urgent Product Recall letters, dated 2/21/2011. The letter identified the affected product and described the issue regarding kinked or bent cannulas. The firm requests that customers discontinue use of the ACCU-CHEK FlexLink Plus infusion set and return the unused stock. Customers are to contact their health care provider to determine if changes to their pump therapy are needed and how to temporarily continue insulin pump therapy without the ACCU-CHEK FlexLink Plus infusion set. The attached Customer Reply Card should be completed and return via mail or fax per the instructions on the card. The letter also lists some additional information pertaining to signs and symptoms. Questions regarding this recall should be directed to ACCU-CHEK Customer Care at 1-800-688-4578, 24 hours a day, 7 days a week. |  | 1454918000000 | 1,454,918 million worldwide | Worldwide Distribution -- US and Canada. | Terminated | July 27, 2011 | 153 | 1454918000000 | Computer
Z-3002-2011 | 59378 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Eclipse Treatment Planning System; |  |  | June 30, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | In affected Eclipse versions, the treatment field energy may change from the current selection to a previous selection without use intervention and without invalidating the MU or calculated dose distribution. | Software design | Varian sent an "Urgent Medical Device Correction/Urgent Field Safety Notice" letter dated June 30, 2011 to all affected consignees. The letter identifies the affected product, problem, and actions to be taken. The letter states that Varian will contact all affected customers to apply a database configuration script to the database and is preparing a correction for Eclipse Version 10 to address the issue. A Varian representative will contact each site to schedule correction when it is available. The letter also provides RECOMMENDED USER ACTIONS to follow. The Letters go on to reiterate caution and warnings from Eclipse instructions for use. All appropriate personnel should be notified of the correction notice. For future reference, this document will be posted to the Varian customer support website: http://www.MyVarian.com. For further clarification, contact your local Varian Customer Support District or Regional Manager. |  | 2436 | 2436 | Worldwide Distribution. | Terminated | May 09, 2014 | 1044 | 2436 | Computer
Z-3111-2011 | 59297 | No | DXJ | Medical cathode-ray tube display. | Cardiovascular | Cardiovascular | 510(k) | Philips Medical: |  |  | July 08, 2011 | 2011 | Philips Healthcare Inc. | Under certain circumstances, edits to an order can result in unintended scheduled interventions on the IntelliVue Clinical Information Portfolio. A patient may receive interventions or additional medication doses that are not intended. | Software design | Philips Healthcare sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated July 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.Customers are asked to follow the guidelines in the "Action to be taken by Customer/User" section of the Urgent Medical Device Correction Notice. Philips has developed a software upgrade to address this problem in the short term. This upgrade will be available immediately. This software upgrade will be performed by Philips service personnel. While customers await the correction of their devices, customers are instructed to follow the instructions for use for the device, which includes the directions to:1. Always use the medication administration record (MAR) in determining patient care.2. Always review orders and pending administrations in the administration record after any changes to an order are stored.For the longer term, Philips is developing safeguards intended to minimize the potential for similar errors resulting from order changes and upgrade to be available in September 2011.If you need any further information or support concerning this issue, please contact your local Philips representative. |  | 454 | 454 units | Worldwide Distribution-USA (nationwide) and the countries of AUSTRALIA, AUSTRIA, BELGIUM, CANADA, CHINA, CZECH REPUBLIC, DENMARK, FINLAND, FRANCE, GERMANY HONG KONG, IRELAND, ITALY, KENYA, LATVIA, MALAYSIA, NETHERLANDS, NORWAY, SAUDI ARABIA, SINGAPORE, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, PROVINCE OF CHINA, THAILAND, UNITED ARAB EMIRATES, and UNITED KINGDOM. | Open, Classified |  |  | 454 | Computer
Z-3098-2011 | 59350 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand 4D Integrated Treatment Console (4DITC) Radiation Therapy, version 8.1 SP2 to 10.0 SP1; Model Number: H51; Reference/PSCA Identifier: CP-03940; Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA |  |  | June 23, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | Varian has scheduled for an upgrade of the device from current version to 10.2. There are circumstances in which the 4DITC may be involved in a failure to correctly record treatment history information. | Software design | Varian Medical Systems, Inc. sent an "Urgent Medical Device Correction/Urgent Field Safety Notice" letter dated June 23, 2011 to all affected customers. The letter identified the product, problem and recommended user actions to follow. Customers were instructed to notify all appropriate personnel of the correction notice. The letter states that the recalling firm will issue a Customer Technical bulletin to inform users of the mandatory upgrade and related information. For future reference, this document will be posted to the Varian customer support website: http://www.MyVarian.com. Any questions regarding this Product Notification letter, contact Varian Medical Systems Help desk at 1-888-827-4265 for USA and Canada and +41417498844 for Europe. |  |  |  | Worldwide Distribution - USA (nationwide) | Terminated | March 15, 2012 | 266 |  | Computer
Z-2733-2011 | 58686 | No | MXC | Medical magnetic tape recorder. | Cardiovascular | Cardiovascular | 510(k) | SJM Confirm, Implantable Cardiac Monitor (ICM) Models: DM2100, DM2102 (IDE). |  |  | March 31, 2011 | 2011 | St Jude Medical CRMD | The recall was initiated because St. Jude Medical made new software available for the SJM Confirm DM2100 implantable cardiac monitor (ICM) that incorporates enhancements to the devices sensing algorithm. These enhancements are intended to provide improved sensitivity and specificity for diagnosing arrhythmias. The user is presented with the option to upgrade new software into the SJM Confirm ICM | Software change control | St. Jude Cardiac Rhythm Management Division sent an "Important Product Correction" letter on March 25, 2011 to all affected customers.The letter explains the reason for the recall and provided recommendations including a software upgrade by SJM Field Clinical Engineer or Technical Services. Customers were recommended to fill out and return a "Physician Acknowledgement Form".For additional information please contact your local St.Jude Medical Sales Representative, Field Clinical Engineer or the Technical Services Department at 800-722-3774 |  | 83 | 83 units | Worldwide distribution: (USA) Nationwide including the States of CA, WA, FL, WV, NY, MD and LA, and the country of Germany | Terminated | July 19, 2011 | 110 | 83 | Computer
Z-3205-2011 | 59473 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | PCR Eleva Radiological Image Processing System with software R.1.2 |  |  | July 27, 2011 | 2011 | Philips Healthcare Inc. | The mirror icon on the bottom of each image could be interpreted as a RIGHT marker for side indication if the image is mirrored again within a PACS system. This could potentially lead to a misdiagnosis | Software design | Philips Healthcare Inc. sent out "Urgent --Medical Device Correction" notices dated June 10, 2011 to all affected customers.The letter described the product and problem. Customers were informed the "burned-in Mirror-Icon" may be misinterpreted as a Right marker. Customers are advised to preferably use lead markers during each X-ray acquisition or alternatively electronic annotations on each image for side indication. The presentation of these markers within each image will clearly show the difference between the intended Left/Right indication and the little mirror ICON. Philips will install improved software that prevents the misinterpretation of the symbol and will schedule the installation of the updated software.For questions on this recall call Phillips Healthcare Inc. at 978-687-1501. |  | 197 | 197 units | Worldwide Distribution: Nationwide distribution; and including the countries of Albania, Angola, Argentina, Australia, Austria, Azerbaijan, Belgium, Bolivia, Brazil, Bulgaria, Canada, Chile, China, Croatia, Denmark, Egypt, France, Germany, Ghana, Greece, Greece, Haiti, Hungary, Indonesia, Iraq, Israel, Italy, Kazakhstan, Kenya, Latvia, Libya, Liechtenstein, Lithuania, Malawi, Malta, Martinique, Mexico, Mozambique, Netherlands, New Zealand, Nigeria, Norway, Oman, Philippines, Portugal, Puerto Rico, Romania, Russia, Saudi-Arabia, Serbia, Slovakia, South Africa, Spain, Sudan, Sweden, Switzerland, Tanzania, Thailand, Tschech. Rep, Turkey, Turkmenistan, UAE, Uganda, and United Kingdom. | Terminated | January 12, 2017 | 1996 | 197 | Computer
Z-0008-2012 | 59213 | No | MRW |  |  | Orthopedic | 510(k) | OSTEOMED SPINE Facet Fixation System PrimaLOK FF Threaded Implant Driver Shaft OsteoMed L.P. 3885 Arapaho Road, Addison, Texas 75001 |  |  | June 07, 2011 | 2011 | Osteomed, Lp | Two of the Implant Driver Assembly tips were reported to break during surgery. | Device Design | OsteoMed LP sent an Urgent --Product Recall letter dated June 7, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their inventory immediately for affected inventory and promptly return any affected inventory to:OsteoMed LPAttan: Rebecca Ellis3885 Arapaho RoadAddison, Texas 75001If the customer further distributed any of the affected product customers were instructed to immediately contact their accounts and advise them of the recall situation and have them return any devices they may have to OsteoMed at the above address.Customers were asked to complete and return the enclosed response immediately in the enclosed, postage paid envelope even if they do not have the device. For any questions please call (800) 456-7779. |  | 20 | 20 units | Nationwide Distribution including PA, TX, CA, IL IA, FL, CT, and MO | Terminated | December 04, 2012 | 546 | 20 | Computer
Z-0006-2012 | 59139 | No | LYJ |  |  | Neurology | PMA | Model 104 VNS Therapy Demipulse Duo Generator Cyberonics, Inc. Houston, Texas, Model 104 Generators are provided in a sterile single use packages as part of the VNS Therapy System. Product labeling may be found at http://us.cyberonics.com/en/vns-therapy-for-epilepsy/healthcare-professionals/vns-therapy/manuals-page. |  |  | May 06, 2011 | 2011 | Cyberonics, Inc | An investigation was initiated based on a report from the field in which an Intensive Follow-up Indicator (IFI) message was unexpectedly received by a medical professional when using Model 250 version 8.0 software to interrogate a patient's Model 103 Generator. | Software design (manufacturing process) | The firm, Cyberonics Inc., initiated containment actions on March 17, 2011 to prevent additional affected devices from being shipped. The firm initiated contact with implanting facilities in the United States on May 6, 2011 and international facilities and distributors on May 10, 2011 to remove product from the field that has not yet been implanted. Additionally, a "Field Safety Alert" letter was sent to all affected medical professionals. The letter identifies the affected product, the issue, and the action(s) to be taken by the physicians. Consignees are instructed to complete and return an attached Effectiveness Check Form. If assistance is needed physicians may contact Clinical Technical Support at 866-882-8804 or via email at clinicaltechnicalservices@cyberonics.com |  | 13 | 13 units | Worldwide Distribution - USA (nationwide) including the states of AL, AZ, CA, CO, FL, GA, ID, IL, KS, OH, AR, MA, MI, MN, MO, NC, NH, NJ, NM, NY, OH, OR, PA, TN, TX, UT, WA, WI and WV. Products were also distributed to WASHINGTON DC and Puerto Rico and the countries of: Austria, Belgium, Finland, Germany, Ireland, Italy, Netherlands, Norway, Spain, Sweden, Switzerland, and United Kingdom. | Terminated | December 04, 2012 | 578 | 105 | Computer
Z-0007-2012 | 59847 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Outlook ES Safety Infusion System, Model 621-100ES, 621-200ES, 621-300ES, and 621-400ES |  |  | July 20, 2010 | 2010 | B Braun Medical, Inc | The Outlook ES Dose Guard Care Area can inadvertently be exited when the pump enters the KVO (Keep Vein Open) state. The same condition can occur when a sequence of HOLD-HOLD key strokes are performed on the pump key panel. | Software design | The firm, B. Braun Medical Inc., sent an "URGENT: FIELD CORRECTION" letter dated September 8, 2011 to its customers. The letter described the product, problem and actions to be taken. The firm informed the customers that a B. Braun representative will be contact ing them to schedule a field software upgrade of the affected pumps. The customers were instructed to complete and return the FIELD CORRECTION INFORMATION FORM via mail using the enclosed envelop and contact B. Braun Customer Service for shipping instructions at 1-800-627-7867 if product is to be returned.Should you require any additional information, please do not hesitate to contact B. Braun Medical Inc., Technical Product Support at 1-800-627-7867 (800-627-PUMP). |  | 11525 | 11,525 pumps | Worldwide distribution: USA (nationwide) including states of: AR, CA, FL, GA, IA, IL, KS, KY, LA, MA, MI, MO, NC, NJ, NY, OH, PA, SC, TX, VA, WA, WI, and WV; and country of: Canada. | Terminated | January 07, 2015 | 1632 | 11525 | Computer
Z-3163-2011 | 59427 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | TomoMobile, Hi-Art System, H-0000-0003, TomoTherapy 1240 Deming Way, Madison, WI 53717 |  |  | July 07, 2011 | 2011 | Accuray Incorporated | TomoTherapy Inc. is sending this Field Safety Notice to make you aware of an anomaly which may affect the performance of the TomoMobile Hi-Art System.TomoTherapy has discovered that when attempting to open the TomoMobile shielding doors, while the door hinges are in the locked position, the hinges may fail allowing the door to disengage from the shielding. | Device Design | The firm, TomoTherapy, sent an "Urgent Medical Device Correction" letter dated July 7, 2011 to its consignees/customers. The letter described the Issue, Product Affected, Recommended Action (advised consignees to not attempt to open the shielding doors when the hinges are in the locked position. The Hinges should always be in the unlocked position when opening the doors) and Resolution. A TomoTherapy field service engineer or an authorized service representative will contact the customers when a resolution is available.If you have any questions, please contact the TomoTherapy Technical Solutions Center by email at support@tomotherapy.com or phone at: North America: 1-866-368-4807; Belgium: 0800 38783; France: 0805 631 565; Germany: 0800 000 7541; Italy: 800 986 399; Netherlands: 0800 020 1364; Spain: 800 300049; Switzerland: 0800 001927; United Kingdom: 0808 238 6035; China/CNC (Northern): 10 800 712 1701; China (Southern): 10 800 120 1701; Hong Kong: 800 967912; Japan: 0044 22 132374; Singapore: 800 1204 683; South Korea: 0079 81 4800 7204, all other locations: +1 608 824 2900 or +32 2 400 44 44. |  | 4 | 4 | Worldwide distribution: USA including states of: AZ, CA and OK; and country of: United Arab Emirates | Terminated | June 29, 2012 | 358 | 4 | Computer
Z-1392-2011 | 53691 | No | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | V5M TEE Transducer, when used with an ACUSON S2000 ultrasound system that has cardiac options installed and is at software version VA10, VA10A, VA10B, VA10C, VA15, VA15A, VA15B, VA16 or VA16A, manufactured by Siemens Medical Solutions, Mountain View, CA. |  |  | August 28, 2009 | 2009 | Siemens Medical Solutions USA, Inc. | If the transducer is not disconnected from the system after an automatic shutdown due to thermal limits being exceeded, then the next time it is used, the thermal safety guards do not take effect. Normally, the thermal limit before automatic image freeze (and therefore power shut off) occurs at 43 degrees Centigrade. | Software design | Siemens sent a Customer Safety Advisory Notification letter on August 28, 2009, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed that their local Customer Service Engineer would install a software update as soon as it becomes available. Customers were instructed to pass the Customer Safety Advisory Notification letter to all who need to be aware of these issues until the corrective action is completed. For questions regarding this recall call 650-969-9112 or 800-422-8766. |  | 97 | 97 units | Worldwide Distribution - USA (nationwide) and internationally to Canada, China, Czech Republic, Denmark, Germany, Great Britian, Hong Kong, Indonesia, Italy, Japan, Norway, Poland, Portugal, Russian Federation, Singapore, South Korea, Taiwan, and Turkmenistan. | Terminated | March 29, 2011 | 578 | 97 | Computer
Z-3280-2011 | 58556 | walker, mechanical | ITJ | Mechanical walker. | Physical Medicine | Physical Medicine | 510(K) Exempt | Dual Blue-Release Adult Walkers - Wheeled, and Blue Release Junior Walkers with 5 inch Fixed Wheels, Model Numbers: 6240-5F and 6240-JR5F. The recalled product is packaged four (4) units per shipping box. |  |  | March 02, 2011 | 2011 | Invacare Corporation | The firm is conducting a recall/correction of their adult and junior wheeled walkers due to a component defect. The firm determined that the caster bolt on the wheeled extension was too short to adequately maintain the Nylex fastener. Consequently, over time, this condition incurs the potential for the caster to loosen and detach, causing the wheels to fall off the walker. This scenario would res | Device Design | The firm, Invacare Corporation notified their customers of this recall/corrective action via an 'URGENT RECALL INFORMATION' letter dated April 7, 2011. The letter was sent to the customers by registered mail or Federal Express. The letter informs the customer(s) of the component defect involved and the inherent possibility that over time, a walker wheel(s) could fall off resulting in instability of the product and potential for serious injury to the user of the device. The letter provides a list of the affected units sold to each facility on an enclosed tracking sheet. It also asks that the customer: 1) Examine their stock for the presence of any of the affected walkers and to quarantine any such products found; 2) Complete an attached recall response form indicating the customers receipt and understanding of the instructions provided in the letter and to fax the completed form to the telephone number listed on the bottom of the form. The letter goes on to state that once the faxed back form is received by Invacare, the company will send the customer replacement wheels kits for all affected walkers purchased; 3) the customer is asked upon receipt of the replacement wheel kits to replace the wheels of any units still in stock and discard all of the old wheel parts removed; 4) Lastly, the customer is asked to contact their sub-account customers and notify them of the recall. If the customer have any questions concerning the instructions, call the recall support line at 1-866-567-7899. |  | 12764 | 12,764 walkers were distributed. | Worldwide Distribution - USA (nationwide: AL, AK, AZ, AR, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VT, VA, WA, WV, including Puerto Rico and countries of Canada, Columbia, and Israel. | Terminated | July 02, 2013 | 853 | 12764 | Computer
Z-3282-2011 | 59527 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | ADVANTAGE WINDOWS (Workstation) WITH FUNCTOOL OPTION |  |  | August 18, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware that image orientation of reference series displayed in Functool Diffusion Tensor/FiberTrak protocol on Advantage Workstation may not match with the orientation annotations displayed in the viewport. This issue may impact patient safety. | Software design | GE Healthcare sent "Urgent Medical Device Correction" letters dated August 18, 2011 to all affected customers.The letter included name of product affected and described problem. Recommended actions where provided to customers and they were notified a GE Healthcare representative will contact them for a software system upgrade.For question of this recall contact the GE Service Representative at (800) 437-1171. |  | 697 | 697 | Worldwide Distribution: Nationwide distribution including the states of AZ, CA, CT, DC, FL, ID, IL, KY, LA, MD, MI, MN, NJ, NY, OR, PA, PR, TN, TX, VA, WA, and WI; and the countries of VIETNAM, MEXICO, UNITED KINGDOM, TURKEY, TUNISIA, THAILAND, TAIWAN, SWITZERLAND, SWEDEN, ST HELENA, SPAIN, SLOVENIA, SINGAPORE, SAUDI ARABIA, RUSSIAN FÂżDÂżRATION, ROMANIA, REUNION, REPUBLIC OF CHINA, QATAR, PORTUGAL POLAND, PARAGUAY, NETHERLANDS, MOROCCO, MALAYSIA, MACEDONIA, LUXEMBOURG, LIBYA, KUWAIT, KOREA, JAPAN, ITALY, ISRAEL, INDONESIA, INDIA, HUNGARY, GREECE, GERMANY, FRANCE, FINLAND, ESTONIA, EGYPT, CZECH REPUBLIC, CROATIA, COSTA RICA, COLOMBIA, CHILE, CANADA, BELARUS, BRAZIL, BOSNIA, BELGIUM, BANGLADESH, AUSTRIA, AUSTRALIA, and ARGENTINA. | Terminated | August 09, 2012 | 357 | 697 | Computer
Z-1388-2011 | 57245 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | S-ROM Distal Femoral Cutting block Guide SZ X- SM/SM, REF 216311001, non-sterile, Depuy International Leeds, England. |  |  | October 19, 2010 | 2010 | Depuy Orthopaedics, Inc. | The firm was notified that the Hinge Block Femoral Box Guide cut being off alignment when using the S-ROM Distal Femoral Cutting Block Guide SZ SML/X-SML pin holes to locate. The firm then found that the same issue was found int the MED block size. | Device Design | The firm, DePuy, sent an "URGENT INFORMATION - DEVICE CORRECTION" email and letter dated October 19, 2010 to all customers. The notices described the product, problem and actions to be taken by the customers. The notices stated that: Depuy is issuing a device correction; however, the product may still be used; the devices do not need to be returned; the distributors were to provide a letter to clinicians that were using the device; the surgical technique will be revised in October 2010; the product will undergo a design change, and the recalled product will be traded out in December 2010 and January 2011.If you have any additional questions, please contact DePuy's Scientific Information office at 1-888-554-2482.On 3/4/2011, The firm sent out a revised URGENT DEVICE RECALL asking that all affected cutting guides be taken out of service immediately. Devices should be returned to DePuy and new cutting guides will be sent out within 2 weeks. |  | 229 | 229 | Worldwide distribution: USA including states of: AZ, AR, CA, CO, FL, GA, IA, ID,IL, IN, KY, KS, LA, MA, MD, MI, MN, NC, NV, PA, TN, TX, VA, WA, and WI; and countries including: Canada and Ireland. Australia, Austria, Canada, Denmark, Ireland, Israel Italy, New Zealand, Russia, Sweden, Switzerland, Spain, and UK. | Terminated | December 03, 2012 | 776 | 448 | Computer
Z-3276-2011 | 59734 | No | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | LOGIX Order Entry (OE) Software, all versions, catalog 2M8410; Baxter Healthcare, Deerfield, IL 60015; The software is included in a number of Baxter's pharmacy compounders. The software assists the pharmacist in management of compounding, creating and authorizing nutritional solutions intended for IV administration. |  |  | August 15, 2011 | 2011 | Baxter Healthcare Corp. | There is the potential for an order entry error during use of the LOGIX Order Entry (OE) software. When adding ingredients to a TPN order, the user can accept the pre-populated "per units" of "Bag" or select other "per units". A dosing error may occur if the user does not select the prescribed "per units" for the added ingredient. | Software design | Baxter Healthcare Corp. sent an "Urgent Device Correction" letters dated August 15, 2011 which included LOGIX OE screen illustrations of a representative neonate and adult patient, via first class mail to all affected customers. The letter identified the product, problem and actions to be taken. The customers were instructed to follow safe compounding practice, which includes verifying that the ingredient per units entered match the TPN prescription prior to preparation. Baxter requests that customers communicate this issue to all personnel responsible for LOGIX order entry and order authorization at their facility. Customers are instructed to retain a copy of the letter at each computer on which LOGIX OE software is installed. The customers were also requested to complete the attached customer reply form and fax it to Baxter at the number provided on the form. Any questions regarding the communication were directed to Baxter's Medical Information Services at 1-800-422-2751. |  | 95 | 95 units | Worldwide Distribution - USA (nationwide) including Puerto Rico and the country of Canada. | Terminated | January 05, 2017 | 1970 | 95 | Computer
Z-3277-2011 | 53897 | No | NSX |  |  | General Hospital | Contact ODE | CareSuite" - Critical Care Manager, PACU Manager and Anesthesia Manager Software, transmission &amp; storage, patient data. |  |  | August 12, 2009 | 2009 | Picis Inc. | Electronic health record software applications malfunctioned in a manner which failed to display the discontinued status for medication orders, as entered by the attending physician | Software design | Picis Inc. sent a customer notification letter dated August 11, 2009 to all affected customers. The letter describes the product, problem, and action to be taken. Customers informed of the potential patient safety error were advised to make the appropriate configuration change, within current software versions 7.1.4 or later. Contact your local Picis support representative for questions regarding this notice. |  | 8576 | 8576 UNITS | Worldwide Distribution-USA (nationwide) and the countries of Australia, Austria, Brussels, Belgium, Denmark, Finland, France, Germany, Malaysia, Netherlands, Norway, Portugal, Singapore, Spain, Sweden, and UK. | Terminated | September 27, 2011 | 776 | 8576 | Computer
Z-3269-2011 | 59664 | No | ODX | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | CombiSetÂż True Flow" 2008K@home Bloodlines |  |  | August 11, 2011 | 2011 | Fresenius Medical Care Holdings, Inc. | Blood lines coiling tabs #1 and #5 on the arterial line are not in the correct positions | Device Design | Fresenius Medical Care sent out an "Urgent" recall notice dated August 11, 2011 to all affected customers .The letter included description of the product name, problem and asked them to examine their stock immediately and quarantine bloodlines from the recalled lots. Customers were asked to return the product to Fresenius.For information on this recall call the Fresenius Customer Service Team at 1-800-323-5188. |  | 118 | 118 cases | Nationwide Distribution. | Completed |  |  | 118 | Computer
Z-3265-2011 | 59303 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Powerheart Automated External Defibrillator (AED), G3 is indicated for emergency treatment of victims exhibiting symptoms of sudden cardiac arrest who are unresponsive and not breathing. Post-resuscitation, if the victim is breathing, the AED should be left attached to allow for acquisition and detection of the ECG rhythm. If a shockable ventricular tachyarrhythmia recurs, the device will charge automatically and be able to deliver therapy. |  |  | June 30, 2011 | 2011 | Cardiac Science Corporation | Cardiac Science Corporation (CSC) has identified two errors, 0x48 and 0x99, that have repeatedly occurred during a device self-test in the G3 installed base over the last 6 years. These self test errors cause the device to enter into a service required state.CSC remind owners/operators of the Powerheart AED devices to verify the "Rescue Readiness" of their equipment and follow maintenance pro | Device Design | Cardiac Science began mailing out the recall notification letter on August 1, 2011, to all their US consignees via USPS first class mail. Cardiac Science's distribution partners will receive a notification letter and be provided with a template letter to send to their customers directly.CSC remind owners/operators of the Powerheart AED devices to check the status of of the AED every day to ensure it is Rescue Ready, and perform the regularly schedule maintenance as outlined in the Powerheart AED "Operator and Service Manual" provided with every AED.Customers can call Cardiac Science at 800-426-0337 or e-mail at aed205@cardiacscience.com. |  | 195739 | 195,739 units in US | Worldwide Distribution - US (nationwide)This action applies only to domestic | Terminated | February 13, 2014 | 959 | 195739 | Computer
Z-1409-2011 | 57487 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Stryker Navigation System - iNtellect Cranial Planning Software - Enable, Part number 6000-652-300. It is part of the product series of the Stryker Navigation System. The system comprises software for pre-operative surgical planning and supports computer assisted surgery based on a wireless optical tracking localization device for the use in navigated cranial surgery. It includes automatic multi-modality correlation. The system provides an easy-to-use work flow concept for image importation, FESS, Craniotomy, Biopsy, etc. Additional, automatic segmentation of brain surface, tumor, vessels, etc. are supported. |  |  | November 10, 2010 | 2010 | Stryker Instruments Division of Stryker Corporation | When users imports more than one image set and changes the viewing direction for both images, the orientation of every other image series might be stored incorrectly, as a mirror image. This may lead to insufficient or inaccurate correlation results or incorrectly shown orientation of every second image series during navigation. | Software design | Urgent: Medical Device Recall Notification letters were sent on 12/3/2010, and software upgrades were to be performed in the field by the end of 2010. Specific Customer Instructions: 1) Locate the units listed in the notification.2) Distribute this notification to all departments affected.3) Return the enclosed business reply form to confirm receipt of this notification and identify how many affected units are currently in your inventory. 4) Once the new software is available, Stryker will dispatch a service tech in order to update the software. Note: Software upgrades for all applications will be available by the end of the year. The upgrades will be scheduled as we receive your business reply form.Note: If you currently have a Time License or Pay Per Use card, a new one will be issued to you at the time of the update.5) If you have loaned or sold any of the units listed in this notification, please, forward a copy of this notice to the new users and advise Stryker of the new users name and their new location.6) If you have disposed of any of the affected units and they are no longer in use, please, advise us of their obsolescence by providing us with their serial numbers.Questions regarding this recall should be directed to Stryker Field Action Team at 800-800-4236, ext 3808, 3584, or 4354 between the hours of 8 a.m. - 5 p.m. Monday - Friday. |  | 1155 | 1,155 total | Worldwide Distribution -- USA, Canada, Mexico, Australia, Argentina, Brazil, United Kingdom, Japan, Poland, S. Korea, Hong Kong, Taiwan, India, and New Zealand. | Terminated | June 01, 2011 | 203 | 1155 | Computer
Z-3221-2011 | 59519 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | TSX-101: Aquilion 32/64; System, X-Ray, Tomography, Computed. |  |  | December 08, 2010 | 2010 | Toshiba American Medical Systems Inc | There is the potential that the current may not be controlled correctly in examinations using the ECG modulation function, which adjusts the x-ray output in synchronization with the ECG waveform. | Software design | Toshiba America Medical Systems (TAMS) issued an Urgent: Medical Device Correction notification to consignees via USPS return receipt mail. The letter identified the affected product, the problem, and corrective action. The firm plans to install revised firmware to prevent the occurrence of the problem. A Toshiba Service Representative will contact customers to schedule an appointment for installation. The firm asked customers to stop using the Power Reduction option during ECG modulated scanning until the new firmware has been installed. Customers are to share the information in the letter with all users, reviewing radiologist, and the clinical engineering or Biomedical group at their facility. Customers are also to complete and return the attached reply form via fax or e-mail to TAMS as proof of notification. Questions pertaining to the letter should be directed to 800-421-1968 or customers' local Toshiba Representative at 800-521-1968. |  | 528 | 528 units | Nationwide Distribution. | Terminated | March 18, 2013 | 831 | 528 | Computer
Z-3206-2011 | 57889 | No | DPZ | Ear oximeter. | Cardiovascular | Anesthesiology | 510(k) | DATEX-OHMEDA TRUSAT PULSE OXIMETER AND ACCESSORIES, MODEL 3500 (GE Healthcare) |  |  | April 08, 2011 | 2011 | GE Healthcare, LLC | The plastic enclosure on the external power supply of the TruSat Pulse Oximeter is prone to breakage, and overheating. The issue may impact patient or user safety. Potential for system or components to start on fire or explode (general case). | Device Design | On 4/8/11 GE Healthcare sent consignees an "Urgent Medical Device Correction" letter dated March 30, 2011. The letter was addressed to Healthcare Administrator / Risk Manager, Chief of Nursing, Director of Biomedical Engineering &amp; Home Healthcare Provider.The letter listed the Safety Issues, Safety Instructions, Affected Product Details, Product Correction and Contact Information. |  | 31846 | 31,846 | Worldwide distribution: USA, ZAMBIA, VIETNAM, VENEZUELA, UNITED KINGDOM, UNITED ARAB EMIRATES, UGANDA, TURKEY, TUNISIA, TRINIDAD &amp; Tobago, THAILAND, UNITED REPUBLIC OF TANZANIA, PROVINCE OF CHINA TAIWAN, SWITZERLAND, SWEDEN, SPAIN, SUDAN, SOUTH AFRICA, SLOVENIA, SLOVAKIA, SINGAPORE, SERBIA, SAUDI ARABIA, RUSSIAN FEDERATION, ROMANIA, QATAR, PORTUGAL, POLAND, PHILIPPINES, PERU, PANAMA, OMAN, NORWAY, NICARAGUA, NEW ZEALAND, NETHERLANDS, NEPAL, MOZAMBIQUE, MEXICO, MALTA, MALAYSIA, MACEDONIA, LITHUANIA, LIBYAN ARAB JAMAHIRIYA, LEBANON, LATVIA, KUWAIT, REPUBLIC OF KOREA, KENYA, JORDAN, JAPAN, ITALY, ISRAEL, IRELAND, IRAQ, INDONESIA, INDIA, HUNGARY, HONG KONG, HONDURAS, GUATEMALA, GUADELOUPE, GREECE, GERMANY, FRANCE, FINLAND, ESTONIA, KOREA, EL SALVADOR, EGYPT, ECUADOR, DENMARK, CZECH REPUBLIC, CYPRUS, CROATIA, COSTA RICA, COLOMBIA, CHINA, CHILE , CANADA, BRITISH INDIAN OCEAN TERRITORY, BRAZIL, BOLIVIA, BENNIN, BELGIUM, BANGLADESH, BAHAMAS, AUSTRIA, AUSTRALIA, ARGENTINA, and ALGERIA. | Terminated | March 04, 2016 | 1792 | 31846 | Computer
Z-3107-2011 | 59470 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | GemStar Li-Ion Docking Station; a lithium ion battery powered docking station for use with GemStar Infusion Pumps; Hospira, Inc., Lake Forest, IL 60045; List Numbers 13075-04-01, 13075-04-03, 13075-04-05, 13075-04-07, 13075-24-01, 13075-24-03, 13075-24-05 and 13075-24-07, Note: list numbers with xxxxx-04-xx are for U.S., Canada, Japan, Netherlands and Singapore, while those with xxxxx-24-xx are for international distribution only. |  |  | August 01, 2011 | 2011 | Hospira Inc. | Hospira has received reports of sparking, smoking, charring and electrical shock when using the GemStar Docking Station due to fluid ingress into the docking station. | Device Design | Hospira Inc. sent an "URGENT DEVICE FIELD CORRECTION" letter dated August 1, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The requests that customers do not hang or place fluid containers over the docking station, not spray fluid directly on the docking station, and to clean the docking station using a cloth dampened with an apporoved cleaning solution. A reply form was attached to the letter for customers to complete and return via fax at 262-577-6921.Contact Hospira at 1-800-441-4100 for questions concerning this notice.*** Hospira sent an update letter dated September 26, 2012 to the accounts, informing them of the availability of the improved design replacement docking stations, and requesting them to complete and return the enclosed reply via fax or e-mail to Hospira, indicating the number of docking stations to be returned for replacement at no charge if the old docking stations are returned within 45 days of receipt of the replacement stations. *** |  | 18000 | 18,000 units | Worldwide Distribution-USA (nationwide) and the countries of Australia, Belgium, Canada, Chile, France, Germany, Hong Kong, Ireland, Italy, Japan, Jordan, the Netherlands, Portugal, Saudi Arabia, Singapore, Spain, Sweden and Switzerland. | Open, Classified |  |  | 18000 | Computer
Z-1495-2011 | 53690 | No | JQW | Pipetting and diluting system for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | PhD System EIA/IFAversion 2.1A software, model number 426-0246, a component of the Bio-Rad PhD System. Manufactured by Bio-Rad Laboratories, France, Distributed by Bio-Rad Laboratories, Hercules, CA. |  |  | September 29, 2009 | 2009 | Bio-Rad Laboratories Inc | Incorrect patient results are being generated from out-of-range samples tested withc ompetitive enzyme immunoassays on the PhD system with software version 2.1A. | Software change control | BIO-RAD sent an URGENT: MEDICAL DEVICE CORRECTION letter dated September 29, 2009, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately examine their customer records to identify all PhD System customers that are using version 2.1A of the EIA/IFA Software and Methods CD-ROM.Customers who are not running competitive immunoassays on the PhD System may continue using the 2.1A software.Customers who run competitive immunoassays must revert to software version 2.0B.Customers were instructed to translate if necessary and add local contact information to the customer letter and distribute to all PhD System customers that are using 2.1A of the EIA/IFA Software and Methods CD-ROM. A Customer Medical Device Correction Response Form must be collected from each customer to ensure that they have received this important communication. Once forms have been collected from all applicable customers, they should be faxed to Bio-Rad CSD Regulatory Affairs Department along with the completed Subsidiary Medical Device Correction Response Form.For questions regarding this recall call 510-741-3954. |  | 262 | 262 units | Worldwide Distribution - USA including AL, CA, CO, CT, DC, FL, IL, MA, MD, MN, MO, MS, MT, NC, NY, OH, PA, SD, TN, TX, UT, and VT and the countries of Australia, New Zealand, Singapore, China, France, Sweden, and Canada | Terminated | March 11, 2011 | 528 | 262 | Computer
Z-3184-2011 | 59797 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Straumann coDiagnostiX, version 8.0, used to guide a clinician in the proper implant selection, oral placement and the proper selection of an oral drill. |  |  | August 04, 2011 | 2011 | Straumann Manufacturing Inc. | Software version 8.0 has a software anomaly, in which the length is not adjusted accordingly when it is changed after the surgical sleeve position planning step has been completed. | Software design | Staumann USA notified accounts by an Urgent Field Safety Notice letter dated August 4, 2011 and delivered via Federal Express Delivery Service.The letter identified the affected product and explained the reason for the recall. Customers were asked to review the letter, take appropriate actions, including completing and faxing the enclosed. Customer Confirmation Form. Straumann released a corrected software upgrade (coDiagnostiX 8.0.2.3672), which is available for download when users go online and activate current software. If the user does not have internet access a replacement CD can be ordered from Technical Support at 866 - 531 -7365. Questions should be directed to Technical Support Monday through Friday from 8:00 am to 8:00 pm Eastern Time. |  | 272 | 272 | Nationwide Distribution | Terminated | May 23, 2012 | 293 | 272 | Computer
Z-3195-2011 | 59690 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | TSX-301B: AQUILION Premium; System, X-Ray, Tomography, Computed |  |  | June 22, 2011 | 2011 | Toshiba American Medical Systems Inc | Toshiba America Medical Systems is initiating a field correction on their TSX-301B: AQUILION Premium; System, X-Ray, Tomography, Computed,TSX-301A: AQUILION One; System, X-Ray, Tomography, Computed,TSX-20IA: AQUlLlON LB; System, X-Ray, Tomography, Computed, and TSX-1O1A: AQUILION 16; System, X-Ray, Tomography, Computed due to several software issues that cause the scanners to function incorr | Software design | Toshiba America Medical Systems (TAMS) sent an URGENT MEDICAL DEVICE CORRECTION letter dated June 22, 2011, to all affected customers. Customers were instructed that a revised software to prevent occurrence of the problem would be installed on their system. Their Toshiba Service Representative would contact them for an appointment to install the software to their system when it became available.Customers were requested to take temporary measures until corrective measures were taken.1) In dual scanoscopy, do not press the short key between the completion of the first scanoscopy and the start of the second scanoscopy.2) In dual scanoscopy, do not click the (Scano Skip) button immediately before the completion of the first scanoscopy.3) When the SplitCase function is set to ON, do not select two or more orders simultaneously from the appointment list.4) When an image comment is to be preset by making a selection from the preset memo, do not select any image comment with more than 10 characters. Customers were asked to share the information with all users and reviewing radiologists as well as clinical engineering or Biomedical group at the facility.Customers were asked to complete and return the attached form and fax it to 877-349-3054.For any questions call (800) 421-1968 or contact your local Toshiba Representative at (800) 521-1968. |  | 155 | 155 units total | Nationwide Distribution | Terminated | August 11, 2014 | 1146 | 155 | Computer
Z-3200-2011 | 59175 | [6 empty] | Philips GEMINI TF64 Diagnostic Imaging Systems, Model Number: 882471, 510(k) #K052640, Serial Numbers: 7502, 7504, 7506, 7514, 7083 and 7164. The recalled GEMINI TF64 system units equipped with software versions: 3.5.1.1 and 3.5.2.1 are subject to recall/software correction. |  |  | May 31, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | The firm is conducting a recall/software correction to several of their GEMINI PET/CT X-ray system units in the field. The firm believes that if these software defects were to recur, it could impact the performance of reliability of the these PET/CT system units. | Software design | The firm , Philips Healthcare, issued a 'Customer Information Letter' dated August 9, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were not required to perform any action. A Philips' Field Service Engineer will visit each customer site and install the necessary software upgrade and Philips will track the progress of the software corrections as they take place.For customers in North America, if you need any further information or support concerning this update, please contact our Customer Care Solutions Center at 1-800-722-9377. Selection option 5 for "All Imaging Systems". Enter your site ID number. If you do not know your site ID, please hold and then select option 5 for "Nuclear Medicine" and finally select option 2 for "PET" support. In all other countries, the local Philips Healthcare office should be contacted at 440-483-3000. |  | 6 | 6 units | Worldwide distribution: USA including states of: FL, MA, NE, NC, OH, TX, WA and WV; and countries of: France, Germany, Italy, Kuwait, The Netherlands, Spain and Turkey. | Terminated | February 22, 2013 | 633 | 19 | Computer
Z-3178-2011 | 59331 | No | LPB |  |  | Cardiovascular | PMA | STOCKERT 70 Radio Frequency Generator, Catalog Number 57001 |  |  | July 07, 2011 | 2011 | Biosense Webster, Inc. | Biosense Webster has initiated a voluntary field notification to all customers that have a Stockert 70 Radio Frequency Generator, to communicate undetected software anomalies in the form of software release note in order to amend the labeling and thereby enhance the safety of the product. | Software design (manufacturing process) | The recall communication was initiated on 7/8/11 and 7/11/11 with Biosense Webster forwarding A Customer Notification Letter, Acknowledgement Form and Software Release Note (Attachments 3, 4 and 1) via FedEx to all U.S. customers who purchased the Stockert 70 Radio Frequency Generator. Two letters and forms were sent to U.S. customers via express mail, addressed to the EP/Cardiology Lab and Risk Management Office, respectively.The Release Note provides tips for the avoidance of potential events when working with the EP-SH UTTLE, STOCKERT 70, STOCKERT J50 or STOCKERT J70 RF generators with SW versions 1.035 / 1.035J /1.035J70 to 1.037/1.037J70. |  | 3397 | 3,397 units total (1,217 in U.S.) | Worldwide Distribution -- Algeria, Argentina, Australia, Austria, Belarus, Belgium, Brazil, Brunei Darussalam, Canada, Chile, China, Colombia, Croatia, Cyprus, Czech Republic, Denmark, Egypt, England, Finland, France, Germany, Greece, Guatemala, Hong Kong, Hungary, Iceland, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Korea, Latvia, Lebanon, Lithuania, Luxembourg, Malaysia, Malta, Mexico, Morocco, Myanmar, Nepal, Netherlands, New Zealand, Norway, Pakistan, Philippines, Poland, Portugal, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Tahiti, Taiwan, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom, United States, Uruguay, and Vietnam. | Terminated | October 12, 2012 | 463 | 3397 | Computer
Z-1473-2011 | 57083 | [6 empty] | GE ADVANTAGE SIM MD versions 5.x, 6.x, 7.x up to and including version 7.7.0 |  |  | November 24, 2010 | 2010 | GE Healthcare, LLC | GE Healthcare has become aware that inconsistent DRR (Digitally Reconstructed Radiographic) visualization may occur when the blended DRR feature or the Isocenter Move mode feature is used in combination with specified other functions of the AdvantageSim application on Advantage Workstation. These inconsistencies may affect patient safety. Other AdvantageSim functions are not affected. | Software design | GE Healthcare sent an "Urgent Medical Device Correction" letter dated November 23, 2010. The letter was addressed to Hospital Administrators / Risk mangers, Oncology Department Managers / Radiation Oncologists, &amp; Radiation Physicists / Dosimetrists / Therapists. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information.For any questions or concerns regarding this recall, please call one of thefollowing phone numbers:United States: 800 437 1171Japan: 0120 - 055 - 919For other countries, please contact your local GE Healthcare Service Representative. |  | 1531 | 1531 | Worldwide Distribution - USA including AL, AZ, CA, CO, CT, DC, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MI, MA, MI, MN, MS,MO, NT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, PR, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, and WY and the countries of VENEZUELA, URUGUAY, UNITED KINGDOM, UNITED ARAB EMIRATES, TURKEY, THAILAND, TAIWAN, SWITZERLAND, SWEDEN, SPAIN, SOUTH AFRICA, SINGAPORE, SAUDI ARABIA, RUSSIAN FEDERATION, REPUBLIC OF KOREA, PORTUGAL, POLAND, PAKISTAN, NOWAY, NIGERIA, NETHERLANDS, MEXICO, MALAYSIA, LITHUANIA, LATVIA, JAPAN, ITALY, ISRAEL, IRELAND, INDIA, HUNGARY, HONG KONG, GREECE, GERMANY, FRANCE, FINLAND, EGYPT, ECUADOR, DOMINICA REPUBLIC, DENMARK, CZECH REPUBLIC, CHINA, CHILE, CANADA, BRAZIL BELGIUM, AUSTRIA, AUSTRALIA. | Terminated | August 03, 2012 | 618 | 1531 | Computer
Z-1474-2011 | 57419 | No | HBC | Electric cranial drill motor. | Neurology | Neurology | 510(k) | Aesculap MicroSpeed Uni Control Box, Part #GD670 |  |  | November 22, 2010 | 2010 | Aesculap Inc | Potential safety issue in that the control box can misinterpret the information sent from the low speed motor if the control magnet inside the motor changes its posthion. | Software design | AESCULAP sent an IMPORTANT CORRECTION NOTICE letter dated November 22, 2010, to all affected consignees. The letter identified the product, the problem, and the action the consignee should take. Consignees were instructed to not use the motors for thread cutting and removal of screws.If any unusual reactions of any cutting tools were observed consignees should check to ensure that the motor is not running in the opposite direction.If any error message is displayed, consigness were instructed to not use the motor.Consignees were also instructed to check both directions of running before use and do not use a motor which produces an error message in one direction. Aesculap would be providing all customers with a software update to the MicroSpeed Uni control box either by the unit being returned to Aesculap for update or by firm representative doing the upgrade on site at their facility.For questions regarding this recall call 800-258-1946 x5254 or 800-258-1946 x5267. |  | 91 | 91 units | Worldwide Distribution - USA including CA, CO, DE, FL, GA, IN, LA, MO, NC, NJ, OK, PA, TN, TX, and VA and the country of Canada. | Terminated | July 05, 2012 | 591 | 91 | Computer
Z-3183-2011 | 59421 | Yes | LWS |  |  | Cardiovascular | PMA | LATITUDEÂż Patient Management System Communicator, Model 6482. |  |  | May 25, 2011 | 2011 | Boston Scientific CRM Corp | Boston Scientific has identified two occurrences where the LATITUDE system did not process a data payload, which is the information uploaded to the LATITUDE system from a remote interrogation of the implanted device. These occurrences resulted in non-displayed payloads because the data is not posted to the website. If the payload is not displayed, the pertinent data will not be available on the | Device Design | Boston Scientific notified all affected customers by telephone on May 25, 2011. The customers were informed that there was an interrogation for one of their patients that was not able to be displayed on the website. Boston Scientific detected that a payload failure had occurred, but could not conclude what information was contained within the payload. Each patient has had previous and subsequent payloads that have processed and displayed on the LATITUDE website. |  | 2 | 2 | Nationwide Distribution-in the states of Florida and Texas. | Terminated | December 04, 2012 | 559 | 2 | Computer
Z-3162-2011 | 58473 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Ratcheting Torque Limiting Handle, Catalog Number 52-1012, Orthofix Spinal Implants, Firebird Spinal Fixation System, Firebird Spinal Fixation System DDD Instrument Case. |  |  | April 01, 2011 | 2011 | Orthofix, Inc. | Upon reaching torque limit, there is a possibility the product may jam or exhibit a ratchet direction switch from forward to reverse. | Device Design | Orthofix, Inc. sent an "URGENT PRODUCT RECALL NOTIFICATION" letter dated April 1, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter instructs customers to identify and remove the affected product from use. Customers were advised to use the combination of the Ratching R Handle and the Set Screw Driver if they experience problems with the Ratching Torque Limiting Handle. A Tracking and Verification form was attached for customers to complete and return to the Orthofix Inc. Regulatory Affairs Department via fax at 214-937-3322.Contact your Account Services Representative at 888-298-5700 for questions regarding the removal and return of the affected product. |  | 287 | 287 units | Worldwide Distribution-USA (nationwide) and Puerto Rico and the countries of Australia and Germany. | Terminated | March 23, 2012 | 357 | 287 | Computer
Z-3160-2011 | 59737 | No | NSX |  |  | General Hospital | Contact ODE | Picis ED PulseCheck-Software, transmission &amp; storage, patient data |  |  | August 01, 2011 | 2011 | Picis Inc. | Software does not display appropriate allergy interaction warning | Software design | Picis issued notification on 08/01/11 to affected clients, explaining the potential for the error, with instructions to monitor documented patient allergies closely when prescribing medication orders in ED PulseCheck, using Muitum drug database content. Picis will track the progress on acknowledgement of notification letter receipt through the implementation of software upgrade, for all potentially affected Customers.For questions regarding this recall call 781-557-3000. |  | 62 | 62 | Nationwide Distribution | Terminated | June 27, 2013 | 696 | 62 | Computer
Z-1522-2011 | 56660 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, ApexPro Telemetry System |  |  | January 28, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware of a potential for loss of patient monitoring when ApexPro Telemetry System is used in conjunction with a mirrored view CARESCAPE CIC Pro in combination monitoring mode (combo mode). The safety issue includes the potential for a delay in annunciation of audible alarms when used in combo mode, and may result in thedelay of treatment of a cardiac event. | Software design | GE sent an "Urgent Medical Device Correction" letter dated December 29, 2010. The letter was addressed to Healthcare Administrator / Risk Manger, Chief of Nursing and Director of Biomedical Engineering. The letter described the Safety Issue, Affected Product Details, Product Correction, Safety Instruction and Contact Information.Customers were instructed to contact Technical Support or their local Service Representative with any questions regarding this medical device.For any questions regarding this recall call 262-573-4122. |  | 2346 | 2346 | Worldwide Distribution - USA including AL, AK, AZ, AR, CA, CO, CT, DC, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO,MT,NE, NJ, NM, NY, NC, OH, OK, OR, PA, PR, SC, TN, TX, UT, VT, VA, WA, WV, WI, WY. and the UNITED KINGDOM, THAILAND, SWITZERLAND, SWEDEN, SPAIN, SOUTH KOREA, RUSSIA, NORWAY, NEW ZEALAND, NETHERLANDS, MALAYSIA, JAPAN, ITALY, ISRAEL, IRELAND, GERMANY, FRANCE, FINLAND, DENMARK, CUBA, CHINA, CANADA, AUSTRIA, AUSTRALIA, POLAND. | Terminated | May 29, 2012 | 487 | 2346 | Computer
Z-1498-2011 | 57303 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | SSH-880CV Aplio Artida; System, Diagnostic Ultrasound |  |  | September 28, 2010 | 2010 | Toshiba American Medical Systems Inc | Toshiba America Medical Systems is initiating a field correction on the SSH-880CV Aplio Artida Diagnostic Ultrasound System due to software issues. | Software design | Toshiba America Medical Systems (TAMS) sent an URGENT MEDICAL DEVICE CORRECTION letter to consignees via USPS return receipt mail. The letter contained a return reply form that is to be faxed to TAMS at 877-349-3054 as proof of notification.Customers were instructed to share the information with all users and reviewing radiologists as well as clinical engineering or biomedical group at their facility.For questions regarding this recall call 800-421-1968 or 800-521-1968. |  | 40 | 40 units | Nationwide Distribution to CA, GA, IA, IL KY, LA, MA, MN, MT, OH, OR, PA, TX, WI, and Puerto Rico. | Terminated | May 02, 2012 | 582 | 40 | Computer
Z-1520-2011 | 57841 | [6 empty] | VCareÂż Vaginal-Cervical Ahluwalia's Retractor-Elevator devices. The VCare IFU is packaged one per carton of eight VCare Pouched Devices. The Instructions for Use is an 8.5" x 12" 12-page booklet and is packed loose in the carton. Only the carton will have to be opened to replace the old IFU; there is no need to open the sterile packaging of the individual VCare pouches. |  |  | February 02, 2011 | 2011 | ConMed Corporation | Product manufactured after 4/11/2008, but before 8/01/2009, contains the earlier Instructions For Use, IFU (PN 17416, Rev. D) which does not contain the revisions on the addition of Warning Notes 10 and 11 and the expansion of Warning Note 6 to the VCareÂż Instructions for Use. The reason for the field correction is to provide a copy of the revised IFU to all customers who may have unexpired VcareÂż | Device Design | ConMed Corporation sent a communication to provide its direct consignees with important supplemental information concerning the Instructions for Use (IFU) for the VCAREÂż Vaginal-Cervical Ahluwalia's Retractor-Elevator uterine manipulators. The direct consignees were instructed to place a copy of the enclosed Instructions for Use, Part No. 14071 (Rev. A), with any inventory in your possession with an expiration date prior to August 2011 (2011-08) and replace any other Instructions for Use in thier possession. A copy of the enclosed Instructions for Use is already packaged with all devices bearing later expiration dates. Any inventory in their consignees possession with an expiration date prior to January 2011 (2011-01) is expired and should not be used. If the devices were further distributed to other facilities or distributors, or their consignees are a distributor, they were instructed to forward this information as appropriate to any known users of these devices.For questions regarding this recall call 315-624-3225. --- The IFU Warning changes are as follows: A new warning statement # 6 asks the user to verify the presence of all the components of the device upon removal of the VCAREÂż device from the patient. Although the prior IFU contained warnings about verifying that all components of the device were properly removed following use during surgery, the new statement expands the earlier, briefer statement. 6. Upon removing VCARE, the surgeon should visually inspect the VCARE device, and the patient, to make sure that the entire VCARE device was properly removed and that no components or fragments of these components were retained in the patient. There are 5 parts/components to the VCARE Cervical Elevator Retractor. These are: 1) the balloon; 2) the forward "cervical" cup; 3) the back or vaginal cup; 4) the locking assembly with thumb screw; 5) the metal shaft and handle with balloon inflation valve. --- A new warning statement # 10 further addresses |  | 31028 | 31,028 VCareÂż Vaginal-Cervical Ahluwalia's Retractor-Elevator devices | Worldwide Distribution - USA including NY, TX, CA, IA, AZ, PA, IN, and KS and the countries of Austria, Jordan, Turkey, China, Portugal, Chile, Belgium, Ecuador, Brazil, Romania, Japan, Malaysia, India Switzerland, Colombia, Netherlands, Spain, Saudi Arabia, France, and Canada. | Terminated | August 30, 2012 | 575 | 31028 | Computer
Z-3121-2011 | 59232 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE Healthcare, 1.5T Optima MR450w |  |  | July 29, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware of two potential safety issues associated with MR products (Optima MR450w, Discovery MR450, Discovery MR750, Brivo MR355, Optima MR360) that may impact patient safety.1) Inline post-processed (i.e. automated post-processing) image locations are not correctly annotated under some conditions. 2) Reference images displayed in Functool FiberTrak do not match with | Software design | GE Healthcare sent an "Urgent Medical Device Correction" letter dated July 27, 2011 to customers with affected systems. The letter described the recalled product, safety issues and recommended actions to be taken. The letter provided immediate mitigation (Safety Instuctions) that address each issue. The letter states that the recalling firm will contact customers to arrange to provide a software upgrade to all affected systems at no cost. For any questions concerning this notification contact your local service representative. |  | 159 | 159 | Worldwide Distribution - USA (nationwide) including the states of: AL, AR, CA, CO, CT, DE, DC, FL, GA, ID, IL, IN, IA, LA,ME, MD, MA, MI, MN, MS, MO, MT, NV, NH, NJ, NM, NY, NC, OH, OK, OR, PA, SC, TN, TX, UT, VT, VA, WA, WV, WT, WY.and the countries of: AUSTRALIA, AUSTRIA, BELGIUM, BRAZIL, CANADA, CROAT IA, CZECH REPUBLIC, DENMARK, EGYPT, FINLAND, FRANCE, GERMANY, GREECE, INDIA, INDONESIA, ISRAEL, ITALY, JAPAN, KUWAIT, MEXICO, NETHERLANDS, NEW ZEALAND, NORWAY, POLAND, REPUBLIC OF KOREA, RUSSIA, SAUDI ARABIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, TUKEYand UNITED KINGDOM. | Terminated | August 07, 2013 | 740 | 770 | Computer
Z-2996-2011 | 59404 | No | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | Titanium Fletcher-style Applicator Set - Defined Geometry - CT Compatible with Titanium intrauterine probe with stopper length 60 mm, 30 degree angle. |  |  | July 13, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | Potential of weld failure of the Titanium intrauterine probe at the cervical stopper. | Device Design | Varian Medical Systems sent an "URGENT MEDICAL DEVICE CORRECTION/URGENT FIELD SAFETY NOTICE" letter dated July 13, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter instructs customers to return all affected product to Varian Brachy Therapy by using the contact information provided in the letter. Replacement product will be provided free of charge to the customers.Customers should contact their local Varian Customer Support District or Regional Manager for questions regarding this notice. |  | 437 | 437 - all sizes | Worldwide Distribution. | Terminated | June 20, 2012 | 343 | 437 | Computer
Z-3003-2011 | 59122 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARTISTE Linear Accelerator with RT Therapist v4.1 system. |  |  | June 03, 2011 | 2011 | Siemens Medical Solutions USA, Inc | A potential safety issue was discovered when a pause for an imaging segment is inserted (where an imaging segment does not exist) with the ARTISTE and RT Therapist 4.1 system will continue to the next image. | Software design | Siemens sent an "Urgent: Medical Device Correction/ Customer Safety Advisory Notice" letter to all affected customers starting June 3, 2011. The letter describes the product, problem, and instructions to prevent mistreatment. The letter states that Siemens is working on an updated version of RTT 4.1 that will fix the issues. Users will be notified as soon as the update for the system is available. The customers were instructed to include Customer Safety Advisory notice in their syngo RT Therapist 4.1 System Owner Manual chapter "Safety Advisory Letters" where it should remain. |  | 223 | 223 active devices | Worldwide Distribution - USA (nationwide) including the states of: DE, FL, KY, LA, MA, MI, MO, NE, NJ, OH, PA, TN, WI, UT, WV and Puerto Rico and the countries of: Australia, Belgium, Brazil, Canada, China, Egypt, France, Germany, Hungary, India, Ireland, Italy, Lebanon, Malaysia, Netherlands, New Zealand, Philippines, Poland, Republic of Korea, Russian Fed, Saudi Arabia, South Africa, Norway, Spain, Turkey and United Kingdom. | Terminated | February 17, 2012 | 259 | 223 | Computer
Z-2976-2011 | 59296 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Allura Xper Systems Models: |  |  | July 08, 2011 | 2011 | Philips Healthcare Inc. | System freezes from 20 seconds up to one minute. No fluoroscopy exposure and no geometry movements during this time period. | Software design | Philips Health Care North America sent a FIELD SAFETY CORRECTION ACTION letter dated July 4, 2011, to all affected customers. The letter identified the product, the problem and the action needed to be taken by the customer.This FCD will be identified as FCD72200200 and will be free of charge. This FCD consists a CD with the correct firmware and a installation procedure to implement the FCD correctly.Please contact your local Philips representative. For further follow up please call (978 ) 687-1501. |  | 1423 | 1423 units | Nationwide Distribution | Terminated | May 14, 2014 | 1041 | 1423 | Computer
Z-2771-2011 | 55164 | No | LIG | Vitamin B[bdi1][bdi2] test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | UniCel DxC 880i, 860i,680i and 660i Synchron Access Clinical Systems |  |  | May 07, 2009 | 2009 | Beckman Coulter Inc. | The recall was initiated after Beckman Coulter confirmed reports of inaccurate Access Intrinsic Factor Ab test results displayed on, or printed from, the UniCel Dxl portion of an integrated workstation listed above. The problem is limited to numerical Intrinsic Factor Ab results produced on integrated systems running UniCel Dxl system software version 4.0, 4.1, or 4.2. The DxC portion of your inte | Software design | Beckman Coulter, Inc. sent out an "Urgent: Product Corrective Action" letter dated May 19, 2009 to all affected customers.The letter explained the reason for recall and provided recommendations including: -System software currently under development will correct this problem, and will be made available soon. Until then, please refer to UniCel DxC reports, user interface information, and LIS transmissions, which contain accurate Access Intrinsic Factor Ab test results. -Refer to the UniCel DxC system documentation for information on displaying and printing results from the DxC portion of an integrated workstation, or from your laboratory information system (LIS) -Past results for these assays only need to be reviewed if they were reported from the Dxl portion of an integrated workstation -Complete and return the enclosed response formFor additional information please contact the Customer Technical Support at 1-800-854-3633, or your local Beckman Coulter representative. |  | 5 | 5 units | Worldwide Distribution: (USA) Nationwide including KY; and the country of Canada. | Terminated | May 07, 2012 | 1096 | 5 | Computer
Z-1733-2011 | 55146 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Abbott brand CELL-DYN Emerald Hematology Analyzer System; List Numbers: 09H39-01; The product is manufactured in France for Abbott Diagnostics Division of Abbott Laboratories, Abbott Park, IL |  |  | March 23, 2010 | 2010 | Abbott Laboratories | Abbott has determined that gain settings for the non-standard specimen types may not match the gain settings for the standard specimen type. There is the potential for impact to product performance. | Software design (manufacturing process) | The firm, Abbott, sent two "Product Correction" letters one dated March 23,2010 and a follow-up letter dated July 1, 2010, via express mail to all its consignees/customers. The letters described the product, problem and actions to be taken. The customers were instructed to perform the following steps, noted in the letter, to print and verify their gain setting. Abbott informed the customers that they would receive a mandatory software update and were instructed to install the CELL-DYN Emerald according to the attached instructions, and to complete and return the Customer Reply form via fax to 1-800-777-0051 or email to QAGCO@abbott.com.Should you have any questions or would like to discuss in more detail, contact Director Quality Assurance and Regulatory Affairs Abbott Hematology at 408-567-3403. |  | 891 | 891 units | Worldwide distribution: USA including: AK, AL, AR, AZ, CA, CO, CT, FL, GA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NJ, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, VA, WA, WI, and WV; and countries including: Australia, Brazil, Canada, Columbia, Dominican Republic, Germany, Puerto Rico, Singapore, South Korea, St. Eustalius &amp; St. Martin, and Turks &amp; Caicos. | Terminated | March 18, 2011 | 360 | 891 | Computer
Z-2866-2011 | 59167 | No | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | MEDRAD Avanta Fluid Management Injection System, Catalog numbers: AVA 500 TABL, AVA 500 PEDL |  |  | June 02, 2011 | 2011 | Medrad Inc | Event was initially intended to be a product enhancement of the MEDRAD Avanta Fluid Management Injection System. Upgrades were initiated in October, 2007. Observance of a possible increase in complaints and adverse event reporting resulted in the firm issuing "Important Product Use Reminder" to remind users of proper procedures for using the system and the need for diligence in reducing air embo | Device Design | Medrad Inc., sent two types of "MEDICAL DEVICE FIELD CORRECTION" notifications on June 2, 2011 to all affected customers. 1. To those customers who Avanta Systems have already undergone market withdrawal and thus have been upgraded, the notification provided customer awareness that prior action should be considered a field correction. The letter stated that NO action is required since their system(s) has already been upgraded. If you have any questions, concerning these activities contact the Director of Quality Systems at (724) 940-6905. 2. To those customers whose Avanta Systems have not yet been upgraded through market withdrawal, the notification informed customers that the recalling firm is undertaking a Field Correction to upgrade their system(s). The letter included a step by step guide that will allow customers to verify that their system(s) has NOT been upgraded. Customers are instructed to complete and fax this form to MEDRAD at (412) 406-0941. If you have any questions, concerning these activities contact the Product Manager at (412) 767-2400, extension 8401. |  | 842 | 842 | Worldwide Distribution -- (USA) (nationwide) including the state of Puerto Ricoand the countries of: AE, AL, AT, BE, BG, BR, CA, CH, CO, CY, DE, DK, DO, ES, FI, FR, GB, HK, HR, HU, IE, IL, IS, IT, JO, JP, KR, KW, LY, MA, MT, MX, MY, NL, PA, PH, PK, PT, QA, RS, RU, SA, SE, SG, SI, TR, VI, and ZA. | Terminated | June 05, 2012 | 369 | 842 | Computer
Z-2867-2011 | 59077 | [6 empty] | Philips brand BrightView Gamma Camera System, |  |  | June 03, 2011 | 2011 | Philips Medical Systems | The gap between the two parts of the table may permit pinches to occur on either patients skin or other body parts; it has the potential for serious injury if a patients finger is caught between the pallet and patient table resulting in broken finger bone. | Device Design | Philips Healthcare issued a Customer Information letter (Field Change Order (#88200402)) dated June 15, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to follow the instruction for use for their system, which states "Monitor the patient and the system vigilantly during any system motions to make sure that ...accessories, equipment, and all ojbects such as hair, jewelry, or neckties-remain clear of any moving camera and imaging table parts. For additional information, customers were istructed to see the BrightView SPECT instructions for Use, section 2.2 - Safety Warnings and Precautions and Safety and Information Labels. Philips Healthcare will notify all affected customers and will install a sticker through an Action for Performance - Proactive Field Change Order (FCO) to resolve the issue. For customers in North America who require further information or support contact Customer Care Solutions Center at 1-800-722-9377, Select Option 5. In all other countries the local Philips Healthcare office should be contacted. |  |  |  | Worldwide Distribution - USA (nationwide) and Puerto Rico and the countries of Australia, Austria, Bahrain, Belgium, Brazil, China, Chile, Columbia, Czech Republic, Denmark, Dominican Republic, Egypt, France, Greece, India, Iran, Israel, Italy, Japan, Lithuania, Malaysia, Mexico, Netherlands, Peru, Portugal, Russia, Saudi Arabia, Singapore, South Korea, Spain, Switxerland, Taiwan, Thailand, Turkey, Ukraine, United Arab Emirates, and the United Kingdom. | Terminated | November 06, 2012 | 522 |  | Computer
Z-1722-2011 | 55875 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | The LIFEPAK 20 defibrillator/monitor and LIFEPAK 20e. Product part numbers: 3202487-xxx, 3202488-000, U3202487-000, 3202488-015, and U3202488-000. |  |  | May 27, 2010 | 2010 | Physio Control, Inc. | A component failure on the AC power supply assembly results in "NO AC POWER" operation.A failure of AC power can result in a delay of defibrillation therapy if backup battery power is depleted. | Device Design | PhysioControl, Inc. began sending an Urgent Medical Device Correction letter via certified receipt to all affected customers describing the potential issue. Customers were advised to keep the defibrillators in service and to follow the daily Operator's Checklist. All power supplies will be updated.For questions customers were instructed to contact Technical Support at 1-800-442-1142.For questions regarding this recall call 425-867-4000. |  | 53550 | 53550 units | Worldwide distribution - USA including AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MI, MD, MI, NC, NY, OH, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, and WY and the coutnries of ANGUILLA, ARGENTINA, AUSTRALIA, BAHAMAS, BAHRAIN, BARBADOS, BOLIVIA, BRAZIL, CAYMAN ISLANDS, CHILE, COLOMBIA, COSTA RICA, CYPRUS, DOMINICAN REPUBLIC, EGYPT, EL SALVADOR, GERMANY, GUAM, GUATEMALA, HONG KONG, INDIA, IRAN, ISRAEL, JAMAICA, JORDAN, KUWAIT, LEBANON, MEXICO, NETHERLANDS, NEW CALEDONIA, NEW ZEALAND, NICARAGUA, PAKISTAN, PANAMA, PARAGUAY, PERU, PUERTO RICO, SAUDI ARABIA, SOUTH AFRICA, SYRIA, TRINIDAD, TRINIDAD AND TOBAGO, TURKEY, UKRAINE, UNITED ARAB EMIRATES, URUGUAY, VENEZUELA, VIRGIN ISLANDS, and YEMEN. | Terminated | September 29, 2014 | 1586 | 53550 | Computer
Z-2854-2011 | 59189 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | VirtuoSaph Plus Endoscopic Vessel Harvesting System , REF VSP550, Sterile, Terumo Cardiovascular System |  |  | June 10, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems received four (4) complaints from a single customer about tunnel breathing/inadequate insufflation with the VirtuoSaph Plus Endoscopic Vessel Harvesting System during cases in a limited production launch situation. Teumo's investigation identified issues with the devices's capability to maintain pressure within the tunnel. | Device Design | Terumo Cardiovascular Systems contacted the one customer to this device on June 10, 2010 by telephone.There was one consignee. The firm contacted this one consignee by phone on 6/10/2010. There were 2 contacts at this one consignee. The recalling firm contacted both of the contacts on 6/10/2010. The consignee was asked to dispose of remaking inventory of devices. The consignee said had disposed of the product.For information on this recall call (734) 741-6173 |  | 35 | 35 units | Nationwide (USA) Distribution including the state of New York. | Terminated | July 27, 2011 | 47 | 35 | Computer
Z-1724-2011 | 57674 | No | JEQ | Powered toothbrush. | Dental | Dental | 510(K) Exempt | W Vital Health Pulsating Clean Rechargeable Toothbrush; Distributed by: Walgreen Co., 200 Wilmot Rd., Deerfield, IL 60015-4616, Made in China; |  |  | January 14, 2011 | 2011 | Brushpoint Innovations, Inc. | The power toothbrush may make a loud noise and the battery cap becomes separated from the main body of the toothbrush. The battery cap can become a projectile that could cause potential injury if it comes in contact with the body. | Device Design | The firm, BrushPoint Innovations, sent a email dated January 7, 2011 notifying their customers of the recall. The email described the product, problem and actions to be taken. The customers were instructed to pull and quarantine the toothbrushes due to potential product quality concerns, and hold their stocks until January 14, 2011, call in the number of toothbrushes quarantined, and then return them to the Walgreen distribution centers. A stop sale was placed on the affected item numbers in their cash register scanners.If you have any questions, call 905-833-5122 or 905-944-1000. |  | 8100 | 8,100 units | Nationwide distribution. | Terminated | December 16, 2012 | 702 | 8100 | Computer
Z-2858-2011 | 59051 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Runthrough NS Tapered PTCA Guidewire, item code number TW*DP418FH |  |  | June 03, 2011 | 2011 | Terumo Medical Corporation | Baloon catheter guidewires may become uncoiled during angioplasty procedures. | Device Design | Terumo Medical Corporation initiated contacting consignees on June 3, 2011 via an in-person visit to hand deliver an "URGENT PRODUCT BULLETIN: TERUMO PRODUCT SAFETY ALERT" letter dated June 3, 2011. The letter emphasized to consignees the importance of following the current instructions and recommendations provided to minimize the potential for damage during use and removal of the recalled product. For further questions, please contact your Terumo Territory Manager or call (800) 283-7866 and ask to speak with the Runthrough NS Guidewire Product Manager, at extension 4129. |  | 70 | 70 units | Nationwide (USA) Distribution -- including the states of: CT, GA, and SC. | Terminated | July 27, 2011 | 54 | 70 | Computer
Z-1734-2011 | 57232 | microtome, cryostat | IDP | Tissue processing equipment. | Pathology | Pathology | 510(K) Exempt | Thermo Scientific HM 550 Cold Disinfection Cryostat, Microm International GmbH, Waldorf, Germany |  |  | September 23, 2010 | 2010 | Thermo Fisher Scientific | Through internal quality review, firm determined the need to conduct a more comprehensive system validation to support the claim of complete bacterial, fungal and viral disinfection efficacy. | Device Design | ThermoFisher sent Advisory Notice letters beginning on September 28, 2010, to all affected customers. The letters stated that the firm needed to conduct more comprehensive system validation data on the Cold-Disinfection sytem option available on the HM550-D cryostat to support the claim of complete bacterial, fungal and viral disinfection efficacy. The letter recommended that the Cold Disinfection option on the HM550-D not be used as the sole disinfection method on the HM550-D until the validation work is completed. The firm provided addtional safety protocols pertaining to the cryostat disinfection. Customers were instructed to complete and fax the Confirmation of Receipt Form to Technical Support, COLD-D Advisory Notice at (269) 372-2436. |  | 100 | 100 units | Worldwide Distribution - USA including the states of LA, OH, FL, MD, PA, TX, MO, WI, KS, AL, CA, VA, MS, IL, NE, MA, and the countries of Australia, Belgium, China, Czech Republic, Denmark, Estonia, Finnland, France, Germany, Hong Kong, India, Italy, Japan, Netherlands, Norway, Poland, Portugal, Russia, Spain, Sweden, Taiwan, and the UK. | Terminated | April 02, 2012 | 557 | 100 | Computer
Z-2834-2011 | 59099 | No | EAZ | Dental operating light. | Dental | Dental | 510(K) Exempt | LuxStar Dental Light have the following Agency model or Catalog Model numbers as ML1000, OL1000, CL1000, WL1000, CB1000, TL1000, and UL1000. |  |  | May 27, 2011 | 2011 | Dental Equipment LLC | Lens heat shield (birdcage) of the Luxstar Dental Light could fall and injure a patient. | Device Design | Dental Equipment, LLC sent an "URGENT MEDICAL DEVICE RECALL" notification letter via USPS 1st class to US consignees and via Federal Express or Fax to foreign consignees on June 24, 2011. The letter informed customers about the product, problem and tether upgrade corrective action. A Field Correction Acknowledgement Form and the Field Correction Completion Form was attached for customers to complete and return via the fax number 503-537-3617. The letter states that customers are to review their records in order to identify where the affected devices are located and which are still in service. Customers can call the Customer Service at 800-304-5332 and order the self-install Upgrade Kits with the Part Number 30792 at no charge. Customers are instructed to complete the Rework Correction Form once all of the lights have been updated and fax or mail to Marcus, attention Bonnie Dundas, at (503) 537-3617. |  | 13226 | 13,226 units distributed worldwide | Worldwide Distribution -- USA (nationwide) including states of: CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NH, NJ, NM, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, WI, WV, and Puerto Rico and the countries of: Australia, Belgium, Canada, France, Germany, Ireland, Italy, Mexico, Norway, Poland, Russia, United Kingdom, Vietnam, and Trinidad. | Terminated | October 24, 2012 | 516 | 13226 | Computer
Z-2818-2011 | 59104 | No | GAP | Natural nonabsorbable silk surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Syneture Sofsilk 3/0 BK 12X18" PRE-CUT WAX-RACE |  |  | June 03, 2011 | 2011 | Covidien LP | Potential for the sterile barrier to be breached compromising the sterility of the product | Equipment maintenance | Covidien sent an URGENT MEDICAL DEVICE RECALL letter dated June 3, 2011 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately identify and quarantine all affected inventory. Distributors who further distributed the affected product were instructed to forward the letter to their customers who received the affected product. Customers were instructed to contact Covidien Customer Service at SDFeedback@Covidien.com or (800) 722- 8772, option 1, to obtain a Return Material Authorization (RMA) prior to returning the affected product. Customers were also asked to complete the attached product recall form in its entirety and fax the form to 203-822-6009 and include a copy with products being returned. If customers did not have any of the affected product they were to simply return the recall form via fax or email indicating they had zero units. For any questions regarding this recall customers were to contact their Covidien Representative or Covidien Quality Assurance at 203-492-5232. |  | 960 | 960 units | Worldwide Distribution - USA (nationwide) and the countries of Australia, Canada, Denmark, France, Germany, Italy, Japan, Korea, and Singapore | Terminated | February 10, 2014 | 983 | 93600 | Computer
Z-2725-2011 | 56016 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | IQecg IQdevices SDK, Part Numbers: 4-100-1403, Software Version 8.3.2 |  |  | May 18, 2010 | 2010 | Brentwood Medical Technology Corp | The recall was initiated because when a user of the IQmanager, IQdevices SDK, and IQiC (Centricity) software versions listed above accesses the QT dialogue box to either accept or edit the QT value, the QTc value is recalculated using a Hodges QTc equation, regardless of patient's age. The software does this without notifying the user of the change from use of the Bazett equation, which is utiliz | Software design | The firm, Midmark, sent an "URGENT: MEDICAL DEVICE FIELD CORRECTIVE ACTION" letter on May 19, 2010 with an attached Corrective Return Material Authorization (RMA) form via US Postal Service Certified Mail to its consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately examine their IQecg, IQmanager, or lQiC Centricity software to determine the software version currently in use. If it is a listed product, it is subject to this field corrective action. A sample software label was attached to aid the consignees/customers in locating and identifying the software CD. For customers who have already installed the software, the version number can be identified in the "About" page.Consignees/customers were instructed to complete the enclosed Field RMA form and return it with the affected software to:Quality Assurance DepartmentMidmark Diagnostics Group1125 West 190 S t.Gardena, CA 90248In addition, if they further distributed this product, identify customers and notify them at once of this product field corrective action.Replacement software was to be shipped to the consignees/customers upon receipt of the returned software CD. If they require a UPS return shipping label or an on-line download of replacement software, contact Midmark Diagnostics Group Support Services at 800-624-8950 extension 5145 or 5108. If you have any questions, call Midmark Diagnostics Group Support Services at 800-624-8950 extension 5145 or QA/RA Manager, Midmark Diagnostics Group 310-516-5116. |  | 8 | 8 units | Worldwide distribution: USA (Nationwide) and countries of: Algeria, Australia, Canada, France and Saudi Arabia. | Terminated | August 22, 2011 | 461 | 1621 | Computer
Z-1748-2011 | 58137 | No | NEZ | Ingestible telemetric gastrointestinal capsule imaging system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Given Imaging Limited, PillCam Express Delivery Device, For use with PillCam SB, endoscopic imaging device. Part numbers FGS-0367 (3-Pak) or FGS-0366 (1-Pak). |  |  | March 08, 2011 | 2011 | Given Imaging Ltd. | The capsule holder can break away from the catheter during endoscopic procedures, necessitating retrieval. | Device Design | Given Imaging, Inc. consignees were notified via FedEx letter on/about 03/08/2011. They instructed to return all affected product. A Customer Response form with return instructions was included to be completed and returned to Given Imaging. |  | 389 | 389 | U.S. Nationwide, Canada, Germany Israel and Australia. | Terminated | December 02, 2011 | 269 | 389 | Computer
Z-2789-2011 | 59055 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ONCOR/PRIMUS Linear Accelerator systems with RTT v2.x AND an ARTISTE Linear Accelerator/RT Therapist v4.1 system with the LANTIS OIS. |  |  | May 25, 2011 | 2011 | Siemens Medical Solutions USA, Inc | Potential safety issue of possible data loss for the transfer of the same patients treatment data between the LANTIS OIS and the RTT v2.x software and the RT Therapist version 4.1 system | Software design | Siemens Customers Service group began providing Update Instructions and installing the software update on affected machines beginning May 25, 2011. A customer information letter was also provided. |  | 82 | 45 customers with affected configuration in US, 37 customers ROW. | Worldwide Distribution - USA including distribution to LA, CA, PA, MA, NE, NJ, DE, WI, NC, NY, WI, MD, FL, KY and to Poland, Netherlands, Ireland, G=Belgium, Germany, New Zealand, P.R. China, Saudi Arabia, Australia, Malaysia, Egypt, Canada, and Spain. | Terminated | March 01, 2012 | 281 | 82 | Computer
Z-1796-2011 | 57177 | No | KWA | Hip joint metal/metal semi-constrained, with an uncemented acetabular component, prosthesis. | Orthopedic | Orthopedic | PMA | ASR 300 Size 46, DEPUY ASR 300 ACETABULAR CUP SYSTEM. |  |  | August 23, 2010 | 2010 | Depuy Orthopaedics, Inc. | As part of our ongoing post-market surveillance of all products, DePuy is continually evaluating data from a variety of sources including national joint replacement registries, published literature, company sponsored clinical trials, internal complaints data and unpublished clinical research reports.DePuy Orthopaedics issued a Field Safety Notice in March 2010 after receiving new data from the | Device Design | DePuy Orthopaedics issued a Field Safety Notice in March 2010 to all affected customers. The Notice identified the product, the problem, and the action to be taken by the customer. Customers were instructed to share the Field Safety Notice with their organization and any organization where the ASR products may have been transferred. Customer were instructed not to implant the ASR devices. Patients who received the ASR System should be informed of this recall and instructed to return for a follow up visit.Patients with radiographic changes indicative of product failure should be addressed according to normal procedures.All other patients should be followed according to the April 22, 2010, and May 25, 2010 UK Medicines and Healthcare products Regulatory Agency (MHRA) Device Alerts. Follow up for all patients implanted with ASR acetabular cups at least annually for five years postoperatively. Beyond five years, follow up in accordance with locally agreed protocols.For patients who are symptomatic or implanted with a cup angle greater than 45 degrees, particulary where a small component has been implanted.:" Consider measuring cobalt and chromium ion levels in whle blood and/or performing cross sectional imaging including MRI or ultrasound scans." If metal ion levels in whole blood are elevated above 7 parts per billion (ppb) for either metal ion, a second test should be performed three months after the first in order to identify patients who require closer surveillance, which may include cross sectional imaging" If MRI or ultrasound scan reveals soft tissue reactions, fluid collections or tissue masses, then revision surgery should be considered.DePuy representatives were to assist with returns of any remaining inventory.For questions regarding this recall call 574-372-7333. |  | 270 | 270 | Worldwide | Terminated | August 29, 2013 | 1102 | 120984 | Computer
Z-2780-2011 | 58969 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Eclipse Treatment Planning System; Release 6.5 through 10 |  |  | May 13, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | In affected Eclipse versions, user editing of the MLC Leaf positions may not automatically invalidate calculated dose distributions. | Software design | Varian Medical Systems, Inc sent a "Urgent Medical Device Correction Urgent Field Safety Notice" letter dated May 13, 2011 to all affected customers. ***12-12-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.***The letter included the description of the product, problem and recommendations and includes the release of a service pack in the future. For additional information please contact the Varian Oncology Help Desk at 888-VARIAN5 (888- 827-4265) |  | 10664 | 10,097***12-12-12 AMENDED TO: 10664 units*** | Worldwide distribution: Nationwide (USA) and the country of Canada. | Terminated | August 15, 2014 | 1190 | 10664 | Computer
Z-2987-2011 | 59367 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Direct Drive Grasper, Model number C4140, 5MM X 45CM DIRECT DRIVE |  |  | June 17, 2011 | 2011 | Applied Medical Resources Corp | The recall was initiated because Applied Medical has confirmed specific lot numbers of the DirectDriveÂż Grasper (model numbers C4130 and C4140) due to a potential weakness in the locking trigger, which could result in the trigger fracturing and the jaw locking in the closed position. | Device Design | Applied Medical sent an "URGENT: MEDICAL DEVICE RECALL" letter dated June 22, to all affected consignees. The letter described the product, problem and actions to be taken. Consignees were instructed to check their inventory for the recalled products and returned immediately as per the instructions. The letter instructs consignees to complete the attached Voluntary Recall Notification Confirmation Form and fax it to Regulatory Affairs at 949-713-8832 or email it to recall60342148@appliedmedical.com. For questions pertaining to product return contact Chris Jacobson, Sales Operations Administrator, at (949) 713-8652 or by email at cjacobson@appliedmedical.com. For regulatory questions contact Susan Fehrenbacher at (949) 713-8041 or by email at sfehrenbacker@appliedmedical.com. |  | 38864 | 38,864 units total (24,868 units in US) | Worldwide Distribution - USA (nationwide) including the state of Puerto Rico and the countries of: Australia, Bahrain, Canada, Chile, Colombia, Ecuador, Europe, India, Israel, Jordon, Kuwait, Lebanon, Libya, Malaysia, Mexico, Peru, Qatar, Singapore, , South Africa, South Korea, Saudi Arabia. Taiwan, Thailand, Turkey, and Venezuela. | Terminated | March 26, 2012 | 283 | 38864 | Computer
Z-2788-2011 | 59006 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Eclipse Treatment Planning System; |  |  | May 20, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | A software anomaly has been identified with the Eclipse Monte Carlo (eMC) algorithm where not-equidistant data points in the open beam profile used during configuration may lead to inaccurate results. | Software design | Varian sent an URGENT MEDICAL DEVICE CORRECTION letter via certified mail dated May 20, 2011, to all affected customers. ***12-20-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.*** The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to remove the product from use and advise users to manually initiate the configuration process for resampling to be applied when a previous calculation model is copied. Customers were asked to sign and return the notice to confirm receipt. For questions call USA and Canada 1-888-827-4265For Europe call +41 41 749 8844 |  | 10356 | 10,356 | Worldwide Distribution - USA (nationwide) | Terminated | July 13, 2012 | 420 | 10356 | Computer
Z-2781-2011 | 54621 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | DICOM Query/Retrieve, version 1.1.4. Individual part # 9733808, or as part of a system order, DICOM Q/R Software Kit, part # 9733933. DICOM Query/Retrieve software is installed on Medtronic StealthStation systems. |  |  | June 02, 2011 | 2011 | Medtronic Navigation, Inc. | Software allows incorrect patient or patient history to be displayed. | Software design | Medtronic Navigation sent a "FIELD SAFETY NOTICE" letter dated June 2, 2011 to all affected customers. The letter identified the product, problem and description of the corrective action. The letter states that Medtronic field representatives will visit each consumer/user and upgrade the version of DICOM Q/R to 1.1.5. Further information is available by contacting Medtronic Navigation Technical Services at 1-800-595-9709. |  | 90 | 90 | Worldwide distribution - USA (nationwide) -- including states of: TX, NC, ID, FL, OH, UT, CA, TN, IL, AZ, NE, LA, PA, OR, AR, MI, NJ, MD and countries of: Canada, Denmark, Finland, Great Britain, Hungary, Ireland, Japan, Korea, Netherlands, Poland, South Africa, Spain, and Taiwan. | Terminated | July 12, 2011 | 40 | 90 | Computer
Z-2767-2011 | 58981 | No | LTW | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | VITEK 2 System, VITEK 2 60 110V (Catalog #27225) and VITEK 2 60 220V (Catalog #27226). |  |  | June 02, 2011 | 2011 | Biomerieux Inc | Cards may not load into the incubator correctly causing them to be loaded in the wrong order which could result in the results being linked to the wrong accession number and reporting results erroneously. | Device Design | Biomerieux, Inc. sent an "URGENT PRODUCT CORRECTION NOTICE" dated June 2, 2011 to all affected customers. The letter describes the products, problem, and actions to be taken by the customers.The letter explains how to recognize the issue and instructs customers to take certain steps to resolve the problem. Additionally, an Acknowledgement Form is attached to the letter for customers to complete and return via fax to 919-314-4539. A firmware update will be applied by Biomerieux, Inc. technical support representatives.Questions regarding this recall are directed to the Customer Service Department at 1-800-682-2666, option 3. |  | 627 | 627 instruments | Worldwide Distribution-- USA (nationwide) and the countries of Austria, Australia, Belgium, Brazil, Canada, China, Colombia, Germany, France, India, Israel, Italy, Japan, Korea, Mexico, The Netherlands, Poland, Portugal, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Thailand, United Arab Emirate, and United Kingdom. | Terminated | June 06, 2013 | 735 | 1734 | Computer
Z-2766-2011 | 58994 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | TrueBeam Linear Accelerators (aka Trilogy Mx) |  |  | May 19, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | Imaging arms of the TrueBeam Accelerator may have loose encoder pulleys that could lead to inaccurate readout of arm geometry. | Device Design | Varian Medical Systems, Inc. Oncology Systems sent an "URGENT MEDICAL DEVICE CORRECTION" letters on May 19, 2011, via certified mail, return receipt requested in the US. The letter identified the product, problem and recommended actions to be taken. For additional information contact Varian Customer Support at 1-888- 827-4265. |  | 77 | 77 | Worldwide Distribution -- USA (nationwide) and countries of: Canada, Denmark, France, Germany, Isreal, Italy, Netherlands, Russia, Saudi Arabia, Spain, Sweden, Switzerland, Turkey, and United Kingdom. | Terminated | March 13, 2012 | 299 | 77 | Computer
Z-2765-2011 | 58566 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Philips Diagnost 94, Diagnost 96, Diagnost 97 Universal X-Ray System |  |  | April 18, 2011 | 2011 | Philips Healthcare Inc. | Updated Instructions For Use: During an examination where the patient is positioned on a -rotatable- chair with knees pointed to the rear of the system, the patients lower legs or ankles may be injured/broken | Device Design | Philips Healthcare notified accounts by letter on April 18, 2011. The letter advised users of the problem and provided updated instructions for proper placement of the patient during the examination. An Updated Instruction For Use was provided.For further information and support customers were instructed to contact the Philips representative. |  | 52 | 52 units | Nationwide Distribution | Terminated | February 15, 2013 | 669 | 52 | Computer
Z-2762-2011 | 58830 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo Dynamics Picture Archiving Communication System |  |  | May 10, 2011 | 2011 | Siemens Medical Solutions, USA, Inc | For systems upgraded to syngo Dynamics version 9.0, if a legacy report generated by version 3.x or earlier is opened in the version 9.0 portal, the legacy report may be deleted and replaced with a new report from the version 9.0 reporting system or a blank report. | Software design | Siemens Medical Solutions, USA, Inc sent a letter entitled "Customer Safety Advisory Notice" to all direct accounts on May 10, 2011. The letter stated the name and version of the software, problem noted, a work-around to be implemented until a service patch is created to fix the problem, and that the letter should be placed in the software's operator's manual to be seen by all who utilize the software. |  | 38 | 38 | Nationwide Distribution | Terminated | July 09, 2012 | 426 | 38 | Computer
Z-2763-2011 | 58974 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARIA, versions 8.0 and 8.1; |  |  | May 17, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | An issue has been detected when using the RT Chart function in ARIA versions 8.0 and 8.1. When users select the "Convert to Setup Field" option, a Set-up field is created, that removes the DMLC plan components in order to create an open collimator field for setup purposes. | Software design | Varian Medical Systems, Inc. sent an "URGENT MEDICAL DEVICE CORRECTION" letters on May 17, 2011, via Certified Mail to US; Federal Express to Canada. Foreign distribution will be handled by Varian Managers. The letter contained recommended actions. For additional information contact Varian Medical Systems, Inc at 888- 827-4265.***12-20-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.*** |  | 357 | 357 | Worldwide Distribution -- USA (nationwide) and countries of Algeria, Australia, Austria, Belgium, Brazil, Chile, China, Canada, Colombia, Denmark, Ecuador, Egypt, Finland, France, Germany, Hong Kong, India, Iraq, Ireland, Italy, Japan Jordan, Kuwait, Malaysia, Morocco, Netherlands, Oman, Pakistan, Panama, Portugal, Romania, Russia, South Korea, Spain, Sweden, Switzerland, Syria, Taiwan, Thailand, Tunsia, Turkey, Ukraine, United Arab Emirates, and United Kingdom, | Terminated | March 15, 2012 | 303 | 357 | Computer
Z-1834-2011 | 54706 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Coherence Dosimetrist, version 2.2. |  |  | November 06, 2009 | 2009 | Siemens Medical Solutions USA, Inc | A problem in the software has the potential to be associated with the dose; when a template is created the user has the possibility to predefine organs. For the Target the user defines the max. dose and the min dose, for the OAR (Organs at Risk) the user only defines the max. dose. The firms engineering team determined that the software defect was repeatable. | Software design | Affected consignees will be notified by the Customer Service Engineer to schedule a software update beginning November 2009. |  | 60 | 60 units | Worldwide distribution: USA including: CA, KY, MI, NJ and TN; and countries of: Canada and Mexico. | Terminated | April 06, 2011 | 516 | 60 | Computer
Z-1832-2011 | 54842 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Acuson S2000 Ultrasound System with software and imaging. |  |  | March 01, 2010 | 2010 | Siemens Medical Solutions USA, Inc. | "Any 2D measurement performed in Exam Review or offline, on a zoomed Retrospective clip capture may be erroneous. This applies to 2D, 2D/PW, 2D/CW and 2D/M-mode." | Software design | Siemens Medical Solutions USA, Inc. sent a Consumer Safety Advisory Notification letter on February 16, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to pass the Customer Safety Notice to all those within their organization who need to be aware of the issue until the corrective action was completed. For questions regarding this recall call 650-694-5580. |  | 1903 | 1903 | Worldwide Distribution - USA (nationwide) and the countries of United Arab Emirates, Australia, Belgium, Bolivia, Brazil, Bahamas, Canada, Switzerland, Chile, China, Colombia, Czech Republic, Germany, Denmark, Ecuador, Eguypt, Spain, Finald, France, Great Britian, Germany, Hong Kong, Croatia, Indonesia, Ireland, Israel, India, Ireland, Italy, Jordan, Japan, Kenya, South Korea, Kuwait, Lithuania, Mexico, Malaysia, Netherlands, Norway, Philippines, Poland, Portugal, Qatar, Romania, Russian Federation, Saudi Arabia, Sweden, Slovenia, Slovak Republic, Thailand, Turkmenistan, Tunisia, Turkey, Taiwan, Ukraine, Venezeula, Yemen, and South Africa.. | Terminated | June 18, 2012 | 840 | 1903 | Computer
Z-1837-2011 | 56840 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | ACUSON S2000 ultrasound systems (all versions) with cardiovascular options installed; |  |  | August 17, 2010 | 2010 | Siemens Medical Solutions USA, Inc., Mountain View, CA | There is an error in the estimation of mean pressure gradient in the Cardiac measurement package. The currently implemented calculation will result in an underestimation of the mean pressure gradient (PGmean) when estimated from a trace of the Doppler waveform of the stenotic flow. | Software change control | Siemens Medical Solutions USA, Inc., sent a Customer Safety Advisory Notification letter on August 17, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to pass the notice on to all those within their organization who need to aware of the issue until the corrective action was completed. For questions customers were to contact their local service support person. For questions regarding this recall call 650-694-5993. |  |  |  | Worldwide Distribution - USA (nationwide)and the countries of Antares distribution includes: United Arab Emirates-AE, Austria-AT, Australia-AU, Azerbaijan-AZ, Brazil-BR, Belarus-BY, Canada-CA, Switzerland-CH, Chile-CL, China-CN, Czech Republic-CZ, Germany-DE, Denmark-DK, Egypt-EG, Spain-ES, France-FR, Great Britain-GB, Guatemala-GT, Croatia-HR, Hungary-HU, Indonesia-ID, Ireland-IE, India-IN, Italy-IT, Jordan-JO, Japan-JP, South Korea-KR, Moldova-MD, Mexico-MX, Malaysia-MY, Netherlands NL, Norway-NO, New Zealand-NZ, Poland-PL, Portugal-PT, Romania-RO, Serbia-RS, Russian Federation-RU, Saudi Arabia-SA, Sweden-SE, Singapore-SG, Slovak Republic-SK, Suriname-SR, Turkey-TR, Taiwan-TW, Ukraine-UA, United States-US, Uruguay-UY, Uzbekistan-UZ, Venezuela-VE, Vietnam-VN, South Africa-ZA S2000 distribution includes: Austria-AT, Australia-AU, Brazil-BR, Canada-CA, Switzerland-CH, Chile-CL, China-CN, Czech Republic-CZ, Germany-DE, Denmark-DK, Spain-ES, Finland-FI, France-FR, Great Britain-GB, Hong Kong-HK, Indonesia-ID, India-IN, Italy-IT, Japan-JP, South Korea-KR, Kazakhstan-KZ, Lebanon-LB, Mexico-MX, Norway-NO, Philippines-PH, Poland-PL, Portugal-PT, Romania-RO, Russian Federation-RU, Sweden-SE, Slovak Republic-SK, Turkmenistan-TM, Turkey-TR, Taiwan-TW, United States-US X300 distribution includes: AE (United Arab Emirates) , AL (Albania) , AN (Netherlands Antilles), AO (Angola), AR (Argentina), AT (Austria), AU (Australia), AZ (Azerbaijan), BA (Bosnia-Herzegovina), BD (Bangladesh), BE (Belgium), BG (Bulgaria), BN (Brunei), BO (Bolivia), BR (Brazil), BY (Belarus), CA (Canada), CH (Switzerland), CL (Chile), CN (China), CO (Colombia), CV (Cape Verde), CZ (Czech Republic), DE (Germany), DK (Denmark), DZ (Algeria), EC (Ecuador), EG (Egypt), ES (Spain), ET (Ethiopia), FJ (Fiji), FR (France), GB (United Kingdom), GE (Georgia), GR (Greece), HK (Hong Kong), HN (Honduras), HR (Croatia), HU (Hungary), ID (Indonesia), IE (Ireland), IN (India), IR (Iran), IT (Italy), JO (Jordan), JP (Japan), KE (Kenya), KR (Korea, South), KW (Kuwait), KZ (Kazakhstan), LT (Lithuania), LV (Latvia), LY (Libya), MD (Moldova), ME (Montenegro), MK (Macedonia), MM (Myanmar), MU (Mauritius), MX (Mexico), MY (Malaysia), NG (Nigeria), NI (Nicaragua), NL (Netherlands), NO (Norway), NP (Nepal), NZ (New Zealand), PE (PerÂż), PH (Philippines), PK (Pakistan), PL (Poland), PS (Palestine), PT (Portugal), QA (Qatar), RS (Serbia), RO (Romania), RU (Russia), RW (Rwanda), SA (Saudi Arabia), SC (Seychelles), SE (Sweden), SG (Singapore), SI (Slovenia), SK (Slovakia), SV (El Salvador), SY (Syria), TH (Thailand), TM (Turkmenistan), TN (Tunisia), TR (Turkey), TT (Trinidad/Tobago), TW (Taiwan, ROC), UA (Ukraine), US (United States of America), UY (Uruguay), UZ (Uzbekistan), VE (Venezuela), VN (Viet Nam), YE (Yemen), ZA and (South Africa) | Terminated | June 18, 2012 | 671 |  | Computer
Z-1831-2011 | 57887 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | The enGenTM Laboratory Automation System consists of 3 main components including Instrument Management Subsystem (Instrument Manager), Automated Track Subsystem (TCAutomation) and Configuration File developed by Ortho-Clinical Diagnostics (OCD). The Instrument Manager is a Windows-based software product that operates on a Pentium-class (or higher) PC processor. The instrument management subsystem is capable of connectivity with clinical analyzers, a Laboratory Information Systems (LIS), barcode label printers and other third party devices. The TCAutomation is a pre-analytical and post-analytical sample management system. The configuration file enables operational support of the site specific layout and instrument selection. |  |  | January 27, 2011 | 2011 | Ortho-Clinical Diagnostics | A software anomaly that can occur with analyzers connected to the Bypass modules listed below on an enGen" Laboratory Automation System using TCAutomation" Software Version 3.1.1 and below. When this anomaly occurs, it is possible for a sample other than the intended sample to be aspirated, potentially leading to test results being assigned to the wrong patient without error notification. | Software design | The firm, Ortho Clinical Diagnostics (OCD), sent an "URGENT PRODUCT CORRECTION NOTIFICATION" letter dated January 27, 2011, by Federal Express overnight mail to all customers. Foreign affiliate consignees/customers were notified by email informing them of the issue on the same day. The letter described the problem, product and actions to be taken. The firm also included temporary procedures to help prevent the occurrence of the software anomaly. The customers were instructed to follow the enclosed Procedure to STOP the Automation System and/or Bypass Module; complete and return the attached Confirmation of Receipt - Important Response Required form by February 4, 2011, via fax to: Ortho Clinical Diagnostics, 1-888-557-3759 or 1-585-453-4110, and display this communication near the enGen Laboratory Automation System to ensure compliance by all operators until updated TCAutomation Laboratory Automation System software is installed. A Technical Bulletin will be sent to customers in the near future to be stored with user documentation. Ortho will implement a TCAutomation Laboratory Automation System software change in the near future to resolve this anomaly.If you have any questions regarding this notification, please call Customer Technical Services at 1-800-421-3311 (Options 1, 2). |  | 63 | 63 units | Worldwide distribution: USA and countries including: Canada, Chile, England, France, Germany, Italy, Mexico, Singapore and Spain. | Terminated | May 09, 2017 | 2294 | 63 | Computer
Z-2740-2011 | 58792 | No | OOI | Instrumentation for clinical multiplex test systems. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | NucliSens EasyQ Enterovirus serotype 68 v1.1: Catalog number: 200300 |  |  | May 10, 2011 | 2011 | bioMerieux, Inc. | The NucliSENS EasyQ Enterovirus Reagents are unable to detect Enterovirus serotype 68, in respiratory samples. | Device Design | BioMerieux Vitek, Inc. sent an "Urgent Product Correction Notice" letter on/about May 10, 2011.The letter included; a description of the product, problem, possible impact to the patient and recommendations. Recommendations included instructing customers to discontinue the use of the kit for the detection of Enterovirus 68 and Rhinovirus 87 serotypes and noted the corrected IFU will be available the first week of April 2011. An Acknowledgement form was attached to be returned via fax.For additional questions please contact your local BioMerieux Clinical Customer representative at (800) 682-2666, option 3, option 2. |  | 42 | 42 kits | (USA) Nationwide distribution including the states of MD, MI, MN, MT, NY and PA | Terminated | October 06, 2011 | 149 | 42 | Computer
Z-2738-2011 | 58547 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Large Roller Pump 6" Diameter, Catalog 801041, Terumo Cardiovascular Systems, Corp., Ann Arbor, MI |  |  | March 31, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | The firm has received multiple reports of the lower roller pump tube clamp mechanism failing to function in a small population of devices, making it difficult to remove or insert the tubing. If the failure occurs, the clamp dial will no longer be able to control opening/closing of the tube clamp mechanism. | Device Design | Terumo Cardiovascular Systems sent an URGENT MEDICAL DEVICE CORRECTION letter to all affected customers dated March 21, 2011. The letter identified the product, the problem and the action needed to be taken by the customer. An additional URGENT MEDICAL DEVICE RECALL correction notification and cover letter was sent to all consignees on XX/XX/XX to explain the reason for the expansion of the recall and will provide an update to the timing of implementation of the field correction.Expansion now includes devices manufactured between 2006 and March 2012.The correction strategy is to issue the customers a letter and addendum to the Operator's Manual that specifies what the user should do in the event of a tube clamp failure for any reason. The letter states that the firm is still investigating the issue and the correction. The letter states that in the coming months when a resolution is identified, consignees may be contacted to make arrangements for a upgrade to the system. There is no specific timeline for the correction to be made.If you have any further questions, please call a Terumo CVS Customer Service at 1-800-521-2818. Customer Service hours Monday - Friday, 8 Am - 6 Pm EST. |  | 111 | 111 devices. | Worldwide Distribution -- USA (nationwide) and the countries of Argentina, Australia, Belgium, Canada, Chile, Columbia, Costa Rica, Dominican Republic, Egypt, Germany, Guatemala, Honduras, Hong Kong, Indonesia, Japan, Jordan, Korea, Kuwait, Malaysia, Mexico, Pakistan, Republic of Georgia, Russia, Saudi Arabia, Singapore, South Africa, Taiwan, Thailand, The Phillipines, Turkey, United Arab Emirates and Vietnam. | Terminated | August 28, 2013 | 881 | 3679 | Computer
Z-2723-2011 | 58536 | No | HXX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Pathfinder NXT Closure top driver - short, REF 3566-2, Zimmer Inc., Warsaw, IN |  |  | May 20, 2011 | 2011 | Zimmer Inc. | Zimmer Spine, Inc. is initiating a voluntary recall regarding an increased risk of rounding or partial rounding of the hex tip on the 3566-1 (long) and 3566-2 (short) Closure Top Driver used for insertion, provisional tightening, and final locking of Closure Tops during the screw and rod construct assembly of the Pathfinder NXT Pedicle Screw System.When the hex tip of the Closure Top Driver is | Device Design | Zimmer, Inc. issued an "URGENT MEDICAL DEVICE RECALL" letters dated April 2011, to all affected customers. The letter identified the product, problem and actions to be taken. The customers were instructed that (THE GOLD CLOSURE TOP DRIVERS 3566-1 and 3566-2 ARE NOT TO BE USED FOR FINAL TIGHTENING); an alternate instrument for final construct tightening will be included in the instrument set 3570-1 short and 3570-2 long (These instruments are silver instead of gold for easy identification and do not include a retention feature for holding the Closure Top); to ensure final drivers are added to each PathFinder NXT set and inspect Closure Top Drivers for rounded hex wear; contact their customer service representative if they do not have both versions; and complete and return all attached forms via fax to: Zimmer Spine at 512-258-0995. The letter states that the firm will revise the surgical technique to show the updated uses of the screwdrivers. Please keep Zimmer Spine informed of any adverse events associated with the device or any other Zimmer spine product by calling 800-326-0635. For any further questions or concerns contact the Vice President Quality Assurance &amp; Regulatory Affairs at 800-655-2614. |  | 135 | 135 devices | Worldwide Distribution -- USA (nationwide) including states of: AZ, CA, WI, MI, KS, OK, TX, GA, TN, FL, NJ, CT, PA and the countries of: Australia and Italy. | Terminated | October 07, 2011 | 140 | 269 | Computer
Z-2879-2011 | 59146 | No | HWE | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | Orthopedic | 510(K) Exempt | Arthrex Reciprocating Saw Attachment |  |  | May 24, 2011 | 2011 | W &amp; H DentalWerk Buermoos GMBH | Reciprocating saw attachment V 600 Sr (REF: 17306210), screws can become loose or come off (the attachment) and fall into the surgical wound and may not be noticed due to its small size. | Equipment maintenance | Madrimed GmbH (Distributor partner) of W&H; Dentalmechanik GmbH sent a letter dated May 24, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were ask to advise their users to check the adapters thoroughly before and after each application, until the conversion of all attachments have been completed and to send all products that need to be converted to the nearest service centers.W&H; decided to convert all products that might be affected in their service centres and equip them with an additional security device.If you have any questions, call +43627462360 or email: regulatory@wh.com. |  | 194 | 194 units | Worldwide distribution: USA (nationwide) and countries of: Austia, Australia, Canada, Czech Republic, Denmark, France, Germany, Italy, Kuwait, Malaysia, Netherlands, Russia, Slovakia, South Africa, Spain, Taiwan, Turkey and United Kingdom. | Terminated | October 09, 2015 | 1599 | 194 | Computer
Z-1704-2011 | 55617 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Stryker Stiletto Electrosurgical Probe (E-Probe); Model number: 0270-070-550, 0270-070-551, 0270-070-552, 0270-070-553, 0270-070-554, 0270-070-555; Manufactured by Stryker Endoscopy Puerto Rico. |  |  | February 25, 2010 | 2010 | Stryker Endoscopy | The heat shrink, PTFE E-Probe Tip were out of specification (OOS) in the longitudinal change (+5% -20%). | Device Design | All Sales reps were notified of the recall on February 25, 2010. Urgent Device Removal advisory notices were also sent to all consignees.Device Removal letters identified the specific part numbers and lot numbers involved. Letters also list the potential risk to patients as a result of the issue. Customers are asked to inspect their inventory for part and lot numbers and quarantine any affected product. If recalled product is found, customers are to contact Stryker Customer Service at 800-624-4422 and request an RMS to send back recalled units. New units will be provided at no charge. The acknowledgement form is to be completed and returned. Questions or concerns should be directed to 408-754-2000 or e-mail at eprobe@stryker.com. |  | 31014 | 31, 014 units | Worldwide Distribution -- US, Australia, Canada, Germany, Greece, Iberia, Korea, Switzerland, and Newbury. | Terminated | March 18, 2011 | 386 | 31014 | Computer
Z-2888-2011 | 59023 | No | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | WarmAir Model 135 Convective Warming Device, |  |  | June 01, 2011 | 2011 | Cincinnati Sub-Zero Products Inc | Due to its design the circuit breaker on the device could trip causing it to shut down and not provide heated air to the blanket. | Device Design | Cincinnati Sub-Zero Products, Inc. sent an 'URGENT Medical Device Recall Field Action Notice" dated July 12, 2011 to all affected customers.The letter includes: description of product, problem, and information to request a repair tubing kit free of charge. It also asks customers to complete and return a "Recall Response Form".For additional information e-mail WA135-2011@cszinc.com or call (800) 989-7373 or (513) 772-8810. |  | 5474 | 5474 units | Worldwide Distribution: (USA) Nationwide, and the countries of Australia, Brazil, Bulgaria, Canada, Chile, Ecuador, Germany, Hong Kong, India, Indonesia, Ireland, Israel, Italy, Japan, Malaysia, Mexico, New Zealand, Oman, Pakistan, Peru, Philippines, Poland, Portugal, Puerto Rico, Qatar, Saudi Arabia, Singapore, South Korea, Spain, Sri Lanka, Taiwan, Thailand, The Netherlands, Turkey, U.A.E., UK, Uruguay, Venezuela, and Vietnam. | Terminated | August 02, 2013 | 793 | 5474 | Computer
Z-2890-2011 | 59036 | No | MDB | Microbial growth monitor. | Microbiology | Microbiology | 510(k) | BD BACTEC(tm) MGI (tm) System Software, Catalog 445922, Lot V5.02A. |  |  | March 28, 2011 | 2011 | Becton Dickinson &amp; Co. | Software contained three additional set configurations which caused the system to reinitialized the AST set carrier defaults. | Software design (manufacturing process) | BD Diagnostic Systems verbally notified all affect customers via the telephone on March 28, 2011. The notification identified the product, the problem, and the action to be taken by the customer. An "URGENT PRODUCT RECALL" follow-up letter was mailed to all customers on April 8, 2011 via UPS. Customers were instructed to complete and fax the attached reply form to acknowledge their receipt of the notification to Regulatory Compliance, 410-316-4256. The letter instructs customers to refer to the modified release notes and IMPORTANT NOTICES outlined in the attached notification letter. If further assistance is needed regarding replacements and all other inquires, please contact BD Technical Services and Support at 1-800-638-8663. |  | 421 | 421 disks | Nationwide Distribution (USA) including states of: AK, AL, AR, CA, CO, CT, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, and WI. | Terminated | July 29, 2011 | 123 | 421 | Computer
Z-2983-2011 | 59347 | No | LMB | Medical image storage device. | Radiology | Radiology | 510(K) Exempt | LANTIS Oncology Information System (OIS) |  |  | June 30, 2011 | 2011 | Siemens Medical Solutions USA, Inc | A potential safety issue exists for data transfer between LANTIS OIS system (containing PRIMEVIEW or PRIMEVIEW Data Gateway) and the Sequencer V&R; system via a third party interface is used. | Software design | Siemens sent an "Update Instructions/Customer Safety Advisory Notice" to all affected users, beginning June 30, 2011.The letter includes description of the product and problem and provides updated instructions to ensure data is stored in the system.If you have questions contact Siemens at (925) 602-8157. |  | 59 | 59 devices | Worldwide Distribution: Nationwide (USA) including the states of MI, DE, WI, FL, CA, TN, NJ, PA, NE, IL, WV , MA and VA; and the countries of Poland, France, Germany, Canada, Netherlands, India, Ireland, Egypt, P.R. china, Australia, Saudi Arabia, Turkey, Republic Korea, Belgium, Hungary, New Zealand, Italy, Malaysia, and Spain. | Terminated | March 09, 2012 | 253 | 59 | Computer
Z-2993-2011 | 59401 | No | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | GammaMed Plastic Needle with Mandrin; 2.0mm diameter, length 320 mm. |  |  | July 13, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | Needle tip may separate from the needle shaft such that the tip may detach in the patient upon withdrawal. | Device Design | Varian sent an "Urgent Medical Device Correction/Urgent Field Safety Notice" letter dated July 13, 2011 to all possibly affected customers. This notification supplements and modifies the previous "Customer Technical Bulletins" originally sent in September 2008. The initial Customer Technical Bulletins informed customers of the possibility that needle tips cans crack and separate. It also reminded users that the product has a limited life and to check needles prior to use as per instructions. Current correction notice letter instruct customers to stop using any plastic needles in inventory with lot numbers C01 to D01 and those with lot numbers starting with H or I. The letter also request that the recalled products be returned to Varian BrachyTherapy for replacement. The letters include information as to where the lot number is located on the device. All appropriate personnel should be notified of the correction notice. Contact information is provided for all geographical areas. USA contact for questions or concerns is 1-800-360-7909. |  | 601 | 601 - all sizes. | Worldwide Distribution - USA (nationwide) and the countries of: Argentina, Austria, Bangladesh, Belgium, Brazil, Canada, Columbia, Cyprus, Czech Republic Denmark, Dominican Republic, Estonia, France, Germany, Greece, Hungary, India, Japan, Kazahastan, Kuwait, Lebanon, Macedonia, Mexico, Moldova, Nepal New Zealand, Norway, Pakistan, Philippines, Poland, Russia, Serbia, Singapore, Slovenia, South Africa, South Korea, Spain Sweden, Switzerland, Taiwan, Thailand, Trinidad, Tobago, Turkey, Ukraine, United Arab Emirates Uruguay, United Kingdom, and Venezuela. | Terminated | June 20, 2012 | 343 | 601 | Computer
Z-1651-2011 | 57534 | No | OOI | Instrumentation for clinical multiplex test systems. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | MagNA Pure LC 2.0 Instrument, catalog number 05197686001, Roche Diagnostics, Indianapolis, IN |  |  | October 26, 2010 | 2010 | Roche Diagnostics Operations, Inc. | If the bar code is printed from the Batch Results Screen, a wrong batch ID may be printed, while the correct batch ID is still assigned to the batch in the data repository of the MPLC 2.0 software. This may result in sample mis identification. If printed from the ordering screens, the bar code is printed with the correct batch ID. | Software design | The firm, Roche, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated November 04, 2010 to its customers. The letter described the product, problem and action to be taken. The customers were instructed to print batch ID barcode labels from the ordering or stage set up screens only; print Cooling Block barcode labels from the Post Elution Edit screen only; complete and return the attached UMDC Faxback Form for MagNA Pure LC 2.0 Instrument [UMDC 10-228] via fax to 1-877-551-7152 and file this correction notice for future reference. The firm plans to issue a software update in first quarter 2011. If you have questions about the information contained in this UMDC, for Clinical customers, please contact Roche Molecular Diagnostics Technical Service at 1-800-526-1247; for Research customers, please contact Roche Applied Sciences Technical Service at 1-800-262-4911. |  | 119 | 119 devices | Worldwide distribution: AL, AZ, CA, CO, DC, FL, GA, HI, IL, KY, LA, MA, MD, ME, MN, MO, MT, NC, ND NE, NH, NJ, NM, NY, OH, OK, OR, SC, SD, TN, TX, UT, VA, VT, WI and Puerto Rico; and country of Europe. | Terminated | August 30, 2011 | 308 | 119 | Computer
Z-2994-2011 | 58744 | No | LNR |  |  | Gastroenterology/Urology | PMA | XT125 PHOTOPHERESIS PROCEDURAL KIT FOR USE WITH THE UVAR XTS INSTRUMENT; XT001 PHOTOPHERESIS PROCEDURAL KIT FOR USE WITH THE UVAR XTS INSTRUMENT |  |  | March 31, 2011 | 2011 | Therakos, Inc. | Increased complaints were received of tubing leaks at the junction between the tubing and Fluid Logic Module plastic ports of the kits. | Device Design | The firm, Therakos Photopheresis, sent a "URGENT PRODUCT CORRECTION NOTIFICATION" letter dated March 31, 2011 to its customers. The letter described the product, problem, and actions to be taken. The customers were instructed to immediately discontinue using the affected products; deface and discard their remaining inventory of affected procedural kits, following their internal procedures for destroying and discarding of medical waste, and complete and return the attached Confirmation of Receipt Form no later than 08-April-2011 via fax to: Safety, Risk Management and Surveillance at 1-888-557-3759 or 1-585-453-4110. If you have any additional technical questions, please contact Therakos Customer Service Center at 1-800-828-6316. If you have technical questions regarding this communication, please call Therakos Technical Support at 1-877-865-6850. |  | 13239 | 13239 kits | Worldwide distribution: USA (nationwide) including states of: AK, AL, AR, AZ, CA, CO, DC, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, NC, NE, NH, NJ, NM, NY, OH, OK, OR, PA, SC, SD,TN, TX, UT, VA, VT, WA, and WI; and countries of: Argentina, Austria, Australia, Belgium, Canada, Denmark, Estonia, Finland, France, Germany, Greece, Hungary, Ireland, Israel, Italy, Korea, Kingdom of Saudi Arabia, Jordan, Netherlands, Norway, Philippines, Poland, Portugal, Russia, Singapore, Spain, Sweden, Switzerland, Turkey and United Kingdom. | Terminated | March 18, 2015 | 1448 | 13239 | Computer
Z-2978-2011 | 59087 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Clinac, Trilogy. Trilogy Tx and Novalis Tx Linear Accelerators. |  |  | May 26, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | On Varian Linear Accelerators where the customer may calibrate the collimator angle position readout in reverse, switching the 90 degree and 270 degree positions. | Device Design | Varian Medical Systems sent an "Urgent Medical Device Correction-Urgent Field Safety Notice" letter dated May 26, 2011 to all affected customers.***12-20-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.*** The letters identifies the product, describes the problem and provides recommendations. Varian recommends maintenance only be performed by qualified personnel with appropriate training, always verify collimator calibration and anytime a collimator posthion readout calibration is performed, and always verify that the collimator angle is not calibrated in reverse by performing the appropriate checks. For questions contact Varian at (888) 827-4265. |  | 5685 | 5685 | Worldwide Distribution: Nationwide (USA) Distribution and the countries of; Algeria, Angola, Argentina, Australia, Austria, Azerbaijan, Bahrain, Bangladesh, Belgium, Belorussia, Bolivia, Bosnia and Herzegoniva, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Cyprus, Czech Rep., Denmark, D.R. Ecuador, Egypt, Estonia, Finland, France, French Poly., Georgia, Germany, Greece, Guam, Guatemala, H.K., Honduras, Hungary, Iceland, India, Indonesia, Iran, Iraq, Ireland, Israel, Italy, Japan, Kazakhstan, Kenya, Kuwait, Latvia, Lebanon, Lithuania, Luxembourg, Macau, Macedonia, Malaysia, Martinique, Mexico, Moldova, Monte Negro, Morocco, Nepal, Netherlands, N.Z., Norway, Oman, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Portugal, P.R., Qatar, Reunion, Romania, Russia, S.A.,Serbia, Singapore, Slovakia, South Africa, S.K., Spain, Sri Lanka, Sweden, Switzerland, Syria, Taiwan, Thailand, Trinidad/Tobago, Tunisia, Turkey, Ukrain, U.A.E., U.K., Uruguay, Venezuela, Vietnam, V.I., and Zimbabwe. | Terminated | July 13, 2012 | 414 | 5685 | Computer
Z-2980-2011 | 58918 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Symbia Radiation Therapy Pallet. |  |  | June 22, 2011 | 2011 | Siemens Medical Solutions USA, Inc. | There a potential patient safety issue when using the Symbia Radiation Therapy Pallet. When a patient is not iso-centered before starting a scan with the Symbia T Series Systems and a non-circular orbit is selected, it is possible that one of the detectors may come into contact with the patient. | Device Design | Siemens Medical Solutions USA, Inc. sent an "Urgent Field Correction Recall" letter dated June 06, 2011 to all affected customers.The letter informs customers of the potential safety issue if the patient is not iso-centered before starting a scan and a non-circular orbit is selected and provided recommendations including ensure the patient is iso-centered before starting a scan. If the patient cannot be positioned at the iso-center of the RTP, instructions were provided. The customers are asked to complete the Mandatory Fax-Back Form acknowledging receipt and understanding of the letter.For questions or assistance contact Siemens Medical Solutions USA, Inc. at (800) 888-7436 (USA). |  | 18 | 18 pallets | Worldwide Distribution: Nationwide (USA) including the states of Florida and Texas; and the countries of Canada, China, Italy, Mexico, Poland, Portugal and Russia. | Terminated | November 27, 2012 | 524 | 18 | Computer
Z-2977-2011 | 59107 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARTISTE/RT Therapist v4.2 and/or Oncologist v4.2 systems. |  |  | June 03, 2011 | 2011 | Siemens Medical Solutions USA, Inc | A potential safety rick related to transferring reference images for offset calculation between a syngo RT Therapist v 4.2 or a syngo RT Oncologist 4.2 and other versions of the COHERENCE/syn go RT Therapist or Oncologist or vice versa. An incorrect offset may be applied for patient positioning and result in dose being delivered to the wrong location. | Software design | Siemens sent an "URGENT: MEDICAL DEVICE CORRECTION/CUSTOMER SAFETY ADVISORY NOTICE" letter dated June 3, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter provides preventive measure instructions for customers to use until an update for the system is available. |  | 121 | 121 devices | Worldwide Distribution-USA (nationwide) including the states of CA, FL, NC, NY, OH, PA, SD, UT, and WI and countries of Australia, Belgium, Croatia, Czech Republic, Germany, India, Ireland, Japan, Poland, Saudi Arabia, South Africa, Spain, Sweden, and United Kingdom. | Terminated | March 30, 2012 | 301 | 121 | Computer
Z-2981-2011 | 55941 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | Aplio XG model SSA-790A |  |  | June 02, 2010 | 2010 | Toshiba American Medical Systems Inc | Toshiba America Medical Systems Inc (TAMS) is initiating a field correction on the the Aplio XG model SSA-790A because it did not meet the requirement of clause 15b, IEC60601-1 which requires that the residual voltage shall not exceed 60V after 1 second from taking off the AC plug. | Device Design | Toshiba America Medical Systems (TAMS) sent an "URGENT:MEDICAL DEVICE CORRECTION' letter dated June 2, 2010 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers.Customers were instructed to be careful not to touch the contacts or prongs on the AC plug within two seconds after unplugging the system from a wall socket. This precaution is to be adhered to until a Toshiba Service Representative replaces the power supply. Customers are to complete and return a Customer Reply Form and fax to 877-349-3054.Customers with questions regarding the notification should contact the Director of Regulatory Affairs at (800) 421-1968 or call their local Toshiba Representative at (800) 521-1968. |  | 78 | 78 | Nationwide Distribution | Terminated | April 20, 2012 | 688 | 78 | Computer
Z-1649-2011 | 57914 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | TomoTherapy Hi-ART System, Model # H-0000-0003 |  |  | January 26, 2011 | 2011 | TomoTherapy Incorporated | It was determined that the Treatment Planning Station (TPS) can potentially under dose. | Software design | Consignees were sent a TomoTherapy "Field Safety Notice Medical Device Correction" letter dated January 26, 2011. The letter was addressed to "TomoTherapy Customer". The letter described the Planning Issue, Related DQA Consideration, Graphical Presentation of Results, Product Affected, Recommended Action and Resolution. The letter suggests to use the information presented in order to avoid creating plans with an increased potential for discrepancies in planned vs. delivered dose. Specifically include measurements for small, off-axis targets that could be susceptible to the effects of this issue. DQA dose measurements should be performed at each unique full fraction dose that will be used for patient treatment, without using DQA scaling.If Customers have any questions they are to contact TomoTherapy Interaction Center by email or phone. |  | 326 | 326 | Worldwide Distribution -- US, including States of: AL, AK, AZ, AR, CA, CO, CT, DC, FL, GA, HI, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MO, MT, NE, NV, NJ, NM, NY ,NC, ND,OH, OK OR, PA, PR, SC, SD, TN, TX, UT, VA, WA, WI, and WY and countries of SPAIN, GERMANY, FRANCE, INDIA, KOREA, SWEDEN, ITALY, MALAYSIA, SWITZERLAND, CANADA, UNITED KINGDOM, SINGAPORE, TURKEY, TAIWAN, AUSTRALIA, JAPAN, BELGIUM, CHINA, MEXICO, NETHERLANDS, POLAND, SAUDI ARABIA, and UNITED ARAB EMIRATES. | Terminated | March 18, 2014 | 1147 | 326 | Computer
Z-2964-2011 | 59426 | No | HJB | Corneal radius measuring device. | Ophthalmic | Ophthalmic | 510(K) Exempt | Lenstar LS 900 EyeSuite Biometry version i2.000, Four (4) Protective Cardboard Containers Per Each System. |  |  | June 20, 2011 | 2011 | Haag-Streit USA Inc | A patient record may be linked to an incorrect/inconsistent set of calculation results under specific circumstances when using EyeSuite Software version i2.000. | Software change control | Haag-Streit USA, Inc. sent an "Urgent Field Safety Notices" dated June 20, 2011 to all affected customers.The letter included a description of the affected product and problem. The firm instructed their customers to make sure all new and current templates are saved and upgrade to EyeSuite version i4.000. Customers are asked to complete and return the Software Update Verification Form.For additional information contact the Haag-Streit at (513) 398-3937. |  | 216 | 216 | Nationwide (USA) Distribution. | Terminated | March 08, 2013 | 627 | 216 | Computer
Z-2965-2011 | 59330 | No | LZS |  |  | Ophthalmic | PMA | STAR S4 IR Excimer Laser System; Software Version 5.00 |  |  | July 18, 2011 | 2011 | AMO Manufacturing USA, LLC | Reports of lower success rates with Iris Registration (IR) image capture with the STAR S4 IR Excimer Laser System v 5.3 that with previous version. | Software design | Abbott Medical Optics (AMO) sent an "ADVISORY NOTICE" dated July 18, 2011 to all customers using the AMO WAVESCAN Wavefront System. The letter identifies the product, problem, and actions to be taken by the customers.The letter instructs customers to create a separate patient file for each eye on the WAVESCAN System (STAR S4 IR Excimer Laser Systems,software versions 5.00-5.30) according to the directions in attachment A. Additionally, customers are to complete and return an attached form to AMO via fax at the regional fax number.Contact an AMO representative for questions regarding this notice. |  | 673 | 673 total v 5.00 - 5.22 | Worldwide Distribution-USA (nationwide) including Puerto Rico and countries of Africa, Asian Pacific, Canada, Europe, Japan, Latin America, and the Middle East. | Terminated | September 10, 2013 | 785 | 673 | Computer
Z-1665-2011 | 56407 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | TSX-301A: AQUILION ONE System |  |  | July 22, 2010 | 2010 | Toshiba American Medical Systems Inc | Toshiba America Medical Systems (TAMS) has found that communication may be cut off in the Aquilion ONE/ Premium Systems, owing to the failure of a part used in the console. | Software design | The firm, Toshiba America Medical Systems (TAMS), sent an "URGENT: MEDICAL DEVICE CORRECTION" letter dated July 22, 2010, to its customers. The letter described the product, problem and actions to be taken. The letter informed the customers that a new relay will be installed to prevent the occurrence of the problem within the next 90 days and that their Toshiba Service Representative will contact them to install the part to their Aquilion ONE/Premium System. The customers were instructed to contact their Toshiba Service Representative, if the problem occurs before the corrective action has been taken; to share the information with all users and review radiologists as well as clinical engineering or Biomedical group at their facility, and to complete and return the attached CUSTOMER REPLY FORM via fax to 877-349-3054 or email to raffairs@tams.com.If you have any questions regarding this letter, please feel free to contact the Director, Regulator Affairs at (800) 421-1968 or your local Toshiba Representative at (800) 521-1968. |  | 43 | 43 units | Nationwide distribution: USA including states of: AL, AR, AZ, CA, CO, FL, GA, IA, KY, LA, MA, MD, MI, MN, MT, ND, NJ, NV, NY, OH, PA, PR, TX, VA, WI, and WY. | Terminated | April 05, 2011 | 257 | 43 | Computer
Z-2953-2011 | 58043 | No | JPA | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(k) | Hemochron JR. Microcoagulation Systems Prothrombin Time (PT). |  |  | May 10, 2011 | 2011 | International Technidyne Corp. | Certain lots of Hemochron Jr. PT Cuvettes (J201) demonstrate a higher bias than historically observed when compared to laboratory reference instruments. | Device Design | International Technidyne Corp. (ITC) sent "Urgent Medical Device Recall" letters dated May 10, May 25 and June 09, 2011 to all affected customers.The letter included a description of the problem, asked customers fill out the attached form, and return affected product for replacement.For questions, please contact ITC Technical Support at (800) 631-5945. |  | 10931 | 6,218 boxes in US ; 4713 boxes outside US | Worldwide Distribution: (USA) Nationwide Distribution including the countries of; Austria, Australia, Canada, Switzerland, Germany, Denmark, Estonia, Spain, Finland, France, United Kingdom, Hong Kong, Ireland, Israel, India, Italy, Korea, Liechtenstien, Libya, Netherlands, Norway, Oman,Portugal, Qatar, Saudi Arabia, Sweden, Turkey, Taiwan, and South Africa. | Terminated | November 02, 2012 | 542 | 10931 | Computer
Z-1675-2011 | 56415 | No | IXR | Radiologic table. | Radiology | Radiology | 510(K) Exempt | Toshiba DUA-450F Kalare Radiographic Tilting Table. |  |  | July 15, 2010 | 2010 | Toshiba American Medical Systems Inc | Top table lateral movement on their DUA-450F Kalare Table cannot be performed, or the table top lateral movement is performed at an extremely low speed. As a result, positioning of the patient cannot be performed by tabletop movement. | Software design | Toshiba America Medical Systems (TAMS) issued Urgent Medical Device Correction notification dated July 8, 2010, to consignees via USPS return receipt mail. The letter contains a return reply form that is to be faxed to TAMS as proof of notification. |  | 238 | 238 systems | US Nationwide | Terminated | April 20, 2012 | 645 | 238 | Computer
Z-2948-2011 | 59252 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | BrightView XCT Gamma Camera, 3/8". |  |  | June 15, 2011 | 2011 | Philips Medical Systems | Non-conformances with BrightView X/XCT systems running version 2.0.1 including data transfer, XCT acquisition, Pixel Mask for image reconstruction and table moving incorrectly. | Software design | Philips Medical Systems, Inc. set a "Customer Information" letter dated July 8, 2011 to all affected customers.The letter informed them of an upcoming software updated (version 2.5).For questions please contact 800-722-9377 (option 5). |  | 213 | 213 units, all systems | Worldwide Distribution: (USA) Nationwide distribution and the countries of Greece, Australia, Belgium, China, Czech Republic, France, Germany, Indonesia, Israel, Japan, Lithuania, Malaysia, Netherlands, New Zealand, Norway, Poland, Portugal, Puerto Rico, Saudi Arabia, Slovakia, South Korea, Spain, Switzerland, Taiwan U.A.E., and the UK. | Terminated | February 15, 2013 | 611 | 213 | Computer
Z-2930-2011 | 59300 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Intuitive Surgical da Vinci Si Surgical System (IS3000); |  |  | July 05, 2011 | 2011 | Intuitive Surgical, Inc. | Potential failure of the retention component of the Master Tool Manipulator (MTM) of the Intuitive Surgical da Vinci Si Surgical System which could cause uncontrolled movement. | Device Design | Intuitive Surgical sent an "URGENT- MEDICAL DEVICE CORRECTION" letter dated July 1, 2011 to all affected customers. Customers were asked to complete the attached response form and return as per the instructions. The letter identifies the product, problem, and actions to be taken by the customers. The letter reinforce to customers to follow the product labeling WARNING. The letter states that all affected customers sites will be receiving a call from a Customer Support staff to schedule a retrofit of the component. A letter was also sent to non affected customers to inform them of the urgent medical device correction and actions to be taken. If you have any questions, please contact Customer Service at 1-800-876-1310 or for Outside United States call +41 21 821 2020. |  | 183 | 183 | Worldwide Distribution -- USA (nationwide) and the countries of: Belgium, Czech Republic, Denmark, France, Germany, Greece, India, Ireland, Israel, Italy, Netherlands, Norway, Philippines, Portugal, Saudi Arabia, Singapore, South Korea, Spain, Switzerland, Turkey and United Kingdom. | Terminated | September 02, 2011 | 59 | 183 | Computer
Z-2931-2011 | 58852 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Pinnacle Pelvic Floor Repair Kit--Anterior/Apical STERILE, and Pinnacle Pelvic Floor Repair Kit--Posterior STERILE. |  |  | May 10, 2011 | 2011 | Boston Scientific Corporation | The device may exhibit low tensile strength between the needle and suture and lead to needle detachment during mesh leg placement. | Device Design | Boston Scientific Corporation sent on May 10, 2011 an "URGENT MEDICAL DEVICE RECALL-IMMEDIATE ACTION REQUIRED" letter to all affected customers.The letter included; reason for recall, instructions for the quarantine, discontinued use, and return of the devices, and information for Medwatch reporting. The letter also includes a form that is to be faxed to the firm.For information call (508) 683-4678. |  | 540 | 540 | Worldwide Distribution; (USA) Nationwide distribution, and the countries of Puerto Rico, France, Great Britain, Ireland, Italy, Spain, Sweden, Canada, and Australia. | Terminated | November 15, 2011 | 189 | 540 | Computer
Z-2918-2011 | 58978 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Aruba Wireless Network Access Points, Model 65 utilized by the Acuity Central Monitoring System. |  |  | May 16, 2011 | 2011 | Welch Allyn Protocol, Inc | Aruba Wireless Network Access Points, Model Number 65, experienced a higher than expected incidence of dropouts due to external RF interference. | Software design | Welch Allyn Protocol, Inc. sent a letter dated May 16, 2011 to affected customer.The letter stated there were concerns over drop-outs in the system recommended the system at this hospital to revert back tothe Symbol wireless infrastructure. For questions call Welch Allyn at 315-685-4100. |  | 12 | 12 units were affected | Worldwide distribution: (USA) Nationwide distribution including the state of California; and the countries of Singapore; Philippines, France, Italy, Malaysia, Poland, UK, England, UAE, Germany, Spain, Ireland, Portugal, New Zealand, Hong Kong, Thailand, Australia, Canada, and the Netherlands. | Terminated | August 12, 2011 | 88 | 12 | Computer
Z-2919-2011 | 59102 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System, XiO versions equal to or greater than 2.4.0 and less than 4.63.0 |  |  | June 09, 2011 | 2011 | Computerized Medical Systems Inc | The XiO Software currently prompts for entry of CT to ED data for a studyset. No indication is given that the Electron Density data should only be used for Photon and Electron calculations and either Relative Stopping Power or Mass Density should be used for Proton calculations (depending on the algorithm). | Software design | The firm, Eletka, Inc., issued an "Important Safety Notice" dated May 18, 2011 via mail to all affected sites. The notice explains the issue and identifies a work around to the problem until the new release is available. The customers were instructed to complete and return the enclosed "Confirmation of Receipt" form via FAX to 314-993-1175, attn: QA; e-mail or mail: CMS Software, Attn: QA, 13723 Riverport Drive, Suite 100, Maryland Heights, MO 63043.The notice states that "this problem has existed since XiO Release 2.4.0 and has been resolved in XiO Release 4.63.0 by adding warning messages and explanations to the XiO On-Line Help. Your site will be notified when this release is available."If you have any question, call 314-993-0003, 800-878-4267 or 408-830-8023. |  | 7 | 7 | Worldwide distribution: USA (nationwide) including: IL, IN, MA, NJ and OK; and country of Germany. | Terminated | April 04, 2013 | 665 | 7 | Computer
Z-1676-2011 | 57932 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX Remote Cardiology Review Station, a cardiovascular information system. Model Number CV7.8 |  |  | February 10, 2011 | 2011 | AGFA Corp. | A discrepancy in the validation testing which resulted in the product not performing as intended. | Device Design | AGFA sent an "URGENT SAFETY NOTICE" letter dated 2/10/11 via FED-EX to the customer. The letter included an acknowledgment which was to be FAX-Back indicating that the information was received and understood. Agfa has discussed with the customer how the de-installation of their CRS Remote product - Software Version 7.8_HL would be performed. |  | 1 | One | TX | Terminated | September 15, 2011 | 217 | 1 | Computer
Z-2914-2011 | 59166 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Vertical Brake Hubs component of the Philips MX8000 Series Dual v. Exp. Computed Tomography System, Serial Numbers: 8008-8413, 80001-800022, Model #728130, 510 (k) #K010817. |  |  | June 03, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | Philips Medical Systems decided to conduct a recall/field correction of the patient support vertical brake hub component in all of their CT MX8000 Dual v. Exp Computed Tomography X-Ray System scanners due to a component defect which could cause an unexpected downward motion of the table (couch) while the patient is on the table. This event could occur because the short key connecting the vertical | Device Design | Philips Medical Systems (Cleveland) Inc. issued an "URGENT - Medical Device Correction - Field Safety Notice" letter dated June 1, 2011 to all affected customers, via Certified Mail for tracking purposes. The notification identified the product, the problem involved, and the action(s) to be taken by the customers with regard to this device issue. The notification instructs the customers to monitor their units and if the recalled product malfunctions, to immediately notify their Phillips service representative. The letter also states that a Philips Field Service Engineer will visit each customer location and perform a brake repair or replacement on their system to resolve the issue. For further information or support, the customers are instructed to contact their local Philips representative or local Philips Healthcare office. For North America and Canada the customers are instructed to contact the Customer Care Solutions Center (1-800-722-9377, option 5: Enter Site ID or follow the prompt(s). |  | 412 | 412 system units were distributed. | Worldwide Distribution -- USA (nationwide) including the states of: AR, AZ, CA, CO, FL, FA, HI, IL, IN, KY, MA, MI, MO, MS, NC, NV, NY, OH, OK, SC, SD, TN, TX, VA, VT, WA and Puerto Rico and the country of Canada. | Terminated | August 07, 2013 | 796 | 412 | Computer
Z-2915-2011 | 59210 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | GEMINI TF Big Bore computer tomograph software version 3.6. Catalog #882476, Gantry S/N 9001-9023. |  |  | August 05, 2009 | 2009 | Philips Medical Systems (Cleveland) Inc | Two non-safety related defects have been found on the Philips GEMINI TF Big Bore running version 3.6; if these defects were to recur, it could impact the performance or reliability of the system. Philips is issuing a software update to all Philips GEMINI TF Big Bore customers to correct these defects. | Software design (manufacturing process) | Philips Medical Systems (Cleveland) Inc. sent an "GEMINI TF BIG BORE CPE CUSTOMER LETTER" dated October 20, 2011 to all affected customers. The letter describes the product, problem, and updates that are available to the customers. The letter provides additional information about acquisition workflows for SUV Calibrations and Validations.For support concerning this upgrade, contact the Customer Care Solutions Center at 1-800-722-9377. |  | 23 | 23 units | Worldwide Distribution-USA (nationwide) including the states of CO, LA, MD, MI, MO, MN, PA, TX, VA, VT and WI and the countries of Belgium, Germany, Italy, Saudi Arabia and Switzerland. | Terminated | September 21, 2012 | 1143 | 23 | Computer
Z-2902-2011 | 59136 | No | CBR | Nitrous oxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Ultraview Spacelabs Medical Multigas Analyzer, Model 91518, V1.00.14. |  |  | June 08, 2011 | 2011 | Spacelabs Healthcare, Llc | Spacelabs received a report of the Model 91518 Multigas Analyzer trending low levels of an anesthetic agent that is not present during a particular case. The error is limited to only the trend table data. | Software design | Spacelabs Healthcare, Llc. sent a "Customer Bulletin" dated June 20, 2011 to all affected customers. The customers were notified of the trend problem and are advised to follow the instructions in the Perioperative section and the ADMIT/DISCHARGE section of the users manual.For questions contact Spacelabs Healthcare at 1-800-522-7025 and select 2 for Technical Support or 1-425-657-7200 x5089. |  | 481 | 481 units worldwide (333 units internationally and 148 units in the US) | Worldwide Distribution: (USA) Nationwide distribution, and the countries of ARGENTINA, BAHRAIN, BOLIVIA, CANADA, CHILE, CHINA, COLOMBIA, CZECH REPUBLIC, ECUADOR, EGYPT, FRANCE, INDIA, MEXICO, NICARAGUA, PARAGUAY, PERU, POLAND, Slovakia, SWITZERLAND, SYRIAN ARAB REPUBLIC, TAIWAN, CHINA, THAILAND, UNITED ARAB EMIRATES, U.K., URUGUAY, VENEZUELA, and VIETNAM. | Terminated | September 21, 2011 | 105 | 481 | Computer
Z-2883-2011 | 59212 | Yes | MSD | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Arrow NextStep Antegrade Chronic Hemodialysis Catheter. |  |  | June 28, 2011 | 2011 | Arrow International Inc | Firm has received complaints indicating breakage and/or separation of the stylet within the Arrow(R) NextStep(R) Antegrade Catheters. | Device Design | An "Urgent Medical Device Recall" letter dated June 28, 2011 was issued to consignees. The letter requested that customers check their stock, cease use and distribution, and quarantine all affected product. Consignees were further directed to complete the "Recall Acknowledgement &amp; Stock Status Form", return the form to their sales representative, and return affected product as instructed on the letter. For questions or concerns, call 1-800-233-3187. |  | 41 | 41 | United States (states of CA, DE, FL, MI, NC, and TN). | Terminated | April 16, 2012 | 293 | 41 | Computer
Z-2848-2011 | 59047 | No | BYD | Heat and moisture condenser (artificial nose). | Anesthesiology | Anesthesiology | 510(K) Exempt | Vital Signs- A GE Healthcare Company - Hygroscopic Condenser Humidifier (HCH)/Anesthesia Breathing Circuit; GE Medical Systems SCS 283 rue de la Miniere 78530 BUC, France p +33.1.30.70.40.40; Vital Signs Colorado Inc 11039 E. Lansing Circle Englewood, CO 80112 P +1.973.956.5300. |  |  | April 29, 2011 | 2011 | Vital Signs Devices, a GE Healthcare Company | The Hygroscopic Condenser Humidifier (HCH) device may have an occlusion that can prevent proper flow of the medical gases or oxygen, possibly resulting in insufficient oxygen delivered to the patient. | Device Design | An "Urgent Medical Device Correction" letter with the return confirmation form was sent to affected customers. Phone calls were also placed to customers. Consignees were informed of the affected device, issue, and actions to take. Customers were instructed to isolate all affected product, arrange return by contacting Customer Service at 1-800-932-0760, and return only unused product. Customers were asked to complete the attached confirmation form and fax back as per the instructions. For questions, contact your local Vital Signs Account Manager or call Customer service at 1-800-932-0760. |  | 222009 | 222,009 | Worldwide Distribution including the United States. | Terminated | August 06, 2014 | 1195 | 222009 | Computer
Z-2895-2011 | 58031 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Aquilion One: Aquilion TSX-301A, 64 System, X-Ray, Tomography. |  |  | January 07, 2011 | 2011 | Toshiba American Medical Systems Inc | Toshiba America Medical Systems (TAMS) is initiating a field correction for the TSX-301A and TSX-301B: Aquilion 64 System, X-Ray, Tomography because communication between reconstruction units in the system console may be broken due to failure of the optical conversion connector, resulting in a break of communication between units. | Software design | Toshiba America Medical Systems, Inc. sent an "Urgent Medical Device Correction" letter dated January 4, 2011 to all affected customers.The letter provides the customers with an explanation of the problem and the firms plan to install preventive measures and provided them with a Customer Reply form to send Toshiba. A representative will contact the customers for an appointment to install the measures to their CT system. For question on this recall call (800) 421-1968 or contact your local Toshiba Representative at (800) 521-1968. |  | 6 | 6 | (USA) Nationwide Distribution including the states of WI, MO, TN, IA, AL, AZ, TX, and GA. | Terminated | August 17, 2012 | 588 | 11 | Computer
Z-2897-2011 | 57856 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Aplio Artida Ultrasound Diagnostic System SSH-880CV. |  |  | October 01, 2010 | 2010 | Toshiba American Medical Systems Inc | Toshiba America Medical Systems, Inc. is initiating a field correction on the Aplio Artida Ultrasound Diagnostic System because the panel operation for the Aplio Artida may become disabled (i.e., system may hang up) when processing to display patient information (heart rate, time, VCP counter, ect.). | Software design | Toshiba America Medical Systems, Inc. sent out an "Urgent Medical Device Correction" dated December 21, 2010 to all affected customers.The letter provides the customers with an explanation of the problem identified and an action to be taken. Customers were notified that Toshiba will issue a field modification to correct the software, and if the problem occurs, they were instructed to shutdown the system holding down the power supply button for 5 seconds for 5 seconds and then restart the system. A form to send back to Toshiba is included. For questions on this recall call (800) 421-1968 or contact your local Toshiba Representative at (800) 521-1968. |  | 46 | 46 units | Worldwide Distribution: (USA) Nationwide distribution including the states of MD, OH, PA, NY, MA, GA, AL, TX, IL, WI, IA, MN, KY, MT, OR, LA, and CA; and the country of Puerto Rico. | Terminated | April 20, 2012 | 567 | 46 | Computer
Z-1691-2011 | 57734 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Liko Sabina EE Lift, Liko, Sweeden. |  |  | January 27, 2011 | 2011 | Hill-Rom, Inc. | The firm has received reports of injuries related to the Sabina Sit-in-Stand lifts. Of six injury reports, two were injuries in patints during use and four injuries were in non-patients when they tripped over the lift device and fell on the unit. | Device Design | The firm, Liko, issued two "URGENT FIELD SAFETY NOTICES" one dated January 27, 2011 and a follow-up notice dated February 9, 2011 to its consignees/customers. The notices described the product, problem and actions to be taken. The customers were instructed to keep the device in a separate unoccupied area when not in use; use extra care and support when lifting patients with the device; not allow unauthorized persons, especially children, to play around or operate the lift; examine the sling hooks daily to make sure they are secure, and to forward a copy of this notice to any other facility personnel they deem appropriate. The notices state the firm is investigating a design change to mitigate the risk of injuries with the device.If you have any questions concerning this request or the procedure as outlined, please contact Hill-Rom Technical Support at 800-445-3720. |  | 1009 | 1,009 all products | Worldwide distribution: USA including: AL, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NE, NJ, NM, NV, NY, OH, OK ,OR, PA, RI, SC, SD, TN ,TX, UT, VA, WA, WI, and WV; and country of: Canada. | Terminated | September 19, 2014 | 1331 | 1009 | Computer
Z-2893-2011 | 59088 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Clinac Linear Accelerators; |  |  | May 24, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | The throat cover on the High Energy Clinac may detach if not properly installed and possibly strike a patient. | Device Design | Varian Medical Systems sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated May 24, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers.The letter provides directions on how to re-install the gantry center throat cover.Customers can contact their Varian Customer Support Manager at 1-888-827-4265 for further clarification. |  | 278 | 278 | Worldwide Distribution-including USA (nationwide) | Terminated | March 14, 2012 | 295 | 278 | Computer
Z-0790-2011 | 56965 | No | DSP | Intra-aortic balloon and control system. | Cardiovascular | Cardiovascular | 510(k) | Arrow Ultra 8 IAB (Intra-Aortic Balloon) Catheter, 8 Fr, 30cc. |  |  | October 11, 2010 | 2010 | Arrow International, Inc., Division of Teleflex Medical Inc. | SuperArrow Flex IAB catheter becomes stuck in the sheath, unable to move the IAB catheter forward or backward, causing a delay in therapy, bleeding or arterial injury. | Device Design | Arrow international initiated a voluntary recall on 12/21/10 by letter requesting users to discontinue use and return all unused 5800 Series IAB with SuperFlex Introducers to Arrow. This field action supercedes the Safety Alert IAB 'Stuck in Sheath' letter issued on 10/8/2010. Questions should be directed towards their local sales representative or the IABP support Line at 1-866-396-2111 or 1-617-389-6400. |  | 5289 | 5289 units | Worldwide Distribution -- USA, Canada, Australia, Belgium, Czechoslovakia, Germany, Russia, Spain, Finland, France, Great Britain, Hungary, Italy, Netherlands, Poland, Saudi Arabia, Switzerland, Brazil, Australia, Chile, Colombia, Ecuador, India, Indonesia, Japan, Korea, Malaysia, Mexico, New Zealand, Panama, Peru, Philippines, Taiwan, Thailand, and Venezuela. | Terminated | February 05, 2015 | 1578 | 69523 | Computer