## Sheet1
Recall Number | Recall Event ID | Device Type | Product Code | Regulation Description | Medical Specialty | Review Panel | Submission Type | Trade Name/Product | Main Name | Recall Class | Date Posted | Year Posted | Recalling Manufacturer | Reason for Recall | FDA Determined Cause | Action | Consumer Instructions | Clean Quantity | Quantity in Commerce | Distribution | Recall Status | Termination Date | Time to Terminate | Merged Quantities | Fault_Class
Z-1747-2013 | 65558 | No | HQL | Intraocular lens. | Ophthalmic | Ophthalmic | PMA | Afinity CQ2015A Intraocular Lens |  |  | June 10, 2013 | 2013 | Staar Surgical Co. | STAAR Surgical Company is recalling a limited number of nanoFLEX CC4204A and Afinity CQ2015A intraocular lenses due to the potential presence of glass particles in the storage vial. | Material/Component Contamination | The firm, STAAR Surgical, sent an "URGENT - MEDICAL DEVICE RECALL" letter dated June 10, 2013 to all the customers that received a limited number of nanoFLEX CC4204A and Afinity CQ2015 Intraocular lenses. The letter described the product, problem and actions to be taken. The customers were instructed to check your inventory; remove any lenses from the attached list; contact STAAR Surgical Customer Service for Returns Goods Authorization Number via phone: 1-800-352-7842; email: customerservice@staar.com; and/or fax: 1-800-952-4923; return listed lenses to STAAR Surgical; complete and return the attached response form by fax to Customer Service at (800) 952-4923, and pass this notice on to all those who need to be aware within your organization or any organization where the devices have been transferred.If you have any questions or concerns regarding the product return, contact STAAR Surgical Customer Service at (800) 352-7842. |  | 3361 | 3,361 units total | Nationwide distribution: USA including states of: AL,AZ, CA, CO, CT, FL, GA, HI, KY, IA, IL, IN, LA, MA, MI, MN, MO, MS, NC, NE, NH, NJ, NM, NV, NY, OH, PA, SC, TN, TX, UT, VA, WI, and WV. | Terminated | October 17, 2013 | 129 | 3361 | Undetermined
Z-2175-2013 | 64150 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Outlook ES pump is intended for use with B. Braun Medical Inc. Horizon Pump IV Sets. Used to regulate the flow of primary and secondary fluids when positive pressure is required. The infusion system is capable of delivering fluid from a negative head height (when the IV fluid container is lower than the pump), and provides clinically accepted volumetric accuracy for all standard IV fluids, including blood, lipid, and Total Parenteral Nutrition (TPN). Model number: 621-100ES, and 621-200ES. To deliver a broad range of drugs with use with the B. Braun Medical Inc. Horizon Pump IV Sets. |  |  | December 05, 2012 | 2012 | B Braun Medical, Inc. | B. Braun Medical Inc. is informing our customers of a field correction previously performed to address the loss of the drug library in the Outlook Pump Models 620-100 and 620-200. This could have occurred after periods of lost DC (battery) power or out of tolerance low power. This could have led to a delay in therapy, or to delivery of an over or under-dose of a therapy if the user was to rely on | Component design/selection | The firm, B. Braun Medical Inc., sent an "URGENT: MEDICAL DEVICE FIELD CORRECTION" letter dated December 05, 2012 to its customers. The letter describes the product, problem and actions taken. The firm finished performing "field upgrades" on March 20, 2010 for this issue.Should you require any additional information, please do not hesitate to contact B. Braun Medical Inc., Infusion Systems Customer Service at 1-800-627-7867 (627-PUMP). |  | 14161 | 14,161 | Worldwide Disribution: USA (nationwide) and country of: Canada. | Terminated | October 24, 2014 | 688 | 14161 | Undetermined
Z-1828-2013 | 65473 | No | JOJ | Sterilization process indicator. | General Hospital | General Hospital | 510(k) | Medline Med-Checks BI Test packs with Instant Readout Integrator, SteriTec Product Code PL 200, Distributor Product Code MDS200500IR |  |  | June 10, 2013 | 2013 | Steritec Products Mfc. Co., Inc. | Mesa Labs, based out of Bozeman, MT, informed SteriTec Products, Inc. of a product recall involving the EZTest Steam and Smart-Read Biological Indicators (BI) because the readout time is not in specification with the label claims. | Nonconforming Material/Component | The firm, SteriTec, sent all customers recall notifications via e-mail. The e-mail included a .pdf file of the recall notification letter entitled "URGENT: MEDICAL DEVICE RECALL" dated June 11, 2013. The letter described the product, problem and actions to be taken. The letter also informed customers of MesaLabs, "URGENT: MEDICAL DEVICE RECALL" letter dated June 7, 2013 to its customers. The customers were instructed to notify SteriTec products of receipt of this notification via email at: Jon@SteriTecproducts.com or fax to: 303-660-4213 Attn: Director Regulatory Affairs; examine the lot number printed on the packaging components; place affected items in hold and do not sell or distribute these lots; record inventory of all items remaining in your warehouse and all items that have been distributed; advise SteriTec of the status of all inventory you received used or unused, if you have distributed the affected product to other locations within your healthcare system, notify them immediately of this recall, ask them to provide inventory status and proceed with destruction or return via instructions provided in letter-If you have shipping questions, please contact SteriTec directly at 303-660-4201, or email us at: Info@SteriTecproducts.com, and also provide this notice to anyone in your facility that needs to be informed.Should you have any questions or concerns, please do not hesitate to contact Director of Regulatory Affairs and Quality Assurance or our offices at 303-660-4201. SteriTec's hours of operation are 7:30am Mountain Standard Time (MST) through 5:00pm MST. |  | 3535 | 3535 units | Worldwide distribution: US (Nationwide) and countries of: Argentina, Australia, Canada, British Virgin Island, Czech Republic, Italy, and Turkey. | Terminated | December 19, 2013 | 192 | 57665 | Undetermined
Z-1819-2013 | 65068 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Custom Spine ISSYS LP Inverted Screw System. 8.5MM X 45 MM ISSYS LP Screw Rx Non-Sterile. Custom Spine, Inc. 9 Campus Dr. Parsippany, NJ 07054 USA. TEL: 973-808-0019 |  |  | March 28, 2013 | 2013 | Custom Spine, Inc. | On March 28, 2013, Custom Spine, Inc. issued a minor removal of a lot of 18 screws that were found to be out of specification. The screw length was 5 mm shorter in length than the required specification for the screw size. The product is laser marked 8.5mm X 45mm, but its length measured 40mm. | Process control | Custom Spine initiated a voluntary removal On 3/28/2013, Custom Spine initiated a voluntary removal of all 10 pieces of the ISSYS LP (8.5mm X 45mm screws) Lot 363. Distributors were contacted via phone and/or email and were asked to immediately return these parts to Custom Spine. |  | 18 | 18 units | Nationwide Distribution including Nebraska and Virginia. | Terminated | October 24, 2013 | 210 | 18 | Undetermined
Z-1830-2013 | 65576 | No | KHO | Fluorometer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AIA Analyzer Pipette Tips used on Tosoh AIA-2000, AIA-1800, AIA-900 and AIA-600II Analyzers, 1000 tips/bag |  |  | May 15, 2013 | 2013 | Tosoh Smd Inc | Visual inspection of the pipette tips indicated that some of the tips were not uniform in size/length which was causing error messages on the analyzers. | Nonconforming Material/Component | Tosoh Bioscience sent an Urgent Recall Notification letter dated May 15, 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to respond on the amount of product that was destroyed by completing the attached form and faxing it to 614-317-1941.Customers with questions were instructed to call Tosoh Technical Support at 1-800-248-6764.For questions regarding this recall call 614-317-1909. |  | 11309 | 11,309 bags | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, and WY.Internationally to Canada, Japan, Mexico, Chile, and Uruguay. | Terminated | November 27, 2013 | 196 | 11309 | Undetermined
Z-1831-2013 | 65084 | No | DCF | Albumin immunological test system. | Immunology | Immunology | 510(K) Exempt | uTASWako DCP, Lot TG340 2013.11, REF 995-60701, Wako Pure Chemical Industries, Ltd. 1-2 Doshomachi 3-chome, Chuo-Ku Osaka 540-8605 Japan Distributed by Wako Diagnostics Wako Life Sciences, Inc. 1025 Terra Bella Ave., Suto A. Mountain View, CA 94043 USA. The product is a single reagent cartridge run on an automated system and each assay performed on a |  |  | April 30, 2013 | 2013 | Wako Life Sciences, Inc. | Product fails to meet the accuracy at the low end of the measuring range. | Material/Component Contamination | Wako Life Sciences, Inc. sent an email and issued a Recall Notice letter via FED EX on April 30, 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue the use of these kits, fill out the attached RMA form, and return unused kits using Wako's FedEx account (provided on RMA form). Partially used reagent cartridges should be disposed of according to your local and state requirements. Wako Life Sciences will relplace your existing stock of uTASWako DCP kits with kits from a different log number that has been tested and has shown to not be affected. If you have any questions regarding the Recall, please submit your inquiries to liver@wakousa.com or call 443-535-0702. |  | 3600 | 36 cartridges -100 tests per cartridge | Distribution to CA, GA, MN, NC, UT and the District of Columbia. | Terminated | August 19, 2013 | 111 | 3600 | Undetermined
Z-1846-2013 | 65646 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | GAMMA3 System Long Nail Kit, R2.0, Ti, Left |  |  | May 21, 2013 | 2013 | Stryker Howmedica Osteonics Corp. | Stryker Osteosynthesis has become aware that LEFT GAMMA3 nails are marked as right nails. | Process control | Stryker sent notification letters and product accountability forms dated May 21, 2013. via Fed Ex with return receipt to all branches/agencies and on May 22, 2013, via Fed Ex with return receipt to all hospital risk management, chief of orthoapedics, and surgeons.Our records indicate that you may have received and/or used the above referenced product(s). It is Stryker'sÂżÂż responsibility as the manufacturer to ensure that customers who may have received and/or used these affected products also receive this important communication.Please contact your StrykerÂżÂż Sales Representative to arrange for return of the product if you have any of the above listed items. Please assist us in meeting our regulatory obligation by faxing back the attached Product Recall Acknowledgment Form as soon as possible to 855- 251-3635.Please note that your signature on the following form only confirms that you received this notification and does not obligate you to take any additional action beyond what is called for in this notification letter.We regret any inconvenience this action may cause you and if you have any questions, please call (201) 972-2100. |  | 5 | 5 units | US Distribution including the states of IL, KY and TN. | Terminated | October 14, 2014 | 511 | 5 | Undetermined
Z-1832-2013 | 65568 | No | GFC | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Exactech Articulated Driver |  |  | May 30, 2013 | 2013 | Exactech, Inc. | When torque is applied to the Articulated Driver at the maximum angle, the cross pin may shift out of position, potentially leading to disassembly of the instrument. | Under Investigation by firm | Exactech sent a" Important Product Recall Notice" dated June 4, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were requested to Immediately cease distribution or use of the products. Extend this information to your accounts that may have this product in their possession. Verify if you have any of the subject ARTICULATE DRIVER (catalog number 181-00-04) in the specified lots. FAx back the attached form.Further questions please call 1-800-392-2832. |  | 57 | 57 | US Distribution including the states of AR, CA, FL, NJ, NY, OK, and TX | Terminated | September 19, 2013 | 112 | 57 | Undetermined
Z-1847-2013 | 65610 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | The Hospira Symbiq Infusion System is intended for the delivery of fluids, solutions, drugs, agents, nutritionals, electrolytes, blood and blood products via parenteral, enteral, intravenous, intra-arterial, subcutaneous, epidural, or irrigation routes of administration. |  |  | February 04, 2013 | 2013 | Hospira Inc. | The devices secondary audio signal fails to activate. | Process control | Hospira sent a Urgent Device Field Correction letter dated February 4, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. In addition, letters were sent by e-mail to ECRI, NRAC, Noblis and ISMP for notification of their members and posting on their websites February 5, 2013. The letter was also posted on the Hospira website February 4, 2013.Customers were requested to return a reply form indicating they have received the letter and notified their users.For further inquiries please contact Hospira Global Complaint Management at1-800-441-4100 ( 8am - 5pm CST, M-F) to report adverse events or product complaints, or Hospira Advanced Knowledge Center at 1-800-241-4002, option 4 for additional information or technical assistance. ( 24hrs a day / 7 days a week. |  | 2791 | 2,791 units | US Distribution including the states of AL, CA, CT, DE, FL, HI, IA, IL, IN, LA, MA, MD, ME, MI, MN, NC, NJ, NY, PA, SD, TX and WA. | Open, Classified |  |  | 2791 | Undetermined
Z-1867-2013 | 65544 | No | FRC | Sterilization process indicator. | General Hospital | General Hospital | 510(k) | Mesa EZtest BIOLOGICAL INDICATOR MONITORING SYSTEM. |  |  | June 04, 2013 | 2013 | Mesa Laboratories, Inc., Bozeman Manufacturing Facility | Mesa EZTest , Self-contained steam biological indicator, are recalled because Mesa recently determined that the read out time is not in specification with the label claims. | Nonconforming Material/Component | Mesa Labs sent an "URGENT: MEDICAL DEVICE RECALL" letter dated June 7, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. For questions contact Mesa Labs at BI-Support@mesalabs.com. |  | 21625 | 865/25 ct boxes, 414/100 ct boxes, 307/500 ct boxes | Worldwide Distribution: US (nationwide) and countries of: Australia, Brazil, Bulgaria, Canada, China, Colombia, Croatia, Egypt, Greece, India,Israel, Japan, Jordan, Malaysia, Mexico, Milanese, Netherlands, Peru, Portugal, Pozan, Romania, Spain, Sweden, Taiwan, Thailand, United Kingdom, and Uruguay. | Terminated | November 22, 2013 | 171 | 532650 | Undetermined
Z-1835-2013 | 65564 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris PC unit, model 8015, with version 9.12-Keyboard Processor and logic board assembly part number TC10006584. |  |  | June 17, 2013 | 2013 | CareFusion 303, Inc. | CareFusion is recalling the Alaris PC units model 8015 (PC unit), version 9.12, because it is operating at an incorrect voltage and it could potentially experience a loss of communication between the PC Unit main Processor and Keyboard Processor. | Process design | The firm, Carefusion, sent an "URGENT MEDICAL DEVICE RECALL NOTIFICATION" letter dated June 17, 2013 to all customers who purchased the Alaris PC unit (model 8015) with version 9.12-Keyboard Processor. The recall letter informs the customers of the problems identified and the actions to be taken. The customers were instructed to either discontinue use of affect PC units until remediated or continue use on patients and follow the steps provided in the letter; fully charge Alaris PC units prior to use; during operation and when not in transport, keep Alaris PC units plugged in at all times; consider having additional back up Alaris PC units available and if you experience the described error code messages on the Alaris PC unit during infusion, you may continue the infusion(s) as programmed unless a programming change is required. Obtain a back-up Alaris PC unit to replace the device. In addition, if the customer experiences a SYSTEM ERROR (111.2000 or 111.2002), or a CHANNEL DISCONNECTED error on the Alaris PC unit contact CareFusion Customer Advocacy at 888-812-3266; 24 hours a day, 7 days a week; email: customerfeedback@carefusion.com. The customers were also instructed to promptly complete and return the enclosed Customer Response Card.CareFusion will provide an update to correct the voltage on the keyboard processor for those affected Alaris PC units. Carefusion will contact all affected customers within 60 days to initiate the scheduling process for the remediation.Customers are instructed to contact CareFusion Support Center at (888) 562-6018 or at supportcenter@carefusion.com for recall related questions; hours 7:00am to 4:00pm PT, Monday-Friday. For adverse reports, customers are instructed to contact Customer Advocacy at (888) 812-3266; 24 hours a day, 7 days a week or at customerfeedback@carefusion.com. For Technical questions, customers are instructed to call (888) 812-3229; hours 6:00pm to 5:00pm PT, Monday-Friday or at dl-us-inf-techsupport@ca |  | 26564 | 26,564 total units (26,220 units for US) | Worldwide distribution: US (nationwide) and countries of: Australia, Canada and Europe. | Terminated | May 20, 2014 | 337 | 26564 | Undetermined
Z-1872-2013 | 65724 | No | FZZ | Skin marker. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Non-sterile Viscot Surgical Skin Markers |  |  | July 10, 2013 | 2013 | Precision Dynamics Corporation | PDC Healthcare initiated this recall of Non-Sterile Viscot Part Number 1437-100 Surgical Skin Markers, becuase they were inadvertently mislabeled as PN STER-VIS Sterile Skin Marker Pen with Labels and Flex Ruter. | Component change control | Precision Dynamics Corp. (PDC) sent a Urgent: Skin Marker Pen Recall letter dated July 10, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.All affected customers have also been contacted by phone and email as of July 12, 2013We want to take every precaution to ensure patient safety.1. Immediately examine your inventory2. Quarantine the product subject to recall3. Complete and return the enclosed response form as soon as possibleYour PDC Healthcare customer service representative will contact you directly to arrange the return of the product. In addition, if you may have distributed this product, please identify your customers and notify them at once of this product recall. Your customer notification may be enhanced by including a copy of this recall notification letter.If you have any questions please call the PDC Healthcare Returns Department at 800-435-4242 ext. 5003.Thank you for your understanding and assistance related to this matter. |  | 1200 | 1,200 pens | US Distribution including the states of SC, TX, TN, WA, OK. | Terminated | January 06, 2014 | 180 | 1200 | Undetermined
Z-0476-2013 | 63187 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Enterprise 8000 Bed; an AC-powered adjustable hospital bed; ArjoHuntleigh AB, Eslov, Sweden; Models 8000, 8001, 8002, 8003, 8500, 8600, 8700 and 8800 |  |  | November 16, 2012 | 2012 | Arjo, Inc. dba ArjoHuntleigh | ArjoHuntleigh has received reports of unintended movement where the backrest and/or thigh sections of the bed have moved without anyone consciously pressing the control button. These incidents have occurred in High Dependency and Intensive Therapy environments where unintended movement of the patient may have serious consequences. | Under Investigation by firm | Please be aware that this is not a new recall. Posting to the Internet was delayed due to a technical error.ArjoHuntleigh sent an Urgent Field Correction Recall letters dated November 16, 2012, to all affected customers, informing the accounts of reports of unintended movement of the backrest and/or thigh sections of the bed without intentional activation of the control buttons, which could have serious consequences due to the patient's links to other life supporting devices. The accounts were requested to disseminate the information to all users of the affected beds, to attach the enclosed warning labels to the outside of both head end safety sides immediately adjacent to the controls, to include a copy of the notice in the Instruction for Use Manuals for their beds, and to complete the Customer Response Form online, using the enclosed card's instructions. Questions were directed to the ArjoHuntleigh at 1-800-323-1245, ext 57985. |  | 26135 | 94 beds in U.S., 31 beds in Canada, 26,010 beds in ROW | Worldwide Distribution-USA (nationwide, including Puerto Rico) and the countries of Algeria, Angola, Australia, Austria, Bahrain, Belgium, Bosnia and Herzegovia, Botswana, Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cuba, Czech Republic, Denmark, Ecuador, Egypt, Falkland Islands, France, Germany, Greece, Hong Kong, Hungary, India, Indonesia, Iran, Iraq, Ireland, Israel, Italy, Jordan, Kenya, Kuwait, Lebanon, Libyan Arab Jamahira, Lithuania, Malaysia, Malta, Mauritius, Mexico, Morocco, the Netherlands, New Zealand, Nigeria, Nordic, Norway, Oman, Panama, Poland, Qatar, Romania, Russian Federation, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Sri Lanka, Switzerland, Syrian Arab Republic, Thailand, Turkey, United Kingdom, United Arab Emirates, Uruguay, Venezuela and Yemen. | Open, Classified |  |  | 29474 | Undetermined
Z-1879-2013 | 65486 | No | LFL |  |  | General &amp; Plastic Surgery | 510(k) | Sonopet Ultrasonic Aspirator Console |  |  | June 24, 2013 | 2013 | Stryker Instruments Div. of Stryker Corporation | A supplied component exhibiting corrosion was built into the power supply boards of the Sonopet Consoles listed above. Theissue is limited to 2 lots of the supplied component. The presence of the corrosion on the relay of the power supply board canpotentially lead to a conduction defect which could cause the irrigation pump to stop working. The purpose of the irrigationpump is to supply sali | Component change control | Stryker sent an URGENT Medical Device Recall Notice letter dated June 24, 2013 to all affected customers. The letter identified the affected product, product issue, risk to health and actions to be taken. Customers were instructed to locate affected product and hold for upgrade, use enclosed response form to document inventory and receipt of notice, and in case of further distribution forward the recall notice. Stryker will dispatch a technician to perform the upgrade to the affected product. For questions regarding this recall please contact Stryker Instruments: Monday-Friday 8am-5pm (EST)Kara Spath, Regulatory Affairs269-389-4518kara.spath@stryker.com. |  | 215 | 215 units total | Worldwide - USA Nationwide including the states of CA, CO, FL, HI, IA, IL, KS, KY, LA, MD, MI, MN, MS, NC, NH, NJ, NV, NY, OH, PA, TX, VA, ITALY, and the countries of AUSTRALIA, CANADA, UNITED KINGDOM, NETHERLANDS, INDIA, JAPAN, NEW ZEALAND, SWITZERLAND, HONG KONG, and TAIWAN. | Terminated | January 30, 2014 | 220 | 215 | Undetermined
Z-1925-2013 | 65152 | No | JIT | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension(R) Mass Creatine Kinase MB Isozyme Calibrator (RC420) |  |  | June 05, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has received complaints that the lyophilized cake is not fully dissolved after the stated time in the MMB calibrator Instructions For Use (IFU). The frequency of this occurrence is low but if it occurs, gel-like clumps may be observed. | Material/Component Contamination | The firm, Siemens, sent an "Urgent Medical Device Correction" letter dated May 2013, to all consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed as follows: Until a permanent solution is identified, please follow the instructions listed in the letter in order to optimize calibration preparation; complete and return the attached Correction Effectiveness Check form via fax to the Siemens Technical Solutions Center at (302) 631-8467 and forward this notification to anyone to whom you may have distributed this product.If you have any questions, please contact the Siemens Technical Solutions Center or your local Siemens technical representative or call 800-441-9250. |  | 7470 | 7,470 | Nationwide distribution: US including states of : AK, AL, AR, AZ, CA, CO, CT, DC, GA, GU, FL, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY. | Terminated | March 18, 2015 | 651 | 7470 | Undetermined
Z-1924-2013 | 65706 | No | MLM | Tacrolimus test system. | Clinical Chemistry | Toxicology | 510(k) | Dimension Tacrolimus Flex reagent cartridge (DF107) and Dimension TACR CAL (DC107) |  |  | July 01, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Siemens initiated a recall due to confirmed complaints of low patient sample recovery with Dimension TACR Flex Reagent cartridge (DF107) lot BB4087 and its linked calibrator (DC 107) lot 3BD029. Falsely low Tacrolimus results across the assay range may lead to alterations in dosing. | Process control | Siemens sent an Urgent Medical Device Recall letter dated July 2013 to affected customers The letter identified the affected product, problem, potential risk, and actions to be taken by customers. Customers were requested to discuss the letter with thier laboratory director to determine if additional follow up is appropriate for patient tests conducted with the affected device lots. Additionally, customers were asked to discard any remaining inventory and to complete and return the field correction effectiveness check for replacement product. They were also asked to forward the letter to anyone wha had received this affected product.For questions contact your Siemens Technical Solutions Center or your local Siemens Technical support representative. |  | 5423 | 5423 | USA Nationwide Distribution including the states of: TX, NC, GA, TN, UT, NJ, MI, VA, AR, LA, MO, FL, AZ, CA, DC, PA, NH, WI, SD, SC, NM, MD, and IN. | Terminated | August 26, 2014 | 421 | 5423 | Undetermined
Z-1923-2013 | 65322 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | BD Vacutainer Multiple Sample and Direct Draw Luer Adapter |  |  | May 01, 2013 | 2013 | Becton Dickinson &amp; Company | BD Vacutainer Multiple Sample and Direct Draw Luer Adapter may have blood leakage and low draw volume. | Under Investigation by firm | BD sent a Product Advisory Notice dated May 3, 2013, via UPS to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discard filled tubes since it is possible they will contain less blood volume (low draws), replace the Luer Adaptor unit being used with a fresh unit, and continue to draw blood samples using the fresh unit. Customers with questions were instructed to contact BD Technical Service at 1-800-631-0174, Option 4, or contact them through website: www.bd.com/vacutainer/contact.For questions regarding this recall call 201-847-6800. |  | 37906400 | 37,906,400 units | Nationwide Distribution | Terminated | December 04, 2015 | 947 | 37906400 | Undetermined
Z-1932-2013 | 64993 | No | MDM | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | T-Handle, Strike Plate, Push Connection, Part Number: SI50023 Rev A |  |  | August 29, 2012 | 2012 | SpineFrontier, Inc. | During an operation the strike plate of a T-Handle loosened and separated from the T-Handle. | Process design | On August 29, 2011, The Supply Chain Team was notified of the recall via email. They were informed that all handles are being pulled from the field. This resulted in the return of all T-Handles to SpineFrontier.Further questions please call (978) 232-3990. |  | 11 | 11 | US Distribution including the states of MA, NV and TX. | Terminated | October 18, 2013 | 415 | 11 | Undetermined
Z-1817-2012 | 61253 | No | LFL |  |  | General &amp; Plastic Surgery | 510(k) | TX1 Tissue Removal System Console. Model: TX1 |  |  | February 22, 2012 | 2012 | American Optisurgical Inc | Two reasons for recall.1. Incorrect pressure relief valve installed inside console, leading to a remote chance of a risk to health. 2. Technical Specifications, under Irrigation, the Operating Pressure value is incorrect. | Nonconforming Material/Component | The firm, American Optisurgical Inc., sent an "URGENT: MEDICAL DEVICE RECALL" letter via email on February 23, 2012. An updated "URGENT: MEDICAL DEVICE RECALL" recall letter dated March 1, 2012 was sent via email on March 1, 2012 to its customer. This updated recall letter described the product, problem and actions to be taken. The letter notified the two reasons for recall: " An incorrect pressure relief valve has been installed inside the console. Should the regulator that controls pressure delivered to the inflation cuff fail, the pressure relief valve would not open until a pressure of 10psi was achieved, which may present a health risk". "The Operator's Manual included with the system, Rev. 1, has a typographical error. Specifically, in Section E. Technical Specifications, under Irrigation, the Operating Pressure value is incorrect". The updated recall included the above second reason for recall. The firm instructed the customer, for replacement of the pressure relief valve, and to return the revision 1 operator manual to the company's address using a pre-paid FedEx account. For questions call 949.580.1266 M-F, 7:00am to 5:00pm PST. |  | 5 | 5units | Nationwide distribution: MN only. | Terminated | June 15, 2012 | 114 | 5 | Undetermined
Z-1935-2013 | 65563 | No | MNI | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | 5.0mm Unit Rod 270mm |  |  | June 05, 2013 | 2013 | Synthes USA HQ, Inc. | DePuy Synthes has issued a Medical Device Correction notice for the 5.0mm Stainless Steel rods. This product was produced using a finishing process not identified as part of the manufacturing specification. No adverse event reports associated with the use of this product were received, and an extensive engineering and medical safety review concluded that the product poses no incremental risk to t | Nonconforming Material/Component | On June 5, 2013, De Puy Synthes notified by letter the direct acc acccounts which were identified to have received 5.0mm stainless steel rods and requested return of the product by calling 1-800-479-6329 to obtain a return authorization number. Completion of the verification section of the letter, including negative responses was also requested. |  | 492 | 492 | Nationwide Distribution. | Terminated | September 21, 2015 | 838 | 492 | Undetermined
Z-1903-2013 | 65617 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Nova Max Plus Monitor System (Kit w/10 Count Vials) |  |  | July 15, 2013 | 2013 | Nova Biomedical Corporation | Reports of false abnormally high glucose readings from test strips | Material/Component Contamination | Nova Biomedical U.S.A issued an Urgent Medical Device Voluntary Recall Letter dated July 19, 2013 via registered/certified letters to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immediately discontinue use of the affected product; contact Nova Diabetes Care Customer Service for replacement strips; do not stop testing blood glucose while waiting for replacement strips; if affected strips must be used follow precautions provided in notification letter. For questions call Nova Diabetes Care Customer Service at 1-800-681-7390.. |  | 1316420 | 1,316,420 kits | Worldwide Distribution - USA Nationwide and the countries of Chile, England , Peru, Nicaragua | Terminated | August 01, 2017 | 1478 | 58957354 | Undetermined
Z-1946-2013 | 65337 | No | IOQ | Powered flotation therapy bed. | Physical Medicine | Physical Medicine | 510(K) Exempt | IsoGel AIR 84, Support Surface. |  |  | July 12, 2013 | 2013 | Stryker Medical Division of Stryker Corporation | Complaints were received from the field indicating that some Air Pumps were flashing and alarming when connected to the IsoGel AIR" surface upon turning the pump on and the setting it to the IsoGel AIR Mode. Upon further investigation, it was determined that the flashing/alarm being experienced was due to the air pressure being out of the expected range; indicating a blockage within the LAL tubi | Nonconforming Material/Component | On 7/12/13, an URGENT MEDICAL DEVICE RECALL NOTIFICATION was mailed to consignees. The notification addressed the affected product and issue, potential hazard, actions needed including faxing (269-488-8691) or emailing (ProductFieldAction@stryker.com) the attached reply form to confirm receipt of notification, and risk mitigation. Please contact the firm at 269-389-8129 if you have questions or concerns regarding this recall. Business hours are M-F 8am-5pm(ET). |  | 586 | 586 units | Nationwide Distribution-USA including the states of OH, ME, IL, CO , FL, TX, MI, GA, and VA. | Terminated | September 11, 2014 | 426 | 586 | Undetermined
Z-1941-2013 | 65828 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | 3.2mm Drill Bit. This is a component within the PediLoc Locking Plate System utilized to pre-drill a hole in bone prior to placement of a bone screw. |  |  | July 24, 2013 | 2013 | OrthoPediatrics Corp | OrthoPediatrics has initiated a voluntary recall of Part # 01-1050-002, 2.5mm Drill Bit, Lot #91300335. Only drill bits that are in their original packaging are being recalled. The description on the package of this drill bit has a wrong size in the description (3.2mm). The recall affects only packaged devices still in original packaging due to misbranding of the product. The drill bit itself has | Process design | On 7/24/13, OrthoPediatrics notified their consignees of this URGENT MEDICAL DEVICE RECALL via phone call. A follow-up email message was then sent with the recall notification, and then a hard copy of the notice was sent via certified letter. The notification addresses the affected product, recall issue, reply form instructions, and directs consignees to return affected product to OrthoPediatrics via a Fed-Ex return prepaid pouch. The firm directs consignees to forward the notification and reply form if the product was further distributed. |  | 68 | 68 | Nationwide distribution, including the states of MO, MI, WA, FL, NY, TN, CO, AZ, VA, TX, CA, MN, LA, NC, and UT. | Terminated | August 26, 2013 | 33 | 68 | Undetermined
Z-1812-2013 | 65619 | No | MLA | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | BD Affirm VPIII Microbial Identification Tests, packaged in kits, 120 test\kit, Catalog # 446257 and 24 test\kit, Catalog # 446252; kits are labeled in part ***Becton, Dickinson and Company, 7 Loveton Circle Sparks, MD 21152 USA, 800-638-8663, www.bd.com/ds*** |  |  | June 13, 2013 | 2013 | Becton Dickinson &amp; Co. | In vitro diagnostic test kit may exhibit invalid internal negative control results. | Under Investigation by firm | Becton Dickinson Diagnostic Systems notified distributors by fax on June 13, 2013, and advised consignees to discontinue use and discard affected product for credit. Distributors were requested to return their distribution list to BD for direct notification by the recaller. End users were contacted by letter on June 13, 2013 and the notification addressed to Microbiology Lab Director/Risk Managers advised that a small number of tests associated with the referenced lots of product may exhibit invalid internal negative control results. If encountered, the issue would result in an invalid test, as described in the package insert. End users were requested to discontinue use of these lot numbers, discard any remaining packages for replacement and return the enclosed response form.Customers with questions were instructed to contact BD Technical Services Department at 1-800-638-8663. For assistance regarding replacements customers were instructed to contact BD Customer Service at 1-800-675-0908.For questions regarding this recall call 410-316-4054. |  | 177 | 177 kits | Nationwide Distribution | Terminated | September 28, 2013 | 107 | 177 | Undetermined
Z-1811-2013 | 65686 | [6 empty] | Sodium Bicarbonate, USP |  |  | June 11, 2013 | 2013 | Church &amp; Dwight Inc | Church and Dwight Co., Inc. has initiated a recall for Sodium Bicarbonate, USP produced by the Old Fort, OH Sodium Bicarbonate Production facility due to metal and other contamination. | Material/Component Contamination | Church and Dwight Inc., issued an Urgent Recall letter dated June 11, 2013, and return response form to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to quarantine and return the affected product. For questions regarding this recall call 609-279-7748. |  | 585000 | 234 sacks at 2500 lbs. each of Dialysis Grade 1.5 ; 880 bags at 50 lbs. each of Dialysis Grade 2 | Nationwide Distribution including IL, PA, IN, OH, MI, TN, TX, MN, MO, NJ, KY,and MD. | Terminated | June 30, 2015 | 749 | 585000 | Undetermined
Z-1807-2013 | 65401 | No | INN | Mechanical chair. | Physical Medicine | Physical Medicine | 510(K) Exempt | Swift Mobil shower chair, Swift Mobil 160 shower chair, and Swift Mobil Tilt shower chair. |  |  | May 23, 2013 | 2013 | Etac Supply Center Ab | Etac is voluntarily conducting a field correction of the Etac Swift Mobil shower chair. If the Swift Mobil shower chair loses a caster and thus becomes instantly unstable, there is potiential risk for harm or injury. | Component change control | Etac sent an "URGENT FIELD CORRECTION" letter dated May 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. |  | 7200 | 7200 internationally | International distribution: Austria, Australia, Belgium, Brazil, Canada, Czech Republic, Denmark, Esonia, Finland, France, Great Britain, Germany, Greece, Hong Kong, Ireland, Isreal, Iceland, Italy, Kingdom of Bahrain, Kuwait, Lithuania, Netherlands, Norge, New Zealand, Spain, Switzerland, Sweden, and United Arab Emirates. | Terminated | October 24, 2014 | 519 | 7200 | Undetermined
Z-1744-2013 | 65355 | light, microscope | IBJ | Microscopes and accessories. | Pathology | Pathology | 510(K) Exempt | Polarizer D Fixed w/Rotating Red 1 lambda plate. |  |  | May 28, 2013 | 2013 | Zeiss, Carl Inc | Carl Zeiss Microscopy, LLC is recalling the Polarizer D component used in All upright ZEISS microscopes (Axio Imager, Axio Scope.A1, Axio Lab.A1) due to Misaligned Full-Wave (Lambda) Plate. | Nonconforming Material/Component | On May 28, 2013, Carl Zeiss Microscopy, LLC recalled a Polarizer Component due to Misaligned Full-Wave (Lambda) Plate. Carl Ziess Microscopy distributed notices to their customers via e-mail. Carl Zeiss Microscopy, LLC requests that all customers immediately discontinue the use of the component for any applications which rely on accurate color-change sequence results. Your service provider will contact you to schedule a visit in which a factory trained service technician will confirm the alignment of your Polarizer D component;s lambda plate. If it is defective, a new Polarizer D component will be installed. There will be no charge for this action. Should you have any questions related to this issue, please do not hesitate to the Carl Zeiss facility in New York by telephone 914-681-7815 or via email at ed.mancini@zeiss.com |  | 10 | 10 | Distributed in the states of MI, TX, SD, FL, GA, NY, and NY. | Terminated | October 29, 2014 | 519 | 10 | Undetermined
Z-1685-2013 | 65503 | No | LRZ | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Integra Hallu-Fix cannulated screw driver (model number P/N 129733ND) and cannulated screwdriver tip (model number P/N 129735ND). |  |  | February 01, 2013 | 2013 | Integra LifeSciences Corp. | The firm has received numerous complaints of the cannulated screw driver tip breaking during surgery. | Nonconforming Material/Component | The firm sent recall notification letters to all consignees via traceable email or traceable fax on 02/01/13. The firm requested that the consignees fill out and return the Recall Acknowledgement and Return Form. |  | 221 | 154 screwdrivers were distributed; 67 screwdriver tips were distributed. | Worldwide Distribution-USA (nationwide) including the states of AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KY, LA, MA, MD, MI, MN, MO, MT, NC, NE, NJ, NV, NY, OH, PA, RI, SC, TN, TX, UT, VA, and WA and the countries of Canada, Austria, Belgium, Switzerland, Czech Republic, Germany, Denmark, Spain, Finland, France, Ireland, Israel, Italy, Lebanon, Luxembourg, Netherlands, Norway, Portugal, Russia, Saudi Arabia, Sweden, United Arab Republic, United Kingdom, and South Africa. | Terminated | July 25, 2013 | 174 | 221 | Undetermined
Z-1725-2013 | 65148 | No | MNR | Breathing frequency monitor. | Anesthesiology | Anesthesiology | 510(k) | Finger Sensor for Remmers Recorder (Formerly SnoreSat Sleep recorder) |  |  | May 08, 2008 | 2008 | SagaTech Electronics, Inc. | The outer jacket is receding away from the black sensor prematurely. | Process design | SagaTech Electronics Inc, sent a letter to their Customers dated May 2008. The letter identified the product the problem and the action needed to be taken by the customer.Please take particular attention when cleaning these sensors by holding the cable somewhere in the middle and clean towards each end. This should become standard practice with all sensors to prolong their lifespan.If you have further questions please call (403) 228-4214 ext. 234. |  | 27381 | 152 units 77 devices from lot 13367 and 75 devices from lot 13710 | Worldwide Distribution: US (nationwide)Âżand countries of: ÂżCanada and Mexico. | Terminated | July 22, 2013 | 1901 | 27381 | Undetermined
Z-1726-2013 | 65670 | No | MQN | Bone plate. | Dental | Dental | 510(k) | Synthes CMF Distraction System |  |  | June 12, 2013 | 2013 | Synthes USA HQ, Inc. | The firm initiated a voluntary recall of the BC Distractor Body which is part of the Craniomaxillofacial (CMF) Distractor System, due to a mis-alignment issue. There have been instances reported within the impacted lots in which the slot in the barrel of the BC distractor was rotated such that it was not possible to attach the B-type and C-type foot plates. | Nonconforming Material/Component | SYNTHES sent an Urgent Notice: Medical Device Recall letter dated June 12, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were requested to examine their inventory for affected product, remove it from use, and return it to Synthes. If they have affected product they were instructed to call Synthes at 1-800-479-6329 to obtain a Return Authorization Number. Complete the Verification Section at the end of the letter and indicate the number of devices found and note the Return Authorization Number and return to Credit/Returns, Synthes, 1101 Synthes Avenue, Monument, CO 80132. If customers did not have the affected product they were asked to complete the Verification Section and faxt to 610-251-9005. Customers with questions were instructed to call 610-719-5450 or e-mail FieldAction@synthes.com.For questions regarding this recall call 610-719-5450. |  | 10 | 10 | Nationwide Distribution including PA, CT, DE, and NY. | Terminated | September 21, 2015 | 831 | 10 | Undetermined
Z-1729-2013 | 65664 | No | FMK | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Safety Lancets - Odyssey, Button Activated; |  |  | July 03, 2013 | 2013 | KooJoo Trading Company | Lancets were not sterilized within the dose range established by sterilization validation. | Process control | The firm, KooJoo, sent a "MEDICAL DEVICE VOLUNTARY T RECALL" letter dated July 3, 2013 via first class mail to its customer. The letter described the product, problem and actions to be taken. The customer was instructed to identify and hold all safety lancets you have in inventory that was received before June 2013; notify your retail customers of this product recall and complete and return the attached Customer Reply Form via email to heegonkim@hotmail.com and also return form along with any unused/undistributed product for replacement.If you have further questions, please email heegonkim@hotmail.com or call 209-403-9191 during 9:00am to 5:00pm M-F PST. |  | 7 | 7, 712, 000 pieces (1928 cartons) | US Distribution GA only. | Terminated | March 13, 2015 | 618 | 7 | Undetermined
Z-1735-2013 | 65088 | No | KIS | Animal and human sera. | Hematology | Pathology | 510(K) Exempt | Various types of Fetal Bovine Serum |  |  | April 12, 2013 | 2013 | Paa Laboratories Inc | Current product labeling band data sheets state that FBS Pharma Grade and FBS/FBS Nativeare pure Fetal Bovine Serum (FBS). These products may contain added adult Bovine SerumAlbumin (BSA)of United States origin, water, and/or cell growth promoting additives. All otherFBS products, as listed below, including FBS Gold and FBS Standard, described as specialtysera', may contain adult BSA of Un | Under Investigation by firm | GE Healthcare notified customers via USPS on April 12, 2013 and April 14, 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customers being contacted by phone as well.Customers were asked to contact their local customer service representative if the product does not meet their needs. Customers were instructed to call (732) 457-8455 or 301-660-0256. For questions regarding this recall call 609-514-6888. |  | 280000 | approximately 280,000 liters | Worldwide Distribution - USA (nationwide), Novia Scotia, and Canada | Terminated | September 14, 2015 | 885 | 280000 | Undetermined
Z-1737-2013 | 64989 | No | MNI | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | KRD1 PedFuse Rods, Model: IM80100-13 Rev B |  |  | December 30, 2011 | 2011 | SpineFrontier, Inc. | One of the rods in each of two test constructs did not meet fatigue requirement as outlined in ASTM F1717-04 dynamic compression testing 5 million cycles. | Vendor change control | SpineFrontier issued a Recall Notification to the affected distributor via telephone. The distributor was instructed to return the affected product for a replacement product. |  | 2 | 2 rods | USA Nationwide Distribution including the state of TX | Terminated | January 08, 2014 | 740 | 2 | Undetermined
Z-1742-2013 | 65456 | No | ERL | Ear, nose, and throat electric or pneumatic surgical drill. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | RemB Electric Universal Driver REF 6400-099-000 For use with the CORE and TPS System sRx Only |  |  | July 01, 2013 | 2013 | Stryker Instruments Div. of Stryker Corporation | The safety margin values detailed in the Engineering Design were entered incorrectly into the programming software. Handpieces that have been programmed with the incorrect parameters could result in unintended activation. There is a potential for injury to the patient or user as a result of unintended forward or reverse activation. There is also the potential that the handpiece may not reach 100 | Process change control | The firm, Stryker Instruments, sent emails beginning May 30, 2013 "URGENT MEDICAL DEVICE CORRECTION" letter dated July 1, 2013 to its customers. The letter identified the device, the reason for the recall, possible adverse events and actions to be taken. The customers were instructed to immediately review this letter, locate the units listed, complete and return the business reply form via FAX to 866-521-2762; or email Kara.spath@stryker.com; if you have further distributed this product, please forward this letter and the attached Business Reply Form to all affected locations. After receiving this form, Stryker will send you a pre-paid shipper to send your affected device(s)back to Stryker to be reprogrammed with the correct specifications. Send back all affected devices using the pre-paid shipper provided to you by Stryker. If you require a loaner, please contact the Stryker Service Department to arrange for a loaner at 888-308-1983. For questions regarding this recall please Stryker Instruments: Monday-Friday 8am-5pm (EST) 269-389-4518 / 269-389-3808 kara.spath@stryker.com / jennifer.mars@stryker.com. |  | 9 | 9 | Worldwide distribution: US (nationwide) including states of: AZ, FL, GA, NC, NY, SC, TX and VA; and countries of: India, Netherlands and Japan. | Terminated | September 16, 2013 | 77 | 20 | Undetermined
Z-1749-2013 | 65492 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Edwards Lifesciences Suction Wand, models S099 and S0998B |  |  | June 19, 2013 | 2013 | Edwards Lifesciences, LLC | Edwards is recalling certain lots of Rigid Suction Wands because they discovered plastic particulate inside the sterile pouch and within the internal diam ether of the wand. | Component design/selection | Edwards sent an Urgent Field Safety Notice Product Recall letter, dated June 19, 2013, via FedEx to all affected customers. Two additional attempts will be made to notify customers to return products. The letter idenditifed the product, the problem, and the action to be taken by the customer. Customers were instructed to review their inventory for any affected device in their inventory, quarantine all affected product from their inventory and return all devices to Edwards. Customers were instructed to acknowledge that they have reviewed the Recall Notice and confirm that they have taken appropriate action by completing, sign and dating the enclosed Recall Response Form and returning the form to Edwards by fax to Edwards Customer Service at 800-422-9329. Customers should contact Customer Service at 800-424-3278 to obtain a Returned Goods Authorization number to return their current product to and receive replacement product. Customers with questions were instructed to call Edwards Customer Service at 800-424-3278. |  | 676694 | S099: 171,587 units and S099B: 505,107 units | Worldwide Distribution - USA (nationwide) and Canada. | Terminated | November 04, 2014 | 503 | 676694 | Undetermined
Z-0629-2013 | 63602 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Medical Surgical Bed, Model 3005S3EX |  |  | December 18, 2012 | 2012 | Stryker Medical Division of Stryker Corporation | Stryker Medical Model 3002S3EX and Model 3005S3EX beds built between March 15, 2012 and May 23, 2012. may have been built with brake rings that do not meet specifications, negatively impacting the holding force of the brakes. | Process control | Stryker sent an Urgent Medical Device Correction letterdated December 18, 2012, to all affected customer. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to locate the affected product and ensure that the brakes are holding properly. Consignees were directed to contact Stryker (1-800-327-0770, option 2) to make arrangements to replace all brake rings and continue to routinely inspect brake functionality. Customers were asked to Fax (269-488-8691) the enclosed acknowledgment and Stryker field service will contact customers to schedule service. Consignees were directed to forward a copy of this notice to the new users and advise us of their new location.For questions regarding this recall call 269-488-8691. |  | 1004 | 1004 | Worldwide Distribution - Nationwide and Canada | Terminated | August 30, 2013 | 255 | 1013 | Undetermined
Z-1753-2013 | 65176 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Hydra Vision DR 60/80 System, Models 700539 and 700540, Liebel-Flarshiem Company. The System is used in radiologic imaging procedures. |  |  | April 25, 2013 | 2013 | Leibel-Flarsheim Company LLC | A complaint was received in which a hospital physicist indicated the dosage was too high on the HUT system. During the service call the systems Maximum Fluoro Dose was found to be 30 R/min which is three times the Maximum Fluoro Dose limit of 10 R/min. | Employee error | The customers received an "Urgent Device Correction" notification letter on May 12, 2013 explaining the product issue and providing recommended actions. Customers were contacted to set up a service call to test and recalibrate equipment as necessary. Questions were directed to 800-778-7898. |  | 12 | 12 units | Nationwide Distribution (US), including the states of TN, KY, CO, KS, CT, HI, and AZ. | Terminated | July 26, 2013 | 92 | 12 | Undetermined
Z-1806-2013 | 65703 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | ProSpace Peek Implant 5 degree x 8.5 x 22 mm |  |  | June 18, 2013 | 2013 | Aesculap, Inc. | Aesculap Implant Systems, LLC initiated a recall of the Aesculap Prospace PEEK 5 degree Implant - 8 x 8.5 x 22 mm (SN038P) because the device is incorrectly etched with an 8 degree lordosis instead of a 5 degree lordosis. The device is labeled correctly, only the etching of the implant may be incorrect. | Employee error | Aesculap notified customers via email and phone on June 18, 2013, and were instructed to check their inventory and return affected product to Aesculap.Aesculap believes that the risk to patient is low.For further questions please call 1-800-234-9179. |  | 9 | 9 | US Distribution including the states of NV and MI | Terminated | January 13, 2014 | 209 | 9 | Undetermined
Z-1752-2013 | 65215 | No | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | ExactaMix 1200 and 2400 Compounders. |  |  | June 04, 2013 | 2013 | Baxa Corporation | Baxter Healthcare is recalling certain ExactaMix automated compounding devices due to the complementary metal oxide semi-conductor battery failure resulting in resetting the internal clock of the device to midnight December 31, 2010. | Nonconforming Material/Component | On 6/04/13, customers were notified by 1st class mailing with mitigating instructions including reminding customers that the battery failure results in a failure of the ExactaMix system to alert the user to replace the disposable set after 24 hours of use. Users should not rely on the alert to trigger the daily replacement of disposable sets. Users should also confirm the correct time and date at the setup of each use. If the compounder displays an incorrect time and date, customers are to contact Baxter Technical Support. |  | 753 | 753 units | Worldwide Distribution-USA (nationwide) including Puerto Rico, and the countries of Canada, Latin America, Europe/Middle East/Africa, and Taiwan. | Terminated | October 16, 2015 | 864 | 753 | Undetermined
Z-1751-2013 | 65647 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | LifeShield; Sterile; Rx only; List No. 19666-28; SYMBIQ PUMP SET, CONVERTIBLE PIN, PIGGYBACK, 2 CLAVES PORTS AND OPTION LOK, NON-DEHP; NOMINAL LENGTH: 97 IN (249 CM); PRIMING VOLUME: 14 ML; For the administration of I.V. fluids; Made in Costa Rica; Hospira, Inc., Lake Forest, IL 60045 USA |  |  | June 18, 2013 | 2013 | Hospira Inc. | Hospira has received reports of customers experiencing air-in-line alarms related to administration sets that have a drip chamber with a ball valve. Preliminary investigation findings found that there is a potential for the ball valve to not seal the drip chamber as intended, leading to air in the line of the administration set. | Under Investigation by firm | Hospira sent a Urgent Device Recall letter dated June 18, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customer were instructed to check their inventories, quarantine any affected products, complete an attached reply form, and return the product to Stericycle for credit using a label included in the mailing. Customers who further distributed the products are requested to forward the recall notification information to their customers.A 100% effectiveness check will be performed for direct Hospira customers who do not return the reply forms and/or the product. Any direct or indirect accounts that return a reply form indicating they have product but do not return it will also be contacted as part of the effectiveness check. In addition, customers who further distributed the product will be requested to provide confirmation that they have notified their customers who purchased the product.Please contact Hospira Customer Care at 1-877-946-7747 ( M-F 7am - 6pm CT) or your Hospira representative regarding replacement product availability and for questions regarding this recall. |  | 14400 | 14,400 units | Worldwide Distribution - US (nationwide) including the states of IA, IL, MN, SD and WA., and Costa Rica. | Open, Classified |  |  | 14400 | Undetermined
Z-1776-2013 | 65141 | No | GDO | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | LIGACLIPÂż Endoscopic Rotating Multiple Clip Applier Convenience Kit (Flex Trays), 10mm M/L |  |  | April 12, 2013 | 2013 | Ethicon Endo-Surgery Inc | Ethicon Endo-Surgery is initiating a voluntary recall for LIGACLIPÂż 10mm M/L Endoscopic Rotating Multiple Clip Applier (ER320) due to potential clip formation and feeding issues which may result in improper clip formation, vessel damage and/or insufficient occlusion of the vessel or other structure | Under Investigation by firm | Ethicon sent an Urgent : Medical Device Recall letter dated April 30, 2013 to all primary and secondary consignees via overnight UPS to allow tracking of the receipt. The letter identified the affected product, problem and actions to be taken. Customers were instructed to examine inventory immediately and return affected products for credit. This credit is intended to facilitate replacement purchase. For questions contact your local representative or call Ethicon Customer Support Center at 1-877-384-4266 option 6. |  | 49357 | 49,357 | Worldwide Distribution - USA Nationwide to all fifty (50) states, and internationally including the following countries: ARGENTINA, ARGENTINA, BAHRAIN, BELGIUM, BRAZIL, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, CYPRUS, ECUADOR, EGYPT, EL SALVADOR, GREECE, HONDURAS, INDIA, INDONESIA, ISRAEL, JAMAICA, JAPAN, JORDAN, KOREA, REPUBLIC OF, KUWAIT, LEBANON, MEXICO, NETHERLANDS ANTILLES, NICARAGUA, PAKISTAN, PANAMA, PARAGUAY, PERU, PHILIPPINES, PUERTO RICO, QATAR, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, TAIWAN, PROVINCE OF CHINA, THAILAND, TRINIDAD AND TOBAGO, TURKEY, UNITED ARAB EMIRATES, URUGUAY, VENEZUELA and YEMEN. | Terminated | May 01, 2014 | 384 | 798201 | Undetermined
Z-1773-2013 | 65565 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | The LIFEPAK 500 Automated External Defibrillator is a semi-automatic defibrillator that uses a patented Shock Advisory System. This software algorithm analyzes the patient's electrocardiographic (ECG) rhythm |  |  | June 12, 2013 | 2013 | Physio Control, Inc. | The LIFEPAK 500 Automated External Defibrillator (AED) could be susceptible to malfunction due to an assembly error. If affected, the malfunction would be the result of internal shorting on a printed circuit board assembly and could render the device inoperable. | Employee error | The firm, Physio-Control Inc., contacted their consignees by telephone or email starting on June 13, 2013 and sent and/or delivered an "URGENT MEDICAL DEVICE CORRECTION" letter dated June 2013 to its customers. The letter described the product, problem and actions to be taken. A Physio-Control representative will be contacting the customers by direct visit, telephone, or email you to coordinate the replacement of their device(s).Customers with questions regarding this notification, please call Physio's Technical Support at 1-800-442-1142, 6:00 a.m. to 4:00. pm. (Pacific), Monday-Friday. |  | 139 | 139 total (125 units in the US and 14 units outside US) | Worldwide distribution: US (nationwide) including states of: IL, MA, NC, ND, NY, PA, RI, TX, WA, and WI; and country of: Canada. | Terminated | September 19, 2013 | 99 | 139 | Undetermined
Z-1784-2013 | 65399 | No | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | HemosIL RecombiPlasTin 2G, Part Number 0020002950 (8 ml size). The product is used for the evaluation of the extrinsic coagulation pathway and the monitoring of Oral Anticoagulant Therapy. |  |  | April 09, 2013 | 2013 | Instrumentation Laboratory Co. | Instrumentation Laboratory is recalling HemosIL RecombiPlasTin 2G (RPT2G), Part No. 0020009580 (8 mL Size) due to being Out of Specification. | Nonconforming Material/Component | Between April 9, 2013 and April 11, 2013, Instrumentation Laboratory distributed "Important Product Correction" notification letters to their U.S. and foreign customers informing them that the on-board instrument stability claim has been reduced to 3 days for all current and future product lots of HemosIL RecombiPlasTin 2G. A second notification to non-respondents was initiated and completed on May 7, 2013. Technical questions were referred to 1-800-678-0710, Option # 2. |  | 5553475 | Kits per product:1)11080, 2)11603, 3)11487, 4)11483, 5)11290, 6)11562, 7)11535, 8)11577, 9)11505, 10)11708,11)11579,12)11580, 13&amp;14) not distributed | Worldwide Distribution- US (nationwide) including states of: MA, ME, CT, NJ, NY, PA, MD, VA, NC, SC, GA, FL, AL, TN, MS, KY, OH, IN, MI, IA, WI, MN, SD, MT, IL, MO, KS, NE, LA, AR, OK, TX, CO, ID, AZ, CA, OR, WA, AK, RI, VT and the countries of: Angola, Argentina, Australia, Austria, Belgium, Bermuda, Brazil, Bulgaria, Canada, Caribbean, Chile, Colombia, Costa Rica, Czech, Denmark, Dubai, Ecuador, Egypt, France, Germany, Gibraltar, Great Britain, Greece, Guatemala, Hungary, India, Indonesia, Iran, Israel, Italy, Jamaica, Jordan, Korea, Kosovo, Kuwait, Lebanon, Lithuania, Malaysia, Malta, Mexico, Morocco, Netherlands, Norway, Panama, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Syria, Taiwan, Thailand, Trinidad, Turkey, Ukraine, Uruguay, Vietnam, and Japan. | Terminated | May 06, 2015 | 757 | 5553475 | Undetermined
Z-1786-2013 | 65573 | apparatus, traction, non-powered | HST | Nonpowered orthopedic traction apparatus and accessories. | Orthopedic | Neurology | 510(K) Exempt | STERIS Fine Traction Device, one unit per box |  |  | May 20, 2013 | 2013 | Steris Corporation | A bearing used in the assembly of the Fine Traction Device did not perform successfully in pre-release reliability testing; a replacement bushing passed reliability testing and was substituted per a change order. The change order authorized use of the substitute bushing in assembly of the Fine Traction Accessory; however, inadvertently the new part was not implemented. During complaint evaluati | Component change control | STERIS Account Managers notified affected customers via onsite visits. All consignees were notified as of 5/29/2013. Customers were instructed to return affected product for a replacement unit. |  | 16 | 16 units | USA Nationwide Distribution in the states of CO, KY, LA, VA, and WI. The affected product was also distributed to Canada. | Terminated | August 28, 2013 | 100 | 16 | Undetermined
Z-1783-2013 | 65679 | No | GAD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Rultract Retractor Ratchet Cover Plate Distributed by Pemco, Inc., Cleveland, OH 44131 |  |  | July 10, 2013 | 2013 | Pemco Inc | The firm became aware of the problem when a ratchet cover plate on a device was returned for repair showed signs of corrosion. | Nonconforming Material/Component | Rultract Skyhook sent a Recall Notification letter dated July 15, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to identify the device and return preferably just the gearbox portion of the device to:Pemco, Inc.Attn: Cover Plate Replacement5663 Brecksville RoadIndependence, OH 44131For customers who are unable to send the gearbox for recall service, Pemco would replace the Cover Plate free of charge.Customers with questions were instructed to call 216-524-2990, ext 103.For questions regarding this recall call 216-524-2990. |  | 87 | 87 units | Worldwide Distribution - USA including AL, CA, CT, FL, GA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MO, MS, NC, NJ, NM, NV, OK, PA, TN, TX, UT, VA &amp; WA. Internationally to Czech Republic, Italy, Japan, Korea &amp; Singapore. | Terminated | February 11, 2014 | 216 | 87 | Undetermined
Z-1759-2013 | 65553 | No | MAL | Vascular graft prosthesis. | Cardiovascular | Cardiovascular | 510(k) | AlboGraft Knitted with Collagen, Bifurcate |  |  | June 17, 2013 | 2013 | LeMaitre Vascular, Inc. | Blood blushing/leaking from the surface of the graft after implantation | Process control | LeMaitre Vascular, Inc.sent an Urgent Field Safety Notice daed June 19, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to identify and return the affected device to MeMaitre Vascular which they will replace with a new unit. Customers with questions were instructed to call 781-221-2266, ext. 183. For questions regarding this recall call 781-425-1670, ext 108. |  | 3 | 3 units | Worldwide distribution: USA including PA and countries of: BRAZIL, BULGARIA, FRANCE, GEORGIA, GERMANY, GREECE, ITALY, LEBANON, NETHERLAND, POLAND, SPAIN, and TURKEY. | Terminated | August 15, 2016 | 1155 | 33 | Undetermined
Z-1799-2013 | 65614 | No | JSJ | Selective culture medium. | Microbiology | Microbiology | 510(K) Exempt | BD BBL Schaedler K-V Agar with 5% Sheep Blood, carton of 20 plates, catalog # 221555, labeled in part ***Becton, Dickinson and Company 7 Loveton Circle Sparks, MD 21152 USA 800-638-8663 www.bd.com/ds*** |  |  | June 06, 2013 | 2013 | Becton Dickinson &amp; Co. | Microbiological identification media may exhibit reduced levels of Vancomycin. | Nonconforming Material/Component | The firm, BD Diagnostic Systems, sent an "URGENT PRODUCT RECALL" letter dated June 2013 to Distributors by fax, US customers via letters by UPS and OUS BD sites via email and teleconference. The letter described the product, problem and actions to be taken. All customers were instructed to discontinue use of the listed product; discard any remaining packages, and complete and return the attached form, whether or not you have any inventory remaining, via Fax to: Attention: Regulatory Compliance, 410-316-4258. BD will issue replacements for the discarded material. It is not necessary to contact BD by telephone to receive replacements.If you have any questions, call 410-316-4000. |  | 92 | 92 cartons | Worldwide distribution: US (nationwide) and countries of:Belgium, Canada, Columbia, Mexico, Singapore, and Taiwan. | Terminated | September 03, 2013 | 89 | 1744 | Undetermined
Z-1951-2013 | 65655 | No | CEM | Potassium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ISE Electrolyte Reference reagent (for Synchron LX Systems and UniCel DxC 600/800 Systems), Part No. A28937, A28945, 467935, 467915. |  |  | September 22, 2010 | 2010 | Beckman Coulter Inc. | Beckman Coulter is recalling the Synchron Systems LX20/UniCel DxC ISE Reference Reagent because it has a potassium (K) response near the instrument calibration flag limits and may exceed the limits on some instruments. | Process control | Beckman Coulter Inc sent an Urgent Product Corrective Action (PCA) letter dated September 22, 2010 to all affected customers. Customers were notified of the the problems identified and the actions to be completed. Customers are instructed to complete and return the enclosed response form within 10 days. Customers with questions are instructed to contact Customer Support Center (Call Center) on the Beckman Coulter website at http://www.beckmancoulter.com/customersupport/support/ or call (800) 854-3633 in the US and in Canada. |  | 718 | 718 units total (132 units in US) | Worldwide Distribution - USA Nationwide and the countries of Belgium, France, Germany, Netherlands, Poland, Spain, Sweden, Switzerland, and United Kingdom | Terminated | August 10, 2013 | 1053 | 718 | Undetermined
Z-1877-2013 | 65369 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Carestream DRX Evolution System, X-Ray System --- |  |  | June 03, 2013 | 2013 | Carestream Health Inc. | Due to potential for an x-ray exposure technique change. | Radiation Control for Health and Safety Act | Carestream will make the correction which includes the installation of a SW patch to the affected customer sites, and to any other affected customer sites utilizing the DRX Evolution with V5. 5 SW outside of the US. --- Once receiving FDA approval on the corrective action plan, the correction will be implemented to all affected devices at identified customer sites: 1. A notification letter was sent to all customers identifying the issue and highlighting the approved procedures of the Corrective Action Plan. The letters were sent on 8/16/13 via Certified Mail, Return Receipt Requested. 2. Carestream Health will contact the customer to schedule a service visit to install software updates to correct this issue. The installation time will be approximately one hour.3. As stated in the customer notification letter, this software update will be provided to the customer at no expense.4. All software updates are estimated to be completed within 1-3 months.5. Carestream Health will provide monthly progress reports on the effectiveness of the plan, including the number of units corrected.Further questions please call (800) 328-2910. |  | 49 | US: 49 units; Foreign: 183 units (revised) | Worldwide Distribution - USA (nationwide) including the countries of Austria, Australia, Belgium, Canada, China, Finland, France, Germany, Italy, Korea, Kuwait, Saudi Arabia, Spain, Sweden, Switzerland, Taiwan, United Kingdom. | Terminated | January 29, 2014 | 240 | 49 | Undetermined
Z-1955-2013 | 65822 | applicator, absorbent tipped, sterile | KXG | Absorbent tipped applicator. | General Hospital | General Hospital | 510(K) Exempt | Fisherbrand Sterile Swab, Calcium Alginate Fiber Tipped, Wood Applicator Swab 100 pouches/boxes, 10 boxes/case. |  |  | June 05, 2013 | 2013 | Fisher Scientific Co | The firm is recalling Fisherbrand Sterile Swabs due to 4 customer complaints reporting the presence of non-pathogenic organisms in two lots. The finding was confirmed in lot 61513 and lot 8710 is also being recalled as a precautionary measure. | Material/Component Contamination | Customers were notified by letter on June 18, 2013 and requested to discontinue use and return all affected product. The firm also asked that any additional affected customer base be contacted and notified of the recall if the lots had been further distributed. If you have any questions, contact Stericycle (reference "Event 4070") at 1-877-410-5558. |  | 1230 | 1230 | Nationwide and the country of Canada. | Terminated | October 20, 2014 | 502 | 1230 | Undetermined
Z-2136-2013 | 65823 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | SC2316, Distal Femoral Augment with Screw |  |  | June 24, 2013 | 2013 | Stelkast Co | The firm became aware of an incident relating to a breach of sterility in the sterility barrier packaging of SC2316, Distal Femoral Augment with Screw. | Process control | Stelkast called and emailed all customers on June 24, 2013, to notify them of the recall. Customers were asked to recover all affected products from their inventory and return them to Stelkast. Customers were instructed to contact Stelkast Customer Service for a Return Authorization (RA) Number prior to shipment to Stelkast. Customers with questions were instructed to call 1-888-273-1583. For questions regarding this recall call 724-941-6368. |  | 11 | 11 | Nationwide Distribution including TX, VA, OK, and PA | Terminated | October 08, 2014 | 471 | 11 | Undetermined
Z-2081-2013 | 64852 | No | JTN | Antimicrobial susceptibility test disc. | Microbiology | Microbiology | 510(k) | Oxoid Antimicrobial Susceptibility Test Discs F300 (Nitrofurantoin), Cat. #CT0036B, packaged in cartridges of 50 discs, 5 cartridges/pack. The firm name on the label is Oxoid Ltd., Basingstoke, Hampshire, England |  |  | March 25, 2013 | 2013 | Remel Inc | Use of affected lots may give false indication of susceptibility to Nitrofurantoin. | Unknown/Undetermined by firm | ThermoFisher sent an Medical Device Recall letter dated March 18, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Please notify any personnel who needs notified of the potential for a false indication of susceptibility to Nitrofurantoin. Please review your inventory for affected product lots and discard the affected product. Return the attached Medical Device Recall Return Response acknowledging the disposal of the affected product.If you have any questions, please contact our Technical Services Department at 800-255-6730 (US) or 913- 888-0939 (International). |  | 60 | 60 packs | US Distribution including the states of AL, AR, CA, FL, GA, IL, KS, LA, MA, MD, NY, OK, PA, RI, TX, VA, and WA. | Terminated | June 03, 2014 | 435 | 60 | Undetermined
Z-2087-2013 | 65890 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes 2.4 VA-LCP 2 CLMN VLR DSTL Radius PL 7H HD/3H Shaft/Right |  |  | July 30, 2013 | 2013 | Synthes USA HQ, Inc. | The possibility exists that 2.4 mm VA-LCP 2 column distal radius plates from a specific lot are etched with "R" for right plate when they should be etched with "L" for left plate. | Process control | Synthes GMBH sent an Urgent Notice: Medical Device Recall letter dated July 30, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to review their inventory and immediately remove the affected product from stock. Customers were advised to complete the Verification Section at the end of the letter. For questions call 610-719-5450 or contact your Synthes Sales Consultant. |  | 11 | 11 | USA Nationwide Distribution in the states of FL, GA, NC, OH, KY, NJ, ID, IN, and MN. | Terminated | May 21, 2015 | 660 | 11 | Undetermined
Z-2099-2013 | 65891 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Stryker |  |  | July 08, 2013 | 2013 | Stryker Howmedica Osteonics Corp. | Stryker received a report that the length of the cannulated screw asnis III, article # 604640 Lot R11482, was incorrect. | Process control | Stryker sent an Urgent Medical Device Removal letter dated July 8, 2013 via Fed Ex to all affected customers. The letter identified the affected product, product issues, potential hazards, risk mitigation and actions to be taken. Customers were instructed to return all affected products available to Stryer Osteosynthesis or contact Stryer customer service and complete and sign the enclosed Business Reply Form and fax a copy to: 1-865-251-3635. Questions can be directed to Stryker at 201-972-2100. |  | 23 | 23 units | USA Nationwide Distribution in the states of: FL. OH, TX, &amp; VA | Terminated | March 31, 2015 | 631 | 23 | Undetermined
Z-2100-2013 | 65913 | No | ELZ | Preformed crown. | Dental | Dental | 510(K) Exempt | 3M ESPE Unitek Primary Stainless Steel Crowns (First Primary UL4), Product code 902150 (Kit- 112 crowns/kit) and 900114 (Refill- 5 crowns/box). |  |  | August 05, 2013 | 2013 | 3M Company/3m Espe Dental Products | 3M ESPE is recalling Unitek Primary Stainless Steel Crowns (First Primary UL4) because these products have been manufactured with the incorrect marginal contour edge. | Process control | 3M ESPE sent a 3M "Voluntary Recall" letter dated August 1, 2013 to all affected customers. Separate letters were addressed to distributors and doctors. The letters described the product affected by the recall, problem and actions to be taken. The doctors were advised to complete the attached inventory checklist and call the 3M ESPE Dental Products Customer Care Center at 1-800-634-2249. Select option #1 to arrange for return and replacement of your product.Distributors were advised to quarantine and return the product and not to further distribute it. For product return they should contact 3M ESPE Dental Products Customer Service at 1-800-237-1650. If the product was further distributed they ask that they inform them of the recall. A letter was prepared that can be used for their communication to customers. for additional questions they should contact 651-733-7767. |  | 3409 | 3409 crowns (2949 US and 460 OUS) | Worldwide Distribution - USA Nationwide in the states of CA, FL, GA, IN, IA, MO, NV, NJ, NY, PA, SC, TN, TX, WA. and the country of CANADA. | Terminated | December 11, 2013 | 128 | 3409 | Undetermined
Z-2097-2013 | 65995 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | E-Z CleanÂż Disposable Laparoscopic Electrode Model Number 0100L and 0100LS. Expiration dates of 05/2018 to 07/2018. Sterile. RX ONLY. This product is packaged individually in a chevron TyvekÂż - polyester peel pouch, 6 pouches per box. |  |  | August 07, 2013 | 2013 | Megadyne Medical Products, Inc. | Megadyne Medical Products is recalling certain lots of E-Z Clean, L-Wire laproscopic Electrode Standard Stem and Split Stem due to potential breach of insulation. | Process control | Megadyne sent an Urgent: Medical Device Recall Field Safety Notice dated August 7, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customer were also notified with follow-up phone calls and email beginning on August 7, 2013. Customers were instructed to examine their stock immediately to determine if you have any of the products from the affected lots on hand. If so, please remove form service/discontinue use and contact Megadyne Customer Service at 1 (801) 576-9669 or (800) 747-6110 (USA) for return authorization and replacement or credit. |  | 1816 | 1816 | Class II Recall - Worldwide Distribution - US including the states of CA, HI, IL, MN, NC, NJ, OH, PA, TX, UT and VA. | Terminated | January 28, 2014 | 174 | 1816 | Undetermined
Z-2104-2013 | 65831 | No | DWE | Cardiopulmonary bypass pump tubing. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Custom Perfusion System. Model Numbers: 1D80R8, 1E34R1, 1F75R1, 1G49R2, 1G75R1, 1P07R8, 1P91R4, 5B15R11, and 5P51R9. |  |  | July 17, 2013 | 2013 | Medtronic Inc. Cardiac Rhythm Disease Management | Medtronic was notified that Edwards Lifesciences has initiated a product recall for the Rigid Suction Wand, model S099B. The Wands are included within certain Perfusion Tubing Packs that are manufactured and distributed by Medtronic. | Process control | Metronic sent an "Urgent Field Safety Notice" letter dated July, 2013 to all affected customers. The letter described the affected product, problem and actions to be taken by User. Customerss were advised to remove and destroy all affected products. At the same time customers were informed that replacement products can be ordered by contacting their Medtronic representative. Customers were requested to complete the Customer Confirmation Certificate and fax it to Medtronic at 651-367-0612. For any questions they can contact Medtronic Lifeline Technical Services at 877-526-7890 or their Medtronic representative. |  | 13223 | 13223 | Worldwide Distribution - USA Nationwide CA, IL, LA, MI and the country of CANADA | Terminated | February 04, 2014 | 202 | 14198 | Undetermined
Z-2117-2013 | 66016 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, TRAM MODULE, High-acuity patient monitoring. |  |  | April 03, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue due to intermittent noise on ECG and/or Respiration waveforms associated with movement of the ECG connector on your DASH, PDM, and/or Tram Module. ECG and Respiration signal noise caused by DASH, PDM, and/or Tram ECG cable connector movement may result in reduced ECG, Respiration and arrhythmia detection performance. | Nonconforming Material/Component | GE Healthcare "Urgent Medical Device Correction" letter dated April 3, 2013 to all affected customers. The letter was addressed to Healthcare Administrator / Risk manger, Chief of Nursing and Director of Biomedical Engineering. The letter described the Safety Issue, Product, Safety Instructions, Affected Product Details, and Product Correction. For questions contact Technical Support at 1-800-558-7044. |  | 4807 | 4807 (2167 US, 2207 OUS, &amp; 434 unknown. | Worldwide Distribution - USA Nationwide including DC and PR except ND and the countries of: ALGERIA, ANGOLA, ARGENTINA, AUSTRALIA, AUSTRIA, BAHAMAS, BAHRAIN, BANGLADESH, BELARUS, BELGIUM, BOSNIA-HERZEGOVINA, BRAZIL, BULGARIA, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, CZECH REPUBLIC, DENMARK, ECUADOR, EGYPT, EL SALVADOR, FINLAND, FRANCE, GERMANY, CHANA, GREECE, HONDURAS, HUNGARY, ICELAND, INDIA, INDONESIA, IRAQ, ISRAEL, ITALY, JAPAN, JORDAN, KAZAKHSTAN, KENYA, KOREA, KUWAIT, LATVIA, LEBANON, LIBYA, LITHUANIA, MACEDONIA, MALAYSIA, MEXICO, MOROCCO, NEPAL, NETHERLANDS, NEW ZEALAND, NIGERIA, NORWAY, PANAMA, PARAGUAY, PERU, POLAND, PORTUGAL, QATAR, ROMANIA, RUSSIA, SAUDI ARABIA, SINGAPORE, SLOVAK IA, SLOVENIA, SOUTH AFRICA, SPAIN, SRI LANKA, SWEDEN, SWITZERLAND, SYRIA, TAIWAN, THAILAND, TUNSIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, BOLIVIA, HONG KONG, JAMAICA, PHILIPPINES, SOUTH KOREA, TRINIDAD TOBAGO, URUGUAY, VENEZUELA. | Open, Classified |  |  | 77509 | Undetermined
Z-2108-2013 | 65793 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Axiom Artis systems. C-arm angiographic x-ray system |  |  | April 26, 2013 | 2013 | Siemens Medical Solutions USA, Inc | The firm became aware of a potential issue with Axiom Artis systems installed by a third party installer that may involve incorrect channel nuts being used. | Nonconforming Material/Component | Siemens sent Update Instructions dated December 2012, to all affected customers. This update was initiated on behalf of the US service organization to re-check all systems installed by MEC Group. A call center representative or designated engineer called each site on April 26, 2013, and described the work being done, amount of time needed and number of engineers that would be on site. At that time the call center or engineer scheduled the update with the customer.For further questions please call (610) 219-6300. |  | 178 | 178 | US Distribution including the states of PA, NJ, DC, VA, MD and DE. | Terminated | November 06, 2014 | 559 | 178 | Undetermined
Z-1948-2013 | 65472 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Colorimetric (CM)/ Rate (RT) incubator wear pads. Product Codes: 1) 6801375, 2) 6801890, 3) 6802445, 4) 6802413 |  |  | June 06, 2013 | 2013 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics is recalling CM/RT wear pads which are a component in the CM/RT incubator due to becoming worn before their replacement. | Nonconforming Material/Component | Ortho Clinical Diagnostics sent an Important Product Correction Notification letter dated June 6, 2013 via FedEx to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were informed that an OCD trained service personnel will contact them and schedule a replacement of the CM/RT wear pads. OCD requested customers to complete and return the enclosed Confirmation of Receipt form no later than June 17, 2013. Customers are asked to place a copy of the notification by each VITROS System in their facility. Foreign affiliates were notified by email on June 6, 2013 of the issue and instructed to notify their consignees of the issue and actions. If you have any questions or need additional information, please call Customer Technical Services at 1-800-421-3311. |  |  |  | Worldwide Distribution - USA Nationwide in the states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY, including Puerto Rico and the countries of Mexico, Singapore, Venezuela, England, Germany, Italy &amp; Spain. | Open, Classified |  |  |  | Undetermined
Z-2135-2013 | 65868 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | ADVIA Chemistry A1c_3 Calibrators, Catalog Number 10491408, IVD --- COMMON/USUAL NAME: ADVIA Chemistry A1c_3 Calibrators |  |  | July 30, 2013 | 2013 | Siemens Healthcare Diagnostics | Siemens confirmed that the ADVIAÂż Chemistry Systems HbA1c method exhibits a positive bias of up to 12% for patient and CAP survey samples due to over-recovery when using A1c_3 Calibrator lots 1MD014, 1BD063, and 2GD014. The established ranges of CAP Survey Samples GH2-04, GH2-05 and GH2-06 and IFCC sample recoveries use Âż 7% from the established mean as their acceptance criteria. This is the crite | Under Investigation by firm | Siemens sent an Urgent Medical Device Correction letter dated July 30, 2013, to all affected customers via FedEx (for US consignees). The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use and discard any of the affected reagent kit lots. Customers werealso requested to complete and return the Effectiveness Check questionnaire attached to the Urgent Medical Device Correction letter within thirty (30) days.Field service personnel were sent an SKB (Service Knowledge Base) describing the issue and instructing them how to deal with customer questions.Customers were instructed to contact their Siemens Technical Solutions Center or their local Siemens technical support representative for questions.For questions regarding this recall call 914-524-2955. |  | 2064 | US: 1546 units; Foreign: 518 units | Worldwide Distribution - USA (nationwide) Puerto Rico and internationally to Argentina, Austria, Brazil, Bulgaria, Canada, Chile, China, Columbia, Croatia, Czech Republic, Denmark, Estonia, Finland, France, Germany, Great Britain, Greece, Guatemala, Hungary, Indonesia, Italy, Kazakhstan, Lithuania, Norway, Poland, Portugal, Romania, Russian Federation, Serbia, Singapore, Slovakia, Slovenia, Spain, Sweden, Switzerland, South Africa, South Korea, Thailand, Turkey, Uruguay, Vatican City, and Vietnam. | Open, Classified |  |  | 2064 | Undetermined
Z-2048-2013 | 65841 | No | DIO | Cocaine and cocaine metabolite test system. | Toxicology | Toxicology | 510(k) | SARKEN STAT! Multi-Drug Test Panel |  |  | July 29, 2013 | 2013 | Sarken, Inc. | These products are being recalled because they were distributed without adequate labeling and instructions for use. Inadequate labeling could lead to misuse of the Test Kits and incorrect test results. | Other | Firm notified customers via telephone, followed up with MEDICAL DEVICE RECALL letter dated July 29, 2013 sent 08/02/13, via USPS regular mail to all affected customers. The recall notification letter is accompanied by a return response, acknowledgement and receipt form. The letter notify customers of the voluntary recall, product description (with prodcut images), reason for recall, potential impact, instructions/actions to follow, and contact information. |  |  | various quantities | Worldwide Distribution - US Nationwide in the country Virgin Islands | Terminated | July 16, 2014 | 352 |  | Undetermined
Z-2139-2013 | 65485 | No | OBC |  |  | Cardiovascular | PMA | LeGooÂż Endovascular Occlusion Gel-0.25 mL LeGoo with 3 Cannula Pouch |  |  | June 21, 2013 | 2013 | Genzyme Corporation, A Sanofi Company | Product contains particulates (fibers) | Material/Component Contamination | Genzyme sent a Medical Device Recall Notification letter dated June 21, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were avised to immediately discontinue use of all lots of LeGoo Endovascular Occlusion Gel and return all LeGoo and return product with the shipping label and packing slip enclosed. syringes.Complete the reply card via fax (414) 459-8791. Customer with questions were instructed to call GENCO at 877-674-2087.For questions regarding this recall call 508-270-2285. |  | 30 | 30 units | Worldwide Distribution - USA (nationwide) and internationally to Austria, Belgium, Canada, Denmark, Egypt. Europe, Finland, France, Germany, Hong Kong, Ireland, Italy, Netherlands, Poland, Saudi Arabia, Slovena, Spain, Sweden, Switzerland, The Netherlands, Turkey, and UK. | Terminated | August 07, 2014 | 412 | 5312 | Undetermined
Z-2134-2013 | 64937 | No | NDN | Polymethylmethacrylate (PMMA) bone cement. | Orthopedic | Orthopedic | 510(k) | CONFIDENCE SPINAL CEMENT SYSTEM KIT, Product Code: 2839-13-000 |  |  | April 17, 2013 | 2013 | DePuy Spine, Inc. | During injection of cement, the water in the hydraulic pump leaks past the piston within the pump body resulting in the loss of pressure and inability to continue to inject cement. | Process control | DePut Synthes Spine sent an Urgent Voluntary Product Recall Notification letter dated April 15, 2013, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the con signee. Consignees were instructed to review their inventory and quarantine for return any products listed in the recall notice., complete the enclosed Business Reply Form, indicating level of inventory of the affected product and return the form to Depuy via fax to 888-943-4897or email to DePuy7862@stericycle.com and to return the product with the enclosed UPS label to Stericycle.For questions regarding this recall call 508-977-6606. |  | 3896 | 3896 units (Overall quantity distributed appears greater than quantities manufactured due to the distribution pathways) | Worldwide Distribution - USA including AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IN, KY, LA, MA, MD, MI, MN, MO, NC, NE, NJ, NY, OH, OK, OR, PA, SC, TN, TX, VA, WA, WV, WY; Worldwide: Algeria, Australia, Austria, Belgium, Brazil, Canada, Colombia, Cyprus, Czech Republic, France, Germany, Hong Kong, India, Israel, Italy, Latvia, Luxembourg, Malaysia, Mexico, Netherlands, Norway, Poland, Portugal, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Turkey, UAE, and UK | Terminated | August 13, 2014 | 483 | 3896 | Undetermined
Z-2137-2013 | 65919 | No | LIT | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific, Small Peripheral Cutting BalloonTM, Sterilized using irradiation: |  |  | July 31, 2013 | 2013 | Boston Scientific Corporation | Boston Scientific is initiating a recall removal of the 4.0mm x 15mm size of Small Peripheral Cutting Balloon (Monorail and Over-the-Wire) Microsurgical Dilatation Devices and Flextome Cutting Balloon (Monorail and Over-the-Wire) Microsurgical Dilation Devices. Boston Scientific has determined that users may experience significant difficulty or inability to remove the protector cap from the devic | Process change control | Boston Scientific sent an Urgent Medical Device Recall Removal letter dated July 30, 2013, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were advised that further distribution or use of the product should cease immediately, and to segregate the product and return it to Boston Scientific. Consignees were asked to complete and return the "Recall Removal Reply Verification Tracking Form". For questions consignees were instructed to contact their local Sales Representative.For questions regarding this recall call 763-494-2042. |  | 1331 | 1,331 | Worldwide Distribution - USA including AL, AZ, AR, CA, CO, CT, DE, DC, FL, GA, ID, IL, IN, IA, KY, LA, MD, MA, MI, MN, MS, MO, MT, NE, NV,NH,NJ, NY, NC, ND, OH, OK, OR, PA, RI, SC,SC, TN, TX, UT, VA, WA, WV, and WI. Internationally to Algeria, Finland, Japan, Spain, Australia, France, Latvia, Sweden, Austria, Germany, Netherlands, Switzerland, Belgium, Great Britain, Norway, Taiwan, Brazil, Greece, Panama, Turkey, Canada, Hong Kong, Poland, Unit.Arab Emir., China, India, Russian Fed., Colombia, Ireland, Saudi Arabia, Czech Republic, Israel, South Africa, Denmark, Italy, and South Korea. | Terminated | January 27, 2015 | 545 | 1671 | Undetermined
Z-2161-2013 | 65970 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | Tec 6 continuous Flow Vaporizer and Tec 6 Plus Anesthesia Vaporizer. The Datex-Ohmeda Tec 6 Plus and Tec 6 Plus NAD Variant is an electronic vaporizer which delivers the anesthetic agent desflurane. The Tec 6 Plus attaches only to Datex-Ohmeda anesthesia machines equipped with the proprietary Selectatec manifold. The Tec 6 Plus NAD Variant attaches only to NAD anesthesia machines equipped with the NAD exclusion type manifold. |  |  | July 29, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue involving the pre-operative check, low pressure leak test at 1% dial setting of the Tec 6 and Tec 6 Plus Vaporizers. The user manual pre-operative check, low pressure leak test may not detect the full range of leaks from seal wear degradation in the vaporizers. If the leak is not detected, it will result in fresh gas leaking to | Component change control | GE Healthcare sent an Urgent Medical Device Correction letter dated August 1, 2013, to all affected customers. The letter described the issue, asked that the Tec 6 and Tec 6 Plus Vaporizers user manual be updated with the attached addendum. Also Users are to perform the user manual recommended pre-operative check which includes the low pressure leak test on every vaporizer with the dial turned to the 12% setting instead of the specified 1% setting as per the attached addendum. If the device does not pass the new low pressure leak at 12%, it should be removed from service and consignee should ontact a GE Healthcare Representative for a repair/replacement. If the vaporizer passes the new low pressure leak test at the 12 % setting, it can continue to be used, however, as detailed in the user manual, continue to routinely perform the pre-operative check which includes the lowpressure leak test.A Medical Device Correction Confirmation form was asked to be sent back to GE.Customers with questions or concerns were instructed to call 1-855-700-4350 or send an email to vaporizer.admin@ge.com. For countries outside the US, customers were instructed to contact their local GE Healthcare Service Representative.For questions regarding this recall call 262-513-4122. |  | 31443 | 31,443 devices (23,947 US, 7496 OUS) | Worldwide Distribution - USA (nationwide) Puerto Rico and internationally to AUSTRALIA, AUSTRIA, BAHAMAS, BAHRAIN, BELGIUM, BOLIVIA, BRAZIL, CANADA, CHILE, CHINA, COSTA RICA, CROATIA local name-Hrvatska, CYPRUS, DENMARK, ECUADOR, FRANCE, GERMANY, GREECE, HONG KONG, INDIA, IRELAND, ISRAEL, ITALY, JORDAN, KOREA-REPUBLIC OF, KUWAIT, LITHUANIA, MALAYSIA, MALTA, MEXICO, NETHERLANDS, NEW ZEALAND, NORWAY, OMAN, PAKISTAN, POLAND, PORTUGAL, QATAR, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SPAIN, SWITZERLAND, TAIWAN, THAILAND, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, and VENEZUELA, | Terminated | April 17, 2015 | 627 | 31443 | Undetermined
Z-2166-2013 | 65993 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Outlook Safety Infusion System Burette Set. |  |  | August 05, 2013 | 2013 | B. Braun Medical, Inc. | Reports were received of an incorrectly assembled Outlook Safety Infusion System Burette Set (Product Code 375038), and the investigation determined that one of the ULTRASITE injection sites and corresponding low pressure backcheck valve was assembled upside down. | Process control | On August 4, 2013, a letter was sent to direct accounts asking them to check their inventory for the affected product. These customers were also instructed to notify their customers if the product was further distributed and to return all affected product and/or the product removal acknowledgement form. |  | 900 | 900 | Distributed in the states of FL, NY, NJ, and TX. | Terminated | September 03, 2014 | 394 | 900 | Undetermined
Z-2163-2013 | 64147 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Outlook ES pump is intended for use with B. Braun Medical Inc. Horizon Pump IV Sets. Used to regulate the flow of primary and secondary fluids when positive pressure is required. The infusion system is capable of delivering |  |  | October 24, 2012 | 2012 | B Braun Medical, Inc. | B. Braun Medical Inc. has become aware of an issue with the Outlook ES Safety Infusion System. The key panel may fail to function as intended if the underlying circuitry fails due to intermittent contact. This issue applies specifically to the HOLD button. If this occurs, the Door Processor (DP) will not receive the desired input from the user. This issue can be observed in pumps with key panels | Component design/selection | B Braun sent an URGENT: MEDICAL DEVICE FIELD CORRECTION letter dated October 24, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact Customer Support at 1-800-627-7867 for questions regarding this notice. |  | 28427 | 28,427 | Nationwide Distribution. | Terminated | July 27, 2015 | 1006 | 28427 | Undetermined
Z-2168-2013 | 65960 | No | GDC | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | OPERON D850 Surgical Table |  |  | April 16, 2012 | 2012 | Berchtold Corp. | Table was not lowering. | Nonconforming Material/Component | Berchtold sent a Customer Notification letter dated April 16, 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter informed customers that BERCHTOLD will upgrade the OPERON Surgical tables with a new firmware version that will eliminate the potential for experiencing problems. For questions call 843-569-6100 ext. 322. |  | 263 | 263 | USA Nationwide Distribution in the states of AL, AZ, CA, CO, CT, FL, GA, HI.IL, IA, KY, LA, MD, MA, MI, MN, MS, NH, NJ, NY, NC, OH, OK, OR, PA, SC, TN, TX, VT, VA, WI and including Canada. | Terminated | September 10, 2013 | 512 | 263 | Undetermined
Z-2167-2013 | 64965 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Xlumena brand NAVIX Access Device, Model/Catalogue Number: NVX-10-03, Part Number: FG-00666, Product is manufactured and distributed by Xlumena, Inc., Mountain View, CA |  |  | April 17, 2013 | 2013 | Xlumena, Inc. | A complaint investigation has found that product fractures can occur at the distal end of the catheter under load. | Process design | Xlumena sent an Urgent Medical Device Recall letter dated April 17, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately discontinue use and segregate the affected product. Customers with questions were instructed to contact dhasker@xlumena.com. For question regarding this recall call 650-961-9901. |  | 148 | 148 units | Worldwide Distribution - USA including AZ, CA, FL, GA, IN, NY,SC, WA, PA, IL, CO, and internationally to Netherlands, Germany, Spain, Denmark, and Switzerland. | Terminated | September 23, 2013 | 159 | 148 | Undetermined
Z-2178-2013 | 65978 | No | NGT | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Monoject 0.9% Sodium Chloride Flush Syringe, 12 mL Syringe with 10 mL Fill |  |  | August 16, 2013 | 2013 | Covidien LLC | Monoject prefill flush syringes may contain non-sterile water and labeled as saline or heparin | Process control | Covidien sent an Urgent Medical Device Recall letter dated August 16, 2013 via Federal Express to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter requested that customers immediately quarantine and discontinue use of the affected products, return any affected product by following the instructions provided in the letter, and complete and return the Recalled Product Return Form to Covidien. The letter states that if the affected products have been distributed, assure that all customers are notified and affected product is removed from the distribution channel as well as from customer stock. Questions or concerns, please do not hesitate to contact your Covidien representative or Covidien Customer Service, Monday through Friday, 8am- 6:30pm ET, at (800) 962-9888, option 1, and then option 2.Firm issued Press on 8/19/13 |  | 3235860 | 3,235,860 units | Worldwide Distribution - USA Nationwide and the country ofBermuda | Terminated | October 21, 2015 | 796 | 3565980 | Undetermined
Z-2085-2013 | 65770 | No | NYQ |  |  | Immunology | PMA | HSV 2 Primers, Catalog #18081020. Analyte Specific Reagent. The HSV-2 primer pair is specific for HSV-2 (Herpes Simplex Virus 2). For use in Clinical laboratories when developing their own diagnostic tests. |  |  | May 13, 2013 | 2013 | Life Technologies Corporation | Firm received complaints that HSV-2 primers failed to amplify HSV-2 DNA. | Process design | Life Technologices sent an Urgent Product Recall letter dated May 13, 2013 by certified mail to all affected customers. The letter identified the product, description of problem, and actions to be taken by the user. Customers were instructed to examine inventory immediately, discontinue use and dispose all affected units. Customers were advised to complete the Customer Response Sheet included for product/quantities consumed, and send e-mail to techsupport@lifetech.com. For questions about the details of this product issue, were instructed to contact Technical Support at 800-955-6288 option 4. |  | 15 | 15 kits | USA Nationwide Distribution in the states of DE, IA, MN, and NC. | Terminated | August 28, 2013 | 107 | 15 | Undetermined
Z-2045-2013 | 65810 | No | OUT |  |  | Neurology | PMA | Pipeline Embolization Device (PED) |  |  | July 19, 2013 | 2013 | Micro Therapeutics Inc, Dba Ev3 Neurovascular | Covidien Neurovascular has initiated the recall of several lots of Pipeline Embolization Device (PED) because the Instructions For Use (IFU) do not include the following contraindication: Patients in whom a pre-existing stent is in place in the parent artery at the target aneurysm location. Additionally, the designation of aspirin and clopidogrel was missing from the following contraindication " | Employee error | Covidien sent an Customer Notification letter dated July 19, 2013, to all affected customers who received affected lots of Pipeline Embolization Device (PED). The customer notification letter informed the customers of the problems identified and the actions to be taken. Customer are instructed to reply and fax response form to Covidien Customer Service at (763) 398-7001. Customers are instructed that Customer Service will issue an RGA number for the return of the affected product back to Covidien. Customers are also instructed that upon receipt of a Returned Goods Authorization (RGA) number, to write theRGA number on the FEDEX slip and send the product back to:CovidienAttention: Product Recall14005 13th AvePlymouth, MN 55441Prepaid and preaddressed FEDEX slips are enclosed for their convenience. Once Covidien receives the returned units, product replacement units will be shipped directly to customers. Local Covidien Neurovascular representative is also available to assist them in expediting the exchange of product.Customers with any questions, are instructed to call (949) 839-3700. |  | 62 | 62 units | US Distribution including the states of AZ, KY, MO, NC, GA, WI, OR, FL, NY, CA, IL, TN, NV, TX, CO, MA, IA, UT and VA. | Terminated | October 17, 2013 | 90 | 62 | Undetermined
Z-1959-2013 | 65857 | No | FLL | Clinical electronic thermometer. | General Hospital | General Hospital | 510(k) | Disposable Temperature Probes, (sterile), Sensor diameter: 3mm/0.12 in; 25 units per box. Vital Signs Inc. a GE Healthcare Company, 20 Campus Road, Totowa, NJ 07512 USA, www.gehealthcare.com/vitalsigns. |  |  | July 18, 2013 | 2013 | Vital Signs Devices, a GE Healthcare Company | GE Healthcare has become aware of a potential safety issue associated with the use of GE General Purpose 9 French Temperature Probe, Disposable, 400 Series [M1024229]. GEs Disposable General Purpose 9FR Temperature Probe M1024229 is intended for oro-esophageal and rectal use. As it pertains to esophageal placement, it has been brought to our attention that some users are inserting the probe | Use error | The firm, GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated July 23, 2013 via Fed Ex beginning on July 29, 2013 to customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure that all potential users in your facility are made aware of this safety notification and the recommended actions. GE will update the Instructions for Use (IFU) and pouch label with new content related to insertion locations. GE has will also place the probe (M1024229) on quality hold until the new labeling changes are implemented.If you have any questions or concerns regarding this notification, please contact GEMSIT Customer Service at 1-800-588-7044 (Domestic) or Vital Signs Customer Service at 1-800-932-0760 (international). Hours of operation: 8:00 am EST to 6:00 pm EST. |  | 12239 | 12,239 units (5,002 units US, 7,237 units OUS). | Worldwide distribution: US (nationwide) and countries of: Australia, Belgium, Canada, Czech Republic, Finland, France, Germany, Italy, Lebanon, Netherlands, Norway, Poland, Saudi Arabi, Singapore, South Africa, Spain, Sweden and United Kingdom. | Terminated | January 02, 2014 | 168 | 12239 | Undetermined
Z-1982-2013 | 65855 | No | ODE | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | EndoGastric Solutions EsophyX2 Device with Serosa Fastener and Accessories. Models R2001 and R2002. |  |  | June 06, 2013 | 2013 | Endogastric Solutions Inc | Endogastric Solutions, Inc. has received a limited number of reports relating to the loss of tissue mold control when operating the R2001 or R2002 EsophyX2 device. In one case, surgical intervention was necessary to remove the device. | Nonconforming Material/Component | Endogastric sent the Safety Alert: EndoGastric Solutions (EGS), EsophyX2 Device letter, dated June 5, 2013, to their consignees.Endogastric sent the second letter URGENT: MEDICAL DEVICE RECALL EsophyX2, dated July 26, 2013. This letter advised customers that the firm is voluntarily recalling EsophyX2 Devices with SerosaFuse Implantable Fastener and Accessories (Models R2001 and R2002) manufactured before February 2012.Customers who have UN-EXPIRED devices in their inventory are advised to discontinue use and fill out the Medical Device Recall Return Response form and return it to Endogastric. The Customer Service will contactcustomers with instructions on how to return the product to the company.Customers who may have EXPIRED devices are advised to with their central supply departments to ensure all identified products have been removed from inventory and destroyed. They should fill out the Medical Device Recall Return Response form with the lot numbers and quantity destroyed and return thecompleted form to the firm.Customers can call the Customer Service and Support at 425-307-9269, Monday through Friday, 8:00AM to 5:00 PM, Pacific Time for questions. Adverse reactions or quality problems experienced with the use of this product may be reported to the FDA's MedWatch Adverse Event Reporting program either online, by regular mail or by fax. |  | 5233 | 5192 units in the US and 41 units outside the US | Distributed nationwide and Italy. | Terminated | April 09, 2014 | 307 | 5233 | Undetermined
Z-1962-2013 | 64711 | No | GNX | Influenza virus serological reagents. | Microbiology | Microbiology | 510(K) Exempt | Remel X/pect Flu A&B; Control Swabs, For In Vitro Diagnostic Use, Ref R246003, each package contains 20 Flu A Control Swabs and 20 Flu B Control Swabs, each swab is individually packaged in a foil pouch. The firm name on the label is Remel Europe, Ltd., Dartford, Kent, United Kingdom. |  |  | February 19, 2013 | 2013 | Remel Inc | Flu A+ Control Swabs incorrectly give Flu A- results when used for quality control testing with the Xpect Flu A&B; test kit. | Employee error | The firm, ThermoFisher Scientific, sent a "MEDICAL DEVICE RECALL" letter dated February 20, 2013 to its customers via regular mail. The letter described the product, problem and actions to be taken. The customers were instructed to review your inventory for affected product lots; discard the affected product; notify any personnel who needs to be notified of the potential for a false negative indication when using the swab, and complete and return the attached Medical Device Recall Return Response Acknowledgment &amp; Receipt Form via Fax to: Attn: Technical Service &amp; Regulatory Affairs at 1-877-428-1924.If you have any questions, please contact our Technical Services Department at 800-255-6730 (US) or 913-888-0939 (International). |  | 1920 | 96 kits of 20 swab sets | Nationwide distribution: AL, AR, CA, CO, CT, FL, GA, IA, IL, KS, MA, MD, MI, MO, NM, NV, NY, OH, PA, TX, WA, and WI. | Terminated | August 14, 2013 | 176 | 1920 | Undetermined
Z-1947-2013 | 64971 | Yes | BTM | Manual emergency ventilator. | Anesthesiology | Anesthesiology | 510(k) | 1) Adult Manual Resuscitator, 2K8004, Oxygen Reservoir Bag, Adult Mask; |  |  | April 22, 2013 | 2013 | Carefusion 2200 Inc | CareFusion has received four (4) customer reports of a component (oxygen connector) within the resuscitation bag being occluded which disables the flow of supplemental oxygen to the resuscitation bag and thus may pose a serious patient safety risk. | Nonconforming Material/Component | CareFusion sent an URGENT RECALL NOTICE letter dated April 22, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers are requested to complete and return an attached Distributor Acknowledgement Form within fifteen (15) days of receipt indicating their acknowledgement of the URGENT RECALL NOTICE and indicate if they have any affected product on hand. They are further requested to perform a 100% physical inventory to verify if any of the affected product codes/lot numbers are in stock within their distribution network/facility. In addition, customers are requested to forward the Customer Acknowledgement Form, the Customer Notification Letter, the Certificate of Destruction Form, the Pictures of Location of Lot Numbers on Case/Product, and the Frequently Asked Questions (FAQs) to all end users which they have distributed the affected product. For questions call 847-473-7097. |  | 14112 | 14,112 units | US Nationwide in the states of AZ, CA, CO, DE, FL, GA, IL, LA, MA, MD, MI, MN, MO, MS, NC, NJ, NM, NY, OH, PA, TX, UT, WA | Terminated | June 16, 2015 | 785 | 14112 | Undetermined
Z-1972-2013 | 65744 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Edwards EndoReturn Arterial Cannula, model ER21B and ER23B |  |  | July 16, 2013 | 2013 | Edwards Lifesciences, LLC | Edwards Lifesciences is recalling certain lots of theEndoReturn Arterial Cannula that were built with an incorrect component. This component can create aperformance problem with the IntraClude IntraAortic Occlusion device. | Nonconforming Material/Component | Edwards sent an Urgent Field Safety Notice letter dated July 16, 2013 via Fed-Ex. The letter identified the affected product, problem and action to be taken. The letter included a Customer Acknowledgement Form to be completed and faxed to 800-422-9329 Attn: Recall Coordinator. If the customer does not send back the acknowledgement form, two additional contacts are made to notify the customer and request return of affected product. This is tracked by Edwards customer service and reported each week. Customer were requested to assess their inventory to assure that all affected product is identified and returned. Returned products will be destroyed. For questions call Edwards Customer Service at 800-424-3278. |  | 855 | 855 | US Nationwide including Canada and VA sites. | Open, Classified |  |  | 855 | Undetermined
Z-1966-2013 | 65732 | Yes | DTM | Cardiopulmonary bypass arterial line blood filter. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Affinity Pediatric Arterial Filter, model number S4014, sterilized using ethylene oxide. |  |  | June 14, 2013 | 2013 | Medtronic Inc. Cardiac Rhythm Disease Management | Medtronic is recalling 148 Affinity Pediatric Arterial Filters from 2 manufacturing lots because a limited quantity of filters may have a small breach in the filter media that may allow unwanted particulate or gaseous emboli to pass through the filter and the outlet of the device, which could result in serious injury to a pediatric patient including neurological damage, or potential patient death. | Employee error | Customer communication was initiated verbally on June 14, 2013, to the 4 affected customers to quarantine any un-used units. An Urgent Medical Device Recall letter, dated June 26, 2013. was mailed to customers on June 28, 2013. The Urgent Medical Device letter referenced that the letter is in follow-up to the previous phone call. It identified the affected product, stated the issue with the 2 affected lot numbers and health risks, and asked that the affected devices be quarantined for return to Medtronic.If you have any questions please contact your Medtronic Sales representative or Lifeline Technical Services at 1-877-526-7890. 1-763-526-7890.Update: Customer communication to the newly identified customer was initiated verbally on August 20,2013, with the customer letter delivery beginning AUG 27, 2013. A Medtronic representative is available to facilitate the return of affected product. |  | 153 | 153 devices. | Worldwide Distribution - US Distribution only to MS and IL, and the countries of Argentina, South Africa and Singapore. | Terminated | March 07, 2014 | 266 | 153 | Undetermined
Z-1970-2013 | 65626 | No | JIS | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ADVIA Centaur Calibrator A. Cat. #04800735, SMN 10285904 (6 pack); Cat. # 04800840, SMN 10285905 (6 pack - REF). |  |  | June 28, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc | Calibrator A lot CA77 is under-recovering to the assigned target for the ADVIA Centaur Total T4 Assay. | Nonconforming Material/Component | An Urgent Medical Device Correction (UDMC) was sent to all affected Siemens Healthcare Diagnostic customers in the United States on June 28, 2013 via FedEx. An Urgent Field Safety Notice (UFSN) was sent out for distribution to all customers outside the United States on June 27, 2013. FedEx Return Receipts will be used to confirm that the customers received the UMDC. If confirmation of receipt by any customer cannot be verified through a return receipt, a Siemens Technical Solutions Center Representative will contact those customers by telephone to ensure that they received the Urgent Medical Device Correction and completed the required actions. These notices inform the customer that they should discontinue the use of Calibrator A, lots ending in 77, for use with the ADVIA Centaur Systems T4 assay. Customers are informed that they may continue to use Calibrator A lot CA77 with the other assays FT3, FT4, Total T3, T-Uptake. For questions or additional information contact your local Technical Support Provider or Distributor. |  | 265 | 265 Cal A 6 packs (US and OUS) | Worldwide Distribution-Nationwide including PR and the states of AL, AR, AZ, CA, CO, CT, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, NC, NJ, NV, NY, OH, PA, TN, TX, VA, VT, and WV and the countries of Slovakia, Germany, Turkey, TW, HK, CH, TH, BR, CN, SG, and ZA. | Terminated | July 25, 2014 | 392 | 265 | Undetermined
Z-1971-2013 | 65126 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | epoc BGEM Test Card (Catalogue number CT-1004-00-00). |  |  | March 13, 2013 | 2013 | Epocal | This recall is being initiated because the affected lots of this product may yield results that are not consistent with stated precision performance criteria at concentration values at the high-end of the Glucose test sensor (&gt;400 mg/dl). Specifically, after a period of storage of three (3) to four (4) months, the Glucose sensor may report low results at the high end of the measurement range. | Nonconforming Material/Component | On March 13th 2013, Epocal placed all CT-1004-00-00 test card on shipment hold. Affected product in Epocal control has been segregated and quarantined. Beginning on March 26 2013, Epocal Inc. plans to begin distribution of the Recall Notification Letter to the 4 direct consignees (customers/distributors) via certified mail. Included with the Recall Notification Letter is a faxback verification form. Notification will include all distribution channels to the end-user level. Direct consignees are instructed to discontinue use and discard all stock on hand, complete the provided faxback verification form to document disposal and contact their consignees to advise them of the notification and disposal of affected product. |  | 828750 | 828,750 Test Card | Worldwide Distribution-USA (nationwide) including the District of Columbia, US Virgin Islands and Puerto Rico, and the states of CA, CO, CT, FL, GA, IA, IL, IN, KS, KY, LA, MD, MI, MN, MS, NC, ND, NE, NH, NJ, NM, NY, OH, PA, RI, SC, SD, TN, TX, UT, VA, VE, WA, WI, and WV and the coutries of India, Argentina, Australia, Azerbaijan, Belarus, Canada, China, Colombia, Egypt, Germany, Guatemala, Israel, Japan, Malaysia, Mexico, Nepal, Netherlands, Peru, Philippines, Russia Fed., Singapore, South Africa, South korea, Thailand, and Vietnam. | Open, Classified |  |  | 828750 | Undetermined
Z-1963-2013 | 65838 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Edwards Lifesciences Anastaflo Intravascular Shunt, Model code IVS12512, IVS1512, IVS17512, IVS2012, IVS2512, IVS3012, IVS1530, IVS2030 and IVS25 |  |  | July 30, 2013 | 2013 | Edwards Lifesciences, LLC | Edwards Lifesciences is recalling certain lots of the Anastaflo intravascular Shunt due to concerns of excessive adhesive on the shunt body that may interfere with suturing of by-pass grafts. | Process control | Edwards Lifesciences sent a Urgent Field Safety Notice letter dated July 30, 2013, via Fed-Ex to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letters are tracked through customer service and as part of the customer acknowledgement letter. Customers are asked to return products.If the customer does not send back the acknowledgement letter, three additional contacts are made to notify the customer and request return of affected product. This is tracked by Edwards customer service and reported each week. We reconcile the shipments made to customers against the returned product to determine impact to distribution. Once you have verified your inventory, please complete the attached acknowledgment form and fax it back to Edwards Customer Service at 800.422.9329 within three days of receipt of this Field Safety Notice. The return of this form allows us to confirm that you have reviewed this notice and have taken appropriate action. Please contact Customer Service at 800.424.3278 to obtain a Returned Goods Authorization number and replacement product.We sincerely regret the inconvenience caused by this action and appreciate your immediate attention to this matter. The Customer Service organization can answer questions about when replacement IVS Shunts will be available. If you have questions that have not been answered by this letter, please call Edwards Customer Service at 800.424.3278 from the hours of 6:00AM - 4:30PM PST or contact your Edwards sales representative concerning the recall. |  | 29618 | 29,618 units | Nationwide and the country of Canada | Terminated | October 14, 2014 | 441 | 29618 | Undetermined
Z-1979-2013 | 65798 | Yes | LWS |  |  | Cardiovascular | PMA | Boston Scientific, ENERGEN DR ICD, Model E143. The device is an Implantable Cardioverter Defibrillator. |  |  | July 15, 2013 | 2013 | Boston Scientific CRM Corp | Boston Scientific CRM manufacturing quality system recently discovered test artifacts (related to shock charge time) in a small number of implantable defibrillators (ENERGEN DR ICD model E143) which require further investigation. While these devices successfully passed all manufacturing tests and met all requirements, they showed a slightly longer charge time than other devices during a manufactu | Under Investigation by firm | Consignee was sent on 7/15/13 a Boston Scientific "Medical Device Retrieval" letter dated July 12. 2013. The letter described the product problem and informed the consignee that they will replace the product removed from their inventory. For further questions, contact Boston Scientific Technical Services at 1 800 227-3422. |  | 1 | 1 (4 devices were implanted and are not part of this action) | US distribution in the state of Tennessee. | Terminated | November 12, 2013 | 120 | 1 | Undetermined
Z-1980-2013 | 65552 | No | INO | Electric positioning chair. | Physical Medicine | Physical Medicine | 510(K) Exempt | CHAMPION Medical Recliner, AC Powered with heat and massage, one chair per carton, or one chair wrapped in a blanket if delivered in person. |  |  | June 18, 2013 | 2013 | Invacare Corporation | There is a possibility of elevated temperatures in the seat pad. In addition, the hospital cord set does not agree with the labeling (information in manual). | Process control | Champion Manufacturing plans to initiate this recall within approximately 10 business days after the submission of this report through a phone campaign and notification letters. |  | 285 | 285 units | Nationwide distribution: USA including states of: CA, FL, IN, MN, NJ, NY, OH, RI, SC, SD, TX, and VA. | Terminated | January 09, 2014 | 205 | 285 | Undetermined
Z-1984-2013 | 65908 | No | HQL | Intraocular lens. | Ophthalmic | Ophthalmic | PMA | Product Name: nanoPOINT Injector System, item number 37-0000-08. |  |  | August 05, 2013 | 2013 | Medicel Ag | The firm initiated this recall of one lot (121402) of nanoPOINT Injectors after they noticed that Tyvek lids might have been damaged and, therefore, the stability of the sterile barrier might be negatively affected. | Use error | Medicel AG sent recall notification to its one direct consignee via email on 08/05/2013. The recall notification letter was titled "URGENT-Medical Device Correction Notice for nanoPOINT Injectors LOT 121402". The recall notification letter informed the customer of the product device description and lot number recalled, reason for recall, instructions to "not use this affected lot...and return them", contact information and an area on the bottom of the letter to acknowledge receipt of the correction notice.The firm's direct consignee who further distributed this product to the end users will send out notifications. |  | 748 | 68 boxes or 680 units | Nationwide Distribution. | Terminated | March 11, 2014 | 218 | 748 | Undetermined
Z-1997-2013 | 65832 | No | KHE | Occult blood test. | Hematology | Hematology | 510(k) | Hemoccult ICT (Immunochemical Fecal Occult Blood Test) is a rapid, visually read, qualitative immunochemical chromatographic method for detection of human hemoglobin in fecal samples. |  |  | July 17, 2013 | 2013 | Beckman Coulter Inc. | Beckman Coulter has received a customer complaint that suggests trending ofpositive results has increased from approximately 4% to 1 0%. As a result of ourinvestigation, Beckman Coulter has confirmed an increase in sensitivity at theborderline positive level." This testing also demonstrated that the issue does not affect results reported asnegative. | Under Investigation by firm | Beckman Coulter sent an Urgent Product Correction letter dated July 9, 2013, to all affected customers. The letter was accompanied with a cover letter and response form, dated July 9, 2013. Customers were informed of the product under recall with product description, actions to take, contact information, the issue, impact, and resolution. Beckman Coulter conducted a risk assessment for this issue. Per the risk documentation, there is only a remote probability that use of the product will cause any adverse health consequences,and the product is not likely to lead to serious adverse health consequences.If you have any questions please call (714) 961-3777. |  | 68116 | 68,116 units | Worldwide Distribution - USA and the country of Canada. | Terminated | October 24, 2014 | 464 | 68116 | Undetermined
Z-1991-2013 | 65879 | No | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | ConvaTec Flexi-Seal SIGNAL +, Rx Only |  |  | June 25, 2013 | 2013 | Convatec Inc. | ConvaTec notified their customers on 6/25/2013 regarding the manufacturing irregularity that occurred related to the packaging of certain lots of Flexi-Seal Signal FMS. In certain lots, the lot number convention was duplicated on the packaging. | Under Investigation by firm | ConvaTec issued a notification letters to their customers dated June 25, 2013 via UPS. The letter informed the customer of the issue with the related product and the action to be taken. If the customer has any question regarding the notification they can contact customer service at 1-800-422-8811. |  | 22896 | 22,896 units | US Nationwide Distribution | Terminated | January 08, 2014 | 197 | 22896 | Undetermined
Z-1985-2013 | 65869 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Edwards Lifesciences Suction Wand Model S099B, packaged in Sorin Group Heart/Lung Perfusion Packs. |  |  | July 08, 2013 | 2013 | Sorin Group USA, Inc. | Sorin Group is recalling certain lots of Edwards Lifesciences Suction Wand Model S099B due to presence of plastic particles inside the pouch. | Material/Component Contamination | On 7/8/13, customers and distributors were sent a "Field Safety Notice Product Recall" letter and requested to return unused Edwards Suction Wands found in the Perfusion Packs to the Sorin Group. Customers are to contact Sorin Group Customer Service at 1-800-650-6502 to receive instructions for returning the affected Edwards Suction Wand to Sorin. Customers were also advised to not discard affected product. |  | 18927 | 18,927 packs | Nationwide Distribution | Terminated | January 28, 2014 | 204 | 18927 | Undetermined
Z-1986-2013 | 65644 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Medtronic MiniMed Paradigm Reservoirs, Model No: MMT-326A (1.8mL) and MMT-332A (3.0mL) |  |  | July 03, 2013 | 2013 | Medtronic MiniMed | Medtronic is recalling certain lots of Medtronic MiniMed Paradigm Reservoirs MMT-326A (1.8mL) and MMT-332A (3.0mL) used with Medtronic Paradigm insulin pumps because they may have increased risk for leaking. A leak in the reservoir may result in delivery of less insulin than intended. In addition, if there is a leaky reservoir and an insulin blockage occurs in the infusion set, the pump may not | Under Investigation by firm | An Urgent Medical Device Recall notice was sent to customers on July 3, 2013, who purchased the Medtronic Paradigm Reservoirs MMT-326A and MMT-332A. The recall letter informed the customers of the problems and the actions to be taken. Medtronic instructed customers that they will replace all reservoirs from the affected lots at no cost. Customers are instructed to refer to www.medtronicdiabetes.com/support/product-updates for a list of frequently asked questions. Customers are instructed to report adverse reactions or quality problems experienced with the use of Medtronic Diabetes insulin pumps, infusion sets or reservoirs to FDA's MedWatch Adverse Event Reporting Program. Customers are instructed to call for replacements at (866) 222-2584 Monday through Friday (8am-6pm) central time. Customers with an urgent issue may use the 24 hour helpline. |  | 669672 | 669,672 units | Worldwide Distribution - USA (nationwide) and the countries of Australia, Bermuda, Brazil, Canada, Chile, Columbia, Ecuador, El Salvador, Germany, Japan, Mexico, Netherland, Paraguay, Puerto Rico, Singapore, Uruguay and Venezuela. | Terminated | January 16, 2014 | 197 | 669672 | Undetermined
Z-1967-2013 | 65797 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | The Discovery NM/CT 570c CT Standalone Scan |  |  | April 10, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of an incident where the wrong part of the patient body wasscanned during a CT standalone exam with the Discovery NM/CT 570c | Radiation Control for Health and Safety Act | GE Healthcare will remedy the issue or bring the product into compliance with each applicable Federal standard. GE Healthcare has provided field modification instructions (FMI) to update software for the Discovery NM/CT570c hybrid system. The software update will correct the incorrect positioning of the table not prescribed by the operator. The FMI includes:A software upgrade with foreign language support to be installed in the affected models listed above,Labels (FMI label and packing label),Software upgrade workflow procedure with functional verification for all modes of operation, andCertification Report, Postponement of Medical Device Correction, Rejection of Medical Device Correction, or Refusal and/or Location Unknown Documentation.For further questions please call (262) 513-4122. |  | 10 | 10 Systems | US Distribution in the state of CT. | Terminated | January 21, 2014 | 286 | 10 | Undetermined
Z-1988-2013 | 65745 | No | FPX | Medical adhesive tape and adhesive bandage. | General Hospital | General &amp; Plastic Surgery | 510(k) | Distributed in USA: Product Number: 40-2894: Reinforced Wound Closure, Latex Free Sterile EO Rx, 1/4" X 4 ", 6 mm X100 mm, (2 Cards/pouch, 1,000 Pouches/box); |  |  | July 18, 2013 | 2013 | Aspen Surgical Products, Inc. | Aspen Surgical is recalling Wound Closures because firm could not confirm sterility, after FDA sampling indicated the product may not be sterile. Using the recalled product could lead to infection which may require further treatment including the use of antibiotics. There have been no reports of injuries or illnesses as of 7/18/2013. | Process design | Aspen sent an Urgent Medical Device Recall letter dated July 18, 2013 via certified mail to all customers with directions to forward the notification and provide Aspen with sub-account information. Aspen has given the customers the option to either 1.) Return the product 2.) Destroy the product with a Certificate of Destruction to be provided to Aspen. Any returned product to Aspen will be put on MRB and dispositioned by Quality. Questions or concens should be addresses to Kelli Jona Directo of QA / RA at 888-364-7004 ext 177 jonask@aspensurgical.com or Michelle Pulaski, Quality System and Compliance Manager 888-364-7004 ext 146 pulaskim@aspensurgical.com 8-5 EST M-F. |  | 3286 | 3286 boxes ( 6,353,000) | Worldwide Distribution - US Nationwide in the state of Florida and in the countries of , Netherlands, and Reditch Worcestershire England | Terminated | April 17, 2014 | 273 | 3286 | Undetermined
Z-1964-2013 | 65839 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Philips Healthcare MobileDiagnostwDR system. Mobile Digital radiography X-Ray System |  |  | June 11, 2013 | 2013 | Philips Healthcare Inc. | The Instructions for Use (IFU) for the MobileDiagnost wDR fails to comply with a Federal standard. The IFU does not provide tolerance levels for several generator values as required by 1020.30(h)(3). | Radiation Control for Health and Safety Act | Phillips Healthcare sent an Important Notification letter dated July 10, 2013 to all affected customer. The notification informed customers that Philips will implement an update to the Instruction for Use (IFU). Philips will, without charge, remedy the defect or bring the product into compliance with each applicable Federal standard in accordance with a plan to be approved by the Secretary of Health and Human Services. For questions contact your local Philips representative or local Philips Healthcare office. |  | 240 | 240 systems US -- 1 system within Puerto Rico | US Nationwide Distribution including Puerto Rico | Terminated | January 12, 2017 | 1311 | 240 | Undetermined
Z-1994-2013 | 65866 | No | FRC | Sterilization process indicator. | General Hospital | General Hospital | 510(k) | Getinge Assure Accufast Biological Indicators, Catalog Number 61301606636 (50 Pack) and 61301606637 (100 Pack), STEAM --- Manufactured for: Getinge USA, Inc. 1777 E. Henrietta Rd., Rochester, NY 14623. |  |  | August 02, 2013 | 2013 | Getinge USA Inc | AccuFast Biological Indicator Lot SR-430 may have been exposed to conditions affecting its viability. This could result in a false negative (or pass) when exposed to sub-lethal sterilization parameters. | Under Investigation by firm | GETINGE sent an DEVICE CORRECTION NOTICE letter dated July 31, 2013, and response forms were sent to customers on August 2, 2013, via Certified Mail Return Receipt Requested. The letters instruct customers to immediately discontinue use, quarantine, and destroy the affected lot. In addition, they instruct customers to: Check the appropriate box and record the quantity of affected unused product on the enclosed Correction Response Form OR If if they do not have any lots of affected unused inventory check the box indicating such. Destroy the unused product in accordance with your institution's policy and instructions for use. Record the method of destruction on the Correction Response Form. Be sure to include the Facility Name, Address, Contact Name and Phone Number. Sign, date and e-mail the form back to SR430responseform@getinge.com within 10 days or sooner of receiving this notice. Credit will be issued upon receipt of the Correction Response Form. Replacement product of unaffected lots is currently available. The warehouse shipping the affected lot is no longer being used (due to business decision, not as result of this recall). The current warehouse is maintaining the product in a controlled temperature segregated room away from any potential sterilants.If you have any questions regarding this notification please call 1-800-475-1400 ext. 5026 |  | 78 | Catalog # 61301606636: 78 Packs (3900 BIs); Catalog # 61301606637: 43 Packs (4300 BIs) | Worldwide Distribution - USA and the country of Canada. | Open, Classified |  |  | 78 | Undetermined
Z-1992-2013 | 65888 | No | KZF | AC-powered medical examination light. | General Hospital | General Hospital | 510(K) Exempt | Stryker Visum 300 Ceiling-mounted Exam Light. |  |  | July 25, 2013 | 2013 | Stryker Communications | Potential incomplete final weld on the drop tube flange. | Process control | The firm decided to recall and notified consignees with an "Urgent Medical Device Correction/Removal" notification letter dated 07/29/13. The product issue was described and recommended actions were provided. Customers may contact Stryker Technical Support at 866-841-5663. |  | 1375 | 1,375 units | Worldwide Distribution: US natiowide and to the country of Canada. | Terminated | April 03, 2015 | 617 | 1375 | Undetermined
Z-1976-2013 | 65777 | No | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Stryker Oasys Midline Occiput Plate-1) Small 2) Medium 3) Large 4) Large Long |  |  | May 30, 2013 | 2013 | Stryker Spine | Stryker has received reports from customers indicating post-operativefracture of the pin that connects the tulip head to the plate body. | Other | Stryker Spine sent a 'Urgent Medical Device Removal Letter'/'Customer Response Form' to Branch Managers, Agent Principals, Quality Contacts, Hospital OR Risk Managers, OR supervisor on 5/30/2013 via Fed Ex. Inventory should be examined and returned.The attached form should be completed and returned as well.Stryker Spine sent 'Urgent Medical Device Removal Letters' to Spinal Implant Surgeons on 6/20/2013 via Fed Ex. |  | 1536 | 1536 units (US) 880 (Foreign) | Nationwide Distribution. | Terminated | December 30, 2014 | 579 | 1536 | Undetermined
Z-1693-2013 | 65090 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius Service Replacement 0-Ring. |  |  | May 24, 2013 | 2013 | Fresenius Medical Care Holdings, Inc. | Incorrect rubber nitrile O-rings distributed instead of the correct material EPDM. (Ethylene Propylene Dine Monomer (M-class) rubber) | Nonconforming Material/Component | Fresenius notified customers on May 24, 2013, by telephone and by certified mail with signature confirmation and fax-back form. Customers were instructed to examine their stock immediately to determine whether they have any of the affected parts on hand. If customers have the affected parts, they will be instructed to immediately discontinue use and place all 0-rings in a secure area for return to Fresenius Medical Care North America.Please contact your Fresenius Medical Care Technical Service Team at 1-800-227-2572, (choose option 4, then option 2) for instructions how to return the recalled product. |  | 725 | 725 units | Nationwide Distribution including the states of AR, CA, FL, GA, KY, LA, MA, MD, NC, NJ, NY, PA, PR, SC, TX and VA. | Completed |  |  | 725 | Undetermined
Z-1692-2013 | 65554 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | GEO-MED ORTHO TOTAL JOINT TRACECART(R), REF 53-1831.07, Rx Only, NON-STERILE |  |  | May 30, 2013 | 2013 | DeRoyal Industries Inc | DeRoyal recalled kits that contained Stryker togas which were recalled for a potential problem with sterilization. | Nonconforming Material/Component | DeRoyal sent an Urgent: Voluntary Product Recall letter dated May 30, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to examine inventory, remove any affected product from their inventory, open its package to prevent use, and dispose in a standard trash receptacle. Customers were asked to complete the enclosed notice of destruction form and return it to DeRoyal by faxing it to 865-362-3716 or emailing it to recalls@deroyal.com. For question call 865-362-1034 or contact your DeRoyal representative. |  | 87 | 87 units | USA Nationwide Distribution in the state of PA | Terminated | July 12, 2013 | 43 | 87 | Undetermined
Z-1598-2013 | 65049 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | TOSHIBA Kalare Diagnostic X-Ray System |  |  | April 03, 2013 | 2013 | Toshiba American Medical Systems Inc | Firm initiated recall because when radiographic conditions with the small focus are set and the radiography ready switch is held down, the large focus is also heated after a few seconds or more. This occurs when the Automatic Exposure Function (AEC) function is used. As a result, a radiographic tube current exceeding the set value flows and a brighter image than expected is acquired. | Component design/selection | Firm initiated recall by sending recall notification letters along with a customer reply form on April 3, 2013 via USPS return mail receipt. A follow-up letter (with customer reply form) was sent on April 11, 2013 to inform customers that they should continue to use the AEC function as usual. The updated letter along with the customer reply form dated April 11, 2013 was also sent via USPS return mail receipt. The letter titled, "URGENT: MEDICAL DEVICE CORRECTION", notified customers of the reason for recall, product description under recall, corrective action, instructions to customer under 1. A new circuit board will be installed in each involved unit by a Toshiba service representative, without any charge to the user.2. In the meantime while waiting for service, the user is to continue usingthe device with the AEC function. If the AEC function is not available, theuser should push the radiography READY and EXPOSURE switches atthe same time."Note to Customer", and contact information "(800) 421-1968 or your local Toshiba Representative at (800) 521-1968". |  | 153 | 153 units | US Nationwide Distribution | Terminated | February 05, 2016 | 1038 | 153 | Undetermined
Z-1440-2013 | 65109 | No | EPT | Surgical microscope and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Zeiss brand OPMI Pentero C, Surgical microscope and accessories, Model: OPMI Pentero C, Product Numbers: 302583-9001-000 (PENTERO C), and 305961-0000-00 (Ceiling Mount for PENTERO C), |  |  | April 18, 2013 | 2013 | Carl Zeiss Meditec, Inc. | Carl Zeiss Meditec is conducting a field corrective action for its OPMI Pentero C surgical microscope system. During a routine inspection, one microscope installation was found to have 3 (three) broken screws out of 6 (six) in the lower section of the ceiling mount. | Under Investigation by firm | Carl Zeiss Meditec, Inc sent a Customer Notification letter to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter stated that firm's sales representative will be scheduling an on-site visit, informing them of the affected product and providing instructions on the recall. For question call 925-580-5371. |  | 6 | 6 units | US Nationwide Distribution. | Terminated | June 26, 2013 | 69 | 6 | Undetermined
Z-1408-2013 | 65150 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Merit Medical PerfOrma Anglographic Catheter 5F Multipack |  |  | May 02, 2013 | 2013 | Merit Medical Systems, Inc. | Merit Medical Systems is recalling various lots of Performa Cardiac Multipacks containing a JR4.0 catheter due to a slight variation in tip shape. | Process control | Merit sent an Urgent Product Recall Notice letter dated May 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were asked to identify, quarantine, discontinue using products, and contact their sales representative to arrange product return and replacement or credit. For questions call 1-801-208-4468. |  | 1465 | 1,465 | USA Nationwide Distribution. | Terminated | July 22, 2013 | 81 | 1465 | Undetermined
Z-1403-2013 | 65067 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | PTS PANELS, Glucose Test Strips for use with CardioChek Brand Analyzers |  |  | April 04, 2013 | 2013 | Polymer Technology Systems, Inc. | Internal investigation demonstrated under-recovery for glucose when tested against a reference method. This could result in a low to moderate health risk. | Under Investigation by firm | Polymer Technology Systems sent a Field Correction Action letters dated April 4, 2013 to distributors and customers. The letter identified the affected product, problem and actions to be taken. The letter directed customers to discontinue use and dispose of remaining inventory of recalled product. Customers were asked to acknowledge and confirm recall directions were follwed by fax to 1-317-870-5608. |  | 7048 | 7048 total, 3499 in USA | Worldwide Distribution - USA Nationwide including Puerto Rico and countries of: UK, CHL, ENG, BEL, ITA, TKY, BEL, LIT, LTV, EGY, SPG, HOK, SWE, MEX, UAE, COR, ISR, MAL, AUS, POR, POL, SLN and STQ. | Terminated | April 16, 2014 | 377 | 7048 | Undetermined
Z-1402-2013 | 65060 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Integrity 3.0 |  |  | April 30, 2013 | 2013 | Elekta, Inc. | Unexpected movement during beam delivery may not be detected if there are incorrect static tolerances for Linac Geometric Axes loaded in Integrity R3.0. | Nonconforming Material/Component | Elekta sent an Important Field Safety Notice (IFSN) letter dated April 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers are instructed to complete and return the FCO Action Notification Report to local Elekta Office or Representative. |  | 31 | 31 | USA Nationwide including the states of CA, FL, IL, MD, MA, NJ, NY, NC, ND, OH, OR, PA, TN, TX, VA, WA, DC and Puerto Rico | Terminated | July 26, 2017 | 1548 | 31 | Undetermined
Z-1372-2013 | 65070 | No | CCQ | Medicinal nonventilatory nebulizer (atomizer). | Anesthesiology | Anesthesiology | 510(K) Exempt | AsthmaNefrin Starter Kit contains ten (10) AsthmaNefrin vials and one (1) EZ Breathe Atomizer. |  |  | April 30, 2013 | 2013 | Nephron Pharmaceuticals Corp. | Possible dislodgement of the "Plate A", if this instance occurs, it presents a potential choking hazard. | Nonconforming Material/Component | The firm, Nephron Pharmaceuticals Corporation, notified its consignees on April 30, 2013, and May 3, 2013, of the problem via a recall notice and a Press Release, respectively. The firm directed consignees to discontinue use of any EZ Breathe atomizer units contained in the Asthmanefrin Starter Kits with the affected lot numbers and also those that were sold individually. Customers were directed to return all affected product. |  | 241538 | 241,538 kits | Worldwide Distribution-USA (nationwide) including DC and Puerto Rico and the states of AL, AZ, CA, CO, CT, FL, IL, IN, IO, GA, KS, KY, LA, MD, MN, MS, MO, NK, NV, NJ, NM, NY, OH, OK, OR, PA, RI, SC, TN , TX, VT, VA, WA, and WV, and the countries of Guatemala, Poland, Panama, and United Kingdom. | Terminated | June 28, 2016 | 1155 | 242892 | Undetermined
Z-1423-2013 | 65089 | No | JTO | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | BBL DrySlide Oxidase Kit, catalog number 231746 |  |  | April 12, 0013 | 13 | Becton Dickinson &amp; Co. | In-vitro diagnostic agents were placed into incorrect packaging for distribution to microbiological laboratories. | Mixed-up of materials/components | BD Diagnostic Systems sent an Urgent Product Recall" letter dated April 2013 to all affected consignees and end users. The letter identified the affected product, problem and actions to be taken. The notice instructed consignees and end users to discontinue use of the products, discard remaining products, and return a completed response form. Consignees were requested to provide contact information for end users who received recalled product in distribution, to allow for recall notice. The letter included instructions to report injuries and/or illnesses to Medwatch, and address recall inquiries to BD Technical Services at 1-800-638-8663. |  | 610 | 610 total test kits (PYR and Oxidase) | Worldwide Distribution - USA Nationwide and tthe countries of Australia, Canada, Europe, Hong Kong, Mexico, New Zealand, Singapore, and Taiwan. | Terminated | September 28, 2013 | 730654 | 610 | Undetermined
Z-1430-2013 | 65118 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AU2700 Clinical Chemistry Analyzer |  |  | April 26, 2013 | 2013 | Beckman Coulter Inc. | Beckman Coulter is recalling AU2700/AU2700 Plus &amp; AU5400 Clinical Chemistry Analyzers becuase it was determined that the tubing/joint connections associated with these MODs may leak concentrated detergent. Beckman Coulter has received complaints that the joint connecting the tubing appears to disintegrate over time, causing a leak. | Component design/selection | Beckman Coulter sent an Urgent Product Correction letter on April 26, 2013 along with a response form. The letter notified customers of the product, problem, and actions to be taken by the customers. Contact the Call Center at 800-854-3633 for questions regarding this notice. |  | 108 | 108 units | Worldwide Distribution-USA Nationwide and countries of: AU, CA, DE, ES, GB, IT, PT, and JP. | Terminated | August 04, 2014 | 465 | 108 | Undetermined
Z-1437-2013 | 49033 | No | BZD | Noncontinuous ventilator (IPPB). | Anesthesiology | Anesthesiology | 510(k) | Fisher &amp; Paykel Healthcare ICON CPAP (continuous positive airway pressure). |  |  | April 22, 2013 | 2013 | Fisher &amp; Paykel Healthcare, Ltd. | Fisher and Paykel Healthcare is recalling the ICON CPAP units because there is a potential power issue on certain ICON CPAP. These ICON CPAP units could potentially become inoperable, which would result in a lack of treatment to the patient's obstructive sleep apnea symptoms until a replacement unit is obtained. | Nonconforming Material/Component | The firm, Fisher &amp; Paykel Healthcare Limited, sent a "Medical Device Recall" letter dated April 22, 2013 to all their customers who purchased the Fisher &amp; Paykel Healthcare ICON CPAP Replacement. The recall letter described the product, notified the customers of the problem identified and the actions to be taken. Customers are instructed to locate and verify if any of the affected products are in stock; check the part and lot numbers of your device; complete and return the enclosed Product Recall Response form, even if you have not received or distributed any affected products via Email: FA-2013-001@fphcare.com; Fax: (949) 453-4001; or Post: Fisher &amp; Paykel Healthcare,15365 Barranca Parkway, Irvine, CA 92618; transfer this notice immediately to all those who need to be aware and if affected product has been distributed to any other organizations or distributors, notify them regarding the recall within 10 business days upon receipt of this letter; prepare any products to be returned to Fisher &amp; Paykel Healthcare and contact the field action coordinator via email at FA-2013-001@fphcare.com or directly at 1-800-792-3912 ext. 1426 to arrange for a return shipping label to be sent to you. If you have any questions related to this recall, contact the field action coordinator via email at FA-2013-001@fphcare.com or directly at (800) 792-3912 ext. 1426. |  | 1407 | 1,407 units | Worldwide distribution: USA (nationwide) and countries of: Canada, Japan, and Taiwan. | Terminated | October 18, 2013 | 179 | 1407 | Undetermined
Z-1438-2013 | 65120 | Yes | DRO | External transcutaneous cardiac pacemaker (noninvasive). | Cardiovascular | Cardiovascular | 510(k) | adult defibrillation, cardioversion, pacing, monitoring Disposable Electrodes, Model P-211-M1 and P-214-M1. |  |  | April 30, 2013 | 2013 | TZ Medical Inc. | TZ Medical Inc is recalling its P-211-M1 and P-214-M1 Medtronic / Physio Control Compatible transparent defibrillation electrodes because it may contain a defective white wire on the connector assembly.If a failure were to occur, the end user would be able to identify the failure by inspecting the white wire 1/4 inch from the Gray Medtronic / Physio adapter. The consignees will notice discolor | Process control | TZ Medical Inc. started notifying their customers by telephone on April 30, 2013. The firm followed up with an Urgent: Medical Device Recall notification letter dated May 14, 2013, to their consignees. TZ Medical Inc., Representatives have will contact all affected customers and arrange for the replacement of the effected product in stock with inspected product. For any questions about this recall, please fax to 503-639-0239. |  | 10510 | US: 10,410 pieces; Outside US: 100 pieces | Worldwide distribution - USA (nationwide) and Chile. | Terminated | September 05, 2013 | 128 | 10510 | Undetermined
Z-1435-2013 | 65085 | Yes | LDF | Cardiovascular permanent or temporary pacemaker electrode. | Cardiovascular | Cardiovascular | 510(k) | Vermed, Inc Radiotransparent Defibrillation Electrodes |  |  | April 26, 2013 | 2013 | Heart Sync, Inc | On 4/18/13 HEARTSYNC became aware of a recall from their wire/connector sub-assembly supplier which is used in the C100-Physio Electrodes. There were reports of equipment damage due to misalignment from tooling installed in August 2012. The result of this damage could render the device incapable of delivering the appropriate energy or shock to the patient. | Process design | HeartSync sent an URGENT MEDICAL DEVICE RECALL notification letter dated April 26, 2013, to all consignees of Heart Sync for C100-Physio Radiotransparent Defrillation Electrodes. On April 29, 2013, an URGENT MEDICAL DEVICE RECALL notification was sent to Vermed Inc for the D304 Defrillation Electrodes . The letters identified the product, the problem, and the action to be taken by the customer. Customers were instructed to return the affected product to the firm for full credit. Customers were asked to complete the form and send back via fax to 734-213-5640 or call 734-213-5530.For questions regarding this recall call 734-213-5530.. |  | 17180 | 17180 units for all part numbers | Worldwide Distribution - USA including AL, GA, MI, MN, NC, ND, OH, PA, SD, VA, VT, OH, TX, MA, and Internationally to Korea. | Terminated | April 16, 2014 | 355 | 17180 | Undetermined
Z-1442-2013 | 65216 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | 8 mm Monopolar Curved Scissors (a.k.a. Hot Shears) |  |  | May 08, 2013 | 2013 | Intuitive Surgical, Inc. | Intuitive Surgical has identified a potential for some units of the Monopolar curved Scissors version 09 and 10 instruments to develop very small cracks near the distal end. This may create a pathway for electrosurgical energy to leak to tissue and potentially cause thermal injury. | Nonconforming Material/Component | Intuitive Surgical sent an Urgent Medical Device Recall letter dated May16, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised that they did not need to quarantine or return the affected device at that time. a Intuitive Surgical representative would contact them to arrange for replacement of the affected product once they became available. In the interim customers were instructed to follow the precautions and warnings outlined in Attachment A in the letter as well as the Instructions for Use. Customers were also asked to ensure all affected personnel were fully informed of the contents of the Recall Notification. The letter should be forwarded to other surgeons and/or managers within their facility. Customers were also asked to complete and return the attached Acknowledgement Form. For further questions customers were instructed to contact Customer Service:North and South America: 800-876-1310 Option 3Japan: 0120-56-5635 or 003-5575-1362Korea: 02-3271-3200Europe, Middle East, Asia and Africa +800 0821 2020 or +41 21 821 2020For questions regarding this recall call 408-523-2244. |  | 132441 | ~132,441 instruments | Worldwide distribution - USA (nationwide) Foreign Distribution to Argentina, Austria, Australia, Belgium, Brazil, Bulgaria, Canada, China, Chile, Czech Republic, Denmark, Finland, France, Germany, Cyprus- Greece, Egypt, Guadeloupe, Israel, Ireland, India, Indonesia, Italy, Japan, Kuwait, Lebanon, Malaysia, Mexico, Monaco, Netherlands, New Zealand, Norway, Pakistan, Philippines, Portugal, Poland, Qatar, Russia, Romania, Slovenia, South Korea, Spain, Sweden, Switzerland, Thailand, Taiwan, Turkey, United Kingdom, Saudi Arabia, Singapore, South Korea, and Venezuela. | Terminated | May 01, 2014 | 358 | 132441 | Undetermined
Z-1376-2013 | 65042 | No | HTW | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | NEXGEN Complete Knee Solution Monoblock Tibial Drill with stop 10.7 mm. |  |  | March 27, 2013 | 2013 | Zimmer, Inc. | Zimmer is initiating a lot specific recall of the Trabecular Metal" Tibia Stop Drill due to the potential of the outside diameter being oversized. As a result, there is a potential for the drill to not pass through the applicable drill guide. | Process control | Zimmer sent an "URGENT MEDICAL DEVICE RECALL LETTER" dated April 19, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. |  | 69 | 69 units | Nationwide Distribution. | Open, Classified |  |  | 69 | Undetermined
Z-1439-2013 | 65255 | No | LRC | Ear, nose, and throat manual surgical instrument. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(K) Exempt | GYRUS ACMI, MYRINGOTOMY KNIFE DISPOSABLE, SICKLE BLADE FLAT HANDLE, REF 130715, Rx only, STERILE EO, GYRUS ACMI, INC., 136 TURNPIKE ROAD, SOUTHBOROUGH, MA 01722-2104 |  |  | May 24, 2013 | 2013 | Gyrus Acmi, Incorporated | lack of sterility assurance | Employee error | Cyrus ACMI notified their direct accounts of the recall by Priority Mail with delivery confirmation on May 24, 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to inspect their inventory and immediately cease any further use of any affected product they may have, remove it from inventory, and quarantine it until they ship it back to the firm. For questions customers were instructed to contact their Olympus customer service representative at 1-800-848-9024 to obtain a Returned Goods Authorization so that they may return the product with no charge to them. Customers were asked to note on the enclosed questionnaire that they have received this information and fax it to 484-896-7128.For questions regarding this recall call 508-804-2739. |  | 600 | 100 boxes of 6 units | Worldwide Distribution - USA (nationwide) and countries of: Canada and United Kingdom. | Terminated | May 30, 2014 | 371 | 600 | Undetermined
Z-1369-2013 | 64803 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Shimadzu Corporation - Mobile DaRt Evolution |  |  | March 18, 2013 | 2013 | Shimadzu Medical Systems | The recall was initiated because Shimadzu Corporation has identified a potential issue in certain manufacturing lots of the Shimadzu MobileDaRt Evolution in combination with CXDI-70C, CXDI-55C, and CXDI-55G. As a result of Shimadzu's investigation, there is a potential risk for screws becoming loose inside the Shimadzu Mobile X-ray systems. | Radiation Control for Health and Safety Act | Shimadzu sent a recall letter dated March 18, 2013, to all customers who purchased the Shimadzu MobileDart Evolution in combination with CXDI-70C, CXDI-55C, or CXDI-55G. The letter informs the customers of the problem identified and the actions to be taken. Customers are instructed to follow the instructions of the recall notice step by step. Customers with questions are instructed to contact their local Shimadzu representatives or contact Shimadzu National Technical Support at (800) 228-1429.A Field Safety Notice dated March 18, 2013, was sent to Service Managers in regards to the Shimadzu Mobile DaRt Evolution in combination with CXDI-70C, CXDI-55C, or CXDI-55G. The Safety Notice informs the Service managers of the problems identified and the actions to be taken. |  | 200 | 200 units | US Nationwide Distribution | Terminated | February 13, 2014 | 332 | 200 | Undetermined
Z-1443-2013 | 65164 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | Canister: CO2 canister with GE part numbers 1407-3200-000, 1407-7004-000, or M1084850. |  |  | May 14, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue involving the reusable CO2 absorbent canister accessory used with Aespire, Aespire View, Avance, Avance CS2, Aisys, ADU, 9100, M900, M904, 9300 systems. Reusable CO2 absorbent canisters accessories may not seal properly. | Component change control | GE Healthcare sent an "URGENT MEDICAL DEVICE CORRECTION" letter April 30, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. A Fax back form was included with the letter requesting the consignees to complete and return. |  | 800 | 800 | Worldwide Distribution-USA (nationwide) including the states of AZ, CA, CO, CT, ID, MK, MI, NJ, NY, OH, TN, TX, and VA, and the countries of INDIA, AUSTRALIA, BAHRAIN, BANGLADESH, BOLIVIA, BOSNIA &amp; HERZEGOWINA, BRAZIL, CANADA, CHILE, CHINA, COLOMBIA, CZECH REPUBLIC, EGYPT, EL SALVADOR, FINLAND, FRANCE, GERMANY, HONG KONG, HUNGARY, INDONESIA, IRAQ,ISRAEL, ITALY, JAPAN, REPUBLIC OF KOREA,LEBANON, MALAYSIA, MEXICO, NETHERLAND,NEW ZEALAND, OMAN, PARAGUAY, PERU., PHILIPPINES, RUSSIAN FEDERATION, SAUDI ARABIA, SERBIA, SINGAPORE, SLOVAKIA, SOUTH AFRICA, THAILAND, TURKEY, UNITED ARAB EMIRATES , UNITED KINGDOM, and VIET NAM. | Terminated | March 12, 2014 | 302 | 800 | Undetermined
Z-1444-2013 | 64826 | No | BSO | Anesthesia conduction catheter. | Anesthesiology | Anesthesiology | 510(k) | I-Flow ON-Q Pain Relief System, T-Block Needle/Catheter Tray, Model #TBTB1089T |  |  | June 01, 2012 | 2012 | I-Flow LLC | I-Flow initiated two voluntary product removals of the ON-Q Pain Relief System: T-Block Needle/Catheter Tray and T-Block Needle/Catheter Kit because they may have potentially malfunctioned. | Use error | IFlow sent the customer notification letter on May 29, 2013, to 28 customers via FedEx along with a copy of the Service Guide to inform customers of the issue that can cause the Sting Ray to disconnect from the Catheter and how to prevent this problem from occuring. Customers are informed that in certain circumstances, ultrasound gel on the catheters significantly reduces the amount of force required to pull the catheter out of the Stingray connector. The In Service Guide clarifies that the method of catheter securement. It also instructs that wiping off any excess ultrasound gel from the catheter prior to inserting it into the Stingray connector will help with connecting the catheter to the Stingray connector. Customers are instructed that there is no action required for the I Flow ON Q Pain Relief System and ON Q T Bloc. The recall letter is only to provide further instructions on successful set up of the SPIROL catheter with the Stingray Connector included in the ON-Q T-bloc trays. Customers with any questions are instructed to call IFlow Customer service at (800) 448-3569 or IFlow 24 hr Tech support at (800) 444-2728.On June 1, 2012, an I-Flow sales representative removed the catheters and connectors from the one consignee/customers.For questions regarding this recall call 949-206-2700. |  | 450 | 450 units | Nationwide distribution: USA including states of: CA, CT, FL, IL, IN, KY, NY, OH, OR, TN, TX, VA and WI. | Terminated | June 25, 2013 | 389 | 525 | Undetermined
Z-1420-2013 | 64631 | No | ETN | Surgical nerve stimulator/locator. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | REF 8229705 x1 NIM TriVantage EMG Endotracheal Tube 5.0 mm I.D. x 6.5 mm O.D. Medtronic Xomed, Inc., 6743 Southpoint Drive North, Jacksonville, FL 32216-0980 USA STERILE Rx Only |  |  | March 04, 2013 | 2013 | Medtronic Xomed, Inc. | In March 2013, Medtronic issued a recall of the NIM TriVantage EMG Endotracheal Tube due to cuff leakage in the intraoperative cuff deflation which then requires re-inflation or replacement of the deflated tube in order to ensure the continued ventilation of the patient. | Use error | The firm, Medtronic Xomed, Inc. sent an "URGENT Product Recall Notification" letter dated March 14, 2013, to its U.S. and International customers. The letter described the product, problem and actions to be taken. The customers were instructed to return any affected devices still in their possession per the following procedure: 1. Check inventories for the affected products listed on the enclosed checklist. 2. Fill-in the "quantity on-hand" column on the checklist. 3. Fax checklist to Medtronic ENT at 1-904-296-2386. 4. If you have any of the subject product, contact Medtronic ENT Customer Service at 1-800-874-5797 to arrange for returns and replacement or credit. 5.When returning products, clearly mark the outside of the container Returned Goods Authorization (RGA) number. If you have any questions regarding the subject action or the content of the letter, contact Gabriela Anchondo at 904-279-7550. |  | 2591 | 2,591 units | Worldwide distribution: AL, AR, AZ, CA, CO, FL, GA, HI, IL, IN, KS, KY, MA, MD, MI, MN, MT, NC, NH, NJ, NY, OH, PA, RI, SC, TX, UT, VA, and WA; and countries of : Australia, Canada, Germany, Italy, Norway, Poland and United Kingdom. | Terminated | January 30, 2014 | 332 | 2591 | Undetermined
Z-1485-2013 | 65258 | No | GFD | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Zimmer Air Dermatome II, REF 00-8851-001-01 |  |  | May 07, 2013 | 2013 | Zimmer Surgical Inc | This removal is in response to complaints alleging that the Zimmer Air Dermatomes and Zimmer Air Dermatome II products would either not operate oroperate intermittently. Zimmer's investigation determined that the planetary gear teeth were broken. This may cause a delay or postponement of the surgery or result in a suboptimal graft retrieval with potential of requiring an additional donor site. | Process control | The firm, Zimmer, sent an "URGENT: DEVICE REMOVAL" notice dated May 7, 2013, to all consignees/customers via FED Ex and Distributors via E-mail. Consignees who had devices repaired received telephone calls as well as the mailed Recall Notice. The Notice identified the devices being recalled; possible adverse events and directions to immediately stop use of and remove affected product, and obtain an RGA (Return Goods Authorization) number by calling (800)348-2759 or email information to dover.repairs@zimmer.com . to return the devices for repair to:Zimmer Surgical, Attn: QA/RA Dept.- Recall, 200 West Ohio Avenue Dover, Ohio 44622 USA; and to complete and return the Response Form-Immediate Response Required via email to CorporateQualitv.PostMarket@zimmer.com. If you have further distributed affected product(s), please provide the customer's information on the Business Response Form to Zimmer.If you have any questions regarding this action, please email us at CorporateQuality. PostMarket@zimmer.com or call 1-330-364-0989. |  | 13 | 13 units | Worldwide distribution: USA Nationwide including Puerto Rico; UAE, TAIWAN, SWEDEN, THAILAND, KOREA, SINGAPORE, HONG KONG CHINA, CHILE, CANADA, AUSTRALIA, JAPAN, ENGLAND, GERMANY, and EGYPT. | Terminated | December 15, 2014 | 587 | 223 | Undetermined
Z-1466-2013 | 65257 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Hill-Rom 1000 Bed, Product No. P1160. The Hill-Rom 1000 Bed is intended for low to moderate acuity patient in the medical surgical area of the hospital. |  |  | April 15, 2013 | 2013 | Hill-Rom, Inc. | Complaints have been reported for brakes disengaging when the bed is subjected to a hard jolt or force to the frame. | Process control | On 4/16/2013 an URGENT MEDICAL DEVICE CORRECTION notification letter was sent to all consignees. The firm plans a visit to all facilities to correct each bed. Call Hill-Rom Technical Support at 800-445-3720 if you have any questions. |  | 8737 | 8,737 devices | Worldwide distribution: US (Nationwide) and countries of: Argentina , Bahamas, Bahrain, Bolivia, Brazil, Canada, Chile, China, Colombia, Costa Rica, Ecuador, Egypt, Ethiopia, France, Ghana, Guam, Hong Kong, India, Indonesia, Japan, Jordan, Kuwait, Lebanon, Libya, Malaysia, Mexico, N. Mariana Islands, Pakistan, panama, Paraguay, Puerto Rico, Qatar, Saudi Arabia, Singapore, South Africa, South Korea, Taiwan, Thailand, United Arab Emirates, United States, US Virgin Islands, Venezuela and Vietnam. | Open, Classified |  |  | 8737 | Undetermined
Z-1481-2013 | 65140 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | BrightView SPECT gamma camera for Emission Computed Tomography, Philips Medical Systems, Cleveland, OH 44143 |  |  | April 25, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | Philips has received one report from the field that there was an unexpected motion of the detector due to a failure of the mechanical assembly. | Other | Philips Healthcare sent an Urgent Medical Device Correction letter dated April 26, 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customer were not required to take any specific action. Customers who need further information or support concerning the affected product, they should contact their local lPhilips representative or local Philips Healthcare office. For North Americ and Canada they should contact the Customer Care Solutions Center (1-800-722-9377, option 5).For questions regarding this recall call 440-483-7600. |  | 505 | 505 units | Worldwide Distribution - USA including AK, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NJ, NM, NV, NY, OH, OK, OR, PA, PR, SC, SD, TN, TX, VA, WA, WI, WV &amp; WY.Internationally to Australia, Austria, Bahrain, Belgium, Brazil, Canada, China, Czech Republic, Denmark, Dominican Republic, Egypt, France, Germany, India, Indonesia, Iran, Italy, Japan, Japan, Korea, Lebanon, Lithuania, Malaysia, Mexico, Portugal, Russian Federation, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Syria, Thailand, Turkey, Taiwan, Ukraine, United Arab Emirates, United Kingdom, Venezuela &amp; Viet Nam. | Terminated | February 12, 2014 | 293 | 505 | Undetermined
Z-1463-2013 | 65168 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | da Vinci 5 mm Instrument Cannulae for use with the da Vinci S and Si Surgical Systems. |  |  | May 06, 2013 | 2013 | Intuitive Surgical, Inc. | Intuitive Surgical is implementing a voluntary Field Action to implement a modified inspection process for the 5 mm Cannula (PN 420011) and 5 mm Flared Cannula (PN 420262). | Use error | Intuitive Surgical sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated May 14, 2013 to all affected customers. The letter identified the product,problem, and actions to be taken by the customers. Contact Customer Service at 800-876-1310 for questions regarding this notice. |  | 2101 | 2101 consignees | Worldwide distribution: US (nationwide) and countries of: Austria, Belgium, Denmark, Finland, France, Germany, Ireland, Netherlands, Norway, South Korea, Sweden, Switzerland, and United Kingdom. | Terminated | April 15, 2014 | 344 | 2101 | Undetermined
Z-1400-2013 | 63367 | No | KHP | Lactic acid test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Synchron LX System(s) and UniCel DxC 600/800 System(s) Lactate Reagent, Part Number: A95550 |  |  | April 17, 2013 | 2013 | Beckman Coulter Inc. | Beckman Coulter initiated the recall for the Lactate Assay (LACT) Lot M206209 as they received a customer complaint of error message (RXN RATE HI), i.e. no numerical values reported for samples containing Lactate concentration greater than 6.5 mmol/L. | Under Investigation by firm | Beckman Coulter sent an Urgent Product Correction letter dated April 17, 2013 to all affected customers. The letter identified the affected product, problem actions to be taken. Customers are instructed to complete and return the enclosed Response Form within 10 days. Customers with questions are instructed to contact Beckman Coulter Customer Support Center (Call Center) on the Beckman Website at http://www.beckmancoulter.com/customersupport/support/ or call (800) 854-3633 in the US and Canada. |  | 8774 | 8,774 units | Worldwide Distribution - USA Nationwide and the countries of Canada, Australia, Belgium, Czech Republic, France, French Polynesia, Germany, Greece, Hong Kong, Ireland, Italy, Japan, Jordon, Kuwait, Lebanon, Malawi, Morocco, Namibia, Netherlands, oman, Panama, Paraguay, Qatar, Reunion, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom | Terminated | October 22, 2014 | 553 | 8774 | Undetermined
Z-1217-2013 | 65036 | No | EHD | Extraoral source x-ray system. | Dental | Radiology | 510(k) | MyRay RXDC dental unit |  |  | August 10, 2012 | 2012 | CEFLA DENTAL GROUP | The RX DC dental x-ray unit was found to be noncompliant with 21CFR 1010,2, 1020.30(e) &amp; 1020.30(h). Specifically, 1. The tube housing did not contain a Certification label.2. The collimator did not contain an identification label.3. The user's manual did not contain leakage technique factors, tube housing cooling curves and rating charts, and technique factors which lead to operation at ma | Radiation Control for Health and Safety Act | The firm, CEFLA Dental Group will update the user's manual to include the information required by 21 CFR 1020.30(h). Future production will ensure that system labels required by 21 CFR 1010.2, and 1020.30(e) are properly applied including the 53 units manufactured awaiting shipment. The updated user's manual will be made available online for free download. Alternatively, system purchasers, subsequent transferees, dealers and distributors may elect to receive a hard copy of the updated user's manual at no charge. All corrections will be made at no cost to the purchaser, subsequent transferee, dealer, or distributor.The above actions will be completed within 4 weeks. For further questions please call 1-800 323-2690. |  | 173 | 173 shipped to US | FOREIGN: CANADA; DOMESTIC: GA, TX, IL, CA, FL, NY, MA, AL, NJ, SC, WA, LA, MD, WY, CT, NM, WI, CO, NC, VA, PA and MI. | Terminated | March 04, 2015 | 936 | 173 | Undetermined
Z-1505-2013 | 65098 | No | OJU | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | FAST Procedure Kit. |  |  | April 30, 2013 | 2013 | American Optisurgical Inc | American Optisurgical, Inc initiated a recall of ChloraPrep Hi-Lite Orange 3mL Applicator, in response to a sub-recall letter from Medline Industries, Inc due to lack of sterility assurance caused by an incomplete seal on the individualunit packaging. | Nonconforming Material/Component | American Optisurgical Inc. sent an "URGENT: VOLUNTARY MEDICAL DEVICE RECALL AND PRODUCT EXCHANGE" letter dated May 17, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at 949-580-1266 for questions regarding this notice. |  | 63 | 63 individual applicators | Nationwide Distribution-USA (nationwide) including the states of AZ, CA, FL, GA, IL, KS, MA, MI, NC, NJ, OH, PA, TN, TX, VA, and WI. | Terminated | March 24, 2014 | 328 | 63 | Undetermined
Z-1318-2013 | 65029 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Innova 2100 IQ, Innova 3100 IQ, Innova 4100 IQ, Innova 2121 IQ, Innova 3131 IQ , Innova 3100, Innova 4100. |  |  | December 20, 2011 | 2011 | GE Healthcare, LLC | While performing a fluoroscopic examination on the GE Innova systems, there is a potential of nonrecoverableloss of imaging modes with no advance warning, due to a relay failure in the power distribution box (PDB). Although no injuries have been reported due to this issue to date, there is riskof complications during the procedure as a result of imaging functionality loss. | Nonconforming Material/Component | GE Healthcare did not send out customer letters. GE Healthcare completed the correction by replacing the faulty relay in all affected systems.As an effectiveness check, they have tracked all service requests for the field action to completion. These service requests recorded the repair of the affected units. Exceptions to completion of the service request were recorded as well. Records of field action completion will be retained per their processes and procedures.Should you have any questions or require additional information, please call (262) 548-2757. |  | 713 | 713 | Worldwide Distribution -- Nationwide including the states of AL, AZ, AR, CA, CO, CT, DC, FL, GA, GUAM, HI, ID, IL , IN, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NH, NM, NY, NC, ND, OH, OK, OR, PA, PR, SC, TN, TX, UT, VT, VA, WA, WY and WI., and the countries of Philippines, Peru, Paraguay, Panama, Mexico, Malaysia, Korea, Japan, Guatemala, Ecuador, Dominican Republic , Columbia, China, Chile, Canada, Brazil.,Argentina and Venezuela | Terminated | March 13, 2014 | 814 | 713 | Undetermined
Z-1273-2013 | 64898 | No | HFW | Obstetric-gynecologic specialized manual instrument. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(K) Exempt | Cypress Medical Products LLC, MediChoice Umbilical cord clamp, sterile; single use; Distributed by Owens &amp; Minor, Mechanicsville, VA. The product is used to clamp the umbilical cord after childbirth. |  |  | April 15, 2013 | 2013 | Cypress Medical Products LLC | Cypress Medical Products LLC is voluntarily recalling one lot of the Owens and Minor (O&M;) Umbilical Cord Clamps because recent complaints indicate the clamps are breaking upon use. The product may not perform as intended. | Unknown/Undetermined by firm | Cypress Medical Products sent an "Urgent Voluntary Medical Device Recall" notification and an associated response form dated April 15, 2013 to their only customer for this device - Owens &amp; Minor. The product problem was identified and recommended actions provided. |  | 150 | 150 Cases (75,000 Units) | US Distribution to the state of VA only. | Terminated | August 10, 2015 | 847 | 150 | Undetermined
Z-1277-2013 | 64935 | No | KNS | Endoscopic electrosurgical unit and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Olympus SurgMaster UES-40 electrosurgical unit ("UES-40") |  |  | November 10, 2009 | 2009 | Olympus America Inc. | Olympus initiated a software upgrade and enhancement in 2009 and completed the action in 2010. The upgrade and enhancement were initiated due to complaints of overheating, smoke vapors and failure to operate associated with the Olympus SurgMaster electrosurgical unit, model UES-40 ("UES-40") when used in conjunction with the Olympus WA22557C electrode. The complaint investigations revealed that | Use error | Olympus America Inc sent a Field Corrective Action letter dated November 10, 2009 to all affected consignees. The letter identified the affected product, problem and actions to be taked. Consignees were directed to contact Gyrus at 888-524-7266 to schedule return of the UES-40 generator for the software upgrade. The letter also directed consignees to contact OAI Customer Service at 800-848-9024 if they needed a new copy of the UES-40 Instruction Manual. For questions call 484-896-5688. |  | 608 | 608 | Nationwide Distribution - USA Nationwide including the states of: AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NV, NY, OH, OR, PA, RI, SC, TN, TX, UT, VA, WA, WV, and WY. | Terminated | July 05, 2013 | 1333 | 608 | Undetermined
Z-1278-2013 | 64913 | No | NBH | Arthroscope. | Orthopedic | Orthopedic | 510(K) Exempt | DePuy Mitek Tissue Liberator -Blade Up |  |  | April 15, 2013 | 2013 | DePuy Mitek, Inc., a Johnson &amp; Johnson Co. | Labeled incorrectly as a Blade Up configuration instead of a Blade Down configuration. | Process control | DePuy Synthes sent an Urgent Voluntary Product Recall letter dated April 15, 2013, to all affected customers. The letter iidentified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately check all inventories to locate and return the affected product. For returns customers were instructed call 1-877-379-4871 to obtain a goods authorization (RGA) number. Affected product should be returned to:Mitek Sports MedicineATTN: Mitek Complaints325 Paramouont DriveRayham, MA 02067For questions customers were instructed to call 1-877-379-4871. |  | 27 | 27 units | Worldwide Distribution - USA including KY, NJ and Internationally to Argnetina, Australia, Belgium,Columbia, India, France, Japan, Jordan, Latvija, Israel, Korea, and Great Britain.For questions regarding this recall call 508-977-3813. | Terminated | March 11, 2016 | 1061 | 27 | Undetermined
Z-1282-2013 | 64778 | No | GXQ | Dura substitute. | Neurology | Neurology | 510(k) | Dural Graft Matrix |  |  | March 28, 2013 | 2013 | Integra LifeSciences Corporation | Due to the process deviation, product lots in question may have been released with higher levels of pyrogens than permitted by the specifications for the products. | Process control | The firm, Integra, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated April 9, 2013, to all US, Canada, and non-EU Integra consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed to verify if they have any lots of the listed product, if so, STOP using that product; remove the product from service; and promptly complete and return the PRODUCT RECALL RETURN ACKNOWLEDGMENT FORM via Email: FCA@integralife.com or FAX to : 1-609-275-9445. Customer Service department will contact you to return and replace the product."However, if you have already implanted or used collagen sponge products affected by this recall, we recommend you monitor the patient for a fever in the immediate postoperative period according to the standard hospital or clinician protocol."For assistance or any other questions that you may have, please contact Customer Service at 1-855-532-1723. |  | 15607 | 15,607 boxes | Worldwide distribution: USA (nationwide) including Puerto Rico and countries of: Arab Emirates, Argentina, Belgium, Brasil, Canada, Chile, China, Costa Rica, Denmark, Egypt, Spain, France, Great Britain, Italy, Japan, Peru, Singapore, Thailand, and South Africa. | Terminated | June 03, 2015 | 797 | 587490 | Undetermined
Z-1281-2013 | 64887 | No | HMX | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | Insight Instruments MIS ( Multi-Port Illumination System) |  |  | March 15, 2013 | 2013 | Insight Instruments, Inc. | Insight Instruments, Inc., Stuart, FL is recalling M.I.S. (Multi-Port Illumination System), Ophthalmic Cannula due to a misclassification of the medical device. | Use error | Insight Instruments, Inc. sent a recall letter dated March 15, 2013, to all affected customers The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to do the following: 1)Perform a physical count of their inventory immediately for any of the recalled lots and complete and return the Return Response Form. 2) If they have any products listed, contact the firm at (800) 255-8354, to arrange for return and prompt replacement. 3) Discontinue using and dispensing all M.I.S. devices of the recalled lots. |  | 776 | 776 | US Distribution IN and IL | Terminated | August 21, 2014 | 524 | 776 | Undetermined
Z-1280-2013 | 64875 | No | EZY | Implanted mechanical/hydraulic urinary continence device. | Gastroenterology/Urology | Gastroenterology/Urology | PMA | AMS 800Âż Urinary Control System. |  |  | March 01, 2013 | 2013 | American Medical Systems, Inc. | American Medical Systems, Inc. is initiating a recall on two components of the AMS 800Âż Urinary Control System, Control Pump &amp; Pressure-Regulating Balloon. These products were released despite not meeting all final product release criteria. | Release of Material/Component prior to receiving test results | American Medical Systems sent two written communications distributed within the United States. The first is a letter that is being delivered to customers who purchased the affected lots using their ship to address from the original sales order. The second is a letter being delivered to implanting physician customers who are known by AMS to be regular users of the devices. Similar letters will be sent to international customers, including both end users and distributors (where applicable). American Medical Systems sent a "Urgent Medical Device Recall" letter dated February 28, 2013, to all affected customers. The letters identified the product, the problem, and the action needed to be taken by the customer. Customers were instructed to return the recalled product and to complete the Product Recall Acknowledgement Form. The letters addressed to "Physicians" advised consignees against prophylactic removal of implanted control pumps and pressure-regulating balloons contained within the scope of the product recall.For further questions please call AMS Customer Service at 1 (800) 328-3881 |  | 77 | 77 | Worldwide Distribution -- Nationwide Distribution including the states of AL, AZ, AR, CA, CO, CT, DE, FL, GA, IL, IN, IA, KS, KY LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NH, NJ, NM, NY,NC, OH, OK, OR, PA, PR, RI, SC, TN, TX, UT, VT, VA, WA and WI., and the countries of AUSTRIA, AUSTRALIA, CANADA, BELGIUM, SWITZERLAND, CZECH REPUBLIC, GERMANY, DENMARK, ESTONIA, SPAIN, FINLAND, FRANCE, GREAT BRITAIN, IRELAND, ITALY, LEBANON, NETHERLANDS, NORWAY, SWEDEN, TURKEY, ARGENTINA, BRAZIL, IVORY COST, COLOMBIA, ECUADOR, INDIA, JAPAN, MEXICO, NEW ZEALAND, PANAMA and VENEZUELA. | Terminated | April 24, 2014 | 419 | 682 | Undetermined
Z-1317-2013 | 65112 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | PadPro(R) Adult Radiotransparent Multifunction Electrodes, REF/Catalog No. 2001M-C. |  |  | May 03, 2013 | 2013 | ConMed Corporation | ConMed was notified by the contract manufacturer that their wire/connector supplier reported damage to the connector tooling which could pinch or compress the lead wire filaments and potentially render the device incapable of delivering the appropriate energy or shock to the patient and/or could cause potential injury to user or patient. | Component design/selection | ConMed sent a URGENT: Field Safety Notice letter dated May 3, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed to please review your inventory for any of the devices listed on Attachment I.Business Reply Forms were sent to the domestic and foreign consignees on May 3, 2013. The domestic notices were mailed by USP Priority Mail pouches. The recall is being conducted to the end user level.We ask that you contact all of those organizations within your facility and any other facilities that you may have supplied or given these affected products to. It is imperative that all end users of these devices receive this notice and respond immediately. If you have questions, please call 1 315-624-3237 orfax to +1 315-624-3225. |  | 600 | Domestic: 49,010; Foreign: 600 | Worldwide distribution: USA (nationwide) and the country of: Saudi Arabia. | Terminated | August 25, 2015 | 844 | 600 | Undetermined
Z-1319-2013 | 65034 | No | JYO | Ear, nose, and throat manual surgical instrument. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(K) Exempt | Summit Medical, BL-2006, Myringotomy Lance (Juvenile) Blade, Sterile P, Summit Medical, Inc., 815 Northwest Parkway, Suite 100, St. Paul, MN 55121 |  |  | March 27, 2013 | 2013 | Summit Medical, Inc. | Summit Medical has initiated a recall of Myringotomy Lance Blade (Juvenile) because the box contains the wrong blade. The box contains Myringotomy Spear Blade (Juvenile). | Process control | Summit Medical sent a Medical Device Recall letter dated March 27, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed to immediately examine your inventory, and quarantine the product subject to this recall. Please complete and return the enclosed "Recall Response Form" as soon as possible. The fax number is (651) 789-3961.For further questions please call (651) 789-3921. |  | 9 | 9 boxes | Distributed in MS, NY, PA and TX. | Terminated | August 06, 2013 | 132 | 9 | Undetermined
Z-1320-2013 | 64650 | No | LKQ | Cytomegalovirus serological reagents. | Microbiology | Microbiology | 510(k) | QIAGEN artus CMV RG PCR ASR (96) (Catalog number 4503225) |  |  | January 30, 2013 | 2013 | QIAGEN Gaithersburg, Inc. | CMV RG PCR kits may contain a mixture of incorrect vials among the correct vials. | Mixed-up of materials/components | Qiagen Gaithersburg, Inc. issued a PRODUCT FIELD ACTION letter dated February 4, 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to inspect the vials prior to use to ensure that the vials are labeled "CMV RG Master" Vials labeled "EBV TM Master" should NOT be used and count the vials and discard per internal laboratory procedures. Qiagen will replace free of charge or credit any kits discovered with incorrect master mix vials. Please complete page 2 of this notice as soon as possible and email or fax it back to us at: (661) 702-3854, techservice-na@qiagen.com. If you have additional questions or concerns polease do not hestiate to call QIAGEN Technical Services at 1-800-344-3631 option 2 or by e-mail techservice-na@qiagen.com. |  | 90 | 90 kits in total | USA Nationwide Distribution including the states of: CA, DC, FL, GA, MA, MD, MI, and IL. | Terminated | July 01, 2013 | 152 | 90 | Undetermined
Z-1274-2013 | 64975 | Yes | NIU |  |  | Cardiovascular | PMA | ZilverÂż PTXÂż Drug-Eluting Peripheral Stent - 6 mm X 80mm; 125 cm, Rx, Sterile |  |  | April 18, 2013 | 2013 | Cook, Inc. | Cook Medical has received a small number of complaints related to the delivery system for the Zilver PTX Drug-ElutingPeripheral Stent involving fractures of the delivery system inner catheter after stent deployment, and separation of theinner catheter tip section (13 complaints with an occurrence rate of 0.043%) Two adverse events, including one death,occurred in cases where a tip separation | Process control | Cook Medical sent an Urgent Medical Device Recall letter dated April 24, 2013 to all affect customers. The letter identified the affected products, description of the problem, and actions to be taken. The letter instructed customers to stop use, quarantine, and return all affected products. Customers were asked to complete product reply form and return affected products for credit to Cook Medical, ATTN: Retrun Goods/RGA # 2013C0003. For questions contact Cook Medical Customer Relations at 1-800-457-4500 or 1-812-339-2235. |  | 1057 | 1,057 units in USA | Worldwide Distribution - USA Nationwide; in the countries of Japan, Europe, Brazil, and Australia | Terminated | December 05, 2013 | 231 | 1592 | Undetermined
Z-1355-2013 | 64931 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Progressive Medical PTFE Coated Disposable Electrosurgical Electrodes -Laparoscopic Electrode PTFE Coated Curved Blade 33cm Length, Product number P0019 Rx Only. Sterilized by Radiation. Manufactured for Progressive Medical, Inc., St. Louis, MO 63128. |  |  | April 12, 2013 | 2013 | Progressive Medical Inc | Product coatings require 510K approval. | Manufacturing material removal | Progressive Medical Inc. sent an Urgent Medical Device Recall letter dated April 12, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their inventory and segregate the affected units prior to return to Progressive Medical, Inc. If product was further distributed customers should notify their accounts of ther ecall and return the products. Customers were asked to complete and return the attached form by fax to 314-961-4535, even if they had no inventory remaining. Customers with questions were instructed to call 1-800-969-6331.For questions regarding this recall call 314-961-5786. |  | 228 | 38 boxes with 6 units each. | Nationwide Distribution including AR, AZ, CA, CO, FL, ID, IL, IN, KS, KY, LA, MA, MI, MN, MO, NE, OH, OK, OR, TN, TX, WA, and WI. | Terminated | May 07, 2014 | 390 | 6317 | Undetermined
Z-1328-2013 | 64910 | No | JJW | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Iris Diagnostics Division iQ Control/Focus Set, Part Number 800-3104, Lot Number 153-12. |  |  | September 18, 2012 | 2012 | Iris Diagnostics | The firm initiated this recall because complaints from customers reported that the recovery compared to assigned assay values is significantly below specification. This results in the instrument not passing controls and the customer not being able to run patient samples. | Nonconforming Material/Component | IRIS sent a Medical Device Recall letter dated September 18, 2012 to all affected customers. The letter identified the affected product, problem, and actions to be taken. Customers were instructed to dispose of all unopened or partially used inventory, complete the attached worksheet certifying the disposal of the affected product and email or fax to Iris Diagnostics. |  | 942 | 942 boxes | Worldwide Distribution - USA Nationwide and the countries of: Canada, Malaysia, Spain, Switzerland, Taiwan, Peru, France, Russia, Japan, United Kingdom, Germany, and Paraguay. | Terminated | January 27, 2014 | 496 | 942 | Undetermined
Z-1325-2013 | 64997 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian High Energy Clinacs, High Energy Accelerator, Radiation Treatment System, Model Numbers: H14, H27, H29, HCX. |  |  | May 31, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | Under certain conditions, photon beams in High Energy Clinacs may experience a gradual change in beam symmetry, potentially reaching an asymmetry of up to about 7% before interlock occurs. At maximum asymmetry, this may result in no more than about 3.5% dose deviation from expected at any point in the beam, and may result in minor injury to the patient or slightly decreased local disease control. | Process design | Varian sent a Customer Technical Bulletin to all affected customers. The bulletin identified the affected product, problem and actions to be taken. For questions, issues, or concerns contact Varian at 1-888-827-4265 or go to http://my.varian.com |  | 3943 | 3943 units | Worldwide distribution - USA (nationwide) including Puerto Rico, and countries of: Algeria, Angola, Australia, Austria, Bahrain, Bangladesh, Belarus, Belguim, Brazil, Bosnia and Hezregovina, Bulgaria, Canada, Chile, China, Columbia, Cyprus, Denmark, Dominican Republic, Ecuador, Egypt, Estonia, Finland, Frances, French Polynesia, Georgia, Germany, Greece, Gaudaloupe, Gaum, Hong Kong, India, Indonesia, Ireland, Italy, Israel, Latvia, Lithuania, Luxemburg, Japan, Kazarhstan, Kuwait, Lebanon, Macau, Macedonia, Malaysia, Martinique, Mauritania, Mexico, Moldova, Monaco, Morroco, Nepal, Netherlands, New Zealand, Norway, Oman, Pakistan, Panama, Peru, Phillipines, Poland, Portugal, Qatar, Reunion, Romania, Russia, Saudia Arabia, Serbia, Singapore, Slovakia, Spain, South Africa, South Korea, Sri Lanka, Syria, Sweden, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Venezula, Vietnam, and Zimbabwe. | Terminated | May 15, 2013 | 715 | 3943 | Undetermined
Z-1330-2013 | 64959 | No | JJY | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Liquichek" ToRCH Plus IgM Control - Positive |  |  | April 15, 2013 | 2013 | Bio-rad Laboratories, Inc. | The Liquichek" ToRCH Plus IgM Control - Positive, contains detectable levels of HCV RNA which originatedfrom the non-specific human IgM component used in the mousemonoclonal antibody conjugation process for this product. | Material/Component Contamination | Bio-rad Laboratories initiated this recall on April 18, 2013, via telephone call and certified mailing to customers. The notification letter titled "Important Product Information Notice of Correction" was accompanied by a "Field Correction Notification Acknowledgement" form that notified customers of the product for recall (description and codes), issue (This product contains detectable levels of HCV RNA which originated from the non-specific human IgM component used in the mouse monoclonal antibody conjugation process for this product), Customers were instructed to immediately discontinue use this lot of Liquichek ToRCH Plus IgM Control. Discard all remaining inventory of this lot of Liquichek ToRCH Plus IgMControl that you have at your location in accordance with therequirements of your local waste management authorities. Handle this product with the same precautions used with patientspecimens. In accordance with good laboratory practice, all humansource material should be considered potentially infectious. Have the appropriate laboratory personnel complete theacknowledgment form and email or fax it to us. It is very important to usto know that you have received the information provided in this Notice.Contact information (1-800-2-BIORAD) |  | 216 | 211 boxes and 5 MiniPaks | Worldwide Distribution - Distributed in San Juan Puerto Rico and the states of NM, NY, MN, OR and WI., and the countries of Australia, Brazil, France, Italy, Korea, and New Zealand. | Terminated | August 01, 2013 | 108 | 216 | Undetermined
Z-1329-2013 | 64800 | No | DIO | Cocaine and cocaine metabolite test system. | Toxicology | Toxicology | 510(k) | PROFILE-V MEDTOXScan Drugs of Abuse Test System. |  |  | April 05, 2013 | 2013 | Medtox Diagnostics Inc | Review of validation information and customer complaints which indicated preliminary screen positive results for cocaine that were not positive on confirmatory testing. | Nonconforming Material/Component | MEDTOX issued "URGENT MEDICAL DEVICE RECALL" letters dated April 3, 2013 to consignees informing them of the recall. The letter requested that they examine their inventory and immediately discontinue all use and distribution of the affected lots. The letter also requested that they immediately inform clinicians of the recall at the testing sites that have utilized the recall product. The letter also included a Customer Verification Form which was to be returned. |  | 221625 | 221,625 | Worldwide distribution: USA (nationwide) and in the country of: Canada. | Terminated | September 12, 2013 | 160 | 221625 | Undetermined
Z-1334-2013 | 64883 | No | JSG | Multipurpose culture medium. | Microbiology | Microbiology | 510(K) Exempt | BBL(tm) CDC Anaerobe 5% Sheep Blood Agar. |  |  | March 11, 2013 | 2013 | Becton Dickinson &amp; Co. | Microbiological media may be contaminated with bacteria. | Material/Component Contamination | BD Diagnostics notified consignees with a "Field Correction Notification" sent on 3/11/13 to distributors, direct accounts, and distributors' customers as available, via fax and email, and followed by a letter sent via UPS. The notice identified the product, the problem, and the action to be taken by customers. BD Diagnostic Systems informed customers of the bacterial contamination of some microbiological media, including BBL Trypticase Soy Agar with 5% Sheep's blood, part number 221239, lot 2348199; part number 221261, lots 2339019, 2348108, 2348208, 2348074, 2348157, 2355145, 2348202, and 2361011; BBL Trypticase Soy Agar with 5% Sheep's blood plus MacConkey II-I plate, part number 221291, lots 2355138 and 2355164; and BBL CDC Anaerobe 5% Sheep blood agar, lot 2361003. The notice instructed customers to follow product insert directions for inspecting media before use to identify and eliminate contaminated media, and to verify Listeria findings with clinicians. The notice included a response form to indicate receipt of the notice. BD Diagnostic Systems contacted distributors and end-users via fax, email, and letter on 4/25/13 to notify them of a recall expansion to include lot 3038371 from part/catalog number 221261. The recall notice included an acknowledgement form to document receipt. The notice instructed customers to follow product insert directions for inspecting media before use to identify and eliminate contaminated media, and to verify Listeria findings with clinicians. |  | 521160 | 521,160 total pieces (all three recalled products) | Worldwide Distribution-USA (nationwide) including Puerto Rico, Virgin Islands, and Guam and the countries of Columbia, Japan, Singapore, Canada, Turks, Calicos Islands. | Terminated | June 24, 2013 | 105 | 521160 | Undetermined
Z-1337-2013 | 64901 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS Immunodiagnostic Products CK-MB Calibrators, REF 128 6293, IVD, Ortho-Clinical Diagnostics High Wycombe HP12 4DP, UK |  |  | April 03, 2013 | 2013 | Ortho-Clinical Diagnostics | Internal testing demonstrated that opened-reconstituted VITROS CK-MB Calibrators stored at -20 degrees Celsius (-4 degrees Farenheit) generated invalid calibrations that were detected by positively biased quality control results. | Process control | Ortho Clinical Diagnostic sent a " Important Product Correction Notification" letter dated April 3, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The customers were informed of the issue and advised they must discontinue storage of opened-reconstituted vials of VITROS CK-MB Calibrators frozen at -20 degrees Celsius (-4 degrees Farenheit). The same day, the Distributor letters were sent via US Postal Service Priority Mail or FedEx overnight courier to US Distributor consignees who were informed of the issue and instructed to notify their customers of the issue and actions. Also on April 3, 2013, the foreign affiliates were informed of the issue by e-mail and instructed to notify their consignees of the issue and actions. Confirmation of Receipt - Important Response Required so that we can complete our records, please return this form to us no later than April 19, 2013.FAX: 1-888-557-3759 or 1-585-453-4110We apologize for any inconvenience this may have caused your laboratory. If you have any additional questions, please call Customer Technical Services at 1-800-421-3311. |  | 3924 | Domestic: 3019 units, Foreign: 905 units | Worldwide Distribution - Nationwide including Puerto Rico, and the countries of Australia, Bermuda, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Singapore, Spain and Venezuela. | Open, Classified |  |  | 3924 | Undetermined
Z-1338-2013 | 64713 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance CT 16-Slice Water, System Code: 728245; Brilliance CT 10-Slice Water, System Code: 728250; Brilliance CT 6-Slice Water, System Code: 728255; MX8000 IDT 16, System Code: 728120 &amp; MX8000 IDT 10, System Code: 728125 Computed Tomography X-ray systems. Distributed by Philips Medical Systems (Cleveland), Inc., 595 Miner Road, Cleveland, OH 44143 |  |  | March 12, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | Philips Healthcare received a report from the field stating when they viewed the CTDIVol information from the PACS system, the calculations were too high. Instead of storing the CTDIVol value for each slice in the DICOM tag, the system took the sum of all CTDIVol values from all the images in the dataset and stored this value in the DICOM tag, resulting in incorrect information. | Use error | The firm, Philips Healthcare, sent an "Urgent - Medical Device Correction" letter dated March 15, 2013, to its consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed to disregard unusually higher CTDIvol value for Type-3 DICOM tag 0018:9345 in the PACS system, if seen. The firm is implementing software update versions 4.2.2 or 2.2.2 to correct the issue. A Philips Field Service Engineer will contact you to schedule the software installation at your site.If you need any further information or support concerning this issue, please contact your local Philips representative or local Philips Healthcare office. For North America and Canada contact the Customers Care Solutions Center (1-800-722-9377, Option 5: Enter Site ID or follow the prompts). |  | 515 | 515 units | Worldwide distribution: USA (nationwide) including states of: AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, TN, TX, VA, VT, WA, WI, WV and WY; countries of: Argentina, Australia, Austria, Belgium, Brazil, Canada, China, Colombia, Denmark, France, Germany, Gibraltar, India, Iraq, Ireland, Israel, Italy, Japan, Mexico, Netherlands, New Zealand, Nicaragua, Norway, Philippines, Poland, Portugal, Singapore, South Korea, Spain, Sweden, Switzerland, Thailand, Turkey, United Kingdom and Venezuela. | Terminated | June 30, 2015 | 840 | 515 | Undetermined
Z-1342-2013 | 64512 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Molift Smart 150 |  |  | February 15, 2013 | 2013 | Moller Vital | The leg spreading mechanism will fail to hoist and collapse if assemble incorrectly by the user. | Employee error | Molift Group AS issued a Field Safety Notice letter to all affected customers. The letter identified the affected products, problem and actions to be taken to the Smart hoists that are already in use. Customers were instructed to make sure that all affected customers receive this information and encorage to increase attention until the affected product has been upgraded. For questions contact your local Molift representative or Customer Service at: groupsales@molift.com. |  | 15110 | 7,555 units subject to review; US: 1,280 units - Foreign: 6,275 units | US Nationwide Distribution: CA, FL, GA, ID, IL, KY, MA, ME, MI, MN, NC, NH, NJ, NM, NV, NY, OH, OK, PA, TN, TX, VT, WA and Hawaii.Foreign Distribution:: Australia, Austria; Belgium, Denmark, Finland, France, Germany, Iceland, Israel, Italy, Japan, Netherlands, New Zealand, Norway, Spain, Sverige, Switzerland, United Kingdom. | Terminated | July 11, 2014 | 511 | 15110 | Undetermined
Z-1179-2013 | 65018 | No | EAP | Intraoral source x-ray system. | Dental | Radiology | 510(k) | Fona Srl IntraOs 70 intra-oral x-ray system |  |  | July 03, 2012 | 2012 | Chicago X-Ray Systems, Inc. | The IntraOs 70 system was found to have missing or incomplete system labels as required by 21 CFR 1010.2 &amp; 1010.3. Also, the user's manual was found to have missing or improperly labeled information as required by 21 CFR 1020.30(h). | Radiation Control for Health and Safety Act | FONA will have upgraded copies of the manuals that will be made available on the firms website for free download and may also be mailed to each customer or end user should they request it. They will also provide users and distributers with corrected system labels to be placed on the x-ray control, beam limiting device, and tube housing assembly along with instructions for the proper application of these labels. They will provide users and distributers with an amended "Technical Data" page for insertion int the user's manual.Further questions please call (847) 459-3889. |  | 1960 | 1960 units | Worldwide Distribution- USA (nationwide) and Latin America. | Open, Classified |  |  | 1960 | Undetermined
Z-1482-2013 | 64991 | No | HWA | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Straight Impactor - Dorado IBC, Part Number: SI02140. The device is used to position implants after they have been inserted into the disc space during implantation of intervertebral body fusion devices. |  |  | May 10, 2008 | 2008 | SpineFrontier, Inc. | During two separate occasions an impactor plate on the handle broke off while being struck by a mallet. | Process design | Firm states they notified consignees of recall between 5/10/08 and 5/22/2008. Zero product is currently in distribution. |  | 10 | 10 | Worldwide distribution: US (nationwide) to states of: CO, MA, and TX; and country of: Jamaica. | Terminated | October 08, 2014 | 2342 | 10 | Undetermined
Z-1503-2013 | 65321 | No | FYA | Surgical apparel. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Stryker Instruments 0400-850-000S2 T4 Toga, XL, w/Inside Tie) Sterile, EO; single use; |  |  | May 23, 2013 | 2013 | Stryker Instruments Div. of Stryker Corporation | Based on an accelerated aging test for this product packaging, there is an indicator of potential for a breach in the sterile barrier in the header bag and potentially the blue sterilization wrap for togas within the second and third year of shelf life. The test method used involved high heat to accelerate the aging of the packaging. Stryker is conducting additional testing and assessing this pr | Nonconforming Material/Component | Stryker sent an Urgent Urgent Recall noitifications letter dated May 7, 2013 via certified US mail and E-mail l to all affected customers, distributors and direct accounts. The letter identified the affected products, problem, risk to health, and actions to be taken. Customers were instructed to check inventory, immediately quarantine any affected product, and open each package before discarding to prevent use. Customers were directed to fill out and return a Business Reply Form to Stericycle, Phone: 888-943-2396 Fax: 888-943-4192, Email: stryker5893@stericycle.com verifying directions and confirming recalled product was destroyed. For questions call 269-389-4354 or 269-389-4518. |  | 2000 | 2,000 units ( 200 boxes) | USA Nationwide Distribution | Terminated | April 14, 2014 | 326 | 599490 | Undetermined
Z-1674-2013 | 65332 | Yes | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius Liberty Cycler Set Single Conn./Ext DL |  |  | June 03, 2013 | 2013 | Fresenius Medical Care Holdings, Inc. | Cassette may leak during set up or during peritoneal dialysis treatment | Unknown/Undetermined by firm | Fresenius Medical Care sent a Urgent Recall letter dated June 11, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to check their stock immediately and discontinue use and return product. Please contact your Fresenius Medical Care Customer Service Team at 1-800-323-5188 for instructions to return the product. |  | 5480 | 5480 units | Nationwide Distribution | Terminated | June 18, 2014 | 380 | 5480 | Undetermined
Z-1617-2013 | 65506 | No | DRF | Electrode recording catheter or electrode recording probe. | Cardiovascular | Cardiovascular | 510(k) | 20-Pole LASSO Nav Catheter Eco, Catalog # D-1349-02S. |  |  | June 19, 2013 | 2013 | Biosense Webster, Inc. | Biosense Webster is recalling the 20-Pole LASSO NAV Catheter Eco because it has been incorrectly calibrated causing them to be recognized and displayed as a 10-Pole LASSO NAV Catheter Eco by the CARTO 3 System. | Use error | An Urgent Field Notice letter was sent to all their customers who purchased the 20-Pole LASSO NAV Catheter Eco. The letter informed the customers of the problem identified and the actions to be taken. Customers are instructed to complete and return the attached Voluntary Field Removal Certification Form in accordance with the instructions listed on the form. Customers are instructed to contact Biosense Webster sales representatives or call (866) 473-7823, Monday-Friday from 7:00AM to 8:00 PM EST if they have any questions.For questions related to the Voluntary Field Removal Certification Form and its return, customers are instructed to contact the Field Action Coordinator, at (909) 839-8468. |  | 10 | 10 units | Distributed in the states of MN, FL, and UT. | Terminated | February 07, 2014 | 233 | 10 | Undetermined
Z-1590-2013 | 65297 | No | NEK |  |  | Orthopedic | PMA | INFUSE(R) Bone Graft, MEDIUM KIT, REF 7510400, Medtronic Sofamore Danek USA, Inc., 1800 Pyramid Place, Memphis, Tennessee 38132 |  |  | May 28, 2013 | 2013 | Medtronic Sofamor Danek USA Inc | Kits contain a collagen sponge recalled by Integra due to possible elevated endotoxin levels resulting from a process deviation. | Nonconforming Material/Component | Medtronic on behalf of Integra LifeSciences Corporation sent an Urgent Voluntary Removal (Recall) letter on May 24, 2013, directly to all Risk Managers and Sales Representatives, with Recall Questionnaires and Confirmation forms included. Foreign consignees were notified by email on May 28, 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to remove the product and discontinue use. The letter should be forwarded to surgeons or attending physicians who have used the affected products. Customers were instructed to use the attached questionnaire to account for product in their possession, document which physicians they notified and report any patient injuries that may have occurred related to the use of the affected product by calling the phone number in the questionnaire. Customers were asked to return the questionnaire using the enclosed self-addressed stamped envelope, via their Medtronic Sales Representative or via fax to 763-367-8134For questions regarding this recall 901-399-3182.Customers were instructed to contact their Medtronic Sales Representative to coordinate the return and replenishment of affected product . Customers with questions should call 800-876-3133, ext 3197. |  | 8884 | 8884 units | Worldwide distribution: USA (nationwide including Puerto Rico) and Internationally to Australia, Brazil, Canada, Chile, Hong Kong, India, Mexico, Netherlands, , New South Wales, Panama, and Paraguay. | Terminated | May 13, 2014 | 350 | 57508 | Undetermined
Z-1519-2013 | 65372 | No | MUE | Full-field digital mammography system. | Radiology | Radiology | 510(k) | GE Healthcare Senographe Essential System. |  |  | January 14, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential issue due to inaccessible Operator Instructions associated with the defective Operator Manual CD of the Senographe Essential. | Radiation Control for Health and Safety Act | GE Healthcare will notify customers with affected systems via letter and enclose a new CD- Rom, along with notification that a GE Representative will schedule an appointment to install the User's Manual along with a GE Service Engineer will go on site to install it on the control station of the Senograph Essential system. 16- H LA 6.Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter.GE Healthcare will secure and summarize documentation that the manual has been installed and is available on each affected system and provide summarize of the documentation to the district recall coordinator.For further questions please call ( 262 ) 513-4122. |  | 49 | 49 installed in the US. | Nationwide Distribution including the states of NJ, CO, LA, AL, OH, MA, NC, ME, WI, GA, PA, TX, IN, NY, MO, KY, KS, CA, MN, FL and IL. | Terminated | November 26, 2013 | 316 | 49 | Undetermined
Z-1599-2013 | 65407 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Multilite Dye Handpieces, an optional accessory of the RevLite SI Laser. |  |  | May 09, 2013 | 2013 | ConBio, a CynoSure Company | An inconsistency was identified related to electronic display identification of the optional Multilite Dye Handpieces of the RevLite SI Laser system that may confuse user. | Nonconforming Material/Component | The firm, Cynosure Inc., sent a "MEDICAL DEVICE FIELD ACTIONS NOTIFICATION" letter dated May 9, 2013 to its customers via UPS next day. The letter described the product, problem and actions to be taken. The customers were instructed to immediately stop using the Revlite SI Multilite Dye Handpieces upon receipt of this letter and complete and return the enclosed Field Corrective Action Acknowledgement Form via e-mail to: mflynn@cynosure.com or FAX to: 978-256-6556. A Cynosure representative will be contacting customers to schedule and arrange an exchange of the affected handpieces.If additional information is required please contact VP of Regulatory Affairs and Quality Systems at 978-367-2350 or via email at ikulinets@cynosure.com. |  | 72 | 12 (6 of model 659-1050; 6 of model 659-1100) | Worldwide distribution: USA (nationwide) including states of: IL, ME, NM and VA; and countries of: Australia and Chile. | Terminated | June 26, 2013 | 48 | 72 | Undetermined
Z-1597-2013 | 65361 | No | KMI | Bed-patient monitor. | General Hospital | General Hospital | 510(K) Exempt | Code Alert Advanced 4-Way Care Solution, CA520 System, |  |  | May 17, 2013 | 2013 | RF Technologies, Inc. | RF Technologies, Inc. has initiated a recall involving CA520 Code Alert Integrated Care Management Advanced 4-Way Care Solution . Due to a firmware error, the alarm may not be transmitted to the system server immediately. | Nonconforming Material/Component | Consignees were sent on 5/17/2013 a "Medical Device Recall" letter dated 3 May 2013. The letter described the problem and the product involved in the recall. The letter advised consignees to discontinue use of the device and return to RF Technologies. |  | 167 | 167 | Nationwide Distribution in US including the states of AK, AZ, CA, GA, IA, KY, MT, NE, NJ, ND, and WI. | Terminated | December 03, 2013 | 200 | 167 | Undetermined
Z-1615-2013 | 65247 | No | DCK | C-reactive protein immunological test system. | Immunology | Immunology | 510(k) | VITROS Chemistry Products CRP Slides (Reagent), REF/Catalog Number 192 6740 and 809 7990, |  |  | May 09, 2013 | 2013 | Ortho-Clinical Diagnostics | While performing interval testing Ortho Clinical Diagnostics (OCD) determined that an excess amount of surfactant was unintentionally added to VITROS Chemistry Products CRP Slides during the manufacturing process. | Employee error | Ortho Clinical Diagnostics sent an Important Product Correction Notification letter dated May 9, 2013 via FedEx or US Priority Mail to all US customers. Foreign Affiliates were notified by e-mail on May 9, 2013 informing them of the issue. Distributors were instructed to notify their customers of the recall. Customers were instructed to discontinue use, discard the affected product, comp[lete and return the attached Confirmation of Recipt Form, and post this notification by each VITROS System. For questions call Customer Service Center at 1-800-828-6316. |  | 923 | Domestic: 4,257; Foreign: 923 | Worldwide Disribution - USA Nationwide including Puerto Rico, and the countries of Australia, Bermuda, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Republic of Panama, Singapore, Spain and Venezuela. | Terminated | June 12, 2015 | 764 | 9262 | Undetermined
Z-1608-2013 | 65464 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Cordis a Johnson &amp; Johnson Company, RadialSource Spring Wire Transradial Access Kit. Manufacturer Greatbatch Medical 2300 Berkshire Lane North, Minneapolis, MN 55441, Distributor Cordis Corporation 14201 North West 60th Avenue, Miami Lakes, Florida 33014. |  |  | May 31, 2013 | 2013 | Greatbatch Medical | Cordis Corporation, Inc and Greatbatch Medical are recalling one lot of the cordis RADIALSOURCE Spring Wire Transradial Access Kit product. Some units were manufactured with an undersized vessel dilator tip, too small to slide over the guidewire. | Process control | Greatbatch Medical and Cordis sent an "Urgent Field Safety Notice" dated May 31, 2012, to all affected customers. The letter identified the product the problem and the action to be taken by the customer. Customers were instructed:1.Read the Description of the problem section.2. Immediately identify and set aside all product listed above in a manner thatensures the affected product will not be used.3. Review, complete, sign and return the enclosed Acknowledgement Form inaccordance with the directions on the form.4. Either return any affected product per the attached instruction, or contact yourlocal sales representative to facilitate return of the affected product. Credit willbe provided.5. Pass on this notice to anyone in your facility that needs to be informed.6. If any product listed above has been forwarded to another facility, contact thatfacility to arrange return.7. Maintain awareness of this notice until all product listed above has beenremoved from the facility and returned to Cordis.8. Maintain a copy of this notice with affected product until product is returned.We apologize for any inconvenience this communication may cause. We knowthat you place high value in our products and we appreciate your cooperation inthis matter. Cordis and Greatbatch are committed to maintaining your confidence in the safety and quality of the products that Greatbatch manufactures and Cordis supplies. Further questions please call (763) 951-8181. |  | 550 | 550 kits (OUS) | Worldwide Distribution - AUSTRIA, BELGIUM, SWITZERLAND, GERMANY, SPAIN, GREECE, HUNGARY, IRELAND, ISRAEL and ITALY. | Terminated | August 04, 2014 | 430 | 550 | Undetermined
Z-1600-2013 | 65365 | No | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | Brand Name: Infusor SV and LV Elastomeric Infusion Devices. |  |  | June 07, 2013 | 2013 | Baxter Healthcare Corp. | Baxter Healthcare Corporation (Baxter) is issuing an Urgent Product Recall for the products listed above due to an increase in rupture complaints. Baxter has implemented product and process improvements which have substantially improved product quality. Complaint rates for ruptures have since decreased. Baxter wants to ensure that no affected product remains in the field. | Nonconforming Material/Component | An Urgent Product Recall notification was sent to all affected consignees on 06/13/2013. The letter states the affected products, problem description, hazard involved and actions to be taken by customer/user, with contact information. The letter was accompanied by a Customer Reply Form and a table of Affected Product List. The recall notification instructs customers to segregate affected product and contact Baxter for instructions for return of product. |  | 3255096 | 3,255,096 total | Distributed Nationwide and in Puerto Rico. | Terminated | June 04, 2014 | 362 | 3255096 | Undetermined
Z-1613-2013 | 65377 | No | KHP | Lactic acid test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | SYNCHRON System(s) Lactate (LACT) Reagent. |  |  | May 28, 2013 | 2013 | Beckman Coulter Inc. | Beckman Coulter has confirmed customer reports of LACT Lot M210077 failing calibration witherror condition OCR LOW. The calibration failures appear to be related to improper shipping orstorage conditions (frozen reagent). Failed calibration of LACT may cause a delay in reporting results. | Under Investigation by firm | Beckman Coulter Inc. sent an Urgent Product Correction letter dated May 28, 2013 to affected customers. The letter identified the affected product, the issue, impact, actions to be taken, and resolution. The letter instructed customers to discontinue use of the affected product and request a replacement product. For questions contact Customer Support Center (Call Center) on the Beckman Coulter website at http://www.beckmancoulter.com/customersupport/support/ or call 800-854-3633 in the United States and Canada. Outside the United States or Canada, contact your local Beckman Coulter Representative. |  | 3353 | 3,353 units | Worldwide Distribution - USA Nationwide and country of Canada. | Terminated | June 03, 2014 | 371 | 3353 | Undetermined
Z-1607-2013 | 65433 | Yes | NKE |  |  | Cardiovascular | PMA | Medtronic, ConsultaÂż CRT-P, Model C4TR01, C3TR01 (not approved for US), and SyncraÂż CRT-P, Model C2TR01, Sterile EO. |  |  | June 07, 2013 | 2013 | Medtronic Inc. Cardiac Rhythm Disease Management | Medtronic recently identified an issue with a subset of ConsultaÂż CRT-P and SyncraÂż CRT-P devices during production. As of May 30, 2013, there have been no reported or confirmed device failures. However, because of the potential for malfunction, Medtronic is requiring the return of non-implanted devices manufactured between April 1 and May 13, 2013 for re-inspection. The issue is unique to specifi | Nonconforming Material/Component | Medtronic sent an Urgent Medical Device Recall letter dated June 7, 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to immediately return non-implanted devices to Medtronic and to register the implanted ones immediately by calling the Medtronic Patient Registration Department at 1-800-848-9300. Customers with questions were instructed to contact their local Medtronic representative or Medtronic Customer Service at 1-800-848-9300.For questions regarding this recall call 763-526-6000. UPDATE: An Urgent Medical Device Correction letter, dated June 2013, was sent to affected implanting and following physicians beginning June 28, 2013. In the US these letters were hand delivered by a Medtronic representative. Outside the US, the notice will be mailed or hand delivered as determined by the geographic center. The letter informed physicians that they have implanted or are following a patient implanted with a device potentially affected by the connector bracket weld issue. The letter described how a loss of hermeticity may affect the device functionality and what symptoms a patient may experience. The letter also offers patient management recommendations. |  | 1912 | 1912 (159 USA &amp; 1137 OUS) UPDATED: 1128 affected | Worldwide Distribution - USA including CA, FL, GA, KY, LA, LMN, MO, NE, NY, OK, PA, TN, TX, WI, and PR and Internationally to Australia, Austria, Belgium, Brazil, Bulgaria, Canada, Chile, China, Colombia, Denmark, Ecuador, Egypt, Finland, France, Germany, Hungary, Ireland, Italy, Japan, Latvia, Lebanon, Netherlands, Norway, Poland, Singapore, South Africa, Spain, Sweden, Switzerland, Taiwan, Tunisia ,United Arab Emirates, United Kingdom, Uruguay , and Yemen. UPDATED Physicians location in USA: AL, AR, FL, GA, IL, IN, KS. LA, ME, MI, MN, MO, NY, OH, PA, SC, TN, TX and VA. OUS: Luxembourg | Terminated | March 07, 2014 | 273 | 1912 | Undetermined
Z-1619-2013 | 65166 | No | MAI | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | GENESYS Cross FT Suture Anchor with Two #2 (5 metric) Hi-FI Registered Trade Mark Sutures Rx Only STERILE EO CONMED LINVATEC REF CFBC-4502. |  |  | March 06, 2013 | 2013 | Linvatec Corp. dba ConMed Linvatec | Incorrect blue/white suture was used. | Mixed-up of materials/components | ConMed Linvatec sent an "URGENT: Medical Device Recall Notification" letter by FedEx next day delivery on March 6, 2013, to notify all direct consignees of the recall. If the account no longer had the product in possess, they were asked to respond via a reply form or to contact ConMed Linvatec Customer Service to obtain a Service Request number for return and credit of unused unopened product. All returned product for this action would be identified, segregated and quarantined per the firm's procedures. Any returned product not part of the recall will be dispositioned in accordance with the firm's procedures.Please accept our sincere apology for any inconvenience this action may have caused . Should you have any questions please contact our Customer Service department at (800) 237-0169. |  | 60 | 60 each of both units | US Distribution including the states of FL, KY, LA, MI, NC, and NY. | Terminated | July 22, 2013 | 138 | 60 | Undetermined
Z-1566-2013 | 65343 | No | OZT |  | Cardiovascular | Cardiovascular | 510(k) | TRUE Dilatation Balloon Valvuloplasty Catheter, |  |  | March 18, 2013 | 2013 | Loma Vista Medical | Slow deflation time: Slower than normal balloon deflation time of greater than 10 seconds. The balloon normally deflates in less than 5 seconds. The device performance specification for deflation time is 10 seconds maximum. | Process design | The firm sent an "URGENT VOLUNTARY DEVICE REMOVAL" letter dated March 18, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. |  | 2379 | 2379 units | Worldwide Distribution-US Distribution in New Jersey and the countries of Germany, Switzerland, Italy, Norway, The Netherlands, Spain, Portugal, Ireland, and Austria. | Terminated | July 17, 2013 | 121 | 2379 | Undetermined
Z-1645-2013 | 65423 | No | DRG | Radiofrequency physiological signal transmitter and receiver. | Cardiovascular | Cardiovascular | 510(k) | Lithium Thionyl non-rechargeable disposable battery pack P/N 350-0294-00 used with ER920 Wireless Series Arrhythmia Event Recorder. |  |  | May 02, 2011 | 2011 | BRAEMAR, INC. | Braemar Inc., became aware of a battery related incident that occurred with an Able battery pack being used in a Braemar ER920W wireless event monitor. The event did not result in any patient related injury or present any clinical impact. | Process control | Braemar began notifying customers by telephone on 5/2/2011. Braemar issued Advisory Notice letter P/N 600-0662-00 RevA dated May 16, 2011. The letter was sent by UPS to all consignees.The letter described the problem and the product involved in the recall. Consignees were advised not to use the battery pack in the Braemar ER920W or Braemar Fusion device and requested consignees to complete and return the Verification Form.For questions regarding this recall call 651-286-8620. |  | 10785 | 10,785 packs | Nationwide Distribution including FL, IL, MD, MN, NY, OH, PA, SC, TN, and TX. | Terminated | July 15, 2013 | 805 | 10785 | Undetermined
Z-1652-2013 | 65410 | No | DRG | Radiofrequency physiological signal transmitter and receiver. | Cardiovascular | Cardiovascular | 510(k) | Lithium Thionyl non-rechargeable disposable battery pack P/N 350-0294-01 used with ER920 Wireless Series Arrhythmia Event Recorder. |  |  | February 24, 2012 | 2012 | BRAEMAR, INC. | Braemar has recently became aware of a battery related incident that occurred with EWT battery packs that are used in a Braemer ER920W wireless event monitor and Fusion device. The event did not result in any patient related injury or present any clinical impact. | Material/Component Contamination | Braemar began notifying customers by telephone on 2/24/2012. Braemar issued Advisory Notice letter P/N 600-066201 RevA on 3/02/2012. The letter described the product, the problem, and the action to be taken by the customer. Customers were requested to only use TESTED batteries that they have receivewd fromBraemar starting on February 23, 2012. Complete and return the attached Verification Form. If customers have any untested product in their inventory, they were instructed to segregate those battery packs and await further instruction from Braemar to coordinate the return of those untested units. Upon receipt of the attached verification form, Braemar will continue to send replacement battery packs that have passed the additional testing protocol. Customers were instructed to contact their Braemar sales representative if they have any questions regarding the notice.For questions regarding this recall call 651-286-8620. |  | 3286 | 3286 packs | Nationwide Distribution including FL, IL, LMD, MA, MN, NY, OH, PA, TN, and TX. | Terminated | July 15, 2013 | 507 | 3286 | Undetermined
Z-1653-2013 | 65286 | No | CAD | Anesthetic vaporizer. | Anesthesiology | Anesthesiology | 510(k) | Delta Vaporizer |  |  | May 14, 2013 | 2013 | Penlon, Ltd. | It has come to Penlon Ltd's attention of an instance where the installation and pre-use checks listed in the User Manual may not have been followed prior to the clinical use of the vaporizer. Penlon Ltd has updated the Installation and daily Pre-Use Check sections of the User Manual to ensure that the instructions are explicit and clear. | Nonconforming Material/Component | Penlon sent an Urgent Field Safety Notice dated May 14, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were reminded and requested to follow the instructions for daily pre-use checks, initial checks, and servicing. Customers were instructed to distribute the notice to all personnel who need to be aware within their organization and to any organization where the devices are used.For questions regarding this recall call 01235 547000. |  | 5303808 | 9208 (of which 576 were within the US) | Worldwide Distribution - USA including FL, KY, PA, ND, MN, IL, NJ, TX, TN, and VA. Internationally to Belgium, China, Czech Republic, France, Germany, Switzerland, Australia, Finland, Holland, Lithunia, Malta, Poland, and Slovakia. | Terminated | May 12, 2014 | 363 | 5303808 | Undetermined
Z-1658-2013 | 64973 | Yes | NAV |  |  | Cardiovascular | HDE - Humanitarian Device Exemption | ***REF SC2101***eMax 2 Plus Console, Single Port, with Irrigation***SN G17307901802*** |  |  | June 24, 2011 | 2011 | The Anspach Effort, Inc. | The RPM of the Motor Handpiece may be less than 80,000 RPM. | Process control | Anspach Effort sent an "URGENT: Medical Device Correction" letter dated June 24, 2011 to their customers via mail. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete a reply form promptly confirming receipt of their letter and return it by fax to the number provided on the form. If any of the affected products were distributed to other services or facilities, please forward this information as appropriate. For questions contact Anspach Customer Support at 800-327-6887 or email customer.support@synthes.com. |  | 61 | 61 units | Worldwide Distribution - USA Nationwide including the states of: AZ, CA, FL, KY, ME, MI, MO, OH, PA, TN, TX , VA, and the countries of: Australia, Egypt, Italy, Japan, Switzerland, Taiwan, and Turkey; | Terminated | December 24, 2013 | 914 | 61 | Undetermined
Z-1663-2013 | 65418 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Siemens MEVATRON series, PRIMART (limited release)PRIMUS and/or ONCOR or ARTIST Linac Systems with BEAMVIEW capability. |  |  | May 29, 2013 | 2013 | Siemens Medical Solutions USA, Inc | The information regarding the FILM Mode calibration may have been considered difficult to interpret when translated. SIEMENS HEALTHCARE CR / RADIATION ONCOLOGY had become aware that this issue may have resulted in mistreatment of patients. | Process change control | Siemens sent an "URGENT: MEDICAL DEVICE CORRECTION" letter and Update instruction TH022/13/S beginning on May 29, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers.July 19, 2013 - TH023/13/S was distributed is now in progress. This is the FILM Mode Physics Primer Addendum. |  | 1956 | 1956 customers | Worldwide Distribution: USA Nationwide and countries of: Algeria, Angola, Argentina, Austria, Australia, Bahamas, Belgium, Bosnia Herzeg, Bulgaria, Brazil, Canada, Chile, Colombia, Croatia, Ecuador, Egypt, El Salvador, France, Germany, Greece, Hungary, India, Indonesia, Italy, Iran, Iraq, Ireland, Japan, Jordan, Kenya, DPR of Korea, Kosovo, Kuwait, Lebanon, Lithuania, Malaysia, Mexico, Nepal, Netherlands, New Zealand, Norway, Pakistan, Palestinian Ter, Peru, Philippines, P.R. China, Poland, Portugal, Romania, Saudi Arabia, Slovakia, Spain, Sri Lanka, South Africa, Sweden, Switzerlan, Taiwan, Thailand, Trinidad Tobago,Turkey, United Kingdom, Venezula, Vietnam, Rep of Yemen, and Zambia. | Terminated | January 14, 2014 | 230 | 1956 | Undetermined
Z-1662-2013 | 65431 | No | NJR | Streptococcus spp. serological reagents. | Microbiology | Microbiology | 510(K) Exempt | BD MAX Instrument, catalog #441916 |  |  | May 31, 2013 | 2013 | Becton Dickinson &amp; Co. | Electrical power module installed on some in vitro diagnostic instruments does not meet industry standards as labeled. | Component change control | Becton Dickinson (BD) notified their customers in the US by phone on May 31, 2013 and by an Urgent Field Correction Action letter dated May 2013. Customers were advised of the affected product, problem and actions to be taken. The letter informed customers that an onsite visit by a BD Service representative is planned to perform a system check so to determine if the instrument is affected. BD's field service team plans to immediately replace the module on site if the affected part is found. For questions contact BD Technical Services Department at 1-800-638-8663. |  | 71 | 71 instruments | Worldwide Distribution - USA Nationwide and in the countries of: Japan, Singapore, Hong Kong, China, Australia, Europe, and Chile. | Terminated | August 19, 2013 | 80 | 71 | Undetermined
Z-1664-2013 | 65432 | No | FSM | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | BASE, S-LIFT Instrument Case |  |  | May 17, 2013 | 2013 | SpineFrontier, Inc. | The S-LIFT Instrument Case contained a bracket with peeling nylon coating. | Component design/selection | Spine Frontier sent an Urgent Advisory Notice letter dated On June 7, 2013 to all consignees via email. The letter identified the affected product, problem and actions to be taken. The letter informed customers that a modification has been implemented for the improved durability of S-LIFT Case Brackets and that SpineFrontier has initiated removal of these cases from the field. An acknowledgement form is to be signed and returned from each recipient of the Advisory Notice. Upon receipt of the acknowledgement form and RMA request for part return, replacement parts will be shipped. For questions call 978-232-3990. |  | 12 | 12 | USA Nationwide including the states of: FL, CO, Wash DC, VA, KS, TX, OH, TN, OK, CA, IL, MO, and MN. | Open, Classified |  |  | 12 | Undetermined
Z-1669-2013 | 65462 | Yes | LDD | DC-defibrillator (including paddles). | Cardiovascular | Cardiovascular | 510(k) | Cardio Medical Products, Inc. Defibrillation/Cardioversion/Monophasic or Biphasic Pacing/ECG Electrode Rx Only Single Patient Use Only |  |  | April 26, 2013 | 2013 | Cardio Medical Products | Covidien received customer reports of Arcing/Sparking on the defibrillation electrode leadwire which may affect Cardio Medical Products Inc. defibrillator electrodes Model Number O101. | Process control | Cardio Medical Products, Inc. issued an Urgent Medical Device Recall letter dated April 8, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review their inventory and segregate any product with the affected lot numbers and return affected product according to the instructions for product credit by contacting the Customer Service group at 1-800-962-9888, option 1, then option 2, and reference the recall letter to request a Return Goods Authorizing (RGA). Customers were instructed to complete the enclosed Verification Form and return it to Covidien by fax to 800-895-6140, even if they did not have any of the affected lots. Customers were also instructed to report any issues to the Quality Assurance Group to ensure proper device reporting procedures are followed. Customers with questions were instructed to call 1-800-962-9888, option 8, then ext 2500.For questions regarding this recall call 973-586-9500. |  | 657 | 657 units | Nationwide Distribution | Open, Classified |  |  | 657 | Undetermined
Z-1672-2013 | 65317 | No | LJE |  |  | Gastroenterology/Urology | 510(k) | NephroMax high Pressure Nephrostomy Balloon Catheter. Material Number/Catalog Number: M0062101400/210-140, M062101440/210144 |  |  | May 28, 2013 | 2013 | Boston Scientific Corporation | Incorrect sized renal sheath was packaged with the NephroMax high pressure nephrostomy Balloon Catheter. | Process control | The firm, Boston Scientific, sent an "Urgent Medical Device Recall Removal - Immediate Action Required" letter dated May 28, 2013, consignees via Federal Express Priority Mail to its consignees/customers.OUS - Customer notification letters has also been sent to consignees/customers in the EU. The letter instructs the customers to Immediately discontinue use of and segregate recalled product; check their inventory; immediately remove all affected product from their inventory, and return it to Boston Scientific at Boston Scientific Corporation, US Distribution Center, Customer Fulfillment Center, 500 Commander Shea Blvd., Quincy, Massachusetts 02171 include original Verification Tracking Form and RGA #. The letter also requests the completion and return of the Account Reply Verification Tracking Form via email to: MapleGroveFieldActionCenter@bsci.com or Fax to: Field Action Center 1-866-213-1806. Follow up communications will be made to the customers that do not respond to this inquiry. Distributors are instructed to forward the recall notification to their customers.If you have any questions call 763-494-1133 or email to MapleGroveFieldActionCenter@bsci.com. |  | 27 | 27 | Worldwide distribution: USA (nationwide) including states of: NJ, NY, OH, IL, MI, MD, MA and countries of: Denmark, Great Britain, and Spain. | Terminated | January 23, 2014 | 240 | 42 | Undetermined
Z-1578-2013 | 65193 | No | FNG | AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Prism Medical C-Series Carry Bar, Class I, model number 360755. |  |  | January 27, 2013 | 2013 | Prism Medical Services USA | The firm replaced the recalled carry bar with a plastic insert with a carry bar with a metal insert due to improper installation of the recalled carry bars. | Use error | The firm, Prism Medical, sent a letter dated January 27, 2013 to is customers.The letter described the product, problem and actions to be taken. The customers were instructed to perform an immediate safety check checking that both the pin and the plug have been inserted into the carry bar as outlined in the attached User Instruction guide. Prism offered to replace the recalled product with a different product with greater functionality.If you have any questions, contact VP-Global Product Development at 416-726-3107. |  | 1131 | 1131 carry bars/New England | Worldwide distribution: US (nationwide) and country of: Canada. | Open, Classified |  |  | 1131 | Undetermined
Z-1577-2013 | 65243 | No | MEG | Piston syringe. | General Hospital | General Hospital | 510(k) | BD SafetyGlide Allergy. REF 305950 Sterile Allergy Treatment Syringe Tray. The device is used for aspiration and injection of medications. |  |  | May 01, 2013 | 2013 | Becton Dickinson &amp; Company | There has been an increase in complaints regarding the needle/safety mechanism disengaging from the syringe. | Process control | Direct consignees were contacted by phone and email on 05/01/2013 and instructed to return the affected product. End users were mailed "Urgent: Product Recall" letters/Response Forms dated 05/01/2013. The notice advises customers of the product issue, identifies affected product and provides instructions on how to handle affected product. |  | 1326000 | 1,326,000 units | Worldwide Distribution including US Nationwide and the countries of Canada and Suriname. | Terminated | September 25, 2013 | 147 | 1326000 | Undetermined
Z-1506-2013 | 65006 | No | HQE | Vitreous aspiration and cutting instrument. | Ophthalmic | Ophthalmic | 510(k) | The Constellation Vision System Pneumatic Handpiece. |  |  | April 17, 2013 | 2013 | Alcon Grieshaber AG | Alcon is conducting a Medical Device Removal for the Constellation Pneumatic Handpiece, DSP, product number 725.01. This action is being taken because the forceps or scissors tips (sold separately) may dislodge from the hand piece, if not attached by the user according to the Directions for Use. | Use error | Alcon sent an "URGENT: MEDICAL DEVICE REMOVAL" letter on April 17, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact your local Alcon Sales Representative for questions regarding this notice. |  | 8455 | 8,455 units | Worldwide distribution: US (nationwide) and countries of: Argentina, Australia, Austria, Belgium, Canada, Chile, China, Egypt, France, Germany, Great Britain, India, Ireland, Israel, Italy, Japan, Kazakhstan, Kuwait, Kyrgyzstan, Lebanon, Korea, Malaysia, Mexico, Pakistan, Peru, Russia, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, Turkistan, Venezuela, and Veit Nam. | Terminated | January 07, 2015 | 630 | 8455 | Undetermined
Z-1532-2013 | 65175 | No | MNR | Breathing frequency monitor. | Anesthesiology | Anesthesiology | 510(k) | embla Embletta Gold (It functions as a Polysomnography Recorder and Ventilator Effort Recorder). |  |  | April 24, 2013 | 2013 | Embla System Llc | Calibration error for Embla Embletta Gold units calibrated through service center between June 15, 2012 to February 27, 2013. This calibration error could cause a clinician to over titrate the patient and prescribe a therapy pressure higher than is necessary. | Nonconforming Material/Component | The firm, Embla Systems, LLC, sent an "URGENT: MEDICAL DEVICE RECALL" letter on April 24, 2013, to their customers via courier. The letter described the product, problem and actions to be taken. The customers were instructed to return all recalled units for recalibration and to contact Technical Support at 1-888-662-7632, option #1, at your earliest convenience; Reference the EBG Field Corrective Action in order to obtain your return number. Embla will provide overnight shipping both ways and a three business day turn around for affected units. If you have transferred any of these products to another location, please forward a copy of this urgent notification to that location and notify Manager, Quality and Regulatory Affairs, at 613-254-8877 ext 1105 or via email at EBGFieldCorrectionProgram@natus.com. |  | 519 | 519 units were calibrated at the service center | Worldwide distribution: US (nationwide) including states of: CO, AK, WA, CA, WI, MI, MA OH, DE, FL CA, CT, TX, IL, GA, PA, AZ, DC, VA, MD, UT, IN, MO, NY, NC, OR, NV, NJ, NH, DE, and country of: Canada. | Open, Classified |  |  | 519 | Undetermined
Z-1487-2013 | 65223 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | AGFA Digital Radiography X-Ray System DX-D 100 |  |  | May 15, 2013 | 2013 | AGFA Corp. | Customers could potentially experience intermittent, unintended and illogical movement when using the product. | Under Investigation by firm | Agfa sent an "URGENT SAFETY NOTICE" letter dated May 15, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers. Customers were instructed if you experience unintended movement, the unit has to be taken out of operation. Please notify you Agfa service contact at once.Please distribute this information within your facility to all those who need to be aware of this. Please complete the feedback form as soon as possible and return to us.For further questions please call 1-877-777-2432. |  | 104 | 104 | Worldwide Distribution - US Nationwide including the states of AL, FL, IL, IN, IA, KY, MA, MI, MN, MS, MO, NE, NJ, NY, NC, OH, OR, SC, TN, VA, WV, WY and the country of Canada | Terminated | December 31, 2014 | 595 | 104 | Undetermined
Z-1458-2013 | 65192 | No | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | Part 510450 Disposable Infusion Pump, ...Elastomeric Pump 600 mL( 10mL/hr ),... found in the following kits |  |  | May 10, 2013 | 2013 | Symbios Medical Products, LLC | Symbios Medical Products, Indianapolis, Indiana, is performing a voluntary recall of all GoPumps found in the GoPump Rapid Recovery System kits and all GoBlock kits manufactured with flow control components assembled prior to July 2012. The recall is being expanded after receiving additional reports relating to excessively high flow rates on the 150 mL reservoir, 2 mL/hr GOPump (PN510031), and the | Under Investigation by firm | Symbios sent UrgentMedical Device Recall notices to all consignees beginning May 10, 2013. A second corrected/amended notice was sent on May 14, 2013 to consignees and a National Press release was issued. May 30,2013, a third amended Urgent Recall Notice was sent to all consignees correcting the pump description and lot number range table on page 5 of the notice. The notices and press release provided identification of the recalled devices including devices in the previous February 2013 recall, described possible adverse events, and included directions to to quarantine product in a secure location, in preparation for return. Consignees were asked to call 317 225-4447 ext. 25, 8:00 a.m to 4:00 p.m. EDT to receive a Returned Goods Authorization Number prior to returning product. Consignees were asked to complete and return the Attached Verification Form. Question and Concerns should be addressed to 317-225-4447 ext 25. |  | 33543 | 33, 543 for all kits | US Nationwide distribution. | Terminated | March 07, 2014 | 301 | 33543 | Undetermined
Z-1436-2013 | 65248 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Brand Name: AFX" Introducer System, Model Number S17-45. |  |  | May 13, 2013 | 2013 | Endologix Inc | Endologix, Inc. initiated this voluntary recall of certain lots of Endologix, Inc. due to reports of dilator breakage during clinical procedures. | Nonconforming Material/Component | Endologix Inc sent an "URGENT-MEDICAL DEVICE RECALL NOTICE" letters dated May 13, 2013 to all affected customers. The letter identified the affected product, reason for recall, and actions to be taken. The letter instructed customers to do not use or further distribute any affected product. The letter instructed customers to share this information with physicians that perform these procedures at the facility. The firm will call to confirm that this device is located at the facility. After confirmation the firm will retrieve the device. For questions call (800) 983-2284. |  | 53 | 53 | Worldwide Distribution - USA Nationwide in the states of: NY, NH, IN, MI, FL and NJ and in the countries of: Europe and Latin America. | Terminated | August 19, 2013 | 98 | 53 | Undetermined
Z-1520-2013 | 65081 | No | HWT | Template for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | DePuy Glenosphere orientation guide |  |  | May 03, 2013 | 2013 | DePuy Orthopaedics, Inc. | The Glenosphere Orientation Guide Instrument Used with the Delta XTEND Reverse Shoulder is being recalled because an arrow that is etched on the instrument was incorrectly placed on the opposite side of the instrument. Using the improperly etched arrow as a guide could potentially lead to incorrect orientation of the glenosphere implant. | Process control | DePuy Synthes sent and Urgent Information - Medical Device Instrument Recall Notice letter dated May 3, 2013 via email to affected customers. The letter identified the affected product, problem, clinical implications, and potential adverse events. Arrangements will be made for all affected Glenosphere Orientation Guides (part #230795000/lot #5120443) to be returned. Customers were asked to return the response form via fax (574-371-4964). Product related questions should be directed to DePuy Synthes Joint Reconstruction (1-888-554-2482). Recall inquiries should be directed to Kim Earle, Recall Coordinator, 574-371-4917 (M-F; 8a.m.-5 p.m. EDT). |  | 20 | 20 units | Worldwide Distribution - USA Nationwide and countries of: Australia, Canada, Czech Republic, England, Germany, Switzerland and Turkey. | Terminated | August 28, 2013 | 117 | 20 | Undetermined
Z-1521-2013 | 65217 | No | NWV | Gynecologic laparoscope and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(K) Exempt | Applied Medicals EpixÂż and Direct Drive Laparoscopic Graspers, Model Numbers C4130 and C4140. |  |  | May 01, 2013 | 2013 | Applied Medical Resources Corp | Applied Medical is conducting a voluntary recall on specific lot numbers of its EpixÂż and Direct Drive LaparoscopicGraspers. When the ratchet trigger is activated, if a great enough force is applied, there is a potential for the triggerto fracture and the jaws to remain in the closed position. The likelihood of this situation to occur and result inpermanent patient injury is highly unlikely; | Nonconforming Material/Component | Applied Medical notified customers of the recall beginning on May 13, 2013. The "URGENT: MEDICAL DEVICE RECALL" letter identified the product, problem, and actions to be taken by the customers. For questions regarding this notice, contact the firm at (949) 713-8041. |  | 127305 | 127,305 units | Worldwide Distribution: USA Nationwide and countries of: Argentina, Austria, Australia, Bahrain, Belgium, Canada, Chile, Columbia, Denmark, Ecuador, Finland, France, Germany, Greece, India, Ireland, Israel, Italy, Jordan, Kuwait, Malaysia, Lebanon, Lybia, Mexico, Netherlands, New Zealand, Pakistan, Peru, Poland, Portugal, Qatar, Saudi Arabia, Serbia, Singapore, Spain, South Africa, Switzerland, Taiwan-Province of China,Turkey, United Arabs Emirates United Kingdom, Uruguay, and Venezula-Bolivarian Republic. | Terminated | February 06, 2014 | 281 | 127305 | Undetermined
Z-1525-2013 | 65188 | No | LTT | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | Siemens brand MicroScan Dried Negative Breakpoint Combo 47 panels (B1017-417), SMN #: 10483099), |  |  | May 02, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | The product has the potential for false negative or delayed reaction for Arginine (ARG) results. In addition, 42 hour identification panel holds may occur due to a weak ARG reaction. | Nonconforming Material/Component | Siemens sent an Urgent Field Safety Notice dated May 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact your local Siemens Healthcare Diagnostics Representative for information concerning this notice. |  | 683 | 683 units | Nationwide Distribution. | Terminated | September 17, 2013 | 138 | 683 | Undetermined
Z-1523-2013 | 65296 | No | LKN |  |  | Gastroenterology/Urology | 510(k) | AMICUS Exchange Kit; |  |  | May 24, 2013 | 2013 | Fenwal Inc | Fenwal has initiated a voluntary Urgent Product Recall for lot FA12C07070 of product code R4R2339, AmicusExchange Kits. Fenwal identified a labeling issue with this batch of Product Code R4R2339 (which was cleared for European use only) in which this batch was distributed within the United States market. The problem is identified as labeling that did not include the written description for the s | Process change control | Fenwal sent an Urgent Product Recall letter dated May 24, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immediately check inventory, discontinue the use of this prod.uct, return the affected product to Fenwal, and complete and return the attached reply form. For questions call 1-800-333-6925. |  | 19 | 19 Cases (114 Units) | USA Nationwide Distribution | Open, Classified |  |  | 19 | Undetermined
Z-1524-2013 | 65046 | No | FZP | Implantable clip. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | BARD Biopsy Systems UltraClip Dual Trigger breast tissue marker. |  |  | April 16, 2013 | 2013 | Bard Peripheral Vascular Inc | Bard Peripheral Vascular (BPV) has confirmed that some devices with the recalled product code / lot numbercombination may have the metal tissue marker separated from the needle prior to use. | Nonconforming Material/Component | Bard Peripheral Vascular, Inc. initiated a recall of certain product code and lot number combinations of Bard UltraClip Dual Trigger Tissue Marker by contacting customers on 04/16/2013. On 04/18/2013, BPV sent a Customer Recall notification Letter via FedEx with proof of delivery notice. The letter included reason for recall, a table of affected product code and lot numbers, instructions (Do not use or further distribute any affected product), a response form (titled Recall and Effectiveness Check Form) enclosed, and contact information. |  | 3085 | 3,085 | Worldwide Distribution: USA Nationwide and Internationally to Europe. | Terminated | January 30, 2014 | 289 | 3085 | Undetermined
Z-1526-2013 | 64276 | No | NEU | Implantable clip. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | BARD Biopsy Systems (A Business Unit of Bard Peripheral Vascular, Inc.), UltraClipÂż Dual Trigger Breast Tissue Marker 17g X 10cm needle, Ultrasound Enhanced Coil Marker 5 Pack; Sterilize, Non-pyrogenic, Rx only, Single Use; Bard Peripheral Vascular, Inc., 1625 West 3rd Street, Tempe, AZ 85281 USA. |  |  | January 23, 2013 | 2013 | Bard Peripheral Vascular Inc | Bard Peripheral Vascular (BPV) is initiating this recall because a combination of BardÂż UltraClipÂż Dual Trigger Tissue Marker's and support catheters (Microsheath and Usher) were inadvertently distributed to customers without completing the sterilization process (non sterile). | Process control | The firm, BARD Peripheral Vascular, sent an "URGENT MEDICAL DEVICE RECALL NOTIFICATION" letter on January 23, 2013, to its customers. Bard EXPANDED its recall to include (MicroSheath and Usher support Catheters) and sent another "URGENT MEDICAL DEVICE RECALL NOTIFICATION" letter dated March 5, 2013 describing the additional products, to its customers. The letter described the product, problem and actions to be taken. The customers were instructed as follows: Do not use or further distribute any affected product; check all inventory locations within your institution; remove any identified product from your shelves and return to Bard Peripheral Vascular, Inc. 1415 W. 3rd Street, Tempe, AZ 85281; if you have further distributed any of the product, immediately contact that location and advise them of the recall and have them return to BPV; If you have used the affected product, consider notifying, educating and monitoring those affected; and complete and return the Recall and Effectiveness Check form via fax to: ATTN; Recall Coordinator in Customer Service at 1-800-994-6772, even if you no longer have possession of the product; if unable to FAX call BPV at 1-800-321-4254 Option #2 Ex 2727 and report verbally.Please call our Recall Coordinator at 1-800-321-4254 Option #2 Ext 2727 (M-F 7am to 4pm MST) or email at HUsilvia.carrillo@crbard.com with any questions. |  | 815 | 815 | Nationwide Distribution including AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, IL, IN, KY, KS, LA, MA, MD, MI, MN, MO, MS, NC, NJ, NY, OH, OK, PA, TX, VA, WA, WV and WY. | Terminated | June 10, 2013 | 138 | 1370 | Undetermined
Z-1533-2013 | 65395 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | BARD PERIPHERAL VASCULAR VACCESSÂż PTA Balloon Dilatation Catheters, 8mm x 4cm x 80cm. |  |  | May 28, 2013 | 2013 | Bard Peripheral Vascular Inc | Bard Peripheral Vascular (BPV) has confirmed that some product code/lot number combination of devices may have the incorrect balloon diameter labeled on the inflation hub. Specifically, the balloon size printed on the hub may read 7mm x 4cm, when the actual balloon size is 8mm x 4cm. All other product labeling is consistent with the actual 8mm x 4cm balloon size. | Mixed-up of materials/components | BARD sent an URGENT MEDICAL DEVICE RECALL NOTIFICATION letters dated May 30, 2013 via FedEx overnight with proof of delivery notice to all affected customers. The letter identified the affected product, reason for recall, actions to be taken, and contact information. Customers were instructed to complete the enclosed Recall and Effectivenss Check Form and fax to the attention of Recall Coordinator in Customer Service at 1-800-994-6772. |  | 415 | 415 | US Nationwide Distribution | Terminated | October 23, 2013 | 148 | 415 | Undetermined
Z-1569-2013 | 64530 | Yes | DWC | Cardiopulmonary bypass temperature controller. | Cardiovascular | Cardiovascular | 510(k) | Sarns " TCM II Temperature Control Module |  |  | May 30, 2013 | 2013 | Terumo Cardiovascular Systems Corporation | Internal testing found certain Sarns" TCM II systems were released with incorrectly assembled heatsink assemblies which could cause the units to fail. | Employee error | Terumo sent beginning May 30, 2013, an URGENT MEDICAL DEVICE RECALL NOTIFICATION were sent to all customers with instructions to notify all customers of the recall requesting all affected units to be quarantined until the heatsink units are replaced Confirm receipt by filling out the response form containing communication by faxing or emailing the attached Affiliate Response Form to the fax number or email address indicated on the form.Questions should be directed to Terumo CVS customer service in the U.S.:Telephone: +1-734-663-4145; Fax: +1-734-741-6149 |  | 7 | 7 units | Worldwide Distribution including the countries of UAE (United Arab Emirates) HONG KONG and THAILAND. Not distributed in USA | Terminated | September 17, 2013 | 110 | 7 | Undetermined
Z-1542-2013 | 65173 | No | DWS | Cardiovascular surgical instruments. | Cardiovascular | Cardiovascular | 510(K) Exempt | QuickDraw Venous Cannula with Insertion Kit. ThruPort Systmes. Models QD22 (22 Fr) and QD25 (25 Fr). The product is packaged sterile and non-pyrogenic in a sealed, peel-type pouch. Two pouches (cannula and accessories) go into one shelf carton. One shelf carton goes in one shipper carton. |  |  | May 14, 2013 | 2013 | Edwards Lifesciences, LLC | Edwards is recalling the QuickDraw Venous Cannula due to complaints of cannula separation during withdrawal of the device when the Percutaneous insertion method is used. | Process design | The firm, Edwards Lifesciences, sent an "URGENT FIELD SAFETY NOTICE PRODUCT RECALL - ACTION REQUIRED" notice dated May 14, 2013 to its customers. The letter described the product, problem, and actions to be taken. The customers were instructed to review our entire inventory for any QuickDraw Venus Cannula that has not expired; quarantine affected product from your inventory; discontinue use and return unused devices; complete and return the acknowledgment form via fax to Edwards Customer Service on 800-422-9329 within three days of receipt of this Field Safey Notice; contact Customer Service at 800-424-3278 to obtain a Returned Goods Authorization number and replacement product; return product to Edwards Lifesciences, Attn: Santosh Bhagat, 12050 Lone Peak Drive, Draper, UT 84020 Attention: RECALL, RGA#xxxxxx.If you have any questions that have not been answered by this notice, please call Edwards Customer Service at 800-424-3278 from the hours of 6:00AM - 4:30PM PST. |  | 22616 | 22,616 | Worldwide distribution: US Nationwide and country of Canada. | Terminated | October 14, 2014 | 518 | 22616 | Undetermined
Z-1541-2013 | 65285 | No | LZS |  |  | Ophthalmic | PMA | STAR S4 IR Excimer System (STAR S4 IR) |  |  | May 28, 2013 | 2013 | Abbott Medical Optics, Inc. | Under specific conditions, there is an unlikely potential to result in inadvertent laser firing of the laser. | Other | AMO Manufacturing sent an Advisory Notice dated May 28, 2013, to all affected customers via Federal Express. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to distribute the attached advisory notice to all affected customers in their assigned region. Customers were instructed to notify each surgeion at their site and complete and return the attached form to AMO via fax as an an acknowledgement of receipt of this advisory notice and that they understand the information contained within the letter. Customers with questions were instructed to contact their regional contact.For questions regarding this recall call 408-273-4016. |  | 1882 | 1,882, total all systems | Worldwide Distribution - USA (nationwide) and Internationally to Australia, Belgium, Canada, China, Denmark, Finland, France, Germany, Great Britian, Greece, India, Ireland, Italy, Netherlands, New Zealand, Norway, Poland, Puerto Rico, Singapore, Sweden, Turkey, Japan, Spain, Chile, Argentina, Colombia, Brazil, Taiwan, Mexico, South Korea, Venezuela, Paraquay, South Korea, Russian Federation, Costa Rico, Portugal, Uruquay, Martinique, Israel, Bolivia, Ecuador, Lebanon, Thailand, Saudi Arabia, Egypt, Dominican Republic, United Arab Emirates, Czech Republic, Indonesia, Kuwait, Ukraine, Guatemala, Phillipines, Trinidad, Tobago, Mongolia, Jordan, Malaysia, Hong Kong, Bulgaria, Oman, Kazakhstan, Bulgaria, Peru, Bolivia, Croatia, Romania, Yemen, Vietnam, Uzbekistan, Libyia, Slovakia, Algeria, El Salvador, and Cyprus.. | Terminated | November 12, 2014 | 533 | 1882 | Undetermined
Z-1449-2013 | 64128 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips BuckyDiagnost Radiographic, a Diagnostic X-ray system. |  |  | January 16, 2013 | 2013 | Philips Healthcare Inc. | An inspection shows several IATDs (Installation Acceptance Test could not be located or different templates have been used | Release of Material/Component prior to receiving test results | Philips Healthcare issued an Urgent Field Safety Notification on January 11, 2013, to advised users of issued and perform corrective action to the units. Field Service Engineer (FSE) will execute corrections and attach record to IATD file.Your proposed corrective action plan (CAP) for the Bucky Diagnost system involves two mandatory field actions including new testing for sites with missing documentation and correcting the table height measurements for systems where the incorrect measurement procedure was used. You have also revised your IATD document to improve instructions to the field engineer for table height measurements. In your March 28, 2013, letter, you proposed a 6 month timeframe for this corrective action, and in your May 9, 2013 email, you communicated that the field corrections were ongoing and would be completed by July 16, 2013. In your May 9, 2013 email, you also clarified that the recalls are to be completed free of charge to any affected customer, whether or not they have a service agreement with Philips Healthcare.Further questions please call (978) 687-1501. |  | 66 | 66 units | Worldwide Distribution - US Nationwide and the country of Canada. | Terminated | October 15, 2013 | 272 | 182 | Undetermined
Z-1447-2013 | 64130 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips Diagnost Stationary Radiographic System |  |  | January 16, 2013 | 2013 | Philips Healthcare Inc. | Installation Acceptance Test Document (IATD) non compliant with regulations IATDs (Installation Acceptance Testcould not be located or different templates have been used | Radiation Control for Health and Safety Act | Philips Heathcare issued an Urgent - Field Safety Notice letter dated January 11, 2013 to all affected customers to inform them of the systems afffected by this Field Action. The letter identified what the issue is, under what circumstances it may occur, the actions to take to avoid or minimize the occurrence of the issue and the actions planned by Philips to correct the issue. Customers were inform that if they have multiple BuckyDiagnost systems, they will be contacted by a Field Service Engineer to clarify which systems are affected. The Field Service Engineer (FSE) performs a full IATD (Installation PHILIPS Acceptance Test Document) and file the IATD record. Your proposed corrective action plan (CAP) for the Bucky Diagnost system involves two mandatory field actions including new testing for sites with missing documentation and correcting the table height measurements for systems where the incorrect measurement procedure was used. You have also revised your IATD document to improve instructions to the field engineer for table height measurements. In your March 28, 2013, letter, you proposed a 6 month timeframe for this corrective action, and in your May 9, 2013 email, you communicated that the field corrections were ongoing and would be completed by July 16, 2013. In your May 9, 2013 email, you also clarified that the recalls are to be completed free of charge to any affected customer, whether or not they have a service agreement with Philips Healthcare. For questions contact your local Philips representative. 1-800-722-9377. |  | 177 | 177 units | Worldwide Distribution - USA Nationwide and the country ofCanada | Open, Classified |  |  | 278 | Undetermined
Z-1556-2013 | 64874 | No | LPS | Urinary homocystine (nonquantitative) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | VITROS Chemistry Products-HCY Reagent packaged in a kit carton. Catalog# 6802002. |  |  | April 01, 2013 | 2013 | Ortho-Clinical Diagnostics | The firm is recalling the specific lot due to lower than expected results. | Process control | Ortho Clinical Diagnostics sent an "Important Product Correction Notification" letter dated April 1, 2013 to all affected customers. Ortho Clinical Diagnostics recommends that their customers immediately discontinue using and discard all VITROS HCY Reagent remaining in your inventory for the affected lot. They have requested that their customers complete and return the attached Confirmation of Receipt form no later than April 12, 2013. Ortho Clinical Diagnostics requested that these notifications be forwarded to all laboratories and posted beside each VITROS System that utilizes VITROS HCY Reagent within their customers facilities. To place a product order for an alternate lot, please contact Ortho Clinical Diagnostics customer service representatives at 1-800-828-6316. The VITROS Chemistry Products HCY Reagent is used on the VITROS 5,1 FS Chemistry System, the VITROS 4600 Chemistry System and the VITROS 5600 Integrated System to quantitatively measure total homocysteine concentration in human serum and plasma. |  | 478 | 478 total (339- Domestic) (139- International) | Worldwide Distribution-USA (nationwide) including the states of AR, AZ, CA, CO, CT, FL, HI, ID, IL, KY, LA, ME, NJ, OK, TN, TX, WA, and WV, and the countries of Colombia, India, Japan, Mexico, Puerto Rico, Republic of Panama, Singapore, Venezuela, England, France, Germany, Italy, and Spain. | Terminated | March 22, 2017 | 1451 | 478 | Undetermined
Z-1553-2013 | 65263 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | BC Hammock Mesh Sling (Large), product number, 537030. |  |  | May 23, 2013 | 2013 | Prism Medical Services USA | The product does not meet the firm's size standards, the product was produced in the incorrect size. | Employee error | Prism Medical contacted their US consignees on May 28, 2013 with an e-mail notification of the recall. The notification explained the Field Safety Notice Letter dated May 22, 2013. The notice identified the affected product, problem and actions to be taken. Customers were instructed to contact Customer Service at Prism Medical to receive a return goods authorization to replace the affected product at no charge. For questions call 905-850-9003 Ext. 249. |  | 5 | 5 slings | Worldwide distribution: USA (nationwide) including states of: CA, IL, OH, and TX. and country of: Canada. | Terminated | November 25, 2013 | 186 | 21 | Undetermined
Z-1557-2013 | 64830 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes 11-Hole, 1/3 Tubular Plate DCL Plate with Collar, Part Number 241.000.011 |  |  | February 22, 2013 | 2013 | Synthes USA HQ, Inc. | The Synthes One-Third Tubular DCL Plate (11-Hole, 1/3 Tubular Plate w/Collar), was released to the United States market in error. This action is not being initiated as a result of adverse events. | Under Investigation by firm | The firm initiated their recall of the product on February 22, 2013 by sending an "Urgent Notice: Medical Device Recall" letter to the firm's sales representatives. The letter requested they examine their inventory, and/or customers inventories for the Part Number and if found, return it to Synthes for appropriate disposition. |  | 10 | 10 | Nationwide distribution. | Terminated | May 21, 2015 | 818 | 10 | Undetermined
Z-1561-2013 | 64856 | No | CAN | Pressure regulator. | Anesthesiology | Anesthesiology | 510(K) Exempt | MEDICYCL-E- Lite Portable Oxygen System |  |  | March 08, 2013 | 2013 | Linde Gas North America Llc | A number of MEDICYL-e Portable Oxygen Systems may have misapplied valve flow knobs. | Nonconforming Material/Component | The firm, The Linde Group, sent a "Voluntary Medical Device Recall" letter dated June 21, 2013 to its customers. The notice described the product, problem and actions to be taken. The customers were instructed to view the enlarged photo on the reverse side of the notice for guidance. If you or your customers have a Unit with lot number 25/07, then please return it with the Product to your local LifeGas fill center as soon as possible. No other lot numbers are subject to this removal action.Customers with questions should contact Linde's Customer Care Center at 1-866-LIFEGAS (543-3427). |  | 52264 | 52,264 units (1106 units) | Nationwide distribution. | Terminated | February 03, 2015 | 697 | 52264 | Undetermined
Z-1563-2013 | 65187 | No | JJX | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Thermo Scientific MAS Bilirubin for Siemens and Dimension Vista Systems, , Product is manufactured and distributed by Microgenics Corporation, Fremont, CA |  |  | May 13, 2013 | 2013 | Microgenics Corporation | It has been confirmed from in-house testing that Direct Bilirubin (DBIL), and Total Bilirubin (TBIL) claimed in Thermo Scientific MAS Bilirubin Systems are displaying reduced stability. | Under Investigation by firm | ThermoFisher Scientific sent a Product Recall Notification letter dated May 13, 2013, by Certified US mail to all affected customer. The letter identified the product the problem and the action needed to be taken by the customers.Customers were instructed1. Determine if you are using or have inventory of the recalled product.2. Discontinue use of the product immediately.3. Evaluate quality controls. They may vary among laboratories and should be evaluated according to laboratory procedures.4. Retain a copy of the letter for laboratory records.5. If any of the affected lots of MAS Bilirubin had been forwarded to another laboratory, please provide a copy of the recall letter to them.6. Complete and return the response form within 5 days.Letters also provide an alternative product - MAS Bilirubin part number BC103.Questions or concerns may be directed to Technical Services at 1-800-232-3342 or 510-979-5417. |  | 188 | 188 | Nationwide Distribution: US including the states of WI, UT, FL NY, PA, IL, CO, AZ and NJ. | Terminated | August 27, 2013 | 106 | 188 | Undetermined
Z-2174-2013 | 65957 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | STERRAD 100NX Cassettes, Product Code: 10144 |  |  | July 25, 2013 | 2013 | Advanced Sterilization Products | The recall of the STERRAD 100NX Cassettes was initiated because Advanced Sterilization Products (ASP) does not have adequate data to support the effectiveness of the packaging's leak indicator for the entire duration of the STERRAD 100 NX System cassette's 15 month shelf life. | Under Investigation by firm | ASP sent an Urgent Medical Device Field Safety Notification letter dated July 25, 2013, to all affected customers. The recall letter informed the customers of the problem identified and the actions to be taken. Customers were instructed to contact Stericycle directly at (866) 562-5937 to report any complaints or suspected problems with their STERRAD 100NX System cassette. For questions regarding this recall call 949-453-6400. |  | 1923 | 1,923 units | Worldwide Distribution - USA (nationwide) and internationally to RU, VE, KW, SA, HU, GB, IT, HK, IL, MX, ZA, PH, GR, AE, IN, CO, KR, AU, IN, ES, SG, MY, TH, TR, TW, SI, BE, PL, EC, CN, JP, PA, DE, VN, CL, and Switzerland, United Kingdom. Israel. Slovenia, Qatar, Panama, Venezuela. | Terminated | November 27, 2013 | 125 | 1923 | Undetermined
Z-2172-2013 | 65806 | No | CGN | Folic acid test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA Centaur Folate (2500 tests), REF SMN 10340209, Cat No. 09132781 |  |  | July 25, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc | Calibration failures and/or significant negative shifts in quality control (QC) and patient results. | Material/Component Contamination | Siemens sent an Urgent Medical Device Correction letter on July 26, 2013, to all affected customers via Federal Express. The letter identified the product, the problem, and the action to be taken by the customer. an Urgent Field Safety Notice was sent to all domestic customers via Federal Express. On July 25, 2013 this notice was sent to all OUS customers via hard copy, email, and/or fax as determined by each country's local regulations and procedures. These notices inform the customer that a portion of Folate ReadyPacks in kit lots ending in 218 are exhibiting calibration failures and/or significant negative shifts in quality control (QC) and patient results. If customers have an alternate lot of the ADVIA Centaur or ADVIA Centaur CP Folate assay, they are instructed that this lot should be used for future testing. However if no alternative lot is available, quality control (QC) results should be obtained for each individual ReadyPack as a QC or calibration is an indication of an affected pack. An affected pack must be replaced with a different Folate ReadyPack where a valid calibration and acceptable QC values are obtained. These notices should be reviewed by the facility's laboratory or medical director, though no look back is required.Customers were instructed that Fedex Return Receipts will be used to confirm that the customers received the UMDC. Customers with questions were instructed to contact their local Technical Support Provider or Distributor.For questions regarding this recall call 508-668-5000. |  | 44 | 44 | Worldwide Distribution - USA including AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA,MD, ME, MI, MN, MO, MS, MT, NC, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, TN, TX, UT, VA, VT, WA, WI, WV, and WY. Internationally to Canada and Mexico. | Terminated | June 12, 2015 | 687 | 6486 | Undetermined
Z-1216-2013 | 60555 | No | OAY |  |  | Gastroenterology/Urology | PMA | KARL STORZ Photodynamic Diagnosis (PDD) D-Light C System, |  |  | August 16, 2012 | 2012 | Karl Storz Endoscopy America Inc | The recall was initiated because Karl Storz has confirmed that the Photodynamic Diagnosis (PDD) D-Light C System instruction manuals have important safety-related changes. | Process control | A Safety Alert letter, dated 8/16/12, was sent to all their customers to inform them that there is a labeling error in which important safety-related changes were made to the KARL STORZ Photodynamic Diagnosis (PDD) D-Light C System instruction manuals. The correct Ethylene Oxide (EtO) Sterilization set point for Relative Humidity should read, =70% RH. The correction applies to the PDD Camera Head instruction manual, PDD Rigid Telescopes instruction manual, and Tricam SL II instruction manual, the three component manuals indicated for use within the KARL STORZ Photodynamic Diagnosis (PDD) D-Light C System. The letter instructed the customers to place each of the three enclosed stickers, in the appropriate instruction manual. Customers are instructed to cover the erroneous information with the provided stickers as indicated. Customers with any questions regarding the proper use of the device, are instructed to contact Karl Storz's Technical Support Department at 800-421-0837, Ext. 5350, during office hours; or Ext. 7222, 24 hours/7 days a week for emergencies. |  | 15 | 15 units | Nationwide distribution: USA including states of: IL, CA, WV, TX, SC, WI, NY, PA, NE, NJ, MD, OH, IA, and NC. | Terminated | May 08, 2013 | 265 | 15 | Undetermined
Z-2187-2013 | 65968 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | StaXxÂż IB Convex Expandable Implant (25 x 12 x 9 mm), Intervertebral Fusion Device, Lumbar; Catalog Number 11 -1902. |  |  | May 14, 2013 | 2013 | Spine Wave, Inc. | A component Superior endplate manufactured from a different polymer than specified | Component change control | Spine Wave notified surgeons verbally on 5/14/13 and followed by written notification on June 6, 2013 of the incorrect polymer used for the endplate and the health risks. |  | 17 | 17 units | Distributed in the states of FL and OR. | Terminated | February 10, 2014 | 272 | 17 | Undetermined
Z-0376-2014 | 66359 | No | MZW | Dental cement. | Dental | Dental | 510(k) | ExciTE F DSC Soft touch Single Dose, Dual-curing Regular Ref. 50 |  |  | September 23, 2013 | 2013 | Ivoclar Vivadent, Inc. | Ivoclar Vivadent, Inc. is recalling a lot of ExciTE F DSC Dental Adhesive found to produce a blue green discoloration when used in combination with cements. | Material/Component Contamination | The firm, Ivoclar Vivadent, Inc., sent an "URGENT MEDICAL DEVICE RECALL" letter dated September 19, 2013 via first class mail to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to check your stock and segregate any material marked batch R59595; contact the Ivoclar Vivadent Customer Service at 800-533-6825 as soon as possible between the hours of 9:00am and 7:00pm EST Monday through Friday and a representative will arrange to have your stock replaced; both unopened and opened packages can be returned; and complete and return the Medical Device Recall Return Response Acknowledgment and Receipt Form.Please communicate locally with Ivoclar Vivadent Customer Service at 800-533-6825 for any questions. |  | 63 | 63 refill packs (50 single use applications per refill) | US Distribution: including states of PA, WA, WI, NY, LA, TX, NJ, CA, OR, IA, SC, IN, IL, and AZ. | Open, Classified |  |  | 63 | Undetermined
Z-0375-2014 | 66657 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker |  |  | September 24, 2013 | 2013 | Stryker Spine | Multiple complaints have reported that the hex tip of the torque wrench fracturing during final tightening. | Process design | Stryker sent an Urgent Medical Device Removal Notification letter dated September 25, 2013, with product accountability forms to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their inventory and hospital locations to identify any affected product, reconcile any Torque Wrenches with Catalog Number with Lot Numbers referenced in the letter. Customers were asked to fax a copy of the Customer Response Form to Regulatory Compliance at 855-632-9049 or you email it to Spine-RegulatoryActions@Stryker.com. Customers were asked to return the affected product using the enclosed mailing label to: RECALL RA 2013-105Stryker Spine59 Route 17 South,Allendale, New Jersey 07401branches/agencies, the field and surgeons via Fed Ex on starting on 9/24/2013.For questions regarding this recall call 201-760-8298.Stryker sent another Urgent Medical Device Removal Notification Letter dated 1/29/2014 which included additional lots and one product. |  | 651 | 220 US units; 431 OUS units | Worldwide Distribution - USA (nationwide) and Internationally to Argentina, Australia, Canada, Chile, France, Canada, Germany, Hong Kong, Ireland, Italy, Japan, Korea, Netherlands, New Zealand, Spain, Iraq, Egypt, and Turkey. | Open, Classified |  |  | 651 | Undetermined
Z-0377-2014 | 66561 | No | HCH | Aneurysm clip. | Neurology | Neurology | 510(k) | Mizuho Sugita Titanium Standard Aneurysm Clip II , |  |  | October 10, 2013 | 2013 | Mizuho America, Inc. | Specific item numbers form Lots 27-J and 29-J of Sugita T2 Aneurysm Clips. Mini Temporary Type, are discolored. | Process control | Mizuho sent a Product Recall Notice dated October 10, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to verify serial number before opening package and conduct a physical examination of the product for discolored Clips from the affected lots. Customers were also asked to fill out the Confirmation of User Receipt, and return to Misuho America. If customers should find affed clips they were instructed to quarantine them an call Mizuho America Customer Service at 800-699-2547 for return instructions. The Clips will be exchanged at no charge. Customers were also asked to ensure that all appropriate persons within the Hospital have been notified of the Recall Notice.For questions regarding this recall call 510-324-4500, ext 122. |  | 32 | 32 | Nationwide Distribution | Terminated | January 13, 2014 | 95 | 67 | Undetermined
Z-0383-2014 | 66739 | No | JJW | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | MAS Urinalysis Control Level 1, Multi-analyte control (Assayed and Unassayed), Catalog/Part Numbers: 024222; Product is manufactured and distributed by Microgenics Corporation, a subsidiary of Thermo Fisher Scientific, Clinical Diagnostic Division, Fremont, CA |  |  | October 29, 2013 | 2013 | Thermofisher Scientive | Via one (1) consumer complaint and in-house testing has confirmed that the product does not meet the 30 day open vial stability specifications for Lot: UA1514011. | Under Investigation by firm | ThermoFisher Scientific, sent an Urgent Medical Device Field Correction letter dated October 29, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately discontinue use, segregate and discard the affected product. Customers with questions were instructed to contact mas.controls@thermofisher.com. For question regarding this recall call 1-800-232-3342. |  | 914 | 914 units | Worldwide Distribution - USA (nationwide) including Puerto Rico and internationally to Hong Kong. | Terminated | March 17, 2014 | 139 | 914 | Undetermined
Z-0381-2014 | 66792 | No | FZX | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | REF 200068 Security (TM) Enclosed Carpal Tunnel System |  |  | October 25, 2013 | 2013 | Biomet, Inc. | Security Blades were manufactured oversized and potentially won't advance or may be difficult to advance through the mating Security Clip. | Nonconforming Material/Component | 10/25/2013 Urgent Medical Device Recall Notifications were sent to each Distributor and Risk/Recall Manager at each direct account identifying the recalled device, possible adverse health consequences and instructing consignees to immediately locate and remove devices from circulation, following instructions on the FAX back form and returning it within 3 business days to FAX 574-372-1683 and returning product to Biomet Return Goods, 56 East Bell Drive Warsaw In 46580. Questions should be directed to 574-372-1570 M-F, 8 a.m. to 5 p.m. |  | 1130 | 1130 | Worldwide Distribution: USA (nationwide) and Internationally to Japan. | Terminated | May 06, 2015 | 558 | 1130 | Undetermined
Z-0386-2014 | 66542 | No | LAS | Clinical toxicology control material. | Toxicology | Toxicology | 510(K) Exempt | Single Analyte Urine DAU Control (5 m L |  |  | October 03, 2013 | 2013 | Lin-Zhi International Inc | Product catalog # A68825 Norbuprenorphine 13 ng/mL Level 2 Controls gives a decreased concentration reading due to the degradation of the NBUP analyte in solution. | Under Investigation by firm | Beckman Coulter sent a Manufacturer Field Action Letter dated November 2013, to the one distributor. The letter identified the product the problem and the action needed to be taken by the customer.Please share this information with your laboratory staff and retain this notification as part of your laboratory Quality System documentation. If you have forwarded any of the affected product(s) listed above to another laboratory, please provide them a copy of this letter.If you have any questions regarding this notice, please contact our Customer Support CenterÂżÂż Via our website, http://www.beckmancoulter.com/customersupport/supportÂżÂż Via phone, call 1-800-854-3633 in the United StatesÂżÂż Outside the United States, contact your local Beckman Coulter Representative.We apologize for any inconvenience. |  | 60 | 60 items | US Distribution in CA. | Terminated | December 26, 2013 | 84 | 60 | Undetermined
Z-0387-2014 | 66421 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Large Buttress Compression Nut For Blade Guide Sleeve TFN |  |  | July 31, 2013 | 2013 | Synthes USA HQ, Inc. | The Large Buttress Compression Nut For Blade Guide Sleeve TFN was recalled because of complaints that the compression nut did not thread completely onto the blade guide sleeve and became stuck. | Process control | Synthes sent an Urgent Notice: Medical Device Recall letter dated July 31, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Action to be taken: Our records indicate you received this product. Synthesis requesting that the product be returned:If you DO have any of the identified devices, please take the following steps:"Call Synthes at 1-800-479-6329 to obtain a Return Authorization Number."Complete the Verification Section at the end of this letter by checking the appropriate box indicating affected product has been located. Also, please indicate the number of devices found and note the Return Authorization Number. Please include your name, title, telephone number and signature in the spaces provided."Return the Verification Form {page 3 of this letter) with the product to:"Credit/Returns, Synthes, .1101 Synthes Avenue, Monument, CO 80132.If you DO NOT have the identified product, please take the following steps:"Complete the attached Verification Section at the end of this letter by checking the appropriate box indication that no affected product has been located. Please include your name, title, telephone number and signature in the spaces provided this return documentation acknowledges your receipt of medical device removal information."Return the documents to Synthes by." Fax: 1-888-731-7952 or" Scan/email: FieldAction@synthes.comIf you have any questions, please call 610-719-5450 or contact your Synthes Sales Consultant. Thank you for your attention to this issue |  | 135 | 135 | US Distribution including the states of TX, PA, NY, GA, CO, IN, AI, AR, IL, MI, TN, CA, AZ, MS and VA. | Terminated | June 01, 2015 | 670 | 135 | Undetermined
Z-0389-2014 | 66755 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Multi Parameter Patient Monitor (with Arrhythmia Detection and Alarms) |  |  | October 07, 2013 | 2013 | Mindray DS USA, Inc. dba Mindray North America | Mindray has identified an issue with V Series Monitor where the V Dock power pins connecting the V Series Monitor to the V Dock may lose spring tension, potentially causing the screen to temporarily go blank, lose communication with the central station monitor (when one is in use) or shut down. No patient related events have been reported to Mindray concerning the issue. | Environmental control | Mindray sent an Urgent Corrective Action letter dated October 7, 2013 to all of their customers. The letter identified the affected product, problem and actions to be taken. Customers should contact their Mindray Service Representative at 1-800-288-2121 Monday through Friday from 830 am to 530 pm to arrange for the replacement of the affected product. |  | 1839 | 1839 units | Worldwide Distribution - USA Nationwide | Terminated | April 27, 2015 | 567 | 1839 | Undetermined
Z-0397-2014 | 66882 | No | JXG | Central nervous system fluid shunt and components. | Neurology | Neurology | 510(k) | CSF-Unitized Shunt, Contoured Regular, Low Pressure, Catalog No. 46022; |  |  | November 18, 2013 | 2013 | Medtronic Neurosurgery | Medtronic Neurosurgery is recalling the CSF-Unitized Shunts, Delta Shunt Assembly Kit, and the Strata II Shunt Assembly Kit because they were found to have label discrepancies. | Other | Medtronic Neurosurgery sent an Urgent Medical Device Recall letter dated November 18, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to provide a copy of the letter to all those who need to be aware of the matter within their organization or to any organization where the potentially affected product may have been transferred. Customers with questions were instructed to call 805-571-8725. |  | 288 | 288 units total (147 units in US) | Worldwide Distribution - USA (nationwide) and Internationally to US and worldwide: Austria, Belgium, Bosnia, Herzegovina, Canada, Croatia, Germany, Japan, Poland, Portugal, Taiwan, United Arab Emirates, United Kingdom. | Terminated | February 06, 2014 | 80 | 327 | Undetermined
Z-0393-2014 | 66551 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Synthes Material Mix (Adjustable Cervical Distractor-Right, Matrix Distractor Rack, T-PAL Spacer Remover, Implant Inserter) |  |  | October 15, 2013 | 2013 | Synthes USA HQ, Inc. | One material batch used for production of the affected parts of the Synthes Material Mix were mixed in with an incorrect material with a lower than specified hardness and corrosion resistance. | Nonconforming Material/Component | SYNTHES sent an Urgent Notice: Medical Device Recall letter dated October 15, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to call Synthes at 1-800-479-6329 to obtain a Return Authorization Number. Complete the Verification Section at the end of the letter by checking the appropriate box indicating affected product was located. Customers were also asked to indicate the number of devices found and note the Return Authorization Number. Return the Verification Form with the product to:Credit/ReturnsSynthes1101 Synthes AvenueMonument, CO 80132Customers should also return the Verification Form if no affected product was found in their inventory. Return all documents to Synthes by fax: 888-679-4875 or Scan/email: FieldAction@synthes.comCustomers with questions were instructed to call 610-719-5450 or contact their Synthes Sales Consultant.For questions regarding this recall call 610-719-5000. |  | 44 | approximately 44 | Worldwided Distribution: USA (nationwide) including IL, MO, TX, WA and Internationally to Canada and Switzerland. | Terminated | August 31, 2015 | 685 | 44 | Undetermined
Z-0401-2014 | 66723 | No | JDP | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes 4.5MM TI VA-LCP CRVD Condylar Plate/6H/159MM/RT-STER |  |  | May 29, 2013 | 2013 | Synthes USA HQ, Inc. | A plate was inadvertently released to a sales consultant that was restricted for sale. | Employee error | Synthes sent a letter dated November 7, 2013, to the implanting surgeon to follow up on a complaint received. The firm requested any additional information the surgeon might have about the use of the plate as well as a status on the patient following use. For questions the surgeon was instructed to call 800-620-7025.For questions regarding this recall call 610-719-5000. |  | 1 | 1 | US Distribution in Colorado. | Terminated | May 21, 2015 | 722 | 1 | Undetermined
Z-0394-2014 | 66580 | No | HTN | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes Spiked Washers 13.5/6.5 |  |  | October 03, 2013 | 2013 | Synthes USA HQ, Inc. | During a retrospective assessment, Spiked Washers 13.5/6.5 were discovered to be mislabeled as Spiked Washers 13.5/5.5. | Finished device change control | SYNTHES sent an Urgent Notice: Medical Device Recall letter dated October 3, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed call Synthes at 1-800-479-6329 to obtain a Return Authorization Number, complete the Verification Section at the end of the letter by checking the appropriate box indicating affected product located. Return the Verification Form to:Credit/ReturnsSynthes1101 Synthes Avenue, Monument, CO 80132Customers were also asked to return the Verification Form indicating they did not have the affected product also. All documents were to be returned to Synthes by:Fax 1-877-270-5855Scan/email: FieldAction@synthes.comCustomers with questions were instructed to call 610-719-5450 or contact their Synthes Sales ConsultantFor questions regarding this recall call 610-719-5000.. |  | 3087 | 3087 | Nationwide Distribution | Terminated | December 18, 2015 | 806 | 3087 | Undetermined
Z-0400-2014 | 66836 | No | OVD | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | LANX Timberline MPF Lateral Modular Plate Fixation System Straight Split Tip Driver - Short ("Driver"). Model Number 8630-0107. |  |  | November 04, 2013 | 2013 | Lanx, Inc. | Lanx, Inc. is recalling various lots of Timberline MPF Straight Split Tip Driver - Short ("Driver") because the width of the most distal end from the Driver handle (the tip of the hex portion of the Driver) is oversized. | Process control | LANX sent an URGENT: Recall letter dated November 5, 2013, to all affected customers via Federal Express, Overnight delivery. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to return all affected product to the firm. Customers with questions were instructed to contact Field Inventory Management.For questions regarding this recall call 303-501-8402. |  | 25 | 25 units | Nationwide Distribution including AZ, PA, WI, NY, TX, PR, and TN. | Terminated | December 29, 2015 | 785 | 25 | Undetermined
Z-0391-2014 | 66416 | No | HXP | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Synthes Matrix 5.5.mm Left and Right Coronal Bender |  |  | June 22, 2013 | 2013 | Synthes USA HQ, Inc. | A complaint was received of the tips of the Matrix 5.5.mm Left and Right Coronal Bender breaking while bending a spinal rod component of the Matrix System. | Component design/selection | Synthes sent an Urgent Notice: Medical Device Recall letter dated June 22, 2013 to users and sales reps. The letter identified the affected product, problem and actions to be taken. For questions call 610-719-5450. |  | 209 | 209 | Worldwide Distribution - USA (nationwide) including GA, FL, CA, UT, IA, AZ, OH and Internationally to Canada. | Terminated | September 10, 2015 | 810 | 209 | Undetermined
Z-0396-2014 | 63230 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Synthes Titanium Minimally Invasive Reduction Screw System Locking Cap, Part Number 04.631.502 |  |  | September 07, 2012 | 2012 | Synthes USA HQ, Inc. | This action is being initiated following a detailed investigation in response to reported complaints of post-operative locking cap loosening. Synthes has concluded that if the construct is mis-aligned during final tightening, the potential exists for locking caps to loosen post-operatively in certain situations. | Use error | Synthes sent an Urgent: Medical Device Product Removal letter on September 7, 2012, to all affected consignees. Consignees were requested to immediately cease using the product, examine their inventory and remove the parts from inventory. If they had the parts, they were further instructed to call Synthes at 1-800-479-6329 to obtain a Return Authorization Number, complete the verification section at the end of the letter and return the verification form. If they did not have the parts, they were requested to complete the verification section at the end of the letter and return the verification form. Consignees with questions were instructed to call 610-719-5450.For questions regarding this recall call 610-719-5000. |  | 745 | 745 | Nationwide Distribution including CA, CO, CT, FL, GA, ID, IL, IN, LA, MA, MD, MI, MN, MT, NC, NJ, NY, OH, OR, PA, SC, TN, TX, UT, VA, WA, and WI. | Terminated | December 18, 2015 | 1197 | 1502 | Undetermined
Z-0405-2014 | 66577 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Synthes Implant Holder for Synfix (TM)-LR |  |  | October 08, 2013 | 2013 | Synthes USA HQ, Inc. | Complaints were received which described the SynFix LR implant holder breaking at the interface between the implant and the holder. | Unknown/Undetermined by firm | Synthes sent an Urgent Notice: Medical Device Notification letter dated October 8, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to returny any of the affect devices, call Synthes at 1-800-479-6329 to obtain a Return Authorization Number, complete the Verification Section indicating the affected product has been located, also indicate the number of devices found and note the Return Authorization Number. Return the Verification Form to:Credit/ReturnsSynthes1101 Synthes AvenueMonument, CO 80132Customers were also instructed to return the Verification Section even if they don't have the affected product. Customers with questions were instructed to call 610-719-5450 or contact their Synthes Sales Consultant.For questions regarding this recall call 610-719-5000. |  | 548 | 548 | Worldwide Distribution - USA (nationwide) and Internationally to Canada. | Open, Classified |  |  | 548 | Undetermined
Z-0428-2014 | 66509 | No | ELZ | Preformed crown. | Dental | Dental | 510(K) Exempt | 3M ESPE, Unitek Stainless Steel Permanent Molar Crowns |  |  | June 05, 2013 | 2013 | 3M Company - Health Care Business | 3M is recalling specific lots of 3M ESPE Unitek Primary Stainless Steel Crowns - Series UR4 contain LL3 crowns which are incorrectly labeled as UR4 crowns. | Process control | 3M Company sent a Voluntary Recall letterdated June 4, 2013, to all affected consignees. The letter described the problem and the product involved in the recall. Advised consignees to review their inventory and quarantine the product and do not further distribute it. For replacement they should call Customer Service at 3M ESP Dental Products at 1-800-237-1650. If they have further distributed the product they requested they contact their customers to inform them of the recall. 3M prepared a letter to be used in their communications with their customers. The customers should contact 1-800-634-2249.For questions regarding this recall call 651-733-5718. |  | 1419 | 1419 (1325 US, 94 OUS) | Worldwide Distribution - USA including IN, TX, PA, FL, NV, SC, and TN and Internationally to Canada and Saudi Arabia. | Terminated | December 11, 2013 | 189 | 1419 | Undetermined
Z-0409-2014 | 66418 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes 4.5 mm LCP Curved Condylar Plate |  |  | July 30, 2013 | 2013 | Synthes USA HQ, Inc. | The Synthes 4.5 mm LCP Curved Condylar Plate was recalled because a Straight Condylar Plate is etched with a Curved Condylar Plate part number. | Process control | Synthes sent an Urgent Notice: Medical Device Notification letter dated October 8, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to returny any of the affect devices, call Synthes at 1-800-479-6329 to obtain a Return Authorization Number, complete the Verification Section indicating the affected product has been located, also indicate the number of devices found and note the Return Authorization Number. Return the Verification Form to:Credit/ReturnsSynthes1101 Synthes AvenueMonument, CO 80132Customers were also instructed to return the Verification Section even if they don't have the affected product. Customers with questions were instructed to call 610-719-5450 or contact their Synthes Sales Consultant.For questions regarding this recall call 610-719-5000. |  | 13 | 13 | Worldwide Distribution - USA including ID, FL, TN, NY, LA and Internationally to Canada. | Terminated | August 31, 2015 | 762 | 13 | Undetermined
Z-0407-2014 | 66419 | No | JDS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes Titanium Cannulated Humeral Nail |  |  | July 30, 2013 | 2013 | Synthes USA HQ, Inc. | The packaging incorrectly depicts the Synthes Titanium Cannulated Humeral Nail as a 3 distal hole nail instead of a 2 distal locking hole nail. | Pending | Synthes sent an Urgent Notice: Medical Device Recall letter dated July 30, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to call Synthes at 1-800-479-6329 to obtain a Return Authorization Number, complete the Verification Section and return to Credit/Returns, Synthes, 1101 Synthes Avenue, Monument, CO 80132. If customers do not have any affected product they should still complete Verification Form and return to Synthes by fax to 1-888-731-7954 or Scan/email: FieldAction@synthes.com.Customers with questions were instructed to call 610-719-5450.For questions regarding this recall call 610-519-5000. |  | 14 | 14 | Worldwide Distribution - USA (nationwide) including SC, IA, NC, TX, KY, VA, AL, PA, and Internationally to Canada. | Terminated | September 10, 2015 | 772 | 14 | Undetermined
Z-0411-2014 | 66654 | No | HXX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Shortened Plug Driver; Low Volume Specialty Device; Qty:1; Rx Only; Non Sterile; Biomet Spine, Parsippany, NJ 07054 |  |  | September 24, 2013 | 2013 | Ebi, Llc | Biomet received a complaint on 5/20/2013 reporting that the tip of the instrument sheared off during a removal case. | Nonconforming Material/Component | Biomet sent an Urgent Medical Device Recall Letters/Fax Back Response Form dated September 24, 2013, via Fed Ex. The customer is asked to examine their inventory and return the recalled product following the attached instructions. The Fax Back Response Form should be completed and returned to 973-887-1347 prior to the return of the product. Any questions please call (973) 299-9300 x 2162. |  | 2 | 2 units | US Distribution to CA | Terminated | May 15, 2015 | 598 | 2 | Undetermined
Z-0429-2014 | 66530 | No | GAD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Synthes Soft Tissue Retractor Small Extendible. |  |  | August 01, 2013 | 2013 | Synthes USA HQ, Inc. | The small extendible Soft Tissue Retractor was recalled due to locking nut (clamping sleeve) malfunction. | Use error | Synthes sent an Field Safety Notification letter dated August 1, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to take the following actions:Exchange the old surgical technique guide with the new guide version F provided with this notification.Review the revised surgical technique guide.Forward this Field Safety Notification to anyone in your facility that needs to be informed.If the technique guide has been forwarded to another facility, contact that facility.Maintain awareness of this notice until all technique guides have been exchanged.Maintain a copy of this notice.Please complete and return this page by either of the following methods: Fax: 877-567-9320 or Scan/email: FieldAction@synthes.comThank you for your attention to this issue. |  | 374 | 374 | Worldwide Distribution - USA (nationwide) and Internationally to Canada and Mexico. | Terminated | April 16, 2015 | 623 | 374 | Undetermined
Z-0373-2014 | 66459 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | GE Vivid E9 ultrasound system. GE Vivid E9 is a Track 3 diagnostic ultrasound system, which is primarily intended for cardiac imaging and analysis, but which also includes vascular and general radiology applications. The Vivid E9 incorporates a variety of electronic array transducers operating in linear, curved linear, sector/phased array or matrix array format, including two dedicated CW transducers and several real time 3D transducers. It consists of a mobile console with keyboard control panel; color LCD/TFT touch panel, LCD color video display and optional image storage and printing devices. It provides high performance ultrasound imaging and analysis and has comprehensive networking and DICOM capability. |  |  | October 11, 2013 | 2013 | GE Healthcare, LLC | GE became aware of a potential safety issue due to a system slow down and system lock up associated with the software of Vivid E9 ultrasound system. During a critical procedure the Vivid E9 Scanner may gradually become unresponsive and eventually lock up, with the result that the vivid E9 Scanner must be rebooted. | Component change control | An Urgent Medical Device Correction letter, dated 10/11/2013, was sent to 3 departments within the affected hospitals. The letter described the safety issue, and identified affected devices. The letter stated to reboot the scanner in an unresponsive /lock up condition. Also if the system is being used for an interventional procedure, the TEE probe should remain in the patient while system is rebooting. A GE healthcare service engineer will install a correction to affected devices. |  | 710 | 710 | Worldwide Distribution-USA (nationwide) and the countries of Algeria, Argentina, Australia, Austria, Belarus, Belgium, Canada, Chile, Colombia, Croatia, Czech Republic, Denmark, Finland, France, Germany, Greece, Hong Kong, India, Indonesia, Italy, Japan, Korea (Republic of), Kuwait, Latvia, Lebanon, Libyan Arab Jamahiriya, Malaysia, Netherlands, Norway, Oman, Peru, Philippines, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Serbia, Singapore, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, and United Kingdom. | Terminated | January 23, 2015 | 469 | 710 | Undetermined
Z-0372-2014 | 66263 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | *** 1) SYMBIQ One-Channel Infuser; |  |  | August 28, 2013 | 2013 | Hospira Inc. | Hospira shipped the incorrect replaceable battery pack (Part No: HSP3566-E10) instead of the correct E09 battery pack (Part No. HSP3566-E09) for the Symbiq Single and Dual Channel Infusers. | Employee error | The firm, Hospira, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated August 28, 2013 to its customers. Hospira contracted with Stericycle, Inc., in Indianapolis, IN for customer notification, response, tracking, returns, effectiveness checks and destruction of returned product starting on August 28, 2013. The letter described the product, problem and actions to be taken. Customers were instructed to: 1) complete and return the attached reply form via fax to 1-866-470-3291 or e-mail to Hospira2921@stericycle.com, even if you do not currently have the affected batteries; 2) contact Stericycle at 1-877-779-7810 (M-F, 8am - 5pm ET) if they require additional response forms; and, 3) return any affected batteries upon receipt of the new (correct) replacement batteries to Stericycle. Customers are also instructed to forward the recall information to any accounts that they may have shipped the affected batteries to and have them contact Stericycle to recieve a reply form and return label.For further inquiries, please contact Hospira Advanced Knowledge Center at 1-800-241-4002, option 6, availabel 24 hours a day/7 days per week. |  | 59 | 59 batteries | US Distribution in states of: : KY, MN, ND, NJ, NY, PA, and SD. | Open, Classified |  |  | 59 | Undetermined
Z-0369-2014 | 66552 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Synthes Titanium Trochanteric Fixation Nail (TFN)-Helical Blade |  |  | October 03, 2013 | 2013 | Synthes USA HQ, Inc. | A potential exists for non conforming Synthes Titanium Trochanteric Fixation Nail (TFN)-Helical Blades to be distributed. | Process control | Synthes sent a Field Safety Notice dated October 3, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the options provided in the Field Safety Notice to clarify and mitigate the potential risk. Customers were asked to prepare a list of the affected part and lot numbers and contact the Complaint Unit at 610-719-5596 to receive a Return Authorization Number. Complete and return the Verification Form with the product and Return Authorization number. Customers were also asked to forward the Field Safety Notice to anyone in their facility that needs to be informed. Customers were asked to maintain a copy of the Field Safety Notice.Customers with questions were asked to call 610-719-5450 or contact their Synthes Sales Consultant.For questions regarding this recall call 610-719-5000. |  | 350195 | 350195 | Worldwide Distribution - USA (nationwide) and Internationally to Mexico and Canada. | Open, Classified |  |  | 350195 | Undetermined
Z-0331-2014 | 66547 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ALTRUSÂż THERMAL TISSUE FUSION, Tissue Fusion Handpiece, 5 mm O.D., 16 cm Shaft, REF 60-9520-001, STERILE -- CONMED CORPORATION 525 FRENCH RD, UTICA, NY 13502 --- The ConMed AltrusÂż Thermal Tissue Fusion System is comprised of a dedicated energy source and disposable handpiece used to ligate (seal) and divide (cut) blood vessels and tissue bundles that fit into the jaws of the handpiece. The AltrusÂż System utilizes a thermal energy platform to achieve the desired clinical effect. The ConMed AltrusÂż Thermal Tissue Fusion handpiece is a single use device which is provided sterile. The device uses scissors to parallel jaw closure mechanism with one flat jaw and one crowned jaw in which the vessels/tissues are grasped and through which pressure and heat are applied. It is a multi-functional device capable of vessel sealing, grasping and dissecting during open general surgical procedures. The two jaw sizes allow for applications in a variety of clinical environments for access in confined areas as well as large open areas. |  |  | October 23, 2013 | 2013 | ConMed Corporation | The affected lots were exposed to a higher than acceptable level of particulate matter during manufacture and subsequently distributed. ConMed evaluated the potential risk that this issue could present to the patient and concluded the risk of injury to be negligible; however exposure could cause a minor, local, self correcting infection. | Material/Component Contamination | The firm, ConMed, sent an "URGENT: VOLUNTARY DEVICE RECALL" letter dated October 23, 2013 and Business Reply Forms to the domestic consignees/customers via USPS Priority Mail. Distribution of the recall letter to the foreign consignees will follow once the letter has been translated into the different foreign languages. The letter describes the product, problem and actions to be taken. The customers were instructed to immediately stop the use of these devices; review your inventory for any of the devices listed; contact all of those organizations within your facility and any other facilities that you may have supplied or given these affected products to; complete and return the ATTACHMENT I Effectiveness Check Business Reply Form via fax to: ConMed Recall Coordinator at 315-624-3225.If you have any of the affected devices, please complete Attachment I and return it with the devices to: ConMed Corporation, 525 French Road, Utica, NY 13502Attn. Ed Kovac, Return via: FedEx Account # 487553646. Please indicate with your returned devices, if you are requesting credit or replacement of the devices. Please do not return used devices. If you do not have any devices to return, please complete Attachment I, indicating you have no devices and fax it to 315-624-3225, Attn: ConMed Recall Coordinator. For questions, please contact ConMed Recall Coordinator at 315-624-3237 or e-mail altrus@conmed.com. |  | 65 | 65 handpieces (Domestic: 41 devices; Foreign: 24 devices) | Worldwide distribution: US (Nationwide) and countries of: Canada, Germany and Italy. | Open, Classified |  |  | 212 | Undetermined
Z-0190-2014 | 66546 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | Sterrad 100NX, Product Code: 10104. |  |  | October 18, 2013 | 2013 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) is recalling the Sterrad 100NX Sterilizer some healthcare workers detected odors/smells emanating from the 100NX Sterilizers, and it may be related to normal breakdown of the oil used in the Sterrad 200 vacuum pumps. | Under Investigation by firm | A recall letter was sent to all customers on 10/18/13 who purchased the Sterrad100NX Sterilizer. The letter informed the customers of the problems identified and the actions to be taken. Customers are instructed to contact ASP Professional Services directly at (888) 783-7723 and select option #2. Customers are instructed to report any complaints or suspected problems with their Sterrad 100NX System to Stericycle at (877) 598-5705. |  | 2106 | 2106 units total (1224 units in US) | Worldwide Distribution-USA (nationwide) and the countries of EMEA, APAC, Japan, Canada. AE, AU, BE, BH, CA, CH, CL, CN, CO, DE, EC, ES, GB, GR, HK, HU, IL, IN, IT, JP, KR, KW, MX, MY, NX, PA, PH, PK, PL, PR, QA, RU, SA, SG, SI, TH, VE, VN, ZA, and Tripoli. | Terminated | January 16, 2014 | 90 | 2106 | Undetermined
Z-0197-2014 | 66120 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | SIGMA SPECTRUM Infusion Pump with Master Drug Library. |  |  | June 10, 2013 | 2013 | Baxter Healthcare Corp. | Baxter Healthcare Corporation issued an Urgent Device Correction for the SIGMA SPECTRUM Infusion Pump with Master Drug Drug Library due to reports of upstream occlusion alarms and air-in-line alarms occurring during infusions. | Process design | On June 10, 2013 Baxter Healthcare Corporation issued an Urgent Device Correction for the SIGMA SPECTRUM Infusion Pump with Master Drug Library. The notices were dated June 5, 2013 and were distributed via first class mail. Until further information is provided, users should continue to follow the SIGMA SPECTRUM Infusion Pump Operator's Manual, which identifies a number of environmental and clinical usage factors that contribute to upstream occlusion and air-in-line alarms. For high-risk patients, consider having additional devices as backup. If the biomedical engineer finds that the sensor is not functioning properly, contact Baxter Healthcare at 800-356-3454 (choose option 1). For clinical questions, please contact Medical Information Services at Baxter at 800-933-0303, Monday through Friday 8:00am to 5:00 pm Central Time. |  | 273835 | 273,835 units | Worldwide Distribution-USA including Puerto Rico and Canada. | Terminated | April 20, 2017 | 1410 | 273835 | Undetermined
Z-0187-2014 | 66574 | Yes | OAD |  |  | Cardiovascular | PMA | Safire Duo Ablation Catheter MediGuide Enabled, REF A700240 and A700238, Rx only, Sterile EO. |  |  | October 11, 2013 | 2013 | St. Jude Medical | St. Jude Medical Cardiovascular and Ablation Technologies Division is performing a voluntary notification regarding selected batched of Safire Duo Ablation Catheter MediGuide Enabled. These units have a cable connector configuration that is different than intended. | Process change control | Consignees were sent a St. Jude Medical "Urgent Medical Device Notice" dated 11 October 2013. The letter was addressed to Cath lab manger / Risk Manger. The letter described the problem and the product involved in the recall. The letter stated that it is acceptable to continue using these devices as long as the product is used in accordance with the IFU. |  | 46 | 46 | Nationwide Distribution including the states of KY, OH, TX, MA, and NC. | Terminated | June 09, 2014 | 241 | 46 | Undetermined
Z-0201-2014 | 66725 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Titan 3T MRI System, MRT-3010A/5. |  |  | October 28, 2013 | 2013 | Toshiba American Medical Systems Inc | Toshiba America Medical Systems, Inc. is recalling due to a potential problem with Specific Absorption Rate (SAR)values when using the Toshiba Titan 3T System. When the system is used in combination with either the QD Head Coil (MJQH-142A) or the Large KneeSPEEDER (MJAJ-182A), SAR values lower than the actual SAR values may be displayed on the operatingmonitor. | Nonconforming Material/Component | Toshiba America Medical Systems, Inc. sent an "URGENT: MEDICAL DEVICE CORRECTION". letter dated October 28, 2013 via USPS return receipt mail. The letter contains a return reply form that is to be faxed to TAMS for retention. The letter listed Device, Serial Number, reason for recall, problem, corrective actions, request to customers, and contact information. |  | 18 | 18 units | USA Nationwide Distribution in the states of: CA, FL, GA, IA, NV, NY, LA, IN, PA, OR, TX, and WI. | Terminated | December 24, 2014 | 422 | 18 | Undetermined
Z-0285-2014 | 64263 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | The Plum A+ Infusion System is a cassette based multifunction infusion system that allows two lines in and one line out and can be used for standard, piggyback, or concurrent delivery. The Plum A+ Infusion System is designed to deliver parenteral, enteral, or epidural infusions over a broad range of infusion rates from multiple fluid container types. The Plum A+ Infusion System includes an infuser, disposable IV sets (hereafter called a set), optional accessories, and this operator's manual. |  |  | February 01, 2013 | 2013 | Hospira Inc. | Plum A+ infusers have the potential for the distal (occlusion) press sensor pin to break. The distal pressure sensor pin is a part of the overall subsystem that measures the pressure within the distal line of the administration set and indicates the presence of a full or partial distal occlusion. A broken distal pin can only be detected via a visual inspection of the distal pressure pin. A broken | Component design/selection | Hospira sent an Urgent Device Field Correction letter dated February 1, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The letter recommends that facilities immediately visually inspect their Plum a+ devices to determine if the distal pressure pin is broken or damaged via instructions included in the letter. If a broken or damaged distal pressure pin is observed, remove the device from service and contact Hospira at 800-441-4100 (Monday - Friday, 8:00 AM - 5:00 PM, CST) to report the issue. If the return of defective pumps significantly impairs a facility's ability to operate normally, Hospira will provide loaner pumps. The letter also encourages customers to insert cassettes into pumps in accordance with the instructions found in the operator's manual. Customers were asked to complete the attached reply form and return it via fax to the number on the form. Customers with questions were instructed to call 1-800-441-4100 or 1-800-241-4002, option 4.The letter also asked direct accounts that further distributed the pumps notify their customers of the URGENT DEVICE FIELD CORRECTION and ask them to contact Stericycle at 1-866-201-9068 to receive a reply form.. For questions regarding this recall call 800-441-4100, |  | 71 | 71 units | Worldwide Distribution - USA (nationwide) including states of: AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY, District of Columbia, Puerto Rico and U.S. Virgin Islands; and countries of: Costa Rica and Dominican Republic. | Open, Classified |  |  | 231084 | Undetermined
Z-0291-2014 | 66397 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access 2 Immunoassay System, Part Number: 81600N |  |  | September 12, 2013 | 2013 | Beckman Coulter Inc. | Beckman Coulter is recalling the Access 2 Immunoassay System because they determined that the fuses in the system may be of the wrong amperage. | Mixed-up of materials/components | A phone call was made on 9/12/13 to customer who received the Access 2 System. The phone script explains the problem identified of a possible wrong fuse and may result in failure of the fuse to operate properly. Customer was instructed that their Field Service Engineer, will call them back the following week to check on their system and replace the fuses if necessary. If you have any questions concerning this information, please call 952-393-8546. |  | 8 | 8 units total (1 unit in the US) | Worldwide distribution: US (nationwide) in the state of: TX and countries of: Colombia, Hong Kong and Korea. | Terminated | November 14, 2013 | 63 | 8 | Undetermined
Z-0280-2014 | 66444 | No | JDZ | Cement mixer for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | Acetabular Pressurizer; single use and six units per box. |  |  | November 01, 2013 | 2013 | Stryker Instruments Div. of Stryker Corporation | On March 14, 2013 a Work Order was opened for 120 units (20 boxes) of the Acetabular Pressurizer (P/N 0206-522-000 Rev. H). On March 15, 2013, when the product was going to be placed in the pouch to be sealed, the manufacturing operator noticed that some assemblies were cracked. As a result of the injection pressure of the cement gun, a cracked tiplet may release from the Hub during cement applica | Nonconforming Material/Component | The firm, Stryker Instruments, sent an "URGENT MEDICAL DEVICE RECALL NOTIFICATION" letter dated 11/01/2013, to all domestic consignees/customers and an email was sent to the International Subsidiaries and Sales Representatives. The letter described the product, problem and actions to be taken. The customers were instructed to do the following:1. Immediately review this Recall Notification2. Immediately check all stock areas and/or operating room storage and quarantine any affected product found.3. Complete the enclosed Business Reply Form to confirm receipt of this notification and identify how many, if any, affected items are currently in your inventory. Please complete and return the Business Reply Form even if youdon't have any affected product on hand.Note: Your signature on the BRF indicates that you received and understand this Notification and have followed the instructions in this Notification.4. If you have further distributed this product, please forward this letter and the attached Business Reply Form (BRF) to all affected locations. Please indicate each location on the BRF.5. Fax the completed Business Reply Form to Stryker Instruments Regulatory Department, 866-521-2762.6. Upon receipt of the Business Reply Form, if you have recalled product on hand, Stryker will email a pre-paid shipper. This shipper can be used to return the recalled Acetabular Pressurizers to Stryker.7. Upon receipt of recalled Acetabular Pressurizers, credit will be applied to your account.For questions regarding this recall please contact Stryker Instruments: Monday-Friday 8am-5pm (EST) at 269-389-2921 or email: kellyjo.whipple@stryker.com. |  | 87 | 87 boxes (522 each) | Worldwide distribution: US (nationwide) including states of: TX, FL, IL, IA, ND, WI, VA, and NJ.; and countries of: CANADA, POLAND and SWITZERLAND. | Terminated | February 10, 2014 | 101 | 87 | Undetermined
Z-0206-2014 | 64091 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Plum A+3 Infusion Pump System; |  |  | February 05, 2012 | 2012 | Hospira Inc. | Hospira has received reports of distal pressure sensor calibration drift on Plum A+ Infusers. The following error codes may be displayed during setup or infusion signifying possible calibration drift: E180/N180, E181/N181, E186/N186, E187/N187 or E346. Additionally, if the pressure sensor calibration has drifted, users may experience false distal occlusion alarms or distal occlusions that may not | Nonconforming Material/Component | Hospira sent an URGENT DEVICE FIELD CORRECTION letter dated February 5, 2013 was sent to all direct accounts notifying them of the issue. For device users who experience these issues, Hospira recommends the following actions be taken: 1) Remove the pump from service and perform a Performance Verification Test (PVT) Distal Occlusion Test to determine if the distal pressure sensor is performing correctly. 2) If the device does not pass the PVT, it may indicate that the distal pressure sensor has drifted. Perform the appropraite troubleshooting and repair activities defined by the user's facility, whihc may include returning the device to Hospira for further diagnosis and servicing. The remediation for this issue is to perform a recalibration of the distal pressure sensor. Hospira also intends to release in 2013 a recommendation of a yearly test to determine if the distal pressure sensor's calibration has drifted. Customers who suspect that their device(s) may be affected by this issue are urged to contact Hospira at 1-800-441-4100 (Monday - Friday, 8:00 AM - 5:00 PM CST) to report the issue. Hospira will provide loaner devices if necessary. Customers are requested to complete and return the reply form included with the notification letter per the instructions on the reply form.For questions regarding this recall call 224-212-2000. |  | 214 | List No. 12348: 214 units | Worldwide Distribution - USA (nationwide) including states of: AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY, District of Columbia, Puerto Rico and U.S. Virgin Islands; and countries of: Croatia, Dominican Republic, Turks and Caicos. | Open, Classified |  |  | 357778 | Undetermined
Z-0275-2014 | 66608 | No | HOZ | Ophthalmic sponge. | Ophthalmic | Ophthalmic | 510(K) Exempt | DeRoyal (R) Surgical Eye Spears, REF 30-049, Sterile EO, Manufacturer: DeRoyal Industries, Inc. 200 DeBusk Lane, Powell, TN 37849, MADE IN GUATEMALA |  |  | August 12, 2013 | 2013 | DeRoyal Industries Inc | One lot of eye spears was labeled as being x-ray detectable, but contained non x-ray detectable product. | Employee error | DeRoyal Industries sent an Urgent Voluntary Recall letter dated August 12, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to cease usage and return the product to:DeRoyal RecallsRA#: Eye Spears1755 Hwy 33 SouthNew Tazewell, TN 37825Ship via UPS Ground using Shipper #312780 Distributors were requested to notify their consignees of the same. The recall is extended to the user level. Customers with questions were instructed to call 1-800-251-9864.For questions regarding this recall call 865-362-2334. |  | 726 | 726 cases (7242 units) | Nationwide Distribution | Terminated | August 04, 2014 | 357 | 726 | Undetermined
Z-0329-2014 | 65352 | No | FEQ | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Olympus Endoscopic Flushing Pump, Model OPF-2. |  |  | May 20, 2013 | 2013 | Olympus America Inc. | There is a possibility that the pump head was not shipped in the locked position. If the pump head is in an unlocked position this may, under certain circumstances, cause the pump to stop and go into power stand-by mode. | Component design/selection | Notification letters were sent to customers on May 20, 2013. |  | 263 | 263 | Worldwide distribution: USA (nationwide) and the country of Canada. | Terminated |  |  | 263 | Undetermined
Z-0367-2014 | 66631 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | GE Healthcare Discovery NMCT 670, Discovery NM 630, Optima NMCT 640, &amp; Brivo NM 615. |  |  | October 25, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue related to a SPECT scan procedure used for pulmonary imaging.Contact between the Detector (GE Healthcare Discovery NM/CT 670, and/or Optima NM/CT 640, and/or DiscoveryNM630, and/or Brivo NM 615 all configurations) and the Patient's elbow may occur during SPECT protocols using 450 angular steps in scenarios when a Patients a | Employee error | GE Healthcare sent an "Urgent Medical Device Correction" letter dated September 25, 2013, to all affected customers. The letter identified the product the problem and the actions to be taken by the customer.We plan to correct all of the affected Discovery NMCT 670, Discovery NM 630, Optima NMCT 640, &amp; Brivo NM 615 systems that were distributedThe plan for public warning is to contact all customers with an affected Discovery NMCT 670, Discovery NM 630, Optima NMCT 640, &amp; Brivo NM 615 system by means of an Urgent Medical Device Correction letter. In the US this will be sent by a traceable means (either US Postal Service with receipt acknowledgement (registered mail) or equivalent courier service such as FEDEX). This letter will be sent to 3 titles within the affected accounts; Managers of Nuclear Medicine, Hospital Administrators / Risk Managers, Managers of Radiology/Cardiology. A GE Healthcare Field Engineer will perform the required software update on each affected system.Contact Information (GE Healthcare Call Center: USA or Canada 800-437-1171 and Other countries:( 262-896-2890). |  | 464 | 464 (143 US, 321 OUS) | Worldwide Distribution - USA (nationwide) including the states of AL, AZ, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, MI, ME, MD, MA, MN, MS,MO,MT, NV, NH, NJ, NY, NC, OH, OK, PA, PR, SC, SD, TN, TX, UT, VA, WV and WI., and the countries of ALGERIA, AUSTRALIA, BELGIUM, BRAZIL, BULGARIA, CANADA, CHINA, COLOMBIA, DENMARK, EGYPT, ESTONIA, FINLAND, FRANCE, GERMANY, GREECE, GUADELOUPE, HONG KONG, INDIA, IRELAND, ISRAEL, ITALY,. JAPAN, REPUBLIC OF KOREA, KUWAIT, MEXICO, NETHERLANDS, NEW ZEAL NAD, NORWAY, PAKISTAN, POLAND,QUATAR, RUSSIA, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN,K THAILAND, TURKEY and UNITED KINGDOM. | Terminated | April 01, 2014 | 158 | 464 | Undetermined
Z-0264-2014 | 66567 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | ProViaÂż CROSSING GUIDEWIRE; Models: 15PROV300HS, 3PROV180HS, 3PROV180SS, 3PROV300HS, 3PROV300SS, 9PROV180HS, 9PROV180SS |  |  | October 21, 2013 | 2013 | Medtronic Vascular | Potential for PTFE (polytetrafluroethylene) coating to delaminate and detach from Guidewire. Medtronic guide wires are steerable guide wires that are used for the introduction and placement of diagnostics or interventional devices inthe coronary and peripheral vasculature, and may be used to reach and cross atarget lesion. Medtronic guide wires are not intended for cerebral vasculature.Medt | Nonconforming Material/Component | The firm, Medtronic, sent a "URGENT MEDICAL DEVICE RECALL" letter dated October 21, 2013 (in the US via UPS overnight service, 1-day delivery) to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately remove and quarantine all potentially affected guidewires that remain in inventory; return the affected product to Medtronic; if replacement product is needed, your Medtronic representative can assist you with identifying suitable replacement product, and to complete and return the Customer Confirmation Certificate confirming receipt of the notification via fax to Medtronic at 651-367-0612 to the attention of Customer Focused Quality (100% customer confirmations will be required).Should have any questions, please contact your Medtronic Representative or call 763-526-2513.Notifications to customers concerning the additional affected product began December 19, 2013 (in the US via UPS overnight service, 1-day delivery). Customers are requested to return a Confirmation Certificate confirmingtheir receipt of the notification. 100% customer confirmations will be required from the additional consignees. |  |  |  | Worldwide distribution: US (Nationwide) including states of: AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, and WV; and countries of: Albania, Algeria, Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, China, Colombia, Croatia, Czech Republic, Denmark, Ecuador, Egypt, Finland, France, Georgia, Germany, Greece, Guyana, Honduras, Hong Kong,Hungary, Iceland, India, Ireland, Israel, Italy, Japan, Jordan, Kuwait, Lebanon, Libya, Luxembourg, Malaysia, Mexico, Montenegro, Nepal, Netherlands, New Zealand, Nepal, Nicaragua, Norway, Oman, Poland, Portugal Puerto Rico, Qatar, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, United Kingdom, Uruguay, and Yemen. | Terminated | December 15, 2014 | 420 |  | Undetermined
Z-0355-2014 | 66558 | No | ONA | Stereotaxic instrument. | Neurology | General &amp; Plastic Surgery | 510(k) | Sterile Disposable Kit ( Contains Skin Tensioner ASY-25067; Three Mark Needle Assembly ASY-25094; Nose Cone Assembly ASY-25098; Dissection 4 mm PRT-25117, follicle Trap PRT 23999): |  |  | October 13, 2013 | 2013 | Restoration Robotics | Reports of discoloration (residue) on the pins of the skin tensioner have been received by several customers. Discoloration due to insufficient cleaning of stainless steel pins during passivation process. The residue hasbeen identified as a compound of iron and chlorine. | Process control | Restoration Robotics sent an Urgent Recall letters dated October 13, 2013 to all distributors via first class mail. Direct accounts were notified by phone as well as fax. Customers were informed of the affected product, problem and actions to be taken. Replacement products are being distributed to each site and will be delivered to each customer. For questions call 408-883-6760 |  | 515 | 515 US, 315 International (all packaging cnfigurations) | Worldwide Distribution - USA Nationwide in the countries of: Korea, Japan, Singapore, Thailand, United Kingdom, Canada, and Taiwan. | Terminated | December 10, 2013 | 58 | 515 | Undetermined
Z-0340-2014 | 66712 | No | NLH | Electrode recording catheter or electrode recording probe. | Cardiovascular | Cardiovascular | 510(k) | Reprocessed Diagnostic Electrophysiology (EP) Catheters and Diagnostic EP Catheter Cables, Model Numbers: D6-DR-252-RT, F6-QF-252RT, 201101, 201104, 401442, 401860, D7-A20-131-RT, 401661, 401972, 401977, 401981, 401982, C6-MR/MST4-SA, CB3434CT. |  |  | October 18, 2013 | 2013 | Stryker Sustainability Solutions | Stryker Sustainability Solutions is recalling the Reprocessed Diagnostic Electrophysiology (EP) Catheters Cables due to a potential sterility breach in the package seal. | Process control | Stryker sent an Urgent Medical Device Removal Notification letter to affected customers. The letter informed customers of the affected product, problem identified and the actions to be taken. Customers are instructed to complete the "Recall Effectiveness Form" and fax to Moira Barton-Varty, Senior Director Regulatory Affairs at (480) 809-4957. Customers with any serious adverse events or product problems are instructed to call Stryker Sustainability Solutions at (888) 763-8803. |  | 46 | 46 units total | US Distribution: in the states of: Arizona, California, and Ohio. | Terminated | December 06, 2013 | 49 | 46 | Undetermined
Z-0351-2014 | 66727 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Spacelabs Healthcare qube Compact Patient Monitor, Model 91390. It is a patient monitor with arrhythmia detection or alarms. |  |  | October 17, 2013 | 2013 | Spacelabs Healthcare, Llc | Spacelabs Healthcare qube Compact Patient Monitor, Model 91390, may fail to power ON due to incorrect programmed microcontroller component. | Nonconforming Material/Component | Spacelabs Healthcare sent the "URGENT -MEDICAL DEVICE CORRECTION Spacelabs Healthcare qube Compact Monitor" letter, dated 01 November 2013, to the US consignees/customers on November 1, 2013. Spacelabs Healthcare will send the customer letter (translated as necessary) to all international subsidiaries and distributors of record via email on November 6, 2013. The letter described the product, problem and actions to be taken. The customers were instructed to advise staff of the potential for the monitor to fail to turn ON.Spacelabs Healthcare will contact their customers at the earliest possible date to schedule a convenient time for Spacelabs to correct all of your facility's affected monitors at no cost. If you have any questions about this corrective action program, please contact Spacelabs Healthcare at 1-800-522-7025 and select 2 for Technical Support. |  | 54 | 54 units total (42 units in the US and 12 outside US) | Worldwide distribution: US (nationwide) including states of: Georgia, Idaho, New York, Oklahoma, Tennessee, Texas, and Wyoming and in countries of: Bahrain, Canada, France, Great Britain, Poland, and The Netherlands. | Terminated | April 02, 2014 | 167 | 54 | Undetermined
Z-0347-2014 | 65584 | No | NEW | Absorbable polydioxanone surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Polyglycolic-Lactic Acid Absorbable Surgical Suture |  |  | December 13, 2012 | 2012 | Riverpoint Medical, LLC | Riverpoint Medical is recalling PGLA Surgical Suture (Vilet) because the label indicates the suture is undyed and it is actually a violet colored suture. | Employee error | Riverpoint sent an Urgent Medical Device Recall letter in December 2012 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were requested to review their inventory and immediately quarantine the product in question, if found. In addition, if they have further distributed the recalled product, the consignees were advised to identify their customers and notify them of this product recall. All affected product is to be returned to Riverpoint Medical. Consignees were asked to call 503-517-8001 to return recalled product and request a credit or replacement.For questions regarding this recall call 503-517-8001. |  | 1932 | 1932 pieces | US Distribution in the states of: AZ, FL, and TX. | Terminated | November 21, 2013 | 343 | 1932 | Undetermined
Z-0365-2014 | 66396 | No | CIG | Bilirubin (total or direct) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Siemens Chemistry Calibrator (CAL) used with the ADVIA Chemistry Direct Bilirubin_2 Reagent, Siemens Material Number (SMN) 10312279, REF 09784096, Part Number T03-1291-62; IVD |  |  | September 24, 2013 | 2013 | Siemens Healthcare Diagnostics | Siemens Healthcare Diagnostics is conducting a field correction for Siemens Chemistry Calibrator Lots 090445B and 090445C on the ADVIA 1200, 1650, 1800, and 2400 Chemistry Systems. The System Specific Values (SSVs) for the Direct Bilirubin Assay (DBIL_2) have been reassigned for these lots. Siemens has determined that the DBIL_2 SSVs currently assigned to these lots may cause an average shift of | Under Investigation by firm | Siemens Healthcare Diagnostics sent an Urgent Medical Device Correction Letters dated September 24, 2013 and Response Forms to the US consignees via FedEx. Urgent Field Safety Notices were also distributed to customers outside the US. Customers are instructed to use the reassigned System Specific Values for the affected lots listed in the letter. Customers are requested to complete and return the Effectiveness Check questionnaire attached to the Urgent Medical Device Recall letter within thirty (30) days. The questionnaire is to be faxed to the Technical Solutions Center at 302-631-7597. On September 30, 2013, Siemens sent a revised customer notifications to correct the Part Number. |  | 4243 | Domestic: 1191 units; Foreign: 3052 units | Worldwide Distribution - USA Nationwide in the states of: Argentina, Austria, Belgium, Bulgaria, Canada, Chile, China, Colombia, Croatia, Czech Republic, Denmark, Egypt, Estonia, France, French Polynesia, Germany, Guadeloupe, Guatemala, Great Britain, Greece, Korea, Indonesia, Israel, Italy, Kazakhstan, Latvia, Liechstenstein, Lithuania, Luxembourg, Macedonia, Martinique, Mexico, Netherlands, Norway, Peru, Poland, Portugal, Russian Federation, Serbia, Singapore, Slovenia, Slovakia, Spain, Sweden, Switzerland, Thailand, Taiwan, Tunisia, Turkey, United Arab Emirates, Uruguay, Vatican City, and Vietnam. | Open, Classified |  |  | 4243 | Undetermined
Z-0358-2014 | 66391 | No | LRJ | General purpose disinfectants. | General Hospital | General Hospital | 510(K) Exempt | Ready to Use Thymo-Cide, Germicidal Detergent, containing 0.051% thymol, packaged in 1-qt. bottles. The firm name on the label is Wexford Labs, Inc., Kirkwood, MO. |  |  | September 10, 2013 | 2013 | Wexford Labs Inc | Expired active ingredient was used in the manufacture of the product. | Nonconforming Material/Component | Wexford Labs, Inc. sent undated letters by e-mail to affected customers beginning September 10, 2013, and follow-up phone calls were initiated requesting return of the affected product. E-mail letters dated September 26, 2013, were issued requesting return of product and asking the customers to notify their downstream customers. Customers with questions were instructed to call 800-506-1146.For questions regarding this recall call 314-966-4134, |  | 7020 | 7,020/1-qt. bottles (1755 gallons) | Nationwide Distribution - USA including FL, SC, TX, MO, CA, OH, OK, MA, IN, and GA. | Terminated | April 23, 2014 | 225 | 7020 | Undetermined
Z-0352-2014 | 66721 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | CAS I/II Absorbers that are parts of the BleaseSirius Anesthesia Workstation, the BleaseFocus Anesthesia Workstation with the following part numbers: 12200900; 12200901; 12200902; 12200903; 14200100, and service kits part number 050-0659-00 and 050-0901-00. |  |  | October 17, 2013 | 2013 | Del Mar Reynolds Medical, Ltd. | Spacelabs Healthcare Ltd. of United Kingdom recalls , and Service Kits, Part Number 050-0659-00 and 050-0901-00. The Bag to-Vent switch in CAS I/II Absorbers may fail due to loose fastening hardware in an absorber. | Process control | Spacelabs Healthcare sent an Urgent Medical Device Correction letter dated November 1, 2013, to all affected customers. Spacelabs Healthcare also sent a customer letter (translated as necessary) to all international subsidiaries and distributors of record via e-mail on November 6, 2013. The letter informed customers of a potential defect in CAS I/II Absorber products employed in the Spacelabs Healthcare BleaseSirius Anesthesia Workstation, and Service Kits, Part Number 050-0659-00 and 050-0901-00.The Bag-to-Vent switch in CAS I/II Absorbers may fail. In the worst case, this would result in a large leak from the absorber circuit and could impede the ability to provide ventilation.Customers were instructed to immediately advise their staff of the problem. Customers with questions were instructed to contact Spacelabs Healthcare at 1-800-522-7025 and select 2 for Technical Support.For questions regarding this recall call 800-522-7025. |  | 37 | 37 units (19 in the US and 18 International) | Worldwide Distribution - USA (nationwide) including Oklahoma and North Carolina and Internationally to Libyan Arab Jamahiriya, Finland, Jordan, Mexico, Papua New Guinea, and Taiwan. | Terminated | April 24, 2014 | 189 | 37 | Undetermined
Z-0366-2014 | 66640 | No | CCW | Rigid laryngoscope. | Anesthesiology | Anesthesiology | 510(K) Exempt | Green Rusch - Lite Disposable Metal Laryngoscope Ridge Blade to visualize upper airway and aid in placement of tracheal tube. |  |  | October 15, 2013 | 2013 | Teleflex Medical | Labeling inconsistency | Other | Recall letters were sent to customers and distributors on 10/15/2013 via courier. The letter indicated that the consignee was to immediately discontinue use and quarantine any products with the affected catalog numbers. The letter also included a Recall Acknowledgment Form which was to be faxed back to the firm. |  | 6200 | 6,200 ea. | Nationwide distribution: AL, AZ, CA, CO, FL, GA, IL, IN, IA, KY, LA, MD, MA, MI, MN, MO, NE, NJ, NM, NY, NC, OH, OK, OR, PA, SC, TN, TX, VA, WA, WV and WI. | Terminated | November 04, 2015 | 750 | 6200 | Undetermined
Z-0137-2014 | 66353 | No | JCA | Bleeding time device. | Hematology | Hematology | 510(K) Exempt | Surgicutt fully automated incision making device for bleeding time determination. ITC 8 Olsen Avenue, Edison, NJ 08820. 732.548.5700. www.itcmed.com. A subsidiary of Thoratec Corporation. |  |  | August 22, 2013 | 2013 | ITC-Nexus Dx | ITC has determined that the blade in their Surgicutt with Blotting Paper (Product Code: SUB50I; Lot: AK301M) devices may be loose. | Nonconforming Material/Component | ITC sent a Medical Device Recall letter dated September 20, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to remove/quarantine affected product in their inventory and to return the affected product to ITC. Customers with questions were instructed to call 1-800-631-5945.Questions regarding this recall call 732-548-7000, ext 4476. |  | 300 | 6 boxes (Each box contains 50 Surgicutt Devices plus 55 sheets of Blotting Paper). | Nationwide Distribution including Arizona andTexas.. | Terminated | June 09, 2014 | 291 | 300 | Undetermined
Z-0435-2014 | 66758 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | VersaSafe Extension Set, Model No. 21000M-007 |  |  | November 08, 2013 | 2013 | CareFusion 303, Inc. | CareFusion is recalling the VersaSafe Extension IntravascularAdministration Set because it may have excessive adhesivewhich will prevent the male spin collar to operate as intended. | Employee error | A customer letter and a distributor letter will be distributed on 11/8/13 to all customers who purchased the VersaSafe Extension Set, Model 21000M-07, Lot Number 13066550. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to complete and return to CareFusion the enclosed, pre-addressed and postage paid, Recall Response Card. |  | 3500 | 3500 units | Nationwide Distribution. | Terminated | April 17, 2014 | 160 | 3500 | Undetermined
Z-0450-2014 | 66759 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | SmartSite Low Sorbing Infusion Set with Texium Close Male Luer, Model No. 24301-0007T |  |  | November 08, 2013 | 2013 | CareFusion 303, Inc. | CareFusion is recalling the SmartSite Low Sorbing Infusion Set with Texium Closed Male Luer because it is not assembled with low-sorbing tubing. | Employee error | The firm, CareFusion, sent an "URGENT: Medical Device Recall Notification" letter dated November 8, 2013 to all customers who purchased the SmartSite Low Sorbing Infusion Set with Texium Closed Male Luer. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to immediately complete and return to CareFusion the enclosed, pre-addressed and postage paid, Recall Response Card and return all recalled product directly to the distributor form whom it was purchased. Customers with questions are instructed to contact their distributor directly or call the CareFusion Support Center at (888) 562-6018.If you have recall related questions or need support, please call CareFusion Support Center at 1-888-562-6018; hours: 7am to 4pm PST or Technical Support at 1-888-812-3229; hours 7am to 5pm PST. |  | 480 | 480 units | US Distribution in CA, WY, OH, UT, and MA. | Terminated | May 27, 2014 | 200 | 480 | Undetermined
Z-0556-2014 | 67014 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | McKesson Radiology 12.0, Picture Archiving Communication System (PACSA). |  |  | October 15, 2013 | 2013 | Mckesson Information Solutions LLC | When merging two patient records into one patient record using McKesson Radiology 12.0, the resulting patient record is missing the Contrast Allergy information for the second patient record. | Other | Consignees were notified via letter (Field Safety Notice) on 10/15/2013. |  | 6 | 6 units | MO, NY, TN, WV and France. | Terminated | June 03, 2014 | 231 | 6 | Undetermined
Z-0533-2014 | 66885 | No | CAZ | Anesthesia conduction kit. | Anesthesiology | Anesthesiology | 510(k) | Integra Disposable, Convenience Kit Single Shot Epidural Tray Reorder Number 340269 Rx Only |  |  | October 28, 2013 | 2013 | Integra LifeSciences Corp. d.b.a. Integra Pain Management | Integra LifeSciences Corporation is recalling certain lots of Integra Kits for Epidural Anesthesia Administration as a result of a Product Advisory Notice from BD Medical for the LOR syringes indicating the plunger of the syringe may stall or stick when travelling within the syringe. | Component design/selection | BD Medical sent a Product Advisory Notice letter dated August 23, 2013 to Iaffected consignees by via traceable courier. The letter identified the affected product, problem and actions to be taken. Distributor have been asked to forward copies of the notification to their customers. For questions call BD Customer Service at 1-888-237-2762. |  | 359250 | 359,250 | USA Nationwide Distribution and one foreign consignee in Mursalat. | Terminated | January 29, 2014 | 93 | 359250 | Undetermined
Z-0528-2014 | 65734 | No | JTN | Antimicrobial susceptibility test disc. | Microbiology | Microbiology | 510(k) | Oxoid Antimicrobial Susceptibility Test Discs, CTO425B Ciprofloxacin CIP 5 mcg, 50 discs per cartridge, 5 cartridges per box. The firm name on the box label is Oxoid Ltd., Basingstoke, UK. |  |  | November 21, 2012 | 2012 | Remel Inc | Boxes labeled as CIP5 may contain cartridges of AMP5. | Mixed-up of materials/components | Thermo Fisher sent a Field Safety Notice dated November 13, 2012, to all affected customers. The letter identified the product, the problem and the action to be taken by the customer. Customers were instructed to check stocks of the affected product if necessary and contact Customer Service or their local Oxoid supplier who will issue a replacement. The notice should be passed on to all who need to be aware within their organization. Customers were also instructed to discard any product labeled as AMP5. For questions, customers were instructed to call +44 (0)1256 841144.discard any product labeled AMP5.For questions regarding this recall call 913-895-4185. |  | 300 | 60/5/50-disc cartridge boxes | Nationwide Distribution including AL, CA, CO, FL, IA, KS, LA, MA, MS, NE, NH, NJ, NY, OK, OR, RI, TN, TX, VA, WA WI, and Puerto Rico. | Terminated | March 12, 2014 | 476 | 300 | Undetermined
Z-0539-2014 | 66989 | No | LNX |  |  | General Hospital | 510(k) | ABACUS TPN Calculation Software, Baxa: 8300-0045 ABACUS CE (Calculator Edition), 8300-0046 ABACUS SE (Single-Workstation Edition), 8300-0047 ABACUS ME (Multi-Workstation Edition) |  |  | December 03, 2013 | 2013 | Baxter Corporation Englewood | Baxter Healthcare Corporation is issuing a voluntary correction for ABACUS Total parenteral Nutrition (TPN) Calculation Software versions 2.0 through 3.0 due to the possibility that a patient could receive another patient's TPN order as a result of data corruption. | Unknown/Undetermined by firm | Baxter sent an Urgent Device Correction letter dated December 4, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check the patient order against the final bag label as part of the TPN acceptance procedure to ensure that the patient's TPN order is correct. Customers will be contacted by Baxter to arrange for installation of a software upgrade specific to their compounder's configuation. Customers were asked to acknowledge the receipt of this notification by completing the attached Customer Reply Form and return it to Baxter by fax to 1-224-270-5457 or scanning and emailing it to fca@baxter.com. If the product was distributed to other facilities customers were instructed to notify end-user customers in accordance with their customary procedures. Customers with questions were instructed to call 1-800-678-2292.For questions regarding this recall call 800-422-9837. |  | 2211 | 2,211 | Nationwide Distribution | Terminated | October 16, 2015 | 682 | 2211 | Undetermined
Z-0535-2014 | 66944 | No | ELW | Impression material. | Dental | Dental | 510(k) | 3M ESPE Imprint II Regular Body , Catalog Number 9379, and Paradigm Regular Body VPS, Catalog Number 5315, Impression Material that are intended to reproduce the structure of a patients teeth. |  |  | November 21, 2013 | 2013 | 3M Company - Health Care Business | Specific lots of Imprint II and Paradigm impression material refill packs are being recalled because the moisture displacement feature in both is greatly reduced which may impact the quality of the dental impression. | Employee error | 3M ESPE sent an Voluntary Recall Notification Letter dated November 21, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.Customers were instructed to:1) Review your inventory and identify any products from these lot numbers2) Quarantine products from these lots and do not further distribute3) If you have shipped any products from these lots to your customers, inform them of this voluntary recall using the attached Dear Doctor Letter and Inventory Checklist4) To return and replace product from these lots, contact 3M ESPE Dental ProductsCustomer Service at 1-800-237-1650If you have any questions, please contact the above number for assistance. Once again, we sincerely apologize for any inconvenience this may cause you. |  | 1977 | 1977 kits | Worldwide Distribution - USA (nationwide) and the countries of Canada, France, Germany, Hong Kong, Italy, Japan, Korea, Mexico, United Arab and Emirates. | Terminated | June 02, 2014 | 193 | 1977 | Undetermined
Z-0537-2014 | 65476 | No | JTO | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | Streptex* kit, Ref R30950501 ZL50, Rapid latex test for Streptococcal grouping, IVD. The kit contains reagent groups A, B, C, D, F, G, a control, and an extraction enzyme. The firm name on the label is Remel Europe Ltd., Dartford, England. |  |  | May 02, 2013 | 2013 | Remel Inc | The product may give weak or false negative test results. | Process control | ThermoFisher sent a Medical Device Recall letter dated April 26, 2013, or May 2, 2013, to all affected customers via first class mail. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to notify any personnel who need to be notified of the potential for a false negative test result, review their inventory for affected product lots and discard the affected product. Customers were also instructed to return the attached Medical Device Recall Return Response acknowledging the disposal of the affected product. Customers with questions were instructed to contact the Technical Services Department at 800-255-6730 (US) or 913-888-0939 (International).For questions regarding this recall call 913-888-0939. |  | 768 | 768 kits | Nationwide Distribution | Terminated | March 12, 2014 | 314 | 877 | Undetermined
Z-0541-2014 | 66363 | No | LPS | Urinary homocystine (nonquantitative) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Siemens Dimension Vista Homocysteine Flex Reagent Cartridge |  |  | September 18, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Dimension Vista HCYS lot 13064MA shows erratic recovery of QC and patient samples. | Unknown/Undetermined by firm | Siemens sent a Urgent Medical Device Recall letter dated September 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.Actions to be taken by Customer:Please discontinue use and discard any remaining inventory of Dimension VistaÂż HCYS FlexÂż reagent cartridge (Catalog# K7044) lot 13064MA.Please complete the attached form and fax it to (302) 631-8467 to facilitate your no charge replacement needs.Please retain this letter with your laboratory records, and forward this letter to those who may have received this product.We apologize for the inconvenience this situation has caused. If you have any questions, please contact your Siemens Customer Care Center- Technical Solutions or your local Siemens technical support representative. For further questions please call (800) 441-9250. |  | 302 | 302 | Worldwide Distribution - US Distribution including the states of OH, MD, CO, NJ, IA, NC, VA, WI and DC., and the countries of Austria, Denmark, France, Italy, Netherlands and Norway. | Terminated | October 01, 2014 | 378 | 302 | Undetermined
Z-0550-2014 | 66842 | No | MRG | Vitamin D test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | FastPack Free T4 Immunoassay |  |  | November 08, 2013 | 2013 | Qualigen Inc | Qualigen initiated this recall because the FastPack kit may cause lower than expected results due to a manufacturing defect. | Process change control | Qualigen, Inc., notified customers of the recall by calling customers via telephone (with a telephone script) on 11/08/2013. The telephone (script) included the reason for recall, product description with codes, instructions and contact information. The customers were instructed to examine your inventory and discontinue use of the kits immediately. Qualigen will send youreplacement kits and provide you with a shipping label to return all of the defective kits including any that were partially used.Feel free to contact System Support at 760.579.6900 if you have anyquestions or require assistance with returning your kits. |  | 6 | 6 | US Distribution to states of: AL, CA, KS, KY, IN, MA, NC, NY, NV, OK, SC, TN, TX, UT, VA, and WA. | Terminated | September 18, 2014 | 314 | 223 | Undetermined
Z-0557-2014 | 66900 | No | LOM |  |  | Microbiology | PMA | VITROS Anti-HBs Calibrators &amp; VITROS Anti-HBs Reagent Packs |  |  | October 25, 2013 | 2013 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics issued a Product Correction Notification for VITROS Immunodiagnostic Products Anti-HBs Calibrators and Anti-HBs Quantitative Reagent Packs due to false positive results. | Other | On October 25, 2013, Ortho Clinical Diagnostics issued product correction notifications to their US customers via US Postal Service Priority Mail and email notices to their foreign affiliates informing them of the recall. Ortho Clinical Diagnostics advised their customers to review previously reported results between &gt; 12 and &lt;30 mIU/mL and consult the Laboratory Medical Director. Customers were instructed to utilize alternative methods for anti-HBs testing. Customers are advised to seek alternative methods for anti-HBs testing. For additional information, please contact the OCD Customer Technical Service representative at 1-800-421-3311. |  | 1920 | 1,920 kits | Nationwide distribution - all states. | Open, Classified |  |  | 1920 | Undetermined
Z-0558-2014 | 66971 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | HT Connect Peripheral Guide Wire |  |  | November 25, 2013 | 2013 | Abbott Vascular, Inc. | Abbott Vascular has initiated a recall of the HT Connect Peripheral Guide Wires due to a small number of devices exhibiting partial delamination of the PTFE coating. | Pending | Abbott Vascular sent an Urgent Device Recall letter dated November A recall communication letter dated November 25, 2013, to all affected customers. The letter informed customers of the problems identified and the actions to be taken. The letter informs the customers of the 9 part numbers and 9 description of the products. Abbott Vascular is instructing the customers that the returned products will be replaced with similar products pending availability. Abbott Vascular has attached a customer response form for customers along with the customer notification letter.For questions regarding this recall call 951-914-2400. |  | 54895 | 54,895 units total (5,356 units in US) | Worldwide Distribution - USA (nationwide) and Internationally to Australia, French Guyana, Macedonia, South Korea, Austria, Germany, Malaysia, Spain, Bahrain, Greece, Martinique, Sweden, Belgium, Hong Kong, Netherlands, Switzerland, Brazil, Indonesia, New Zealand, Taiwan, Bulgaria, Iran, Norway, Thailand, Canada, Iraq, Poland, Turkey, China, Ireland, Portugal,United Arab Emirates, Czech Republic, Italy, Saudi Arabia, United Kingdom, Denmark, Latvia, Singapore, Estonia, Libya, Slovakia, Finland, Lithuania, Slovenia, France, Luxembourg, South Africa. | Terminated | July 14, 2014 | 231 | 54895 | Undetermined
Z-0569-2014 | 66713 | No | CGN | Folic acid test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | IMMULITE/IMMULITE 1000 Systems FOL Folic Acid, REF/Catalog Number LKFO1 (100 tests), Siemens Material Number (SMN) 10380902, IVD; for in vitro diagnostic use with the IMMULITE and IMMULITE 1000 Analyzers --- Siemens Healthcare Diagnostics Products Ltd. |  |  | November 06, 2013 | 2013 | Siemens Healthcare Diagnostics | potential for an accelerated decline in counts per second (CPS) for three IMMULITE/IMMULITE 1000 Folic Acid kit lots. This may lead to a failed adjustment before the labeled expiration date is reached. | Unknown/Undetermined by firm | Siemens sent an Urgent Medical Device Correction notices dated November 7 2013, via FedEx to the US customers. The letter identified the product the problem and the action needed to be taken by the customers.Customers are instructed that If the adjustment slope is less than or equal to 1.8 and quality control results are within the established ranges, the kit may be used to generate patient results. If the adjustment slope is greater than 1.8, customers are instructed to discontinue use and discard the kit lots. Ex-US distribution is determined at the country level. Customers are requested to complete and return the Effectiveness Check questionnaire attached to the Urgent Medical Device Correction letter within thirty (30) days. Field service personnel were sent a support document describing the issue and instructing them how to deal with customer questions.For further questions please call CUSTOMER CARE CENTER NUMBER: (302) 631-7597. |  | 1122 | Domestic: 442 kits; Foreign: 680 kits | Worldwide Distribution - USA (nationwide) including Puerto Rico and the countries of Argentina, Bosnia-Herzegovina, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Cyprus, Denmark, Dominican Republic, Estonia, France, Germany, Great Britain, Greece, Guatemala, Hungary, India, Italy, Latvia, Mexico, Netherlands, Panama, Paraguay, Poland, Portugal, Qatar, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, Spain, Tunisia, United Arab Emirates, Uzbekistan and Venezuela. | Terminated | March 23, 2017 | 1233 | 1122 | Undetermined
Z-0525-2014 | 66496 | No | LTK | Tumor-associated antigen immunological test system. | Immunology | Immunology | 510(k) | VITROS CA 125 II Calibrators |  |  | October 09, 2013 | 2013 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics issued a Product Correction Notification for VITROS Immunodiagnostic Products CA 125 II Reagent Pack and Calibrators due to a positive shift in patient results. | Under Investigation by firm | The firm, Ortho Clinical Diagnostics, sent a "IMPORTANT PRODUCT CORRECTION NOTIFICATION" letter dated October 9, 2013 to their clients/customers informing them of a shift in results that may be observed when performing patient sample correlation tests between VITROS CA 125 II Calibrators and Reagent Packs. The letter described the product, problem and actions to be taken. The customers were informed that it is acceptable to continue using any remaining inventory until its expiration date providing that quality control results are within acceptable limits (Note: Lots 1220 and 1230 are expired). The customers were also instructed to consider the need to review patient results using the affected lot(s). If you suspect that a previously reported result may have been affected, provide this information to your Laboratory Medical Director and the requesting physician or health care provider so that appropriate actions may be taken for the patients; Complete and return the attached Confirmation of Receipt form no later than October 18, 2013 via FAX TO: Ortho Clinical Diagnostics at 1-888-557-3759 or 1-585-453-4110, and Forward the information in this notification if you have distributed this product outside of your facility.If you have any additional questions, please contact Customer Technical Services at 1-800-421-3311. |  | 601 | 601 units | Worldwide Distribution: US (nationwide) including states of: AK, CA, CT, CO, DC, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, ME, MI, MN, MO, MS, MT, NC, NE, NJ, NM, NY, OH, OK, OR, PA, SC, TN, TX, VA, VT, WA, WI, WV, WY, and Puerto Rico, and countries of: Australia, Brazil, Canada, Chile, China, Colombia, India, Japan, Mexico, Republic of Panama, Singapore, Venezuela, England, France, Germany, Italy &amp; Spain. | Open, Classified |  |  | 601 | Undetermined
Z-0572-2014 | 66916 | No | LIP | Treponema pallidum treponemal test reagents. | Microbiology | Microbiology | 510(k) | Captia Syphilis-G Test Kit |  |  | August 27, 2013 | 2013 | Clark Laboratories, Inc. (dba,Trinity Biotech USA) | Trinity Biotech recalled Trinity Biotech Captia Syphilis IgG kits due to false negative results. | Process design | Trinity Biotech issued an Urgent Product Information Request letter to their customer on August 27, 2013. The letter identified the affected product, problem and actions to be taken. The request was to determine if their customer did or did not backfill microtiter-strips during the analytical process. If a customer backfilled, they were advised to return the completed request form. On September 11, 2013 Trinity Biotech issued a notice to their customers advising them to re-evaluate any negative patient results for those who perform backfilling. For questions contact Trinity Biotech Help Desk at 800-325-3424, option 2. |  | 38 | 38 kits | USA Nationwide Distribution in the state of WA | Open, Classified |  |  | 38 | Undetermined
Z-0581-2014 | 67042 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | REFLECTION(R), MICROSTABLE(R), ACETABULAR LINER, 20o OVERHANG, 28 MM ID, 54-56 OD, QTY: (1), STERILE EO, REF 71742854, Smith &amp; Nephew, Inc. Memphis, TN 38116 USA. |  |  | November 25, 2013 | 2013 | Smith &amp; Nephew Inc | One batch of the Reflection(TM) Polyethylene Acetabular Liners may be out of tolerance causing the liner not to fully seat into the shell. | Under Investigation by firm | All affected sales staff was notified of problem and the recall via email on 11/25/2013. There were no US hospitals affected by this recall. |  | 38 | 38 units | Distributed in China and Spain. | Open, Classified |  |  | 38 | Undetermined
Z-0603-2014 | 67064 | No | MNJ | Cardiopulmonary bypass blood reservoir. | Cardiovascular | Cardiovascular | 510(k) | Terumo Cardiovascular Custom Cardiovascular Procedure Kits containing an Overpressure Safety Valve, LN130B, in the Arterial Recirculation Line |  |  | June 04, 2013 | 2013 | Terumo Cardiovascular Systems Corporation | Over-Pressure Safety Valves not cleared with an indication for use on the arterial side of the extracorporeal bypass circuit. | Component design/selection | Terumo Cardiovascular issued Customer notification letters on June 4, 2013. Accounts advised to: Review this notice and decide to continue to use or not use the Cardiovascular Procedure Kits as configured until a replacement kit with a one-way valve indicated for use on the arterial side of the perfusion circuit is available. If you have questions, contact Terumo CVS Customer Service: 800-521-2818. |  |  | 24,227. | NationwideCanada | Terminated | October 01, 2014 | 484 |  | Undetermined
Z-0587-2014 | 66499 | No | FZS | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | DIST LAT FEM LOCK PLT, LT |  |  | November 05, 2013 | 2013 | Zimmer, Inc. | Based on a review of historic packaging validations, it was determined that a subset of product packaged at a Zimmer facility required testing to ensure that previously packaged product from this facility met the established requirements for sterile barrier integrity. This testing determined that a packaging tray configuration which consists of an inner and outer sterile tray, resulted in failure | Process change control | On 11/5/2013, Urgent Medical Device Recall Notifications were sent to affected distributors, hospital Risk Managers, and doctors informing them of the failure. All distributors were notified via electronic mail. Hospital risk managers, surgeons and distributors with affected inventory within the expiration date of the device will also be notified via courier. The letter identifies the issue, health risks, and their responsibilities. Questions concerning the recall are directed to 1-877-946-2761Communications outside of the United States will occur approximately two weeks after the United States communications.Zimmer will conduct effectiveness checks in the following way:Distributors/Hospitals/Surgeons100% of the notifications will be tracked to ensure delivery of the notifications. An inventory certification form will need to be returned from each distributor that certifies that all available affected product from their territory has been returned.Additional notifications via email and FedEx will be sent to those who have not completed the required certification form or returned affected product.Accounts will be deemed unresponsive after 3 attempts. |  | 42 | 42, 065 units total | Worldwide Distribution. | Terminated | May 05, 2015 | 546 | 42 | Undetermined
Z-0584-2014 | 67012 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips Essenta DR, a Stationary x-ray system, X-ray system |  |  | December 05, 2013 | 2013 | Philips Medical Systems, Inc. | X-ray system C-arm casting may loosen and fall | Nonconforming Material/Component | Philips Medical Systems issued a Urgent Field Safety Notice-Warning Label and Inspection, dated November 21, 2013, advising of potential of C-arm detaching and requesting a daily inspection and a safety warning label . If you need any further information or support concerning this issue, please contact your local Philips representative. 1-800-722-9377. |  | 424 | 424 units | Worldwide Distribution - USA (Nationwide) including Puerto Rico and the countries of : Afghanistan, Argentina, Australia, Brazil, Bulgaria, Canada, Chile, China, Colombia, Czech Republic, Egypt, France, Germany, Greece, Hungary, India, Indonesia Iran,Italy, Jordan, Kazakstan, Kenya, Liberia, Lithuania, Morocco, Netherlands, New Zealand, Nigeria, Poland, Portugal, Russia, Saudi Arabia, Singapore, Slovakia, South Korea, Spain, Sweden, Thailand, Turkey, Ukraine, United Arab, United Kingdom and Viet nam. | Open, Classified |  |  | 424 | Undetermined
Z-0605-2014 | 67036 | No | FRC | Sterilization process indicator. | General Hospital | General Hospital | 510(k) | VERIFYÂż Dual Species Self-Contained Biological Indicator, Distributed by STERIS Corporation, Mentor, OH. |  |  | September 10, 2013 | 2013 | Steris Corporation | STERIS has identified that the population of the G. stearothermophilus spores that monitor STEAM sterilization cycles does not meet specifications for Lot # 131004 of the Verify Dual Species Self-Contained Biological Indicators. As a result the spore count on the affected product is lower than the labeled minimum and the product may incorrectly indicate that a 5-log spore kill was achieved. | Use error | On September 10, 2013, the firm sent Urgent Voluntary Recall Notices to their customers. |  | 1238 | 1,238 boxes | Worldwide Distribution-USA (nationwide) including the states of AK, Al, AR, AZ, CA, CO, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV and WY, and the countries of BAHRAIN, CANADA, COLUMBIA, GERMANY, ITALY, MALAYSIA, PHILIPPINES, ROMANIA, SPAIN, and THE NETHERLANDS. | Terminated | February 14, 2014 | 157 | 1238 | Undetermined
Z-0604-2014 | 66981 | No | FRC | Sterilization process indicator. | General Hospital | General Hospital | 510(k) | STERRAD Cyclesure Biological Indicator (BI, Part No. 14324. |  |  | November 26, 2013 | 2013 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) is recalling the STERRAD Cyclesure Biological Indicator (BI) because ASP has identified a microscopic intermittent and low frequency flaw in the outer vial of the STERRAD Cyclesure 24 Biological Indicator. | Nonconforming Material/Component | A customer notification letter was sent to all customers on 11/26/2013 who purchased the STERRAD CYCLESURE Biological Indicator (BI). The letter informs the customer of the problems identified and the actions to be taken. The customer notification letter includes a table with the products on recall which includes the manufacturing dates, lot numbers, expiration dates, and quantity. |  | 4048 | 4048 units total (4014 units in US) | Worldwide Distribution. | Terminated | March 05, 2014 | 99 | 4048 | Undetermined
Z-0607-2014 | 66979 | No | NEP | Intravascular administration set. | General Hospital | General Hospital | 510(K) Exempt | IntelliFill i.v. Pharmacy Compounding Device (PCD). |  |  | November 08, 2013 | 2013 | Baxter Healthcare Corp. | Following a retrospective review for Baxter Daytona Complaint CR-954, it was identified that the IntelliFill i.v. Pharmacy Compounding Device (PCD) (product code: IFIV) has the potential to continue to run and fill syringes with the HEPA filtration disabled. | Nonconforming Material/Component | A Service Bulletin was sent to all IntelliFill i. v. customers February 22, 20 I 0 by USPS mail.The Service Bulletin informs customers that IntelliFill i.v. operators are to be trained to listen for the running HEPA filtration fan motor and observe/feel for airflow coming from the automation deck.The field service engineer dispatched to the site replaced the damaged sensor on 11-Feb-2010 and a Service Bulletin was sent to all customers regarding the function and use of the HEPA filter on 22-Feb-2010.A new HEPA sensor was designed and installed into all current IntelliFill i.v. devices between 21-July-2011 to 9-Feb-2012. |  | 28 | 28 devices | Nationwide Distribution including the states of CA, CO, CT, FL, IL, LA, MA, MD, MI, NC, NE, NY, OH, OR, PA, TN, TX, and WV. | Terminated | March 23, 2015 | 500 | 28 | Undetermined
Z-0613-2014 | 67035 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart MRx Monitor/Defibrillator |  |  | April 10, 2013 | 2013 | Philips Medical Systems, Inc. | Device Operating on Battery Power May Shutdown without Warning if exposed to elevated levels of electromagnetic interference from other radio frequency (RF) energy sources and | Environmental control | Philips issued a Medical Device Correction letter April 2013 to all affected customers along with an IFU Addendum titled, HeartStart MRx Electromagnetic Compatibility. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed that the condition could be avoided by operating the device on external AC or external DC power. Customers were instructed to follow the device instructions when using battery power to minimize or eliminate potential interference from other RF sources. Customers with questions were instructed to contact their local Philips representative.For questions regarding this recall call 978-687-1501. |  | 64972 | 64,972 | Worldwide Distribution - USA (nationwide) and Internationally to the following countries:AFGHANISTANALBANIAARGENTINAARUBAAUSTRALIAAUSTRIAAZERBAIJANBAHRAINBANGLADESHBELGIUMBOLIVIABOSNIA AND HERZEGOWINABRAZILBRUNEI DARUSSALAMBULGARIACAMBODIACANADACHILECHINACOLOMBIACOSTA RICACROATIA (local name: Hrvatska)CYPRUSCZECH REPUBLICDENMARKECUADOREGYPTEL SALVADORESTONIAETHIOPIAFINLANDFRANCEGERMANYGREECEGUADELOUPEHONG KONGICELANDINDIAINDONESIAIRAQIRELANDISRAELITALYJAPANJORDANKAZAKHSTANKENYAKOREA, REPUBLIC OFKUWAITLAOSLATVIALEBANONLIBYAN ARAB JAMAHIRIYALITHUANIALUXEMBOURGMACAUMACEDONIAMALAYSIAMALTAMEXICOMONGOLIAMOROCCONAMIBIANEPALNETHERLANDSNETHERLANDS ANTILLESNEW CALEDONIANEW ZEALANDNICARAGUANIGERIANORWAYOMANPAKISTANPALESTINEPANAMAPAPUA NEW GUINEAPARAGUAYPERUPHILIPPINESPOLANDPORTUGALPUERTO RICOQATARREUNIONROMANIARUSSIAN FEDERATIONSAUDI ARABIASERBIASINGAPORESLOVENIASOUTH AFRICASPAINSRI LANKASWEDENSWITZERLANDTAIWAN, PROVINCE OF CHINATANZANIA, UNITED REPUBLIC OFTHAILANDTRINIDAD AND TOBAGOTUNISIATURKEYTURKMENISTANUGANDAUKRAINEUNITED ARAB EMIRATESUNITED KINGDOMUZBEKISTANVENEZUELAVIET NAMYEMEN | Terminated | April 13, 2017 | 1464 | 64972 | Undetermined
Z-0485-2014 | 66886 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | FreeStyle Blood Glucose Test Strips; |  |  | November 18, 2013 | 2013 | Abbott Diabetes Care, Inc. | Certain lots of FreeStyle and FreeStyle Lite Blood Glucose Test Strips produce erroneously low blood glucose results when using FreeStyle Blood Glucose Meters, FreeStyle Flash Blood Glucose Meters and the FreeStyle blood glucose meter built into the OmniPod system. Erroneously low results that are not recognized may pose significant risks to your health. | Other | Abbott Diabetes Care sent an Urgent Product Recall letter in November 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers who were affected were instructed to take the following steps:Call Abbott Diabetes Care Customers Service immediately at 1-888-736-9869 for replacement of the affected product. While waiting for replacement strips to arrive, use an alternate method to measure their blood glucose.If only test strips available to customers were from affected lots, customers should not stop testing their blood glucose. Customers were instructed to dispose of the test strips in accordance with their local and state disposal regulations. Customers with questions were instructed to call Abbott Diabetes Care Customers Service at 1-888-736-9869.For questions regarding this recall call 510-239-2775.Abbott issued a press release at the request of CDRH on november 27, 2013. |  | 141219 | 141,219 cartons for human use - all lot numbers | Worldwide Distribution - USA (nationwide and Puerto Rico) and Internationally to Australia, Austria, Belgium, Denmark, Egypt, Finland, France, Germany, Jamaica, Ireland, Israel, United Kingdom, Luxembourg, Netherlands, Norway, Sweden, Switzerland, UAE, and Canada.**Center Recommended Depth - Consumers/User** | Terminated | August 03, 2015 | 623 | 141219 | Undetermined
Z-0527-2014 | 66862 | No | ECQ | Dental hand instrument. | Dental | Dental | 510(K) Exempt | Broncho Flexible Fiberscope, models: 11001BN1, 11002BD1, 11004BC1, and 11009BC1. |  |  | November 13, 2013 | 2013 | Karl Storz Endoscopy America Inc | KARL STORZ Endoscopy America, Inc. is recalling the Flexible Broncho Fiberscope or a Flexible Intubation Fiberscope with a suction port because the material used in the suction channel portion was thermoplastic polyurethane (TPU), which is not listed in STERRAD systems claims for interior endoscope lumens. | Under Investigation by firm | The firm, KARL STORZ, sent a "A SAFETY ALERT-ATTENTION NEEDED" letter dated November 8, 2013 to all their customers on November 13, 2013. The letter describes the product, problem and actions to be taken.The letter informs the customers that the endoscopes incorporated a lumen material that is only approved for use as an external surface material and not included in the STERRAD NX and 100NX FDA-cleared claims for use in flexible endoscope lumens. The customers were instructed to follow the recommended processing methods included in the letter.Customers with questions about the letter are instructed to contact the undersigned and the Technical Services at (800) 421-0837, ext 5350. |  | 141 | 141 units | Nationwide Distribution. | Terminated | March 27, 2014 | 134 | 1657 | Undetermined
Z-0441-2014 | 66828 | No | FED | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Navigator HD l3F/15F x 36cm Ureteral Access Sheath Set Product No M0062502280, RDF Catalog No 250-228, Sterile. |  |  | November 05, 2013 | 2013 | Boston Scientific Corporation | The 13F/15F Navigator HD Ureteral Access SheathSet are being recalled after receipt of 5 reports where the tip of the dilator has separated fro the body of the dilator. | Component design/selection | Urgent recall notification letters were sent via Federal Express Priority mail on November 5, 2013. The notification instructs the account to check their inventory, remove any affected product from their inventory, and return it toBoston Scientific. The notification also requests the completion of a Reply Verification Tracking Form asa method of documenting the presence or absence of affected product in their inventory. Follow-up communications will be made to the accounts that do not respond to this inquiry. |  | 939 | 939 | Worldwide Distribution-USA (nationwide) including Puerto Rico and the countries of Australia, Ireland, Denmark, Italy, France, Canada, Great Britain, Switzerland, South Africa, Germany, and the Netherlands. | Terminated | December 08, 2014 | 398 | 1662 | Undetermined
Z-0427-2014 | 66674 | No | OPT | Near Infrared (NIR) Brain Hematoma Detector. | Ophthalmic | Ophthalmic | 510(k) | Cutera brand truSculpt infrared System (optional), Massager, Vacuum, radio frequency induced heat, Product is manufactured and distributed by Cutera, Inc., Brisbane, CA |  |  | October 16, 2013 | 2013 | Cutera, Inc. | Cutera received two adverse events associated with the edge of the electrode, when it was pressed against neck skin over the thyroid cartilage and due to the absence of adequate underlying conductive tissue, increased skin heating occurred. Both patients developed linear welts and blisters consistent with partial thickness thermal burns to the skin. | Employee error | Cutera sent an IMPORTANT SAFETY INFORMATION letter dated October 16, 2013, to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. It is important that you share this correspondence and the updated Operator Manual andvtreatment guidelines with all staff at your office that use the truSculpt, so they are aware of this new safety information. Please immediately remove and dispose of the current truSculpt Operator Manual and clinical guidelines and replace them with the new and revised truSculpt Operator Manual and clinical guidelines.Customers with questions were instructed to call 415-657-5500 or 415-657-5568. |  | 451 | 451 | Worldwide Distribution - USA (nationwide) and Internationally to Canada | Terminated | December 09, 2013 | 54 | 451 | Undetermined
Z-0451-2014 | 66757 | No | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | Abbott i-STAT PT/INR cartridges Abbott Point of Care Inc., Princeton, NJ 08540 |  |  | October 28, 2013 | 2013 | Abbott Point Of Care Inc. | Abbott Point of Care Inc (APOC) has determined that i-Stat PT/INR cartridges have the potential to exhibit incorrectly elevated INR. | Nonconforming Material/Component | Abbott Point of Care Inc. (APOC) issued an Urgent Recall Notice on October 28, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use of cartridges from PT/INR lots in this letter. All unused cartridges would be credited and are to be returned as instructed on the enclosed Business Reply Card. If customers provided cartridges from the lot number listed above to another institution they were instructed to provide a copy of this letter to them.Customers with questions were instructerd to contact Abbott Point of Care Technical at 1-800-366-8020, Option 1.For questions regarding this recall call 800-366-8020. |  | 914400 | 794,928 units US; 119,472 units OUS | Worldwide Distribution - USA (nationwide) and Internationally to Argentina, Australia, Brazil, Chile, China, Finland, France, Germany, Greece, Hong Kong, Ireland, Italy, South Korea, Kuwait, Mexico, Netherland, New Zealand, Nigeria, Peru, Poland, Portugal, Qatar, Russia, Saudi Arabia, Singapore, south Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, UAE, and United Kingdom, | Terminated | May 29, 2014 | 213 | 914400 | Undetermined
Z-0455-2014 | 66538 | No | OVE | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Synthes Zero-P VA Implant 7 MM Height Parallel Sterile |  |  | September 30, 2013 | 2013 | Synthes USA HQ, Inc. | A potential non conformance was identified with the Zero-P VA plate where the screw could back out if the blocking mechanism is not visually confirmed to properly cover the head of the screw. | Process design | Synthes sent an email message dated March 23, 2012, to the sales representative who received the affected product. The email identified the product, the problem, and the action to be taken. The letter requested the sales rep examine inventory, stop using the product, and quarantine the affected product to ensure that it will not be used. Synthes instructed the customer not to remove any product from their account or return any product without further instruction. For questions they were instructed to call 610-719-5364.For questions regarding this recall call 610-719-5000. |  | 1 | 1 | Nationwide Distribution - only AZ. | Terminated | April 16, 2015 | 563 | 1 | Undetermined
Z-0453-2014 | 66896 | No | IOR | Mechanical wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Sunrise Brand Zippie Voyage, Chair with Casters, Early Intervention Device (Stroller), Model: EIZ14; Product is manufactured and distributed by Sunrise Medical (US) LLC., Fresno, CA. |  |  | November 11, 2013 | 2013 | Sunrise Medical (US) LLC | There is a remote possibility for the plastic latch to deform over the lifetime of use. When this happens there is the possibility for the seat to become dislodged from the base. And when not using the positioning strap, there is a risk that the occupant could slide out of the seat and sustain injury. | Nonconforming Material/Component | Urgent Medical Device Field correction letters were sent to Sunrise Suppliers/Dealers and Zippie Customers on November 2013 by first class mail. Foreign notification was supplied to suppliers and regulatory affairs director in Canada. |  | 345 | 316 units in US, 29 units international | Worldwide Distribution-USA (nationwide) and the countries of Canada, Australia, Hong Kong, and China. | Terminated | March 18, 2014 | 127 | 345 | Undetermined
Z-0432-2014 | 66407 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | TB-0545PC: THUNDERBEAT 5 mm , 45 em, Pistol Grip TB-0535PC: THUNDERBEAT 5 mm, 35 em, Pi stol Grip TB-0545IC: THUNDERBEAT 5 mm, 45 em, In line Grip TB-0535IC: THUNDERBEAT 5 mm, 35 em, Inline Grip TB-0520IC: THUNDERBEAT 5 mm, 20 em, In line Grip TB-051 OIC: THUNDERBEAT 5 mm, I 0 em, In line Grip |  |  | September 27, 2013 | 2013 | Olympus America Inc. | Complaints of damage to various models of Thunderbeat Hand Instruments during surgical procedures, including cracking, breaking, or deformation of the components of the probe tip or jaw. | Use error | Olympus sent an Urgent Medical Device Field Correction Action letter dated September 27, 2013, to all affected customers. The letter identified the product the problem and the actions needed to be taken by the customer.Please make the attached Procedural Tips and Techniques instructions available to all users of the THUNDERBEAT hand instruments. Please also maintain a copy with the Instruction Manual(s) fot"the device.OCA requires that you indicate on the enclosed questionnaire your acknowledgement of receipt of this corrective action and the contact information for your facility.This additional user instruction is also being added to the Instruction Manual(s) for the THUNDERBEAT hand instruments. If you would like a copy of the updated Instruction Manual(s), please contact OAI Customer Service at 1-800-848-9024.For further questions please call (484) 896-5688. |  | 8709 | 8709 | Worldwide Distribution - USA (nationwide) including Puerto Rico and Internationally to Canada. | Terminated | February 02, 2015 | 493 | 8709 | Undetermined
Z-0452-2014 | 66417 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Titanium Trochanteric Fixation Nail (TFN)System |  |  | July 30, 2013 | 2013 | Synthes USA HQ, Inc. | Two lots of Titanium Trochanteric Fixation Nail (TFN) System are missing the second bend in the nail. | Nonconforming Material/Component | Synthes sent an Urgent Notice: Medical Device Recall letter dated July 30, 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to call Synthes at 1-800-479-6329 if they have any of the affeccted product to obtain a Return Authorization Number, Complete the Verification Section of the letter by checking the appropriate box indicating the affected product has been located. Return the Verification Form with the product to:Credit/ReturnsSynthes1101 Synthes AvenueMonument, CO 80132.Customers should complete and return the Verification Frorm even if they do not have any of the affected product and return by fax to 1-888-386-2077 or Scan/email: FieldAction@synthes.com. Customers with questions were instructed to call 610-719-5450 or contact their Synthes Trauma Sales Consultant.For questions regarding this recall call 610-719-5000. |  | 17 | 17 | Nationwide Distribution including IL, CA, AZ, FL, TX, UT, VA, NY, NM, and MT. | Terminated | September 10, 2015 | 772 | 17 | Undetermined
Z-0457-2014 | 66431 | No | GCX | Vacuum-powered body fluid suction apparatus. | General Hospital | General Hospital | 510(K) Exempt | Product is labeled in part: LIST NO. 43449; RECEPTAL CANISTER; DO NOT FILL 900 ABOVE THIS LINE; Product of China; Hospira, Inc.,; Lake Forest, IL 60045 USA. |  |  | July 17, 2013 | 2013 | Hospira Inc. | Hospira has recieved reports that the vacuum needed to create suction with the 1L Receptal System cannot be achieved due to the canister and liner not properly seating. | Process control | An URGENT DEVICE RECALL letter dated July 17, 2013 was mailed to direct consignees. The letter requested the consignees: 1) quarantine any affected products, 2) complete and return the reply form included with the letter, 3) notify healthcare professionals within the consignee's organization of the recall, 4) return the affected products to Stericycle in accordance with the instructions in the letter. Consignees are further requested to forward a copy of the notification materials to any customers that they may have distributed the products to. |  | 7576 | 3,560 units distributed from June 2004 - December 2012 | Worldwide Distribution-USA (nationwide) including the states of AL, AZ, CA, CO, FL, GA, IL, IN, KS, MD, MI, MN, MO, MT, NC, ND, NE, NJ, NY, PA, SD, TX, UT, VA, and WA and countries of: Canada and EMEA. | Open, Classified |  |  | 7576 | Undetermined
Z-0458-2014 | 66627 | No | NEW | Absorbable polydioxanone surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Ethicon Stratafix Symmetric PDS Plus Knotless Tissue Control Device Ethicon LLC Guaynabo, Puerto Rico 00969. |  |  | September 25, 2013 | 2013 | Ethicon, Inc. | Ethicon Stratafix Symmetric PDS Plus Knotless Tissue Control Device has a small number of tab failures and fascial dehiscences in lower abdominal incisions. | Unknown/Undetermined by firm | Ethicon Inc. sent recall notifications/business reply forms on 9/25/2013 via UPS next day certified mail. |  | 37164 | 37,164 eaches | Worldwide Distribution-USA (nationwide) and countries of: Caribbean, Chile, Singapore and Malaysia. | Completed |  |  | 37164 | Undetermined
Z-0466-2014 | 66948 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Vasoview Hemopro Endoscopic Vessel Harvesting System Maquet Cardiovascular LLC 45 Barbour Pond Drive Wayne, NJ 07470 |  |  | October 30, 2013 | 2013 | Maquet Cardiovascular, LLC | Maquet has received an increase in the number of complaints for "self-activated/remains activated/overheats" for the Hemopro VH-30000, "Hemopro 1" | Component design/selection | Maquet Inc. sent an Urgent - Medical Device Recall (Removal) Immediate Action Required Letter dated October 30, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Please examine your inventory immediately to determine if you have any of the affected VASOVIEW HEMOPRO Endoscopic Vessel Harvesting System (HEMOPRO 1) VH-3000 with the lot number(s), as specified above. The VH-3000 catalog number and lot number are located on the package label. Please remove the affected VASOVIEW HEMOPRO Endoscopic Vessel Harvesting System and place in a secure location.NOTE: VASOVIEW HEMOPRO Endoscopic Vessel Harvesting System HEMOPRO 1 lot numbers not listed are not affected by this recall and should not be removed. HEMOPRO 2 is not affected by this recall.Your MAQUET Cardiovascular Sales Representative will contact you shortly to assist with the completion of the enclosed Acknowledgement Form ( Fax Back Form), the return of any affected VASOVIEW HEMOPRO Endoscopic Vessel Harvesting System HEMOPRO 1 product and arrange for necessary replacement.We apologize for any inconvenience this may cause. If you have any questions, please contact your local MAQUET Cardiovascular Sales Representative or MAQUET Customer Service at 1-888-880-2874, Monday through Friday, between the hours of 6:00 a.m. and 5:00 p.m. PST.Thank you for your cooperation and immediate assistance. |  | 12259 | 12,259 units | Worldwide Distribution - USA (nationwide) and the countries of Brazil, Czech Republic, Denmark, France, Hong Kong, Italy, Middle East, Netherlands, Singapore, Switzerland, and the United Kingdom. | Terminated | April 24, 2015 | 541 | 12259 | Undetermined
Z-0468-2014 | 66586 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Arrow Transradial Artery Access products |  |  | October 09, 2013 | 2013 | Arrow International Inc | There is a risk that the introducer needles packaged within the kits can be partially or fully occluded, which can cause delayed treatment or an adverse event when used on a patient. | Process control | Arrow International sent an Urgent Medical Device Recall Notification letter dated October 14, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately discontinue use and quarantine any affected product,to return product complete the enclosed Recall Acknowledgement Form and fax to 1-800-343-2935, ATTN: Customer Service. Customers were asked to completed the Recall Acknowledgement Form even if they have none of the affected product.Customers with questions were instructed to contact their local sales representative or Customer Service at 1-800-343-2935.For questions regarding this recall call 610-378-0131. |  | 294430 | 294430 | Worldwide Distribution - USA (Nationwide) and Internationally to Australia, Brazil, Canada, China, Ireland, New Zealand, Singapore, and Thailand. | Terminated | August 12, 2015 | 672 | 294430 | Undetermined
Z-0470-2014 | 66826 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Coulter LH 750 Hematology Analyzer, Part No. 6605632, |  |  | November 08, 2013 | 2013 | Beckman Coulter Inc. | Beckman Coulter is recalling the LH 750 and LH 780 Hematology Systems because the optical degradation may potentially cause a delay in reporting results for nucleated red blood cells (nRBC), white blood cell differential and/or reticulocytes. | Under Investigation by firm | Beckman Coulter sent an customer notification letter dated 11/8/13 to all customers who purchased the LH 750 and LH 780 Hematology Systems. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to complete and return the enclosed response form within 10 days. Customers with questions are instructed to contact Beckman Coulter's Customer Support Center at http://www.beckmancoulter.com/customer/support or call (800) 526-7694 in the US and Canada. Outside the US and Canada, customers are instructed to contact Beckman Coulter Representative.On 2/17/14, Beckman Coulter expanded the recall to inform Coulter LH 750 and Coulter LH 780 Hematology Analyzer customers of the issue related to LH Cleaner and Coulter Clenz. |  | 4155 | 4155 units total (1676 units in US) | Worldwide Distribtion - USA Nationwide including the countries of Algeria, Andorra, Angola, Australia, Austria, Bahrain, Bangladesh, Belgium, Bermuda, Bosnia and Herzegovina, Bulgaria, China, Colombia, Croatia, Czech Republic, Ecuador, Egypt, France, Germany, Greece, Guatemala, Honduras, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Korea, Republic of, Kuwait, Latvia, Lebanon, Libyan Arab Jamahiriya, Lithuania, Macao, Malaysia, Mayotte, Mexico, Morocco, Myanmar, Netherlands, New Zealand, Pakistan, Panama, Philippines, Poland, Portugal, Puerto Rico, Qatar, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Tunisia, Turkey, United Arab Emirates, United Kingdom, Uruguay, Venezuela, Viet Nam, Virgin Islands, U.S., and Zimbabwe. | Terminated | December 25, 2014 | 412 | 4155 | Undetermined
Z-0517-2014 | 65051 | No | MLM | Tacrolimus test system. | Clinical Chemistry | Toxicology | 510(k) | Dimension(R) TACR Flex(R) reagent cartridge (DF107) |  |  | April 25, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Siemens has confirmed that the TACR method may demonstrate reduced on-board stability which may result in imprecise and inaccurate QC and patient results. | Component design/selection | Siemens sent an Urgent Medical Device Correction letter dated April 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. A recall notification letter was sent to consignees in April 2013. Customers using the affected product should load Flex reagent cartridges on board immediately prior to use and remove the Flex after 8 hours. Do not load more Flexes than the lab will use in 8 hours and do not pre-hydrate TACR Flexes. Customers were asked to complete the attached Effectiveness Check form and fax it to the Siemens Technical Solutions Center at 302-631-8467. Customers were asked to forward the communication to anyone to whom may have distributed the product. Customers with questions were instructed to contact the Siemens Technical Solutions Center or their local Siemens technial support representative. For questions regarding this recall call 800-441-9250. |  | 13751 | 13,751 | Nationwide Distribution including AK, AL, AR, AZ, CA, DC, FL, GA, IA, IL, LA, MD, MI, MO, MS, NC, NH, NJ, NM, OH, OR, PA, PR, SC, SD, TN, TX, UT, VA, WA, and WI. | Terminated | November 12, 2014 | 566 | 13751 | Undetermined
Z-0477-2014 | 66717 | No | HSN | Ankle joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer Trabecular Metal Total Ankle Tibial Base Component Sizes 1-6 |  |  | November 12, 2013 | 2013 | Zimmer, Inc. | As part of routine manufacturing process monitoring by Zimmer, it was discovered that the Tibial Base and Talar Components from the Total Ankle system have the potential of containing elevated levels of manufacturing materials residue on the implants. The specific material was found to be mass media utilized during the manufacturing process to provide the uniform surface finish. This mass media | Manufacturing material removal | Zimmer sent an URGENT MEDICAL DEVICE RECALL notification letter dated November 12, 2013 to all consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to locate and remove the affected product from their inventory and notify their Zimmer representataive. The Zimmer representative will remove the recalled product from their facility. For patients that previously had the affected product implanted, consignees were instructed to continue their operative follow up routine. Consignees with questions were instructed to call 1-877-946-2761.For questions regarding this recall call 800-613-6131. |  | 50310 | 390 units (US) 129 units (outside of US) | Worldwide Distribution - USA (nationwide) and Internationally to Australia, Canada, China, Finland, Germany, and Italy. | Terminated | October 15, 2014 | 337 | 50844 | Undetermined
Z-0474-2014 | 66409 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Synthes Trochanteric Fixation Nail |  |  | July 30, 2013 | 2013 | Synthes USA HQ, Inc. | Certain lots of Trochanteric Fixation Nail were distributed to the field with the incorrect labeling of the length and diameter of the nail on the packaging. | Employee error | Synthes sent an Urgent Notice: Medical Device Recall letter dated July 30, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to take the following actions:Call Syntehs at 1-800-479-6329 to obtain a Return Authorization Number.Complete the Verification Section checking the appropriate box indicating affected product has been located.Return the Verfication Form with the product to:Credit/Returns, Synthes, 1101 Synthes Avenue, Monument, CO 80132.If no affected product found customers were instructed to take the following steps:Complete the attached Verification Section at the end of the letter by checking the appropriate box indicating no affected product has been located.Return the documents to Synthes by fax to 610-251-9005 or Scan/email: FieldAction@synthes.com. Customers with questions were instructed to call 610-719-5450.For questions regarding this recall call 610-719-5000. |  | 12 | 12 | Nationwide Distribution including WA, FL, LA, TX, GA, IN, OH, CA, MI, and PA. | Terminated | September 10, 2015 | 772 | 12 | Undetermined
Z-0482-2014 | 66834 | No | JOZ | Automated platelet aggregation system. | Hematology | Hematology | 510(k) | VerifyNow IIb/IIIa Test, 10-Test Kit, Catalog No. 85310, |  |  | November 14, 2013 | 2013 | Accumetrics Inc | Accumetrics is recalling the VerifyNow IIb/IIIa 10-Test because it contains incorrect information for the amount of time in which to perform the test following collection of the patient sample. The correct instruction is to perform the test within 15 minutes of collecting the patient sample. | Other | Accumetrics sent an Urgent Medical Device Correction letter dated November 14, 2013 to end users for the VerifyNow IIb/IIIa 10-Test. The letter informs the customers of the problems identified and the actions to be taken. The customers are instructed to call Accumetrics Customer Support at (800) 643-1640 option 2 or email at support@accumetrics.com. Customers are instructed to complete and return the enclosed Customer Account Tracking Form as soon as possible.For questions regarding this recall call 858-643-1600. |  | 69 | 69 units | Nationwide Distribution including NY, PA, DC, CT, IN, NJ, AR, TX, AZ,and NM. | Terminated | December 23, 2013 | 39 | 69 | Undetermined
Z-0481-2014 | 66978 | No | NJM | Orthodontic plastic bracket. | Dental | Dental | 510(k) | Opal Orthodontics by Ultradent. Avex CX2 Ceramic Orthodontic Brackets |  |  | November 19, 2013 | 2013 | Ultradent Products, Inc. | Ultradent Products Inc. is recalling various lots of Avex CX2 Orthodontic Brackets because the design control process was not complete prior to distribution. | Process control | Ultradent Products notified consignees via email on November 19, 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to return all unused product. Sales representatives will pick up remaining unplaced brackets or returns will be sent back using pre-paid envelopes back to Ultradent recall coordinator for quarantine. Consignees were asked to complete evaluation forms to be returned by fax to 800-793-5206.For questions regarding this recall call 801-553-4420. |  | 217 | 217 units | Nationwide Distribution including UT, CA, WA, VA, OH, and TN. | Terminated | October 06, 2014 | 321 | 217 | Undetermined
Z-0484-2014 | 66790 | No | FSY | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Helion S Exam Light or H300 |  |  | February 19, 2013 | 2013 | Trumpf Medical Systems, Inc. | Possible fatigue failure of the plastic joint may occur after an average use of seven years. | Nonconforming Material/Component | TRUMPF sent an Urgent Recall Notice dated February 19, 2013, to all US affected customers. The letter identified the product the problem and the action needed to be taken by the customers. The recall notice instructed customers to verify if they still have these light systems, and if they do have the light system verify if they have a metal or plastic ring at the joint of concern (per the manufacturer's Safety Notice). TRUMPF will directly work with affected customers to replace the affected arm system.For further questions please call ( 888 ) 474-9359. |  | 287 | 287 | US Distribution including the states of AR, CA, CO, KS, MA, MI, MO, NJ, NY, NC, OH, OK, PA, SD, TX and WA. | Open, Classified |  |  | 287 | Undetermined
Z-0492-2014 | 66536 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Aesculap Columbus Revision CoCr 15mm distal and posterior spacers, sizes F5-F7 |  |  | September 19, 2013 | 2013 | Aesculap, Inc. | The Columbus Revision CoCr 15 mm distal and posterior implant spacers, sizes F5-F7, were not compatible. | Component change control | Domestic consignees were contacted by phone and a recall notification letters was also sent to these Users on/about September 26, 2013. |  | 24 | 24 | Distributed in FL and MD. | Terminated | December 22, 2014 | 459 | 24 | Undetermined
Z-0518-2014 | 66859 | No | CAZ | Anesthesia conduction kit. | Anesthesiology | Anesthesiology | 510(k) | Kimberly-Clark Single Shot Epidural Pain Management Trays containing the BD LOR Syringe (BD Code 405198) |  |  | October 24, 2013 | 2013 | Kimberly-Clark Corporation | Kimberly-Clark Single Shot Epidural Pain Management Trays contain the BD 7 mL Epilor plastic Luer-Lok LOR Syringes under recall by Becton Dickinson (BD) because the product may stall or stick when traveling within the barrel of the syringe. | Component design/selection | Kimberly Clark Corporation sent a Product Advisory Notice letter to affected consignees. The letter identified the affected product, problem and actions to be taken. For questions call 770-587-8393 or via email |  | 389300 | 389,300 units | U.S. Nationwide Distribution | Open, Classified |  |  | 389300 | Undetermined
Z-0513-2014 | 66626 | No | LFW | Tobramycin test system. | Toxicology | Toxicology | 510(k) | Thermo Fisher Scientific/Microgenics brand CEDIA Tobramycin II Assay, Catalog Number 100018, Product is manufactured and distributed by |  |  | October 16, 2013 | 2013 | Microgenics Corp | Microgenics Corp., part of Thermo Fisher Scientific, is recalling CEDIA Tobramycin II Assay (Lot Number: 60169764) due to calibration stability problem relating to two complaints received. | Under Investigation by firm | Microgenics Corp., part of Thermo Fisher Scientific sent an Urgent Medical Device Correction letter dated Octobere 16, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately discontinue use and discard. Customers with questions were instructed to contact mgc@thermofisher.com. For question regarding this recall call 1-800-232-3342. |  | 370 | 370 units | Worldwide Distribution - USA including WI, UT, ND, MT, WA, CA, OK and Internationally to Australia, Canada, and Germany. | Terminated | February 03, 2014 | 110 | 370 | Undetermined
Z-0519-2014 | 66521 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access Immunoassay Systems Reaction Vessels. |  |  | October 02, 2013 | 2013 | Beckman Coulter Inc. | Beckman Coulter is initiating a recall for the Access Immunoassay Systems Reaction Vessels, including the Acess/Access 2 Immunoassay Systems, SYNCRON LXi 725 Clinical System, and UniCel DxC 600i SYNCHRON Access Clinical System because Beckman Coulter has observed an intermittent increase in Access System "Wash Carousel Motion Errors" with specific lots of Access SystemReaction Vessels (RVs). I | Pending | Beckman Coulter sent an Urgent Product Correction (PCA) letter dated October 1, 2013, along with a customer response form to all affected customers on October 2, 2013. The customer letter informed customers of the recall with the issue, impact, action (If your Access Immunoassay System experiences Wash Carousel Motion Errors that cannot be corrected with normal troubleshooting measures as outlined in the Access 2 Reference Manual section 6.3, contact Beckman Coulter Customer Technical Support.If you do not have Access System RVs with the lot numbers listed above, or if your Access Immunoassay System is not experiencing Wash Carousel Motion Errors, no additional action is required.), resolution, and contact information, website: http://www.beckmancoulter.com. Customers in the United States and Canada with questions were instructed to call 1-800-854-3633, outside the United States and Canada, customers were instructed to contact their local Beckman Coulter representataive.For questions regarding this recall call 714-961-4941. |  | 16154 | 16,154 units | Worldwide Distribution - USA (nationwide) and Internationally to Canada, Japan, India, Hong Kong, Guatemala, China, Brazil, Brunei Darussalam, Costa Rica, El Salvador, JordanMalaysia, Mexico, Panama, Philippines, South Africa, Taiwan, Thailand, Trinidad and Tobago, United Arab Emirates, Venezuela, and Viet Nam. | Terminated | October 23, 2014 | 386 | 16154 | Undetermined
Z-0193-2014 | 66624 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Sarns" Soft-FlowÂż Aortic Cannula without Suture Flange, angled tip, wire-reinforced with luer port, aortic cannula, 8.0 mm (24 Fr) OD with 3/8" connector, 14" (36 cm) long. |  |  | October 23, 2013 | 2013 | Terumo Cardiovascular Systems Corporation | During an in-process inspection, Terumo Cardiovascular Systems (Terumo CVS) identified the presence of loose fiber particulate that exceeded finished product specifications on certain product lots of Sarns" Soft-FlowÂż Aortic Cannulae and Sarns" Venous Return Cannulae. | Employee error | On 10/23/13, Terumo Cardiovascular Systems sent an URGENT MEDICAL DEVICE RECALL LETTER to their consignees. The letter identified the reason for the device removal, identified the affected product and associated potential hazards, and provided recall instructions to their customers regarding affected product return. Questions regarding this recall are directed to CVS Customer Service M-F, 8am-6pm at 1-800-521-2818. |  | 90 | 90 units | Worldwide Distribution-USA including DC and the states of MO,TX,OK, LA, MA, CA, WI, and MI and the country of Canada. | Terminated | July 18, 2014 | 268 | 230 | Undetermined
Z-0188-2014 | 66726 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Spacelabs Healthcare elance Vital Signs Monitor, Model 93300, with Option S. It is a patient monitor with arrhythmia detection or alarms. |  |  | October 17, 2013 | 2013 | Spacelabs Healthcare, Llc | Spacelabs Healthcare elance Vital Signs Monitor with Option S, model 93330, is recalled because the monitor will not permit connection of any of the Spacelabs TruLink line of SpO2 sensors. | Nonconforming Material/Component | Spacelabs Healthcare sent the "URGENT -MEDICAL DEVICE CORRECTION Spacelabs Healthcare elance Vital Signs Monitor with Option S" letter, dated 01 November 2013, to the US consignee on November 1, 2013.Spacelabs Healthcare will send the customer letter (translated as necessary) to all international subsidiaries and distributors of record via email on November 6, 2013.USERS ARE ADVISED THAT:-They may not be able to monitor their patient's oxygen saturation because the monitor will not allow the TruLink SpO2 sensor to be connected. The clinician will be required to find an alternate method to monitor oxygen saturation.-They will be contacted at the earliest possible date to schedule a convenient time and Spacelabs will correct all of your facility's affected monitors at no cost.Customer in the US can call Spacelabs Healthcare, Inc. of Snoqualmie, Washington at 800-522-7025.If you have any questions about this corrective action program, please contactSpacelabs Healthcare at +86 512 87178304 for Technical Support. |  | 89 | 89 units total (1 in US and 88 International) | Worldwide Distribution-Distributed in Puerto Rico and the countries of France, Indonesia, Korea, Nicaragua, Philippines, and Taiwan. | Terminated | February 25, 2014 | 131 | 89 | Undetermined
Z-0182-2014 | 66598 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Prostiva RF Therapy Hand Piece Kit, REF 8929, Rx only, sterilized using Ethylene Oxide. |  |  | October 21, 2013 | 2013 | Urologix, Inc. | A Hand Piece high deployment force exceeded the maximum specification. | Process change control | Consignees were sent a Urologix Urgent Medical Device Recall letter dated October 21, 2013. The letter described the problem and the device involved in the recall. Advised consignees to examine their inventory and remove the recalled product. Requested consignees to complete and fax the "Customer Acknowledgement Form" to Hope Przybilla at 763-475-1443 or call her for questions. The consignees local representatives will be personally assisting them with the return process and ensuring that replacement product is offered and received promptly. |  | 142 | 142 | Nationwide Distribution-including the states of AZ, CA, FL, GA, IL, KY, MA, MN, NC, NJ, NV, NY, OH, OR, TN, TX, and UT. | Terminated | April 21, 2014 | 182 | 142 | Undetermined
Z-2295-2013 | 65821 | No | LKK |  |  | General Hospital | PMA | MedStream Refill Kit, U.S: 91-4287 |  |  | August 13, 2013 | 2013 | Codman &amp; Shurtleff, Inc. | Drug flow rates exceeding programmed flow rates. | Use error | Codman &amp; Shurtleff, Inc. sent a Medical Device Field Safety Notice letter dated August 13, 2013 to US Consignees via FedEx overnight delivery and International Consignees were notified of the recall via email. The letter identifed the affected products, potential clinical impact, and updated IFU pump filling warning. The letter informed consignees that the Instructions for Use (IFU) and product training materials had been updated by adding a warning statement about potential over infusion if air gets inside the pump reservoir and additional clarifications to reinforce the proper filling technique. Consignees are asked to complete the Field Safety Notice Acknowledgement Form and return it to your local Codman Neuro Sales Representative or fax to: 508-977-6665.Consignees are asked to contact your Codman Neuro representative if they have questions, or contact Codman Neuro Clinical Support at 1-800-660-2660. Please report any malfunction or adverse event related to this device to 1-866-491-0974, Option 2. |  | 5 | 5 units | Worldwide Distribution - USA Nationwide in the states of NJ, NC, TX, FL and in the countries of Austria, Belgium, Egypt, Germany, Greece, Ireland, Israel, Italy, Luxemborg, Portugal, Russia, Saudi Arabia, Slovenia, Spain, Sweden, Switzerland, Turkey, and UK. | Terminated | September 11, 2014 | 394 | 4116 | Undetermined
Z-2284-2013 | 66355 | No | ETA | Total ossicular replacement prosthesis. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | Universal Titanium Prosthesis, Round Head, with Flex H/A titanium shoe. The ear prosthesis is implantable and indicated for hearing loss. |  |  | August 23, 2013 | 2013 | Medtronic Xomed, Inc. | One lot of the product was assembled using the Cam Head version of the same prosthesis instead of the Round head Shaft. | Process control | An "URGENT PRODUCT RECALL" letter was faxed to customers following the initial telephone notification. The letter described the product issue and provided recommended actions. Questions were directed to customer service at 800-874-5797. |  | 60 | 60 | Nationwide Distribution (US), including the states of CA, NE, GA, and IL. | Terminated | November 07, 2013 | 76 | 60 | Undetermined
Z-2274-2013 | 66070 | No | CAZ | Anesthesia conduction kit. | Anesthesiology | Anesthesiology | 510(k) | BD Epilor Syringe |  |  | August 23, 2013 | 2013 | Becton Dickinson &amp; Company | BD has received several reports indicating that the plunger of the 7 mL BD Epilor plastic BD Luer-Lok Loss of Resistance Syringe may stall or stick within the syringe preventing the user from detecting a loss of resistance upon entering the epidural space. | Unknown/Undetermined by firm | BD sent a product advisory notice to distributors/customers dated August 23, 2013 via UPS 2nd day mail. Consignees will be contacted by phone email and letter. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to be aware of the issue with the product and to follow the standard routine procedure to ensure that there is no resistance in the syringe prior to use. Any syringe that shows sticking or stalling should be discarded. Customers who distributed this product, should identify their customers and notify them at once of the product advisory. Customers with questions were instructed to call BD Customers Service at 1-800-237-2762.For questions regarding this recall call 201-847-6800. |  | 1335885 | 1,335,885 units | Worldwide Distribution - USA including FL, VA, TX, NY, IL, AZ, MI, RI, NJ, SC, MO, TN, CA, MN, UT, GA, NE, OH, MD, PA, MT and internationally to Brazil, Australia, Canada, and Puerto Rico.. | Terminated | March 19, 2014 | 208 | 1335885 | Undetermined
Z-2281-2013 | 66152 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | CareFusion Gravity Set, Model #44000-07 |  |  | August 30, 2013 | 2013 | CareFusion 303, Inc. | CareFusion is recalling the Gravity Set (Model 44000-07) because of an incorrect expiration date. The affected lot numbers of the Gravity Set are labeled witha 5 year expiration instead of three (3) years. | Employee error | CareFusion sent an Urgent Medical Device Recall Notification letter dated August 30, 2013, to all affected customers to inform them that CareFusion is recalling the Gravity Set, Model 44000-07, Lot Numbers 12086930, 12106215, 13016408, 13016834, 13025446 and 13025672 as a result of an incorrect expiration date. The letter informs the customers of the problem identified, issues, potential risk, and the immediate actions to be taken. Customers are instructed to complete the customer response form and return the form to CareFusion. Customers are instructed to return recalled products directly to distributors. A distributor letter will also be sent to distributors on August 30, 2013, to inform them of the problems identified and the actions to be taken. Distributors are instructed to contact the Customer Support Team, return customer response form, and if they have questions then they are instructed to contact the CareFusion Support Center at 1-800-562-6018.For questions regarding this recall call 858-617-4000. |  | 2850 | 2,850 units | IL | Terminated | December 26, 2013 | 118 | 2850 | Undetermined
Z-2285-2013 | 66151 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | SmartSite Low Sorbing Infusion Set, Model# 10015048. Intravascular administration set. |  |  | August 30, 2013 | 2013 | CareFusion 303, Inc. | CareFusion is recalling the SmartSite Low Sorbing Infusion Set, Model# 10015048, because there is a potential for separation at the tubing sleeve below the drip chamber. | Process design | A customer letter dated August 30, 2013 will be sent to customers to inform them that CareFusion is recalling the SmartSiteÂż Low Sorbing Infusion Set Models 10015048, Lot Numbers 130153883 and 13025342 as a result of a potential for separation at the tubing sleeve below the drip chamber.The letter informs the customers of the problem identified, issues, potential risk, and the immediate actions to be taken. Customers are instructed to complete the customer response form and return the form to CareFusion. Customers are instructed to return recalled products directly to distributors. A distributor letter will also be sent to distributors on 8/30/13 to inform them of the problems identified and the actions to be taken. Distributors are instructed to contact the Customer Support Team, return customer response form, and they are instructed to contact the CareFusion Support Center at 1-800-562-6018 with any questions. . |  | 15860 | 15860 units | Nationwide in US: PA, OK, TX, FL, NC, CA, IN, WI | Terminated | April 16, 2014 | 229 | 15860 | Undetermined
Z-2291-2013 | 66242 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE Healthcare, Discovery MR750 3.0T, Optima MR450w 1.5T, and Discovery MR450 1.5T . |  |  | August 13, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue due to exposed yellow foam on the end bells of Discovery MR450, Discovery MR750, and Optima MR450w products.The yellow foam strip used between the end bell covers and patient bore of certain GE MR Products may extend past the surface of the covers. The foam itself is not hazardous. However, any fluid or body substance absorbed | Component design/selection | GE Healthcare sent and Urgent Medical Device Correction letter dated August 13, 2013, to all affected consignees. The letter was addressed to Hospital Administrators / Risk Mangers, Radiology Department Mangers and Radiologists. The letter described the Safety Issue, Safety Instructions, Affected Product details, Product correction and contact Information. Customers were instructed that the system may continue to be used as normal even with the yello foam protruding in the bore. Customers were told not to attempt to remove the yellow foam. Customers with questions were instructed to contact their local service representative if they have any questions concerning the notification.For questions regarding this recall call 262-513-4122. |  | 828 | 828 (443 US, 385 OUS) | Worldwide Distribution USA including DC, all states except HI, ND, RI and WY and Internationally to Australia, Bahrain, Belarus, Belgium, Brazil, Canada, Chile, China, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, Hong Kong, India, Indonesia, Israel, Italy, Japan, Korea (Republic of), Kuwait, Lebanon, Martinique, Mexico, Netherlands, New Zealand, Norway, Oman, Poland, Qatar, Russia, Saudi Arabia, Serbia, Singapore, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, United Arab Emirates, United Kingdom, Uruguay, .OUS: AUSTRALIA, AUSTRIA, BAHRAIN, BELARUS, BELGIUM, BRAZIL, CANADA, CHILE, CHINA, CROATIA, CZECH REPUBLIC, DENMARK, EGYPT, FINLAND, FRANCE, GERMANY, GREECE, HONG KONG, INDIA, INDONESIA, ISRAEL, ITALY, JAPAN, REPUBLIC OF KOREA, KUWAIT, LEBANON, MARTINIQUE, MEXICO, NETHERLANDS, NEW ZEALAND, NORWAY, OMAN, POLAND, QATAR, RUSSIA, SAUDI ARABIA, SERBIA, SINGAPORE, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY. | Terminated | April 25, 2014 | 255 | 828 | Undetermined
Z-2292-2013 | 65043 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Solitaire-35 Peek Solitaire Anterior Spinal System. The Solitaire-35 is an intervertebral body fusion device. |  |  | April 09, 2013 | 2013 | Ebi, Llc | Several complaints were recorded regarding one of the three screws driving through the cage of the Solitaire 35 implant. | Process control | Biomet LLC sent "Urgent Medical Device Recall" letters/fax back response forms on Tuesday April 9, 2013 via Fed Ex to all distributors. The notification identified the problem, product involved, and risk factors. In addition, the notification provides instructions to the representatives on how to return the product to EBI, LLC. Surgeons that have implanted the product were contacted on April 12, 2013. |  | 1730 | 1730 units | Nationwide distribution: US including states of: CA, GA, IN, KY, MO, NY, OR, TX, and WI. | Terminated | June 17, 2014 | 434 | 1730 | Undetermined
Z-2273-2013 | 66219 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Product Name: AVEAÂż Ventilator. |  |  | September 05, 2013 | 2013 | Carefusion 211 Inc dba Carefusion | CareFusion has identified a potential risk associated with AVEAÂż ventilators when used at higher altitudes. A proactive complaint review identified an error in the barometric pressure compensation.CareFusion is voluntarily initiating a field correction of the affected devices to correct the error. | Component design/selection | CareFusion sent customer notification letters on September 05, 2013 via USPS mail return receipt carrier. The customer notification letter dated September 3, 2013 was titled "URGENT PRODUCT RECALL", and was addressed to Biomedical Engineering, Respiratory Care, and Risk Management. The customer notification letter informed customers of the product under recall with product name, description, and serial numbers enclosed as list unique to each customer. The customer notification letter also states the Problem and Affected Devices: Issue, Potential Risk, Actions to be Taken by CareFusion, Action to be Taken by the Customer, and Contact Information.Care Fusion has initiated a project to develop a new version of software that will resolve this issue. Once this software is available to download you will receive a notification letter.The customers were instructed to promptly return the enclosed Response Card to expedite the correction process and acknowledge receipt of this Notification and follow the listed instructions noted in the letter.For Recall Related and Support contact CareFusion Recall Support Center at 888.562.6018 or email: SupportCenter@carefusion .com Questions. For Product Technical Support Adverse Event Reporting contact Support/Customer Advocacy at 800.231.2466, 714.283.2228 or email to support.vent.us carefusion.com. from (6:30am -5:00pm POST).UPDATED:Firm issued press release of the recall on 09/30/2013. |  | 19546 | 19,546 units | Worldwide distribution: US (nationwide) including Puerto Rico and countries of: Albania, American Samoa, Argentina, Australia, Austria, Bahrain, Barbados, Bangladesh, Belgium, Bermuda, Belize, Bolivia, Bosnia and Herzegovina, Brazil, Canada, Chile,China, Colombia, Costa Rica, Croatia, Cyprus, Czech Rep, Dominican Republic, Ecuador, Egypt, Estonia, El Salvador, Finland, France, Germany, Georgia, Greece, Guatemala, Hong Kong, Honduras, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Korea, Kuwait, Latvia, Lebanon, Lithuania, Lybia Jamahiriya, Malaysia, Mexico, Moldova, Morocco, Netherlands, Nepal, Nicaragua, Norway, Oman, Pakistan, Palastine, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Qatar, Romania, Russian Federation, Rwanda, Saudi Arabia, Singapore, Slovenia, Slovakia, Spain, Switzerland, Taiwan, Thailand, Tunisia, Turkey, United Arab Emirates, and United Kingdom. | Terminated | June 04, 2015 | 637 | 19546 | Undetermined
Z-2302-2013 | 66334 | No | LOX | Percutaneous Transluminal Coronary Angioplasty (PTCA) Catheter. | Cardiovascular | Cardiovascular | 510(k) | NC Sprinter; Rapid Exchange Balloon Dilatation Catheter; |  |  | September 16, 2013 | 2013 | Medtronic Vascular, Inc. | Compliance chart included in lot 207002011 of the NC Sprinter RX 2.25 x 12 mm products lists the size of the device as 2.5mm instead of the correct 2.25mm size. | Mixed-up of materials/components | Medtronic sent an Urgent Medical Device Recall letter beginning September 16, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Our records indicate that your facility has received affected NC Sprinter RX Balloon Dilatation Catheters as noted in the attached customer notification detail report. Consequently, Medtronic is asking you to take the following actions:1. Remove and quarantine all affected catheters that remain in your inventory.2. Return the affected catheters to Medtronic. A Medtronic representative can assist in facilitating the return and replacement of this product as necessary.3. Complete the attached Customer Confirmation Certificate and fax it to Medtronic at 651-367-0612 to the attention of Customer Focused Quality.Medtronic has taken the necessary steps to prevent any future shipment of mislabeled product. Regulatory agencies will be notified about this recall as applicable.Please share this notification with others in your organization as appropriate. If any NC Sprinter RX Balloon Dilatation Catheters within the scope of this recall have been forwarded to another facility, please notify that facility accordingly and facilitate the retrieval of the affected product.We appreciate your cooperation and apologize for the inconvenience that this issue may cause. Should you have any questions, please contact your local Medtronic representative. |  | 122 | 122 | Worldwide Distribution - USA and the countries of Australia, Belgium, Canada, Hong Kong, Malaysia, New Zealand, Singapore, Spain, Switzerland, Taiwan and Thailand. | Terminated | February 10, 2014 | 147 | 122 | Undetermined
Z-2303-2013 | 66108 | No | KHP | Lactic acid test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Lactate Membrane units for E7077 Electrode Radiometer Medical ApS. The intended use is in vitro testing of samples of whole blood. |  |  | October 08, 2012 | 2012 | Radiometer America Inc | RADIOMETER became aware that some membranes may have enzyme residue on the outer membrane. The enzyme residue may cause an initial negative bias on the reported Lactate result upon replacement of the Lactate membrane. | Process control | On 11/07/2012, the firm sent "Urgent Field Safety Notices to their customers. The product issue was described and recommended actions were provided. Customers may contact 1 (800) 736-0600, option 4 if questions. |  | 54528 | 13,632 boxes/4 membranes | Nationwide Distribution, including the states of: AK, AL, AR, AZ, BM, CA, CO, CT, DC, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI and WV. | Terminated | December 24, 2013 | 442 | 54528 | Undetermined
Z-0006-2014 | 66128 | Yes | BSK | Inflatable tracheal tube cuff. | Anesthesiology | Anesthesiology | 510(k) | H&H; Emergency Cricothyrotomy Kit (Part No. HHECT01/ HHECTK01, NSN 6515-01-573-0692), vacuum sealed and packaged in a 5 mil poly bags, labeled in part ***H &amp; H Associates, Inc. CAGE 1NNH6 www.gohandh.com Ordinary, VA 23131 (800) 326-5708***. |  |  | August 27, 2013 | 2013 | H &amp; H Associates | Medical device tubing used for endotracheal airway application may become weakened and rendered ineffective. | Process change control | H&H; Medical Corporation notified their customers by "Urgent Medical Device Recall" letter dated August 27, 2013 and advised that the cuff balloon wall may become weakened and rendered ineffective. The firm advised their customers to discontinue use and take steps to remove affected units from further use. The manufacturer will replace the current endotracheal airway and cuff within the kits identified with a new cuffed airway in its own sterile packaging. The recaller included two response forms for customers to return: The Acknowledgement of Recall Form is provided for those who can quarantine the Cricothyrotomy Kits indicated in the recall. The form confirms notification of this recall and discontinued use of affected devices and provides details regarding the accounting and quarantine of the product lots in question for return and correction. The Product Recall Replacement Part Request Form is only required when inventory quarantine is not feasible. This form acknowledges the receipt of this notification and the undersigned accepts the responsibility for affixing the replacement airways along with the instructions that H&H; Medical Corporation will provide. The undersigned also assumes the responsibility to inform users regarding the subject of this recall, requires the use of the replacement airways over the airways provided in the affected kits and requires the disposal of the airways within the kits upon opening. For questions regarding the corrective request, please contact Ms. Michelle Morgan at 804-642-3663 or at mmorgan@gohandh.com during the hours of 8:00 a.m.-4:30 p.m. Eastern Standard Time, Monday through Friday. |  | 6619 | 6,619 kits | Nationwide distribution. | Terminated | January 13, 2015 | 504 | 6619 | Undetermined
Z-2279-2013 | 66150 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | SmartSite Low Sorbing Infusion Set, Model#72313E |  |  | August 30, 2013 | 2013 | CareFusion 303, Inc. | CareFusion is recalling the SmartSite Low Sorbing Infusion Set, Model 72313E, Lot Number13016314, because of an incorrect filter was assembled on this lot. The correct filter is a 1.2 micronfilter (blue color) but a 0.2 micron filter (clear color) may have been added instead. | Employee error | CareFusion sent an Urgent Medical Device Recall Notification dated August 30, 2013, to all affected customers. The letter informed the customers of the problem identified, issues, potential risk, and the immediate actions to be taken. Customers were instructed to immediately check their inventory for the affected product, if any affected product is encountered, customers were instructed not to use. Customers were instructed to complete the customer response form and return the form to CareFusion with the affected product for replacement.. A distributor letter was also sent to distributors on August 30, 2013, to inform them of the problems identified and the actions to be taken. Distributors were instructed to contact the Customer Support Team, return customer response form, and if they have questions then they were instructed to contact the CareFusion Support Center at 1-800-562-6018.For questions regarding this recall call 858-617-4000. |  | 600 | 600 units | CA, FL, AZ, and TN | Terminated | October 28, 2013 | 59 | 600 | Undetermined
Z-2319-2013 | 66347 | No | JCX | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | DeVilbiss Vacu-Aide QSU Suction Unit (Disposable Suction Container/Filter), 800 ML DISP Container, Unassembled, 800 ML DISP Container w/6' patient tubing, and 800 ML DISP Container w/6' patient tubing, assembled. |  |  | March 07, 2013 | 2013 | DeVilbiss Healthcare LLC | End users of the DeVilbiss Disposable Suction Container/Filter had exposed the bacterial filter cartridge to fluid, thereby occluding the filter and rendering it unusable. | Use error | DeVilbiss Healthcare sent End users a field corrective action letter dated March 7, 2013, to all affected customers to reinforce the proper cleaning method.This letter is to inform you that we have come across cases where end users had exposed the bacterial filter cartridge in the disposable suction container to fluid, thereby occluding the filter and rendering it unusableSpecifically what do you have to do when you receive a correction/recall notice?" The US FDA states in its correction/recall regulations that: "Consignees that receive a recall communication should immediately carry out the instructions set forth by the recalling firm and, where necessary, extend the recall to its consignees.""Enclosed you will find two Urgent Medical Device Correction letters addressed to DeVilbiss Dealer/Provider" and "DeVilbiss User" that we ask you read carefully."It is important that you follow the actions outlined in these letters to ensure we can effectively disseminate this information to end-users."To reach your direct accounts we ask that you send a copy of the "DeVilbiss User" letter to each of these end-users.We sincerely appreciate all of your efforts to ensure the safety of your customers that depend on DeVilbiss products. We thank you in advance for your assistance and apologize for any inconvenience that you may experience as a result of this situation. Should you have any questions or concerns, please contact DeVilbiss Healthcare Customer Service at 1-800-338-1988. |  | 112849 | 112,849 | Worldwide Distribution - USA and the countries of Mexico and Canada. | Terminated | November 06, 2014 | 609 | 112849 | Undetermined
Z-0005-2014 | 64454 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | The product is labeled in part "***Surgical Laser Technologies, Inc. Montgomeryville, PA***SLT Fiber Delivery System***Sterile***Non-Pyrogenic***. |  |  | January 04, 2013 | 2013 | PhotoMedex, Inc. | Firm became aware that some products within the specified lot may not have been sealed appropriately, thus compromising the sterile barrier. | Process change control | The firm initiated their recall of these products beginning on January 4, 2013 by sending notification letters to their consignees via electronic mail and/or facsimile. Customers were asked to return the affected product for a replacement. |  | 549 | 549 | Worldwide Distribution-USA including the states of AR, KS and OH, and the countries of Germany, South Africa and the United Kingdom. | Terminated | April 16, 2014 | 467 | 549 | Undetermined
Z-0010-2014 | 66250 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | GE Healthcare, Avance, Avance CS2, Aisys, Aespire View, and Engstrom. |  |  | August 28, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue involving unresponsive buttons on the display of the EngstrÂżm ventilator and the Aespire View, Aisys, and Avance anesthesia machines.As a result of a manufacturing issue, the buttons on the left, right, and bottom keypads may not always detect button presses. This may result in the inability to access certain menu functions wh | Nonconforming Material/Component | Consignees were sent on 8/28/2013 a GE Healthcare "Urgent Medical device Correction" letter dated August 28, 2013. The letter was addressed to Chief of Anesthesia, Health Care Administrator / Risk Manager and director of Biomedical / clinical engineering. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information (1-800-345-2700 US). |  | 911 | 911 (252 US, 659 OUS) | Worldwide Distribution-USA (nationwide) including the states of US: AK, AZ, CA, CO, DE, FL, GA, IL, KY, LA, MD, MI, MN, MS, MT, NV, NY, NC, OH, OR, PA, SC, TX, WY, and WI and the countries of AUSTRALIA, AUSTRIA, BANGLADESH, BELGIUM, BOLIVIA, BRAZIL, CANADA, CHILE, CHINA, COST ARICA, CROATIA, CZECH REPUBLIC, EGYPT, FINLAND, FRANCE, GERMANY, GREECE, INDIA, INDONESIA, IRAQ, ITALY, JAPAN, KAZAHHSTAN, REPUBLIC OF KOREA, KUWAIT, LEBANON, LITHUANIA, MAURITITUS, MEXICO, MOROCCO, NETHERLANDS, NEW ZEALAND, NIGERIA, PAKISTAN, PERU, PHILIPPINES, POLAND, RUSSIAN FEDERATION, SAUDI ARABIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, VIETNAM, and VENEZUELA. | Terminated | January 30, 2015 | 520 | 911 | Undetermined
Z-0019-2014 | 66237 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | MX 16-Slice CT Scanner System. |  |  | July 22, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | A pinch point at the end of the table top can cause injury. | Use error | Field Safety Notice (FSN 72800601_72800603, will be sent to all US consignees via certified mail. Philips Key Market representatives will distribute the letter internationally. Additionally, a Field Service Engineer will contact each site to schedule a time to implement the Field Change Order. |  | 326 | 326 | Worldwide Distribution-USA including the states of AL, AR, CA, FL, GA, IL, IN, KY, MS, NC, NJ, NV, NY, OH, PA, and VA; and countries of: Albania, Algeria, Argentina, Australia, Austria, Belgium, Bolivia, Brazil, Cambodia, China, Ghana, Greece, India, Indonesia, Ireland, Italy, Jordan, Latvia, Lebanon, Libya, Lithuania, Malawi, Mexico, Nepal, Netherlands, Oman, Palestine, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Spain, Switzerland, Thailand, Turkey, Uganda, Ukraine, Uzbekistan, Venezuela and Vietnam. | Terminated | September 25, 2014 | 430 | 326 | Undetermined
Z-0012-2014 | 66277 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Precision 500D, Legacy, RFX, and SFX X. |  |  | August 30, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue involving the spotfilm device power assist handle.There is a potential for the internal bolts on the power assist handle to be loose or to loosen and/or fall out. This issue, if present, can result in a loss of calibration and offset of the power assist handle and/or, under extreme condition, sticking of the handle in the ON p | Employee error | GE Healthcare sent an "Urgent Medical Device Correction" letter dated August 30, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were requested to inspect your power assist handle for either of the two conditions listed below. Immediately contact your local GE Healthcare Service Representative if any issue is found.1. If you observe that the power assist handle plastic grip is loose relative tothe mount, discontinue use.2. If plastic grip is secure, but you observe erratic performance, (power assistunresponsive, drifting of spotfilm device with no command, or lock releasealways ON), you may continue to use the device with care. Use all safetyprecautions relative to application of the myelographic stop and the inhibit,table horizontal stop OFF.GE Healthcare will correct all affected systems at no cost to you. A GE Healthcare service representative will contact you to arrange for this correction.If you have any questions or concerns regarding this notification, please call one of the following phone numbers:United States: 800 437 1171Japan: 0120 055 919For other countries, please contact your local GE Healthcare Service Representative. |  | 810 | 810 (783 US, 27 OUS) | Worldwide Distribution - USA (nationwided) including DC., and the countries of AUSTRALIA, CANADA, CHILE, GERMANY, CHANA, GUATEMALA, ICELAND, ISRAEL, ITALY, JAMAICA, JAPAN, RUSIAN FEDERATION, UNITED ARAB EMIRATES and UNITED KINGDOM. | Terminated | May 22, 2014 | 265 | 810 | Undetermined
Z-0017-2014 | 64808 | No | JGQ | Total protein test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension Vista (R) Urinary/Cerebrospinal Fluid Protein (UCFP) Flex (R) Reagent Cartridge Lot 11349BB |  |  | December 19, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Siemens has confirmed an increase on absorbance errors when calibrating Dimension Vista Urinary/Cerebrospinal Fluid Protein (U CFP) Flex Reagent Cartridge lot 11349 BB. | Process control | The firn, Siemens, sent an "Urgent Medical Device Correction" letter (dated December 2012) was provided to all Dimension Vista (R) UCFP customers who have received UCFP lot 11349BB. The letter described the product, problem and actions to be taken. The customers are instructed to discontinue use and discard any remaining inventory of Dimension Vista UCFP lot 11349BB; provide instructions for product replacement; and complete and return the attacted REQUEST FOR REPLACEMENT FORM via fax to (302) 631-8467. Siemens will replace any un-used inventory of this lot, at no charge.If you have any questions, please contact your Siemens Technical Solutions Center. your local Siemens Technical Support Representative or 800-441-9250. |  | 3395 | 3395 | Worldwide distribution: US (nationwide) including sates of: AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IL, ID, KY, KS, LA, MA, MD, MI, MN, MO, MS, MT, PA, NC, NE, NH, NJ, NY, OH, PA, SC, TN, TX, UT, VA, VT, WA, WI, WV, and PR; and country of: UK. | Terminated | August 27, 2014 | 616 | 3395 | Undetermined
Z-0018-2014 | 66000 | No | MEC | Medical washer-disinfector. | General Hospital | General Hospital | 510(K) Exempt | Getinge 46-Series Medical Washer-Disinfectors, Listing # E208196 |  |  | July 30, 2013 | 2013 | Getinge Disinfection | Getinge Disinfection AB initiated a voluntary field safety correction for the Getinge 46-Series electrically-heated Washer Disinfectors due to possible overheating. | Nonconforming Material/Component | Getinge USA, Inc. sent a Device Correction Notices letter dated July 26, 2013 to all affected customers via US Postal Certified Mail within the U.S. and Canada. The letter identified the affected product, problem and actions to be taken. The letter indicate that all customers will be contacted by a Getinge representative for guidance if any machine within the serial number range of affected product is within their possession. If you should have any questions or concerns regarding this field correction, please call 1-800-475-9040, press "1" and enter extension 5036. |  | 143 | 143 units (142 domestically &amp; 1 unit internationally) | USA Nationwide in the states of AL, AR, AZ, CA, CO, FL, GA, HI, IA, ID, IL, IN, KS, LA, MA, MD, MI, MN, MO, MS, NC, NE, NH, NJ, NM, NV, NY, OH, OK, PA, PR, SC, SD, TN, TX, UT, VA, WA, WI, WV, WY and internationally in Canada. | Terminated | May 11, 2015 | 650 | 143 | Undetermined
Z-0025-2014 | 66171 | Yes | LWP |  |  | Cardiovascular | PMA | Boston Scientific COGNIS CRT-D, Cardiac Resynchronization Therapy Defibrillator, models N118, N119. The following models are not available in the US: models N106, N107, P106, P107. Sterilized using ethylene oxide. Boston Scientific, 4100 Hamline Avenue North, St. Paul, MN 55112-5798 USA. |  |  | August 29, 2013 | 2013 | Boston Scientific CRM Corp | The performance of a low voltage capacitor in a subset of COGNIS CRT-Ds and TELIGEN ICDs manufactured prior to December 2009, may be compromised over time, causing increased current drain that can lead to premature battery depletion. | Component design/selection | An Important Medical Device Information letter dated August 2013 was sent to implanting and following physicians. The letter provided a description of the issue, describes clinical implications, provides the rate of occurrence, and gives recommendations for monitoring patients. |  | 38500 | 38,500 total worldwide | Worldwide Distribution-USA (nationwide) including Guam and Puerto Ricoand the countries of Argentina, Aruba, Australia, Austria, Belgium, Bermuda, Brazil, Canada, Chile, Hong Kong, Colombia, Costa Rica, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, El Salvador, Finland, France, Gabon, Germany, Greece, Hungary, Iran, Ireland, Israel, Italy, Japan, Jordan, Kuwait, Latvia, Lebanon, Malaysia, Mexico, Netherlands, Netherlands Antilles, New Zealand, Norway, Panama, Peru, Poland, Portugal, Saudi Arabia, Singapore, Si Lanka, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Syria, Taiwan, Thailand, Trinidad, UK, United Arab Republic, and Venezuela. | Terminated | April 10, 2014 | 224 | 38500 | Undetermined
Z-0021-2014 | 66254 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Advantx, Innova 2000, Innova 2100IQ, Innova 3100, Innova 3100IQ, Innova 4100, Innova 4100IQ. |  |  | August 16, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has become aware of a potential safety issue involving missing screws in the overhead Video Monitor Suspension of the Advantx, Innova 2000, Innova 2100IQ, Innova 3100, Innova 3100IQ, Innova 4100, Innova 4100IQ Cardiovascular X-ray imaging systems. | Process control | Consignees were sent on 8/12/2013 a GE Healthcare "Urgent Medical device Correction" letter dated August 12, 2013. The letter was addressed to Hospital Administrators / Risk Mangers, Mangers of Radiology /Cardiology /Surgery Department and Radiologists /Cardiologists /Surgeons.The letter described the Safety Issue, safety Instructions, Affected Product details, Product Correction and Contact Information (US 800-437-1171, Japan 0120 055 919). |  | 2184 | 2184 (1063 US. 1121 OUS) | Worldwide Distribution-USA (nationwide) including DC and Puerto Rico, and except ND, SD and VT, and the countries of ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA, BELARUS, BELGIUM, BOLIVIA, BRAZIL, CANADA, CHILE, CHINA, COLOMBIA, COSTA ARICA, CROATIA, CZECH REPUBLIC, DOMINICAN REPUBLIC, ECUADOR, EGYPT, FINLAND, FRANCE, GEORGIA, GERMANY,GREECE, GUADELOUPE, GUATEMALA, HONDURAS, HONG KONG, HUNGARY, INDIA, INDONESIA, IRAN,IRELAND, ITALY, JAMAICA, JAPAN,JORDAN, KAZKHSTAN, REPUBLIC OF KOREA, KUWAIT, LEBANON, LIBYA ARAB JAMHIRIYA, LITHUANIA, MALAYSIA MALTA, MARTINIQUE, MEXICO, MOROCCO, NETHERLANDS, NEW ZEALAND, NICARAGUA, NIGERIA, NORWAY, PAKISTAN, PANAMA, PARAGUAY, PERU, PHILIPPINES, POLAND, PORTUGAL ROMANIA, RUSSIA, SAUDI ARABIA, SERBIA, SLOVAKIA, SPAIN, SWEDEN, SYRIA, TAIWAN, THAILAND, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, UKRAINE, UNITED KINGDOM, and VENEZUELA. | Terminated | May 30, 2014 | 287 | 2184 | Undetermined
Z-2280-2013 | 66209 | No | CIC | Calcium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ABX PENTRA Calcium CP Reagent, Model # A11A01633 |  |  | September 06, 2013 | 2013 | Horiba Instruments, Inc dba Horiba Medical | HORIBA Medical is informing all PENTRA 200 and PENTRA 400 customers of a potential for reportingincorrect calcium results when using the ABX PENTRA CALCIUM CP reagent. | Under Investigation by firm | Horiba Medical sent an Urgent Safety Notice dated September 6, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The customer notification letter dated September 6, 2013 will be sent to all PENTRA 200 and PENTRA 400 customers of a potential for reporting incorrect calcium results when using the ABX PENTRA CALCIUM CP reagent (Reference #: A11A01633) Lot#: 014664503. The recall letter informs the customers that the initial functional tests indicate that there is no interference identified with regards to the patient results. However, with time, there may be an increase in the quantity of precipitates and this could potentially interfere with the performance of this reagent. Customers are informed of the actions to be taken and to complete the customer response form within 10 days.Further questions please call (949) 453-0500 Ext. 208 |  | 2658 | 2,658 units ( 1240 US 1418 OUS ) | Worldwide Distribution - USA including the country of Canada. | Terminated | August 23, 2014 | 351 | 2658 | Undetermined
Z-2268-2013 | 66234 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Smith Medical, Deltec, REF 21-2386-24, 11 Fr Introducer Set with Peel-Away Sheath, Sterile EO, Rx Only |  |  | September 10, 2013 | 2013 | Smiths Medical ASD, Inc. | Smiths Medical is conducting a voluntary recall of PORT-A-CATHÂż and PORT-A-CATHÂż II Implantable Venous and ArterialAccess Systems and Introducer Sets.Smiths Medical has become aware of a small number of customer reports regarding the introducer sheath and the introducer opening being too small for the width of the catheter with the use of these Sets.Smiths Medical has received no reports | Mixed-up of materials/components | Smiths Medical ASD, Inc. sent an Urgent Field Safety Notice dated September 6, 2013, to all affected customers. The letter described the problem and the product involved in the recall. The letter included the "Advice on Action to be Taken by the User" which instructed consignees to inspect and remove affected devices and to complete and return the "Urgent Field Safety Notice Confirmation Form. The letter also included information on the "Transmission of this Urgent Field safety Notice" which included information where customers should report issues related to the product to Smiths Medical's Global Complaint Department at 1-866-216-8806 or globalcomplaints@smiths-medical.com. For questions on the information provided they should contact 800-258-5361 (Smiths Medical's Customer Service Department).For questions regarding this recall call 1-800-332-0178. |  | 190 | 190 | Worldwide Distribution - USA including AZ, AR, CA, GA, IL, IN, KY, MD, MA, MI, NV, NJ, NY, NC, OH, OR, PA, RI, SC, TN, TX, VA, WA, WV, WY.Internationally to DENMARK, BELGIUM CANADA. | Terminated | October 21, 2014 | 406 | 562 | Undetermined
Z-0100-2014 | 65722 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | pocH-100i; sysmex; Automated Hematology Analyzer; 7.3" x 13.8" x 18.1" / 30.8 lbs.; Sysmex Corporation; 1-5-1 Wakinohama-Kaigandori, Chuo-ku, Kobe 651-0073, Japan; Sysmex America, Inc; One Nelson C. White Pkwy, Mundeline, IL 60060, U.S.A. |  |  | May 15, 2013 | 2013 | Sysmex America Inc | The data printed on the thermal printer may omit a digit or decimal point in the results or sample identification number. | Nonconforming Material/Component | The firm, Sysmex, sent a "Product Notification-FIELD CORRECTION" letter dated May 2013 to all customers in the U.S. and Canada on May 15, 2013 by U.S. Mail. The letter describes the product, problem and actions to be taken. The customers will be contacted by a Sysmex Service Representative within the following 90 days to schedule an appointment to replace the Printed Circuit Board (PCB) on the internal printer. The repair will be performed during an on-site service visit. The customers were instructed to, until the repair can be made, use the thermal printer for reporting results; to compare the sample identification, values and decimal points on the printout with the data on the LCD screen for accuracy.If you have any questions concerning this information, please email the Sysmex Technical Assistance Center (TAC). Simply click on email TAC button on the home page of the CRC, complete the form that appears and then click the submit to send your question to TAC. Sending an email to TAC is recommended for non-urgent requests.For urgent request, call the Technical Assistance Center at 1-888-879-7639. |  | 193 | 193 units | Worldwide distribution: US (Nationwide) including Puerto Rico and states of: AK, AL, AR, CA, CO, CT, FL, GA, IA, IL, IN, KS, KY, LA, MA, MI, MN, MO, MT, NC, ND, NE, NH, NJ, NV, NY, OH, OK, SD, TN, TX, UT, VA, VT, WA, WI, and WY; and country of Canada. | Open, Classified |  |  | 193 | Undetermined
Z-2213-2013 | 65657 | No | ILQ | Powered communication system. | Physical Medicine | Physical Medicine | 510(K) Exempt | Responder 5 is a nurse call communication system that connects patients with staff in a hospital. The nurse call system also connects team members and departments from radiology, physical therapy, transport, environmental services and more. |  |  | May 10, 2013 | 2013 | Rauland-Borg Corp | A counterfeited electronic component was identified in the products affected. The component, a data transceiver (transmitter/receiver), is associated with the intercommunication of these devices in the Responder 5 system across the L-Net data path. Failure of a data transceiver results in the loss of nurse call functionality in the associated room, and may disrupt communications on the L-Net run | Counterfeit | Tech Bulletins sent a letter dated May 10, 2013, to all affected customers notifying them of the counterfeit parts and requesting that customers uninstall and return affected products to Rauland-Borg for replacement. A subsequent recall notification letter dated August 5, 2013 was mailed to customers. The letter included instructions to both direct accounts and end users on the process of replacing affected units with units manufactured with authentic parts.For additional information please call 847-590-5920. |  | 385 | Product Code 351000: 385 units; Product Codes 352000 and 352020: 14,808 units (total) | Worldwide distribution - US: (nationwide) including the states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IL, IN, KS, KY, LA, MA, MI, MN, MO, NC, NE, NV, NY, OH, OK, OR, PA, SD, TN, TX, UT, VA, WA, WI and WV., and the countries of : Australia, Canada, Columbia, Lebanon, Mexico and Saudi Arabia. | Terminated | May 06, 2015 | 726 | 385 | Undetermined
Z-2122-2013 | 66058 | No | LON | Fully automated short-term incubation cycle antimicrobial susceptibility system. | Microbiology | Microbiology | 510(k) | Microscan Synergies Plus Negative Urine Combo 2 |  |  | August 21, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Confirmed increase in false susceptible and false intermediate misreads for Imipenem and/or meropenum on Synergies plus Negative and rap ID/S plus Negative panels read on Walk Away System. | Process control | Siemens sent an Urgent Field Safety Notice letters dated August 2013 to all US customers via over night delivery, receipt requested. Letters for foreign customers are being translated and will be sent as soon as possible. The letter identified the affected product, reason for field correction, risk to health problem and actions to be taken by customer. Customers were instructed to suppress all reporting of sensitive and intermediate results for imipenem and meropenem. Siemens also recommends that the customer consider the need to review previous test results, conduct patient follow-up, and/or repeat testing. Customer were requested to complete the attached Field Correction Effectiveness Check form and fax it to 302-631-8467. For questions contact your Siemens Technical Support Provider or Distributor. |  | 27804 | 1,324 boxes, 26,480 panels | Worldwide Distribution - USA Nationwide in the states of: AK, SC, IA, MO GA, WI, TX, KS, CA, IL, CO, AZ, SD, IN, MD, FL, MN and the countries of:Japan, France, Algeria, Germany, Austria, Romania, Russian Federation, South Africa, Kazakhstan, Belgium, Great Britain, Turkey, Spain, Italy, Saudi Arabia, Qutar, Poland, Iraq, Unite Arab Emrites, Algeria, Oman, Greece, Portugal, and Ireland. | Terminated | May 15, 2014 | 267 | 200172 | Undetermined
Z-2192-2013 | 66031 | No | FMK | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | BD Microtainer Quikheel Lancet 1.0 mm incision Depth 2.5 mm Incision Length REF 368101 Rx Only Becton, Dickinson and Company, Franklin Lakes, NJ 07417 USA Made in USA. |  |  | July 10, 2013 | 2013 | Becton Dickinson &amp; Company | The Lancet will not retract after activation thus creating a potential for exposure to a used sharp. | Component change control | BD notified their customers and distributors by phone, e-mail,and/or letter/notice of return/ customer packing list on 7/10/2013. |  | 58200 | US - 39,800 units ; Canada - 18,400 units | Distributed in the states of CT, KY, IL, NY, PA, TX, and VA and the country of Canada. | Terminated | March 19, 2014 | 252 | 58200 | Undetermined
Z-2193-2013 | 66123 | No | GCB | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | *** 1) Cat. 4341B; Qty. 10; Thoracentesis Tray; With Catheter; Sterile; CareFusion; 1500 Waukegan Road; McGaw Park, IL 60085 USA; |  |  | August 28, 2013 | 2013 | Carefusion 2200 Inc | Catheters included in Thoracentesis Tray and Safe-T Thoracentesis Tray - Drug Free may occlude, preventing drainage of excessive fluid from the pleural cavity leading to a potential delay in the medical procedure and potentially serious patient complications. | Process change control | Recall notification letters dated 8/27/13 were distributed via courier to direct accounts. Consignees were/are requested to confirm receipt of the notification, examine their inventory for the recalled products and to destroy the products. After the products are destroyed, the consignees are requested to submit a completed response form to CareFusion for replacement with functional devices. Consignees who cannot destroy the affected products are requested to contact CareFusion for additional information. |  | 197913 | *** 1) 174,580 units; *** 2) 23,330 units | Worldwide Distribution-USA (nationwide) including the states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MD, MT, NC, NE, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, WI, and WV and Puerto Rico, and the country of Iceland. | Terminated | June 16, 2015 | 657 | 197913 | Undetermined
Z-2160-2013 | 66024 | No | KXJ | Radiologic table. | Radiology | Radiology | 510(K) Exempt | Juno DRF System. Mobile X-ray system |  |  | August 02, 2013 | 2013 | Villa Radiology Systems LLC | It has been discovered that the Air Kerma rate information is inconsistent in Continuous Fluoro modes 30x30 (12"x12") and 15x15 (6"x6). This could result in a deviation of the displayed values of Air Kerma Rate and Cumulated Air Kerma from the actual values greater then +/-35% and then in a failure to comply with the requirements of 21 CFR 1020.32(k)(6 | Radiation Control for Health and Safety Act | The Center for Devices and Radiological Health (CDRH) understands that the Villa Sistemi Medicali (U.S. Agent: Villa Radiology Systems) is responsible for the correction of all 22 noncompliant units distributed within the U.S. Should the CAP prove ineffective in bringing the noncompliant unit(s) in the field into compliance with the applicable performance standard, the CDRH can require more stringent measures to obtain compliance.The system will be modified by upgrading the generator power firmware, which will bring Juno DRF x-ray system into compliance. You provided a verification and validation test report (Juno DRF System AK fix-DAM Prom V&V; Test Report) to FDA demonstrating that these proposed upgrades addressed the problem. Field service engineers will schedule an appointment with customers to install the new software.2. The customer notification letter (URGENT: Electronic Product Radiation Warning), dated August 02, 2013, includes the statement that "the manufacturer will, without charge, remedy the defect or bring the product into compliance with each applicable Federal standard in accordance with a plan to be approved by the Secretary of Health and Human Services, the details of which will be included in a subsequent communication to you."3. The CAP will be completed within a six month timeframe (by March 1, 2014).4. The manufacturer will notify the Secretary on monthly basis providing progress with the field correction.CDRH approves the CAP subject to the following conditions:Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR 1003.21 and 1003.22.For further questions please call (203) 262-8836. |  | 22 | 22 units within US | US Distribution including the states of AZ, MI, IN, UT, CT, AL, NY, PA, IL, MD, CA, MA, IA, OH, MS and MO. | Completed |  |  | 22 | Undetermined
Z-2196-2013 | 65980 | No | JJW | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | DiaScreen Liquid Urine Controls: Positive and Negative Assayed Liquid Quality Controls for use with Dipsticks in Routine Urinalysis; Product Number D76004 (4x12 mL bottles) consists of: 2x12mL Positive controls [and] 2x12mL Negative controls; Manufactured for ARKRAY USA, Inc.; Minneapolis, MN 55439 |  |  | July 16, 2013 | 2013 | Assuramed | The products may have been stored at temperatures outside their required storage conditions. | Storage | Independence Medical (im) sent an Urgent Product Recall letter dated July 16, 2013, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to identify and immediately cease the distribution of the recalled products, quarantine the recalled products, and contact the recalling firm at 800-860-8027 for instructions on how to return the recalled products. Consignees who further distributed the recalled products were requested to notify their customers of the recall. Consignees were also instructed to complete the Product Recall Acknowledgement Form and return the document by email to recall@indemed.com indicating receipt of the notification.Customers with questions were instructed to contact the product manufacturer for ProTime Disposable Cuvettes at 800-579-2255, DiaScreen Liquid Urine Controls at 800-818-8877, and Hemoccult II SENSA elite Dispensapak at 800-854-3633.For questions regarding this recall call 800-860-8027. |  | 10 | 10 units | Nationwide Distribution including AR, AZ, CA, CO, FL, IL, KY, LA, MI, NJ, NY, OH,OK, OR, SC, TX, UT, and WA. | Open, Classified |  |  | 154 | Undetermined
Z-2200-2013 | 66010 | No | HXX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | T8 REMOVAL HEAD, REF 8910641, QTY: 1 EA, |  |  | August 12, 2013 | 2013 | Medtronic Sofamor Danek USA Inc | One lot of the T8 Removal Head had incorrect hardness specification, which could result in the T8 Removal head failing during removal of the CD Horizon Spire and Spire Z Spinal System implant. There is a potential for tissue damage to occur and/or minor surgical delay to remove the remaining fragments of the tip. | Nonconforming Material/Component | Medtronic sent an Urgent Voluntary Market Removal Recall letter dated August 9, 2013 with Recall Questionnaires and Confirmation forms included to Risk Managers and Sales Representatives. The letter identified the affected product, problem and action to be taken. Customers were instructed to remove the affected product, discontinue use and contact your Medtronic Sales Representative to coordinate the return and replenishment of affected product. For questions call your Medtronic Sales Representative or Global Quality Department at 800-876-3133. |  | 9 | 9 units | USA Nationwide Distribution in the state of GA, AK, WA, and CA. | Terminated | December 27, 2013 | 137 | 9 | Undetermined
Z-2210-2013 | 65691 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance CT 6, M/N 728256; Brilliance CT 16 M/N 728246; Brilliance CT 40, M/N 728235 &amp; Brilliance CT 64 M/N 728231, Philips Healthcare, Cleveland, OH 44143. |  |  | July 05, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | The patient support may move in an unintended manner if the footswitch cover impinges on the footswitch. | Component design/selection | Philips Healthcare sent an "URGENT-MEDICAL DEVICE CORRECTION" letter dated June 4, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the Customer Care Center at 1-800-722-9377, option 5 for information concerning this issue. |  | 2913024 | M/N: 728256, 4 units; M/N: 728246, 89 units; M/N: 728235, 9 units &amp; M/N: 728231, 207 units | Worldwide Distribution-USA (nationwide) including DC and the states of AL, AR, AZ, CA, CO, DE, FL, GA, IN, KY, MA, MD, MI, MN, MS, NC, NM, NV, NY, OH, OK, PA, SC, TN, TX, VT, WI &amp; WV, and the countries of Australia, Austria, Bosnia &amp; Herzegovina, Chile, China, Croatia, Denmark, Equatorial Guinea, France, Germany, Greenland, Iceland, Indonesia, Israel, Japan, Kazakhstan, Netherlands, Nigeria, Norway, Poland, Romania, South Africa, South Korea, Sweden, Switzerland, Taiwan, United Arab Emirates, United Kingdom &amp; Vietnam. | Terminated | September 15, 2015 | 802 | 2913024 | Undetermined
Z-2211-2013 | 66208 | No | FWP | Chin prosthesis. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Device Name: Terino Square Chin-Style II-Medium; Catalog Number: TSCII-M. |  |  | August 30, 2013 | 2013 | Implantech Associates Inc | Implantech initiated a voluntary recall of certain lots of Terino Square Chin-Style II-Medicum (Lot # 848488 and 850156), after determining that the devices from the lot numbers do not conform to the dimensional requirements for the TSCII-M. | Nonconforming Material/Component | Recall was initiated by sending recall notification letters to customers on 08/30/2013 via a combination of 1st class mail, email, and telephone. The recall notification letter, dated August 30, 2013, informs customers of the reason for recall, product description and codes, instructions on what to do with the recalled product and contact information. Within the recall notification letter, a section is provided for responses and signatures. |  | 8 | 8 | Worldwide distribution: US (nationwide) including states of NY, KY, and CO, and the country of Israel. | Terminated | February 06, 2014 | 160 | 8 | Undetermined
Z-2212-2013 | 66176 | No | JCX | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Zeppessis' reprocessed single-use manifolds (filters) for the Stryker Neptune 1 and Neptune 2 Waste Management Systems: |  |  | August 16, 2013 | 2013 | Zeppessis Reprocessing, LLC | Zeppessis' reprocessed single-use manifolds (filters) for the Stryker Neptune 1 and Neptune 2 Waste Management Systems are recalled due to lacking 510K and and without a validated cleaning protocol for reprocessing the used manifolds. | Reprocessing Controls | Zeppessis Reprocessing, LLC sent an Urgent Medical Device Recall letter dated September 4, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers are advised to stop using the Zeppessis reprocessed single-use manifolds (filters) for the Stryker Neptune 1 and Neptune 2 Waste Management Systems immediately and instructed to destroy the products in a proper manner.Customers should fill out the Customer Response form and fax or email it back to Zeppessis.For any questions about this recall, please call Zeppessis Reprocessing, LLC at 208-559-7064. |  |  | undetermined the distributed amounts | Distributed in the states of California, Idaho, Oregon, and Washington. | Terminated | November 08, 2013 | 84 |  | Undetermined
Z-2252-2013 | 66060 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Central Dressing Change Kit; AMS-8464CP- 6. Vygon Central Dressing Change Kit contains: 1 Pair Nitrile Gloves W/Wallet, Powderfree 1 Face Mask 1 Polylined Drape, 13" X 18" 1 ChloraPrep***, 3ml Applicator W/Insert 1 Pkg. Alcohol Swabsticks (3/Pkg.) 1 14" Extension Set W/Bionector*** 2 Dressing Sponges, 2" X 2" 1 Foam Strip, 8/5cm X 1cm 1 3M*** Tegaderm *** IV Dressing W/CHG, 2 1/2" X 4 1/2" *** Vygon MFG Dover, NH 03820***" |  |  | August 12, 2013 | 2013 | Churchill Medical Systems, Inc. | Microbore tubing disconnecting from male luers on the extension sets. | Under Investigation by firm | "Urgent Medical Device Recall" letters were sent to each customer via courier service on August 12, 2013. Consignees were provided with a description of the issue and instructions on how to handle product that they received and that was further distributed by them. Questions were directed to customer service at 1-800-473-5414. |  | 3040 | 3040 | Nationwide Distribution (US), including the states of NY, MN, NJ, PA, FL, MD, and VA. | Terminated | June 16, 2014 | 308 | 16140 | Undetermined
Z-2270-2013 | 65348 | No | ITI | Powered wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Invacare Power Wheelchair, one per box |  |  | June 04, 2013 | 2013 | Invacare Corporation | Unintended slowing or deceleration, and on rare occasion,unanticipated and unattended acceleration of power wheelchairs equipped with SPJ and MPJ type joysticks. | Process control | Invacare sent an Urgent Patient Safety Notification letter to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to contact their wheelchair povider to have their joystick replaced. Customers with questions were instructed to call their provider at Invacare's recall support line at 877-413-6008. Go to www.invacare.com/joystickrecall for more information regarding this recall |  | 166230 | 166,230 | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY. Internationally to Canada, Australia, Bermuda, Brazil, China, Costa Rica, Germany, France, United Kingdom, Guam, Hong Kong, Israel, India, Kuwait, New Zealand, Puerto rico, Qatar, Saudi Arabia, Singapore, Thailand, Trinidad, Tobago, and South Africa. | Terminated | January 04, 2016 | 944 | 166230 | Undetermined
Z-2242-2013 | 66052 | bandage, cast | ITG | Prosthetic and orthotic accessory. | Physical Medicine | Physical Medicine | 510(K) Exempt | 3M Scotchcast, Wet or Dry Cast Padding, Catalog Numbers WDP2, WDP3, WDP4, and WDP6. |  |  | August 05, 2013 | 2013 | 3M Company - Health Care Business | 3M is conducting a voluntary product recall of selected lots of 3M" Scotchcast" Wet or Dry Cast Padding. Upon investigation of a health care provider feedback, and confirmed with internal testing, it was found that the cast padding material does not consistently repel water. Lack of repellency may result in prolonged cast dry times after exposure to wet conditions. There have been no repo | Nonconforming Material/Component | 3M sent a Voluntary Product Recall Notification letter dated July 31, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to examine their inventory and set aside any affected lots subject to the recall notice. Customers were asked to inform all departments of any affected lots and provide them with a copy of the Notification letter that instructs them to contact their 3M distributor for alternative options. All affected products were to be returned to the 3M distributor they were purchased from. Customers with questions were instructed to contact the 3M Customer Helpline at 1-800-228-3957. For questions regarding this recall call 651-733-1110. |  | 24140 | 1207 cases (4 puches per case, 20 rolls per pouch) | Worldwide Distribution - USA including AL, AZ, AR, CA, CO,CT, FL, GA, HI, IL, IN,KS,KY, LA, MD, MA, MI, MN, MO, NE, NV, NJ, NY, NC, OH, OR,PA, TN, TX, UT,VA, WA, and WI. Internationally to GERMANY, CANADA, BELGIUM, NETHERLANDS, HONG KONG, NEW ZEALAND, AUSTRALIA, INDIA, KOREA, SINGAPORE AND UNITED ARAB EMIRATES, JAPAN, JAMAICA.. | Terminated | June 16, 2014 | 315 | 24140 | Undetermined
Z-2258-2013 | 65999 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | VisuALIF,System consisting of a 'U' shaped PEEK block in two footprint configurations, four heights and three lordosis angles., part number: 0301-1410. |  |  | July 31, 2013 | 2013 | Spine Smith Holdings, LLC | The product has the potential to be laser marked as a Medium when it actually is a Small. | Process control | The firm decided to remove this lot and verify all of the markings. A stock sweep for mis-marked parts was performed at the time and it is not known why this lot is identified, however an additional stock sweep of all PEEK products from all product lines was performed beginning 07/31/2013 and no additional lots were identified. |  | 32 | 32 | Distributed in FL. | Terminated | March 18, 2014 | 230 | 32 | Undetermined
Z-1538-2013 | 62624 | No | REA |  |  |  |  | Laserworld RGB &amp; Piko Nano RGB Models |  |  | July 19, 2012 | 2012 | Laserworld Usa Inc | A potential lack of certification, and/or missing laser certification, laser safety warning and administrative labels, and potentially incomplete Performance Feature set and incomplete User Information were discovered. | Radiation Control for Health and Safety Act | CDRH approves the CAP subject to the following conditions:1. Laserworld shall expand the scope of the CAP to include all projectors distributed between 2008 and 20l2, including, but not limited to, the Proline Series AG, Revolution Series AG, Club Series AG, Club Series, Pure Light Series, and Revolution Series AG, and Laserworld will continue to evaluate whether these additional projector models have similar defects.2. Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR 1003.21 and 1003.22. This office and the Food and Drug Administration (FDA) district office coordinator noted below are to be included in the notification.3. Laserworld will maintain records to document the number of projectors returned under the CAP, and will document the specific items of non-compliance addressed for each product tested, repaired, relabeled and/or replaced.4. Laserworld will provide CDRH with quarterly status updates on the progress of the CAP.5. The correction of the non-compliant products must be made at no cost to thecustomers.You may proceed with implementation of the CAP (subject to the conditions |  | 1004000 | Total distribution of more than 500 units between July1, 2008, and June 30, 2012 | US Distribution including the states of CA, CO, FL, KS and PA. | Open, Classified |  |  | 1004000 | Undetermined
Z-2259-2013 | 66276 | No | EBF | Tooth shade resin material. | Dental | Dental | 510(k) | NX3 Try-In Gel. The product is used as a tooth shade resin material. |  |  | September 13, 2013 | 2013 | Kerr/Pentron, DBA Kerr Corporation and Pentron Clinical | Kerr Corporation is voluntarily recalling one lot of NX3 Try-In Gel, because some of the NX3 Try-In Gel syringes in the affected lot contain a different product material. The material in the affected syringe does not match the shade of the cement as it is labeled. | Mixed-up of materials/components | Kerr Corporation initiated the voluntary recall by sending out notifications on 09/13/2013 via USPS 1st class mail. The recall notification letter, titled "URGENT: MEDICAL DEVICE RECALL", informed customer of the recall's product description, reason for recall, potential risk, instructions to customers and contact information. |  | 197 | 197 units | Worldwide Distribution, including Nationwide in the US and the countries of Canada, Great Britain, Australia, France, Germany, Italy, Denmark, Hong Kong, China, Brazil and Saudi Arabia. | Terminated | February 11, 2014 | 151 | 197 | Undetermined
Z-2261-2013 | 65981 | Yes | DTM | Cardiopulmonary bypass arterial line blood filter. | Cardiovascular | Cardiovascular | 510(k) | Terumo Pall AL6X Arterial Blood Line Filter, labeled in part Terumo Cardiovascular Systems Corporation, packaged as part of a Terumo Cardiovascular Procedure Kit or as a single, sterile filter unit |  |  | July 26, 2013 | 2013 | Terumo Cardiovascular Systems Corporation | Medical device used in cardiovascular procedures is defective. | Nonconforming Material/Component | Terumo Cardiovascuolar Services sent recall letters to Chief of Perfusion, Director of Operating Room Services, Director of Biomedical Services, and Risk Management departments on July 26, 2013. The notifications advised that filters are being recalled because some of the filters may exhibit a partial or full occlusion of an outlet or purge port. Terumo CVS advised users to stop using Terumo AL6X arterial filters from the affected lot numbers for replacement. A supply of replacement filters for customer's immediate use accompanied the notification. For users ordering the AL6X filters as part of a Cardiovascular Procedure Kit, Terumo CVS is providing a preliminary supply of replacement Terumo AL6X arterial filters in which users may follow the instructions to replace the filter or users may return the Cardiovascular Procedure Kit and Terumo CVS will replace the filter for them. Customers with questions were advised to contact Terumo CVS Customer Service at 800-521-2818.For questions regarding this recall call 800-521-2818. |  | 15817 | 15,817 filters | Nationwide Distribution including WA, IN, OK, TX, MI, CA, NY, CT, AZ, FL, WV, NC, VA, PA, TN, GA, AL, MA, SC, OH, MS, MO, SD, WI, AR, AND NV. | Terminated | July 28, 2015 | 732 | 15817 | Undetermined
Z-2262-2013 | 66049 | Yes | DTZ | Cardiopulmonary bypass oxygenator. | Cardiovascular | Cardiovascular | 510(k) | QUADROX iD HMOD 2030 Diffusion Membrane Oxygenator; QUADROX-i ADULT HMO 70000 / HMO 71000; QUADROX-iD Pediatric HMOD 30000. The QUADROX oxygenators are intended for use in extracorporeal circulation during cardiopulmonary bypass in cardiac surgery. |  |  | August 15, 2013 | 2013 | Maquet Cardiovascular Us Sales, Llc | The manufacturer has conducted a review of the results of product testing performed during the last twelve months and has identified specific lots of product which may not meet specifications for bond strength on the inlet and/or outlet connector. | Nonconforming Material/Component | Maquet sent a Urgent Product Recall letter/Response Form, dated 8/15/2013 via Fed Ex. On 8/26/2013 Maquet sent a letter to their customers informing them that all shipments of oxygenators and stainless steel hold clamps have been placed on a temporary hold due to an agency review in connection with FDA clearance. Customer service may be contacted at 1-888-627-8383, (option 2 followed by option 2) Monday though Friday between the hours of 8:00 am and 6:00 pm EST.Maquet issued a Certificate of Medical Neccesity (CMN) on September 17, 2013 to notify Quadrox customers that they may continue to receive and use their current inventory of Quadrox oxygenators and clamps by acknowledging the conditions provided in the CMN. |  | 1182 | 1,182 units | Worldwide Distribution, including Nationwide (US) and foreign countries. | Terminated | July 11, 2017 | 1426 | 1182 | Undetermined
Z-2263-2013 | 66247 | No | NZA | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Reliance Endoscope Processing System (REPS), STERIS Corporation, 5960 Heisley Road, Mentor, OH 44060-1834 USA www.steris.com. |  |  | August 29, 2013 | 2013 | Steris Corporation | The air pipe that delivers unfiltered ambient air from the procedure room into the water bottle on the listed FUJINON G5 &amp; G8 type flexible GI endoscopes does not connect to the main air/water channel within the light guide (LG) connection head and, as a result, the air pipe does not receive forced flow during the Reliance EPS processing cycle. | Use error | Steris sent an "URGENT PRODUCT SAFETY NOTICE" dated September 5, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. |  | 266 | 266 units | Worldwide Distribution-USA (nationwide) including the states of AK, CO, FL, GA, HI, ID, IL, KS, LA, MD, MN, MO, MT, NE, NM, OH, PA, SC, UT, &amp; WY and the countries of Australia, Canada &amp; Spain. | Terminated | July 03, 2014 | 308 | 266 | Undetermined
Z-2264-2013 | 66063 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | ENDOPLUS; Precision Laparoscopic Instruments; REF 346VA; Description 5mm Atraumatic Grasper S/A, Standard; Non-sterile; Manufactured by Endoplus; 750 Tower Rd., Suite A, Mundelein, IL 60060 USA. |  |  | July 15, 2013 | 2013 | Endoplus, Inc. | Multiple complaints reporting that the moveable jaw has broken during laparoscopic procedures. | Process control | Letters were sent on 7/15/2013 to all customers notifying them of the recall. The letters requested that customers return the recalled devices via private courier using an enclosed account number. Replacement devices will be made available upon customer request. |  | 81 | 81 units | Distributed in the states of FL, TX, and VA. | Terminated | February 08, 2016 | 938 | 90 | Undetermined
Z-2267-2013 | 65883 | No | DIO | Cocaine and cocaine metabolite test system. | Toxicology | Toxicology | 510(k) | DRI Cocaine Metabolite Assay, labeled as Multigent Cocaine Assay for Abbott Laboratories. |  |  | July 23, 2013 | 2013 | Microgenics Corp | Surveillance of reagent stability for specific lots of DRI Cocaine Metabolite Reagent show an unexpected decreasing change in the reaction rate, indicating a faster than expected reagent degradation. | Under Investigation by firm | Thermo Fisher Scientific sent an Urgent Medical Device Recall letters dated July 23, 2013, except Abbott labs who was mailed a letter on August 2, 2013. Abbott had previously been notified by telephone of the recall.The letter identified the product the problem and the action needed to be taken by the customer.Customers advised users on reagent stability issue with specific lots and the fact that it indicated a faster than expected degradation of the reagent. Microgenics asked that all use of the affected lots be stopped. Customers were informed that since each patient run is monitored with the controls at +/-25% of the cutoff, a possibility of reporting incorrect patient results is minimal to nonexistent. The use of appropriate controls addresses an risks of incorrect results.In addition to stopping use of the affected reagent, customers are asked to contact Thermo Fisher Scientific for replacement and to keep a copy of the recall letter for laboratory files. If the affected lots of reagent were forwarded to any other laboratories, they should be contacted immediately.. The enclosed acknowledgement form is to be completed and returned indicating that the letter was received, how much product, if any was on hand and destroyed.Questions and concerns may be addressed to Microgenics Technical Service at 1-800-232-3342 or 510-979-5417. |  | 1473 | 1473 units in distribution. | US Distribution including the states of DC, MA, NJ, TX. FL, CA, WA, NY, TX, CT, KY, LA, AR, TN, AL, GA, CT, VA, NC, KS, MD, OH and IL | Terminated | December 02, 2013 | 132 | 1473 | Undetermined
Z-0015-2014 | 66294 | No | LLH | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | Product Name: OSOM C. Difficile Toxin A/B Test |  |  | September 16, 2013 | 2013 | Sekisui Diagnostics LLC | Sekisuki Diagnostics has issued a Product Correction for certain lot numbers of the OSOM C. difficile Toxin A/B Test Kit, because it has been determined that there is potential for false negative results. | Other | Sekisui Diagnostic sent an Customer Recall Notification letter dated September 16, 2013, via email and/or FedEx. The letter identified the product the problem and the action needed to be taken by the customer. The customer notification letter informs the customers that Sekisui Diagnostics has issued a Product Correction for certain lots of the OSOM C. difficile Toxin A/B Test Kit (P/N 173). The customer notification lists the reason for recall, instructions to the customer, product description, required actions, and contact information. Customer Notification Letter lists the specific lot numbers being recalled for that catalog number. For each of these two catalog numbers, there is one letter directed to a distributor and one letter directed to an end user. Customer letters will be sent by email and/or FedEx. The letters have REQUIRED ACTIONS outlined in a box. If no response is received within 30 days, there will be directed follow-up communication. For further questions please call ( 858) 777-2600. |  | 391 | 391 kits | Worldwide Distribution - US (nationwide) and the countries of UK, Saudi Arabia, Dubai, Kuwait, Poland, Spain, India and Romania | Terminated | May 05, 2014 | 231 | 391 | Undetermined
Z-0030-2014 | 66310 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | Perkin Elmer 226 Sample Collection Device, Serial numbers: 06132903 - 06198900. |  |  | September 05, 2013 | 2013 | Perkin Elmer Health Sciences, Inc. | Some devices may be contaminated on the back side of the device with glue in the blood application area. | Material/Component Contamination | The firm, Perkin Elmer, sent an "URGENT - MEDICAL DEVICE RECALL" letter dated September 5, 2013 via email to its consignees. The letter describes the product, problem and actions to be taken. The customers were instructed to immediately notify all submitting facilities and personnel that have received the PerkinElmer 226 Sample Collection Device of the potential glue contamination and to distribute the enclosed record of actions form to all submitting facilities/personnel and encourage them to perform the inspection activities as listed in the letter; and complete and return the enclosed Recall Response Form return it to PerkinElmer Health Sciences via fax to: 864-299-8797 or email to: Rick.Haines@perkinelmer.com.Please contact PerkinElmer Health Sciences Greenville facility Quality Manager for further information 864-299-8787 (ext 104). |  | 65984 | 65,984 devices | US distribution in OH | Terminated | July 17, 2014 | 315 | 65984 | Undetermined
Z-0189-2014 | 66392 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | SIGMA SPECTRUM Infusion Pump with Master Drug Library component (SHARP LCD screen, 55074-2 revision B) |  |  | September 24, 2013 | 2013 | Baxter Healthcare Corp. | Baxter Healthcare Corporation is recalling the LCD display screen on the SIGMA SPECTRUM Infusion Pump with Master Drug Library for not meeting standards for withstanding an electrostatic discharge. | Under Investigation by firm | Baxter Healthcare Corporation sent an Urgent Product Recall letter dated September 24, 2013 to all affected customers via first class mail. The letter identified the affected product, problem and actions to be taken. Customers were instructed to locate the affected product, remove them from use and acknowledge your receipt of the recall notification by completing the provided Customer Reply Form and fax to 224-270-5467. Customers were asked to contact Baxter Healthcare Medina at 800-356-3454 (choose option 1) Monday through Friday during the hourse of 7:00 am and 7:00 pm Eastern Time to order replacement device(s). For technical questions contact Baxter Healthcare Medina at 800-356-3454, (choose option1) and for clinical questions contact Medical Information Services at Baxter 800-933-0303. |  | 346 | 346 units | Worldwide Distribution - USA Nationwide in the states of: AR, AZ, CA, CO, CT, FL, HI, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, NC, NH, NM, NY, OH, OK, OR, PA, TN, UT, VA, WA and WV and the country Canada. | Open, Classified |  |  | 346 | Undetermined
Z-0155-2014 | 66587 | No | JEA | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ORBIT PT. |  |  | July 24, 2013 | 2013 | Trumpf Medical Systems | An odor and smoke developed - The cause was determined to be a power supply unit sitting in the controller column. | Component design/selection | The user facility was contacted via phone in early August. They were contacted 9/17/2013 to schedule the replacement of the power pack because the new power pack became available. |  | 1 | One | US Distribution: NJ only. | Terminated | June 03, 2014 | 314 | 1 | Undetermined
Z-0136-2014 | 66101 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | Bayer brand A1CNow+ MULTI-TEST A1C SYSTEM TEST KITS (PART NUMBER 3021/08842610); Hemoglobin A1c Monitor; Glycosylated Hemoglobin Assay; Product is manufactured and distributed by |  |  | August 26, 2013 | 2013 | Bayer Healthcare Llc | Test results from these lots may produce a high-bias reading (by up to 20%) that is outside of limits defined for stability and product accuracy when product is stored at room temperature four (4) to five (5) months after release. | Pending | Bayer Healthcare sent an Urgent Medical Device Recall letter dated August 26, 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their inventory for the affected product and quarantine any affected product. Complete the attached Recall Response Form and fax to Bayer Customer Service. Customers were instructed to use the attached letter to notify affected next level customers of the recall. For questions regarding this recall call 408-524-2255, ext 290. |  | 138 | 138 Kits in US (Total 2452) | Worldwide Distribution - USA (nationwide) and Internationally to China, Argentina, Uruguay, Japan. | Terminated | January 14, 2014 | 141 | 138 | Undetermined
Z-0036-2014 | 66335 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Xoran Mini CT Scanner |  |  | July 27, 2010 | 2010 | Xoran Technologies, Inc. | A description of the defect in the product or the manner in which the product fails to comply with an applicable Federal standard; The MiniCAT CT Scanners subject to this notification failed to meet Xoran's specifications for dose rate of 125 ÂżÂż 3 kVp. These 18 complaints received include dose rates both on the high and low end of Xoran's specifications. See ATTACHMENT 1 for details. (e) A n evalu | Radiation Control for Health and Safety Act | All of the MiniCAT CT Scanners subject to these complaints were serviced by Xoran technicians and the out-of-specifications devices were brought into specification. This was accomplished by either readjusting the settings or by installing a new component. |  | 320 | 320 | Worldwide Distribution - US nationwide | Terminated | September 11, 2014 | 1507 | 320 | Undetermined
Z-0041-2014 | 66344 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Xoran MiniCAT CT Scanner ID ENT00001 |  |  | December 14, 2009 | 2009 | Xoran Technologies, Inc. | Failure of the gantry to rotate. The complaint alleged that the scanner would emit radiation but would not rotate. | Radiation Control for Health and Safety Act | The service report in response to this complaint stated that the Head select relay board, chip and fan were replaced. The machine (ENTOOOO I) was subsequently tested. All tests were successful. |  | 320 | 320 | Worldwide Distribution - US nationwide | Terminated | September 11, 2014 | 1732 | 320 | Undetermined
Z-0038-2014 | 66338 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Xoran Minicat00202 |  |  | July 27, 2010 | 2010 | Xoran Technologies, Inc. | The MiniCAT CT Scanners subject to this notification failed to meet Xoran's specifications for dose rate of 125 Âż 3 kVp. These 18 complaints received include dose rates both on the high and low end of Xoran's specifications. See ATTACHMENT 1 for details. (e) A n evaluation of the hazards reasonably related to defect or the failure to comply with the Federal standard; An evaluation of the health ha | Radiation Control for Health and Safety Act | All of the MiniCAT CT Scanners subject to these complaints were serviced by Xoran technicians and the out-of-specifications devices were brought into specification. This was accomplished by either readjusting the settings or by installing a new component. Xoran's Research &amp; Development group is evaluating the current specification to see if the field specifications are appropriate and to ensure that proper testing protocols are communicated and are in place for use by 3 rd party testing organizations. In addition, Xoran has initiated CAPA # 18 0 that is looking at the source block with regard to evaluating the stability of the dose rating. |  | 320 | 320 | Worldwide Distribution - US nationwide | Terminated | September 11, 2014 | 1507 | 320 | Undetermined
Z-0040-2014 | 66340 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Xoran MiniCAT00202 |  |  | July 27, 2010 | 2010 | Xoran Technologies, Inc. | The MiniCAT CT Scanners subject to this notification failed to meet Xoran's specifications for dose rate of 125 Âż 3 kVp. These 18 complaints received include dose rates both on the high and low end of Xorans specifications. | Radiation Control for Health and Safety Act | All of the MiniCAT CT Scanners subject to these complaints were serviced by Xoran technicians and the out-of-specifications devices were brought into specification. This was accomplished by either readjusting the settings or by installing a new component. Xoran's Research &amp; Development group is evaluating the current specification to see if the field specifications are appropriate and to ensure that proper testing protocols are communicated and are in place for use by 3 rd party testing organizations. In addition, Xoran has initiated CAPA # 18 0 that is looking at the source block with regard to evaluating the stability of the dose rating. |  | 320 | 320 | Worldwide Distribution - Nationwide | Terminated | September 11, 2014 | 1507 | 320 | Undetermined
Z-0037-2014 | 66337 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Xoran MiniCAT CT Scanner |  |  | July 27, 2010 | 2010 | Xoran Technologies, Inc. | A description of the defect in the product or the manner in which the product fails to comply wvith an applicable Federal standard; The MiniCAT CT Scanners subject to this notification failed to meet Xoran's specifications for dose rate of 125 Âż 3 kVp. These 18 complaints received include dose rates both on the high and low end of Xoran's specifications. See ATTACHMENT 1 for details. (e) A n evalu | Radiation Control for Health and Safety Act | All of the MiniCAT CT Scanners subject to these complaints were serviced by Xoran technicians and the out-of-specifications devices were brought into specification. This was accomplished by either readjusting the settings or by installing a new component. Xoran's Research &amp; Development group is evaluating the current specification to see if the field specifications are appropriate and to ensure that proper testing protocols are communicated and are in place for use by 3 rd party testing organizations. In addition, Xoran has initiated CAPA # 18 0 that is looking at the source block with regard to evaluating the stability of the dose rating. |  | 320 | 320 | Worldwide Distribution - Nationwide | Terminated | September 11, 2014 | 1507 | 320 | Undetermined
Z-0146-2014 | 66570 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | da Vinci Vision System Cart, a component of the |  |  | October 21, 2013 | 2013 | Intuitive Surgical, Inc. | Intuitive has identified that the metalized coating on the IS3000 HD lamp module may be incompatible with the illuminator control board, leading to an erroneous red Preventive Maintenance Message that shows up on monitors. | Nonconforming Material/Component | Intuitive Surgical sent an Urgent Medical Device Correction letter dated October 23, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to open the illuminator and replace any metal-finish lamp modules with plastic finish lamp modules, if customers do not have a replacement plastic-finish lamp module available, they may continue to use the metal-finish lamp module until they receive a replacement. If the message "illuminator lamp module error: Please reseat or replace lamp module" is displayed on the patient side monitor and the surgeon console, contact intuitive Surgical Customer Service. Customers were also instructed to examine any remaining or spare inventory to identify any metal-finish lamp modules and return them to Intuitive Surgical for credit. Customers were instructed to ensure all affected personnel are fully informed of this Urgent Medical Device Corrrection letter. Customers should complete the attached Acknowledgement Form and return it to Intuitive Surgical as instructed. Customers with questions should contact Intuitive Surgical Customer Service.For questions regarding this recall call 408-523-2244. |  | 684 | Standalone: 161 units; With Vision Starter Kit: 240 units; With IS3000 Vision Cart: 283 units | Worldwide Distribution - USA (nationwide) and Internationally tAustralia, Austria, Belgium, Brazil, *Bulgaria, Canada, Chile, China, CzechRepublic, Denmark, Finland, France, Germany, Great Britain, India, Israel, Italy, Mexico, Monaco, Netherlands, Norway, Pakistan, Panama, Qatar, Romania, Russian Fed., Saudi Arabia, Singapore, South Korea, Sweden, Switzerland, Taiwan, and Turkey. | Terminated | September 04, 2014 | 318 | 684 | Undetermined
Z-0147-2014 | 66382 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | UniCel DxH 800 and DxH 600 Coulter Cellular Analysis Systems |  |  | September 24, 2013 | 2013 | Beckman Coulter Inc. | Beckman Coulter is initiating a recall for UniCel DxH 800 Coulter Cellular Analysis System (PN 629029) and UniCel DxH 600 Coulter Cellular Analysis System (PN B23858) because the company's internal investigation indicates that specific lots of check valves used in several locations within the DxH 800 and DxH 600 systems may fail. A failure may result in an air or liquid leak at the VCS Waste Chamb | Under Investigation by firm | The firm, Beckman Coulter, sent an "URGENT PRODUCT CORRECTION" letter dated September 12, 2013, its customers. The letter includes the products descriptions and codes, issue, impact, action, resolution, instructions to the customers, and contact information. The customers were instructed to review the three samples analyzed prior to error message,if you experience a "Vent/overflow chamber detected overflow' error; to avoid physical contact with a possible leak by continuing to use personal protective equipment as recommended in the Instructions for Use and by national regulations; share this information with your staff and retain this notification as part of your Quality System documentation; if you have forwarded any of the affected products, provide them a copy of this letter; and to complete and return the enclosed RESPONSE FORM within 10 days via fax to: 786-639-7500 /7501/7502/7504; email: Regulatory.Notifications@beckmancoulter.com (note: please include "DxH system check valve letter dated 09/12/2013" in email subject field) and/or by mail to: Beckman Coulter, Inc., Attn: Regulatory Affairs Mail Code 31-B06, 11800 S.W. 147th Avenue, PO Box 169015, Miami Fl 36116-9015. .Your service representative will contact you to schedule an on-site visit for check valve inspection and replacement of any affected valves.If you have any questions concerning this notification, please contact our Customer Support Center: via our website, http://www.beckmancoulter.com/customersupport/support, via phone, call 800-526-7694 in the United States and Canada, and Outside US and Canada, contact your local Beckman Coulter Representaitive. |  | 1025 | 1,025 units | Worldwide distribution: US (nationwide) including Puerto Rico, and countries of: Australia, Belgium, Canada, China, Croatia, Czech Republic, France, Germany, Hong Kong, India, Ireland, Italy, Japan, Macao, Netherlands, New Zealand, Panama, Poland, Portugal Qatar, Russian Federation, Spain, Sweden, Switzerland, Taiwan, Thailand, Trindad and Tobago, Turkey, United Arab Emirates, and United Kingdom.. | Terminated | April 21, 2014 | 209 | 1025 | Undetermined
Z-0145-2014 | 66457 | No | FJL | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Individual units labeled in part: REF: 4630022; for S(a)line resectoscopes; Vaporization electrode, sterite; 4mm telescope 30 25 / 12: only for continuous irrigation sheath 24,5 Fr.; LOT 426111; 2011/06; 2016/06; STERILE; RICHARD WOLF; |  |  | September 18, 2013 | 2013 | Richard Wolf Medical Instruments Corp. | Richard Wolf Medical Instrument Corporation recieved a complaint stating that a 4630024 vaporization electrode was packaged in a 4630022 pouch. | Process control | The firm, Richard Wolf, sent an "Urgent: Medical Device Recall" letter dated September 16, 2013, to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to inspect their inventory of vaporization electrodes from lot 426111, contact Richard Wolf Medical Instrument Corporation at 800-323-9653 x 344 or ext 225 or ext. 380 to inform Richard Wolf Medical Instrument Corporation if the customer has or does not have any devices from this lot, return the affected devices with the return authorization provided by Richard Wolf Medical Instrument Corporation for credit, and complete and return the attached Recall Response form via fax to: 847-913-0924; email: dclark@richardwolfusa.com; or mailed to: Richard Wolf Medical Instrument Corporation, 353 Corporate Woods Parkway, Vernon Hills, IL 60061-3110.Customers who require urgent replacement of the affected devices can contact their customer service representative for replacement.If you have any questions, call 800-323-9653 x 344. |  | 78 | 78 cartons (234 units) | Nationwide distribution: US including states of: CA, GA, IL, KY, LA, MI, MD, NY, OR, TX, VA and Puerto Rico. | Terminated | April 24, 2017 | 1314 | 78 | Undetermined
Z-0167-2014 | 66461 | No | INN | Mechanical chair. | Physical Medicine | Physical Medicine | 510(K) Exempt | Aquatec Ocean VIP, Aquatec Ocean Dual-VIP, Aquatec Ocean E-VIP mobile shower and toilet commode with tilting function. |  |  | August 31, 2013 | 2013 | Invacare Corporation | The snap-on fixtures of the backrest cushion on the Mobile Shower and Toilet Commode Aquatec Ocean VIP and Ocean Dual VIP do not withstand a required pull out force of 500 N, when the product is used in a reclined position and or tilted position. This may result in the snap-on fixture coming unlatched. | Nonconforming Material/Component | Invacare sent an Urgent Medical Device Recall Correction letter to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to acknowledge their receipt and understanding of the urgent recall notice and indicate the number of affected product in their inventory on the enclosed Provider Response Card. Customers were instructed to examine their existing stock and quarantine any affected product and fax the completed Provider Response Card immediately to Invacare Corporation at 440-326-3544. Customers were responsible for ensuring that the recall was conducted to the customer level. Correction kits were provided to the customers with the communication. Customers were instructed to document the correction details on the enclosed Product Tracking Sheet and fax to Invacare Customer Service at 440-326-3544.For questions regarding this recall call 440-329-3544. |  | 538 | 538 (US only) 3317 (Total) | Worldwide Distribution - USA including PA and Internationally to Canada. | Terminated | September 15, 2014 | 380 | 538 | Undetermined
Z-0130-2014 | 66545 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | GENESIS(R) II, NON-POROUS TIBIAL BASE, SIZE 5, LEFT, QTY: (1), REF 71420168. |  |  | October 07, 2013 | 2013 | Smith &amp; Nephew Inc | Swapped product. Non-porous tibial base size 5 L was mis-packaged as size 6 R and vice versa. | Process control | Smith &amp; Nephew sent an Urgent - Product Recall 1st Notification letter to all affected sales staff via email and FedEx on October 7, 2013. The hospital was notified by sales rep on October 9, 2013 and the notification letter was sent to the hospital via Fed Ex on October 11, 2013. The letter identified the affected product, problem and actions to be taken. Customers were instructed to inspect inventory, discontinue use, quarantine affected products immediately and return to Smith &amp; Nephew. Customers were asked to complete the attached Inventory Return Certification Form and fax to 901-566-7975. |  | 4 | 4 units | USA Nationwide Distribution in the states of: AZ, NC, IA, PA, OH, MD, and TN. | Open, Classified |  |  | 4 | Undetermined
Z-0151-2014 | 66465 | No | LTI |  |  | Gastroenterology/Urology | PMA | Adjustable Gastric Band, packaged alone in transparent PETG inner and outer blisters with Tyvek lids. |  |  | September 30, 2013 | 2013 | Ethicon Endo-Surgery Inc | During aging studies, the firm determined that the gastric belts were unable to pass the test requirements established for Force to Lock. | Unknown/Undetermined by firm | Ethicon sent an Urgent: Medical Device Recall letter dated September 30, 2013 to affected consignees via overnight, UPS to allow tracking of the receipt. The letter identified affected products, problem and actions to be taken. Each consignee is requested to respond to a provided Business Reply Form (BRF), indicating whether they have affected product, how much and whether they will/have returned the product or whether they no longer have any of the recall product. For questions contact your local sales representative or call our Customer Support Center 1-877-384-4266 option 6. |  | 1248 | Product Code: BD2XV - 1,248 units | Worldwide Distribution: USA (nationwide) including the states of: Alabama, Connecticut, Indiana, Michigan, North Carolina, South Carolina, Alaska, Delaware, Kansas, Minnesota, Ohio, Tennessee, Arizona, Florida, Kentucky, Mississipp,i Oklahoma, Texas, Arkansas, Georgia, Louisiana, Nebraska, Oregon, Virginia, California, Idaho, Maryland, Nevada, Pennsylvania, Washington, Colorado, Illinois, Massachusetts, New Jersey, Rhode Island, and West Virginia; and the countries of: Argentina, France, Malaysia, Spain, Australia, Georgia, Martinique, Sweden, Belgium, Germany, Mexico, Switzerland, Brazil, Greece, Netherland, Taiwan, Chile, Hungary, Panama, Tunisia, China, India, Philippines, United Arab Emirates, Czech Republic, Israel, Poland, United Kingdom, Egypt, Italy, Portugal, Venezuela, Estonia, Kenya, Russia, Finland, Kuwait, and Saudi Arabia. | Terminated | March 26, 2014 | 177 | 7438 | Undetermined
Z-0153-2014 | 66253 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ISE REF Valve, Part No. MU7638 for the AU480, AU680, and AU5800 Chemistry Analyzer. |  |  | September 09, 2013 | 2013 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed that some AU Clinical Chemistry Analyzer ISE units can experience premature ISE REF (Reference) valve failure. Premature valve failure can be detected by the presence of air bubbles in the ISE REF solution tubing on the ISE unit. If air bubbles are present, accuracy of the ISE test results may be affected. | Nonconforming Material/Component | A customer notification letter dated 9/9/13 was sent to customers who received the ISE REF Valve, Part Number MU3226 for the AU400, AU640, AU2700, AU5400 and the ISE REF Valve, Part Number MU7638 for the AU480, AU680, AU5800 Chemistry Analyzers. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to complete and return the enclosed Response Form within 10 days. Customers with questions are instructed to contact Beckman Coulter Customer Support Center via http://www.beckmancoulter.com/customersupport/support, via phone at (800) 854-3633 in the US and Canada. Outside of the US and Canada, customers are instructed to contact their local Beckman Coulter Representative. |  | 4584 | 4584 units total (1173 units in US) | Worldwide Distribution-USA (nationwide) and the countries of Canada, Vietnam, Zambia, and Zimbabwe. | Terminated | September 30, 2014 | 386 | 4584 | Undetermined
Z-0164-2014 | 66351 | No | CAZ | Anesthesia conduction kit. | Anesthesiology | Anesthesiology | 510(k) | PortexÂż Saddleblock Tray with Drugs  25g Quincke; 4795-20 |  |  | September 23, 2013 | 2013 | Smiths Medical ASD, Inc. | Visual particulate in the glass ampules of 5% Lidocaine HCL and 7.5% Dextrose Injection, USP, 2 mL Ampoule, NDC # 0409-4712-01, Hospira Lot Number 23-227-DK. These ampules are included in certain Portex Spinal Anaesthesia Trays. | Material/Component Contamination | Smiths Medical sent an Urgent Recall Notice dated September 23, 2013, to all affected customers. The notice identified the product, the problem, and the action to be taken by the customer. Customers were instructed to inspect their inventory for the affected product and remove from use. Complete the Confirmation Form and return by fax to 603-358-1017 or by email to spinal@smiths-medical.com. Upon receipt of the completed form, a customer service representative would contact them to arrange for exchange of their unused affected devices for credit or replacement. Customers were also instructed to forward the notice to all personnel who need to be aware within their organization. Customers with questions were instructed to contact Smiths Medical's Customer Service Department at 800-258-5361.For questions regarding this recall call 800-258-5361. |  | 222 | 222 | Nationwide Distribution including AK, AR, AZ, CA, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NJ, NY, OH, OK, OR, PA, RI, SC, TX, UT, VA, VT, WA, WI, WY. | Open, Classified |  |  | 7549 | Undetermined
Z-0154-2014 | 66532 | Yes | NIQ |  |  | Cardiovascular | PMA | Resolute Integrity Zotarolimus-eluting Coronary Stent: |  |  | October 07, 2013 | 2013 | Medtronic Vascular, Inc. | A small number of units in lot number 0006573585 of the Resolute Integrity Zotarolimus-Eluting Coronary Sten System may contain a compliance chart incorrectly referencing a 2.25 mm diameter instead of the correct 3.0mm. | Process control | Urgent Medical Device Recall Letters were sent on October 7, 2013 via UPS 3-day service. |  | 201 | 201 units | Nationwide Distribution. | Terminated | February 11, 2014 | 127 | 201 | Undetermined
Z-0141-2014 | 66138 | [6 empty] | Harmony Equipment Management System, Model #603X/80XX, STERIS Corporation, Mentor, OH |  |  | August 15, 2013 | 2013 | Steris Corporation | DrÂżger Medical GmbH has notified STERIS of the potential for a break in the drive screw contained in the motorized ceiling supply unit, a component of the Harmony EMS system. The ceiling supply unit controls the motorized adjustment of the height and rotation of the supply column. Should a break occur in the drive screw, there is potential for the adjustable arm to drop down, causing the EMS sys | Component design/selection | STERIS sent an Urgent Voluntary Field Correction Notice dated August 15, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.STERIS action- Drager Medical has provided STERIS instructions for inspection of the alignment of the mounting plate of the motor unit. If it is found that the mounting plate is bent, the drive screw within the motor unit must be replaced. STERIS will visit your facility to inspect the motorized ceiling supply unit of your Harmony EMS system(s) and if damage is observed, will replace the drive screw within the motor unit. For further information or if you have any questions regarding our visit to your facility, please contact STERIS Field Service Dispatch at 1-800-333-8828. |  | 127 | 127 units | Nationwide distribution: USA including the states of Alabama, Arkansas, Florida, Georgia, Illinois, Indiana, Kansas, Kentucky, Louisiana, Maryland, Massachusetts, Michigan, Minnesota, Missouri, Nebraska, New Jersey, New York, Ohio, Oklahoma, Oregon, Pennsylvania, Texas, Virginia, Washington and Wisconsin. | Terminated | April 28, 2014 | 256 | 127 | Undetermined
Z-0143-2014 | 66249 | No | OBK | Subcutaneous, implanted, intravascular infusion port and catheter. | General Hospital | General Hospital | 510(k) | Dual Luer Lock Cap |  |  | September 10, 2013 | 2013 | Baxter Healthcare Corp. | The firm is recalling lots 10043 and 10044 due to loose particulate matter found in the packaging. Particulate matter entering the fluid path from the Luer lock caps may result in embolic events. | Employee error | Baxter sent an Urgent Product Recall notification dated September 16, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were advised to :1. Locate and remove all affected product from your facility. If you distribute these products to other facilities or departments within your institution, please forward a copy of this communication to ensure that they locate and remove affected product from potential use.The product code and lot number can be found on the shipping carton label and individual product package.2. Acknowledge your receipt of this recall notification by completing the attached Customer Reply Form (Attachment 1) and returning it to Baxter by either faxing it to 1-224-270-5457 or scanning and emailing it to fca@baxter.com. Returning the Customer Reply Form promptly will prevent you from receiving repeat notifications.3. The affected product should be returned for credit by contacting Baxter Healthcare Center for Service. The Center for Service can be reached at 1-888-229-0001 between the hours of 7:00 am and 6:00 pm Central Time. Please have your Baxter eight-digit ship-to account number ready when calling. Please have the product code and lot number(s) available.4.If you are a dealer, wholesaler,or distributor/reseller that distributed any product to other facilities, please forward this communication. Baxter distributed this product to customers between June 19, 2013 and August 20, 2013.The Center for One Baxter is available at 1-800-422-9837, Monday throughFriday during the hours of 8:00 am to 5:00 pm Central Time. |  | 628992 | 628,992 units | Worldwide Distribution - USA (nationwide) and the country of Canada. | Terminated | June 05, 2015 | 633 | 628992 | Undetermined
Z-0181-2014 | 66267 | No | DLZ | Phenobarbital test system. | Toxicology | Toxicology | 510(k) | Emit 2000 Phenobarbital Assay and Syva Emit 2000 Phenobarbital Assay |  |  | August 30, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Imprecision with certain lots | Nonconforming Material/Component | Siemens sent an Urgent Medical Device Recall letter dated August 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately cease use of the affected product. Siemens will replace unused affected product at no charge. For replacements customers were instructed to incude this information on the Effectiveness Check form included with the letter. The Effectiveness Check form must be returned even if no replacement reagents are needed. The recall letter was to be retained with customers laboratory records and forwarded to those who may have received this product. Customers with questions were instructed to contact Siemens Technical Solutions or their Siemens technical support representative.For questions regarding this recall call 800-441-9250. |  | 4963 | 4963 | Worldwide Distribution - USA (nationwide) and Mexico. | Terminated | August 27, 2014 | 362 | 4963 | Undetermined
Z-0175-2014 | 66475 | ph buffer | JCC | General purpose reagent. | Pathology | Pathology | 510(K) Exempt | pH 7.0 Buffer Solution Pint and Quart Bottles, Part Number 02.0030 for Pints and 02.0031 for Quarts. Packaged in Plastic Bottles. Can be sold individually or in cartons with 6 bottles per box (for quart bottles only). |  |  | September 13, 2013 | 2013 | Mesa Laboratories, Inc. | Mesa Laboratories, Inc. is recalling certain lots of pH 7.0 Buffer Solution due to mold growth. | Material/Component Contamination | An email letter was sent to consignees on October 2, 2013 asking all customers to discard the product and a replacement product and tubing will be issued. |  | 58466 | 58,466 bottles | Worldwide Distribution-USA (nationwide) and the country of Canada. | Open, Classified |  |  | 58466 | Undetermined
Z-0191-2014 | 66145 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | DePuy Mitek FMS Intermediary Tubing with One-Way valve |  |  | August 29, 2013 | 2013 | DePuy Mitek, Inc., a Johnson &amp; Johnson Co. | Pillow valve included in the FMS Duo outflow tubing set and the FMS Solo intermediary tubing set may not perform as intended. may result in backflow of irrigation fluid into the One Day Set which may potentially lead to patient cross-contamination | Component change control | DePuy Mitek inititaed telephone notification on Ausgust 29, 2013 and sent an Urgent Voluntary Product Recall letter dated September 5, 2013 to US and OUS affected customers/affiliates. The customers were notified of the affected products, problem and actions to be taken. The customers were instructed to immediately check all inventories to locate and return affected product following the enclosed instructions. Stericycle is cooridianting the returns. If you have any questions or concerns in regards to this recall, please contact Stericycle directly at 1-866-737-1928. |  | 368 | 368 pk | Worldwide Distribution - USA (Nationwide) and in the countries of: Australia , Austria , Belgium , Canada , Czech Republic , Denmark , Finland , France, Germany , India, Ireland , Israel, Italy , Latvia , Netherlands , Norway , Nepal, Poland , Portugal , Spain , Sweden , Switzerland , United Kingdom. | Open, Classified |  |  | 29648 | Undetermined
Z-0135-2014 | 66112 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Edwards Lifesciences Dispersion Aortic Perfusion Cannula with Duraflo Coating, Sterile, RX only For Single Use Only, |  |  | August 26, 2013 | 2013 | Edwards Lifesciences, LLC | Edwards Lifesciences is recalling two lots of Duraflo Cannulas due to reported complaints involving cracks in the t-connector of the devices which were confirmed during product evaluation. | Process control | Edwards sent an Urgent Field Safety Notice Product Recall letter dated August 26, 2013 to all affected customer, via Fed-Ex. The product is sold directly to hospitals and distributors. The distributors involved will communicate with their customers directly. The letter identified the affected product, problem and actions to be taken. Customers were instructed to review their entire inventory, quarantine and return the affected products. An acknowledgement form is included for completion to assist in the assessment of the inventory. For questions contact Edwards Customer Service at 800-424-3278. |  | 140 | 140 units | Nationwide Distribution. | Open, Classified |  |  | 440 | Undetermined
Z-0133-2014 | 66504 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | The EkoSonic Endovascular System employs high frequency (2-2.5 MHz), low power (0.585 watt average power per transducer) ultrasound to facilitate the delivery of physician-specified fluids in the peripheral vasculature and to deliver solutions into the pulmonary arteries. |  |  | October 02, 2013 | 2013 | EKOS Corporation | The EkoSonic Control Unit is recalled because it had the incorrect power entry module installed. It is missing an AC filter to reduce the noise and electrical interference. | Mixed-up of materials/components | EKOS notified their customer by email on October 2, 2013, about the recall. EKOS has requested the customer to return of EkoSonic Control Unit PT3B-783, so that it can be replaced.CORRECTIVE ACTION:This single affected unit will be reworked with the correct power entry module and released into finished goods as a refurbished control unit.EKOS did not distributed any affected devices in the US.For any questions, please call EKOS Corporation at 425-415-3100. |  | 1 | one unit | Worldwide Distribution in Germany only. | Terminated | November 22, 2013 | 51 | 1 | Undetermined
Z-0031-2014 | 65952 | No | GAD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | UNITRAC Pneumatic Retraction Arm. |  |  | July 16, 2013 | 2013 | Aesculap, Inc. | Changes were implemented to the UNITRAC instructions for use to ensure safe use and potentially extend the life of the product. This included recommending the use of Sterile Cover JG901 and new Quick Release Coupling RT02OR to eliminate the need to sterilize the UNITRAC and a functional check was added prior to use to ensure optimal functionality of the UNITRAC system. | Process change control | A letter dated July 16, 2013 was sent requesting users to review the new instructions for use in their entirety, sign the attached Hospital Acknowledgment form and return to Aesculap. An Aesculap representative would be in contact to provide an in-service regarding the process for set-up of the Sterile Cover and required Functional Check. Discontinuation of use and repair of the device was recommended if the joints did not move freely or if the set position drifts during the Functional Check. |  | 217 | 217 | Nationwide Distribution. | Terminated | August 26, 2014 | 406 | 217 | Undetermined
Z-0087-2014 | 66290 | [6 empty] | chromID Salmonella Agar, Reference No. 43621. |  |  | September 06, 2013 | 2013 | bioMerieux, Inc. | Complaints for Salmonella strains growing on this medium producing an uncharacteristic white colored colony instead of the expected pale pink to mauve colored colonies. This could potentially lead to Salmonella colonies not being detected. | Nonconforming Material/Component | Recall letters were issued to the US customers by Biomeriux, Durham, NC on 9/6/2013. |  | 90 | 90 | Nationwide Distribution. | Terminated | June 24, 2014 | 291 | 90 | Undetermined
Z-0049-2014 | 66304 | No | LZS |  |  | Ophthalmic | PMA | STAR Excimer Laser System, Model STAR S4 IR, Catalog Number: 0030-4864, Product is manufactured and distributed by AMO Manufacturing USA, LLC |  |  | September 16, 2013 | 2013 | Abbott Medical Optics, Inc. | This corrective action will address an identified failure of the lamp and illumination system within the laser unit caused by an unexpected failure of a thyristor component on the AC Power Distribution Printed Circuit Board (PCB). | Nonconforming Material/Component | The firm, Abbott Medical Optics, Inc., (AMO), sent an "ADVISORY NOTICE and a COVER LETTER FOR ADVISORY NOTICE" dated September 16, 2013 to all affected customers. The letters identified the product, problem, and actions to be taken by the customers. The customers were instructed to complete and return the Star Notification Acknowledgement form via fax to 1-714-247-4530 within 3 business days or mail to Abbott Medical Optics, 510 Cottonwood Drive, MS 347118, Milpitas, CA 95035, and share this advisory notice with your staff.If you have any questions related to operation of STAR Excimer Laser System or documentation requirements associated with this letter, please call AMO Technical Support at 1-800-511-0911 and you will be directed to an AMO Representative. |  | 834 | 834 units in total | Worldwide distribution: US (nationwide) including Puerto Rico and countries of: Asia Pacific, Canada, Europe, Africa, Middle East, Japan, and Latin America. | Terminated | November 12, 2014 | 422 | 834 | Undetermined
Z-0044-2014 | 66317 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | STERRAD 100NX Cassettes |  |  | September 18, 2013 | 2013 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) is recalling the three lots (13F049, 13F050, 13F051) STERRAD 100NX Cassettes because ASP has determined that the leak indicator on the cassettes is effective for only 12 months of the cassette's labeled 15-month shelf-life. | Under Investigation by firm | Advanced Sterilization Products (ASP) sent an Urgent Medical Device Recall letter on September 18, 2013, to all affected customers. The letter informed the customers of the problems identified and the actions to be taken. Customers were instructed to examine their inventory and return all affected cassettes using the enclosed prepaid UPS return lave to Stericycle. Customers were asked to count their inventory and record the data on the enclosed Business Reply Card and Packing Slip. Customers were asked to provide the notice to anyone in their facility that needs to be informed. Customers were also asked to maintain a copy of the notice with the affected product. Customers with questions or suspected problems were instructed to contact Stericycle directly at 1-877-257-7119.For questions regarding this recall call 949-453-6400. |  | 2995 | 2,995 units | Nationwide Distribution | Terminated | January 13, 2014 | 117 | 2995 | Undetermined
Z-0043-2014 | 66328 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | STERRAD 200 Sterilizer (Single Door), Part Number: 10201. |  |  | September 18, 2013 | 2013 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) is recalling the STERRAD 200 Sterilization System because it may emit an odor or smell into the surrounding environment that is not typical to normal operating conditions. | Under Investigation by firm | A customer notification letter was sent on September 18, 2013, to all customers who received the STERRAD 200 Sterilization System. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to indicate if they have received a STERRAD 200 System and to report any complaints or suspected problems to Advanced Sterilization Products (ASP) Professional Services directly at (888) 783-7723 and select option #2.An alternate design for the oil mist filter is being developed and qualification of new vacuum pump oil is in process. This action is to modify the PM interval to service the filter and oil more frequently. This will serve as an immediate/interim solution to reduce incidence of odor/smell complaints until the new oil mist filter design and vacuum pump oil are released.A Customer Letter will be sent to all STERRAD 200 customers to notify them of the issue and the corrective actions we are taking to resolve this issue. |  | 1305 | 1,305 units total (654 units in US , 651 OUS ) | Worldwide Distribution - USA (nationwide) and the countries of EA (Europe Middle East and Asia), APAC (Asia Pacific), Japan and Canada | Terminated | April 11, 2014 | 205 | 1305 | Undetermined
Z-0045-2014 | 66400 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Merit Medical Performa Angiographic Catheter 5F Multipack. Part numbers 7560-43 and 7560-G3. |  |  | September 13, 2013 | 2013 | Merit Medical Systems, Inc. | 4 lots of Performa Cardiac Multipacks contain a JR4.0 catheter with a slight variation in tip shape. | Process control | Consignees were notified on 9/13/13 via Urgent Product Recall Notice, email, and product retrieval forms. Actions required of customers:1. Please immediately determine if any of the devices are within your facility, quarantinethem and discontinue use.2. Ensure that all personnel to whom the devices were distributed are made aware of this field action.3. Contact your Merit Sales Representative in order to arrange product return and replacement or credit.4. If you have any questions concerning this communication, please don't hesitate to contact Guido Sandulli at 1-801-826-4046. |  | 200 | 200 | US Distribution including the states of CO, VA, and MS. | Terminated | January 03, 2014 | 112 | 200 | Undetermined
Z-0054-2014 | 65742 | No | GWM | Intracranial pressure monitoring device. | Neurology | Neurology | 510(k) | Integra Licox Brain Tissue Oxygen Monitoring; 1 x Complete Brain Probe Kit |  |  | June 11, 2013 | 2013 | Integra LifeSciences Corp. | Integra LifeSciences Corporation has determined that there is a possibility that some lots of Licox CMP Brain Tissue Oxygen and Temperature Catheter Kits contain a hex wrench that does not loosen the set screw on the drill bit. | Nonconforming Material/Component | Integra LifeSciences Corporation sent a recall letter/return response form on 6/11/2013 by traceable courier service, traceable e-mail or facsimile. |  | 21 | 21 kits | Worldwide distribution: US (nationwide) and countries of: AR, BR, CH, CL,CR. DE, ES, FI, FR, GB, IL, IT, MX, PA, PL, PT, RU, TW, and ZA. | Terminated | February 18, 2014 | 252 | 1663 | Undetermined
Z-0093-2014 | 65897 | No | GCC | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | PIPELINE LS ACCESS SYSTEM, 15 Dilator, 5 mm, Insulated |  |  | July 31, 2013 | 2013 | DePuy Spine, Inc. | Dilators and Holding Clip are being used for a secondary function of neuromonitoring, for which it is not 510(k) cleared. | Under Investigation by firm | DePuy Synthes sent an Urgent Voluntary Product Recall Notification letter dated July 30, 2013, to all affected consignees via Fed Ex overnight delivery.Consignees were informed that DePuy Spine, Inc. is voluntarily recalling all lots of the PIPELINE LS ACCESS SYSTEM Dilators (1 st, 2nd, 3rd, and 4th) and Holding Clip. They were asked to review inventory located at their facility and quarantine any products listed in this recall notice. Consignees were instructed to contact their local DePuy Synthes Spine Sales Consultant to return any inventory located within their facility.Consignees with questions were instructed to call 508-828-3583. |  | 118 | 118 units | Nationwide Distribution including the states of AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MT, NC, ND, NE, NH, NJ, NY, OH, OK, OR, PA, SC, SD, TN, TX, VA, WA, WI, WV, and WY. | Terminated | June 29, 2015 | 698 | 623 | Undetermined
Z-0053-2014 | 66394 | No | DIF | Clinical toxicology control material. | Toxicology | Toxicology | 510(K) Exempt | Alere TriageÂż TOX Drug Screen Control 1 |  |  | September 23, 2013 | 2013 | Alere San Diego, Inc. | Alere San Diego, Inc. initiated a voluntary recall of Alere Triage TOX Drug Screen Control 1 (PN 94413, Lot C2681A) because the CODE CHIP that was provided with this lot causes the control to fail when run in the QC Sample Mode using Alere Triage TOX Drug Screeen products. The Alere Triage TOX Drug Screen test devices and patient results are not affected. | Under Investigation by firm | Alere San Diego, Inc. sent an Urgent Medical Device Notice dated September 23, 2013, to all affected customers via FedEx and outside US (OUS) consignees on September 24, 2013, via email. The customer notification letter titled "URGENT MEDICAL DEVICE NOTICE", informed customers of the recall, product description with part number and lot code (Alere TriageÂż TOX Drug Screen Control 1 PN 94413 Lot C2681A), reason for recall and impact (The purpose of the letter was to inform customers that the CODE CHIP" that was provided with this lot causes the control to fail when run in the QC Sample Mode using Alere TriageÂż TOX Drug Screen products. Customers were instructed to discontinue use of and discard the CODE CHIP contained within the Alere Triage TOX Drug Screen Control 1 PN 94413 Lot C2681A kit box. Customers were also instructed to provide a copy of the notice to their customers and instruct them to contact Alere for a replacement code chip. Customers were asked to complete and fax the enclosed Verification Form to confirm their receipt of the notice. Customers with questions were instructed to contact:(Alere San Diego, Inc.9975 Summers Ridge RoadSan Diego, CA 92121, U.S.A.Phone: 877 308 8287FAX: 858 805 8457E-mail: Responses.ts@alere.com).For questions regarding this recall call 858-805-2000, ext 3015. |  | 38 | 38 | Worldwide Distribution - USA (nationwide) and countries of: South Africa and Canada. | Terminated | September 26, 2014 | 368 | 38 | Undetermined
Z-0065-2014 | 66121 | No | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | VARISOURCE Breast Template System for high-dose rate Brachytherapy. |  |  | August 09, 2013 | 2013 | Varian Medical Systems, Inc. | The recall was initiated after Varian Medical became aware the Lexan templates of the Breast Bridge Template System could become bent and turn opaque after steam sterilization. | Nonconforming Material/Component | Urgent Medical Device Correction, Urgent Field Safety Notice were mailed by FedEx tracked delivery to all consignees on August 12, 2013. |  | 3 | 3 sets | Worldwide distribution: US (nationwide) distribution in CA and in the country of India. | Terminated | April 10, 2014 | 244 | 3 | Undetermined
Z-0083-2014 | 66322 | No | MIH |  |  | Cardiovascular | PMA | Valiant Thoracic Stent Graft with Captivia Delivery System, |  |  | September 19, 2012 | 2012 | Medtronic Vascular, Inc. | Firm received reports of difficulty in flushing the graft cover of the delivery system during device preparation as required by IFU. | Component design/selection | A Medical Device Information letter was sent to all affected customers on September 19, 2012. Additionally, all Medtronic Field Representatives were notified of the issue via a Field Memo |  | 1000 | ~1000 | International distribution:Austria, Belgium, Cyprus, Czech Republic, Denmark, Finland, Germany, Greece, Hungary, Iceland, Ireland, It lay, Latvia, Lithuania, Luxembourg, Netherlands, Norway, Poland, Portugal, Slovakia, Slovenia, Spain, Sweden, Switzerland, Turkey, United Kingdom. | Terminated | October 25, 2013 | 401 | 1000 | Undetermined
Z-0092-2014 | 66352 | No | DZJ | Bone cutting instrument and accessories. | Dental | Dental | 510(k) | REF 5100-088, TPS Micro Driver, For Use with Stryker Total Performance System |  |  | October 02, 2013 | 2013 | Stryker Instruments Div. of Stryker Corporation | It was discovered that the silicone potting failed to properly cure after undergoing the prescribed curing possess. The potting material is used to insulate, stabilize and provide a moisture barrier to the TPS MicroDriver main Printed Circuit Board Assembly (PCBA). Inadequate cure of the silicone potting could potentially allow it to leak to the outside of the handpiece or cause an electrical ma | Nonconforming Material/Component | Stryker sent an Urgent Medical Device Correction Notification letter via overnight mail and interntion consignees via E-mail on October 2, 2013. The Urgent Notice identified the recalled product, the reason for the recall, potential adverse risks to health associated with the recall. Customers were instructed to:1. Immediately review this Recall Notification2. Locate, and quarantine, the units listed in this notification. Do not use the recalled TPS MicroDriver until it has been repaired. A list, of units shipped to your facility, is included on the attached Business Reply Form.3. Return the enclosed Business Reply Form to confirm receipt of this notification and identify how many affected units are currently in your inventory.4. If you have further distributed this product, please forward this letter and the attached Business Reply Form (BRF) to all affected locations. Please indicate each location on the BRF.5. Fax the completed Business Reply Form to Stryker Instruments Regulatory Department, 866-521-2762.6. Upon receipt of the Business Reply Form, Stryker will send you a pre-paid shipper to send your affected device(s) back to Stryker to be repaired. Unit will be repaired and returned to your facility7. Send back all affected devices using the pre-paid shipper provided to you by Stryker.8. If you require a loaner, please contact the Stryker Service Department at 888-308-1983.For questions regarding this recall call 269-389-4518. |  | 67 | 67 | Worldwide distribution: US (nationwide) and countries of: Australia, Canada, France, Germany, Italy, Japan, Spain, Switzerland, and United Kingdom. | Terminated | April 15, 2014 | 195 | 67 | Undetermined
Z-0085-2014 | 66368 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Baylis Medical TorFlex Transseptal Guiding Sheath, Rx Only, non-pyrogenic, sterile. |  |  | September 23, 2013 | 2013 | Baylis Medical Corp * | The TorFlex Transseptal Guiding Sheath Kits are being recalled due to the possibility that the sheath could remove particulate of less than 5 mm from the dilator. The particulate could potentially migrate into a patients bloodstream. | Component design/selection | Baylis Medical sent an Urgent Device Recall Notification letter dated September 23, 2013, via overnight courier to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete the response form provided and fax to 905-602-5671 ATTN: Quality Department to arrange for a free of charge replacement of the affected product. The form requires consignees to return the affected product. For questions contact Quality Assurance at 800-276-4416. |  | 370 | 370 devices | Worldwide Distribution - USA Nationwide in the states of : NE, NH, NC, FL, PA, VA, MI, CA, AL, and TN and the countries of China, Canada, Saudi Arabia, and Italy. | Terminated | March 31, 2015 | 554 | 370 | Undetermined
Z-0088-2014 | 66434 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | da Vinci Si Surgical System IS3000 with the da Vinci Surgeon Console which contains a high Resolution Stereoscopic Viewer. |  |  | October 17, 2011 | 2011 | Intuitive Surgical, Inc. | Improper restraints during transportation of the da Vinci system could cause the HRSV Monitor to become loose from its mount. | Use error | Advisory letters were sent to all affected customers on 10/17/11 informing them to contact Customer Service prior to moving the system outside of the hospital. |  | 1090 | 1,090 consoles | Worldwide Distribution-USA (nationwide) including Puerto Rico and the countries of Australia, Austria, Belgium, Canada, Chile, China, Denmark, Egypt, Finland, France, Greece, Germany, India, Indonesia, Israel, Ireland, Italy, Kuwait, Mexico, Netherlands, Norway, Philippines, Poland, Portugal, Qatar, Romania, Saudi Arabia, Singapore, Slovakia, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, and United Kingdom. | Terminated | October 28, 2013 | 742 | 1090 | Undetermined
Z-0084-2014 | 66379 | No | FCX | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | PENTAX Endoscopic Accessories. OF-B194. Made in Japan. |  |  | September 16, 2013 | 2013 | Pentax Medical Company | PENTAX Medical has become aware of a product issue associated with the OF-B194 Gas/Water Feeding Valve. A manufacturing defect may cause improper stoppage of CO2 gas during an endoscopic procedure with the potential for serious hazard to and/or user. | Use error | PENTAX Medical sent an "Urgent Device Recall" notification letter with /Response forms dated September 16, 2013, via UPS Certified delivery to all affected customers. . The notification identified the affected product; informed the customers of the nature of the issue; advised of potential adverse patient consequences that may occur. The customer was asked to immediately examine their inventory and quarantine product subject to the recall. The customer was instructed to identify and notify their customers if they further distributed the product. The notification instructs users to return the OF-B194 Valves for replacement with corrected product. Customers were asked to complete the attached response form and fax back to Pentax Medical at 1.201.799.4063. An enclosed UPS Airway Bill is to be used to return recalled product. A contact phone number and email address is provided if the customer has questions (1.800.431.5880 xt2064).For questions regarding this recall call 800-431-5880.Pentax issued a press release on 10/31/2013.Pentax posted their press release on their website at www.pentax.com. |  | 76 | 76 units | Nationwide Distribution including Arizona, California, Florida, Idaho, Indiana, Kansas, Maryland, Missouri, Missippi, North Carolina, New York, Pennsylvania, Virginia, Wisconsin and Wyoming. | Terminated | July 09, 2014 | 296 | 76 | Undetermined
Z-0099-2014 | 66238 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Baxter CLEARLINK System Non-DEHP Three Lead Extension Set. 6.0" (15 cm), 1.0 mL, Luer Activated Valve, male Luer Lock Adapter, Sterile, nonpyrogenic, Baxter Healthcare Corporation, Deerfield IL 60015. |  |  | September 04, 2013 | 2013 | Baxter Healthcare Corp. | The firm is issuing a voluntary recall due to the potential for pinholes in the packaging. Pinholes are a breach in the sterile barrier and could lead to potential contamination. | Under Investigation by firm | Baxter sent an Urgent Product Recall letter dated September 9, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Actions to be taken by Customer/User:1. Locate and remove all affected product from your facility. If you distribute these products to other facilities or departments within your institution, please forward a copy of this communication to ensure that they locate and remove affected product from potential use. The product code and lot number can be found on the shipping carton label and individual product package.2. Acknowledge your receipt of this recall notification by completing the attached Customer Reply Form (Attachment 1) and return it to Baxter by either faxing it to 224-270-5457 or scanning and emailing it to fca@baxter.com. Returning the Customer Reply Form promptly will prevent you from receiving repeat notifications.3. The affected product should be returned for credit by contacting Baxter Healthcare Center for Service. The Center for Service can be reached at 888-229-0001 between the hours of 7:00 am and 6:00 pm Central Time. Please have your Baxter eight-digit ship-to account number ready when calling.Please have the product code and lot number(s) available. |  | 3200 | 3,200 units | US Distribution including the states of OK, GA, CA, MN, TX and HI. | Terminated | April 19, 2017 | 1323 | 3200 | Undetermined
Z-0117-2014 | 66243 | No | GWM | Intracranial pressure monitoring device. | Neurology | Neurology | 510(k) | Brand Names: IntegraÂż Camino 110-4 Intracranial Pressure (ICP) Monitoring |  |  | September 10, 2013 | 2013 | Integra Limited | Integra LifeSciences has identified through an internal evaluation that in order to comply with US FDArequirements, certain CaminoÂż Intracranial Pressure Monitoring Kits require additional labelling to indicatethey are UNSAFE in an MR environment.Only the following CaminoÂż Intracranial Pressure Monitoring Kits are affected:110-4B, 110-4BC, 110-4G, 110-4H, 110-4HC, 110-4L, 110-4LC. | Under Investigation by firm | Integra sent an Urgent Medical Device Correction Field Safety Notice dted September 10, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review their inventory and determine if they have any of the affected product. If so, customers were asked to ensure that the information contained in the letter be distributed to the appropriate personnel. Complete the attached form return it by email or fax as indicated on the form. Customers with questions were instructed to contact Customer Service at 1-855-532-1723.For questions regarding this recall call 1-800-997-4868, outside of the US 888-980-7742. |  | 79887 | 79,887 | Nationwide Distribution | Terminated | June 04, 2014 | 267 | 79887 | Undetermined
Z-0112-2014 | 66144 | No | FDS | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fujinon/Fujifilm Ultrasonic Endoscope |  |  | July 31, 2013 | 2013 | Fujinon Inc | A limited number of endoscopes were shipped to FMSU-ESD with a manufacturing defect which included missing set screws used to more securely attach the Forceps Inlet (biopsy inlet port) to the endoscope. | Unknown/Undetermined by firm | The firm, FUJIFILM Medical Systems, USA, Inc. (FMSU-ESD) sent an "URGENT-MEDICAL DEVICE CORRECTION (UMDC)" and Customer Acknowledgement forms dated July 31, 2013 to their customers. The letter described the issue with the affected product; instructed customers on how to identify affected units; instructed customers to immediately cease use of the affected units; return affected units to FMSU-ESD for repairs and quality control inspection; and complete and return the attached "Customers Acknowledgement" form to FMSU-ESD, Attn: Regulatory Compliance, via fax to (973) 686-2616 or email: gwalljasper@fujifilm.com.If you have any questions regarding this notice, please do not hesitate to contact Director, Quality and Regulatory Compliance, at (973) 686-2636 or by email: gwalljasper@fujifilm.com. |  | 12 | 12 Units | Worldwide distribution: US (nationwide) including states of: NJ, PA and VA; and country of Canada. | Terminated | March 04, 2014 | 216 | 12 | Undetermined
Z-0115-2014 | 66410 | No | INI | Motorized three-wheeled vehicle. | Physical Medicine | Physical Medicine | 510(k) | Breeze C Scooter is a motorized electric scooter for elderly people. Brand name Breeze, model C. It is intended to provide increased mobility for elderly and/or disabled person and it is also an aid to independent living. |  |  | September 27, 2013 | 2013 | S.A.E. AFIKIM | Breeze C Scooter, motorized 3-wheeled electric scooter for elderly people, is recalled because it may become unresponsive during operation due to an electrical contact which may fault due to extensive wear, which potentially could result in the driver not being able to stop the scooter during operation except by turning off the ignition key. | Process control | Afikim Electric Vehicles sent an URGENT: MEDICAL DEVICE RECALL letter, dated September 27, 2013, to all affected customers and distributors. .Customers are advised to remove the ignition key and contact Afikim Electric Vehicles via telephone Monday through Friday at 1-800-809-3010 between the hours of 9AM and 5PM (EST). Consumers may also contact the company via e-mail at dov@afiscooters.com.Customers are instructed to return the Recall Response Form by fax it to 800-466-1171 or email to ofer@afiscooters.com.Consumers should report any adverse event or quality problems to FDA's MedWatch Adverse Event Reporting program either by online, fax, or regular mail. |  | 69 | 69 units | Nationwide Distribution: US including the states of California, Florida, Georgia, Illinois, Maine, Maryland, Massachusetts, Michigan, Minnesota, New Jersey, New York, Ohio, Oklahoma, Pennsylvania, Texas, Virginia, and Wisconsin. | Terminated | July 23, 2014 | 299 | 69 | Undetermined
Z-0123-2014 | 66266 | No | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | Globus Medical MicroFuse Putty, 2.5cc and 10cc. |  |  | August 28, 2013 | 2013 | Globus Medical, Inc. | The sterility of this product to a Sterility Assurance Level (SAL) of 10(-6) cannot be assured. | Process control | Globus sent an Urgent: Medical Device Recall Notification letter, dated August 28, 2013 to customers/users. The letter identified the affected units, issue, potential risk, product, and actions to be taken. Customers were requested to complete the attached response card and return of the affected product to Globus Medical. For questions and support contact Globus Medical by fax 1-610-300-1342 or email:recall@globusmedical.com. |  | 193 | 193 | USA Nationwide Distribution. | Terminated | November 14, 2014 | 443 | 193 | Undetermined
Z-0091-2014 | 66305 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Edwards Lifesciences EMBOL-X Glide Protection System; Model Numbers: EXGF24D, EXGF24LLD, EXGF24MMD, EXGF24SSD, EXGF24XLD, EXGF24XSD, EXGFXS2D, EXGF24SS2D, EXGF24MM2D, EXGF24LL2D and EXGF24XL2D. |  |  | September 16, 2013 | 2013 | Edwards Lifesciences, LLC | Edwards Lifesciences is recalling certain lots of EMBOL-X Glide Protection System due to deformation of the cannula tip. | Nonconforming Material/Component | The firm, Edwards Lifesciences, sent an "URGENT FIELD SAFETY NOTICE PRODUCT RECALL-ACTION REQUIRED" letter dated September 16, 2013 via Fed-Ex to its consignees/customers. Consignees were asked to review inventory for the cannulae and remove the cannulae from any kits received; quarantine affected material at site and return product to Edwards LIfesciences, Attn: Santosh Bhagat, 12050 Lone Peak Drive, Draper, UT 84020, Attention: RECALL, RGA #; call Edwards Customer Service at 800-424-3278 to obtain RGA# and replacement product (note: the filter is not affected and does not have to be returned with the cannulae); and complete and return the attached acknowledgment form via fax to Edwards Customer Service at 800-422-9329 within three days of receipt of this Field Safety Notice.If you have any questions, please call Edwards Customer Service at 800-424-3278 from the hours 6:00am - 4:30pm PST or contact your Edwards' sales representative concerning the recall. |  | 12034 | 12,034 | Worldwide distribution: US (nationwide) and country of: Canada. | Terminated | March 10, 2017 | 1271 | 12034 | Undetermined
Z-0113-2014 | 66383 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | COULTER LH 750 |  |  | September 24, 2013 | 2013 | Beckman Coulter Inc. | The check valves may fail on the Coulter LH 750 Hematology Analyzer (PN 6605632 and A85570) and Coulter LH 780 Hematology Analyzer (PN 723585 and A90728). A failure may result in an air or liquid leak at the connection to the Vacuum Overflow Tank waste line or at the Backwash Tank drain line. | Component change control | Beckman Coulter sent a Urgent Product Correction letter dated September 12, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. . The customer notification letter dated September 12, 2013, titled "URGENT PRODUCT CORRECTION", provided product description with codes, the issue (reason for recall), impact, action, resolution, and contact information.Customers were instructed to:If you experience a "Backwash Tank Not Full" error message, review manual mode sample results since the last acceptable QC run. System safeguards that assist with detection of erroneous results include Startup, QC, system messages and flagging (e.g., R), delta checks, flagging from limit sets (H, L, a, c), definitive messages and decision rules that would trigger further review.Avoid physical contact with a possible leak by continuing to use personalprotective equipment as recommended in the Instructions for Use and by national regulations.If you have any questions concerning this notification, please contact our Customer Support Center: Via our website, http://www.beckmancoulter.com/customersupport/supportVia phone, call 800-526-7694 in the United States and Canada.Outside US and Canada, contact your local Beckman Coulter RepresentativeWe apologize for the inconvenience this issue has caused your laboratory. |  | 1697 | 1,697 units | Worldwide Distribution - Nationwide Distribution and the country of Canada.. | Terminated | October 09, 2014 | 380 | 1697 | Undetermined
Z-1272-2013 | 64699 | No | LTT | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | Siemens Healthcare Diagnostics Inc. MIcroScan MICroSTREP plus Type 1 B1027-201 panels. |  |  | March 06, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Customer complaint investigation confirmed the issue of false skipped wells and false susceptible misreads with A. agalactiae affecting multiple antimicrobial agents on MicroScan Microbiology Systems. | Under Investigation by firm | Siemens Healthcare sent a Urgent Field Safety Notification letter dated March 6, 2013, to all affected customers of MicroScan MICroSTREP plus type 1, who process panels in a Walk Away System. The letter identified the product the problem and the action needed to be taken by the customer.As stated in our product Instructions for Use, test results should be interpreted in conjunction with the patients medical history, clinical presentation and other findings. We recommend discussing the content of this letter with your laboratory director regarding the need to review previous test results, conduct patient follow-up, and/or repeat testing for testing conducted in the last twelve months.Confirmation of receipt of this letter is required. Attached you will find a form indicating you have received and understood the information. We would greatly appreciate your assistance in notifying us that you have received the information by taking a moment to complete and return the form below.If you require further information or assistance, please contact your local Siemens Healthcare Diagnostics Representative. Please forward this notification to anyone to whom you may have distributed this product.We apologize for the inconvenience that this situation has caused. For further questions please call 1-800-677-7226 Option 1 (USA/Canada)Thank you for your patience and continued support. |  | 432160 | 21,608 boxes = 432,160 panels | US Nationwide Distribution including Puerto Rico. | Terminated | October 28, 2013 | 236 | 432160 | Undetermined
Z-0615-2014 | 65007 | No | HWE | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | Orthopedic | 510(K) Exempt | Synthes Hand Switch - Electric Pen Drive, Part Number 05.001.012 |  |  | April 15, 2013 | 2013 | Synthes USA HQ, Inc. | Synthes received a complaint from a customer in which the Hand Switch for the Electric Pen Drive was not labeled correctly. The "on" and "lock" labels were interchanged. | Process design | Synthes sent an Urgent Notice: Medical Device Recall letter dated April 15, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their inventory for the affected product, remove from inventory, and return them to Synthes immediately for replacement. Customers were instructed to call Synthes at 1-800-479-6329 to obtain a Return Authorization (RA) Number. Complete the Verification Section at the end of the letter by checking the appropriate box indicating affected product has been located. Return the Verification Form with the product to: Credit/Returns, Synthes1101 Synthes AvenueMonument, CO 80132If customers do not have the affected product they should still complete the Verifcation Section at the end of the letter by checking the appropriate box and faxing the document to 610-251-9005 or email to: FieldAction@synthes.comCustomers with questions were instructed to call 610-719-5450 or email FieldAction@synthes.com.For questions regarding this recall call 610-719-5000. |  | 290 | 290 | Worldwide Distribution - USA (nationwide) and Internationally to Mexico and Canada.. | Open, Classified |  |  | 290 | Undetermined
Z-0762-2013 | 63997 | No | MWA | Nucleic acid-based in vitro diagnostic devices for the detection of Mycobacterium tuberculosis complex in respiratory specimens. | Microbiology | Microbiology | 510(k) | COBAS TaqMan Mycobacteria tuberculosis (MTB) Test FOR IN VITRO DIAGNOSTIC USE. |  |  | November 28, 2012 | 2012 | Roche Molecular Systems, Inc. | The COBAS TaqMan MTB Test Definition Files (TDF) v3.0, for both CE-IVD and JPN-IVD tests, being used with AMPLILINK software version 3.3 series and COBAS TaqMan MTB Test Kits, both CE-IVD and JPN-IVD, have an incorrect Target Elbow Max cutoff value of 48 creating the potential for false negative results to be generated for patient samples with Target Elbows in the range of 48 to 50. | Component change control | Roche issued Product Advisory Notice (PAN) 2012-11 and Urgent Field Safety Notice PAN 2012-11 by email to Local Safety Officers (LSO) and General/Country Managers from all countries on 28-Nov-2012. The notice identifed the affected product, problem and actions to be taken. LSOs are to provide the information in the Field Safety Notice (FSN) to each customer and return a Feedback Form to Roche (email address: rmdsafeb@roche.com) once actions in FSN are executed. |  | 383 | 335 CE-IVD TDFs; 48 JPN-IVD TDFs | Worldwide Distribution - Foreign (22 EU countries - Austria, Belgium, Bulgaria, Croatia, Cypress, Czech Republic, Denmark, Germany, Finland, France, Greenland, Hungary, Iceland, Italy, Luxembourg, Netherland, Norway, Sweden, Switzerland, Spain, Turkey, UK); (41 Rest of World countries - See Attachment for full country distribution + Japan, Portugal and Mexico) | Terminated | August 06, 2015 | 981 | 383 | Undetermined
Z-0959-2013 | 64509 | apparatus, traction, non-powered | HST | Nonpowered orthopedic traction apparatus and accessories. | Orthopedic | Neurology | 510(K) Exempt | Allen Traction Boot, Modular, OT 1000 Traction Device Boot Assembly, Fracture Table Orthopedic Leg Holders, non-invasive medical devices designed to attach to orthopedic fracture tables to aid clinical staff in applying traction to the leg of a patient during surgical procedures. |  |  | February 28, 2013 | 2013 | Allen Medical Systems, Inc. | Some units may have been assembled incorrectly which could result in the boot breaking free from the mounting base and allowing the patient's leg to fall. | Process control | Allen Medical Systems sent a "Urgent Medical Device Recall" letter dated February 28, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken. Customers were instructed to: 1. Immediately remove from service your Allen Traction Boot with the serial number(s) XXXXX 2. Contact Allen Medical Customer Service at (800) 433-5774 for an RMA number and instructions to return your Allen Traction Boot for inspection and repair if necessary. There will be no cost to you. 3. Return your Allen Traction Boot to Allen Medical as soon as possible. In order to minimize the inconvenience, we will repair your Allen Traction Boot and return the device to you within 48 hours of receipt. |  | 24 | 24 devices | Worldwide Distribution- USA including the states of TX, NV, KY, GA, OH, PA, NY, AL, WA, IL and the countries of New Zealand and Japan. | Terminated | August 07, 2013 | 160 | 24 | Undetermined
Z-0945-2013 | 64401 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance 6, 10, 16, 40, 64, Big Bore and the iCT, iCT SP, Ingenuity CT and MX8000 Dual v. EXP are Computed Tomography X-Ray Systems. |  |  | December 20, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Philips Healthcare received reports from the field stating the patient table on a system had an unexpected downward movement while a patient was on the table. Philips has determined the cause was a malfunction of the vertical brake. | Component design/selection | Philips Healthcare sent a "Urgent Medical Device Correction" letter dated December 20, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were informed Philips received reports from the field stating that the screws holding the vertical brake together failed.A Field Service Engineer will contact each site to schedule a time to implement the Field Change Order. If you need any further information please contact your local Philips representative or local Philips Healthcare office. For North America and Canada Contact the Customer Care Solutions Center at ( 1-800-722-9377, option 5: Enter Site ID or follow the prompts. |  | 4797 | 4797 Units | Worldwide Distribution - Nationwide Distribution including the states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV and WY. | Terminated | April 17, 2014 | 483 | 4797 | Undetermined
Z-0805-2013 | 63244 | No | JXG | Central nervous system fluid shunt and components. | Neurology | Neurology | 510(k) | Aesculap(R) - Miethke Shunt System, FT040T - Control Reservoir Set w/Distal Catheter and FT075P 0 Ventric Cathet, w/Stylet and RT-ANG Guide |  |  | June 04, 2012 | 2012 | Aesculap, Inc. | The deflector may not move freely on the ventricular catheter. The deflector comes pre-positioned at the 5 cm mark on the catheter. In some cases the surgeon will want to adjust the deflector to a different location on the catheter before placing it in the ventricle. There is a small percentage of cases where aggressively attempting to adjust the deflector has a potential to cause the catheter | Process control | Aesculap sent an Important Recall Notification letter dated June 4, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to return the contact their local sales representative or customer service department at 1-800-282-9000 to return the affected product. Customers were also instructed to complete the Inventory Sheet attached to the letter with the quantity being returned. If customers could not locate the affected product they should provide an explanation as to why the inventory will not be returned. Customers with questions should call 610-984-9251 or 610-984-9291.For questions regarding this recall call 1-800-234-9179, ext 5067. |  | 11 | 11 | Nationwide Distribution including MI, PA and TX. | Terminated | January 08, 2014 | 583 | 11 | Undetermined
Z-1109-2013 | 64755 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | IS2000 USER MANUAL for Intuitive Surgical da Vinci S Surgical System, Model IS2000. |  |  | March 15, 2013 | 2013 | Intuitive Surgical, Inc. | Revised instructions for users of the da Vinci S, Si and Si-e systems with additional information regarding sterile adapter engagement and instrument installation. | Nonconforming Material/Component | Intuitive Surgical sent an "Urgent Medical Device Correction" letter dated April 1, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at 800-876-1310 for questions regarding this letter. |  | 1642 | 1642 customers | Worldwide Distribution.-USA (nationwide) and countries of: Austria, Belgium, Denmark, Finland, France, Germany, Ireland, Netherlands, Norway, South Korea, Sweden, Switzerland, and United Kingdom. | Terminated | September 04, 2014 | 538 | 1642 | Undetermined
Z-0808-2013 | 63556 | No | JTN | Antimicrobial susceptibility test disc. | Microbiology | Microbiology | 510(k) | Oxoid Antimicrobial Susceptibility Discs, Cefoxitin 30ug REF CT0119B, packaged with 50 discs per cartridge, 5 cartridges per box. The responsible firm name on the label is Oxoid Ltd., Basingstoke, Hants, UK. |  |  | August 21, 2012 | 2012 | Remel Inc | The firm is recalling the product because the cartridges may contain discs that have an insufficient dose of antibiotic. | Nonconforming Material/Component | The recalling firm sent a recall notification letter dated 8/17/ 2012 and 8/21/2012. |  | 23 | 23 units | US Nationwide Distribution including the states of: CA, CT, GA, IA, KS, OK, RI, TX, WA, WI, | Terminated | June 04, 2013 | 287 | 23 | Undetermined
Z-1106-2013 | 64577 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS 5,1 FS Chemistry System, Product Code 6801375 and VITROS 5,1 FS Chemistry System Refurbished, Product Code 6801890. |  |  | February 07, 2013 | 2013 | Ortho-Clinical Diagnostics | Ortho-Clinical Diagnostics, Inc. (OCD) determined that extension springs on some VITROSÂż Systems may beout of its manufacturing specifications. Extension springs are part of the slide alignment guide assembly. During an internal evaluation in its manufacturing facility, it was discovered that an extension spring had insufficient force. If the extension spring does not provide sufficient force, MicroSlides may not be properly positioned within theMicroSlide incubator which has the potential to cause imprecise results. Extension springs are just onecomponent in the slide alignment assembly. OCD's evaluation determined that the probability of an out ofspecification extension spring alone adversely affecting results is unlikely. | Nonconforming Material/Component | URGENT MEDICAL DEVICE CORRECTION letters were sent to the consignees on 2/07/13. On/about 2/07/13, OCD trained service personnel started to perform an assessment of all the VITROSÂż System(s) in End-User facilities worldwide to determine if the extension springs are out of manufacturing specifications. In addition, the slide alignment guides will also be inspected. If the extension springs are out of specification they will be replaced. Consignees will also be provided with a communication that explains the issue. |  | 2075 | VITROS 5,1 FS: USA: 670 units, Foreign: 1249 units; VITROS 5,1 FS Refurbished: USA: 52 units, Foreign: 2 units | Worldwide Distribution- USA (nationwide) including Puerto Rico and countries of: Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Republic of Panama, Singapore, Spain, and Venezuela. | Terminated | April 20, 2017 | 1533 | 3811 | Undetermined
Z-0672-2013 | 63698 | No | HMX | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | Bausch and Lomb 27G sterile cannula packed in the Bausch and Lomb Amvisc 1.2% Sodium Hyaluronate ( Model 59051, 59081, 59051L, 59081L) and Amvisc Plus 1.6% Sodium Hyaluronate (Model 60081, 60051, 60051L, 60081L) ophthalmic viscosurgical device. |  |  | November 15, 2012 | 2012 | Bausch and Lomb, Incorporated | Some disposable cannulas (with 4-digit lot numbers) provided with the Amvisc or Amvisc Plus OVD are not properly engaging to the Luer-Lock on the Amvisc and Amvisc Plus sterile glass syringes. The cannulas may leak viscoelastic material or detach from the syringe during injection. In rare incidents, detachment has resulted in serious patient injury. | Nonconforming Material/Component | The firm, Bausch &amp; Lomb, initially notified their customers via telephone November 14, 2012 and sent a follow-up letter entitled "URGENT VOLUNTARY PRODUCT RECALL AND FIELD CORRECTION" on November 15, 2012 to their customers. The letter described the product, problem and actions to be taken. The customers were instructed to quarantine the product; identify the affected cannulas; and complete and return the attached Recall and Field Correction Acknowledgment Form via Fax to (800) 822-8918 or email at: Luis.Leyva@bausch.com. Note: Based on your form selections, replacement cannulas will either be expedited directly to your facility or your Bausch &amp; Lomb Sales Representative will contact you shortly to furnish replacements and properly dispose of the cannulas with 4-digit lot numbers in your product inventory.Please feel free to contact Customer Service with any questions or concerns at (800) 338-2020 (option 2, followed again by Option 2). |  | 2515760 | 2,515,760 | Worldwide distribution: USA including DC and PR; and countries of: Canada and Mexico accounts not available at this time. | Terminated | December 04, 2013 | 384 | 2515760 | Undetermined
Z-0697-2013 | 63840 | No | MRG | Vitamin D test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | 25-Hydroxy Vitamin D EIA, Catalog #AC-57F1. |  |  | November 07, 2012 | 2012 | IDS (Immunodiagnostic Systems Ltd.) | Immunodiagnostic Systems is recalling the 25-Hydroxy Vitamin D EIA, AC-57F1 because the QC report was found to have an erroneous calibrator value assigned. The value of 24.4 ng/mL for calibrator 4 was stated on the QC report, however the correct value should have read 28.4 ng/mL. | Nonconforming Material/Component | IDS sent a recall notification letter on November 7, 2012 to all customers that received the the 25-Hydroxy Vitamin D EIA (AC-57F1). The letter informed the customers of the problem identified and the action to be taken.Customers are instructed to complete the attached Recall and Acknowledgement Form and fax back to Immunodiagnostic Systems Inc. Customers with questions are instructed to contact their local IDS office. |  | 675 | 675 units | US Nationwide Distribution | Terminated | May 31, 2013 | 205 | 675 | Undetermined
Z-0954-2013 | 64452 | No | HXX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Biomet Spine Polaris 5.5 Button Lock Screw Inserter. Catalog Number 14-500178. The Polaris Spinal System is a non-cervical, pedicle screw spinal fixation device. |  |  | February 07, 2013 | 2013 | Ebi, Llc | Increased risk of failure at the driver tip during screw insertion or removal. | Process control | Biomet Spine sent a "Urgent Medical Device Recall" notices dated February 6, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed to immediately locate and remove affected items from circulation and return affected product to EBI, LLC.Please follow the instructions on the enclosed "FAX Back Response Form, "and fax a copy of the Response Form to (973) 257-0232, prior to return of the items.Questions related to this notice should be directed to (973) 299-9300 ex. 2322.Monday through Friday, 8am to 5pm. |  | 59 | 59 units. | Nationwide Distribution including the states of AL, IN, MO, NE TN and TX. | Terminated | December 07, 2015 | 1033 | 59 | Undetermined
Z-0957-2013 | 63767 | No | MLM | Tacrolimus test system. | Clinical Chemistry | Toxicology | 510(k) | Siemens Dimension(R) Tacrolimus Flex(R) reagent cartridge (DF107). The TACR method is used as an in vitro diagnostic test. |  |  | November 27, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Siemens has confirmed customer complaints of low patient sample recovery with Dimension(R) TACR Flex(R) lot FA3085. Quality Control (QC) materials have not exhibited low recovery of the same magnitude.Internal method regression testing versus LC/MS has shown an average 25% low bias compared to LC/MS across the assay range. Individual patients, particularly samples with lower TACR values | Process control | The firm initiated their recall of this product on November 27, 2012 by sending a letter to all consignees. The letter included a recall effectiveness check form to be completed by the consignee and returned to the firm. The firm expanded their recall on January 29, 2013 by sending another letter to consignees. This letter also included a recall effectiveness check form to be completed by the consignee and returned to the firm. |  | 2976 | 2,976 | Product was distributed within the United States to AL, AZ, CA, DC, FL, LA, MD, MI, MO, NJ, NM, NC, OH, OR, PA, TN, TX, UT, VA, and WA. | Terminated | August 04, 2014 | 615 | 2976 | Undetermined
Z-0952-2013 | 64330 | No | HXI | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Extremity Medical - Disposable Orthopedic Surgical Instrument - 1.6 mm Guidewire - Reference Number 101-00006 - Non-Sterile. Contained in Extremity Medical Surgical Kits including CompressX, Hallu.X, IO Fix and XMCP. Used in the selection of screw size when used with appropriate depth gauge. |  |  | January 10, 2013 | 2013 | Extremity Medical LLC | The 1.6 mm Guidewire is undersized up to 2 mm in length. When used with the depth gauge, the measurement reading will not correspond with the desired screw length. The selected screw size based on the depth gauge reading will be about 2 mm longer than intended. | Component change control | Extremity Medical sent an "URGENT: FIELD SAFETY NOTICE" dated January 10, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. A Response Form was included for customers to complete and return. Contact Customer Service at 1 (973) 588-8980 for questions regarding this notice.On April 10, 2013, Extremity sent notifications to their customers at the surgeon level. They were asked to complete the attached product field action response form. |  | 238 | 238 units | Worldwide Distribution-USA including the states of Florida, Pennsylvania, New Hampshire, Oklahoma, California, Maryland, Arizona, Ohio, North Carolina, Texas, Virginia, Illinois, New York and Tennessee, and the countries of Australia, Belgium, Israel, Netherlands and United Kingdom. | Terminated | July 18, 2013 | 189 | 238 | Undetermined
Z-0967-2013 | 55155 | No | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | ILLICO Straight CP TI Rod, -5.5MM x 110M; REF#73601-110; Qty 1; LBL-005; REV.B; RX ONLY; Alphatec Spine, Inc. The ILLICOÂż MIS Posterior Spinal Fixation System is intended to facilitate the surgical correction of non-cervical spinal deformities by providing temporary internal fixation and stabilization during bone graft healing and/or fusion mass development. |  |  | July 06, 2011 | 2011 | Alphatec Spine, Inc. | The rods were incorrectly laser etched as ILLICOÂż Straight CP Tl Rods, 110mm, when they were in fact manufactured to 100mm in length. | Process control | Alphatec Spine sent an Urgent Medical Device Recall letter dated July 6, 2011, to all affected customers. This notification described the product, problem and actions to be taken. The consignees were instructed to review their inventory to verify the affected device; abstain from use; immediately contact Alphatec Spine Sales Administration (salesadministration@alphatecspine.com) for instructions on how to return this affected product; and complete, sign and return confirmation form to Alphatec Spine, Inc., 5818 El Camino Real, Carlsbad, CA 92008.If you have any questions, call 760-431-9286. |  | 7 | 7 | Worldwide Distribution - USA including CO and NC; Internationally to Belgium. | Terminated | March 20, 2013 | 623 | 7 | Undetermined
Z-1101-2013 | 64670 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | MRT-1504/A4 - Titan X MRI System |  |  | October 15, 2012 | 2012 | Toshiba American Medical Systems Inc | It was found that there is a problem with the LCD monitor used in the console. The LCD monitor maybecome blank during use. Specifically, the backlight of the LCD monitor may turn off while the systemis in use, causing the LCD monitor to become blank. However, the other parts of the system continue tooperate normally, and the risk of a serious health hazard is therefore very low. The LCD Moni | Nonconforming Material/Component | Toshiba sent an Urgent Medical Device Correction letter dated October 1, 2012 via USPS return receipt mail to all affected customers . The letter identified the affected product, problem and actions to be taken. The letter contains a return reply form that is to be faxed to TAMS for retention. For questions contact your local Toshiba representative at 800-521-1968. |  | 27 | 27 units | USA Nationwide Distribution including the states of : IL, NH, MA, MI, MO, PR, NV, NM, MN, ID, NY, OH, NC, AZ, MS, LA, OR, PA, CA, and CT. | Terminated | November 06, 2013 | 387 | 27 | Undetermined
Z-0699-2013 | 63958 | No | JSS | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | Becton, Dickinson and Company, BBL Crystal Enteric/Nonfermenter ID Kit, Catalog # 24500, cartons of 20 kits, In-vitro diagnostic reagents for microbiological testing. |  |  | December 13, 2012 | 2012 | Becton Dickinson &amp; Co. | In-vitro diagnostic reagents for microbiological testing may be reversed in well positions in the test kit and when utilized may lead to bacteria misidentification. | Mixed-up of materials/components | Becton Dickinson initiated the recall by letter to direct accounts on 12/13/2012. Ex-US BD sites were notified via email, distributors were notified by fax and customer letters were delivered by UPS ground. The notification advised that the recalled kit contained citrate (CIT} and malonic acid (MLO} reagents in the reversed prong/well positions. Customers were advised to discontinue use of identified lot numbers and discard any remaining packages for replacements, and also to return the enclosed reply form. The recalling firm requested distributors to submit their customer list for direct BD notification to all end users. Contact the BD Customer Service Department at 1-800-638-8663 if further assistance is required. |  | 67820 | 67,820 units | Worldwide Distribution-USA including the states of MA, TN, KY, IN, GA, NC, TX, NY, FL, AZ, WA, NJ, MI, and CA, and the countries of Uruguay, Taiwan, Singapore, Australia, South Korea, Mexico, Hong Kong, Guatemala, Colombia, Chile, Canada, India, Japan, Brazil, China and Belgium. | Terminated | June 03, 2013 | 172 | 67820 | Undetermined
Z-1116-2013 | 64841 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | The GemStar Pump is a small and lightweight, single channel infusion device designed for use in the home, in the hospital, or anywhere electronic infusion is required. The GemStar Pump can be powered by AC mains adaptor, rechargeable battery pack, docking station, or two disposable AA alkaline batteries. When powered by batteries, The GemStar Pump is ideal for ambulatory patients. |  |  | March 21, 2013 | 2013 | Hospira Inc. | During infusions of 2.0 mL/hr the motor assembly may rotate backwards capturing additional medication resulting in over-infusion. | Nonconforming Material/Component | Hospira sent an Urgent Device Field Correction letter dated March 21, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to refrain from using the affected produc and contact Hospira at 1-800-441-4100 to arrange for return of their device for repair. Customers were directed to notify their customers if they further distributed the product and confirm they have done so by return of the reply form to Stericycle.For questions regarding this recall call 224-212-4892. |  | 38801 | 1) List Number 13000: 25,119 pumps; 2) List Number 13100: 185 pumps; 3) List Number 13150: 13,497 pumps | Worldwide Distribution - USA (nationwide) and Internationally to Australia, New Zealand, China, Hong Kong, Taiwan, Japan, Korea, Philippines, Thailand, Malaysia, Canada, Austria, Bahrain, Belgium, Croatia, Denmark, Egypt, Finland, France, Germany, Gibraltar, Greece, Hungary, Iceland, Ireland, Israel, Italy, Jordan, Kuwait, Lebanon, Libya, Luxembourg, Malta, Netherlands, Norway, Oman, Portugal, Qatar, Romania, Saudi Arabia, Slovenia, South Africa, Spain, Sweden, Switzerland, Turkey, UK, United Arab Emirates, Brazil, Chile and Colombia. | Open, Classified |  |  | 107955 | Undetermined
Z-0969-2013 | 64471 | Yes | KOC | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Sterile Access &amp; Return Line Set, 36" Length, Part # UM-872-36. |  |  | June 19, 2012 | 2012 | Molded Products Inc | The recalled product leaks at the connection site. | Process control | The firm's one customer was notified of the recall by phone call on 5/8/2012. The customer returned the affected product once the replacement product was received. |  | 1200 | 1200 sets | Distributed to one customer in Michigan. | Terminated | August 14, 2013 | 421 | 1200 | Undetermined
Z-0700-2013 | 63972 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | GreenLight HPS (High Performance System) |  |  | December 11, 2012 | 2012 | AMS Innovative Center - San Jose | During electrical recertification testing it was found that the Voltage Scaling Board was non-conforming to the IEC 60601-1-2 standard. | Nonconforming Material/Component | AMS sent "NOTICE OF DEVICE FIELD CORRECTIVE ACTION" letters dated December 20, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at 1-408-456-3320 for questions about this notification. |  | 389 | 389 devices, all model numbers. | Worldwide Distribution. | Terminated | February 13, 2014 | 429 | 389 | Undetermined
Z-1097-2013 | 64678 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Iconos R200 C20, Iconos R200 T20, Iconos R100, Iconos MD, Iconos RHD, Luminos RF Classic, Luminos Fusion or Luminos Select. |  |  | January 15, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Siemens has become aware of a potential malfunction and possible hazard to patients and staff when using the Iconos R200 C20, Iconos R200 T20, Iconos R100, Iconos MD, Iconos RHD, Luminos RF Classic, Luminos Fusion or Luminos Select. This malfunction is possible in cases where the table vertical lift movement is obstructed by a solid object which could potentially result in collision. This can re | Process design | Siemens sent a Customer Safety Advisory letter dated January 15th, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The SIEMENS service representative will investigate if parts need to be exchanged in your product, an if so, which parts need to be exchanged to restore the system safety and function. Please kindly confirm receipt of the Field Safety Notice by sending back the attached form.For further questions please call (610) 448-3237. |  | 51 | 51 | Nationwide Distribution including the states of CA, FL, GA, IA, IN, KS, LA, ME, MI, MO, NC, NH, NJ, NY, NV, OH, PA, TN, TX, VA and WA. | Terminated | April 23, 2013 | 98 | 51 | Undetermined
Z-0698-2013 | 63912 | No | MRG | Vitamin D test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Abbott Laboratories, ARCHITECT 25-OH Vitamin D Reagent Kit, 03L52-25 (ROW) and 03L52-27 (U.S.) (1 x 100 tests), 03L52-35 (ROW) and 03L52-37 (U.S.) (5 x 100 tests); in vitro diagnostic. |  |  | December 12, 2012 | 2012 | Abbott Laboratories | The Architect 25-OH Vitamin D assay has the potential to generate falsely elevated or falsely depressed results when used with certain lots of Architect Reaction Vessels. | Mixed-up of materials/components | Abbott Laboratories sent a "Product Correction Immediate Action Required" letter dated December 12, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The customers were instructed to immediately discontinue use and destroy any remaining inventory of the reaction vessel lots listed; to clear the ARCHITECT RV hopper of impacted RV lots following the instructions provided, prior to continued use of the ARCHITECT 25-OH Vitamin D assay; to review the patient results obtained with the affected lots with the treating physician or heath care provider to ensure they are aware of the potential analytical issue and can determine whether re-evaluation of the patient is necessary; to forward a copy of the letter to any laboratory to whom they may have forwarded the reagent; and to retain a copy of the letter for their records and to complete and return the Customer Reply Form via fax to: 1-800-777-0051 For further questions please call your local Customer Service representative at 877-422-2688. |  | 28600 | 28,600 total | Worldwide Distribution -- USA (nationwide) and the countries of Austria, Barbados, Belgium, Canada, China, Colombia, Cyprus, Czech Republic, Denmark, France, Germany, Greece, Hungary, India, Ireland, Italy, Japan, Latvia, Libya, Malaysia, Netherlands, New Zealand, Poland, Portugal, Russia, Saudi Arabia, Serbia, Slovakia, Spain, Sweden, Switzerland, Turkey and United Kingdom. | Open, Classified |  |  | 28600 | Undetermined
Z-0964-2013 | 64456 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Howmedica Osteonics Corp. Stryker Orthopaedics Specialty Triathlon Left and Right A//P Sizer Adjustment Housing PER File K2401. For use in knee prosthesis implantation. NON-STERILE |  |  | January 24, 2013 | 2013 | Stryker Howmedica Osteonics Corp. | The standard Triathlon Femoral Stylus does not fit into the Specialty Sizers and the "R" and "L" markings on the Right Sizer are reversed. | Process design | A Notification Letter/Product Accountability Form dated 01/28/2013 was sent via FedEx on 01/28/2013 to the sole US branch receiving the affected devices. The branch was instructed to also contact the hospitals in its territory that have the affected product to arrange return of the product. Affected products to be returned to Stryker Orthopaedics, 325 Corporate Drive, Mahway, NJ 07430. |  | 8 | 4 units (total of 2 sets) | 1 Branch (NJ) with one consignee (DE). | Terminated | May 06, 2015 | 832 | 8 | Undetermined
Z-0705-2013 | 63665 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | DePuy Spine VIPER POLY SCREW 5.5 8 X70MM Tl |  |  | November 07, 2012 | 2012 | DePuy Spine, Inc. | Polyaxial screws 8x 80 and larger may disassociate when inserting in the Ilium., VIPER large diameter long screws in high-torque circumstances, the screw shank has the potential to separate from the polyaxial head | Component change control | DePuy Synthes Spine sent an Urgent Voluntary Product Recall Notification letter dated November 7, 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to examine inventory, quarantine the material and return remaining inventory. Customers were advised to complete the enclosed Business Reply Form and fax back to 508-828-3762 or email to lregan3@its.jnj.com. For questions call Linda Regan @508-977-6606. |  | 34 | 34 units | Worldwide Distribution - US (nationwide) including the states of: AZ, CA, FL, HI, IL, MD,MI, MO,NC, TN, TX and the country of Australia. | Terminated | December 24, 2013 | 412 | 169 | Undetermined
Z-0962-2013 | 64072 | No | NDN | Polymethylmethacrylate (PMMA) bone cement. | Orthopedic | Orthopedic | 510(k) | DePuy CONFIDENCE SPINAL CEMENT SYSTEM 11C Plus KIT |  |  | January 15, 2013 | 2013 | DePuy Spine, Inc. | Cement mixer used to prepare cement for surgery may not turn properly and may jam or stuck | Process control | DePuy Synthes Spine issued notification letter dated 1/11/13 issued via UPS Next Day Air Saver on 17 January 2013. Accounts requested to review inventory, quarantine any products listed in this recall notice and to complete and return the enclosed Business Reply Form. Contact the firm at 508-977-6606 for questions relating to this notice. |  | 7773 | 7773 | Worldwide Distribution-USA (nationwide) and the countries of Argentina, Armenia, Kazakhstan, Russia, Ukraine, Australia, Austria Belgium, Luxembourg , Brazil, Canada, China, Colombia, Cyprus, Jordan Lebanon, Saudi Arabia, UAE, Czech Republic, Finland ,France, Germany, Greece, Hong Kong, Hungary, India, Ireland, Israel, Italy, Latvia, Malaysia Mexico, Netherlands, New Zealand, Norway, Sweden, Poland, Portugal Singapore, Slovakia, Slovenia, South Africa, Spain, Switzerland, Turkey, and UK. | Terminated | June 30, 2014 | 531 | 16013 | Undetermined
Z-0966-2013 | 64410 | lens, condensing, diagnostic | HJL | Diagnostic condensing lens. | Ophthalmic | Ophthalmic | 510(K) Exempt | 20 Diopter Binocular Indirect Ophthalmoscope (BIO) Lens |  |  | October 25, 2012 | 2012 | Volk Optical Inc | The defect was with mis-engraving of the lot # directly on the product and its packaging. Mixed up device engraving of V20LC (lot # BB09213) with that from V60C (lot # BB09205). | Employee error | Volk sent a Voluntary Medical Device Field Correction letter dated October 25, 2012, to all affected customers. The letter identifed the product, the problem, and the action to be taken by the customer. Customers were asked to correct their records with the correct lot number and to notify their customers of the voluntary field correction and request their compliance in updating their records. For questions customers were instructed to call 440-510-0800 or 440-942-6161.For questions regarding this recall call 440-942-6161. |  | 40 | 40 devices | Worldwide Distribution - USA including the states of AZ, CA, GA, MN, NJ, and NM. Internationally to Japan. | Terminated | June 11, 2013 | 229 | 40 | Undetermined
Z-0799-2013 | 64175 | Yes | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius Liberty Cycler Set Single Conn. Ext. DL. Part Number: 050-87216. |  |  | October 12, 2012 | 2012 | Fresenius Medical Care Holdings, Inc. | There is a small risk of a leak in Cycler Sets and if the Cycler Set is punctured, the dialysate may become contaminated potentially resulting in peritonitis. | Process control | Fresenius Medical Care began contacting the affected customers by telephone and mailing an "URGENT: FMCNA LIBERTY CYCLER SET RECALL" letter and Return Receipt form dated October 31, 2012. The letters identified the specific part number and product lot and request that customers discontinue use of the lots immediately and to return unused product to FMCNA. Instructions are provided for returning the product, asking customers to isolate and discontinue use of the lot number 12BR08009 and place the product in a secure, segregated area. Consignees were instructed to contact their Fresenius Medical Care Customer Service Team at 1-800-323-5188 for instructions on how to return the recalled product. |  | 560 | 560 cases, totaling 5590 units | Nationwide Distribution. | Completed |  |  | 560 | Undetermined
Z-0801-2013 | 63584 | No | JSK | Supplement for culture media. | Microbiology | Microbiology | 510(K) Exempt | The product is a lyophilized growth supplement for use with dehydrated culture media to prepare BCYE Medium used for isolation of Legionellae from clinical and environmental samples. |  |  | October 10, 2012 | 2012 | Remel Inc | The product may contain low level microbial contamination which could result in incorrect results. | Material/Component Contamination | ThermoFisher Scientific notified customers on October 10, 2012, with a Field Safety Notice letter to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were advised use of this lot could result in incorrect results reporting. This notice needs to be passed on to all who need to be aware within your organization.The firm request the information form be faxed back confirming receipt of notice.For further information please contact Technical Support via our main switchboard on +44 (0) 1256 841144. |  | 10 | 10 units | Distribution included the states of IL, IN and MI. | Terminated | June 04, 2013 | 237 | 10 | Undetermined
Z-0749-2013 | 63761 | Yes | BTM | Manual emergency ventilator. | Anesthesiology | Anesthesiology | 510(k) | Giraffe and Panda Bag and Mask Resuscitation System |  |  | October 15, 2012 | 2012 | GE Healthcare, LLC | Oxygen and air wall inlet fittings on the back panel of the resuscitation systems were reversed during assembly. | Employee error | GE Healthcare sent an Urgent Medical Device Correction letter dated October 15, 2012, with a return receipt to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer along with the product correction. Required Action for Users: It is recommended that a qualified biomedical technician or respiratory therapist identify the affected resuscitation system and locate the air and oxygen fittings and labels at the back of the unit. With the table provided, identify the appropriate fitting for their country and location and verify the air fitting is located on the left side and the oxygen fitting is located on the right. Verify the air label is located on the left side and the oxygen label is on the right and verify that the color of the air and oxygen labels match the picture on table provided. If the unit does not meet all of the criteria, the unit will require repair. If the unit in use with patient, disconnect the wall supply gas and switch to tank supply. Once the patient no longer requires respiratory support, transfer the patient to another bed and remove the unit from service. A Field Engineer will be deployed to your site to complete the required repair. If the unit NOT in use with Patient, take the unit ouf of service. A Field Engineer will be deployed to your site to complete the required repair. For questions or concerns regarding this notification, please call the phone number 1-800-345-2700. For questions regarding this recall call 262-513-4122. |  | 30084 | 276 units total, 109 of those nationwide. | Worldwide Distribution - Worldwide Distribution - USA including AL, AZ, CA, CO, FL,GA, IA, ID, IL, KY, MA, MD, MI, MN, MO, MS, MT, NC, ND ,NE, NM, NV, NY, OH, OR, PA ,TX, VT, WA, WI and WV; and the countries of Australia, Belgium, Canada, China, France, New Zealand, Poland, Saudi Arabia, Switzerland and United Kingdom. | Terminated | January 18, 2017 | 1556 | 30084 | Undetermined
Z-1075-2013 | 56584 | No | JPD | Hemoglobin A[bdi2] assay. | Hematology | Hematology | 510(k) | G7 HPLC. |  |  | August 18, 2010 | 2010 | Tosoh Bioscience Inc | The Beta-thalessemia Mode of the G7 Analyzer-Beta Thalassemia Mode is an automated High Performance Liquid Chromatography (HPLC) system that separates and reports HbF and HbA2 quantitative percentages in whole blood. The firm received complaints that the A2 peak could not be detected with the G7 Beta-thalessemia columns. | Component change control | Please be aware that this is not a new recall. The firm has taken action. The firm stated that they have notified consignees via mail, with the list of notified customers maintained by the Regulatory Affairs/Quality Assurance department. The firm stated that they have notified distributors in the same way that they have notified other consignees, with a required confirmation to the firm of corrective actions. |  | 633 | 633 | Worldwide Distribution-USA (nationwide) including DC, Puerto Rico, Virgin Islands and the states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NV, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI and WV, and the countries of Canada, Columbia, Chile, Brazil, and Jamaica. | Terminated | May 13, 2014 | 1364 | 633 | Undetermined
Z-0947-2013 | 64178 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Advantx, Innova2000, Innova3100, Innova4100, Innova2100IQ, |  |  | January 21, 2013 | 2013 | GE Healthcare, LLC | A locking nut might loosen out of the overhead Video Monitor Suspension yoke affecting Advantx, Innova 2000, Innova 2100IQ, Innova 3100, Innova 3100IQ, Innova 4100, and Innova 4100IQ Cardiovascular X-ray imaging systems. | Nonconforming Material/Component | GE Healthcare sent "URGENT MEDICAL DEVICE CORRECTION" letters dated January 18, 2013 and February 11, 2013 to all affected customers. The letters were addressed to Hospital Administrators / Risk Managers, Mangers of Radiology / Cardiology and Radiologists / Cardiologists. The letters described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information.The second letter dated February 11, 2013 was sent out to consignees which incorporated an additional model number (2270677-2) which was initially missed. Contact the firm 800-437-1171 for questions regarding these letters. |  | 297 | 297 | Worldwide Distribution-USA (nationwide) including the states of AL, AZ, CA, CO, CT, FL, GA, IL, IN, IA, KS, KY, ME, MD, MA, MN, MS,MO, NV, NH, NJ, NM, NY, NC, OH, OK, OR, PA, SC, TN, TX, UT , VA, WA, WV, WI and the countries of VENEZUELA, UNITED KINGDOM, TURKEY, THAILAND, TAIWAN, SPAIN, SAUDI ARABIA, RUSSIAN FEDERATION, PORTUGAL, POLAND, NORWAY, NIGERIA, NICARAGUA, MOROCCO, MALTA, LEBANON, KAZAKHSTAN, JAPAN, ITALY, INDONESIA, INDIA, HUNGARY, GREECE, FRANCE, EGYPT, CHINA, CANADA, BRAZIL, BELGIUM, and NEW ZEALAND. | Terminated | January 16, 2014 | 360 | 297 | Undetermined
Z-0940-2013 | 64423 | No | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | GOPump Elastometric infusion PumpKit with/ and without accesssories all containing Dual GOPump, part number 510076: One infusion pump comprising two, 150 mL elastomeric reservoirs, each of which connects to its own catheter. Liquid flow rate is labeled at 2 mL/hr per elastomer. |  |  | February 25, 2013 | 2013 | Symbios Medical Products, LLC | Received 2 complaints out of 458 distributed pumps in this lot, where flow restrictor bead became displaced, which were reported to permit fast flow of contents. | Pending | Symbios sent an Urgent Medical Device Recall letter dated February 25, 2013, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to follow the enclosed 'Steps for Voluntary Field Action' in their entirety. These steps contain information about: I) identifying the product to be returned, 2) obtaining a Returned Goods Authorization number, and 3) returning the recalled product. If product(s) is (are) located, please call 317-450-0145 from 7:00 a.In '0 5:00 p.m. EDT to receive a Returned Goods Authorization umber ("RGA"). Please indicate the RGA number on the outside of the package. Attached is a Verification Form. Customers were instructed to complete and return the form even if they do not have any product to return. Their local sales representative can assist them in completing this form. Distributors of the affected product were instructed to contact Symbios at the number provided and forward the recall notice to all their affected accounts.For questions regarding this recall call 317-225-4447. |  | 458 | 458 | Nationwide Distribution including PA, CA, OH, AL, FL, WV, TN, NJ, NC, TX, CO, NY, TX, MO, MS, SC, and GA. | Terminated | February 13, 2014 | 353 | 458 | Undetermined
Z-1076-2013 | 64559 | No | MZW | Dental cement. | Dental | Dental | 510(k) | Kerr Maxcem Elite, self-etch/self-adhesive resin cement; Intro Kit Part No. 34417; DMC 0446 Rev 0. |  |  | February 22, 2013 | 2013 | Kerr/Pentron, DBA Kerr Corporation and Pentron Clinical | Kerr Corporation is initiating this recall because raw material used to produce the Maxcem Elite is causing the material to prematurely polymerize in the syringe. The product could be difficult to extrude or demonstrate a shorter work-time as the material may prematurely set prior to placing restorations on the tooth or core build-up. As a result, the restoration may not be seated properly. | Nonconforming Material/Component | The firm, Kerr Corporation/Sybron Dental Specialties, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated February 22, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers are instructed to determine if you have any of the affected products in your inventory; If you have the affected lots, please contact Kerr Customer Care at 1-800-537-7123 to receive an RMA number for quick return; complete and return the Maxcem Elite Acknowledgement and Recall Return Form via fax to 714-288-4609 regardless of whether you have any product in your inventory. If you are an authorized Kerr Corporation distributor, identify and contact those customers that may have been shipped the affected product within 48 hours of receipt of this field correction notification.If you have any question, contact Kerr Customer Care at 1-800-537-7123. |  | 40819 | 40,819 kit and syringe units | Worldwide distribution: USA (nationwide) and countries of: Australia, Canada, Belgium, Czech Republic, France, Great Britain, Poland, Albania, Austria, Bulgaria, Switzerland, Cyprus, Germany, Denmark, Estonia, Spain, Finland, Croatia, Hungary, Ireland, ITtaly Luxembourg, Latvia, Netherlands, Norway, Portugal, Romania, Serbia (Republic of Serbia), Sweden, Slovenia, Turkey, Ukraine, United Arab Emirates, Congo, The Democratic Republic of the (formerly Zaire), Kazakhstan, Lebanon, Russian Federation, South Africa, Mexico, Thailand, Brazil, China, Hong Kong, Israel, India, Japan, Malaysia, Panama, Singapore, Viet Nam. | Terminated | January 22, 2014 | 334 | 40819 | Undetermined
Z-0934-2013 | 64274 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Innova 2100 IQ, Innova 3100 IQ, Innova 4100 IQ. Cardiovascular X-ray imaging |  |  | February 04, 2011 | 2011 | GE Healthcare, LLC | Exposed high voltage wires inside the cabinet . On a cable located in the system cabinet, two of the shrink tubings covering the 400V live cable leave an area potentially unprotected. | Nonconforming Material/Component | No consignee letter sent out. GE Healthcare's recall strategy was to replace the impacted cable on all affected Innova 2100 IQ, Innova 3100 IQ, Innova 4100 IQ systems that were distributed. At this time, deployment of this recall is considered complete as of 9/6/2011. |  | 48 | 48 | Worldwide Distribution-USA including the states of FL, IL, IN, NJ, NC, OK, TX, and WI, and the countries of UNITED KINGDOM, TUNISIA, SERBIA, RUSSIAN FEDERATION, NORWAY, NIGERIA, LITHUANIA, LEBANON, KOREA, REPUBLIC OF KOREA, ITALY, INDONESIA, INDIA, FRANCE, FINLAND, COLOMBIA, CHINA, and BRAZIL. | Terminated | January 27, 2014 | 1088 | 48 | Undetermined
Z-0929-2013 | 64015 | No | HQL | Intraocular lens. | Ophthalmic | Ophthalmic | PMA | AMO Sensar Intraocular Lenses (IOLs), Model AR40e, |  |  | December 20, 2012 | 2012 | Abbott Medical Optics Inc (AMO) | Two production orders of these IOLs that contain mislabeled IOLs. | Process control | AMO sent an Urgent Field Safety Notice letter to all affected customers. The letter identified the affected product, problem, and actions to be taken. Customers were instructed to stop using and remove affected product from inventory. The letter advice customers to complete and return the enclosed Facsimile Form. For questions contact AMO Customer Service Representative. |  | 155 | 155 units | Worldwide Distribution - USA Nationwide including the state of OH and the countries of: Germany, Switzerland, Sweden, Italy, China, France, Singapore, UK, Croatia, Hungary, Turkey and Brazil. | Terminated | January 08, 2014 | 384 | 155 | Undetermined
Z-0941-2013 | 64159 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Large Combination Clamp/Large External Fixation System, Part Number 390.005 |  |  | January 11, 2013 | 2013 | Synthes USA HQ, Inc. | The L1 hex nuts, used to tighten the Large Combination Clamp, were oversized and the 11 mm sockets and wrenches will not fit on the hex nuts. | Nonconforming Material/Component | Synthes sent an Urgent Notice: Medical Device Recall letter dated January 11, 2013 to all affected customers. The letter identified the affected product, problem, and actions to be taken. Customers were instructed to examine their inventory for the affected product(s), remove them from inventory and return them to Synthes. For questions call 610-719-5450 or email FieldAction@synthes.com |  | 136 | 136 | USA Nationwide Distribution including the states of AL, AR, AK, CA, FL, ID, KS, LA, MD, MA, MO, NJ, NY, OH, PA, TN, TX, and VA. | Terminated | August 31, 2015 | 962 | 136 | Undetermined
Z-1130-2013 | 64478 | No | JJX | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | LifeScan brand OneTouchÂż Select Control Solution, Part Number/Model #: 02168902. |  |  | February 25, 2013 | 2013 | Lifescan Inc | When OneTouchÂż Ultra and OneTouchÂż Select test strips are tested with Ultra and Select control solution the meter may display a result outside the control solution test range printed on the strip vial label even though the system is working properly. | Nonconforming Material/Component | LifeScan sent an "URGENT MEDICAL DEVICE RECALL" letter dated February 25, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers may contact OneTouch Healthcare Professional Line at 866-461-0089 for questions about this notice. |  | 132301 | 132,301 units total | Worldwide Distribution-USA (nationwide) including DC and Puerto Rico and the states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, 10, IL, IN, KS, KY, LA, MA, MD, ME, Ml, MN, MO, MS, MT, NC, NO, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SO, TN, TX, UT, VA, VI, VT, WA, WI,WV, and WY, and the countries of Russia and Ukraine. | Terminated | October 23, 2013 | 240 | 132301 | Undetermined
Z-1128-2013 | 64698 | No | FRW | Patient scale. | General Hospital | General Hospital | 510(K) Exempt | SR Instruments FG3105 Patient Scales Used on various Joerns Healthcare Patient Lifters. Model# FG3105 Patient Scales. |  |  | December 13, 2012 | 2012 | Sr Instruments Inc | SR Instruments Inc. is implementing a field correction for the Joerns Healthcare Patient Lifts after they received two incident reports regarding two field failures of a tensile type load cell involving the FG3105 Patient Scale. The failures resulted in patient falls. | Use error | SR Instruments, Inc. sent a Notification letter dated December 13, 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. For questions call 716-693-5977. |  | 806 | 806 patient scales | USA Nationwide Distribution in the state of Texas. | Terminated | March 12, 2015 | 819 | 806 | Undetermined
Z-0663-2013 | 63231 | No | NGT | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Merit Medical's Custom Procedural Tray or Kits K12T-03162B &amp;amp; K12T-02641F.Custom Procedural Tray or Kits is a combination of legally marketed medical devices placed into one container for the convenience of the user and used within the limits of the product's intended uses. |  |  | September 14, 2012 | 2012 | Merit Medical Systems, Inc. | Custom Procedural Trays/Kits contain 0.9% Sodium Chloride Injection which were recalled because one confirmed customer report where four separate particulate issues were identified in four individual flexible containers. The particular matter may result acutely in local inflammation, phlebitis and/or generalized low-level allergic response to the particulate and/or embolize to other organs in the | Under Investigation by firm | Merit Medical sent an "URGENT PRODUCT RECALL NOTICE" dated September 14, 2012 to all affected customers. The letter identified the affected product, description of problem and recommended user actions to be taken. Consignees were requested to ensure that all appropriate personnel are notified of the affected Merit trays so that they can take immediate action. They were instructed to quarantine all affected kits, remove the bag of 0.9% Sodium Chloride and return the affected product to Merit for credit or replacement. Consignees were requested to complete and sign the Customer Response Form and immediately fax a copy of the completed Customer Response Form to: 1-804-416-1031 and mail the completed original Customer Response Form using the provided postage-paid, self-addressed envelope to Merit. Consignees were given product return instructions. |  | 422 | 422 total kits/trays | Distributed in Mississippi and New Jersey. | Terminated | January 14, 2013 | 122 | 422 | Undetermined
Z-1136-2013 | 64395 | No | ODE | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | EndoGastric Solutions EsophyX2 Plus Device with Serosa Fastener and Accessories Part number C02033-01. Label reads in part "EsophyX2 2-LINK Device Qty: (1) Fastener Delivery System ***" |  |  | February 04, 2013 | 2013 | Endogastric Solutions Inc | EsophyX2Plus device was not validated under actual or simulated use conditions prior to distributing the product to physicians. | Release of Material/Component prior to receiving test results | EndoGastric Solutions, Inc. sent an Urgent Medical Device Recall letter dated January 31, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to segregate and record lot numbers for any affected product. If no product remain in inventory customers were instructed to report such to customer service. Upon receipt of the product a credit will be issued for each device returned. For any questions customers were instructed to call 425-307-9226.For questions regarding this recall call 425-307-9200. |  | 190 | 190 | Nationwide Distribution including NC, VA, WI, TN, KY, IN, FL, MI, IA, IL, TX, LA, AR, HI, CA, NV, OR, and WA, | Terminated | July 02, 2013 | 148 | 190 | Undetermined
Z-0664-2013 | 63892 | No | DLZ | Phenobarbital test system. | Toxicology | Toxicology | 510(k) | Emit(R) 2000 Phenobarbital Assay (OSR4D229) and Syva(R) Emit(R) Phenobarbital Assay (4D019UL). |  |  | December 10, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Firm has confirmed higher imprecision with the recalled lots on Beckman Coulter AU(R) Systems and the Syva(R) Emit(R) 2000 Phenobarbital (4D019UL) lots E1, E2, and E3 as an application if the AU(R) System is used, the issue may be experienced. Firm has found that within-run imprecision of the recalled lots frequency of outliers that differ from the expected value by 20%. This is an expansion of | Nonconforming Material/Component | Siemens sent an "URGENT MEDICAL DEVICE RECALL" notice letter dated November 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immediately discontinue use and discard any remaining inventory of the affected product. Customers were asked to share the content of the letter with their laboratory director regarding the need to review previous Phenobarbital results, conduct patient follow up and/or repeat testing. For a no charge replacement of remaining inventory customers were instructed to contact Siemens Customer Service Center at 888-588-3916. For questions contact Siemens Technical Solutions Center at 800-227-8994. |  | 1979 | 1979 cartons | Distributed in the states of CA, FL, IL, KS, MA, MI, MN, MS, NM, NY, NC, OH, OR, PA, TN, TX, UT, and WA. | Terminated | August 27, 2014 | 625 | 1979 | Undetermined
Z-1132-2013 | 64669 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | TRIGEN (TM) L-P SCREW, 4.5 MM X 37.5 MM, TI-6AL-4V, QTY: (1), STERILE R, REF 71645437, Smith &amp; Nephew, Inc. Memphis, TN 38116 USA |  |  | March 06, 2013 | 2013 | Smith &amp; Nephew Inc | One batch of 4.5 mm diameter Trigen Low-Profile Bone Screws were incorrectly anodized with a gold coating (indicating a 5 mm diameter) instead of a gray coating (indicating a 4.5 mm diameter). | Process control | Smith &amp; Nephew sent an Urgent Product Recall Notification letter dated March 6, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to cease use and distribution of the affected product and quarantine for return. Non-responding consignees were notified again on March 15, 2013, by email. Surgeons were to be notified by letter beginning on March 20, 2013.For questions regarding this recall call 978-749-1330. |  | 23 | 23 units | Nationwide Distribution including CA, FL, GA, MA, MI, NC, SC, OH, and MO. | Open, Classified |  |  | 23 | Undetermined
Z-1020-2013 | 64734 | No | EAP | Intraoral source x-ray system. | Dental | Radiology | 510(k) | DeGotzen XRay unit XGenus |  |  | September 26, 2012 | 2012 | DeGotzen | It was discovered at the WEAC Labs method sample did not meet requirements for labeling and certification | Radiation Control for Health and Safety Act | DeGotzen will add new labeling or change existing labeling to correct for the 5 mentioned deficiencies. This correction of the noncompliant products must be made at no cost to the customers.For further questions please call 800-289-6367 ext. 390. |  | 159 | 159 units | US (nationwide) | Open, Classified |  |  | 159 | Undetermined
Z-0666-2013 | 64007 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Sarns flexible arterial cannula, 6.7 mm (20Fr) OD with 3/8 in. connector, suture ring, 9.5 in (24 cm) long |  |  | January 07, 2013 | 2013 | Terumo Cardiovascular Systems Corporation | Flexible Arterial Cannula and High-Flow Aortic Arch Cannula connectors were identified to contain flash on the exterior of the cannula connector that exceeded established specifications. Because the flash has the potential to detach in the operative field, TCVS has decided to remove all affected lots that were manufactured with the affected connector part. | Process control | TERUMO sent an URGENT MEDICAL DEVICE RECALL notification to all consignees on January 7, 2013. Terumo CVS is advising users to discontinue use and return all affected cannulae in inventory. Terumo CVS will issue credit for returned cannulae. Terumo CVS providied a list of marketed devices that can be used as alternatives. Contact Terumo CVS Customer Service with questions or concerns: 800- 521- 2818 Monday  Friday, 8 AM - 6 PM ESTFor questions regarding this recall call 734-741-6173. |  | 6830 | 6830 total units for all products | Nationwide and worldwide: AL , AR , AZ , CA , CO , CT , FL , GA , HI , IA , IL , IN , KS , KY , LA , MA , ME , MI, MO , NC , ND , NV , NY , OH , OK , PA , SD , TN , TX , UT , VA , WA , and WI. Bahrain, BRAZIL, CANADA, CHILE, Honduras, Singapore, Taiwan, Trinidad and Tobago, and UNITED ARAB EMIRATES (UAE). | Terminated | August 26, 2013 | 231 | 6830 | Undetermined
Z-1138-2013 | 64486 | No | MBP | Resorbable calcium salt bone void filler device. | Orthopedic | Physical Medicine | 510(k) | OsteoSelect DBM Putty |  |  | February 21, 2013 | 2013 | Bacterin International, Inc. | The firm's retesting procedure was not validated. | Process control | Bacterin sent a Voluntary Product Recall Notification letter via e-mail to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers are being asked to return any remaining product by obtaining a Return Authorization number from them. Bacterin is also asking customers to return the verification form mailed with the recall letter regardless of whether or not they had any product on hand. For questions contact Bacterin sales representative or Bacterin Customer Service. |  | 165 | 165 | USA Nationwide Distribution | Terminated | July 01, 2013 | 130 | 165 | Undetermined
Z-1126-2013 | 64712 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Accuray CyberKnife System Medical Physicist, Robotic Radiosurgery System, Medical charged-particle radiation therapy system, Model: Iris Variable Aperture Collimator P/Ns 028986 and 031777. The CyberKnife System is indicated for radiation treatment. |  |  | March 19, 2013 | 2013 | Accuray Incorporated | A defect on the master-side connector block of the pneumatic tool changer could introduce a tilt in the collimator housings, resulting in a potential shift to the central axis of the radiation beam. The issue could pose significant clinical impact to plans created using the smallest collimators. | Component design/selection | On 3/19/13, an "Urgent Device Correction Letter" was distributed to all affected users, with a description of the problem and user corrective action steps. US customers can contact 1-877- 668-8667 if questions. |  | 176 | 176 units | Worldwide distribution, including Nationwide in the US and the countries of Korea, Myanmar, Canada, Belgium, England, Czech Republic, Finland, France, Germany, Greece, India, Italy, Poland, Russia, Spain, Switzerland, Turkey, Columbia, Mexico, Japan and Saudi Arabia. | Terminated | October 23, 2013 | 218 | 176 | Undetermined
Z-0813-2013 | 63654 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, Dash 3000, Dash 4000 and Dash 5000 Physiological Patient Monitor and accessories. |  |  | June 13, 2012 | 2012 | GE Healthcare, LLC | GE has become aware of multiple issues affecting the Dash 3000/4000/5000 Patient Monitor. These issues may occur when the Ethernet controller in the XPC823 CPU on the Dash and DashPort 2motherboard functions intermittently. GE Healthcare has identified that these issues occur most often at sites with large Unity Network installations. It should be noted that the Dash 3000/4000/5000 patient m | Nonconforming Material/Component | GE Healthcare sent an Important Product Information letter dated June 13, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were told the otherboard would be replaced, on request, should the customer continue to expeience the issue with the affected product. A second letter dated August 30, 2012, was sent on September 5, 2012, informing consignees that the motherboard replacement did not fully correct the issues. Therefore, the motherboard replacement field action had been suspended. It further explained that the correction will be provided when it is available. Customers with questions were instructed to call 1-800-558-7044.For questions regarding this recall call 262-513-4122. |  | 119262 | 119,262 | Nationwide Distribution - ALL STATES IN CONTINENTAL USA INCLUDING PR and DC excluding AK, DE, ND, and VT.ALGERIA, AUSTRALIA, ANGOLA, ARGENTINA, AUSTRIA, BAHAMAS, BAHRAIN, BANGLADESH, BELGIUM, BOSNIA HERZEGOVINA, BRAZIL, BRUNEI, BULGARIA, CAMBODIA, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, CROATIA, CUBA, CYPRUS, CZECH REPUBLIC, DENMARK, ECUADOR, EGYPT, EL SALVADOR, ESTONIA, FINLAND, FRANCE, GEORGIA, GERMANY, GHANA, GREECE, GUATEMALA, HONDURAS, HONG KONG, HUNG ARIA, ICELAND, INDIA, INDONESIA, IRAQ, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KAZAHSTAN, KENYA, KOREA, KUWAIT, LATIVIA, LEBANON, LIBYA, LITHUANIA, MACEDONIA, MALAYSIA, MEXICO, MALTA, MOROCCO, NEPAL, NETHERLAND, NEW ZEALAND, NIGERIA, NORWAY, PAKISTAN, PALESTINIAN, PANAMA, PARAGUAY, PERU, PHILIPPINES, POLAND, PORTUGAL, QATAR, ROMANIA, RUSSIA, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SRI LANKA, SWEDEN, SWITZERLAND, SYRIA, TAIWAN THAILAND, TRINIDAD AND TOBAGO,K TUNISIA, TURKEY, UGANDA, UKRAINE, UNITED ARB EMIRATES , UNITED KINGDOM, URUGUAY.PUERTO RICO | Open, Classified |  |  | 119262 | Undetermined
Z-1118-2013 | 63945 | No | DLZ | Phenobarbital test system. | Toxicology | Toxicology | 510(k) | Clinical Chemistry Phenobarbital, list number 1E08 |  |  | November 16, 2012 | 2012 | Abbott Laboratories, Inc | Abbott Diagnostics is recalling Clinical Chemistry Phenobarbital reagent lots 62299UN12 and 85773UN12 due to imprecision issues which increase with the age of the reagent lot. The imprecision of the Phenobarbital assay is d 7% Total CV%; however, internal testing confirmed CVs up to approximately 12% for in-date reagent lots over five months old. Although lot 52803UN12 is now expired, it exhibit | Nonconforming Material/Component | Abbott Laboratories sent an Immediate Action Required Product Recall letter dated November 16, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use and destroy any remaining inventory of the affected product in accordance with their institution's policies and procedures, order replacement lot 27426UN12 which will expire 18 DEC 2012, and is the only lot currently available (Additional lots will be available in December and January), and complete and return the enclosed Customer Reply Form. If lots were forwarded to another laboratory, a copy of the letter should be forwarded to them. For questions regarding this recall call 1-877-4ABBOTT. For customers outside of the US they should contact their local area Customers Service. |  | 1289 | 1289 | Worldwide Distribution - USA (nationwide) and Internationally to Internationally distributed to Argentina, Australia, Austria, Bahamas, Brunei, Canada, Chile, Curacao, Denmark, France, Germany, Greece, Ireland, Italy, Lebanon, Mexico, Netherlands, Norway, Oman, Paraguay, Portugal, Puerto Rico, Russia, Saudi Arabia, Serbia, South Korea, Spain, Switzerland, Turkey, and United Kingdom. | Terminated | June 18, 2013 | 214 | 1289 | Undetermined
Z-0936-2013 | 64326 | No | OEI | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Bacterin Elutia, Hubless Coated Silicone Surgical Wound Drain 24 Fr Round Hubless, Full-Fluted Coated with Silver Sulfadiazine STERILE, Box of 10 Part #7118. Indicated for use in surgical wound drainage. |  |  | January 28, 2013 | 2013 | Bacterin International, Inc. | A portion of a silicone wound drain was found to be cracked. | Nonconforming Material/Component | Consignees were informed of the recall via e-mail beginning 1/28/13 with a request for a return response. This was followed by a "Voluntary Product Recall Information" letter dated 2/13/13 that described the product issue and provided recommended actions. The firm asked for all product to be returned to them. |  | 60 | 15 boxes (10/box) to 4 consignees in US and 81 boxes (10/box) to 1 foreign consignee | 13 US consignees located in CA, CO, MT, TX, UT, and VA. This represents 4 customers and 9 other individual contacts in USA that received marketing samples. Plus 4 Bacterin representives nationwide received samples.Foreign Consingee | Terminated | July 08, 2013 | 161 | 60 | Undetermined
Z-0937-2013 | 64458 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | Churchill Medical Systems, Inc. A Vygon Company - All sizes of CMS-CT6T225 6F Triple Lumen CT PICC with M.S.T. Components, Peripherally Inserted Central Catheter. |  |  | December 17, 2012 | 2012 | Vygon Corporation | Product was packaged with the incorrect introducer needle. The kit label indicates a "safety" needle is included; the kits were packaged with a non-safety needle. | Component change control | Churchill Medical Systems sent a letter to their consignees on December 19, 2012. |  | 210 | 210 | Product was distributed to 4 hospitals (under one hospital network) within the State of Pennsylvania | Terminated | January 14, 2014 | 393 | 210 | Undetermined
Z-0810-2013 | 63991 | No | JEA | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Mizuho Orthopedic Table, model 5855, sub assembly 5855-901 Orthopedic Table Spar. |  |  | December 17, 2012 | 2012 | Mizuho OSI | Possible instability of Orthopedic Table Top. If the epoxy bond fails, the aluminum plug may migrate out of the carbon fiber tube. If this occurs, the Orthopedic Table Spar may need to be replaced to ensure stability of the Orthopedic Table Top. | Process control | Mizuho sent a Recall Notification letter dated December 28, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The Orthopedic Table Spar will be replaced at no cost by Mizuho technicians. Customers were informed the potential failure and were advised inspection of tables prior to use. If tables are experiencing the described failure, contact Mizuho Resource Group at 1-800-777-4674, option 2. . |  | 869 | 869 devices | USA Distribution -- Worldwide Distribution including the states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WI, WV and WY., and the countries of Australia, Canada, Egypt, Finland, France, Germany, China, Israel, Italy, Japan, Netherlands, New Zealand, P.R. China, Saudi Arabia, Singapore, Slovenia, South Korea, Spain, Taiwan R.O.C., UAE and UK. | Terminated | April 15, 2016 | 1215 | 869 | Undetermined
Z-1028-2013 | 64453 | No | OOI | Instrumentation for clinical multiplex test systems. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Roche Molecular Systems cobas x 480 System 50 mL reagent reservoir, Catalog number 05232732001. Used with the cobasÂż CT/NG 4800 System, for in vitro diagnostics. |  |  | March 01, 2013 | 2013 | Roche Diagnostics Operations, Inc. | Certain lots of Reagent reservoirs, may not have a separation in the lower part of the reservoir. As a consequence, pipetting errors can occur due to uneven distribution of reagent. These pipetting errors may lead to delayed results and wasted reagent. | Process control | Roche sent a Urgent Medical Device Correction letter via UPS to all affected customers ( receipt required) The letter identified the product the problem and the action needed to be taken by the customers. Customers were instructed to visually inspect affected product for the molding defect and to verify equal distribution of reagent in both chambers of the reservoir after adding reagent. All customers are requested to fill out the FAX Back form to schedule shipment of defective part. Non-responding accounts will be monitored on an ongoing basis and follow-up attempts will be made. Questions and concerns are to be addressed to Roche Molecular Diagnostics Technical Support 24 hours/day 1-800-526-1247. |  | 189 | 189 units | Nationwide Distribution including Puerto Rico and the states of MO, NE, CO, MT, NJ, FL, WA, TX, GA, NY, KS, NH, VA, HI, PA, MI, IA, CA, NC, AZ, UT, LA, IL, TN, ME, MN, KY, CT, OR, NV, AL and NM. | Terminated | June 04, 2014 | 460 | 189 | Undetermined
Z-0688-2013 | 63771 | No | JLA | Creatine test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | NEONATAL TOTAL GALACTOSE KIT |  |  | November 26, 2012 | 2012 | Perkinelmer | Truncated V-Bottomed Microplate Wells used in assays contain defective wells. | Process control | Perkin Elmer initiated the recall on November 23, 2012. The Recall Letter will be delivered to the end customer by email facilitated through our PerkinElmer Technical Support group located in Akron Ohio. Technical Support will be asked to make sure the customers receive the Recall Letter and implement the recommended precautions:- Inspect the wells of row A and row H of the affected microplate lots for possible holes and cracks prior to use.- If a hole or crack is visually detected in a well of the microplate, the well should not be used for testing.- If a hole or crack is not visually detected in a well of the plate, it is required during use of the well, the volume of solution in the well should be visually confirmed to be consistent before and after incubation.Please contact PerkinEimer Technical Support for further information (800) 321 -9632, select prompt 2. |  | 30290010 | 3029-0010 =1562 kits (10 plats/kit) and 3029-001B=13 kits (50plates/kit) | Nationwide Distribution. | Terminated | March 07, 2014 | 466 | 30291065 | Undetermined
Z-0933-2013 | 64063 | No | JMO | Iron-binding capacity test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens Total Iron Binding Capacity (IBCT) Flex Reagent Cartridge. |  |  | December 20, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | The analytical sensitivity claim of 6 ug/dL (1.1 umol/L) in the Instructions for Use (IFU) is no longer valid. Siemens is conducting a voluntary corrective action for Dimension Total Iron Binding Capacity (IBCT) to modify the analytical sensitivity claim to 36 ug/dL (6.44 umol/L). | Process design | Siemens Healthcare sent a "Urgent Medical Device Correction" letter dated Dec 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were notified that theanalytical sensitivity of 6 ug/dL (1.1 umol/L) in the Dimension(R) IBCT FLex(R) Reagent Cartridge instructions for us (IFU) is no longer valid. Please make adjustments to middle wave products or Laboratory Information Systems ( LIS ) as appropriate for your Laboratory.Please complete the attached form and fax it to (302) 631-8467, to indicate you have received this information.If you have any questions please contact your Siemens Technical Solutions Center or your local Siemens technicial support representative at (302) 631-6311. |  | 19503 | 19, 503 cartons | Nationwide Distribution | Terminated | August 26, 2014 | 614 | 19503 | Undetermined
Z-0938-2013 | 64173 | No | MNH | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Synthes Spine 5.2mm Titanium Click ' X Pedicle Screw Dual Core 40 mm Thread Length, Part Number 487.072. |  |  | January 11, 2013 | 2013 | Synthes USA HQ, Inc. | Lot number 7542279, of the 45 mm Click 'X Pedicle Screw (Part number 487.072), was incorrectly etched as "Click 'X Pedicle Screw 05.2 L40 TAN violet" although the actual length of the screw is 45 mm. | Employee error | The firm initiated a recall of this product on January 11, 2013 by sending an Urgent Notice: Medical Device Recall letter, dated January 11, 2013, to all consignees. The letter identified the affected product and the issue, as well as requested that consignees examine their inventory for the product(s) with the identified Part and Lot Numbers, remove them from use, and return them to Synthes. Customers were to follow the steps provided in the letter regarding if they had or did not have any of the identified affected product. Questions were to be directed to 610-719-5450 or Fieldaction@synthes.com. |  | 18 | 18 | Distributed in the states of AL, CA, MS, MO, NJ, OR, and TN. | Terminated | September 10, 2015 | 972 | 18 | Undetermined
Z-0942-2013 | 64287 | No | JEY | Bone plate. | Dental | Dental | 510(k) | Synthes Titanium Midface Distractor, MR Conditional, Part number 487.982. For temporary stabilization and gradual lengthening of the cranial or midfacial bones. |  |  | January 11, 2013 | 2013 | Synthes USA HQ, Inc. | There is the potential for the anterior footplates to not fit onto the Titanium Midface Distractor Assembly as the tab of some distractors may be oversized. | Employee error | The firm initiated their recall of this product on January 11, 2013 by sending an Urgent Notice: Medical Device Recall letter to all consignees. The letter identified the affected product and the problem. Also, the firm requested consignees to examine their inventory for products(s) with the specified part number, remove them from use, and return them to Synthes. Customers were to follow the instructions provided on what to do if they have or do not have any identified affected product. Questions should be directed to 610-719-5450 or FieldAction@synthes.com. |  | 596 | 596 | Worldwide Distribution -- USA, including the states of CA, LA, MA, NY, and PA, and the countries of Canada and Switzerland. | Terminated | August 31, 2015 | 962 | 596 | Undetermined
Z-0946-2013 | 64479 | No | GWL | Physiological signal amplifier. | Neurology | Neurology | 510(k) | Nicolet EEG Wireless Amplifier 32/64 Channel. Catalog/Part Numbers : 515-015400 for the 32 channel amplifier and 515-015500 for the 64 channel amplifier. |  |  | February 27, 2013 | 2013 | Natus Neurology Incorporated | Natus Neurology Incorporated is recalling the Nicolet EEG Wireless Amplifier which may overheat and become uncomfortable to touch around the area of the on/off button. The overheating may result in discomfort ranging from too hot up to the potential for a burn. | Nonconforming Material/Component | Natus sent an "Urgent Medical Device Correction and Removal Notification' letter dated February 27, 2013 to all affected customers. The letter described the affected product, problem and actions to be taken. The letter advised customers that a Natus Field Service technician will remove and replace the affected product at no charge to the users. The letter also advised if any of the affected products have been transferred to another location to please forward a copy of the letter and to notify Natus. For questions call 800-356-0007 ext. 5129. |  | 114 | 114 (31 USA; 83 OUS) | Worldwide Distribution - USA Nationwide including the states of CA, FL, LA, MD, MS, MO, NJ, NY, PA, TX and the country of AUSTRALIA.BAHRAINBRAZILCANADACountry DENMARKEGYPTGERMANYHONG KONGINDIAINDONESIAIRELANDITALYJAPANKOREAKUWAITMALAYSIAMEXICONETHERLANDSROMANIARUSSIASINGAPORESPAINSWITZERLANDTAIWANTHAILANDTURKEYUNITED KINGDOMURUGUAYVIET NAM | Terminated | April 25, 2014 | 422 | 114 | Undetermined
Z-1123-2013 | 64700 | No | FEB | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | EVOTECH Endoscope Cleaner &amp; Reprocessor System, P/N 50004 |  |  | March 22, 2013 | 2013 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) has determined that a limited number of EVOTECH ECRs manufactured between 2011 and 2012 did not receive sufficient High Potential (HIPOT) DC voltage testing. | Nonconforming Material/Component | An Urgent Medical Device Field Correction letter, dated 3/22/13, was sent to customers who purchased the EVOTECH Endoscope Cleaner and Reprocessor (ECR) HIPOT Testing. The letter informs the customers of the problem identified and the actions to be taken. Customers are instructed on what to do if electrical shock occurs. Customers are instructed that they will be contacted in the next 2 weeks on behalf of ASP to schedule HIPOT testing. Customers with questions are instructed to contact Stericycle at (888) 965-5816. Customers are instructed to report complaints or suspected problems with their EVOTECH ECR system by contacting ASP Professional Services directly at (888) 783-7723 and select option #2. |  | 273 | 273 units | Nationwide Distribution-USA including DC and the states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, ME, MD, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, and WI. | Terminated | November 18, 2013 | 241 | 273 | Undetermined
Z-0679-2013 | 63768 | Yes | KPO | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius GranuFlo Naturalyte Dry Acid Concentrate For Bicarbonate Dialysis 2.0K 2.25 Ca 1Mg .16.5GAL |  |  | November 26, 2012 | 2012 | Fresenius Medical Care Holdings, Inc. | GranuFlo Dry Acid Concentrate may show discoloration | Nonconforming Material/Component | Fresenius Medical Care North America notified consignees by telephone on November 26, 2012 and follow-up letter on November 30, 2012 of the product recall. On January 10, 2013, the firm expanded there recall to an additional 32 lots (34 lots total). Customers were instructed to examine their inventory immediately, discontinue use, place all units in quarantine and return to Fresenius Medical Care North America. Customers were advised to contact their Fresenius Medical Care Customer Service Team for instructions on how to return the recalled product. For questions contact the GranuFio Recall Line at 1-888-803-0509 |  | 13409 | 13,409 units | US Nationwide Distribution | Open, Classified |  |  | 158107 | Undetermined
Z-0806-2013 | 63558 | No | LLH | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | Remel Xpect Clostridium difficile Toxin A/B, IVD, 20 clostridium difficile toxin A/B tests, REF 24650. |  |  | August 20, 2012 | 2012 | Remel Inc | The firm is recalling the products due to a potential for false negative test results. | Material/Component Contamination | Thermo Fisher sent a recall notification letter with details of the recall on August 20, 2012 to all affected customers. The letter identified the affected products, problem and actions to be taken. Customers were requested to follow the instructions provided and fax back the attached "Acknowledgement &amp; Receipt Form" to 1-877-428-1924 (toll free in the US) or 1-913-895-4190 (International). For questions contact the Technical Services Department at 800-255-6730 (US) or 913-888-0939 (International). |  | 260797 | 493 of lot 082505; 529 of lot 131988; and 448 of lot 131994 | Worldwide Distribution - USA Nationwide and the countries Canada and the United Kingdom | Terminated | June 04, 2013 | 288 | 260847 | Undetermined
Z-1078-2013 | 64398 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | GE Healthcare Quasar Nuclear Medicine System, Hawkeye Option and Hawkeye 4 Option for Dual-Head Variable Angle Gamma Camera. Nuclear Medicine imaging systems. Recall includes all Infinia systems, all configurations. |  |  | March 06, 2012 | 2012 | GE Healthcare, LLC | Accelerated fatigue of the lateral axis motion subsystem of the Infinia Computed Emission Tomography System could result in mechanical failure causing the detector to slide; may impact patient and operators. | Process control | GE Healthcare sent an " IMPORTANT PRODUCT INFORMATION" letter dated March 6, 2012. The letter was addressed to Managers of Nuclear Medicine, Hospital Administrators / Risk Managers and Managers of Radiology/Cardiology. The letter described the Issue, instructions, Affected Product Details, and Product Correction. For questions contact your local GE Healthcare Service Representative or call the GE Healthcare Call Center: United States or Canada: 800-437-1171 and Other Countries 262-896-2890. |  | 2180 | 2180 | Worldwide Distribution - USA (nationwide) including DC and PR and the countries of UKRAINE, TAIWAN, TURKEY, TUNISIA, THAILAND, SINGAPORE, SWEDEN, SAUDI ARABIA, RUSSIAN FEDERATION, PORTUGAL, POLAND, PAKISTAN, PHILIPPINES, PANAMA, NEW ZEALAND, NORWAY NETHERLANDS, MALAYSIA, MEXICO, MALTA, MARTINIQUE, MOROCCO, LATVIA, LITHUANIA, LEBANON, KUWAIT, KOREA, JAPAN, JORDAN, ITALY, IRAN, INDIA, ISRAEL, IRELAND, INDONESIA, HUNGARY, CORATIA, HONG KONG, GREECE, GEORGIA, UNITED KINGDOM, FRANCE, FINLAND, SPAIN, EGYPT, ECUADOR, ALGERIA, DOMINICAN REPUBLIC, DENMARK, GERMANY, CZECH REPUBLIC, CYPRUS, COLUMBIA, CHILE, SWITZERLAND, CANADA, CHINA, BRAZIL, BELGIUM, AUSTRALIA, AUSTRIA, ARGENTINA, SOUTH AFRICA, VENEZUELA, URUGUAY, ALGERIA. | Terminated | January 27, 2014 | 692 | 2180 | Undetermined
Z-0970-2013 | 64529 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | CR MD1.0 General Set, x-ray imager image plate. CR 10-X Image Plate 35CM X 43CM. An image plate, contained in an X-Ray cassette, used to capture images from X-Ray exposures. The cassette with the image plate, once exposed, is inserted in the Agfa Digitizer CR 10-X where the image plate is taken out of the cassette and then scanned so that the scanned image can be displayed for viewing. |  |  | February 07, 2013 | 2013 | AGFA Corp. | The cassettes were shipped with the wrong IP (image plate) size bar code. | Nonconforming Material/Component | On February 7, 2013, AGFA called dealers involved to alert them of the safety issue. Also, an 'URGENT FIELD SAFETY NOTICE" letter dated February 26, 2013 was sent to the customers. The letter described the affected product, problem and actions to be taken. Customers were requested to check their inventories for the affected product and block any identified product within their inventory. Customers were instructed to complete the attached Feedback form and return via FAX or email. For questions call 877-777-2432 Prompt 5. |  | 17 | 17 | USA Nationwide Distribution including the states of: AL, CA, FL, IN, MN, NE, NV, OH, SC and UT. | Terminated | June 06, 2013 | 119 | 17 | Undetermined
Z-1003-2013 | 64708 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Philips Healthcare Computed Tomography X-Ray System. |  |  | April 05, 2010 | 2010 | Philips Healthcare Inc. | Philips Healthcare discovered the customer was performing CTDI measurements for facility accreditation and received results that were 15 to 25% higher than Philips protocol parameters stated. | Radiation Control for Health and Safety Act | Philips Healthcare proposed corrective action plan (CAP) is comprised of the following elements:* Updating procedures in manufacturing to verify, either through software calibration or through gage-pin measurement, that A-plane collimation is appropriate to the design specification. It is our understanding that these procedures have been implemented.* For refurbished PET/CT and SPECT/CT systems distributed by Philips, themanufacturing process now requires that CTDI testing be conducted in addition to air testing and that each system meet updated CTDI tolerance limits to be published in the product's instructions for use.* Reassessing and correcting the correlation between air dose and CTDI specifications for all CT systems or sub-systems, resulting in updates of the CTDI specifications (tolerances) to be published in the instructions for use. It is our understanding that this reassessment and correction have been implemented, that the updated CTDI specifications will be made available as new CTDI tolerance tables with the customer information letter to be distributed immediately, and that these updated CTDI specifications will also be distributed as an addendum to the instructions for use manualsto the installed base during the next 12 months.* For all Brilliance CT, Gemini PET/CT, and Precedence SPECT/CT systems, in the installed base as well as in forward production, instructions for field service will be updated to provide clear directions on how to verify that the A-plane collimator has been appropriately adjusted to ensure expected product performance after its replacement and/or the replacement of an x-ray tube.This plan appears to adequately address the problem, and it is hereby approved. For further questions please call (978) 687-1501. |  | 6125 | 2089 US 4036 ROW | Worldwide Distribution including USA (nationwide) | Open, Classified |  |  | 6125 | Undetermined
Z-0781-2013 | 62943 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Synthes(R) Spine Vectra, Vectra-T and Vectra-One Anterior Cervical Plating for Spinal Fusion - 2.5mm Drill Bit, graduated 12mm - 26mm adjustable depth, Part Number 03.613.011. |  |  | August 06, 2012 | 2012 | Synthes USA HQ, Inc. | Although the reported incidence is low, there is the potential for the drill stop to slip on the drill bit resulting in an adverse event. | Component design/selection | Synthes sent a "NOTICE: MEDICAL DEVICE RECALL" letter dated August 6, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at 610-719-5450 for questions related to this notice. |  | 63 | 63 | Nationwide Distribution-including Puerto Rico and the states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV and WY. | Terminated | August 31, 2015 | 1120 | 111 | Undetermined
Z-1056-2013 | 64494 | No | FMH | Specimen transport and storage container. | Pathology | General Hospital | 510(K) Exempt | BD Vacutainer C&S; Transfer Straw Kit C&S; Preservative Plus Urine Tube. 4.0 ml, 13 X 75 mm STERILE REF 364953. Becton Dickinson &amp; Company, Franklin Lakes, NJ USA, |  |  | February 14, 2013 | 2013 | Becton Dickinson &amp; Company | The BD Vacutainer Urinalysis Transfer Straw Kit #364991, Lot 1335484 has preprinted polybags which are incorrectly labeled as Catalog #364953. BD advises any Customer who purchased this Kit that the Tube in this Kit is not a C&S; tube. If a Customer relies solely on the polybag labeling and uses the Tube inside the Kit for the purpose of collecting urine for culture and sensitivity, it is possible | Employee error | BD Diagnostics sent a "Product Advisory Notice" dated February 14, 2013 via UPS Second Day Air to all end user customers. The notice advises customers of the product issue, identifies affected product and provides instructions on how to handle affected product.For questions regarding this recall call 201-847-5033. |  | 236400 | 236,400 units | Nationwide Distribution including Arizona, California, Colorado, Georgia, Florida, Illinois, Iowa, Kentucky, Louisiana, Maryland, Massachussets, Michigan, Minnesota, Missouri, New Jersey, New Mexico, New York, North Carolina, Ohio, Oregon, Pennsylvania, Tennessee, Texas, Utah, Virginia, Washington DC, Washington, West Virginia, and Wisconsin. | Terminated | August 06, 2013 | 173 | 236400 | Undetermined
Z-0993-2013 | 64370 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | Animas(R) 2020 Insulin Infusion Pump |  |  | January 03, 2013 | 2013 | Animas Corporation | Animas has identified a component issue affecting a small supply of the Animas(R) 2020 insulin pumps. The component issue may trigger the pumps to sound a false alarm or warning related to one of the following: - Loss of prime; -Occlusion; or -No cartridge detected. | Nonconforming Material/Component | Animas sent an Urgent Recall letter dated December 17, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter informed customers that Animas will send a replacement pump free of charge, along with instructions for returning their pump. For questions or concerns contact at 866-796-6373. |  | 7157 | 7,157 (494 US pumps affected) | Worldwide Distribution - USA (nationwide) including the states of: AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV and Puerto Rico and the countries of: Australia, Austria, Bermuda, Canada, Czech Republic, Finland, France, Germany, Ireland, Israel, Italy, Mexico, New Zealand, Norway, Spain and the United Kingdom. | Terminated | May 14, 2015 | 861 | 7157 | Undetermined
Z-0516-2013 | 61870 | No | RFZ |  |  |  |  | Motorola Solutions PCS3000/CS3070 Laser |  |  | April 27, 2012 | 2012 | Motorola Inc | It was discovered the CS3000/CS3070 was manufactured using the laser power settings for a Class 2 laser device instead of the Class 1 power levels as specified on the product label. | Radiation Control for Health and Safety Act | Motorola Solutions, Inc. agrees to provide instructions for downloading and installing a software update from the internet, which will reduce the laser power to IEC Class 1 AELs. Purchasers will be provided an internet site and instructions for notifying Motorola Solutions, Inc. that the software installation and power reduction were successful, thereby completing the corrective action.corrections to all units is anticipated to be completed by May 2013.I. Notifications to all dealers and purchasers must be made, and the FDA NY District Office Recall Coordinator is to be included in the notifications. 2.Correction of all units to be completed by May 2013.3.All products will be brought into compliance with the laser performance standard at no cost to the purchaser. For questions regarding this recall call 631-738-3474. |  | 16961 | 16,961 units domestically | Nationwide distribution | Open, Classified |  |  | 16961 | Undetermined
Z-1022-2013 | 64501 | No | LCO | Platelet factor 4 radioimmunoassay. | Hematology | Hematology | 510(k) | Gen-Probe LIFECODES PF4 Enhanced assay, Catalog No. X-HAT45 |  |  | February 04, 2013 | 2013 | Genetic Testing Institute,inc | Lifecodes PF4 Enhanced assay, lot number 3000389, is being recalled because of observed inconsistent replicate values when testing the kit Positive Control and patient samples. The variation in reactivity can lead to invalid assays. If an invalid assay is obtained, those results should not be reported and the test should be re-run. An invalid assay would not result in significant harm to a pati | Process control | Hologic/Gen-Probe sent an Urgent Medical Device Recall letter dated February 4, 2013, via FedEx overnight deliver yto all affected customers. The letter identified affected product, stated the issue, and asked that product from the affected lot stop being used. Reported results for all samples tested using the affected lot should be reviewed. The referring physician should be notified that any negative patient result may be erroneous. Remaining inventory should be returned to Gen-Probe. A response form was also asked to be returned. For questions customers were instructed to contact Technical Support at 262-754-1000 or waukeshatechsupport@gen-probe.com.For questions regarding this recall call 262-754-1026. |  | 386 | 386 kits (370 US, 16 OUS) | Worldwide Distribution - USA (nationwide) and Internationally to Canada, China, Slovakia, and Spain. | Terminated | March 21, 2016 | 1141 | 386 | Undetermined
Z-0751-2013 | 64129 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | 45Âż Angle Driver / Z-H, Greatbatch Medical 10 000 Wahrie Drive, Clarance, NY 14031. |  |  | January 16, 2013 | 2013 | Greatbatch Medical | Greatbatch Medical has initiated a global voluntary recall and market withdrawal of its entire Small Angled Driver 45Âż product line. Greatbatch has received reports of noisy and / or rough operation when using these Small Angled Drivers. | Nonconforming Material/Component | Greatbatch sent an "URGENT MEDICAL DEVICE RECALL" letter dated January 15, 2012 to all affected customers. The letter described the Reason For Recall, Risk to Health, and Instructions to Customers. Advised consignees to examine their inventory and quarantine the product in question. If they have further distributed the product to notify their customers and to complete and return the "Recall Response Form". Contact the firm at 1-763-951-8235 for questions relating to this notice. |  | 913 | 913 | Worldwide Distribution-USA including the states of CA, IN, KS, MA, NJ, PA, TN, and TX, and the countries of AUSTRALIA, BRAZIL, CANADA, DENMARK, FRANCE, GERMANY, ITALY, JAPAN, PORTUGAL, and SWITZERLAND. | Terminated | September 16, 2014 | 608 | 913 | Undetermined
Z-0783-2013 | 63793 | No | HTT | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Anspach Effort, Inc., 98-0061 MA-15C Bearing Sleeve with Removable Bur Guard. A reusable device used with dissection tools; designed for Transphenoidal and Skull base procedures. |  |  | September 05, 2012 | 2012 | The Anspach Effort, Inc. | Contactbetween the bur and bur guard could generate metal fragments that may or may not be visible to thesurgeon and can potentially remain in the surgical site. The materials used to fabricate the bur guard are not traceable, design validation did not effectively evaluate adequate protection of adjacent tissue and inspection results for each of these bur guards were not documented. | Nonconforming Material/Component | The firm, Anspach Effort, Inc. sent an "URGENT: Medical Device Removal" letter dated November 7, 2012, to its customer. The letter identified the product, problem and actions to be taken. The customer was instructed to do the following: 1) Screen their inventory and remove and return all products immediately. 2) Complete and return the attached Customer Reply Form via fax or email to the address provided on the form. Should the customer have any questions, please contact Anspach Product Support at (800) 327-6887. |  | 8 | 8 devices | Distributed in the state of Massachusetts | Terminated | August 13, 2014 | 707 | 8 | Undetermined
Z-0753-2013 | 64004 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | TUNNELOC ROUND HEAD INTERFERENCE SCREW |  |  | January 03, 2013 | 2013 | Biomet, Inc. | Thru hole at the tip of the screw is undersized. | Reprocessing Controls | Biomet sent an URGENT MEDICAL DEVICE RECALL letter dated January 3, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed locate and discontinue use of the affected product and return it to Biomet. recall notification was sent to all consignees on 1/4/2013. The action requires the immediate location and discontinued use of the product and its return to Biomet. Questions related to this notice should be directed to (574) 372-1570, Monday through Friday, 8a.m. to 5 p.m.For questions regarding this recall call 574-267-6639, ext 1676. |  | 44 | 44 units distributed | Worldwide Distribution - USA including CA, FL, IL, AL, OH, KY, TX , and Internationally to Australia, Columbia, Mexico, Netherlands, and Chile. | Terminated | August 02, 2013 | 211 | 44 | Undetermined
Z-1019-2013 | 64623 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Archer 0.035 inch (0.89mm) Super Stiff Guidewire, Models ARCSJ200W, ARCSJ260W, and ARCDJ260W. Sterile using ethylene oxide. |  |  | March 13, 2013 | 2013 | Medtronic Inc. Cardiac Rhythm Disease Management | Medtronic has identified an issue involving specific lot numbers of the Medtronic Archer 0.035 inch (0.89mm) Super Stiff Guidewire, Models ARCSJ200W, ARCSJ260W, and ARCDJ260W, where a change in the manufacturing process has been identified as the cause for potential kinking and breaking at the proximal weld end of the outer spring coil during in-vivo use. | Component change control | Medtronic sent an Urgent Medical Device Recall letter dated March 2013 to all affected customers. The letter identifies the affected product, problem and actions to be taken. Customers were instructed to immediately quarantine all affected products and returned to Medtronic. Customers were requested to complete the attached Customer Confirmation Certificate and fax it to Medtronic to the attention of Customer Focused Quality. For question contact your local Medtronic representative or Medtronic Customer Service at 1-888-283-7868. |  | 27682 | 27,682 devices | Worldwide Distribution - USA (nationwide) including Puerto Rico and the countries of: Argentina, Armenia, Australia, Austria, Belgium, Bolivia, Canada, Chile, China, Colombia, Croatia, Czech Republic, Ecuador, El Salvador, Estonia, Finland, France, Germany, Greece, Hong Kong, Hungary, Iceland, India, Indonesia, Iran, Israel, Italy, Jordan, Korea, Kuwait, Latvia, Lebanon, Libya, Malaysia, Montenegro, Morocco, Netherlands, New Zealand, Norway, Oman, Philippines, Poland, Portugal, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom, Venezuela, Viet Nam. | Terminated | February 06, 2014 | 330 | 27682 | Undetermined
Z-1031-2013 | 64497 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Medtronic Prostiva RF Therapy Generator. |  |  | February 27, 2013 | 2013 | Medtronic Neuromodulation | Medtronic has found through device testing that if the optional footswitch is not connected to the Model 8930 RF generator during use, an electrostatic discharge (i.e. static electricity) to the footswitch connector port may toggle the RF output on or off. If this were to occur, it could result in unintentionally starting or stopping the lesion process, depending on the current status of the equi | Nonconforming Material/Component | Medtronic sent a "Urgent Medical Device Correction" letter dated January 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Medtronic has found through device testing that if the optional footswitch (Model 60883) is not connected to the Model 8930 RF generator during use, an electrostatic discharge (i.e. static electricity) to the footswitchconnector port may toggle the RF output on or off. If this were to occur, it could result in unintentionally starting or stopping the lesion process, depending on the current status of the equipment. The status messages on the display screen and associated audible alarms operate as intended, and would notify the clinician of any change in status as described in chapter 2 of the system manual.Enclosed you will find an insulating cap that should be placed over the footswitch connector port when the Model 8930 RF Generator is used without the footswitch. The footswitch connector port is located on the rearpanel of the Model 8930 RF Generator as shown below.For further questions In the United States, contact Medtronic Neuromodulation Technical Services at 1-800-707-0933 or your local product representative. Outside of the United States, contact your local product representative. |  | 848 | 848 (696 US, 152 OUS) | Nationwide Distribution including Puerto Rico and all States in continental USA except ID, MT, and NH. | Terminated | November 21, 2013 | 267 | 848 | Undetermined
Z-1017-2013 | 64620 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Archer Super Stiff Guidewire 0.035 inch (0.89mm), Model ARCSJ200W. |  |  | March 13, 2013 | 2013 | Medtronic Inc. Cardiac Rhythm Disease Management | Medtronic has identified an issue with a specific lot of the Archer 0.035 inch (0.89mm) Super Stiff Guidewire, model ARCSJ200W, where the PTFE (Teflon) coating has the potential to separate from the wire during in-vivo use. | Mixed-up of materials/components | Medtronic sent an "URGENT MEDICAL DEVICE RECALL" letter dated March 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Medtronic requested any unused guidewires from the lot number to be quarantined, and returned by contacting the customer service phone number at 1-888-283-7868. A Customer Confirmation Certificate was requested to be faxed back to Medtronic at 651-367-0612. |  | 225 | 225 units | Worldwide Distribution-USA including the states of FL, OH, and MN and the countries of Belgium, France Germany, Netherlands, Turkey, and United Kingdom. | Terminated | November 21, 2013 | 253 | 225 | Undetermined
Z-0764-2013 | 63485 | No | LZH | Infusion pump. | General Hospital | General Hospital | 510(k) | Infinity Enteralite Ambulatory Feeding Pumps, List number: INFKIT2 - Addendum to Operator's Manual |  |  | October 12, 2012 | 2012 | MOOG Medical Devices Group | The firm is adding an addendum to the operator's manual for Infinity Enteralite Ambulatory Feeding Pumps pertaining to the Infinite Dose mode and the inability of the Pump to detect air in line. | Nonconforming Material/Component | An Urgent Field Safety Notice was sent to all original consignees by certified mail beginning on Monday, October 15, 2012. Consignees were provided with an Addendum to the Operator's Manual as part of the Notice. Consignees were instructed to assure pump users are aware of the information in the addendum and to store the addendum with the Operator's Manual. Consignees were also requested to complete and return a Medical Device Recall reply form to the firm. For additional information or clarification, consignees are to contact Holly Imes, Manager, Regulatory Affairs at 801-264-1001, ext. 230 or via email at: himesstark@moog.com. |  | 92488 | 92,488 units | Worldwide Distribution and US Nationwide in the countries of (Germany, Sweeden, Croatia, Lithuania, and Dubai). | Terminated | April 10, 2014 | 545 | 92488 | Undetermined
Z-0757-2013 | 60906 | No | NTE | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Brand Name: RX Accunet Embolic Protection System |  |  | November 30, 2011 | 2011 | Abbott Vascular | The recall was initiated because Abbot Vascular has discovered that the identified lots of the RX Accunet Embolic Protection System may exhibit difficult removal of the peel away sheath due to higher than normal wall thickness. | Nonconforming Material/Component | Abbott Vascular sent an "URGENT DEVICE RECALL" letter dated November 30, 2011 to all affected customers. The letter provides the customers with an explanation of the problem identified and an action to be taken. Customers were instructed to work with their local account representative to review their inventory, complete the attached Field Action Reconciliation/ Effectiveness Check Form and return any unused identified products to Abbott Vascular. |  | 1253 | 1253 units (962 units in the US &amp; Puerto Rico) | Worldwide Distribution-USA (nationwide) and the countries of Austria, Belgium, China, Germany, Spain, France, United Kingdom, Italy, Pakistan, Poland, Russia, Slovakia, Taiwan, and South Africa. | Terminated | January 31, 2013 | 428 | 1253 | Undetermined
Z-0758-2013 | 64055 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Ethicon Endo-Surgery Inc. The ENSEALÂż G2 Tissue Sealers are indicated for bipolar coagulation and mechanical transection of tissue during laparoscopic and open procedures. |  |  | November 06, 2012 | 2012 | Ethicon Endo-Surgery Inc | Potential for self activation while the device is in use, which may result in unintended thermal damage. | Process control | Ethicon Endo-Surgery Inc. sent a "URGENT MEDICAL DEVICE RECALL" letter dated November 28, 2012, to all affected customers. The letter identified the the product the problem and the action needed to be taken by the customer.The customers were instructed to:1. Examine your inventory immediately to determine if you have affected product on hand and remove the affected product(s).2. Complete the Business Reply Form and fax it to Stericycle at 1-877-787-0367 within 3 business days, even if you do not have affected product. Please reference Event 3436. If you have product to be returned, keep a copy of this form for your records. 3.To return affected product, photocopy your completed Business Reply Form, place it in the box with your product, and affix the pre-printed UPS shipping label to return to Stericycle. Note: Returned product will be credited.Should you have any questions, please contact your local representative or call our Customer Support Center, 1-800-873-3636, Option 6. |  | 58338 | 58,338 Instruments | Worldwide Distribution - USA including the states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, VA, WA, WI, &amp; WV. and the countries Arab Emirates, Australia, Belgium, Canada, Chile, China, Colombia, Cyprus, Hong Kong, Indonesia, Israel, India, South Korea, Kuwait, Malaysia, Panama, Philippines, Puerto Rico, Russia, Saudi Arabia, Singapore, Slovenia, Thailand, Turkey &amp; Vietnam. | Terminated | April 11, 2014 | 521 | 58338 | Undetermined
Z-0745-2013 | 64081 | No | JIS | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension(R) IRON Calibrator (DC85), Siemens material number 10445010. |  |  | December 20, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has received complaints of IRON calibrations outside of acceptable calibration slope (m) values (0.97 - 1.03) when using IRON calibrator lots 2DN060 and 2FN063. Lots 2GN068 and 2JN074 may also be impacted but no complaints have been received. | Nonconforming Material/Component | Siemens sent a "CUSTOMER NOTIFICATION" letter dated December 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact Siemens Technical Solutions Center at 1-800-441-9250 or your local Siemens Healthcare Diagnostics representative for assistance with this issue. |  | 5999 | 5999 | Nationwide Distribution. | Terminated | January 08, 2014 | 384 | 5999 | Undetermined
Z-1023-2013 | 64579 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | 20-Pole Eco Cable, Catalog #EM-5050-060 |  |  | March 05, 2013 | 2013 | Biosense Webster, Inc. | Biosense is recalling the 20-Pole Eco Cable because it was incorrectly calibrated. When a catheter is connected to an affected to an affected 20-Pole Eco Cable, the catheter will not be recognized by the impacted 54 cable serial numbers within the field. | Employee error | Biosense Webster sent an Urgent Field Notice Medical Device Letter dated March 5, 2013 to all affected customers. The letter identified the affected product, problem and the actions to be taken. Customers with questions are instructed to contact Biosense Webster sales representatives or call (866) 473-7823, Mon-Friday from 7amto 8pm EST. |  | 54 | 54 units | Worldwide Distribution - USA (Nationwide) including the states of: DE, NL, CH, FR, GB, DK. | Terminated | October 17, 2013 | 226 | 54 | Undetermined
Z-0774-2013 | 63531 | No | MRA |  |  | Orthopedic | PMA | Trilogy Acetabular AB Shells with Cluster Holes. |  |  | September 14, 2012 | 2012 | Zimmer, Inc. | Some units of Trilogy shells manufactured since March 2009 may be below the lower range of the pore size specification (as low as 28% vs. 30% porosity specification). On 10/31/2012, the firm expanded the recall action to include lots of Uni-Hole Trilogy Acetabular System shells. | Process control | Zimmer sent an "URGENT MEDICAL DEVICE NOTIFICATION" letter dated September 20, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers.The firm expanded the scope of the recall in January 2013 to include devices that use a secondary mean of fixation. New recall letters will be submitted to distributors and surgeons. |  | 275952 | 275,952 units | Worldwide Distribution-USA (nationwide) and the countries of Denmark, Germany, Greece, Hungary, Ireland, Italy, Netherlands, Spain, Sweden, UK, Poland, France, Australia, Chile, Canada, India, Japan, New Zealand, Peru, South Africa, Colombia, Ukraine, and UAE. | Terminated | June 20, 2014 | 644 | 275952 | Undetermined
Z-1024-2013 | 64448 | No | OPP | Infusion pump. | General Hospital | General Hospital | 510(k) | Z-800 Volumetric Infusion System. |  |  | June 13, 2011 | 2011 | Zyno Medical LLC | Update Instruction For Use for Preventative Maintenance (PM) Schedule and addition of Odometer software to monitor accumulative fluid volume infused by the pump | Process design | Zyno Medical initiated recall 6/13/11 to accounts advising of a Preventative Maintenance (PM) Schedule and a software upgrade Odometer to monitor counts of accumulative fluid volume infused by the unit. On 3/1/13, firm issued correction notice to users that have not had the software implemented on their unit. Units will be serviced at their scheduled PM. The notice identified the product, problem, and actions to be taken by the customers. |  | 2589 | 2589 units | Nationwide Distribution. | Completed |  |  | 2589 | Undetermined
Z-0759-2013 | 61879 | No | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | Rectal Retractors, stainless steel, item # AL07289000, a component of Varian Medical Ring and Tandem Applicator Sets; Intended for use in Brachytherapy. |  |  | April 12, 2012 | 2012 | Varian Medical Systems, Inc. | A medical device used in brachytherapy may become contaminated and can cause infections, toxicity or allergic reactions. | Component design/selection | Varian Medical Systems an an Urgent Medical Device Correction letters dated April 11, 2012 and May 8, 2012 to all affected customers. The April notification advised of an anomaly with titanium rectal retractors, and the correction was expanded in May to include stainless steel retractors. In both notifications, users were instructed to discontinue use and quarantine their cleaned and sterilized devices for collection by Varian field service engineers. Users were requested to return an enclosed reply form to respond with number of units in inventory. |  | 344 | 344 pieces | US Nationwide Distribution | Terminated | April 10, 2014 | 728 | 496 | Undetermined
Z-0785-2013 | 64044 | No | IOR | Mechanical wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Solara 3G Custom Manual Wheelchair |  |  | September 10, 2012 | 2012 | Invacare Corporation | The potential exists for the wheel to rotate freely despite engagement of hub brake. | Process control | Invacare sent an URGENT: RECALL INFORMATION Letter dated September 10, 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to 1) Complete the Provider Response Card indicating your receipt and understanding of these instructions. 2) Contact your customers to inform them of the recall. It is recommended that you provide them with the enclosed Customer Letter. Set up an appointment to correct the customer's wheelchair as soon as possible. 3) Please examine the wheelchair to confirm it has one of the serial numbers listed above. The serial number can be found on the front lower crossbar located on the right while sitting in the chair. 4) Using the enclosed retrofit kit, make the necessary correction. Installation instructions can be found on page 109 of the Solara service manual, which is enclosed.5) Document the correction of the wheelchair on the enclosed tracking sheet.6) Using the enclosed pre-paid UPS mailing label and original box, return the replaced parts to the address listed on the label. 7) Fax or email the Provider Response Form and tracking sheet to 330-329-3458 or recall@invacare.com. For questions please contact Invacare Customer Service at 800-333-6900. |  | 3 | 3 units | US Nationwide Distribution including the states of CA &amp; MT. | Terminated | May 13, 2014 | 610 | 3 | Undetermined
Z-1014-2013 | 64543 | No | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Smiths Medical, BCIÂż Remote Alarm Cable (Normally Open &amp; Normally Closed), Reorder Numbers: WW3398NO &amp; WW3398NC. For use with respiration carbon dioxide measurement equipment. Product Usage: The BCIÂż Remote Alarm Cable is designed for use with the 9004 CapnocheckÂż Capnography (Plus or Sleep) system to transfer alarm signals from the CapnocheckÂż Capnography system to a remote alarm system. The BCIÂż Remote Alarm Cable and the 9004 CapnocheckÂż Capnography systems are intended for use in all critical environments including ventilatory applications, patient transport, and anesthesia. |  |  | February 06, 2013 | 2013 | Smiths Medical ASD, Inc. | In rare instances, the BCIÂż Remote Alarm Cables (BCIÂż Cable) are not transferring alarms when used with some remote nurse alarm systems. | Nonconforming Material/Component | Smiths Medical sent an "Urgent Medical Device Field Safety Notice-Recall" letter dated February 96, 2013 via UPS to all affected customers. The letter described the affected product, problem and actions to be taken. Customers were instructed to inspect their inventory and remove all affected devices from use, complete and return the attached "Confirmation Form", by fax to 1-800-237-8033 or email to BCIRemote.AlarmCable@smiths-medical.com. The letter stated that a Customer representative will contact customers with instructions for returning affected products for replacement or credit. For questions contact Smiths Medical's Customer Service Department at 10-800-258-5361. |  | 39 | 39 | USA Nationwide Distribution - including the states of: CO, IL, IN, MA, SC and WI. | Terminated | July 12, 2013 | 156 | 39 | Undetermined
Z-0777-2013 | 63651 | No | INM | Mechanical chair. | Physical Medicine | Physical Medicine | 510(K) Exempt | PREMIER LIFECARE RECLINER 5400 |  |  | September 24, 2012 | 2012 | Winco Mfg., LLC | The performance of the chemicals in the foam and vinyl are not consistent enough to meet the TB 133 fire protection mandate required by California and 4 other states. | Component design/selection | Winco sent a Chair Field Correction Notification dated September 24, 2012, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. The consignees were instructed to forward the "Field Correction Form - REPAIR" form back to Winco and the firm would send a Winco Representative to the consignees' facility to correct the chairs in the field.Customers with questions were instructed to contact Winco at 800-237-3377 or email customerservice@wincomfg.com.For questions regarding this recall call 352-854-2929, ext 110. |  | 95 | 95 units (all models) | Nationwide Distribution including CA, FL, IL, MA, MD, MN, NM, NY, OH, OK, and TX. | Terminated | May 23, 2013 | 241 | 95 | Undetermined
Z-0765-2013 | 64163 | No | EKS | Dental hand instrument. | Dental | Dental | 510(K) Exempt | Sybron Endo Lightspeed: LSX Files, 25MM, Part Numbers: 835-2525; LSX Files, 25MM Assorted Kit, Part Numbers: 835-2455; LSX Files, 30MM, Part Numbers: 835-2031. |  |  | December 12, 2012 | 2012 | Ormco/Sybronendo | The length markings on the Lightspeed LSX files were wrong. | Process control | Sybron Endo sent a Urgent: Medical Device Recall letter dated December 21, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.It has come to our attention that the length rings laser-etched on several lots of the Lightspeed LSX Files are inaccurate. We are asking for the product to be returned since use of the file could be a safety risk due to misadjustment of the file length. Please contact SybronEndo Customer Care at 1-800-346-3636 directly to handle the arrangements of a quick return and replacement. |  | 740 | 740 units | Worldwide Distribution - USA and the country of Canada. | Terminated | November 07, 2013 | 330 | 740 | Undetermined
Z-1048-2013 | 64570 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens Dimension(R) EXL(TM) 200 (Siemens Material Number 10636929). |  |  | February 27, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Through investigation of customer complaints, Siemens Healthcare Diagnostics has confirmed an increase in System Check failures due to high System Check LOCI results. Upon further investigation, Siemens has determined that many of these errors are due to the current limit set for System Check LOCI results. The errors have no impact on results, but they may cause customer inconvenience. Siemens h | Under Investigation by firm | Siemens initiated their recall of this product on February 27, 2013 by sending an Important Customer Notification letter to consignees by FedEx. The letter identified the product, problem, and actions to be taken by the customers. |  | 2207 | 2207 total | Worldwide Distribution. | Terminated | January 14, 2014 | 321 | 2207 | Undetermined
Z-0768-2013 | 64057 | No | CAW | Portable oxygen generator. | Anesthesiology | Anesthesiology | 510(k) | Li-Ion battery used in the Oxus Portable Oxygen Concentrator (POC), RS-00400. |  |  | January 14, 2013 | 2013 | Oxus America, Inc. | Due to past battery inventory storage conditions and deficiencies in the maintenance procedure, it was possible that some batteries were over-discharged to point of damaging the cells. | Nonconforming Material/Component | Oxus, Inc. sent an "URGENT: MEDICAL DEVICE RECALL" letter dated January 11, 2013 to all affected customers.The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at 1-888-475-1568 for information related to this recall. |  | 1724 | 1724 total units | Worldwide Distribution-USA including the states of New York, Michigan, Colorado, Texas, Iowa, Florida, Georgia, Ohio, Wisconsin, Oregon, Alabama, Illinois, Utah, Oklahoma, West Virginia and California and the countries of Canada, European Union, South Africa, Korea and Brazil. | Terminated | July 17, 2013 | 184 | 1724 | Undetermined
Z-0756-2013 | 64010 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | BD Vacutainer Buffered Sodium Citrate (9NC) Plus Blood Collection Tubes; Sterile; IVD; REF 363083; 13 X 75 mm, 2.7 mL |  |  | November 12, 2012 | 2012 | Becton Dickinson &amp; Company | BD is conducting a recall of the BD Vacutainer 2.7mL Plus Citrate Blood Collection Tube due to extended Partial Thromboplastin Time (aPTT) test results reported in two customer complaints. | Material/Component Contamination | BD sent "ATTENTION: PRODUCT RECALL" letters dated November 12, 2012 via UPS to distributors and customers instructing them to examine their inventory immediately and discontinue the shipment of the affected product. All customers were instructed to return all products from the lot affected, by following the instructions on the enclosed packing slip,and BD will send them replacement products at no charge. Contact the firm at 10201-847-4267 for assistance. |  | 2944700 | 2,944,700 units | Worldwide Distribution-USA (nationwide) and the countries of Canada, India, Japan, Nepal, Malaysia, Singapore, Brazil, Costa Rica, and Virgin Islands. | Terminated | September 06, 2013 | 298 | 2944700 | Undetermined
Z-0766-2013 | 64099 | No | ILQ | Powered communication system. | Physical Medicine | Physical Medicine | 510(K) Exempt | P10 Tobii battery box (product number 510410) |  |  | January 14, 2013 | 2013 | Tobii Assistive Technology, Inc. | Tobii ATI customers P10 external battery pack, which contains sealed lead-acid batteries, malfunctioned while recharging and produced smoke with a chemical smell. | Nonconforming Material/Component | The recalling firm, Tobii Assistive Technology, Inc. sent a notification letter to all customers who received the recalled product via certified mail. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immediatley discontinue the use of the P10 battery pack and provded with instructions to receive a complimentary replacement solution. The letter instructed customers to dispose of their P10 External Battery Pack in a manner stipulated with local laws. An acknowledgment form is included to be completed by customers. For questions contact Tobii ATI at 1-800-793-9227 or P10batterypack@tobiiati.com. |  | 483 | 483 units | US Nationwide Distribution | Terminated |  |  | 483 | Undetermined
Z-0747-2013 | 63918 | No | JJY | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Alere Cholestech LDX Calibration Verification, 4 vial set, Catalog Number: 88770 |  |  | October 22, 2012 | 2012 | Alere San Diego, Inc. | Out of range (increased) HDL Cholesterol results. | Component design/selection | A recall letter was sent on 10/22/12 to all customers who purchased the Alere Cholestech LDX Mutianalyte Control, 2 vial set, Catalog Number 88769, Lot #C2510 &amp; Lot #C2510A. The letter informed the customers that these lots of controls may yield results which are out of range high for the Level 2 control when tested with the Alere Cholestech LDX HDL test. Customers are instructed of the actions to be taken.On 11/16/12, the recall expanded to include the Alere Cholestech LDX Calibration Verification, 4 vial set, Catalog 88770, Lot#271097. The problem identified is that these lots of calibration verification materials may yield results which are out of range high when tested with the Alere Cholestech LDX HDL test. Customers are instruted of the actions to be taken.On 3/6/13, the recall expanded to include the Alere Cholestech LDX Calibration Verification, 4 Vial set, Catalog # 88770, Lot#298148 and 301886. The recall expanded to include the Alere Cholestech LDX Multianalyte Control, 2 vial Set, Catalog # 88769, Lots#C2572, C2572A, C2593 and C2593A.Customers are instructed to complete and fax the enclosed verification form within 10 days.Customers with questions are instructed to contact Alere San Diego, 9975 Summers Ridge Rd, San Diego, CA 92121, (877) 308-8289, fax (858) 805-8457, and responses.ts@alere.com. |  | 614 | 614 kits | Worldwide Distribution - USA (nationwide) Puerto Rico and South Korea. | Terminated | September 26, 2013 | 339 | 11780 | Undetermined
Z-1059-2013 | 64537 | No | KHE | Occult blood test. | Hematology | Hematology | 510(k) | Test Strips packaged in the OC-Light Manual iFOBT Kit, Catalog No. FOB50. |  |  | February 12, 2013 | 2013 | Polymedco, Inc | The test strips in the affected product lots of the OC Light immunological fecal occult blood test are being recalled because they may not be stable through their labeled shelf-life. There may be a possibility of obtaining false negative results when using the affected lots. | Process control | Polymedco initiated phone calls to clients on 2/12/13. When the contact was made with the client, an URGENT DEVICE RECALL letter and Fax Form were sent by e-mail/fax. Mass mailing of the URGENT DEVICE RECALL letter dated February 12, 2013 and Fax Form was done by First Class USPS on 2/14/13-2/18/13. The recall affects the test strips only. For questions, please contact Polymedco Technical Support at 1-800-431-2123 ext 285. |  | 355350 | 355,350 test strips | Worldwide Distribution-USA (nationwide) including Puerto Rico, and the countries of Canada and Guam. | Terminated | February 10, 2016 | 1093 | 584950 | Undetermined
Z-0738-2013 | 64054 | No | MDB | Microbial growth monitor. | Microbiology | Microbiology | 510(k) | BacT/ALERT SA Blood Culture Bottle. |  |  | December 25, 2012 | 2012 | bioMerieux, Inc. | BacT/ALERT SA Culture Bottle lot 1031866 may contain bottles with uncured sensors. | Nonconforming Material/Component | bioMerieux sent an "URGENT PRODUCT REMOVAL NOTICE" and Acknowledgement Form on January 4, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers.Contact your local bioMerieux Customer Service Representative for additional assistance. |  | 1690 | 1690 cases | Worldwide Distribution.USA including the states of AR, AZ, CA, CO, DE, FL, GA, IL, IN, KY, LA, MD, MI, MN, MO, NC, NJ, NY, OH, OR, PA, RI, SC, SD, TN, TX, VA, WI, and WV, and the countries of China, Curacao and Japan. | Terminated | June 03, 2014 | 525 | 1690 | Undetermined
Z-1068-2013 | 64673 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Toshiba TSX-30IC AQUlLlON ONE System, Computed Tomography X-Ray. This system supports Whole body scanning. |  |  | November 14, 2012 | 2012 | Toshiba American Medical Systems Inc | Due to a problem with the control software of the x-ray high-voltage generator, if scanning is performed with a tube current less than 30mA, the control software may detect an error and perform error processing even though the x-ray output is normal. | Process design | Toshiba American Medical Systems Inc. sent a Urgent Medical Device Correction letter dated November 14, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customer was informed a New software to prevent the occurrence of this issue will be installed on your system. Toshiba Service Representative will contact you for an appointment to install the New Software on your Aquilion One System.Please fax the return reply form to (877) 349-3054.We thank you for your urgent attention to this matter. If you have any further questions regarding this letter please call (800) 521-1968. |  | 1 | 1 unit | Distributed in the state of MD. | Terminated | April 08, 2013 | 145 | 1 | Undetermined
Z-1085-2013 | 64754 | No | OYC |  |  | Clinical Chemistry | PMA | Medtronic MiniMed Paradigm Insulin Infusion Pumps |  |  | March 22, 2013 | 2013 | Medtronic MiniMed | Medtronic is recalling the Medtronic Paradigm Insulin Pump because the pump's drive support cap may become detached from the pump case and protrude from the lower right side of the pump. When the pump is expose to water it may result in damage to the pump's internal electronics. This moisture damage can prevent the pump's buttons from working properly or can cause the pump to alarm. | Nonconforming Material/Component | Medtronic sent an Urgent Medical Device Safety Notification letters to all affected customers. The recall letter informs the customers of the problem identified and the actions to be taken. The letter informs the healthcare professionals that Medtronic will inform their patients about the potential safety issues. Customers are instructed to find online information at www.medtronicdiabetes.com/support/product-updates. Customers are instructed to call Medtronic at 1.888.204.7616 (option 1 for loose drive support cap related questions and option 2 for water damage related questions). Distributors are instructed to provide Medtronic with an Excel spreadsheet list of all end users to whom they have shipped any of the listed serialized products. |  | 428000 | 428,000 units | Worldwide Distribution - USA (nationwide) the countries of: AE, AR, AT, AU, BA, BH, BM, BR, BS, CA, CH, CL, CN, CO, CR, CY, CZ, DE, DK, DO, DZ, EC, EE, EG, ES, FI, FR, GB, GR, GT, HK, HR, HU, IE, IL, IN, IQ, IT, JO, JP, KR, KW, LB, LC, LV, MK, MX, MY, NL, NO, NZ, OM, PA, PE, PL, PR, PY, QA, RO, SA, SE, SG, SI, SK, SV, SY, TF, TH, TR, TT, TW, UA, UY, VE, ZA. | Terminated | November 04, 2013 | 227 | 428000 | Undetermined
Z-0955-2013 | 64511 | No | DTL | Cardiopulmonary bypass adaptor, stopcock, manifold, or fitting. | Cardiovascular | Cardiovascular | 510(K) Exempt | Vascular Solutions, Inc., Guardian II and Guardian II NC Hemostasis Valve, used in cardiopulmonary bypass. |  |  | February 27, 2013 | 2013 | Vascular Solutions, Inc. | Risk that air may be introduced into the device which may lead to an air embolism. | Vendor change control | Consignees were sent a Vascular Solutions "Urgent Medical Device Recall" letter dated February 28, 2013. The letter described the problem and the product involved in the recall. In addition, consignees were advised to identify the location of and immediately remove the recalled product from their inventory and secure it. Consignees were requested to complete and return the VSI Account Inventory Form. The letter also stated that VSI Customer Service would contact customers within 1-2 business days after receiving the completed VSI Account Inventory Form. All returned affected product will be replaced upon receipt. Customers are to contact their local Sales Representative if they have any questions or concerns regarding this recall. |  | 79869 | 79,869 | Worldwide Distribution -- USA, including the states of AL, AZ, AR, CA, CO, CT, DE, DC, FL, GA, HI, IL IN, IA, KS, KY, LA, MD, MA, MI, MN, MS, MO, NE, NV, NJ, NY, NC, ND, OH, OK, OR, PA, SC, SD, TN, TX, UT , VT, VA, WA, WV, and WI, and countries of GREECE, DENMARK, AUSTRIA, UNITED KINGDOM, SWITZERLAND, IRELAND, INDIA, TURKEY, POLAND, HUNGARY, ITALY, EGYPT, KOREA, PORTUGAL HONG KONG, GERMANY, SPAIN, MEXICO, SWEDEN, NORWAY SLOVENIA, NETHERLANDS, and BELGIUM. | Terminated | August 05, 2013 | 159 | 79869 | Undetermined
Z-0740-2013 | 63827 | No | MGB |  |  | Cardiovascular | PMA | Cordis, a Johnson and Johnson Company, ExoSeal(TM) 6F Vascular Closure Device, cat. No EX500, EX600, and EX700. Used for femoral artery puncture site closure, reducing times to hemostasis and ambulation in patients who have undergone diagnostic or interventional procedures. |  |  | October 29, 2012 | 2012 | Cordis Corporation | Affected lots of Cordis ExoSeal (TM) VCD have been sterilized using a process that did not achieve the sterility assurance level that is required by the standards of this class of product. Model Numbers EX500, EX600, and EX700. | Material/Component Contamination | Cordis Corporation sent an "Urgent Medical Device Recall (Removal ) letter dated October 29, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed to immediately identify and set aside all product listed in a manner that ensures the affected product will not be used. A product recall response form was attached for customers to review, complete, sign and return the enclosed acknowledgement form directly to Cordis.For further questions please call (786) 313-2000. |  | 98675 | 98,675 units distributed in US. | US (nationwide) including the states of AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV and Hawaii. | Terminated | November 06, 2013 | 373 | 98675 | Undetermined
Z-0793-2013 | 64132 | No | DHB | Radioallergosorbent (RAST) immunological test system. | Immunology | Immunology | 510(k) | OPTIGEN Food 30; Catalog number: 85035. For allergen testing. |  |  | December 21, 2012 | 2012 | Hitachi Chemical Diagnostics Inc | There is no evidence that a visual inspection was performed on one tray of coverslips. Possibility of a low fill volume which could lead to a false negative results. | Process control | Hitachi sent an Urgent Product Recall Notification letter dated December 21, 2012, to the affected customer. The letter identified the product, the problem, and the action to be taken by the customer. The customer was instructed to discontinue distribution and destroy any remaining inventory of the affected product at their location per their local waste ordinances, determine if any of their customers are using or have inventory of the affected product, rquest that their customers review all patient results reported with the use of the affectted product. The customer was instructed to tell their customers to discontinue use and destroy any remaining inventory per their local waste ordinances. The customer was instructed to retain a copy of the notification letter for their records. The customer was also instructed to send confirmation of the destruction within 5 days Customers were instructed to complete and return the attached Certificate of Destruction. For questions customers were instructed to contact their local area Customer Service.For questions regarding this recall 650-961-5501. |  | 20 | 20 kits | Worldwide Distribution - France only | Terminated | May 16, 2013 | 146 | 20 | Undetermined
Z-1079-2013 | 64418 | bath, incubators/water, all | JTQ | Microbiological incubator. | Microbiology | Microbiology | 510(K) Exempt | CO2 Incubators (General Purpose Incubator), model numbers 370, 371, 380, 381, 3307, 3308, 3310, 3311, 3540, 3541, 3542, 3543, 3584, 3585, 3598, and 3599. |  |  | November 29, 2012 | 2012 | Thermo Fisher Scientific (Asheville) LLC | Incubators manufactured prior to mid 2009 have coated metal hinges that have reportedly failed in the field. | Component design/selection | Thermo Fisher sent an Urgent Field Correction Notice letter dated November 19, 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to locate all affected products and determine if it is within the recall. If so, they are to return the Acknowledgement Form in order to get replacement hinges and instructions. Upon receipt of the new hinges they are to replace them according to the instructions. |  | 15567 | 15,567 incubators | Worldwide Distribution - USA (nationwide) including the states of: AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, SN, TN, TX, UT, VA, VT, WA, WI, WV &amp; WY. and the countries: AFGHANISTAN, ARGENTINA, AUSTRALIA, AUSTRIA, BELARUS, BELGIUM, BRAZIL, CANADA, CHILE, COSTA RICA, DENMARK, EGYPT, ENGLAND, FINLAND, FRANCE, GERMANY, GREECE, HONG KONG, ICELAND, INDIA, INDONESIA, IRELAND, ISRAEL, ITALY, JAMAICA, JAPAN, JORDAN, KOREA, LEBANON, MALAYSIA, MEXICO, NEPAL, NETHERLANDS, NEW ZEALAND, NORWAY, PALESTINE, PEOPLES REPUBLIC OF CHINA, PERU, PHILIPPINES, PORTUGAL, QATAR, REPUBLIC OF PANAMA, ROMANIA, SINGAPORE, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TUNISIA, TURKEY, UNITED ARAB EMIRATES, VENEZUELA, VIETNAM &amp; WEST INDIES. | Terminated | April 14, 2015 | 866 | 15567 | Undetermined
Z-0796-2013 | 64084 | No | FCH | Manual gastroenterology-urology surgical instrument and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Custom Breast Biopsy Tray, Catalog # B 9-16261, packaged in bags, 1 tray/bag, 20 bags/case. |  |  | November 26, 2012 | 2012 | Custom Healthcare Systems, Inc. | Surgical procedure trays are packed with pharmaceuticals for injection that may contain particulate matter. | Material/Component Contamination | Custom Healthcare Systems, Inc. sent an Urgent Recall Letter dated November 19, 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to check inventory, quarantine affected product and return affected product to Stericycle using the label provided with the letter. Customer were also requested to complete and return the enclosed recall response form. For questions contact Hospira Customer Care at 1-877-946-7747. |  | 80 | 80 trays | USA Nationwide Distribution including the states GA, OH, GA, NC and ME. | Terminated | February 08, 2013 | 74 | 80 | Undetermined
Z-1058-2013 | 64751 | No | BSJ | Anesthetic gas mask. | Anesthesiology | Anesthesiology | 510(K) Exempt | Respironics SealFlex Single Port Ribbed Mask (Infant). |  |  | September 12, 2011 | 2011 | Engineered Medical Systems,Inc | The mask inside the package has been identified as "Neonatal" in size and not "Infant" as the labeling indicates. | Process change control | Consignees were notified 9/12/2011 with a Medical Device Voluntary Recall letter identifying the product and the problem. Consignees were directed to quarantine, cease further distribution, and destroy the affected product, returning a proof of confirmation/disposal letter, to be submitted for credit. Correspondence should be sent to Tamara Lefevers, Quality Manager, at tlefevers@ engmedsys.com or fax 317-246-5501. |  | 1595 | 1500pcs, 75cs {20pcs per cases) | Distributed in the states of PA, IN, NY, CO, and CA. | Terminated | August 02, 2013 | 690 | 1595 | Undetermined
Z-1082-2013 | 64427 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | The LIFEPAK CR Plus and LIFEPAK EXPRESS defibrillators. Product part's numbers: 3200731-xxx, 3202177-xxx, U3200731-xxx, and U3202177-006. |  |  | February 08, 2013 | 2013 | Physio Control, Inc. | LIFEPAK CR Plus or LIFEPAK Express Automated External Defibrillators (AEDs) are recalled due to the potential to experience an early internal battery depletion issue. | Component design/selection | The firm start sending the "URGENT MEDICAL DEVICE CORRECTION LIFEPAK CR Plus AND LIFEPAK EXPRESS AEDs" letter, dated February 2013, to their consignees.Consignees are advised to check the readiness of the AEDs and should contact Physio if the OK symbol is NOT visible and either the ATTENTION or WRENCH symbols are present.In the event a device is returned to Physio-Control for assessment, the firm will analyze it to determine if it is affected by the early internal battery depletion issue. If it is verified the device is affected, the consignee will be provided with a permanent replacement device.Customers with questions can visit the firm's website www.physio-control-notices.com/CHARGE-PAK or call Technical Support at 1-800-442-1142. |  | 54197 | 54,197 units in US and 78,627 units internationally (total 132, 824 units) | Worldwide Distribution. | Terminated | November 05, 2015 | 1000 | 54197 | Undetermined
Z-1073-2013 | 63804 | No | GGW | Partial thromboplastin time tests. | Hematology | Hematology | 510(k) | Dade(R) Actin(R) FSL Activated PTT Reagent. |  |  | March 11, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Siemens has observed that the listed lot numbers show an increase in Heparin sensitivity over the shelf life which is demonstrated by testing with Dade(R) Citrol(R) Heparin Control, low and high. | Unknown/Undetermined by firm | Siemens sent an "URGENT FIELD SAFETY NOTICE" dated February 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. The safety notice instructed consignees to discontinue use of the affected lots and discard any remaining inventory. Contact the firm at 1-800-242-3233, option 1 for questions about this notice. |  | 13083 | 13,083 | Worldwide Distribution-USA (nationwide) including the states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WV and Puerto Rico and the country of Canada. | Terminated | December 24, 2013 | 288 | 13083 | Undetermined
Z-1071-2013 | 63622 | No | FMQ | Protective restraint. | General Hospital | General Hospital | 510(K) Exempt | Posey Biothane Restraint Cuffs, Catalog Number 2900. |  |  | October 31, 2012 | 2012 | J T Posey Company | The firm voluntarily recalled Biothane Restraint Cuffs, Catalog No. 2900, that did not meet the required specifications for tensile strength. Failure of these rivets could allow the cuff to become disconnected from the connecting strap and allow patients to regain free use of the limb. | Process design | J. T. Posey sent an Medical Device Recall (Removal) letter to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letter instructed the customer that Posey Company is voluntarily recalling all units of Posey BiothaneQD Cuffs, Catalog Number 2900. All lots of this product are being recalled. Reason for recall was listed:"Posey has identified rivets on the Biothane~ Restraint Cuffs, Catalog No. 2900, that did not meet therequired specifications for tensile strength. Failure of these rivets could allow the cuff to become disconnected from the connecting strap and allow patients to regain free use of the limb. No patient injuries or deaths have reported."Please immediately discontinue further distribution or use of Posey Biothane~ Cuffs, Catalog Number 2900.Please Complete and return the Acknowledgement and Receipt Form to Posey within 7 days using the provided self-addressed, stamped envelope, by fax to 626-443-5418 or via e-mail to RegulatoryAffairs@Posey.com. Return this form even if you do not have any devices in your facility.Please return directly to Posey ill! units you have of Biothane~ Cuffs, Catalog Number 2900, using the enclosed pre-paid shipping label containing an RMA number. Posey will provide replacement product or credit upon receipt of the product.If you have any further questions please call (626) 443-3143. |  | 1673 | 1673 units | Nationwide Distribution | Terminated | July 17, 2014 | 624 | 1673 | Undetermined
Z-0739-2013 | 64008 | No | JSC | Microbiological assay culture medium. | Microbiology | Microbiology | 510(K) Exempt | BBL(TM) Fildes Enrichment, catalog number 211866, 10 tubes/pack, labeled in part ***Becton Dickinson and Company, Sparks, MD 21152 USA*** |  |  | December 10, 2012 | 2012 | Becton Dickinson &amp; Co. | Enrichment media intended to enhance the cultivation of various microorganisms may fail to support the growth of Haemophilus species. | Unknown/Undetermined by firm | Becton Dickinson sent an Urgent Product Recall letter dated December 2012 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were requested to discontinue use and discard any remaining product for replacement. Customers were also advised that due to an anticipated extended delay for the availability of Fildes Enrichment, BD will substitute DIFCO" Supplement B with Reconstituting Fluid as the replacement product. Customers were advised to ensure that this replacement product is appropriately validated in accordance with their laboratory's procedures. Customers with questions should contact BD Technical Services Department at 1-800-638-8663.For questions regarding this recall call 410-316-4000. |  | 11710 | 11,710 tubes | Worldwide Distribution - USA (nationwide) and internationally to Australia, Belgium, Canada, South Korea, Taiwan, Hong Kong, Japan, Mexico, and Singapore. | Terminated | July 01, 2013 | 203 | 11710 | Undetermined
Z-0968-2013 | 64535 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Spacelabs BleaseSirius Anesthesia Workstation, BleaseFocus Anesthesia Workstation, and CAS I/II Absorber contained in Service Kits Part Number 12200902, 050-0659-00 and 050-0901-00 and any workstation installed with the CAS I/II Absorber. |  |  | January 15, 2013 | 2013 | Del Mar Reynolds Medical, Ltd. | A defect in CAS I/II Absorbers in the Spacelabs Anesthesia Workstations and Service Kits may cause a condition leading to an increase in the carbon dioxide concentration within the inspired gas being delivered to the patient. | Process control | Spacelabs sent an URGENT-MEDICAL DEVICE CORRECTION letter, dated 16 January 2013, to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers until a Spacelabs representative replaces the affected product: 1) Immediately cease the use of BleaseSirius and BleaseFocus Anesthesia Workstations that shipped after 31 October 2012. 2) Immediately cease use of Service Kits Part Number 12200902, 050-0659-00 and 050-0901-00 and any workstation containing a CAS I/II Absorber from such Service Kit. 3) Advise your staff of this situation Customers can call Spacelabs Healthcare at 1-800-522-7025 and select 2 for Technical support. |  | 108 | 108 total (17 units in the US and 91 units outside the US) | Worldwide Distribution - USA Nationwide including the states of New Mexico, North Carolina and the countries of Argentina, China, Mexico, and Taiwan. | Terminated | May 14, 2013 | 119 | 108 | Undetermined
Z-0975-2013 | 64485 | No | HTW | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Intraosseous Fixation System (IFS) 1.7 mm, 2.0 mm &amp; 2.5 mm Cannulated Drill Bit. |  |  | October 30, 2012 | 2012 | Orthohelix Surgical Designs Inc | The firm was notified by a customer that there was a burr on the inside end of the drill bit resulting in the drill bit not fitting over the associated guide wire. | Nonconforming Material/Component | OrthoHelix sent a "FIELD REMOVAL NOTIFICATION" letter dated November 7, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at 330-247-1444 for questions concerning this notice. |  | 5040415514 | Lot # 2399041, 97 pcs; Lot # 2399051, 143 pcs; Lot # 2399051R, 4 pcs &amp; Lot # 2399061, 119 pcs | Nationwide Distribution including the states of CA, CO, IL, MI, MD, NC, NM, OH, OK, PA, SD, TX, UT, VA, WA and WI. | Terminated | December 05, 2013 | 401 | 5040415514 | Undetermined
Z-1051-2013 | 64189 | No | FMZ | Neonatal incubator. | General Hospital | General Hospital | 510(k) | Giraffe OmniBedÂż and Giraffe Incubator p/n 6600-0723-700 (rev A and B) labeled in part: ***GE Healthcare 8880 Gorman Road Laurel, MD 20723 U.S.A. www.gehealthcare.com***. |  |  | January 08, 2013 | 2013 | GE Healthcare, LLC | Neonatal beds and incubator display boards may randomly overwrite stored information and radiant heater and oxygen set points upon power up or recovery from a power failure. | Process control | GE Healthcare sent an "Urgent Medical Device Correction" letter, dated January 8, 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customers were given the safety issue and instructions for continued use of the Giraffe OmniBed and Incubators, as well as given the affected product details. GE Healthcare will correct all affected Giraffe Incubators and Giraffe OmniBeds at no cost, and will contact customers to arrange for this correction. If customers have any questions or concerns regarding this notification, please call: United States: 1-800-345-2700. For other countries, please contact your local GE Healthcare Service Representative. |  | 679 | 679 units | Worldwide Distribution -- USA in the states of AR ,CA, CO, DC, FL, GA, HI, ID, IL, IN, KS, MA, MD, MI, MO, MS, NC, NE, NJ, NY, NV, OH, PA, TN, TX, VA, WA, WI, and WV, and the countries of Arab Emirates, Albania, Australia, Belgium, Brazil, Canada, Chile, China, Croatia, Czech., Denmark, Ecuador, Estonia, Finland, France, Germany, Hungary, Italy ,Korea, Kurdistan, Lebanon, Macedonia, Malaysia, Mexico, Norway, Oman, Poland, Russia, San Salvador, Saudi Arabia, Singapore, South Africa, Sweden, Turkey, Uruguay, Venezuela, and United Kingdom. | Terminated | July 10, 2013 | 183 | 679 | Undetermined
Z-0976-2013 | 64310 | No | NJR | Streptococcus spp. serological reagents. | Microbiology | Microbiology | 510(K) Exempt | BD MAX PCR Cartridges Catalog #437519, box of 24 cartridges labeled in part:***BD MAX PCR Cartridges, GeneOhm Sciences Canada, Inc. 2555 boul du Parc-Technologique, Quebec, Canada G1P 4S5***BD Max Group B Streptococcus (GBS) Assay is an in vitro diagnostic (IVD) test for detection of GBS DNA in Lim Broth cultures. |  |  | December 11, 2012 | 2012 | Becton Dickinson &amp; Co. | Cartridges are not sealing properly during PCR testing and may cause false results, indeterminates and/or invalid runs. | Process control | BD Diagnostic Systems issued an Urgent Product Recall letter via UPS delivery on December 11, 2012. The letter identified the product, the problems, and the action to be taken by the customer. Customers were instructed to take the following actions: Discontinue and discard any affected product. BD would send replacements for the discarded material. Complete the attached form whether or not they have any inventory remaining so that the firm may acknowledge their receipt of the notification. Fax the form to: Attention: Regulatory Compliance, Fax: 410-316-4258 or Email to RegulatoryComplianceFax@BD.com.. If further assistance is needed regarding replacements, please contact BD Customer Service Department at 1-800-675-0908. For all other inquiries, contact BD Technical Services Department at 1-800-638-8663. The BD site in Belgium was notified via email.For questions regarding this recall call 410-316-4054. |  | 17496 | 27 boxes/648 cartridges | Worldwide Distribution - USA including CA, FL, IN, IL, KS, NY, OH, OR, PA, TX, and WI and the country of Belgium | Terminated | July 22, 2013 | 223 | 17496 | Undetermined
Z-0988-2013 | 63494 | No | JIT | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens Dimension Vista(R) Alkaline Phosphatase (ALP) Calibrator |  |  | October 10, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Correlation slope bias and accuracy shift | Nonconforming Material/Component | Siemens sent a Urgent Medical Device Correction letter dated October 16, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed to calibrate with newer lots of Dimension Vista(R) ALP Calibrator as they become available, beginning with lot 2GD011. Customers were also asked to discontinue use and discard any remaining inventory of lot 2AD052. Customers further instructed to complete the attached response form and fax to (302) 631-8467, to indicate they had received the information and to facilitate their no charge replacement.Further questions please call (302) 631-6311. |  | 1985 | 1985 cartons | Nationwide distribution: USA including Puerto Rico and the states of AL, OH, MO, GA, IL, CT, IA, SC, VA, TX, FL, NY, NC, CA, NM, AZ, CA, MI, NJ, PR, MS, NE, WV, PA, MT, MA, MD, KS, NH, TN, VT, LA, CO, ID, WA, NH, DE and OR. | Terminated | January 08, 2014 | 455 | 1985 | Undetermined
Z-0792-2013 | 64115 | No | LKA | Culture medium for antimicrobial susceptibility tests. | Microbiology | Microbiology | 510(k) | Diagnostic Hybrids, Inc., RhMK, Model #:49-0600, 49-0600A, 49-0102A, and 49-T075A. Culture media, used in the cultivation and amplification of viruses, such as respiratory viruses, enteroviruses, measles, and poliovirus from patient specimens. |  |  | December 31, 2012 | 2012 | Diagnostic Hybrids Inc | RhMK product fungal contamination. | Process control | Diagnostic Hybrids sent a Urgent: Medical Device Recall letter dated December 31, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The customers were instructed to immediately review your inventory and determine if you have any of the affected product.If so, please remove it from inventory. Complete the enclosed Product Recall Fax-Back form. Upon receipt of Fax-Back form, Diagnostic Hybrids will send you replacement product or issue credit.Representatives are available to assist you in this process and answer any questions you may have about this recall, and how to obtain replacement product. Please contact Technical Support at technicalsupport@quidel.com or by calling 800.874.1517, Monday through Friday, 8:00 a.m. to 5:00 p.m. EST. |  | 13005 | 13,005 units | Nationwide Distribution including the states of AR, AZ, CA, CO, CT, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NJ, NM, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, VA, WA &amp; WI. | Terminated | December 05, 2013 | 339 | 13005 | Undetermined
Z-1067-2013 | 64667 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | TSX-IOI: AQUILION 32/64; System, X-Ray, Tomography, Computed. |  |  | November 30, 2011 | 2011 | Toshiba American Medical Systems Inc | The firm initiated this recall due to a potential software issue. The DLP (Dose Length Product) value may be displayed incorrectly when the vHP option is used. | Process design | Toshiba sent "URGENT: MEDICAL DEVICE CORRECTION" letters to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. The letter contained a return reply form that is to be faxed to TAMS for retention. Contact the firm at (800) 421-1968 for questions relating to this letter. |  | 23 | 23 units | Nationwide Distribution including the states of WA, OH, CA, TX, WV, IL, AZ, CA, WI, ME, WI, NC, FL, MT, and HI. | Terminated | June 17, 2013 | 565 | 23 | Undetermined
Z-0702-2013 | 63868 | No | RCA |  |  |  |  | Juno DRF , X-ray system, diagnostic, Fluoroscopic, general-purpose |  |  | December 13, 2012 | 2012 | Philips Healthcare Inc. | Values of fluoroscopy time, Air Kerma Rate andCumulated Air Kerma may not be displayed "atthe fluoroscopist's working position" as per 21 CFR 1003.21 | Radiation Control for Health and Safety Act | Philips Medical issued a notification by letter dated December 13, 2012. via Certified Mail to affected customer. The letter identified the affected product, problem and actions to be taken. Philips will contact affected customers to schedule an upgrade of the equipment . For questions or support concerning this issue, please contact Dominic Siewko at 978-659-7936. |  | 16 | 16 units | US Nationwide Distribution including the states of AL, CA, CO, CT, IL, MD, MI, MO, NY, OH, PA | Open, Classified |  |  | 16 | Undetermined
Z-0744-2013 | 64135 | No | IWZ | Radiographic film. | Radiology | Radiology | 510(K) Exempt | KODAK INSIGHT DENTAL FILM in SureSoft Packets, Size 2 Single Packet Film, Model Number IP21S, REF/Catalog Number 8118226 --- COMMON/USUAL NAME: Dental Intraoral Radiographic Film -- 20 cartons of 80 packets. |  |  | January 17, 2013 | 2013 | Carestream Health, Inc. | Missing symbols and manufacturing information on the clear carton label. The expiration dating symbol is missing but the expiration date is not missing. The Lot symbol is missing but the lot number is not missing. The Manufacturer symbol, name and address including country of origin are missing. All of the remaining symbols and information are on the case label. Missing information is only a | Process change control | Carestream sent an Urgent Customer Notification and Field Corrective Action Letters dated January 17, 2013 and response form to the domestic consignees via UPS Overnight Mail. The letter identified the affected product, problem and actions to be taken. In addition, the consignees were instructed to complete the response form and fax it back to Carestream to fax # 585-323-7643. For questions or concerns customers were instructed to contact the Carestream Customer Care Center in the U.S. at 1-800-933-8031. |  | 202 | 101 cartons: Domestic - 95 cartons; Foreign - 6 cartons | US Nationwide Distribution including the states of California, Florida, Texas, Illinois, Indiana, Maryland, Pennsylvania and Canada. | Open, Classified |  |  | 202 | Undetermined
Z-1064-2013 | 64527 | No | DSP | Intra-aortic balloon and control system. | Cardiovascular | Cardiovascular | 510(k) | Arrow International FiberOptix Intra-Aortic Balloon Catheter Kit (IAB) 8 Fr 30 cc; |  |  | November 27, 2012 | 2012 | Arrow International Inc | Potential to have an internal leak at the fill port of the stainless steel tip, which may necessitate removal and replacement ofthe catheter. | Process control | Arrow International issued Urgent Medical Device Recall notification on 11/27/12. The letter identified the affected product and the reason for the recall. Accounts are requested to check their stock for the products included within the scope of this recall. Also, customers should cease use and distribution, and quarantine all affected product immediately. Contact Arrow's Customer Service Department at 800-523-8446 for a Return AuthorizationNumber. Customers are to complete and fax the Recall Acknowledgement &amp; Stock Status Form to the number provided, and return the affected product. If customers have questions or need clarification regarding this issue, they should contact Arrow's Customer Service Department at 1-800-523-8446. |  | 188 | 188 units | Worldwide Distribution -- USA, Belgium, Great Britain, Italy, Kazakhstan, Latvia, Russia, Slovenia, Turkey, and South Africa. | Terminated | June 02, 2015 | 917 | 825 | Undetermined
Z-0790-2013 | 63555 | No | JTN | Antimicrobial susceptibility test disc. | Microbiology | Microbiology | 510(k) | Oxoid Antimicrobial Susceptibility Test Discs, Aztreonam 30 ug, IVD, REF CT0264B, Oxoid Ltd., UK, oxoid@thrmofisher.com (Both Remel &amp;amp;amp;amp;amp; Oxoid, Ltd are a part of Thermo Fisher Scientific) |  |  | August 21, 2012 | 2012 | Remel Inc | The firm is recalling the product due to some cartridges may contain individual discs that are not impregnated with antibiotic. The firm notified their customers of the recall with a notification letter and requested a confirmation fax back of receipt. | Nonconforming Material/Component | ThermoFisher sent an Important Medical Device Product Recall Notice dated August 17, 2012, to all affected customers. The letter identifed the product, the problem, and the action to be taken by the customer. Customers were instructed to review results reported using this lot and consider restesting and/or seek approproiate expert advice. Customers were requested to destroy any remaining inventory of the affected lot and contact Customer Services or their local Oxoid supplier who will issue a replacement. Customer were asked to return the attached "Acknowledgement &amp; Receipt Form by FAX to 1-877-428-1922 or 1-913-895-4190. Customers with questions were instructed to call the Technical Support Department at 800-255-6730 (US) or 913-888-0939 (International).For questions regarding this recall call 913-888-0939. |  | 30 | 30 | Nationwide Distribution including WI, MO, MA, CT, AL, IL, WA, CA, RI, MD, NJ, NC and TX. | Terminated | June 04, 2013 | 287 | 30 | Undetermined
Z-1061-2013 | 64655 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker Orthopaedics Baseplate Impactor Extractor TRIATHLON Instruments. Howmedica Osteonics Corp. 325 Corporate Drive, Mahwah, NJ 07430 USA A subsidiary of Stryker Corporation. Made in USA. |  |  | February 23, 2013 | 2013 | Stryker Howmedica Osteonics Corp. | Stryker received reports from the field indicating that the handle detached from the cam of the Baseplate Impactor/Extractor Assembly. | Nonconforming Material/Component | Stryker sent Notification Letters/Product Accountablility Forms dated March 1, 2013 via FedEx to their Branches/Agencies on March 1, 2013.. Notification Letters/Product Recall Acknowledgement Forms dated March 4, 2013, were sent via FedEx to Hospital Risk Management on March 4, 2013. The Notification letters identified the product, the problem, and instructed customers on how to identify affected product; informs customers of the issues and potential hazards with the affected product and instructs them to contact their representative to arrange for return/replacement of any affected product. Customers are also asked to fax back the provided "Urgent Medical Device Removal Notification Acknowledgement Form" to Stryker within 5 days of receipt.Customers with questions were instructed to call 201-972-2100.For questions regarding this recall call 201-831-5523. |  | 871 | 871 units | Worldwide Distribution - USA (nationwide) and internationally to Canada, Japan, Australia, China, New Zealand, France, Italy, Netherlands, Poland, Spain, Sweden, Switzerland and the United Kingdom. | Terminated | June 15, 2015 | 842 | 871 | Undetermined
Z-0995-2013 | 64689 | No | FRG | Sterilization wrap. | General Hospital | General Hospital | 510(k) | KIMGUARD Container Filters, 7.5in (19cm), Round and 9in. X 9in. (23cm x 23cm), Square. |  |  | March 06, 2013 | 2013 | Kimberly-Clark Corporation | Some individual container filter units might contain thin areas that may potentially compromise the ability of the filter to maintain a sterile barrier. | Component design/selection | All US distributors and clinics were sent an Urgent: Voluntary Product Recall letter via Federal Express, and the Canadian importer/distributor was also notified with the same letter via Federal Express and e-mail. Distributors were instructed to contact their customers by sending a copy of the letter. The letter identified the affected product and stated the reason for the recall. The recall notification includes a Response Sheet that allows the customer (end-user) to indicate the quantity of affected product they have in their inventory or indicate whether they do not have any affected product. The customer is instructed to immediately cease further distribution of the product, physically quarantine and destroy the product, and fax the completed Response Sheet to the Kimberly-Clark Recall Coordinator. |  | 144717 | Code 68507 - 6420 cases; Code 68509 - 1281 cases | Worldwide Distribution -- USA, including the states of AL, AK, AZ, AR, CA, CO, FL, GA, ID, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MS, MO, NV, NJ, NM, NY, NC, OH, OK, OR, PA, TN, TX, UT, VA, WA, and WI, and the countries of Canada, Australia, Belgium, Hong Kong, Japan, Portugal, and Singapore. | Terminated | September 16, 2013 | 194 | 144717 | Undetermined
Z-0789-2013 | 63823 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | ConMed Linvatec Biomaterials, PopLok bone Punch PKL-35M, 3.5mm |  |  | November 15, 2012 | 2012 | Linvatec Corp. dba ConMed Linvatec | ConMed Linvatech recalled one lot of their PopLok bone Punch 3.5mm because the metal shaft diameter is 4.5mm when it should be 3.5mm. | Employee error | ConMed Linvatech issued a "URGENT MEDICAL DEVICE RECALL NOTIFICATION" to its customers on November 15, 2012 for a recall that was conducted at the end user/distributor level. All recall notifications were sent FedEx next day delivery or equivalent method to notify all direct consignees. The notifications identified the affected product along with reason for the recall health risk. All accounts were asked to immediately check their facility for the affected product, stop using any of the product with specified lot number, and to segregate the product. The letter also asked customers to return all affected product to ConMed Linvatec and respond using a Reply Form. If customers have any additional questions, they should contact the Customer Service department at 800-237-0169.For questions regarding this recall call 727-399-5205. |  | 8 | 8 | Worldwide Distribution - USA including California and Internationally to Spain and Korea | Terminated | May 22, 2013 | 188 | 8 | Undetermined
Z-1063-2013 | 64420 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Covidien Endo GIA AutoSuture Universal Loading Unit, 30mm-2.0mm |  |  | February 11, 2013 | 2013 | Covidien LP | Single Use Loading Unit contained two staples loaded in each cartridge pocket and may result in difficulty firing and removing the device from the application site, which may require medical intervention | Process control | Covidien notified customers with an Urgent Medical Device Recall notification on February 11, 2013. The letter identified the affected product and the reason for the recall. Accounts were requested to quarantine and discontinue use of the affected devices, and notify Universal Loading Unit 30mm-2.0mm users of this recall. Customers should follow the instructions provided in order to return the affected product. Customers are to complete and return the Product Return Form regardless if they have any affected product or not. For questions regarding the RGA I return process, please contact Covidien Customer Service, M-F, 8 am - 6:30 pm ET at (800) 962-9888, option 1, and then option 2. |  | 6751 | 6,751 units | Worldwide Distribution -- USA, Asia Pacific, Canada, Europe, and Japan. | Terminated | March 20, 2014 | 402 | 6751 | Undetermined
Z-0818-2013 | 63927 | No | GXY | Cutaneous electrode. | Neurology | Neurology | 510(k) | Foam electrodes with Silver-Carbon construction: |  |  | November 30, 2009 | 2009 | NAImco Inc dba Richmar Inc | Electrode Socket Terminal pulling away from the electrode lead wire when the electrical stim lead wire was removed. | Unknown/Undetermined by firm | Rich-Mar sent "URGENT: MEDICAL DEVICE VOLUNTARY RECALL" letters to affected customers beginning on 11/30/2009. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at 423-648-7730 ext. 214 for questions relating to t his recall. |  | 56980 | 56980 electrodes | Distributed in the states of GA and TN. | Terminated | September 03, 2013 | 1373 | 422100 | Undetermined
Z-1026-2013 | 64692 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | STERRAD 100NX Sterilization System, P/N 10104 |  |  | March 22, 2013 | 2013 | Advanced Sterilization Products | The recall was initiated because Advanced Sterilization Products (ASP) has determined that the STERRAD 100NX and STERRAD 200 Sterilization Systems may emit an odor or smell into the surrounding environment that is not typical to normal operating conditions. | Nonconforming Material/Component | Advanced Sterilization Products sent an Urgent Medical Device Field Safety Notification letter dated March 22, 2013 to all customers. The letter identified the affected products, problem and actions to be taken. Customers are instructed to report any complaints or suspected problems with their STERRAD 100NX and/or STERRAD 200 System to ASP Professional Services directly at (888) 783-7723 and select option #2. |  | 1500 | 1500 units | US Nationwide Distribution | Terminated | September 30, 2013 | 192 | 2538 | Undetermined
Z-0644-2013 | 63842 | No | CGJ | Follicle-stimulating hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Laparoscope, general; plastic surgery |  |  | November 05, 2012 | 2012 | Applied Medical Resources Corp | Applied Medical is conducting a voluntary recall on a specific lot number of its 12x150mm Kii shielded bladed access systems. This is due to a potential for the blade to remain exposed after shield deployment. Further use of any affected product should cease immediately.The affected model number is CTB71, and the affected lot is 1179107. | Nonconforming Material/Component | Applied Medical sent an Urgent Medical Device Recall to all affected customers in November 2012. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their inventory for the affected product, complete the attached Recall Notification Confirmation Form to acknowledge the recall and indcated any product their facility is returning. The Recall confirmation form should be returned to Applied Medical by emailing to recall60494844#applied medical.com or fax to 949-713-8832. Return affecte4d product and a copy of the confirmation form to Applied Medical. Customers with questions were instructed to call 949-713-8662. For questions regarding this recall 949-713-8662. |  | 72 | 72 units | Worldwide Distribution - USA (nationwide) and France. | Terminated | January 16, 2013 | 72 | 72 | Undetermined
Z-1158-2013 | 64715 | No | KNX | Urine collector and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Belly Bag Urine Collection Bag with Hip Belt. |  |  | March 28, 2013 | 2013 | Teleflex Medical | Sterile packaging may be compromised. | Nonconforming Material/Component | Teleflex Medical sent an Urgent Medical Device Recall Notification dated March 28, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed if you have any of the affected stock, immediately discontinue use and quarantine any products with the catalog numbers listed.If you and your customers have no affected stock please complete the enclosed Recall Acknowledgement Form and fax it to 1 (866) 804-9881 Attn: Customer Service. This will allow us to document your receipt of this letter.To return product complete the enclosed Recall Acknowledgement Form and fax it to 1 (866) 804-9881 Attn: Customer Service. This will allow us to document the amount of product you have on hand for return.If you have any other questions feel free to contact your local sales representative or Customer Service at 1 (866) 246-6990. |  | 190103 | 190,103 | Worldwide Distribution - USA and the countries of Singapore, Canada, New Zealand, Germany, Ireland, Japan and Australia. | Terminated | May 01, 2014 | 399 | 190103 | Undetermined
Z-0886-2013 | 64357 | No | KNF | Unipolar endoscopic coagulator-cutter and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Integra(R) Jarit(R) Monopolar Cable (Catalogue No. 600-290). |  |  | February 12, 2013 | 2013 | Integra LifeSciences Corp. | Addendum to the Instructions for Use.The addendum emphasizes the warnings and precautions already contained in the device's IFU and also provides additional warnings and precautions that are applicable to the use of electrosurgical equipment in health care facility environments. | Other | Integra sent an Urgent Medical Device Correction Field Safety Notice letter dated February 12, 2013 and a copy of the Instructions for Use addendum to all affected customers. The letter identified the product, nature of the issue and actions to be taken. Customers were requested to examine their inventory for the affected product and include a copy of the addendum to that inventory. Customers were instructed to complete and return the Acknowledgement Return Form and provide a copy of the IFU addendum to any of their customers that had purchased the Jarit(R0 Monopolar Cable. For questions call Integra Surgical Customer Service at 800-431-1123. |  | 22500 | 22,500 | Worldwide Distribution - USA Nationwide including Canada. | Terminated | December 24, 2013 | 315 | 22500 | Undetermined
Z-1178-2013 | 64701 | No | DCK | C-reactive protein immunological test system. | Immunology | Immunology | 510(k) | The Alere Cholestech LDXÂż high sensitivity C-Reactive Protein (hsCRP) Test Cassette, Model Number 12-807. |  |  | March 06, 2013 | 2013 | Alere San Diego, Inc. | The Alere Cholestech LDXÂż hsCRPcassette (PN 12-807) may have increased imprecision relative to the performance data in the package insert,which may cause differences in quantitative results for hsCRP compared to a reference method orrepeat testing on the same patient sample. | Nonconforming Material/Component | Alere sent an Urgent Medical Device Recall letter dated March 6, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to block all stock of the affected lot in their warehouses; discontinue use and discard all stock on hand, complete the provided Customer Verification form to document disposal, contact their primary consignees and advise them of the required discontinuation and disposal of the affected lot. For question call 877-308-8289. |  | 2333 | 2333 kits | USA Nationwide Distribution | Terminated | March 12, 2014 | 371 | 2333 | Undetermined
Z-0648-2013 | 63852 | No | JLW | Thyroid stimulating hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Immunoradiometric assay. |  |  | November 28, 2012 | 2012 | Siemens Healthcare Diagnostics | The recall was initiated because Siemens Healthcare Diagnostics confirmed an under-recovery of dose values with some vials of Coat-A-Consta (CAC) IRMA TSH calibrators TSI4 (level B), TSI5 (level C), TSI6 (level D), TSI8 (level F), and TSI9 (level G) lot 035. The affected calibrator vials can be identified by a darker yellow color than normal. The TSI3 (level A), TSI7 (level E), and TSIX (level H) | Nonconforming Material/Component | Siemens sent an Urgent Medical Device Recall letter dated November 2012 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use of the affected product, complete and return the Urgent Medical Device Recall Effectiveness Check form included with the letter by fax to 302-631-7597. Customers were also instructed to determine their replacements needed, and to forward the recall letter to whomever they may have distributed the affected product. For questions customers were instructed to contact their Siemens Technical Solutions Ceter or their local technical support representative. For questions regarding this recalla call 914-524-2955. |  | 8 | Lot 742= 8 units all foreign consignees) Lot 743=7 units (all foreign consignees) | Worldwide Distribution - USA including CA, GA, IL, IN, MI, MN, NY, and Puerto Rico.Internationally to Agentina, Bolivia, Brazil, Canada, Chile, Guatemala, South Korea, Philippines, Paraguay, Taiwan, and Uraguay, | Terminated | October 28, 2013 | 334 | 1987 | Undetermined
Z-0786-2013 | 63392 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Mx8000 IDT Whole Body Computed Tomography X-ray system Distributed by Philips Medical Systems, Highland Heights, OH 44143 |  |  | September 18, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | The low contrast pins in the Image Quality (IQ) phantom deteriorates over time. This prevents the customers from doing the Daily Image Quality Checks as defined in the current Instructions for Use (IfU). | Radiation Control for Health and Safety Act | Philips sent a Urgent Medical Device Recall Notice letter dated September 19, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken.This issue does not effect the performance of the Scanner. The Customer Information is intended to inform you about:1. What the problem is and under what circumstances it can occur.2. The action you as a customer can take to minimize the effect of this problem.3 The action planned by Philips to correct this problem.If you need any further information or support concerning this issue please contact your local Philips Representative. In the US and Canada please contact the Philip's Healthcare customer care solutions at 1-800-722-9377, and follow the recorded menu options to reach a Customer Solutions Engineer. In all other countries please contact your local Philips Healthcare office. |  | 1474 | 1,474 Units | Worldwide Distribution - USA ( nationwide) including the states of AK, AL, AR, AZ, CA, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, TN, TX, VA, WA, WV and WY, and the countries of Albania, Algeria, Argentina, Australia, Austria, Belarus, Belgium, Bolivia, Brazil, Canada, China, Colombia, Denmark, Dominican Republic, Ecuador, Egypt, France, Germany, Ghana, Gibraltar, Greece, Hungary, India, Indonesia, Iran, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kuwait, Lebanon, Libyan Arab Jamahiriya, Lithuania, Malaysia, Mexico, Morocco, Netherlands, New Zealand, Nicaragua, Nigeria, Norway, Panama, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Singapore, Slovakia, South Korea, Spain, Sri Lanka, Sweden, Switzerland, Syrian Arab Republic, Taiwan, Thailand, Turkey, Turkmenistan, United Arab Emirates, United Kingdom, Venezuela, Viet Nam &amp; Yemen. | Terminated | September 02, 2015 | 1079 | 1474 | Undetermined
Z-1161-2013 | 64739 | No | LOM |  |  | Microbiology | PMA | VITROSÂż Immunodiagnostic Products Anti-HBs (total antibody to hepatitis B surface antigen (anti-HBs)) Reagent Packs |  |  | November 01, 2012 | 2012 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics is recalling certain lots of the Vitros Anti-HBs Reagent packs an inability to obtain a valid calibration. | Under Investigation by firm | Ortho Clinical Diagnostics (OCD) sent an Urgent Product Correction Notification letter dated November 1, 2012, to all affected customers via USPS. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed to discontinue and discard any remaining recalled lots of VITROS Anti-HBS Reagent Packs. Foreign affiliates were informed by email on November 1, 2012, of the issue and instructed to notify their consignees of the issue and actions . On March 22, 2013, OCD sent customer letters via FedEX overnight and/or via USPS to US consignees for two additional lots of product being recalled. The letter provided similar recall instructions as the initial notification. Ortho Clinical Diagnostics has requested that their clients discontinue and discard any remaining inventory of VITROS Anti-HBS Reagent Packs subject to this recall. OCD has also requested that recall product recipients also complete a Confirmation of Receipt and return that document to OCD, fax to: 1-888-557-3759 or 585-453-4110.We apologize for any inconvenience this may cause your laboratory. If you have technical questions please contact Customer Technical Services at 1-800-421-3311. If you have questions about your replacement order, please call our Customer Service Center at 1-800-828-6316. |  | 1464 | 1464 units | Worldwide Distribution- USA (nationwided) and countries of: Australia, Brazil, Canada, Chile, Chao Yang District, Beijing, Columbia, England, France, Germany, India, Italy, Japan, Mexico, Republic of Panama, Singapore, Spain and Venezuela. | Terminated | March 30, 2017 | 1610 | 1464 | Undetermined
Z-1166-2013 | 64783 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | NeuViz Dual series CT Scanner System, Part number (PN): 989605651321. This is a computed X-ray, Tomography system. |  |  | March 04, 2013 | 2013 | Philips And Neusoft Medical Systems Co., Ltd. | User error of the NeuViz 16 or NeuViz Dual may cause personnel injuries. There was a report where a finger was pinched when the table was moved backward while the tabletop was controlled by another person from the gantry left control panel. | Use error | The firm, Neusoft Medical Systems Co., Ltd., sent an "URGENT DEVICE CORRECTION" NeuViz Dual and NeuViz 16 letter, dated Mar. 4, 2013, to their US consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed to review this information with all members of their staff and retain a copy of the letter with the equipment Instruction For Use.Customers with questions should contact the Service Support Department by email to nms-service@neusoft.com |  | 25 | 25 units | Nationwide distribution: USA including states of: California, Connecticut, North Carolina, Nebraska, Ohio, South Carolina, Texas, and Puerto Rico. | Terminated | November 22, 2013 | 263 | 48 | Undetermined
Z-0651-2013 | 63226 | No | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | Ultrafill DBM labeled in part: TISSUENET Sterilized by T10^6 Description: UltraFill DBM Size: 1cc Product Code: RT53001 Donated Human Tissue Allograft SINGLE PATIENT USE ONLY FOR INTERNATIONAL USE ONLY See Package Insert for Additional Information |  |  | March 14, 2012 | 2012 | Surgical Tissue Network, Inc. | Surgical Tissue Network Inc., DBA TissueNet Inc. recalled their UltraFill DBM Putty-Porcine (1 cc; 5cc; 10cc) because products may contain trace amounts of 316L Stainless Steel particulates. | Employee error | Surgical Tissue Network sent a Notification of Voluntary Tissue Recall dated March 22, 2012, to all affected customers. The firm issued an additional notification letter on August 30, 2012. The letter identified the product, the problem, and action to be taken by the customer. Consignees were asked to return all listed products that remain in inventory. If the product was further distributed, they were asked to forward the recall information to their customers.Customers were instructed to contact TissueNet's Customer Service Department at 800-465-8800 x283 to coordinate return/replacement of the affected product. Customers with questions were instructed to call 800-465-8800 x230.For questions regarding this recall call 407-380-2424. |  | 25 | 25 units | Worldwide Distribution including Turkey, Greece, &amp; Mexico. | Terminated | December 02, 2014 | 993 | 393 | Undetermined
Z-0832-2013 | 64204 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | CareFusion SmartSite Gravity Blood Set Model 42082E. Used to administer blood and blood products through a needle or catheter inserted into the patient's artery or vein. |  |  | December 20, 2012 | 2012 | Carefusion 303 Inc | CareFusion is recalling the SmartSiteÂż Gravity Blood Set as a result of a potential for separation at the connection of the drip chamber and tubing . The separation may be observed at priming or observed during use. If a separation between the drip chamber and tubing occurs, this could cause a delay in treatment. | Nonconforming Material/Component | Carefusion sent an "URGENT: Medical Device Recall Notification" letter dated January 21, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The customers were instructed to return all recalled product directly to the distributor from whom it was purchased immediately, complete and return the enclosed pre-addressed and postage paid mandatory Customer Recall Response Card and to contact your distributor directly, or call the Carefusion Support Center, if you have any questions.For questions Recall Related Questions contact CareFusion Support Center at 888-562-6018; 7am to 4pm (Pacific). For technical questions contact Technical Support at 888-812-3229, 7am to 5pm (Pacific). |  | 5050 | 5,050 units | Nationwide Distribution including the states of AK, CA, CO, IL, MO, NV, NY, TX and UT. | Terminated | July 25, 2013 | 217 | 5050 | Undetermined
Z-0908-2013 | 64417 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Medtronic DPL Single Stage Venous Cannula with Metal Tip, used in cardiopulmonary bypass procedures. |  |  | February 19, 2013 | 2013 | Medtronic Inc. Cardiac Rhythm Disease Management | Potential for uncharacteristically rough metal edges to occur along the inner diameter of the metal tip side ports of the cannula. | Nonconforming Material/Component | Consignees were sent a Medtronic "Urgent Medical Device Recall" letter dated February 19, 2013. The letter described the problem and the product involved in the recall. Advised consignees to quarantine affected product, and return the product to Medtronic. Requested consignees the complete and return the attached Customer confirmation certificate. If product has been forwarded to another facility, requested consignees to notify the facility of the issue. |  | 13691 | 13,691 (6711 USA; 6980 OUS) | Worldwide distribution, including USA nationwide, Albania, Algeria, Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, Colombia, Croatia, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, Finland, France, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Jamaica, Japan, Kazakhstan, Korea, Libya, Lithuania, Macedonia, Malaysia, Mexico, Morocco, Netherlands, New Zealand, Norway, Peru, Philippines, Poland, Portugal, Qatar, Romania, Russian Federation, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom, and Viet Nam. | Terminated | November 21, 2013 | 275 | 13691 | Undetermined
Z-0885-2013 | 64213 | No | IXI | Radiation therapy beam-shaping block. | Radiology | Radiology | 510(k) | Elekta, Apex add-on Micro MLC, collimator for radiation therapy. |  |  | January 18, 2013 | 2013 | Elekta, Inc. | If the Apex collimator is incorrectly installed on the radiation head, a touchguard inhibit activates the view restrictions dialog box only shows the top 12 restrictions. | Employee error | Elekta sent an Important Field Safety Notice Field Safety Notice FCO Ref: 20001407029 "Important information for the safe and satisfactory operation of Apex" to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to obey all the applicable instructions, warning, and cautions in the user documentation for each piece of equipment.For questions regarding this recall call 770-300-9725. |  | 1 | 1 | Nationwide Distribution including TX | Terminated | May 08, 2014 | 475 | 1 | Undetermined
Z-1188-2013 | 64889 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | GE, VCT for Discovery VCT, Discovery RX VCT, and Discovery PET/CT 690, Lightspeed VCT, Discovery CT750HD. |  |  | July 25, 2011 | 2011 | GE Healthcare, LLC | Fastener bolts on the heat exchanger were not properly torqued and part of the heat exchanger can detach from the gantry and hit the inside of the covers and other components of the gantry possibly resulting in the oil conduit breaking, spilling oil within the gantry and outside the system. This issue affects GE Healthcare VCT for Discovery VCT, Discovery RX VCT, and Discovery PET/CT 690, Lightspe | Process control | No consignee letter was sent out. At customer site Field Engineer replaced heat exchanger, the tube and the pump. Field Engineer also verified torque on all bolts. Field action was initiated on 7/25/2011 and has been completed as of 3/29/2013. |  | 166 | 166 | Worldwide Distribution-USA (nationwide) including DC and the states of AZ, CA, CO, CT, FL,GA, IL, IN, KY, MA, MI, MS,MO, NH, NJ, NM, NY, NC, OK, OR, PA, SC, SD, TX, UT, VA, WA, and WI and the countries of ALBANIA ARGENTINA AUSTRALIA BELGIUM BRAZIL CANADA CHINA COLOMBIA COSTA RICA CZECH REPUBLIC DENMARK ECUADOR EGYPT FINLAND FRANCE GREECE HONG KONG HUNGARY INDIA IRELAND ITALY JAPAN KOREA, REPUBLIC OF MEXICO MOROCCO NIGERIA NORWAY POLAND RUSSIAN FEDERATION SINGAPORE SPAIN SWEDEN SWITZERLAND TURKEY UKRAINE UNITED KINGDOM. | Terminated | February 04, 2014 | 925 | 166 | Undetermined
Z-0654-2013 | 64029 | No | REX |  |  |  |  | Lockheed Martin Gyrocam Systems ( LMGS) LMGS 15" camera systems. Laser pointer and illuminator used for law enforcement surveillance from aircraft. Models DNV, TS and DS |  |  | January 21, 2010 | 2010 | Lockheed Martin Gyrocam Systems, Inc. | These camera systems were manufactured by LMGS and found to fail to comply with the FDA laser performance standard during an FDA inspection. These camera systems incorporate Class IIIb surveying, leveling, and alignment lasers and were entered into US commerce without a variance. Additionally, the manufacturer did not submit a product report to FDA for the camera system. | Radiation Control for Health and Safety Act | Please note that LMGS is responsible for the correction of all noncompliant products.1. Notifications to all purchasers must be made, and the FDA district office Recall Coordinator is to be included in the notifications. 2. Correction of all units to be completed by November 8, 2013.3. All products will be brought into compliance with the laser performance standard at no cost to the purchaser.4. LMGS will submit a product report for the new camera system prior to installation of the replacement lasers as per 21 CFR 1002.10 and certify that the new camera systems comply with the general performance standard in 21 CFR 1010 and the laser performance standard found in 21 CFR 1040.10 and 1040.115. Updated operator manuals will be provided to purchasers.Further questions please call (941) 355-3206. |  | 33 | 33 units | Nationwide Distribution | Terminated | February 24, 2015 | 1860 | 33 | Undetermined
Z-0912-2013 | 64404 | No | OCL | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | Cardiovascular | 510(k) | Atricure Inc., Max1 Isolator Transpolar Pen. Model A000163. Used to ablate cardiac tissues during cardiac surgery using radiofrequency energy when connected to the AtriCure Ablation and Sensing Unit or for temporary cardiac pacing, sensing, recording, and stimulation during the evaluation of cardiac arrhythmias during surgery when connected to a temporary external cardiac pacemaker or recording device. |  |  | January 14, 2013 | 2013 | Atricure Inc | Products from lot no. 41865 may have been sent out with Incorrect Instructions for Use (IFU) labeling. | Employee error | AtriCure sent a Urgent Medical Device Recall/Correction letter dated January 14, 2013, to all affected customers. The letter identified the product the problem and the action to be taken by the customer.AtriCure has identified that products from lot no. 41865 may have been sent out with Incorrect Instructions for Use (IFU) labeling. If the MAX1 is used as intended there is no risk to the patient as a result of this issue.The correct instructions for Use for the MAX1 Pen is included with this notification. AtriCure will replace any affected product if requested by the facility.Please review and fill out the attached " Acknowledgement Form" and either fax to AtriCure at (513) 644-4753. If you need to contact customer service, their number is 1-866-349-2342. |  | 5 | 5 devices | Nationwide Distribution including the states of MD, PA and TX. | Terminated | June 11, 2013 | 148 | 5 | Undetermined
Z-0631-2013 | 63719 | No | LPL | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic | 510(k) | AVAIRA (enfilcon A) Soft (Hydrophilic) Daily Wear Toric Contact Lenses |  |  | November 09, 2012 | 2012 | Cooper Vision Caribbean Corp. | The lenses exceeded the acceptance limit for silicone oil residual. | Material/Component Contamination | The recall was conducted to the wholesale/distributor and Eye Care Practitioner level. The company contacted the Authorized Distributors by telephone, followed by written notification dated November 9, 2012. All written notifications to first consignees, for both revenue and Fit Sets, were sent by overnight mail with signature required as proof of receipt and the envelopes were marked "Urgent". The notifications described the issue and provided recommended actions. Customers may contact 800-341-2020 if questions. |  | 246 | 246 | US Nationwide, including Puerto Rico, and Canada. | Terminated | May 09, 2014 | 546 | 246 | Undetermined
Z-0914-2013 | 64464 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Lifeguard Infusion Set with Smartsite, 20Ga, 1, CLGY-2010 |  |  | January 21, 2013 | 2013 | Churchill Medical Systems, Inc. | The label on the device states that the item number is CLGY-2210 when the item number on the device label should read CLGY-2010. This is a 20Ga Huber needle set; however, the product label states the pouch contains a 22Ga set. | Under Investigation by firm | Vycon sent an Urgent Medical Device Recall letter dated January 21, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed check their stock for the affected product and immediately cease use and distribution and quarantine all affected product immediately. Count and document their affected inventory on the Recall Acknowledgement and Inventory Return Form. Customers were instructed to fax or email the form to Vygon including customer contact information. All affected product would be credited or replaced at no cost to the customer. Customers with questions were instructed to call 1-800-473-5414, or by e-mail at customerservice@vygonus.com.For questions regarding this recall call 603-743-5988. |  | 2340 | 2340 | Nationwide Distribution including MO, IL, FL, PA, VA, MO, NY, NE, NH, and OH. | Terminated | September 10, 2013 | 232 | 2340 | Undetermined
Z-1155-2013 | 64809 | No | JAI | Powered radiation therapy patient support assembly. | Radiology | Radiology | 510(k) | RoboCouch Patient Positioning System |  |  | March 27, 2013 | 2013 | Accuray Incorporated | A5 gearboxes responsible for pitch (head up and head down) and roll (patient left and right) motions of the RoboCouch manipulator may be defective. | Nonconforming Material/Component | Accuray issued an Urgent Advisory Notification letter dated March 27, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter stated that Field Service Engineers will be replacing the affected product. Questions and concerns are to be directed to Accuray Customer Support at 1-877-668-8667 |  | 16 | 16 devices are subject to correction. | Worldwide Distribution - USA Nationwide including the states of: OH, MO. OK, CA and VA and the countries of: Finland, France, Luxembourg, Poland, Russia, Turkey and Japan. | Terminated | August 20, 2013 | 146 | 16 | Undetermined
Z-0909-2013 | 64425 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Hickman 9.0 French Dual Lumen Catheter with SureCuff Tissue Ingrowth Cuff with Peel-Apart Percutaneous Introducer system (pediatric product code 0600580; adult product code 0600600). |  |  | February 21, 2013 | 2013 | Bard Access Systems | Bard Access Systems is recalling two lots of Hickman 9.0 French and one lot of Leonard 10.0 French Dual Lumen Catheters with SureCuff Tissue Ingrowth Cuff with Peel-Apart Percutaneous Introducer Systems due to the wrong size introducer being included with the kit. | Mixed-up of materials/components | Bard sent an Urgent Recall Notification letters dated February 2013 to all affected customers. The letter identified the affected product, problem, and actions to be taken. Customers were instructed to examine their inventory for the affected product, return any remaining in stock and complete the enclosed Reply Form and Inventory Reconciliation Form. For questions contact Bard Access Systems Customer Service at 1-800-290-1689. |  | 195 | 195 units. | USA Nationwide Distribution and the countries Europe, Belgium, Australia, and Latin America. | Terminated | January 03, 2014 | 316 | 268 | Undetermined
Z-1192-2013 | 64919 | applicator, absorbent tipped, non-sterile | KXF | Absorbent tipped applicator. | General Hospital | General Hospital | 510(K) Exempt | MediChoiceÂż Rayon Tipped OB/GYN Applicator, 8 inch, Plastic Shaft, Non-Sterile, Item Numbers WOD5001 and WOD5002. |  |  | March 26, 2013 | 2013 | Owens &amp; Minor Distribution, Inc. | Owens &amp; Minor Distribution, Inc. is conducting a market withdrawal of MediChoice Applicator OB-GYN 8 inch product that does not meet our quality standards because the rayon may loosen or unravel from the applicator during use. | Process change control | Owens&Minor; sent an Urgent: Medical Device Recall letter to the first set of customers on March 26, 2013 and to their second set of customers on April 9, 2013. The letter identified the affect product, problem and actions to be taken. Customers were instructed to examine their inventory and quarantine all affected products, contact service representative to discuss whether affected products should be returned or destroyed, and complete and return the Recall Reply Form. For questions call 804-723-7417. |  | 4298.5 | 4298.5 cases | US Nationwide Distribution. | Terminated | October 31, 2014 | 584 | 4298.5 | Undetermined
Z-0835-2013 | 64321 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | VACUETTE Safety Blood Collection Set + Luer Adapter, 23x12" Safety Blood Collect LA Product, Item number 450096. Usage: The Safety Blood Collection Set is used in routine venipuncture procedures. |  |  | February 01, 2013 | 2013 | Greiner Bio-One North America, Inc. | The VACUETTE Safety Blood Collection Set + Luer Adapter may display leaking (small droplets) during use due to a crack on the luer connector piece. | Other | Greiner Bio-One sent a notification letter dated January 31, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to stop distributing the affected product, complete the attached Product Disposition form and fax to GBO at 800-726-0052 or email to patech@gbo.com. GBO will contact customers for the return of the affected product after receiving the attached Product Disposition form. Customers were requested to provide the names of customers who may have received the product. For questions call 800-515-8112. |  | 192000 | 192,000 pieces (192 cases) | US Nationwide Distribution in the states of FL, IL, KY, NY, NC, PA, OH, TX, VA | Terminated | October 29, 2013 | 270 | 192000 | Undetermined
Z-0837-2013 | 64076 | No | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | Pro Time Microcoagulation System/ Pro Time 3 Cuvette, Pro3-4, Pro3-25, and Pro3-50. |  |  | September 28, 2012 | 2012 | International Technidyne Corp. | Some Pro Time3 Test Cuvettes within a specified lot range may recover lower than expected Prothrombin Time/International Normalized Ratio (PT/INR) results. | Process change control | ITC sent an 'Urgent Medical Device Recall' letter/Health Professional Account Tracking Form dated October 2, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to complete the form whether or not they have remaining inventory of these lots and return the form promptly. All unused returned product will be replaced. Any questions can be directed to ITC Technical Support at 800-631-5945 Extension 4707 or e-mail ProTime@itcmed.com.For questions regarding this recall call 800-631-5945, ext. 4707. |  | 269690256 | 11,742 boxes US (25 cuvettes per box); 22,968 boxes OUS (25 cuvettes per box) | Worldwide Distribution - USA (nationwide) and Internationally to Bermuda, Canada, Chile, Denmark, Germany, Italy, Latvia, Lithuania, Malaysia, Philippines, South Africa, Spain, Switzerland, Taiwan, and United Arab Emirates. | Terminated | January 09, 2014 | 468 | 269690256 | Undetermined
Z-1191-2013 | 64439 | No | MRU |  |  | Neurology | HDE - Humanitarian Device Exemption | Medtronic DBS Therapy for Dystonia Kit, models 3317, 3319, 3337 and 3339. Sterile and Non-Pyrogenic. |  |  | February 08, 2013 | 2013 | Medtronic Neuromodulation | There is a potential for lead damage due to the use of the lead cap provided in DBS and Dystonia kits. Medtronic has received reports of DBS leads being damaged at the connector end of the lead when the lead cap is used. The connector end of the lead is the end of the lead connected to the lead extension. Tightening or loosening of the setscrew may twist the setscrew connector block and may damag | Process control | Medtronic sent an Urgent Medical Device Correction letter dated February 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to follow the modified instructions if the lead cap is used during the implant procedure. For questions call Medtronic Neuromodulation Technical Services 1-800-707-0933. UPDATE: Medtronic mailed an Urgent: Medical Device Removal notification to customers on June 07, 2013, who may have unused DBS lead kits that were manufactured prior to the process change to correct the issue. On June 10, 2013 Medtronic representatives began visiting the identified accounts to facilitate removal of any unused product. The Medtronic Rep documented the hospital visits on an Account Specific Customer Confirmation Form and left a copy with the Hospital. |  | 30000 | 30,000 | Worldwide Distribution - USA Nationwide and the countries of: Algeria, Australia, Austria, Belarus, Belgium, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, El Salvador, Estonia, Finland, France, Germany, Greece, Hong Kong, Hungary, Iceland, India, Ireland, Israel, Italy, Japan, Jordan, Korea, Latvia, Lebanon, Lithuania, Luxembourg, Malaysia, Malta, Mexico, Morocco, Netherlands, New Zealand, Norway, Peru, Philippines, Poland, Portugal, Puerto Rico, Romania, Russian Federation, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Venezuela, Viet Nam. | Terminated | May 22, 2014 | 468 | 30000 | Undetermined
Z-0868-2013 | 64215 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Vascular Solutions, SuperCross Microcatheter, Sterilized using ethylene oxide, Rx Only, Model numbers 5300 - 5307, 5324, 5325, 5329, 5332, 5336, 5340, and 5341. |  |  | January 29, 2013 | 2013 | Vascular Solutions, Inc. | Vascular Solutions, Inc has discovered a potential problem with a number of lots of Supercross microcatheters. A potential for surface contamination on aged catheters was observed during in-house testing. | Material/Component Contamination | Vascular Solutions sent an "Urgent Medical device Recall" letter dated January 28, 29013 to all affected customers. The letter identified the affected product, described the problem and actions to be taken. The letter advised customers to remove the product from their inventory and to complete and return the VSI Account Inventory Form. For questions contact your local Sales Representative. |  | 4731 | 4731 | Worldwide Distribution - USA Nationwide including the states of AL, AZ, AR, CA, CO, CT, DE, CD, FL, GA, HI, ID, IL ,IN, IA, KY, LA, ME, MD, MA, MI,MN, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, SC, SD, TN, TX, UT, VT VA, WA, WV, WI, WY and AUSTRIA, CANDA, DENMARK, FINLAND, FRANCE, GERMANY, GREECE, HUNGARY, ISRAEL,. ITALY, MEXICO, NETHERLANDS, NORWAY, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, TURKEY, UNITED KINGDOM. | Terminated | June 25, 2013 | 147 | 4731 | Undetermined
Z-1171-2013 | 64851 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Newport Medical Instruments HT70 |  |  | April 05, 2013 | 2013 | Newport Medical Instruments Inc | Newport Medical Instruments is conducting a voluntary recall on certain Newport Medical Instruments HT70and HT70 Plus ventilator Power Pac batteries due to customer reports of Newport HT70 and HT70 Plus ventilators alarming and going to internal backup battery sooner than expected while the ventilator is being operated on Power Pac battery. | Component design/selection | Newport Medical Instruments sent an Urgent Medical Device Recall letter dated April 4, 2013, via FedEx to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers who sold the affected product to another party, or placed the affected product with a patient or healthcare provider were instructed to immediately inform the other party, patient, or healthcare provider of this action and provide them with a copy of the letter. If customers were using the affected product, they should ensure that an alternate source of power is available nearby. Customers were instructed to report any issues with Newport" HT70 and HT70 Plus ventilators or Power Pac batteries to the Technical Support group to ensure proper device reporting procedures are followed. Please call 800-451-3111, extension 500, for U.S. customers; and 714-427-5811, extension 500 for international customers and you will be given further instructions).For questions regarding this recall call 714-427-5811, ext 344.A press release was issued by the firm on 04/18/13. |  | 4392 | 2528 ventilators and 1864 power pac batteries | Worldwide Distribution - USA including US (AL, AR, AZ, CA, CO, CT, DE, FL, GA, lA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MT, NC, NH, NJ, NV, NY, OH, OK, OR, PA, Puerto Rico, RI, SC, TN, TX, UT, VA, WA, and WI. Internationally to Argentina, Australia, Bangladesh, Belgium, BoliVia, Brazil,Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Croatia, Ecuador, Egypt, France, Greece, Hong Kong, Hungary, India, Indonesia, Israel, Ireland, Italy, Japan, Kazakhstan, Kosova, Lebanon, Lithuania, Malaysia, Mexico, Nambia, Nepal, Norway, Oman, Paraguay, Peru, Philippines, Poland, Qatar, Saudi Arabia, Singapore, South Africa, South Korea, Sri Lanka, Switzerland,Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Venezuela, and Viet Nam. | Terminated | March 10, 2015 | 704 | 4392 | Undetermined
Z-1183-2013 | 64504 | Yes | FKQ | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Diacap(R) Ultra Dialysis Fluid Filter, Catalog Number 7107367. |  |  | March 01, 2013 | 2013 | B. Braun Medical, Inc. | There is the potential of the residual moisture of the Diacap Ultra membrane being reduced over time, which results in the reduction of the membrane permeability (ultrafiltration coefficient). | Process design | B. Braun Medical Inc. sent and "URGENT RECALL NOTICE" dated February 28, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the Customer Support Department at (800) 848-2066 for questions regarding this issue.This recall was expanded on/about June 18, 2013 to include additional lot #s 305010113 and 305000113. Customers were notified via certified mail on/about June 14, 2013. |  | 1580 | 1580 | Distribution and manufacturing dates for the expansion lots was reported as 5/23-6/14/2013 and 1/22-24/2013 respectively. Nationwide Distribution including the states of AZ, CA, DE, FL, GA, HI, IA, ID, IL, KS, KY, LA, MN, MO, NJ, NM, NY, OH, OR, PA, SC, SD, TN, TX, UT, WA and WV. | Terminated | September 03, 2014 | 551 | 1580 | Undetermined
Z-0840-2013 | 63692 | No | JOW | Compressible limb sleeve. | Cardiovascular | Cardiovascular | 510(k) | STRYKER Sustainability Solutions, AC Power Adaptor for Restep Deep Vein Thrombosis (DVT) System, an external compression system. PUMPKIT (Pump Kit), and RSP159-01 (AC Power Adaptor). |  |  | August 30, 2012 | 2012 | Stryker Sustainability Solutions | Reports that the AC Power Adaptor may not operate due to wires coming loose in the unit. | Process control | Stryker Sustainability Solutions sent a Urgent Medical Device Removal Notification letter dated September 5, 2012, to all affected customers. The letter identified the product, problem and the action to be taken by the customer. The firm notified that affected AC Power Adaptors would be replaced with improved Adaptors. Any affected lots would be removed and shipped back to Stryker Sustainability Solutions for final disposition. Customers were instructed to fill out the attached Recall Effectiveness Form even if no affected product was found and fax to Stryker Sustainability Solutions at (480) 809-4957 |  | 1311 | 1,311 | Distributed in Arizona and New Jersey. | Terminated | September 30, 2014 | 761 | 1311 | Undetermined
Z-0895-2013 | 62804 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Orthofix ISKD Limb Lengthener. ISKD - Intramedullary Skeletal Kinetic Distractor. |  |  | August 06, 2012 | 2012 | Orthofix, Inc | There is a possibility that the ISKD limb lengthener may stop distracting post-operatively during treatment, which may result in premature bone consolidation (limb not achieving the desired length) leading to revision surgery to remove and/or replace the device. | Nonconforming Material/Component | Orthofix sent an Urgent Product Recall Notification letter dated August 6, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue distribution of devices, identify the affected and remove from their inventory, and returned to the firm. Credit will be issued upon receipt of the returned product. Customers were asked to take the actions listed in the enclosed Product Return Instructions immediately. Customers with questions should contact Orthofix Customers Service by telephone at 800-266-3349 or 214-937-2047.For questions regarding this recall call 214-937-2509. |  | 1224 | 1224 | Worldwide Distribution - USA (nationwide) Internationally to Australia, Hong Kong, Saudi Arabia, Korea, Turkey, Greece, Finland, South Africa, Spain, Singapore, Switzerland, Brazil, Germany, France, Japan, Netherlands, New Zealand, Israel, Canada and Mexico | Terminated | May 17, 2013 | 284 | 1224 | Undetermined
Z-0866-2013 | 64202 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Merit Medical Systems Inc., Resolve Biliary Locking Drainage Catheter, Models RBC and RBDC, used for drainage of bile within the biliary system. |  |  | January 16, 2013 | 2013 | Merit Medical Systems, Inc. | A consumer complaint that a catheter was observed to be in several pieces under patient fluoroscopic visualization. | Nonconforming Material/Component | Merit initiated the recall on January 16, 2013. A voicemail script notification to users and Q&A; document was provided to those sales representative that had received and or distributed product. On January 16, 2013, Merit sales representatives began contacting each physician and/or user facility that had been provided product.Written communication distribution began on 22 January 2013. Affected sales representatives, physicians, and hospitals were provided written identification information, instruction to not distribute or continue placing these devices, and on how to return any unused product if applicable.All product on hand at Merit was inventoried and remains under quarantined status.Further questions please call (801) 208-4301, |  | 387 | 387 units | Worldwide Distribution - USA including the states of NJ, FL, GA, MO, AZ, UT and the country of Denmark | Terminated | September 16, 2015 | 973 | 387 | Undetermined
Z-0899-2013 | 64389 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Philips MX 16-Slice CT Scanner System, a whole-body computed tomography X-ray system. |  |  | December 19, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | The four screws that secure the patient table top to the carrier became loose, which caused the table top to detach. | Component design/selection | Philips Healthcare sent an Urgent - Medical Device Correction letter dated December 14, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to have BioMedical Engineering personnel inspect the patient table top screws that secure it to the carrier. Customers were instructed that If it is found that the table top screws are improperly seated, they should stop using the system and contact their local Philips Field Service Engineer for repair. A Philips Field Service Engineer will schedule a time to inspect and repair all the affected systems. For further information customers were instructed to contact their Philips Healthcare Customer Care Solutions Center at 1-800-722-9377 and follow the recorded menu option to reach a Customer Solutions. For questions regarding this recall call 800-722-9377, ext 5. |  | 412 | 412 Units | Worldwide Distribution - USA including AL, AR, CA, FL, GA, IL, IN, KY, MI, MS, NC, NJ, NV, NY, OH, PA, PR and VA. | Terminated | December 03, 2013 | 349 | 412 | Undetermined
Z-0649-2013 | 63896 | No | DTS | Cardiopulmonary bypass cardiotomy return sucker. | Cardiovascular | Cardiovascular | 510(K) Exempt | Sterile Rigid Tip Suction Wand |  |  | November 02, 2012 | 2012 | California Medical Laboratories Inc | California Medical Laboratories, Inc. has identified an issue with the Rigid Tip Suction Wand. Due to insufficient adhesive application, the distal tip has the potential to come loose from the wand and fall into the patient's chest cavity. | Nonconforming Material/Component | CalMed sent an Urgent Field Safety Notice dated November 2, 2012, to all affected customers by certified mail and additionally followed up via electronic mail. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to contact Sorin Group USA's Customers Service Department at +1-800-650-2623 to return any unused product and order replacements. Customers were also asked to communicate the important safety information to all potential users within their organization or any other organization where the affected product may have been transferred and to maintain awareness of this communication for an appropriate period of time to ensure effectiveness of the corrective action. Customers were asked to return all unused product to CalMed Labs. Customers with questions were instructed to contact California Medical Laboratories at +1-714-556-7365.For questions regarding this recall call 714-556-7365. |  | 1430 | S112462=995 units &amp; S111480 1430 units | Worldwide Distribution - USA including AZ, FL, CA, LA, CO, and MS.Internationally to Canada, Israel, Greece, Hong Kong, India, Ireland, Jordan, Lebanon, Malaysia, Netherlands, Russia, Saudi Arabia, South Africa and Spain. | Terminated | March 04, 2013 | 122 | 357891 | Undetermined
Z-0889-2013 | 64228 | No | NCD |  |  | Microbiology | PMA | Vacuette Cellestis 1 ml QuantiFERONÂż TB ANTIGEN (HA) High Altitude: Shake 10x; REF 454413. HE0396. greiner bio-one; 13x75 50pcs.,1200 pcs in total. |  |  | September 27, 2012 | 2012 | Cellestis Inc | Cellestis is recalling QuantiFERONÂżÂż - TB Gold TB Blood Collection Tubes and HA TB Blood Collection Tubes due to potential contamination. | Process design | Cellestis, inc., sent a "Notification of market withdrawal of QuantiFERONÂż" letter dated September 27, 2012 to its domestic and international customers. The letter described the product, problem and actions to be taken. The customers were informed to discontinue use of these products and any product sequestered such that it cannot be used or distributed. Cellestis will provide replacement materials, if and when a retest is required.Please do not hesitate to contact your local QuantiFERON sales representative or customer service representative at +613 8527 3500 or 661-775-7480 should you have any questions or concerns. |  | 29690 | 29,690 units | Worldwide distribution: USA (nationwide) and countries of: Argentina, Australia, Austria, Belgium, Bosnia and Herzegovina, Bulgaria, Czech Republic, Denmark, Estonia, Finland, France, Germany, Great Britain, Ireland, Israel, Jordan, Luxembourg, Mexico, Netherlands, New Zealand, Peru, Poland, Saudi Arabia, Slovakia, Slovenia, South Africa, Switzerland, and Turkey. | Terminated | April 25, 2013 | 210 | 339090 | Undetermined
Z-0887-2013 | 64407 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Aesculap Bipolar Generator Foot Pedal, Catalog No. GK226, for electrosurgical cutting and coagulation. |  |  | January 18, 2013 | 2013 | Aesculap, Inc. | The bipolar energy did not stop after release of the foot pedal. | Nonconforming Material/Component | Aesculap AG sent an Important Recall Notification letter dated January 21, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to examine inventory and return for a replacement product or credit. The letter included an inventory sheet to be completed by the customer and returned to the recalling firm. For question call 610-984-9265 or 610-984-9291. |  | 59 | 59 | USA Nationwide Distribution including the states of: AR, CA, FL, IN, IA, KS, MD, MN, MO, OR, TN, TX, VA and WA. | Terminated | January 08, 2014 | 355 | 59 | Undetermined
Z-1184-2013 | 64868 | Yes | DRO | External transcutaneous cardiac pacemaker (noninvasive). | Cardiovascular | Cardiovascular | 510(k) | Medi-Trace Cadence Adult Radiotransparent Defibrillation Electrode |  |  | April 08, 2013 | 2013 | Covidien LLC | Defibrillation Electrodes may arc/spark result in thermal damage to the leadwire, which could render the device incapable of delivering the appropriate energy or shock to the patient | Process control | Covidien sent an "URGENT MEDICAL DEVICE RECALL" letter dated April 8, 2013 to all affected customers. The letter identifies the product, problem and requests that customers immediately stop using and segregate any affected product from inventory. Questions:Call 1-800-962-9888 option 8, then extension 2500. |  | 133470 | 133470 (sets of 2) | Worldwide Distribution-USA (nationwide) and the countries of Korea, Japan, Canada, Australia, Brazil, Belgium, Singapore, and South Africa. | Terminated | November 19, 2015 | 955 | 140955 | Undetermined
Z-0839-2013 | 64288 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Maquet ELS Cannula 15 FR. Used for extracorporeal membrane oxygenation (ECMO) for the treatment of respiratory insufficiency in infants. |  |  | December 26, 2012 | 2012 | Maquet Cardiovascular Us Sales, Llc | ELS 15 Fr Cannula could potentially contain an occluded arterial lumen. | Nonconforming Material/Component | Maquet sent Urgent Device Recall Letters/Recall Acknowledgement Forms to their customers on 12/26/2012. |  | 257 | 41 units - US Distirbution; 216 units - OUS Distribution | Worldwide Distribution-USA including the states of AZ, CO, FL, TN, and TX and the countries of Australia, Canada, Czech Republic, Dubai/Bahrain/Kuwait/Lebanon/Saudi Arabia/Yemen, France, Germany, Hong Kong, Italy, Japan, Netherlands, Singapore, South Africa, Sweden, Switzerland, Thailand, Turkey, and UK. | Terminated | May 08, 2014 | 498 | 257 | Undetermined
Z-0841-2013 | 64328 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco RTP System. |  |  | February 11, 2013 | 2013 | Elekta, Inc. | The beam is displayed at the "gantry" angle, but the dose is calculated at the "rotation_start" angle. | Nonconforming Material/Component | Elekta provided customers with an "Important Safety Notice - Dose Display Not Matching Beam Geometry - Monaco RFC 35748". The notice identified the product, problem, and actions to be taken by the customers. Contact the form at 800-878-4267 for questions relating to this notice. |  | 13 | 13 | Nationwide Distribution-including the states of AK, CA, FL, NJ, NC, OH, OR, TN, TX, WA, and WI. | Terminated | January 21, 2014 | 344 | 13 | Undetermined
Z-0635-2013 | 63809 | No | JJW | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens Chek-Stix Positive and Negative Control Strips for Urine Chemistry |  |  | November 21, 2012 | 2012 | Siemens Healthcare Diagnostics Inc | When these lots of control material are reconstituted, the positive control solution leukocyte result may become negative if the solution is held for more than two hours above 25ÂżC (77ÂżF). | Nonconforming Material/Component | Siemens issued An Urgent Field Safety Notice via e-mail to all affected SiemensHealthcare Diagnostics Regional offices on November 21, 2012 both in the United States and foreign accounts. This notice informed customers of the issue with the CLINITEK AtlasÂż and Chek-StixÂż Positive Urinalysis Control Solutions When these lots of control material are reconstituted, the positive control solution leukocyte result may become negative if the solution is held for more than two hours above 25ÂżC (77ÂżF). You can use the reconstituted positive control solution for up to eight hours if the solution is held at 25ÂżC (77ÂżF) or lower. If the operating temperature is greater than 25ÂżC and lower than 30ÂżC (77ÂżF and 86ÂżF), the solution can be used for up to two hours.For questions regarding this recall call 781-269-3158. |  | 7865 | 7865 | Worldwide Distribution - USA (nationwide) and internationally to Canada, Aruba, Brazil, Denmark, India, Mexico, New Zealand, Singapore,Thailand, Bahamas, South Wales, and Hong Kong, | Terminated | November 06, 2014 | 715 | 37839 | Undetermined
Z-0642-2013 | 63030 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart FRx 861304 Automated External Defibrillator Manufactured by Philips Medical Systems, Seattle, WA 98121 USA; and Laerdal branded FRx 861305 |  |  | July 26, 2012 | 2012 | Philips Medical Systems North America Co. Phillips | Philips determined that it is important to clarify information in the Owner's Manual and keep customers informed about the maintenance of their FRx (models 861304 and 861305), OnSite (models M5066A and M5067A), or HeartStart Home (model M5068A) automated external defibrillators (AEDs) shipped between December 2005 and July 2012. Philips had determined there is a need to emphasize the importance o | Nonconforming Material/Component | Philips Healthcare sent an Emergency Care and Resuscitation letter Customer Information Letter dated July 26, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to contact their Philip's representative to obtain the tools and information they need to ensure that their device is functioning properly.For questions regarding this recall, please call AED Recall Support at 1-800-263-3342 (Option 6). |  | 201598 | 201,598 | Worldwide Distribution - USA (nationwide) and the countries of Japan and Malaysia. | Open, Classified |  |  | 814598 | Undetermined
Z-0898-2013 | 64235 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | REF CMS-CT6T225, 6F Triple Lumen CT PICC with M.S.T. Components |  |  | December 13, 2012 | 2012 | Medical Components, Inc dba MedComp | The affected product was packaged with the incorrect introducer needle. The kit label indicates a "safety" needle is included; the kits were packaged with a non-safety needle. | Nonconforming Material/Component | Medcomp sent a notification letter dated December 13, 2012 via email to all affected customers. The letter identified the affected product, problem and actions to be taken. For questions or concerns call 215-256-4201 |  | 785 | 785 | USA Nationwide Distribution including the state of PA. | Terminated | July 31, 2013 | 230 | 785 | Undetermined
Z-1139-2013 | 64671 | No | CKD | Alanine amino transferase (ALT/SGPT) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Alere Cholestech LDX ALT AST Test Cassette, Model #12-788. |  |  | March 06, 2013 | 2013 | Alere San Diego, Inc. | The recall was initiated because Alere San Diego has an update for the Alere Cholestech LDX ALT AST cassette, and it has the potential for humidity changes to impact results for certain analytes on the Alere Cholestech LDX System. | Other | Alere sent an "URGENT MEDICAL DEVICE RECALL UPDATE" letter dated March 6, 2013 to all affected customers. The letter informed the customers of the problem identified and the actions to be taken. Customers with any questions about the recall notification, are instructed to contact Alere San Diego 9975 Summers Ridge Rd, San Diego, CA 92121; (877) 308-8289; Fax: (858) 805-8457. Customers are instructed to complete the customer response form even if they do not have any involved product. |  | 38571 | 38,571 kits | Worldwide Distribution-USA (nationwide) and the countries of Canada, South Africa, Netherlands, Italy, Brazil, Singapore, Russia, Hong Kong, Argentina, South Korea, Philippines, El Salvador, Indonesia, and Trinidad. | Terminated | August 13, 2013 | 160 | 38571 | Undetermined
Z-0842-2013 | 64158 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | NEUROPATH 6Fx90 GUIDING CATH MPC |  |  | January 11, 2013 | 2013 | Codman &amp; Shurtleff, Inc. | Sterility of device may be compromised due to lack of packaging integrity | Component design/selection | Codman sent an Urgent Voluntary Product Recall Notification letter dated January 11, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to check all inventory to determine if they have any affected product, use the included instructions to report inventory status, return and obtain replacement product. Customers were instructed to contact your local Codman Representative or Stericycle at 1-866-201-9077.For questions regarding this recall call 508-880-8100. |  | 867 | 867 units | Worldwide Distribution - USA (nationwide) and Internationally to Austria, Belgium, Canada, Chile, Costa, Rica, Czech Republic, Denmark, France, Germany, Great Britian/UK, Ireland, Italy, Lithuania, Malaysia, Netherlands, New Zealand, Norway, Poland, Portugal, Singapore, Slovakia, Spain, and Switzerland. | Terminated | August 10, 2015 | 941 | 5911 | Undetermined
Z-1180-2013 | 64972 | No | FRC | Sterilization process indicator. | General Hospital | General Hospital | 510(k) | STERRAD CYCLESURE 24 Biological Indicator (BI) |  |  | April 22, 2013 | 2013 | Advanced Sterilization Products | Advanced Sterilization Products is recalling certain lots of STERRAD CYCLESURE 24 Biological Indicator product because it may not have adequate data to support the entire duration of the labeled shelf-life. | Under Investigation by firm | Advanced Sterilization Products sent an Urgent Medical Device Recall letter dated April 22 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers that have installed and validated STERRAD system between Feb 2012 and March 2013, are instructed to contact Stericycle at (888) 943-4901 to determine if their STERRAD System is eligible for an optional revalidation. If customers would like to have their system validation repeated, ASP will revalidate their system without charge. Customers are instructed that the revalidation can occur as part of a routine Planned Maintenance (PM) within 12 months. Customers are also notified that the optional procedure will require that their Sterrad System will be unavailable for use for approximately 31 hours (depending on their system) while revalidation is performed. Customers are instructed to contact Stericycle at the number above with their STERRAD System serial number by May 31 2013. Customers are instructed to contact Stericycle at (888) 943-4898 or contact their ASP representative. Stericycle's hours of operations are Mon-Fri , 8am-5pm, ET. |  | 69994 | 69,994 cases | Worldwide Distribution - Nationwide Distribution including the states of AR, AL, AK, AZ, CA, CT, CO, DE, FL, GA, HI, IA, IL, IN, ID, KY, LA, MA, ME, MD, MI, MS, MN, MO, MT, NC, ND, NE, NJ, NY, NV, NH, NM, OH, OR, OK, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI and WV., and the countries of AR, AU, BE, BR, CA, CL, CN, CO, DO, EC, GT, HK, ID, IN, JP, KR, MV, MX, PA, PE, PH, PR, RU, SG, TH, TW, UY, VE and VN. | Terminated | August 01, 2013 | 101 | 69994 | Undetermined
Z-1152-2013 | 64707 | No | DLZ | Phenobarbital test system. | Toxicology | Toxicology | 510(k) | VITROS Chemistry Products PHBR (Phenobarbital) Slides, REF 822 1384. |  |  | March 14, 2013 | 2013 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics (OCD) received customer complaints of intermittent imprecision, outliers, or shifts in accuracy and control fluid values outside of expected intervals when using VITROS Chemistry Products PHBR Slides. OCD's investigation confirmed that positively biased values for patient samples could be obtained that may not be detected by quality control fluids. | Unknown/Undetermined by firm | On 3/14/13, URGENT PRODUCT CORRECTION NOTIFICATION letters (dated 3/14/13), Questions &amp; Answers, and Response Forms were sent to consignees via FedEx courier or US Postal Service Priority Mail. Consignees were informed of the issue and advised to immediately discontinue using VITROS PHBR Slides. On 3/14/13, distributor letters were sent via US Postal Service Priority Mail to the distributors. They were advised of the issue and instructed to discontinue distributing and discard their remaining inventory of VITROS PHBR Slides. On 3/14/13, foreign affiliates were informed by e-mail of the issue and instructed to notify their consignees of the issue and actions. |  | 11758 | Domestic: 11,164 units, Foreign: 594 units | Worldwide Distribution-USA (nationwide) and Puerto Rico and the countries of Australia, Bermuda, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Republic of Panama, Singapore, Spain, and Venezuela. | Terminated | June 23, 2015 | 831 | 11758 | Undetermined
Z-0621-2013 | 63794 | No | BZO | Ventilator tubing. | Anesthesiology | Anesthesiology | 510(K) Exempt | Fisher &amp; Paykel Healthcare reusable breathing circuit. |  |  | November 19, 2012 | 2012 | Fisher &amp; Paykel Healthcare, Ltd. | FPH has become aware that one batch of tubes supplied to FPH may contain hole damage. These tubes were incorporated by FPH into certain 900MR068 reusable circuits with lot numbers 110810 and 111020. If this hole damage is not detected during the standard leak test before patient use, it could potentially result in a gas leak in the breathing system, which may lead to a loss of pressure for the int | Nonconforming Material/Component | Fisher &amp; Paykel sent an Urgent Medical Device Recall letter dated November 19, 2012, with a Product Recall Response Form to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to destroy and discard any affected circuits in their possession. Customers were instructed to complete and return the Product Recall Response form to the address indicated on the form, even if they have no affected product. Customers were asked to forward the recall notice to all those who need to be aware within their organization. Customers with questions should call 1-800-792-3912.For questions regarding this recall call 949-443-4095. |  | 200 | 200 units total, 26 units US | Worldwide Distribution - USA including IL, IN, AR and GA. | Terminated | July 11, 2013 | 234 | 200 | Undetermined
Z-1215-2013 | 64444 | No | JTY | Culture medium for pathogenic Neisseria spp. | Microbiology | Microbiology | 510(k) | BD ProbeTec Neisseria gonorrhoeae (GC) Qx Amplified DNA Assay Reagent Pack, REF 441124, contains 12 each Microwells labeled in part ***BD ProbeTec Neisseria gonorrhoeae (GC) Qx Amplified DNA Assay Reagent Pack REF 441124 ***For use with the BD Viper System *** Becton, Dickinson and Company 7 Loveton Circle, Sparks, MD 21152 USA www.bd.com/ds EC REP Benex Limited, Pottery Road, Dun Laoghaire, Co. Dublin, Ireland. |  |  | February 08, 2013 | 2013 | Becton Dickinson &amp; Co. | Increased levels of false positives in BD Probe Tech Neisseria gonorrhoeae Q Amplified DNA assay kits. | Environmental control | BD Diagnostic Systems sent an "URGENT PRODUCT RECALL" letter via UPS overnight delivery on February 8, 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to verify that these kits have been consumed, returned or destroyed by completing the enclosed form. The firm will issue a no charge replacement for customers' product. Any product remaining in customers' inventory should be discarded or destroyed per their facility procedures. Complete the attached form regardless if any inventory remains, so that the firm may acknowledge customers' receipt of this notification. It is not necessary to contact BD by telephone to receive replacements. Complete and fax the form to: Attention: Regulatory Compliance, Fax: 410-316-4258 OR E-mail to RegulatoryComplianceFax@BD.com. If further assistance is needed regarding replacements, please contact BD Customer Service Department at 1-800-675-0908. |  | 77 | 77 kits | Distributed in the states of CA, FL, GA, NJ, and TX. | Terminated | May 13, 2013 | 94 | 77 | Undetermined
Z-0878-2013 | 64294 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Versafitcup Impacting Ring |  |  | January 15, 2010 | 2010 | Medacta Usa Inc | Medacta USA is recalling Versafitcup Impacting Ring because the impaction ring may not fit properly in the implant shell in all cases. This could result in surgical delay potentially increasing possible complications under anesthesia, having to ream up to a larger shell size leaving less bone than planned and repositiong/removal of the shell after initial impaction. | Process control | Medacta USA sent notification letters dated 1/15/10 to Sales Representatives for customers that purchased the Versafitcup Impacting Ring. The letter informed the customers of the problem identified and the actions to be taken. Customers are informed that Medacta USA personnel will contact them by telephone to provide them with a RMA number. Customers are informed that they will receive a new set of redesigned/reworked impacting rings and a second near the end of the January 2010. Customers with questions are instructed to contact the firm at (805) 437-7085. |  | 293 | 293 units | Nationwide Distribution including the states of ID, CO, TX, WI, NY, CT, WA, CA, UT, and AZ. | Terminated | February 23, 2013 | 1135 | 293 | Undetermined
Z-0925-2013 | 64354 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | REF S313152 Straight Acetabular Inserter Handle Ringloc |  |  | February 05, 2013 | 2013 | Biomet, Inc. | The Straight Ringloc Button Latch Inserter Handle may have been manufactured at the higher end of the tolerance causing interference with the Ringloc cup. The inserter handle may lock onto the cup during impaction. If alternative inserters are not available this could cause a surgical delay greater than 30 minutes. | Process control | BIOMET sent an Urgent Recall Notice beginning February 13, 2013, to all affected customers. The notice included possible adverse events and directions to locate and remove identified devices by returning a Fax Back Response Form to 574-372-1683 and obtaining an RGA # prior to return of product to Biomet, Inc. 56 East Bell Drive Warsaw, IN 46580. Questions should be directed to 574 372-1570 M-F, 8 a.m.-5.p.m.For questions regarding this recall call 574-267-6639, ext 1676. |  | 58 | 58 | Worldwide Distribution - No USA distribution -Internationally to Canada, Japan, and Chile | Terminated | September 30, 2013 | 237 | 58 | Undetermined
Z-0927-2013 | 64348 | No | LSR | Treponema pallidum treponemal test reagents. | Microbiology | Microbiology | 510(k) | ANTI-BORRELIA (LYME) MICROPLATE EIA Kit, Catalog # 32507 Kit Lot Number 120960. Package Insert Supplement: Calibrator-1206 U/mL, Positive Control-228 - 532 U/mL, and Negative Control-&lt;120 U/mL; Manufactured by Innominata dba GenBio for Bio-Rad. The kit is used as an aid in diagnosis of Lyme disease. |  |  | February 11, 2013 | 2013 | Innominata Dba Genbio | Multiple incidents of low-reacting assay plates. | Nonconforming Material/Component | The firm, Innominata dba GenBio, sent an "Urgent Product Correction" letter dated February 11, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to review the listed questions to determine sites risk and take the appropriate actions. For questions regarding this issue, customers may contact Bio-Rad Technical Support at : 1-800-224-6723, option #2, and then option #3 for assistance. |  | 195 | 195 | Nationwide distribution: CA, IL, MA, MD, ME, NC, ND, NJ, NY, OR, VA and WI. | Terminated | August 27, 2013 | 197 | 195 | Undetermined
Z-0921-2013 | 64421 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | 3.75mm x 12mm Fixed Cervical Screws. Components of the Genesys Spine Anterior Cervical Plate System |  |  | February 20, 2013 | 2013 | Genesys Orthopedic Systems, LLC | Manufacturing residue may be present on the cervicalscrews. Use of this product may result in patient infection, delayed fusion or non-fusion (pseudarthrosis), inflammatory response, or revision surgery. | Employee error | Genesys Orthopedic Systems, LLC sent a Recall Notification letter dated February 20, 2013 to all affected customers. The letter identified the affected product, problem and necessary actions to be taken. Customers were instructed to immediately return the affected product to the Genesys Spine office and complete and return the enclosed Response Form as soon as possible. For questions call 512-381-7070. |  | 9 | 9 cervical screws | USA Nationwide Distribution including the states of: CA and TX. | Terminated | September 09, 2013 | 201 | 9 | Undetermined
Z-0920-2013 | 64318 | No | LBK |  |  | Neurology | 510(k) | Integra CUSA excel. CEM Nosecone for CUSA EXcel System. |  |  | January 07, 2013 | 2013 | Integra LifeSciences Corp. | There is a potential for erosion of the CUSA EXcel tip used with the Nosecone. Specifically, the potential for tip erosion occurs when using a CUSA CEM Nosecone if the side of the tip, rather than the annulus of the tip, is in contact with the tissue. This erosion can potentially lead to tip breakage. There were no reports of a patient injury associated with these complaints. | Nonconforming Material/Component | Integra sent a Urgent Medical Device Recall Field Safety Notices dated January 7, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed of the nature of the issue and how if affects CUSA top erosion/breakage; advises customers and provided a copy of the CEM Nosecone IFU additional Warning; and requess customers to complete and return the Acknowledgement Return Form.Please feel free to contact our service hotline at ( 888 ) 772-7378. |  | 29450 | USA 6,190, EU 22,972 and ROW 288 units | Worldwide Distribution - USA and the countries of Austria, Australia, Bangladesh, Belgium, Bulgaria, Brazil, Switzerland, Canada, China, Costa Rica, Czech Republic, Germany, Dominican Republic, Egypt, Emirates, Spain, Finland, France, Great Britian, Hong Kong, Hungary, Indonesia, Ireland, Israel, India, Italy, Japan, Korea (South), Kuwait, Lebanon, Sri Lanka, Latvia, Macedonia, Mexico, Malaysia, Netherlands, Philippines, Pakistan, Poland, Portugal, Qatar, Serbia, Saudi Arabia, Singapore, Slovenia, Thailand, Tunisia, Taiwan, United Arab, Viet Nam and South Africa. | Terminated | August 06, 2013 | 211 | 29450 | Undetermined
Z-0658-2013 | 62950 | No | DRM | External cardiac compressor. | Cardiovascular | Cardiovascular | 510(k) | AutoPulse Resuscitation System Model 100, |  |  | August 08, 2012 | 2012 | Zoll Circulation, Inc. | Complaints associated with an unexpected stop in AutoPulse compressions. | Employee error | Zoll sent Customer Notification letters dated August 8, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact Zoll Technical Support Line at 1 (800) 348-9011 for further assistance. |  | 5727 | 3881 in US, 1846 ROW. | Worldwide distribution. | Terminated | May 12, 2014 | 642 | 5727 | Undetermined
Z-1212-2013 | 64885 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Optima MR430s MRI Scanner; 5000-0002 |  |  | March 28, 2013 | 2013 | GE Healthcare, LLC | Coolant supply lines may leak liquid onto the floor resulting in a slippery surface and fall hazard. | Process control | GE Healthcare sent an Urgent Medical Device Correction letter dated April 23, 2013, to all affected customers. The letter informed customers of a potential safety issue due to leakage from gradient coolant lines. Consignees were instructed to monitor the floor for leakage and continue to use if leakage is not observed. If leakage is observed consignees are instructed to stop using the system and notify their local service representative. GE Healthcare will correct all affected scanners by providing an update at no cost.Customers with questions were instructed to contact their local service representataive.For questions regarding this recall call 262-513-4122. |  | 81 | 81 | Worldwide Distribution - USA including AL, CA, CO, FL, ID, IL, MA, NH, NJ, NY, OH, PA, SC, SD, TN, TX, UT, VA, WA, WI; Worldwide: Argentina, Australia, Austria, Brazil, Canada, Chile, Denmark, Egypt, Finland, Germany, India, Italy, Japan, Korea, Malaysia, Netherlands, Norway, Peru, Poland, Russia, Singapore, South Africa, Spain, Sweeden, Switzerland, UK, Ukraine, and UAE. | Terminated | October 25, 2013 | 211 | 302 | Undetermined
Z-0926-2013 | 64256 | No | OXO | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Medtronic O-arm Imaging System |  |  | January 31, 2013 | 2013 | Medtronic Navigation, Inc. | Potential failure of the braking system that controls the O-arm Imaging System gantry movement in the vertical direction | Component design/selection | Medtronic Navigation issued a Field Safety Notice letter dated January 18, 2013 to all affected customers. The letter identified the affected product, problem, and actions to be taken. The letter instructed customers not to discontinue use and that a Medtronic Navigation Representative will be contacting them shortly to schedule a time to conduct the motion control box upgrade. For questions contact your local Medtronic Navigation representative or call Technical Services Group at 800-595-9709 or (720) 890-3200, or email at rs.navtechsupport@ medtronic.com. |  | 463 | 463 units | Worldwide Distribution - USA Nationwide and the countries of: Canada, AUSTRIA, BELGIUM, BRAZIL, BULGARIA, CHILE, CHINA, CROATIA, CZECH REPUBLIC, DENMARK,FINLAND, FRANCE, GERMANY, GREECE, INDIA, IRELAND, ISRAEL, ITALY, JAPAN ,KOREA, REPUBLIC OF KUWAIT, LEBANON, LIBYAN ARAB JAMAHIRIYA, MEXICO, NETHERLANDS, New Zealand, POLAND, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SLOVAKIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, Taiwan(R.O.C.) , THAILAND, TURKEY, UNITED ARAB EMIRATES. and UNITED KINGDOM. | Completed |  |  | 463 | Undetermined
Z-0660-2013 | 63850 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | PT113950 PT HYBRID Glenoid Post |  |  | December 06, 0012 | 12 | Biomet, Inc. | Biomet is recalling Part Number PT-113950 PT Hybrid Glenoid Post, following an investigation which identified that the male thread of the post may be oversized. This oversized condition can vary in degree and may lead to the following three events:1) If the PT Hybrid Glenoid Post is not fully seated into the Hybrid Base and the implant construct isimplanted, then a gap of 1-3 mm will be prese | Nonconforming Material/Component | Biomet Inc. sent a Urgent Medical Device Recall Notices dated December 6, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The notice identified the reason for the recall including directions to immediately locate, discontinue use of the device. Surgeons were asked to continue monitoring of patients implanted. Customers were asked to confirm receipt of the notice by calling 1-888-348-9500 extension 1570. Customers with product to be returned are asked to obtain RGA # by calling 574-372-6677 . A FAX Back Response Form is to be filled out and sent within 3 days of receipt to 574-372-1683. Biomet, Attention Return Goods, Building B Recalls, 56 East Bell Drive, Warsaw, IN 46580. Questions related to the notice should be directed to 574-371-1570, Monday- Friday 8 a.m. to 5 p.m. |  | 1050 | 1050 | Worldwide Distribution -- US (nationwide) including the countries of , Canada, Japan, Argentina, Australia ,Netherlands, Europe and Chile. | Terminated | November 08, 2013 | 730822 | 1050 | Undetermined
Z-1151-2013 | 64596 | No | KXT | Digoxin test system. | Toxicology | Toxicology | 510(k) | Siemens Dimension Vista(R) System Vista Drug 4 Calibrator. The device is used as an in vitro diagnostic product. |  |  | February 21, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Siemens has confirmed complaints of low QC and patient recovery with the Dimension Vista(R) LOCI(R) Digoxin(DIGXN) Flex(R) reagent cartridge when calibrating DIGXN with either Drug 4 Cal (KC460) lot 2KD052 or Drug 4 Cal (KC460A) lot 2KD053. Internal testing has shown an average 15% low bias in the therapeutic range of 0.9 to 2.0 ng/mL. | Process control | The firm initiated their recall of this product by sending an Urgent Medical Device Recall letter dated February 2013 to all Dimension Vista Drug 4 Calibrator customers who received the affected lots. The letter notified them of the issue, the potential risk to health, and actions to be taken by customers. |  | 1638 | 1,638 | Distributed within the US to AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PR, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY. | Terminated | January 07, 2014 | 320 | 1638 | Undetermined
Z-0922-2013 | 63551 | No | NBH | Arthroscope. | Orthopedic | Orthopedic | 510(K) Exempt | Endotrig Trigger Release Sterile Micro Hook Blade |  |  | September 18, 2012 | 2012 | Instratek, Incorporated | Instratek, Inc.has received multiple complaints of the 1052 Trigger Finger Release Blades which could not be advanced into the cannula through which it is used during trigger finger release procedures. | Nonconforming Material/Component | The firm decided to recall the product and sent notification letters to consignees dated 09/18/2012. |  | 235 | 235 | USA Nationwide Distribution including the states of: AZ, MI, NE, NV, and TX. | Terminated | April 10, 2013 | 204 | 235 | Undetermined
Z-0819-2013 | 64017 | No | GNX | Influenza virus serological reagents. | Microbiology | Microbiology | 510(K) Exempt | Quidel, Sofia Influenza A + B FIA, Kit #20218 (25-Test Kit), in vitro diagnostic. |  |  | December 03, 2012 | 2012 | Quidel Corporation | Quidel Corporation has initiated a voluntary recall of specific lots of the Sofia Influenza A+B FIA kits because of false positive results which may lead to misdiagnosis. | Nonconforming Material/Component | Quidel Corporation sent a Medical Device Recall dated December 3, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their inventory for the affected product and remove the affected product from their inventory. Customers should contact Quide Corporation for instructions on how to obtain replacement product. Customers were instructed to fax the enclosed inventory Form to Quidel at 858-552-7905.For technical support customers should call 800-874-1517 or 858-552-1100.For questions regarding this recall call 858-552-1100, ext 7922. |  | 13090 | 13,090 kits (327,250 tests) | Worldwide Distribution - USA (nationwide) and internationally to Korea, Norway, Germany, Chile, Hong Kong, Switzerland, Netherlands, Thailand, Japan, China, and the UK | Terminated | January 31, 2014 | 424 | 13090 | Undetermined
Z-1140-2013 | 64499 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Duane Reade TRUEtrack monitor kit, UPC# 6 39194 01679 7 Part # A4006-81 |  |  | February 25, 2013 | 2013 | Nipro Diagnostics, Inc. | Meter kit outer boxes containing 10 ct. test strips state incorrect storage temperature range. | Environmental control | NIPRO Diagnostics sent a Voluntary Product Field Action Notice dated August 12, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check with all Duane Reade retail locations and Duane Reade Disgtribution Center to determine if they have any of the affected product ath their location. If product is located customers were instructed to call 1-800-803-6049 for a Return Authorization Number and instructions for returning the affected product.For questions regarding this recall call 954-677-9201. |  | 1627 | 1,627 Kits | Nationwide Distribution only in NY. | Terminated | July 08, 2013 | 133 | 1627 | Undetermined
Z-1131-2013 | 64848 | No | KST | Vacuum-assisted blood collection system. | Hematology | Hematology | 510(K) Exempt | Heritage Labs Path Study Extra Supplies Kit, Product # A 2376, kit contains 3 zippered plastic bags which hold support supplies for health care providers when treating patients. Bag #3 contains 5 different blood collection tubes (5 each) and 2 different gauge needles--all for drawing blood. One of the tubes is the recalled BD blue top citrate tube (5 tubes). These tubes were manufactured by BD, Franklin Lakes, NJ. |  |  | February 06, 2013 | 2013 | Heritage Labs International LLC | The product has a Lithium Heparin contaminant which could adversely affect coagulation results potentially causing erroneous lab results. | Material/Component Contamination | The recalling firm contacted 25 of their 26 customers by phone on February 5, 2013, to notify them of the recall and sent a follow up memo on February 6, 2013, During the call, the firm reminded their customers how many recalled kits they recieved and instructed them to go through the kits and pull the recalled lots. The recalled kits are to be returned to the recalling firm and the firm sent a follow-up memo on February 6, 2013, with the shipping label their customers to place on the returning kit shipping box. This memo was sent to the customers by overnight Fed-ex mail and the ship label for the returned product was an overnight UPS label. On February 7, 2013, the final customer was notified by phone and instructed to return the recalled kits.For further questions please call (913) 764-1045. |  | 25 | 25 kits | Nationwide Distribution including the states of AL, CA, GA, IL, IN, KY, MA, MD, NC, NY, OK and VA. | Terminated | November 25, 2013 | 292 | 25 | Undetermined
Z-1148-2013 | 64790 | No | KTN | Infectious mononucleosis immunological test system. | Immunology | Immunology | 510(k) | IMO-402 Mononucleosis Rapid Test Device (Whole Blood/Serum/Plasma). Catalog #IMO-402/94201 |  |  | February 26, 2013 | 2013 | Alere San Diego, Inc. | Alere San Diego is recalling the Clearview Mononucleosis Cassette because the positive control included in the rapid test may produce false negative results. | Material/Component Contamination | Alere sent an Urgent Medical Device Recall letter dated February 26, 2013, to all affected customers. The letter informs the customers of the problems identified and the actions to be taken. Customers with unused product remaining are instructed to immediately discard it. Investigations of customer complaints indicate that the positive control included with this lot may produce false negative results. Customers are instructed to return the attached Verification form. Customers that have consumed all product are instructed to indicate zero under "Quantity Discarded" and complete and fax back the form. Replacement product will be provided for any unused portion of the affected lot based on completion and return of the attached Verification Form. Customers with any questions are instructed to contact Alere Technical Services at (866) 216-0094 or by email at later.flow.support@alere.com.For questions regarding this recall call 858-805-2000, ext 3015. |  | 78 | 78 kits | Nationwide Distribution including TN, WA, UT, KS, KY, MN, GA, MA, PA, NY, VT, and the Dist. of Columbia | Terminated | August 29, 2013 | 184 | 78 | Undetermined
Z-0919-2013 | 64306 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Command Instrument Broach Handle. Used in Total Hip Arthroplasty. Catalog Number: 6266-0-100 |  |  | December 20, 2012 | 2012 | Stryker Howmedica Osteonics Corp. | Reports of handle fracture and subsequent disassociation of cam assembly subcomponents during use. | Process control | Stryker Orthopaedics sent an "URGENT MEDICAL DEVICE RECALL NOTIFICATION" dated December 20, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at (201) 972-2100 for questions regarding this issue. |  | 48 | 48 units | Nationwide Distribution. | Terminated | September 30, 2015 | 1014 | 48 | Undetermined
Z-1150-2013 | 64781 | No | FEB | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | EVOTECH Endoscope Cleaner and Reprocessor System, P/N 50004. |  |  | March 28, 2013 | 2013 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) is recalling the EVOTECH Endoscope Cleaner and Reprocessor (ECR) because they have received reports of smoke emanating from the system due to a circuit board malfunction that may occur in the Channel Control Manifold sub-assembly. | Under Investigation by firm | An Urgent Medical Device Field Correction letter dated 3/28/13 was sent to all customers who purchased the EVOTECH Endoscope Cleaner and Reprocessor (ECR). The letter informed the customers of the problems identified and the actions to be taken. Customers are informed that the failure can potentially cause the following outcomes for the cycles listed below:1. In the Full Reprocessing Cycle (Cleaning and HLD): there is a potential for residual rinse water and/or alcohol to remain in the endoscope channels after the cycle completes. 2. In the HLD-Only Cycle: there is potential for an adequate amount of high-level disinfectant solution to be sent through some endoscope channels.If customers observe smoke, then they are instructed to:a. Power OFF the device.b. Report your incident to ASP Professional Services at (888) 783-7723 and select option #2.c. Do not use the device until it has been serviced by ASP.Customers are instructed to return attached buyer return card acknowledging their receipt of this communication. Customers are instructed to contact Stericycle at (866) 761-9566 for any questions related to the field correction. |  | 761 | 761 units | Nationwide Distribution | Terminated | August 22, 2013 | 147 | 761 | Undetermined
Z-1153-2013 | 64638 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Nova StatStrip Xpress Glu-Test Strips (Glucose test strips) |  |  | February 25, 2013 | 2013 | Nova Biomedical Corporation | Glucose test strips reports no result message when tested | Process design | The firm, Nova Biomedical, sent an "Urgent Field Safety Notice 01-13SS" dated March 18, 2013 and US customers notification by phone contact and to the International customers including a Faxback Form to be communicated to and sent to Nova Biomedical Subsidiary by two methods of delivery (Email and Phone Call). The customers were instructed to INVALIDATE the Nova StatStrip Glucose Test Strip lots indicated in the list on all StatStrip meters within your institute; REPLACE all Nova StatStrip Glucose Test Strips from the indicated lots with replacement test strips provided by Nova; REMOVE all unused boxes of the Nova StatStrip Glucose Test Strip lots indicated from all locations within your institute; RETURN all unused boxes of the Nova StatStrip Glucose Test Strips lots indicated to Nova Biomedical U.K. or your local Nova distributor following the country specific instructions provided by your authorized Nova representative, and complete and return the Fax-Back Form via email or FAX to your Regional Service Manager at 781-8914806.On April 08, 2013, the firm expanded the recall notification to include one additional test strip lot #0313021309 which was shipped under two catalog numbers(Catalog # 42214 and Catalog# 51493). US Notification by phone contact with key customer contact facilitated by Senior Manger of Technical Support Services. International - Field Safety Notice and Faxback Form is being communicated to and sent to Nova Biomedical Subsidiary by two methods of delivery (Email and Phone Call).Questions regarding this notice should be directed to the following Nova Biomedical EU Authorized Representative at: Nova Biomedical Innovation House, Aston Lane South Runcorn,WA7 3FY Telephone: + 44 1244 287087Email: Europe@novabiomedical.co.uk. or US at (781) 894-0800. |  | 7433 | 7433 cases | Worldwide distribution: USA (nationwide) including states of: CA, IL, KY,NY, WV and countries of: United Kingdom and Italy. | Terminated | September 26, 2014 | 578 | 9633 | Undetermined
Z-0814-2013 | 63662 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, Dash Port 2 Docking Station. |  |  | June 13, 2012 | 2012 | GE Healthcare, LLC | GE has become aware of multiple issues affecting the Dash 3000/4000/5000 and DashPort 2monitor. This report of correction encompasses the DashPort 2 device, in which the DashPort 2 stopscommunicating with the Dash, resulting in Loss of Monitoring at the CIC. The DashPort 2 status LEDsin the back of the unit indicate Unity Network and DashPort 2 to Dash communication has stopped. | Nonconforming Material/Component | GE Healthcare sent an Important Product Information Letter dated June 13, 2012 via FedEx to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter indicated that the motherboard would be replaced, on request, should the customer continue to experience the issue. A second letter dated August 30th 2012 was issued informing consignees that the correction that was initially proposed did not fully correct the issues. Therefore, the motherboard replacement field action had been suspended. It further explained that the correction will be provided at no charge when it is available. For questions contact Technical Support at 1-800-558-7044. |  | 10497 | 10,497 | Worldwide Distribution and US (nationwide) including DC and the countries of ALGERIA, ARAB EMIRATES, AUSTRALIA, AUSTRIA, BANGLADESH, BELGIUM, BRAZIL, CANADA, CHINA, CYPRUS, CZECH REPUBLIC, DENMARK, EGYPT, FRANCE, GERMANY, INDIA, ISRAEL, ITA LY, JAPAN, JORDAN, KUWAIT, MOROCCO, NETHERLANDS, NEW ZEALAND, NORWAY, POLAND, PORTUGAL, RUSSIA, SAUDI ARABIA, SINGAPORE SLOVENIA, SOUTH KOREA, SPAIN SWEDEN SWITZERLAND, TAIWAN, TURKEY, UNITED KINGDOM. | Open, Classified |  |  | 10497 | Undetermined
Z-1146-2013 | 64229 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | Pressure Monitoring sets with VAMP Plus closed blood sampling system model # T001691A, PXVP0566, PXVP2260, PXVP0826, PXVP0827, PXVP2284AT3, PXVP23X3AT3, T398807D, T100605B, T001762A, T391T00A, VP2. |  |  | January 30, 2013 | 2013 | Edwards Lifesciences, LLC | The device is being recalled because the firm has received multiple complaints related to incorrect assembly that may result in dilution of blood samples. | Process change control | The firm initiated a recall in Europe via customer letters starting January 30, 2013 to notify the recall of model number T001691A lot number 59326333. The recall was expanded and a worldwide recall via customer letters was initiated on February 11, 2013. The letter identified the product, problem, and actions to be taken by the customers. |  | 30467 | 30,467 | Worldwide Distribution-USA (nationwide) including the states of AR, CA, FL, HI, ID, IN, KS, KY, LA, MA, MI, MN, MO, NC, ND, NM, NV, NY, OH, OK, PA, SD, TN, TX, VA, WA, and WI) and Europe. | Terminated | March 10, 2016 | 1135 | 30467 | Undetermined
Z-0923-2013 | 64477 | No | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | Genesis Straight Esthetic 4.0 mm 1mm Cuff Abutment |  |  | July 05, 2011 | 2011 | Keystone Dental Inc | Mislabeled: Genesis Straight Esthetic 5.0mm Contour ,3mm cuff labeled as Genesis Straight Esthetic 4.0 mm, 1 mm cuff abutment, | Mixed-up of materials/components | Keystone issued Market Withdrawal notice via fax on July 5, 2011, to affected customers. The notice identified the product, the problem, and the action to be taken by the customer. Customers were asked to complete and return the fax-back form to facilitate whether they hae any affected product and indicate whether they would like to receive a replacement of a full refund. Customers ewre asked to contact Keyston'es Customer Relations Department to coordinate shipment of the returned product. For questions regarding this recall call 781-328-3390. |  | 35 | 35 units | Nationwide Distribution including CA, IN, MA, MD, MN, MI, , NH, NY, NC, PA, and WI. | Terminated | March 11, 2013 | 615 | 57 | Undetermined
Z-1248-2013 | 53487 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Mercury Spinal System DIA 5.5 x 50 MM Lordosed Rod, |  |  | April 17, 2009 | 2009 | Spinal Elements, Inc | Spinal Elements has initiated a recall for the Mercury Spinal System due to a raw material issue. | Nonconforming Material/Component | Spinal Elements sent Urgent Medical Device Recall notification letters to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the Chief Technology Officer at Spinal Elements for questions about the recall letter. |  | 68 | 68 units total | Distributed in OH and PA. | Terminated | May 07, 2013 | 1481 | 68 | Undetermined
Z-1267-2013 | 64962 | No | DWS | Cardiovascular surgical instruments. | Cardiovascular | Cardiovascular | 510(K) Exempt | Codman(R) Collins Radioparent Sternal Blade Nylon,4 3/4" (121 mm) long, 3" (76 mm) wide, 1 1/2" (38 mm) deep, REF 50-8081 |  |  | April 10, 2013 | 2013 | Symmetry Medical/SSI | There is a possibility that the Collins Radioparent Sternal Blades Nylon may break during use. | Under Investigation by firm | Symmetry Surgical sent an Urgent Medical Device Recall letter dated April 8, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their inventory immediately to determine if they have any of the affected product. If so, customers were instructed to discontinue use of the product immediately and promplty call 800-251-3000 for a return material authorization (RMA) number and to identify if they would like credit or an alternative product. Customers were asked to fill out the attached Recall Response Form and fax the completed form to 615-964-5567. Customers with questions were instructed to call 615-964-5515.For questions regarding this recall call 615-964-5515. |  | 5240 | 524 boxes of 10 | Worldwide Distribution - USA including CA, MS, FL, PA, TN, CO, IL, TX, OH, MI, MA, IL and internationally toBelgium, Canada, Germany, UK | Terminated | May 12, 2014 | 397 | 5240 | Undetermined
Z-1145-2013 | 64840 | No | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | ACIST Medical Systems Inc. |  |  | March 15, 2013 | 2013 | Acist Medical Systems | ACIST Medical Systems is issuing a voluntary recall notice for three specific lots of BT2000 Automated Manifold Kits. For these specific lots, ACIST has received field reports related to ingress of air into the manifold body. This has been detected by the Air Column Detect Sensor halting the injection, and normally replacement of the BT2000 Automated Manifold Kit is necessary in order to proceed w | Nonconforming Material/Component | ACIST Medical Systems sent an Urgent Medical Device Recall letter dated March 15, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letter advised customers that if they had inventory of any of the lots under the scope of the recall, to discontinue use or further distribution of the product. If you possess inventory containing any of the lot numbers listed above, please discontinue use or further distribution of these products. Please return this inventory to ACIST. Upon receipt of the returned BT2000 Automated Manifold Kits, ACIST will arrange for a shipment of replacement kits at no charge. To order additional or immediate replacement product, please contact BRACCO customer service at 1-877-BRACCO9 (1-877-272-2269), option 1.Please also complete and return the enclosed Medical Device Recall Form (see enclosure). Even if you do not possess any of the affected product in your inventory, the form needs to be completed and returned, indicating that you have no remaining units containing these lot numbers in your inventory. |  | 6640 | 6640 units | Nationwide Distribution including the states of AL, AK, AZ, AR, CA, CO, CT, DC, GA, HI, ID, IL, IN, IA, KS, KY, MD,MI, MN, MS, MO, NV, NH, NJ, NM, NY,NC, OR, OK, OH, PA, TN, TX, VA and WA. | Terminated | August 07, 2013 | 145 | 6640 | Undetermined
Z-1244-2013 | 64705 | No | MQF | Assisted reproduction catheters. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Wallace Sure-ProÂż Single Stage Embryo Replacement Catheter |  |  | March 28, 2013 | 2013 | Smiths Medical ASD, Inc. | Unable to pass catheter through outer sheath during embryo transfer | Process control | The firm, Smiths Medical, sent an "URGENT FIELD SAFETY NOTICE" dated March 28, 2013 to its customers. The notification describes the product, problem and actions to be taken. The customers were provided with two options for proceeding: 1) continue to use product if the product passes a pre-test; or 2) return affected product to Smiths Medical for replacement or credit. In addition, the customers were instructed to inspect their inventory and quarantine all affected products; perform a count of affected products currently in inventory; complete and return the attached Confirmation Form via fax to +1-866-258-6068 or email to pamela.harrell@smiths-medical.com; send a copy of this notice to your customers identified having received affected product; and provide email confirmation of your customers notifications to Smiths medical at pamela.harrell@smiths-medical.com. Customers should report any issues with these products to Smiths Medicals Global Complaint Department at +1-866-216-8806 or globalcomplaints@smiths-medical.com. If you should have any questions regarding this information, please contact Smiths Medical's Customer Service Department at +1-800-258-5361. |  | 800 | 800 units | Worldwide distribution: USA (Nationwide) and countries of: Arab Emirates (AE) Ireland (IE), Albania, Israel (IL), Argentina (AR),Italy (IT) Armenia (AM) , Japan (JP), Austria (AT) Jordan (JO), Bahrain (BH) Kuwait (KW) Belgium (BE), Lebanon (LB), Belarus , Netherlands (NL), Brazil (BR),Nigeria (NG), Bulgaria (BG) , Pakistan (PK), Canada (CA) Poland (PL), China (CN) Romania (RO), Columbia (CO), Russia (RU), Cyprus(CY) , Saudi Arabia (SA) Czech (CZ) , Singapore (SG), Denmark (OK), South Korea (KR), Egypt (EG), Spain (ES), Finland (FI), Sri Lanka (LK),Germany (DE) Turkey (TR) Ghana (GH) , Uganda (UG), Great Britain (GB) Ukraine (UA), Greece (GR) Uruguay (UY), and Hong Kong (HK). | Completed |  |  | 136349 | Undetermined
Z-1000-2013 | 64589 | No | EAP | Intraoral source x-ray system. | Dental | Radiology | 510(k) | CareStream Dental CS 2200 X-ray system |  |  | September 18, 2012 | 2012 | Trophy Sas | There have been approximately 2% of the x-ray generators manufactured between February 2012 and November 2012 returned due to errors. These generators showed the type of defect that indicates poor contact of HT transformer power pin. The defect is a bent connector associated with the generator. | Radiation Control for Health and Safety Act | Trophy, A Carestream Dental Company, - A Subsidiary of Carestream Health Inc. The Center for Devices (CDRH) understands that the corrective action plan (CAP) involving 167 units will consist of implementing rework procedure TL008, which describes the procedure for a bending the metallic strip the female connector to insure strong connection with the male counterpart.The CAP appears to adequately address the problem and is hereby approvedFor further questions please call (585) 781-1997. |  | 167 | 167 installed in the US | USA Nationwide Distribution | Open, Classified |  |  | 167 | Undetermined
Z-0761-2013 | 64226 | No | JAC | Radiographic film marking system. | Radiology | Radiology | 510(K) Exempt | GE Healthcare Brivo XR385 DR-F/ X-Ray System |  |  | August 03, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has become aware of a software issue on the interface of the Brivo XR385 DR-F/ Model 5271997 radiographic product related to Visual indication as mandated by the US Code of Federal Regulations (21CFR) and IEC 60601-1-3. | Radiation Control for Health and Safety Act | GE Healthcare will, without charge, remedy the issue or bring the product intocompliance with each applicable Federal and IEC standard in accordance with a planto be approved by the Secretary of Health and Human Services, the details of which will be included in a subsequent communication to you or through a GE Healthcare field engineer site visit. A GE Healthcare Service Representative will update the software on the system to address the issue.Contact Call Center phone number: 800-437-1171 (United States). For other countries, please contact your local GE Healthcare Service Representative if you have any questions or concerns regarding this notification |  | 9 | 9 units installed in US | Nationwide Distribution including NM, WY, FL, MO, MA, and MI. | Terminated | December 05, 2013 | 489 | 9 | Undetermined
Z-0930-2013 | 64231 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | ACUSON S2000 Automated Breast Volume Scanner (ABVS); Model number 10042541 |  |  | January 18, 2013 | 2013 | Siemens Medical Solutions USA, Inc. | On some ACUSON S2000 ABVS systems, the inside of the transducer pod may contain sharp edges. | Process design | Siemens sent a Customer Safety Advisory notification letter to Siemens Service Personnel on January 18, 2013 and which would be provided to all affected sites during their visit to correct the problem. The notification identified the product, the problem and the action to be taken. The local Customer Service Engineer will perform the test at their site. If the system passes the test, no further action is required. If the ABVS pod fails the inspection, the Service Engineer will correct the sharp edge.For questions regarding this recall call 650-694-5398. |  | 214 | 214 affected devices | Worldwide Distribution - USA including WV and OH; Internationally to Australia, Austria, Bahrain, Belgium, Brazil, China (P.R.) Estonia, France, Germany, Greece, India, Indonesia, Israel, Italy, Japan, Jordan, Kuwait, Lebanon, Macedonia, Malaysia, Mexico, Netherlands, Peru, Philippines, Poland, Portugal, Qatar, Rep of Korea, Romania, Sweden, Taiwan, Thailand, Turkey, Unite Arab Emirates, United Kingdom and Venezuela. | Terminated | July 16, 2013 | 179 | 214 | Undetermined
Z-1147-2013 | 64770 | No | DXG | Single-function, preprogrammed diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | HCS3011 AV Loop Kit For use with Transonic CO status System Only, -HCS3021 AV Loop Kit For use with Transonic CO status System Only. |  |  | March 28, 2013 | 2013 | Transonic Systems Inc | Transonic Systems Inc. have received complaints of blood leakage involving the AV Loop Kit. | Component design/selection | Transonic Systems Inc. sent an Urgent: Medical Device Recall dated March 27, 2013 to all affected customers. The letter identified the recalled product, problem and actions to be taken. Customers were instructed to examine inventory, quarantine, discontinue use and return any recalled product via FedEx. Transonic Systems Inc. has offered to reimburse the firm for the return of these recalled products by check or credit memo for the returned goods and postage. Transonic Systems Inc. has requested that if a firm has further distributed any of the above lots, that they immediately contact their accounts, inform them of the recall, and have them return their outstanding recalled stocks to you. Transonic Systems Inc. have included copies of the product labeling and a Recall Return Response Form in their packet. Completed Recall Return Response Forms should be faxed to 607-257-5310. |  | 195 | 195 units total (140 units Domestic, 55 units Foreign) | Worldwide Distribution - USA Nationwide including the states of AZ, CA, FL, GA, NM, NY, TX, VA and WA and countries of Netherlands, Canada and Russia. | Terminated | March 05, 2014 | 342 | 195 | Undetermined
Z-1981-2013 | 65833 | No | JDG | Hip joint femoral (hemi-hip) metallic cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Moore Hip Prosthesis Long Fenestrated Stem, several head Diameter, 203 mm Stem length. |  |  | July 16, 2013 | 2013 | Zimmer, Inc. | Zimmer recently conducted a review of historic packaging validations completed. Based on this review, it was determined that a subset of product packaged at a single Zimmer facility required testing to ensure that previously packaged product from this facility met the established requirements for sterile barrier integrity. This testing determined that a single packaging tray configuration, which | Packaging process control | Zimmer is removing the devices from the field including distributor and hospital inventories. A notification was issued to known users of the product on 07/16/2013. |  | 1710 | 1,710 units are in distribution | Worldwide Distribution-USA (nationwide) including the states of CT, FL, GA, IL, IN, KS, KY, MA, MI, MN, MT, NC, NJ, NY, OH, OR, PA, SC, TN, TX, UT, VI, and WI, and the countries of Singapore, Canada, Mexico, Arab Emirates, Saudi Arabia, and Hong Kong. | Terminated | April 30, 2015 | 653 | 1710 | Not_Computer
Z-1983-2013 | 65785 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | GYRUS ACMI, REF 3740, Everest BiCOAG Bipolar Dissecting Forceps 5mm/45 cm, Sterile EO. |  |  | July 24, 2013 | 2013 | Gyrus Medical, Inc | On April 29, 2013 Olympus Surgical Technologies America R&D; received results from an accelerated aging study conducted as part of a new market submission. The study found that one of 60 Dissector units that underwent 3 year accelerated aging failed the dye penetration test indicating that the seal on sterile packaging was compromised. The dissector products are labeled with a 3 year shelf life. | Packaging process control | Olympus sent a "Medical Device Recall" letter dated July 24, 2013, to all affected customers. The letter identified the problem the product and the action needed to be taken by the customer.. Customers were advised to immediately cease any further use of the product and remove it from inventory and quarantine it until is shipped back to the company. To return the product, call 1-888-524-7266 Olympus customer service representative. Customers were requested to complete and return by fax the enclosed questionnaire to 484-896-7128. If the product was further distributed then they need to identify their customers and notify them at once of the product recall. |  | 1140 | 1140 | Worldwide Distribution - USA (nationwide) including the states of AL, AZ, CA, FL, GA, IL, IN, KS,KY, MD, MA, MI, MN, MS, MO, NY, NC, OH, PA, SD, TN, TX, VA and WI., and the countries of BRAZIL, CHINA, AUSTRALIA, CHINA, GERMANY and INDIA. | Terminated | June 20, 2014 | 331 | 1140 | Not_Computer
Z-1001-2013 | 64585 | No | OJU | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Sepp, Compound Benzoin Tincture Packaged and Tincture 3,000 Bulk. |  |  | March 04, 2013 | 2013 | CareFusion 213, LLC | An internal review of CareFusion's labeling for preoperative skin prep product codes was recently conducted. The labeling of the product codes identified above state that the contents of the package are sterile and does not clarify that this sterility claim is only applicable to the applicator. The antiseptic solution that is applied to the skin is a cutaneous disinfectant, and is not sterilized. | Labeling False and Misleading | CareFusion sent an Urgent Recall Notice dated March 4, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immediately check inventory, quarantine and return the affected product for replacement. For questions call (913) 345-3504. |  | 13749000 | 13,749,000 | USA Nationwide Distribution | Terminated | May 12, 2014 | 434 | 13749000 | Not_Computer
Z-0109-2014 | 66222 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Vented Spike Adapter Product Code 2C0471, Sterile, nopyrogenic. Manufactured by an affiliate of Baxter Healthcare Corporation, Clintec Nutrition Division, Deerfield IL 60015 USA. |  |  | August 30, 2013 | 2013 | Baxter Healthcare Corp. | Potential tears in the pouch of six lots of Vented Spike Adapter Product Code 2C0471 | Packaging | The firm, Baxter, sent an "Urgent Product Recall" letter dated September 10, 2013 to its consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed to locate and remove all affected product from your facility; If you distributed the products to other facilities or departments, forward them a copy of the letter; complete and return the attached Customer Reply Form via fax to 1-224-270-5457 or scan and email to fca@baxter.com; and contact Baxter Healthcare Center for Service at 1-888-229-0001 between hours of 7:00 am and 6:00pm Central Time to return affected product and receiving credit.For additional information contact the Center for One Baxter at 1-800-422-9837, Monday thru Friday during the hours of 8:00am to 5:00pm Central Time. |  | 49344 | 49,344 units | Nationwide distribution. | Terminated | April 19, 2017 | 1328 | 49344 | Not_Computer
Z-1938-2013 | 65698 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | C-QUR V-Patch Mesh (All sizes and shapes). |  |  | July 19, 2013 | 2013 | Atrium Medical Corporation | Additional Instructions for Use and Storage Conditions as Coated mesh can adhere to the inner packaging liner due to exposure to high humidity conditions | Package design/selection | Atrium Medical issued Recall Letter via UPS and e-mail on 7/19/13 to the accounts and field representatives. The notification identifies the problem, product, and risk factors. If the Product is exposed to excessive humidity for an extended period of time, then the increased humidity occurring inside the pouch can potentially cause the coating on the mesh to strongly adhere to the inner handling sleeve. A reply for is requested to be completed to acknowledge receipt to the notification.Additional language to the instructions for use (IFU) to include:Prolonged exposure to high humidity may result in increased rate of adherence of the C-QUR mesh to its handling sleeve. Store in a Controlled Room Temperature (25 C / 77 F ) or less. Brief exposure to up to 40 C (104 F ) is acceptable.Questions please contact Atrium Medical Customer Service at 1- 800-528-7486 Monday through Friday 9:00 am to 5:00 pm EDT. |  | 95286 | 95,286 units | USA (nationwide) including Puerto Rico and the countries of Australia Austria Bahrain Brazil Canada Chile Colombia Dominican Republic Ecuador El Salvador France Germany Great Britain Greece Honduras Hong Kong India Ireland Israel Italy Japan Jordan Korea Malaysia Mexico Netherlands New Zealand Nicaragua Norway Panama Peru Portugal Romania Saudi Arabia Singapore South Africa Spain Sri Lanka Switzerland Taiwan Thailand Turkey and Venezuela. | Terminated | March 18, 2016 | 973 | 145251 | Not_Computer
Z-0987-2013 | 63946 | No | LOM |  |  | Microbiology | PMA | ARCHITECT HBsAg Qualitative Assay, List Number 4P53 and ARCHITECT HBsAg Qualitative Confirmatory Assay, List Number 4P54 |  |  | December 18, 2012 | 2012 | Abbott Laboratories, Inc | Abbott has identified that assay files for the ARCHITECT HBsAg Qualitative and HBsAg Qualitative Confirmatory assays (LNs 4P53 and 4P54) on the ARCHITECT i1000SR System were mistakenly made available on abbottdiagnostics.com in October 2012. Additionally, some i1000SR customers received notification through AbbottLink in November 2012 that assay files for LNs 4P53 and 4P54 were available. The AR | No Marketing Application | The firm decided to correct the issue and and sent out Urgent Field Safety Notices to their customers dated 10/18/2012. |  | 12 | 12 | Nationwide Distribution including AZ, CA, FL, GA, HI, LA, MA, MD, MO, NJ, NY, PA, RI, TN, TX, UT, and WA. | Terminated | July 31, 2013 | 225 | 12 | Not_Computer
Z-0111-2014 | 66524 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Single Site Curved Needed Driver |  |  | October 02, 2013 | 2013 | Intuitive Surgical, Inc. | A limited number of FDA-cleared instruments for da Vinci Single Site use, Bipolar Maryland and curved needle Driver, were shipped without the correct revision of the user manual which provided added instruction for use of these additional Single Site instruments including indication, warnings and precautions. | Labeling Change Control | The firm, Intuitive Surgical, sent an " Urgent Medical Device Correction" letter dated October 2, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure all affected personnel are fully informed of this notice; after receiving the updated manual, locate and discard all previous versions of the Single Site Instruments manual and replace with the new manual; complete and return the attached Acknowledgement Form via FAX to Intuitive Surgical, Inc., ATTN: REGULATORY COMPLIANCE, U.S. (408) 716-3040, or Scan and Email to: isi.compliance@intusurg.com.; Inform affected personnel when the correction has been completed; and Retain a copy of this Notice and the Acknowledgement Form for your files.If you have any questions, contact your Clinical Sales Representative or contact Intuitive Surgical Customer Service at the numbers listed below:" North and South America: 800-876-1310 Option 3 (6 am to 5 pm PST)" Japan: 0120-56-5635 or 003-5575-1362 (9 am to 6 pm JST)" South Korea: 02-3271-3200 (9 am to 6 pm KSTJ)" Europe, Middle East, Asia and Africa: +800 0821 2020 or +41 21 821 2020 (8 am to 6 pm CET). |  | 1118880 | 504 devices for a total of 2,220 affected clinical product. | Worldwide distribution: US (nationwide) and countries of:Australia, Belgium, France, Italy and United Kingdom. . | Terminated | September 04, 2014 | 337 | 1118880 | Not_Computer
Z-0875-2013 | 64322 | No | NUC | Condom. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | KY TINGLING JELLY Personal Lubricant. The product is packaged into 5 oz tubes with the UPC printed on it. Distributed By: McNeil-PPC Inc. |  |  | January 25, 2013 | 2013 | Mc-NEIL-PPC, Inc. | This product is not cleared for marketing in the US. | No Marketing Application | Johnson and Johnson Healthcare Products sent recall notification letters dated 1/25/2013 to its customers (retailers and wholesalers). The letter identified the product, problem, and actions to be taken by the customers. Customers were asked to return all cases of affected product in their possession. The letter included delivery confirmation and a return packing slip, return authorization label and business reply card. Contact J &amp; J Sales &amp; Logistics at 866-912-8414 for questions regarding the return of affected product. Contact your J &amp; J Sales &amp; Logistics Company Sales Representative for questions about the product. |  | 425304 | 425,304 units | Nationwide Distribution. | Terminated | April 11, 2014 | 441 | 2361232 | Not_Computer
Z-0101-2014 | 65920 | No | GKR | Automated hemoglobin system. | Hematology | Hematology | 510(k) | HemoCueÂż Hb 201 Microcuvettes, carton/4 boxes/25 pouches of Microcuvettes, HemoCue AB Angelholm, Sweden |  |  | July 25, 2013 | 2013 | HemoCue AB | HemoCue AB has become aware of a production related problem involving punctured HemoCue Hb 201 Microcuvettes single pack pouches. | Packaging | On 8/13/2013 the firm sent Field Safety Notices to their customers. |  | 169320 | 42,330 packages/4 boxes/25 microcuvetes | Worldwide distribution which includes the following states: AK, AL, AS, AZ, CA, CO, CT, FL, GA, IA, IL, IN, KS, KY,MD, ME, MI, MN, MO, NC, ND, NE, NJ, NM, NY, OH, OK, OR, SC, SD, TN, TX, UT, VA, WA, WI, VT, WV &amp; WYand the countries of Armenia, Australia, Austria, Belgium, Cyprus, Czech Republic, Denmark, El Salvador Finland, France, Germany, Great Britain, Guatemala, Hong Kong, Israel, Ivory Coast, Netherlands, New Zealand, Norway, Peru, Poland, Russia, South Africa, Sweden, Switzerland and Taiwan. | Terminated | August 11, 2014 | 382 | 169320 | Not_Computer
Z-0338-2014 | 66763 | No | DXZ | Intracardiac patch or pledget made of polypropylene, polyethylene terephthalate, or polytetrafluoroethylene. | Cardiovascular | Cardiovascular | 510(k) | Vascu Guard, Peripheral Vascular Patch with Apex Processing. |  |  | October 28, 2013 | 2013 | Synovis Surgical Innovations, Inc. | Synovis, a subsidiary of Baxter International Inc., is recalling certain shipments of VASCUGUARD and DURA-GUARD because Synovis, a subsidiary of Baxter International Inc., is recalling certain shipments of VASCUGUARD and DURA-GUARD because the shipping configurations in which the 6 fluid ounce units containing a specific revision of the Instructions for Use (IFU) shipped in a divided cell shipping | Packaging change control | A customer notification letter dated 10/28/13 was sent to all customers who purchased the Vascu Guard and the Dura Guard. The letter informs the customers of the problem identified and the actions to be taken. Customers with any questions regarding the recall are instructed to contact Customer Service unit at (800) 423-2090. If customers are distributors or if they have have distributed the recalled products to any other health care provider, they are instructed to notify their accounts and provide them with a copy of the recall notification. If customers would like Baxter to perform the recall task for them, then they are instructed to contact Stericycle at 1-866-912-8417. |  | 6014 | 6014 units | Nationwide Distribution. | Terminated | June 03, 2014 | 218 | 6460 | Not_Computer
Z-0098-2014 | 66362 | No | KWI | Elbow joint radial (hemi-elbow) polymer prosthesis. | Orthopedic | Orthopedic | 510(k) | Small Bone Innovations (SBi) rHead Radial Stem, rHead Recon Radial System, rHead Radial Stem Lateral Assembly, rHead Radial Stem 6mm Collar, rHead Recon Radial Stem 6mm Collar, rHead Radial Stem Lateral Assembly 6mm Collar, rHead Standard Extended Stem, rHead Radial Stem Plasma Coated, rHead Recon Radial Stem Plasma Coated, rHead Radial Stem 6mm Collar Plasma Coated, rHead Recon Radial Stem 6mm Collar Plasma Coated, uHead Ulnar Stem Standard, uHead Ulnar Stem 20mm Collar packaged in a sterile pouch inside a shrink-wrapped box. |  |  | September 06, 2013 | 2013 | Small Bone Innovations, Inc. | The recall was initiated due to an increased risk of implants breaching the sterile pouches inside the shrink-wrapped box. | Packaging change control | Email correspondence and a notification letter, dated September 6, 2013, were sent to users requesting the immediate return of affected devices. Another notification letter, dated December 20, 2013 was also sent for recall of the additional devices after the recall was expanded. |  | 2036 | 2036 devices | Worldwide Distribution-USA (nationwide) and the countries of Canada and Mexico. | Terminated | March 30, 2015 | 570 | 2036 | Not_Computer
Z-1987-2013 | 65678 | No | KDQ | Vacuum-powered body fluid suction apparatus. | General Hospital | General Hospital | 510(K) Exempt | PleuraGuide Disposable Chest Tube Kit; Product codes: 17100, 17150, and 17200. |  |  | July 16, 2013 | 2013 | Atrium Medical Corporation | Individual component's IFUs are not included with the kit. | Labeling Change Control | Notice of correction was sent on July 18, 2013 via UPS and e-mail to the field representatives that represent these accounts. The notification identifies the problem, product, and risk factors. It provides instructions how to handle the product that may be affected. Additional labeling will be added to the kit to include instructions for use for the additional components. Atrium Medical will confirm through tracking of the UPS letters that each consignee has received the letter of correction. This shall be 100%. Where a UPS delivery cannot be confirmed, Atrium Medical will make good faith effort to contact the consignee to ensure that the letter was received. |  | 6516 | 6,516 units | Nationwide Distribution. | Terminated | July 15, 2014 | 364 | 6516 | Not_Computer
Z-1928-2013 | 65370 | No | MBL | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | R3 (TM) XLPE ACETABULAR LINER UHMWPE, 20o OVERHANG, 36 MM ID, 62 MM OD, QTY: (1), STERILE H2O2, REF 71335762. |  |  | May 29, 2013 | 2013 | Smith &amp; Nephew Inc | One batch of 60 mm 20o R3 XLPE liner may be laser etched and packaged as a 62 mm 20o liner and one batch of 62 mm 20o R3 XLPE liner may be laser etched and packaged as a 60 mm 20o liner. | Packaging process control | All affected Smith &amp; Nephew Sales Representatives were notified of problem and the recall via email on 5/29. They were instructed to cease use and distribution of the product and quarantine for return. Surgeons were notified by letter beginning on 6/12. International consignees were notified on 05/29 via email. |  | 53 | 53 units | Worldwide Distribution-USA including the states of PA, WA, TX, FL, IA, MI, WI, TN, KY, IN, NY, MO and OR, and the countries of India, Canada, Chile, Venezuela, Dubai, South Africa, Australia, Poland, Romania, Belgium, Germany and United Kingdom. | Terminated | March 22, 2016 | 1028 | 53 | Not_Computer
Z-0028-2014 | 66325 | No | MDB | Microbial growth monitor. | Microbiology | Microbiology | 510(k) | BBL (tm) MGIT (tm) Mycobacteria Growth Indicator Tube, part # 245122, packaged in 100 ct cartons and labeled in part |  |  | September 09, 2013 | 2013 | Becton Dickinson &amp; Co. | An in-vitro diagnostic reagent may be labeled with an incorrect barcode, which when scanned on associated diagnostic instrument may cause a false negative result. | Error in labeling | BD sent an Urgent Product Recall letter dated September 2013 to customers identified as microbiology lab managers. The letter identified the affected product, problem and actions to be taken. The letter advised that referenced lot numbers may contain an overlabel barcode that may begin to lift, exposing the original incorrect barcode pictured in the notification. Customers were instructed to complete the attached form and fax to Regulatory Compliance 410-316-4258. For further assistance regarding replacements contact BD Customer Service Department at 1-800-675-0908 and all other inquiries contact BD Technical Services Department at 1-800-638-8663. |  | 187000 | 187,000 tubes | USA Nationwide Disatribution in the states of: AL, AR, AZ, CA, CO, CT, FL, GA, IL, KY, MA, ME, MN, MO, NC, ND, NJ, NV, NY, OH, PA, TX, UT, VA, WA, WI | Terminated | October 29, 2013 | 50 | 187000 | Not_Computer
Z-0876-2013 | 64303 | No | EJF | Orthodontic appliance and accessories. | Dental | Dental | 510(K) Exempt | Avex CX Suite, Avex CX 0.22 UL4,5 HK REF #224-017, |  |  | February 08, 2013 | 2013 | Ultradent Products, Inc. | Orthodontic brackets distributed was incorrectly labeled. | Packaging | Ultradent Products, Inc. sent an Urgent Medical Device Correction letter dated February 8 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were asked to return affected products for replacement. Customers with questions can call firm representative, Corey Jaseph at (801) 553-4220. |  | 124 | 124 | USA Nationwide Distribution including the states of CA, AZ, CO and WA. | Terminated | June 07, 2013 | 119 | 124 | Not_Computer
Z-1013-2013 | 64327 | No | LTW | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | bioMerieux VITEK 2 AST-ST01 Streptococcus Susceptibility Card, REF 410 028. |  |  | February 06, 2013 | 2013 | Biomerieux Inc | The product package did not contain the correct version of the package insert listing the limitations of use and performance requirements. | Error in labeling | bioMerieux sent an "PRODUCT SAFETY CORRECTION NOTICE" dated February 13, 2013 to the affected customer. The letter identified the product, problem, and actions to be taken by the customer. Contact the firm at (800) 682-2666 for questions regarding this notice. |  | 4 | 4/20-card cartons | Distributed only in HI. | Terminated | October 30, 2013 | 266 | 4 | Not_Computer
Z-0051-2014 | 65969 | No | HMX | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | Synergetics 25 gauge Soft Tip Cannula, packaged in 6 units per box. Catalog number: 30.08.25. |  |  | July 30, 2013 | 2013 | Synergetics Inc | One lot of 25 gauge Soft-Tip Disposable Cannulas are co-mingled with 23 gauge cannulas; the product is labeled as 25 gauge cannulas. | Labeling Change Control | The firm sent recall notification letters to their consignees on 7/30/13. |  | 48 | 48/6-cannula boxes | Nationwide Distribution-USA including the states of MO, CO, KY, TN, NC, PA, IL CA, NY, IN, MI, LA, NJ and OH. | Terminated | April 07, 2014 | 251 | 48 | Not_Computer
Z-0787-2013 | 64028 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Exactech NOVATION CROWN CUP, NEUTRAL LINER, Use with Crown Cup Only, GXL UHMWPE, SN 0123456, REF 130-32-53. |  |  | November 05, 2012 | 2012 | Exactech, Inc. | Recall of Novation Crown Cup Connection GXL Neutral Liners Catalog No. (130-32-53) and Novation Crown Cup Connection GXL + 5mm Lateralized Liners Catalog No. (136-32-53). Exactech discovered the device labeling was inadvertently transposed. | Labeling mix-ups | Exactech, Inc. notified consignees by telephone on 11/05/12 of the affected devices and sent an "IMPORTANT PRODUCT RECALL NOTICE" letter dated 11/06/12. The letter describes the purpose of the recall, the product affected, the problem and actions to be taken by the consignees. A Product Recall Response Form was attached for the consignees to contact the Exactech inventory representative within 5 business days to confirm quantities of product in their inventory and return them to Exactech. Consignees with any questions relating to these devices should call 1-800-392-2832. |  | 26 | 26 Units | Distributed in the states of TN, FL, VA, WA and NV. | Terminated | December 24, 2013 | 414 | 26 | Not_Computer
Z-0086-2014 | 65918 | No | KQO | Automated urinalysis system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | HemoCueÂż Urine Albumin Microcuvettes, 25 Microcuvettes HemoCue AB Angelholm, Sweden |  |  | July 16, 2013 | 2013 | HemoCue AB | HemoCue AB has become aware of a production related problem involving punctured HemoCueÂż Urine Albumin Microcuvettes single pack pouches. | Packaging | HemoCue sent an Urgent Field Safety Notice dated July 24, 2013. to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed to :1. Quarantine any HemoCueÂż Urine Albumin (110608); lot numbers 1208848 to 1304847 you may have in your possession2. Return the Field Safety Notice Verification Form to: fax: 562-668-5794We apologize for the inconvenience this advisory has caused you. Please feel free to contact Technical Support at 800-426-7256 with any questions. |  | 517800 | 10,356 boxes/50 microcuvettes | Worldwide Distribution - Nationwide Distribution including the states of AL, AR, AZ, CA, CA ,CT, FL, GA, HE, ID, IL, IN, LA, MA, ME, MI, MN, MO, MS, NC, NH, NM, NY, OH, PA, SC,TX, UT, VA, VT, WA and WI., and the countries of Australia, Bahrain, Belarus, Canada, Czech Republic, Denmark, Finland, France, Germany, Great Britain, Hong Kong, India, Indonesia, Italy, Kenya, Lesotho, Malawi, Mauritius, Mexico, Myanmar, Nigeria, Norway, Omen, Poland, Qatar, Romania, South Africa, South Africa, South Korea, Spain, Sweden, Taiwan, Tanzania, Ukraine, Yemen and Zambia. | Terminated | April 28, 2014 | 286 | 517800 | Not_Computer
Z-0983-2013 | 63951 | No | FSM | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Needle Counters; 12.2 x 10.5 x 3.5 cm; 40 Count/70 Capacity; Sterile; Removable Numbered Foam Block; Adhesive Area (Cat. No. 3FA70SSA) |  |  | December 18, 2012 | 2012 | Cardinal Health | Small holes (approximately 1 mm or less) found in device unit level packaging potentially compromising product sterility. | Package design/selection | Cardinal Health sent an Urgent Product Recall letter dated December 18, 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immedaitely stop using affected product and discard any product that may be on hand, return the enclosed acknowledgement form, notify any customers to whom you may have distributed product affected by this recall, and contact the appropriate Customer Service group to arrange for credit of any affected product discarded. For questions contact Cardinal Health Professional Services at 800-292-9332. |  | 54912 | Cat. No. 3FA70SSA: 23,808; Cat. No. 3FM70SSA: 54,912 | Worldwide Distribution - USA (Nationwide) and the countried of Australia, Canada and Japan. | Open, Classified |  |  | 316752 | Not_Computer
Z-0890-2013 | 64316 | No | FRO |  |  | General &amp; Plastic Surgery | 510(k) | Atrapro Antipruritic Hydrogel, 1.5 oz gel tube, 4 Net wt/ 1.5 oz.. |  |  | February 12, 2013 | 2013 | Oculus Innovative Sciences Inc | Out of specification result - stability test failure of free available chlorine (FAC) which functions as a preservative for the hydrogel. The 13 month test result obtained was 4 ppm FAC and the stability specification is &gt;10 ppm FAC. | Incorrect or no expiration date | Oculus sent an Urgent Medical Device Recall letter dated February 12, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately discontinue distribution and segregate the affected product. Immediately remove the affected lots of Recall product from the inventory of sales representative, sub contracted distribution and or fulfillment centers. Segregate the affected product in a secure location for return or destruction. If the product was further distributed customers were instructed to notify healthcare professionals of this recall. Customers were also instructed to completed and return the Business Reply Card and fax to 1-415-462-5163. Once the firm receives the affected product the customer would be contacted by an Oculus Representative and provided a Returns Goods Authorization (RGA) number.For questions customers were instructed to call 1-707-559-2445.For questions regarding this recall 707-559-7191, ext 14. |  | 56567.5 | 56,566 1.5 oz tubes | Nationwide Distribution - One distributor in PA | Terminated | November 25, 2013 | 286 | 56567.5 | Not_Computer
Z-0336-2014 | 66670 | No | DSY | Vascular graft prosthesis. | Cardiovascular | Cardiovascular | 510(k) | Vascutek Gelsoft Vascular Graft (peripheral) |  |  | October 23, 2013 | 2013 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (Terumo CVS) became aware of a customer complaint that a 15mm x 8mm Gelsoft graft was received which the customer believed was a 15mm x 6mm Gelsoft graft unit. The manufacturer confirmed a batch of product code 631508 was incorrectly labeled. | Error in labeling | Terumo CVS called affected consignees (via telephone)on 10/23/2013 of this device removal. Terumo CVS advised users to discontinue use and return all affected grafts in inventory. Terumo will replace or issue credit for returned grafts. Terumo provided consignees with the reason of the removal, potential hazard, and description of affected product. Terumo addressed any questions from the consignees. Consignees with questions were instructed to contact Terumo CVS Customer Service at 800-521-2818.For questions regarding this recall call 734-741-6173. |  | 9 | 9 | US Distribution in the states of: NC and CA. | Terminated | February 27, 2014 | 127 | 9 | Not_Computer
Z-0870-2013 | 64329 | No | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | MEDRAD Hand Controller Sheath, Catalog number AVA 500 HCS/3010903, for cardiac catheterization procedures. The Sterile Sheath is specifically intended for use in the x-ray angiography environment to cover the Medrad Avanta Sterile Hand Controller to prevent cross contamination between patients. |  |  | November 05, 2012 | 2012 | Medrad Inc | The seal between the white polyethylene and the clear plastic may be missing on the Hand Controller Sheath package, resulting in a potential breach of the sterility of this accessory. | Packaging | Medrad sent a Recall letter dated November 5, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customer were instructed to review their inventory and quarantine any packages with an expiration date prior to April, 2017. Customers were also given the option of conducting a visual inspection on site to determine if the product should be returned. Alternatively, the customer has the option of returning the affected product. If the customer has affected product, they are instructed to contact Medrad Customer Support to make arrangements for the return and replacement of the product.For further questions please call our Customer Support Team at 1-877-229-3767. |  | 832175 | 832,175 | Worldwide Distribution including USA nationwide. | Terminated | March 10, 2015 | 855 | 832175 | Not_Computer
Z-1958-2013 | 64970 | No | FSM | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Case Lid and Base (Inspan, Invue, SLIFT) Caddy Lid (Inspan, Arena-C); Inspan Lid: 12-60039A, 11-60042B, 11-60043B, 11-60045B, 11-60046C, 11-60045A, 11-60046A, 11-60053A; Invue Lid: SI70095A; S-Lift Lid: SI50070A; S-Lift Base: SI50067A; Arena-C Lid: 11-72013D, 11-72014D; Generics Lid: 12-99005A. |  |  | April 22, 2013 | 2013 | SpineFrontier, Inc. | Non-toxic inks, used on the silk screen graphics of the case bases and lids, and caddy lids, were found to be either delaminating or running off from the substrate after repeated washing and sterilization cycles. | Labeling design | On 2/15/2013, SpineFrontier emailed distributors informing them that they needed specific parts returned to work towards completing the conversion of product graphics. |  | 258 | 258 units (30 in distribution at the time of the recall) | Worldwide Distribution-USA including the states of FL, OH, MA, and TX, and the countries of Mexico, and Dominican Republic. | Terminated | November 06, 2015 | 928 | 258 | Not_Computer
Z-1018-2013 | 64510 | No | HDW | Contraceptive diaphragm and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Omniflex Diaphragm - Size 65 (Milex Wide Seal Silicone Vaginal Diaphragm) Model MXWF65 |  |  | February 26, 2013 | 2013 | Cooper Surgical, Inc. | Incorrect size was printed on the outer product packaging. | Packaging | CSI notified consignees affected by this recall by a Recall Notification letter, dated February 26, 2013, sent via Federal Express with confirmed delivery receipt. The letter identified the affected product, and asked customers to contact the firm's Customer Service Department at 1-800-243-2974, if they had any of the affected lot in their inventory. Customers also have the option of completing the form attached to the letter and faxing it back to 800-262-0105. At CooperSurgical's expense, arrangements will made to replace the affected products that customers have on hand. The firm will provide further instruction and disposition of the product at the time of replacement. |  | 80 | 80 individual diaphragms | Worldwide Distribution -- USA, including the states of NE, MA, FL, TN, CA, ME, SC, CA, NH, OK, TX, NM, MO, NY, PA, and AZ, and countries of Germany, Italy, and Canada. | Terminated | March 04, 2014 | 371 | 80 | Not_Computer
Z-0029-2014 | 66105 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Model BP7507 OPTI LION E-Plus Cassettes 25 Cassettes with Sample Probes, OPTIMedical |  |  | August 09, 2013 | 2013 | OPTI Medical Systems, Inc | OPTI LION E-Plus Cassettes were labeled with an incorrect barcode. The incorrect barcode could perfect the performance of one or more parameters as the product ages. Also a Customer complaint from a distributor reported that a customer was getting QC failures. | Error in labeling | The firm, OPTIMedical, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated August 9, 2013 via email and fax to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately discontinue use of the product; return any unused product to OPTI Medical System; distributors outside the U.S. may destroy unused product and certify the instructions on the enclosed Recall Response Card; contact your distributor or OPTI Medical System Technical Support at 1-800-490-6784 or technicalsupport@optimedical.com to obtain authorization number and shipping instructions; and complete and return the Recall Response Card to technical support via fax to 1-770-510-4447; mail to: OPTI Medical Technical Support, OPTI Medical Systems, Inc., 235 Hembree Park Drive, Roswell, GA 30076 or email: technicalsupport@optimedical.com Please contact OPTI Medical Systems, Technical Support department at 1-800-490-6784 if you have questions related to this notice. |  | 9025 | 361 boxes of 25 cassettes each | Worldwide distribution: US (nationwide) including states of: CA, MA, MN, and VA; and countries of: China, Guatemala, India, Iraq, South Korea, and Untied Arab Emirates. | Terminated | December 12, 2013 | 125 | 9025 | Not_Computer
Z-1960-2013 | 65005 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Biomet 3i Dental Implant, Parallel Walled 6mm (D) x 10mm (L) |  |  | May 10, 2013 | 2013 | Biomet 3i, LLC | A small number of outer implant boxes may be mislabeled. The inner tray, containing the implant, is label correctly. This condition, if present and not recognized could potentially result in the dental implant not being able to be placed into the osteotomy or seated to the desired depth if placement is attempted. | Labeling mix-ups | The firm issued "Urgent Medical Device Recall" notification letters to their customers dated May 10, 2013. The notification described the issue and provided recommendations actions regarding affected product. Customers with questions may contact 1 800-342-5454. |  | 51 | 51 | Distribution to the state of Massachusetts, and the foreign countries of Australia, Belgium, Spain, Denmark, France, Luxembourg, Switzerland, Germany, Bonaire, Sweden, and Italy. | Terminated | August 13, 2014 | 460 | 84 | Not_Computer
Z-0879-2013 | 64250 | No | GXI | Radiofrequency lesion probe. | Neurology | Neurology | 510(k) | BMC RF Cannula Curved Sharp RadiOpaque, Non-Pyrogenic, Active Tip: 10mm, Gauge 16 Ga, Length: 100mm. The BMC RF Cannula is sealed in a Tyvek pouch (primary packaging) and then placed into a white shelf-box (secondary packaging) with the instructions for use. A predetermined number of white shelf-boxes are then placed into shipping cartons (tertiary packaging). The BMC RF Cannula is intended for use in radiofrequency heat lesion procedures for relief of pain. |  |  | October 04, 2012 | 2012 | Baylis Medical Company Inc. | A report of an event was received indicating that upon opening a box containing 10 sterile cannula by a nurse, the pouch package of 5 of the 10 individually packaged cannula had not been sealed along one edge. Thus the pouch was open and the device could fall out of the pouch and on to the floor when retrieved from the shelf box. The cannulas were not used on any patients, as the unsealed packages | Packaging | A recall was initiated on 10/04/2012 and a letter was sent out to the distributor with a Fax-back Acknowledgment Form. The letter described the issue and provided recommendations and a contact number. |  | 1780 | 1780 | International distribution to the country of Australia. | Terminated | May 23, 2013 | 231 | 1780 | Not_Computer
Z-0081-2014 | 66483 | No | DTK | Cardiovascular intravascular filter. | Cardiovascular | Cardiovascular | 510(k) | Cordis OPTEASE Vena Cava Filter |  |  | September 27, 2013 | 2013 | Cordis Corporation | Cordis identified a printing error on one unit of the Optease Vena Cava Filter, in which the orientation arrow for the femoral approach was printed in the incorrect direction. The error resulted in the filter being implanted upside down, requiring and additional percutaneous procedure to retrieve the filter. | Error in labeling | Cordis sent an Urgent Medical Device Recall letter dated October 8, 2013, to all affected customers. The letter identified the product (Cordis OPTEASEÂż Vena Cava Filter Catalog Numbers 466F220A and 466F220B), the problem, and the action to be taken by the customer. Customers were instructed to:1) Read this Urgent Medical Device Recall letter.2) Immediately identify and set aside all product listed below in a manner thatensures the affected product will not be used.3) Review, complete, sign and return the enclosed Acknowledgement Formdirectly to Cordis at the fax number on the form.4) Return any affected product with reference to the RA# on the form. ContactCustomer Service (see Available Assistance section) if assistance is needed for return of the affected product. Credit will be provided or consignment levels reduced.5) Share this letter with others in your facility that need to be made aware of thisrecall.6) Contact any other facility to arrange the return of OPTEASEÂż if any productlisted below has been forwarded to them.7) Maintain awareness of this notice until all affected product has been returnedto Cordis.8) Keep a copy of this notice with the affected product.Please contact the special Customer Service number (800) 551-7683,Monday through Friday from 8:30 AM to 6:30 PM ET.o Return instructions are provided on the Acknowledgment Form.Questions related to the Acknowledgement Form" Please contact Cordis QA at (786) 313-8730.For questions regarding this recall call 1-800-551-7683. |  | 29395 | 29,395 units | Worldwide Distribution - US (nationwide) including Hawaii and Puerto Rico and states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY; and countries of: Canada and Panama. | Terminated | September 23, 2014 | 361 | 29395 | Not_Computer
Z-1016-2013 | 64271 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Synthes Dens Instrument Set (Dens Graphic Case), Part Number 687.030 (also known as the Anterior Column Retractor Set) |  |  | January 11, 2013 | 2013 | Synthes USA HQ, Inc. | The Dens Instrument Set has been designed to facilitate the anterior internal fixation of the Dens (odontoid process). There are currently no screws specifically indicated to treat Dens fractures via anterior screw fixation therefore the affected product will be removed. | No Marketing Application | Synthes sent an Urgent Notice Medical Device Recall letter dated January 11, 2013, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to examine their inventory, remove the recalled product from use and return it to Synthes. Customers were asked to completed the attached verification Section at the end of the letter by checking the appropriate box indicating that no affected product had been located and fax to 610-251-9005.For questions customers were instructed to call 610-719-5450 or email at FieldAction@synthes.com.For questions regarding this recall call 610-719-5000. |  | 148 | 148 | Nationwide Distribution including AL, AR, AZ, CA, CO, CT, DE, FL, GA, IL, KS, KY, LA, MD, ME, MI, MN, MO, NC, NE, NH, NJ, NV, NY, OH, OR, PA, RI, TN, TX, VA, VT, WI, WV, and WY. | Terminated | September 10, 2015 | 972 | 148 | Not_Computer
Z-0784-2013 | 63390 | pack, hot or cold, reuseable | IME | Cold pack. | Physical Medicine | Physical Medicine | 510(K) Exempt | THERA PEARL 3-in-1 BREAST THERAPY, 2 hot and cold packs in plastic box. |  |  | August 14, 2012 | 2012 | Lansinoh Laboratories Inc | The Spanish translation on the Instructions for Use (IFU) insert and product package for hot and cold packs has incorrect heating instructions which would cause product to over heat. | Error in labeling | Lansinoh Laboratories, Inc. sent an Urgent Product Recall letter dated August 14, 2012, to all affected customer. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to stop distributing and quarantine the affected product. Customers were informed that a credit would be issued for the affected product and they were asked to mail the postage paid Business Reply Card even if they do not have the affected product. If necessary call 1-877-366-1182 to arrange pickup. For questions customers were instructed to call the Customer Service Center at 1-800-292-4794.For questions regarding this recall call 703-299-6477. |  | 51472 | 51,472 units | Nationwide Distribution, including the states of TX, CA, NJ, NY, SD, RI, WV, IL PA, NC and MN. | Terminated | March 08, 2013 | 206 | 51472 | Not_Computer
Z-1004-2013 | 64531 | No | HHW | Vaginal pessary. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Cooper Surgical Incontinence Ring Flexible Pessary #4, 2 3/4", 70mm. Milex |  |  | February 26, 2013 | 2013 | Cooper Surgical, Inc. | Pessary mislabeled outer package may not reflect the size of the product contained in the package | Packaging process control | Cooper Surgical initiateda recall by letter on February 26, 2013, advising accounts to examine inventory and report quantity back to the firm for replacement. For questions customers should call Customer Service at 1-800-243-2974.For questions regarding this recall call 203-601-9825. |  | 80 | 80 units | Worldwide Distribution - USA (nationwide) and internationally to Australia, Canada, Colombia, Demarkt, France, Holland, New ZeaJand, andThailand. | Terminated | March 04, 2014 | 371 | 750 | Not_Computer
Z-0342-2014 | 66688 | No | JEB | Manual operating table and accessories and manual operating chair and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | TruSystem Surgical Table (TS7000)  Leg Section One Part Light HU |  |  | October 28, 2013 | 2013 | Trumpf Medical Systems, Inc. | The coupler point (hood) of the Leg Section One Part Light may become detached from the leg section frame. | Packaging | TRUMPF sent an Urgent Safety Information letter dated October 28, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.Customers were instructed to inspect Leg Section One Part Light for damage to the coupler point (hook) between each case. If any damage is noticed to the coupler, they were instructed to stop use immediately and contact TRUMPF Technical Service. If they drop or collide the Leg Section One Part Light into anything then stop use immediately and contact TRUMPF Technical Service. Review the User Instruction for proper use and care of the device. The leg section should be used only as intended as part of the TRUMPF TS7000 surgical table. The intended use is to support and position the patients legs during the pre-operative and post-operative stages of surgery. The leg section should be removed during surgery. TRUMPF will be exchanging the hooks, counter plates and screws of the coupler point to increase the area that the load is distributed over. This will allow the screws securing the hood to the leg frame to withstand external forces. TRUMPF anticipates parts to be available the first week of November. TRUMPF trained technicians will complete the TRUMPF Medical Systems, Inc. for the part exchange. Completion is anticipated by November 11, 2013.Further questions please call (843) 534-0606 |  | 15 | 15 components | US Distribution including the states of CA, IL, NJ, NY, NC, TX, WA and SC. | Terminated | September 16, 2014 | 323 | 15 | Not_Computer
Z-0064-2014 | 66248 | No | MBB | Polymethylmethacrylate (PMMA) bone cement. | Orthopedic | Orthopedic | 510(k) | SmartSetÂż Endurance GMV, Gentamicin, REF 3105-040, Medium Viscosity Antibiotic Bone Cement, Sterile. |  |  | September 16, 2013 | 2013 | DePuy Orthopaedics, Inc. | The SmartSetÂż GMV Endurance Gentamicin Bone Cement, product ref: number 3105-040 is intended for use in indications that have not been cleared by the US FDA. The intended use stated in the IFU for the SmartSetÂż GMV Endurance Gentamicin Bone Cement is for the fixation of prostheses to living bone in arthroplasty procedures of joints in which infection by gentamicin-sensitive organisms is a potentia | No Marketing Application | The firm DePuy Synthes, sent an "URGENT INFORMATION-MEDICAL DEVICE RECALL NOTICE" dated September 16, 2013 to its consignees via E-mail, hand delivery or fax depending on the most appropriate delivery method for each consignees. The Urgent recall notice described the recalled product, explained the reason for the recall, and directed consignees to cease further distribution, complete the attached Reconciliation form and quarantine and return all affected products through the DePuy Orthopaedics Sales consultant, Fax 1-787-287-7681 or Email to vbaez1@its.jnj.com. Consignees were directed to maintain a copy of the notice and forward the notice if product was further distributed. Clinical questions should be addressed to DePuy Orthopaedics, Inc's Scientific Information Office at 1-888-554-2482 (M-F; 8 a.m. - 5 p.m. EDT). Recall questions or concerns should be addressed to the Quality &amp; Compliance Supervisor, at 1-787- 272-1900 x3114 or vbaez1@its.jnj.com (MF; 8 am- 5 pm EST). |  | 16040 | 16,040 units | Distributed in Puerto Rico and US Virgin Islands. | Terminated | February 24, 2015 | 526 | 16040 | Not_Computer
Z-0074-2014 | 65799 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | NexGenÂż Femoral Component (CR-FLEX GSF POR FEM, C-L; CR-FLEX GSF POR FEM, C-R; CR-FLEX GSF POR FEM, C-L MINUS; CR-FLEX GSF POR FEM, C-R MINUS; CR-FLEX GSF POR FEM, D-L; CR-FLEX GSF POR FEM, D-R |  |  | August 30, 2013 | 2013 | Zimmer, Inc. | The low density polyethylene (LDPE) bag used to package implants adheres to the highly polished surface of the devices. | Package design/selection | August 30, 2013 Urgent Recall Notices were sent to Distributors, hospital risk managers and surgeons identifying the issue and directing consignees to make all affected personnel aware of the contents of the notification. The notice identified the products involved, the issue, possible adverse events associated with the issue and recommended that implants found with LDPE bag adhesion not use the product and return it to Zimmer sales representatives. Questions or concerns should be addressed to Zimmer customer call center at 1-877-946-2761 between the hours of 8 a.m. and 5 p.m. EST. |  | 11694 | 11694 | Worldwide Distribution: US (nationwide) and countries of: Argentina, Australia, Bolivia, Brazil, Canada, Chile, China, Colombia, Costa Rica, Dominican Republic, Ecuador, El Salvador, Germany, Guatemala, Honduras, Hong Kong, India, Israel, Jamaica, Japan, Korea, Malaysia, Mexico, New Zealand, Nicaragua, Panama, Peru, Singapore, Taiwan, Thailand, and Venezula. | Terminated | September 25, 2014 | 391 | 204220 | Not_Computer
Z-0360-2014 | 66720 | No | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Salter Labs EtCO2 Sample line 15' (4.6 M) tube, male luer connectors. |  |  | November 01, 2013 | 2013 | SALTER LABS | It has been determined that three lots of model number 4515-15-25 CO2 Sample Line 15' tube, male Luer lock connectors may contain a 30' tube instead of the 15' length. | Labeling Change Control | Salter Labs called the five customers that received product from the three lots, then follow-up immediately by sending a recall notification letter directly to each customer. Each letter will identify all shipments of subject product that they received.All customers letters were sent certified mail, return receipt requested. |  | 91 | 91 cases (25 units per case) | Distributed in the states of MS, TX FL,, WI, MI, VA, NY, MO, NC, and MS. | Terminated | December 17, 2013 | 46 | 91 | Not_Computer
Z-0114-2014 | 65997 | No | JDQ | Bone fixation cerclage. | Orthopedic | Orthopedic | 510(k) | Synthes Application Instrument Sternal Zip Flex System |  |  | August 13, 2013 | 2013 | Synthes USA HQ, Inc. | When using the 1st generation Zip Fix Application Instrument it is possible to cut the implant while the tensioning trigger is being squeezed which could lead to loosening of the implant. | Labeling False and Misleading | Synthes sent a Urgent Medical Device Notification dated August 13, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Action to be taken: This notification is to make you aware of the issue and the importance of completely releasing the tensioning trigger before cutting the implant. Please take the following action:"To ensure your receipt and understanding of this information, Please complete the verification section of this letter and return to Synthes by either of the following methods:" Fax: 1-888-352-9616 or" Scan/email: FieldAction@synthes.comIf you have any questions, please call 610-719-5450 or contact your Synthes Sales Consultant. |  | 319 | 319 | Worldwide Distribution - US (nationwide) and country of Mexico. | Open, Classified |  |  | 319 | Not_Computer
Z-0042-2014 | 66433 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Brand: The ForeCYTE Breast Health Test, Mammary Aspiration Specimen Cytology Test (MASCT), MASCT System Kits, Clarity System Kits, MASCT Patient Sample Kits &amp; Clarity Patient Sample Kits Instrument, Biopsy. |  |  | September 19, 2013 | 2013 | Atossa Genetics, Inc. | Atossa Genetics is recalling ForeCYTE Breast Health Test, Mammary Aspiration Specimen Cytology Test (MASCT) due to concerns raised by the U.S. Food and Drug Administration (FDA) in a warning letter received by Atossa in February 2013. The FDA raised concerns about (1) the current Instructions for Use (IFU); (2) certain promotional claims used to market these devices; and (3) the need for FDA cl | No Marketing Application | Atossa Genetics Inc. issued a press release to announce a voluntary recall on October 4, 2013.The firm started sending out the URGENT - VOLUNTARY PRODUCT RECALL NOTIFICATION ForeCYTE Breast Health Test; Mammary Aspiration Specimen Cytology Test (MASCT) RESPONSE REQUIRED letter to consignees on October 7, 2013. The letter identified the affected product, problem and actions to be taken. The firm initiated a voluntary recall to remove the ForeCYTE Breast Health Test and the Mammary Aspiration Specimen Cytology Test (MASCT) device from the market. The MASCT device has not been cleared by the FDA for the screening or diagnosis of breast cancer. In addition, the ForeCYTE Breast Health Test has not been cleared or approved by the FDA for any indication. The ForeCYTE Breast Health Test and the MASCT device are not a replacement for screening mammograms, diagnostic imaging tests, or biopsies. Patients should follow the recommendations and instructions of their physician with respect to breast cancer screening and diagnosis.Distributors are instructed to notify downstream customers and physicians are instructed to notify patients who have been tested using the MASCT and/or ForeCYTE systems. Consignees should return the Response Form back to Atossa.Distributors and customers should call 800-219-4629 to obtain a Returned Materials Authorization (RMA) number to return affected products. Customers with questions about this voluntary recall or Atossa's products, or would like assistance with the recall documentation may contact Atossa at 1-888-219-4629.Any problems experienced with the use of this product may be reported to the FDA at http://www.fda.gov/Safety/MedWatch/HowToReport/default.htm,Online (form available to fax or mail), or call FDA 1-800-FDA-1088. |  | 58277 | 577 MASCT System kits, 101 Lab Kits, and 2557 Patients Kits (5 test per kit) | USA Nationwide Distribution in the states of: AL, CA, FL, GA, IL, IN, KY, MD, ME, MI, MO, NC, ND, NJ, NV, NY, OK, PA, SC, TN, TX, UT, VA, WA, Washington DC, and Puerto Rico. | Terminated | August 19, 2014 | 334 | 58277 | Not_Computer
Z-1965-2013 | 65829 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Arrow International, Inc. Two-Lumen CVC Kit with Blue Flex Tip Catheter. The multiple lumen catheter permits venous access to the central circulation of patients. |  |  | July 31, 2013 | 2013 | Arrow International Inc | Arrow is recalling the product due to a labeling inconsistency. One portion of the lidstock states the kit contains a 3-Lumen catheter while the rest of the lidstock correctly identifies the catheter as 2-Lumen. | Labeling Change Control | An "Urgent Medical Device Recall Notification" letter dated July 31, 2013 was sent to direct accounts requesting immediate discontinuation of use and quarantine of all affected product on hand. Negative responses were also requested and the immediate return of product to Arrow was authorized. Questions were directed to Customer Service at 800 343-2935. |  | 156 | 156 | AL, PA, CT, NC, VA, KS, RI, AZ, NJ, and IL. | Terminated | April 28, 2015 | 636 | 156 | Not_Computer
Z-0128-2014 | 66423 | No | EWL | Laryngeal prosthesis (Taub design). | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | Blom-Singer Adjustable Tracheostoma Valve II |  |  | October 01, 2013 | 2013 | Helix Medical LLC | InHealth Technologies (IHT) is recalling the Blom-Singer Adjustable Tracheostoma Valve II (ATSV II) because the body component of the device is missing from the packaging. The ATSV II Valve product cannot be used as intended without the product body. | Packaging change control | A phone script and a customer notification letter are provided. The recall was initiated on 10/1/13 and customers will be notified that InHealth Technologies, a division of Helix Medical, LLC has initiated a recall of the following 3 lots of the Blom-Singer Adjustable Tracheostoma Valve (BE8025H and BE8025H-R3): 1057760, 1057919 and 1059222. The phone script provides the problem identified and the actions to be taken. Customers are instructed to complete the Customer response form and return by email to belinda.jackson@helixmedical.com, fax to (805) 566-5381 or by mail to InHealth Technologies, 1110 Mark Avenue, Carpinteria, CA 93013. Customers are also instructed to Contact UPS at 1-800-742-5877 and provide In Health Technologies UPS account no., 846718. UPS will come to your site, package the recalled product(s) and ship it to InHealth Technologies at InHealth Technologies expense. Customers are instructed to respond within 10 days. |  | 47 | 47 units | Worldwide Distribution-USA (nationwide) and the countries of Belgium, Canada, Croatia, Germany, Ireland, and Italy. | Terminated | January 21, 2014 | 112 | 47 | Not_Computer
Z-2243-2013 | 64406 | No | HTY | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Synthes Cannulated Distal Femoral Nail (Trauma Nail), Part Number 450.851S, 420MM-Sterile. Orthppedic implant device. |  |  | April 18, 2013 | 2013 | Synthes USA HQ, Inc. | Firm discovered that a nail that was within the scope of a previous Trauma Nail recall had been distributed. | Packaging process control | A letter dated April 18, 2013 was sent to one consignee who was identified as receiving the affected device. A review of inventory and return of any affected product on hand was requested. Negative responses were also requested. |  | 1 | 1 | US distribution to MA. | Terminated | March 30, 2015 | 711 | 1 | Not_Computer
Z-1995-2013 | 65834 | No | KKX | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Blue Belt Technologies, Inc. Monitor Drape REF: PFSD-V0001, Rx only, Sterile EO. |  |  | July 19, 2013 | 2013 | Blue Belt Technologies MN | Blue Belt Technologies is recalling sterile Monitor Drapes used with their Navio PFS system during orthopedic knee surgery. The drapes are used to cover the Navio PFS system computer monitor and arm assembly and provide a sterile barrier in the surgical suite. The defect is evident when the computer monitor is being draped. The front section of the drape separates from the rest of the drape at | Package design/selection | Blue Belt Technologies contacted consignees by telephone using a script on July 19, 2013. The consignees were informed of the problem and the product involved in the recall. Consignees were advised to immediate stop using the product and to locate and quarantine all unused product. Additionally, an "Urgent Medical Device Recall" letter dated July 24, 2013 was issued on July 24, 2013 and August 12, 2013. The letter includes the description of the Issue, description of safety concerns, affected product, and action to be taken by user. For questions contact Blue Belt Technologies, Lindsey Bring (+1 763.452.4922) or Robin Gonzalez (+1 763.452.4910) and get a Return Authorization number. Lindsey or Robin will arrange a pickup of the product. BBT will send replacement product at no charge. |  | 2400 | 60 cartons (20 pieces/carton) 40 in EU, 20 in US. | Worldwide Distribution and US Nationwide in the states of: CA and PA.United Kingdom. | Terminated | November 14, 2013 | 118 | 2400 | Not_Computer
Z-2158-2013 | 66029 | No | HTY | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Steinmann Pin, packed in a clear tube with end caps, tubes packed in a sterile barrier pouch, 6 pouches/sales unit. Pouches are labeled in part ***MicroAire Surgical Instruments 3590 Grand Forks Blvd Charlottesville, VA LLC 22911*** |  |  | August 12, 2013 | 2013 | MicroAire Surgical Instruments, LLC | A medical device used in surgical procedures was packaged in a pouch in which a defective seal may compromise sterility | Packaging process control | MicroAire Surgical LLC sent an Urgent Medical Device Recall letter dated August 12, 2013, to all affected customers. The letter advised consignees recover and quarantine the affected product so that it does not make it into the surgical process. Customers were instructed to either return affected product to MicroAire for inspection or they could inspect the product for the defect by looking at the seal area on the inside sterile package (short non-chevron side). The recalling firm further advised distributors to contact all of their customers to recover and quarantine the product so that it does not make it into the surgical process. Customers with questions were instructed to call 434-975-8370.For questions regarding this recall call 434-975-8370. |  | 15786 | 15786 total units | Worldwide Distribution - USA (nationwide) and internationally to Australia, Ireland, Canada, Netherlands, Sweden, Japan, and United Kingdom. | Terminated | February 12, 2014 | 184 | 15786 | Not_Computer
Z-0958-2013 | 64162 | No | CGX | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Creatinine, Part Numbers OSR6178, OSR6678 |  |  | December 28, 2012 | 2012 | Beckman Coulter Inc. | Beckman Coulter is including additional information and instructions in the Creatinine (OSR6x78) IFU and Setting Sheet, relating to calibration stability and quality control frequency to assure product performance. | Labeling design | Beckman Coulter sent an Urgent Product Correction letter to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter informed the customers that Beckman Coulter is including additional information and instructions in the Creatinine (OSR6x78) IFU and Setting Sheet, relating to calibration stability and quality control frequency to assure product performance. Customers are instructed to complete and return the enclosed response form. Customers with questions about the notification letter are instructed to contact the call center (hotline) at (800) 854-3633 in the US and Canada. |  | 57962 | 57,962 units (50,452 units in the US) | Worldwide Distribution - USA Nationwide and the countries of: Canada, Brazil, Colombia, Mexico, Puerto Rico, Venezuela, Guam, Honduras, Peru, Ecuador, Costa Rica. | Terminated | October 17, 2013 | 293 | 57962 | Not_Computer
Z-0186-2014 | 65739 | No | ONW | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Combiset Blood Line Hemodialysis use |  |  | August 29, 2013 | 2013 | Fresenius Medical Care Holdings, Inc. | Updated Instructions for Use: Bloodline connections not tightened per instructions in the operator's manual may disconnect and lead to blood loss | Labeling design | Fresenius issued a Field Notification Letter dated 8/29/13 to Patient Clinics/ associated at home patients and provides additional detail instructions to ensure the connection to the transducer protector bloodline female luer interface is secure.This field advisory notice sent by Certified Mail, Return Receipt Requested. Accounts requested to return both the receipt and the completed reply form (enclosed) to confirm receipt of the notice.If you have any questions regarding this field advisory, please contact FMCNA Customer Service Team at 1-800-323-5188. Customer Service will be available Monday-Friday 7am-6pm CST. |  | 508029 | 508,029 cases | Worldwide Distribution-USA (nationwide) and the country of Canada. | Open, Classified |  |  | 540759 | Not_Computer
Z-0917-2013 | 64367 | No | GEY | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | D3000I Advantage Drive System, 230v, surgical instrument motor. |  |  | January 24, 2013 | 2013 | Linvatec Corp. dba ConMed Linvatec | The Membrane Switch Panel (referred to as "front panel" of the component part number A34-003-207, used on both the D3000 and D3000I Advantage Drive System, contained the incorrect symbology per third edition IEC60606-01 Medical Electrical Equipment: General requirements for basic safety and essential performance. | Error in labeling | Linvatec Corp sent an "URGENT: Medical Device Recall Notification" letter dated January 24, 2013 to its customers after initial phone notifications. The letter describes the affected product, problem and actions to be taken. Customers were instructed to immediate check facility's inventory for the affected product, use the enclosed FedEx air-bill when returning the product, and complete and return to the Recall Reply Form. For questions contact your Customer Service department at 800-535-8536 or fax at 727-319-5701 or email Custserv1@linvatec.com. |  | 23 | 23 units | Worldwide Distribution - USA (nationwide) including the states of: CA, FL, GA, IL, KS, MA, MD, MO, NE, NY, and TX and the countries of: Canada, Germany, Romania, Spain, Palestine, South Africa, France, Sweden and Singapore | Terminated | July 05, 2013 | 162 | 54 | Not_Computer
Z-0800-2013 | 63433 | No | MBP | Resorbable calcium salt bone void filler device. | Orthopedic | Physical Medicine | 510(k) | Bacterin International, Inc., Osteoselect Demineralized Bone Matrix DMB Putty; Graft ID# B090056-664 Donor# B090056 2.5cc Expiration Date: 12/25/2011; LBL 210-0010-1 REV: 040109; and Graft ID# B090056-675. Bone void filler for voids or gaps in bone. Indicated for treatment of surgically created osseous defects. |  |  | April 18, 2012 | 2012 | Bacterin International, Inc. | Expired product was shipped to two customers. | Incorrect or no expiration date | Bacterin International phoned the two customers and they were instructed to return any of the unused product.For questions regarding this recall call 406-388-0480.The 3rd expired product was discovered by Bacterin Sales Rep 7/12/2012 at the consignee. |  | 3 | 3 packages each containing 2.5cc of putty | Nationwide Distribution in TX , GA and CA. | Terminated | May 13, 2013 | 390 | 3 | Not_Computer
Z-2165-2013 | 66082 | No | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Baxa Vial Adapter, 14mm - Luer Lock, Order No. REF: 90, Quantity: 50 each, Single Use. Product Code H93890. |  |  | August 22, 2013 | 2013 | Baxa Corporation | Baxter Healthcare Corporation is voluntarily recalling one lot of Baxa Vial Adapters, 14mm - Luer Lock, due to incorrect expiry date. | Labeling Change Control | An Urgent Product Recall notification was mailed to one customer via USPS First Class mail on 8/22/13. The one customer affected by this action was instructed to locate and remove the affected product from inventory and to return to Baxter for credit and/or replacement. Any customer returns received will be discarded per the Corporate Hold. |  | 34 | 34 cases | Distributed to One Customer in NC. | Terminated |  |  | 34 | Not_Computer
Z-1269-2013 | 64696 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | A) Product Code 2N3371: Non-DEHP Y-Type Catheter Extension Set, 5.7", Vol. 0.8 mL, 2 Injection Sites, Male Luer Lock Adapter, Single use only |  |  | March 28, 2013 | 2013 | Baxter Healthcare Corp. | Customers have reported that when separating an individual package from its attached grouping, the adjacent package has opened, compromising its sterile barrier properties. | Packaging process control | Baxter sent an "Urgent Product Notification" dated March 28, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taked. Customers are advised to locate and remove all of the affected products from their inventory. If the products were further distributed, consigneees are requested to notify their customers of the recall. Products can be returned for credit by contacting Baxter. An acknowledgement form is required to be compelted by customers. |  | 228200 | A) Product Code 2N3371: 228,200 units | Worldwide Distribution - USA Nationwide and the countries of Puerto Rico, Brazil, Canada and Columbia. | Terminated | July 31, 2017 | 1586 | 844300 | Not_Computer
Z-0138-2014 | 65782 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Stereotactic Circular Collimator - 3D Line stereotatic Hardware Accessories |  |  | July 23, 2013 | 2013 | Elekta, Inc. | Potential for clinical errors. | No Marketing Application | Elekta sent an Important Field Safety Notice dated 7/2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were advised of the issue and recommendation to resolve either by upgrading their system or to discontinue a clinical use of the device.For further questions please call ( 770 ) 300-9725. |  | 89 | 89 | Nationwide distribution: USA including the states of AZ, CA, FL, GA, MO, NJ, NM, NC, OR, TN and WV. | Terminated | November 18, 2014 | 483 | 89 | Not_Computer
Z-0918-2013 | 64343 | No | GAG | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | PSX PROXIMATE Skin Staple Extractor, |  |  | October 04, 2012 | 2012 | Ethicon Endo-Surgery Inc | Damage to the packaging may have compromised the sterility of the device. | Package design/selection | On 10/28/2012, the firm sent Urgent: Device Recall, Event 5928 notification letters to their customers. The letters identified the affected product, the reason for the recall, the action to be taken by customers, as well as the actions taken by the firm. Customers are to examine their inventory immediately and remove the affected product. The Business Reply Form should be completed and faxed to Stericycle at 1-877-650-0360 within 3 business days, even if there is no affected product on hand. Customers must complete the "Replacement Purchase Order Number Field" on the Business Reply Form to ensure adequate identification of the return product in their facility. Customers should follow the instructions provided in order to return the product. The information should be shared with all appropriate staff at the customer's facility. EES planned to replace any unopened an unexpired recall product that was returned by December 31, 2012. Questions regarding the issue should be directed to Stericycle at 1-877-650-5412. |  | 5894975 | 5,894,975 pieces | Worldwide Distribution -- USA, including the states of AK, AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, &amp; WY; and, the countries of ARGENTINA, AUSTRALIA, AUSTRIA, BAHRAIN, BELGIUM, BRAZIL, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, CZECH REPUBLIC, DOMINICAN REPUBLIC, ECUADOR, EGYPT, GERMANY, GREECE, GUATEMALA, HONG KONG, HUNGARY, INDIA, INDONESIA, ISRAEL, ITALY, JAMAICA, JAPAN, JORDAN, KUWAIT, LIBYAN ARAB JAMAHIRIYA, MALAYSIA, MEXICO, NETHERLANDS ANTILLES, NICARAGUA, NICOSIA, PAKISTAN, PANAMA, PERU, PHILIPPINES, POLAND, PORTUGAL, PROVINCE OF CHINA, PUERTO RICO, REPUBLIC OF KOREA, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SWITZERLAND, THAILAND, TRINIDAD AND TOBAGO, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY, VENEZUELA &amp;, YEMEN. | Terminated | June 11, 2013 | 250 | 5894975 | Not_Computer
Z-0943-2013 | 64342 | No | JDS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | INTERNAL HEX RECON SCREW, TI-6AL-4V, REF 7162380, TRIGEN (R) INTERNAL HEX CAPTURED RECON SCREW, QTY: (1), 6.4 MM, 80 MM, TITANIUM, Smith &amp; Nephew, Inc. Orthopedic. Bone screw for intramedullary rod fixation. |  |  | February 07, 2013 | 2013 | Smith &amp; Nephew Inc | 75 mm screws were packaged and labeled using an 80 mm screw outer box and label, while 80 mm screws were packaged and labeled using a 75 mm screw outer box and label. | Labeling mix-ups | Smith &amp; Nephew sales representatives and Smith &amp; Nephew international distributors were notified of problem and the recall via email on February 7, 2013. They were instructed to cease use and distribution of the product and quarantine for return. Non responding consignees were notified again on February 28, 2013 by email. Hospitals were notified by letter beginning on February 7, 2013. Smith &amp; Nephew also contacted the surgeons who implanted the screws by letter the week of February 11th.For questions regarding this recall call 978-749-1330. |  | 69 | 69 units | Worldwide Distribution - USA including NM, TN, TX, NC, CA, MI, KY, MA, NJ, OH, and HI. Internationally to Canada, United Arab Emirates, Mexico, Argentina, Japan, and Chile. | Terminated | January 11, 2016 | 1068 | 159 | Not_Computer
Z-0803-2013 | 64177 | No | ILO | Water circulating hot or cold pack. | Physical Medicine | Physical Medicine | 510(K) Exempt | BREG Cold Therapy Combination Units containing Sterile Polar Pads. 02356 Rev M, PAD M/U XL STER, Mfg.Date: 112012. Model #s: 10903, 10703, 09101, 09111, 09131, 09731, 09611, 09621, and 09631. Local anesthetic effect. |  |  | January 14, 2013 | 2013 | Breg Inc | The recall was initiated because Breg has determined that some Sterile Polar Pads products manufactured from January 2010 to October 2012 may have sustained damage to the product packaging which may compromise product sterility assurance. | Packaging | The firm, BREG, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated January 14, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately examine your inventory and quarantine the product; complete the Return Response Form by email, Fax or mail; and contact Breg for a Return Authorization for the affected product and request replacement by contacting Breg Customer Care at 800-321-0607. Note: response is required even if you have no affected inventory.Should you have any questions regarding this communication or need to report an adverse event, please contact Breg Customer Care at 800-321-0607. |  | 44883 | 44,883 units | Worldwide Distribution-USA (nationwide) and the countries of Australia, Singapore, Chile, Latvia. | Terminated | July 01, 2013 | 168 | 44883 | Not_Computer
Z-2277-2013 | 65998 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco RTP System Radiation Therapy Treatment Planning System. |  |  | May 22, 2013 | 2013 | Elekta, Inc. | Patient needs to be positioned as indicated by the treatment plan. | Labeling design | Computerized Medical Systems issued an Important Safety Notice dated May 22, 2013 to all affected customers. This letter idenitfies the problems the products and the actions needed to be takem by the custom. |  | 72 | 72 | Distributed in Puerto Rico and the states of AK, AZ, CA, CO, CT, FL, IL, IN, MA, MD, MI, MN, MO, NC, NJ, NM, NY, OH, OK, OR, PA, RI, SC, TN, TX, VA, and WA. | Terminated | December 23, 2013 | 215 | 72 | Not_Computer
Z-2278-2013 | 65009 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes 3.5 MM LCP(R) Distal Humerus System |  |  | March 08, 2013 | 2013 | Synthes USA HQ, Inc. | Recall is being initiated due to the part being mislabeled (Part Number 241.267 was mislabeled with Part Number 241.276). | Labeling mix-ups | Synthes sent an Urgent Notice Medical Device Recall letter dated March 8, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to obtain a Return Authorization Number, complete the Verification Section at the end of the letter by the checking the appropriate box indicating affected product located. Return the Verification Form with the product to: Credit/Returns, Synthes, 1101 Synthes Avenue, Monument, CO 80132. Customers with questions were instructed to call 610-719-5450 or email Fieldaction@synthes.com.For questions regarding this recall call 610-719-5000. |  | 2 | 2 | Nationwide Distribution including CA,and VA.. | Terminated | May 21, 2015 | 804 | 2 | Not_Computer
Z-2189-2013 | 65984 | No | DRS | Extravascular blood pressure transducer. | Cardiovascular | Cardiovascular | 510(k) | "***Flo Trac Sensor Model MHD85***" |  |  | July 31, 2013 | 2013 | Edwards Lifesciences, LLC | The Edwards Lifesciences FloTrac Sensor is being recalled due to a manufacturing non-conformance related to incorrect packaging configuration where the IFU was placed on top of the Tyvek instead of being placed at the bottom of the box. There is no evidence available at this time that could validate the sterility of this incorrect packaging configuration. | Packaging process control | Edwards Lifesciences initiated a US recall via an Urgent Product Recall letter dated July 31, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised that an Edweards Lifesciences representative would come to their facility to retrieve the affected device. A device replacement would be provided to the customer. Customers were asked to complete the enclosed Confirmation Form and give it to the Edwards Lifesciences representative that will come to their facility. For customers with questions were instruct4ed to call Edwards Customers Service at 1-800-424-3278.For questions regarding this recall call 972-410-7100. |  | 35 | 35 units | Nationwide Distribution including MO and MI. | Terminated | April 14, 2017 | 1353 | 35 | Not_Computer
Z-0809-2013 | 64110 | No | IOR | Mechanical wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Chenica Gunnell Rehab Series of Wheelchairs (GRSW); Mechanical Wheelchair, Sold as Rehab Tough and Tilt (TNT), Rehab Recline and Mobility (RAM), Rehab Kidster (Kidster). Provide mobility to individuals limited to a seating position. |  |  | January 18, 2013 | 2013 | CHENICA INC. | The wheelchairs may not meet the specifications of WC-19 (crash testing), therefore the WC-19 verbiage and corresponding logo are removed from the flyers on the Gunnell Rehab Series of Wheelchairs (GRSW). | Labeling False and Misleading | Gunnell Inc. sent a "RE: URGENT MEDICAL DEVICE CORRECTION/RECALL" letter dated January 18, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm 1-800-551-0055 for questions relating to this letter. |  | 53 | 53 | Nationwide Distribution-including the states of Florida, Michigan, South Carolina, California, Texas, Maryland, New York, Tennessee, Virginia, Illinois, Ohio, Massachusetts, Arizona and Washington. | Terminated | April 15, 2014 | 452 | 53 | Not_Computer
Z-0795-2013 | 63858 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | The Lumenis VersaCut Tissue Morcellator |  |  | January 02, 2013 | 2013 | Lumenis Limited | Lumenis Ltd is recalling certain models of VersaCut Morcellator devices to perform corrective labeling as a result of three reported events of air embolism secondary to incorrect reversal tubing system hook-up. | Labeling design | Lumenis sent an Urgent Safety Advisory Notice letter dated January 2, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to re-label the devices with the enclosed labels provided, replace the Operation Manual with the current revised Manual, complete the information requested on the Confirmation Card and return the Confirmation Card using the enclosed pre-paid shipping label by UPS. For question contact VersaCut Morcellator field correction administrator by either Cell: +972-(54)-9599196, Tel: +972-(4)-9599296 or E-mail: VersaCutMorcellator@lumenis.com. |  | 931 | 931 units | Worldwide Distribution - US Nationwide including: CA, WI, NC, FL, PA, IN, NM, NJ, IA, WA, OH, FL, MD, MA, WY, NY, AZ, NH, MN, TN, OK, KY, TX, AK, MI, VT. Internationally: Japan, Nepal, India, UK, South Korea, Netherlands, Germany, Australia, China, Canada, Turkey, Spain, Italy, Egypt, Mexico, Bulgaria, Utd. Arab, Emir., Poland, Kuwait, Hong Kong, Saudi Arabia, Yemen, Qatar, Czech Republic, New Zealand, Malaysia, Portugal, France, Argentina, Singapore, Romania, Chile, Colombia, Switzerland, Bahrain, Brazil, Serbia/Montenero, Taiwan, Israel, Indonesia, Peru, Pakistan, Nepal, Belarus, Latvia, Sri Lanka, Russian Fed, Thailand, Azerbaijan, Belgium, and Denmark. | Terminated | May 20, 2013 | 138 | 931 | Not_Computer
Z-0176-2014 | 66191 | No | NGX | Powered muscle stimulator. | Physical Medicine | Physical Medicine | 510(k) | Pinook Mini Massager |  |  | May 23, 2013 | 2013 | Pinook-Usa | Failure to submit a premarket submission and gain approval of a medical device. | No Marketing Application | On October 18, 2013 the firm sent the following notification to their customers:This is to inform you of a product recall involving: Pinooks Mini-massager and Pinooks Micro Massager model number PT-11 See enclosed product label (for ease in identifying the product at retail/user level). This recall has been initiated due to failure to submit a premarket submission and gain approval.We began shipping this product on 11/27/2012. Immediately examine your inventory and quarantine product subject to recall. In addition, if you may have further distributed this product, please identify your customers and notify them at once of this product recall. Your notification to your customers may be enhanced by including a copy of this recall notification letter.Pinook USA is removing the products (mini and micro massager) from the market until receiving FDA approval to market them again under a new 510(k).If you have stock on your end please return them to our office:901 C FL PKWY STE A6, Orlando FL, 32824.Before returning, please contact Sheva Ben-Shushan: accounting@pinookusa.com in order to receive an RMA number.Once you have RMA number, place that number on the package label.This recall should be carried out to the retail level. Your assistance is appreciated and necessary to prevent patient harm. If you have any questions, call Dvir Lev-Ran, telephone number: 407-854-0055 Ext 304. This recall is being made with the knowledge of the Food and Drug Administration. |  | 28994 | 28,994 units (Mini and Micro) | Worldwide Distribution-USA (nationwide) including Puerto Rico and Guam and the states of AZ, CA, CO, FL, IL, MA, MD, MO, MT, NC, ND, NJ, NV, NY, TX, VA, and WA and the country of Canada. | Open, Classified |  |  | 28994 | Not_Computer
Z-2205-2013 | 66143 | No | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | C2244.4308 CONELOG Titanium Base CC Âż4.3GH2 0.8 inkl HS |  |  | August 28, 2013 | 2013 | Camlog Usa | Product not cleared for sale/use in the United States. | No Marketing Application | Camlog Usa notified four US consignees of the recall by telephone on August 5, 2013. The notification identified the recalled product, which did not have proper marketing clearance for distribution in the United States and asked customers to return all unused product. Cusotmers were sent an overnight UPS label via e-mail and asked customers to return the device to CAM:LOG USA Attn: Debby Callahan, 1822 Aston Avenue Carlsbad, CA 92007. |  | 6 | 6 to US customers | USA Distribution in the states of NH, MD, CO, and IL. | Terminated | January 29, 2014 | 154 | 12 | Not_Computer
Z-2209-2013 | 66155 | No | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | Blanketrol III, Model Number 233, Cincinnati Sub-Zero Products, Inc., 12011 Mosteller Road, Cincinnati, OH 45241 |  |  | September 03, 2013 | 2013 | Cincinnati Sub-Zero Products Inc | During a label review it was discovered that both the Blanketrol III Operations (57201-N) &amp; Operations &amp; Technical (56201-Q) manuals had verbiage added back into the manual during ECN M111-4260 (released 12/29/2011) which was previously removed for clarification of the intended use and subsequent 501(K) approval. | Labeling Change Control | Cincinnati Sub-Zero (CSZ) sent an Urgent Medical Device Field Action letter dated September 5, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to contact CSZ at 513-772-8810 or 1-800-989-7373 to request new versions of the manuals. At that time customers are instructed to provide the quantity and serial numbers of the affected devices that are at their facility. Upon receipt of the new versions of the manual, immediately collect and destroy/discard all the affectred devices Operation and Technical Manuals their facility is currentlyin possession of and replace the discarded manuals with the new versions of the manuals provided by CSZ. After the manuals have been discarded and replaced, customers were asked to complete and return the enclosed response form as soon as possible to acknowledge receipt of this notification and to inform CSZ that they have performed and completed the requested actions. Customers should return the form by fax to 513-772-9119 scan and email the information to B3manual@cszinc.com or mail to Cincinnati Sub-Zero, 12011 Mosteller Road, Cincinnati, Ohio 45241. Customers with questions were instructed to contact their Blanketrol III Model 233 field action coordinator at 1-800-989-7373 or 513-772-8810.For questions regarding this recall call 513-772-8810, ext 5295. |  | 1126 | 115V - 396 units; 230V - 385 units | Worldwide Distribution - USA including AK, AR, AZ, CA, CO, DC, DE, FL, GA, IA, ID, IL, IN, KY, LA, MA, MD, ME, MI, MN, MT, , NC, ND, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, SC, TN, TX, UT, VA, WA, WI &amp; WV. Internationally to ARGENTINA, AUSTRALIA, AZERBAIJAN, BELGIUM, CANADA, CHILE, CHINA, COSTA RICA, CROATIA, GERMANY, HONDURAS, HONG KONG, INDIA, INDONESIA, ISRAEL, KOREA, KUWAIT, MEXICO, NEW ZEALAND, PERU, SAUDI ARABIA, SINGAPORE, SPAIN, TAHITI, THAILAND, THE NETHERLANDS, THE PHILIPPINES, UNITED ARAB EMERITES &amp; UNITED KINGDOM. | Terminated | August 22, 2014 | 353 | 1126 | Not_Computer
Z-2202-2013 | 66027 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | Sonogage EyeScan, A-Scan Biometer, |  |  | August 09, 2013 | 2013 | Sonogage Inc | During an FDA inspection on 7/15-26/2013, it was discovered that the firm lacked registration and 510(k) clearance for the device. | No Marketing Application | Sonogage sent an Urgent: Medical Device Recall letter dated August 9, 2013 to their customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immediately examine inventory and quarantine the affected product for return. Customers were asked to complete and return the enclosed response form. For question call 216-464-1119 or 800-798-1119. |  | 564 | 564 units | Worldwide Distribution - USA Nationwide in the following states: AK, AZ, CA, FL, GA, IL, KY, LA, MA, MD, MN, NE, NJ, NM, NY, OH, OR, PA, TX, VA and WA. Product was also shipped to Canada and Mexico. | Terminated | September 19, 2015 | 771 | 564 | Not_Computer
Z-0166-2014 | 66425 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | OmniPod Insulin Management Product Demonstration Kit; Model Reference Number: (PDK-ZXP490) |  |  | September 27, 2013 | 2013 | Insulet Corporation | Alcohol Prep Pads contained within the kit are not properly labeled. | Labeling design | Insulet sent an Important Corrective Action Notification letter dated September 27, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers are asked to discontinue use and properly dispose of the affected product. For questions contact Customer Care at (866) 941-0294. |  | 5050 | 5,050 kits | Worldwide Distribution - USA Nationwide and the country of Israel. | Terminated | March 04, 2014 | 158 | 6769 | Not_Computer
Z-2260-2013 | 66186 | No | HGB | Contraceptive tubal occlusion device (TOD) and introducer. | Obstetrics/Gynecology | Obstetrics/Gynecology | PMA | Cranial Access (CRAK) Kits, Product / Catalogue No.s: INSHITH, INSHITHND, INSHITHRZN |  |  | September 04, 2013 | 2013 | Integra LifeSciences Corp. d.b.a. Integra Pain Management | Integra is recalling certain lots of Cranial Access Kit due to a possibility that the seal integrity of the sterile package may become compromised during shipping. | Packaging | Integra sent an Urgent Voluntary Medical Device Recall letter dated September 4, 2013, to all affected customers that have been shipped unexpired affected product lots of IntegraÂż Cranial Access Kits, Catalogue No.'s INSHITH, INSHITHND, INSHITHRZN.The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to review their inventory and immediately stop the use or distribution of any affected product. Customers were instructed to identify and report to Integra if they do or do not have any of the affected product lots. Customers were advised how to return any affected product lot numbers in their possession.Effectiveness Checks will be conducted to determine the level of customers who:- Integra has verified, have been notified, and / or acknowledged the recall either by returning the acknowledgement form.- Indicate they will be returning / have returned unexpired affected product lots of kits.Any affected product lot numbers that are returned from consignees will be forwarded to the Integra Salt Lake City, UT facility for final disposition per its site QMS processes.Customers with questions were advised to contact Customers Service at 1-855-532-1723.For questions regarding this recall call 801-886-9505. |  | 8648 | 8,648 | Nationwide Distribution including CA, CO, FL, IN, KS, NC, NV, NY, OK, TN, TX, VA, WA, and WI. | Terminated | January 28, 2014 | 146 | 8648 | Not_Computer
Z-0811-2013 | 64272 | No | ELW | Impression material. | Dental | Dental | 510(k) | Ultradent Products, Chromaclone PVS, Regular Set heavy body, Bubble Gum Flavor, 50 ml, Impression Material, Part No. 5012, Lot No. B7PPC. Used in prosthodontics. |  |  | February 04, 2013 | 2013 | Ultradent Products, Inc. | Incorrect packaging of regular set and fast set cartridges. | Packaging process control | Ultradent Products, Inc. sent an Urgent Medical Device Correction letter dated February 4, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their inventory for the affected products. They may continue to use product or dispose unwanted cartridges and reply back to firm indicating acknowledgement of letter and if replacement is needed. Consumers with questions can call the firm at 801-553-4420.For questions regarding this recall call 801-553-4420. |  | 32 | 32 cartridges | Nationwide Distribution | Terminated | May 30, 2013 | 115 | 132 | Not_Computer
Z-2215-2013 | 66037 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | BIORAPTOR 2.9 mm Suture Anchor with 2 ULTRABRAID Suture; Part Number: 72200775 |  |  | August 06, 2013 | 2013 | Smith &amp; Nephew, Inc. Endoscopy Division | Pinholes detected in a small number of primary pouches that may constitute a breach of the sterile barrier. | Packaging | Smith&Nephew; sent an Urgent Product Recall Notification letter dated August 6, 2013, via Federal Express to all affected customers. Addendum customer letters for 4 additional part numbers were sent via Federal Express on August 19, 2013, The letters identified the product, the problem, and the action to be taken by the customer. Customers were asked to inspect their inventory, locate unused devices and quarantine them. Customers were also asked to complete the last three columns in the inventory Return Certification form indicating the quantities that need to be returned, and include their contact information in the spaces provided. Customers were instructed to contact Smith &amp; Nephew's Demo Service Departmentat 800-343-5717 for a Return Authorization (RA#), and provide the information response form. Replacement product would be provided. For questions regarding this recall call 978-749-1073. |  | 43385 | 43385 US | Worldwide Distribution - USA (nationwide) Canada, Europe: Austria, Belgium, Cyprus, Czech Republic, Denmark, Germany, Greece, Finland, France, Italy, Ireland, Luxembourg, Netherlands, Norway, Poland, Portugal, Russia, Slovenia, Spain, Sweden, Switzerland, Turkey, Ukraine and United Kingdom, Asia-pacific: Australia, China, Dubai, Hong Kong, India, Japan, Malaysia, New Zealand, Philippines, Sri Lanka, Singapore, South Korea, Taiwan, Thailand, Argentina, Bermuda, Brazil, Chile, Colombia, Costa Rica, Israel, Mexico, Puerto Rico, South Africa, and Venezuela. | Terminated | April 20, 2016 | 988 | 224222 | Not_Computer
Z-0939-2013 | 64447 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Draeger Infinity Acute Care System (IACS) Monitoring Solution. Catalog Number MS20783 (Infinity MCable-Dual Hemo) |  |  | December 13, 2012 | 2012 | Draeger Medical, Inc. | The displayed value for the invasive blood pressure measurement of the IACS was different than the actual patient values. Liquid ingress in the cable connection may cause inaccurate measurements. This can happen if the Dual Hemo MCable is used in a non-vertical position, a vertical position but upside down, or laying it on the floor or on the patient bed. | Labeling design | Drager sent an "URGENT MEDICAL DEVICE RECALL" letter dated December 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at 1-800-543-5047 for questions relating to this notice. |  | 260 | 260 | Distributed in the states of CA, IL, and MA. | Terminated | May 20, 2013 | 158 | 260 | Not_Computer
Z-2255-2013 | 66185 | No | CAE | Oropharyngeal airway. | Anesthesiology | Anesthesiology | 510(K) Exempt | King LT-D Oropharyngeal Airways |  |  | September 03, 2013 | 2013 | King Systems Corp. | Report that size 5 KLTD airways were packaged in pouches indicating a size of 2.5 and shipped to your facility. The correct size of the KLTD airway is clearly marked on the actual airway product and the physical size difference between 2.5 and 5 is distinct and obvious. | Labeling mix-ups | King Systems notified one consignee who received the affected product verbally on August 7, 2013, via email and an Urgent Medical Device Recall - Immediate Action Required notification on September 3, 2013. The Urgent Medical Device Recall - Immediate Action Required notification described the product and issue as well as includes a response form for the consignee to complete and send back to King Systems via FAX to 317-776-5175, or email to wnet@ambu.com. Questions regarding the notice are directed to: King Systems customer service representative at 800.642.5464, Monday through Friday 8:00am- 5pm EDT.For questions regarding this recall call 317-776-6823. |  | 54 | 9 cases; 45 products | Distributed in NY. | Terminated | April 15, 2014 | 224 | 54 | Not_Computer
Z-0932-2013 | 62993 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Cynch Lordotic T-PLlF Implant. Intervertebral fusion device with bone graft, for use with supplemental fixation. |  |  | August 24, 2012 | 2012 | Spine Smith Partners LP | Implants were incorrectly laser marked as 23mm instead of 28mm. | Labeling mix-ups | Spine Smith sent a Field Action Notice dated August 24, 2012 to its affected customer. The letter identified the product, problem, and actions to be taken by the customer. Contact the firm at 512-637-2097 for questions concerning this notice. |  | 3 | 3 units | Distributed in Texas. | Terminated | May 17, 2013 | 266 | 3 | Not_Computer
Z-0173-2014 | 66132 | No | CEW | Parathyroid hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA Centaur Intact PTH Calibrator (6 pack); Catalog Number: SMN 10492387 |  |  | August 16, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc | PTH Calibrator packaged with only high calibrator. | Packaging | Packaging error was brought to Siemens' attention through escalations received from (3) customers in the United States. Internal Investigation confirmed the issue was limited to the five (5) kits distributed to these customers. As the customers were aware of the issue and non affected product was provided to the customers to use during Siemens' investigation, no further communication was warranted. |  | 595 | 595 kits total (5 kits with packaging error) | Nationwide Distribution including MD, PA, and CT. | Terminated | February 13, 2014 | 181 | 595 | Not_Computer
Z-2254-2013 | 66223 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | POWERWAND Safety Introducer with Extended Dwell Catheter; |  |  | August 16, 2013 | 2013 | Access Scientific Inc | Access Scientific is recalling the POWERWAND Saftey Introducer with Extended Dwell Catheter because it is mislabeled. | Labeling Change Control | On 8/16/13, Access Scientific visited 2 consignees and initiated a field correction. This is the communication given to the two(2) accounts by our field personnel: "Access Scientific has inadvertently shipped mislabeled devices to your facility. The products affected were 5 Fr. POWERWANDs, Catalog # 92006, lot # I666473."The consignees were instructed to "Quarantine the devices until Access Scientific personnel arrived with the appropriate labeling and performed the corrective action of applying the correct labels to the devices and inserting the DFU into the boxes." |  | 160 | 160 units | Distributed in MI and TX. | Terminated | September 24, 2013 | 39 | 160 | Not_Computer
Z-0174-2014 | 66597 | Yes | OAD |  |  | Cardiovascular | PMA | US Safire Duo Ablation Catheter MediGuide Enabled, REF A700240 and Cool Path Duo Ablation Catheter, MediGuide Enabled REF A700244, Rx only, Sterile EO. |  |  | October 11, 2013 | 2013 | St. Jude Medical | St. Jude Medical Cardiovascular and Ablation Technologies Division is performing a voluntary notification regarding selected batches of Coolpath Duo Ablation Catheter MediGuide Enabled products and Safire Duo Ablation Catheter MediGuide Enabled. These units were inadvertently shipped with the incorrect version of the Instructions for Use (IFU). | Error in labeling | The firm, St. Jude Medical (SJM), sent a "Urgent Medical Device Notice" dated 11 October 2013 to its consignees. The letter described the problem, the product involved in the recall and the actions to be taken. The consignees were instructed to discard the IFU received with the batches and replace it with the IFU received in conjunction with the notification letter. SJM has determined that it is acceptable to continue using these devices as long as the product is used in accordance with the IFU.If you have any questions, contact Vice President, Quality, at 651-756-5402. |  | 59 | 59 | US distribution: KS, MA, UT,TX, KY, OH. | Terminated | March 10, 2016 | 881 | 59 | Not_Computer
Z-2162-2013 | 65905 | No | FQZ | Clinical color change thermometer. | General Hospital | General Hospital | 510(K) Exempt | Tempa DOT Single-Use Clinical Thermometer |  |  | July 10, 2013 | 2013 | Medical Indicators, Inc. | TempaDot products may have been mislabeled. | Labeling mix-ups | The firm, Medical Indicators Inc.sent an "URGENT" Recall letter dated July 10, 2013 and Acknowledgement Form to all distributors/customers via email and Priority Mail with delivery confirmation. The letter described the product, problem and actions to be taken. The customers were instructed to check full &amp; partial cases of the tempadot; open and inspect all 5122NS cases; pull out and quarantine any of the lots found in the 5122NS cases, pending return to medical indicators; if individual Tempadot 100-count boxes of 5122 and 5122NS are stored in different warehouse follow the instructions provided in the letter; if Tempadot 100-count boxes of 5122 were sold to a customer with the affected lots, contact the customer and have them return the product; and complete and return the Recall Return Response Acknowledgement and Receipt Form via mail to: 16 Thomas J Rhodes Industrial Drive, Hamilton NJ, 08619; fax: (609) 587-5636 or email to: www.medicalindicators.com or General Manager at: pbaker@medicalindicators.com.Contact information for any question(s) related to this recall may be addressed to General Manager of Medical Indicators, Inc. at 609-737-1600, extension 204 or email to: pbaker@medicalindicators.com, Mondays through Fridays from 8:00am to 5:00pm EST. |  | 111800 | 111,800 boxes (5,590 full boxes) | Nationwide distribution: USA including states of: CA, GA, FL, IN, NY, NV OH, PA, TN,TX, VA, and WI. | Terminated | January 16, 2014 | 190 | 111800 | Not_Computer
Z-0956-2013 | 64194 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Variable Angle Locking Compression Plate/VA-LCP Dorsal Distal Radius Plate, Part Number 02.115.431. |  |  | January 11, 2013 | 2013 | Synthes USA HQ, Inc. | It was discovered that one of the VA-LCP Dorsal Distal Radius Plates was incorrectly etched as Part Number 02.115.431 (110 degree) instead of the correct Part Number 02.115.231 (90 degree). | Labeling mix-ups | The firm initiated their recall of this product on January 11, 2013 by sending an Urgent Notice: Medical Device Recall letter to consignees. The letter identified the affected product and the issue. Synthes requested customers to examine their inventory for the affected product, remove them from use, and return them to Synthes. Additionally, customers were to follow the steps provided in the event that they did or did not have any affected product on hand. Questions should be directed to 610-719-5450 or FieldAction@synthes.com. |  | 4 | 4 | Distributed in the states of IN, OH, PA, and TX. | Terminated | August 31, 2015 | 962 | 4 | Not_Computer
Z-0027-2014 | 66346 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | P21x/5-1 MHz transducer, Part Number P07698-23 |  |  | September 11, 2013 | 2013 | SonoSite, Inc. | FUJIFILM SonoSite Inc. remove P21x/5-1 Transducer because the firm discovered the Biopsy Bracket Sensor is not working as intended. | Error in labeling | SonoSite started calling their consignees on September 13, 2013. For customers who could not be reached via phone, a letter Urgent -Medical Device Removal will be sent on September 17, 2013. Affected customers will be offered a replacement transducer (P07698-24) free of charge. Consignees in the U.S. can call 1-877-657-8118, Monday-Friday, 6:00 a.m. - 5:00 p.m. (PST) or send E-mail to service@sonosite.com. Customers outside the US can call 425-951-1200. |  | 10 | 10 units | Worldwide Distribution - USA Nationwide in the states of Florida, Georgia, Oklahoma, Pennsylvania, Tennessee, Texas, West Virginia, and in the country Japan and Italy. | Terminated | December 11, 2013 | 91 | 10 | Not_Computer
Z-2272-2013 | 66184 | No | CAI | Anesthesia breathing circuit. | Anesthesiology | Anesthesiology | 510(K) Exempt | Universal Flex2 Breathing Circuit |  |  | August 23, 2012 | 2012 | King Systems Corp. | King Systems received a customer complaint indicating that the customer had ordered a breathing circuit with a Latex breathing bag. The product they received contained a Latex breathing bag, but was labeled as non-latex. | Labeling mix-ups | Ambu, Inc. sent a letter dated June 24, 2013, informing the customer that they have decided to discontinue the manufacture and sales of the King Systems latex breathing bags and to conver all customers to their King Systems latex-free breathing bags. The customer was advised to adjust their inventories accordingly and transition to the latex free part number as their inventory levels reach their minimums. For questions the customer was advised to call 317-403-8677.For questions regarding this recall call 317-776-6823. |  | 41 | 1 case; 40 units | Nationwide Distribution - NE only. | Terminated | November 18, 2013 | 452 | 41 | Not_Computer
Z-0891-2013 | 64434 | No | GDX | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Technocut Plus, Sterile Single Use Safety Scalpel, Distributed by: Myco Medical. |  |  | February 07, 2013 | 2013 | Myco Medical Supplies Inc | Pouch labeling/printing errors and defective packaging resulting in possible compromised sterility. | Packaging | MYCO Medical sent an Advisory Notice dated February 7, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers wsere instructed to review their inventory for the affected product and immediately quarantine the affected product per their internal procedures. Customers were also instructed to complete the attached Material for Return form and return to MYCO via email QA@mycomedical.com or fax to 919-800-3919. Customers were have the returned material packaged and ready for UPS to pick up at their location. If product was shipped to other customers, customers were asked to contact a MYCO Customer Service Representative with the additional customer's contact and ship-to information. Customers with questions were instructed to call 919-460-2535.For questions regarding this recall call 919-460-2535, ext 105. |  | 110768 | 110,768 scalpels in individual pouches | Worldwide Distribution - USA including CA, FL, IL, MI, MO, NY, OH, PA, SC, TN, WI and Internationally to Canada | Terminated | July 09, 2013 | 152 | 110768 | Not_Computer
Z-1998-2013 | 65413 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | 00597001101 Femur CR PRECOAT FEM COMP SIZE AML, Rx, Sterile; |  |  | June 07, 2013 | 2013 | Zimmer, Inc. | Surgeons are implanting components of the NexGen Cruciate Retaining (CR) Complete Knee System in combinations that are not approved (micro with standard). Includes CR Micro Femurs, Articular Surface, Patellae, instruments and provisionals combined with standard components. A March 2012 recall reiterated the proper combination labeling, but the mis match problem continues. Zimmer is removing the d | Labeling design | Beginning June 7, 2013 all distributors will be sent Urgent Device Recall notices via electronic mail, with distributors, hospital risk managers with affected inventory also receiving the notice via courier. The Notice outlined the products recalled; the reason for the recall, possible adverse events related to the recall and directed consignees to locate, quarantine and count all affected product recording the information on the Inventory Return Form acknowledging that recall directions were followed. Further questions should be directed to 1-877-946-2761. Affected product is to be returned to Zimmer Product Service Department , 1777 West Center St. Warsaw, IN 46580 or internation return at Zimmer International Logistics GmbH, Max - Immelmann Allee 12 79427 Eschbach Germany."Communications outside of the United States will occur approximately two weeks after the United States communications. |  | 192355 | 192,355 all devices | Worldwide Distribution-USA (nationwide) and the country of Japan. | Terminated | December 15, 2014 | 556 | 192355 | Not_Computer
Z-0129-2014 | 66196 | No | NDG | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Ti-Magna Fx Washer |  |  | August 28, 2013 | 2013 | Zimmer, Inc. | Zimmer recently conducted a review of historic packaging validations completed. Based on this review, it was determined that a subset of product packaged at a single Zimmer facility required testing to ensure that previously packaged product from this facility met the established requirements for sterile barrier integrity. The number of samples required for packaging testing, to provide statisti | Packaging process control | The firm, Zimmer Inc., sent a letter entitled "Voluntary Recall of Ti-VFxII Condylar Tube Plates and Ti-Washers" dated September 12, 2013 to its customer. The letter described the product, problem and action to be taken. The customer was instructed to immediately discontinue use of the listed products. A Sales representatives from Zimmer K.K. or agencies will be in contact with you. Zimmer Inc. is planning to resume product manufacturing upon the completion of package validation review. Zimmer K.K. will contact you when a detailed schedule has been determined.If you have any questions, contact Zimmer K.K. Safety Office (9:00 to 17:00; closed weekends and public holidays) TEL: 0120-015-180. |  | 2632 | 2,632 devices | International distribution: Japan only. | Terminated | April 30, 2015 | 610 | 2632 | Not_Computer
Z-0973-2013 | 64633 | No | NDP | Endosseous dental implant accessories. | Dental | Dental | 510(K) Exempt | PrimaConnex Tapered Implant TC, SD Internal Connection 3.5 x 10mm |  |  | March 11, 2013 | 2013 | Keystone Dental Inc | Mislabeled: Cover screw that was in the package was a WD cover screw and not a SD cover screw as labeled | Packaging process control | Keystone Dental initially called customers on 3/11/13 and forms were completed to document these calls. In addition, a letter sent via Federal Express/or hand delivered to all distributors and customers. Product requested to be returned. |  | 120 | 120 units | Worldwide Distribution-USA including DC and the states of CA, FL, IN, MA, MN, NJ, NY, NH, PA, WI, and NY, and the countries of Republic of China, Germany, and France. | Terminated | September 26, 2013 | 199 | 120 | Not_Computer
Z-1996-2013 | 65738 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | GENESIS(R) II CRUCIATE RETAINING ARTICULAR INSERT, TRIAL, SIZE 3-4, A/P 48 MM, M/L 68 MM, QTY 1, NON-STERILE, REF 71430490, |  |  | July 08, 2013 | 2013 | Smith &amp; Nephew Inc | The products were produced using the wrong color material, resulting in the devicesbeing red (used for deep dished inserts) instead of blue (used for cruciate retaining inserts). | Labeling mix-ups | Smith &amp; Nephew sent an Urgent Recall Notification letter dated July 8, 2013 to all affected customers. Surgeons were notified by letter beginning on July 16, 2013. The letter identified the affected product, problem and actions to be taken. Customers were instructed to inspect inventory, locate unused devices, quarantine immediately, and return affected product to the address provided in the notification. For questions email to FieldActions@smith-nephew.com. |  | 9 | 9 units | International only: Germany, France, United Arab Emirates, and United Kingdom. | Open, Classified |  |  | 9 | Not_Computer
Z-0974-2013 | 64435 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | INTIO Inc., ClearStart-SVM Segmentation and Volumetric Measurement System, Users Manual. Part #: UM-06-20-0003. ClearStart-SVM is used for patient assessment using CT (computed tomography) images, in order to plan potential therapies or monitor tumor response to treatment. |  |  | February 26, 2013 | 2013 | Intio Inc | INTIO Inc. distributed a SVM User's Manual and it has an editing mistake that could be interpreted as promoting an unapproved use. | Labeling False and Misleading | INTIO sent a Recall Notice to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. INTIO has corrected this version of the SVM User's Manual. An INTIO Inc. employee will contact the customers to replace the current manual with the correct one. Consignees were notified by letter that an INTIO Inc. employee will contact you within the next 30 days to replace the current manual with the correct one. The new SVM User's Manual will be provided to you once INTIO Inc. has retrieved the incorrect version of the SVM User's Manual. Customers with questions were instructed to email support@intio.us or call 303-396-1777.For questions regarding this recall call 303-396-9291. |  | 7 | 7 | Nationwide Distribution including MN, FL, PA, WI, NY, and RI. | Terminated | April 17, 2013 | 50 | 7 | Not_Computer
Z-0020-2014 | 66047 | No | CEO | Phosphorus (inorganic) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Clinical Chemistry Phosphorus Reagent |  |  | August 12, 2013 | 2013 | Abbott Laboratories, Inc | Reagent lots 64736UN12, 03991UN13, 51611UN13, 74805UN12, and 03992UN13 have incorrect expiration dates assigned by Supplier. | Error in labeling | The firm, Abbott Laboratories, sent a "PRODUCT CORRECTION" letter dated August 12, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to update existing inventory and manually track the implanted lots with the corrected expiration dates listed in the letter; retain a copy of letter and if you have forwarded any of the lots listed to another laboratory, provide a copy of the letter to them.If you or any of your healthcare providers you serve have any questions regarding this information, U.S. Customers please contact Customer Service at 1-877-4ABBOTT (422-2688). Customers outside the U.S., please contact your local area customer service. |  | 7213 | 7213 | Worldwide distribution: US (nationwide) including Puerto Rico and countries of: Albania, Algeria, Argentina, Austria, Australia, Azerbaijan, Bangladesh, Bahamas, Belgium, Bosnia &amp; Herzegovina, Brazil, Bulgaria, Burkina, Canada, Cayman Islands, Chile, China, Columbia, Croatia, Curacao, Cyprus, Czech Republic, Denmark, Dominican Republic, El Salvador, Egypt, Estonia, Ethiopia, Finland, France, Germany, Greece, Guatemala, Hong Kong, India, Iraq, Ireland, Israel, Italy, Ivory Coast, Jamaica, Jordan, Kazakhstan, Kenya, Kuwait, Latvia, Lebanon, Libya, Lithuania, Luxembourg, Malaysia, Mali, Mauritania ,Mexico, Moldova, Morocco, Netherlands, New Zealand, Norway, Oman, Pakistan, Panama, Paraguay, Philippines, Poland, Portugal, Qatar, Republic of Yemen, Romania, Russia, Saudi Arabia, Senegal, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Tajikistan, Taiwan, Tanzania, Thailand, Trinidad &amp; Tobago, Tunisia, Turkey, Turks and Caicos, United Arab Emirates, United Kingdom, Uruguay and Vietnam. | Terminated | April 06, 2015 | 602 | 7213 | Not_Computer
Z-0014-2014 | 65955 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Synthes Midfoot Fusion Bolt 6.5 mm. It is indicated for fracture, fixation, osteotomies, nonunions, and fusions of large bones in the foot and ankle. |  |  | March 08, 2013 | 2013 | Synthes USA HQ, Inc. | A post market safety review following several customer complaints indicated additional clarification was necessary in the Midfoot Fusion Bolt Technique Guide to reduce potential for bolt back-out or migration by emphasizing the need for supplemental fixation. | Labeling design | An "Urgent Notice: Medical Device Recall" letter dated March 8, 2013 was sent to customers asking them to check their inventory, remove all affected product technical guides from use, and return them to Synthes along with the Verification form. Customer questions were directed to (610) 719-5450. |  | 62 | 62 | Nationwide Distribution, including the states of MI, PA, SC, IA, MA, CA, WI, IL, and KS. | Terminated | August 12, 2015 | 887 | 62 | Not_Computer
Z-2091-2013 | 65513 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Free Lock Femoral Hip Fixation System Compression Tube/Plate, different length (59, 75, 91, 107, 123 mm), holes (2, 3, 4, 5, 6) and degree angles (130, 135, 140, 145 and 150). Item codes starting with 00-1181-130, 00-1181-135, 00-1181-140, 00-1181-145, 00-1181-150). |  |  | June 19, 2013 | 2013 | Zimmer, Inc. | Single packaging configuration used for the VERSA-FX-II Femoral Fixation System failed the leak test compromising the sterility of the product. | Packaging process control | The firm, Zimmer, sent an "URGENT MEDICAL DEVICE NOTIFICATION" letter dated July 11, 2013 to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to inspect the device according to the insert instructions; not use the device if damage to the corner of the inner tray is observed, and contact the firm for return and questions. Customers should also keep this notification in mind and inform Zimmer of product with this condition using zimmer.per@zimmer.com.If after reviewing this notification you have questions or concerns, please contact Zimmer at 1-877-946-2761. |  | 85 | 85, 794 units in total | Worldwide distribution: US Nationwide (including PR and USVI), and countries of: Israel, Argentina, Brazil, France, Mexico, Saudi Arabia, Beirut, China, Korea, San Salvador, Honduras, Canada, Australia, Switzerland, Germany, Arab Emirates, UK, Belgium, Egypt, Spain, France, Italy, Iran, Lebanon, Jordan, Iraq, Canary Islands. | Terminated | June 16, 2015 | 727 | 85 | Not_Computer
Z-0276-2014 | 66450 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | 3.5MM LOCKING SCREW, T15 HEXALOBE, SELF TAPPING, 34MM |  |  | September 26, 2013 | 2013 | OrthoPediatrics Corp | The firm identified, through complaints that the 3.5MM LOCKING SCREW packaging material was marked as a 3.5mm screw and the part number was correct but that the description indicated a 7.0mm screw. Upon investigation, it has been determined that a lot number 009U2 was packaged and mislabeled internally by OrthoPediatrics. | Labeling mix-ups | OrthoPediatrics sent an Urgent Medical Device Recall letter dated September 26, 2013, to all affected customers.The letter identified the product the problem and the actions needed to be taken by the customerThe sales representatives and distributors who have been identified via shipping records as having received one or more discrepant devices affected by this Recall were requested to quarantine the devices identified in this Recall by telephone on 18-September-2013. If you or any of your customers have any of these devices in the original packaging please take possession of them and quarantine them in preparation for returning them to OrthoPedlatrics. They are misbranded. If the Screw has already been removed from the original packaging and the original packaging has been discarded and the Screw placed in a sterilization tray, it need not be returned.For further questions please call ( 574) 268-6379. |  | 7 | 7 units shipped | Worldwide Distribution - US including Michigan and the countries of Ireland and Australia | Terminated | December 26, 2013 | 91 | 7 | Not_Computer
Z-2096-2013 | 65874 | [6 empty] | OLA Buffer |  |  | July 12, 2013 | 2013 | Celera Corporation | The package insert for FG OLA Buffer, GPR displayed the wrong concentration of ÂżM instead of the correct mM unit. | Error in labeling | Celera sent a customer notification letter dated July 12, 2013 by e-mail and mailed on July 19, 2013 via first class mail to all affected customers. The letter identified the affected product, problem and actions to be taken. The information is being provided for records and assures customer that there were not quality issues with the product and it may be used as usual in laboratory test procedures. For questions call 1-810-749-4200 or contact your local authorized Celera representative. |  | 43 | 43 units | USA Nationwide in the state of CA | Terminated | October 01, 2013 | 81 | 43 | Not_Computer
Z-0902-2013 | 64383 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Century Console Base, Model #043625051b. |  |  | January 31, 2013 | 2013 | Salyer Prn Biomedical | Concerns that the addition of the UPS Battery Backup to the Century Console Base Shelf did not comply with the FDA premarket notification regulations. | No Marketing Application | An Urgent: Medical Device Recall letter, dated 1/31/13, was sent to all the customers that received the Century Console Base Shelf option with the Linear/1.5k 115v UPS Battery Backup for the Century Heart Lung Machine. The letter informed the customers that Salyer PRN Biomedical will have a Salyer Biomedical technician return the customers' Century Console Bases to their original configuration without the Shelf option and restore the Linear/1.5k UPS Battery Backup to its original condition as a stand-alone unit with its own wheel base for ease of use. A refund or credit will be issued for the return of the product. Customers are instructed to contact Jeff Poland at (314) 497-5232. |  | 9 | 9 units | Distributed in the states of MA,CA, and WI. | Terminated | April 15, 2013 | 74 | 9 | Not_Computer
Z-2058-2013 | 65631 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Synthes Trauma Nail System. The devices are indication for bone fixation. |  |  | November 02, 2012 | 2012 | Synthes USA HQ, Inc. | Recall was initiated due to the possibility that the outer pouch was compromised. The product is packaged in two pouches; a sterile, internal pouch within an external pouch. It is the outer, external pouch that may be compromised. | Packaging change control | The firm initiated their recall of this product on November 2, 2012 by sending an "URGENT NOTICE: MEDICAL DEVICE RECALL" letter to all consignees. The letter explained the problem and gave instructions for consignees to examine their inventory and urged consignees to return the identified devices. The letter stated If returning the product would adversely impact their ability to provide necessary medical care to patients, they should re-sterilize the product per Pre-Vacuum Sterilization instructions contained in the Instruction for Use. |  | 1011783 | 1,011,783 | Nationwide distribution. | Open, Classified |  |  | 1011783 | Not_Computer
Z-0794-2013 | 64069 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | SeriScaffold Surgical Scaffold |  |  | January 08, 2013 | 2013 | Allergen Medical | Product is dual packaged in an inner and outer pouch, the outer pouch seal may be compromised and sterility may not be assured | Packaging process control | Allergan sent an "URGENT MEDICAL DEVICE RECALL" letter dated January 8, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers are requested to quarantine and return any unused devices in inventory . In addition, customers are asked to complete and return the acknowledgement form. Customers with questions should contact Allergan at 781-658-2800, Mon- Fri, 9 a.m. to 5 p.m. EST. |  | 225 | 225 units (US) and 81 (Outside US) | Worldwide Distribution-USA including the states of CA, FL, GA, MI, MN, NJ, NY, OH, OR, RI, TX, VA, and WA, and the countries of Germany, Spain, and UK. | Terminated | July 29, 2013 | 202 | 225 | Not_Computer
Z-0277-2014 | 66199 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | VerSysÂż Cemented Revision Calcar (CRC) Hip System Femoral Stems |  |  | August 28, 2013 | 2013 | Zimmer, Inc. | Based on a review of historic packaging validations, it was determined that a subset of product packaged in one of the clean rooms at a single Zimmer facility required testing to ensure that previously packaged product from this facility met the established requirements for sterile barrier integrity. The number of samples required for packaging testing (seal integrity and seal strength) to provid | Packaging process control | Removal of the subject devices from the field including distributor and hospital inventories. Distributors and Risk Managers and Surgeons at hospitals that received a direct shipment from Zimmer or its Distributors received a notification on 09/19/2013. |  | 12866 | 12866 units | Worldwide distribution: US (Nationwide) including Puerto Rico and Virgin Islands; and countries of: Australia, Algeria, Belgium, Bolivia, Canada, Canary Islands, Chile, Croatia, Finland, France, Germany, Greece, Guatemala, Honduras, India, Israel, Italy, Japan, Jordan, Lima-Peru, Malaysia, Morocco, Nicaragua, Portugal, Saudi Arabia, Seoul Korea, Serbia, Singapore, South Africa, Spain, Switzerland, Turkey, Utd.Arab.Emir., United Kingdom, and Venezuela. | Terminated | April 30, 2015 | 610 | 12866 | Not_Computer
Z-2054-2013 | 65787 | No | GXP | Methyl methacrylate for cranioplasty. | Neurology | Neurology | 510(k) | Stryker Part Numbers 79-43903 (3cc), 79-43905 (5cc), 79-43910 (10 cc), 79-43915 (15cc), 397003 (3cc), 397005 (5cc), 397010 (10 cc), 397015 (15 cc) HydroSet Injectable HA Bone Substitute Sterile R/EO, nonpyrogenic, Product of Ireland |  |  | July 30, 2013 | 2013 | Stryker Craniomaxillofacial Division | Stryker received reports of the desiccant bag in the kit leaking silica gel granulate. This could increase exposure to anesthesia during surgical delay up to 60 minutes or inadvertent implantation of silica gel granulate which could cause foreign body reaction including inflammation and might lead to the need for revision surgery. | Package design/selection | Stryker sent an Urgent Medical Device Recall Notices with attached Product Correction Bulletin dated, July 30, 2013, describing the product, problem and possible adverse events associated with the recall were sent to all affected customers. Customers were instructed to pass the notice and bulletin to all individuals who may be using the device, provided instructions on procedures which would help mitigate risk and return the customer response form to our Customer Service Team by email, fax or letter.Email: CMF-custserv@stryker.comFax: 1-877-648-7114Address: 750 Trade Centre Way, Suite 200Portage, MI 49002 Customers responses will be tracked and non-responders will be contacted.Further questions please call (877) 534-2464. |  | 12719 | 12,719 | Worldwide Distribution - USA (nationwide) including the countries through Stryker Distribution: Austria, Belgium, Denmark, Estonia, Finland, France, Germany, Ireland, Italy, Latvia, Luxembourg, Netherlands, Norway, Poland, Portugal, Romania, Serbia, Spain, Sweden, Switzerland, United Kingdom, Argentina, Australia, Brazil, Jordan, Kingdom of Saudi Arabia, Kuwait, Latin America (others), New Zealand, Russia, Singapore, South Africa, and United Arabian Emirates. | Terminated | April 17, 2014 | 261 | 12719 | Not_Computer
Z-0001-2014 | 66365 | No | OFA | Cardiovascular surgical instruments. | Cardiovascular | Cardiovascular | Enforcement Discretion | Ultima OPCAB System, Sterile, Rx Only, |  |  | September 20, 2013 | 2013 | Maquet Cardiovascular, LLC | MAQUET has identified a potential issue with the ULTIMA OPCAB (part numbers OM-2003S &amp; OM-2001D). Labeling for the 2 blade lots had been incorrectly applied. The standard blades from Lot #25052176 were incorrectly labeled as deep blades, and the 20 deep blades from Lot #25052200 were incorrectly labeled as standard blades. | Error in labeling | MAQUET sent an "Urgent Medical Device Recall" letter and "Field Recall Response Forms" dated September 20, 2013 to customers via Federal Express. The notification identifies the affected product; informs the customers of the nature of the issue; and advises of potential adverse patient consequences that may occur. The customer is asked to complete and return the Field Recall Response Form. If the customer has questions they are directed to contact their field representative or MAQUET Technical Support (888.627.8383 between 8 am and 5 pm EST Monday through Friday). |  | 20 | 20 units (Lot 25052176); 12 units (Lot #25052200) | USA Nationwide Distribution in the states of: Florida, Maryland, New Jersey, New York and Texas. | Terminated | June 10, 2014 | 263 | 20 | Not_Computer
Z-0003-2014 | 66026 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Cordis ADROIT 6F Guiding Catheter |  |  | August 12, 2013 | 2013 | Cordis Corporation | Due to a labeling error related to the conversion of the inner diameter of the catheter from inches to millimeters. The English dimension is correct and is the primary dimension. | Error in labeling | Cordis sent an Urgent Medical Device Recall letter dated August 12, 2013 Acknowledgement Form to all affected customers. The letter identifed the affected product, problem and actions to be taken. Customers were instructed to immediately identify and set aside all affected product in a manner that ensures the product will not be used. Customers were asked to review, complete, sign and return the enclosed Acknowledgement form directly to Cordis to the fax number on the form. For questions call the Customer Service number at 800-561-7683. |  | 1238 | 1238 units | Worldwide Distribution - USA Nationwide in the states of AZ, CA, FL, ID, LA, MA, MD, MI, MN, MO, NC, NH, NY, OH, OK, PA, SC, TX, VA, WI, and WV and the countries of: Austria, Belgium, Czech Republic, France, Germany, Greece,Hungary, Israel, Italy, Latvia, Netherlands, Poland, Portugal, Russian Federation, Slovenia, Spain, Sweden, Switzerland, United Arab, and United Kingdom | Terminated | September 30, 2014 | 414 | 1238 | Not_Computer
Z-0004-2014 | 66357 | No | MXE |  |  | General &amp; Plastic Surgery | 510(k) | SHOPKO Waterproof Adhesive Pads Antibacterial. |  |  | August 21, 2013 | 2013 | ASO, LLC | No antibacterial bandages were packed in boxes labeled as antibacterial. | Labeling Change Control | A letter was sent via e-mail and mail to the only customer on August 23, 2013 that a recall of two lots was in progress and to examine their inventory and stop distribution of the product. Product was to be either place in quarantine or destroyed. A credit for all returned product or by returned completed certificate of destruction. |  | 37 | 37 cases (72 retail boxes in each case) | Distributed in the states of ID, WI, IN, and NE. | Terminated | August 04, 2014 | 348 | 37 | Not_Computer
Z-2107-2013 | 65914 | No | GAA | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | SpineNet Bone Marrow Aspiration Needle Kit (which is the Ranfac Bone Marrow Aspiration Needle #74150-01M). |  |  | February 04, 2013 | 2013 | SpineNet | SpineNet was marketing this device to be used with SpineNet ACC device without FDA clearance or approval. | No Marketing Application | SpineNET sent a "VOLUNTARY PRODUCT RECALL" letter dated 02/04/2013 to Halifax Medical Center. The letter identified the product, problem, and actions to be taken by the customer. Contact the firm at 407-539-2483 for questions regarding this notice. |  | 100 | 100 units | Distributed to one customer in Florida. | Terminated | November 15, 2013 | 284 | 100 | Not_Computer
Z-2103-2013 | 65917 | No | ODP | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | SpineNet Anterior Cervical Cage, Lordotic, Peek (originally submitted as Daytona Anterior Cervical Cage) |  |  | February 05, 2013 | 2013 | SpineNet | The recall decision has been made because because it has recently been determined that the instructions for use (IFU) had certain statements that should have been removed prior to initial product marketing. | Error in labeling | SpineNET sent a Customer Notification letter dated 02/05/2013 to their only customer (Halifax Medical Center). The letter identified the affected product, problem and actions to be taken. Customers were informed to remove any copies of the original IFU-ACC-01 form from their records and replace with the IFU-ACC-02 form attached to the letter. For questions contact SpineNET office directly. |  | 400 | 400 | USA Nationwide Distribution in the state of Florida | Terminated | May 02, 2014 | 451 | 400 | Not_Computer
Z-0002-2014 | 66189 | No | CAE | Oropharyngeal airway. | Anesthesiology | Anesthesiology | 510(K) Exempt | King LTS-D Oropharyngeal Airways |  |  | December 13, 0012 | 12 | King Systems Corp. | On December 13, 2012, King Systems initiated a voluntary recall of one (1) case(s) (5) individual products) from one (1) lot [(IV1V3) 62 cases, 310 products]of King LTSD Oropharyngeal Airways because the package label indicated that the airway was a size 5 when in fact the airway in the package was a size 4. When King manufacturers a lot, all of the components are the same size, part number etc. I | Error in labeling | King Systems notified the only one (1) consignee who received the affected product of the problem via a customer complaint. Five (5) of the five (5) affected products were destroyed by consignee. |  | 6 | 1 case; 5 units | USA Nationwide Distribution in the state of NY | Terminated | November 08, 2013 | 730815 | 6 | Not_Computer
Z-2311-2013 | 65697 | No | DSY | Vascular graft prosthesis. | Cardiovascular | Cardiovascular | 510(k) | Advanta SST Grafts &gt;or= 6mm |  |  | July 12, 2013 | 2013 | Atrium Medical Corporation | Inner tray label may adhere to wall of outer tray which make it difficult to remove the tray onto the sterile field | Package design/selection | Atrium Medical Corporation issued an Urgent Medical Device Safety Alert letter on July 16, 2013, via UPS to customers and e-mail to the field representatives that represent these accounts. The notification identifies the problem, product, and risk factors. In addition, it provides instructions how to handle product that may be affected.For questions customers were instructed to contact their Atrium Medical Customer Service at 603-880-1433 Prompt 1 or by e-mail at returns@atriummed.com |  |  |  | Worldwide Distribution - USA (nationwide) and internationally to the following countries:American Samoa AntillesArgentina Australia Austria Bangladesh Belgium Bolivia Brazil Canada Chile Colombia Egypt El Salvador Finland France Germany Great Britain Greece Guatemala Hong Kong India Indonesia Ireland Italy Jamaica Japan Korea Lebanon Libya Malaysia MauritiusMexico Neal New Zealand Nicaragua Norway PakistanPanama Peru Philippines Poland Portugal Puerto Rico Romania Russia Saudi Arabia Singapore Slovenia South Africa Spain Switzerland Taiwan Thailand the Netherlands Trinidad Tunisia | Terminated | July 15, 2014 | 368 |  | Not_Computer
Z-2101-2013 | 65850 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | ZIMMER PERIARTICULAR LOCKING PLATE SYSTEM |  |  | July 30, 2013 | 2013 | Zimmer, Inc. | Zimmer is initiating a lot specific voluntary recall of the Distal Medial Humeral Locking Plate, Short due to the incorrect information printed on the Chinese label. The original label indicates that implants are RIGHT sided, but the Chinese label indicates that the implants are LEFT sided. | Labeling mix-ups | Zimmer sent an Urgent Medical Device Recall Notification letter dated July 30, 2013 to Distributors and Hospitals in China. The letter described the affected product, problem, labeling issue, directions on how to respond if affected product is discovered, and potential health risks associated with the labeling error. An inventory return certification form is included with the recall notification. Questions or concerns are directed to Zimmer Shanhai at 021-2211 5196, Monday to Friday, between the hours of 9 a.m. and 5 p.m. |  | 11 | 11 units | Internationally Distributed in China only. | Terminated | April 15, 2014 | 259 | 11 | Not_Computer
Z-2113-2013 | 65764 | No | NKC | Laparoscopic insufflator. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(K) Exempt | DeRoyal Insufflation Tubing sets: |  |  | July 15, 2013 | 2013 | DeRoyal Industries Inc | DeRoyal Industries is recalling their product Insufflation Tubing Sets due to materials incompatibility which could possibly result in occlusion at or near the filter housing which could prevent the flow of gas through the device. | Packaging process control | The firm notified customers via letters dated 7/15/2013 sent UPS and via DeRoyal Representatives. Customers were asked to notify DeRoyal of any remaining inventory to return for replacement product or credit. Non-responders will receive additional notifications. Effectiveness checks will be performed by contacting non-responders by phone. |  | 524486 | 506,632 sets in US; 17,854 sets International | Worldwide Distribution-USA (nationwide) including DC and the states of AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, and the countries of Dubai, Arab Emirates, Quebec, Canada, Japan, and Ireland. | Terminated | May 20, 2015 | 674 | 524486 | Not_Computer
Z-0144-2014 | 66040 | No | HNJ | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | MLS Torpedo Mini-Light Ophthalmic single-port illumination. |  |  | June 27, 2013 | 2013 | Insight Instruments, Inc. | The product was registered and distributed as a Class I medical device and was subsequently determined to be a Class II medical device which requires FDA market clearance. | No Marketing Application | Insight Instruments, Inc. sent a letter dated April 8, 2013, via USPS priority mail to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue using and dispensing all affected product, perform a physical count of their inventory and complete the Return Response Form and fax to 772-219-9342. Customers were also instructed to call 800-255-8354 to arrange for return and prompt reimbursement.For questions regarding this recall call 727-219-9393. |  | 10212 | 10212 | Worldwide Distribution - USA (nationwide) including AZ, CA, FL, IN, MA, MD, TX, and WA and Internationally to Czech Republic, Germany, Netherlands, Spain, Switzerland, and Turkey. | Open, Classified |  |  | 10212 | Not_Computer
Z-0825-2013 | 63933 | No | GEH | Cryosurgical unit and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | HealthTronics Endocare 3.8mm RenalCryo Cryoprobe, R3.8L. The Endocare CryoProbe is an accessory probe intended for use with the CryoCare Surgical System, for freezing and ablation of tissue. |  |  | December 20, 2012 | 2012 | Healthtronics, Inc. | The Endocare 3.8mm RenalCryo Cryoprobe is being recalled because the label on the outer box has the incorrect expiration date. | Incorrect or no expiration date | Healthtronics sent an Urgent Device Recall letter dated December 20, 2012, to all affected customers that allows tracking. The letter included a suggested format/content for a correction status/acknowledgement letter. The customer was requested to complete the verification section of the recall letter and return all affected products to Healthtronics. A return material authorization (RMA) and pre-paid FedEx return labels was provided for all customers. All affected Healthtronics field personnell were notified of the recall via email.For questions regarding this recall call 512-439-8355. |  | 50 | 50 units | Nationwide Distribution including CA, CO, FL, GA, NC, NJ, OH, TX, and WI | Terminated | April 22, 2013 | 123 | 50 | Not_Computer
Z-0272-2014 | 66689 | No | LFD |  |  | Dental | 510(k) | Biotene Moisturizing Mouth Spray, with xylitol, For dry mouth symptom relief, 44.3ml/1.5 FL. OZ., Sugar Free Sweetened with xylitol and sorbitol, GlaxoSmithKline, Consumer Healthcare, L.P., Moon Township, PA 15108. |  |  | October 17, 2013 | 2013 | GlaxoSmithKline, LLC. | Label indicates, "Sweetened with xylitol and Sorbitol", however product does not contain sorbitol. | Labeling Change Control | GlaxoSmithKline sent an Urgent Medical Device Recall letter dated October 17, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to return all affected product. If product was further distributed customers were instructed notify their customers of the recall and the actions required. Customers were asked to complete and return the enclosed business replay card even if they have no affected product in their inventory. Customers should complete the return goods packing slip, enclose it with their return goods shipment and affix the return label to the package. Customers with questions should call Customers Service at 1-800-743-4014.For questions regarding this recall call 919-269-1194. |  | 483204 | 483,204 bottles | Nationwide Distribution . | Terminated | July 26, 2017 | 1378 | 483204 | Not_Computer
Z-0200-2014 | 65863 | No | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | Product Catalog (part number 20-0020) utilized with the Electric and Pneumatic Motor Systems including: MicroMax Motor Systems, XMax Motor Systems, EMax2 Motor Systems, and Emax2Plus Motor Systems, all models and serial numbers. |  |  | May 06, 2011 | 2011 | The Anspach Effort, Inc. | Anspach Effort, Inc. Palm Beach Gardens, FL initiated a correction of its medical Device labeling for its Pneumatic Systems Operating Manual, Cutting Burr Chart and Product Catalog due to a recent review of their regulations and compliance requirements that determined the claim that the device is capable of dissection of biometals, bioplastics and cement has been verified, but is is not part of th | Labeling Change Control | The firm, Anspach Effort, Inc., sent an "Urgent Medical Device Labeling Correction" letter dated June 27, 2011 to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to do the following: 1) Destroy or delete any items in their inventory immediately and replace with the attached revised Operating Manual, Product Catalog and/or Cutting Bur Chart. 2) Complete and return the attached Customer Reply Form indicating disposition of the labeling materials and confirming their receipt of this letter, and fax to 1-800-327-6661. 3) If they distributed any of the products to other services or facilities, please forward this information as appropriate. For additional copies, contact Anspach Customer Support at (800) 327-6887. |  | 2020 | 2,020 | Worldwide Distribution: USA Nationwide and the countries of: Egypt, India, Germany, Turkey, Canada, Saudi Arabia, Greece, United Kingdom, Cayman Islands, Kazakhstan, Republic of Korea, Taiwan, Hong Kong, China, Lebanon, Argentina, Denmark, Belgium, Netherlands, Mexico, France, Australia, Panama, Italy, Guam, United Arab Emirates, Israel, Polan, Venezuela, Austria, Hungary, Lithuania, Norway, Malaysia, Colombia, Costa Rica, Ecuador, Jamaica, Peru, Thailand, Italy, Czech Republic, Cameroon, South Africa, Portugal, Finland, Sweden, Nigeria, Bahamas, Cambodia, Dominican Republic, Switzerland, Chile, Singapore, Spain, Syrian Arab Republic, Jordan, Japan. | Terminated | January 28, 2016 | 1728 | 2020 | Not_Computer
Z-2190-2013 | 65854 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | HemoCueÂż Glucose 201 Microcuvettes, 25 Microcuvettes Individually Packaged. |  |  | July 15, 2013 | 2013 | HemoCue AB | During a complaint investigation HemoCue AB discovered punctured HemoCue Glucose 201 single pack pouches. | Packaging | HemoCue sent an Urgent Field Safety Notice dated July 12, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to quarantine any affected product and Return the Field Safety Notice Verification Form to e-mail: 2013glufsn@hemocue.com or fax 562-668-5794 Customers would be contacted to coordinate return and replacement of the affected microcuvettes. Customers with questions were instructed to call 800-426-7256.For questions regarding this recall call 440-925-3485. |  | 2221500 | 22,215 boxes of 100 cuvettes distributed in US, 34, 841 distributed outside the US | Worldwide Distribution - USA (nationwide) and internationally to AUSTRALIA, BELGIUM, CANADA, COTE D'IVOIRE, DENMARK, EGYPT, FINLAND, FRANCE, GEORGIA, GERMANY, GREECE, IRELAND, ITALY, KENYA, NETHERLANDS, NEW ZEALAND, NORWAY, OMAN, POLAND, QATAR, RUSSIAN FEDERATION, SPAIN, SWEDEN, SWITZERLAND, UNITED KINGDOM, VIET NAM &amp; YEMEN.. | Terminated | September 02, 2014 | 414 | 2221500 | Not_Computer
Z-1276-2013 | 64988 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | ExoShape Soft Tissue Fastener 12 mm x 30 mm REF 1101-00-1230. |  |  | April 04, 2013 | 2013 | Medshape Solutions, INC. | Expiration date on the patient label and date on the product carton were not the same. | Error in labeling | Medshape sent an "URGENT: MEDICAL DEVICE RECALL" letter to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. |  | 26 | 26 | Nationwide Distribution including the states of CO, FL, GA, IA, MD, MO, OR, TX, VT, VA, and WI. | Terminated | August 07, 2013 | 125 | 26 | Not_Computer
Z-1887-2013 | 65482 | No | DTK | Cardiovascular intravascular filter. | Cardiovascular | Cardiovascular | 510(k) | Cordis OPTEASE Retrieval Inferior Vena Cava (IVC) Filter***REF Cat. No.***466F220A and 466F220B***Vena Cava Filter and Introduction Kit***(BRITE TIP Catheter Sheath Introducer and Angiographic Vessel Dilator)***Cordis***a Johnson and Johnson Company***ENDOVASCULAR*** |  |  | March 29, 2013 | 2013 | Cordis Corporation | Correction to provide clarification and additional labeling to minimize likelihood on implanting filter backwards. | Labeling design | Cordis Corporation sent an "URGENT Medical Device Correction" notification to their consignees in the US by letter dated April 3, 2013 and an e-mail was sent to the Canadian consignees. The letter identified the affected product, problem and actions to be taken. Consignees were instructed to read the description and recommendations sections carefully, then sign and return the enclosed Acknowledgement Form directly to Cordis to their fax number(908) 429-8287 or Scan and e-mail to MiamiQA@crdus.jnj.com. The letter states to pass on the notice to anyone in the consignees' facilities that need to be informed. Additionally, the notification request consignees to maintain awareness of this notification until the information has been incorporated in the Cordis OPTEASE Vena Cava Filter labeling. For questions related to the issue contact Cordis Endovascular Medical Information Line 877-338-4235. |  | 33000 | 33,000 units | Worldwide Distribution - USA Nationwide in the states of AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, Ny, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY, including Hawaii and Puerto Rico and the countries of Canada &amp; Panama | Terminated | October 27, 2014 | 577 | 33000 | Not_Computer
Z-1728-2013 | 65505 | No | GBA | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | 27023WU Balloon Catheter, 3 Fr., 2/pkg, sterile, contains latex. |  |  | May 31, 2013 | 2013 | Karl Storz Endoscopy America Inc | Karl Storz Endoscopy-America has initiated the recall of the 27023WU Balloon Catheter because the packaging may be compromised, and therefore the sterility of the product could be affected. | Packaging | The firm, Karl Storz, sent a "A SAFETY ALERT-ATTENTION NEEDED" letter dated May 31, 2013 to all customers who purchased the 27023WU Balloon Catheter. The letter described the product, problem and actions to be taken. The customers were instructed to check their inventory and to return the recalled product back to Karl Storz.. The customers were also instructed to complete and fax back the Recall Response Form attached with the safety alert letter at Fax #424-218-8559. If you have any questions, please feel free to call 424-218-8289. |  | 40 | 40 units | US distribution to: AZ, FL, IL, IN and TX. | Terminated | July 15, 2013 | 45 | 40 | Not_Computer
Z-1426-2013 | 65072 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Thyroidectomy Indication for the da Vinci Surgical Systems; |  |  | October 13, 2011 | 2011 | Intuitive Surgical, Inc. | Promotional literature for use of the da Vinci system for thyroidectomy, which had not been approved with a change to the 510(k). | No Marketing Application | Intuitive sent a notification letter on October 17, 2011 by email to all consigness whose email address were on file. Also a copy of the letter was sent by registered mail on February 8, 2012 to all surgeons whose addresses were on file. The letters informed consignees that Intuitive Surgical were in the process of reviewing labeling for the da Vinci System in regard to use for thyroidectomy procedures. During this period, Intuitive will not provide or facilitate any promotional, training, case observations, proctoring or in case procedural support for use of the da Vinci System in conjunction with thyroidectomy procedures. The letters recommend consulting the institution policy regarding use of approved or cleared medical devices for procedures that are not specified in the labeled indication for use. Questions may be directed to customer service at 1-877-408-3872. |  | 2423 | 2,423 consignee | USA Nationwide distribution | Terminated | May 30, 2013 | 595 | 2423 | Not_Computer
Z-0464-2014 | 66830 | [6 empty] | Centurion Sterile # 84 Rubber Bands Reorder EB84, Caution: This product contains natural rubber latex which may cause allergic reactions. |  |  | October 28, 2013 | 2013 | Centurion Medical Products Corporation | Package labeling indicates both "latex free" and "contains natural rubber latex" . The rubber bands do contain natural rubber latex. This could cause a significant risk to users with latex allergies. | Labeling False and Misleading | Centurion sent a Urgent Recall Notification letter via Certified Mail October 31, 2013, return receipt to all affected customers. The affected Centurion Medical Products Corporation sales representatives were notified via email on October 28, 2013. Customers were instructed to destroy all implicated product and complete the accountability record included with the notice and fax to 517-546-3356. Customers were asked to forward a copy of the notice if product was further distributed. Additional notices will be mailed to non-responsive customers via Certified Mail ReturnReceipt, and will be documented in the recall file. For further questions please call (517) 546-5400. |  | 500 | 500 lots | US Distribution including the states of GA, LA and NY. | Terminated | January 29, 2014 | 93 | 500 | Not_Computer
Z-0467-2014 | 66761 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | 1) Accolade II 127 Neck Angle Hip Stem, Size # 3 |  |  | October 25, 2013 | 2013 | Stryker Howmedica Osteonics Corp. | Stryker received a report from the field indicating that an Accolade II Hip Stem Size #3 was identified in packaging with an Accolade II Hip Stem Size #5. | Packaging process control | Stryker sent an e-mail to their domestic locations on 10/11/2013. Notification letters and product accountability forms were sent via Fed Ex on 10/25/2013 with return receipt to branches/agencies and hospital risk management, chief of orthopaedics and surgeons. |  | 20 | 20 units | Nationwide Distribution including states: AL, GA, IL, and WA. | Terminated | March 31, 2015 | 522 | 20 | Not_Computer
Z-1424-2013 | 65094 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | User Manual addendum, TORS for da Vinci Surgical System IS 1200, IS 2000 and IS 3000. |  |  | January 24, 2013 | 2013 | Intuitive Surgical, Inc. | Change to User Manual to remove one indication for use and amend warnings to emphasize restrictions for use. | Labeling False and Misleading | Intuitive sent an Important Medical Device Correction letters dated January 24, 2013 via Federal Express to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to locate and discard the existing TORS addendum, place the new TORS addendum included with this letter with your System User Manual, and complete and return the attached Acknowledgement Form to Intuitive Surgical. For questions call Intuitive Surgical Customer Service at 1-800-876-1310. |  | 1374 | 1,374 | US Distribution Only | Terminated | May 30, 2013 | 126 | 1374 | Not_Computer
Z-1660-2013 | 64078 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Custom Procedural Trays: Merit's Custom Angio Kits |  |  | November 21, 2012 | 2012 | Merit Medical Systems, Inc. | Custom Procedural trays contain Medrad Hand Controller Sheaths which were recalled because the seal between the white polyethylene and the clear plastic may be missing on the package resulting in a potential breach of the sterility of this accessory. | Packaging | Merit Medical sent an "URGENT PRODUCT RECALL NOTICE" letter dated November 21, 2012, to affected customer. The letter identified the product, the problem, and the action to be taken by the customer: Ensure all personnel within their institution to which the affected Merit trays were distributed are notified, quarantine all affected trays and return the affected product to Merit - Maastricht. Complete and sign the Customer Response Form attached to this letter and perform the following steps:(1) Immediately fax a copy of the completed Customer Response Form to: 1-804-416-1031. (2) Mail the completed original Customer Response Form from (1) above using the provided postage-paid, self-addressed envelope to Merit. Product Return Instructions: Please return the affected products by shipping them back to Merit via Federal Express 2nd Day Air (Acct #1129-5275-6) to: RGA Department, RGA #NC-54948 Merit Medical Maastricht, Amerikalaan 42, 6199 AE Maastricht - Airport, The Netherlands. NOTE: A copy of the Customer Response Form must accompany the shipment. If you have any questions, please call 1-804-416-1048. |  | 134 | 134 trays | U.S. Military Hospital on base in Landstuhl, Germany | Terminated | July 12, 2013 | 233 | 134 | Not_Computer
Z-1428-2013 | 65083 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Flush Port (used in almost all da Vinci instruments) |  |  | September 19, 2011 | 2011 | Intuitive Surgical, Inc. | Clarification of existing labeling information and operating procedures. | Labeling False and Misleading | Intuitive Surgical sent an "IMPORTANT PRODUCT NOTIFICATION" letter dated October 17, 2011 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact Customer Service at 1-800-876-1310 for questions regarding this notice. |  |  |  | Worldwide Distribution- USA Nationwide and countries of: Argentina, Australia,Canada, Chile, China, Czech Republic, Egypt, India, Israel, Italy,Korea-Republic of, Poland, Qatar, Romania, Saudi Arabia, Singapore, Spain, Thailand, Turkey, and Venezuela. | Terminated | May 30, 2013 | 619 |  | Not_Computer
Z-0473-2014 | 66582 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes TomoFix Plates |  |  | October 09, 2013 | 2013 | Synthes USA HQ, Inc. | A labeling correction was initiated related to the Surgical Technique Guide J5606-C "TomoFix Osteotomy System" due to the potential for a surgical delay if the surgeon inserts the superior screws at a lesser angle (4 or 6 degrees) and is not aware that those angles are not present. | Labeling False and Misleading | Synthes sent an Urgent Field Safety Notification/Medical Device Labeling Correction letter dated October 9, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to exchange the old surgical technique guide version J5606-C with the new version J5606-D provided with this notification. Review the revised surgical technique guide (Pages 25 &amp; 26). Forward this Field Safety Notification to anyone in their facility that need to be informed.If the surgical technique guide has been forwarded to another facility, contact that facility.Maintain awareness of this notice until all surgical technique guides have been exchanged.Maintain a copy of this notice. Customers with questions were instructed to contact their SynthesSales Consultant. For questions regarding this recall call 610-719-5000. |  | 2912 | 2912 | Worldwide Distribution - USA (nationwide) Canada and Mexico | Terminated | December 18, 2015 | 800 | 2912 | Not_Computer
Z-1654-2013 | 64888 | No | LRY | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | HealthLink***BIOPSY PUNCHES*** |  |  | March 18, 2013 | 2013 | Clorox Healthcare Holdings, LLC (dba HealthLink) | Due to lack of sterilization. | Packaging change control | On March 19, 2013, HealthLink sent their customers letters instructing them to do the following: 1) Screen their inventory and quarantine the product. 2) Notify all customers that received the affected products and ask them to return the lots to their firm. 3) Complete the attached form and return the lots to HealthLink Distribution center. For further assistance, customers were directed to call customer service at 1-800-638-2625. |  | 1044600 | 1,044,600 units | Worldwide Distribution-USA (nationwide) including Puerto Rico and the states of AZ, CA, CO, FL, GA, IA, IL, IN, MI, MO, MS, MT, NC, NJ, NY, OH, SC,TX ,VA ,WA, WI, and HI, and the countries of Canada, Mexico, and New Zealand. | Terminated | February 25, 2014 | 344 | 2849900 | Not_Computer
Z-0551-2014 | 66949 | No | JXG | Central nervous system fluid shunt and components. | Neurology | Neurology | 510(k) | Integra Cranial Access Kit, Model Number HITHSP10, Rx Only. |  |  | November 26, 2013 | 2013 | Integra LifeSciences Corp. d.b.a. Integra Pain Management | Integra LifeSciences is voluntarily recalling anyunexpired lots of HITHSP10 Cranial Access Kits due to identified labeling discrepancies. kit contains DuraPrep and Monocryl sutures whose labeling does not align with the intended use and labeling of the Cranial Access Kit. | Labeling mix-ups | Integra sent an Urgent Voluntary Medical Device Recall letter dated November 26, 2013, to all affected consignees. The letter identified the product, the problem, the action to be taken by the consignee and included a Recall Acknowledgement and Return form. Consignees were instructed to review their inventory and determine if they have any affected product, stop using the affected product immediately and remove from service. Consignees were instructed to complete the attached form even if they do not have the affected product in their inventory and return by email or fax as indicated on the form. When the form is received Customer Service will send an RMA number and directions to return the product. Customers with questions were instructed to contact Customer Service at 1-855-532-1723 or their Integra sales representative at 909-201-4099.For questions regarding this recall call 609-936-6822. |  | 731 | 731 | Nationwide Distribution | Terminated | April 09, 2014 | 134 | 731 | Not_Computer
Z-0483-2014 | 66908 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Stryker Orthopaedics Cancellous Bone Screw |  |  | October 01, 2013 | 2013 | Stryker Howmedica Osteonics Corp. | Stryker received a report from the field indicating that a 5.5 mm Cancellous Bone Screw 40 mm was identified in packaging associated with a 6.5 mm Cancellous Bone Screw 25 mm. | Labeling mix-ups | Stryker sent notification letters/accountability forms via Fed Ex on October 1, 2013, with return receipt. to branches/agencies and hospital risk management, chief of orthopedics and surgeons.Domestic branch locations were notified by e-mail on 9/16/2013.The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to fax the Product Recall Acknowledgement Form to 855-251-3635, return the affected product to the attention of Regulatory Compliance, Stryker Orthopaedics, 325 Corporate Drive, Mahwah, New Jersey, 07430. Customers were instructed to attach the fluorescent organge PRODUCT REMEDIATION sticker to their return, indicating the Produc Remediation # RA 2013-145. Mark the outer box with the words "Product Recall."Customers with questions were instructed to call 201-972-2100.For questions regarding this recall call 201-831-5000. |  | 2232619272 | Catalog Number 2030-6525-1- 4 units; Catalog No: 2029-5540-1-7 units | Nationwide Distribution | Terminated | July 15, 2015 | 652 | 2232619272 | Not_Computer
Z-1623-2013 | 65205 | No | EMF | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | -Instratek Carpal Tunnel Release Blades, (1) Triangle Blade, (1) Hook Blade. REF R2056 - LOT 122279 |  |  | May 03, 2013 | 2013 | Surgical Instrument Service And Savings, Inc. | The seal which maintains a sterile barrier for reprocessed medical devices became compromised to the point where product may fall out of the pouch. | Packaging process control | MEDISISS sent a Correction Immediate Action Request letter dated May 2, 2013, to all affected customes. The letter identified the product, the problem, and the action to be taken by the customer. The letter directed the customer to inspect the products prior to use for seal integrity as specified on the product labeling: "Sterile unless package has been damaged or opened. Do not use if package is damaged."This will be followed up by a visit from the customer's sales representatives who are tasked with inspecting all the lots that may be remaining in the customer's inventory. The sales representative is directed to complete a "Verification of Seal Inspection" form and return it to MEDISISS Customer Relations. Devices with unacceptable seals are to be returned to MEDISISS. The sales representative is responsible for obtaining a shipping label from customer service to enable the devices to be returned to the recalling firm via Fed Ex.Customers with questions were instructed to call 866-866-7477. |  | 56 | 56 units | Nationwide Distribution. | Terminated | July 11, 2013 | 69 | 4540 | Not_Computer
Z-0531-2014 | 66930 | No | LZP | Intraocular fluid. | Ophthalmic | Ophthalmic | PMA | Viscocel and Viscocel Plus. |  |  | November 25, 2013 | 2013 | C.L.R. Medicals International, Inc. | CLR Medicals International Inc., initiated a nationwide recall of Viscocel and Viscocel Plus, in response to FDA inspection findings of deficiencies in the Quality System Regulations for Medical Device Manufacturers, 21 Code of Federal Regulation (CFR) Part 820 and lack of Premarket Approval (PMA) for devices sold or distributed in domestic commerce. The Viscoelastic products, intended for human | No Marketing Application | CLR Medicals initiated a nationwide recall on November 25, 2013, of Viscocel and Viscocel Plus, by issuing a press release,in response to FDA inspectional findings of deficiencies in the Quality System regulations for Medical Device Manufacturers, 21 Code of Federal Regulation (CFR) Part 820 and lack of Premarket Approval (PMA) for devices sold or distributed in domestic commerce. The Viscoelastic products, intended for human and animal use, have been determined to be a public health concern due to distribution in the US without FDA approval, which evaluates the safety and efficacy of the product, and violations of the QSR regulations related to sterility, and biocompatibility that may result in the release of contaminated products. CLR Medicals International, Inc. sent a Medical Device Recall letter dated December 6, 2013, to all affected customers. The letter informed customers of the reason for recall, product description with manufacturing date range, possible health hazard, instructions on what to do with recalled product, and contact information for the firm and to report any adverse events to FDA MedWatch. The recall notification letter was accompanied by a response form. Customers with questions were instructed to call 909-398-1880.For questions regarding this recall call 909-398-1880. |  | 249592 | 249,592 units | Worldwide Distribution - USA (nationwide) and Internationally to Egypt, Mexico, and Canada. | Terminated | December 05, 2014 | 375 | 249592 | Not_Computer
Z-1162-2013 | 64508 | No | NAW |  |  | Radiology | PMA | Sirtex SIR-yttrium 90 microspheres (Radionuclide); SIR-Y001. |  |  | February 27, 2013 | 2013 | Sirtex Medical Limited | Incorrect expiration date | Labeling False and Misleading | Sirtex notified medical end users on 2/27/13 via email correspondence with a copy of the recall letter attached. This was done to ensure advice was received at the same time that the goods were arriving. Hard copies of the recall letter, which were stamped "copy", were also sent to the individual facilities via UPS, processed for delivery Thursday, 28 February, 2013. Medical end users were advised in the recall letter to contact the US Customer Service staff at (888) 474-7839 x 702 to arrange for the pick-up if they decide to return the unused product. |  | 21 | 21 doses | Nationwide Distribution-USA (nationwide) including DC and the states of AZ, CA, CT, GA, ID, IL,KY, LA, MD, ME, MO, MT, NJ, NV, WA, and WV. | Terminated | March 30, 2015 | 761 | 21 | Not_Computer
Z-1618-2013 | 65202 | No | KMK | Intravascular catheter securement device. | General Hospital | General Hospital | 510(K) Exempt | StatLock CV Plus w/Pigtail, Catheter Stabilization Device Sterile EO Single Use Only, C.R. Bard, Inc. The StatLockÂż device provides stabilization for compatible medical tubes and catheters. |  |  | June 10, 2013 | 2013 | Bard Access Systems | BARD Access Systems is recalling lot JUWJF283 of StatLock CV Plus Stabilization Butterfly Devices due to incorrect pad assembly used during manufacture. The absence of the hydrocolloid strip on the bottom may pose an incremental risk including localized skin irritation. Missing directional arrow may lead to mis-orientation of the StatLock stabilization which may impact the securement of the cathet | Packaging | On 6/18/13, an "Urgent Kit Recall" notification letter was mailed to all customers informing them of the recall and instructing the consignee to remove and return any affected product lot numbers to Bard Access Systems through return goods authorization. Customers may contact 800-290-1689 if questions. |  | 3875 | 3,875 units. | National and Global distribution (Europe, European Economic Area, Argentina, China, and Brazil). | Terminated | May 21, 2014 | 345 | 3875 | Not_Computer
Z-1111-2013 | 64782 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Edwards Lifesciences Duraflo coated Femoral Cannula. |  |  | March 14, 2013 | 2013 | Edwards Lifesciences, LLC | Two unsealed pouches were found at an Edwards' distribution location in Japan. Incorrectly sealed pouches may result in a sterility barrier breach. | Packaging process control | Edwards Lifesciences sent an Urgent Field Safety Notice letter dated March 28, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to discontinue use of any affected product remaining in inventory and to return any unused devices to Edwards. A completed Acknowledgement Form should be faxed to Edwards at 800-422-9329 within three days of receiving the notice. Questions regarding the recall should be directed to Edwards Customer Services at 800-424-3278, 6:00AM - 4:30PM PST. |  | 379 | 379 units | USA Nationwide Distribution including the states of: FL, MA, MI, NY, PA, TX. | Terminated | March 08, 2016 | 1090 | 379 | Not_Computer
Z-0465-2014 | 66823 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Axillent Debakey Grasping Forceps, Part No. 33821DYD |  |  | November 06, 2013 | 2013 | Karl Storz Endoscopy America Inc | KARL STORZ is recalling the Axillent Debakey Grasping Forceps because the 33821WTD Wavy grasper was mislabeled as 33821DYD Axillent Debakey Grasping Forceps. | Labeling mix-ups | Karl Storz Endosocopy sent a Safety Alert letter dated November 6, 2013, to all customers who received the 33821DYD Axillent Debakey Grasping Forceps. The letter informed the customers of the problems identified and the actions to be taken. Customers were asked to set aside the mislabeled product and return the mislabeled instruments with the RMA (Return Material Authorization) number issued to them. Customers were instructed to fax or email completed recall response form to (424) 218-8519. Customers with any questions were instructed to call (800) 421-0837 ext 8289.For questions regarding this recall call 746-708-8440. |  | 7 | 7 units | Nationwide Distribution including states of: TX, PA, and NY. | Terminated | January 13, 2014 | 68 | 7 | Not_Computer
Z-1667-2013 | 65449 | Yes | BTO | Tracheostomy tube and tube cuff. | Anesthesiology | Anesthesiology | 510(k) | Tracoe Mini 3.0mm Tube, for pediatric use. |  |  | April 17, 2013 | 2013 | Bryan Medical Inc | The pediatric tracheostomy tubes were mislabeled as the smaller infant tracheostomy tubes. | Labeling Change Control | The consignees were contacted by telephone and asked to return the recalled product. |  | 5 | 5 units | Distributed in KY and AL. | Terminated | October 29, 2013 | 195 | 5 | Not_Computer
Z-0461-2014 | 66809 | No | HWT | Template for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | NexGen Complete Knee Solution |  |  | November 19, 2013 | 2013 | Zimmer, Inc. | This is an expansion of the June 2013 recall of NexGen Micro components. Five (5) additional instruments have been identified that are associated only with the NexGen Micro components. There is the potential for these devices to be inadvertently utilized during surgery, resulting in a delay. The associated implants would not be available due to the previous recall. | Labeling design | Zimmer sent an Urgent Medical Device Recall Notice dated November 19, 2013, to all affected customers. The Notice outlined the products recalled; the reason for the recall, possible adverse events related to the recall and directed consignees to locate, quarantine and count all affected product recording the information on the Inventory Return Form acknowledging that recall directions were followed. Further questions should be directed to 1-877-946-2761. Affected product will be removed by a Zimmer Sales representative, Question should be directed to 1-877-946-2761 8a.m - 5 p.m. EST.For questions regarding this recall call 800-613-6131. |  | 638 | 638 units | Worldwide Distribution - USA (nationwide) and Internationally to Australia, Belgium, Canada, Denmark, Ecuador, Germany, India, Italy, Japan, Korea, Malaysia, Mexico, Netherlands, Peru, Singapore, South Africa, Spain, Switzerland, UAE, and UK. | Terminated | May 06, 2015 | 533 | 3027 | Not_Computer
Z-0717-2013 | 63873 | No | HIS | Condom. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Lifestyles Pleasure Collection |  |  | November 05, 2012 | 2012 | Ansell Healthcare Products LLC | One lot of Lifestyles Pleasure Collection Lubricated Condoms is misbranded. Ultra Thin was included rather than Skyn. | Packaging | Ansell Healthcare Products LLC sent an "URGENT: VOLUNTARY CONDOM RECALL" letter/return response form dated November 5, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm's Customer Service Department at 1-800-305-1054 for questions or if additional information is needed. |  | 445 | 445 cases (10680 consumer units) | Nationwide Distribution. | Terminated | July 31, 2013 | 268 | 445 | Not_Computer
Z-0579-2014 | 67003 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | OrthoPediatrics ACL Reconstruction System is a pediatric and small stature adult ACL reconstruction system containing Sleeve Installation Tool Set , Part 01-1008-0040/0047, 6mm; 6.5mm, 7mm, 7.5mm, 8mm, 8.5mm, 9mm, and 10mm. |  |  | November 07, 2013 | 2013 | OrthoPediatrics Corp | Part # 01-1008-0040/0047, ACL Sleeve Installation/ Insertion Tool, shipped November 4, 2013 - November 6, 2013 may be etched with the wrong size. | Labeling design | November 7, 2013 all consignees were notified of the Urgent Device recall via telephone calls and E-mail messages describing the product and reason for the recall,. The notice included directions to identify and quarantine the recalled devices returning product to OrthoPediatrics with completed return form to OrthoPediatrics via mail, email or courier service, or byphotographing with a smart phone and emailing to logistics@orthopediatrics.com or by answering therecall specific questionsin an email to mfox@orthopediatrics.com |  | 40 | 40 | Nationwide Distribution including the states of CA, KY, OH, RI, and TN. | Terminated | May 09, 2014 | 183 | 40 | Not_Computer
Z-1405-2013 | 65119 | No | DQR | Catheter cannula. | Cardiovascular | Cardiovascular | 510(k) | AngioDynamics Micro Introducer Kits under the following labeling: |  |  | April 30, 2013 | 2013 | Angiodynamics, Inc. | AngioDynamics Inc., is recalling Micro Introducer Kits and VenaCure EVLT NeverTouch Procedure Kits, because the potential exists that the micro Introducer (sheath/dilator) packaged within the kit could be the incorrect French size. | Labeling mix-ups | AngioDynamics sent an Urgent Medical Device Recall letter dated April 29, 2013, to all customers by Fed Ex. The letter identified the affected product, problem and actions to be taken. Customers were instructed to lmmediately remove the affected product from inventory and call AngioDynamics, Inc. Customer Service at 1-800-772-6446 to obtain a replacement or credit for your returned product and to obtain UPS Account Number to return the package via second day delivery. Customers are instructed to segregate and return all affected devices to: AngioDynamics 603 Queensbury Avenue Queensbury, NY 12804 Attn: Micro Introducer and VenaCure EVLT NeverTouch Procedure Kits Recall Coordinator. Write the RMA number on the shipping box (Provied on the Recall Verification Tracking Form). Promptly complete, sign and return the enclosed Reply Verification Tracking Form (even if you do not have any product to return); following the directions on the notification form and the Reply Verification Form. Send the Reply Verification form to AngioDynamics either via Fax or Emai. Fax Reply Verification Tracking Form to: Attn: Micro Introducer and VenaCure EVLT NeverTouch Procedure Kits Recall Coordinatory Fax number 1-518-798-1360 or Email Reply Verification Tracking Form to: rdenino@angiodynamics.com. |  | 806 | Domestic(1,760 units): 1)260 units/26 boxes, 2)310 units/31 boxes,3)260 units/26 boxes,4)880 units/88 boxes,5)60 units/6 boxes,Foreign:10 units/1 box | USA Nationwide Distribution including the states of: IL, GA, CA, NY OH, AZ, TX, NE, FL, PA, NC, WI, AL, IN, CO, MO, MT, NE, MA, VA, AK, IN, MI, NJ, MD, TN, CT, WA, MN | Completed |  |  | 1557 | Not_Computer
Z-1401-2013 | 65048 | No | JHI | Human chorionic gonadotropin (HCG) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Alere" hCG Combo Cassette (20/10 mIU/mL) Kit. |  |  | April 16, 2013 | 2013 | Alere San Diego, Inc. | The firm initiated recall of Alere" hCG Combo Cassette (20/10 mIU/mL) kit PN 92215, lots hCG1110133 and hCG1110135, because an investigation has indicated a portion of these Alere" hCG Combo Cassette (20/10 mIU/mL) PN 92215 kit lots were incorrectly packaged with the Alere" hCG Cassette (20 mIU/mL) PN 92210 Product Insert. The Alere" hCG Cassette (20 mIU/mL) Product Insert contains information for | Packaging | Alere initiated this recall by sending recall notification letters, titled "URGENT MEDICAL DEVICE RECALL", dated April 16, 2013 via via fax, email or direct mail. The letter informed customers of the product recalled, reason for recall, customer required action, and contact information, "Alere Technical Services at 866-216-0094 or by e-mail at lateral.flow.support@alere.com.". A verification form was attached to the notification letter. |  | 3975 | 3975 kits | United States Nationwide Distribution | Terminated | October 18, 2013 | 185 | 3975 | Not_Computer
Z-0448-2014 | 66333 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | GENESIS(R) II, PS HIGH FLEXION ARTICULAR INSERT, SIZE 5-6, 11 MM QTY: (1), REF 71421516, STERILE EO, Smith &amp; Nephew, Inc. Memphis, TN 38116 USA |  |  | August 27, 2013 | 2013 | Smith &amp; Nephew Inc | The Tyvek(TM) lid on one, or both, of the two trays that form adouble barrier around the insert may not have been adequately sealed. | Package design/selection | Smith &amp; Nephew sent an Urgent Recall Notification letter dated August 27, 2013, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the customer. They were instructed to cease use and distribution of the product and quarantine for return. Customers were instructed to contact Smith &amp; Nephew via email at FieldActions@smith-nephew.com or fax 901-566-7975 to obtain a return authorization (RA) number. Customers with questions were also instructed use this email address. Customers with affected product were instructed to return it along with the completed inventory Return Certification Form to the address listed on the form. Non-responding consignees were notified again on September 27, 2013, by email.For questions regarding this recall call 901-396-2121. |  | 46 | 46 pieces | Worldwide Distribution - USA including Nevada and Internationally to United Arab Emirates, Italy, Spain, Belgium, Bulgaria, Poland, France, Japan, Netherlands, Australia, Turkey, Russian Federation, Germany and United Kingdom. | Open, Classified |  |  | 227 | Not_Computer
Z-1518-2013 | 65108 | No | LFL |  |  | General &amp; Plastic Surgery | 510(k) | TX1 Tissue Removal System disposable handpieces. |  |  | April 30, 2013 | 2013 | American Optisurgical Inc | American Optisurgical, Inc. is recalling certain lots of TX1 Tissue Removal System disposable handpiece due to inventory being released into distribution without a cleared FDA 510(k) to cover recent product changes. Use of this product may lead to an underperformance of the device. | No Marketing Application | American Optisurgical Inc. (AOI) sent an Urgent Medical Device Recall letter along with a response sheet via USPS registered mail to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immediately examine your inventory and quarantine product subject to recall. Once the product has been identified, return all products by FedEx to AOI using the AOI FedEx account 30655659-2, and referencing the Returned Goods Authorization (RGA) number:RGA# (to be assigned). For questions contact American Optisurgical Inc. Customer Service 949-580-1266. |  | 1436 | 1,436 units | Worldwide Disitribution - USA Nationwide in the states of Alabama, Arkansas, Arizona, California, Colorado, Delaware, Florida, Georgia, Iowa, Illinois, Indiana, Kansas, Kentucky, Massachusetts, Maryland, Michigan, Minnesota, Missouri, North Carolina, Nebraska, New Hampshire, New Jersey, New Mexico, Nevada, New York, Ohio, Oklahoma, Oregon, Pennsylvania, South Carolina, Tennessee, Texas, Utah, Virginia, Vermont, Washington, Wisconsin, West Virginia, Wyoming and the countries of Quezon City, Philippines, Singapore, and Panama City, Panama. | Terminated | March 24, 2014 | 328 | 1436 | Not_Computer
Z-1084-2013 | 64749 | No | DZD | Orthodontic appliance and accessories. | Dental | Dental | 510(K) Exempt | Bracket Buccal Tube U7L DB/UL22 Roth MT HK 0/-14/0 |  |  | March 08, 2013 | 2013 | Ormco/Sybronendo | Ormco Corporation is voluntarily recalling one (1) lot of Bracket Buccal Tubes as a result of a packaging error. More specifically, the BracketBuccal Tubes product package correctly identifies the product as gingival (0Âż) offset brackets, butactually contains distal (10Âż) offset brackets. Use of the affected product may cause unintended rotationof the molars. Although reversible, a delay in | Packaging | Ormco sent an Urgent Medical Device Recall letter dated March 8, 2013 to all affected customers, via USPS 1st class mail. The letter identified the affected product, problem and actions to be taken. The consignees were instructed to complete the Return Form and return any affected product identified in stock. For questions contact Ormco Customer Care. |  | 8210 | 8210 | Worldwide Distribution - USA (nationwide) including the country of EU. | Terminated | September 04, 2013 | 180 | 8210 | Not_Computer
Z-1675-2013 | 65439 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | AngioDynamics RITA Main Cable, 9 Pin to 14 Pin, REF 700-101892. RITA Main Cables are indicated for use with the RITA RF Generators and RITA electrosurgical devices. |  |  | June 11, 2013 | 2013 | Angiodynamics, INC | The RITA Main Cable product contains the incorrect Instructions for Use. | Packaging process control | Consignees were notified via certified mail on 06/11/2013 with an "Urgent Medical Device Field Safety Correction" notification. This included the correct Instructions for Use and the recommended actions. |  | 13 | 13 units | Worldwide Distribution, including the US, Austria, China, Italy, Russia and the UK. | Terminated | June 12, 2015 | 731 | 13 | Not_Computer
Z-0559-2014 | 66535 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | B. Braun Introcan Safety IV Catheter PUR 24G, 0.7X19mm-US, catalog number 4251601-02. |  |  | September 27, 2013 | 2013 | B. Braun Medical, Inc. | One lot of Introcan Safety IV Catheters may not contain the anti-needlestick safety clip which is used for the prevention of needlestick injuries. | Labeling False and Misleading | B. Braun sent an Urgent Medical Device Recall Notification letter dated September 27, 2013 to all affected customers. The letter identifed the affected product, problem and actions to be taken. Customers were instructed to utilize the attached "Product Removal Acknowledgement Form" to record the number of individual units in their inventory and fax the completed form to 610-848-1197. To ensure proper credit for affected product, customers were instructed to return the product to B. Braun Medical Inc. For questions contact Customer Support Department at 800-227-2862. |  | 55200 | 55,200 units | USA Nationwide Distribution including Puerto Rico. | Terminated | October 08, 2014 | 376 | 55200 | Not_Computer
Z-1676-2013 | 65290 | No | HSD | Shoulder joint humeral (hemi-shoulder) metallic uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Bigliani/FlatowÂż The Complete Shoulder Solution, Humeral Stem Provisional 12mm x 170mm |  |  | June 11, 2013 | 2013 | Zimmer, Inc. | The firm is initiating a removal of one lot of the Bigliani/Flatow Humeral Provisional Stem (00-4301-012-17, lot 62283991) as the stems manufactured under the lot are 14mm x 170mm devices incorrectly etched and packaged as 12mm x 170mm devices. | Packaging process control | Zimmer sent an URGENT MEDICAL DEVICE RECALL notification dated June 11, 2013, to all consignees via email and courier describing the affected device and problem and the risks associated with using the affected device. The notifications provided instructions for returning the affected device to Zimmer Product Service Department, 1777 West Center Street, Warsaw, IN 46580. Any questions or concerns concerning the recall, please call the customer call center at 1-877-946-2761. Customers may report adverse events or quality problems experienced with the use of these products to Zimmer by email at zimmer.per@zimmer.com.For questions regarding this recall call 800-613-6131. |  | 1 | 1 | Nationwide Distribution including IL and MO. | Terminated | January 30, 2014 | 233 | 1 | Not_Computer
Z-1683-2013 | 65183 | No | LJH | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Two Lead Arthroscopic Irrigation Set, for use with ARTHROMATIC Plastic Containers, Product Code 2C4030 - Approximate Length 98" (2.5. m) |  |  | May 15, 2013 | 2013 | Baxter Healthcare Corp. | A firm's internal investigation led to the discovery of pinholes in the packaging of the products. The pinholes are not likely to be detected by the end user prior to use and could compromise the sterile barrier properties of the packaging. | Packaging | Baxter sent an Urgent Product Recall letter dated May 15, 2013 to all affected custorners. The letter identified the affected product, problem and actions to be taken. Customers were requested to remove all affected products and return for credit by contacting Baxter Healthcare Center for Service at 888-229-0001. Customers were instructed to complete the attached Customer Reply Form and return it to Baxter by either faxing to 224-270-5457 or scanning and e-mailing the reply form to fca@baxter.com |  | 530488 | 530,488 units | US Nationwide Distribution | Terminated | June 16, 2015 | 762 | 1136792 | Not_Computer
Z-1671-2013 | 65347 | No | KNQ | Esophageal dilator. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | RigiflexTM 11 Single-Use Achalasia Balloon Dilator, Material Number (UPN) M00554510 |  |  | May 28, 2013 | 2013 | Boston Scientific Corporation | Product labeled with incorrect expiration date. | Incorrect or no expiration date | Boston Scientific sent an Urgent: Medical Device Recall letter to all affected customers via Federal Express. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to immediately discontinue use, check their inventory, remove any affected product from their inventory, and return it to Boston Scientific. The letter also requests the completion of a Reply Verification Tracking Form as a method of documenting the presence or absence of affected product from their inventory. Follow-up communications will be made to the accounts that do not respond. For questions contact your local Sales Representative. |  | 91 | 91 | Worldwide Distribution - USA Nationwide including the state of NY, NE, MD, OH, TX, PA, VA, PR, CA, NJ, NY, CA and the countries of: AT, FR, DE, GB, IL, IT, LB, NL, ES, SE, TN, Argentina, and Japan. | Terminated | February 10, 2014 | 258 | 91 | Not_Computer
Z-1409-2013 | 64586 | No | MNI | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | 45 mm Screw, Item #BG7045 |  |  | February 20, 2013 | 2013 | Spinal Solutions, LLC | The recall of the Loaner Bin and Sterilization Trays for implants and instruments was initiated by Spinal Solutions because the directions for use and surgical techniques were not being provided to every physician and for every surgery. | Labeling design | Spinal Solutions sent a recall letter dated March 8, 2013 to all customers who purchased loaner bins and sterilization trays for implants and instruments from Spinal Solutions. The letter identified the affected product, problem identified and actions to be taken. Customers are instructed to call the office at the number (951) 304-9001 if they are requesting paper copies of any of the literature. |  | 24 | 24 units | USA Nationwide Distribution in the states of MS: TX, NV, CA, MD, and WI. | Terminated | March 26, 2014 | 399 | 503 | Not_Computer
Z-0560-2014 | 66934 | No | HTD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Edwards Lifesciences Ergonic Double Acting DeBakey Forceps with Narrow Straight Tip. Model Number ERGF01330. |  |  | November 20, 2013 | 2013 | Edwards Lifesciences, LLC | Edwards Lifesciences is recalling one lot of Ergonic instrument due to mislabeling. Ergonic scissors standard curved labeled as a Double Acting DeBakey forcep. Edwards control confirmed that the laser etching on the device was incorrect. | Labeling mix-ups | Edwards sent an " Urgent Field Safety Notice" to all affected customers dated November 20, 2013. The letter identified the product the problem and the action needed to be taken by the customers.Action to be taken by user:Our records show that you received one of these instruments. Our sales representative will be on-site to remove this instrument and provide a replacement. Please assist our sales representative in reviewing your entire inventory for the instrument listed in the recall letter. The form has been pre-populated with the inventory from our records. If you have additional inventory of this affected lot within your control, please add the quantity to the section on the form provided. Please remove the instrument from use, quarantine the device at your site and support the Edwards Sales Representative to return this product to Edwards. An acknowledgment form is included to assist you in the assessment of your inventory.While our sales representative is on-site, please complete the attached acknowledgment form and fax it back to Edwards Customer Service at 800.422.9329. The return of this form allows us to confirm that you have reviewed this notice and have taken appropriate action. Please contact Customer Service at 800.424.3278 if you have questions about the return and crediting process.The Customer Service organization can answer questions about tracking of the replacement instrument provided by the Edwards Sales Representative. If you have questions that have not been answered by this letter, please call Edwards Customer Service at 800.424.3278 from the hours of 6:00AM - 4:30PM PST . |  | 3 | 3 | Distributed to 3 customers in FL only. | Terminated | February 10, 2014 | 82 | 3 | Not_Computer
Z-1176-2013 | 64706 | No | NEW | Absorbable polydioxanone surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ORTHOCORD One Violet and One Blue Braided Composite Suture with OS-6 Reverse Cutting Needles, Product Code: 223116. |  |  | March 22, 2013 | 2013 | DePuy Mitek, Inc., a Johnson &amp; Johnson Co. | The IFU provided with the free strand ORTHOCORD sutures currently indicates that approximately 100% suture strength remains at six weeks after implantation. While this statement is true for Violet ORTHOCORD sutures, it is an inaccurate statement for Blue ORTHOCORD free strand sutures. Based on in-vivo strength retention study, a Blue ORTHOCORD sutures maintain approximately 80% of strength at six | Labeling Change Control | On March 22, 2013, firm notified consignees of recall via letter. Consignees were informed of the following: The IFU provided with the free strand ORTHOCORD sutures currently indicates that approximately 100% suture strength remains at six weeks after implantation. While this statement is true for Violet ORTHOCORD sutures, it is an inaccurate statement for Blue ORTHOCORD free strand sutures. Based on in-vivo strength retention study, a Blue ORTHOCORD sutures maintain approximately 80% of strength at six weeks after implantation. As Blue and Violet ORTHOCORD free strand sutures share the same IFU, IFU-108147 will be revised. Consignees were requested to confirm the receipt of the letter by March 29, 2013. |  | 46501 | 46501 | Nationwide Distribution. | Terminated | February 24, 2014 | 339 | 449009 | Not_Computer
Z-1115-2013 | 64810 | No | GFA | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Product is distributed by Zimmer (Warsaw, Indiana)/Synvasive Technology, Inc. (El Dorado Hills, CA) &amp; manufactured under Synvasive Technology, Inc. El Dorado Hills, CA |  |  | April 02, 2013 | 2013 | Synvasive Technology Inc | Various reciprocating and oscillating saw blades were incorrectly labeled. The labeling errors include: (1) The pouch says oscillating saw blade when it actually contains a reciprocating saw blade; (2) The box bar code date does not match the human readable expiration date; and (3) The label incorrectly states use with Zimmer/Synthes handset and not the Zimmer Universal Power System. | Packaging process control | Zimmer sent a Urgent Device Removal letter dated April 2, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers that received affected product from both Synvasive and from Zimmer Surgical will receive two customer letters, as each letter sets out specific item and lot numbers dependent upon distribution entity. Zimmer Distributors will also receive an e-mail notification. All notifications include a response form for both Synvasive and Zimmer distributors as well as customer consignees to account for product disposition (disposal in the field per facilities' sharps protocol).(1) Notifications will be mailed via Fed Ex to Risk Managers and Zimmer Distributors with inventory.(2) Notifications will be emailed to international geographies.(3) Notifications will be emailed all Zimmer Distributors.(4) A second Fed Ex mailing will be sent to domestic non-responding consignees.(5) Effectiveness phone calls will be made to any remaining domestic non-responding consignees. (6) If a customer reports they have further distributed the product, the secondary consignee identified will receive a notification from Zimmer.The scope of the recall is limited to certain confirmed lots which were shipped to the affected consignees..E-mails, certified mail receipts, and telephone records will be the firm's course of action in monitoring effectiveness checks. All returns will be stored in a separate area away from regular inventory. Final disposition of returned product is yet to be determined.The firm's recall strategy appears to be adequate.For further questions please call (916) 939-3913. |  | 3442 | 3442 units per the part numbers listed | Worldwide Distribution - USA (nationwide) and countries of: Austria, Australia, Belgium, Czech Republic, Denmark, Finland, France, Germany, Italy, Japan, Netherlands, Norway, Panama, Spain, Sweden, Switzerland, Tunisia, United Kingdom and Yemen. | Terminated | July 15, 2014 | 469 | 3442 | Not_Computer
Z-0692-2013 | 64039 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Eterna Giovinezza System, Powered Laser Surgical Instrument. |  |  | November 07, 2012 | 2012 | Quanta System, S.p.A. | Quanta System S.p.A, Italy, is notifying U.S. consignees that their website contained promotional materials for products distributed in the U.S. which did not specify that some of the claims made therein were not meant to apply to the U.S. distributed devices. | Labeling False and Misleading | Quanta System sent a letter dated November 7, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to perform the following: quarantine and destroy all promotional literature printed from the manufacturer, website or obtained from the manufacturer. Provide any promotional literature independently produced by the initial importer or distributor containing information about Quanta System products. Provide URL of initial importers or distributors website if information about Quanta System products is contained therein. Provide a written certification that all actions have been taken. Quanta System has or will: review all promotional materials and websiteinformation provided by initial importers and distributors; advise initial importers and distributors of any actions that need to be taken to change promotional materials and/or website information.For questions regarding this recall call 303-223-4336. |  | 12 | 12 units | Class II Recall - Worldwide Distribution - USA including TX, AZ, OR, and CO.Only US product is affected by this recall. | Terminated | January 23, 2013 | 77 | 109 | Not_Computer
Z-1611-2013 | 65397 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | PinnacleÂż DestinationÂż Peripheral Guiding Sheath, 8 French, 90cm, Straight Tip, Cross Cut Valve. |  |  | May 30, 2013 | 2013 | Terumo Medical Corporation | Medical devices were incorrectly labeled with extended expiration dates. | Error in labeling | The recalling firm notified consignees by letter on 5/30/2013. The letter was flagged as an urgent product bulletin addressed to risk managers. The notification advised customers to immediately discontinue use and quarantine any available stock for subsequent instructions to be provided for pick-up and replacement. |  | 20 | 20 units | Nationwide Distribution. | Terminated | October 30, 2013 | 153 | 570 | Not_Computer
Z-1543-2013 | 65374 | No | DRS | Extravascular blood pressure transducer. | Cardiovascular | Cardiovascular | 510(k) | Deltran Disposable Pressure Transducer Kit (DPT Kit) |  |  | June 04, 2013 | 2013 | Utah Medical Products, Inc | Utah Medical Products, Inc. is recalling certain Deltran Disposable Pressure Transducer Kits due to complaints related to packaging integrity of the sterile packaged devices. | Packaging | Utah Medical Products (UTMD) sent a Recall Notice letter dated June 3, 2013 via fax and email to all affected customers and will be completing follow up telephone calls by June 7, 2013. The letter identified the affected product, problem, health risk, and actions to be taken. UTMD instructed customers to locate and return affected products for exchange. For questions call Customer Service at 800-533-4984. |  | 1800 | 1,800 kits | Worldwide Distribution - United States Nationwide in the states of (MO, NE, KY, FL, OH, NY), and countries of Columbia, Brazil, Venezuela. West Indies, Latin America, Ecuador are distributed through a US direct consignee exporter. | Terminated | September 19, 2013 | 107 | 1800 | Not_Computer
Z-1486-2013 | 65314 | No | GAM | Absorbable poly(glycolide/l-lactide) surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Digitex Absorbable Polydiaxanone Suture, Size 2-0. |  |  | May 13, 2013 | 2013 | Coloplast Manufacturing US, LLC | Coloplast is conducting a recall on three lots of Digitex Absorbable Polydiaxanone Suture, Size 2-0 because it is mislabeled with a three year expiration date. | Incorrect or no expiration date | The firm, Coloplast, sent a letter dated May 8, 2013 to its customers. The letter described the problem, the product and actions to be taken. The customers will be contacted by Coloplast Customer Service with instructions for the handling and return of any unused product.If you have any additional questions, the dedicated Customer Service number regarding this effort is 800-258-3476. |  | 772 | 772 units | US Distribution to: KY, MI, OH, PR, TX, and VA. | Terminated | April 10, 2014 | 332 | 772 | Not_Computer
Z-0665-2013 | 63941 | No | DKB | Clinical toxicology calibrator. | Toxicology | Toxicology | 510(K) Exempt | Siemens Emit(R) tox(TM) Serum Calibrators, catalog number 7B019UL |  |  | December 06, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Incorrect expiration date was assigned to the D1 lot of Emit(R) tox(TM) Serum Calibrators. The correct expiration date is 2013-12-31. The incorrect assigned expiration date is 2014-12-31. | Error in labeling | Siemens sent a Customer Notification letter dated December 2012 to all affected customers of Emit(R) tox(TM) Serum Calibrators lot D1, informing them that Siemens Healthcare Diagnostics, Inc. has confirmed that the expiration of 2014-12-31 is incorrect. Siemens requested customers to discontinue use and discard any remaining inventory of the D1 lots. Customers were instructed to complete the Effectiveness Check form included with the letter even if no replacement reagents were needed and fax to 302-631-8467. Siemens would replace any unused kits of the affected lots at no charge.Customers with questions were instructed to contact their Siemens Technical Solutions Center or their local Siemens technical support representative.For questions regarding this recall call 302-631-6311. |  | 1998 | 1,998 cartons | Nationwide Distribution to the following states: AL, AR, CA, CO, FL, GA, IL, IN, KY, MA, MD, MI, MO, NC, NY, OH, PA, TN, TX and WI. | Terminated | December 24, 2013 | 383 | 1998 | Not_Computer
Z-0530-2014 | 65477 | No | JXB | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | remel RapID Yeast Plus System, 20 tests, Ref: R8311007, IVD, containing a 15-ml bottle of RapID Yeast Plus Reagent A and a 10-ml bottle of RapID Yeast Plus Reagent B. |  |  | April 04, 2013 | 2013 | Remel Inc | The two reagents supplied in the RapID Yeast Plus System kit (Reagent A and Reagent B) are mislabeled. The labels were reversed. | Labeling mix-ups | Thermo Fisher Scientific sent a Medical Device Recall letter dated April 3, 2013 to affected customers. The letter identified the affected product, problem and actions to be taken. The recalling firm instructed customer to review their inventory for affected lot and discard the product. A Medical Device Recall Return Response was enclosed for the customer to indicate they understood the instructions and to report the amount of product they discarded. For questions call Technical Services Department at 800-255-6730 (US) or 913-888-0939 (International). |  | 150 | 150 kits | Worldwide Distribution - USA Nationwide in the states of AL, AZ, CA, CT, FL, ID, IN, KY, MA, MI, MS, NJ, OH, PA, TX, WI and the countries of Canada, UK, and Japan. | Terminated | March 12, 2014 | 342 | 150 | Not_Computer
Z-1549-2013 | 33728 | microscope, phase contrast | IBM | Microscopes and accessories. | Pathology | Pathology | 510(K) Exempt | Aperio ScanScope System. Aperio ScanScope Models, XT, AT, AT Turbo, CS and CSO. |  |  | February 22, 2013 | 2013 | Aperio Technologies Inc | Aperio initiated this recall because the User Guides incorrectly state the devices had been cleared for specific applications. The User Guides should state the devices are for Research Use Only and are not for use in diagnostic procedures. | No Marketing Application | Aperio sent an Urgent Medical Device Recall letter dated April 22, 2013 to all affected consignees via certified USPS mail. The letter included the reason for the recall, recall product information, instructions, acknowledgement form, contact information ( Aperio Technical Support at 1-866-478-3999), and product labeling. The consignees were requested to acknowledge receipt of the Medical Device Correction letter. In addition, an Acknowledgement Form is required to be signed and e-mailed back to the Manufacturer asevidence of letter receipt.For questions regarding this recall call 760-304-6211, ext 6211. |  | 364 | 364 | Nationwide Distribution - including all states, none internationally. | Terminated | September 25, 2013 | 215 | 364 | Not_Computer
Z-0523-2014 | 66964 | No | HTD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | outer box: Ref 906768, CurvTek Eye Needle 22 MM X-large Stainless Steel, Single Use Gamma Sterilized, Rx |  |  | November 26, 2013 | 2013 | Biomet, Inc. | Inner label of PN: 906768 Lot: 055450 CurvTek Eye Needle 22mm X-Large states incorrectly PN: 906760 7mm Medium. | Labeling mix-ups | Biomet sent an Urgent Medical Device Recall Notice dated November 26, 2013, to all affected customers. The notice identifies the product, problem, and actions to be taken by the customers. The letter included a Customer Response Form for customers to complete and return to the firm by FAX 574-372-1683. Customers with questions were instructed to call 574-372-1570.For questions regarding this recall call 574-372-1570. |  | 21 | 21 | Nationwide Distribution including MI, IA, IL, IN, and TN. | Terminated | July 07, 2015 | 588 | 21 | Not_Computer
Z-1544-2013 | 64216 | No | EKR | Dental hand instrument. | Dental | Dental | 510(K) Exempt | SybronEndo Elements Obturation Unit; REF 973-0322; SN 91-1234; 2005-10; SybronEndo, 1332 S. Lone Hill Ave. Glendora, CA 91740; EU Rep: Kerr Italia, S.r.l, Via Passanti, 332, 1-84018 Scafati (SA), Italy |  |  | January 17, 2013 | 2013 | Ormco/Sybronendo | A recall has been issued by SybronEndo because a Canadian Standards Association (CSA)-directed cautionary statement label was inadvertently omitted from the Elements Obturation Unit and its Replacement A/C Power Cord. | Labeling mix-ups | The firm SybronEndo sent a "FIELD CORRECTION ACTION" letter dated January 30, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were requested to review and determine if they have any of the affected equipment in their inventory; Note: the equipment itself has not been modified in any way, and therefore there is no change in the performance; customers should continue to use the equipment in the same manner they have in the past:; Note: SynbronEndo has enclosed the cautionary label along with instructions for proper placement of the affected equipment; the customers should apply the enclosed cautionary statement label to Elements Obturation Unit and/or Replacement AC Power Cord as per instructions and complete and return the FIELD CORRECTION ACKNOWLEDGEMENT FORM ,whether or not they have any affected equipment in their inventory, via Fax to: 909-962-5605.For questions regarding this recall call 714-516-7602.Note: If you are an authorized SybronEndo international distributor, we request that you identify those customers that may have been shipped the affected lots of equipment and contact these customers to inform them of the issue within forty-eight (48) hours of receipt of this notification.If you have any questions, contact SybronEndo Customer Care at 1-800-346-3636. |  | 3615 | 3615 units total | Worldwide Distribution - USA (nationwide) and the countries of Australia, Azerbaijan, Belgium, Canada, China, Cyprus, Czech Republic, Egypt, Estonia, Finland, France, Germany, Greece, Hong Kong, Iceland, India, Italy, Japan, Kuwait, Latvia, Lebanon, Malaysia, Netherlands, New Zealand, Pakistan, Poland, Romania, Russia, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Thailand, Turkey, Ukraine, United Arab Emirates, and United Kingdom. | Terminated | November 07, 2013 | 294 | 3615 | Not_Computer
Z-1142-2013 | 64884 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Innova 2100IQ, Innova 3100IQ, Innova 4100IQ, |  |  | October 21, 2011 | 2011 | GE Healthcare, LLC | As part of GE Innova IQ table introduction, a label was designed to be affixed on the front Cover of theGantry Pivot section. This safety label is missing on systems manufactured between March 2010 and July 2011. In the situation a staff member (doctor, nurse, operator, anesthetist, etc&amp;) stands in the region between the C-arm gantry pivot and the head section of the table, there is a risk of co | Labeling design | A field action was initiated on October 21, 2011, and completed March 9, 2012, . The issue was corrected to all affected systems by applying the missing label onto the gantry. No recall letter was issued.For further questions please call (262) 513-4122. |  | 48 | 48 | Worldwide Distribution - Nationwide Distribution including the states of FL, IL, IN, LA, MI, MN, NJ, NM, NY, OH, PR, SC, TX and VA and the countries of Vietnam, Latvia, Italy, Israel, Indonesia, Hong Kong, United Kingdom, Finland, Spain, Czech Republic, Canada, Bahrain and Australia. | Terminated | February 04, 2014 | 837 | 48 | Not_Computer
Z-1135-2013 | 64747 | Yes | NPT |  |  | Cardiovascular | PMA | Edwards Lifesciences Crimpers, Model 9100CR26. |  |  | March 26, 2013 | 2013 | Edwards Lifesciences, LLC | Holes in the sterile barrier pouch may be caused by handling of the device during manufacturing or once pouches have been removed from the unit box and stored on the shelf by customers. | Package design/selection | Edwards Lifesciences notified customers via letter and instructed them to maintain the device in its protective unit carton until ready for use and to inspect the pouch for holes or tears prior to use as instructed in the product's labeling. If any holes or tears are found, customers are instructed to contact Edwards Customer Service to arrange for return and replacement of damaged product. |  | 6117 | 6,117 units | Nationwide Distribution-USA (nationwide) including DC and the states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, ME, MD, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, and WI. | Terminated | September 16, 2013 | 174 | 11589 | Not_Computer
Z-1507-2013 | 65182 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | FLEXSTENT Biliary Self-Expanding Stent System. |  |  | April 23, 2013 | 2013 | Cordis Corporation | Holes and/or tears can be introduced in the FLEXSTENTÂż inner pouch under some conditions, and to a lesser degree, in the outer pouch. | Package design/selection | Cordis sent an Urgent Medical Device Correction and Potential Removal letter dated April 24, 2013 e-mail to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immediately identify and set aside all affected products to prevent further use of the affected product, until inspected. Âż Use only product identified with a blue dot next to the lot number on the carton face label, carton end label, and outer pouch label, or product with upgraded sterile barrier system (not yet available, as of the date of this letter), review, complete, sign and return the enclosed Acknowledgement Form and maintain a copy of this notice with the affected product. For questions related to the issue, please contact our FSS/Cordisrepresentative, or please contact Cordis Direct at (800) 781-0282, Mondaythrough Friday from 9:00 AM to 5:00 PM EDT. |  | 23 | 23 units | Worldwide Distribution - USA (nationwide) and the states of: AZ, CA, CT, FL, GA, IA, IL, LA, MD, MI, MS, NC, NJ, NY, OH, PA, RI, SC, TX, WA, WI, WV, and the countries of: Czech Republic, Italy, Netherlands, and New Zealand | Terminated | June 27, 2014 | 430 | 2724 | Not_Computer
Z-1515-2013 | 65224 | No | JIT | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ADVIA CentaurÂż iPTH Calibrator 6 pk; REF: 10492387, 10630628. |  |  | May 14, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc | iPTH Assay- reagent and calibrator lots may have demonstrated higher than anticipated variability | No Marketing Application | Siemens sent an Urgent Medical Device Correction letter dated May 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. For questions, contact your Siemens Technical Solutions Center or your local Siemens technical support. |  | 198 | 198 (6 pk) | Nationwide Distribution. | Terminated | March 05, 2015 | 660 | 3908 | Not_Computer
Z-1127-2013 | 64646 | No | HSD | Shoulder joint humeral (hemi-shoulder) metallic uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | The Oxford Partial Knee; Oxford Knee Resection Procedure 3 Pack CEMENTED |  |  | March 14, 2013 | 2013 | Synvasive Technology Inc | Biomet part # 506076, lot 928182 was received from Synvasive Technology containing the incorrect blade. | Packaging change control | An Urgent Medical Device Recall Notice letter dated March 21, 2013 was sent to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to discontinue use; remove the products from inventory and complete and return the attached Response Form. For questions call 574-372-1570. |  | 100 | 100 kits | USA Nationwide Distributor including the state of IN | Terminated | January 13, 2014 | 305 | 100 | Not_Computer
Z-1517-2013 | 65289 | No | DXE | Embolectomy catheter. | Cardiovascular | Cardiovascular | 510(k) | lnvatec SpA brand Diver C.E., Model/Catalog Numbers: DVR010014145, DVR011014145, |  |  | March 04, 2009 | 2009 | Invatec Llc | Devices labeled for an intended use was not included in the existing 510(k) or PMA, in that, the Indications for use was modified. | Labeling False and Misleading | Invatec sent an Urgent - Device Field Action letter dated March 4, 2009, to all affected customer. The letter identified the product, the problem, and the action to be taken by the customer. Customers were required to return a confirmation form, acknowledging the action taken, and identifying affected units in their inventory for return of the product for replacement. The firm indicated that this recall was fully executed in 2009, and this 806 captures the events that took place at that time.For questions regarding this recall call 707-566-1143. |  | 177 | 177 units | Nationwide Distribution | Terminated | June 10, 2013 | 1559 | 177 | Not_Computer
Z-0546-2014 | 66589 | No | CAZ | Anesthesia conduction kit. | Anesthesiology | Anesthesiology | 510(k) | Continuous Peripheral Nerve Block Catheter Kit and Set |  |  | October 09, 2013 | 2013 | Arrow International Inc | A labeling inconsistency was discovered in which the lidstock states that the kit contains a stimulator cable with 1.5mm socket, however, the kit contains a stimulator cable with 1.5mm pin. | Labeling mix-ups | The firm, Arrow International, Inc., sent an "Urgent Medical Device Recall Notification" letter dated October 14, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately discontinue use and quarantine any of listed affected products/lots you have in stock; return affected products, customer service representative will contact you with RGA# and instructions for the return of the product, and complete enclosed Recall Acknowledgement form and fax to 1-800-343-2935, Attn: Customer Service, even if you have not affected product.If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-800-343-2935. |  | 121 | 121 | Worldwide distribution: US (nationwide) including states of: MA, LA, CA, GA, KS, TX, PA, NY, MO, WI, NJ, MS, CT, ID, AL, CO, and PR; and Singapore. | Terminated | September 21, 2015 | 712 | 121 | Not_Computer
Z-1143-2013 | 64538 | No | MCW | Intraluminal artery stripper. | Cardiovascular | Cardiovascular | 510(k) | Elana Catheter Sizer 2.0*** |  |  | February 14, 2013 | 2013 | Elana, Inc. | Distribution of an unapproved device into interstate commerce. | No Marketing Application | Elana Inc. issued an Urgent: Medical Device Recall letter on February 14, 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to return the Elana Catheter Sizer 2.0 as soon as possible. The Sizer 2.0 was intended as an optional accessory to the Elana Surgical Kit. The Elana Surgical Kit may be used properly without the Elana Catheter Sizer 2.0. Please acknowledge receipt of these instructions immediately by Email: Claartje@elana.com or Fax: 301-215-7038. The return shipping address for the Sizer 2.0 is Elana, Inc., 10480 Little Patuxent Parkway, Suite 400, Columbia, MD 21044 USA. Contact information for questions: Phone: 301-215-7010 or 443-538-4923 or +1 31 30 2537276 www.elana.com. |  | 3 | 3 devices | Distributed in the states of Pennsylvania, Texas, and New York. | Terminated | May 09, 2013 | 84 | 3 | Not_Computer
Z-1550-2013 | 65245 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | IDEV Technologies SUPERA VERITAS Interwoven Self-Expanding Nitinol Stent Transhepatic Biliary System. Indicated for palliative treatment of biliary strictures produced by malignant neoplasms |  |  | May 29, 2012 | 2012 | IDEV Technologies, Inc. | There was an error on the outer labeling of this device. The inner label was correct. | Labeling mix-ups | The firm decided to rework the products that were in their control at the time and the one that was not in their control was corrected by an IDEV employee. That hospital was called and notified of the situation in May 2012. |  | 3 | 3 | Nationwide Distribution to the states of FL, NY, and PA. | Terminated | July 29, 2013 | 426 | 3 | Not_Computer
Z-1465-2013 | 65209 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Active Articulation E1. |  |  | April 26, 2013 | 2013 | Biomet, Inc. | Biomet is recalling Active Articulation E1 Hip Bearing implants. Biomet has initiated this action following an investigation which identified that outer label indicates two conflicting O.D. sizes while the inner labels indicate an incorrect O.D. size.The implant itself is etched with the correct O.D. size. | Labeling mix-ups | Biomet sent an Urgent Medical Device Recall letter dated April 26, 2013 to consignees detailing that the outer label indicates two conflicting O.D. sizes while the inner labels indicate an incorrect O.D. size. This action requires the immediate location and discontinued use of the product and its return to Biomet. Ciustomers were advised to:Immediately locate and remove the identified device(s) listed from circulation .-Carefully follow the instructions on the enclosed "FAX Back Response Form" .- Fax a copy of the Response Form to 574-372-1683 prior to return of product.-Use priority carrier for your shipment.- If you have further distributed this product, you MUST notify hospital personnel of this action via the enclosed "Dear O.R. Manager" notice. This letter MUST be given to hospital personnel responsible for receiving recall notices. However, you are charged with the location and return of these products. |  | 8 | 8 units | Distributed in New York. | Terminated | August 29, 2013 | 125 | 8 | Not_Computer
Z-0656-2013 | 61718 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Accolade TMZF Plus 127 degree Neck Angle V40 Hip Stem #2.5 and Accolade TMZF Plus 127 degree Neck Angle V40 Hip Stem #3.5, Howmedica Osteonics Corp. The intended use is for reconstruction of the head and neck of the femoral joint. |  |  | October 23, 2012 | 2012 | Stryker Howmedica Osteonics Corp. | Stryker has received a report that single size 3.5 stem was packaged as size 2.5 stem resulting in a possible product mix. | Labeling mix-ups | Stryker sent "Urgent Medical Device Recall Notification" letters and product accountability forms via Fed Ex with return receipt on 10/23/12 to branches/agencies and to hospitals and physicians on 10/24/12. The letter described the issue and provided recommended actions to mitigate the risk. Customers can call (201) 972-2100 if questions. |  | 18107 | 6021-2530- four units; 6021-3535- | Nationwide distribution, including the states of NY, NJ, CA and FL. | Terminated | September 11, 2013 | 323 | 18107 | Not_Computer
Z-1477-2013 | 64164 | No | GQO | Herpes simplex virus serological assays. | Microbiology | Microbiology | 510(k) | Plastic bottle containing 10.5 mL of IgG Conjugate |  |  | January 23, 2013 | 2013 | Bion Enterprises Ltd | Affected products have a labeled expiration date of one year too long. | Error in labeling | MBL Bion sent an Urgent Product Recall letter dated January 28, 2013 to all affected customers via email. The letter identified the affected product, problem and actions to be taken. Customers were requested to destroy the affected product and to notify Bion of the number of units consumed and destroyed. For questions call 847-544-5044. |  | 207 | 207 units | Worldwide Distribution - United States (nationwide) in the states of IN, SC and the countries of Canada (Toronto); South Africa (Randburg) and Brazil (Rio de Janeiro) | Terminated | May 05, 2015 | 832 | 912 | Not_Computer
Z-1596-2013 | 65138 | No | MED | Liquid chemical sterilants/high level disinfectants. | General Hospital | General Hospital | 510(k) | Quick Connect QLC1677E, QRC1699E, QPC1721E and QFC1683E. STERIS Corporation, Mentor, OH 44060 |  |  | March 04, 2013 | 2013 | Steris Corporation | During a routine label review, the firm identified four SYSTEM 1E Quick Connects that will be updated to reflect the removal of obsolete device models. This labeling review process also highlighted the opportunity to combine two existing quick connects into a single product while eliminating the obsolete models. | Labeling Change Control | STERIS sent a Customer Notification letter dated February 19, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to replace their existing labels with labels they provided in their customer notification package.We apologize for any inconvenience this may have caused you and as always STERIS Corporation is dedicated to support our products and valued Customers. If you have any questions regarding this matter, please contact your STERIS Capital Account Manager. For further questions please call (440)392-7348 |  |  | QLC1677E: 38; QRC1699E: 128 &amp; QPC1721E/QFC1683E: 0 | Nationwide Distribution including the states of AK, AZ, CA, CO, FL, IL, IN, KY, LA, MA, MD, ME, MI, MO, MS, NC, NE, NH, NY, OH, OK, OR, PA, SD, TN, TX, VA, WA, WI and WY. | Terminated | July 02, 2013 | 120 | 0 | Not_Computer
Z-1080-2013 | 64594 | No | GWQ | Electroencephalograph. | Neurology | Neurology | 510(k) | Harmonie software versions 5.1 uo to 6.2e : |  |  | April 06, 2009 | 2009 | Stellate Systems | The firm is initiating this recall because the implied or stated indications in their labeling associated with Harmonie software have not yet been cleared by the FDA for sale in the USA. | No Marketing Application | Stellate sent an "DEVICE CORRECTION" letter dated April 6, 2009, to all affected customers. The letter identified the product, the problem and the action needed to be taken by the customer. Customers were instructed to check the product code referenced in the service manual to verify if your system is affected and check in "Help" &amp; "About Harmonie" in the menu to determine the software version on these systems. Stellate will shortly release updates to disable the noted features that were present in your particular version of the Harmonie software and revise labeling and Instructions for Use. All customers in possession of an affected system will be contacted by our Customer Support team to set up an appointment for installation of the updated version.If you have any questions or concerns regarding this notification, please do not hesitate to call our customer support line at 1-888-742-1306. |  | 2660 | 2660 Software licenses | Worldwide distribution: USA (nationwide) including states of: AL, AR, AZ, CA, CT, DE, FL, HI, IA, IN, IL, KS, KY, LA, MA, MD, ME, MI, MN, MO, NC, ND, NJ, NY, OH, OR, PA, RI, SC, TX, VA, VT, WI, WV, and WY; and the countries of: Argentina, Australia, Bahrain, Brasil, Canada, China, Denmark, Finland, France, Germany, Greece, Haiti, Hong Kong, Hungary, Indonesia, Iran, Israel, Italy, Japan, Korea, KSA, Mexico, Netherlands, Norway, Pakistan, Portugal, Rep. DE Panama, Saudi Arabia, Spain, Switzerland, Taiwan ROC, Thailand, Turkey, and United Kingdom. | Terminated | April 22, 2013 | 1477 | 2660 | Not_Computer
Z-1119-2013 | 64610 | No | MRG | Vitamin D test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA Centaur Vitamin D Assay (100 tests); Reference Number: 10491994. |  |  | March 13, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc | Negative bias in patient samples for Vitamin D | No Marketing Application | Siemens Healthcare issued Urgent Medical Device Recall was sent to all affected Siemens Healthcare Diagnostics customers in the United States on March 13th, 2013. This notice informs the customer of the recall and the actions that need to be taken. The Urgent Medical Device Recall instructs customers to discontinue the use of ADVIA Centaur VitD Assay Kit Lots (REF 10491994 and 10631021) and ADVIA Centaur Systems VitD Cal (REF 10493589 and 10630911) released prior to the recalibration of the assay and request replacement lots. This notice also directs customers to review the contents of the recall notice with their laboratory or medical director. A confirmation fax-back form is included to ensure that all customers in the United States have been notified and have complied with the mandatory action. |  | 3871 | 3871 | Nationwide Distribution. | Terminated | March 05, 2015 | 722 | 8489 | Not_Computer
Z-1157-2013 | 64578 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | APLIF Implants and Instruments, |  |  | March 08, 2013 | 2013 | Spinal Solutions, LLC | Spinal Solutions is recalling the APLIF system because it is not supported by adequate testing and documentation to demonstrate that it meets performance or safety standards. These inadequacies might result in product performance failures that could cause patient harm due to implant breakage, movement, or inadequate sterilization. | No Marketing Application | Spinal Solutions LLC sent an Urgent Medical Device Recall letter dated April 12, 2013, to all affected customers. The letter informed the customers of the problem identified and the actions to be taken. Customers were instructed to discontinue use of the affected product and return to the firm. Customers were also instructed to complete the bottom of the form and to return it by fax to (858) 764-9739.For questions regarding this recall call 951-304-9001. |  | 220 | 220 units | Nationwide Distribution including NV, WI, MD,and CA. | Terminated | March 26, 2014 | 383 | 220 | Not_Computer
Z-1451-2013 | 65160 | No | FXX | Surgical apparel. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Comfort brand Ear-Loop Face Mask, Item #: 3100, Model Numbers: # 3100B, # 3100P, # 3100G, # 3100Y, # 3100W; Product is manufactured and distributed by Maytex Corporation, CA |  |  | May 15, 2013 | 2013 | Maytex Corp | Recent Lab testing conducted in March 2013 indicates that the product does not meet the standard of "Differential Pressure of only 1.6 Delta P (mm H2O/cm2)" as indicated on the box/package of the product; and also product manufactured between 2008 - 2012 cannot be guaranteed as Fluid Resistant. | Labeling False and Misleading | The firm, Maytex Corporation, sent a "MEDICAL DEVICE SAFETY NOTICE" dated May 15, 2013, to all its customers. The notice described the product, problem and actions to be taken. The customers were instructed to immediately contact your medical professional for treatment and contact Maytex Corporation to file a complaint should any illness or injury occur; discontinue use of the product if you detect any potential health hazards; if this product has been further sold or distributed, notify your customers to the retail level and send copies of safety notice to customer/users; and complete and promptly return the enclosed Medical Device Safety Notice Return Response Acknowledgement and Receipt form via mail to: Maytex Corporation 23521 Foley Street, Hayward CA 94545; fax at; 510-786-0209; email: www.maytex18@aol.com or by delivery to Maytex Corporation.Maytex Corporation has temporarily taken down their Website at www.maytexcorp.com to correct the wrong product specifications as described in the notice.If you have any questions, contact Maytex Corporation at 1-800-462-9839 or email: maytex18@aol.com Monday through Friday, 8:30 AM to 5:00 PM Pacific Time. |  | 43915 | 43,915 Cases | Worldwide distribution: USA (nationwide) including states of: AL, AZ, CA, CO, FL, GA, HI, ID, IL, IN, MA, MD, ME, MI, MN, MO, MT, NE, NH, NJ, NM, NY, OK, OR, PA, PR, TN, TX, UT, WA, and WI; and country of: Mexico. | Terminated | October 29, 2013 | 167 | 43915 | Not_Computer
Z-0493-2014 | 66967 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | G7 Acetabular Screw 6.5mmx35mm |  |  | November 26, 0013 | 13 | Biomet, Inc. | Investigation identified that the outer label, inner label, and patient label all have the incorrect diameter size of 10.5mm. The screw is actually a 6.5mm diameter and correctly corresponds to the items number on the label. | Labeling Change Control | Beginning November 25 Urgent recall electronic notices were sent by Biomet UK to all distributors with hospitals outside the USA notified by ground mail November 27,2013. The notification identified the product and reason for the recall, asking for immediate location and discontinued use of the product and return to Biomet UK, Ltd. Questions can be addressed to Biomet USA at 01 574-372-1570 M-F 8 am to 5 pm. |  | 532 | 532 in USA all under Biomet USA control | Worldwide Distribution - USA Nationwide and in the countries of Europe, Italy, UK, Australia, Japan. all product in USA is under Biomet control. | Terminated | September 29, 2014 | 730792 | 532 | Not_Computer
Z-0516-2014 | 66894 | No | GGN | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(K) Exempt | HemosIL High Abnormal Control 3 UNASSAYED |  |  | November 15, 2013 | 2013 | Instrumentation Laboratory Co. | HemosIL Low, Abnormal, Normal Controls UNASSAYED (no-value assigned) lacks 510(k) approval | No Marketing Application | Instrumentation Laboratory issued Important Correction Notice via USPS to US customers dated November 2013 advising users in the US that unassayed controls are not acceptable for use in their unassayed (non-value assigned) format. IL will no longer sell or distribute controls without assayed values for the US market as of November 15, 2013, and are converting US customers to an assayed control format. Additional contact with customers via phone calls and a reminder email.An enclosed response tracking form provided to be returned. For technical questions, contact the IL Technical Support Center at 1-800-678-0710, Option # 6 (ACL Coagulation Products).For replacement product if value assignments are not available for your product lot, please contact your IL Customer Service Representative at 1-800-955-9525, Option # 6.For questions regarding this recall call 781-861-4467. |  | 84561 | 84561 kits | Nationwide Distribution | Terminated | March 25, 2014 | 130 | 171384 | Not_Computer
Z-0490-2014 | 66735 | [6 empty] | Nephros Dual Stage Ultra Filter (DSU) Brochure |  |  | October 28, 2013 | 2013 | Nephros Inc | Promotional materials for non-medical water filtration products were determined to promote claims which constitute marketing the product as a medical device. | Labeling False and Misleading | Nephros Inc., sent an Urgent Medical Device Recall letter dated October 29, 2013 to all affected customers via UPS 2nd Day Air. The letter identified the affected product, problem and actions to be taken. Customer are asked to provide a customer information list, immediately remove and discard all affected product labeling materials and any independent labeling materials they have created for the affected products. Customers are instructed to complete and sign the attached Response Form to acknowledge receipt, understanding and compliance with the instruction. For questions contact the Nephros representative at 201-343-5202 x 100.The letter was also posted at : http://www.nephros.com/download/news/DSU_Recall_Letters.pdf |  | 3367 | 3,367 untis with affected labeling | USA Nationwide Distribution | Terminated | July 06, 2016 | 982 | 3367 | Not_Computer
Z-1125-2013 | 64568 | Yes | BTL | Powered emergency ventilator. | Anesthesiology | Anesthesiology | 510(k) | Instrumentation Industries BE183-SU and BE 183 SU-12 Manual Jet Ventilators; BE 183-SUR and BE 183 SUR-12 Manual Jet Ventilators with Regulator and Gauge; BE 183-2 4 ft. Small Bore Tubing Assembly w/Inline Filter; and BE 183-R Regulator |  |  | September 07, 2012 | 2012 | Instrumentation Industries Inc | Firm initiated an update to Directions for Use necessitating replacement of previously distributed Directions for Use. | Error in labeling | Instrumentation Industries, Inc. initiated a recall of this product on September 7, 2012 by sending a letter via priority mail to all consignees. The letter identified the product, the problem, and the action to be taken by the customer. Consignees were instructed to replace the package insert of any affected product with the enclosed corrected package insert (now at Rev. H). The Directions for Use can alsbe downloaded from their website if they should need additional copies. If the product was further distributed consignees were asked to notify their consignees about this Field Correction.For questions regarding this recall call 412-854-1133. |  | 2998 | 2998 | Worldwide Distribution - USA including AK, AL, AZ, CA, CO, CT, DC, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NH, NJ, NV, NY, OH, OK, OR, PA, PR, SC, SD, TN, TX, UT, VA, VT, WA, WI and WV. Shipped internationally to Canada, Ireland, Hong Kong, the Netherlands, New Zealand, and Thailand. | Terminated | July 05, 2013 | 301 | 2998 | Not_Computer
Z-1565-2013 | 65254 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Baxter, MiniCap with Povidone-Iodine solution, 5C4466P, Baxter Healthcare Corporation. For dialysis. |  |  | May 21, 2013 | 2013 | Baxter Healthcare Corp. | leaking pouches | Packaging change control | The firm, Baxter Healthcare Corporation sent an "URGENT PRODUCT RECALL" letter dated May 21, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. For general questions regarding this communication, contact The Center for One Baxter at 1-800-422-9837. |  | 629280 | 629,280 units | Nationwide Distribution. | Completed |  |  | 629280 | Not_Computer
Z-1564-2013 | 65458 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Medtronic, Launcher Guiding Catheter, Model # LA8EBU45SH, Lot # 0006722866, 8F and Model # LA6IMAD, Lot # 0006724224, 6F |  |  | June 12, 2013 | 2013 | Medtronic Inc. Cardiac Rhythm Disease Management | Medtronic has identified a labeling issue for 24 units of the Launcher coronary Guide Catheter. The two affected catheter lots are mislabeled and do not accurately reflect the catheters contained within each package. | Labeling mix-ups | Medtronic "Urgent Field Safety Notice" letter dated June 2013 to all affected customers. The letter described the affected product, problem and actions to be taken. The letter instructed customers to remove and quarantine all affected product within their inventory and return the affected product to Medtronic. For questions contact your local Medtronic representative. |  | 15 | 15 | Internationally distributed: to countries of Germany and Poland. | Terminated | November 19, 2013 | 160 | 15 | Not_Computer
Z-1568-2013 | 65362 | No | KYI | Wrist joint carpal trapezium polymer prosthesis. | Orthopedic | Orthopedic | 510(k) | CMC Saddle |  |  | June 03, 2013 | 2013 | Integra LifeSciences Corp. | As a result of a quality review, labelling content discrepancies were identified between the Instruction for Use (IFU) distributed in the US with the Ascension CMC and the content that is cleared by FDA. | Labeling Change Control | Integra sent an Urgent: Medical Device Recall letter to all affected product. The letter identified the affected product, problem and actions to be taken. Customers were instructed to examine inventory, discontinue use, and complete and return the attached form. When the form is received, Customer Service Returns and Repairs will send an RMA number to return the affected inventory. For quwstions contact your regional managers or call 512-520-7697. |  | 467 | 467 devices | USA Nationwide Distribution | Terminated | October 08, 2014 | 492 | 467 | Not_Computer
Z-1572-2013 | 65417 | No | FGY | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | MAQUET GETINGE GROUP Arterial Cannula OD |  |  | May 14, 2013 | 2013 | Maquet Cardiovascular Us Sales, Llc | The specified part numbers and lot numbers of Maquet Pediatric Arterial Cannula Packs may contain an incorrectly sized connector. | Packaging process control | Maquet sent an "Urgent Product Recall"/"Recall - Device Corrective Action Response" forms dated May 14, 2013, via Federal Express to all affected customers. The notification identifies affected product; informs customers of the nature of the issue; advises of potential adverse patient consequences that may occur; and advises customers to follow instructions on the "Recall-Device Corrective Action Response" forms included as part of the notification.If you have any further questions please contact your local MAQUET representative or call Customer Service at (888) 627-8383. |  | 100 | 100 units (60 Domestic; 40 International) | Worldwide Distribution - US Distribution including the states of Delaware, Illinois, Minnesota , New York., Ohio and Pennsylvania and the countries of Germany, Turkey and France. | Terminated | September 18, 2014 | 492 | 100 | Not_Computer
Z-0543-2014 | 66844 | No | MVN | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic | 510(k) | ClearSight(TM) Toric 1 day (ocufilcon D), 30 daily disposable contact lenses for astigmatism, STERILE, Rx only --- CooperVision Hamble, SO31 4RF, UK - Scottsville, NY 14546 USA - Juana Diaz, PR 00795 USA |  |  | November 20, 2013 | 2013 | CooperVision Inc. | Lots were labeled with incorrect expiration dates on secondary package (carton). The primary package (blister) is marked with the correct expiration date. | Incorrect or no expiration date | CooperVision sent an Urgent Medical Device Recall Letters dated November 20, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers. Our records indicate you received some of the affected product. Enclosed you will find a list of affected carton lot numbers, with the correct expiration date, that were shipped to your account." Please immediately examine your inventory," Stop any further distribution of carton(s) subject to the recall," Quarantine carton(s) with incorrect expiration date, and" CooperVision will make arrangements to remove affected, quarantine product.As soon as possible, please complete the attached response form and fax it to 1.866.484.1048 or scan and email to info@coopervision.com. If you have product to return please contact customer service for instructions and to schedule pick-up and removal of affected, quarantined product. Your account will be credited once the product is returned and processed.We appreciate your assistance and apologize for any inconvenience you or your patients experienced by this removal. If you have any further questions, please feel free to contact us at 800-341-2020. |  | 45 | 45 cartons | Worldwide Distribution - USA (nationwide) and the countries of Aruba, Canada, Colombia and Trinidad. | Terminated | July 22, 2015 | 609 | 198 | Not_Computer
Z-0714-2013 | 63959 | No | FIE | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Medisystems 15 Gauge x 3/5 inch Buttonhole Needle Set |  |  | December 20, 2012 | 2012 | Medisystems a NX Stage Company | A total of 5 (ea) single needle pouches are mislabeled as 15 Ga x1" instead of 15 Ga x 3/5". | Labeling mix-ups | Medisystems sent a notification letter dated December 20, 2012 to affected customers. The letter identified the affected products, problem and actions to be taken. Customers were instructed to immediately open and inspect inventory for the affected products. Results of the inspection are to be provided to Medisystems and the product is to be returned to Medisystems. For questions call (978) 332-5973. |  | 26 | 26 casesx 250 needles (6500needles) | US Nationwide Distribution in the states of NY, PA and TN. | Terminated | October 10, 2013 | 294 | 42 | Not_Computer
Z-0716-2013 | 64043 | No | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Abbott Point of Care i-Stat cTnI cartridge. |  |  | November 01, 2012 | 2012 | Abbott Point of Care Canada Limited | APOC has determined that 10 boxes of i-Stat cTnI cartridges from lot P12179 have an incorrect barcode applied to the cartridge portion pack. | Error in labeling | APOC sent an "IMPORTANT PRODUCT INFORMATION" letter dated October 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact Abbott Point Of Care Technical Support at 800-366-8020, Option 1 or your Abbott Point Of Care representative for questions regarding this information. |  | 560 | 560 boxes (25 cartridges in each box) | Worldwide Distribution. | Terminated | June 14, 2013 | 225 | 560 | Not_Computer
Z-0661-2013 | 63877 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | STERRAD NX Cassettes, P/N 10133. |  |  | September 20, 2012 | 2012 | Advanced Sterilization Products | Advanced Sterilization Products is recalling three lots of STERRAD NX System Cassettes because it does not have adequate data to support the entire duration of the product's labeled shelf-life. | Labeling design | Advanced Sterilization Product sent an Urgent Medical Device Recall (Removal) letter dated December 4, 2012, to all affected customers. The letter informed the customers that the Sterrad NX Cassettes do not have adequate data to support the entire duration of the product's labeled shelf-life. Customers are instructed of the problems identifed and the actions to be taken. Customers are instructed to notify their customers if they have further distributed the recalled products and to have them contact Stericycle at (877) 650-7686 to arrange for return of the product. Customers with questions are instructed to call Stericycle at (877) 650-7686 Mon-Fri from 8am to 5pm, ET.For questions regarding this recall call 949-453-6400. |  | 12280 | 12,280 units | Worldwide Distribution - USA (nationwide) and internationally to the following countries:AR, CO, TW, SI, AU, MX, MY, HU, PH, BE, FR, DE, IE, IT, ES, NO, CH, and GB. | Terminated | October 17, 2013 | 392 | 12280 | Not_Computer
Z-1347-2013 | 64618 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | PROGRESSIVE MEDICAL Laparoscopic Electrode PTFE Coated L Hook 33cm. |  |  | February 15, 2013 | 2013 | Unimed Surgical Products, Inc. | Firm is recalling due to concerns regarding changes that were made to the PTFE coating materials. | No Marketing Application | Unimed Surgical Products, Inc. issued a Medical Device Correction and Removal letter dated February 15, 2013 to its customers. Customers were instructed to quarantine any undistributed product, and notify its customers of the recall to retrieve any unused product. For further information, contact Unimed Surgical Products Inc. at 1-800-531-2627. |  | 120416 | 120,416 units | Distributed in the states of MO, TN, and PA. | Terminated | August 07, 2013 | 173 | 120416 | Not_Computer
Z-1304-2013 | 64908 | No | KDD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Curity Staple Remover Kit; Product Code: 66701. |  |  | April 12, 2013 | 2013 | Covidien LLC | On March 26, 2013 Covidien received a recall notice from one of their suppliers CareFusion. CareFusion is recalling a variety of OTC drug products including the 10% Povidone Iodine SEPP Applicator which is utilized in specific convenience kits. | Error in labeling | US Customers were notified of the recall by letter sent by Federal Express on April 12, 2013. The letter requests that they immediately stop using any affected product, quarantine it, and return it to Covidien. Consignees were asked to return a verification form and return it to Covidien via fax. |  | 52363 | 52363 distributed | Worldwide distribution: USA (nationwide) and in the country of: Panama. | Terminated | June 02, 2015 | 781 | 70103 | Not_Computer
Z-1796-2013 | 65635 | No | JPA | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(k) | TEG 5000 Thromboelastograph Coagulation Analyzer |  |  | June 20, 2013 | 2013 | Haemonetics Corporation | Research Use Only (RUO) parameters were included in TEG 5000 User Manual. | Labeling False and Misleading | Haemonetics sent a Medical Device Correction letter dated June 20, 2013 to all affected customers. The letter identified the affected product, problem and action to be taken. The letter instructs customers to discard the current manuals and replace with revised manuals included with the Customer Notification Letter. A Consignee Response Form has been included with the letter, which is to be signed and returned for effectiveness purposes. Follow up will be made to customers who don't return the Response Form. For questions contact TEG Technical Support team at 1-800-438-2834. |  | 1590 | 1590 devices | US Nationwide | Terminated | June 04, 2015 | 714 | 1590 | Not_Computer
Z-0406-2014 | 66413 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes 2.4 mm LC-DCP Plate 6 Holes/51mm |  |  | July 30, 2013 | 2013 | Synthes (USA) Products LLC | A complaint was reported regarding a screw being found in a package labeled for 2.4 mm LC-DCP Plate. | Labeling mix-ups | Synthes sent an Urgent Notice: Medical Device Recall letter dated July 30, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to review their inventory, immediately remove the affected product and return to Credit/Returns, Synthes. For questions call 610-719-5450. |  | 13 | 13 | Worldwide Distribution - USA (nationwide) including IN, NY, WA, NE, MD, TX, AL, MI, and Internationally to Canada. | Terminated | August 17, 2015 | 748 | 13 | Not_Computer
Z-1341-2013 | 64943 | No | OTO | Surgical mesh. | General &amp; Plastic Surgery | Obstetrics/Gynecology | 510(k) | Rostorelle L, Polypropylene Mesh, Catalog number 501440, Manufacturer Coloplast A/S 3050 Denmark. |  |  | April 12, 2013 | 2013 | Coloplast Manufacturing US, LLC | Coloplast is recalling one lot of Restorelle L prosthesis because it is mislabeled. The incorrect Instructions For Use (IFU) pamphlet is included in the packaging. | Labeling mix-ups | Coloplast sent a letter to dated April 12, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The letter included an attachment with the correct IFU to replace the incorrect IFU. Customers with questions were instructed to call Customer Service at 800-258-3476.For questions regarding this recall call 612-302-4992. |  | 111 | 111 | Nationwide Distribution including AL, CA, CO, DC, IL, IN, KY, MD, MA, MO, NH, OH, RI, and VA. | Terminated | January 21, 2014 | 284 | 111 | Not_Computer
Z-0410-2014 | 66414 | No | OAR | Cement dispenser. | Orthopedic | Orthopedic | 510(K) Exempt | Vertebroplasty Needle Kit 10g Diamond Tip |  |  | July 30, 2013 | 2013 | Synthes USA HQ, Inc. | Complaints were received of inconsistencies in the expiration dates shown on the packaging of the Vertebroplasty Needle Kit 10g Diamond Tip device. | Incorrect or no expiration date | Synthes sent an Urgent Notice: Medical Device Notification letter dated October 8, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to returny any of the affect devices, call Synthes at 1-800-479-6329 to obtain a Return Authorization Number, complete the Verification Section indicating the affected product has been located, also indicate the number of devices found and note the Return Authorization Number. Return the Verification Form to:Credit/ReturnsSynthes1101 Synthes AvenueMonument, CO 80132Customers were also instructed to return the Verification Section even if they don't have the affected product. Customers with questions were instructed to call 610-719-5450 or contact their Synthes Sales Consultant.For questions regarding this recall call 610-719-5000. |  | 40 | 40 | Nationwide Distribution including IN, MT, HI, and FL. | Terminated | April 16, 2015 | 625 | 40 | Not_Computer
Z-1810-2013 | 65600 | syringe, irrigating (non dental) | KYZ | Irrigating syringe. | General Hospital | General Hospital | 510(K) Exempt | Medtronic Preimplantation Test Kit, Catalog No. 21047. The kit is indicated for use with Medtronic CSF-Flow Control Valves and Delta Valves, when a simple preimplantation test is desired to verify that the valve conforms to labeled product specifications. |  |  | June 28, 2013 | 2013 | Medtronic Neurosurgery | Medtronic Neurosurgery is recalling the Medtronic Preimplantation Test Kit because the outer carton had expiration dates that were incorrect. The date on the outer carton is later than the expiration date for one or more of the individually packaged sterile components in the kit. | Incorrect or no expiration date | An "Urgent Field Safety Notice" recall letter dated 6/28/13 was sent to customers who purchased the Medtronic Preimplantation Test Kit, catalog number 21047, to inform them of the recall due to the incorrect outer carton expiration dates. The recall letter informed the customers of the problems and actions identified. Customers were instructed to contact the Quality Department at (805) 571-8725 for questions. |  | 915 | 915 units | Nationwide Distribution in the US. | Terminated | September 23, 2013 | 87 | 915 | Not_Computer
Z-1213-2013 | 64859 | No | OIY | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | 3DKnee e+ Tibial Insert. The insert is used along with the other components that comprise the 3DKnee System. |  |  | April 01, 2013 | 2013 | Encore Medical, Lp | U.S. Agent contacted Director Commercial Logistics, notifying him of incorrect color coded labels on two packages of e+tibial inserts. There were 8 e+ tibial inserts of the same lot number in Finished Goods as the two from the agency. These devices were inspected and all were incorrectly labeled. | Labeling mix-ups | The firm sent their US customer a "Urgent Field Safety Notice" letter on April 1, 2013. The letter explained the product problem and provided recommended actions as well as customer contact numbers. |  | 5 | 5 | US Distribution to the state of California and International Distribution to the country of Venezuela. | Terminated | July 01, 2013 | 91 | 5 | Not_Computer
Z-1808-2013 | 65654 | No | NJE |  |  | Neurology | HDE - Humanitarian Device Exemption | Stryker brand Neuroform 3 Microdelivery Stent System, 4.5 mm X 20 mm, REF Catalog Number: E345020, Model/Product Number: M003450200, Product is manufactured by Boston Scientific Corporation, Fremont, CA and distributed by Stryker Neurovascular, Fremont, CA |  |  | June 24, 2013 | 2013 | Stryker Neurovascular | The device is labeled incorrectly, in that, the device Model Number on the outer carton and inner pouch labels did not match, although both referenced the same lot number. The Model Number on the outer box states UPN M003E345020 (4.5mm x 20 mm Neuroform 3 Stent) while the pouch label states UPN M003E3450300 (4.5mm x 30 mm Neuroform 3 Stent). | Packaging process control | Stryker Neurovascular sent an Urgent Field Safety Notice on June 28, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately check their internal inventory and quarantine all subject devices pending return to Stryker. Circulate the Field Safety Notice internally to all interested/affected parties. Maintain awareness of this notice internally until all requred actions have been completed within their facility. Inform Stryker if any of the subject devices were distributed to other organization. Complete the attached customer response form and return to their Stryker Representataive.For questions regarding this recall call 510-413-2500. |  | 23 | 23 units | International Distribution including Austria, China, and Germany. | Terminated | September 30, 2013 | 98 | 23 | Not_Computer
Z-0382-2014 | 66569 | No | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | Tri-Nex Healing Abutment, 0.50 mm x 5 mm height, Catalog No. HA-L-50-5. |  |  | September 13, 2013 | 2013 | Southern Implants, Inc | Southern Implants is recalling the Tri-Nex healing abutment because the package was incorrectly labeled and packaged as the Tri-Nex healing abutment, 05.0 mm x 5 mm height (HA-L-50-5). The package contains the Tri-Nex healing abutment 05.0 mm x 6.0 mm Flare x 5 mm Height (HA-L-50W-5). | Labeling design | Southern Implants, Inc. sent an Urgent Product Action letter dated September 11, 2013, to all affected customers. The letter identiied the product, the problem, and the action to be taken by the customer. Customers were instructed to fill out the customer response form and fax/email the completed form to Customer Support at (866) 903-9272 or customersupport@keystonedental.com.For questions regarding this recall call 866-903-9272. |  | 197 | 197 units | Worldwide Distribution - USA (nationwide) and Internationally to Canada, Guatemala, Taiwan. | Terminated | January 21, 2014 | 130 | 197 | Not_Computer
Z-0384-2014 | 66415 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Synthes Matrix Threaded Persuader |  |  | July 31, 2013 | 2013 | Synthes USA HQ, Inc. | Synthes became aware that during a procedure using the Matrix Threaded Persuader, the Reduction Insert malfunctioned causing pieces to break off. Once there is a breakage to the Reduction Insert it will not allow the instrument to function as required. | Labeling design | The firm, Synthes, sent an "URGENT NOTICE: MEDICAL DEVICE RECALL" letter dated July 31, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to review your inventory and immediately remove the affected lots from stock; If you do have the affected devices, contact Synthes at 1-800-479-6329 to obtain the RA# and return the product to with completed form to Credit/Returns, Synthes, 1101 Synthes Avenue, Monument, CO 80132; complete and sign the Return Receipt Requested form via fax at 1-888-731-7955 or email to: FieldAction@synthes.com, even if you do not have the affected product.If you have questions, please call 610-719-5450. |  | 634 | 634 | Worldwide distribution: US (nationwide) and countries of Canada and Switzerland. | Terminated | August 31, 2015 | 761 | 634 | Not_Computer
Z-1214-2013 | 64820 | Yes | KPO | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | NxStage PureFlow-B Solution, Model/Catalog Number: RFP-401, packaged in a case |  |  | March 18, 2013 | 2013 | NxStage Medical, Inc. | Product may be mislabeled. | Labeling mix-ups | NxStage sent a letter dated March 18, 2013, via email to all consignees instructing them to check their entire inventory and to segregate and quarantine lot F023444 pending further instruction from NxStage.On 4/2/13 NxStage sent a letter via email to all consignees requesting a return of all RFP-400 and RFP-401 labeled with lot F023444.. Consignees were instructed to complete the recall reply form and provide a list of customers (name address, phone number) and fax them to NXStage Medical, Inc.Customers with questions were instructed to call 1-866-697-8243.For questions regarding this recall call 1-866-697-8243. |  | 1480 | 1480 cases (2 bags per case) | Nationwide Distribution including IL, SC, TX, VA, FL, AR, and NY. | Terminated | September 17, 2014 | 548 | 1480 | Not_Computer
Z-1809-2013 | 65391 | No | LXM |  |  | Physical Medicine | 510(k) | Advanced Orthogonal Percussion Adjusting Instrument |  |  | March 06, 2013 | 2013 | Advanced Orthogonal Equipment, Incorporated | Unapproved medical device | No Marketing Application | The firm, Advanced Orthogonal Equipment, sent a "CORRECTION NOTICE" to its consignees/customers. The notice described the product, problem and actions to be taken. The customers were instructed that " none of these devices be used pending final clearance by FDA".Questions and concerns regarding this notice should be directed to the President of Advanced Orthogonal Equipment by phone at: 727-528-8700 or email: Atlasortho@aol.com. |  | 54 | 54 | Nationwide distribution: US including states of: AL, CA, CO, FL, GA, ID, KY, MD, NC, NJ, NY, OH, PA, TN, VA, and WA. | Open, Classified |  |  | 54 | Not_Computer
Z-1395-2013 | 63204 | No | OAR | Cement dispenser. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker Orthobiologics Aliquot Delivery System Manual Surgical Instrument - Aliquot Reamer, 8", Part Number 2110-0525 |  |  | November 10, 2012 | 2012 | Orthovita, Inc., dBA Stryker Orthobiologics. | There have been a few reported cases of the catheter breaking when removing the Aliquot catheter from the needle during vertebroplasty procedures. This has been attributed to bending of the catheter during the retraction from the needle. It was determined to be use errors. The correction is limited to a revision of the Directions for Use. | Labeling design | Stryker sent an Urgent Medical Device Correction letter dated November 10, 2012 to all affected customers. The letter identified the affected product, problem issue, potential hazards and actions to be taken. Customers were instructed to immediately review this notification and the attached IFU. Customers were asked to complete and sign the enclosed Business Reply Form and fax a copy to 1-484-323-8803. For questions contact Stryker Orthobiologics. |  | 5 | 5 | USA Nationwide Distribution in the states of AZ, CA, CO, CT, DE, FL, GA, HI, IL, IN, KS, KY, LA, MD, ME, MI, MO, NC, ND, NH, NJ, NY, OH, OK, PA, RI, SC, TN, TX, UT, VA, WA, and WI. | Terminated | October 29, 2014 | 718 | 1413 | Not_Computer
Z-1057-2013 | 64593 | No | DYG | Flow-directed catheter. | Cardiovascular | Cardiovascular | 510(k) | Arrow/Teleflex Medical Berman Angiographic Balloon Catheter and Reverse Berman Angiographic Balloon Catheter. |  |  | March 04, 2013 | 2013 | Arrow International Inc | The product specifications sheet, part of the Instructions for Use (IFU), for the specified product numbers, reference Renografin-76 contrast medium (viscosity of 8.4 centipoise), which is no longer available. Firm is in the process of updating the labeling for these products. In the meantime, when choosing an alternative contrast medium, health care providers are instructed to select a contrast | Labeling Change Control | Arrow sent an Urgent Medical Device Customer Notification letter dated March 4, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Arrow is in the process of updating the labeling; in the meantime they wanted to instruct all customers to select a contrast medium with a viscosity of 8.4 centipoise or lower for use with the pressure and flow settings specified in the current IFU. In addition, they are placing an insert card in kits for the specified product codes alerting them that Renografin-76 is no longer available and to select a contract medium with a viscosity of 8/4 centipose or lower for use with the pressure and flow setting specified in the current IFU.Arrow remains committed to providing high quality, safe and effective products. If you require additional information or clarification regarding this matter, please contact your Sales Representative or Arrow's Customer Service Department at (866) 396-2111. |  | 99627 | 99627ea | Worldwide Distribution - USA including Puerto Rico and the states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI and WV. and the countries of Australia, Dominican Republic, Uruguay, Ecuador, Chile, Taiwan, Pakistan, Colombia, Argentina, Korea Republic, New Zealand, Guatemala, Thailand, Peru, Bahamas, Philippines, Trinidad and Tobago, Japan, Ireland, India, Singapore, China, Canada and Mexico. | Terminated | December 21, 2013 | 292 | 99627 | Not_Computer
Z-1821-2013 | 65668 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Synthes Synapse SD Drill Guide, Part Number SD389.477, Lot Number 6913900 |  |  | October 12, 2012 | 2012 | Synthes USA HQ, Inc. | A complaint was received from the field, that the drill guide was etched incorrectly, 12mm etch instead of 18mm drill guide. The firm identified that this constituted a field action and retrospectively reported this action. | Labeling Change Control | The firm, Synthes USA, sent an email to its customer. The email identified the product, problem and the actions to be taken. The customer was instructed to do the following: do not use this product; do not remove any product from your account or return any product without further instructions. If you have identified that you have the product, please have the product located in an area that ensures it will not be used. Further instructions will be sent out with regards to any action that needs to be taken next.Please call Synthes if you have any questions at 610-719-5364. |  | 2 | 2 | US distribution: PA only.. | Terminated | June 11, 2015 | 972 | 2 | Not_Computer
Z-0392-2014 | 66389 | No | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Baxter, VIAL-MATE Adaptor, 2B8071, Rx Only, STERILE R, Baxter Healthcare Corporation, Deerfield, IL 30015 USA |  |  | September 25, 2013 | 2013 | Baxter Healthcare Corp. | Incomplete foil seal on one lot of sterile product. | Packaging process control | Baxter Healthcare sent an Urgent Product Recall letter dated September 25, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to locate and remove all affected product from their facility, contact Baxter's Center for Service at 1-888-229-0001 to arrange for return and credit. Complete the attached Customer Reply Form and return it to Baxter by fax or email. Customers with questions were instructed to call 1-800-422-9837. For questions regarding this recall call 224-948-4770. |  | 38400 | 38,400 units | Nationwide Distribution including WV, KY, IA, AR, OH, IL, IN, MI, KS, TN, WI, MN, MO. | Terminated | July 22, 2014 | 300 | 38400 | Not_Computer
Z-1336-2013 | 64468 | No | LPL | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic | 510(k) | ACUVUE ADVANCE Brand Contact Lenses with HYDRACLEAR, Soft, disposable contact lenses. |  |  | December 21, 2012 | 2012 | Johnson &amp; Johnson Vision Care, Inc. | Limited number of individual contact lens packages may not have been completely sealed. | Packaging | The firm, Johnson &amp; Johnson, Vision Care, Inc. sent a letter dated January 9, 2013 to all their consignees stating the following actions:" Please contact your patients that may have received any of the affected product and instruct them to discontinue use immediately and return product to you for replacement " Please return the enclosed Business Reply Card (BRCl confirming for our records that you have received this notice, and indicating whether or not you have in stock, or have distributed to patients,any product from the lots noted above. " Please return any remaining affected product in your possession using the enclosed pre-paid mailer. " Pass this notice on to anyone in your practice that needs to be informed " Customer service will arrange for the return and replacement of any affected product. If you have any questions and/or if replacement lenses are needed, please contact customer Service (1-800-87 4-5278). |  | 3546 | 3546 lenses | Worldwide distribution: USA (nationwide) including states of: AL, AR, AZ, CA, CO, CT, DC, FL, GA, IL, IN, KY, LA, MA, MD, MI, MO, NC, NJ, NY, OH, OK, PA, SC, TN, TX, UT, VA, WA, WI and Puerto Rico; and countries of: Austria, Belgium, Canada, Costa Rica, Denmark, France, Germany, Hungary, India, Ireland, Israel, Italy, Korea, Japan, Luxembourg, Netherlands, Portugal, Norway, South Africa, Slovenia, Spain, Sweden, and United Kingdom. | Terminated | July 05, 2013 | 196 | 39876 | Not_Computer
Z-1331-2013 | 65078 | No | FYE | Surgical apparel. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | DePuy LifeLiner Stick &amp; Cut Resistant Gloves Rt, Lg |  |  | May 08, 2013 | 2013 | DePuy Orthopaedics, Inc. | LifeLiner Stick and Cut Resistant Gloves are being recalled because cases and individual glove packages were mispackaged, with right handed, large sized glove packaging containing left handed, large sized gloves. | Labeling mix-ups | DePuy Synthes Joint Reconstruction, a division of DePuy Orthopaedics, Inc.,sent a " Urgent Information Medical Device Recall Notice" dated May 8, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Please cease further distribution or use of the affected gloves immediately. Our records indicate your hospital purchased the affected lot. The purpose of this communication is to inform you of this recall and requestacknowledgement of receipt of this letter by signing and returning the Hospital/User Facility Reconciliation Form within 10 business days of receiving this notice.Response cards should be faxed to (574) 371-4964. Please contact your DePuy Synthes Joint Reconstruction Sales Representative for assistance returning any unused affected gloves.For product related questions, please contact your local DePuy Synthes Joint Reconstruction Sales Consultant. For clinical questions from surgeons, please contact DePuy Synthes Joint Reconstructions Scientific Information Office at 1-888-554-2482 (M-F; 8 a.m.- 5 p.m. EDT). For questions about device recall information provided, please call (574)371-4917 (M-F; 8 a.m. - 5 p.m. EDT). |  | 180 | 180 gloves (120 gloves distributed in the US) | Worldwide Distribution - US (nationwide) including the states of IL, OH, NC, SC, MN, NV, TN and WA., and the countries of UK, Netherlands and Canada. | Terminated | January 14, 2014 | 251 | 180 | Not_Computer
Z-1054-2013 | 64704 | No | KFX | External assembled lower limb prosthesis. | Physical Medicine | Physical Medicine | 510(K) Exempt | LeMaitre Stent Guide: 1102-00 (100 strips); 1102-50 (50 strips); 1102-20 (20 strips). |  |  | March 18, 2013 | 2013 | LeMaitre Vascular, Inc. | Devices were not sealed correctly during the manufacturing process, and the sterility of these products have been compromised. | Packaging process control | LeMaitre sent an "URGENT:VASCUTAPE RADIOPAQUE TAPE DEVICE FIELD SAFETY NOTICE" dated March 20, 2013 to affected customers. The recall letter contains a form that requests customers to return to LeMaitre Vascular, Inc. as a record of notification and reconciliation. The letter provides instructions on how the customer can inspect the defective tape seals and/or return the affected products for replacement/credit. |  | 6970 | 6970 | Worldwide Distribution-USA (nationwide) and the countries of Canada, EU, and Asia. | Terminated | December 03, 2014 | 625 | 20530 | Not_Computer
Z-0614-2014 | 67128 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | Catheter, Continuous Flush |  |  | December 06, 2013 | 2013 | Merit Medical Systems, Inc. | Merit Medical Systems, Inc. is conducting a recall for 3 lots of Maestro Microcatheters with a graphicdiscrepancy on the carton/box label end panel. The catheter graphic on the front of the cartonbox (correct) and the graphic on the end panel (incorrect) do not match. | Packaging process control | Merit Medical Systems sent an Urgent Product Recall Notice dated December 6, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately discontinue use of the Maestro Microcatheters identified and isolate them in a quarantined area until the product is over-labeled by their Merit Sales Representative Customers were instructed to ensure that all personnel to whom the devices were distributed were made aware of the field action. Customers with questions were instructed to contact their sales representative or call 1-801-826-4062.For questions regarding this recall call 801-253-1600. |  | 59 | 59 | Worldwide Distribution - USA (nationwide) and Internationally to Canada. | Terminated | January 27, 2014 | 52 | 59 | Not_Computer
Z-0403-2014 | 66815 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | STERRAD NX, Product Code: 10033 |  |  | November 11, 2013 | 2013 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) is recalling the STERRAD NX and STERRAD 100NX sterilization systems because ASP has determined that the bronchoscopes may contain a polyurethane material in the suction channel location that is not listed in the current claims for interior channels (lumens) for processing in STERRAD systems. | Labeling design | Advanced Sterilization Products sent an Urgent Medical Device Field Safety Notificaiton letter on November 11, 2013, to all affected customers. The letter informed customers of the problems identified and the actions to be taken. Customers were instructed to provide the notice to anyone in their facility that needs to be informed and post this notice where appropriate. Customers were instructed to contact ASP at 1-888-783-7723 for any suspected problems. ASP is collaborating with KARL STORZ Endoskope to determine the sterilization impact of the suction port lumen material. Customers are instructed to contact KARL STORZ Endoskope Technical Support at (800)421-0837, ext 5350 for alternative reprocessing instructions.For questions regarding this recall call 949-453-6400. |  | 4381 | Total 4,381 units | Nationwide Distribution | Terminated | January 22, 2014 | 72 | 4381 | Not_Computer
Z-0390-2014 | 66666 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | 3.7mm and 5.0mm Dynamic Locking Screw (DLS) |  |  | July 22, 2013 | 2013 | Synthes USA HQ, Inc. | All lots of 3.7mm and 5.0mm Dynamic Locking Screw (DLS) were recalled due to complaints of breakages at the distal tip of the pin identified after successful healing and during planned removal of the device. | Labeling False and Misleading | Synthes sent an Urgent Notice: Medical Device Recall letter to Sales Reps and End Users on July 22, 2013. The letter identified the affect product, problem and actions to be taken. For questions call the Recall center at 610-719-5450. |  | 40248 | 40248 | Worldwide Distribution - USA (nationwide) and internationally to Canada. | Open, Classified |  |  | 40248 | Not_Computer
Z-1053-2013 | 64528 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | LEEP REDIKIT Product No. 6060 |  |  | February 26, 2013 | 2013 | Cooper Surgical, Inc. | LEEP RediKit Products (for use in Loop Electrosurgical Excision Procedures) have Xylocaine (Lidocaine) HCL (2%) listed on the tray lid label but this is not included in the contents of the kit, which contained Polocaine (Mepivacaine) HCL (3%) instead (which was not listed on the tray label). | Labeling mix-ups | CooperSurgical sent a Recall Notification dated February 26, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.CSI intends to track and quarantine returned product in a segregated location at its manufacturing facility to ensure its separation from acceptable product and prevent its misuse. CSI intends to render used product useless by physical destruction and to rework unused product by correcting the labeling for redistribution.If you have any one of these affected lot numbers in your current inventory, please contact our Customer Service Department at 800-243-2974, or simply complete the attached form and fax to 800-262-0105. Arrangements will be made to replace product in your possession at CooperSurgical's expense. Further instruction and disposition of product will be communicated to you at that time.CooperSurgical deeply regrets the inconvenience that this problem has caused. |  | 4395 | 4395 kits | Nationwide Distribution including Puerto Rico | Terminated | June 12, 2014 | 471 | 4395 | Not_Computer
Z-0755-2013 | 63934 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Ortho Clinical Diagnostics VITROSÂż Chemistry Products Vapor Adsorption Cartridge REF/CAT No. 680 0100. VITROSÂż Vapor Adsorption Cartridges are used to remove vapor contaminants that may potentially interfere with Signal Reagent used with MicroWell Assays on the VITROSÂż 3600, VITROSÂż 5600, and VITROSÂż ECi/ECiQ Systems. |  |  | December 13, 2012 | 2012 | Ortho-Clinical Diagnostics | Some VITROSÂż Chemistry Products Vapor Adsorption Cartridge CAT No. 6800100 cartons incorrectly contain a Particulate Cartridge Part No. J19612. In some instances, the Particulate Cartridge has been inadvertently installed on the VITROSÂż System instead of the intended VITROSÂż Vapor Adsorption Cartridge which may bias results. | Packaging process control | Ortho-Clinical Diagnostics sent URGENT PRODUCT CORRECTION Letters (dated 12/13/12) along with Q&A; and Confirmation of Receipt Forms to distributors and consignees on 12/13/12. They were advised of the issue and instructed to discard their current inventory of VITROSÂż Vapor Adsorption Cartridges. On 12/13/12, foreign affiliates were informed by e-mail of the issue and instructed to notify their consignees of the issue and actions. |  | 5237 | Domestic: 15,891; Foreign: 5237 | Worldwide Distribution, including Nationwide (USA) and the countries of AU, GM, FR, IT, SP, BR, BM, AR, CA, JA, MX, UK, CH, CL, CO, IN, EC, PE, PY, PA, PT, NZ, SG, VE and UY. | Open, Classified |  |  | 5237 | Not_Computer
Z-0608-2014 | 66968 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | CPT Hip System, Parts 00-8114-002-10 and 00-8011-004-00 |  |  | December 10, 2013 | 2013 | Zimmer, Inc. | Zimmer is initiating a lot specific recall of the CPT Hip Stem, Versys Build-Up Block, and Trabecular Metal Reverse Glenosphere medical devices due to reports of the products being packaged in the incorrect outer carton. Product labels and patient labels are correct. | Labeling mix-ups | Zimmer sent an Urgent Medical Device Recall letter dated December 10, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the notification and ensure affected personnel were aware of the contents. If affected product was found, it should be quarantines and their Zimmer representative notified. The Zimmer sales representative will remove the recalled product from their facility. Customers with questions were instructed to call 1-877-946-2761. . |  | 36 | 36 | Worldwide Distribution -USA including OH, FL, WI, MI NJ, SC, CA, NC, PA, OR, WA, TN, KY, OK, KS, AZ, TX, CO, NV and Internationally to Germany.There have been 71 distribution events. Note that this product is lot controlled and may be distributed multiple times if it was returned and returned to the shelf. 7 units have been returned, including three through the complaint process, which leaves 64 affected units in the field subject to removal. | Terminated | September 11, 2014 | 275 | 64 | Not_Computer
Z-0404-2014 | 66581 | No | JDS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes 14.5 mm Outer Protection Sleeve for Suprapatellar |  |  | October 14, 2013 | 2013 | Synthes (USA) Products LLC | Complaints were received reporting that a product packaged as a 14.5 mm Outer Protection Sleeve for Suprapatellar contained a 12.0 mm Outer Protection Sleeve for Suprapatellar. | Labeling mix-ups | Synthes sent an Urgent Notice: Medical Device Recall letter dated October 14, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to call Synthes at 1-800-479-6329 if they have any of the affected product to obtain a Return Authorization Number, complete the Verification Section at the end of the letter and return the Verification Form with the affected product to:Credit/ReturnsSynthes1101 Synthes AvenueMonument, CO 80132If customers do not have the affected product they should complete the attached Verification Section and return all documents to Synthes by:Fax: 866-229-7778Scan/email: FieldAction@synthes.comCustomers with questions were instructed to call 610-719-5450.For questions regarding this recall call 610-719-5000. |  | 55 | 55 | Nationwide Distribution | Terminated | August 18, 2015 | 673 | 55 | Not_Computer
Z-1344-2013 | 49026 | No | JHI | Human chorionic gonadotropin (HCG) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Cardinal Health hCG Cassette Rapid Test, Part Number: B1077-22 (Mfg PN: 4582105026) |  |  | January 21, 2013 | 2013 | Alere San Diego, Inc. | Alere San Diego Inc. is recalling the Cardinal Health hCG Cassette Rapid Test because some of them were incorrectly packaged into Cardinal Health hCG Combo Rapid test kit boxes. The kit box incorrectly indicates the product may be used for both urine and serum sample types. | Labeling mix-ups | The firm, Alere, sent an "URGENT MEDICAL DEVICE RECALL" letter dated January 21, 2013 to all their customers. The letter describes the product, problems identified and the actions to be taken. Customers are instructed to discontinue use and/or sale of the product; immediately discard all units of the the product in accordance with your local regulations; share this information with your laboratory staff and retain this letter for your records; if you have forwarded the product to another laboratory, provide a copy of this letter to them and include the name, address and a contact for the laboratory on the enclosed Customer Verification Form; complete and return the Customer Verification Form within 10 days via fax to: Technical Service at 858-805-8457 or e-mail to responses.ts@alere.com. Customers with any questions are instructed to contact Alere Technical Services at (866) 216-0094 or by email at lateral.flow.support@alere.com. |  | 2167 | 2,167 kits | Nationwide distribution: USA including states of: AL, AK, AR, AZ, CA, CO, CT, DE, HI, IA, ID, IL, IN, FL, GA, KY, KS, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, UT, TN, TX, VA, WA, WI, WV AND WY. | Terminated | August 28, 2013 | 219 | 2167 | Not_Computer
Z-0379-2014 | 66663 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Product Brand Name(s): Alere Cholestech LDXÂż Lipid Profile"GLU Test Cassette, Model 10-991. |  |  | October 23, 2013 | 2013 | Alere San Diego, Inc. | Alere initiated a recall of Cholestech LDX Lipid Profile-GLU test cassettes, Catalog Number 10-991, Lot # 320167, becuase an incorrect lot number was printed on the outer 10 pack kit box label. The incorrect Lot # number printed was 321067. | Labeling Change Control | Alere initiated this recall by sending out notification letters to customers via fax, email, or direct mail. The letter titled "URGENT MEDICAL DEVICE NOTICE", dated October 23, 2013 is accompanied by a response form/Verification form. The notification informed customers of the reason for recall, product description with codes, required actions by the customers, and contact information. |  | 3621 | 3,621 kits | Worldwide Distribution: USA (nationwide) and Internationally to Canada. | Terminated | September 25, 2014 | 337 | 3621 | Not_Computer
Z-1778-2013 | 65488 | No | FTC | Ultraviolet lamp for dermatologic disorders. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Dermalight 80 Model Number UVB-080, one unit per box. |  |  | April 09, 2013 | 2013 | National Biological Corp | The incorrect manual may have been packaged with this product. UVB-080 devices may have included the manual for a UVA-083 device. | Labeling mix-ups | The firm, National Biological, sent an "VOLUNTARY PRODUCT RECALL NOTIFICATION" letter dated May 10, 2013 to its consignees/customers by UPS or other delivery service with a signature required. The letter described the product, problem and actions to be taken. Consignees/customers who do not respond to the letter will be contacted via telephone. The customers were instructed to check within your household to determine if you have any of the listed product; immediately destroy the incorrect User Manual, it is not necessary to return the product; record the Customer Name and or the Dermalight-80 serial number and complete and return the enclosed Verification Form, even if you do not have any User Manuals to destroy.Please contact your National Biological Corporation sales representative if you have any questions regarding this recall, any of our products, or would like assistance with the recall call 216-831-0600 or 800-338-5045 email: infor@PhototherapyExperts.com. |  | 124 | 124 units | Worldwide distribution: US (nationwide) including states of: AL, CA, CO, CT, FL, GA, IA, ID, IL, LA, MA, MN, MO, MS, NE, NJ, NY, OH, PA, SC, TN, TX, VA, WA, and WI; and countries of: Canada, Brazil, Malaysia, and Mexico. | Terminated | February 24, 2014 | 321 | 124 | Not_Computer
Z-1181-2013 | 64833 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | VitreaÂż CT Brain Perfusion made available in various configurations under the names Vitrea, Vitrea fX, VitreaAdvanced and Vitrea Enterprise Suite. |  |  | March 15, 2013 | 2013 | Vital Images, Inc. | Vital Images, Inc. (Vital Images) received one report about potential user confusion when viewing a brain perfusion scan on Vitrea CT Brain Perfusion 2D when the scan was performed using a wide coverage shuttle acquisition technique. This may result in patient treatment delays, misdiagnosis or mistreatment. | Labeling design | Consignees were sent on 3/15/2013 a VITAL Images, Inc., "Urgent Field Safety Notice" dated March 15, 2013. The letter described the problem and the product involved in the recall. The letter listed the limitations of the product and the steps to take. Consignees were requested to return the Product Recall Effectiveness Check form. Questions should be directed to Vital Images Customer Support at 1-800-208-3005 or via e-mail, support@vitalimages.com. |  | 1978 | 1,978 | Worldwide Distribution -- USA, excluding RI and SD and including DC and PR, and the countries of ALGERIA, ANGOLA, ARGENTINA, AUSTRALIA, AUSTRIA, BELGIUM, BELARUS, BOLIVIA, BRAZIL, BULGARIA, CAMBODIA, CANADA, CHILE, CHINA, COLOMBIA, CYPRUS, CZECH REPUBLIC, DENMARK, EGYPT, FINLAND, FRANCE, GEORGIA, GERMANY, CHANA, GREECE, GUAM, HONDURAS, HONG KONG, ICELAND, INDIA, INDONESIA, IRAQ, IRELAND, ISRAEL, ITALY, JAPAN,JORDAN, KOREA, KUWAIT, LEBANON, LITHUANIA, MALAYSIA, MEXICO, MOSCOW, MYANMAR, NAMIBIA, NETHERLANDS, NEW ZEALAND, NIGERIA, NORTHERN IRELAND, NORWAY, OMAN, PAKISTAN, PANAMA, PERU, PHILIPPINES, POLAND, PORTUGAL, QATAR, ROMANIA, RUSSIA, SAUDI ARABIA, SERBIA, SERBIA AND MONTENEGRO, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SRI LANKA, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TUNISIA, TURKEY, UKRAINE, UNITED ARAB EMIRATES, UNITED KINGDOM, UZBEKISTAN, VENEZUELA, VIETNAM, and YEMEN. | Terminated | April 08, 2015 | 754 | 1978 | Not_Computer
Z-1045-2013 | 64569 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | Zimmer NexPosure, retractor portals of the NexPosure MIS Access System, Part Numbers or REF numbers: N1830R80-100, N1830R80-090Âż, N1830R80-080Âż, N1830R80-070Âż, N1830R80-060, N1830R50-100, N1830R50-090, N1830R50-080, Âż N1830R50-070, N1830R50-060, N1830L80-100, N1830L80-090, N1830L80-080, N1830L80-070, N1830L80-060, N1830L50-100, N1830L50-090, N1830L50-080, N1830L50-070, N1830L50-060, Âż N182450-085, N182450-075, N182450-065, N182450-055, N182450-045.Âż Retractor access system that allows placement of posterior spinal fixation implants. |  |  | February 08, 2013 | 2013 | Zimmer, Inc. | Possibility that the inner pouch may be compromised. The product is packaged in two pouches - a sterile inner pouch within an external pouch. NexPosure Retractors contained in a breached inner pouch have an increased risk of contamination (i.e.; loss of product sterility) if proper aseptic technique is not followed when transferring the inner package into the sterile field. Patient infection may o | Package design/selection | Consignees were sent on 2/8/13 a Zimmer Spine "Urgent Medical Device Recall" letter dated February 08, 2013. The letter was sent to Surgeons and Facilities using the NexPosure MIS Access System. The letter described the product involved in the recall and the problem. Advised consignees to discontinue the use of the product and to contact their sales representative for removal.A consignee letter dated February 8, 2013 was also sent to Distributors, Sales Representatives, and Distribution Operation Managers Distributing the NexPosureÂż MIS Access System. The letter contained instructions if they had the affected product or not and the steps to follow. |  | 58 | 58 | Nationwide Distribution including the states of AZ, CO, DE, FL, ID, LMA, MI, NH, NJ, NM, PA, RI, TX, and VA. | Terminated | December 09, 2014 | 669 | 58 | Not_Computer
Z-1093-2013 | 64776 | No | LRR |  |  | General &amp; Plastic Surgery | Enforcement Discretion | LHB, Perishable Items Pack Type IV. Kit number: 2114016-100. |  |  | March 22, 2013 | 2013 | Lighthouse For The Blind | The product contains an iodine 10% Ampule (10% Povidone Iodine swab) which is being recalled by the firm's distributor. The labeling of the iodine 10% Ampule labels state that the contents of the package are sterile and does not clarify that this sterility claim is only applicable to the applicator. The antiseptic solution that is applied to the skin is a cutaneous disinfectant, and is not sterili | Labeling False and Misleading | The firm sent "URGENT DEVICE RECALL" letters dated March 22, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at 800-542-3697 for questions regarding this notice. |  | 425 | 425 kits | Nationwide Distribution including DC and the states of AK, AL, AR, AZ, CA, CO, DE, FL, GA, ID, IL, IN, KS, KY, MI, MD, MN, MO, MT, NC, NJ, NM, NY, NV, OH, OK, PA, SC, TN, TX, UT, VA, WA and WY. | Terminated | November 25, 2013 | 248 | 108939 | Not_Computer
Z-1922-2013 | 65781 | No | FPD | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Nestle Nutrition, Compat, Enteral Feeding Tube, sterile, 8 French (REF 08310400), 10 French (REF 08310600), and 12 French (REF 08310500). |  |  | July 24, 2013 | 2013 | Nestle HealthCare Nutrition | Nestle HealthCare Nutrition, Inc. is initiating a recall on some of Compat NG Tubes because they have been delivered in an insufficiently sealed pouch, thus compromising the sterility claimed on the label. | Packaging process control | Nestle Health Science sent a "Medical Device Recall" letter dated July 24, 2013. The letter was addressed to Risk Manger. The letter described the product involved, problem and actions to be taken. Customers were advised that products currently in use should not be returned. Customers were requested to examine their inventory, quarantine the product, carry the recall to the user facility level and to complete and return the Acknowledgment / Response Form.For questions call Nestle HealthCare Nutrition Customer Service at 1-877-463-7853. |  | 37310 | 37,310 | US Nationwide Distribution - including the states of: AK, AZ, CA, CO, CT, FL, GA, HI, IA, IL, IN, KS, KY, MA, MD, MI, MN, MO, NC, NE, NJ, NM, NV, NY, OH, OK, PA, RI, SC, TN, TX, UT, VA, WA. | Terminated | June 13, 2014 | 324 | 37310 | Not_Computer
Z-0577-2014 | 66935 | [6 empty] | ProPlan CMF Guide(s) Model(s) |  |  | November 08, 2013 | 2013 | Materialise USA LLC | Patient Specific Guide contained the incorrect case report. The case report contains patient specific surgical plan information. | Labeling Change Control | Customer/ ordering physician was made aware of the issue11/08/2013. Revised accurate case report was immediately provided. |  | 1 | 1 | Distributed in Canada. | Terminated | January 07, 2014 | 60 | 1 | Not_Computer
Z-1748-2013 | 65303 | No | OOY |  |  | Anesthesiology | PMA | Alair Bronchial Thermoplasty Catheter, Bronchial Thermoplasty System; Material Number: M005ATS25010, Catalog Number: ATS 2-5. Product is manufactured and distributed by Boston Scientific Corporation (BSC), Sunnyvale, CA |  |  | May 28, 2013 | 2013 | Boston Scientific Corporation | Boston Scientific has discovered an inconsistent "Use By" expiration date on a single lot/batch of Alair Bronchial Thermoplasty Catheters. The expiration date indicated on the inner tray, 2012-04, is incorrect. The correct expiration date is 2014-04 as indicated on the outer box label. | Error in labeling | The firm, Boston Scientific, sent a letter dated May 28, 2013, via express mail to all its direct consignees. The letter described the product, problem and actions to be taken. The customers were informed that you can continue to use the device until the expiration date on the outer box lablel, 2014-04; No product is being recalled and the you are not required to return product to Boston Scientific. The customers were instructed to read the enclosed information carefully and post the information in a conspicuous location near the affected product and/or take such other steps, consistent with your internal policies and procedures, to ensure that the information is easily accessible to all users of this device; and to complete and return the Customer Acknowledgment Form even if you no longer have any units from the affected lot/batch via email: MapleGroveFieldAc@bsci.com or Fax to: Field Action Center - 1866-213-1806. If you have any questions, call 763-494-1133 or email: MapleGroveFieldAc@bsci.com . |  | 35 | 35 units | Worldwide distribution: US (nationwide) including states of: CA, FL, IA, ID, MI, OK, TA and WA; and countries of: Canada and Thailand. | Terminated | August 28, 2013 | 92 | 35 | Not_Computer
Z-1888-2013 | 65819 | No | MGO | Hydrophilic wound dressing. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | CVS Pharmacy Wound Closure Adhesive Surgical Tape Strips |  |  | July 26, 2013 | 2013 | ASO, LLC | This recall has been initiated by Aspen Surgical Products, Inc., due to the inability to confirm sterility. | Package design/selection | ASO sent a letter on July 26, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately examine your inventory and stop using/distributing the product. Place in quarantine to be returned or destroyed. ASO customers will receive credit for returned product or a returned completed Certificate of Destruction. Please use Return Authorization Number 1675 when returning the product.Please complete and return the enclosed response form as soon as possible. Return the form to:Address:ASO LLCRegulatory Affairs300 Sarasota Center Blvd.Sarasota, FL 34240Fax:941-378-9040Attn: Regulatory AffairsFor questions regarding this recall call 888-364-7004, ext 146. |  | 2689560 | Total 2,689,560 pouches (2 cards/pouch) | Nationwide Distribution including FL, IA, IL, MI, NC, NY, OH, OR, PA, RI, TX, WA and WI. | Terminated | August 04, 2014 | 374 | 2689560 | Not_Computer
Z-1377-2013 | 64886 | No | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | ***REF S-3D*** 3mm Diamond Ball. |  |  | March 04, 2011 | 2011 | The Anspach Effort, Inc. | Due to three batches of product may have diamond grit size larger than intended applied to the head of the cutting burr. | Labeling mix-ups | Anspach sent an URGENT: Medical Device Correction letter dated March 4, 2011 to its customers identifying the affected product and actions to be taken by the customers. Customers questions or concerns should be directed to Anspach Customer Support at (800) 327-6887. |  | 1827 | 1,827 burrs | Worldwide Distribution-USA (nationwide) including DC and the states of OH, TX, IL, CA, GA, KY, FL, WI, NC, VA, WI, MN, MD, KS, OK, PA, OR ,HI, AZ, WA, TN, AR, IA, MI, CO, LA, NJ, IN, WV, MO, AL, NY, VT, and SC and the countries of France, Canada, Great Britain, Mexico, Australia, Sweden and Netherlands. | Terminated | December 17, 2013 | 1019 | 1827 | Not_Computer
Z-1378-2013 | 64588 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Femoral component, CR, cemented, #1, left2103-1310 |  |  | February 20, 2013 | 2013 | Orthopedic Alliance LLC | The recall of the Loaner Bin and Sterilization Trays for implants and instruments was initiated by Orthopedic Alliance because the directions for use and surgical techniques were not being provided to every physicianand for every surgery. | Labeling design | A recall letter dated 3/8/13 was sent to all customers who purchased loaner bins and sterilization trays for implants and instruments from Orthopedic Alliance. The letter informed the customers of the problem identifiedand the actions to be taken. Customers are instructed to call the office at the number (951) 304-9001 if they are requesting paper copies of any of the literature. |  | 136 | 136 units | Nationwide Distribution including the states of CA, MD, TX, NV, and WI. | Terminated | March 26, 2014 | 399 | 821 | Not_Computer
Z-1206-2013 | 64736 | No | OFP | Anesthesia breathing circuit. | Anesthesiology | General Hospital | Enforcement Discretion | Presource PBDS, Knee Arthroscopy, Kit, Circulator |  |  | March 26, 2013 | 2013 | Cardinal Health, Medical Products &amp; Services | Various Presource Kits contain a pre-assembled Filter and Anesthesia Circuit that may contain outer plastic packaging on one or more components. If the packaging is removed without disassembling the components, remnants of the plastic from the packaging material may become lodged in the filter potentially causing an obstruction of airflow. | Packaging process control | Affected consignees were notified with a letter, dated March 26, 2013, sent via FedEx Priority, of this action. Cardinal Sales Representatives will also be advised and will verbally communicate the action to affected customers. The letter identified the affected product and the problem. Customers are to locate any affected product and quarantine them. In addition, clinicians should be informed of the issue; and, warning labels should be affixed to the front of each kit instructing users to remove and discard the Anesthesia Circuit and Filter Assembly. In order to confirm receipt of this action and completion of labeling the product, the enclosed Acknowledgement Form should be completed and returned to the number provided. If customers do now wish to accept overlabeled product or to overlabel the product in their possession, they are to contact Presource Sales Operations at 800-766-0706, or their sales representative for further instruction. Customers are to contact Cardinal Health Quality Systems at 800-292-9332, if they have any questions or need special assistance. |  | 4571 | 4,571 Kits total | Distributed in the states of AR, MS, SC, and TN. | Terminated | April 12, 2017 | 1478 | 4571 | Not_Computer
Z-0573-2014 | 66914 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Indus Invue Screws: IM71058-XX: 04.2mm, SelfTapping, Tapered. |  |  | November 20, 2013 | 2013 | SpineFrontier, Inc. | Mismarked and unmarked screws | Packaging process control | On 11/20/2013, consignees were notified of the email via letter. Consignees were instructed to return the screws listed in the recall notification letter and return the Recall Acknowledgement form to SpineFrontier via fax or email. |  | 993 | 993 | Nationwide Distribution- including DC and the states of OH, TX, KS, MA, IN, CO, NY, AL, and MI. | Open, Classified |  |  | 1915 | Not_Computer
Z-1305-2013 | 55358 | No | FJL | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Electrode, cutting, 24FR (27050G/6). |  |  | June 01, 2009 | 2009 | Karl Storz Endoscopy America Inc | Possible damage to the sterile packaging (blister tray) of some Karl Storz sterile products, which may result in airborne contamination. | Packaging | A Safety Alert dated May 29, 2009 was sent to affected customers. The letter informed them of possible damage to the sterile packaging (blister tray) of some Karl Storz sterile products which may result in airborne contamination. Karl Storz requested consignees to inspect their inventory of the subject products thoroughly and look for small holes in the blister trays. Any identified units should be returned to Karl Storz, and a replacement package (6 units per package) will be provided to at no charge. The Karl Storz Customer Support number 1-800-421-0837 ext. 7577 was provided to obtain a Return Authorization Number (RMA) and arrange replacement. Questions and comments were forwarded to Technical Support at 1-800-421-0837 ext. 5350. |  | 9432 | 9432 units | Distributed Nationwide and in Canada. | Terminated | May 13, 2013 | 1442 | 9432 | Not_Computer
Z-1366-2013 | 61315 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Pinnacle Hip Solutions ALTRX Polyethylene Acetabular Liner, Ref 1221-36-452 |  |  | March 07, 2012 | 2012 | DePuy Orthopaedics, Inc. | DePuy Orthopaedics, Inc. (DePuy) is issuing a voluntary recall of one lot of the AltrX" liner because the product was incorrectly labeled on the box. The box was labeled for a 36IDX52OD +4 Neutral Liner, but the product enclosed was a 36IDX56OD Neutral Liner. Three ( 3 ) complaints were received in January 2012. | Packaging process control | DePuy Orthopaedics sent an URGENT INFORMATION- MEDICAL DEVICE RECALL NOTICE dated March 7, 2012 via email to all affected distributors and customers. The letter identified the affected product, problem and actions to be taken. Distributors and customers are instructed to cease further distribution or use or of recalled product and to return them to DePuy. The sales representatives are expected to aid customer in returning of recalled product. Distributors and sales representatives are to return the recalled product through the normal DePuy Returns process. Reconciliation form should be completed and returned to your DePuy Sales Representative or faxed to 574-372-7567. For questions about recall information contact the Manager of Customer Quality 574-372-7333. |  | 20 | 20 liners | Worldwide Distribution - USA Nationwide including the state of VA and the country of Ireland. | Terminated | May 24, 2013 | 443 | 20 | Not_Computer
Z-0606-2014 | 66902 | No | CAI | Anesthesia breathing circuit. | Anesthesiology | Anesthesiology | 510(K) Exempt | RTC 25-C Inline Aerosol Tee Adapter. |  |  | November 13, 2013 | 2013 | Instrumentation Industries Inc | Some packages of the RTC 25-C may have been shipped to customers with incomplete Installation and Usage Directions due to a printing error. | Labeling Change Control | A recall notification letter, dated November 13, 2013, was sent to Users. |  | 220 | 220 | Nationwide Distribution. | Terminated | July 09, 2014 | 238 | 220 | Not_Computer
Z-1218-2013 | 64682 | No | GCX | Vacuum-powered body fluid suction apparatus. | General Hospital | General Hospital | 510(K) Exempt | Stryker Irrigator Reusable Tips: |  |  | March 27, 2013 | 2013 | Stryker Endoscopy | The pre-vacuum steam sterilization (wrapped method) temperature in the Instructions for Use incorrectly listed 121-123ÂżC , instead of the correct 132-133ÂżC. | Error in labeling | Stryker sent "Urgent Device Correction" letters dated March 27, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. |  | 147921 | 147,921 - all sizes | Worldwide Distribution: USA (nationwide) and countries of: Mexico, Portugal, Latin America, India, Greece, Taiwan, Singapore, New Zealand, Phillipines, Netherlands, France, Australia, Canada, Japan, Scandinavia, Iberia, Brasil, Romania, South Africa, Chile, Korea, China. | Terminated | September 24, 2013 | 181 | 147921 | Not_Computer
Z-1781-2013 | 65542 | No | GXH | Cryogenic surgical device. | Neurology | Neurology | 510(k) | Myoscience brand iovera Smart Tip, Myoscience Cryo-Touch IV, a Cryogenic Surgical device, REF Catalog Number: STA0311-10. |  |  | June 21, 2013 | 2013 | Myoscience Inc | The outer box label of the product was labeled with a down revision label which did not include the US labeling requirements, specifically the symbol descriptors and also the "Caution: Federal Law (USA) restricts this device to sale by or on the order of a physician." This product was previously labeled for EU distribution only. | Labeling mix-ups | Myoscience representatives visited both affected sites on June 14, 2013 and relabeled the product in person. Neither site had used the product and was unaffected by the relabeling. No formal communication was issued. |  | 3 | 3 units | US distribution including CA and ID. | Terminated | July 15, 2014 | 389 | 3 | Not_Computer
Z-1876-2013 | 65757 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Mullins Transseptal Adult 7F, REF 008551. |  |  | July 08, 2013 | 2013 | Medtronic Inc. Cardiac Rhythm Disease Management | Medtronic identified a specific labeling issue for specific lots of Mullins Transseptal Catheter Introducer Sheath, model 008551, in which packages labeled as Mullins 7F (French) may actually contain an 8F(French) dilator. Medtronic is requesting the return of potentially mis-labeled Mullins Transseptal Catheter Introducer 7 French Sheaths. Due to the differences in size, an 8F dilator will not | Packaging change control | Consignees were sent on 7/8/2013 a Medtronic, Inc " Urgent Medical Device Recall" letter dated 8 July 2013. The letter was addressed to Healthcare Professionals. The letter identified the problem and the product involved in the recall. Advised consignees to remove and quarantine the product from their inventory and return the product. Requested consignees to complete the attached Customer Confirmation Certificate and fax it to Medtronic at 651-367-0612. If the product has been further distributed advised consignees to please notify them of the recall. |  | 269 | 269 | Worldwide Distribution-USA (nationwide) including DC and the states of AZ, AK, CA, DE, GA, IL, KY, MA, MI, MN, MS, NV, NY, NC, OH, PA, TN, and TX, and the countries of AUSTRALIA, BRAZIL, CHILE, INDIA, and SINGAPORE. | Terminated | March 06, 2014 | 241 | 269 | Not_Computer
Z-1779-2013 | 65494 | No | FTC | Ultraviolet lamp for dermatologic disorders. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Dermalight 80 Model Number UVB-083, one unit per box. |  |  | April 08, 2013 | 2013 | National Biological Corp | The incorrect manual may have been packaged with this product. UVB-083 devices may have included themanual for a UVA-080 device. | Labeling mix-ups | The firm, National Biological, sent an "VOLUNTARY PRODUCT RECALL NOTIFICATION" letter dated May 10, 2013 to its consignees/customers by UPS or other delivery service with a signature required. The letter described the product, problem and actions to be taken. Consignees/customers who do not respond to the letter will be contacted via telephone. The customers were instructed to check within your household to determine if you have any of the listed product; immediately destroy the incorrect User Manual, it is not necessary to return the product; record the Customer Name and or the Dermalight-80 serial number and complete and return the enclosed Verification Form, even if you do not have any User Manuals to destroy.Please contact your National Biological Corporation sales representative if you have any questions regarding this recall, any of our products, or would like assistance with the recall call 216-831-0600 or 800-338-5045 email: infor@PhototherapyExperts.com. |  | 6 | 6 units | US distribution including the states of: AL, CA, MN, and NC. | Terminated | February 24, 2014 | 322 | 6 | Not_Computer
Z-1772-2013 | 65504 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Endopath Electrosurgery Probe Plus II, twelve (12) different devices, including hooks, spatulas, right angles, curved dissectors, suction/irrigation devices, accessory ports, pool/sump devices and stone retrieval cannula, packaged and sterilized individually in a transparent flexible blister sealed to a Tyvek lid. 6 flexible packages are placed into a sales unit carton, with 4 cartons per corrugated shipper. |  |  | April 08, 2013 | 2013 | Ethicon Endo-Surgery Inc | The firm is recalling the ENDOPATH Probe Plus II Shafts because the possibility that a tear in the Tyvek packaging may affect the sterility of the device. | Packaging | The firm, Ethicon Endo-Surgery, Inc., sent an "URGENT: MEDICAL DEVICE RECALL" letter dated April 22, 2013 to all consignees/customers via overnight UPS to allow tracking of the receipt. The letter described the product, problem and actions to be taken. The consignees/customers were instructed to do the following: EFFECTIVE IMMEDIATELY-DO NOT USE ANY OF THE FOLLOWING TWELVE (12) PRODUCT CODES within the noted Expiration Dates provide in the letter; examine your inventory immediately to determine if you have affected product on hand; remove the affected product; complete and return the Business Reply Form (BRF) and fax to: 1-888-912-7089 or email to: ees3118@stericycle.com within 3 business days, even if you do not have affected product; return affected product with a photocopy of the completed BRF; and share this information with all of the appropriate staff at your facility. If you need additional shipping labels or a communications package, contact the Stericycle's support center at 1-888-943-4896 and reference Event 3118."Returned ENDOPATH Probe Plus II Shafts will be replaced with corrected ENDOPATH Probe Plus II Shafts once supply becomes available. If replacement product is not available and/or has been discontinued, a credit will be issue."If you need clinical or product support, please contact your local representative or call our Customer Support Center, 1-877-ETHICON (877-384-4266). |  | 327648 | 327,648 instruments | Worldwide distribution: US (nationwide) including states of: AK, AL, AR, AZ, CA, CO, CT, FL, GA, lA, ID, IL, IN, KS, KY, LA, MA, MD, ME, Ml, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, &amp; WY and countries of: Canada, Argentina, Australia, Austria, Bahrain, Belgium, Brazil, Canada, Chile, China, Columbia, Czech Republic, Dominican Republic, Ecuador, Egypt, India, Indonesia, Israel, Italy, Jamaica, Japan, Korea, Kuwait, Libya, Malaysia, Mexico, Pakistan, Panama, Philippines, Poland, Puerto Rico, Russia, Singapore, Slovenia, South Africa, Spain, Saudi Arabia, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Uruguay, &amp; Venezuela. | Terminated | December 24, 2013 | 260 | 327648 | Not_Computer
Z-0183-2014 | 66610 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Applied Medical's InziiÂż 12/15mm Retrieval System. |  |  | October 21, 2013 | 2013 | Applied Medical Resources Corp | Applied Medical is conducting a voluntary recall of the InziiÂż 12/15mm retrieval system. During shipment, the retrieval system packaging has the potential to become punctured with small holes, which could compromise the sterile barrier. The likelihood of this situation to occur and result in patient harm is highly unlikely; however, out of an abundance of caution for patient safety and a commitmen | Package design/selection | Customer notifications of the recall began the week of 10/21/2013 via UPS overnight with letters requesting that all affected products be returned to Applied. The notification letter is dated October 21, 2013 and titled "URGENT: MEDICAL DEVICE RECALL". The notification was accompanied by a Recall Notification Confirmation Form. The letter informed customers of the recall by providing the product description with codes (Model and Lot Numbers). reason for the recall, actions to take and contact information. |  | 29506 | 29,506 units | Worldwide Distribution: US (nationwide) and countries of: Australia, Austria, Belgium, Brazil, Canada, Chile, Colombia, Czech Republic, Denmark, Ecuador, Finland, France, Germany, Hong Kong, India, Ireland, Israel, Italy, Japan, Lebanon, Malaysia, Mexico, New Zealand, Netherlands, Norway, Peru, Poland, Portugal, Saudi Arabia, Singapore, Spain, South Africa, Sweden, Switzerland, United Arab Emirates, and United Kingdom. | Terminated | September 15, 2014 | 329 | 29506 | Not_Computer
Z-0593-2014 | 67052 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart MRx Monitor/Defibrillator |  |  | June 07, 2013 | 2013 | Philips Medical Systems, Inc. | Error in the printing of the therapy energy setting labels on Philips HeartStart M3536A MRx Monitor/Defibrillators. The first energy setting is labeled 10 when it should be labeled 1-10 | Labeling Change Control | Philips sent a Customer Information Letter dated June 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.During an EI of the firm, the firm had not reported the correction to FDA.The Customer Information letter informs the customers what the behavior is and under what circumstances. The correction consisted of a replacement label that was provided free of charge to all units affected by this issue. Philips Healthcare representatives have been contact customers with affected devices to arrange for installation of the replacement label.To correct this issue, Philips will replace the label free of charge. A Philips Healthcare representative will contact customers with affected devices to arrange for installation of the replacement label. This issue does not affect current product shipments.If you need any further information or support concerning this issue, please contact your local Philips representative at 1-800-722-9377.Philips apologizes for any inconvenience caused by this problem. |  | 93 | 93 units | Worldwide Distribution - US Distribution including the states of IN, MS, TX and VA., and the countries of DENMARK, FRANCE, GERMANY, SPAIN and SWITZERLAND. | Terminated | March 22, 2017 | 1384 | 93 | Not_Computer
Z-0380-2014 | 66411 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | 2.4 and 2.7 MM Locking Screw SLF-TPNG with T8 Stardrive Recess 12 MM |  |  | July 31, 2013 | 2013 | Synthes USA HQ, Inc. | Specific lots of the SS 2.4mm and SS 2.7mm Locking Screw (self-tapping) due to a mislabeling issue. The SS 2.7mm Screw was labeled as SS 2.4 mm Screw. | Labeling mix-ups | Synthes sent an URGENT NOTICE: MEDICAL DEVICE RECALL letter dated July 31, 2013, to all affected customer. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to call Synthes at 1-800-479-6329 if they have any of the affected product to obtain a Return Authorization Nuber, complete the Verification Section at the end of the letter by checking the appropriate box indicating affected product has been located, also indicate the number of devices found and note the Return Authorization Number. Fax the documents to Synthes at 1-866-792-6446 or email to FieldAction@synthes.com. Customers with questions were instructed to call 610-719-5450 or contact their Synthes Sales Consultant.For questions regarding this recall call 610-719-5000. |  | 98 | 98 | Nationwide Distribution including PA, CA, FL, MO, KS, IN, IL, WV, KY, LA, NM, SC, IA, TX, AZ, GA, NC, OK, NY, TN, MI, CO, and MT. | Open, Classified |  |  | 98 | Not_Computer
Z-0148-2014 | 66339 | Yes | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius LibertyÂż Cycler |  |  | October 14, 2013 | 2013 | Fresenius Medical Care Holdings, Inc. | Fluid leaking may be observed inside the pump compartment of the Liberty Cycler and may result in dialysate becoming contaminated, potentially causing peritonitis | Device Design | Fresenius Medical Care contacted all affected consignees by telephone on 10/14/13 to inform them of the product safety alert. Follow-up Urgent Safety Alert letter was sent on 10/18/13 and 10/23/13 by Certified Mail, Return Receipt requested. The letter identified the affected product, problem and actions to be taken. Customers were instructed to report any events associated with this safety alert to Customer Service Team at 1-800-323-5188. |  | 16600 | 16,600 | USA Nationwide Distribution | Open, Classified |  |  | 16600 | Computer
Z-0618-2014 | 66468 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | S3 Med/Surg Bed |  |  | November 26, 2013 | 2013 | Stryker Medical Division of Stryker Corporation | Stryker Medical has identified a potential issue associated with Secure II and S3 Med/Surg beds built between April 2nd, 2005 and July 26th, 2011. The affected beds have the potential to experience damage to the Foot End Cover and/or Power Coil Cable due to reduced clearance between the Foot End Cover and Foot End Lift Header. | Device Design | Stryker Medical sent an Urgent Medical Device Recall Notification letter on November 26, 2013, to all affected customers. The notification included the products being recalled, the product description, product issue, potential hazards, and actions needed. Customers were instructed to locate and inspect the beds listed in the notice and ensure that no damage has occurred to the foot-end cover, power coil cable, or sensor coil cable. If the beds were not damaged customers may continue to use them. If damage occurred, customers wsere instructed to remove the beds from service and contact Stryker at 1-800-327-0770, Option 2, then Option 2 to arrange a service call. Continue to inspect the foot-end cover and coil cables as part of their routine preventive maintenance. Customers were instructed to fax the enclosed business reply form to 269-488-8691 or email productfieldaction@stryker.com. Customers were instructed to forward the letter if the product was loaned or sold to other customers. If any beds were disposed of and were no longer in use, customers should advised Stryker by providing them with the serial number. Questions or concerns are directed to Stryker Customer Service at 1-800-327-0770. Consignees are to fax (269-488-8691) or email (productfieldaction@stryker.com) the business reply form to confirm receipt of this notification.For questions regarding this recall call 269-324-6609. |  | 108769 | 108,769 Total | Worldwide Distribution - USA (nationwide) and Internationally to Canada, Chile, China, Argentina, Brazil, Germany, Korea, Mexico, Singapore, UK, Europe, Middle East, Africa, and Latin America. | Terminated | May 29, 2015 | 549 | 108769 | Computer
Z-0872-2013 | 62334 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Liko Overhead Rail System, Ceiling Mounted Rail Systems for all Liko Stationary Lifts including Likorall, Multirall and Masterlift Systems |  |  | October 31, 2012 | 2012 | Hill-Rom, Inc. | Liko/Hill-Rom to date has received two reports from facilities alleging that the overhead rail system has failed to the extent that the overhead rails have fallen down while in use in high humidity (swimming pool) areas. Analysis of the incidents suggest that the components (hardware) used to secure the rail to the ceiling had corroded and may be unsuitable for environments with an atmosphere con | Device Design | Liko, a Hill-Rom company, sent an URGENT MEDICAL DEVICE RECALL notice dated October 31, 2012 to all affected customers. Letter identified the affected product, problem and actions to be taken until a new design is finalized and the field corrective action implemented. The letter instructed customers to complete and return the Response Form immediately even if you do not have any rail systems in a chlorinated environment. For questions concerning this request or the procedure as outlined, please contact Hill-Rom Technical Support at 800-445-3720. |  | 3816 | 3816 in total | Worldwide Distribution: USA (nationwide) including states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, and WY; and CANADA LOCATIONS: AB, BC, MB, NB, NS, ON, and SK. | Terminated | March 09, 2016 | 1225 | 3816 | Computer
Z-0616-2014 | 67002 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | SYSTEM 1E Liquid Chemical Sterilant Processing System, Distributed by STERIS Corporation, Mentor, OH |  |  | November 21, 2013 | 2013 | Steris Corporation | Customers indicated fluctuations in environmental conditions sites and gradual degradation/deformation of check valve o-rings following frequent contact with S40 sterilant use solution result in nuisance aborts. | Device Design | STERIS sent an Urgent Voluntary Field Correction Notice dated November 21, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed that a STERIS Field Service Representative would contact their facility to schedule installation of the new hardware and updated software on their SYSTEM IE Liquid Chemical Sterilant Processing System(s). Customers were instructed that they may continue to operate the SYSTEM IE processor prior to the completion of the update. Customers with questions were instructed to contact STERIS Field Service Dispatch at 1-800-333-8828.For questions regarding this recall call 440-392-7519. |  | 6028 | 6,028 units | Nationwide Distribution - USA (nationwide) and the District of Columbia and Puerto Rico.Product was also shipped to the following countries: Afghanistan, Bahrain, Hong Kong, Italy, Saudi Arabia, Turkey &amp; United Kingdom. | Terminated | June 16, 2014 | 207 | 6028 | Computer
Z-0540-2014 | 66987 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | MedfusionÂż 4000 Syringe Infusion Pumps |  |  | November 22, 2013 | 2013 | Smiths Medical ASD, Inc. | Smiths Medical is conducting a recall involving all MedfusionÂż 4000 Syringe Infusion Pumps.Smiths Medical has identified an issue with electrical shorting at the AC power receptacle. Shorting can occur when saline solution or other conductive fluids (e.g., potassium chloride) enter the AC power receptacle. When this happens, sparking and/ or flames, accompanied by smoke, may appear at the AC po | Device Design | Smiths Medical sent an "Urgent Medical Device Correction" letter dated November 22, 2013, to all affected customers.The letter identified the product the problem and the action needed to be taken by the customers.All customers will be provided with instructions on inspection, cleaning, and prevention regarding this issue, for the continued safe and effective use of these pumps. Additionally, customers will be instructed to:1. Return damaged pumps for servicing or damaged AC power cords for replacement to Smiths Medical.2. Return the completed Confirmation Form to acknowledge receipt and understanding of the information provided in the Urgent Medical Device Correction letter, via fax or email, to Smiths Medical for processing."The Instructions for preventing this issue", "Mitigations", Conclusions", " Action to be taken by the User", and "Transmission of this Urgent Medical Device Correction". The consignees were also asked to complete the "Urgent Medical Device confirmation Form" and fax it to 1-800-237-8033 or by sending an electronic copy via e-mail to medfusion4000nov2013@smiths-medical.com.For questions regarding this information, consignees should contact their local Sales Representative, Smiths Medical's Customer Service Department at 1-800-258-5361, or via e-mail at medfusion4000nov2013@smiths-medical.com.UPDATE: Consignees were sent a follow-up letter dated 13 January 2014 from the one dated November 22, 2013. The letter was addressed to Risk / Safety Managers, Biomedical Professionals, Clinicians who oversee the use of Medfusion Pumps, Distributors, and other users of these devices. The "Urgent Medical Device Recall" letter included Advise on Action to be Taken by the User. This included customers to have their Pumps undergo a free correction. The correction consists of the installation of a new AC power Receptacle seal. Smith Medical provided options for completing the correction process which included: 1) Consignees were a |  | 9560 | 9560 (8601 US, 959 OUS) | Worldwide Distribution - USA (nationwide ) including the states of : AZ, NY, TN, MS, AR, WV, VA, VT, PA, DE, DC, OH, CO, NH, MI, CA, FL, CT, WA, MD, NM, RI, MA, TX, NC, IL, KY, WI and MI., and the countries of CANADA, SAUDI ARABIA and AUSTRALIA. | Terminated | March 25, 2015 | 488 | 9560 | Computer
Z-0545-2014 | 66983 | No | EIA | Dental operative unit and accessories. | Dental | Dental | 510(k) | SEAC Advanced Dental System, Model A591561 |  |  | November 05, 2013 | 2013 | ASI Medical, Inc. | ASI Medical is initiating a field correction for the SEAC Advanced Dental System because the miniature plastic air tank can potentially rupture. | Device Design | ASI Medical sent an Field Correction Action and Safety Notice letter dated November 5, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.REQUIRED FIELD CORRECTIONOur field correction will consist of sending you a warning label to be applied to the back service access panel of the cart and an ASME safety pop off valve that will open and exhaust if pressure should reach over 100 psi. This safety pop off valve is designed be easily installed in line between the air compressor and air tank in your system. We are requesting your cooperation in making this field correction. |  | 337 | 337 | Worldwide Distribution - USA (nationwide) including the countries of Austria, Aruba, Saudi Arabia, Canada, Germany, Spain and West Indies. | Terminated | August 11, 2014 | 279 | 337 | Computer
Z-0544-2014 | 66698 | No | MNI | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | PedFuse Reset Screw; Model: 01-80115-55 |  |  | October 18, 2013 | 2013 | SpineFrontier, Inc. | During routine servicing, it was discovered that one 8mmx55mm Reset screw was assembled with an 8mm tulip head and a 9mm screw. The implant should haveconsisted of an 8mm tulip head and an 8mm screw. A recall was initiated to remove all remaining screws of the same lot from the field to investigate the nonconformity. | Software design (manufacturing process) | SpineFrontier sent an Voluntary Recall Notification dated November 5,2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.All affected implants are to be returned to SpineFrontier for inspection. SpineFrontier is requiring confirmation that you have received this notice and that you will initiate the immediate return of the implants listed above.IMMEDIATE RESPONSE REQUIRED.Please feel free to contact us at 978-232-3990 Monday through Friday8:30AM to 5:30PM EST with any questions |  | 35 | 35 Devices | US Distribution to VA and TX. | Open, Classified |  |  | 35 | Computer
Z-0119-2014 | 66059 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Xario Diagnostic Ultrasound System SSA-660A |  |  | August 05, 2013 | 2013 | Toshiba American Medical Systems Inc | The firm initiated this recall of certain Toshiba Diagnostic Ultrasound Systems because it has been found that under certain conditions, the temperature of the transducer head may increase in 2D+PW mode or 2D+CDI (POWER, ADF) +PW mode. When the system is used under both conditions listed below, the temperature of the transducer head connected to the system may increase in temperature and possibly | Device Design | Toshiba sent an Urgent Medical Device Correction letter dated August 5, 2013. along with a Reply Form via USPS return receipt mail. The letter identified the product, the problem, and the action to be taken by the customer. Until the software is corrected customers were instructed to change the preset settings as described in the letter. Customers were asked to complete and return the attached form and fax it to the toll free number. The form also can be sent via email to raffairs@tams.comCustomers with questions were instructed to call 800-421-1968 or contact their local Toshiba Representative at 800-521-1968.For questions regarding this recall call 714-730-5000. |  | 17 | 17 (1352 Total) | Nationwide Distribution including and US Territory Puerto Rico | Open, Classified |  |  | 17 | Computer
Z-0823-2013 | 64242 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, CARESCAPE Monitor B650. |  |  | September 30, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware of potential issues associated with the CARESCAPE Monitor B650. There have been no reported illnesses or injuries that have occurred with the use of the device. | Software design | GE Healthcare sent an Important Product Notification letteron September 30, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to keep the patient under close observation or use an alternate monitoring device. Customers should not disconnect the affected USB devices from the patient monitor during an active patient case. If the license is enabled, the alarm limit adjustments from the iCentral Client has an L-NETM-00 license should be disabled. GE Healthcare would contact customers to schedule the update. Customers with questions should contact Technical Support or their local representative.For questions regarding this recall call 262-513-4122. |  | 2813 | 2813 | Worldwide Distribution - USA including AL, AZ, AR, CA, CO, CT, FL, GA,IL,IN, IA, KS, KY,LA, MD, MA, MI, M, MO, NE, NE,NJ, NM, NY, NC, OH, OK PA, SC, TN, TX, UT, VA, WI. Internationally to UNITED KINGDOM, UNITED ARAB EMIRATES, TURKEY, THAILAND, SWITZERLAND, SWEDEN, SPAIN, SOUTH AFRICA, SLOVENIA, SLOVAKIA, SINGAPORE, SAUDI ARABIA, ROMANIA, PORTUGAL, POLAND, PERU, NORWAY, NEW ZEALAND, NETHERLAND, MALAYSIA, LITHUANIA, KUWAIT, JAPAN, ITALY, ISRAEL, IRELAND, INDIA, ICELAND, HONG KONG, GERMANY, FRANCE, FINLAND, EGYPT, DENMARK, CZECH REPUBLIC, COLOMBIA, CHILE, CANADA, BELGIUM BAHRAIN, AUSTRIA, AUSTRALIA. | Terminated | January 27, 2014 | 850 | 2813 | Computer
Z-0634-2013 | 63746 | No | JJH | Clinical sample concentrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Abbott m2000sp Instrument; an automated fluid handling system for performing sample preparation for nucleic acid testing; Abbott Molecular Inc., Des Plaines, IL 60018 USA; Abbott m2000sp Instrument; an automated system for performing sample preparation for nucleic acid testing; Abbott Molecular Inc., Des Plaines, IL 60018 USA; list 09K14 |  |  | December 04, 2012 | 2012 | Abbott Molecular | An issue with the m2000sp Instrument software has been identified whereby, after a series of steps, if changes are made to the control or calibrator information on the Sample Extraction Assay Details screen, the changes are not saved. | Software design | Abbott Molecular Inc. sent a Field Correction Recall/Urgent Field Safety Notice letter dated December 4, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The customer was instructed to follow the instructions identified until an upgrade to their m2000sp software can be installed. Abbott Molecular estimates that this software version will be available the first quarter of 2013. They were requested to complete the enclosed reply sheet and fax it to (847) 775-6728 at prior to December 18, 2012 to acknowledge receipt and understanding of the letter.For questions customers were instructed to contact their local Abbott Molecular representative.For questions regarding this recall call 224-361-7274. |  | 215 | 215 units - U.S., 693 units - ROW | Worldwide Distribution - USA (nationwide) and internationally to Algeria, Argentina, Australia, Austria, Azerbaijan, Belgium, Bolivia, Brazil, Burkina Faso, Burundi, Cameroon, Canada, Chad, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Denmark, Ethiopia, Estonia, Finland, France, Gaza &amp; Jericho, Germany, Greece, Guinea, Haiti, Honduras, Hong Kong, India, Indonesia, Ireland, Israel, Italy, Japan, Kazakhstan, Kenya, Libya, Lithuania, Luxembourg, Malawi, Malaysia, Mali, Mexico, Mozambique, Netherlands, New Zealand, Norway, Pakistan, Panama, Paraguay, Peru, Poland, Portugal, Romania, Russia, Saudi Arabia, Singapore, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Turkey, Uganda, Ukraine, United Arab Emirates, United Kingdom, Venezuela, Vietnam and Zambia | Terminated | March 08, 2016 | 1190 | 215 | Computer
Z-0168-2014 | 66239 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | COBAS INTEGRA 800 Analyzer with software 9864.C2 |  |  | September 16, 2013 | 2013 | Roche Diagnostics Operations, Inc. | Roche COBAS INTEGRA 800 system, Software Version 9864.C2 does not perform necessary ISE Service Actions. Whenever the ISE rack is removed, the ISE counters in the software are reset to zero and the corresponding Service Actions are not conducted automatically by the system. Additionally, certain maintenance activities that have to be performed manually by the operator are not requested on the sof | Software Design Change | Roche sent an Urgent Medical Device Correction letter dated September 16, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Roche Diagnostics recommends tracking all ISE counters and manually requesting the corresponding Service Actions listed in the notification be monitored until updated software becomes available during Q4 2013. If the affected product was moved to another site, the notification is to be forwarded. consignees were asked to complete the attached faxback form and fax it to 1-877-266-0997 and file the Urgent Medical Device Correction (UMDC) for future reference. Question should be directed toRoche Support Network Customer Support Center 24 hours a day, seven days a week at 1-800-428-2336For questions regarding thisr ecall call 317-576-3911. |  | 527 | 527 units | Nationwide Distribution and Puerto Rico | Terminated | May 12, 2015 | 603 | 527 | Computer
Z-0127-2014 | 66456 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | Dekompressor Percutaneous Discectomy Probe. |  |  | November 04, 2013 | 2013 | Stryker Instruments Div. of Stryker Corporation | Stryker has identified a potential for the auger and/or cannula to fracture when the Dekompressor is used contrary to the warnings in the Instructions For Use (IFU). Specifically, manually bending the cannula can weaken the device and may result in a fracture of the auger and/or cannula. This could result in the need for additional medical intervention to remove the broken piece, pain or loss of m | Device Design | On 11/4/13, an URGENT MEDICAL DEVICE Recall Notification was sent to all domestic consignees via mail and foreign by email. An advisory was sent to every physician that uses any of the products listed. Every physician thatuses any of these products must print their name and sign below which means you acknowledge that youve received a copy of the attached advisoryand that you understand the information regarding the proper usage of the Decompressor.Once the form has been signed by all physicians, the form mustbe completed and signed by the recall coordinator. Please fax to866-521-2762 or email to kara.spath@stryker.com. |  | 20385 | 20,385 total distributed since 2004 | Nationwide Distribution including the states of AR, AZ, CA, CO, FL, GA, IL, KS, MS, ND, and NE. | Terminated | April 29, 2014 | 176 | 20385 | Computer
Z-0116-2014 | 65471 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | The Centricity PACS Workstation is intended for use as a primary diagnostic and analysis tool for diagnostic images by trained healthcare professionals, including radiologists, physicians, technologists, clinicians and nurses. It is also intended for use as a clinical review workstation throughout the healthcare facility. The workstation interface provides the user with a means to display, manipulate, archive, print and export images when connected with the Centricity PACS infrastructure. |  |  | June 11, 2013 | 2013 | GE Healthcare It | GE has recently become aware of potential issues due to out of context issues associated with specific workflows of the Centricity PACS RA1000 workstation. These issues are as follows: 1) In Centricity RIS-IC, and a RIS-Driven integration, the "close" button closes the exam in RIS-IC, but does not close the exam in Centricity PACS RA1000 Workstation. 2) Interrupted workflow could break the | Software design | URGENT MEDICAL DEVICE CORRECTION letters dated June 11, 2013 was mailed to all customers alerting them to these potential safety issues. The software was modified to correct these defects and a GE Healthcare service representative will contact each customer to make arrangements for installation of the new software version at no charge to the customer. Customers with concerns or questions about the market action can contact their local GE Healtchcare IT service representative or GE Healthcare's Remote Online Center - Customer Care Center at 1-800-437-1171. |  | 623 | 623 units | Worldwide Distribution: US (nationwide) including states of: AK, AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI and WV) and the countries of the UK, Taiwan, Switerzland, Sweden, Spain, Singapore, Saudi Arabia, Netherlands, Malaysia, Korea, Japan, Italy, Israel, India, Hong Kong, Germany, France, Denmark, China, Canada, Bermuda, Beligum and Australia. | Open, Classified |  |  | 623 | Computer
Z-0827-2013 | 64248 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ELECTROLASE(R) HYFRECATOR(R) ELECTRODES - BLUNT, REF/Catalog Number 7-101-8BX, Disposable Hyfrecator(R) Tips, Sterile, Rx Only --- CONMED CORPORATION USA. |  |  | January 28, 2013 | 2013 | ConMed Corporation | Some devices had broken through the seal of the sterile pouch. ConMed confirmed instances where the pouch seal was compromised on the affected products. | Equipment maintenance | ConMed sent Urgent Device Recall Letters (dated 1/28/13) and Business Reply Forms to all domestic customers via USPS Priority Mail. International notices were sent via FedEx International Priority to the foreign consignees. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review their inventory and complete Attachment II response form and return it with unused affected product to: ConMed Corporation 525 French Road, Utica, NY 13502. Customers were asked to return Attachment II even if they don't have any affected product. For questions regarding this recall call 315-624-3533. |  | 104250 | Domestic Distribution: 97,450 units; Foreign Distribution: 6,800 units | Worldwide Distribution-USA (nationwide) and the countries of China, Dubai and Indonesia. | Terminated | April 11, 2017 | 1534 | 104250 | Computer
Z-0132-2014 | 65586 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | DRX Revolution, Catalog #'s 1)1019397, 2) 1023415 (w/ installation warranty, US-only) and 3) 1023423 (w/ parts warranty, US-only). |  |  | June 19, 2013 | 2013 | Carestream Health Inc. | Carestream Health Inc. is conducting a recall for the Column end Cover present on the DRX Revolution Mobile X-Ray System due to the cover dislodging from the column. | Device Design | The firm, Carestream Health Inc., sent a "URGENT: Medical Device Recall" letter dated June 17, 2013 via courier service to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure that there is no opportunity for interference between the End Cover and potential obstructions as the column is descending. Carestream will deploy a field service representative to apply a warning label to the column, and to provide an addendum to the Safety and Regulatory Information manual. Your Carestream service representative will contact you to answer any questions and schedule a visit as soon as possible. If you have any questions or concerns, please contact the Carestream Customer Care Center at 1-800-328-2910 from 8am through 8pm. |  | 331 | 331 units | Worldwide distribution: US (nationwide) including states of: AL, AR, AZ, CA, CO, FL, GA, IA, ID, IL, IN, IO, KS, LA, MA, MD, MI, MN, MO, NC, NY, OH, PA, SC, TN, TX, UT, VA, WI, and WV; and countries of: Australia, Brazil, Canada, Dubai, Germany, Japan, Kuwait, Malaysia, New Zealand, Saudi Arabia, Singapore, South Africa, Taiwan, and Thailand. | Open, Classified |  |  | 331 | Computer
Z-0567-2014 | 66976 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Ziehm Vision R C-Arm Interventional Fluoroscopic X-Ray System. |  |  | November 21, 2013 | 2013 | Ziehm Imaging Inc | Engineering change in 2006 resulted in a minor non-conformity of their devices. Specifically the emergency switch was changed from a standard component to just an option on the devices. | Device Design | Ziehm Imaging sent an "URGENT RECALL NOTICE" letter dated November 22, 2013, to all affected customers. The letter identified the products the problem and the action needed to be taken by the customers. Customers were advised and instructed to:Your Model and Serial number can be found on the product label attached to the side of the mobile C-arm Stand. Our records indicate that your facility has one or more of the products listed above that are being used in your facility.Recall of the product was initiated because according to our manufacturing documentation, the referenced device installed at your facility was not equipped with an emergency stopping device for vertical travel of the c-arm imaging assembly. To fully comply with applicable standards the referenced device must be retrofitted with an emergency stopping device identification and label.The issue:According to our manufacturing documentation, the referenced device installed atyour facility was not equipped with an emergency stopping device.To fully comply with applicable standards the referenced device must be retrofitted with an emergency stopping device. The manufacturer Ziehm Imaging GmbH advised customers they may continue to operate the affected device with caution until the correction has been carried out on-site. During this time Ziehm Imaging GmbH advises the operator to be aware of the following safetymeasures which may resolve the possibility of potentially hazardous situations by:" Operating the power-off switch located either on the C-arm Stand top panel or on the monitor cart" Removing the power cable from the wall outlet or," Activate the push-buttons for reverse vertical movement and keeping the pushbutton pressed until patient /operator is no longer in the risk zone or the device stops Please alert all staff members working with the device to these potential hazardous situations, and familiarize them with all possibilities to resolve the hazardous situation |  | 103 | 103 devices | USA (Nationwide) Distribution including the states of AK, AL, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MT, MA, MD, MI, MN, MN, MO,MS, NC ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY and Hawaii. | Terminated | June 23, 2014 | 214 | 386 | Computer
Z-0570-2014 | 67046 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart MRx Monitor/Defibrillator with CO2/EtCO2 Measurement |  |  | December 10, 2013 | 2013 | Philips Medical Systems, Inc. | MRx defibrillator displays a -?- for EtCO2 and does not display EtCO2 values when patient CO2 level falls below 7.6 mmHg | Software design | Philips Healthcare issued an Customer Information letter dated February 20, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer with the addendum to the Instructions for Use.If you have a ? in the parameter block and no CO2 waveform on the display, the waveform source is invalid. If invalid, check patient, confirm airway status and examine the FilterLine to make sure it is connected to the device and not kinked or pinched. For more information on the EtCO2 Option seeMonitoring Carbon Dioxide chapter in the HeartStart MRx Instructions for Use.Should you have any questions or concerns, please contact your local Philips representative at 1-800-722-9377. |  | 18529 | 18,529 units | US Nationwide and OUS. | Terminated | April 18, 2017 | 1225 | 18529 | Computer
Z-0571-2014 | 66754 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Instrument Manager driver : Driver sysu100i that connects to the Sysmex UF-100 analyzer |  |  | November 06, 2013 | 2013 | Data Innovations, Inc. | When connected to the Sysmex UF-100, theInstrument Manager driver sysu100i driver software versions 7.00.0001, 7.00.0002, 7.00.0003, 7.00.0004, 7.00.0005, 8.00.0001, 8.00.0002, 8.00.0003,and 8.00.0004, and is configured with option "Strip Zeros from Result" set to "Yes" may result in sample error | Software design | Data Innovation notified consignees by Important Medical Device Recall Notification dated 11/5/13 to all list serv recipients of the problem and posted critical alert on Customer Web Portal. A certified letter was issued to impacted customers and consignees if no response was recieved from e-mail. |  | 24 | 24 systems | NationwideCanada and Britain | Terminated | December 05, 2014 | 394 | 24 | Computer
Z-0561-2014 | 65131 | No | JMG | Ferritin immunological test system. | Immunology | Clinical Chemistry | 510(k) | Dimension(R) Chemistry Wash (RD701), Siemens Material Number (SMN) 10445052 |  |  | May 06, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics received complaints of Chemistry Wash causing QC and patient sample result shifts on CTN1, LTN1, FT4, TSH, PBNP and LPBN when using Chemistry Wash lots RD23031, RD23111, RD23241, RD23311, RD23391, and RD23461. These lots shipped between 12/12/2012 and 3/22/2013. Internal investigation has confirmed shifts can be encountered in two scenarios: 1 - When an affected l | Software design | Siemens sent an Urgent Medical Device Recall letter dated May 2013 to all affected customers. The letter identified the product and lots, the reason for the field action, the risk to health and the action(s) to be taken. Customers were instructed to discontinue use and discard inventory of the affected product. Siemens will replace any unused inventory of the affected lots at no charge. For questions contact your Siemens Technical Center or your local Siemens technical support representative. |  | 17719 | 17,719 | US Nationwide and OUS. | Terminated | August 27, 2014 | 478 | 17719 | Computer
Z-0580-2014 | 66980 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | steroEOS Workstation with software versions: 1.3, 1.4 and 1.5 used in processing 3D X-ray images |  |  | November 29, 2013 | 2013 | EOS Imaging | Error copying information to the clinical file and to the report when two 3D models are opened simultaneously on the sterEOS workstation | Software design | EOS imaging issued a safety information letter dated November 29, 2013, to the US customer sites. The letter identifed the affected product, problem and actions to be taken. Users of the sterEOS workstation are asked not to simultaneously open several pairs of images in sterEOS 3D. The letter informed customers that EOS imaging will, at it's own expense, corect all systems affected by means of a software update. An EOS imaging representative will contact affected customers to make an appointment for this update once it is available. Users of the sterEOS workstation are asked not to simultaneously open several pairs of images in sterEOS 3D. For questions call +33 (0) 155 25 6060. |  | 36 | 36 units | USA Nationwide Distribution | Terminated | January 13, 2015 | 410 | 36 | Computer
Z-0645-2013 | 63829 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ABX PENTRA ML, Software Versions: V8.0.x and Below; V.9.0.1 and V9.0.2 |  |  | September 20, 2012 | 2012 | Horiba Instruments, Inc dba Horiba Medical | Horiba Medical is recalling ABX PENTRA ML because they have confirmed an issue with the ABX PENTRA ML data management and validation system. An incorrect patient result is transmitted to the Laboratory Information System. This result is from a previous run of the same patient that is stored in the ML database. This issue occurs in a very rare sequence of specific events. The affected software | Software design | .A recall letter dated 11/09/2012 was sent to all their customers who received the ABX PENTRA ML products. The letter informed the customers of the problem identified and the action to be taken. Customers are instructed to complete and return the enclose response form within 10 days. Customers with questions about the letter are instructed to contact their local HORIBA Medical representative.Recall Letter, as produced per FCD-0377, Field Safety Notice &amp; Acknowledgement will be mailed via FedEx to the ten (10) customers noted on the list the week of November 12, 2012. Effectivity will be evaluated by receipt of the Field Safety Notice Acknowledgment Form. Follow-up communications with customers will be conducted, as needed. |  | 10 | 10 units | Nationwide in the US, including the states of CA, MI, IN, FL, and GA. Center Recall Depth is Retail. | Terminated | March 12, 2013 | 173 | 10 | Computer
Z-0838-2013 | 64246 | No | HIS | Condom. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | LifeStyles ZERO 10 uber-thin lubricated latex condoms. Mfd. for Ansell Healthcare Products LLC, Dothan, AL 96303. |  |  | January 14, 2013 | 2013 | Ansell Healthcare Products LLC | Out of an abundance of caution, Ansell is implementing a voluntary recall of the Zero Latex Condom (regular) product. The condom has met or exceeded test and quality standards required by the Food and Drug Administration. However, we are experiencing a number of complaints concerning the difficulty in unrolling and subsequent breakage, and so we are asking to have all products returned. | Device Design | Ansell HealthCare Product LLC sent an Recall letter dated January 14, 2013, via registered mail to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The recall notification informs customers about the problem with the affected product and provides instructions to the customers on how to handle affected product. Customers are asked to complete the Response Form and return via fax (1-800) 722-8155.If you have further questions, please contact our Customer Service Department at 1-800-952-9916. |  | 469134 | 469,134 units | .Nationwide Distribution including the states of AL, AZ, CA, CO, DE, FL, GA, IA, IN, IL, KS, KY, LA, MI, MN, MO.MS, NC, NH, NJ, NY, OH, OR, PA, RI, SC, TN, TX, VA, WI and UT. | Terminated | August 06, 2015 | 934 | 469134 | Computer
Z-0016-2014 | 65161 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Siemens RAD Fluoro Uro Systems - Luminos dRF, Ysio and Uroskop Omnia with software versions VB10D, VB10F, VB10C and VB10E. |  |  | May 06, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Siemens has discovered through product monitoring a potential malfunction when using system operating with software versions VB10D, VB10F, VB10C and VB10. During a RAD examination the possibility exists that radiation may continue longer than expected and may result in unnecessary dose applied during an acquisition. Additionally, it is not possible for the user to stop the radiation manually; ho | Software design | Siemens initiated this correction on May 6, 2013 by sending Update InstructionXP005/13/S. Firm subsequently issued a letter on June 3, 2013 to all accounts.1. Siemens Field Modification of the Luminos dRF, Ysio &amp; Uroskop Omina running software versions VB10D, VB10F, VB10C and VB10E, which improves the signal transmission for the automatic exposure control, is available as of May 2013, and systems will be updated within one year.2. The revised software fix for models Luminos dRF, Ysio &amp; Uroskop Omina running software versions VB10D, VB10F, VB10C and VB10E, will be provided to all identified customers at no charge.CDRH approves the CAP subject to the following conditions:1. Notification of all dealers and purchasers is to be made within 15 workingdays of receipt of this letter in the manner specified in 21 CFR 1003.21 and1003.22. This office and the Food and Drug Administration (FDA) districtoffice coordinator noted below are to be included in the notification.2. Siemens Medical Solutions USA, Inc., will notify the FDA district office inPhiladelphia, Pennsylvania, with reports of progress regarding this fieldcorrection. |  | 23 | 23 | Nationwide Distribution | Terminated | November 06, 2014 | 549 | 23 | Computer
Z-0125-2014 | 66473 | No | LGW |  |  | Neurology | PMA | Medtronic Activa PC, model 37601, and Activa RC, model 37612, Multi-program Rechargeable Neurostimulator for Deep Brain Stimulation. |  |  | September 30, 2013 | 2013 | Medtronic Neuromodulation | Medtronic notified customers regarding over stimulation or stimulation in the wrong area related to select Medtronic neurostimulators. Upon a specific set of conditions, typically related to device recovery from an over discharge, there is a potential for over stimulation or stimulation directed to a lead electrode other than what was intended. | Software design | The firm, Medtronic, sent an "Urgent: Medical Device Correction" letter dated September 2013 entitled "Loss of Stimulation and Over Stimulation" to customers associated with the affected Medtronic neurostimulators. This letter described the product, problem and actions to be taken. The letter informed customers of the potential for these issues to occur and provided recommendations to significantly reduce the likelihood of the issues. The customers were instructed to avoid over discharge; if a rechargeable device becomes over discharged, contact Medtronic Patient Services at (1-800-510-6735) to resolve the POR and reduce the likelihood of over stimulation.If you have questions, please contact Medtronic Neurostimulation Technical Services at 1-800-707-0933 weekdays 7am-6pm CST. |  | 38294 | 38,294 | Worldwide distribution: US (Nationwide) including Puerto Rico and countries of: Algeria, Aruba, Austria, Belarus, Belgium, Brazil, Bulgaria, Chile, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, El Salvador, Estonia, Faroe Islands, Finland, France, French Polynesia, Georgia, Germany, Greece, Guadeloupe, Hungary, Iceland, India, Iran, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kuwait, Latvia, Lebanon, Libya, Lithuania, Luxembourg, Malta, Martinique, Mexico, Morocco, Netherlands, Netherlands Antilles, New Caledonia, Norway, Oman, Pakistan, Panama, Peru, Poland, Portugal, Qatar, Reunion, Romania, Russian Federation, San Marino, Saudi Arabia, Serbia, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay, and Venezuela. | Terminated | November 20, 2015 | 781 | 102457 | Computer
Z-0836-2013 | 64307 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | Animas(R) 2020, Animas(R) IR 1250 (1200 plus), Animas(R) IR 1200 Insulin Infusion Pumps |  |  | December 17, 2012 | 2012 | Animas Corporation | Due to an internal calendar date limit (December 31, 2015) these pumps will cease to operate on January 1, 2016, and will result in a call service alarm that is displayed on the pump display. | Device Design | The firm initiated their recall of this product on December 17, 2012 by sending a letter to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the Animas Customer Support Department at (866) 796-6373 for questions or concerns. |  | 37998 | 37,998 | Worldwide Distribution-USA (nationwide) including Washington, DC, Puerto Rico, US Virgin Island and the states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY, and the countries of Austria, Australia, Bermuda, Canada, New Zealand, Czech Republic, Italy, France, Finland, Hungary, Israel, United Kingdom, Spain, Sweden, Norway, Denmark, Germany, Mexico and Ireland. | Terminated | December 24, 2013 | 372 | 37998 | Computer
Z-0090-2014 | 65805 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | DRX Revolution Mobile X-Ray System, Model DRXR-1, Verona Street Rochester, NY 14608 |  |  | July 26, 2013 | 2013 | Carestream Health, Inc. | During an audit, Carestream discovered a DRX-1 System that captured an image which displayed an artifact. This issue was discovered in manufacturing. It occurs when the DRX product console software is updated to v5.6b software and the detector is not calibrated before use. | Software design | Carestream sent a Urgent-Medical Device Recall letter dated July 26, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Carestream Health Inc. recently discovered an image artifact issue due to a detector not being recalibrated with the 5.6B software upgrade. As our installation instructions did not require the user to recalibrate during an upgrade, there is a possibility your detector could do the same. When this occurs an image artifact could be evident on images captured in the Extended Exposure mode. These artifacts could potentially contribute to misdiagnosis or the necessity of patient re-exposure.We request that you take a moment to verify that your detector(s) are calibrated. At our next visit, Carestream will also verify that your detector was re-calibrated and, if not, re-calibrate the detector for you. Please note this only affects systems upgraded in the field.For questions or concerns, please contact the Carestream Customer Care Center at 1-800-328-2910, 8 am through 8 pm EST in the US. |  | 23 | 23 units installed total (11 domestically &amp; 12 internationally) | Worldwide Distribution - US (nationwide) including the states of New York, Massachusetts, Pennsylvania, Illinois, Texas, California and Louisiana., and the countries of Australia, Canada, Finland, Italy, Switzerland and United Kingdom. | Open, Classified |  |  | 23 | Computer
Z-0126-2014 | 65469 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Carestream DRX-Revolution Mobile X-ray System. |  |  | June 06, 2013 | 2013 | Carestream Health, Inc. | Carestream Health Inc. is recalling the power plug on their DRX-Revolution Mobile X-Ray System due to the item generating sparks and becoming warm to touch. | Device Design | On June 6, 2013 Carestream Health Inc. initiated a recall for the power plug on their DRX-Revolution Mobile X-Ray System due to the item generating sparks and becoming warm to touch. Carestream Health Inc. will replace the power plug in all potentially impacted DRX-Revolution Mobile X-Ray Systems. Your Carestream service representative will contact you to answer any questions and schedule a visit to replace the plug of your Mobile X-Ray System as soon as possible. If you have any questions or concerns, or if you experience any issue with your plug prior to our visit, please contact the Carestream Customer Care Center at 1-800-328-2910, 8AM through 8pm Eastern Standard Time. |  | 180 | 180 units | Nationwide Distribution including the states of NY, TX, PA, OH, WI, CA, IL, CO, FL, IN, OK, NJ, GA, LA, AL, KS, AZ, MA, SC, MI, NC, IA, UT, VA, TN, and MN. | Open, Classified |  |  | 180 | Computer
Z-0655-2013 | 63938 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | FlowGate Balloon Guide Catheter; |  |  | December 20, 2012 | 2012 | Concentric Medical Inc | Firm received complaints of resistance and possible collapse of the distal tip of the FlowGate Balloon Guide Catheters during use for aspiration. | Device Design | Concentric Medical sent a Urgent Medical Device Recall Notification letter dated December 19, 2012, to all affected customers. The letter identified the product the problme and the action needed to be taken by the customer.Customers were instructed to follow these instructions.1. Immediately locate the subject devices and quarantine them.2. Distribute this notice to all affected Departments in your facility.3. Inform Concentric Medical (Stryker Neurovascular) if any of the subjectdevices have been distributed to other organizations.a. Please provide contact details so that Concentric Medical (StrykerNeurovascular) can inform those recipients appropriately.4. Complete and sign the attached Customer Response Form and: Fax it to: 650-237-5230 Please complete this form even if you do not have any subject devices to return to prevent the need of follow up notices.5. Keep a copy of the completed, signed Customer Response Form for yourrecords. Upon receipt of the returned customer response form, a Stryker Representative will contact you to arrange for credit of returned devices.On behalf of Stryker Neurovascular, we sincerely thank you for your help and support and regret any inconvenience and appreciate your assistance in completing this recall.Should you have any questions concerning this matter please do not hesitate to contact the undersigned. |  | 13 | 13 units | US (nationwide) including the states of CA, FL, GA, OR and TN. | Terminated | February 05, 2013 | 47 | 13 | Computer
Z-0602-2014 | 67094 | No | GWM | Intracranial pressure monitoring device. | Neurology | Neurology | 510(k) | Integra Camino Advanced Monitor with Waveform Display Model CAM01 Integra Neurosciences, 311 Enterprise Dr Plainsboro, NJ 08536 USA |  |  | November 26, 2013 | 2013 | Integra LifeSciences Corp. | Integra became aware through a complaint from a single non-US customer that if the optional SYS/DIAS mode is used on CAM01 monitor that contains Mod 4 Firmware, the displayed values for the systolic and diastolic intracranial pressure (ICP) are reversed. But, if the CAM01 is used with an auxiliary bedside monitor, the bedside monitor displays these values correctly. | Device Design | Integra sent an Urgent Medical Device Correction letter and Correction Acknowledgement form dated November 26, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Our records indicate that you received one or more of the affected serial numbers that will require this firmware update. Integra asks that you do the following:1. Review your inventory and verify that you have Camino Advanced Monitors with a serialnumber formatted as MExxxxxx04. If so, you can continue to use the CAM01monitor(s), being aware that if you choose to use the optional SYS/DIAS mode, thevalues on the CAM01 display will be reversed.2. Complete the attached form. If you do not have the product with the affected serialnumber in service at your facility, check the box:I do not have any product with a serial number of the form MExxxxxx04.3. If you do have the affected serial number, check the box:I do have the product with a serial number of the form MExxxxxx04, and pleaseindicate the serial numbers you have.4. Complete the other information and return it by email or fax as indicated on the form.Complete the form and return it by email or fax as indicated on the form. Should you have any questions regarding these instructions, please contact Customer Service at 1-855-532-1723.We apologize for any inconvenience this may cause and thank you for your cooperation in this effort. |  | 276 | 275 CAM01 Monitors | Worldwide Distribution - USA ( Nationwide ) and the countries of : UK, Saudi Arabia, South Africa, Brunei, China, France, Poland, Turkey, Greece, Sweden, United Arab Emirates, Russia, Spain, Pakistan, Finland, Israel, Italy, Slovenia, Japan, Poland, Malaysia, Germany and Panama. | Terminated | February 13, 2015 | 444 | 276 | Computer
Z-0867-2013 | 64319 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Eclipse Treatment Planning System for Radiotherapy; Model Number: H48. In addition, the Eclipse Proton Eye algorithm is specifically indicated for planning proton treatment of neoplasms of the eye. |  |  | January 28, 2013 | 2013 | Varian Medical Systems, Inc. Oncology Systems | The calculation of dose dynamic plans containing different energy fields is incorrect when the Pencil Beam Convolution 11.0 algorithm (PBC 11.0) is used. | Software design | Varian Medical Systems, Inc. Oncology Systems sent an Urgent Medical Device Correction Letter dated January 30, 2013 to all affected customers. The Letter identified the affected product, problem and actions to be taken. Varian Medical Systems is developing a technical correction for this issue, Its customers will be contacted by a Varian Customer Service representatives when this correction is available to schedule its installation. The letter provides customers with help desk contact information. |  | 1168 | 1168 units | Worldwide Distribution - USA Nationwide | Terminated | April 04, 2014 | 431 | 1168 | Computer
Z-0552-2014 | 66956 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Plan Parameters workspace within |  |  | November 18, 2013 | 2013 | Varian Medical Systems, Inc. | An anomaly has been traced to a software deficiency within the Plan Parameters workspace. When a dose relevant parameter is changed and "do no clear MUs and Reference Point dose" is selected, the Monitor units may change. | Software design | Varian Medical Systems will provide all affected consignees with an Urgent Medical Device Correction Letter commencing on November 18, 2013 by federal express. . The Letter will also be distributed to the Varian Sales, Marketing, and Service organizations, informing them of the reported issue. |  | 6701 | 6701 | Worldwide Distribution. | Terminated | March 13, 2015 | 480 | 6701 | Computer
Z-0830-2013 | 63783 | No | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | LPS Diaphyseal Sleeves |  |  | January 04, 2013 | 2013 | DePuy Orthopaedics, Inc. | DePuy Orthopaedics is initiating a voluntary recall of the LPS Diaphyseal sleeve because the product has the potential for fracture at the taper junction. The LPS Diaphyseal Sleeve to Diaphyseal Sleeve Base taper connection may not be sufficient to accommodate potential physiologic loads that may be transferred to the junction during normal gait activities by some patients. | Device Design | DePuy Othopaedics sent an Urgent Information Medical Device Recall Notice letter dated January 4, 2013 to all affected customers. The letter identified the affected products, reason for recall, clinical implications and actions to be taken. The letter instructed customers to check inventory, discontinue use, cease further distribution and return all inventory on hand for credit. Customers were requested to complete and return the Reconciliation Form. For product-related questions contact your local DePuy Othopaedics Sales Representative, clinical questions contact DePuy's Scientific Information Office at 1-888-554-2482 and recall notice information call 574-372-7333. |  | 162 | 162 total units (US) and 102 units total (OUS) | Worldwide Distribution - US Nationwide in the states of AL, AZ, CA, CO, FL, GA, ID, IL, MA, MD, MI, MN, NC, NM, NV, NY, OH, PA, SC, SD, TN, TX, VA, WA, WI, and WV and the countries of Holland, Australia, Austria, South Africa and Canada including Veteran Administration hospitals. | Terminated | May 20, 2014 | 501 | 162 | Computer
Z-0599-2014 | 66895 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Asante Conset Infusion Set with 60cm tubing &amp; 9mm cannula packaged in 5-packs,singles, or inside a Combination Kit (FG-5229-4); |  |  | November 19, 2013 | 2013 | Asante Solutions, Inc. | Reports of the infusion tubing disconnecting from the connector, resulting in leakage of insulin from the infusion set and providing patient with less insulin than intended. | Device Design | Asante sent urgent medical device recall notifications letters to distributors and customers on November 19, 2013 by US mail, return receipt requested. |  | 1979 | 1979 kits | Nationwide Distribution. | Terminated | April 24, 2014 | 156 | 6977 | Computer
Z-0583-2014 | 66909 | No | HWA | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker Orthopaedics Ball Impactor Tip |  |  | September 30, 2013 | 2013 | Stryker Howmedica Osteonics Corp. | Stryker has received complaints associated with cracks and/or fracture of the Ball Impactor Tip and Rim Impactor Tip Instruments. | Device Design | Stryker sent an Urgent Medical Device Correction letter dated October 1, 2013 to all affected customers via Fed Ex. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete and fax the attached Product Correction Acknowledgement form to 855-251-3635. |  | 1024 | 1,024 units | USA Nationwide Distribution | Open, Classified |  |  | 2714 | Computer
Z-0118-2014 | 64730 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Centricity PACS IW by GE Healthcare Dynamic Imaging Solutions is a device that receives medical images (including mammograms) and data from various imaging sources. Images and data can be stored, communicated, processed and displayed within the system or across computer networks at distributed locations. Lossy compressed mammographic images and digitized film screen images may only be interpreted using an FDA approved monitor that offers at least 5 M pixel resolution and meets other technical specifications reviewed and accepted by the FDA. Typical users of this system are trained professionals, including but not limited to physicians, radiologists, nurses, medical technicians an assistants. |  |  | June 15, 2012 | 2012 | Ge Healthcare It | Multiple issues were identified with the GE Centricity PACS-IW system that could pose a safety risk to patients. These issues include:1) With Auto-Fetch in use another patient's image may appear without user's intention. 2.a) In the case where dose administration time and image acquisition time fall on either side of midnight, the value will be under reported and close to zero for all pixe | Software design | GE Healthcare sent an IMPORTANT PRODUCT INFORMATION letter dated June 15, 2012 to direct accounts informing them of the Auto-Fetch functionality issue and advising them to discontinue the use of the Auto-Fetch feature and to change the system setting to not allow users privileges to use Auto-Fetch. Customers were requested to contact their Field Engineer or GE Service Representative with any questions or to call the GE Healthcare Remote Online Center Customer Care Center at 1-800-437-1171. GE Healthcare sent an URGENT MEDICAL DEVICE CORRECTION letter dated March 12, 2013 to all affected customers informing them of all three of the safety issues. The letter again adivsed customers to discontinue the use of the Auto-Fetch feature and to change the system setting to not allow users privileges to use Auto-Fetch. The letter further infomed customers that all three errors can be resolved through the installation of a new version of the software. Customers were requested to contact their Field Engineer or GE Service Representative with any questions or to call the GE Healthcare Remote Online Center Customer Care Center at 1-800-437-1171. |  | 847 | 847 units | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, WY, District of Columbia, Puerto Rico, and the U.S. Virgin Islands.Internationally to Austria, Azerbaijan, Bahrain, Belgium, Bosnia, Brazil, Canada, Columbia, Egypt, France, Germany, Ghana, Hungary, Ireland, Israel, Italy, Japan, Kuwait, Lebanon, Libya, Lithuania, Mexico, Netherlands, Nigeria, Palestine, Poland, Qatar, Russia, Saudi Arabia, Slovenia, South Africa, Spain, Switzerland, Turkey, United Arab Emirates, United Kingdom, Uzbekistan, and Yemen.Center Recommended Depth - Consumers/User** | Open, Classified |  |  | 847 | Computer
Z-0142-2014 | 66464 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | S5 Heart-lung machine, Stockert S5 System. Catalogue No: 48-30/40/50-00. |  |  | September 24, 2013 | 2013 | Sorin Group Deutschland GmbH | Sorin Group is recalling certain models of S5 Perfusion Systems due to malfunction of cardioplegia control. | Software Manufacturing/Software Deployment | Sorin Group sent an Urgent Field Safety Notice dated September 24, 2013, to all affected customers. The letter informed users of the potential for failing of automatic cardioplegia delivery function and notified them that they would be contacted by the Sorin Service Group to arrange an appointment to update the firmware on site. Customers were asked to complete the Response form to confirm they had received, read and understood the Field Notice. For questions regarding this recall call 303-467-6527. |  | 84 | 84 | Worldwide Distribution - USA including WA, NY, VA, TN, SC, TX, CA, CO, MA, WI, MD, AL, ND, and Puerto Rico and Internationally to Italy, Finland, Germany, Sweden, and Great Britain. | Terminated | July 10, 2014 | 289 | 84 | Computer
Z-0554-2014 | 66873 | No | MNI | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | PedFuse Reset Screw Inserters; PedFuse Screw Inserter 05-7 REspond Part# 11-80111 |  |  | November 20, 2013 | 2013 | SpineFrontier, Inc. | Screw inserters may not mate properly with hex interface of the screws. | Device Design | The firm, SpineFrontier, Inc., sent an "Urgent PedFuse Screw Inserter Voluntary Recall Notification" letter on November 20, 2013 to its consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed to initiate the immediate return of the recalled products and to complete a recall ACKNOWLEDGEMENT FORM via fax or email to: SpineFrontier, Inc., 500 Cummings Center, Suite 3500, Beverly MA 01915, Attn: Customer Service; Fax: 978-232-3991; email: customerservice@spinefrontier.com.Please feel free to contact at 978-232-3990 Monday thru Friday 8:30AM to 5:30PM Eastern with any questions. |  | 25 | 25 | US Distribution to states of:Texas, New York, Illinois, Ohio, Virginia, Washington D.C., and Florida.. | Open, Classified |  |  | 42 | Computer
Z-0633-2013 | 63838 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | PENTRA C200, Part Number: A11A01629. |  |  | October 10, 2012 | 2012 | Horiba Instruments, Inc dba Horiba Medical | Horiba Medical is recalling the ABX PENTRA AST CP because during the AST application validation on the Pentra C200 using an Audit Systems linearity kit, it was determined that two higher values for AST (expected at 800 U/L and 1000 U/L) were recorded at 280 and 199 U/L without alarm. | Software design | Horiba Medical sent a Field Safety Notification and Acknowledgement Letter dated November 20, 2012, to all affected customers. The letter notified the customers of the problems identified and the actions to be taken. Customers were instructed to complete and return the enclosed response form within 10 calendar days via mail, email or fax. Customers with questions were instructed to contact their local representative.For questions regarding this recall call 714-273-9254. |  | 384 | 384 units | Nationwide Distribution including CA, OK, NY, FL, AZ, NY, VA, and TX. | Terminated | September 12, 2013 | 337 | 67270 | Computer
Z-0534-2014 | 66865 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris PC unit model 8000 and 8015 with Alaris PCA module model 8120 and Auto ID barcode scanner model 8600. |  |  | December 05, 2013 | 2013 | CareFusion 303, Inc. | CareFusion is recalling the Alaris PC unit (model 8000 and 8015) because the Alaris PC unit will produce an audible alarm and the Alaris PC unit will display a System Error message. | Software design | CareFusion sent an Urgent Medical Device Recall Notification letter dated December 5, 2013, to all affected customers. The letter identifed the affected product, problem and actions to be taken. Customers were instructed to contact CareFusion Customer Advocacy at (888) 812-3266, if they experienced a SYSTEM ERROR on the Alaris PC unit. CareFusion will contact all affected customers within 60 days to initiate the scheduling process for the remediation.Customers are instructed to complete and return the enclosed Customer Response Card to expedite the corrective action process. Customers with questions were instructed to contact CareFusion Support Center at 888-562-6018, or for Technical Support call 888-812-3229. |  | 4824 | 4824 total units (4242 units in US) | Worldwide Distribution - USA (nationwide) and Internationally to Canada, Netherlands, and Australia. | Completed |  |  | 4824 | Computer
Z-0139-2014 | 66526 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Patient Cart used in conjunction with the da Vinci S surgical and SI System. |  |  | November 30, 0012 | 12 | Intuitive Surgical, Inc. | Increased number of complaints regarding fluid entering the patient card base of the da Vinci S and SI systems. | Device Design | Intuitive Surgical Inc sent an Product Notification to all affected customers via email on November 30, 2012.E-mail notifications advised users of an addendum to the User Manual for Covering the Patient Cart column that provided a recommended method for covering the da Vinci patient cart column to prevent splashing on the cart during a procedure. The addendum was attached to the e-mail and is available online.If you have any questions, please contact Customer Service at 1-800-876-1310. Outside the United States call +41 21 821 2020 or +800.0821.2020 |  | 1754 | 1754 patient carts | Worldwide Distribution - USA including Puerto Rico and the countries of Australia, Austria, Argentina, Belgium, Bulgaria, Brazil, Canada, Chile, China, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, India, Indonesia, Ireland, Israel, Japan, Kuwait, Mexico, Netherlands, Norway, Pakistan, Panama, Philippines, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom, Uruguay and Venezuela. | Terminated | November 04, 2013 | 730824 | 1754 | Computer
Z-0627-2013 | 63878 | No | FWZ | Manual operating table and accessories and manual operating chair and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Maquet Yuno Extension Table Device |  |  | November 02, 2012 | 2012 | Maquet Inc. | There is a potential issue with the screw tension of the extension device 1433.62A1 used with the YUNO table. | Device Design | Maquet sent Urgent Product Field Correction letters/Field Correction Confirmation Form dated November 2, 2012 on November 15, 2012, via Fed Ex to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed that a MAQUET authorized Service Technicians would contact them to arrange an appointment to assess their device and to replace the release lever at no cost to them at their facility. Customers were asked to completed and return the enclosed confirmation form, acknowledging their receipt and understanding of the communication, in the enclosed postage prepaid envelope, by e-mailing a scanned copy to FieldActions@maquet.com or via fax at 973-709-6508.For questions customers were instructed to call MAQUET Technical Support at 800-777-4222, Press option2, followed by option 2.For questions regarding this recall call 973-709-7660. |  | 63 | 63 devices | Worldwide Distribution - USA including AR, GA, SC, TX, TN and internationally to France, United Arab Emirates, Australia, Brazil, Canada, Switzerland, China, Germany, Spain, United Kingdom, Italy, Japan, Sweden, Thailand, Turkey, Serbia, and South Africa. | Terminated | August 21, 2013 | 292 | 63 | Computer
Z-0512-2014 | 66888 | No | MJN | Vascular clamp. | Cardiovascular | Cardiovascular | 510(k) | Stryker brand SUPER COMPLIANT 7MM X 15MM, Catalog Number: |  |  | November 04, 2013 | 2013 | Stryker Neurovascular | Products were shipped labeled with "use by" dates that exceed currently available shelf-life data. All products remain within the earliest expiration date of April 2015. | Software Manufacturing/Software Deployment | Stryker Neurovascular sent an Urgent Medical Device Recall letter dated November 4, 2013, all affected users. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately discontinue use and segregate the affected products. Customers with questions were instructed to contact via NeuroComplaints@stryker.com. Or Fax the response form to 1-800-876-4355.For questions regarding this recall call 510-413-2500. |  |  |  | Worldwide Distribution - USa (nationwide) and Internationally to:Austria, Belgium, Czech Rep, Germany, France, Great Britain, Hungary, Italy, Lithuania, Netherlands, Norway, Poland, Slovakia, Spain, Turkey, Australia, Canada, Chile, Hong Kong, India, Korea, Malaysia, Russia and Japan | Terminated | August 27, 2014 | 296 |  | Computer
Z-0536-2014 | 66963 | No | JJH | Clinical sample concentrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | MagNA Pure 96 System |  |  | November 25, 0013 | 13 | Roche Diagnostics Operations, Inc. | This issue is only relevant for customers importing xml files into the MagNA Pure 96 IVD instrument software for sample identification. There is a potential for sample mismatch when using the MagNA Pure 96 IVD instrument (catalog number 06541089001). Importing an order file in xml format to the MagNA Pure 96 IVD instrument software may result in a wrong sequence of samples in the graphical overvi | Software design | Roche Diagnostics sent an URGENT Medical Device Correction Notice dated November 25, 2013, to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers.Customers were advised to discontinue use of the Pipetting Report and to use the sample list file instead. The MagNA Pure 96 Instrument Operators Guide is being updated with an addendum addressing this issue. The letter included a Customer Response Form for customers to complete and return to the firm by FAX : 1-888-670-4135. Please contact the Roche Support Network Customer Support Center at 1-800-526-1247, 24 hours a day, seven days a week, if you have questions about the information contained in this UMDC. |  | 27 | 27 | US Distribution including the states of AL, AZ, CA, CT, GA, NY, NC, NJ, OH, TX, VA and WA. | Terminated | September 04, 2014 | 730768 | 27 | Computer
Z-0357-2014 | 66439 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Trinica Anterior Lumbar Plate (ALP) system, part 07.01058.001. Non Sterile, Rx only. The Trinica Anterior Lumbar Plate System is a temporary supplemental fixation device consisting of a variety of shapes and sizes of plates and screws. The Trinica Anterior Lumbar Plate System is used as an implant for the correction and stabilization of the spine. This system provides temporary stabilization and augments the development of a solid spinal fusion. Additionally, this system provides the surgeon with the ability to supplement an interbody device with anterior plate fixation. The Trinica Anterior Lumbar Plate System components can be locked into a variety of configurations and each construct may be customized to individual cases. The plates are low profile and anatomically designed to provide optimal fit from either anterior or anterior-lateral approach. This system also features anti-migration locking caps to help secure the fixation screws. All Trinica Anterior Lumbar Plate System implant components are made from titanium alloy (Ti-6AI-4V). |  |  | October 09, 2013 | 2013 | Zimmer, Inc. | Zimmer Spine, Inc. is initiating a correction regarding the Trinica ALP Instrument Tray, part 07.01058.001 (all lots). The Ratchet Handles, part 07.00438.001 may not be effectively sterilized if left in their designated location(s) in the instrument tray. This is due to the current bracketing in the tray potentially interfering with complete steam penetration into the cannulated handle. | Device Design | Zimmer Spine sent a Medical Device Correction letter dated October 9, 2013 via FedEx A Medical Device Correction letter, dated October 09, 2013 via FedEx to distributors, Hospitals and Surgeons who have received and/or used the device. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the notification and ensure affected personnel were aware of its contents. Only sterilize the Ratchet Handles external to the instrument tray. Customers were to be notified when replacement trays were available. Customer with questions were instructed to call the Zimmer Customer Call Center at 866-774-6368.UPDATE: July 17, 2014 Zimmer issued another letter to consignees to inform them of a redesigned tray is available and will replace the current tray. A Zimmer representative will conduct the swap-out of the current Rev A tray for the new Rev B design tray. A Return Confirmation Form is asked to be emailed back to CorporateQuality.Postmarket@zimmer.com. For assistance or questions, contact Zimmer Spine Customer Service at 866-774-6368. |  | 94 | 94 Tray sets | Worldwide Distribution - USA (nationwide) and Internationally to Nationwide and OUS countries include: Australia, Austria, Belgium, Canada, France, Germany, Italy, New Zealand, Spain, Switzerland, and the UK. | Terminated | February 24, 2015 | 503 | 94 | Computer
Z-0743-2013 | 63866 | No | CEO | Phosphorus (inorganic) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Alfa Wassermann System Diluent 750 mL |  |  | October 10, 2012 | 2012 | Alfa Wassermann, Inc. | Inorganic phosphorus assays using System Diluent Lot 155599 may not product acceptable calibration results. | Device Design | Alfa Wassermann sent an important advisory notice/product response form to their distributors on October 10, 2012. The letter identified the product, the problem, and the action to be taken by the customer. The letter informs customers about the affected lot and provide instruction regarding the receipt of replacement product. Any recalled product should be destroyed. If there is further distribution, those customers should provide Alfa with a distribution list so that they may be notified. Any questions can be directed to (866) 419 - 2532. |  | 4104 | 4,104 kits | Worldwide Distribution - USA (nationwide) and the countries of Canada and Cayman Islands | Terminated | October 25, 2013 | 380 | 4104 | Computer
Z-0426-2014 | 65985 | [6 empty] | Custom Ultrasonics System 83 Plus 2 and Plus 9 Endoscope Washer/Disinfector. |  |  | March 12, 2013 | 2013 | Custom Ultrasonics, Inc. | Certain System 83 Plus 2 and System 83 Plus 9 units running the Windows 7 operating system were not fully validated prior to distribution. | Software design | A notification letter, dated February 26, 2013, was sent to direct accounts to inform them of the recall and that technicians would visit facilities to replace the Windows 7 operating system with an appropriately validated Windows XP operating system. |  | 118 | 118 | Worldwide Distribution: USA (nationwide) and Internationally to: Canada, Slovakia Republic, Sydney and Peru. | Terminated | June 26, 2014 | 471 | 118 | Computer
Z-0292-2014 | 65572 | No | LDP | Acetaminophen test system. | Toxicology | Toxicology | 510(k) | VITROS Chemistry Products ACET Slides, 90 slides/pack; catalog number 1037274 |  |  | June 12, 2013 | 2013 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics is recalling Slide Dispense cartridges used with some VITROS Chemistry Product Slides due to slide dispense issues. | Equipment maintenance | Ortho Clinical Diagnostics sent an Important Product Correction Notification letter dated June 12, 2013, to all affected customers via FedEx regarding Slide Dispense issues with various VITROS Chemistry Product Slides. VITROS ALT, CL-, GGT, LDH, K+, Na+ and TP slides use a different ABP, they are not affected by this issue. OCD provided instructions that may allow their customers to reposition the ABP if a jam occurs. OCD will provide you with a credit for any cartridge where the ABP becomes misaligned within the cartridge body in such a way that you are unable to reposition the ABP successfully to resolve the slide dispense issue. In addition to the Confirmation of Receipt form, OCD also enclosed three (3) copies of a Request for Credit form for you to complete when you require credit for affected cartridges. Complete and return the Confirmation of Receipt by June 21, 2013. If you experience slide dispense condition codes more than once a week, call our Customer Technical Support at 1-800-421-3311.For questions regarding this recall call 585-453-4224. |  | 16732 | 16732 units (12208 units domestically &amp; 4524 internationally) | Worldwide Distribution - USA (nationwide) including AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY including Puerto Rico and Internationally to Australia, Brazil, Canada, Chile, China, Colombia, India, Japan, Mexico, Panama, Singapore, Venezuela, England, Germany, France, Italy, and Spain. | Terminated | April 11, 2017 | 1399 | 16732 | Computer
Z-0408-2014 | 66412 | No | HTW | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Slipped Capital Femoral Epiphysis System |  |  | July 31, 2013 | 2013 | Synthes USA HQ, Inc. | There is a possibility for the Guide Wires of the Slipped Capital Femoral Epiphysis System to stick in the cannulation of the drill bit and the drill bits to break during surgery. | Device Design | Synthes sent an Urgent Notice: Medical Device Recall letter dated July 30, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to review their inventory, immediately remove the affected product and return to Credit/Returns, Synthes. For questions call 610-719-5450. |  | 166 | 166 | USA Nationwide Distribution | Terminated | June 01, 2015 | 670 | 166 | Computer
Z-0425-2014 | 65356 | No | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | 13mm addEASE(TM) Binary Connector with 17 Ga. Needle for use with PAB(R) Containers, BC1300 |  |  | April 04, 2013 | 2013 | B. Braun Medical, Inc. | Possible solution leakage at the interface of the addEASE(TM) Binary Connector (Product Codes BC2000 and BC1300) and the PAB(R) bag. Based on extensive root cause investigation, it has been determined that the leakage is the result of a misaligned cannula (needle) on the addEASE(TM) Binary Connector, which pierces the PAB(R) bag stopper to access the solution. The misaligned cannula may result i | Equipment maintenance | B Braun sent an Urgent - Medical Device Removal Notification letter dated April 4, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by he customer. Customers were instructed to review the Product Recall Notification in its entirety and ensure that all users in their organization are informed about this voluntary product recall. Determine their current inventory of the affected lots within their distribution center. Contact their customers that have received the affected lots and inform them of this voluntary product recall. Combine both their inventory, and the inventory reported by their customers, and utilize their specific "Product Removal Acknowledgement" form to record the total number of individual units (within partial cases) and the number of full-unopened cases. If no inventory remaining enter zero (0). Return the completed Acknowledgement Form to BBMI Quality Assurance in the self addressed envelope or fax to 610-849-1197. Customers should call Customer Support Department at 800-227-2862 to arrange for eturn and replacement product or for questions.For questions regarding this recall call 800-227-2862. |  | 4000 | 4,000 units | Nationwide Distribution | Terminated | August 26, 2014 | 509 | 1750079 | Computer
Z-0330-2014 | 64806 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | The Centricity Laboratory System is intended to be an information system designed to support the clinical and administrative activities associated with the provision and utilization of clinical laboratory services and facilities, e.g., the storing and delivering of analytical results. It is a specially designed data program application (software), which is supplied for installation in existing mainframe or decentralized computers or a computer network. This software system will not be the software directly used to run, steer, or control any specific laboratory analyzer or equipment. It is intended for use in laboratory facilities, including a central laboratory or in a multiple laboratory environment servicing satellite laboratories, reference laboratories, clinics, etc. Centricity Laboratory System is not intended for direct patient contact. |  |  | March 12, 2013 | 2013 | GE Healthcare It | GE Healthcare has become aware of a potential safety issue associated with the use of GE Centricity Laboratory where panel reflex does not occur if DIP.Specimen is blank. Panel reflexing fails to occur when ad-hoc re-routing functionality is used and the following are true: 1) The ordered panel is a slave panel, and; 2) The panitem specimen group field of one or more of the ordered items on the ma | Software design | GE Healtchare sent an Urgent Medical Device Correction letter dated March 12, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The letter requested that customers review their workflows to determine whether this series of events could happen at their sites. Customers who self-identified this series of events as a possible workflow are recommended to contact their GE service representative for a software correction. These customers may also elect to identify master panels that may be re-routed ad hoc and then to ensure that the specimen group field is updated to NULL if currently set to blank.Customers with any questions about this issue are advised to contact their GE service representative or the GE Healthcare IT Help Desk: North America at tgg.nahelpdesk@med.ge.com.For questions regarding this recall call 847-277-4468. |  | 7 | 7 units | Worldwide Distribution - USA (nationwide) including GA and Internationally to Canada and the United Kingdom. | Open, Classified |  |  | 7 | Computer
Z-0742-2013 | 63800 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Mindray DPM5 Monitor, a vital signs monitor used on human patients in healthcare facilities. |  |  | September 05, 2012 | 2012 | Mindray DS USA, Inc. d.b.a. Mindray North America | Mindray has identified an issue with the DPM5 monitors where ST and Arrhythmia features may not be available. | Device Design | Mindray DS USA Inc. sent an Urgent Product Corrective Action letter on September 5, 2012 via Fed Ex with delivery confirmation to all users. The letter identifies the affected product, problem and actions to be taken. Customers were instructed to contact their Mindray Service Representative at 1-800-288-2121 Monday through Friday 8:30-5:30 pm, to schedule a software upgrade. |  | 13 | 13 units | US Nationwide Distribution including the states of: VA, TN, FL, &amp; TX | Terminated | October 22, 2013 | 412 | 13 | Computer
Z-0333-2014 | 66635 | No | HJO | AC-powered slitlamp biomicroscope. | Ophthalmic | Ophthalmic | 510(k) | Zeiss brand IOLMaster, Version 1-5, Biomicroscope, Slit-Lamp, AC-Powered, |  |  | October 23, 2013 | 2013 | Carl Zeiss Meditec, Inc. | The user database delivered for download contained an erroneous value. Specifically the "aO" constant of the lens AMO Sensar AR40E is listed as "0.0" in the internal lens database of the IOLMaster. The correct value is -2.420. This could lead to the implantation of lens with a wrong diopter value. | Software in the Use Environment | The firm, Zeiss, sent an "Urgent Medical Device Correction" letter dated October 23, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customers. Customers were instructed to immediately update the Laser Interference Biometry (ULIB) lens constants, provided as an Attachment 1 and complete and return Attachment 2- Confirmation Form via mail to: Stericycle, Inc., 6026 Lakeside Blvd., Indianapolis, IN 46278; Attn: Zeiss IOLMaster.If you have any questions, call the following number for assistance: 1 (866) 891-1922. |  | 4907 | 4,907 units | Nationwide distribution. | Terminated | April 17, 2014 | 176 | 6108 | Computer
Z-0430-2014 | 65408 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension Vista(R) 500 and Dimension Vista(R) 1500 - Reloaded Vista Vials - limited to systems using software versions 3.4 and 3.5.1. |  |  | May 20, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | There is the potential for the vial volume of "in use" Vista calibrators, controls, or sample diluent to incorrectly revert to full volume when removed and reloaded. Dimension Vista(R) 500 and Dimension Vista(R) 1500 - Reloaded Vista Vials - limited to systems using software versions 3.4 and 3.5.1 | Software design | A notification letter was sent to consignees in May 2013. |  | 776 | 776 | Worldwide Distribution: USA (nationwide) and iinternationally to Canada. | Terminated | April 10, 2014 | 325 | 776 | Computer
Z-0431-2014 | 66705 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Easy Diagnost Eleva |  |  | October 30, 2013 | 2013 | Philips Medical Systems, Inc. | The delivered application software was an outdated version to the model of equipment. | Software design | Philips Medical Systems sent an Urgent - Field Safety Notice dated October 30, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Cusromers were informed that the delivered application software was an outdated version to the model of their expipment and that a Philips Field Engineer will be implementing an upgrade of the application software. Customers were instructed to review the information contained in the letter with all memebers of their staff who need to be aware of its contents. Customers with questions were instructed to contact their Philips representataive: Technical Support Line 1-800-722-9377.For questions regarding this recall call 1-800-722-9377. |  | 19 | 19 devices | Nationwide Distribution including MN, NH, DE, OR, MA, TX, OH, FL, CA, WA, and SC. | Terminated | January 09, 2017 | 1167 | 19 | Computer
Z-0791-2013 | 64160 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Medacta Quadra Trial Broach size 0 part # 01.10.10.045, for use with the Medacta Total Hip Prosthesis System. |  |  | December 17, 2012 | 2012 | Medacta Usa Inc | The tip of the Quadra Trial Broach size 0, broke during a surgery. | Device Design | Medacta sent a Urgent Safety Notification letter dated December 13, 2012 to all affected customers. The letter identified the recalled product, problem and actions to be taken. Customers were instructed to check inventory for affected product and contact QA@medacta.us.com for an RMA# to return items to Medacta USA. A new version of the recalled product will be provided to all affected customers. For questions call 805-437-7085 x26. |  | 26 | 26 pieces U.S.A. (226 pieces Worldwide) | US Nationwide Distribution including the states of Arizona, California, Idaho, Illinois, North Carolina, Ohio, Utah, Washington and US Virgin Islands | Terminated | February 07, 2013 | 52 | 26 | Computer
Z-0438-2014 | 66724 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | SIEMENS brand ARTISTE Linear Accelerator Linac systems with component: any photon unflat beam option (Multiple X feature) in combination with the IMRT or mARC option. A family of linear accelerator systems is to deliver X-ray photon and electron radiation for the therapeutic treatment of cancer. The ARTISTE/ONCOR systems are manufactured by: SIEMENS AG, MEDICAL SOLUTIONS, GERMANY and distributed SIEMENS Healthcare IM CR RO, CR Radiation Oncology, Martinez, CA |  |  | October 23, 2013 | 2013 | Siemens Medical Solutions USA, Inc | A safety risk regarding a potentially existing dark current radiation phenomenon on the Linear Accelerator (LINAC) in combination with IMRT or mARC (rotational IMRT) treatments using unflat (flattening filter free) beams. In case there are many long periods of field shape changes in a treatment plan, the dose due to dark current radiation may sum up to a significant value. | Device Design | Siemens sent an Urgent Medical Device Correction letter dated October 23, 2013 including the release of update instruction TH024/13/S to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. |  | 26 | 26 units | Worldwide Distribution: USA (nationwide) including states of: FL, NC, NY, PA, TN, UT, and WI; and countries of: Angola, Brazil, Canada, P.R. China, Germany, India, Japan, Republic Korea, Poland, and Thailand. | Terminated | January 22, 2014 | 91 | 29 | Computer
Z-0335-2014 | 66601 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance iCT Computed Tomography X-Ray System, |  |  | October 11, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | If the couch, with the footrest extension attached, is positioned where the table travel is within the bore of the gantry, and the user begins to lower the couch, the footrest extension can contact the gantry cover or come in contact with the user due to the length of the accessory. | Device Design | Philips sent an Urgent - Medical Device Correction letter dated October. 11, 2013, to their customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to discontinue the use of the footrest extension. Philips also, stated they will replace the old footrest extension with the current one in use. For questions contact your local Philips representative or local Philips Healthcare office. |  | 9 | 9 units | Worldwide distribution: USA (nationwide) Distribution in the states of DC, GA, IL, IN, NY and OH and the countries: Germany, Israel and Netherlands. | Terminated | August 04, 2015 | 662 | 9 | Computer
Z-0439-2014 | 66764 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | C-Series: Clinac, Trilogy, Trilogy Tx. , Novalis high energy linear accelerators and UNIQUE single energy linear accelerator; Versions 7, 8 and 9. The UNIQUE is not sold in the US. |  |  | October 24, 2013 | 2013 | Varian Medical Systems, Inc. | This correction is to notify users that a solution to a previous correction has been developed and Varians reps will be contacting locations to schedule installation. | Device Design | Varian sent an Urgent Medical Device Correction letter, P/N CP-05645 Rev D dated October 24, 2013 to their affected customers. The letter identified the affected product, problem and actions to be taken. The letter stated all affected sites will be contacted by Varian representatives to schedule an upgrade to their devices. Questions may be directed to 888-827-4265 for USA and Canada. Europe +41 41 749 8844 |  | 3809 | 3,809 devices | Worldwide Distribution - USA Nationwide and in the countries of Algeria, Angola, Argentina, Australia, Austria, Azerbaijan, Bangladesh, Belgium, Belorussia, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, Estonia, Finland, France, French Polynesia, Georgia, Germany, Greece, Guadeloupe, Guam, Guatemala, Hong Kong, Hungary, India, Indonesia, Iraq, Ireland, Israel, Italy, Japan, Kazakhstan, Kenya, Kuwait, Latvia, Lebanon, Uthuania, luxembourg, Macau, Macedonia, Malaysia, Martinique, Mauritania, Mexico, Moldova, Monaco, Morocco, Netherlands, New Zealand, Norway, Oman, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Puerto Rico, Qatar, Reunion, Romania, Russia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, SriLanka, Sweden, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Venezuela, Vietnam, and Zimbabwe. | Terminated | January 13, 2016 | 811 | 3809 | Computer
Z-0436-2014 | 66866 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Assembly, Patient Side Manipulator on da Vinci Surgical Systems: |  |  | November 11, 2013 | 2013 | Intuitive Surgical, Inc. | Reports of friction within certain instrument arms can interrupt smooth instrument motion. This can be felt by the surgeon as resistance in the movement of the master.In this situation, the instrument can stall momentarily and then suddenly catch-up to the master position if the surgeon pushes through the resistance. | Device Design | The firm, Intuitive Surgical, sent an "Urgent Medical Device Recall" letter dated November 11, 2013 all affected customers via traceable mail. The letter described the product, problem and actions to be taken. The customers were instructed to take the following actions: 1. Ensure all appropriate personnel are fully informed of this Notification. Forward this letter to all Surgeons and OR staff who perform da Vinci surgical procedures, Risk Manager, OR Director, Purchasing, and Biomedical Engineering staff. 2. Complete the attached Acknowledgement Form and return it to Intuitive Surgical via Fax to: Intuitive Surgical, Inc., ATTN: REGULATORY AFFAIRSU.S. +1 (408) 716-3040, or Scan and Email: isi.compliance@intusurg.com.3. An Intuitive Surgical Representative will contact your site to schedule the inspection of the instrument arms on your system(s) as indicated above. Please work with the Intuitive Surgical Representative to schedule this activity.4. Please inform appropriate personnel when the correction has been completed.5. Please retain a copy of this letter and the Acknowledgement Form for your files.If you need further information or support concerning this medical device correction, then please contact your Clinical Sales Representative or contact Intuitive Surgical Customer Service at the numbers listed below. Âż North and South America: 800-876-1310 Option 3 (6 am to 5pm PST) Âż Japan: 0120-56-5635 or 03-5575-1362 (9am to 6pm JST) Âż South Korea: 02-3271-3200 (9am to 6pm KSTJ) Âż Europe, Middle East, Asia and Africa: +800 0821 2020 or +41 21 821 2020 (8 am to 6 pm CET). |  | 1386 | 1,386 | Worldwide Distribution: US (nationwide) and to countries of: Australia, Austria, Belgium, Brazil, Bulgaria, Canada, China, Columbia, Denmark, France, Germany, India, Israel, Italy, Japan, S. Korea, Luxembourg, Mexico, Monaco, Netherlands, Norway, Romania. Russian Federation, Singapore, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, and United Kingdom. | Terminated | February 18, 2014 | 99 | 1386 | Computer
Z-0434-2014 | 66502 | No | FRG | Sterilization wrap. | General Hospital | General Hospital | 510(k) | STERITITE Containers: STERITITE Container Kit, 1/4 size, 3" high, product code: 13902; STERITITE Mini Mid Size, 3" high, product code: 13903; STERITITE Container Kit, 1/2 size, 4" high, product code: 13904; STERITITE Container Kit, 3/4 size, 4" high, product code: 13905; STERITITE Container Kit, Full size, 4" high, product code: 13906. |  |  | October 03, 2013 | 2013 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) is recalling the SteriTite containers because they do not have consistent data to show that sterility assurance level of 10(-6) is achieved during the half-cycle test runs on new, unprocessed containers. | Device Design | Advanced Sterilization Products (ASP) sent a Urgent Medical Device Field Safety Notification letter dated October 3, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers who received the SteriTite Containers.Customers were instructed to:Provide this notice to anyone in your facility that needs to be informed.Maintain a copy of this notice with the affected product.Maintain awareness of this notice.Customers are instructed to contact Stericycle at (888) 912-7092 for additional questions related to the recall letter. Customers are instructed to report any complaints or suspected problems related to the recall by contacting ASP at (888) STERRAD (888-783-7723). |  | 267 | 267 units total (252 units in US) | Worldwide Distribution - USA (nationwide) and the countries of Australia, Vietnam, Thailand, Malaysia, China, Korea (South), Mexico, Spain, Singapore, Hong Kong and Portugal. | Terminated | January 13, 2014 | 102 | 267 | Computer
Z-0274-2014 | 65760 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | S-LIFT Extension Shims |  |  | August 28, 2013 | 2013 | SpineFrontier, Inc. | Fit variation among all lots of S-LIFT extension shims with corresponding retractor blades. The firm's investigation revealed that extension shims may not fit down all retractor blades. | Device Design | SpineFrontier sent an Urgent Voluntary Recall letter dated August 28, 2013, to all affected customers. The letter identified teh product, the problem, and the action to be taken by the customer. Customers were asked to initiate the immediate return of the parts noted and complete an acknowledgement form and return it to SpineFrontier via fax or email. Customers with questions were instructed to call 978-232-3990. For questions regarding this recall call 978-232-3990. |  | 67 | 67 | Nationwide Distribution including FL, TX, OK, MO, IL, OH, VA, OR, CO, KS, CT | Open, Classified |  |  | 67 | Computer
Z-0433-2014 | 65147 | No | NSX |  |  | General Hospital | Contact ODE | syngo(R) Lab Data Manager - Version VA11B and VA12A Systems |  |  | May 10, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Under certain conditions the system may not perform as intended, causing the release of results to the Laboratory Information System that should have been held for manual review due to auto-verification rules or the delay/omission of result transmission to the LIS. | Software design | Siemens sent an Urgent Medical Device Correction letter in May 2013 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the letter in its entirety and ensure that all operators understand the information presented. Specific actions needed to be implemented by the customer were outlined in the letter. Customers were instructed to retain the letter with their laboratory records, and forward the letter to those who may have received the affected product. Customers with questions were instructed to contact their Siemens Technical Solutions Center or their local Siemens technical support representative.For questions regarding this recall call 800-441-9250. |  | 262 | 262 | Worldwide Distribution: USA (nationwide) and internationally to Canada. | Terminated | August 26, 2014 | 473 | 262 | Computer
Z-0172-2014 | 65795 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Artis zee systems |  |  | January 21, 2013 | 2013 | Siemens Medical Solutions USA, Inc | The firm became aware of a potential issue with Artis zee systems bearing model numbers 10094135, 10094137, and 10094141 equipped with a card collimator from a specific lot where a dislocation of the scatter radiation filter inside the card collimator may occur. | Device Design | Siemens sent an letter to all affected customers in January 2013, notifying them of the potential issue and urging customers to make appointments with Siemens to have software update AX 027/12/S realized in order to resolve the problem. The software update was made available from January 22, 2013, on. Customers were also asked to pass along this information if the device was no longer in their possession.The Center for Devices and Radiological Health (CDRH) understands that the CAP involving 17 unit(s) will consist of the following:1. The firm sent out a customer notification letter in January 2013 notifying them of the issue and highlighting the approved procedures of the Corrective Action Plan.2. Siemens directed in the customer notification letter, for the end-user to make an appointment with the Customer Service organization to have safety update AX 027/12/S realized.3. As stated in the safety update AX 027/12/S and in the e-mail response the update to correct this issue will be provided to the customer at no expense.4. Siemens indicated in the customer notification letter that the AX 027/12/S update will be made available from January 22, 2013 and on.5. The AX 027/12/S update provides general information, work instructions and a completion certificate on how to correct the dislocated scatter radiation filter inside the collimator.You may proceed with implementation of the CAP (subject to the conditions noted above). For further questions please call (610) 219-6300. |  | 17 | 17 | US Distribution including the states of AZ, KY, PA, IA WV, FL, MO, OK, CA, AR, MT, NY, KS, and CT. | Terminated | October 29, 2014 | 646 | 17 | Computer
Z-0140-2014 | 66314 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Siemens RAD Fluoro Uro system (Luminos dRF and Uroskop Omnia with software version VD10A/G |  |  | September 03, 2013 | 2013 | Siemens Medical Solutions USA, Inc | When using systems operating with software versions VD10A/G during a RAD examination, a malfunction can result in radiation continuing longer than required resulting in an overexposed image that is not of diagnostic quality, prompting repetition of the examination. | Software design | Siemens sent an Field Safety Notice dated September 3, 2013, to all affected customers.The letter identified the product the problem and the action to be taken by the customer. Action to be taken: 1. It is not possible for the user to prevent this malfunction. The exposure time with the use of automatic exposure control is limited to 500ms, which is the maximum duration of radiation in this mode. The exposure can be interrupted by the user at anytime by releasing the exposure button.2. There are two possibilities to solve this issue:a.) Switch off the system if the message "wait for FLC" in status bar does not disappear approximately. 30s after the OGP selection. The system gets ready after switching the system on and reselecting the Torno OGP.b.) Press Shift+F10. The pop up message has to be confirmed and you have to wait approximately.40s until! the "i" icon in the status bar stops blinking. After that a RAD OGP has to be selected and then the Torno OGP. Thereafter, the system gets ready.3. It is not possible for the user to prevent this malfunctionThe Center for Devices and Radiological Health (CDRH) understands that yourCorrective Action Plan (CAP) will involve 59 units and consist of the following elements:1. You have notified affected customers with Field Safety Notice XP032/13/S.2. You have prepared a field modification of the Luminos dRF and Uroskop Omniasystems with software version VD10A/G.3. Distribution of this field modification will began in September 2013 as UpdateInstruction XP015/13/S at no cost to the user.For further Customer Support call (610) 219-6300. |  | 59 | 59 | Worldwide Distribution - USA (nationwide) and Internationally to Canada. | Terminated | March 09, 2015 | 552 | 59 | Computer
Z-0346-2014 | 66069 | No | JGS | Sodium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA 2400 System Software V4.01. |  |  | September 03, 2013 | 2013 | Siemens Healthcare Diagnostics | Siemens Healthcare Diagnostics issued a Urgent Medical Device Correction due to Ration Parameters Flagging Behavior for All Software Versions utilized on ADVIA 1200, 1650, 1800 or 2400 Chemistry Systems. | Software design | On September 3, 2013, Siemens Healthcare Diagnostics issued an Urgent Field Safety Notice for the ADVIA 1200, 1650, 1800 and 2400 Chemistry Systems due to a software issue. The software issue has caused the Chemistry Systems to generate reports without error flags. Notices were sent to domestic customers via U.S. and outside the US the distribution is determined at the country level. The customers were instructed to contact Siemens Healthcare Diagnostics Technical Solution Center if they have technical questions regarding this issue at 877-229-3711 (prompt 18). |  | 876 | 876 units | Worldwide Distribution-USA (nationwide) and the countries of Sweden, Switzerland, Taiwan, Turkey, United Kingdom, and Vietnam. | Open, Classified |  |  | 3759 | Computer
Z-0659-2013 | 63559 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | The TM Ardis Interbody System Inserter Stainless Steel, Rubber, Silicone Non Sterile Rx only. |  |  | September 20, 2012 | 2012 | Zimmer, Inc. | Zimmer Spine is has received complaints of difficulty in turning the implant release knob of the TM Ardis Inserter after final positioning of the TM Ardis interbody spacer implant. There has been two reports where the inserter instrument was inadvertently removed from patient with the implant still engaged. The correction includes revised directions explaining the technique to remove the implant | Device Design | Zimmer sent a "MEDICAL DEVICE CORRECTION" letter dated September 26, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. A Risk Manager Confirmation Form was attached for customers to complete and return. Questions and concerns should be addressed to Zimmer Spine Customer Service at 866-774-6368.When re-design is approved, Zimmer will remove and replace all affected devices. |  | 170 | 170 | USA nationwide; Europe, Sweden | Terminated | April 23, 2014 | 580 | 170 | Computer
Z-0337-2014 | 64805 | No | HGM | Perinatal monitoring system and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Centricity Perinatal and Centricity Intensive Care is intended to be used in clinical departments healthcare delivery systems such as a Labor and Delivery, Postpartum Maternal Care, newborn Nursery, critical care units including Neonatal Intensive Care Units (NICU), and may also be used in physicians offices and outpatient clinics. This product is primarily intended to serve the purpose of electronic documentation of clinical data and is designed to accept, transfer, display, calculate, store and manage clinical data. System capabilities provide the user with the ability to acquire data from the medical devices and to document, annotate, display, store, print, export and retrieve patient clinical practice information. This product is intended for professional use only. |  |  | March 11, 2013 | 2013 | GE Healthcare It | 1) Under rare network conditions, a Centricity Perinatal process may maintain an incorrect count of Patient slots in use which may lead to an inability to access the patient roster. The user is temporarily unable to review or edit form and chart information. This could result in a potential delay of care; however active fetal strips and permanent patient data storage are not affected. 2) When | Software design | GE Healthcare sent an URGENT MEDICAL DEVICE CORRECTION letter dated March 11, 2013, to all affected customers. The letters requested that direct accounts ensure that all potential users of the software were made aware of the potential problems and the recommended actions. Direct accounts were asked to return a customer response form and to return it to GE Healthcare via fax at 847-939-1522 or e-mail at QARAFMl1@ge.com.Version 6.94 of the software has been modified to correct these defects.GE Technical Support: 800-433-2009. |  | 1132 | 1,132 units | Worldwide Distribution - USA (nationwide) including Puerto Rico and the countries of: Australia, Canada, Denmark, Germany, Hong Kong, Korea, South Korea, Thailand and United Kingdom. | Open, Classified |  |  | 1132 | Computer
Z-0370-2014 | 66056 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Medcare Products Max-500 Stand, Total Support, patient stand, model #400006. The responsible firm name listed in the Operations Manual is Medcare Products, Burnsville, MN. |  |  | October 30, 2013 | 2013 | Ergosafe Products LLC DBA Prism Medical Services USA | The lift might fail with the labeled 500-lb. load rating. | Device Design | The firm, Medcare Products, sent an "URGENT FIELD CORRECTIVE ACTIONS" letter dated October 21, 2013 via certified mail on October 30, 2013 to its customers. The letter described the product, problem and action to be taken. The customers were instructed to peel off the boom stickers and serial number label and replace with the new ones supplied; review the revised Manual and dispose of the original; and complete the action report form once the products have been relabeled. Please refer to your Operations Manual or visit our website at www.medcarelifts.com to view the manual online.If we can answer any questions that you may have regarding the testing standards or the relabeling of the products, please do not hesitate to contact us at 1-800-695-4479. |  | 97 | 97 units | Nationwide distribution. | Terminated | May 17, 2017 | 1295 | 330 | Computer
Z-0767-2013 | 64127 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Phillips Easy Diagnost Eleva DRF. Model Number: 706034 |  |  | January 16, 2013 | 2013 | Philips Healthcare Inc. | The mirror icon on the bottom of each image could be misinterpreted as a RIGHT lead marker for side indication if the image is mirrored within a PACS system. | Software design (manufacturing process) | Phillips issued an Urgent Field Safety Notice to all customers on January 14, 2013, to notify users that the Phillips EasyDiagnost Eleva systems, release 3.0 has a mirror icon that could be misinterpereted and could potentially lead to misdiagnosis. Philips will upgrade the system to version 3.1 which involves software and hardware upgrades of the system. This software provides a replacement icon that reduces the risk of the potential misinterpretation of the original mirror icon. Additionally, Philips will provide updated Instructions for Use (IFU) for the Eleva workspot and the EasyDiagnost 3.1 system and an addendum to the EasyDiagnost system IFU. Users will be contacted by Phillips to schedule the installation. For questions contact your local Philips representative: 1-800-722-9377. |  | 14 | 14 units | US Nationwide Distribution including the states of FL, CA, WA, AZ, NC, KY, DE, AZ and MO. | Open, Classified |  |  | 14 | Computer
Z-0385-2014 | 66605 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ACCU-CHEK Aviva Combo Blood Glucose Meter (as a part of the ACCU-CHEK Aviva Combo System)\ |  |  | October 25, 2013 | 2013 | Roche Diagnostics Operations, Inc. | There is a software synchronization issue with the ACCU-CHEK Combo system. In rare cases, when the Manual Pump option is chosen on the ACCU-CHEK Aviva Combo meter, there is a possibility of receiving an incorrect bolus advice recommendation that may cause a temporary under delivery of insulin. | Software design | The firm, Roche Diagnostics, sent an "URGENT MEDICAL DEVICE CORRECTION (UMDC)" letter dated October 29, 2013, via UPS to their consignees/customers. The letter described the product, problem and actions to be taken. The UMDC informed them of the software synchronization issue with the ACCU-CHEK Combo system that may cause a temporary under delivery of insulin. The customers were instructed to ensure they properly deliver the correction bolus when utilizing the Manual Pump option, as instructed in the Owners Booklet; and complete and return the Business Reply Card via fax to 1-877-282-0223. The Distributors were instructed to provide the UMDC to their direct customers. The HCPs were instructed to remind their patients to follow the instructions in the Accu-Chek Aviva Combo Meter Advanced Owners Booklet when utilizing Manual Pump option.Questions about the information contained in this UMDC are directed to ACCU-CHEK Pump Support, 24 hours a day, seven days a week at 1-800-688-4578. |  | 6595 | 6,595 units | Nationwide Distribution. | Terminated | January 14, 2015 | 446 | 6595 | Computer
Z-0368-2014 | 66662 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | EPIQ 7 Ultrasound System, Model 795200, 795201, 795202. |  |  | October 17, 2013 | 2013 | Philips Medical Systems North America Inc. | A problem has been detected with Philips Model EPIQ 7 Ultrasound System, software version 1.0 and 1.0.1. Using the International System Date Format may result in calculation errors in Gestational Age Gestational Age (GA) and Established Due Date (EDD). | Software design | Affected customers will be notified via the Urgent-Field Safety Notice (FSN79500286), Philips Model EPIQ 7 Ultrasound system, dated 2013 October. Customers are informed that using the International System Date Format in the Philips Model EPIQ 7 Ultrasound System, software version 1.0 and 1.0.1. may result in calculation errors in Gestational Age and Established Due Date (EDD). Philips will upgrade customer systems to software version 1.1, which corrects the issue, for a free of charge. |  | 306 | 306 units total (206 international; 100 in the US). | Worldwide Distribution-USA (nationwide) including the states of CT, HI, IL, IN, KY, MA, ME, MI, NC, OR, PA, TX, and WA, and the countries of Australia, Austria, Belgium, Chile, Denmark, Finland, France, Germany, India, Ireland, Republic of Korea, Latvia, Netherlands, New Caledonia, New Zealand, Norway, Poland, Portugal, Romania, Saudi Arabia, Slovakia, Spain, Sweden, Switzerland, and United Kingdom. | Terminated | July 16, 2014 | 272 | 306 | Computer
Z-0741-2013 | 64117 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | HAMILTON-T1 Ventilator with software versions 1.1.2 or lower. |  |  | December 19, 2012 | 2012 | Hamilton Medical Inc | Analysis of a customer complaint has shown that during ventilation of small pediatric patients with high airway resistance and low lung compliance, the oxygen consumption of a Hamilton T1 ventilator with software versions 1.1.2 and lower must be calculated using a larger margin than originally expected. | Software design | Hamilton sent a Field Safety alert letter dated January 11, 2013 to all customers. The letter identified the affected product, explained the potential problem and a included Medical Device Safety Alert and Corrective Action document. Hamilton is working on a new software version for this issue. Customers are asked to update their operator manual with the update oxygen consumption formula page which provided revised information for calculating oxygen capacity. When the new software is available, Hamilton will provide loaner devices to each location while their unit is being upgraded.Questions should be addressed to Hamilton Medical at 800-426-6331 ext 215 |  | 6693 | 97 total ventilators; 69 of part number 161005 and 28 of part number 161006 | US Nationwide Distribution including the country of Switzerland. | Terminated | April 22, 2014 | 489 | 6693 | Computer
Z-0359-2014 | 66300 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | BIOMET Spine, 5.5 TI 5.5x50mm Translation Screw Polaris 5.5 TI Spinal System. Rx Only. Non-Sterile. |  |  | August 22, 2013 | 2013 | Ebi, Llc | Polaris Translation Screw Failures were recently reported to Biomet Spine. These failures are generally described as dissociation of the screw head from the shaft of the Polaris Translation Screw. | Device Design | Urgent Product Information notices dated 8/22/2013 were mailed via FedEx to distributors on 8/23/2013. Follow-up emails with a copy of the UPI notice were sent to distributors on 8/23/2013. The UPI notice informed the customers of the issue with the product and how to identify the affected product. Customers were advised that the Polaris Translation Screw System and the SpF Implantable Spine Fusion Stimulator not be used in conjunction with any implants containing cobalt chrome. Firm contacts were provided in case customers had any questions. EBI, LLC had a conference call with their domestic distributors on 8/23/2013 informing them they are required to notify their customers. |  | 7092 | 7,092 units | Nationwide distribution including Puerto Rico. | Terminated | August 31, 2016 | 1105 | 7092 | Computer
Z-0363-2014 | 66137 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Model FL14E Rose |  |  | October 23, 2013 | 2013 | Stryker Medical Division of Stryker Corporation | It was reported that the siderail may appear latched when it is not, due to lack of proper preventive maintenance. This might result in bruises and/or lacerations to a patient who is leaning against the siderails if the siderail releases. Additionally, the siderail may not effectively constrain a patient to prevent them from rolling or falling off the bed. | Equipment maintenance | Stryker Medical sent a Medical Device Notification letter dated October 23, 2013, to all affected consignees. The letter described the recalled product, potential injuries associated with the side rail sticking, as well as providing preventive measures to be taken to mitigate the risk. Consignees were directed to locate the beds listed in this notice, ensure that the siderails have been inspected and adjusted, as required, continue to routinely inspect the siderails and functionality using the enclosed, enhanced product maintenance information. Problems and concerns should be addressed to Stryker at (1-800-327-0770, option 2). Consignees were asked to Fax (269-488-8691) or e-mail (ProductFieldAction@stryker.com ) the enclosed return form and directed to forward the notice if product was further distributed and advise Stryker of the beds new location.For questions regarding this recall call 269-324-6609. |  | 24575 | 4126 units USA; 20449 inits Internationally | Worldwide Distribution - USA (nationwide) and Internationally to Canada; Japan, South America, Latin America and China. | Terminated | May 16, 2014 | 205 | 24575 | Computer
Z-0364-2014 | 66639 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Juggerknotless Drill Bit, REF 110003173, |  |  | October 10, 2013 | 2013 | Biomet, Inc. | Investigation identified that the design of the drill bit may make the suture anchor difficult to insert into the bone. The implant may skip across the hole or miss the hole during insertion. | Device Design | Biomet sent an Urgent Medical Device Recall Notice letter dated October 10, 2013 to all consignees. The letter identified the affected product, reason for recall, and actions to be taken. The letter stated customers will be contacted on an individual basis in an attempt to reconcile the affected product. Customers were directed to locate and remove implicated devices and fax a copy of the response form to 574-372-1683 prior to returning product. For questions call 574-372-1570. |  | 25 | 25 | USA Nationwide Distribution in the states of VA, CA, Ky, and IN | Terminated | July 16, 2014 | 279 | 25 | Computer
Z-0361-2014 | 66534 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Aesculap Arcadius XP L Trial Insertion Instrument (ME020R/ME020R-US) |  |  | September 20, 2013 | 2013 | Aesculap, Inc. | The trial insertion instrument may disassemble during use due to a potentially weak weld on then trial inserter shaft block area located near the thumb wheel. | Device Design | An Important Product Removal Notification dated September 20, 2013 was sent to customers. |  | 30 | 30 | Worldwide distribution: USA (nationwide) GA, TN, CA, WY, NV, TX, OK, OH, and IL; and internationally to Canada. | Terminated | December 16, 2014 | 452 | 30 | Computer
Z-0388-2014 | 66420 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Synthes Tensioning Device |  |  | July 31, 2013 | 2013 | Synthes USA HQ, Inc. | All lots of the Synthes Tensioning Device were recalled due to the potential for the reaction fork to come loose from the main body of the device and not allow the wire to release, nor the device to apply tension. | Device Design | Synthes sent an Urgent: Medical Device Recall letter dated July 31, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer."Please remove this product from your inventory and return it to Synthes. A replacement product will be sent to you when available. If you DO have any of the identified devices, please take the following steps:"Call Synthes at 1-800-479-6329 to obtain a Return Authorization Number."Complete the Verification Section at the end of this letter by checking the appropriate box indicating affected product has been located. Also, please indicate the number of devices found and note the Return Authorization Number. Please include your name, title, telephone number and signature in the spaces provided."Return the Verification Form (page 3 of this letter) with the product to:" Credit/Returns, Synthes, 1101 Synthes Avenue, Monument, CO 80132.If you DO NOT have the identified product, please take the following steps:"Complete the attached Verification Section at the end of this letter by checking the appropriate box indicating that no affected product has been located. Please include your name, title, telephone number and signature in the spaces provided. This return documentation acknowledges your receipt of medical device removal information."Return the document to Synthes by: Fax: 877-272-3517"Email: FieldAction@synthes.comIf you have any questions, please call 610-719-5450 or email Fieldaction@synthes.comWe apologize for any inconvenience this may cause and thank you in advance for your attention and cooperation. |  | 369 | 369 | Nationwide Distribution. | Terminated | June 01, 2015 | 670 | 369 | Computer
Z-0108-2014 | 66166 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Siemens RAPIDLabÂż 1265 system blood gas analyzer |  |  | September 17, 2012 | 2012 | Siemens Healthcare Diagnostics Inc | RAPIDPointÂż 400 Series, RAPIDPointÂż 500, orRAPIDLabÂż 1200 Series blood gas analyzer, users can inadvertently enter a wrong patient ID on the demographics screen, which ultimately could lead to a misidentification of the sample | Software Manufacturing/Software Deployment | Siemens Healthcare sent an Urgent Field Safety Notification letter dated September 2013 via Federal Express, to affected Siemens domestic customers. Effected Siemens customers outside the US was provided a copy of the Urgent Field Safety Notice via hard copy, e-mail, and/or fax as determined by each countries local regulations and procedures. Previously Siemens Healthcare had issued a Customer Notification on August 23, 2012. The letter identified the affected product, issue, information to mitigate the issue and actions to be taken. The firm will issue an upgrades software to users. For questions contact your local Customer Care Center - Technical Support Representative at 1-877-229-3711. |  | 2939 | 2939 units | Worldwide Distribution - USA Nationwide and in the countries of Algeria, Argentina, Australia, Austria, Bahrain, Bangladesh, Belgium, Bosnia Herzeg., Botswana, Brazil, Brunei, Bulgaria, Chile, Colombia, Croatia, Czech Republic, Denmark, Egypt, Estonia, Finland, France, Fren.Polynesia, Germany, Greece, Hong Kong, Hungary, Iceland, India, Indonesia, Iraq, Ireland, Israel, Italy, Ivory Coast, Japan, Kazakhstan, Kenya, Kuwait, Latvia, Lithuania, Luxembourg, Macedonia, Malaysia, Mayotte, Netherlands, New Caledonia, New Zealand, Norway, P.R. China, Philippines, Poland, Portugal, Qatar, Rep. of Yemen, Republic Korea, Romania, Russian Fed., Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, Turkmenistan, U.A.E., United Kingdom, Vatikancity, Venezuela, Vietnam, | Terminated | April 01, 2015 | 926 | 11006 | Computer
Z-0362-2014 | 66609 | component, external, limb, ankle/foot | ISH | External limb prosthetic component. | Physical Medicine | Physical Medicine | 510(K) Exempt | Fixed Offset Adapter, A-642100. |  |  | October 24, 2013 | 2013 | Ossur North America Inc | Ossur inititated a voluntary recall of the A-642100 Fixed Offset Adapter because it can develop fatigue cracks in normal use over an extended period of time and ultimately separate completely. | Device Design | Ossur sent an "Safety Notice Voluntary Product Recall" letter dated October 29, 2013, via regular mail. Follow-up telephones are also made. The letter identified the product the problem and the actions to be taken by the customer.Âżssur is committed to providing superior customer service and we will do our utmost to minimize any inconvenience this may cause your practice. To ensure patient safety we kindly ask you to take the following action:" Please identify the users of each of the A-642100 adapters, see attached list." Also, please check your inventory for any A-642100s you may have on hand but have not fitted on users. A follow-up phone call will be made a few days later to coordinate replacements. A new replacement component will be provided to the customer. For further questions please call (949) 360-3634. |  | 66 | 66 units | US Distribution including the states of AL, TX, NY, CA, VA, GA, WI, NC, ID and NJ. | Terminated | June 17, 2014 | 236 | 66 | Computer
Z-0750-2013 | 63875 | No | OAF |  |  | Ear Nose &amp; Throat | PMA | Esteem Programmer Model 6001, US and ISA, Model 3001 Dell E54XX, part of the Esteem Totally Implantable Hearing System. |  |  | January 09, 2013 | 2013 | Envoy Medical Corporation | Envoy Medical is conducting a voluntary correction of a limited number of Esteem Programmers, part of the Esteem Totally Implantable Hearing System, to reduce their susceptibility to noise interference. | Software design | Envoy Medical Corporation sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated December 7, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at 949-233-1204 for questions relating to this notice. |  | 44 | 44 | Worldwide Distribution-USA including the states of AL, AZ, CA, CT, FL, GA, IL, MA, MI, MN, NC, OK, RI, and TX, and the country of GERMANY. | Terminated | April 09, 2013 | 90 | 44 | Computer
Z-0374-2014 | 66231 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | (Siemens Rad Fluoro Uro Systems) Ysio, Luminos Agile, Luminos dRF, and Uroskop Omnia systems with software version VC10 |  |  | September 03, 2013 | 2013 | Siemens Medical Solutions USA, Inc | A potential malfunction and hazard to patients exists when using the Ysio, Luminos Agile, Luminos dRF, and Uroskop Omnia systems with software version VC10, which may result in radiation continuing longer than required and overexposed images that are not of diagnostic quality. | Software design | Siemens Healthcare sent a Field Safety Notice dated September 3, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to forward the Field Safety Notice to any new users and provide the new owner information to Siemens Healthcare. Customers were asked to send back the attached form confirming receipt of the Field Safety Notice. Customers with questions should call 610-850-4634.For questions regarding this recall call 610-219-6300. |  | 357 | 357 | Worldwide Distribution - USA (nationwide) and internationally to Canada. | Terminated | February 24, 2015 | 539 | 357 | Computer
Z-0752-2013 | 63805 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Extremity Medical, IO FiX System Accessory. Catalog Number 101-00010. |  |  | March 22, 2012 | 2012 | Extremity Medical LLC | Lag screw measurement is short (approximately 5 mm) when using current screw measuring device (p/n 101-00010) per the IO FIX Surgical Technique Rev C. which can lead to possible inappropriate implant size selection (e.g. shorter screw size selected). | Device Design | Extremity Medical LLC sent e-mail notification dated March 22, 2012, to all affected customers. The email identified the product the problem and the action to be taken by the customers.The email includes a buleeting which informes the customer that an upgrade to the existing instrumentation is being released. The upgrade will replace the existing gauge. Updated pdf versions of the technique guide are also included. Customers are instructed to return the recalled device after swapping out the recalled device with the replacement unit. Customers will be contacted by Extremity and Sales Reps should contact their principal.For further questions please call (973) 588-8980. |  | 93 | 93 units | Worldwide Distribution -- Nationwide Distribution including the states of AR, AZ, CA, FL, GA, IL, IN, MD, MT, NC, NE, NH, NJ, NY, OH, PA, TN, TX, WI and WA. | Terminated | June 13, 2013 | 448 | 93 | Computer
Z-0754-2013 | 61611 | No | KPQ | Radiation therapy simulation system. | Radiology | Radiology | 510(k) | AdvantageSim MD versions 7.4 through 7.6, Model 5160092-2 Version 7.4, 5160092-3 Version 7.5, 5160092-4 version 7.6. |  |  | May 24, 2012 | 2012 | GE Healthcare, LLC | It was reported by a customer site that when using GE AdvantageSim MD on Advantage Workstation, the series could be incorrectly labeled in image view when multiple series of an exam are loaded simultaneously in Advantage Sim MD and if their series dates are different. A mismatch of series label for structure sets may lead to under-treatment of a tumor due to too small coverage of volume treated. | Software design | GE Healthcare sent an "Urgent Medical Device Correction" letter dated May 11,2012, to all consignees on May 24, 2012. The letter identified the product, the problem, and the action to be taken by the consignee. The letter addressed the Safety Issue, Safety Instructions, Affected Product,, Product Correction and Contact Information. GE indicated that a GE Healthcare service representative would contact they would correct all affected systems by providing a software upgrade at no cost to the consignee. For questions customers were instructed to call one of the following numers:USA 800-437-1171Japan 0120 - 055 - 919For questions regarding this recall call 262-513-4122. |  | 65 | 65 | Worldwide Distribution - USA including AL, CT, FL, IL, MA, MO, NJ, NY, PA, TX VA, WA, WI and Internationally to TURKEY, TAIWAN, SWITZERLAND, SWEDEN, SOUTH AFRICA, NIGERIA, MALAYSIA, KOREA, JAPAN, INDIA, GERMANY, FRANCE, EGYPT, DENMARK, CHINA, CHILE, CANDA, BRAZIL, AUSTRIA, and GREAT BRITAIN. | Terminated | January 16, 2014 | 602 | 65 | Computer
Z-0341-2014 | 66529 | No | GAD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Synthes Soft Tissue Retractor Small Extendible |  |  | September 24, 2013 | 2013 | Synthes USA HQ, Inc. | The potential exists for the locking screw nut to break on the Soft Tissue Retractor if it is overt-tightened. | Device Design | Synthes sent an Urgent Notice: Medical Device Recall letter dated July 31, 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review their inventory and immediately remove the affected lots from stock, call Synthes at 1-800-479-6329 to obtain a Return Authorization Number. Complete the Verification Section at the letter by checking the appropriate box indicating affected product has been located and return the Verification Form with the product to Credit/Returns, Synthes, 1101 Synthes Avenue, Monument, CO 80132. If customers do not have the identified product, they were instructed to complete the Offset blade facilitates easy preparation of the epipereosteal cavity for percutaneous plate insertion.e attached Verification Section and return to Synthes. |  | 374 | 374 | Worldwide Distribution - USA (Nationwide) and Canada. | Terminated | June 01, 2015 | 615 | 374 | Computer
Z-0348-2014 | 66773 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Candela Laser GentleLase Pro, Powered Laser Surgical Instrument for dermatological use - Model number: 9914-00-9015. |  |  | November 06, 2013 | 2013 | Candela Corporation | Unintended single pulse maybe emitted to handpiece prior to pressing of trigger or footswitch and cause injury. | Software design | Candela notiifed Distributors on 11/8/13 through e-mail, Certified return receipt request, or FedEx package containing distributor letter, customer letter for translation and customer lists. A software update to Version 2.0 will be issued to the user by a field service representative. |  | 1258 | 1258 units | Worldwide distribution: US (nationwide) and Internationally to: Azerbaijan, Bulgaria, Chile, Czech Republic, Greece, Israel, Kuwait, Qatar, Saudi Arabia, Serbia, Turkey, and the United Arab Emirates and to Subsidiaries in Australia, France, Germany, Italy, Japan, Portugal, Spain, and the United Kingdom. | Terminated | June 08, 2015 | 579 | 1258 | Computer
Z-0415-2014 | 66864 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | SIGMAÂż LCSÂż HIGH PERFORMANCE" |  |  | November 18, 2013 | 2013 | DePuy Orthopaedics, Inc. | DePuy Orthopaedics, Inc. is issuing a voluntary device correction of the HP M.B.T. Keel Punch Knee Instrument because of the potential for tabs on the connection end to fracture and be left in the patient. | Device Design | DePuy Orthopaedics sent an Urgent Information - Medical Device Correction letter to all affected customers on November 18, 2013. Immediately after initiation on November 18, the company halted initiation with our US distributors and international affiliates because the surgical technique referenced and images from the surgical technique were not correct. No surgeons received the November 18 surgeon letter. The recall was re-initiated on November 20, 2013. The communication to the surgeons who used the affected devices will be delivered by Distributors either by mail or in person and includes the surgeon letter.The device correction includes awareness, reiteration of surgical techniques and instrument care, and visual check of the device as outlined in the Field Notice and Surgeon letter. The devices are not being removed from the market, but the company is investigating potential design changes to the HP M.B.T. Keel Punch Knee Instrument to reduce the potential for the tabs to fracture. All DePuy United States Distributors and DePuy affiliate contacts for Canada, Latin America and International will be notified via email. The US Distributors will identify surgeons who use the affected instruments on Distributor Response Card #1. The US Distributor will notify surgeons either by mail or in person with the written communication. Once all device correction activities are complete, the US Distributor will complete and return Distributor Card #2 to DePuy Orthopaedics, Inc.The US distributors are the US consignees and are responsible for notifying all new HP M.B.T. Keel Punch Knee Instrument surgeon users of this device correction.All territories are required to ensure the HP M.B.T. Keel Punch Knee Instrument Device Correction is promptly initiated. Customers were asked to review the instructions included, along with the attached documents, whichcontain important information regarding notification of surgeons affected by this Device Correction.Customers |  | 58703880 | 29,235 units in total / 2008 through November 2013 | Worldwide Distribution - USA (nationwide) including AR,AZ,CA,CO,FL,GA,HI,IA,IL,IN,KY,LA,MA,MD,ME,MI,MN,MT,NC,NV,NY,OH,PA, TX,VA,WA,WI and Internationally to: AUSTRALIA AUSTRIA BELGIUM BRAZIL CANADA CZECH REPUBLIC DENMARK EGYPT FRANCE GERMANY GREECE HOLLAND HUNGARY INDIA IRELAND ISRAEL ITALY JAPAN KOREA MALAYSIA NEW ZEALAND NORWAY P R CHINA PAKISTAN PANAMA POLAND PORTUGAL PUERTO RICO RUSSIA SAUDI ARABIA SINGAPORE SLOVENIA SOUTH AFRICA SPAIN SWEDEN SWITZERLAND TURKEY UAE UK | Terminated | February 24, 2015 | 463 | 58703880 | Computer
Z-0279-2014 | 66506 | No | MHY |  |  | Neurology | PMA | Medtronic Activa PC, model 37601, and Activa RC, model 37612, Activa SC, models 37602 and 37603. Multi-program Rechargeable Neurostimulator for Deep Brain Stimulation. |  |  | September 30, 2013 | 2013 | Medtronic Neuromodulation | Medtronic notified customers regarding Loss of Stimulation of specific Implantable Neuro-Stimulators under the following specific conditions: Switching between groups with Multiple programs or creating program groups. Therapy can be restored with the Physician Recharge mode (RPM) of the Implantable Neurostimulator Recharger (INSR). | Software design | In September 2013 Medtronic issued a medical device correction letter titled "Loss of Stimulation and Over Stimulation" to customers associated with the affected Medtronic neurostimulators. This letter informed customers of the potential for these issues to occur and provided recommendations to significantly reduce the likelihood of the issues. Medtronic field representatives will be updating the model 8870 software application card to version AAR/01, which is contained in the physician programmer. A Physician Reply Form was asked to be signed and returned by fax or mail for acknowledging the receipt of the Urgent Medical Device Correction letter. When the Medtronic field rep updates the software application card of the physicains' programmer, they will complete a Rep Reply For: 8870 Card Update to AAR/01. |  | 62634 | 62,634 | Worldwide distribution: US (Nationwide) including Puerto Rico; and countries of: Algeria, Aruba, Austria, Belarus, Belgium, Brazil, Bulgaria, Canada, Chile, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, El Salvador, Estonia, Faroe Islands, Finland, France, French Polynesia, Georgia, Germany, Greece, Guadeloupe, Hungary, Iceland, India, Iran, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kuwait, Latvia, Lebanon, Libya, Lithuania, Luxembourg, Malta, Martinique, Mexico, Morocco, Netherlands, Netherlands Antilles, New Caledonia, Norway, Oman, Pakistan, Panama, Peru, Poland, Portugal, Qatar, Reunion, Romania, Russian Federation, San Marino, Saudi Arabia, Serbia, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay, and Venezuela. | Terminated | November 20, 2015 | 781 | 126797 | Computer
Z-0282-2014 | 66537 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Draeger Movita 603x/DVE 803x Ceiling Supply Unit. It is intended for use in operating theaters and intensive care |  |  | September 30, 2013 | 2013 | Draeger Medical, Inc. | Cases were encountered with the Draeger Movita 603x/DVE 803x Ceiling Supply Units where the drive screw in the motor unit broke and the lift arm fell down approximately 60 cm (23.6 inches). | Equipment maintenance | An "Urgent Medical Device Recall" notification letter was sent to users in September 2013. The issue was described and recommended actions were provided. Customers may contact (800) 543-5047 and select 1, then 2. |  | 139 | 139 | Worldwide distribution: US (Nationwide)and in the country of Canada. (note: the ceiling supply units are not classified as a medical device by Health Canada). | Terminated | September 03, 2014 | 338 | 139 | Computer
Z-0443-2014 | 66632 | No | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | Blanketrol III Data Export Software Version 1.10 |  |  | May 15, 2013 | 2013 | Cincinnati Sub-Zero Products Inc | During a customer inquiry regarding a beta version of the Bill Data Export Software, Model USB-127, it was discovered that the wrong software version was released to a customer. | Software Manufacturing/Software Deployment | Cincinnati Sub-Zero sent a Urgent Medical Device Notification letter dated July 15, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.They instructed their customers to to do the following:1) Upon receipt of the corrected Blanketrol III Data Export Software disc, immediately collect and destroy all previous Blanketrol III Data Export Software disc from this lot # currently in possession and replace all downloaded Blanketrol III Data Export Software with the correct version provided by CSZ.2) After the Blanketrol Data Export Software Version 1.00 have been destroyed and downloaded versions replaced with version 1.10, complete and return the enclosed response form as soon as possible to acknowledge receipt of this notification and to inform CSZ that you have performed and completed the requested actions. Return the form by fax to (513) 772-9119 scan and e-mail the information to BIIIExportSoftware@cszinc.com, or mail to:Cincinnati Sub-Zero12011 Mosteller RoadCincinnati, OH 45241 U.S.A.For further questions please call 1-(800) 989-7373 or (513) 772-8810. |  | 12 | 12 Units | Worldwide Distribution - US Distribution including the following states AK, CO, FL, MA, MI, and NC. and Internationally to Canada, Germany and United Kingdom. | Terminated | May 13, 2014 | 363 | 12 | Computer
Z-0281-2014 | 66600 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Ray Station Radiation Therapy Treatment Planning System, Model Numbers 3.0, 3.5 and 4.0 |  |  | October 23, 2013 | 2013 | RAYSEARCH LABORATORIES AB | Ray Search Laboratories found an issue with the dose calculation for Elekta motorized wedges in Ray Station 3.0, 3.5 and 4.0. The error presents when a wedge is used for a beam and the wedge angle differs from 60 degrees. | Software design | Medical Device Correction Letters dated 10/23/13 were e-mailed to the user accounts on 10/23/13. |  | 195 | 195 units domestically | US Distribution in the states of: California, Colorado, Delaware, Florida, Maryland, Missouri, New Jersey, New York, Ohio, Pennsylvania, Tennessee, and Virginia. | Open, Classified |  |  | 195 | Computer
Z-0489-2014 | 66970 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | HGE Healthcare Patient Monitor B30, Procare Monitor B20, Procare Monitor B40, and Monitor B40 |  |  | October 11, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue due to ECG filter settings associated with the B20/B40 V1, B30 and B40 V2 Patient Monitors.The B20, B30 and B40 Patient Monitors set the ECG filter setting incorrectly if the monitor starts with impedance respiration set in the ON position. Instead of the user selected filter, an additional 0.5 Hz high pass filter is used. The | Software design | GE Healthcare sent an Urgent Medical Device Correction letter dated October 10, 2013, to all affected consignees. The letter described the safety Issue, Safety Instructions, Affected Product, Details, Product Correction and Contact Information (Technical Support at 1-800-558-7044).Consignees were asked to retain a copy of the instructions with the user manual.For questions regarding this recall call 262-513-4122. |  | 16061 | 16,061 (343 US, 15,718 OUS) | Worldwide Distribution - USA (nationwide) to CA, CT, FL, GA, MO, NC, NJ, PA, TX, WI and Internationally to:FRANCE, PORTUGAL, ALGERIA, ANGOLA, ARGENTINA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BANGLADESH, BARBADOS, BELGIUM, BOLIVIA, BOSNIA 7 HERZEGOVINA, BRAZIL, BRUNEI DARUSSALAM, BULGARIA, NEW ZEALAND, CHILE, CHINA, COLOMBIA, CONGO, COSTA RICA, COTE DLVOIRE, CROATIA, CYPRUS,. CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, ECUADOR, EGYPT, EL SALVADOR, ESTONIA, FINLAND, GABON, GERMANY, GREECE, GUATEMALA, HONDURAS, HUNGARIA, INDIA, INDONESIA, IRAQ, IRELAND, ISRAEL, ITALY, JAMAICA, KENYA, REPUBLIC OF KOREA, KUWAIT, LATVIA, LEBANON, LITHUANIA, MALAYSIA, MEXICO, MOLDOVA, MOROCCO, MYAMMARA, NAMBIA, NAPAL, NETHERLANDS, NICARAGUA, NIGERIA, NORWAY, OMAN, PAKISTAN, PANAMA, PARAGUAY, PHILIPPINES, POLAND, PORTUGAL ROMANIA, RUSSIA, SENEGAL SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SIR LANKA, SWEDEN, SWITZERLAND, THAILAND, THE NETHERLANDS, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, UNITED ARAB EMIRATES , UNITED KINGDOM, URUGUAY, UZEBEKISTAN, VENEZUELA, and VIETNAM. | Terminated | February 03, 2016 | 845 | 16061 | Computer
Z-0472-2014 | 66638 | No | BSF | Carbon dioxide absorber. | Anesthesiology | Anesthesiology | 510(K) Exempt | Multi Absorber Original, Disposable |  |  | September 19, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue due to air leakage associated with the CO2 Multi Absorber. | Device Design | GE issued an Urgent Medical Device Correction letter dated September 19, 2013, to all affected customers via Fed Ex overnight mail. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to not use the affected product, isolate all affected product and return the attached form via fax to 800-535-7923. Once the information is received, customers would be contacted with an RMA and replacement information. If product was distributed further the notice should be forwarded to those customers. For questions customers should call 1-800-345-2700, option 2 followed by ooption 2. For foreign customers 1-800-932-0760, option 2.For questions regarding this recall call 262-513-4122. GE issued a press release on 12/30/2013. |  | 444115 | US: 59,721 boxes, 358,326 canisters; OUS: 3,724 boxes, 22,344 canisters | Worldwide Distribution - USA (nationwide) Puerto Rico and Internationally to Canada, Costa Rica, Mexico, and Venezuela. | Terminated | March 10, 2015 | 537 | 444115 | Computer
Z-0349-2014 | 64832 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | The GemStar Pump is a small and lightweight, single channel infusion device designed for use in the home, in the hospital, or anywhere electronic infusion is required. The GemStar Pump can be powered by AC mains adaptor, rechargeable battery pack, docking station, or two disposable AA alkaline batteries. When powered by batteries, The GemStar Pump is ideal for ambulatory patients. |  |  | March 15, 2013 | 2013 | Hospira Inc. | The proximal and distal pressure sensor calibration can drift resulting in the pump failing the Proximal or Distal Occlusion Operational Test, as described in the GemStar Technical Service Manual, or reporting one of the following errors during device setup or infusion: 1) Cassette Check - D; 2) Cassette Check - P; 3) Proximal Occlusion; 4) Distal Occlusion; 5) Pressure Calibration Error; 6) Bad Pressure Sensor Event; 7) Bad Pressure State; 8) Distal Pressure is Out of Range; 9) Proximal Sensor is Out of Range. A pump with this issue may, instead of reporting an error, not detected occlusions or issue false occlusion alarms, which will stop the infusion and invoke visual and audible warnings to the user. | Device Design | Hospira sent an Urgent Device Field Correction letter dated March 15, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The letter asks that customers immediately perform proximal and distal occlusion tests as defined i nthe GemStar Technical Service Manual. If the device fails either of the tests, remove it from clinical service and contact Hospira at 1-800-441-4100 (Monday - Friday, 8:00 AM - 5:00 PM) to report the issue and arrange for the return of the device for recalibration. The letter further reminds customers that the proximal and distal occlusion tests should be performed on an annual basis as a part of the yearly GemStar maintenance schedule. Customers were also requested to complete and return the attached reply form and return it to Hospira by fax to 1-888-714-5077 or by e-mail to Hospira7906@stericycle.com.Customers were directed to notify their consignees if they further distributed the product and confirm they have done so by return of the reply form to Stericycle. For questions regarding this recall call 800-441-4100. |  | 4132 | 1) List Number 13000: 2,162 pumps; 2) List Number 13100: 15 pumps; 3) List Number 13150: 1,955 pumps | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY. Internationally to Australia, New Zealand, China, Hong Kong, Taiwan, Japan, Korea, Philippines, Thailand, Malaysia, Canada, Austria, Bahrain, Belgium, Croatia, Denmark, Egypt, Finland, France, Germany, Gibraltar, Greece, Hungary, Iceland, Ireland, Israel, Italy, Jordan, Kuwait, Lebanon, Libya, Luxembourg, Malta, Netherlands, Norway, Oman, Portugal, Qatar, Romania, Saudi Arabia, Slovenia, South Africa, Spain, Sweden, Switzerland, Turkey, UK, United Arab Emirates, Brazil, Chile and Colombia. | Terminated | March 07, 2017 | 1453 | 47972 | Computer
Z-0685-2013 | 62719 | No | HXC | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | DePuy Spine EXPEDIUM Offset Torque Wrench |  |  | July 24, 2012 | 2012 | DePuy Spine, Inc. | A component within the handle of the Expedium Offset Torque Wrench is prematurely wearing, potentially resulting in the instrument recording torque values outside of the specification. | Device Design | DePuy Spine issued an Urgent - Voluntary Product Recall Notification dated July 24, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to disconstinue the use of the affected product and to remove the EXPEDIUMÂż Offset Torque Wrenches from service and return to DePuy Synthes Spine until direction can be provided by DePuy Synthes Spine regarding a replacement torque limiting device.For questions regarding this recall call 508-828-3769. |  | 140 | 140 units | Worldwide Distribution - USA (nationwide) and internationally to Belgium, Egypt, Ireland, Italy, Portugal, and the UK. | Terminated | December 30, 2013 | 524 | 140 | Computer
Z-0180-2014 | 66067 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | CARESTREAM DRX -Mobile Retrofit Kit, Catalog number: 1019397. Made in U.S.A. by Carestream Health, Inc. 150 Verona Street Rochester, NY 14608 |  |  | August 29, 2013 | 2013 | Carestream Health Inc. | Carestream Health Inc. has conducted a voluntary field action which includes the installation of a software patch on the following: DRX-Revolution Mobile X-Ray, DRX-Mobile and DRX Transportable Systems with DirectView Version 5.5 and 5.6 console software when used with DRX-1 Detector version 69 firmware. Carestream recently received reports of a non-diagnostic image on exams when a Normal Exposure | Software design | The firm, Carestream Health Inc., sent an "Urgent: Medical Device Recall" letter dated August 29, 2013 to its customers. The letter described the product, problem and actions taken. Carestream has developed a software patch that will allow Normal Exposure Mode and Extended Exposure Mode images to be taken in any order and will install the patch in all DRX-Revolution, DRC-Mobile and DRX-Transportable Systems with software version 5.5 or higher and Detector firmware version 69 or higher. No DRX Transportables were affected. Your Carestream service representative will install the patch at your next service call. If you have questions or concerns, or if you experience an issue prior to our visit, please contact the Carestream Customer Care Center at 1-800-328-2910, 8am through 8pm Eastern Standard Time. |  | 364 | 364 units | Nationwide distribution:AL, AZ, AR, CA, CO, CT, DC, FL, GA, ID, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MS, MI, MO, NE, NH, NJ, NM, NY, NC, OH, OK, OR, PA, RI, SC, TN, TX, UT, VT, VA, WA, WI, and WV. | Open, Classified |  |  | 554 | Computer
Z-0520-2014 | 66918 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Mega Needle Driver, 8MM assembly of the da Vinci S IS1200 and da Vinci Si, IS2000. |  |  | November 15, 2013 | 2013 | Intuitive Surgical, Inc. | Firm has become aware of the potential for detachment of the jaw insert for Large Needle Drivers and Mega Needle Drivers. | Software design (manufacturing process) | Intuitive Surgical sent a Field Safety Notice dated November 15, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to ensure all appropriate personnel are fully informed of the Notification. Discontinue use of affected product, Use the attached instructions to identify and segregate affected product. Return affected product to Intuitive Surgical following the established Returned Materials Authorization (RMA) process. Complete the Acknowledgement Form and return to Intuitive Surgical. Customers with questions were instructed to contact Intuitive Surgical Customer Service. Customers were asked to retain a copy of this letter and the acknowledgement form for their files.For questions regarding this recall call 800-876-1310, option 3.Japan 0120-56-5635 or 03-5575-1362South Korea 02-3271-3200Europe, Middle East, Asia and Africa +800 0821 2020 +41 21 821 2020For questions regarding this recall call 408-523-2100.Urgent Medical Device Recall notification letters were sent to all affected sites on November 15, 2013. |  | 3540 | 15, 236 UPDATED 12/19/13 to 11,235 | Worldwide Distribution - USA (nationwide and Puerto Rico) and Internationally to Finland, Belgium, Netherlands, UK, Czech Republic, Sweden, Germany, Austria, Switzerland, Norway, Denmark, South Korea, France, Ireland, and Slovakia. | Terminated | September 04, 2014 | 293 | 95754 | Computer
Z-0675-2013 | 63393 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Computed Tomography X-Ray System Brilliance Workspace Portal, Model #728269 Philips Medical Systems, Highland Heights, OH. |  |  | September 11, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Philips was notified that when performing post processing with the Brilliance Workspace Portal platform with software versions 2.6.1.4, the coronary arteries in the rendered image are segmented automatically. However, the center line is not extracted or highlighted upon hovering and making it impossible to activate the vessel and choose the appropriate name (label) from the drop-down menu. | Software design | Philips Healthcare sent a "CUSTOMER INFORMATION -URGENT-MEDICAL DEVICE CORRECTION" letter dated September 13, 2012. The letter identified the product, problem, and actions to be taken by the customers. Philips Service Engineers will contact customers for implementation of the software update. Contact your local Philips representative or local Philips office for further information concerning this notice. |  | 108 | 108 units | Worldwide Distribution-USA including the states of AZ, CA, CO, IL, IL, IN, MA, ME, MI, MO, NC, NJ, NY, OR, TN, TX, VA and WI and the countries of Austria, Canada, China, France, India, Ireland, Italy, Japan, Netherlands, Norway, Spain, Switzerland, Thailand and United Kingdom. | Terminated | February 14, 2014 | 521 | 108 | Computer
Z-0674-2013 | 63817 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance 64 and Ingenuity CT Computed Tomography X-Ray Scanning Systems: Brilliance 64 - Model #728231; and Ingenuity CT - Model # 728326, Distributed By: Philips Healthcare, Cleveland, OH. The Brilliance 64 and Ingenuity CT are Computed Tomography X-Ray Systems intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. These devices may include signal analysis and display equipment, patient, and equipment supports, components and accessories. |  |  | December 03, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Software anomalies. Philips Medical Systems recalled their Brilliance 64 running software v4.0.0xx516 and their Ingenuity CT running software v4.0.0xx518 due to software issues detected in these CT scanning systems. If these issues were to re-occur they could pose the risk of injury to patients undergoing scans; device users; or device service personnel. | Software design | The firm, Philips Healthcare, sent an 'URGENT- Medical Device Correction" letters' dated November 15, 2012, and December 3, 2012, to their customers. Both letters describe the product, problem and actions to be taken. The letters state that Philips is notifying all affected customers of this information via a 'Field Safety Notice and will be installing a software update on the affected Brilliance 64 and Ingenuity CT system in order to address these issues. A Philips' Field Service Engineer will be contacting each customer site and make arrangements to install the necessary software update on these device systems. The customers should contact their local Philips representative or local Philips Healthcare office. In North America and Canada needing further information or support concerning this update, should contact our Customer Care Solutions Center at 1-800-722-9377. Selection option 5, and enter their site ID or follow the prompts. |  | 6 | A total of nine (9) units were distributed (3 units in the U.S &amp; 6 units overseas) | Worldwide distribution: USA including IN, OH and OR; and the countries: Australia, Belgium, China, Israel, and Sweden. | Terminated | June 18, 2013 | 197 | 6 | Computer
Z-0178-2014 | 66331 | No | CJY | Calcium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Calcium Arsenazo, CALA-0250, For the Quantitative Determination of Total Calcium in Human Serum. |  |  | September 04, 2013 | 2013 | SEPPIM SAS | Calcium Arsenazo Reagent, reference CALA-0250, is recalled due to complaint of plasma specimens report high results than serum. | Device Design | ELITech Group of Bothell, WA sent an Urgent: Product Recall Notice dated September 9, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Actions to be taken:" Laboratories must not use plasma for the determination of total calcium when using this product. The product can still be used to determine total calcium with serum sample." ELITech Clinical Systems SAS has revised the package insert and kit box label for the CALCIUM ARSENAZO reagent reference CALA-0250. All mentions of the use of the plasma will be removed from the documentation. If you have inventory of CALCIUM ARSENAZO, reference CALA- 0250. All lots. in your laboratory, please discard the insert that you have and replace it with the insert enclosed with this notice." If you have transferred any of this product to another location, please forward a copy this bulletin to that location." All users receive CALCIUM ARSENAZO, CALA-0250, from distributors. To be certain everyone who is possibly affected by this change is contacted, we are sending this communication to all registered users. As confirmation of receipt of this notice, please return the attached receipt confirmation to Technical Support by fax (401-642-9003) within 48 hours of receipt.We apologize for the inconvenience that this situation can cause in your laboratories; we remain at your disposal for further information. If you have any questions, please contact Technical Support at 1-855-ELITECH (855-354-324)We thank you for your cooperation and the attention which you will give to this bulletin. |  | 95 | 47 kits in US and 48 kits in Sweden | Worldwide Distribution - US including the states of KY and UT and the country of Sweden. | Terminated | December 12, 2013 | 99 | 95 | Computer
Z-0171-2014 | 63859 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Handset component of the Handicare USA's Minivator Bath Bliss Reclining Bath Lifts (Bath Bliss), Part Number BP300-100/BP300-400RSUS (handset) - Model Numbers BBR0311USAB (blue) and BBR0311USSS (silver) |  |  | November 29, 2012 | 2012 | Handicare Accessibility | Firm became aware that two handsets in the affected batches became inoperable while the user was in the tub. | Device Design | Handicare sent an Urgent Medical Device Recall letter dated November 29, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately advise all of their customers affected by the recall letter to immediately remove all affected product from use. Upon receipt of affected product, replacement handsets would be sent to customers. Customers with questions were instructed to call 866-276-5438 or email Customerservice.usa@handicare.com. For questions regarding this recall call 610-266-5260. |  | 3301 | 3301 handsets | Worldwide distribution: USA (Nationwide) and country of: Canada. | Terminated | December 08, 2014 | 739 | 3301 | Computer
Z-0169-2014 | 66659 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | VitreaCoreÂż software. Versions subject to the limitation: |  |  | October 21, 2013 | 2013 | Vital Images, Inc. | Vital Images, Inc. has found an issue in VitreaCore software though internal testing. There is a potential for incorrect and possibly reversed image orientation for certain snapshots taken in VitreaCore software versions 6.0, 6.1, 6.2, 6.3 (except 6.3.1), 6.4 (except 6.4.4 and 6.4.5), as well as their upgrades, and 6.5. | Software design | Vita Images, Inc. sent an Urgent Medical Device Recall letter on October 21, 2013, to all affected customers via USPS First class mail. The letter described the Reason for Voluntary Recall (Nature of software issue, Potential scenarios which exhibit the issue), Risk to Health (How to confirm if the Snapshots are correct), Actions to be Taken by Customer (Recommended Next Steps which includes to complete and sign the enclosed Effectiveness Check Form and send to fieldnotices@vitalimages.com or fax to 1-952-540-3717), and Actions Being Taken by Vital Images (Vital Images Contact Information: If you have questions, direct Vital Images customers and customers of Toshiba America Medical Systems (TAMS) should contact Vital Images Customer Support at +1-800-208-3005, or by e-mail to support@vitalimages.com; customers through a distributor will be contacted by that distributor).For questions regarding this recall call 952-487-9574. |  | 598 | 598 | Nationwide Distribution - all states in continental USA including DC but not OK or RI.OUS: Australia, Austria, Belgium, Brazil, Canada, China, Denmark, Finland, France, Germany, Honduras, Hong Kong, Iceland, Ireland, Italy, Korea, Netherlands, New Zealand, Norway, Pakistan, Poland, Portugal, Qatar, r Saudi Arabia, Serbia, Singapore, Slovenia, South Africa,Spain , Sweden, Switzerland, Thailand, Turkey, Ukraine. | Terminated | January 16, 2015 | 452 | 598 | Computer
Z-0662-2013 | 62121 | No | DSP | Intra-aortic balloon and control system. | Cardiovascular | Cardiovascular | 510(k) | CARDIOSAVE Intra-Aortic Balloon Pump |  |  | March 27, 2012 | 2012 | Maquet Cardiovascular, LLC | In the Cardiosave Intra-aortic Balloon Pump, there is a possibility that the coiled cord assembly, a prominent connector cord between the Cardiosave IABP housing and the dual display head which secured atop the IABP may fail as a result of mishandling. | Device Design | Maquet Cardiovascular sent an Urgent Product Recall Letter dated March 27, 2012, and Urgent Product Recall Acknowledgement to all affected customers . The letter identified the product, the problem, and the action to be taken by the customer. Datascope Corp determined that the best course of action was to replace the affected product. Customers were instructed that they would be contacted by a service representative to make arrangements for the removal and replacement of their affected product. Customers with questions were instructed to contact the US Call Center at 1-800-777-4222.For questions regarding this recall call 973-709-7498. |  | 29 | 29 units (total) | Worldwide Distribution - USA including ID and Fl and Internationally to United Arab Emirates, France, Luxembourg, Poland, Spain, Sweden, Germany, Italy, Netherlands, Republic of China, Spain, Sweden, and Thailand | Terminated | September 30, 2015 | 1282 | 29 | Computer
Z-0170-2014 | 66408 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | HAMILTON-C2 Ventilator: |  |  | September 26, 2013 | 2013 | Hamilton Medical, Inc. | Manufacturer defined limits for improperly functioning loudspeakers of the HAMILTON C2 alarm system were exceeded. | Device Design | Hamilton Medical sent an Medical Device Safety Alert and Corrective Action letter dated September 27, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The corrective action does not require any action by the operator. In the case of an alarm situation as described in the letter, users should follow the instruction for use (Section 8 - responding to alarms in the operator manual) and have the ventilator serviced by an authorized service technician. An addendum to the Operators Manual as a result of the field correction was provided to customers with the Safety Alert Letter.Manufacture will develop and provision a test protocol for the service tech which test the proper functioning of the loudspeaker during the annual service inspection.We appreciate your support in this matter and sincerely regret any inconveniences that this action may cause you. We consider this action as necessary to ensure that our customers receive only safe and effective products with high quality.Further questions please call (775) 426-6331 Ext. 210 |  | 316 | 316 devices distributed in the US. | Nationwide distribution: USA including Puerto Rico and the states of IL, MT, IN, ND, AZ, MT, KY, NC, OH, MA, MO, VA, CA, NV, TX, WA, NY, GA, SC,MN, WY, PA, ND, UT, NE, CO and MI. | Open, Classified |  |  | 316 | Computer
Z-0524-2014 | 66992 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | InstaTrakÂż 3500 Plus, ENTrakTM Plus, InstaTrakÂż 3500, ENTrakTM Computer-Assisted Navigation Systems (USA), |  |  | November 26, 2013 | 2013 | GE OEC Medical Systems, Inc | GE OEC Medical Systems Inc is recalling InstaTrakÂż 3500 Plus, ENTrakTM Plus, InstaTrakÂż 3500, ENTrakTM Computer-Assisted Navigation Systems (USA) due to several potential safety issues related to the use of these products. | Software Design Change | GE Healthcare sent an An Urgent Medical Device Correction letter dated November 26, 2013, to all affected customers. The letter identified the product, the problem, and provided instructions for immediate mitigation of the issue. No product is being returned. The affected systems in the field will be corrected.Customers with questions were instructed to contact the service team at 800-874-7378.For questions regarding this recall call 801-536-4952. |  | 654 | 654 | Nationwide Distribution | Terminated | April 17, 2017 | 1238 | 654 | Computer
Z-0657-2013 | 63355 | No | FSY | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | (l) Visum Halogen Surgical Light, (2) Visum LED Surgical Light, (3) Stryker Flat Panel and Navigation Arm |  |  | September 25, 2012 | 2012 | Stryker Communications, Inc. | Stryker has become aware that there is a low likelihood of incomplete engagement of the Spring Arm Circlip during service or replacementof the Spring Arm. If the Spring Arm Circlip is not fully seated, the Spring Arm could separate and fall along with the surgical lights or monitors mounted to the end of the Spring Arm. To date, there have been no serious injuries reported from the incomplete e | Equipment maintenance | Stryker sent an Urgent Medical Device Correction letter dated September 25, 2012, to all affected customers. The letter identified the product, the prolem, and the action to be taken by the customer. Customers were instructed to immediately inspect their Spring Arm. If customers experienced any concerns with their Spring Arm they were instructed to contact Stryker immediately at 1-866-841-5663 to schedule a service technician to inspect and correct the Spring Arm circlips.For questions regarding this recall call 972-410-7100. |  | 1281 | United States: 1,082 units and Outside United States: 199 units | Worldwide Distribution - USA (nationwide) and Internationally to Australia, Japan, Taiwan, Canada, United Kingdom, France, Netherlands, Spain, Mexico, India, Korea, Singapore, New Zealand, Hong Kong, Philippines and Brazil. | Terminated | April 22, 2013 | 209 | 1281 | Computer
Z-0487-2014 | 66736 | [6 empty] | Nephros SafeSpout |  |  | October 28, 2013 | 2013 | Nephros Inc | Promotional materials for non-medical water filtration products were determined to promote claims which constitute marketing the product as a medical device. | Device Design | Nephros sent an Urgent Medical Device Recall letter dated October 28, 2013, with a return response form to all affected customers via UPS 2nd Day Air. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to cease shipment of the affected product immediately. Obtain an RGA number to return the new and used filters for credit. In addition, customers were instructed to immediately remove and discard all copies of the labeling materials that pertain to the affected product. Remove all electronic copies of these materials from any publicly accessible locations (e.g. website). Customers with questions were instructed to contact their Nephros representative or call 201-343-5202, ext 100.For questions regarding this recall call 201-343-5202.The letter was also posted at:http://www.nephros.com/download/news/SafeSpout_SafeShower_Recall_Letters.pdfNephros sent a new recall letter dated 2/20/204 as they determined they have additional customers. |  | 3742 | 3742 units | Nationwide Distribution | Terminated | June 28, 2017 | 1339 | 8946 | Computer
Z-0822-2013 | 64111 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Viterion TeleHealth Monitor Model Number V500; Product Codes 08832984 (US) and 02080468 (UK) - Physiological patient monitor (without arrhythmia detection or alarms). |  |  | November 15, 2012 | 2012 | Viterion TeleHealthcare Llc | Viterion telehealth monitors do not transmit "HI" (above the meter's measurable range) and "LO" (below the meter's measurable range) blood glucose readings from Bayer's Contour Elite XL blood glucose meters, or from Roche's Accu-Chek Advantage 03 and 04 meters to healthcare professionals. | Software design | Bayer Healthcare sent "IMPORTANT RECALL INFORMATION" to all affected customers.The letter identified the product, problem, and actions to be taken by the customers. Contact Viterion Support at 1-800-866-0133 option 1 for questions relating to this issue. |  | 369 | Domestic - 296 units, Foreign - 73 units | Worldwide Distribution-USA (nationwide) and the United Kingdom. | Open, Classified |  |  | 27440 | Computer
Z-0824-2013 | 64070 | No | DXN | Noninvasive blood pressure measurement system. | Cardiovascular | Cardiovascular | 510(k) | Philips Healthcare SureSigns VS2+ |  |  | January 14, 2013 | 2013 | Philips Healthcare Inc. | VS2+ screen may lock up and stop monitoring during patient use | Software design | Philips issued an Urgent Medical Device Correction letter on January 14, 2013. The field action will take place in two phases: Phase 1 Philips sent customer an Urgent Medical Device Correction Notification/Field Safety Notice. This notification explains the issues and gives the customers instructions on actions to take while they await the hardware correction.Phase 2 Once the hardware correction Philips will contact customers and perform the correction free of charge. Until the correction is available.Affected products may continue to be used while waiting for the correction from Philips if the unit is not used in NBP interval mode.A Philips Healthcare representative will contact you to arrange for the correction of all affected devices. If you need any further information or support concerning this issue, please contact the Philips Customer Care Center at 1-800-722-9377.On 3/14/13, firm stated Initially the fix would be a hardware upgrade but have since changed that to a software update. Customers within the US will be sent a new Field Safety Notice as well as a USB software update package and will be required to notify us upon the upgrade of their affected units. They will be performing the updates themselves |  | 3248 | 3248 units | Worldwide Distribution - US Nationwide and the countries of: AUSTRALIA, AUSTRIA, BELGIUM,CANADA, CHILE, DENMARK, FRANCEGERMANY, ICELAND, INDIA, IRELAND,ISRAEL,ITALY, JORDAN, KENYAKUWAIT, LEBANON, MALAYSIA, NAMIBIA, NETHERLANDS, NEW ZEALANDNORWAY,PALESTINE, POLAND,PORTUGAL, SAUDI ARABIA, SINGAPORESLOVENIA, SOUTH AFRICA, SOUTH KOREA,SPAIN,SWEDEN, SWITZERLANDTHAILAND, TRINIDAD AND TOBAGO, UNITED ARAB EMIRATESUNITED KINGDOM | Open, Classified |  |  | 3248 | Computer
Z-0522-2014 | 65915 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | KayserBetten model IDA beds |  |  | April 11, 2013 | 2013 | mobility unlimited inc | Some joints on the doors of a limited number of IDA beds may loosen or separate over time which can cause the doors to drop outward and cause injury. | Device Design | Mobility Unlimited sent a notification letters dated April 12, 2013 to affected customers The letter identified the affected product, problem and actions to be taken. The letter informed customers that a Retrofit Kits along with illustrated instructions will be sent. For questions call 1-800-574-7880. |  | 33 | 33 beds | Worldwide Distribution - USA Nationwide in the states of AL, AZ, CA, FL, GA, IL, MN, MO, NY, PA, TN, TX, VA, and Canada. | Terminated | October 08, 2014 | 545 | 33 | Computer
Z-0479-2014 | 66533 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Aesculap GNI161 Bipolar Foot Control |  |  | September 10, 2013 | 2013 | Aesculap, Inc. | Complaints were received for the GN161 Bipolar Foot Control reporting that the bipolar energy did not stop after release of the foot control. | Device Design | Aesculap sent an Important Recall Notification letter dated September 23, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to return affected product and receive replacement product or credit. An Inventory Sheet is attached and must be completed. For questions call 610-984-9265 or 610-984-9414. |  | 85 | 85 | Worldwide Distribution - USA Nationwide and countries of: Austria, Germany, Japan, and Taiwan. | Terminated | October 20, 2014 | 405 | 85 | Computer
Z-0694-2013 | 63931 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Spacelabs Smart Disclosure System, Model 92810, is an integrated component of the Intesys Clinical Suite (ICS) G2, Version 4.03. |  |  | December 05, 2012 | 2012 | Spacelabs Healthcare, Llc | Displaying or printing of an incorrect waveform for the time indicated in the printout or display in the Spacelabs Healthcare Intesys Clinical Suite (ICS) G2, Version 4.03, Clinical Access portion of Smart Disclosure, Model 92810. | Software design | Spacelabs has sent the Urgent -Medical Device Correction Spacelabs Healthcare ICS G2 Version 4.03 Model 92810 Smart Closure letter, dated 19 December 2012, to all U.S. customers. The firm plans to send a customer letter (translated as necessary) via e-mail to all international subsidiaries and distributors of record on December 27, 2012.The letter is addressed to Hospital Administrator/Biomedical Manager/Office Manager.Users are advised on the following actions:1) Advise your staff of this potential failure and advise them to always click the"Refresh" button when scrolling to new waveform data prior to printing or viewing waveform information or storing saved events.2) We will contact you at the earliest possible date to schedule a convenient time for your hospital to upgrade your affected software at no cost.Customers with questions can contact Spacelabs Healthcare at 1-800-522-7025 and select 2 for Technical Support. |  | 44 | 27 in US and 17 International | Worldwide Distribution-USA including Puerto Rico and the states of California, Florida, Iowa, Illinois, Louisiana, Michigan, Minnesota, North Carolina, South Carolina, Texas, and Washington, and the countries of Argentina, Mexico, France, and India. | Terminated | July 08, 2013 | 215 | 44 | Computer
Z-0480-2014 | 66714 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | NexGenÂż Complete Knee Solutions Stemmed Tibial Component Precoat |  |  | October 31, 2013 | 2013 | Zimmer, Inc. | Two complaints were received in which it was reported that the stem would not lock into the tibia. Evaluation of the returned device found that the stemmed portion of the tibial plate would not hold the stem as the connection was too loose. The taper diameter was found to be oversized by .010 . | Device Design | All distributors will be notified via electronic mail. Hospital risk managers and distributors with affected inventory will also be notified via courier. Surgeons that have implanted this manufacturing lot will be identified and notified by the distributor. Distributors will be sent a letter identifying the issue and their responsibilities. These responsibilities include locating and removing the affected product in their territory, as well as identifying and notifying surgeons that have implanted this manufacturing lot. Distributors will return on-hand affected products to Zimmer and ensure all of their products are accounted for using Attachment 1 provided in the letter. Distributors are to complete and return Attachment 2 indicating which surgeons have implanted the product and to which they have provided notice to. Hospitals that received direct shipments from Zimmer will be sent a letter identifying the issue and their responsibilities. These responsibilities include locating the affected product and returning it to their Zimmer sales representative.Communications outside of the United States will occur approximately two weeks after the United States communications.Zimmer will conduct effectiveness checks in the following way:Distributors/Hospitalso 100% of the FedEx notifications will be accounted foro An inventory certification form will need to be returned from each distributor that certifies that all available affected product from their territory has been returned.o Additional notifications via email and FedEx will be sent to those who have not completed the required certification formo Accounts will be deemed unresponsive after 3 attempts.Further questions please call 800-613-6131. |  | 128 | 128 units | Worldwide Distribution - USA (nationwide) including the states of CA, IN, OH, GA, WI, IA, NE, NJ, PA, NY, MD, NC, AZ, NM, TX, CO, IL, AL, MI, WA, FL and KS, and the countries of China, Japan and Taiwan. | Terminated | April 30, 2015 | 546 | 128 | Computer
Z-0270-2014 | 66550 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | AW VolumeShare 4 (ADVANTAGE WORKSTATION 4.3.), AW VolumeShare 5 (AW VOLUMESHARE 5 WITH ANGIOVIZ OPTION), AW Server. |  |  | September 19, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue due to the image orientation within GSI viewer on AW products.The GSI Viewer application running on AW/Volume Share or on the AW Server when making a batch and rotating the batch group clockwise 180 degrees (Inferior to Superior) will display images in the Superior orientation. The orientation annotation on the image is correc | Software design | The firm, GE Healthcare, sent an "Urgent Medical Device Correction" letter dated September 19, 2013 to its consignees/customers. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information (800-437-1171 for USA and Canada or contact local GE Healthcare Applications Specialist). The customers were instructed to "Please ensure that all potential users in your facility are made aware of this safety notification and the action taken". GE Healthcare will provide a software update for all affected systems to address the issue. A GE Healthcare service representative will contact you to arrange for this correction. This activity will be performed at no cost to you.A second updated letter was sent on October 22, 2013 to correct the Affected Product Details communicated in the previous letter dated September 19, 2013. The new statement reads "All versions of GSI Viewer prior to version 2.00-OM on AW and prior to version 2.20- OB on AW Server are affected".If you have any questions, please contact us at 800-437-1171 (USA and Canada) or contact local GE Healthcare Applications Specialist. |  | 1458 | 1458 (336 US, 1122 OUS) | Worldwide distribution: US (nationwide) including states of: AL, AZ, CA, CT,DE, FL,GA, HI, IL, IN, IA, KS, MD,MA, MI,MN, MS, O. NV, NH, NJ, NY, NC, OH, OK, OR, PA, SD, TN, TX , UT, VA, WA, WI. OUS: ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA, BELGIUM, BRAZIL, BYELO RUSSIAN, CANADA, CHINA, DENMARK, ESTONIA, FINLAND, FRANCE, GERMANY, GREECE, HONDURAS, HON KONG, HUNGARY, INDIA, INDONESIA, IRAQ, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KOREA, KUWAIT, LATVIA, LIBYA, LITHUANIA, LUXEMBOURG, MEXICO, NETHERLANDS, NEW ZEALAND, NORWAY, POLAND, PORTUGAL, RUSSIA, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TURKEY, UKRAINE, and UNITED KINGDOM.. | Terminated | July 31, 2014 | 315 | 1458 | Computer
Z-0718-2013 | 63981 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ELECTROLASE(R) HYFRECATOR(R) ELECTRODES - SHARP, Disposable Hyfrecator(R) Tips, REF/Catalog No. 7-100-8CS, Sterile, Rx ONLY -- CONMED CORPORATION Electrosurgical, Cutting and Coagulation and Accessories. |  |  | December 27, 2012 | 2012 | ConMed Corporation | ConMed received complaints of some devices which had broken through the seal of the sterile pouch. ConMed confirmed instances where the pouch seal was compromised on the affected products. In no instance has it been reported to ConMed that a compromise in the sterile barrier has resulted in illness or injury. The compromised seals were discovered prior to use. | Equipment maintenance | ConMed sent an Urgent Field Safety Notice dated January 7, 2013, and Business Reply Form to all affected customers via USPS Priority Mail. International notices were sent via FedEx International Priority on December 27, 2012. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review their inventory and complete Attachment II and return it with unused affected product to:ConMed Corporation525 French Road, Utica, NY 13502Customer were asked to return Attachment II even if they don't have any affected product. For questions regarding this recall call 315-624-3533. |  | 313190 | 313,190 (total for all devices) | Worldwide Distribution - USA (nationwide) Puerto Rico and internationally to Belgium, Colombia, Czech Republic, Dubai, Guatemala, Hong Kong, India, Israel, Italy, Japan, Kenya, Kuwait, Malaysia, Mexico, New Zealand, Norway, Oman, Pakistan, Portugal, Saudi Arabia, Thailand, and Uruguay. | Open, Classified |  |  | 313190 | Computer
Z-0454-2014 | 64762 | No | LMD | Medical image communications device. | Radiology | Radiology | 510(K) Exempt | The GE Centricity Web; Software Version 3.0x, is intended for use under the direct supervision of a licensed healthcare practitioner. The Centricity Enterprise Web is an image and information distribution system for the clinical review of medical images and reports. The Centricity Enterprise Web is not intended for primary diagnosis. The Centricity Enterprise Web is available as a supplemental sub-system to Centricity PACS or as a stand alone Web-based image and information distribution system. The Centricity Enterprise Web provides both the server software and a client application, which utilizes off the shelf browser technology. Centricity Enterprise Web supports several interface protocols which allows the system to be integrated with any image or information systems such as a hospital information system (HIS), radiology information system (RIS), or electronic medical record (EMR), which supports one of the supported integration protocols. |  |  | April 19, 2012 | 2012 | Ge Healthcare It | GE Healthcare has become aware of an issue associated with rejected images using the WADO protocol within GE Centricity Web that may impact patient safety. The WADO (Web Access to DICOM Objects) interface allows for embedding DICOM images as JPEGs on a web page and is available as a licensed operation on Centricity Web. When rejected images are archived by Centricity PACS to Enterprise Archive, an | Software design | Urgent Medical Device Correction letters dated April 19, 2013 were sent to customers. The letter requests that customers ensure that their system is configured to disable the WADO interface using the included instructions until the software revision can be installed at their site. The letter further informs customers that updated software will be sent to them for installation.The software has been modified to correct these defects. The modified software will be released and installed in user sites through GE Healthcare FMI# 85206R. |  | 53 | 53 units | US Distribution incluidng states of: CA, FL, GA, IA, IL, LA, MO, NJ, NY, OK, TN, WA, and WI. | Open, Classified |  |  | 53 | Computer
Z-0219-2014 | 66306 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Carestream Vue PACS; Virtual Enterprise SW StandAlone |  |  | September 16, 2013 | 2013 | Carestream Health Inc. | Carestream Health, Inc. has issued a recall when using the CARESTREAM Client Workstation where under very specific conditions, some images from multi-slice studies such as CT and MR may not load into the viewer application due to a software problem. | Software design | Carestream sent an Urgent Medical Device Recall letter dated September 13, 2013, to all affected customers via Fed Ex on September 16, 2013, informing them that a software problem may have the potential to occur under certain conditions when using the CARESTREAM Client Workstation. When certain conditions are met, it appears that the child background process completes its tasks before the parent process and as a result the parent process considers its work done and terminates. In the event that a software problem does occur, the scan may not be fluent or the image numbers may not be sequential. A software patch is currently being deployed by Carestream Health. Customers were instructed that they would be contacted by a Carestream Service Engineer to coordinate the deployment. Customers with questions were instructed to call 1-866 737 2487 or they may call the Carestream Customer Care Center at 1-800-328-2910, 8AM through 8pm Eastern Standard Time.For questions regarding this recall call 1-800-328-2910. |  |  |  | Worldwide Distribution - USA including AR, AZ, CA, CO, CT, FL, IA, IL, MA, MD, ME, MN, MO, NC, NH, NV, NY, OH, OK, PA, RI, TN, UT, VA WA, and WI and Internationally to Canada, Japan, and Kuwait. | Terminated | December 21, 2015 | 826 |  | Computer
Z-0271-2014 | 65792 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Siemens syngo.plaza image processing system |  |  | August 08, 2011 | 2011 | Siemens Medical Solutions USA, Inc | The firm became aware of an unintended behavior when using the syngo.plaza. The "Rename" functionality can cause an unintended patient merge if the patient is sent to another system. This issue affects customers who transfer patient data between two or more syngo.plaza or syngo Imaging XS systems. | Software design | Siemens sent a Customer Safety Advisory Notice dated August 3, 2011, to all affected customer. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to forward the Customer Safety Advisory Notice to their customers. If necessary, customers should have the letter translated technically into the local language and copy it themselves. A letter was sent to direct accounts to inform them of the issu and provide instructions to avoid encountering the problem.For questions regarding this recall call 610-219-6300. |  | 19 | 19 | Nationwide Distribution including MO, MA, NY, WI, OH, NY, CA, FL, NC, MN, and WA. | Terminated | November 06, 2014 | 1186 | 19 | Computer
Z-0701-2013 | 63949 | No | NJR | Streptococcus spp. serological reagents. | Microbiology | Microbiology | 510(K) Exempt | Becton, Dickinson and Company, BD MAX (tm) (6 channel) Instruments, Catalog number 441916, clinical diagnostic instrument. |  |  | September 18, 2012 | 2012 | Becton Dickinson &amp; Co. | Clinical diagnostic instrument may exhibit minor defects, which could lead to erroneous test results. | Device Design | Becton Dickinson sent an Urgent Field Corrective Action letter September 18, 2012,by fax. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised that BD Field Service Associates will contact them within 7-10 days of the notification to determine if their instrument is impacted. Customers were requested to return a response form indicating the serial number on their instrument. If customers needed further assistance they were instrycted ti contact BD Technical Services Department at 1-800-638-8663.For questions regarding this recall call 410-316-4054. |  | 86 | 86 Instruments | USA including CT, FL, IL, KS, MD, MI, MO, NY, PA, RI, TX and WI. Internationally to Canada, Europe and Japan. | Terminated | February 04, 2013 | 139 | 86 | Computer
Z-0202-2014 | 66485 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Synthes 8.5 MM Medullary Reamer |  |  | October 03, 2013 | 2013 | Synthes USA HQ, Inc. | Complaints were received for intraoperative reamer head breakages which can allow for unretrieved fragments of non-implant grade material. | Device Design | Synthes sent an Urgent Notice: Voluntary Medical Device Recall Notification letters, dated October 3, 2013, to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to review their inventory and immediately remove the affected lots from stock. For questions call 610-719-5450 or contact your Synthes Sales Consultant. |  | 2979 | 2979 | Worldwide Distribution: USA (Nationwide) and countries of: Canada and Mexico. | Open, Classified |  |  | 2979 | Computer
Z-0269-2014 | 66560 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Philips brand IntelliSpace PACS 4.4, Image Management System; The system is a software package used with general purpose computing hardware to acquire, store, distribute, process and display images and associated data throughout a clinical environment. Product is manufactured and distributed by Philips Healthcare Informatics, Inc. Foster City, CA |  |  | October 14, 2013 | 2013 | Philips Healthcare Informatics, Inc. | Philips Healthcare became aware of a software problem with the IntelliSpace PACS system, in that a timing issue has been identified that may result in a condition where the device does not store additional images received before processing is complete. The missing images may contain pathology which may result in a different diagnosis if they had been available at time of interpretation. | Software in the Use Environment | Philips sent an Urgent Field Safety Notice dated October 14, 2013, to all affected customers. The letter identified the product, problem, and actions to be taken by the customers.Please review the following with all radiology technicians who use lntelliSpaceBY CUSTOMER/USER PACS:" Assure you and your colleagues are aware of the problem." Verify that all images sent for each study are present on the iVaultbefore the images are purged from the modality.o Use the image count from the lower right corner of thethumbnails as seen in Enterprise only. Single images (such asCR, DX and MG) show no image count, so the user shouldcount those thumbnails without image counts as one. Add all thenumbers up to come up with the total number of viewableimages, stored to lntelliSpace." Compare number of images sent from the modality and those sent fromother 3rd party source to image count in IntelliSpace PACS." If the image count on the modality or 3 party source indicates there aremore images to the study than are present on the iVault, the studyshould be resent from the modality and/or 3 party source." Resending the studies will typically correct the problem but thetechnician should again verify the image count is correct.If you need any further information or support concerning this issue, pleaseINFORMATION AND contact your Philips representative or contact Care at isitecare@philips.com. Please call (650) 293-2300 for further questions. |  | 71 | 71 units | Worldwide Distribution - USA (nationwide) and country: Canada. | Terminated | June 11, 2014 | 240 | 71 | Computer
Z-0210-2014 | 66449 | No | HBL | Neurosurgical head holder (skull clamp). | Neurology | Neurology | 510(k) | Composite Series Skull Clamp, one clamp per tote. |  |  | August 29, 2013 | 2013 | Integra LifeSciences Corp. | An investigation of an adverse trend of complaints for the Skull Clamp index locking knob not easily unlocking or disengaging and malfunctions of the 80 lbs. torque knob was initiated. None of the complaints resulted in a report of patient injury, adverse health consequence, or a delay / prolongation of the surgical procedure. | Device Design | Integra sent an Urgent: Voluntary Medical Device Recall Notification dated August 29, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Your local Integra NeuroSpecialist will set up a time with you to discuss replacement options but in the interim, Integra is asking that you do the following:1. Review your inventory and determine if you have any MAYFIELDÂż Composite Series Skull Clamps (A3059). If so, stop using them immediately and remove them from service to prevent any inadvertent use of the Composite Series Skull Clamp.Either while the Integra NeuroSpecialist is on-site or independently;2. Complete the attached form. If you do not have affected product, check the box: I do not have Composite Series Skull Clamps.3. If you do have affected product, check the box: I do have Composite Series Skull Clamps. Record the total quantity and Serial No.(s) of all Skull Clamps in your inventory. The Serial No. is located on the clamp upright comprised of two letters (SS) followed by six numerals.4. Complete the other information and return by email or fax as indicated on the form, or give it to the Integra NeuroSpecialist who will return it for you.For assistance returning and exchanging these products or any other questions that you may have, please contact Integra Customer Service at 1-855-532-1723.We sincerely apologize for any inconvenience this may have caused and thank you for your cooperation. |  | 37 | 37 units | US Distribution including the states of AL, AR, CA, CT, FL, GA, IA, IL, MS, NY, OH, SC, TN, and TX. | Terminated | January 14, 2014 | 138 | 37 | Computer
Z-0273-2014 | 66664 | No | EBW | Dental handpiece and accessories. | Dental | Dental | 510(k) | Osseocare Pro Console &amp; Set Osseocare Pro, Model 1600870-001, 1700470-001, and 1700471-001. |  |  | October 24, 2013 | 2013 | Nobel Biocare Usa Llc | The recall was initiated because there is a potential for malfunction of the OsseoCare Pro when used with the OseeoCare Pro iPad application software release 1.2.0.7. | Software Manufacturing/Software Deployment | Bien Air Dental send and Urgent Field Safety Notice dated October 24, 2013, to all affected customers. The letter notified customers of the recall, description of the problem, description of the product (OsseoCare Pro iPad app software release 1.2.0.7), action to be taken by user, an alternative solution, and contact details. The letter was also accompanied by an acknowlegement form. Customers with questions were instructed to call 1-800-433-2436.For questions regarding this recall call 714-282-4800. |  | 1082 | 1082 (391 US) | Worldwide Distribution - USA (nationwide) United States and Internationally to Austria, Belgium, Brazil, Canada, Croatia, Czech Republic, Denmark, Finland, France, Germany, Hong Kong, Hungary, India, Ireland, Italy, Japan, Latvia, Lithuania, Luxembourg, Mexico, Netherlands, Norway, Philippines, Poland, Portugal, Romania, Russia, South Africa, Spain, Sweden, Switzerland, Turkey, and the United Kingdom. | Terminated | April 10, 2014 | 168 | 1082 | Computer
Z-0802-2013 | 60695 | No | DRY | Cardiopulmonary bypass on-line blood gas monitor. | Cardiovascular | Cardiovascular | 510(k) | Terumo CDI 101 Hematocrit/Oxygen Saturation Monitoring System. Monitors blood/oxygen saturation, hematocrit and hemoglobin levels. |  |  | November 01, 2010 | 2010 | Terumo Cardiovascular Systems Corporation | Inaccurate readings. Software contains unused (turned off) accuracy correction factors for slope (gain) and offset (bias) for Hematocrit/Saturation calculations. The accuracy on any cuvette size did not meet the specifications. | Software design | Terumo sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated November 1, 2010 to all affected customers. The letter identified the product, problem, and the actions to be taken by the customers. Contact Terumo CVS Customer Service at 1-800-521-2818 for questions or concerns. |  | 622 | 622 units | Worldwide Distribution. | Terminated | November 08, 2013 | 1103 | 622 | Computer
Z-0763-2013 | 64155 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | St. Jude Medical, AGA Medical corporation, AMPLATZERÂż TorqVueÂż FX Delivery System. Model #'s: 9-ITVFX06F45/60, 9-ITVFX07F45/60, 9-ITVFX007F45/80, 9-ITVFX08F45/60, 9-ITVFX08F45/80, 9-ITVFX09F45/80, 9-ITVFX10F45/80 9-ITVFX12F45/80, 9-ITVFX13F45/80. Cardiac surgical instrument, used to close atrial septal defects between the right and left atriums. |  |  | January 18, 2013 | 2013 | St Jude Medical Inc | The distal end of the core wire of the TorqVue FX Delivery System could potentially fracture when exposed to a combination of certain cardiac anatomies and usage conditions. | Device Design | St. Jude Medical sent an "URGENT MEDICAL DEVICE RECALL NOTICE" dated January 18, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers were advised to discontinue use of the device and remove it from inventory. Customers should contact their SJM sales representative for questions relating to this recall. |  | 635 | 635 | Nationwide Distribution-including DC and the states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NH, NJ, NV, NY, OH, OK, OR, PA, RI, SC, TX, VA, WA, and WI. | Terminated | May 21, 2013 | 123 | 635 | Computer
Z-0456-2014 | 66808 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | iba COMPASS SW Version 3.0, Catalog Number CS10-100 (medical linear accelerator) |  |  | August 20, 2013 | 2013 | Iba Dosimetry Gmbh | The development partner of IBA informed that there is problem in a SW component used in COMPASS when converting Region of Interest (ROI) geometries to dose grid representation. Under certain conditions this problem generates too large ROI voxel volumes causing erroneous dose statistics and DVHs. | Software design | Iba sent an Field Safety Notice on August 20, 2013, by email and letter delivered by UPS, Please distribute this notice to all those who need to be aware within your area of responsibility where the affected software have been installed (e.g. colleagues, satellites).Please maintain awareness of this notice and resulting action until we have provided you with a new software version.We would kindly ask you to acknowledge the receipt of this Field Safety Notice. Please sign and send back the attached Confirmation of Receipt within four weeks after the receipt of this Field Safety Notice.We deeply apologize for the inconvenience. We are working on a new software version with corrected functionality. This should become available during next fall and we will ship it to you as soon as it is released, together with the installation and verification instructions.Should you have questions or require additional information, please contact the Customer Service Team at: Service Department IBA Dosimetry GmbH Bahnhofstrasse 5DE-90592 Schwarzenbruck, GermanyTel.:+49 9128 607 38Fax: + 49 9128 607 26Email: service@iba-group.com. |  | 33 | 33 units | Worldwide Distribution - US Distribution including the states of CT, MA, OH, CA, SC, NY and OK, and the countries of International distribution (from Germany) Poland, Germany, GB, The Netherlands, Spain, Turkey, Canada, Mexico, China, Hong Kong and Singapore. | Terminated | February 24, 2014 | 188 | 33 | Computer
Z-0469-2014 | 66596 | research use only/pathology | OTU |  |  | Pathology | Enforcement Discretion | Digital Pathology System (DPS) Software version 1.1 |  |  | October 01, 2013 | 2013 | Omnyx Llc | There is an issue associated with Digital Pathology System (DPS) Software, in which the annotation tools within the DPS version 1.1 Image Viewer will display incorrect measurements and area calculations for a region of interest on a scanned slide, under certain conditions. | Software design | Omnyx sent a Urgent Field Safety Notice letter dated October 1, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.Please discontinue use of the measurement and area calculation features of the annotation tools (line, rectangle/square, ellipse/circle, and freeform) in the DPS Image Viewer at this time. If any of these tools have been used for measurements or area calculations on clinical cases, verify measurements and calculations on glass slides.Please confirm receipt of these instructions by completing the attached form and scanning/emailing it to FieldAction@ omnyx.com or faxing it to 001-412-894-2137Product Correction: An updated version of the DPS SW will be released and installed on your system as soon as it becomes available.Please note that the appropriate regulatory authorities have been informed of this field action.We apologize for any inconvenience that this may have caused you and are working diligently to address this issue.If you have any additional questions or concerns, please contact you GE Healthcare account representative. Further questions please call (412) 891-2100. |  | 2 | 2 | International Distribution to Canada and Switzerland only. | Terminated | April 16, 2014 | 197 | 2 | Computer
Z-0696-2013 | 63617 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Roche COBAS INTEGRA 400 and 400 plus Analyzer, |  |  | November 19, 2012 | 2012 | Roche Diagnostics Operations, Inc. | A software security issue with Oracles TNS-Listener component has been identified. TNS-Listener is supplied with the Roche COBAS INTEGRA 400/400 plus Analyzers. The Oracle TNS-Listener can be considered as an authentication and redirection component between the client and the database inside the Roche product. This means an attacker could register their own non-Roche database via remote access | Software design | Roche sent an Urgent Medical Device Correction letter on November 19, 2012, to all affected customers via UPS Ground (receipt required). The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to complete the attached form and return by fax and file this Medical Device Correction for future reference. For technical support customers were instructed to contact Diagnostics Technical Support at 1-900-428-2336.For questions regarding this recall call 317-521-3911. |  | 845 | 845 Units | Nationwide Distribution including AL, AR, AZ, CA, CT, DC, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WY and GUAYNABO, PR. | Terminated | January 13, 2014 | 420 | 845 | Computer
Z-0798-2013 | 64114 | Yes | MSX | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Mindray DS USA, Inc., DPM Central Monitoring System with ambulatory Telepack. Part numbers: 300BF-PA00002, and 300BF-PA00003 (DPM Central Monitoring System) 115-006908-01, -03, -04 (DPM software upgrade kit). The DPM Central Monitoring System transfers information between DPM Central station and other networked devices. |  |  | November 28, 2012 | 2012 | Mindray DS USA, Inc. d.b.a. Mindray North America | Mindray has identified a software anomaly that may where, in the presence of an asystole, the Telepack ECG waveform may produce a unique step pattern (instead of a flat line) if Lead 1 or V is being utilized for detecting cardiac arrhythmia. | Software design | Mindray North America/ Mindray DS USA Inc. sent an "URGENT PRODUCT CORRECTION ACTION" letter dated November 28, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact your Mindray Service Representative at 1-800-288-2121 for questions relating to this issue. |  | 51 | 51 units | Nationwide Distribution. | Terminated | August 15, 2013 | 260 | 51 | Computer
Z-0673-2013 | 63319 | No | GZT | Self-retaining retractor for neurosurgery. | Neurology | Neurology | 510(k) | Vycor Medical ViewSite Brain Access System 17mm/ 11mm / 5cm Device***STERILE*** |  |  | August 21, 2012 | 2012 | Vycor Medical, Inc. | Vycor Medical recalled their VBAS (Vycor Viewsite Brain Access System) because an unidentified black fiber was found on the device. | Equipment maintenance | Vycor Medical, Inc. notified all customers of the product recall by telephone and follow-up letter. Customers were instructed to quarantine affected products and hold the shipment(s) in quarantine unti further notice. The units will be replaced with inspected and passed units. All distributors were instructed to notify their customers of the recall and ask to have them return their stock. For question call 561-558-2020. |  | 3000 | 3000 devices | Worldwide Distribution - USA (Nationwide) in the states of TX, IL, NJ, FL, OH, PA and (International) in the country of Japan. | Terminated | May 13, 2013 | 265 | 3000 | Computer
Z-0695-2013 | 63863 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Hoffman LRF Wire Bolt, Catalog number 4933-1-001 through -003. |  |  | October 25, 2012 | 2012 | Stryker Howmedica Osteonics Corp. | Two complaints have been filed where Hoffman LRF Wires have broken during load bearing application by patients. | Device Design | Stryker sent "Urgent Medical Device Removal" notification letters and acknowledgement forms dated 10/25/2012 via Fed Ex to Branches/Agencies with return receipt. Stryker sent notification letters/acknowledgement forms dated 10/26/202 via Fed Ex to Hospital Risk Management, Hospital Chief of Orthopaedics and Surgeons. The letters described the device issue and provided recommended actions. Questions were directed to Stryker Customer Service. |  | 476 | 476 units | Nationwide Distribution, including the states of PA, OH, NY, CA and IL. | Terminated | December 01, 2015 | 1132 | 476 | Computer
Z-0478-2014 | 66874 | No | FLE | Steam sterilizer. | General Hospital | General Hospital | 510(k) | AMSCOÂż EvolutionÂż steam sterilizer Transfer Carriage. |  |  | October 18, 2013 | 2013 | Steris Corporation | The firm discovered that unlatching can occur when overloaded transfer carriages are pulled or maneuvered incorrectly or with excessive force. | Device Design | Steris sent an Urgent Voluntary Field Correction Notice letter dated October 18, 2013 to their customers. The letter identified the affected product, problem, and actions to be taken. The letter informed customers that a Steris Service Representative will contact affected facility to arrange for Steris to upgrade the Transfer Carriage(s). For questions contact Steris Field Service Dispatch at 1-800-333-8828. |  | 936 | 936 units | Worldwide Distribution - USA Nationwide in the states of: AL, AZ, AR, CA, CO, CT, DE, FL, GA, ID, IL, IN, KS, KY, LA, ME, MD, MA, MI, MN, MI, MO, MT, NE, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, WY and the District of Columbia including the countries of: Brazil, Canada, Chile, Guam, Mexico, Saudi Arabia, United Arab Emirates &amp; Venezuela. | Terminated | May 13, 2014 | 207 | 936 | Computer
Z-0871-2013 | 64201 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart XL+ Defibrillator/Monitor |  |  | January 28, 2013 | 2013 | Philips Healthcare Inc. | Philips HeartStart XL+ Defibrillator/Monitor may become locked out of clinical use | Software design | Philips Healthcare sent an Urgent Medical Device Recall letter dated January 2013 on January 28, 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised if the Philips HeartStart XL+ Defibrillator/Monitor is first powered on and then there is disconnection or an interruption in AC power, the device may be locked out of clinical mode until the user enters the service mode password and performs an Operational Check. Users are provided with the Safety Notice that provides information on how to identify affected devices and instructions on actions to be taken. Customers were advised to follow the ACTION TO BE TAKEN BY CUSTOMER / USER section of the notice. To correct this issue Philips will provide a software upgrade to customers with affected units free of charge.If you need any further information or support concerning this issue, please contact your local Philips representative at 1-800-722-9377.For questions regarding this recall call 978-687-1501. |  | 3785 | 3785 units | Worldwide Distribution - USA (nationwide) and Internationally to ALGERIA, ANGOLA, AUSTRALIA, AUSTRIA, BAHRAIN, BELGIUM, BULGARIA, COLOMBIA, CROATIA (LOCAL NAME HYVATSKA), CYPRUS, CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, EGYPT, ESTONIA, FINLAND, FRANCE, GABON, GERMANY, HONG KONG, HUNGARY, ICELAND, INDIA, INDONESIA, IRAN (ISLAMIC REPUBLIC OF), IRAQ, IRELAND, ISRAEL, ITALY, JORDAN, KENYA, REPUBLIC OF KOREA, KUWAIT, LATVIA, LEBANON, LIBYAN ARAB JAMAHIRIYA, LITHUANIA, MALAYSIA, MALTA, MAURITIUS, MOROCCO, NAMIBIA, NETHERLANDS, NEW ZEALAND, NORWAY, OMAN, PALESTINE, POLAND, PORTUGAL, ROMANIA, SAUDI ARABIA, SINGAPORE, SLOVAKIA (SLOVAK REPUBLIC), SLOVENIA, SOUTH AFRICA, SPAIN, SRI LANKA, SWEDEN, SWITZERLAND, SYRIAN ARAB REPUBLIC, TANZANIA, THAILAND, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, AND VIET NAM. | Open, Classified |  |  | 3785 | Computer
Z-0949-2013 | 64580 | No | IOQ | Powered flotation therapy bed. | Physical Medicine | Physical Medicine | 510(K) Exempt | TotalCare SpO2RT2 bed. |  |  | February 26, 2013 | 2013 | Hill-Rom, Inc. | In certain situations a software problem with the TotalCare SpO2RTÂż 2 bed may occur that can lead to inadvertent loss of sufficient surface cushionpressure to adequately support the patient. This may increase the likelihoodof skin irritation for patients at higher risk of developing pressure-related skin conditions. A correction to the software program is currently being developed andtested, | Software design | Hill-Rom sent an Urgent Field Safey Notice letter dated February 26, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter advised customers to follow steps provided until the updated software is available. Customers were requested to forward a copy of this letter to any other facility personnel you deem appropriate. Questions should be directed to Hill-Rom Technical Support at 800-445-3720 |  | 309 | 309 | USA Nationwide Distribution including the state of Canada | Terminated | April 16, 2014 | 414 | 309 | Computer
Z-0869-2013 | 64185 | No | FWZ | Manual operating table and accessories and manual operating chair and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | MAYFIELDÂż Composite Series Base Unit Standard, Model (A3101), Rx ONLY, Integra LifeScience Corporation, 4900 Charlemar Drive, Building A, Cincinnati, OH 45227 |  |  | December 21, 2012 | 2012 | Integra LifeSciences Corp. | Integra determined there is a possibility that the Base Units may break during use. | Device Design | Integra LifeSciences Corporation sent an Urgent Medical Device Recall letter dated December 21, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to stop using the affected product immediately and also remove them from service to prevent any further use of the affected product. Integra would provide an alternative base unit to replace those customers could no longer use. Customers were asked to complete the Recall Acknowledgement and Return Form to ensure they received this information. For assistance customers were instructed to call 1-513-533-7944 or 1-513-533-7946 |  | 96 | 96 units | Worldwide Distribution - USA including AL, AR, AZ, CA, CO, CT, FL, GA, IL, IN, KY, MA, MD, MI, MN, MO, MS, NC, NE, NJ, NM, NY, OH, OK, OR, PA, TN, TX, UT, VA, WA &amp; WI. Internationally to Belgium. | Terminated | July 10, 2013 | 201 | 96 | Computer
Z-1681-2013 | 65320 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | ***REF C7122***QTY 1***APEX Arthroscopy Tubing Set. |  |  | May 30, 2013 | 2013 | Linvatec Corp. dba ConMed Linvatec | Conmed Linvatec received complaints that their products C7120 APEX Arthroscopy Tubing Set, One Connection and C7122 APEX Arthroscopy Tubing Set, Two Connection, contained a scratch/grooved area on the inside diameter of the pressure fitting that may cause the APEX Irrigation (pump) Console to sense a pressure drop and increase fluid flow to the surgical site. | Equipment maintenance | ConMed Linvatec sent an "URGENT: Medical Device Recall Notification" letter by FedEx next day delivery or equivalent on 5/30/2013 to all domestic direct consignees. ConMed Linvatec is translating customer notifications letters into the necessary languages and will send International notification letters no later than 6/4/2013. |  | 4691 | 4,691 units, total of each | Worldwide Distribution-USA (nationwide) including the states of AR, CA, FL, GA, IL, IN, KS, MO, NC, NJ, NM, NY, OH, OK, PA, TN, TX, UT, VA, and WI, and the countries of Austria; Belgium, Canada, China, Czech Republic, Finland, France, Great Britain, Italy, and Japan. | Terminated | September 04, 2013 | 97 | 4691 | Computer
Z-1665-2013 | 65429 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | S-LIFT Inserter Outer Sleeve |  |  | May 24, 2013 | 2013 | SpineFrontier, Inc. | The outer sleeve would not fit over the S-LIFT Inserter assembly because the hole at the top of the sleeve handle was too small. | Device Design | Spine Frontier sent an Urgent Advisory Notice letter dated On June 7, 2013 to all consignees via email. The letter identified the affected product, problem and actions to be taken. The letter informed consignees that a modification has been implemented to the Outer Sleeve component of the S-LIFT Inserter and that SpineFrontier has initiated removal of the affected product from the field.. An acknowledgement form is to be signed and returned from each recipient of the Advisory Notice. Upon receipt of the acknowledgement form and RMA request for part return, replacement parts will be shipped. For questions call 978-232-3990. |  | 12 | 12 devices | USA Nationwide Distribution including the states of: FL, TX, OK, MO, NJ, MA | Terminated | October 08, 2014 | 502 | 12 | Computer
Z-1666-2013 | 65442 | No | HXX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | InSpan Inserter, Right Assembly; InSpan Inserter, Left Assembly |  |  | May 07, 2013 | 2013 | SpineFrontier, Inc. | Potential for set screw to be come deformed. | Device Design | Spine Frontier sent an Urgent Advisory Notice letter dated May 30, 2013 to all affected consignees. The letter identified the affected product. problem and actions to be taken. The letter informed consignees that a correction was implemented for InSpan Inserters. Consignees were instructed to remove parts from the field immediately for modification and that replacement Inserters will be shipped upon receipt of completed acknowledgement form and RMA request for part return. For questions call 978-232-3990. |  | 61 | 61 individual inserter assemblies | Worldwide Distribution - USA Nationwide: in the states of: FL, CO, VA, KS, TX, OH, TN, OK, CA, IL, MO, MN, Wash DC, including PR and the countries of Dominican Republic, Medico, Panama | Terminated | January 11, 2016 | 979 | 61 | Computer
Z-1096-2013 | 64725 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | GE Healthcare LOGIQ S8 (with software revisions R1.1.1 and R1.5.1) and LOGIQ S7 (with software revisions R1.0.1, R1.0.2 and R1.0.3) diagnostic ultrasound system with the S4-10-D Ultrasound Probe. The device is intended for use in ultrasound evaluation. |  |  | February 08, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has become aware of a potential safety issue due to the S4-10-D Ultrasound Probe used in conjunction with the LOGIQ S8 and LOGIQ S7 Ultrasound systems. It was also discovered that this issue also leads to a specification issue of the acoustic output of this probe. There is a potential for a skin burn when the flow model in the cardiac application is activated with the S4-10-D Ultraso | Software design | Consignees were sent on 2/8/2013 a GE Healthcare " Urgent Medical Device Correction" letter dated February 8, 2013. The letter was addressed to Hospital Administrators / Risk Manager, Biomedical Engineering and Managers of General/Diagnostic Imaging. The letter described the Safety Issue, Safety Instructions, Affected Product details, Product correction and Contact Information. US customer contact number is (1)800-437-1171. |  | 2061 | 2061 | Worldwide Distribution, USA, including the states of AL, AZ, AR, CA, CO,K CT, DC, FL, GA, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MS, MO, MT NE, NV N NJ, NM, NY, NC, ND, OH, OK, PA, RI ,SC, SD, TN, TX, VA, WA, WI and the countries of AUSTRALIA, AUSTRIA, BRAZIL, CANDA, CHILE, CHINA, COSTA ARICA, CROATIA, CZECH REPUBLIC, DENMARK, FINLAND, FRANCE, GERMANY, GREECE, HONG KONG, HUNGARY INDIA, INDONESIA ITALY JAPAN, KOREA, LATVIA, LEBANON, MALAYSIA, MEXICO, MOLDOVA, MOROCCO, NETHERLANDS, NEW ZEALAND, NORWAY, PHILIPPINES, POLAND, PORTUGAL, ROMANIA, RUSSIA, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA , SOUTH KOREA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, VIETNAM. | Terminated | November 21, 2013 | 286 | 2061 | Computer
Z-1670-2013 | 65292 | No | HQC | Phacofragmentation system. | Ophthalmic | Ophthalmic | 510(k) | SovereignÂż Compact Phacoemulsification System, v 5.3, Part Number: CMP680300. The system is used for cataract surgery. |  |  | May 23, 2013 | 2013 | Abbott Medical Optics, Inc. | Abbott Medical Optics Inc. (AMO) initiated this Urgent Notice of Medical Device Correction and Advisory Notice to all customers who use the AMO SovereignÂż Compact Phacoemulsification System, Version 5.3, because it has the potential to lose the vacuum tone mode during surgery. | Software design | AMO issued an "Urgent Notice of Medical Device Correction and Advisory Notice" on May 23, 2013 via email, to all customers with a SovereignÂż Compact Phacoemulsification System v 5.3 to inform them of the issue and provide them with actions to take in the event of sound loss during use of the SovereignÂż Compact Phacoemulsification System, v 5.3. |  | 534 | 534 systems | Worldwide Distribution, including Nationwide (US) and countries in Latin America, EMEA, and APAC. | Terminated | October 06, 2014 | 501 | 534 | Computer
Z-1668-2013 | 65430 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Medtronic MiniMed Paradigm Insulin Infusion Sets |  |  | June 07, 2013 | 2013 | Medtronic MiniMed | Medtronic is recalling the Medtronic MiniMed Paradigm Medtronic is recalling the Medtronic MiniMed Paradigm Insulin Infusion sets because of a potential safety issue that can occur if insulin or other fluids come in contact with the inside of Medtronic Paradigm infusion sets. Under certain conditions, the infusion set may malfunction and deliver an incorrect volume of insulin. | Device Design | Medtronic, Inc. sent an Urgent Medical Device Safety Notification letter dated June 10, 2013, to all affected consignees. The letter identified the product and notified healthcare professionals of the potential for over or under delivery of insulin if insulin or other fluids contact the inside of Medtronic Paradigm Tubing Connectors. The letter informed the consignees that it applies to all Medtronic infusion sets designed for use with Medtronic Paradigm family infusion pumps. Paradigm Insulin Pump users, distributors, and healthcare professionals were all instructed to contact Medtronic's 24 hour helpline at (888) 204-7616.For questions regarding this recall call 818-567-4700. |  | 11000000 | 11,000,000 units | Worldwide distribution: US (nationwide) distribution: Algeria, Argentina, Australia, Bahamas, Bahrain, Bermuda, Boliva, Bosnia and Herzegovina, Brazil, Canada, Chile, China, Columbia, Costa Rica, Croatia, Cyprus, Denmark, Dominican Republic, Ecuador, Egypt, El Salvador, Estonia, Finland, France, French Southern Territories, Germany, Great Britain, Greece, Guatemala, Hong Kong, India, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kuwait, Latvia, Lebanon, Macedonia, Malaysia, Mexico, Morocco, Netherlands, Norway, New Zealand, Oman, Panama, Paraguay, Peru, Poland, Province of China, Taiwan, Qatar, Republic of Korea, Romania, Russia Federation, Saint Lucia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Syrian Arab Republic, Thailand, Trinidad and Tobago, Turkey, Ukraine, United Arab Emirates, Uruguay, andÂżVenezuela. | Terminated | June 10, 2015 | 733 | 11000000 | Computer
Z-1099-2013 | 64853 | No | OCX | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Torrent" irrigation tubing. |  |  | March 07, 2013 | 2013 | US Endoscopy Group Inc | The firm was notified by their customers that there was a water leakage from the white tubing connector. | Device Design | US Endoscopy Group sent an Urgent Recall Notification letters to their customers. The letter identified the affected product, problem and actions to be taken. They instructed their customers to remove any of the affected products from their inventory and quarantine them. They also requested each customer complete the attached Customer Response Card and return it to US Endoscopy via fax at 1-440-639-4495 or by email at recall@endoscopy.com as soon as possible. For questions contact Customer Service at 1-800-769-8229. |  | 21189 | 21,189 units | USA Nationwide Distribution - including the states of: AK, AL, AZ, CA, CO, CT, DE, FL, GA, ID, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NE, NH, NM, NV, NY, OH, OR, PA, SC, TN, TX, UT, VA, WA, WI &amp; WV. | Terminated | March 20, 2014 | 378 | 21189 | Computer
Z-1678-2013 | 65577 | No | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter's elastomeric infusion pumps are non-electronic, ambulatory, disposable devices that deliver continuous medication for periods of 30 minutes up to 7 days. They are designed for medication therapies requiring continuous infusion. |  |  | July 31, 2009 | 2009 | Baxter Healthcare Corp. | There are reported incidents of leakage on the INFUSOR and FOLFUSOR pumps at the connection of the blue winged cap and the distal male Luer. Investigation of these reports indicates the incidence of leakage has only occurred after pumps are filled, primed, re-capped (using the blue winged Luer cap) and stored for a period of time, typically overnight. In addition, the reported instances of leakage | Equipment maintenance | Baxter sent an "Important Product Information" notification dated July 31, 2009, with a follow-up "Important Product Information" notification mailed on March 5, 2010. The notifications informed customers of the potential issue, and included instructions to ensure that the winged Luer cap is tightly secured after filling and priming. Customers were also instructed to follow normal clinical practices to reduce contact with any leaking medication, and to store pumps with pre-dispensed medication in a sealed pouch. If customers noticed a leak they were instructed to place the pump in a plastic bag, quarantine the device and call Baxter's Corporate Product Surveillance at 1-800-437-5176 for further instructions. Customers were asked to return the attached reply form via FAX to 1-847-270-5457. If customers further distributed the affected product they were instructed to forward the notification to other facilities. Customers with questions were instructed to call 1-800-933-0303.For questions regarding this recall call 224-270-4667. |  | 112860 | Code 2C1700KP: 7,128 units; Code 2C1702KP: 105,732 units | Worldwide Distribution - USA (nationwide), Puerto Rico and Internationally to Australia, Austria, Belgium, Brazil, Bulgaria, Canada, Chile, Columbia, Czech Republic, Denmark (including Faroe Islands), Ecuador, Finland, France (including Martinique and Reunion Islands), Germany, Greece, Iceland, India, Ireland, Isreal, Italy, Luxemburg, Japan, Malaysia, Mexico, Netherlands, New Zealand, Norway, Panama, Portugal, Qatar, Russia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Korea, Spain, Sweded, Switzerland, Thailand, Turkey, United Arab Emirates, and United Kingdom. | Terminated | March 08, 2016 | 2412 | 559758 | Computer
Z-1083-2013 | 64759 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | VMX, VMX Plus Mobile Radiographic Unit with the following models: |  |  | January 23, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has become aware of an issue associated with the base casting for the articulated arm of your VMX andVMX Plus Mobile X-ray systems that may impact Operator or Patient safety. | Device Design | GE Healthcare sent and "Urgent Medical Device Correction" letter dated January 23, 2012. The letter was addressed to Hospital Administrators / Risk Managers, Radiology Department Managers and Radiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details and Product Correction. For questions contact your local GE Healthcare Service Represenative. |  | 1296 | 1296 | Worldwide Distribution - USA (nationwide) including the states of: AR, CA, GA, IL, IN,IA,LA, MA, MN, MT, NY, OH, PA, TX, VA and the countries of BELGIUM, BURKINA FASO, BRAZIL, CHINA, CANADA, SWITZERLAND, CHILE, COLOMBIA, CYPRUS, CZECH REPUBLIC, GERMANY, DENMARK, DOMINICAN REPUBLIC, EGYPT, SPAIN, FRANCE, UNITED KINGDOM, CHANA, GUADELOUPE, GREECE, GUYANA, HONG KONG, CROATIA, HUNGARY, INDONESIA, ISRAEL, ITALY, KUWAIT, LEBANON LUXEMBURG, LIBYAN, MOROCCO, MARTINIQUE, MEXICO, NETHERLANDS, NORWAY, PANAMA, PERU, POLAND, PORTUGAL , REUNION, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SUDAN, SWEDEN, SINGAPORE, SLOVENIA, SLOVAKIA, SIERRA LEONE, THAILAND, TUNISIA, TURKEY, VENEZUELA, MACEDONIA, PALESTINE, SOUTH AFRICA. | Terminated | February 04, 2014 | 743 | 1296 | Computer
Z-1100-2013 | 64763 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE Healthcare, Discovery MR750w 3.0T System. |  |  | April 23, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential issue due to image flip in frequency encoding direction associated with a specific workflow of your Discovery MR750w system.Image flip issue in frequency encoding direction (including S/I, A/P and R/L) can result in incorrect localization ofanatomy and pathological findings. | Software design | GE Healthcare sent an Important Product Information letter dated April 23, 2012. The letter described the Issue, Instructions, Affected Product Details and Product Correction. For questions contact your local service representative. |  | 46 | 46 | Worldwide Distribution - USA (nationwide) including the states of: AZ, CA, CO, FL, IL, MD, NJ, NY, OH, OK, TX. and the countries of Australia, Belarus, Canada,France, Germany, India, Japan, Lebanon, New Zealand, Saudi Arabia, Singapore, Spain, Sweden, Switzerland and United Kingdom. | Terminated | March 14, 2014 | 690 | 46 | Computer
Z-1695-2013 | 65459 | No | HRY | Knee joint femorotibial metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | REF 154200 Oxford Fixed Bearing Partial Knee |  |  | June 07, 2013 | 2013 | Biomet, Inc. | Biomet investigation found the design of the tibial bearing creates the potential for anterior impingement between the bone and the poly bearing component in terminal extension. | Device Design | The firm, Biomet, sent an "URGENT MEDICAL DEVICE RECALL NOTICE" letter dated June 7, 2013 to its consignees/customers. The letter identified the device being recalled, the reason for the recall, potential adverse events associated with the recall and actions to be taken. The consignees/customers were instructed immediately locate and remove the identified device(s) listed in the letter; carefully follow the instructions on the enclosed "FAX Back Response Form"; complete and return the FAX Back Response Form to 574-372-1683 prior to return of product within three (3) business days; return Goods to Biomet, Inc 56 East Bell Drive Warsaw, IN 46580. To confirm receipt of this notice by call 574-372- 1570. Obtain an RGA # by phone (574)372-6677 email rqarequest@biomet.com or through FAST; use priority carrier for your shipment; and if you have further distributed this product you MUST notify personnel of the letter. Questions should be addressed to 574-372-1570 M-F 8 a.m. to 5 p.m. |  | 26 | 26 | Worldwide distribution: USA including NC, IN, PA, and OH; and countries of: Canada and Netherlands. | Terminated | July 17, 2014 | 405 | 794 | Computer
Z-1686-2013 | 65590 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | da Vinci Si Surgeon Side Cart, model number SS3000; |  |  | June 08, 2013 | 2013 | Intuitive Surgical, Inc. | Factory testing on da Vinci Si Surgical Systems may not be in compliance with UL Standards as one of the testing devices was found to be working incorrectly. | Software design | Intuitive Surgical sent an Urgent Device Correction notice dated July 15, 2013, to all affected customers by Federal Express.The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to ensure that all affected personnel were fully informed of the notification and that the notification should be forwarded to other managers within their facility. Customers were also instructed to complete and return the attached Acknowledgment Form to acknowledge receipt of the notification. Customers with questions were instructed to call Customer Service at 800-876-1310, Option 3.For questions regarding this recall call 408-523-2244. |  | 30 | 30 total, all devices | Worldwide Distribution - USA including UT, CA, AZ, CO, FL, HI. MT, NY,TX and WA and one customer in Russia. | Terminated | January 22, 2014 | 228 | 30 | Computer
Z-1081-2013 | 64855 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Siemens brand ACUSON SC2000 Ultrasound System, Model 10433816, hardware that contains version 4 of the MPI board 10434200; |  |  | March 28, 2013 | 2013 | Siemens Medical Solutions USA, Inc. | When using the V5M transducer, the display of the temperature on the ACUSON SC2000 system can be lower than the actual temperature of the V5M transducer. This occurs only if the system hardware contains revision 4 of the MPI Board. There is a potential to cause esophageal burns in patients. | Device Design | Siemens Medical Solutions USA Inc. sent a CUSTOMER SAFETY ADVISORY NOTIFICATION letter dated March 28, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.A return response form is included for customers to complete and mail to the recalling establishment. Siemens stated additional follow up action will occur, once material is available, where Siemens field service engineers will visit each unit and replace the affected MPI 4 board.For further questions please call (650) 694-5747. |  | 120 | 120 units | Worldwide Distribution - USA and countries of Australia, Austria, Bosnia Herzegovina, Belgium, Finland, France, Germany, India, Italy, Japan, Kuwait, Malaysia, Netherlands, Norway, P.R. China, Portugal, Rep.of Korea, Romania, Saudi Arabia, Spain, Sweden, Switzerland, Taiwan and United Kingdom. | Terminated | February 14, 2014 | 323 | 120 | Computer
Z-1694-2013 | 65487 | No | MUE | Full-field digital mammography system. | Radiology | Radiology | 510(k) | Siemens Mammomat Inspiration - full-field digital mammography |  |  | February 28, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Firm became aware of an unintended behavior when using the Mammomat Inspiration. The possibility exists that a malfunction may occur during Stereo or Tomo acquisitions if the tomo clutch was incorrectly adjusted. Although the possibility is rare, the system could move while a patient's breast is compressed resulting in a possible bruise. Firm released Update Instructions XP031/12/S for this iss | Device Design | Siemens issued an Update Instructions XP031/12/S by letter to the affected customers. For affected sites, a Siemens Service Engineer will inspect the unit and make any adjustments needed. Following the inspection and adjustments, a proper calibration will be completed. For questions call 610-448-3237 and 610-446-4634. |  | 3 | 3 | United States Nationwide Distribution in the states of ND, OH and SC. | Terminated | February 21, 2014 | 358 | 3 | Computer
Z-1727-2013 | 65163 | No | JMM | Hydroxyproline test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | B-CrossLaps/serum (B-CTx in serum) |  |  | April 25, 2013 | 2013 | Roche Diagnostics Operations, Inc. | Roche Diagnostics internal investigations in R&D; have shown that the claims for biotin interference and the sample stability given in the Elecsys Âż-CrossLaps (Âż-CL) assay package insert version 5, could not be reproduced. The Âż-CrossLaps test results contribute to the monitoring of antiresorptive therapy. False low or false normal results may lead to delayed diagnosis and treatment of osteoporosi | Device Design | The firm, Roche, sent an "URGENT MEDICAL DEVICE CORRECTION" (UMDC)letter dated May 2, 2013, to its consignees/customers via UPS Ground. The letter describes the product, problem and actions to be taken. The consignees/customers were instructed to apply the correct sample storage conditions as stated in the UMDC and to inform all relevant parties about the correct biotin interference claim. The consignees/customers were instructed to keep the UMDC on file for reference until the revised package insert is available. The consignees/customers were asked to complete and return a form via fax (1-888-943-4205). Consignees/customers were instructed to call Roche Diagnostics Technical Support, 24 hours a day, seven days a week, at 1-800-428-2336 if they had any questions about the recall. Non-responding accounts are monitored on an ongoing basis. |  | 3056 | 3056 units | Nationwide distribution: USA including states of: AL, CA, GA, FL, IL, IN, KS, KY, MA, MD, MI, MN, MO, NC, NY, OH, PA, TX, VA, VT, and WA. | Terminated | September 27, 2013 | 155 | 3056 | Computer
Z-1052-2013 | 64308 | No | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | ShapeMatch Cutting Guides |  |  | November 08, 2012 | 2012 | Stryker Howmedica Osteonics Corp. | With regards to ShapeMatch Cutting Guides using MRI Imaging, Stryker Orthopaedicshas become aware that there are potential issues associated with internal processesfor planning cases and producing ShapeMatch Cutting Guides. | Device Design | Stryker sent Product Correction Bulletins on 11/8/2012 dated 11/9/2012 and 11/19/2012 to all affected customers. Stryker sent an updated Urgent Medical Device Product Notification letters and Acknowledgement letters dated 1/24/2013 through 1/28/2013 to all affected customers via Fed Ex with return receipt: The letter identified the affect product, problem, potential hazards, risk mitigating factors and actions to be taken. For questions call 201-831-6878.Stryker issued press release on 4/10/2013. |  | 7868 | 7868 units | USA Nationwide Distribution | Open, Classified |  |  | 7868 | Computer
Z-1074-2013 | 64624 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | All OneTouch Verio IQ Blood Glucose Meters sold as: |  |  | March 11, 2013 | 2013 | Lifescan Inc | The Verio IQ meter will shut off and revert to set up mode at glucose values above 1023 instead of displaying EXTREME HIGH GLUCOSE. | Software design | Lifescan sent an Urgent Medical Device Voluntary Recall letter dated March 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to identify and hold all affected product in inventory and discontinue distributing. The letter states that a Lifescan sales representative will contact customers to arrange for pick up and replacement. |  | 710413 | US 710,413 | USA Nationwide Distribution | Terminated | August 03, 2015 | 875 | 710413 | Computer
Z-1661-2013 | 65493 | No | JOW | Compressible limb sleeve. | Cardiovascular | Cardiovascular | 510(k) | VenaFlow Elite 30B-B and 30BI-B |  |  | May 28, 2013 | 2013 | DJO, LLC | DJO, LLC has recently identified a product safety issue with the battery assembly for the VenaFlow 30B-B and 30 BI-B products. There is a potential for battery leakage which could result in thermal meltdown (fire). | Device Design | DJO sent a Customer Notification letter dated May 28, 2013 to Domestic customers via USPS certified return receipt and to International customers via UPS on June 14, 2013. The Customer Notification informed customers of the product safety issue, product description, and instructions to follow. DJO plans to install replacement batteries at no charge. For questions contact your local DJO Customer Service Representative. |  | 2491 | 2491 | Worldwide Distribution - USA Nationwide and countries of: Australia, Belgium, China, Guatemala CentroAmeria, France, Hong Kong, India,Italy, Korea, Lebanon, Malaysia, Morocco, Pakistan, Paraguay, The Netherlands, Romania, Singapore, Spain, Sweden, Thailand, Turkey, United Arab Emirates, United Kingdom and Venezula. | Terminated | February 13, 2014 | 261 | 2491 | Computer
Z-1104-2013 | 64693 | No | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | Aplio 500/400/300 Diagnostic Ultrasound System (TUS-500/400/300). |  |  | December 11, 2012 | 2012 | Toshiba American Medical Systems Inc | Toshiba Medical Systems has received reports of startup issues with the Toshiba TUS-500 andTUS-300 ultrasound systems. The startup issue is predominant when attempting to bring thesystem up from standby mode when the system is used remotely. It has been determined that thismay potentially cause damage to the system. | Device Design | Toshiba sent an "URGENT: MEDICAL DEVICE CORRECTION" letter dated December 11, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at (800) 421-1968 for questions regarding this letter. |  | 233 | 233 units | Nationwide Distribution. | Terminated | April 10, 2014 | 485 | 233 | Computer
Z-1103-2013 | 63856 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Integra XKnifeÂż Stereotactic Radiosurgery and Radiotherapy Treatment Planning, |  |  | December 10, 2012 | 2012 | Integra LifeSciences Corp. | Depending on the system configuration, a software error message in versions 5.0.1 and 5.0.2. occurs if a beam plan is transmitted from XKnife using DICOM-RT. | Software design | Integra initiated a voluntary recall on December 10, 2012 by providing a written Urgent: Medical Device Recall Notification to US consignees who have experienced the software issue due to the configuration of their system. The letter identifies the affected product, as well as, advises customers of the nature of the issue and how it affects treatment records.- Advises customers that they will be contacted when V 5.0.3 software, where the software error has been addressed, is available to schedule an upgrade.- Requests that customer consignees complete and return the Product Recall Acknowledgment Form.If customer have additional questions, they should contact the firm's service hotline at 1-888-772-7378. |  | 12 | 11 XKnife V5.0.1 and 1 XKnife V5.0.2 | Distributed in the states of IL, ID, LA, MA, NJ, NY, OH, PA, and SD. | Terminated | August 01, 2013 | 234 | 12 | Computer
Z-1102-2013 | 64547 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Intellispace Portal (ISP) Philips Medical Systems (Cleveland), Inc., |  |  | February 12, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | Philips was notified of a problem using the IntelliSpace Portal with software version 4.0.2 with the URL PACS integration. Philips was informed that when loading a study to an analysis application while using URL PACS integration configuration, after another study was opened by a review application, the Analysis monitor is updated with the study of the current patient, but the Review monitor conti | Software design | Philips sent a Urgent Medical Device Correction letter dated February 12, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.They advised their customers that when reviewing two disparate image datasets, not to only be focused on the actual image content but to also be alerted to the patient's study name and/or ID prior to interpretation.They recommend to manually exit the Review and Analysis applications before loading a new study to Intellispace Portal v4.0.2 from the PACS.For North America and Canada contact the Customer Care Solutions Center at (1-800) 722-9377. |  | 313 | 313 units | Worldwide Distribution - (USA) Nationwide Distribution including the states of AL, AR, AZ, CA, FL, GA, IA, IN, LA, MI, MN, MO, NC, NJ, NY, OH, OK, OR, PA, TN, TX, WA, WI &amp; WV. and the countries of Australia, Austria, Belgium, Canada, China, Czech Republic, Denmark, Egypt, Finland, France, France, Germany, Greece, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Latvia, Lebanon, Monaco, Netherlands, Norway, Poland, Portugal, Russia, Slovakia, South Africa, South Korea, Spain, Sweden, Switzerland, Thailand &amp; United Kingdom | Terminated | November 04, 2013 | 265 | 313 | Computer
Z-1647-2013 | 65207 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | 1) VITROS 5,1 FS Chemistry System (product code 6801375) and 2) VITROS 5,1 FS Chemistry System Refurbished (product code 6801890) |  |  | May 06, 2013 | 2013 | Ortho-Clinical Diagnostics | There is evidence that an electrical short within the systems connector has caused an accumulation of heat which has radiated back through the wiring harness causing the wires and connectors to melt. | Device Design | Ortho Clinical Diagnostics (OCD) sent an Urgent Product Correction Notification letters dated May 6, 2013 and response forms to all affected customers. Ortho Clinical Diagnostics informed their customers of the issue and advised if they detect the smell of melting plastic or observe smoke originating from your VITROSÂż System, to immediately power down the system and contact OCD for assistance. They were also advised to consider performing the backup procedure more frequently to help prevent the loss of data on their Systems. In the near future, an OCD Field Engineer will contact customers who are affected and perform a Modification (MOD) procedure that will replace their potentially affected Wire Connector(s). Foreign affiliates were informed by email on May 6, 2013 of the issue and instructed to notify their customers of the issue and actions. OCD also found that this wire connector is commonly used in many types of computers (i.e., medical devices, personal computers, etc.). On 29 May 2013, OCD expanded their recall to incllude ORTHO Assay Software (OAS) Servers. As a precaution, OCD sent notifications to encompass all ORTHOÂż Assay Software (OAS) Servers that potentially contain this type of wire connector. The letters (Ref. CL13-154) were sent via US Priority Mail to all customers who were shipped an affected OAS Server. Ortho Clinical Diagnostics has requested that customers with any questions or in need of additional information, please call Customer Technical Services at 1-800-421-3311. |  | 101 | Total 101 units (84 units for # 6801375- 26 units domestic &amp; 58 units foreign) and (17 units for # 6801890-15 units domestic, 2 units foreign) | Worldwide Distribution - USA Nationwide including the states of AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NY, OH, OK, OR, PA, SC, SD, TN, TX, VA, WA, WI &amp; WV, and countries of: Australia, Brazil, Canada, Chile, China, Colombia, India, Japan, Mexico, PR, Panama, Singapore, Venezuela, England, France, Germany, Italy, Spain | Open, Classified |  |  | 264 | Computer
Z-1646-2013 | 65139 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Computed Tomography X-ray systems |  |  | April 15, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | Philips Healthcare received reports from the field that certain Brilliance CT 64, Ingenuity Core, and Ingenuity Core128 systems running software versions 2.6.1 or 3.5.4 exhibited intermittent swirl-like ring artifacts that may appear on reconstructed images, which may impair the diagnostic quality of affected images. | Software design | Philips Healthcare sent an Urgent Medical device Correction letter dated April 15, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter states that Philips Healthcare is implementing software update versions 2.6.2 or 3.5.5 to resolve the issue and a Philips Field Service Engineer will contact each site to schedule the software installation. For questions contact your local Philips representative or local Philips Healthcare office. |  | 2348 | 2348 total units: 2182 Brilliance CT 64 units; 95 Ingenuity Core units; and 71 Ingenuity Core128 units. | USA Nationwide Distribution - including the states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY. | Terminated | January 12, 2015 | 637 | 2348 | Computer
Z-1105-2013 | 64737 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Siemens syngo WorkFlow SLR, Radiological Information System, Version VA31A SP1/SP2/SP3 |  |  | October 05, 2012 | 2012 | Siemens Medical Solutions USA, Inc | Under certain conditions, when using the Search functionality, the potential exists for the wrong patient's Interactive Documents to display when viewing a signed report in the single patient view. This can occur when the Interactive Document is viewed or auto displayed, dismissed by the user, and then a different patient and signed exam is selected in single patient view. The Notes Viewer will | Software design | Siemens Medical sent a Customer Safety Advisory Notice letter dated October 5, 2012 , to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.There will be a Hot Fix available.If you have any further questions please contact Siemens Customer Product Support at (1-888) 767-8326. |  | 37 | 37 | Worldwide Distribution - USA including the states of AL, CA, CO, DC, DE, FL, GA, HI, IL, IN, KS, KY, LA, MA, MD, MO, NC, NE, NH, NJ, NV, NY, OH, OR, PA, TN, TX, VA, WI, WV, and WY. and the countries of Canada and Australia. | Terminated | July 16, 2013 | 284 | 37 | Computer
Z-1620-2013 | 65301 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Bigliani/FlatowÂż The Complete Solution Fukuda Retractors Large Nonsterile 4305-19-40. |  |  | May 23, 2013 | 2013 | Zimmer, Inc. | This is a recall expansion from the 2010 recall which now includes seven additional lots that were not previously included. Zimmer, Inc. is initiating a recall of small and large Fukuda-type retractors due to fractures occurring at the grooves around the perimeter of the blades while being used during surgery. | Device Design | In 2010, Zimmer initiated a recall of the Bigliani/FlatowÂż Fukuda Retractors due to fractures occurring at the grooves around the perimeter of the blades. Zimmer is expanding this recall to include seven lots that were not previously included. On 5/23/13, an Urgent Medical Device Recall notification letter went to all consignees explaining the recall expansion and risk associated with the recall. All consignees are to locate and quarantine the device immediately. Please return the product along with the completed inventory certification form to Zimmer. If you have any questions regarding this recall, call Zimmer at 877-946-2761. |  | 722 | 722 units (includes 2010 market withdrawal) | Worldwide Distribution-USA (nationwide) including the states of NY, MI, MN, OH, IN, IL, VA, MD, PA, WA, KY, CT, RI, NH, TX, MT, and CA and the countries of Chile, China, South Korea, and Australia. | Terminated | December 18, 2014 | 574 | 722 | Computer
Z-1622-2013 | 65159 | No | LGX | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General Hospital | 510(K) Exempt | MAQUET Getinge Group Mobile Operating Table YUNO OTN GA 1433.01 EN 05. |  |  | April 26, 2013 | 2013 | Maquet Cardiovascular Us Sales, Llc | MAQUET has identified a potential issue with the extension device handle (part number 1433.62A1) and will make enhancements to the YUNO table (part number 1433.01F0).Identification of the Extension device issue: By fully actuating the lever of the extension device handle, the operator's skin can enter a gap between the handle lever and a recess in the handle bar. When the lever is released, | Device Design | MAQUET Getinge Group issued an Urgent Medical Device Recall letter/"Field Action Response Form dated April 26, 2013, via Federal Express to all affected customers. The recall notification identifies the product, issue with the product, and measures to be taken by the user. Customers were instructed that the firm would contact them to arrange for a loaner YUNO table and extension device while the table and extension device are upgraded. Upon completion of the upgrades, MAQUET would reinstall their upgraded YUNO table and extension device and remove the loaner device. Customers were instructed to complete and return the enclosed Field Action Response Form, acknowledging their receipt and understanding of the communication by e-mailing a scanned copy to FieldAction@maquet.com or via fax at 973-807-1658. Customers with questions were instructed to contact their MAQUET representative or MAQUET Technical Support at 888-627-8383, option 3, followed by option 1.For questions regarding this recall call 973-709-7691. |  | 37 | 13 Tables; 24 Extension devices | Worldwide Distribution - USA including Arizona, Georgia, North Dakota, South Carolina, Texas and Tennessee, Internationally to United Arab Emirates, Austria, Brazil, Canada, Switzerland, China, Germany, Spain, France, United Kingdom, Italy, Japan, Singapore, Thailand, Turkey, Serbia, and South Africa. | Terminated | April 29, 2014 | 368 | 37 | Computer
Z-1616-2013 | 65022 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Innova 3100 and Innova 4100 systems. |  |  | May 16, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare reported a potential safety issuedue to temporary loss of X-Ray Imaging Modes during Collimator Auto-Calibration which affects Innova 3100 and Innova 4100 Systems. | Software design | A Customer Letter was not originally sent. GE Healthcare has retrospectively re-assessed the reportability of the field action and has determined that the action should be reported to the FDA.GE Healthcare has already updated all of the affected Innova 3100 and Innova 4100 systems that were distributed.As an effectiveness check, they have tracked all service requests for the field action to completion. These service requests have recorded the repair of the affected units. Exceptions to completion of the service request have been recorded as well. Records of field action completion have been retained per their processes and procedures.If you have any questions, please call 262-548-2757.For questions regarding this recall call 262-513-4122. |  | 48 | 48 | Worldwide distribution - USA including AZ, CA, CO, FL, ID, IL, ME, NV,NY, NC, OK, TN, UT, and WI; and Internationally to Canada, Chile, France, Germany, Italy, Japan, Korea, Norway and Spain. | Terminated | March 13, 2014 | 1032 | 48 | Computer
Z-1609-2013 | 65364 | No | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | Brand Name: Coiled Tube Infusors (Portable Elastomeric Infusion Systems). Baxter's Infusor SV Elastomeric Infusion Device |  |  | June 07, 2013 | 2013 | Baxter Healthcare Corp. | Due to an increase in complaints for leaks at the distal male Luer and Luer cap. Baxter implemented enhancements to the blue winged cap to improve the seal between the cap and the Luer. | Device Design | Baxter sent an Urgent Produd Recall notification letter dated June 13, 2013, to all affected consignees via USPS.The letter was accompanied by a Customer Reply Form and Affected Product List (table). The letter stated affected products, problem description, risk to health (or health risks), actions to be taken by customers/users and contact information.The recall notification instructs customers to segregate affected product from current inventory and contact Baxter for instructions to return affected product. Customers with questions were instructed to call Baxter at 1-800-437-5176,For questions regarding this recall call 847-270-5556. |  | 601230 | 601,230 total | Worldwide Distribution - USA (nationwide), Puerto Rico, and Internationally to Brazil, Colombia, Panama, Mexico, Singapore Japan, Australia, New Zealand, Hong Kong, India, Thailand, Taiwan, and Korea. | Terminated | June 04, 2014 | 362 | 601230 | Computer
Z-1124-2013 | 64679 | No | NRA |  |  | Orthopedic | PMA | REF 32-422760 |  |  | March 20, 2013 | 2013 | Biomet U.K., Ltd. | Investigation of complaints found excessive use of the instrument could result in cracking on the weld between the head and the handle, which could lead to part of the head becoming detached from the handle. | Device Design | BIOMET sent an Urgent Medical Device Recall Notice on March 20, 2013, to all affected customers with directions to forward the notice if product was further distributed to Hospital Operating Room managers. The notice identified the recalled device, reason for the recall, and possible adverse events directing consignees to locate and discontinue use of the device. A FAX back Response form was to be returned to 574-372-1683 within 3 days of receipt prior to return of the product to Biomet, USA. Questionsregarding this recall should be addressed to 574 372-1570 M-F, 8a.m. -5 p.m. |  | 430 | 430 | Worldwide Distribution - USA (nationwide), Europe, Japan, Australia, Canada, Singapore, South Africa, and Costa Rica. | Terminated | May 04, 2015 | 775 | 430 | Computer
Z-1112-2013 | 64653 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris PC unit model 8015 with software version 9.12 |  |  | March 06, 2013 | 2013 | Carefusion Corporation | The recall was initiated because Carefusion has identified potential risk associated with bolus programming with the Alaris PC unit (model 8015) software version 9.12. Automating infusion pump parameter input from either the electronic medical record (EMR) or the Alaris Auto-ID module for a continuous infusion with a bolus dose option may result in a loss of information from the Guardrails Data s | Software design | Carefusion sent an Urgent Medical Device Recall Notification letter dated March 6, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers are required to confirm receipt of the notification by returning the Recall response card to Carefusion by postage-paid, self-addressed mail, fax, or email. The letter stated that a Carefusion representatives will contact customers by telephone to schedule field remediation. For questions call Carefusion Support Center at 888-562-6018, Customer Advocacy at 888-812-3266, and/or Technical Support at 888-812-3229. |  | 4090 | 4,090 total units | Worldwide Distribution USA Nationwide and the countries of Australia and Canada. | Terminated | May 08, 2014 | 428 | 4090 | Computer
Z-1113-2013 | 64460 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ACCU-CHEK Inform II Blood Glucose Monitoring System bar code scanner. |  |  | March 06, 2013 | 2013 | Roche Diagnostics Operations, Inc. | Roche investigation into complaints found Accu-Chek Inform II and CoaguChek XS Pro System barcode reading meters had the potential of intermittent erroneous decoding of patient identification if barcodes being read fail to meet size and/or quality standards. | Device Design | Consignees of both devices were notified by Urgent Medical Device Correction letters sent via UPS Ground (receipt required) beginning 3/06/2013. The notice contained recommendation not to use EAN-13 barcode system and ensure barcodes used meet quality standards; use operator and patient lists (downloaded from a data manager) in order to check any decoded identification against the respective list. The notice directed forwarding of the notice if the product was further distributed. Customers were asked to complete the attached fax form. Contact the Customer Care Service Center, 24 hours a day, seven days a week at 1-800-440-3638 for Accu check questions and 1-800-428-4674 for CoaguChek questions about the information contained in this UMDC. |  | 4467 | Both devices 4467 | Nationwide Distribution. | Terminated |  |  | 4467 | Computer
Z-1580-2013 | 65443 | No | LKK |  |  | General Hospital | PMA | Medtronic SynchroMed EL Programmable pumps, Models 8626, 8626L, 8627, 8627L, (The SynchroMed EL Pumps are supplied in 10 ml or 18 ml reservoir size.) Sterilized using ethylene oxide. The pump is part of an infusion system that stores and delivers a prescribed drug to a specific site. The drug is stored in the pump reservoir. Per a programmed prescription, the drug moves from the pump reservoir, through the pump tubing, catheter port, and catheter, to the infusion site. |  |  | June 03, 2013 | 2013 | Medtronic Neuromodulation | Medtronic Neuromodulation is providing Healthcare Providers with safety information and patient management recommendations related to the potential for electrical shorting internal to they SynchroMed infusion pump. An electrical short circuit in a feedthrough may present as a motor stall or low battery reset/alarm and lead to a loss of or reduction in therapy which may result in the return of und | Device Design | Medtronic sent an Urgent Medical Device Correction letter dated May 2013, to all affected customers. The letter described the issue of the pump's potentially shorting, how to identify if a pump is affected, and provided ongoing patient management recommendations. The Medtronic Rep would complete a Physician Notification Confirmation Form. A copy of the recall notification letter and the completed Confirmation Form is to be left with the physician. Customers with questions were instructed to call 1-800-707-0933.For questions regarding this recall call 763-526-9738. |  | 261109 | 261,109 total devices estimated implanted Worldwide | worldwide | Terminated | November 16, 2015 | 896 | 261109 | Computer
Z-1730-2013 | 60480 | No | DRY | Cardiopulmonary bypass on-line blood gas monitor. | Cardiovascular | Cardiovascular | 510(k) | CDI Blood Parameter Monitoring System 500 with Arterial and Venous blood parameter modules and Hematocrit/Saturation probe. |  |  | February 27, 2013 | 2013 | Terumo Cardiovascular Systems Corporation | Terumo discovered through its own testing that CDI 500 Blood Parameter Monitoring System software contains inaccuracies associated with a software algorithm, and a software update has been developed to address this issue and a premarket notification (K130425) has been submitted to the agency. Once clearance is received for this premarket notification from the US FDA, Terumo CVS will implement the | Software design | On 2/27/13, all consignees will receive a recall notification letter via Federal Express. Terumo CVS will correct all affected CDl Monitors. Correction to the devices will be completed concurrently with BPM replacements related to Z-2185/2189-11 RES 57194.Terumo continues to manufacture and ship devices affected by this recall. The long term corrective actions are associated with Recall# Z-2185/2189-11 RES 57194. As committed previously, we will re-notify all customers once we are ready to implement the fix for the BPM in the field.All units produced prior to the change in manufacturing will be included in thescope of this correction. Terumo continues to manufacturer units affected by this correction.All units returned for this correction will be tagged with a service traveler to indicate reason for correction and units will be properly segregated to prevent unintended re-shipment. |  | 15888 | 3972 units since 4/26/12 | Worldwide Distribution-USA (nationwide) including DC and Puerto Rico and the states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, and WV and the countries of Jordan, CANADA, AUSTRALIA, Japan ,Hong Kong ,UNITED ARAB EMIRATES (UAE), BELGIUM, South Korea, Malaysia, Mexico, Philippines, Singapore, Taiwan, and Thailand. | Terminated | February 17, 2016 | 1085 | 15888 | Computer
Z-1739-2013 | 63109 | No | LGW |  |  | Neurology | PMA | Eon Mini - 3721 (Charger). |  |  | July 26, 2012 | 2012 | St. Jude Medical | St. Jude Medical has informed your doctor that a number of cases have been reported in which discomfort associated with heating occurred at the device site while patients were using the Charging System to charge their spinal cord stimulator. In three cases, patients received a burn to the skin (one second-degree and two first-degree burns) at the implant site. As of June 30, 2012, St. Jude Medica | Device Design | ST. Jude Medical sent an Important Medical Device Safety Information dated July 26, 2012 to affected customers. The letter identifed the affected product, problem and actions to be taken. For questions or concerns contact your ST. Jude Medical representative or St. Jude Medical Neuromodulation Division at 1-866-240-6741. |  | 70638 | 70,638 (combined total of the Eon and the Eon Mini systems) | Worldwide Distribution - USA Nationwide Including Puerto Rico and countries of: Argentina, Australia, Belgium, Brazil, Canada, China, Chile, Colombia, Ireland, Israel, India, Malaysia, Mexico, Saudi Arabia, Singapore, South Africa, South Korea, Switzerland and United Kingdom. | Terminated | November 18, 2015 | 1210 | 70638 | Computer
Z-1864-2013 | 64625 | No | KTI | Bronchoscope (flexible or rigid) and accessories. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | Inspira AIR Balloon Dilation system. |  |  | January 11, 2013 | 2013 | Acclarent, Inc. | Labeling correction for all sizes of the Inspira AIR Balloon Dilation System to include additional language and warnings as well as to clarify some instructional steps. | Device Design | Urgent Medical Device Labeling correction letters will be sent to all consignees through a delivery system that can track delivery. After CDRH review and corrections, Acclarent began issuing customer notification on August 7, 2013 |  | 1217 | 1217 | Worldwide Distribution-USA (nationwide) and the countries of Belgium, Canada, and Saudi Arabia. | Terminated | October 25, 2013 | 287 | 28660 | Computer
Z-1834-2013 | 65440 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AU5800 Clinical Chemistry Analyzer with Automation Ready Connection (All Software Versions), Model Numbers AU5811, AU5821, AU5831 and AU5841. This system is for in vitro diagnostic use only. |  |  | June 03, 2013 | 2013 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has identified a software error on AU5800 analyzers with the Automation Ready (AR) Sampler Kit. This software error results in the failure of the concentrated wash solution pump to stop, creating an overflow or leak on each analyzer unit. | Software design | An Urgent Product Correction letter dated 5/29/13 was sent on the week of June 3, 2013 to all the customers who received the AU5800 Clinical Chemistry Analyzer with Automation Ready (AR) Connection (All Software versions). The letter informed the customers of the problem identified and the actions to be taken. Customers were instructed to complete and return the enclosed Response Form within 10 days and to share the information with their laboratory staff. Customers with questions are instructed to contact the 24 hours Call Center (Hotline) at (800) 854-3633 in the US and Canada. |  | 76 | 76 units | Worldwide Distribution, including Nationwide (US) and the countries of China, Italy, Japan, Canada, Turkey, United Kingdom, Singapore, Israel, and the Netherlands. | Terminated | August 23, 2014 | 446 | 76 | Computer
Z-1791-2013 | 65579 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Discovery NM 630, dual detector free-geometry integrated nuclear imaging system. |  |  | June 13, 2013 | 2013 | GE Healthcare, LLC | GE became aware of an incident at a VA Medical Center facility in the US in which a patient died due to injuries sustained while being scanned on an Infinia Hawkeye 4 Nuclear Medicine System. UPDATE: On July 03, 2013 GE notified hospitals that they are recalling several Nuclear Medicine Imaging Systems because serious injuries or deaths could occur due to the failure mode associated with this re | Device Design | GE sent an Urgent Medical Device Correction letter dated June 17, 2013 to all affected customers. The letter identified the affected product, recommended that qualified service personnel maintain the equipment and that Preventative Maintenance procedures are executed according to labeling. Also the Safety Chapter Sections should be re-reviewed with personnel to ensure proper operation of the equipment. GE notified customers again on July 03, 2013, via Urgent Medical Device Recall letter (including confirmation of delivery for US customers) and follow-up telephone call. Customers were instructed to cease using the affected system until GE Healthcare Field Engineer can complete an inspection of the system and perform any necessary repairs at no cost. A GE Healthcare representative will contact the hospitals to arrange for the inspection. |  | 108 | 108 | Worldwide Distribution - US Nationwide including Guam, Puerto Rico and the countries of : Afghanistan, Algeria, Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, Finland, France, Georgia, Germany, Greece, Hong Kong, Hungary, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jordan, Korea (Republic of), Kuwait, Latvia, Lebanon, Lithuania, Malaysia, Malta, Martinique, Mexico, Morocco, Namibia, Netherlands, New Zealand, Norway, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Puerto Rico, Romania, Russia, Saudi Arabia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Ukraine, United Kingdom, Uruguay, Venezuela, Vietnam. | Terminated | June 19, 2014 | 371 | 3454 | Computer
Z-1062-2013 | 64200 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Flexible Grip which is part of the Synthes Universal Nail System. |  |  | January 11, 2013 | 2013 | Synthes USA HQ, Inc. | Synthes is initiating a voluntary recall of the Flexible Grip (Part number 355.28) which is part of the Universal Nail System, due to the potential lack of the ability to thoroughly clean the instrument. | Device Design | Synthes sent an "URGENT NOTICE: MEDICAL DEVICE RECALL" dated January 11, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. The letter requested consignees examine inventory, remove the recalled products from use and return them to Synthes. If you have any questions, contact the firm at 610-719-5450. |  | 328 | 328 | Nationwide Distribution including the states of AZ, CA, CO, FL, IA, ID, GA, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, NC, NH, NJ, NV, NY, OH, PA, RI, TN, TX, UT, VT, WA, WI, and WV. | Terminated | August 31, 2015 | 962 | 328 | Computer
Z-1813-2013 | 65328 | No | DSQ |  |  | Cardiovascular | PMA | Heartware HVAD Pump Implant Kit (Heartware Ventricular Assist Device) |  |  | May 16, 0013 | 13 | HeartWare Inc | Potential to electrostatic discharged (ESD) | Device Design | HeartWare sent an Urgent Medical Device Correction Safety Notification letter on May 16, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to educate their patients about the HeartWare Controller and instructions for equipment handling stated in the IFU and recognized practices to avoid electrostatic discharge. The direct accounts were also told to scan and return the signed Acknowledgment Form via fax to 305-364-2665 or e-mail to quality@heartwareinc.com.Customers with questions were instructed to contact their HeartWare Representative or HeartWare Customers Service.For questions regarding this recall call 305-364-1575. |  | 76 | 76 (US) and 82 (International) | Worldwide Distribution - USA including AL, AZ, CA, DC, DE, FL, GA, IA, IL, IN, KY, LA, MA, MD, MI, MN, MO, NC, NJ, NY, OH, OR, PA, TN, TX, UT, VA, WA, and WI. Internationally to:Australia, Austria, Belguim, Brazil, Canada, Chile, Croatia, Czech Republic, France, Germany, Greece, India, Israel, Italy, Japan, Lebanon, Lithuania, Luxemburg, Malaysia, Netherlands, New Zeland,Norway, PolandSingapore, South Africa, Switzerland, Turkey, and United Kingdom. | Open, Classified |  |  | 76 | Computer
Z-1818-2013 | 65718 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | ViSi Mobile Chest Sensor (3 lead-wire ECG, AAMI), Catalog No. 93-10014 |  |  | May 23, 2013 | 2013 | Sotera Wireless, Inc. | Sotera Wireless, Inc. is recalling the ViSi Mobile Monitoring System due to the potential for the Chest Sensor to cause a blister wound. | Device Design | Sotera sent an Urgent Product Recall letter dated May 23. 2013, to all affected customers. The letter informed the customers of the problems identified and the actions to be taken. Customers were instructed to report any adverse reactions experienced with the use of this product and or quality problems to the FDA's MedWatch Program by: www.fda.gov/medwatch, (800) FDA-1008, fax (800) FDA-0178. Customers were also instructed to contact Sotera Wireless Technical Support at (866) 794-5526 for questions. |  | 75 | 75 units | US distribution including CA and UT. | Terminated | September 09, 2013 | 109 | 165 | Computer
Z-1814-2013 | 65666 | No | NSX |  |  | General Hospital | Contact ODE | Picis ED Pulsecheck - EMR Software Application - 2125, Software Versions: 5.2 and 5.3. The application stores patient information in a database, and it may analyze and/or display the data in different formats for evaluation by healthcare professionals for informational purposes. |  |  | June 21, 2013 | 2013 | Picis Inc. | Notes associated with prescription are not printed to the prescription or to the patient chart. | Software change control | Initial customer notifications were sent via email on June 21, 2013 informing consignees of the recall and providing further instruction regarding the software solution. Consignees were provided with recommended actions until they receive the necessary update. |  | 35 | 35 | Nationwide Distribution, including the states of: AK, AR, AZ, CA, CO, DC, DE, FL, GA, ID, IN, MA, MD, MO, NH, NJ, OH, OR, SC, TN, WA, and WV. | Terminated | December 19, 2014 | 546 | 35 | Computer
Z-1815-2013 | 63935 | No | IEO | Automated tissue processor. | Pathology | Pathology | 510(K) Exempt | Leica ASP6025 Vacuum Tissue Processor |  |  | November 13, 2012 | 2012 | Leica Microsystems, Inc. | The Instruction for Use 1v9 Ref F and all previous versions specify incorrect dimensions for the Ready To Use (RTU) bottles from other suppliers than Leica. In case of a pressure failure of the device in combination with a bottle with incorrect dimensions, overflow can cause contamination of reagents. Additionally, the current instructions for use does not highlight the importance to check the filling levels of all reagent bottles (RTU and system bottles) prior to every process run.Without this important information, risk of tissue damage or loss is possible. Therefore the missing information will be distributed to all customers by a Field Service Notice. | Device Design | Leica Biosystems sent an Urgent Field Safety Notice dated December 12, 2012, to all affected customers. The notification included a description of the problem and a recommendation that the personnel who use the equipment only use Leica Ready To Use (RTU) bottles that are already supplied as a standard delivery item with the ASP6025 machine. The personnel are further reminded that reagent bottle levels need to be checked and corrected when necessary. Customers with questions or questions regarding this recall call 847-405-5413. |  | 128 | 128 machines | Worldwide Distribution - USA (nationwide) and internationally to Australia, Belgium, Brazil, Chile, France, Germany, Ireland, Italy, Japan, Malaysia, Mexico, Netherlands, Philipines, Singapore, South Africa, Spain, Switzerland, Thailand, Turkey, and United Kingdom. | Terminated | January 05, 2017 | 1514 | 128 | Computer
Z-1050-2013 | 64386 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ABX Micros IM2, Part# M60CSB110EN04 and D00A00202 |  |  | February 15, 2013 | 2013 | Horiba Instruments, Inc dba Horiba Medical | Horiba Medical is recalling ABX Micros IM2 Data Management System because it may display incorrect patient test results under unanticipated operating conditions. | Device Design | Horiba Medical sent an Urgent Safety Notice dated February 15, 2013, to all affected customers. The letter informed the customers that Horiba Medical is informing all ABX Micros IM2 Data Management System customers of their intention to remove the ABX Micros IM2 product from the market. The letter informed the customers of the problem identified and the action to be taken. Customers were instructed to complete and return the enclosed response form within 10 days.For questions regarding this recall call 949-453-0500, ext 208. |  | 180 | 180 units | Nationwide Distribution | Terminated | August 11, 2014 | 542 | 180 | Computer
Z-1690-2013 | 65420 | No | LKK |  |  | General Hospital | PMA | MedStream Programmable Infusion Pump- Implantable for the intrathecal delivery of Baclofen or Morphine |  |  | June 07, 2013 | 2013 | Codman &amp; Shurtleff, Inc. | Miscalibrated Fill Level Sensor may affect dosing level | Device Design | Codman issued a Medical Device Correction Notification to Medical Facilities dated June 7, 2013 via FedX to inform them of the potential problem and a Worksheet (instructions) to Identify Pumps with Miscalibrated FLS during each patient's next scheduled refill session or sooner if patient is symptomatic. If a pump is identified as having a miscalibrated FLS, notify Codman Neuro using the same worksheet. In addition a sample letter was provided to Physi cans to provide to patients on the system.Follow the recommendations for managing patients who are identified as having a pump with a miscalibrated FLS.For additional information,contact your Codman Neuro representative, or contact Codman Neuro Clinical Support at 1-800-660-2660. Please report any malfunction or adverse event related to this device to Codman Neuro at 1-866-491-0974, Option 2. |  | 737 | 2 units US and 735 OUS | Worldwide Distribution-USA (nationwide) including the states of FL, NC, NJ, and TX, and the countries of Austria, Belgium, Spain, Great Britain, Germany, Italy, Ireland, Luxemburg, Russia, Saudi Arabia, Portugal, Greece, Egypt, Switzerland, and Turkey. | Terminated | April 21, 2015 | 683 | 1102 | Computer
Z-1836-2013 | 63108 | No | LGW |  |  | Neurology | PMA | Eon Mini - 3788 (IPG), |  |  | July 26, 2012 | 2012 | St. Jude Medical | As part of St. Jude Medical's routine tracking and product monitoring, they have received a total of 214 reports, out of 34,617 Eon Mini IPGs, that lost the ability to communicate or recharge due to an inner battery weld issue resulting in loss of pain relief and subsequent explant. Explant surgery, as with any surgery, presents a risk to patient health. Adverse events associated with an unplanne | Device Design | St. Jude Medical sent a Important Medical Device Recall Advisory dated July 26, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.To further assist in your patient care we are providing you with a list of all serial numbers we show have been distributed to you. ( See Attachment A ).Further questions please call (972) 309-2154. |  | 34617 | 34,617 | Worldwide Distribution--USA (nationwide) including Puerto Rico and the countries of Argentina, Australia, Austria, Belgium, Brazil, canada, Colombia, Cyprus, Czech Republic, Denmark, Finland, France, Germany, Greece, India, Ireland, Israel, Italy, Japan, Mexico, Netherlands, Saudi, South Africa, Spain, Sweden, Switzerland, Turkey and United Kingdom. | Terminated | August 21, 2015 | 1121 | 34617 | Computer
Z-1833-2013 | 65755 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | GE Healthcare, VOLUSON E6, E8, E8 EXPERT, E10 DIAGNOSTIC ULTRASOUND SYSTEM. |  |  | July 13, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has recently become aware that during use of the Voluson E6 or E8 systems, that an error message Out of Memory can be displayed which could cause the system to freeze. This is an intermittent behavior that also requires a certain workflow. This issue can be resolved by rebootingthe system. | Software design | No consignee letter was sent out.Starting on July 13, 2012, Field Engineers visited consignees and carried on instructions as per Field Modification Instruction (FMI 79044A). The Engineer returned the system to the customer and ensure acceptance. Before leaving the customer site have the staff operate the system. Ensure that: all settings are customized, all peripherals are functioning and all communications are working properly. We plan to update the software on all of the affected GE VOLUSON E6, E8, E8 EXPERT, E10 DIAGNOSTIC ULTRASOUND SYSTEM that were distributed.As an effectiveness check, we will track all service requests for the field action to completion. These service requests will record the repair of the affected units. Exceptions to completion of the service request will be recorded as well. Records of field action completion will be retained per our processes and procedures.Further questions please call 262-513-4122. |  | 375 | 375 | Worldwide Distribution - US Distribution including the states AK, AZ, AR, CA, CO, CT, DE, FL, GA, IL, LA, MD, MI, MN, MO, NE, NV, NM, NY, NC, OH, OR, PA, SC, SD, TN,TX, VA, WA and WI., and the countries of ALGERIA, AUSTRALIA, AUSTRIA, BAHRAIN, BELGIUM, BRIL, CANADA, CHINA, COLOMBIA, DENMARK, EGYPT, FINLAND, FRANCE, GERMANY, INDIA, IRAQ, ISRAEL, ITALY, JAPAN, JORDAN, KOREA, REPUBLIC OF MOLDOVA, NETHERLANDS, NORWAY, POLAND, ROMANIA, RUSSIAN FEDERATION, KINGDOM OF SAUDI ARABIA, SLOVAKIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, UZBEKISTAN, VENEZUELA. | Terminated | November 10, 2014 | 850 | 375 | Computer
Z-1837-2013 | 65498 | No | BXB | Powered exercise equipment. | Physical Medicine | Physical Medicine | 510(K) Exempt | Omnicycle Elite; US Importer and Distributor: Accelerated Care Plus, Reno NV. Manufactured by Medica Medizintechnik GmbH, Hochdorf, Germany. The product is indicated for powered exercise equipment intended for medical purposes. |  |  | June 12, 2013 | 2013 | Accelerated Care Plus Corporation | An anomaly in the firmware that may cause Omnicycle Elite to operate at a speed other than what was previously set when in reverse mode. | Software design | Letters were issued on June 12, 2013 that included a memory stick with self loading software so consumers can update the device. Letters were sent by trackable method. Customers with questions or concerns should call Customer Support at 800-350-1100. |  | 1001 | 1001 units | Nationwide Distribution. | Terminated | October 23, 2013 | 133 | 1001 | Computer
Z-1820-2013 | 65591 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | LPS Lower Extremity Dovetail Intercalary component. |  |  | July 11, 2013 | 2013 | DePuy Orthopaedics, Inc. | The LPS Lower Extremity Dovetail Intercalary component has the potential for fracture of the female component when exposed to certain physiological loads. | Device Design | On 7/11/13, URGENT MEDICAL DEVICE RECALL NOTIFICATIONS were sent to consignees: Depuy Distributors, hospitals, and surgeons. End-user Hospital Customers will submit reconciliation forms to DePuy. Possible clinical implications of device failure are stated in these notifications, including poor mechanics and loss of function, pain, component dislocation, adverse tissue reaction, bone damage, potential need for revision surgery and the risks associated with this type of surgery. A patient letter will be included with the surgeon letter to assist in surgeon's notification and discussion with patients. The devices will be returned to DePuy Warsaw through the normal DePuy Returns process, to attention of Returns and marking H13-13 on the outside of the box.If a patient presents with a fractured LPS" Lower Extremity Dovetail Intercalary component with well-fixed proximal and distal stems and the surgeon determines LPS" Lower Extremity Dovetail Intercalary component is the best treatment option, DePuy plans to make the LPS" Lower Extremity Dovetail Intercalary component available through authorized approvals. Directions on how to purchase the component are provided in the Recall Communications.Any questions or concerns about the recall, please contact Kim Earle, Recall Coordinator, at 574-371-4917 (M-F; 8am-5pm EDT) |  | 2262 | 58 U.S. / 39 OUS | Worldwide Distribution-USA including the states of FL, GA, ID, KS, KY, MI, MO, NC, NV, NY, OH, OK, PA, TN, TX, VA, and WA and the countries of AUSTRALIA, CANADA, ENGLAND, GERMANY, INDIA, COUNTY CORK, IRELAND, and TURKEY. | Terminated | July 10, 2014 | 364 | 2262 | Computer
Z-1859-2013 | 65603 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo Dynamics Information System (sDIS) versions 9.5 and VA10A. |  |  | June 10, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Siemens initiate this recall due to a potential issue when using the sDIS matching rules of the syngo Dynamics Information System version 9.5 or VA10A with modality device configured to end study at association close. Images or DICOM SR objects sent from a modality device to syngo Dynamics may not be saved when the modality is configured at syngo Dynamics to end study at association close, and whe | Software design | Siemens sent a Customer Safety Advisory Notice dated June 4, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. 1) With sDIS implemented, use the DICOM Modality List (DMWL), without sDIS matching rules enabled OR 2) Do not implement sDIS. Update instructions were also provided. The reported issue will be resolved in a modification to syngo Dynamics version 9.5 and syngo Dynamics VA10A which will be released in the near future.We regret any inconvenience that this may cause , and we thank you in advance for your understanding.For further questions please call (610) 219-6300. |  | 9 | 9 | US Distribution only including the states of PA, OH, UT, CA, NY, MO, and TX. | Terminated | February 24, 2015 | 624 | 9 | Computer
Z-1869-2013 | 64987 | No | FSM | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Caddy - PedFuse, Screw Caddy, 8MM and 9MM, Part Number: 11-80009 Rev. B |  |  | October 05, 2012 | 2012 | SpineFrontier, Inc. | Two sections that are designated for the 8mm and 9mm screws were switched; the holes that hold the 8mm screws are designed for 9mm screws, and the holes that hold the 9mm screws are designated for 8mm screws. | Device Design | The firm, SpineFrontier, sent an Advisory Notice entitled "URGENT May 1, 2013, PedFuse Screw Caddy Removal" dated May 1, 2013 to their consignees/customers. The notice described the product, problem and action to be taken. Consignees/customers were informed that a removal has been initiated to remove the parts from the field to ensure that all sets contain correctly marked screw caddies. Consignees/customers wre instructed to complete and return the URGENT PedFuse Screw Caddy Removal ACKNOWLEDGEMENT FORM via fax to: SpineFrontier, Inc., 978-232-3991 or email: customerservice@spinefrontier.com. Replacement caddies will be shipped upon receipt of completed acknowledgement form and RMA request for part return.Please feel free to contact SpineFrontier at 978-232-3990 with any questions. |  | 11 | 11 caddies | US distribution in states of: CO, DC, MA and TX. | Terminated | July 29, 2015 | 1027 | 11 | Computer
Z-1797-2013 | 65468 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart MRx Monitor/Defibrillator, |  |  | June 19, 2013 | 2013 | Philips Healthcare Inc. | The MRx may deliver a non-synchronized cardioversion shock resulting in the delivery of incorrect therapy, which may induce ventricular fibrillation. | Software design | Philips sent an Urgent Medical Device Correction letter on June 19, 2013, to all affected customers. The letter informed customers of the issue, identified details of the units affected, gave instructions on actions to be taken by the customer and identified what action Philips planned to take to remedy the issue.Philips will perform a software upgrade and replace the label on the unit that outlines the steps for Synchronized Cardioversion with a revised label for all affected units at no charge to the customer.Philips is asking customers to follow the Action to be Taken by Customer/User section of the Urgent Medical Device Correction notification/Field Safety Notice: You can continue to use your MRx prior to receiving the software upgrade and device label revision, provided that you follow the instruction provided.Customers with questions were instructed to contact their local Philips representataive or call 1-800-722-9377.For questions regarding this recall call 978-687-1501. |  | 66150 | 66,150 units | Worldwide Distribution - USA (nationwide) and countries of: AFGHANISTAN ALBANIA ARGENTINA ARUBA AUSTRALIA AUSTRIA AZERBAIJAN BAHRAIN BANGLADESH BELGIUM BOLIVIA BOSNIA AND HERZEGOWINA BRAZIL BRUNEI DARUSSALAM BULGARIA CAMBODIA CHILE CHINA COLOMBIA COSTA RICA CROATIA (local name: Hrvatska) CYPRUS CZECH REPUBLIC DENMARK ECUADOR EGYPT EL SALVADOR ESTONIA ETHIOPIA FINLAND FRANCE GERMANY GREECE GUADELOUPE HONG KONG ICELAND INDIA INDONESIA IRAQ IRELAND ISRAEL ITALY JAPAN JORDAN KAZAKHSTAN KENYA KOREA, REPUBLIC OF KUWAIT LAOS LATVIA LEBANON LIBYAN ARAB JAMAHIRIYA LITHUANIA LUXEMBOURG MACAU MACEDONIA MALAYSIA MALTA MEXICO MONGOLIA MOROCCO NAMIBIA NEPAL NETHERLANDS NETHERLANDS ANTILLES NEW CALEDONIA NEW ZEALAND NICARAGUA NIGERIA NORWAY OMAN PAKISTAN PALESTINE PANAMA PAPUA NEW GUINEA PARAGUAY PERU PHILIPPINES POLAND PORTUGAL PUERTO RICO QATAR REUNION ROMANIA RUSSIAN FEDERATION SAUDI ARABIA SERBIA SINGAPORE SLOVENIA SOUTH AFRICA SPAIN SRI LANKA SWEDEN SWITZERLAND TAIWAN, PROVINCE OF CHINA TANZANIA, UNITED REPUBLIC OF THAILAND TRINIDAD AND TOBAGO TUNISIA TURKEY UGANDA UKRAINE UNITED ARAB EMIRATES UNITED KINGDOM UZBEKISTAN VENEZUELA VIET NAM and YEMEN. | Open, Classified |  |  | 66150 | Computer
Z-1795-2013 | 65604 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Axiom Aristos MX Solid State Flat Panel Digital Imager, Model number 5895003. The device is used for diagnostic X-ray imaging. |  |  | June 05, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Siemens issued a Field Safety Notice about the potential hazard to patients or operators when using the function "table top float" of the AXIOM Aristos MX. When repositioning a patient lying on the table top and releasing the table top, it is possible that the fingers of the patient or operator are between the table top and the detector tray and if the table top is moved, the fingers can be squee | Device Design | On 6/5/13, consignees were notified by letter of a "Field Safety Notice" and received instructions about how to avoid the potential risk of this issue while using the device. Siemens also rolled out a modification of AXIOM Aristos MX that will avoid these potential injuries in the future. |  | 401 | 401 | Nationwide Distribution. | Terminated | February 24, 2015 | 629 | 401 | Computer
Z-1794-2013 | 65259 | No | IOQ | Powered flotation therapy bed. | Physical Medicine | Physical Medicine | 510(K) Exempt | TotalCare Bed used with AccuMax Quantum VPC mattress |  |  | May 30, 2013 | 2013 | Hill-Rom, Inc. | When the bed is put into the chair egress position, the magnets on the mattress may not hold the bottom of the mattress flat to the bed frame, and the mattress foot section will extend away from the bed at an angle. A patient attempting to get back into the bed may miss the seat and sit on the foot section instead. The patient can slide down to the floor, or fall, with possible injury. | Device Design | The firm, Hill-Rom, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated May 29, 2013 was sent to all consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed to do the following: Until the correction is implemented, after placing the bed into chair position, regularly check the foot section of the mattress. If necessary, push it back against the bed frame until the magnets hold it in place before allowing the patient to get back into or out of bed. In addition, forward a copy of this letter to any other facility personnel you deem appropriate.If you have any questions concerning this notification, please contact your distributor, your Hill-Rom representative, or Hill-Rom Technical Support at 800-445-3720 (Mon-Thurs, 8 am-6:30 pm; Friday, 8 am-6 pm). |  | 4121 | 4,121 Units | Worldwide distribution: USA (nationwide) and countries of: Canada, Colombia, Nigeria, Kuwait and Saudi Arabia. | Terminated | April 17, 2014 | 322 | 4121 | Computer
Z-1782-2013 | 65651 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | ARTIS zee and zeego x-ray, angiographic system with Software Artis VC20x/VC21A that have Patch 7 and a 3D and DSA license installed. |  |  | June 26, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Siemens issued a customer safety advisory notice and field safety corrective action for a potential issue on Artis systems with Artis VC20x/VC21A that have Patch 7 and a 3D and DSA license installed. The issue affects-DSA rotational acquisition. The restriction is related to any rotational 3D-acquisition program with real-time subtraction (e.g. 3D DSA). During the fill phase of the rotational aqui | Software Design Change | Siemens sent a Customer Safety Advisory Notice dated June 26, 2013, to all affected customers. The letter identified the product, the problem, and the action needed to be taken by the customer. Customers were strongly recommend to carefully evaluate one of the following options before running 3D rotational acquisition programs that include real-time subtraction until a patch is installed on their system: 1) position the monitor in a way such that contrast advancement can be monitored during the fill phase on the live monitor in the control room.2) For applications not requiring actions based on the contrast flow during the fill phase, the contrast may not need to be monitored during the fill phase. Although the monitor is not showing the expected subtraction, the acquisition is working correctly. There is no need to interrupt, but continue and finish the workflow.3) Perform 3D rotational programs that include real time subtraction from the exam room. The display on the live monitor in the exam room is not affected by this problem. To resolve the issue, the update AX028/13/P (VC20x/VC21B), will be made available around July 2013, and scheduling an appointment for implementation of the update with Siemens Customer Service organization was recommended. If this device/equipment is no longer in theirr possession, customers were asked to forward this safety notice to the new owner and inform Siemens of the new user's identity. |  | 29 | 29 | Nationwide Distribution and Puerto Rico | Terminated | February 05, 2015 | 589 | 29 | Computer
Z-1066-2013 | 64505 | No | OQG | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stelkast Cross-over Acetabular Shell and Liner Hip System, Part Numbers SC3458-52, SC3474-52, SC3459-52, and SC3460-52. |  |  | January 24, 2013 | 2013 | Stelkast Co | Firm was made aware of incidents relating to the difficulty of properly engaging and securing the Cross-Over 36mm liner into the Cross-Over shell. | Device Design | The firm initiated their recall of this product on January 24, 2013 by notifying their consignees by telephone informing them of the recall. |  | 254 | 254 | Nationwide Distribution including the states of NV, OK and PA. | Terminated | October 08, 2014 | 622 | 254 | Computer
Z-1785-2013 | 63947 | No | LGW |  |  | Neurology | PMA | Eon and Eon Mini charging system, model numbers Eon -3701 (Wall Charger), 3711 (Portable Charger) and Eon Mini -3721 (Charger); This is a continuation of the firm's July 2012 recall. |  |  | November 30, 2012 | 2012 | St. Jude Medical | There is issues of warmth or heating at the implant site during charging for the Eon Mini implantable pulse generators (IPGs). In the July 2012 letter St. Jude Medical informed customers that they would be implementing design improvements to the charger to address possible increased energy dissipation when the charger is misaligned with the IPG or the IPG is implanted too near the surface of the | Device Design | St Jude Medical sent an Important Medical Device Safety Information letter dated November 30, 2012 to affected customers. The letter identified the affected product, problem and actions to be taken. The letter informed customers that a replacement charger for the Eon Mini IPGs is now available and will be provided at no charge. For questions or concerns call St. Jude Medical at 877-650-0367. |  | 75663 | 75,663 | Worldwide Distribution - USA Nationwide including Puerto Rico and countries of: Austria, Australia, Belgium, Brazil, Canada, Chile, China, Colombia, Cyprus, Czech Republic, Denmark, Finland, France, Germany, Greece, Hong Kong, India, Ireland, Israel, Italy, Japan, Luxembourg, Martinique, Mexico, Netherlands, Poland, Russian Federal, Saudi Arabia, Singapore, South Africa, South Korea, Spain, Sweden, Switzerland, Turkey and United Kingdom. | Terminated | November 04, 2015 | 1069 | 75663 | Computer
Z-1070-2013 | 64656 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX CV Reporting Cardiac Catheterization module and IMPAX CV Outbound Reporting. |  |  | March 15, 2013 | 2013 | AGFA Corp. | Baseline pulmonary capillary wedge (PCW) pressure values from IMPAX CV Reporting Cardiac Catheterization module were not correctly transferring into the hospital's electronic health record (EHR) system. | Software design | AGFA sent an "URGENT FIELD SAFETY NOTICE" letter dated March 15, 2013 to the consignees who use the IMPAX CV Reporting tool described for this issue. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at 877-777-2432, referencing HQ_10020006 for questions regarding this notice. |  | 26 | 26 | Nationwide Distribution including the states of AK, GA, IL, KY, MA, NY, NC, RI, SC, TX, and VA. | Terminated | January 09, 2014 | 300 | 26 | Computer
Z-1780-2013 | 65622 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Radiohead Access Point 1.4 GHz Radiohead Access Point |  |  | July 08, 2013 | 2013 | Philips Healthcare Inc. | IntelliVue Smart-hopping 1.4 GHz Access Point may experience unexpected signal loss of severalminutes impacting wireless connectivity | Software Design Change | Philips sent an Urgent Medical Device Correction letter dated July 2013 to affected customers. The letter informs customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer. Philips will perform a firmware upgrade (version C.00..07 for all affected customers at no cost. Philips is asking customers to follow the Action to be Taken by Customer/User section of the Urgent Medical Device Correction notification/Field Safety Notice.If you need any further information or support concerning this issue, please contact Philips Customer Support at 1-800-722-9377 |  | 20105 | 20105 units | USA Nationwide Distribution | Terminated | May 17, 2017 | 1409 | 20105 | Computer
Z-1069-2013 | 64654 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX CV 7.8 SU3 - OCR Service. |  |  | March 15, 2013 | 2013 | AGFA Corp. | Software design error in IMPAX CV 7.8 SU3 | Software design | AGFA Healthcare sent an "URGENT FIELD SAFETY NOTICE" letter dated March 15, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed to distribute this information within your facility to all individuals who need to be aware of this. Please fax back to (864) 421-1664 that the information was received and understood was requested from the customers.If you have any questions about this matter please contact your local AGFA Healthcare organization, or call (864) 421-1754. |  | 2 | 2 | Distribute in the state of FL and KY | Terminated | April 01, 2014 | 382 | 2 | Computer
Z-1774-2013 | 65688 | No | HQF | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | PASCAL and PASCAL Streamline (532 nm.577 nm); |  |  | March 27, 2012 | 2012 | Topcon Medical Laser Systems, Inc | When using the "Favorites" function for saved settings, there is potential for a mismatch between displayed power and the actual power output. This may result in over or under treatment. | Software design | Initial notification was sent on March 27, 2012 to all TMLS Distribution Partners and Sales Representative. This was followed by a follow up notification on June 4, 2012. A final notice was sent on February 11, 2013 to all those who had not sent documentation that the field action had been implemented. |  | 40 | 7 US, 33 foreign | Worldwide Distribution-USA and the countries of China, Hong Kong, Brazil, France, Japan, Netherlands, Pakistan, Malaysia, United Kingdom, South Korea, Thailand, Dubai, and South Africa. | Terminated | July 19, 2013 | 479 | 40 | Computer
Z-1771-2013 | 65630 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Spacelabs ARKON Anesthesia Workstation. |  |  | June 07, 2013 | 2013 | Del Mar Reynolds Medical, Ltd. | Spacelabs ARKON Anesthesia System is recalled due to software defect. The System Status Computer may incorrectly determine that a communication error has occurred. If this situation occurs, a buzzer will activate and a yellow triangle will be displayed on the System Status Computer screen that alerts the user of this error. | Software design | Spacelabs Healthcare sent an Urgent Medical Device Correction" letter, dated June 20, 2013, to all affected consignees. The letter identified the product, th problem, and the action to be taken by the consignee. Consignees were instructed to advise their staff of the situation. Consignees were also advised that Spacelabs Field Service personnel would be contacting them at their facility to schedule installation of a software update to resolve the issue at not cost to the consignee. Consignees with questions were instructed to call Technical Support at 1-800-522-7025 and select option 2.For questions regarding this recall call 800-522-7025. |  | 9 | 9 units (4 in the US and 5 outside US) | Worldwide Distribution - USA including South Carolina and internationally to Canada and Great Britian. | Terminated | August 29, 2013 | 83 | 9 | Computer
Z-1072-2013 | 64664 | No | JAI | Powered radiation therapy patient support assembly. | Radiology | Radiology | 510(k) | Precise Treatment Table. |  |  | March 14, 2013 | 2013 | Elekta, Inc. | The Column rotation can accidentally be moved to a different position which in the locked position. | Device Design | Elekta sent an "IMPORTANT FIELD SAFETY NOTICE" dated March 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact your local Elekta office for questions about this notice. |  | 443 | 443 | Nationwide Distribution and Puerto Rico. | Terminated | March 17, 2015 | 733 | 443 | Computer
Z-1743-2013 | 65652 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | TDHistology/Cytology |  |  | June 20, 2013 | 2013 | Technidata S.A. | A recall was initiated by Technidata because use of symbols (less than or equal to or greater than or equal to or any other symbol) in Microsoft Word reports when using electronic transmissions of results by of HPRIM (including HPRIM Medecin via Kermit) or HL7 must be absolutely prohibited. Such symbols are not supported by the Microsoft Word interface used by the Histology/Cytology software and | Software design | TD Technidata Medical Software sent a Medical Device Recall letter dated June 20, 2013, to one affected customer who received the TD-Synergy Histology/Cytology module. The customer was informed of the problems identified and the actions to be taken. Customers were instructed to complete the attached Reply Form and return it to the firm by fax to 514-270.2770. The customer were also instructed to contact Technidata Canada at (514) 270-7777 or at support-canada@technidata-web.com. |  | 1 | 1 site | US Distribution CA only. | Terminated | July 29, 2013 | 39 | 1 | Computer
Z-1745-2013 | 65561 | No | DZL | Intraosseous fixation screw or wire. | Dental | Dental | 510(k) | Multi Vector Distractor Pin Holding Clamp, limited bone stock |  |  | May 31, 2013 | 2013 | Synthes USA HQ, Inc. | Synthes is initiating a voluntary Medical Device removal for specified part numbers and lot numbers of the Mutis Vector Distractor Pin Holding clamp, limited bone stock. The firm received complaints that the clamp and nut construct would not engage the distractor arm threads and subsequently bind. It is not likely that a serious adverse event will result due to a nonconforming clamp, however there | Device Design | Synthes sent a Urgent Notice: Medical Device Removal letter dated May 31, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers. The firm received complaints that the clamp and nut construct would not engage the distractor arm threads and subsequently bind. It is not likely that a serious adverse event will result due to a nonconforming clamp, however, there is a remote potential that additional intervention would be required due to a non-functioning clamp.Synthes requested the product to be removed from inventory and returned to Synthes by calling 1-800-479-6329 to obtain a Return Authorization Number. Direct Accounts were asked to complete and return the Verification Section at the end of the letter indicating the quantity of affected product was found. Negative responses were also requested.If you have any questions please call (610) 719-5450. |  | 121 | 121 | US Nationwide Distribution including the states of MA, TX, MN, SC, UT, VA, and NY. | Terminated | August 31, 2015 | 822 | 121 | Computer
Z-1750-2013 | 65620 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Siemens AXIOM Artis zee/zeego systems equipped with generator A100 Plus |  |  | January 31, 2013 | 2013 | Siemens Medical Solutions USA, Inc | During the course of product monitoring, Siemens became aware of a tolerance problem in the power supply of the generator control above a specific value, thus, the generator may fail. | Device Design | Siemens sent an Urgent Customer Information letter to their consignees. The letter identified the affected product, problem , update Instructions and actions to be taken. Siemens began to equip affected systems with a modified board assembly in January, 2013. |  | 23 | 23 | United States Nationwide Distribution - including the states of: CA, CO, FL, IA, IL, IN, MA, MD, MI, NC, NY, PA, TX, VA and WA and Puerto Rico. | Terminated | November 12, 2014 | 650 | 23 | Computer
Z-1736-2013 | 65406 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Integra PANTA Arthrodesis Nail System. |  |  | May 09, 2013 | 2013 | Integra LifeSciences Corporation | As a result of the investigation into an adverse trend of complaints for reports of misalignment of the calcaneal or tibial screws during use of the Panta Nail Support Device, Integra has decided to issue this voluntary medical device correction. The misalignment can result in the inability to properly insert the screws into the Panta Nail and create the potential need for revision surgery. | Device Design | Integra initiated a voluntary correction on May 9, 2013 by providing a written notification and copy of the revised Surgical Technique Manual as well as the specific highlighted sections of the manual addressing calcaneal and tibial screw alignment to all consignees that Integra records indicate have been shipped a PANTA Nail. The written notifications and Surgical Technique Manual have been delivered to US Consignees by traceable courier service and by traceable email or facsimile for Canada and Non-US Consignees. The notification informs customers of the nature of the issue; advises of potential adverse patient consequences that may occur; advises customers to review highlighted revisions of the Surgical Technique Manual to minimize the potential for misalignment of the screws and requests customers to complete and return the attached Acknowledgement Return Form. |  | 692 | 692 | Worldwide Distribution-USA (nationwide) and the countries of Austria, Australia, Belgium, Canada, Czech Republic, Denmark, Finland, France, Germany, Great Britain, Ireland, Israel, Italy, Netherlands, Portugal, Russia, Saudia Arabia, South Africa, Spain and United Arab Emirates. | Terminated | November 07, 2013 | 182 | 692 | Computer
Z-1077-2013 | 63974 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ELEKTA MOSAIQ SW version below 2.41 |  |  | December 20, 2012 | 2012 | Siemens Medical Solutions USA, Inc | A potential safety risk with the MOSAIQ OncologyInformation System that has been identified by the manufacturer IMPAC Medical Systems, Inc. | Software design | Siemens sent an Urgent Medical Device Correction letter with update instructions to all affected customers. The letter identified the affected product, problem and actions to be taken. For questions contact your responsible Siemens Service Engineer or Application Specialist. |  | 101 | 101 | Worldwide Distribution | Terminated | July 19, 2013 | 211 | 101 | Computer
Z-1588-2013 | 65137 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance CT 64, Ingenuity CT, Ingenuity Core, and Ingenuity Core128 scanners are whole body Computed Tomography X-ray systems. |  |  | March 21, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | If the operator selects a patient name from the worklist using the mouse and then selects a different patient name using the keyboard (arrow keys then press enter) the patient name that populates the exam information page is the patient name selected with the mouse. Also, it is possible to begin a normal clinical scan while the anti-virus scan is running. However, if the user is running a bol | Software design | Philips Medical Systems sent an "URGENT-MEDICAL DEVICE CORRECTION" letter dated Mar 21, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Additionally, a Field Service Engineer will contact each site to schedule a time to implement the Field Change Order. |  | 50 | 50 units | Nationwide Distribution-including the states of AR, AZ, DE, IN, KS, LA, MA, MN, MS, MT, NJ, NY, NY, OH, OR, PA, TN, TX, VT, WV, and WY. | Terminated | November 27, 2013 | 251 | 50 | Computer
Z-1570-2013 | 65135 | No | LKK |  |  | General Hospital | PMA | Medtronic SynchroMed II, Model 8637, (The SynchroMed II Pump is supplied in 20 ml or 40 ml reservoir size.) Sterilized using ethylene oxide. |  |  | June 03, 2013 | 2013 | Medtronic Neuromodulation | Medtronic is providing Healthcare Professionals with important safety information and patient management recommendations regarding the unintended delivery of drug during the priming bolus function for the SynchroMed implantable infusion pump. The unintended delivery of drug can contribute to patient overdose or underdose symptoms which may be clinically relevant. | Device Design | Medtronic sent an Urgent Medical Device Correction letter dated May 2013, to all affected customers. The letter described the background and nature of the issue and provided recommendations for patient management and monitoring after initiation of Intrathecal Therapy. The letter directed customers to continue to use the priming bolus precedure; to monitor all patients following start or restart of intrathecal therapy; consider priming the pump prior to implant; and educate caregivers and family members to recognize the signs and sypmtoms associated with intrathecal drug therapy complications. For questions contact your local Medtronic field representative or Medtronic Neuromodulation Technical Services at 1-800-707-0933. |  | 261109 | 261,109 total devices estimated implanted Worldwide | Worldwide Distribution - USA Nationwide | Terminated | May 14, 2015 | 710 | 261109 | Computer
Z-1586-2013 | 65411 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | DC power cable (p/n 3010556) supplied with the MEDRAD(R) Veris MR Vital Signs Monitor: |  |  | April 15, 2013 | 2013 | Medrad Inc | The firm is recalling these power cables due to a latent design reliability issue and the potential for shorting which can result in heating/melting of the cable jacket. | Device Design | Medrad Inc sent an Urgent Medical Device Component Recall letter dated April 15, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter states that replacement products will be available in June and shipped in the coming months along with additional instructions for returning the affected product. Customers were instructed to complete the enclosed recall response form and fax it back to Bayer at 412-406-0942. For questions contact customer support team at 1-877-229-3767 and select option 2. |  | 1066 | 1,066 | Worldwide Distribution - US Nationwide including the states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY and the countries of: AE, AU, AZ, BN, BR, CA, CH, CL, CO, DE, DK, DO, DZ, EG, ES, FR, GB, HK, HN, HR, HU, IN, IT, JO, JP, KR, LB, MU, MX, MY, NL, PH, PS, SA, SG, TH, TR, TW, UY, VE, and ZA. | Terminated | April 03, 2014 | 353 | 1066 | Computer
Z-1159-2013 | 64839 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | GemStar Pump - Battery |  |  | March 18, 2013 | 2013 | Hospira Inc. | Battery level lower then 2.4 volts results in corrupt history log and loss of data for preview therapy's parameters and causes an 11/004 error resulting in the pump not being able to be used. | Device Design | Hospira sent an Urgent Device Field Correction letter dated March 18, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to notify their consignees if they further distributed the product and confirm they have done so by return of the reply form to Stericycle. Lithium batteries that are older than three (3) years should be replaced. Customers should contact their Hospira Advanced Knowledge Center at 1-800-241-4002, option 4, to determine if their battery need to be replaced and if necessary, arrange for the return of their device to perform battery replacement. Customers with questions should call 1-877-907-7516. For questions regarding this recall call 224-212-4892. |  | 38801 | List Number 13000: 25,119 pumps; List Number 13100: 185 pumps; List Number 13150: 13,497 pumps | Worldwide Distribution - USA (nationwide) Australia, New Zealand, China, Hong Kong, Taiwan, Japan, Korea, Philippines, Thailand, Malaysia, Canada, Austria, Bahrain, Belgium, Croatia, Denmark, Egypt, Finland, France, Germany, Gibraltar, Greece, Hungary, Iceland, Ireland, Israel, Italy, Jordan, Kuwait, Lebanon, Libya, Luxembourg, Malta, Netherlands, Norway, Oman, Portugal, Qatar, Romania, Saudi Arabia, Slovenia, South Africa, Spain, Sweden, Switzerland, Turkey, UK, United Arab Emirates, Brazil, Chile and Colombia. | Terminated | March 07, 2017 | 1450 | 107955 | Computer
Z-1368-2013 | 64850 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | CobasÂż 8000 modular analyzer series software version 03-01, 02, 03 and 04, impacting the cobas e 602 module only. |  |  | April 04, 2013 | 2013 | Roche Diagnostics Operations, Inc. | When using the cobas e 602 module, Diluent Multiassay may be used for autodilution of certain assays instead of Diluent Universal. This may cause incorrect results under specific circumstances if all of the following conditions are met simultaneously: More than a total of 2000 reagent kits or diluents have been registered on the cobas e 602 modules within one system core, and two different diluent | Software design | Roche sent an Urgent Medical Device Correction (UMDC) letter dated April 4, 2013 via UPS to all affected customers. The letter idenified the affected product, problem and instruction to be followed. Customers were instructed to consult with the physician or pathologist at their facility to determine specific clinical implications for patients. A future software version will correct this issue. Customers were asked to complete and fax back the UMDC attached fax form to 1-88-628-0730. Questions should be addresssed to Roche Diagnostics Technical Support 24 hours a day at 1-800-428-2336. |  | 72 | 72 units | USA Nationwide Distribution | Terminated | February 24, 2015 | 691 | 72 | Computer
Z-1367-2013 | 64718 | No | IFB | Cytocentrifuge. | Pathology | Pathology | 510(K) Exempt | SLIDERPREP PLUS SLIDE PREPARATION CENTRIFUGE. |  |  | March 13, 2013 | 2013 | Separation Technology, Inc. | Separation Technology, Inc. recalled SlidePrep Plus Cytology Centrifuges manufactured prior to January 1, 2013 for having the potential to start up unexpectedly when encountering an electrostatic discharge. | Device Design | On March 13, 2013, Separation Technology, Inc. sent their customers letters instructing them to complete and return the attached Acknowledgement &amp; Receipt form via fax to (407) 788-3677 and requesting the return of the SlidePrep Plus Cytology Centrifuges to STI for correction. For any questions regarding this recall, call STI Customer Service at (800) 777-6668. |  | 102 | 102 | Worldwide Distribution-USA (nationwide) including the states of PA, NC, IL, KS, and FL, and the countries Canada and Turkey. | Terminated | December 24, 2013 | 286 | 102 | Computer
Z-1370-2013 | 64941 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker Orthopaedics Triathlon Femoral Distal Augment- Left; 1 single unit per package |  |  | March 25, 2013 | 2013 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics received a report that a label on the external packaging (box) of a Triathalon Femoral Distal Augment 10mm- Size 5 Left depicted a glyph of an acetabular shell rather than a femoral distal augment. In addition the label was missing a reference to the "10mm" augment thickness which per the label print is required to appear below the size reference. | Software design | The firm, Stryker Orthopaedics, sent an "URGENT MEDICAL DEVICE REMOVAL" letter dated March 25, 2013 and Product Accountability Forms via Fed Ex to its Consignees/ customers. The letter described the product, problem and actions to be taken. The customers were instructed to complete and return the attached Notification Acknowledgment Form within 5 days via fax to 855-251-3635; contact the hospitals in your territory that have the affected product to arrange return of product; and return affected product to the attention of Regulatory Compliance Stryker Orthopaedics, 325 Corporate Drive, Mahwah, NJ 07430, attach flourescent orange PRODUCT REMEDIATION sticker to your return indicating the PR #RA2013-011 and mark outer box with the words "Product Recall".If you have any questions, feel free to contact Manager, Divisional Regulatory Compliance at (201) 972-2100. |  | 11 | 11 units | Nationwide distribution: USA including states of: HI, IN, MS, NC, NJ, OK, TX, and WI. | Terminated | May 06, 2015 | 772 | 11 | Computer
Z-1375-2013 | 65031 | No | IQA | Powered environmental control system. | Physical Medicine | Physical Medicine | 510(K) Exempt | QuietCare. |  |  | April 04, 2013 | 2013 | Intel-GE Care Innovations LLC | Customers at 4 facilities reported receiving duplicate alerts using the QuietCare-Networked product. | Software design | Care Innovations contacted their customers by phone and e-mail regarding the potential problems and fixes on April 4, 17, and 23, 2013. The information provided stated the product, problem, and actions to be taken by the customers. Contact Technical Support at 855-885-2273 for questions regarding this issue. |  | 99 | 1722S; 99 Out of US | US and Great Britain | Terminated | June 04, 2013 | 61 | 145 | Computer
Z-1182-2013 | 64944 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Software correction applies to customers who use the RT Therapist RTT4.2.108 and RT Therapist Connect RTT 4.2.108 |  |  | April 04, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Software update to fix multiple safety related issues. | Software design | Siemens sent updated instructions of TH004/13/S on April 4, 2013. The SW Update will be delivered by the Siemens Customer Service Group.For questions regarding this recall call 925-293-5442. |  | 245 | 245 active devices | Worldwide Distribution - USA (nationwide) and internationally to Poland, India, Canada, Germany, United Kingdom, New Zealand, Belgium, Malaysia, Republic Korea, Ireland, Australia, Norway, Japan, Turkey, Croatia, Saudi Arabia, Colombia, P.R. China, Brazil, Thailand, Romania, Lebanon, Mexico, Italy, Sweden, Czech Republic, Russian Federation, and South Africa. | Terminated | August 12, 2013 | 130 | 245 | Computer
Z-1177-2013 | 64645 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Maquet Getinge Group battery modules for use with Maquet SERVO-i ventilator systems (Part #6487180). |  |  | March 07, 2013 | 2013 | Maquet Cardiovascular Us Sales, Llc | When using the SERVO-i on battery power, in a very small number of cases, some battery modules distributed after January 31, 2010 have a shorter battery run time than expected. | Device Design | The firm, Maquet Getinge Group" sent "Field Safety Notices"/"Field Safety Notice Confirmation Forms" dated March 7 &amp; 11 2013 via Federal Express and an updated Field Safety Notices dated April 15, 2013 to its customers. The Field Safety Notices identifies the problem, product, risk factors and actions to be taken by the customer. The customers were instructed to reconnect the ventilator to mains power immediately or add additional battery modules in case of 'low battery voltage' alarm; ensure that at least two fully charged batteries are installed at all times, and that sufficient resuscitation equipment is accessible to the ventilator; identify and immediately replace battery modules with date codes 1005-1243, that have generated either: a "low battery voltage" alarm or a "replace battery" message, prior to the expiry date; and complete and return the enclosed Field Safety Notice Confirmation Form by e-mailing a scanned copy to FieldActions@maquet.com or via fax at (973) 807-1832.If you have any questions or require replacement batteries, please contact your local MAQUET representative, or MAQUET Customer Service at 888-627-8383 (press option 2, followed by option 3) Monday through Friday, between the hours of 8:00 a.m. and 6:00 p.m. EST. |  | 90000 | 90,000 units (batteries) | Worldwide distribution: USA (nationwide) and countries of: ALGERIA, ANDORRA, ANTILLES (Netherlands), ARGENTINA, ARUBA, AUSTRALIA, AUSTRIA, AZERBADJAN, BANGLADESH, BELARUS, BELGIUM, BOLIVIA, BOTSWANA, BRAZIL, BRUNEI, BULGARIA, CANADA, CHILE, CHINA,COLOMBIA, CONGO, COSTA RICA, CROATIA, CUBA, CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, ECUADOR, EGYPT, EL SALVADOR, ESTONIA, YUGOSLAVIA (Former), FIJI, FINLAND, FRANCE, GEORGIA, GERMANY, GREECE, GUATEMALA, GUYANA, HONDURAS, HONG KONG,HUNGARY, ICELAND, INDIA, INDONESIA, IRAN, IRAQ, IRELAND, ISRAEL, ITALY, JAMAICA, JAPAN, KENYA, KUWAIT, KYRGYZSTAN, LATVIA, LIBYA, LITHUANIA, LUXEMBURG, MACAU, MACEDONIA, MALAYSIA, MALDIVES, MARTINIQUE (French), MEXICO, MOLDAVIA, MOROCCO, NEPAL, NETHERLANDS, NEW CALEDONIA (French), NEW ZEALAND, NICARAGUA, NIGERIA, NORWAY, OMAN, PAKISTAN, PALESTINE, PANAMA, PAPUA NEW GUINEA, PARAGUAY, PERU, PHILIPPINES, POLAND, POLYNESIA (French), PORTUGAL, QATAR, REUNION (French), ROMANIA, RUSSIA, SAUDI ARABIA, SERBIA, SEYCHELLES, SINGAPORE, SLOVAK REPUBLIC, SLOVENIA, SOUTH AFRICA, SOUTH KOREA, SPAIN, SRI LANKA, SUDAN, SURINAM, SWEDEN, SWITZERLAND, SYRIA, TADJIKISTAN, TAIWAN, TANZANIA, THAILAND, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, UKRAINE, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY, VATICAN city state, VENEZUELA, VIETNAM and YEMEN. | Terminated | August 11, 2016 | 1253 | 90000 | Computer
Z-1404-2013 | 65082 | No | HKO | Ophthalmic refractometer. | Ophthalmic | Ophthalmic | 510(K) Exempt | Canon Full Auto Ref-Keratometer RK-F2, Catalog Number 6937B002AA. |  |  | December 05, 2012 | 2012 | Canon Inc. | Canon Inc. had received 17 claims from our foreignsales companies between October to November2012 that measurement value of RK-F2s refractivepower (SPH value) sometimes indicated abnormalvalue which occurs approximately 5% in probability. | Software design | Canon initiated recall by sending letters on November 22, 2012 to all affected customers. The letter identified the affected product, problem, actions to be taken and contact information. |  | 384 | 384 units | Worldwide Distribution - USA Nationwide and countries of: Australia, Japan, Korea, Singapore and United Kingdom. | Terminated | September 20, 2013 | 289 | 384 | Computer
Z-1585-2013 | 64789 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | The Centricity Laboratory System is intended to be an information system designed to support the clinical and administrative activities associated with the provision and utilization of clinical laboratory services and facilities, e.g., the storing and delivering of analytical results. It is a specially designed data program application (software), which is supplied for installation in existing mainframe or decentralized computers or a computer network. This software system will not be the software directly used to run, steer, or control any specific laboratory analyzer or equipment. It is intended for use in laboratory facilities, including a central laboratory or in a multiple laboratory environment servicing satellite laboratories, reference laboratories, clinics, etc. Centricity Laboratory System is not intended for direct patient contact. |  |  | November 20, 2012 | 2012 | GE Healthcare It | If environment variables MICRO_CORRECT_MSG_TYPE and CORRECT_MSG_TYPE are set to different values, in Microbiology result entry when released results (Final Report released) are changed and the Corrected Result messages are edited on the free-text form, some of the originally entered panel free-text is lost. For this issue to occur the original result must be amended and the environmental variabl | Software design | GE Healthcare sent an Urgent Medical Device Correction letter dated November 20, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Software correction is available. Please contact your GE service representative to request or install this update. Please contact your GE Service Representative with any question regarding this issue. For other questions please contact GE Helpdesk.Help Desk: North America: 1-888-778-3375 Help Desk: Europe: + 44 1603 877595 Help Desk: Asia Pacific + 61 3 5327 4000 |  | 9 | 9 Units | Worldwide Distribution - US Distribution including the states of OH, NY, GA, TX, TN and MN and the countries of Australia, Qatar and England | Open, Classified |  |  | 9 | Computer
Z-1407-2013 | 65171 | No | GWF | Evoked response electrical stimulator. | Neurology | Neurology | 510(k) | Natus neurology Dantec KEYPOINT.NET 2.20. |  |  | April 09, 2013 | 2013 | Alpine Biomed Aps | Alpine Biomed ApS is recalling Natus neurology Dantec KEYPOINT.NET ver. 2.20 Software due to the presence of a software bug. | Software design | Alpine Biomed ApS sent a Recall Notification/ Field Safety Notice dated April 8, 2013 to all affected customers via e-mail and courier. Alpine Biomed ApS advised their customers to upgrade all installations of Keypoint.NET v2.20 to Keypoint.NET v2.21. Alpine BioMed ApS has recommended that their customers complete and return the verification form to QA/RA Engineer Michael Jacobsen at e-mail michael.jakobsen@natus.com. These actions and the return of the verification form to Alpine Biomed ApS should be completed within 3 months of receipt of the recall notifications. If any assistance is required to complete these actions, please contact Technical Support Engineer Michael Jensen via e-mail michael.jensen@natus.com, alternatively QA/RA Engineer Michael Jacobsen by e-mail michael.jakobsen@natus.com. |  | 871 | 871 licenses installed | Worldwide Distribution - US Nationwide and the countries of: Germany, France, Bulgaria, India, Colombia, Tunisia, Saudi Arabia, Finland, Denmark, Brasil, Italy, Estonia, South Korea, Russia, Greece, Thailand, Israel, Poland, Turkey, Slovakia, Portugal, Switzerland, Hungary, Great Britain, Netherlands, China, Indonesia, Australia, Sweden, Spain, Iran, Singapore, Norway, Austria and Czech Republic. | Terminated | January 06, 2014 | 272 | 871 | Computer
Z-1419-2013 | 65136 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | RayStation Therapy Treatment Planning System, Model Numbers 2.0, 2.5, and 3.0; CLASSIFICATION NAME: System, Planning, Radiation Therapy Treatment. |  |  | February 15, 2013 | 2013 | RAYSEARCH LABORATORIES AB | Under some circumstances, imported CT, MR and PET images are offset from their true positions by one pixel in one or two directions. Users must be aware of this information to avoid incorrect dose calculations during treatment planning. | Software change control | RaySearch Laboratories sent an Urgent Field Safety Notice, Medical Device Correction letter dated January 28, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to educate their planning staff and all users about the recommended workarounds contained in the letter. Customers were also asked to inspect their product and identify all installed units that were infected and then complete the attached Reply Form and e-mail or scan a copy of the completed form to the firm. Customers were instructed to pass along this information to all those who need to be aware of this Medical Device Correction Customers with questions were instructed to contact the RaySearch support email at support@raysearchlabs.For questions regarding this recall call +46 8 545 06130. |  | 97 | Domestic: 70 units; Foreign: 27 units | Worldwide Distribution - USA including Connecticut, Florida, Massachusetts, Missouri, New York, Ohio, and foreign distribution to(Canada, Italy, Australia, Germany, Netherlands, and Switzerland. | Open, Classified |  |  | 97 | Computer
Z-1421-2013 | 65165 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Leucadia Autolok" Polyaxial Pedicle Screw System. |  |  | May 06, 2013 | 2013 | Alphatec Spine, Inc. | Alphatec Spine initiated this medical device recall for the following reason: due to tolerancing, the flange of the Phygen Leucadia" Autolok set screw may interfere with the top of the screw head prior to fully locking the rod. If this condition occurs, the polyaxialscrew head and rod may fail to provide a rigid construct under loading. To date there have beenno reports of patient injury, adve | Device Design | Alphatec Spine initiated recall by contacting customers via telephone and/or email on 05/06/2013. An Urgent Medical Device Recall Notification letter along with a response form were sent via FedEx to all affected accounts. The letter identified the product, problem, and actions to be taken by the customers. |  | 1814 | 1,814 | Nationwide Distribution. | Terminated | September 25, 2014 | 507 | 1814 | Computer
Z-1425-2013 | 65073 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Tip Cover Accessory, 8 MM, MCS, box of 10 |  |  | September 23, 2011 | 2011 | Intuitive Surgical, Inc. | Clarification of existing labeling and related cautions for the tip cover accessory of the da Vinci Surgical System. | Device Design | Intuitive Surgical sent an Important Product Notification letters dated October 10, 2011 to all affected customers. A second letter was sent to those that had not responded on December 4, 2011. The letters identified the affected product, problem, and actions to be taken, The letter included recommendations and suggestions for proper use and prevention of damage to the Tip cover during surgery. Also attachments include recommendations and further information for the following areas: avoid instrument collisions; carefully install tip cover; straighten wrist prior to removal; inspect cannula prior to use and do not exceed maximum monopolar cautery settings.Attached is a list of validated ESU settings and suggestions and instruction on being aware of critical anatomy in contact with the instrument and to change the tip cover in cases with extended cautery use.Questions may be referred to Customer service at 1-800-876-1310 in the US and at + 41 21 821 2020 or +800-0821-2010. |  | 500000 | 50,000 boxes of 10 tip covers | Worldwide distribution - USA Nationwide and the countries of Australia, Austria, Belgium, China, Czech Republic, Denmark, Finland, France, Germany, Greece, India, Ireland, Israel, Italy, Japan, Malaysia, Mexico, Netherlands, New Zealand, Norway, Pakistan, Philippines, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Slovenia, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom. | Terminated | May 30, 2013 | 615 | 500000 | Computer
Z-1431-2013 | 65194 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | 8mm Precise Bipolar Forceps used in conjunction with the da Vinci Surgical System, Model IS1200. |  |  | May 08, 2013 | 2013 | Intuitive Surgical, Inc. | Reports of dislodging of the Cautery Plug Insert from the Cautery Plug Riser, preventing a bipolar cautery cable from being properly connected which results in loss of cautery function of Bipolar Instruments. | Device Design | Intuitive Surgical sent an "Urgent Medical Device Correction" letter dated May 8, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact Customer Service at 800-876-1310 for questions regarding this notice. |  | 697 | 697 units - all products | Worldwide Distribution-USA (nationwide) including the states of AL, CA, CO, FL, IL, KY, MI, MS, NJ, NY, PA, TX, UT, VA, WA, and WI, and the countries of Austria, Australia, Belgium, Canada, Czech Republic, Denmark, France, Germany, Italy, Netherlands, Norway, Qatar, Romania, Spain, Sweden, Switzerland, Taiwan, United Kingdom, Saudi Arabia, Singapore, and Venezuela. | Terminated | February 18, 2014 | 286 | 697 | Computer
Z-1167-2013 | 64731 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AU5800 Clinical Chemistry Analyzer |  |  | March 20, 2013 | 2013 | Beckman Coulter Inc. | Beckman Coulter, Inc. initiated this recall because there is a software error that has the potential to cause the reporting of erroneous patient results because of misidentification by the AU5800 system. | Software design | Beckman Coulter sent an Urgent Product Correction letter during the week of March 18, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were informed of the software limitation and provides instructions to be taken (temporary workaround) until the software is upgraded and released. Customer contact information provided: "Call Center Hotline at 800 854-3633. New customers will receive a copy of the customer notification at the time of installation by the Field Service Personnel until revised software that corrects this issue is released. |  | 490 | 490 | Worldwide Distribution - USA, including Canada and Puerto Rico and the states of NJ, TX, CA, UT, KS, MA, NY, MI, GA, VA, IN, WI, IL, OH, NV, NC and PA. and the countries of Australia, Belgium, Czech Republic, Japan, Germany, Brazil, China, Italy, Ireland, Israel, India, Hong Kong, Malaysia, Netherlands, Spain, Switzerland, France, Taiwan, Turkey, United Kingdom and Vietnam. | Terminated | September 16, 2014 | 545 | 490 | Computer
Z-1365-2013 | 65158 | No | OUG | Medical device data system. | General Hospital | General Hospital | 510(K) Exempt | MEDITECH Bi-Directional interactive Communication Analyzer Interface; |  |  | May 07, 2013 | 2013 | Medical Information Technology, Inc. | Incorrect transmission of laboratory results. | Device Design | Beginning on May 7, 2013, MEDITECH distributed a notification to MAGIC and Client Server customers who have the MEDITECH Bi-Directional Interactive Communication Analyzer Interface in use with Vitros 3600, 4600, or 5600 Analyzer in TEST and/or LIVE. This notification was made via remailed task updates that can be printed by the customer. Task updates sent via e-mail are immediately transmitted to the customer.This notification informs customers of the potential patient safety issue that could result in reporting inaccurate qualitative results. Customers will receive a corrective change, in which the previous test's results will no longer be associated with the test not performed. Until customers receive the corrective change, they are instructed to monitor all No Result values from their analyzer and confirm that the result in MEDITECH matches. For further questions please call 781-821-3000, ext. 5742. |  | 56 | 56 devices | Worldwide Distribution - USA, including Puerto Rico and the country of Canada. | Terminated | October 01, 2014 | 512 | 56 | Computer
Z-1187-2013 | 64812 | No | DQA | Oximeter. | Cardiovascular | Anesthesiology | 510(k) | Rad-8 Pulse Oximeters |  |  | March 13, 2013 | 2013 | Masimo Corporation | Masimo has identified a very small number (0.02%) of Rad-8 devices with 20-pin patient cable connector that, when moved, can power off without the operator pressing the power button. | Device Design | Masimo Corporation sent an Urgent Recall Notification letter dated March 8, 2013, to all affected customers The letter identified the product the problem and the action needed to be taken by the customer.1.Customers were instructed to please promptly remove all Rad-8 device(s) from use and your inventory and segregate these devices.2. Please call 1-800-326-4890 and select option 2 for Technical Services. Determine with Technical Services whether you will return the affected device(s) to Masimo for repair or if eligible, receive a repair kit. If you will return the device(s) for repair, obtain a Return Material Authorization Number (RMA number).3. Complete the Tracking/Verification Form in Attachment 2 and fax it to Masimo at 1-949-297-7700, even if you do not have or intend to use the affected device(s).4. If you are eligible and intend to repair units onsite fax the completed original form in Attachment 2 to Masimo at 1-949-297-7700.5. If you have elected to return the affected device(s), and have an RMA number include the completed original Form in Attachment 2 with the returned device(s) and indicate the RMA number on the outside of the return-shipping box.We apologize for the inconvenience. Please be assured that Masimo is committed to consistently providing high quality products and services to you, our customers. We thank you for your patience and cooperation while we actively work to resolve this issue... |  | 42908 | 42,908 units | Worldwide Distribution - USA (nationwide) including the countries of Africa, Asia, Australia/New Zealand, Europe, Canada, Latin America, and Middle East / SW Asia. | Terminated | June 10, 2015 | 819 | 42908 | Computer
Z-1189-2013 | 64860 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | HeartStart MRx Monitor/Defribillator; Product Codes: M3536A, M3536J, M3536M, M3536MC, M3536M2, M3536M4, M3536M5, M3536M6. |  |  | April 08, 2013 | 2013 | Philips Healthcare Inc. | Device may fail to deliver defibrillation therapy in either Manual Defib or AED mode. | Software design | The firm, Philips Healthcare, sent an "Field Safety Notice-URGENT - MEDICAL DEVICE RECALL" letter dated April 2013 to its customers. The letter describes the product, problem and actions to be taken. In addition, letter provides instructions for using the MRx prior to receiving the software upgrade and informs the customer that Philips is initiating a correction consisting of a software upgrade for all units affected. The customers were instructed to follow the "ACTION TO BE TAKEN BY CUSTOMER/USER" section of the Field Safety Notice; send a copy of the attached package including customer letter and Field Safety Notice to any customer to whom you have distributed one of the affected device; and provide your local Philips organization with the names and addresses of customers to whom you have sold affected devices, so that arrangements can be made to perform the software updates.Should you have any questions or concerns about this device correction, please contact your local Philips representative at 1-800-722-9377.Philips updated the Field Safety Notice dated May 2013 to users who have not had the software upgrade performed. It provided additional information in the "Problem Description" and "Action to Be Taken By Customer/User". |  | 64972 | 64972 | Worldwide distribution: USA (nationwide) and countries of: Afghanistan, Albania, Argentina, Australia, Austria, Azerbaijan, Bahrain, Bangladesh, Belgium, Bolivia, Bosnia and Herzegowina, Brazil, Brunei Darussalam, Bulgaria, Canada, Cambodia, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, El Salvador, Estonia, Ethiopia, Finland, France, Germany, Greece, Guadeloupe, Hong Kong, Iceland, India, Indonesia, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Korea, Republic Of, Kuwait, Laos, Latvia, Lebanon, Libyan Arab Jamahiriya, Lithuania, Luxembourg, Macau, Macedonia, Malaysia, Malta, Mexico, Mongolia, Morocco, Namibia, Nepal, Netherlands (Antilles), New Caledonia, New Zealand, Nicaragua, Nigera, Norway, Oman, Pakistan, Palestine, Panama, Papua New Guinea, Paraguay, Peru, Philippines, Poland, Portugal, Puerto, Rico, Qatar, Reunion, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovenia, South Africa, Spain, SRI Lanka, Sweden, Switzerland, Taiwan (Province of China), Tanzania, Thailand, Trinidad and Tobago, Tunisia, Turkey, Turkmenistan, Uganda, Ukraine, United Arab Emirates, United Kingdom, Uzbekistan, Venezuela, Veit Nam, and Yemen. | Terminated | January 10, 2017 | 1373 | 64972 | Computer
Z-1346-2013 | 64930 | No | OOI | Instrumentation for clinical multiplex test systems. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | cobasÂż 4800 system z480 instrument. |  |  | April 01, 2013 | 2013 | Roche Molecular Systems, Inc. | In rare instances, channel shifted results have been generated with cobasÂż 4800 assays run on v1.1.1 of the cobasÂż 4800 system. Raw data review indicates that, in these cases, all signals are shifted by one channel:channel 1 results report as channel 2,channel 2 results report as channel 3,channel 3 results report as channel 4, andchannel 4 results report as channel 1.In the reported | Software design | Roche sent an "Urgent Medical Device Correction" (UMDC) notification with response form dated April 1, 2013, to all affected customers via UPS Ground on April 1, 2013. The letter identified the product, the problem, and the action to be taken by the customer. A contact number (Roche Diagnostic Technical Support) is provided for technical support 24 hours a day, seven days a week at phone number: 1-888-912-7090. |  | 425 | 425 instruments | Nationwide Distribution. | Terminated | August 07, 2014 | 493 | 425 | Computer
Z-1345-2013 | 56447 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | TD-Synergy. |  |  | July 29, 2010 | 2010 | Technidata S.A. | Technidata has discovered that using the F8 key in the ERM session to add a combined test can result in the modification of the value of a result that has already been entered for a test of the request. | Software design | The firm, Technidata, sent an "P1 DEFECT INFORMATION" notice dated July 2, 2010 to their consignee/customers. The notice describes the product, problem and actions to be taken. The notice directs them to find the correction to the software issue on the Technidata website where they are able to install the correction kit depending on the version and operating system used. Note "You must back up the database and the installation directory before installing this correction".Should you need any further information, do not hesitate to contact the Support team at +33 (0) 4 76 04 13 00. |  | 10 | 10 | Worldwide Distribution-USA (nationwide) including Washington DC and the states of California, Nevada, Kentucky, Alabama, Rhode Island, Pennsylvania, and Louisiana and the country of Italy. | Terminated | December 05, 2013 | 1225 | 10 | Computer
Z-1339-2013 | 65096 | Yes | LWS |  |  | Cardiovascular | PMA | LATITUDEÂż Patient Management System, Model 6488, Version 7.3. |  |  | April 10, 2013 | 2013 | Boston Scientific CRM Corp | Boston Scientific CRM is conducting a recall on the LATITUDE Patient Management System Model 6488 Version 7.3 because between March 4 and 27, 2013, a Siebel release caused ~5000 distributed PMR clinic faxes to display with illegible characters and not the required, Patient non-compliance information that was intended. | Software change control | Boston Scientific CRM Corp began re-faxing the 4,912 previously illegible PMR notifications on April 10, 2013 and completed on April 15, 2013, successfully fulfilling labeling expectations that these notifications be communicated to the clinic.For questions regarding this recall call 651-582-4000. |  | 1 | 1 | Nationwide Distribution including DC and PR | Terminated | November 26, 2013 | 230 | 1 | Computer
Z-1186-2013 | 64621 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | SL3 SOFT-TISSUE Laser (Model LR2002); Manuf: Discus Dental, LLC, 8550 Higuera Street, Culver City CA 90232. |  |  | July 21, 2011 | 2011 | Discus Dental LLC | The firm is initiating this recall of the SL3 Soft Tissue Laser equipped with a fiber optic Unifiber system which contains a fragile glass filament because the fiber broke near the hand piece. | Device Design | DISCUS Dental, sent an "URGENT MEDICAL DEVICE NOTIFICATION" letter dated July 21, 2011. The letter described the product, problem and action to be taken. Customers were instructed tocontact Discus Dental at 1-888-427-9279 for a no-cost replacement.For questions regarding this recall call 310-845-8200. |  | 150 | 150 units | Nationwide Distribution including AL, AZ, CA, CO, FL, GA, IA, ID, IL, IN, KS, MA, MD, MI, MO, MT NH, NJ, NM, NV, NY, OH, PA, SD, TX, UT, VA, and WA. | Terminated | April 29, 2013 | 648 | 150 | Computer
Z-1340-2013 | 64647 | No | KQM | Surgical camera and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Ref: VP1600 HD Still Capture System ConMed Linvatec . |  |  | March 18, 2013 | 2013 | Linvatec Corp. dba ConMed Linvatec | The VP1600 HD Still Capture Image System is recalled due to an anomaly between the software code in the VP1600 console and the communications interface for various keyboard key strokes (Caps lock and Num Lock) which may lock up and requires the user to reboot to continue. | Software design (manufacturing process) | The firm, ConMed, sent an "URGENT: MEDICAL DEVICE RECALL NOTIFICATION" letter dated 3/8/2013 to its customers after initial phone notifications. The letter identified the product, problem, and actions to be taken by the customers. |  | 267 | 267 units | Worldwide Distribution-USA (nationwide) including the Virgin Islands and the states of CA, CT, FL, GA, IN, KY, KS, LA, MD, MI, MN, MO, MS, MT, NE, NC, NJ, NY, OH, OR, PA, SC, TN, TX, UT, VA, and WI, and the countries of Austria, Belgium, Canada, Denmark, Hong Kong and Zambia. | Terminated | November 06, 2013 | 233 | 267 | Computer
Z-1326-2013 | 65028 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Fluoroscan Mini C-arm InSight, software versions 1.x, 2.x, 3.x |  |  | April 25, 2013 | 2013 | Hologic, Inc. | The audible alarm to alert the operator when cumulative irradiation time reaches five minutes is not accurate and the total cumulative irradiation time is not accurate. | Software design | The firm, Hologic, Inc. sent an "Urgent: Medical Device Recall" letter dated April 25, 2013, via FedEx to consignees/customers notifying them of the software defect and providing them with corrected user installable software version. The customers were instructed to complete and return the Enclosure to Customer Notification, Confirmation of Notification form via fax to Hologic at 866-652-8674 or mail the document to Hologic using the enclosed pre-addressed, pre-paid return envelope.Should you have any questions regarding this recall, require further information or assistance with the software installation call Hologic's Help Desk at 1-800-321-4659. |  | 287 | 287 | Worldwide distribution: USA (nationwide) and countries of: DE, RU, AU, IT, GB, ES, FR, KR, ZA, CA, AR, LT, TW, CN, BE, AT, JP, QA, SA, SG, HK, SI, NL, and SE. | Terminated | August 21, 2014 | 483 | 3067 | Computer
Z-1170-2013 | 64837 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | The GemStar Pump is a small and lightweight, single channel infusion device designed for use in the home, in the hospital, or anywhere electronic infusion is required. The GemStar Pump can be powered by AC mains adaptor, rechargeable battery pack, docking station, or two disposable AA alkaline batteries. When powered by batteries, The GemStar Pump is ideal for ambulatory patients. |  |  | March 18, 2013 | 2013 | Hospira Inc. | Pump shutting off during use without warning. | Device Design | Hospira sent an Important Device Information letter dated March 18, 2013, to all affected customer. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to inspect the affected product for signs of leakage, corrosion or other damage prior to each use. In addition, each time the batteries are replaced the battery compartment should be inspected for damage. If a device exhibits damage caused by leakage, it should immediately be removed from clinical service and Hospira should be contacted at 1-800-441-4100 to arrange for return of the device for repair. If the product has been further distributed they should notify their consignees and confirm they have done so by returning the reply form to Stericycle.For questions regarding this recall call 224-212-4892. |  | 69604 | List Number 13086: 286 pumps; List Number 13087: 45,376 pumps; List Number 13088: 23,942 pumps | Worldwide Distribution - USA (nationwide) and internationally to Australia, New Zealand, China, Hong Kong, Taiwan, Japan, Korea, Philippines, Thailand, Malaysia, Canada, Austria, Bahrain, Belgium, Croatia, Denmark, Egypt, Finland, France, Germany, Gibraltar, Greece, Hungary, Iceland, Ireland, Israel, Italy, Jordan, Kuwait, Lebanon, Libya, Luxembourg, Malta, Netherlands, Norway, Oman, Portugal, Qatar, Romania, Saudi Arabia, Slovenia, South Africa, Spain, Sweden, Switzerland, Turkey, UK, United Arab Emirates, Brazil, Chile and Colombia. | Terminated | March 10, 2017 | 1453 | 108405 | Computer
Z-1208-2013 | 64918 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE Healthcare, GE 1.5T SIGNA HDe MR SYSTEM. |  |  | June 08, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has recently become aware of a Reformat Image Flip Issue associated with GE MR Products.The axial and coronal reformatted images from a sagittal 3D series (such as with applications of FSPGR-3D T1, BRAVO T1, or CUBE T1) of over 126 slices may display a left-to-right image flip .In addition, the acquired sagittal images may be annotated with left-right location annotation revers | Software design | GE Healthcare sent an "Important Product Information" letter dated June 8, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers should contact their local service representative for questions regarding this notice. |  | 822 | 822 (154 US; 668 OUS) | Worldwide Distribution-USA (nationwide) including DC and Puerto Rico, excluding the state of RI and the countries of Yugoslavia YEMEN Viet Nam Venezuela Uzbekistan Uruguay United Arab Emirates Tunisia Thailand Tanzania Taiwan Turkey Syria Switzerland Sweden Spain South Korea South Africa Slovenia Slovakia Singapore Saudi Arabia Russia Romania Reunion Qatar Portugal Poland Philippines Peru Pakistan Paraguay Philippines Oman Norway Nigeria New Zealand Netherlands Morocco Mexico Martinique Malaysia Macedonia Mauritania Luxembourg Lithuania Libya Lebanon Latvia Kuwait Kenya Korea Korea (Republic Of [South] Korea) KUWAIT Kazakhstan Jordan Japan Jamaica Italy Israel Ireland Iran Iraq Indonesia India Hungary Hong Kong Honduras Guadeloupe Greece Germany Georgia, Republic of French Polynesia France Finland Estonia Ecuador Egypt Denmark Dominican Republic Czech Republic Croatia Colombia China Chile Canada Bulgaria Brazil Bosnia And Herzegovina Belgium Belarus Barbados Bahrain Austria Australia Argentina Arab Emirates Algeria and United Kingdom. | Terminated | November 03, 2014 | 878 | 12048 | Computer
Z-1210-2013 | 64942 | No | HTW | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker Howmedica Osteonics Corp. AxSOS System Instrument. |  |  | March 28, 2013 | 2013 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics has received a report indicating that the scale of the AxSOS Calibrated Drill Bit 4.3,,X263mm, flat , manufactured by Stryker Osteosynthesis, is incorrect. The incorrect scale could result in the surgeon selecting a screw that is 10mm longer than intended. | Device Design | Stryker Orthopaedics sent a Urgent Medical Device Removal letter dated March 28, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.1. Please inform users of this Medical Device Removal and pass thisnotice to all those individuals who need to be aware within yourorganization.2. Return all affected products available at your location toStryker Osteosynthesis or Contact Stryker Customer Service.3. Complete and sign the enclosed Business Reply Form and fax a copyto: 1-865-252-3635 .4. Keep a copy of the completed and executed Business Reply Form foryour records.Report any adverse events or product quality problems to StrykerOrthopaedics: 1-866-0R-ASSIST. (1-866-672-7747).As we strive for products that meet your expectations for quality and reliability, please do not hesitate to contact us, in case you have any further questions at (201) 831-5970.. |  | 107 | 107 units | Nationwide Distribution including the states of AL, AZ, IL, KY, LA, NJ, NC, FL, CA, NM, MO, TN, VA, CO, MA, TX, IN and MT. | Terminated | November 17, 2015 | 964 | 107 | Computer
Z-1324-2013 | 64986 | No | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | Segmented Cylinder Applicator Set, Model # GM11004150, when used for Pulsed Dose Rate (PDR) |  |  | April 11, 0013 | 13 | Varian Medical Systems, Inc. | Segmented cylinder applicator set may slip during treatment, causing the delivery of radiation to areas outside the target. | Device Design | Varian Medical Systems, Inc. notified customers of the recall with 'Urgent Medical Device Correction Urgent Field Safety Notice," dated 4/11/13, and delivered by mail to customers. The notice identified the product, problem, and instructions for the customers. Varian Medical Systems, Inc. informed customers of the possible malfunction of the GM11004150 Segmented Cylinder Applicator Set for pulsed dose rate brachytherapy treatment, with codes H640383, H640277, H640384, H640788, H640D126, H640385, H640372, H640303, H640387. The notice instructed customers to cease use of the Segmented Cylinder Applicator Set and return the devices along with a completed proof of notification form (included with the recall notice) to the recalling firm. |  | 21 | 21 devices | Worldwide distribution: USA (nationwide) and in the countries of: Belgium, France, Martinique, Slovenia, Sweden, and the United Kingdom. | Terminated | January 13, 2014 | 730762 | 21 | Computer
Z-1323-2013 | 64390 | No | DSQ |  |  | Cardiovascular | PMA | Heartware HVAD Pump Implant Kit (Heartware Ventricular Assist Device) |  |  | December 27, 2012 | 2012 | Heartware Inc | HeartWare, Inc. of Miami Lakes, FL is recalling their Ventricular Assist System due to the loosening of the HVAD Pump driveline connector. | Device Design | HeartWare, Inc. sent an Urgent Medical Device Correction letter to all their customers on February 7, 2013 instructing that at each routine clinic visit, they should inspect the patient's driveline connector as described in the attached Technical Bulletin. They were asked to sign and return the acknowledgement form to HeartWare. For questions contact your HeartWare representative or Customer Service at 1-877-367-4823. |  | 3544 | 3,544 | Worldwide Distribution-USA (nationwide) including the states of MN, IL, MI, FL, WI, TX, MO, CA, MA, TN, OH, NC, GA, VA, UT, KY, MD, IN, PA, NY, NJ, LA, OR, WA, AL and CO, and the countries of Poland, Australia, Austria, Italy, Germany, New Zealand, South Africa, Czech Republic, France, Brazil, Switzerland, Singapore, United Kingdom, Saudi Arabia, Netherlands, China, Chile, Belgium, Canada, Norway, Lebanon, Israel, Malaysia, Turkey, India, Argentina, Greece, Japan and Lithuania. | Terminated | July 09, 2013 | 194 | 3544 | Computer
Z-1322-2013 | 64221 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Proplege Coronary Sinus Catheter model PR9, 11 French Introducer, kitted with PR9. |  |  | January 25, 2013 | 2013 | Edwards Lifesciences, LLC | Edwards Lifesciences is recalling certain Proplege Coronary Sinus Catheter, model PR9 due to confirmed reports of blood leakage through the hemostasis valve in the introducer that is kitted with the catheter. | Device Design | Edwards Lifesciences sent an Urgent Field Safety Notice letter dated February 1, 2013. The letter identified the affected product, problem and actions to be taken. Customers were instructed to review entire inventory for the affected product, quarantine and return to Edwards. Contact Customer Services at 800-424-3278 for any question, to obtain a Returned Goods Authorization number and replacement product. |  | 807 | 807 | Worldwide Distribution - USA Nationwide including the states of: MD, NY, PA, CA, FL, AZ, AL, TX, VA, OH, MO, NC, GA, TN, CO, OR, SC, ME, IL, HI, NJ, WA, WI, CT, KY, LA and Canada. | Open, Classified |  |  | 807 | Computer
Z-1238-2013 | 64771 | No | KWI | Elbow joint radial (hemi-elbow) polymer prosthesis. | Orthopedic | Orthopedic | 510(k) | Synthes Radial Head Prosthesis: Trial Radial Head |  |  | February 22, 2013 | 2013 | Synthes USA HQ, Inc. | The Trial Head may come loose from the implant stem during manipulation of the arm during surgery. To date, no adverse events have been reported related to this issue and this recall is not being initiated as a result of adverse events. | Device Design | Synthes sent an Urgent Notice: Medical Device Recall letter dated February 22, 2013, to all affected consignees. The letter requested consignees examine their inventory, remove them from use and return them to Synthes. Customers were asked to complete the attached Verification Section at the end of the letter indicating whether or not they have any of the affected product. For questions customers were instructed to call 610-719-5450 or email FieldAction@synthes.com.For questions regarding this recall call 610-719-5000. |  | 417 | 417 | Nationwide Distribution: USA including states of: IL, IN, MA, MI, MO, NJ, NY, OH, OR, PA, SC, TX, VA, WA and WI. | Terminated | August 31, 2015 | 920 | 417 | Computer
Z-1245-2013 | 64876 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | FreeStyle InsuLinx Blood Glucose Monitoring System: |  |  | April 15, 2013 | 2013 | Abbott Diabetes Care, Inc. | AT rare, extremely high glucose levels (1024 mg/dL and above), the FreeStyle InsuLinx Blood Glucose Meter may provide an inaccurate reading. | Software design | Abbott Diabetes Care sent an Urgent Product Recall letter dated April 15, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to take one of the following options immediately to address the issue. 1) Customers can access a software update to resolve the issue at www.freestyleinsulinx.com/swupdate.2) Customers can contact Abbott Diabetes Care Customers Service at 1-866-723-2697 to expedite return and replacement of the affected product at no charge. For questions customers were instructed to call Abbott Diabetes Care Customers Service at 1-866-723-2697.For questions regarding this recall call 510-239-2775. |  | 544609 | 335,273 meters; 209,336 remain at firm. | Nationwide Distribution | Terminated | August 03, 2015 | 840 | 544609 | Computer
Z-1168-2013 | 64694 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Ventri Discovery NM 530c, Ventri with Premium Table Discovery NM 530c, models H3000ZW &amp; H3000SA. The system is used to perform nuclear imaging procedures. |  |  | January 09, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of an incident which caused a patient injury during an unload following an exam on a Ventri scanner. During a patient unload following a scan, the patients fingers may be pinched between the moving cradle edge and the stationary base of the table while the table is moving outward automatically. | Device Design | Consignees were sent on 1/9/2013 a GE Healthcare "Urgent Medical Device Correction" letter dated January 9, 2013. The letter was addressed to Managers of Nuclear Medicine, Hospital Administrators / Risk Managers andManagers of Radiology / Head of Biomedical Engineering. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. |  | 556 | 556 | Worldwide Distribution - US, including the states of AL, AZ, AR, CA, CO, CT, DE, DC, FL, IL, IN, IA, KY, LA, ME, MD, MA, MI, ,N,MS, MO, NE, NV, NH, NJ, NM, NY,NC, OH, OK,OR, PA, RI, SC, TN, TX, UT, VA, WA, WV, WI and the countries ofTAIWAN, TURKEY, SINGAPORE, SWD EN, PORTUGAL, POLAND, NEW ZEALAND, NORWAY, NETHERLANDS, MALAYSIA, MEXICO, SOUTH KOREA, JAPAN, ITALY, INDIA, ISRAEL GREAT BRITAIN, FRANCE, DENMARK, GERMANY, SWITZERLAND, CANADA, BRAZIL BELGIUM, AUSTRALIA, and ARGENTINA. | Terminated | August 06, 2013 | 209 | 556 | Computer
Z-1441-2013 | 65191 | No | HLI | Ophthalmoscope. | Ophthalmic | Ophthalmic | 510(K) Exempt | Optovue iVue with Normative Database (NDB) with Software Version 3.0 and 3.1. |  |  | May 14, 2013 | 2013 | Optovue, Inc. | Colors from the NDB comparison for Ganglion Cell Complex thickness in the singular OU report in iVue version 3.0 (US) and 3.1(international) are different from those in the individual iWellness ID and OS reports. | Software design | Optovue sent a Field Correction Notice letter dated May 16, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. |  | 329 | ~329 users | Distributed Nationwide and in Canada. | Terminated | October 24, 2013 | 163 | 329 | Computer
Z-1163-2013 | 64828 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand RPM Respiratory Gating System, RPM Gating v1.7.5 with 3D Option, Model Number: HS2 |  |  | March 26, 2013 | 2013 | Varian Medical Systems, Inc. Oncology Systems | Varian has identified an anomaly with the RPM Respiratory Gating System version I.7.5 with 3D Option. While operating with the 6-dot marker block, this version of RPM Gating may produce an inverted respiratory waveform in amplitude and phase-gating sessions. | Device Design | Varian sent an Urgent Medical Device Correction - Urgent Field Safety Notice Letter dated March 26, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to retain a copy of this document along with their most current product labeling, advise the appropriate personnel of the content of this letter, and for future reference this document will be posted to the Varian Medical Systems customer support website: http://www.MyVarian.com. The letter stated that a Varian Medical Systems has developed a technical correction and a Customer Service representative will contact customers to schedule installation. For question contact your local Varian medical Systems Customer Support District or Regional Manager. |  | 2508 | 2508 units | Worldwide Distribution - USA (Nationwide) and countries of: Australia, Austria, Bangladesh, Belgium, Brazil, Bulgaria, Canada, China, Columbia, Costa Rica, Cyprus, Czech Republic, Denmark, Dominican Republic, Estonia, Finland, France, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Kuwait, Latvia, Lebanon, Lithuania, Luxembourg, Macau, Malaysia, Mexico Morocco, Netherlands, New Zealand, Norway, Peru, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Singapore, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, Ukraine, United Arab E, and United Kingdom. | Terminated | July 09, 2014 | 470 | 2508 | Computer
Z-1546-2013 | 64424 | apparatus, traction, non-powered | HST | Nonpowered orthopedic traction apparatus and accessories. | Orthopedic | Neurology | 510(K) Exempt | Zimmer Patient Helper Bed Frame Adapter Assembly |  |  | February 14, 2013 | 2013 | Zimmer, Inc. | The bed frame adapter assemblies may have a weak weld where the patient helper adapter mounting bars attach to the tube. | Device Design | Zimmer sent an Urgent Device Removal letter dated February 14, 2013, to all affected consignees. The letter identified the product, the problem,and the action to be taken by the consignee. Consignees were advised to review the notification letter, follow the instructions contained in the letter, remove from use and quarantine the affected product. Consignees were also instructed to complete and return the fax back from to 877-787-0375 or zimmer3632@stericycle.com. Consignees were instructed to return the affected product back using the shipping label provided. If affected product was further distribution, consignees were asked to provide the customer's information to the firm on the back of the form. For questions regarding this recall call 866-201-9081. |  | 344 | 344 Units | Worldwide distribution: USA (nationwide) including states of: AL, AR, AZ, CA, CT, FL, HI, IL, KY, MA, MD, MN, MO, NE, NJ, NY, OH, OK, PA, TN, TX, WA, and WI; and country of Canada. | Terminated | May 01, 2014 | 441 | 344 | Computer
Z-1584-2013 | 65246 | No | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | CoaguChek XS System (HCP) |  |  | May 16, 2013 | 2013 | Roche Diagnostics Operations, Inc. | Roche has confirmed the potential for an undetected elevated INR test result with the CoaguChek XS, CoaguChek XS Plus, and CoaguChek XS Pro meters. In rare cases, instead of a value, an ERROR 6 message is displayed. | Software design | Rochet sent an Urgent Medical Device Correction letter dated May 16, 2013, via UPS Ground to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were issued the following instructions: If you are using Warfarin therapy in combination with antibiotics and/or chemotherapy drugs, which could potentially lead to an extremely high test result (&gt; 10 INR), you may in rare cases receive repeated ERROR 6 messages instead of a test result. If the ERROR 6 message is displayed repeatedly on your CoaguChek XS meter, immediately contact your treating physician.The CoaguChek XS PT Test Strip package insert; the online version of the CoaguChek XS PST system users manual; and the CoaguChek XS systems are being updated to provide additional ERROR 6 message information and actions. Updated versions will be available on the www.poc.roche.com website by the end of June 2013.Health Care Professionals are instructed to use an alternative method to determine INR results if an ERROR 6 message is displayed repeatedly on the CoaguChek XS, CoaguChek XS Plus, or CoaguChek XS Pro meters for the same patient. Self testing patients are instructed to contact their Health Care provider immendiately if the ERROR 6 message is displayed repeatedly on the CoaguChek XS.Please contact your Roche Account Manager or Roche Diagnostics Point of Care Technical Service, 24 hours a day, seven days a week at 1-800-428-4674 if you have questions about the information contained in this UMDC.For questions regarding this recall call 317-521-3711. |  | 86553 | 86553 meters | Nationwide Distribution | Terminated | April 16, 2014 | 335 | 89896 | Computer
Z-1587-2013 | 65371 | No | PAA |  |  | Radiology | PMA | Soma v Automated Breast Ultrasound System (ABUS), Automated Breast Ultrasound. |  |  | May 16, 2013 | 2013 | U-systems Inc | Devices labeled for an intended use not included in the existing 510(k) or PMA, in that, the user manual contains the Indication for Use of the soma v Platinum ABUS [a 510(k) product]. | Software design | On 6/5/13, an "URGENT MEDICAL DEVICE CORRECTION" letter was distributed to all affected users, with a description of the problem and user corrective action steps. One consignee received an Urgent Notice to Customer - Medical Device Correction letter on 5/16/13. Contact the firm at 408-328-7925 for questions regarding this notice. |  | 3 | 3 units | Distributed in California and Illinois. | Terminated | November 04, 2013 | 172 | 3 | Computer
Z-1574-2013 | 65342 | No | LKK |  |  | General Hospital | PMA | Medtronic Intrathecal Catheter Pump Segment Revision Kit, model 8596SC. Contents: 60-cm pump segment with attached sutureless pump connector, Spinal Segment Strain-relief sleeves, Pump segment strain-relief sleeves, Connector pin. Contents of inner package are STERILE. |  |  | June 03, 2013 | 2013 | Medtronic Neuromodulation | The Sutureless Connector (SC) Intrathecal Catheter connector has been redesigned to reduce the potential for occlusion at the catheter to pump interface. Medtronic is removing the unused products from the market that were manufactured with the previous design, and recommend the previous design no longer be used due to greater potential for misalignment and subsequent occlusion. | Device Design | The firm, Medtronic, sent an "Urgent: Medical Device Removal" letter dated May 2013 to its customers. The letter described the product, problem and actions to be taken. Representatives (Rep) are visiting all locations, beginning June 3, 2013, to retrieve devices with a Use By date of 2014 08 14 (August 14, 2014) or sooner. The Rep is leaving a letter with the hospitals to tell them that the Sutureless Connector Intrathecal Catheter connector has been redesigned and that they are removing unused devices. They also do not recommend using any of the old design. An Account Specific Customer Confirmation Form will be completed by the Rep. A copy of the completed form will be left with the Hospital along with the Urgent Medical Device Removal letter.If you have questions, please contact Medtronic Neuromodulation Technical Services at 1-800-707-0933 weekdays 7am - 6pm CST. |  | 115722 | 115,722 total | Worldwide distribution: US (nationwide) and countries of: Aruba, Australia, Austria, Belgium, Canada, Chile, China, Colombia, Costa Rica, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Finland, France, Georgia, Germany, Greece, Guadeloupe, Hong Kong, Hungary, Iceland, India, Ireland, Israel, Italy, Jordan, Korea, Kuwait, Lebanon, Lithuania, Luxembourg, Malta, Martinique, Morocco, Netherlands, Netherlands Antilles, New Zealand, Norway, Panama, Poland, Portugal, Puerto Rico, Qatar, Romania, Russian Federation, San Marino, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Thailand, Trinidad and Tobago, Turkey, United Arab Emirates, United Kingdom, Uruguay, and Venezuela. | Terminated | April 07, 2014 | 308 | 115722 | Computer
Z-1098-2013 | 64728 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris PC unit model 8015 with software version 9.12 |  |  | March 06, 2013 | 2013 | Carefusion Corporation | The recall was initiated because Carefusion has received reports of a communication error on the Alaris PC unit model 8015 with software version 9.12 when the Alaris EtCO2 module or Alaris SpO2 module is attached. | Software design | CareFusion sent an Urgent Medical Device Recall Notification letter dated March 6, 2013 to all affected customers. The letter identified the affected product, problem and action to be taken. Customers are required to confirm receipt of the notification by returning the Recall Response card to Carefusion by postage-paid, self-addressed mail, fax, or email. Carefusion representatives will contact customers by telephone to schedule field remediation. For questions contact CareFusion Support Center at 888-562-6018, Customer Advocacy at 888-812-3266 and/or Technical Support at 888-812-3229. |  | 9241 | 9,241 units | Worldwide Distribution - USA (nationwide) and the countries of Australia and China. | Terminated | September 02, 2014 | 545 | 9241 | Computer
Z-1562-2013 | 65435 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARTISTE with syngo RT Therapist version 4.3 |  |  | May 29, 2013 | 2013 | Siemens Medical Solutions USA, Inc | The potential for a safety issue, for the image quality degradation, that is addressed by the Customer Safety Advisory Notice when verifying the patient position based on 2D images that were acquired as single-exposure MV or In-Line kView images using the syngo RT Therapist 4.3. | Device Design | Siemens sent an Urgent: Medical Device Correction Customer Safety Advisory Notice letter to all affected customers. The letter identified the affected product, problem, what preventive measures can the user take and actions to be taken. Customers were advised to perform the preventive measures provided and inform all affected personnel immediately. The letter states that Siemens is currently working on a new version of the software. Customers are asked to include the Customer Safety Advisory Notice in the System Owner Manual chapter "Safety Advisory Letters". For question call 925-602-8157. |  | 8 | 8 units | Worldwide distribution | Terminated | November 26, 2013 | 181 | 8 | Computer
Z-1567-2013 | 65379 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | CARESCAPE Monitor B850 |  |  | July 20, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare become aware of 4 potential issues associated with the CARESCAPE Monitor B850. These issues include CARESCAPE Monitor B850 uDOM failure, CARESCAPE Monitor B850 unstable USB communication, CARESCAPE Monitor B850 remote printing issue with currently printing PRN 50 recorder and CARESCAPE Monitor B850 reboots due to an X server freeze. | Software design | GE Healthcare sent "Important Product Information" letter dated 12 July 2012 to all affected customers. The letter described the Subject (issues) , Instructions, Affected Product Details, Product Correction and actions to be taken. For questions contact Technical Support 1-800-558-7044 or your local Service Representative. |  | 12779 | 12,779 | Worldwide Distribution - USA Nationwide including the states of: AL, AK, AZ, AR, CA, CO, CT, DC, FL, GA, ID, IL, IN, IA, KS, KY, IA, ME, MD,MA, MI, MN, MS, MO, MT, NE,NV,NH, NJ, NM, NY, NC, OR, OK OH, PA, PR, SC, TN, TX, UT, VA, WA, WY, WI, WY, and the countries of: ARGENTINA, AUSTRALIA, AUSTRIA, BAHRAIN, BELGIUM, BRAZIL, CANADA, CHILE, COLOMBIA, COSTA RICA, CYPRUS, CZECH REPUBLIC, DENMARK,, EGYPT, ESTONIA, FINLAND, FRANCE, GERMANY, HONG KONG, INDIA, IRAQ, ISRAEL, ITALY, JAPAN, KOREA, KUWAIT, LEBANON, LITHUANIA, MALAYSIA, MALTA, MEXICO, NETHERLAND, NEW ZEALAND, NORWAY, PANAMA, PERU, POLAND, PORTUGAL, QATAR, SAUDI ARABIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY. | Terminated | June 30, 2014 | 710 | 12779 | Computer
Z-1558-2013 | 65129 | No | CCW | Rigid laryngoscope. | Anesthesiology | Anesthesiology | 510(K) Exempt | GlideScope Video Laryngoscope (GVL) 3, 4, 5 and GlideScope (AVL) 2, 3, 4 and 5 reusable blades |  |  | May 08, 2013 | 2013 | Verathon, Inc. | Verathon GlideScope Reusable GVL and AVL Blades are being recalled due to the potential risk of premature failure/breakage of the blade tip that may not be readily visible during routine inspection before or after intubation. | Device Design | Verathon issued an Urgent - Safety Alert and Medical Device Recall letter, dated May 10, 2013, to all affected customers The letter identified the affected product, problem and actions to be taken. The letter advises customers to discontinue the use of the specific serial numbers listed in the letter and asked that the affected product be returned to Verathon for replacement. The customers are also directed to fill out a return response form (included in the letter) regarding this recall.Consignees with questions are directed to contact their Verathon representative or Verathon Customer Care at 800-331-2313 (US and Canada) and 1-425-867-1348 for those consignees outside of US and Canada. the support team will be available Monday through Friday from 6 AM - 5 PM PST for the US and from 6 AM to 4:30 pm PST outside the US. Alternately they can e-mail Verathon at customerservice@verathon.com |  | 4784 | 4784 units worldwide (3,381 in the US and 1403 international) | Worldwide distribution: USA (nationwide) including states of : AK, AL, AZ, CA, CO, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY, Washington DC, Guam, and countries of : Argentina, Australia, Brazil, Canada, China, Ecuador, France, Germany, Hong Kong, India, Indonesia, Italy, Japan, Japan, Malaysia, Panama, Peru, Philippines, Singapore, South Korea, Taiwan, Thailand, The Netherlands, United Kingdom, and Venezuela. | Terminated | December 17, 2013 | 223 | 4784 | Computer
Z-1560-2013 | 65331 | No | JQW | Pipetting and diluting system for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Tecan Freedom EVO with Touch Tools Suite Version 3.0 |  |  | February 19, 2013 | 2013 | Tecan US, Inc. | WIth TouchTools 3.0 after the execution of a Rich User Prompt in Display and Wait mode, door lock errors are ignored. If an operator presses "Continue" following a Rich User Prompt and the door has been left open, the script will progress (continue without warning). This could lead to the instrument moving unexpectedly and potentially cause an injury to the operator if they are physically inside t | Software change control | The firm, Tecan, sent an "Urgent Field Corrective Action" letter dated February 15, 2013 to its consignees/customers and also contact them by phone February 19, 2013. The letter describes the product, problem and actions to be taken. The customers were instructed to not use Rich User Prompts in Display and Wait mode until TouchTools 3.0 Patch 1 or a higher version has been installed. Tecan is currently testing and releasing a patch that will correct this issue. The patch will be available by the end of February. Tecan will contact you upon its release as installation of the patch is required. If you have any further questions, please contact your local Tecan Helpdesk for assistance or call 919-361-5200 ext 19519. |  | 12 | 12 units | Worldwide distribution: US (nationwide) including states of: CA, KY, NC, OH, TX, and WI; and countries of: France, Germany and Italy. | Terminated | June 20, 2013 | 121 | 12 | Computer
Z-1559-2013 | 65438 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Echelon, Echelon Oval and Oasis MR imaging device systems; one unit per box. |  |  | January 31, 2013 | 2013 | Hitachi Medical Systems America Inc | Hitachi discovered a software error that can occur when simultaneously scanning a patient while processing images previously acquired from a different patient. The software error can occur randomly. The error causes an image in the processed dataset to have the same unique identifier (UID) as in the scanning dataset. The processed image is stored in the place of the original image, which is lost. | Software design | Hitachi sent a Device Correction Notice beginning on January 31,2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Software correction is targeted to start by March 1, 2013. All sites will be visited by Hitachi Service to perform corrections. All letters are recorded by tracking number and date received by the customers.For further questions please call 800-800-3106 x 3720. |  | 45 | 45 system units | Worldwide Distribution - US Nationwide including the states of AZ, CA, CO, DE, FL, IL, IN, KS, KY, MD, NC, ND, NY, OH, OR, PA, TX, VA, WA, and WI. One (1) system unit was also distributed to Mexico. | Terminated | September 20, 2013 | 232 | 45 | Computer
Z-1551-2013 | 65229 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, PET VCAR. |  |  | May 09, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue with respect to the automated PERCIST response criteria calculation applied in the PET VCAR application on Advantage Workstation or AW Server. | Software design | GE Healthcare sent an "Urgent Medical Device Correction" letter dated May 14, 2013, to all affected consignees. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. Customers with questions were instructed to call one of the following numbers:United Sates 800-437-1171Japan 0120-055-919For other countries they were instructed to contact their local GE Healthcare Service Representative.For questions regarding this recall call 407-475-4717. |  | 231 | 231 (38 US; 193 OUS) | Worldwide Distribution - USA including CA, CO, CT, FL, GA, IL, NY, NY NC OH, TX, VA, WV, WI.Internationally to: AUSTRALIA, AUSTRIA, BAHARIN, CANADA, CHILE, CHINE, CZECH REPUBLIC, DENMARK, FINLAND, FRANCE, GEORGIA, GERMANY, HONG KONG, INDIA, INDONESIA, ITALY, JAPAN, KOREA, MEXICO, MOROCCO, NIGERIA, POLAND, ROMANIA, RUSSIA,SAUDI ARABIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TURKEY, UNITED KINGDOM. | Terminated | January 21, 2014 | 257 | 231 | Computer
Z-1548-2013 | 65212 | No | PAO | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Superior Approach Resection Guide Assembly. |  |  | May 15, 2013 | 2013 | Biomet, Inc. | Biomet has initiated this action following an investigation which identified that the slot on the resection guide is offset in the wrong direction.When the slot on the resection guide is offset in the wrong direction, the Steinman pin that is intended to hold the guide into place may not come into contact with the bone. | Device Design | On 5/15/2013 an URGENT MEDICAL DEVICE RECALL notification was sent to all consignees detailing to immediately locate and discontinue use of the product and return to Biomet.The following actions are REQUIRED:-Immediately locate and remove the identified device(s) listed below from circulation .-Carefully follow the instructions on the enclosed "FAX Back Response Form" .-Fax a copy of the Response Form to 574-372-1683 prior to return of product.-Use priority carrier for your shipment.- If you have further distributed this product. you MUST notify hospital personnel of this action via the enclosed "Dear O.R. Manager" notice. This letter MUST be given to hospital personnel responsible for receiving recall notices. |  | 17 | 17 | Worldwide Distribution-USA (nationwide) including the states of UT, TX, AL, KY, and WI, and the countries of Australia, The Netherlands, Korea, and Canada. | Terminated | June 09, 2014 | 390 | 17 | Computer
Z-1537-2013 | 65376 | Yes | MNT | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Respironics V60 Ventilator, Model #V60 |  |  | June 04, 2013 | 2013 | Respironics California Inc | Respironics California, Inc. has initiated a recall on the V60 ventilator because of an issue with the software on the V60 Power Management Board Assembly. If a component fails on the Power Management Board Assembly, it may cause ventilator support to be lost with potentially no audible alarm from the ventilator. | Software design | The firm, Philips Healthcare, sent a "MEDICAL DEVICE CORRECTION" letter dated June 3, 2013 to all customers who received the V60 Ventilators. The letter describes the product, problem identified and the actions to be taken. The customers were informed that the V60 ventilator may continue to be used in accordance with its directions for use, pending the completion of the update. In addition, the customers were instructed to refer to their manual for additional information on warnings. A Philips Field Service Engineer, Approved Sevice Provider or Distributor will be contacting the customers to schedule a no-cost update and replacement of the Power Management Board Assembly Software on all V60 ventilators shipped from the manufacturer prior to April 1, 2013.If you need any further information or support concerning this issue, contact your local Philips Respironics representative at (800) 722-9377. |  | 19200 | 19,200 units total (12,060 units in US) | Worldwide distribution: US (nationwide) and countries of: Argentina, Australia, Austria, Bahrain, Belgium, Bolivia, Brazil, Canada, Chile, China, Croatia, Czech Republic, Denmark, Dominican Rep., Ecuador, Egypt , Finland, France, Germany, Hong Kong, Iceland, India, Indonesia, Iran, Ireland, Israel, Italy, Kenya, Kuwait, Lebanon, Libya, Malaysia, Malta, Mexico, Morocco, Namibia, Netherlands, New Zealand, Norway, Oman, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Puerto Rico, Qatar, Reunion, Russian Fed., Saudi Arabia, Singapore, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom, Utd.Arab.Emir. and Vietnam. | Terminated | April 24, 2014 | 324 | 19200 | Computer
Z-1133-2013 | 64748 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare Muse v7 cardiology information system. Model number 2026443-001: MUSE 7.0 app cd, Model Number 2026443-015: MUSE 7.0.2 app cd, Model number 2026443-017: MUSE 7.1 app cd, Model number 2026443-029:MUSE 7.1.1 A app cd, and Model Number 2026443-031: MUSE 7.1.1 app cd. |  |  | January 04, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare conducted a recall on the Muse v7 because incorrect medications were displayed on the MUSE, which could result in incorrect treatment. | Software design | Consignee letters were not sent for this recall. GE Healthcare's recall strategy was to update all affected systems that were distributed. Updating of the devices was carried out from February 2012 to August 2012 by GE Field personnel. |  | 963 | 963 | Worldwide Distribution -- USA, including all states except ME, and the U.S. Virgin Islands; and, the countries of AUSTRALIA, AUSTRIA, BELGIUM, CANADA, CHILE, CHINA, CZECH REPUBLIC, DENMARK, ESTONIA, FINLAND, FRANCE, GERMANY, GREECE, INDIA, IRELAND, ISRAEL, ITALY, JAPAN, KOREA, KUWAIT, MALAYSIA, NETHERLAND, NEW ZEALAND, NORWAY, PHILIPPINES,S QATAR, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA SPAIN, SWEDEN, SWITZERLAND, TAIWAN, TURKEY, UNITED ARAB EMIRATES, and UNITED KINGDOM. | Terminated | July 19, 2013 | 562 | 963 | Computer
Z-1547-2013 | 65210 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Biomet Trauma Phoenix Retrograde Femoral Connecting Bolt Trauma Fixation Systems. |  |  | April 29, 2013 | 2013 | Biomet, Inc. | Biomet Trauma ("Biomet") has initiated a recall of Retrograde Femoral Connecting Bolt, which involves Part Number 14-442021. The connecting bolt has an undersized diameter specification that may cause an interference fit with the 4mm hex driver (Part Number: 41024) near the edge of the tolerance. If the 4mm hex driver becomes stuck in the connecting bolt it cannot be taken apart and a delay in su | Device Design | Biomet sent an URGENT MEDICAL DEVICE RECALL notification letter dated April 29, 2013 to all affected customers. The letter identifed the affected product, problem and actions to be taken. The letter instructed customers to: immediately locate and remove the identified affected device(s), carefully follow the instructions on the enclosed "FAX Back Response Form", fax a copy of the Response Form to 574-372-1683 prior to return of product, use priority carrier for your shipment, if you have further distributed this product, you MUST notify hospital personnel of this action via the enclosed "Dear O.R. Manager" notice. For questions call 574-372-1570. |  | 75 | 75 | Worldwide Distribution - USA Nationwide: CA, TX, NY, PA, NJ, FL, OH, MI, CO, SD, IN, MT, GA, and WI. and the countries of Netherlands, Costa Rica, Japan, and San Juan. | Terminated | May 06, 2015 | 737 | 75 | Computer
Z-1536-2013 | 65091 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | 90 Degree Cannulated Infant Blade Plate 35mm x 5mm x 3 Hole |  |  | April 19, 2013 | 2013 | OrthoPediatrics Corp | Post Market Surveillance of the 90 degree Cannulated Infant Blade Plate indicated that there have been five reports concerning bending of the plate intra-operatively. | Device Design | OrthoPediatrics sent an Urgent Voluntary Recall notifiation via telephone and email to all affected consignees. The letter instructed consignees to take possession of any affected devices and quarantine immediately. A Mandatory Reply Form is included with the notification letter to be completed and returned. Please contact Greg Teghtmeyer if you have any questions (574-268-6379). |  | 130 | 130 devices were shipped | USA Nationwide Distribution | Terminated | September 19, 2013 | 153 | 130 | Computer
Z-1164-2013 | 64804 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Healthcare, Xeleris 2 Processing &amp; Review Workstation. |  |  | October 06, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware of a software issue in your Xeleris 2.1 workstation (that is used for processing Discovery 530/570 data), that may impact patient safety.On imported CT scans, Myocardial Perfusion Imaging (MPI) studies may create artifacts in attenuation corrected images. This may lead to a potential for misdiagnosis. | Software design | GE Healthcare sent an Urgent Medical Device Correction letter dated October 5, 2011, to all affected consignees. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. Consignees were instructed to contact their local GE Healthcare Service Representative if they have questions or concerns. They may also contact GE Healthcare Call Center , United States or Canada at 800-437-1171.Other countries 262-896-2890.For questions regarding this recall call 262-513-4122. |  | 11 | 11 | Worldwide Distribution - USA including GA, and NY; and countries of: Canada, Israel, Italy and Switzerland. | Terminated | February 04, 2014 | 852 | 11 | Computer
Z-1137-2013 | 64900 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Healthcare, AW Server. |  |  | June 28, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue with respect to the Results Viewer in the AW Server.After starting the Results Viewer with the selected series of a patient, the pixel and certain annotation data from a different patient is displayed in one of the viewports. This occurs on an occasional basis, when multiple users are simultaneously accessing the Results Viewe | Software design | GE Healthcare sent an "Important Product Information" letter dated June 27, 2012. The letter was addressed to Hospital Administrators , Radiology Managers and Risk Managers. The letter described the Subject, Instructions and Affected Product Details. GE Healthcare will correct all affected product by providing a software upgrade at no cost. For questions call: US 800-437-1171, Japan 0120-055-919, and/or local GE Healthcare Service Representative. |  | 429 | 429 | Worldwide Distribution - USA Nationwide including the states of: AK, AZ, CA, CO, CT, DE, DC, FL, GA, HI, IL, IN, IA, LA, MD, MI, ,N, ,S, MO, MT, NJ, NM, NY, NC, OH, OK, PA, SC, SD, TN,TX, UT, VA, WA, WI. and the countries of : AUSTRALIA, AUSTRIA, BELGIUM, BOSNIA, BRAZIL, CANADA, DENMARK, EGYPT, ESTONIA, FINLAND, FRANCE, GERMANY, GREAT BRITAIN, GREECE, GUADELOUPE, HONG KONG, INDIA, ISRAEL, ITALY, JAPAN, KOREA, LIBYA, MALAYSIA, NETHERLANDS, NEW ZEALAND, NORWAY, PANAMA, POLAND, QATAR, REUNION, ROMANIA, RUSSIA, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, THAILAND, TURKEY and UNITED ARAB EMIRATES. | Terminated | February 04, 2014 | 586 | 429 | Computer
Z-1530-2013 | 65330 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ECHELON 60mm Endoscopic Linear Cutter Reloads Black in Flexible Blister pack with Tyvek lid. |  |  | May 13, 2013 | 2013 | Ethicon Endo-Surgery Inc | Ethicon Endo-Surgery is initiating a voluntary worldwide recall for ECHELON 60mm Black Reload (ECR60T) due to the potential for incomplete staple line formation from reload damage during the firing sequence. | Device Design | Ethicon sent a Urgent: Medical Device Recall letter dated May 29, 2013, to all affected customers. The letter identified the product the problem and the action to be taken by the customer.Customers were instructed1.To examine your inventory immediately to determine if you have affected product manufactured by Ethicon and remove the affected product(s).See Attachment A for assistance in identifying affected product.2. Complete the Business Reply Form (BRF) Attachment C and fax or email it to Stericycle at 1-888-943-4190. Please reference Event 4706. If you have product to be returned, keep a copy of this form for your records.Your JJHCS number has been pre-populated on the BRF. Replacement product will be issued for all returned recalled product.3. All affected product must be returned by August 30, 2013, to receive replacement product. To return affected product, photocopy your completed Business Reply Form, placeit in the box with your product, and affix the pre-printed UPS shipping label to return to Stericycle. Returned ECHELON 60mm Black Reload (ECR60T) will be replaced with corrected ECHELON 60mm Black Reload (ECR60T) as soon as supply becomes available.4. If you do not have affected product in your inventory, you should still complete and return the Business Reply Form, indicating you have no affected product.5. Please share this information with all of the appropriate staff at your facility.6. If you need assistance identifying an alternative product code, please speak directly with your Ethicon Sales Representative. If you need clinical or product support, please contact your local sales representative or call ourCustomer Support Center, 1-877-ETHICON (877-384-4266 OPTION 6 to speak with a registered nurse). If you need additional shipping labels or a communications package, contact the Stericycle support center at 1-888-943-2397 and reference Event 4706. |  | 57540 | 57,540 Units | Worldwide Distribution - US Nationwide all fifty states and including the countries of Canada, Belgium, India, Malaysia, Saudi Arabia, Brazil, Israel, Mexico, Singapore, Jordan, New Zealand, Taiwan, Chile, Korea, Puerto Rico, Turkey, Columbia, Kuwait, Qatar, United Arab Emirates, Egypt, Lebanon and Russia. | Terminated | February 11, 2014 | 274 | 57540 | Computer
Z-1531-2013 | 65341 | No | LZF | Infusion pump. | General Hospital | General Hospital | 510(k) | VIA Medical Collection Bag; The collection bag is an accessory waste bag used in conjunction with the sterile LVM blood gas sensor with the LVM Monitor. The collection bag is packaged in a tyvek pouch which is included in either replacement kit (777-3103) or LVM sensor kit (777-3102). |  |  | May 24, 2013 | 2013 | International Biomedical, Ltd. | International Biomedical has received reports of interference between the collection bag and the extension set connectors. Evaluation of the affected components indicates that the collection bag connector can become occluded due to an over-tightening of the luer lock by the end user. To alleviate the occurrence of this issue in the fixture, the connector on the collection bag has been improved to | Device Design | International Biomedical decided to recall and sent "Product Bulletin" notification letters to their consignees on 05/24/2013 or 05/30/2013. The letter described the product issue and provided customers with recommended actions. Questions were directed to (512) 873-0033. |  | 7000 | 7,000 | Nationwide Distribution. | Terminated | July 29, 2013 | 66 | 7000 | Computer
Z-1149-2013 | 64641 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | MAQUET CARDIOHELP Support System: CARDIOHELP-I Software (up to version 3.3.0.0) |  |  | March 07, 2013 | 2013 | Maquet Cardiovascular, LLC | It has come to the attention of MAQUET that there have been episodes of brief unexpected shutdown, followed by an automatic device restart of the human machine interface (touch screen) of the CARDIOHELP device. The human machine interface (HMI) is the central display of the CARDIOHELP device, where measured values, settings and alarms are displayed and adjusted/confirmed, with the exception of bl | Software change control | MAQUET sent an Urgent Medical Device Field Correction letters dated March 7, 2013 via Federal Express to all affected customers. The letter identifies the affected product, problem and actions to be taken. Customers are asked to complete and submit the included Customer Fax-back Form. Maquet Service will exchange all affected products with the updated version. For questions MAQUET Customer Service at 888-627-8383 (press option 2, followed by option 2). |  | 154 | 154 units affected | Worldwide Distribution - USA Nationwide and the following countries: Australia, Austria, Belarus, Belgium, Canada, Columbia, Croatia, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, Hong Kong, Iceland, Iran, Ireland, Italy, Jordan, Kazakhstan, Korea, Kuwait, Luxembourg, The Netherlands, Norway, Oman, Poland, Portugal, Russia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Thailand, Turkey, United Arab Emirates and the United Kingdom. | Terminated | August 22, 2013 | 168 | 154 | Computer
Z-1144-2013 | 64683 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Vitrea CT Multi-Chamber Cardiac Functional Analysis application on Vitrea version 5.2 (Vitrea Enterprise Suite 1.3 and Vitrea fX 3.1) and subsequent versions of VitreaWorkstation, Vitrea fX, VitreaAdvanced and Vitrea Enterprise Suite, through version 6.4. |  |  | March 15, 2013 | 2013 | Vital Images, Inc. | Vital Images has made a decision to perform a field correction to the Vitrea CT Multi-Chamber Cardiac Functional Analysis application. Vital Images has identified a defect in which numerical values may be incorrect in the Results Table under the Cardiac Analysis portion of the user interface. | Software design | Vital Images sent an Urgent Field Safety Notice dated March 15, 2013, to all affected consignees. The letter described the problem and the product affected by the recall. Letter provided a series of steps to follow which included to cease use of the device and to send the signed enclosed Effectiveness Check form to fieldnotices@vitalimages.com or fax to 1-952-540-3717. For questions customers were instructed to call Customer Support at 1-800-208-3005.For questions regarding this recall call 952-487-9500. |  | 102 | 102 | Worldwide Distribution - USA including AL, AR, CA, CO, IA, MI,MN, MO, NY, NC, OH, OR, PA, TX, and UT.Internationally to AUSTRIA, BELGIUM, CANADA, CHINA, COLOMBIA, GEORGIA, JAPAN, KOREA, MACEDONIA, NETHERLANDS, POLAND, PORTUGAL, RUSSIAN FEDERATION, SINGAPORE, SWEDEN, SWITZERLAND, TURKEY, UKRAINE, UNITED KINGDOM. | Terminated | February 27, 2015 | 714 | 102 | Computer
Z-1522-2013 | 65283 | No | NME | Dental bur. | Dental | Dental | 510(k) | StarDental Classique(TM) Diamond Dental Bur. |  |  | May 28, 2013 | 2013 | Dental EZ Stardental Division | Recall was initiated because a manufacturing issue was found that prevented the diamond coating from adhering properly to the shank of the diamond bur for these lots. | Equipment maintenance | Dental EZ Inc. sent an "URGENT: MEDICAL DEVICE RECALL" letter on May 28, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at 866-383-4636 ext.4350 for questions regarding this recall. |  | 3560 | 3,560 | Worldwide Distribution-USA (nationwide) including the states of CA, FL, IA, IN, NJ, NV, NY, PA, TN, TX, and WI, and the country of Japan. | Terminated | March 18, 2014 | 294 | 3560 | Computer
Z-1141-2013 | 64723 | No | JGS | Sodium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | HORIBA ABX PENTRA 400 I.S.E. Module, Model Numbers: P400ISE110EN02 and P400ISE110US02, All Serial Numbers, Software Versions 5.0.7 and Below. |  |  | March 20, 2013 | 2013 | Horiba Instruments, Inc dba Horiba Medical | HORIBA Medical is informing all ABX PENTRA 400 Analyzer with I.S.E. Module customers that there are noLinearity Limit Flags for the I.S.E. assays currently programmed into the ABX PENTRA 400 Analyzer software. Allother assays run on this analyzer do have Linearity Limit Flags to alert the user to values that fall below or above thevalidated linearity range for each assay. | Software design | A recall letter, Field Safety Notice &amp; Acknowledgement mailed to the customers March 20, 2013. The letter included the product information, reason for recall, instructions on what to do with the recalled product, and contact information:"If you have any questions or concerns, please contact your local HORIBA Medicalrepresentative. We sincerely apologize for any inconvenience this may cause your laboratory. Please contact your Customer Service Representative at 1-888-903-5001 (option 3) for any additional questions related to this customerletter." |  | 905 | 905 units | Nationwide Distribution. | Terminated | September 26, 2014 | 555 | 905 | Computer
Z-1488-2013 | 65062 | No | GDC | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Hand Controls for Operon D750, D752, D760, D820 and D850 |  |  | March 12, 2013 | 2013 | Berchtold Corp. | The firm received complaints for unintended movements. | Device Design | Berchtold sent an Urgent Medical Device Recall letter in March 2013, to all customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were recommended to clean the affected product and use the instructions to reduce the possibility of unintended movements until the firm could replace the affected product. Customers with questions were advised to contact the Berchtold Operon Pendanct Recall Coordinator at 800-243-5135, or 843-569-6100, ext 306.For questions regarding this recall call 843-569-6133. |  | 3763 | 3763 | Nationwide Distribution | Terminated | May 08, 2017 | 1518 | 3763 | Computer
Z-1516-2013 | 65113 | No | LFL |  |  | General &amp; Plastic Surgery | 510(k) | American Optisurgical's TX1 Tissue Removal System Console. The console is a software unit that is indicated for use in surgical procedures. |  |  | March 30, 2013 | 2013 | American Optisurgical Inc | The recall has been initiated due to an overheating protection voltage suppressor on the redundant safety protection circuit board. Secondary to the recall, American Optisurgical, Inc. will be making a correction to add spacers between the valve and cassette to correct a small irrigation leak that can occur after the priming cycle. Use of this product may cause the device to lose power and stop wo | Device Design | "Urgent Medical Device Recall" notifications letters were sent to customers on May 20, 2013 via USPS registered mail. Included in the Recall Notification Letter were the recommended actions and instructions for customers. Questions were referred to (949) 580-1266. |  | 194 | 194 units | Worldwide distribution, US, including the states of United States: Alabama, Arkansas, Arizona, California, Colorado, Florida, Georgia, Iowa, Illinois, Indiana, Kansas, Kentucky, Massachusetts, Maryland, Michigan, Minnesota, Missouri, North Carolina, Nebraska, New Jersey, New York, Ohio, Pennsylvania, South Carolina, Tennessee, Texas, Utah, Virginia, Washington Wisconsin, West Virginia and the countries of Singapore and the Philippines. | Terminated | March 24, 2014 | 359 | 194 | Computer
Z-1480-2013 | 64927 | No | OSI | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | E-Z WIDER BARIATRIC CHAIR BED; Single Motion Expandable Width Technology |  |  | August 02, 2012 | 2012 | Kreg Medical Inc. | An EZ Wider bed (capacity of 1000 lb. safe working load collapsed during patient use, leaving the patient in a forward tilt position. | Device Design | Kreg Medical sent an Urgent Voluntary Medical Device Recall letter dated May 10, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Kreg Medical started replacing potentially defective beds with refurbished beds in August 2012. All of the potentiall defective beds were refurbished by November 20, 2012. Customers were requested to send back the Medical Device Recall Return Response to confirm receipt of the notification letter.For questions regarding this recall call 312-275-7070. |  | 59 | 59 beds | Nationwide Distribution including IL and IN. | Completed |  |  | 59 | Computer
Z-1472-2013 | 65093 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Versaport Bladeless Optical Fixation Cannula - 5mm |  |  | May 06, 2013 | 2013 | Covidien LP | Seals may disengage from the cannula which may result in a component inadvertently disengaging into the patients cavity | Device Design | Covidien sent an Urgent Medical Device Recall letter dated May 6, 2013,via federal express to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use of the product and return all inventory to Covidien Surgical Solutions, 195 McDermott Road, North Haven, CT 06473. Customers were requested to return completed Response Form.Customers with questions were instructed to contact their Covidien representative or Covidien Customer Service, Monday through Friday, Sam - 6:30pm ET, at (800) 962-9888, option 1, and then option 2.For questions regarding this recall call 203-492-8165. |  | 57461 | 57461 untis | Worldwide Distribution - USA (nationwide) and internationally to the following countries:CanadaAustriaBahrainBenoniCenturionChatsworthCzech RepublicDenmarkEgyptFinlandFranceGermanyIrelandIsraelItalyJapanKuwaitNetherlandsNorwayPortugalQATARSaudi ArabiaSerbiaSingaporeSouth AfricaSpainOMANSwedenSwitzerlandUAEUnited KingdomVenezuela | Terminated | August 12, 2014 | 463 | 218221 | Computer
Z-1476-2013 | 64550 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | SlingBar Standard. |  |  | April 30, 2013 | 2013 | Hill-Rom, Inc. | Retrospective review found 4 sling bars that may present a potential for impalement if a patient, caregiver or bystander were to fall onto the hooks based on a simulated test method. In effort to eliminate any risk of injury, Hill-Romis requesting products be removed from field use andreplaced with an updated design. | Device Design | On 4/30/13 an URGENT MEDICAL DEVICE RECALL CORRECTION was issued to all consignees detailing the hazards and actions that are needed. Included with this notification is a Response Form/Receipt and a Replacement Guide, complete and return the form to Hill-Rom as soon as possible. Hill-Rom will send you the applicable number of sling bars to you at no cost. After you receive the new sling bars, we request that you discard the old ones. |  | 726 | 726 | Worldwide Distribution-USA (nationwide) including the states of AZ, CA, CO, FL, IL, IN, KS, LA, MA, MI, MN, MO, NC, NH, NJ, NV, OH, OK, PA, SC, TN, UT, WA, and WI, and the countries of Canada, Australia, Austria, Belgium, Switzerland, Germany Denmark, Spain, Finland, France, United Kingdom, Greece, Hong Kong, Ireland, Israel, Italy, Japan, Netherlands, Norway, New Zealand, Portugal, Sweden, and Turkey. | Terminated | November 28, 2014 | 577 | 4710 | Computer
Z-1156-2013 | 64584 | No | MNI | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Orthopedic Alliance Spine System (alias "Blue &amp; Gold") |  |  | March 07, 2013 | 2013 | Orthopedic Alliance LLC | Orthopedic Alliance is recalling Blue &amp; Gold Implants and Instruments because the Blue &amp; Gold system is not supported by adequate testing and documentation to demonstrate that it meets performance or safety standards. These inadequacies might result in product peformance failures that could cause patient harm due to implant breakage, loosening, or inadequate sterilization. | Device Design | A recall letter dated March 7, 2013 was sent to one customer who purchased the Blue &amp; Gold Implants and instruments. The letter informed the customer that the Blue &amp; Gold system is not supported by adequate testing and documentation to demonstrate that it meets performance or safety standards. These inadequacies might result in product performance failures that could cause patient harm due to implant breakage, loosening, or inadequate sterilization. Customer is informed of the actions to be taken. Customer is instructed to complete the customer response form and return it by FAX to (858) 764-9739. |  | 3300 | 3,300 units | Distributed in Texas. | Terminated | March 26, 2014 | 384 | 3300 | Computer
Z-1446-2013 | 64968 | cassettes, tissue | IDZ | Tissue processing equipment. | Pathology | Pathology | 510(K) Exempt | LPC Fine Mesh Cassette for the Leica LPC Printer. Part No: 14060546848; 14060546846; 14060546845; 14060546933; 14060546932; 14060546852; 14060546850; 14060546847. The intended use of this device is for storage of specimens/tissues during processing. |  |  | April 04, 2013 | 2013 | Leica Microsystems, Inc. | An LPC Fine Mesh Cassette failed to stay closed during processing operations. | Device Design | A Medical Device Correction letter dated April 17, 2013 was sent to customers via USPS 1st class mail. The product issue was described and recommended actions were provided. Customers were requested to acknowledge receipt of the letter by fax-back. Customer questions were directed to (847) 405-5413. |  | 20160 | 20,160 units | Worldwide distribution: US (nationwide) to states of AZ and KY; and the country of New Zealand. | Terminated | July 21, 2017 | 1569 | 20160 | Computer
Z-1049-2013 | 64463 | No | CEO | Phosphorus (inorganic) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens Dimension(R) Phosphorus (PHOS) Flex(R) reagent cartridge (DF61) |  |  | February 22, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has initiated a voluntary field action for Dimension(R) PHOS Flex(R) Reagent due to reagent interaction causing falsely depressed Creatinine results when EZCR is processed from open wells that are in close proximity to open wells of PHOS reagent. | Device Design | Siemens sent an Urgent Medical Device Correction letter dated February 22 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Action to be taken by Customer:1. Customers with two or more Dimension systems should perform testing of PHOS and EZCR on separate Dimension systems.2. Customers with only one Dimension system should discontinue use of either PHOS or EZCR on the Dimension system. Please contact your local sales representative to discuss alternate solutions that will best meet the needs of your laboratory.Please complete the attached Effectiveness Check form and fax it to the Siemens Technical Solutions Center at (302) 631-8467, to indicate you have received the information. If you have any questions, please contact the Siemens Technical Solutions enter or your local Siemens technical support representative. We apologize for the disruption this situation will cause your laboratory. We are working diligently to resolve this issue with a permanent solution.. |  | 60674 | 60,674 | Nationwide Distribution including the states of AK, AL, AR, AZ, CA, CO, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WV. | Terminated | January 29, 2015 | 706 | 60674 | Computer
Z-1870-2013 | 65033 | No | FZY | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | S-LIFT Slap Hammer, Part Number: 11-50204, Revision A |  |  | April 25, 2013 | 2013 | SpineFrontier, Inc. | Malfunction of slap hammer. | Device Design | The firm, SpineFrontier, sent an advisory notice entitled "URGENT April 25, 2013 S-LIFT Slap Hammer Removal" dated April 25, 2013 via email to its customers. The notice described the product, problem and action to be taken. The notice requested the removal of all distributed SLIFT Slap Hammer with listed lot number. The customers were instructed to complete and return the URGENT S-LIFT Slap Hammer Removal ACKNOWLEDGEMENT FORM via fax to: SpineFrontier, Inc., 978-232-3991 or email to customerservice@spinefrontier.com.Please feel free to contact SpineFrontier at 978-232-3990 with any questions. |  | 3 | 3 | US distributions to states of: KS and TX. | Terminated | January 09, 2014 | 259 | 3 | Computer
Z-0877-2013 | 63957 | No | KMI | Bed-patient monitor. | General Hospital | General Hospital | 510(K) Exempt | QuietCare-Networked product, Facility Server/model: QC101601; |  |  | December 16, 2012 | 2012 | Intel-GE Care Innovations LLC | Care Innovations detected a software anomaly: During the engineering investigation into the cause of why a Wander Alert was not generated, a defect in the implementation of the Wander and Night Motion ADLs was discovered. | Software design | Care Innovations sent a "DEVICE CORRECTION NOTIFICATION" letter dated 12/20/12 to all affected customers. The letter identified the product and problem. Correction was performed remotely by the firm. Contact the firm at 917-232-4775 for questions regarding this notice. |  | 19 | 19 units | Nationwide Distribution. | Terminated | February 26, 2013 | 72 | 19 | Computer
Z-2257-2013 | 66126 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Healthcare, Seno Advantage, 5183685, 5183685-2, 5183685-3, 5183685-4, 5183685-5, 5183685-6, 5183685-7, 5183685-8, 5183685-9, 5342425, 5342425-2, 5342425-3, 5342425-4, 5342425-5. |  |  | September 20, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has recently become aware that patient mismatch without a warning message could occur on Seno Advantage (SA) workstations under certain circumstances. This issue may impact patient safety. | Software design | GE Healthcare sent an Important Product Information letter dated September 19, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. GE Healthcare has recently become aware that patient mismatch without a warning message could occur on Seno Advantage (SA) workstations under certain circumstances. This issue may impact patient safety.The Contact Information (US 800-437-1171, Japan 0120-055-919, Other countries contact local GE Healthcare Service Representative). |  | 179 | 179 | Worldwide Distribution - USA (nationwide) including the states of : AL AZ, CA, CO, CT, FL, GA, IL, IN, MD,MI, MO, NJ, NY, NC, OK, PA, TX and WI.,and the countries of : AUSTRIA, BELGIUM, CANADA, FINLAND, FRANCE, GERMANY, GREAT BRITAIN, IRAQ, JAPAN, LEBANON, NORWAY, PORTUGAL, SPAIN, SWEDEN and TAIWAN. | Completed |  |  | 179 | Computer
Z-2214-2013 | 66098 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Cypher MIS Screw System, Biomet Spine Parsippany, NJ 07054. The Cypher MIS Screw System is a non-cervical spinal fixation device. |  |  | August 13, 2013 | 2013 | Ebi, Llc | Several complaints were reported to exhibit screw head splay in the Cypher MIS Screw System. | Device Design | Biomet Spine &amp; Bone Healing Technologies sent "Urgent Medical Device Recall Notice" letters with Fax-Back Response Forms dated 8/13/2013 to distributors and OR managers via Fed Ex. Follow-up e-mails with a copy of the Recall Notice were sent to distributors on 8/13/2013. Customers with questions may contact 973-299-9300 x2162 Monday to Friday from 8:00 am to 5:00 pm EST. |  | 427 | 427 units | Nationwide Distribution (US), including the states of IN, MD, MS, NC, OH, PA, SC, and TX. | Terminated | March 11, 2015 | 575 | 427 | Computer
Z-2241-2013 | 66111 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Healthcare, Advantage Workstation, READY View Application. |  |  | March 27, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has become aware that when analyzing fMRI images using the READY View application software, incorrectfMRI activation maps may be computed and displayed. | Software design | GE provided Urgent Medical Device Correction letters to consignees in the following countries: Canada, Finland, France, Germany, Guadeloupe, Italy, Japan, Latvia, Martinique, Netherlands, Norway, Poland, Reunion, Romania, Saudi Arabia, Singapore, Spain,Sweden, Switzerland, Turkey and the United Kingdom.The letters were dated March 27, 2012 and were addressed to Hospital Administrators, Radiology Mangers and Risk Managers. The letters described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information (US 800-437-1171, Japan 0120-055-919 and for other countries contact your local GE Healthcare Service Representative).*Letters to other consignees is being planned for September 2013. |  | 318 | 318 (28 USm 290 OUS) | Worldwide Distribution-USA including the states of IL, MN, NJ, NY, and VA, and the countries of ALGERIA, AUSTRALIA, BELGIUM, BRAZIL, CANADA, DENMARK, FINLAND, FRANCE, GERMANY, GUADELOUPE, INDIA, INDONESIA, ISRAEL, ITALY, JAPAN,KOREA, LATVIA, MARTINIQUE, NETHERLANDS, NORWAY, PHILIPPINES,S POLAND, REUNION, ROMANIA, RUSSIA, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SPAIN, SWEDEN SWITZERLAND, THAILAND, TUNISIA, TURKEY, UNITED KINGDOM, and VIETNAM. | Terminated | November 21, 2014 | 969 | 318 | Computer
Z-0931-2013 | 64165 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Elekta Synergy XVI. Radiation Therapy Digital Imager. Product XVI R3.5.1, R4.2.1, and R4.5.1 |  |  | December 19, 2012 | 2012 | Elekta, Inc. | The DICOM tags do not contain sufficient information for XVI to identify the different volumes that have the same name. | Software design | Elekta sent an "IMPORTANT FIELD SAFETY NOTICE" dated December 2012 to all affected customers. The letter identifies the product, problem, and the actions to be taken by the customers. The IFSN informs the user in the TPS, to give the different volumes different structure names. After importing the patient data set from the TPS, make sure that the structure contours are correct. Users are requested to return the signed acknowledgement. Contact your local Elekta office for questions regarding this notice. |  | 164 | 164 | Nationwide Distribution-DC, Puerto Rico and the states of AL, AK, AZ, CA, CO, CT, FL, GA, IL, IA, KY, MD, MA, MI, MN, MS, MT, NE, NV, NH, NJ, NM, NY, NC, OH, OK, OR, PA, SC, TN, TX, VT, VA, WA, WV, and WI. | Open, Classified |  |  | 164 | Computer
Z-2248-2013 | 65893 | No | GKP | Coagulation instrument. | Hematology | Hematology | 510(k) | ACL TOP 700 CTS; PN 0000280050 |  |  | July 19, 2013 | 2013 | Instrumentation Laboratory Co. | Potential carryover issue that can can cause shortened APTT clotting times on the instrument. | Device Design | Instrumentation Laboratory sent an Important Product Correction letter on July 19, 2013, to all affected customers. On July 23, 2013, the same communication was circulated by the IL International Group in Milan, Italy to the affected international distributors and dealers. The interim field notification was followed by the release of corrective Software V5.1.0 and ParametersP-16.3.00 (Phase II) on September 30,2013.Consignees were advised of a potential carryover issue and a workaround was provided. Consignees were asked to fax the Mandatory Response Tracking Form to 781-861-4207 or email to ra-usa.ilww.comFor technical questions customers were instructed to contact the IL Technical Support Center at 1-800-678-0710, option 2. For product availability questions, customers were instructed to contact their IL Customer Service Representative at 1-800-955-9525, Option #2.For questions regarding this recall call 781-861-4467. |  | 185 | 185 | Worldwide Distribution - USA including AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WV. Internationally to Canada, Argentina, Australia, Austria, Belarus, Belgium, Brazil, Brunei, Chile, China, Colombia, Costa Rica, Czech Republic, Denmark, Dubai, Finland, France, Germany, Great Britian, Greece, Guatemala, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Korea, Kuwait, Luthuania, Luxembourg, Malaysia, Mexico, Netherlands, New Zealand, Norway, Panama, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Uruguay, Venezuela, and Viet Nam. | Terminated | October 27, 2014 | 465 | 4799 | Computer
Z-2253-2013 | 66168 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Discovery IGS 730 Cardiovascular X-Ray Imaging System. |  |  | June 19, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue with respect to Discovery IGS 730 Imaging Systems.A patient data management error may prevent the Discovery IGS 730 Imaging System from booting on start-up or re-booting during procedure. The boot issue may occur when a largeamount of patient data remains in the browser and is not deleted. The inability to boot the system ma | Software design | The firm, GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated June 19, 2013 to its Consignees/customers. The letter described the Safety Issue, Safely Instructions, Affected Product Detail, Product Correction and Contact Information. The consignees were instructed to please ensure that all potential users in your facility are made aware of this safety notification and the recommended actions. GE Healthcare will correct all affected systems at no cost to you. A GE Healthcare service representative will contact you to arrange for this correction.If you have any questions or concerns regarding this notification, please call : United States: 800 437 1171. For other countries, please contact your local GE Healthcare Service Representative. |  | 4 | 4 (2 US, 2 OUS) | Worldwide distribution: US (nationwide) including states of : PA and TN; countries of: FRANCE and SINGAPORE. | Terminated | October 25, 2013 | 128 | 4 | Computer
Z-2256-2013 | 66148 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, CARESCAPE Monitor B650. |  |  | July 12, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has recently become aware of a number of potential issues associated with the CARESCAPE Monitor B650.No injuries or illnesses have been reported. | Software design | GE Healthcare sent an Important Product Information letter dated July 12, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.GE Healthcare will provide a software and hardware correction at no charge once it is available. We will contact you to arrange for this correction.If you have any questions regarding this notification, please contact Technical Support (1-800- 558-7044) or your local Service Representative |  | 9466 | 9466 (3261 US, 6205 OUS) | Worldwide Distribution - USA (nationwide) including the states of : AL, AZ, AR, CA, CO, CT, FL, GA, ID, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MS, MO, NE, NV,NH, NH, NM, NY, NC, OH, OK OR, PA, PR, SC, TN, TX, VA, WA, WV, WI and WY., and the countries of : ARGENTINA, AUSTRALIA, AUSTRIA, BAHRAIN, BANGLADESH, BELGIUM, BOLIVIA, BRAZIL, BULGARIA, CANADA, CHILE, COLOMBIA, COSTA RICA. CROATIA, CZECH REPUBLIC, DENMARK, EGYPT, FINLAND, FRANCE, GERMANY, HONG KONG, HUNGARIA, ICELAND, INDIA, INDONESIA, IRAQ, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KOREA, KUWAIT, LEBANON, LITHUANIA, MALAYSIA, MOROCCO, NETHERLAND, NEW ZEALAND, NICARAGUA, NORWAY, OMAN, PAKISTAN, PANAMA, PARAGUAY, PERU, POLAND, PORTUGAL, QATAR, ROMANIA, RUSSIA, SAUDI ARABIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, SYRIA, THAILAND, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM and VENEZUELA. | Terminated | April 21, 2014 | 648 | 9466 | Computer
Z-0928-2013 | 64344 | Yes | MRM |  |  | Cardiovascular | PMA | Isoline Implantable Defibrillation Lead, Model Isoline 2CR5, 2CR6 and 2CT6. |  |  | January 28, 2013 | 2013 | Sorin Group USA, Inc. | Isoline defibrillation leads distributed may have internal insulation breach. | Device Design | Sorin C.R.M., srl sent an Urgent Medical Device Field Safety Notice dated January28, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were requested to return any unused products by certified mail with signed acknowledgment form, patients should continue to be monitored by their physician field reps during notification visit, and customers should make sure all personnel involved in the management of patients implanted with affected product are aware of the information described in the Field Safety Notice. Customers with questions should contact their local Sorin representative at 1-610-304-9728.For questions regarding this recall call 303-467-6222. |  | 1376 | 1376 Isoline leads | Worldlwide Distribution - USA (Nationwide) and worldwide to Argentina, Australia, Austria, Belgium, Brazil, Canada, Cuba, Czech Republic, Denmark, Egypt, France, Germany, Greece, Hong Kong, India, Iran, Ireland, Israel, Italy, Japan, Kuwait, Lebanon, Morocco, The Netherlands, Portugal, Russia, Saudi Arabia, Slovenia, Spain, Sweden, Switzerland, Tunisia, Ukraine, Uruguay and the United Kingdom. | Terminated | December 09, 2015 | 1045 | 1376 | Computer
Z-0915-2013 | 64440 | No | FLE | Steam sterilizer. | General Hospital | General Hospital | 510(k) | AMSCO C and AMSCO 400 Steam Sterilizers designed for sterilization of heat and moisture-stable materials used in healthcare facilities. |  |  | February 12, 2013 | 2013 | Steris Corporation | The control boards alarm set points and temperature tolerances allowed for erroneous activation of the AD (analog/digital) alarm as temperature fluctuated during operation. When this alarm occurs, the sterilizer locks up and cannot be used until serviced, causing Customer nuisance. | Software design | Steris Corporation initiated their recall on February 12, 2013. They will notify customers via onsite visits by STERIS Field Service Technicians to each affected Customer site.For further questions, please call (440) 392-7601. |  | 395 | 395 Units | Worldwide Distribution - USA including Canada and the states of AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NM, NY, OH, OK, OR, PA, SC, SD, SK, TN, TX, UT, VA, VT, WA, WI, and WV. | Terminated | December 05, 2013 | 296 | 395 | Computer
Z-2266-2013 | 65992 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE 1.5T Signa HDxt, 3.0T Signa HDxt, Signa Vibrant, HD23 Software and the following model numbers: 2395001, 2395001-2, 2395001-3, 2395001-4, 5269997-2, 5269997-4, 2377062-61, 5127452-3, 5127452, 5337894-3, 5337894-4, 5911000-4, 5911000-5, 53085735308573-2. |  |  | April 09, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue due to non-unique series identifiers embedded in image headers of certain GE MR Products.For ClariView and Screen Save from GVTK (Gems Visualization ToolKit) Viewer, the MR Product can generate image headers with non-unique series identifiers, which when exported to some PACS systems, results in the series being associated wit | Software design | Customers were sent two GE Healthcare "Urgent Medical Device Correction" letters one dated April 9, 2013, and the other May 22, 2013, each addressing one of the 2 groups of MR Products involved in the recall. The letters were addressed to Hospital Administrators / Risk managers, Radiology Department Mangers and Radiologists. The letters described the Safety Issue and provided Safety Instructions, Affected Product Details, Product Correction and Contact Information.GE Healthcare will correct all affected systems by providing a software update at no cost to you. A GE Healthcare service representative will contact you to arrange for this correction.Contact your local service representative if you have any questions concerning this notification.For further questions please call (262) 513-4122. |  | 1246 | 1,246 (521 US and 725 OUS) | Worldwide Distribution - All states in continental USA including DC except AK, AZ, RI and WY. and the countries of : AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BELARUS, BELGIUM, BRAZIL, CANADA, CHILE, CHINA, CZECH REPUBLIC, DENMARK, EGYPT, FINLAND, FRANCE, GERMANY, GREECE, HONG KONG, HUNGARY, INDIA, INDONESIA, ISRAEL, ITALY, JAPAN, KOREA, KUWAIT, LEANON, MARINIQUE, MEXICO, NETHERLANDS, NEW ZEALAND, NORWAY, POLAND, QATAR, ROMANIA, RUSSIA, SAUDI ARABIA, SERBIA, SINGAPORE,SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, LUXEMBOURG, PORTUGAL, VIETNAM and URUGUAY. | Terminated | May 05, 2014 | 391 | 1246 | Computer
Z-2271-2013 | 66211 | cart, emergency, cardiopulmonary (excluding equipment) | BZN | Cardiopulmonary emergency cart. | Anesthesiology | Anesthesiology | 510(K) Exempt | Stryker Vision Mounting Arm, for use with other Stryker Endoscopy medical device products for human patients. Model number 0240-095-200, manufactured by Stryker Endoscopy San Jose |  |  | August 30, 2013 | 2013 | Stryker Endoscopy | When positioning the display located on top of the video cart, it is possible for the user to apply excessive force and push the display up and out of the mounting arm. This can result in the display falling upon the user/patient. | Device Design | Stryker Endiscopy sent an Urgent Device Correction letter dated August 30, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to inform individuals within their organization who needed to be aware of this Medical Device Field Correction, contact Stericycle at 1-877-546-0482 or email stryker4645@stericycle.com to schedule a time for Stericycle to perform the correction, or perform the correction themselves by requesting the materials and following the correction instructions included. Customers were also instructed to report all adverse events or product quality problems to Stryker Endoscopy at 1-800-624-4422, press 3 for Customer Service. --UPDATE--On September 17, 2013, the recall was expanded to include additional customers who may have received the screw and washer included with the design change update. Notification was sent to these consignees on September 17, 2013. . |  | 14271 | 14271 units | Worldwide Distribution - USA including AL, AK, AZ, AR, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, WY and internationally to Australia, Canada, Chile, China, Argentina, Guam, Taiwan, India, Japan, Korea, Latin America, Philippines, Mexico, New Zealand, Hong Kong, Singapore, and Switzerland. | Terminated | April 30, 2014 | 243 | 14271 | Computer
Z-2275-2013 | 66051 | No | OUG | Medical device data system. | General Hospital | General Hospital | 510(K) Exempt | TEGRIS system |  |  | August 19, 2013 | 2013 | Maquet Cardiovascular Us Sales, Llc | MAQUET has identified a potential issue that might occur because of the lack of an emergency stop button on a secondary touch screen display on the TEGRIS System. | Device Design | Maquet sent an Urgent Medical Device Recall Letters/Field Action Response Forms dated August 19, 2013, via Federal Express to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were informed that a MAQUET authorized service technician would contact them to arrange an appointment to equip their system with a second emergency stop button that would be located in close proximity to the secondary touch screen. Customers were instructed to complete and return the enclosed Field Action Response Form acknowledging their receipt and understanding of the communiction by emailing a scanned copy to FieldActions@maquet.com or via fax at 973-807-1658. Customers with questions were asked to contact their MAQUET representative or MAQUET Technical Support at 888-627-8383, option 3.For questions regarding this recall call 973-709-7361. |  | 21 | 21 units | Worldwide Distribution - USA including NY and PA and internationally to Australia, Brazil, Canada, Germany, Hong Kong, Japan, Sweden. | Terminated | April 17, 2014 | 241 | 21 | Computer
Z-2276-2013 | 65996 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco RTP System |  |  | July 02, 2013 | 2013 | Elekta, Inc. | Upgrade installation script is designed to overwrite the same pouch files in the Treatment Couch Library with the original default values. | Device Design | Monaco 3.3 upgrades have been temporarily halted and Important Safety Notice (LINMON0007) has been issued to effected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to ensure that couch files are checked for releases 3.20.00 and 3.30.00 to verify that the correct relative electron density values are being used in the dose calculation. If customers have edited the sample couch files, and previously upgraded to Monaco 3.20.00 or 3.30.00, their couch files have already been overwritten. Open the couch files, make the necessary updates to the couch parameters, and use the "Save As Treatment Couch" option to save the modified sample couch files to a new, unique name.Please note that the couch files can also be edited in Focal and then the patients opened in Monaco. Comparable Focal releases would be Focal4.64.00(Monaco 3.10.00), Focal4.70.00(Monaco3.20.00) and Focal4.80.00 (Monaco 3.30.00).For questions regarding this recall call 404-993-5886. |  | 45 | 45 | Nationwide Distribution including AK, CA, CO, CT, FL, IL, IN, MI, MN, MO, NC, NJ, NY, OH, OK, OR, PA, SC, TN, TX, WA and Puerto Rico. | Terminated | May 08, 2014 | 310 | 45 | Computer
Z-2283-2013 | 65922 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | The FL23SE Bed is an AC-powered adjustable hospital bed with four built-in electric DC motors and remote controls that is intended for medical purposes. It can be operated by the patient or the caregiver to adjust the height and surface contour of the bed. The bed includes movable and latchable siderails. The FL23SE Bed is intended for use with patients for procedures, therapy, recovery in healthcare environments, and to transport patients between bays and procedural rooms. |  |  | August 19, 2013 | 2013 | Stryker Medical Division of Stryker Corporation | The nurse call cable (Part Number: QDF23-0571) on the FL23SE Bed is located right under the opening of the bed head box. If an accessory (i.e. IV pole) is inserted into this slot, the circuit directing the nurse call system can be damaged. | Device Design | Stryker sent an Urgent Medical Device Recall Notification letter dated August 19, 2013, to all affected consignees. The notification described the affected product, reason for recall, potential hazards, and recall customer actions and reply instructions. A Stryker Representative will contact the consignees to arrange a time to repair the unit(s) and eliminate the risk of the potential interference . Urgent questions or concerns are directed to Blaine Burnett (269)389-6954. Stryker Medical business hours are Monday-Friday 8 a.m-5 p.m. (ET).For questions regarding this recall call 269-324-6609. |  | 61 | 61 units | Worldwide Distribution - USA including MI, WI, VA and internationally to Canada. | Terminated | January 29, 2014 | 163 | 61 | Computer
Z-0911-2013 | 64013 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | BD Vacutainer PST Gel and Lithium Heparin (LH); 83 Units; Plus Blood Collection Tubes; Sterile; IVD; REF 367962; 4.5 mL, 13 X 100 mm; BD, Franklin Lakes, NJ USA; Made in USA. |  |  | March 05, 2012 | 2012 | Becton Dickinson &amp; Company | BD is conducting a recall of the BD Vacutainer Plus Plastic Plasma Separator Tube (PST) with Light Green/BD Homogard Closure due to a deformed top of the tube reservoir reported in two customer complaints. | Equipment maintenance | BD sent advisory letters via UPS dated 3/5/2012 to customers. The letter identified the affected product and the reason for recall. BD has taken steps to modify the existing manufacturing and assembly processes to help reduce the potential for this condition. Questions regarding this issue should be directed to BD Global Technical Services at 1-800-631-0174 or email at www.bd.com/Vascutainer/contact. |  | 225000000 | 225 million units | Nationwide Distribution | Terminated | July 20, 2015 | 1232 | 225000000 | Computer
Z-2201-2013 | 66156 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, CARESCAPE Monitor B850. |  |  | July 20, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has recently become aware of an issue associated with the CARESCAPE Monitor B850 when used with the iCentral network. The iCentral must have the alarm management software license (L-NETM-00) and the alarm limit change option enabled to experience this issue. When changing the HR, SpO2 or Invasive pressure alarm limits from the iCentral over the network, the alarm priority on th | Software design | Consignees were sent on 7/20/2011 a GE Healthcare "Important Product Information" letter. The letter was addressed to Healthcare Administrator / Risk Manager, chief of Nursing and Director of Biomedical Engineering. The letter described the Issue, Affected Product Details, Instructions, Product Correction, and Contact Information (Technical Support 1-800-558-7044). |  | 636 | 636 (53 US, 583 OUS) | Worldwide Distribution-USA (nationwide) including the states of CA, CO, CT, HI, IL , IA, KY, LA, MA, MI, NE, NY,OH, OK, TX, and WI, and the countries of UNITED KINGDOM, UNITED ARAB EMIRATES, TURKEY, TUNISIA, THAILAND, SYRIAN ARAB REPUBLIC, SWITZERLAND, SWEDEN, SPAIN, SOUTH AFRICA, SLOVAKIA, SINGAPORE, SAUDI ARABIA, RUSSIAN FEDERATION, REPUBLIC OF SERBIA, PORTUGAL, POLAND, PARAGUAY, NORWAY, NETHERLANDS, NEPAL, MEXICO, MALAYSIA, KUWAIT, REPUBLIC OF KOREA, JORDAN, JAPAN, ITALY, ISRAEL, IRELAND, IRAQ, INDONESIA, INDIA, HONG KONG, GREECE, GERMANY, FRANCE, FINLAND, ESTONIA, EGYPT, DENMARK, CZECH REPUBLIC, COSTA RICA, COLOMBIA, CHINA, CHILE, CANADA, BOSNIA AND HERZEGOVINA, BELGIUM, BAHRAIN, AUSTRIA, and AUSTRALIA. | Terminated | October 25, 2013 | 828 | 636 | Computer
Z-0935-2013 | 64291 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Optima MR450w, Discovery MR450, Discovery MR750, Signa HDx, Signa HDxt, Signa HDi, Signa Vibrant, and GE 1.5T Signa HDe MR System. |  |  | May 10, 2011 | 2011 | GE Healthcare, LLC | An error on the SR Viewer Reporting Tool in which edits can be made on a report and can be saved into incorrect reports. | Software design | GE Healthcare sent an Urgent Field Safety Notice dated May 5, 2011, to all affected customers. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information. Customers were instructed to open only one patient report at a time when editing a report in SR Viewer.US Customers did not receive a customer letter, however all US systems have been updated.For questions customers were instructed to contact their local Sales/Service Representative.For questions regarding this recall call 262-513-4122. |  | 4026 | 4026 (1672 US; 2354 OUS) | Worldwide Distribution - USA (All States in continental USA including PR except RI.). Internationall to:ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BRAZIL, BELARUS, BELGIUM, BULGARIA, BOSNIA &amp; HERZEGOVINA, CANADA, CHILE, CHINA, COLOMBIA, CROATIA, CZECH REPUBLIC, DENMARK, EGYPT, ESTONIA, FINLAND, FRANCE, FRENCH POLYNESIA, GEORGIA, GERMANY, GREECE, GUADELOUPE, HONDURAS, HUNGARY, INDIA, INDONESIA, IRAN, ISLAMIC REPUBLIC OF IRAN, IRAQ, ITALY, IRELAND, ISRAEL, JAPAN, JORDAN, KENYA, KOREA, KUWAIT, KAZAKHSTAN, LATVIA, LEBANON, LIBYA, LITHUANIA, LUXEMBOURG, MALAYSIA, MEXICO, MACEDONIA, MARTINIQUE, MOROCCO, NETHERLANDS, NORWAY, PAKISTAN, PERU, PHILIPPINES, REUNION, NEW ZEALAND, OMAN, POLAND, PORTUGAL, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SERBIA, SOUTH AFRICA, SPAIN, SUDAN, SWEDEN, SINGAPORE , SLOVAKIA, SOUTH KOREA, SWITZERLAND, SYRIA, TAIWAN, THAILAND, TURKEY, TUNISIA, UNITED ARAB EMIRATES, UNITED KINGDOM, UZBEKISTAN, VENEZUELA, URUGUAY, VIET NAAM, and YUGOSLAVIA. | Open, Classified |  |  | 4026 | Computer
Z-2207-2013 | 66204 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Osseotite 2 Certain Prevail 5mm(D)X 4.1mm (P) X X 11.5mm (L) Dental Implant |  |  | June 10, 2013 | 2013 | Biomet 3i, LLC | One of the dental implants do not have the internal hex feature. One of the dental implants do not have the internal thread feature. | Device Design | Biomet 3i sent an Urgent: Medical Device Recall letter dated June 10, 2013 to affected customer via email, facsimile, or postal mail. The letter identified the affected product, problem, and actions to be taking. Customer were instructed to check their respective inventory for the affected product and return to Biomet 3i. Biomet 3i will issue replacement product. For questions or concerns call 1-800-342-5454. |  | 215 | 215 | Worldwide Distribution - USA Nationwide in the states of MA, CT, NJ, DE, NY, FL, OK, TX, PA, DC, IL, KS, LA, MD, CA, RI, TN, MN, and SC and the countries of Canada, Japan, Australia, Italy Germany, France, Switzerland, Apain, Brazil, Uruguay, Singapore, Dominican Republic, Taiwan, Austria, Portugal, and the United Kingdom. | Terminated | November 06, 2014 | 514 | 393 | Computer
Z-2199-2013 | 66057 | No | HTZ | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | AEQUALIS Reversed II or AEQUALIS Reversed Fracture Impactor. These products are contained within the Aequalis reversed II Instrument set tray no.YKAD83 and the Aequalis reversed fracture instrument set tray no.YKAD95. The items affected are MWD21, MWD023, MWD024 and MWD025. |  |  | August 13, 2013 | 2013 | Tornier, Inc | Following the identification of the disassociation of an impactor tip during surgery, Tornier is initiating a voluntary recall of all lots of Aequalis Reversed II and Reversed Fracture impactors. Upon review of the event, Tornier determined that a redesign of the instrument was warranted. These products are contained within the Aequalis Revered II instrument set tray no YKAD83 and the Aequalis R | Device Design | Tornier sent an "Urgent Medical Device Recall" letter dated August 12, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customer were advised to cease further use and distribution of the product and to contact all end users. Also requested the completion and return of the attached questionnaire via fax to 952-236-4007, e-mail to CustomerServices@Tonier.com or using the envelope provided.We apologize for the incovenience and thank you in advance for your cooperation in this matter. If you have any questions please contact our Customer Service Department at 1-888 494-7950. |  | 2298 | 2,298 (1045 US, 1253 OUS) | .Worldwide Distribution - USA including Puerto Rico and the states of AL, AZ, CA, CO,CT, FL, GA, IL, KY, MD, MA, MI, MN, NE, NJ, NC, OH, OR, PA and TX., and the countries of CANADA, MEXICO, HONG KONG, SOUTH AFRICA, ARGENTINA, COLOMBIA, ISRAEL, AUSTRALIA, BELGIUM, FINLAND,FRANCE, GERMANY, ITALY, NETHERLANDS, SLOVENIA, SPAIN, SWITZERLAND and UNITED KINGDOM | Terminated | March 03, 2015 | 567 | 2298 | Computer
Z-2188-2013 | 64152 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Outlook ES pump is intended for use with B. Braun Medical Inc. Horizon Pump IV Sets. Used to regulate the flow of primary and secondary fluids when positive pressure is required. The infusion system is capable of delivering fluid from a negative head height (when the IV fluid container is lower than the pump), and provides clinically accepted volumetric accuracy for all standard IV fluids, including blood, lipid, and Total Parenteral Nutrition (TPN). Model number: 621-300ES and 621-400ES. To deliver a broad range of drugs with use with the B. Braun Medical Inc. Horizon Pump IV Sets. |  |  | November 14, 2012 | 2012 | B Braun Medical, Inc. | This notice is being provided in follow up to the field correction initiated by B. Braun Medical Inc. on August 26, 2011 due to an issue in which the Outlook ES Safety Infusion System may halt infusion but the "RUN" Light Emitting Diodes (LEDs) on the front display continue to advance as if the pump were infusing. The pump emits a backup alarm, but there are no visual indicators that the infusion | Software design | The firm, B Braun Medical Inc, sent an "Urgent Expanded Medical Device Field Correction" letter dated November 14, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed that pumps that have been upgraded (software version 151599 or higher), no additional action is required on their part at this time. For the pumps that have not yet been upgraded, within three weeks of the letter, customers should confirm the serial numbers using the enclosed response form or contact B. Braun Medical, Inc. Infusion Systems Customers Support at 1-800-627-7867 (800-627-PUMP) to schedule their upgrade.For questions regarding this recall call 972-245-2243. |  | 6919 | 6,919 | Nationwide distribution: USA including states of: AR, CA, FL, IL, LA, NC, NJ, NY, MD, ME, MN, SC, TN, TX, VA, WI, and WV. | Terminated | March 06, 2015 | 842 | 6919 | Computer
Z-2194-2013 | 66173 | No | OAT | Endosseous dental implant. | Dental | Dental | 510(k) | Brand Name: Vector TAS 8mm Screw. |  |  | July 12, 2013 | 2013 | Ormco/Sybronendo | Ormco Corporation is voluntarily recalling one lot of 8 MM Vector TAS Screws because some screws of this lot were machined undersized at the top of the threads may potentially cause a weakening of the screw. Use of the undersized product may result in breakage of the screw during insertion or removal when torqued, possibly requiring surgical removal of any remaining pieces. | Device Design | Ormco sent an "Urgent Medical Device Recall letter dated July 12, 2013, via USPS 1st class mail, along with a return form to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were informed of the recall, with product description, product codes, ship dates, contact information and instructions to return any affected product identified in stock, Additionally, doctors are instructed to proceedwith treatment if the screw has already been placed and to take special care not to over-torque the screw upon removal. Periodic effectiveness checks will be conducted in order to follow up with non-responders.Please fax this acknowledgement form to the the following number to confirm receipt of this notification whether or not you have any affected products. (714) 288-4603 |  | 4786 | 4,786 | Worldwide Distribution - USA (nationwide) including the countries of Canada, Australia, New Zealand, EU and ROW consignees. | Terminated | September 11, 2014 | 426 | 4786 | Computer
Z-2195-2013 | 66110 | No | DYJ | Orthodontic appliance and accessories. | Dental | Dental | 510(K) Exempt | AOA Mini RPE Screw Assembly , a component of |  |  | June 06, 2013 | 2013 | Sybron Dental Specialties | The firm initiated the recall of AOA Mini RPE Screw Assembly, because it may be missing a complete weld joining one side of the arm with the screw body. Herbst, MARA and Fixed Metal devices may be affected, because they contain an AOA Mini RPE Screw Assembly. | Device Design | AOA sent a recall notification letter titled "Urgent-Medical Device Correction", on June 6, 2013 via USPS 1st class mail. The letter identified the affected product, reason for recall, images, and actions to be taken. Customers were instructed to complete and return the attached Acknowledgement Form to AOA. For questions contact the AOA Customer Support Department at 1-800-262-2551. |  | 1091 | 1,091 | Worldwide Distribution - USA Nationwide and the countries of Australia, Canada, and EU. | Terminated | September 13, 2013 | 99 | 1091 | Computer
Z-1993-2013 | 66046 | No | MNO |  |  | Cardiovascular | PMA | Cardiogenesis Corporation - Cardiogenesis Transmyocardial Laser. |  |  | May 01, 2011 | 2011 | Cardiogenesis Corporation | Following acquisition of Cardiogenesis, we began receiving complaints of sparks, burns, or smoke incidents related to the handpieces. | Device Design | Cardiogenesis is responsible for the correction of all noncompliant products. CDRH understands that the corrective action plan (CAP), involving 2,437 Sologrip III handpieces and 61 PEARL 5.0 handpieces, consists of the following:1. Notification to purchasers of the handpiece removal and replacement procedures,2. Replace existing hospital inventory of handpieces with a redesigned handpieceCDRH approves the CAP subject to the following conditions:1. Notification of all distributors and purchasers will be successfully implemented, and2. Corrections will be made at no cost to the purchasers by October 1, 2013.For further questions please call (949) 420-1800. |  | 2498 | 2498 | US Distribution | Terminated | September 15, 2014 | 1233 | 2498 | Computer
Z-2191-2013 | 66133 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Vision One Laser System Model GA-0025020, Serial No: 10050, 10051, 10053. |  |  | August 19, 2013 | 2013 | Lumenis, Inc. | Lumenis has initiated a recall on certain models of Vision One System due to a potential for unintended laser exposure to the user. | Software design | Customer's were notified via letter on 8/19/13. Service visits to consignees were scheduled with anticipated completion on 8/26/13 to replace the control board. Consignee monitoring was performed by use of return reply verification tracking cards returned by the Lumenis service engineer after completion of CPU board replacement. |  | 3 | 3 units | Distributed in the states of NJ, PA, and IL. | Terminated | December 19, 2013 | 122 | 3 | Computer
Z-2173-2013 | 66036 | No | MAQ |  |  | Microbiology | PMA | KIT cobas 4800 HPV AMP/DET 240T / 960T US-IVD, CE-IVD |  |  | July 25, 2013 | 2013 | Roche Molecular Systems, Inc. | During an internal investigation, a formulation error that occurred was identified during the manufacture of the cobasÂż 4800 HPV Master Mix reagent that was packaged in four (4) cobasÂż 4800 AMP/DET kits, US- and CE-IVD (240 and 960 test kit configurations). Examination of the batch production record of the material revealed there was an error in the tare weight used for the carboy during the fina | Device Design | Roche Molecular Diagnostics sent an Urgent Medical Device Correction letter dated July 25, 2013, to all affected customers via UPS return receipt. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to:1. Immediately discontinue use of the product lots listed above.2. Discard affected product according to local guidelines.3. To receive replacement inventory, complete the entire attached faxback form and fax it to 1-866-918-8741.4. The faxback form must be completed and returned even if you do not have affected inventory.5. File this Urgent Medical Device Correction (UMDC) for future reference.Please contact Roche Molecular Diagnostics Technical Support, 24 hours a day, seven days a week at 1-800-526-1247, if you have questions about the information contained in this UMDC. |  | 791 | US-IVD: 350 Kits; CE-IVD: 441 KITS | Worldwide Distribution - USA (nationwide) including the countries of Austria, Australia, Belgium, Brazil, Czech Rep, Canada, Croatia, Columbia, Denmark, Dominican Republic, Germany, Ecuador, France, Japan, Greenland, Mexico, Italy, New Zealand, Netherlands, Peru, Norway, Singapore, Romania, South Africa, Spain, South Korea, Sweden, Switzerland and UK. | Terminated | August 06, 2015 | 742 | 791 | Computer
Z-2177-2013 | 64601 | No | FSO | Surgical microscope and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Surgical microscope |  |  | December 18, 2012 | 2012 | Leica Microsystems (Schweiz) Ag | The brake in the Z axis and/or brakes in the AB and C axes failed during installation. | Device Design | The firm, Leica Microsystems, sent an "URGENT MEDICAL DEVICE RECALL" letter dated December 19, 2012 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to DISCONTINUE USE until your system software is updated and to complete and return the attached Acknowledgement Form via email at LMGRA@leica-microsystems.com or fax to: 847-236-3747 or +41 71 726 32 49 within 10 working days after receipt of field safety notice. Your Leica Mycrosystems will contact you to make the necessary arrangements to complete an upgrade at no charge to the customers. If you have any further questions, contact Regulatory Affairs/Quality Assurance Manager at 847-405-5413 or +41 71 726 32 16. |  | 28 | 28 units | Worldwide distribution: USA (Nationwide) including states of: AK, CA, CO, NY, and OR; and countries of: Belarus, China, Hong Kong, India, Indonesia, Italy, Norway, South Africa, Spain, Sweden, Turkey, and Vietnam. | Open, Classified |  |  | 33 | Computer
Z-0948-2013 | 64376 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Siemens Somatom CTs equipped with software syngo CT2010A, syngo CT2010B, syngo CT2011A and CT2012B. |  |  | November 27, 2012 | 2012 | Siemens Medical Solutions USA, Inc | Sporadic hang-up malfunctions concerning 3D reconstructions during activated "Preview Image" functionality. | Software design | Siemens sent a "CUSTOMER SAFETY ADVISORY NOTICE" dated November 27, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. The correction for this problem will be available with Service Pack 1 for syngo CT2011A And Service Pack 4 for syngo CT2012B. Both Service pack are planned to be released starting February 2013. Systems still running on syngo CT2010A or syngo CT2010B are already scheduled to be upgraded to syngo CT2011A or syngo CT2012B respectively. Contact your local Siemens Application Specialist for additional information. |  | 896 | 896 | Nationwide Distribution. | Terminated | January 31, 2014 | 430 | 896 | Computer
Z-0950-2013 | 64436 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance 64 and Ingenuity Computed Tomography X-Ray Systems |  |  | February 15, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | This field change order is being released to update software and customer release notes to the affected installed base. | Software design | Philips Healthcare sent an Urgent Medical Device Correction letter dated February 13, 2013, via certified mail to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to review the information contained in the letter with all members of their staff who need to be aware of the contents of the letter. Customers were instructed to retain a copy with the equipment Instruction for Use. For questions customers were instructed to contact their local Phillips representative or local Philips Healthcare office. For North Americal and Canada contact Customer Care Solutions Center at 1-800-722-9377, option 5: Enter Site ID or follow the promptsFor questions regarding this recall call 1-800-722-9377. |  | 42 | 42 Units | Worldwide Distribution - USA including AR, AZ, IN, KS, LA, MA, MN, MS, MT, NJ, NY, NY, OH, OR, PA, TN, TX, VT, WV, WY, and worldwide including the contries of Australia, Belgium, China, France, Germany, Israel, Japan, Norway, Sweeden, and Switzerland. | Terminated | April 11, 2014 | 420 | 42 | Computer
Z-0951-2013 | 64320 | No | GJT | Factor deficiency test. | Hematology | Hematology | 510(k) | Vial Label: |  |  | January 28, 2013 | 2013 | Diagnostica Stago, Inc. | Diagnostica Stago has confirmed that sometimes the delta OD measured can be changed leading to a lowered percentage result. | Software design (manufacturing process) | Diagnostica Stago Inc. sent e-mail notifications on 1/28/2013 to all consignees (laboratory supervisor). The written field safety notice, entitled Important Field Safety Notice, with its response form was attached. The letter identified the affected product and gave a description of the defect. Customers were asked to confirm all the results obtained in the abnormal range by either of the two methods provided. Users are to complete and return the enclosed response form as soon as possible to the Stago Hotline via fax, 973-644-9348. Questions should be directed to the Stago Hotline at 800-725-0607. |  | 2901 | 2901 kits | Worldwide Distribution - USA and Canada. | Terminated | July 19, 2013 | 172 | 2901 | Computer
Z-2164-2013 | 64146 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Outlook ES pump is intended for use with B. Braun Medical Inc. Horizon Pump IV Sets. Used to regulate the flow of primary and secondary fluids when positive pressure is required. The infusion system is capable of delivering |  |  | October 31, 2012 | 2012 | B Braun Medical, Inc. | This notice is being provided in follow up to the field correction initiated by B. Braun Medical Inc. on 26 August, 2011 due to an issue in the Outlook ES Safety Infusion System where the Care Area can inadvertently be exited thereby removing DoseGuard dose limit protection. This occurs when a sequence of HOLD - HOLD key strokes are performed on the pump key panel while the user is in the Dose Dat | Software design | The firm decided to conduct a field correction to correct the issue and carried it out through routine field services. The firm was notified by FDA that this is a reportable event, so they notified their consignees via letters dated 10/31/2012. |  | 1970 | 1,970 pumps added for this expansion | Nationwide Distribution. | Terminated | February 03, 2016 | 1190 | 1970 | Computer
Z-2157-2013 | 66048 | No | IXI | Radiation therapy beam-shaping block. | Radiology | Radiology | 510(k) | DMLC IV-ERGO - Beam shaping block for Radiation Therapy. |  |  | August 15, 2013 | 2013 | Elekta, Inc. | Clinical mistreatment may occur when there are errors in the clinical workflow, or settings are incorrect. | Device Design | Important Field Safety Notices 200 01 406 030 (Changes to your system to improve the safe operation of DMLC) and 200 01 602 031 (Changes to your system to improve the safe operation of 3DLINE DMLC) letters were sent to customers on 8/15/2013 advising them of the issue and recommendations to resolve either by upgrading their system or to discontinue a clinical use of the device. |  | 31 | 31 | Nationwide distribution: AL, CT, FL, GA, IA, MS, MO, NJ, NY, ND, OH, OK, PA, RI, TX, VA and Puerto Rico. | Terminated | July 28, 2015 | 712 | 31 | Computer
Z-0953-2013 | 64465 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Optetrak Logic Comprehensive Knee Prosthesis System, Proximal Tibial Spacer and screws, size 3.5, 8mm, Catalog # 02-012-42-3508. |  |  | February 01, 2013 | 2013 | Exactech, Inc. | A dimensional mismatch was identified which can potentially allow for interference between devices. | Device Design | Exactech Inc. sent a *** Important *** Product Market Recall Notice dated January 30, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.1. Customers were instructed to immediately cease distribution or use of these products. 2. Extend this information to your accounts that may have this product in their possession.3. Verify if you have any of the subject Optetrak Logic Proxima Tibial Spacer (PTS) in the specified lots.Please complete the attached fax back form and return it to Exactech within the next 5 working days. Thank you for your prompt attention to this matter. For further questions please call (352) 377-1140. |  | 123 | 123 devices | Nationwide Distribution including the states of FL, AL, LA, PA, VA, IL, OH, NY, NC, KS, TN, NJ, CO, GA, CA, MD, MN, IN, &amp; Puerto Rico. | Terminated | September 22, 2014 | 598 | 123 | Computer
Z-2282-2013 | 64236 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Synthes Oracle Spacer System Slap Hammer, Part Number 03.809.690, Lot Number 3723847 |  |  | January 11, 2013 | 2013 | Synthes USA HQ, Inc. | This recall is being initiated in response to complaints received where it was reported that the proximal end of the Slap Hammer broke during use. | Device Design | SYNTHES sent an Urgent Notice Medical Device Recall letter dated January 11, 2013, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. The letter requested consignees examine their inventory, remove the recalled product from use and return them to Synthes. Customers were also instructed to completed the attached Verification Section at the end of the letter indicating whether or not they had any of the affected product and the number of devices found. Customers with questions were instructed to call 610-719-5450 or email Fieldaction@synthes.com.For questions regarding this recall call 610-719-5000. |  | 27 | 27 | Worldwide Distribution - USA including CO, PA, TN, VA, TN, WA, and internationally to Czech Republic. | Terminated | August 31, 2015 | 962 | 27 | Computer
Z-0913-2013 | 64000 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Beckman Coulter AU5800 Clinical Chemistry Analyzer. An automated chemistry analyzer that measures analytes such as Glucose, Magnesium, and Potassium. |  |  | November 28, 2012 | 2012 | Beckman Coulter Inc. | Due to a software data storage limitation which only occurs under certain conditions, measurements of the Reagent Blank (RB) or Calibration (CAL) sample results are not saved and the software reverts to a previously stored, acceptable RB/CAL result. | Software design | Beckman Coulter sent an "URGENT PRODUCT CORRECTION" letter dated November 28, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the Call Center Hotline at (800) 854-3633 for questions regarding the notice. |  | 350 | 28 in U.S.; 4 Canada; 318 to other foreign countries. | Worldwide Distribution-USA (nationwide) and the countries of Argentina, Australia, Austria, Brazil, China, Czech Republic, France, Germany, Hungary, India, Ireland, Israel, Italy, Korea Republic, Malaysia, Netherlands, Russian Federation, Spain, Switzerland, Taiwan, Turkey, United Kingdom, and Viet Nam. | Terminated | February 24, 2014 | 453 | 350 | Computer
Z-1849-2013 | 65759 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | GE Healthcare, LOGIQ e Ultrasound System. |  |  | November 09, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare became aware of a potential issue associated with the LOGIQ e system. Image data from one patient has the potential to be linked to the patient ID of another patient. | Software design | Starting on November 9, 2012, GE Healthcare, will plan to update the software on all of the affected GE LOGIQ e Ultrasound systems that were distributed at no cost to the customer.As an effectiveness check, we will track all service requests for the field action to completion. These service requests will record the repair of the affected units. Exceptions to completion of the service request will be recorded as well. Records of field action completion will be retained per our processes and procedures.No Customer letter was sent.Further questions please call (262) 513-4122. |  | 588 | 588 | Worldwide Distribution - USA (nationwide) including the states of AL, AK, AZ, AR, CA, CO, DE, FL, GA, ID, IL , IA, MD, MA, MI, MS, MO, NJ, NM, NY, NC, OH, PA, RI, SC, TN,TX, UT,VT, VA and WI., and the countries of ARAB EMIRATES, AUSTRALIA, AUSTRIA, BELGIUM, DENMARK, FINLAND, FRANCE, GERMANY, GREAT BRITAIN, HONG KONG, INDIA, ITALY, JAPAN, NETHERLANDS,NEW ZEALAND, NORWAY, POLAND, RUSSIA, SAUDI ARABIA, SINGAPORE,SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND and UNITED EMIRATES. | Terminated | September 15, 2014 | 675 | 588 | Computer
Z-0880-2013 | 63822 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Philips Medical Systems, Brilliance Computed Tomography (CT) X-Ray Scanners, Big Bore, Model #728243; and Brilliance Computed Tomography (CT) X-Ray Scanners, Big Bore Oncology, Model #728244. |  |  | October 24, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Software anomaly. Philips determined that the TumorLOC software program for the referenced PET/CT X-Ray scanning systems has the capability to generate Intensity Projection data sets from respiratory gated data. A circumstance was identified by the firm whereby these generated datasets are 'flipped' and incorrectly labeled left-to-right when the are saved to disk. If the flipped Intensity Project | Software design | Philips Medical issued an URGENT - Medical Device Correction letter to their customers concerning the Tumor LOC IPD issue on October 25, 2012. The letter informs the customer of the nature of this software problem and names the Brilliance and GEMINI scanning systems affected by this issue. The letter goes on to describe the hazard inherent to this software issue and provides 'Workaround procedures which the customers can use in order to help mitigate the effects of this problem. Customers with questions should contact their Customers Care Solutions Center at 1-800-722-9377, option 5.For questions regarding this recall call 1-800-722-9377. |  | 765 | 765 units | Worldwide Distribution - USA including AK, AK, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, IL, IN, KS, KY, OA MA, MD MI, MN, MO, MS, NC, D, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WV; Internationally to Canada and other countries.. | Terminated | July 23, 2015 | 1002 | 871 | Computer
Z-0013-2014 | 66296 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance CT Big Bore Oncology, Brilliance CT Big Bore Radiology and Brilliance CT 16 (Air) |  |  | August 30, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | There are artifacts appearing on the Tracker images during threshold-triggered bolus scans with protocols using a Rotation Time of 0.4 seconds. | Software design | Philips sent an "Urgent - Medical Device Correction" letter dated August 30, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. This Field Safety Notice is intended to inform you about:" what the problem is and under what circumstances it can occur" the actions that should be taken by the customer / user in order to prevent risks for patients or users" the actions planned by Philips to correct the problem If you need any further information or Support concerning this issue, please contact your local Philips representative or local Philips Healthcare office.For North America and Canada, contact the Customer Care Solutions Center (1 -800-722-9377, Option 5: Enter Site ID or follow the prompts).This notice has been reported to the appropriate Regulatory Agency.Philips apologizes for any inconveniences caused by this problem |  | 258 | 258 | Worldwide Distribution - USA (nationwide) including Puerto Rico and the states of AR, AZ, CA, CO, CT, GA, HI, IL, IN, KS, MA, MD, MI, MO, MS, NC, NH, NJ, NV, NY, OH, OK, OR,TN, TX, UT, VA, VT and WV., and the countries of Australia, Austria, Bangladesh, Belgium, Brazil, Chile, China, Denmark, Egypt, France, Germany, Honduras, India, Indonesia, Italy, Kazakhstan, Latvia, Lebanon, Libyan Arab Jamahiriya, Mexico, Netherlands, Norway, Palestine, Philippines, Poland, Russia, Saudi Arabia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, Ukraine, United Kingdom and Vietnam. | Terminated | May 14, 2014 | 257 | 258 | Computer
Z-0704-2013 | 64134 | No | JAC | Radiographic film marking system. | Radiology | Radiology | 510(K) Exempt | GE Healthcare Definium 5000/5220493-2 X-Ray System |  |  | September 26, 2012 | 2012 | GE Healthcare, LLC | When quickly switching the protocols from AEC to Fixed under the same view, if the user selection process time is less than 0.5 seconds, the User Interface (UI) may display different exposure techniques than selected, and may not display AEC setting. The exposure can still be made as commanded, however it may not be properly indicated. | Software design | GE Healthcare will remedy the issue or bring the product into compliance with each applicable Federal standard. A GE Healthcare Service Representative will update the software on the system to address the issue.Call Center phone number: 800-437-1171 (United States). For other countries, please contact your local GE Healthcare Service Representative if you have any questions or concerns regarding this notification.Please be assured that maintaining a high level of safety and quality is our highest priority. If you have any questions, please contact us immediately per the contact information above. |  | 312 | 312 installed in US | Nationwide Distribution | Terminated | December 05, 2013 | 435 | 312 | Computer
Z-0034-2014 | 64787 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Centricity Cardio Imaging 5.0 is a software only Picture Archiving and Communication System (PACS). |  |  | October 02, 2012 | 2012 | GE Healthcare It | GE Healthcare is aware of a potential safety risk where an image labeling situation exists in the Image Viewer software of the Centricity Cardio Imaging product, version 5.0 SP1. Multiple studies may be displayed side by side by selecting "Compare Study" from the Study List, or by opening prior studies in the Patient Folder. In these cases, the Viewer differentiates the studies wityh a label (Curr | Software design | GE Healthcare IT mailed an IMPORTANT PRODUCT INFORMATION letter dated October 2, 2012 to the only consignee of this product. The letter notified the customer of the issue, provided temporary instructions on how to use safely use the software until a new software version can be installed, and advises that a software correction will be made available in the near future. Customers who have any questions can contact their local GEHC-IT service representative or call GE Healthcare's Remote Online Center (ROC) Customer Care Center at 1-800-437-1171. |  | 1 | 1 unit | USA Nationwide Distribution in the state of OK | Open, Classified |  |  | 1 | Computer
Z-0033-2014 | 66312 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Philips MR Systems: Asset 0.5T, Apollo 0.5T, Infinion 1.5T, Eclipse/Polaris 1.5T, Panorama 0.6T |  |  | August 07, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | There is a potential for water to collect in the vent pipe elbow in the magnet venting system. Water may freeze, blocking the venting system. | Device Design | Philips Healthcare sent an Urgent Medical Device Correction letter via certified mail on August 8, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the information with all of their staff who need to be aware of the contents of this communication. Philips Healthcare will perform any necessary corrective actions prior to the customer's next cryogen refill. Customers with questions were instructed to contact their Philips field service representative: Technical Support Line 1-800-722-9377.For questions regarding this recall call 1-800-722-9377. |  | 17885 | Total 511 (Asset 0.5T,: 35 units Apollo 0.5T: 15 units, Infinion 1.5T: 116 units, Eclipse/Polaris 1.5T: 290 units, Panorama 0.6T: 55 units) | Worldwide Distribution - US (nationwide) including FL, CO, KY, MS, PA, NC, NJ, TX and Puerto Rico and internationally to Italy. | Terminated | March 26, 2014 | 231 | 17885 | Computer
Z-0035-2014 | 66424 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX 5.2 DB Server (running Oracle 10.1.0.4.0), medical imaging Picture Archiving and Communication System (PACS). |  |  | September 19, 2013 | 2013 | AGFA Corp. | System downtime and/or slow performance may result if software in use is an old version of IMPAX and is not at current supported levels. | Software in the Use Environment | An "URGENT FIELD SAFETY NOTICE" letter was sent via email to the consignees on 9/19/2013, Acknowledgement, via FAX-Back or email, that the information was received and understood has been requested from the consignees. The letter specifically describes the safety alert for the device, informs consignees of the problem and recommends the upgrades for IMPAX and Oracle. |  | 4 | 4 | Worldwide distrbution: US(nationwide) distribution including NY and OK, and country of Canada. | Terminated | October 24, 2013 | 35 | 4 | Computer
Z-0032-2014 | 66293 | No | LOM |  |  | Microbiology | PMA | FSE Image Upgrade Kit, Catalog # 92019C |  |  | August 20, 2013 | 2013 | Bio-Rad Laboratories | The FSE Image Upgrade kit (part 92019C) is recalled due to issue related to the export of data from the EVOLIS analyzer to Laboratory Information Systems (LIS) when using the MONOLISA Anti-HBs EIAQuantitative Determination in EVOLIS APF version 6.0. | Software change control | BioRad has sent an Urgent Product Correction letter dated August 19, 2013, to customers who had the EVOLIS FSE Upgrade Image Kit installed on the EVOLIS at their site by a Bio-Rad Laboratories Field Service Engineer (FSE).This issue impacts laboratory only if you are running the MONOLISA Anti-HBs EIA Quantitative Determination method with EVOLIS APF version 6.0 and TimeLiner Data Management Software. If you use this Quantitative Determination method with EVOLIS APF version 6.0, please contact Bio-Rad Laboratories for additional information or assistance.Bio-Rad Laboratories is currently developing and qualifying a software update CD consisting of the EVOLIS APF V6.0b and Timeliner Data Management Software to correct the issue. The software update CD will be installed by Bio-Rad Field Service Engineers at each of the customer sites wheninstalling the EVOLIS APF V6.0b.Customers should contact Bio-Rad Laboratories at 1-800-224-6723, Option2 and then Option 3 for any questions about this product correction. |  | 23 | 23 kits | US Distribution including the states of FL, IA, ME, MO, MS, NE, NY, TN, TX, and WA | Terminated | November 13, 2013 | 85 | 23 | Computer
Z-0050-2014 | 66361 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Solea 1.0 Laser Surgical instrument for use in general and plastic surgery and dermatology. |  |  | September 12, 2013 | 2013 | Convergent Dental | Specific error condition could lead to unintended laser emission during laser activation. | Software design | The following Notice of Correction letter was sent out to all customers via email on 9/12/13, and followed up with a direct phone call to confirm receipt and review the contents of the letter: Dear Dr. [Customer name],This email is to inform you of a pending notice of correction regarding your Solea system, serial #[Customer serial number].We have identified a specific error condition that could lead to unintended laser emission during laser activation. The software defect that caused the error condition has been identified and tested.Please do not use your Solea system until further notice. Resolution of this issue will require the installation of an updated release of the Solea software. A representative from Convergent Dental will be in touch with you to schedule the upgrade.Convergent Dental is committed to product quality and patient safety. If you have questions about this issue, please contact us. Software fix and testing:New software version (Release 2.0.4.0) was created by the software development group. Convergent Dental has implemented two software changes to eliminate this problem. The first is a defect fix to correct the handling of fatal errors, specifically laser communication timeout errors, so that these errors are not able to be cleared. This will prevent the user from being able to re-enter the workflow after a fatal error. The second is a defect fix to not set the laser state to disconnected when a laser timeout error condition occurs. This will ensure that the command to stop laser emission is always sent to the laser control module anytime the laser cutting state is exited due to error condition or the normal workflow event of stepping off the foot pedal.The software development team ran full Unit and Integration Tests on the new software build. The software verification activities included execution of the full verification test plan (Doc #790-00001), which has been revised to include specific defect |  | 5 | 5 | US distribution including states of: MA, CT, NY. | Terminated | November 15, 2013 | 64 | 5 | Computer
Z-2301-2013 | 66195 | No | MUE | Full-field digital mammography system. | Radiology | Radiology | 510(k) | Selenia Full Field Digital Mammography System with software versions 5.0.x and 5.1.x |  |  | September 06, 2013 | 2013 | Hologic, Inc. | Software: When the Selenia upon meeting the hard drive threshold of 4,000 images, may delete the most recently dated images | Software design | Hologic sent an Urgent Medical Device Recall letter dated September 5, 2013, by Certified Mail to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were advised how to avoid loss of images until Hologic corrects the problem. A software upgrade to customers who are connected to Hologic Connect who have already been connected remotely and customers who have already been connected via a Service Call will receive notification of the problem and the date the system was remotely corrected notified by Certified Mail.A Dealer notification issued by certified mail to inform Dealers of the need to notify their customers of the problem and to correct the identified Selenia system using Hologic Technical Bulletin TB-00072.Selenia software version 5.2 will implement a double safeguard to both guarantee a 28 day retention period, as well as ensure images are removed in a least recently used order.If there are any questions regarding this field correction, please call Hologic Customer Support at 877 371-4372. |  | 62 | 62 systems | Worldwide Distribution - USA including the states of AL, AR, CA, CO, FL, IL, KY, LA, NE, NC, NJ, NY, OH, PA, TN, TX, UT, WA and WV., and the countries of Canada, Denmark, Iran, Italy, Mexico and Taiwan. | Terminated | December 13, 2013 | 98 | 62 | Computer
Z-0052-2014 | 65812 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Acuson SC2000 volume imaging ultrasound systems at software version 3.0. |  |  | July 10, 2013 | 2013 | Siemens Medical Solutions USA, Inc. | Software issue results in inaccurate 3D image may result in ablation in wrong location. | Device Design | Siemens sent an Customer Safety Advisory Notification letter to all affected customers on July 10, 2013.The letter identified the product the problem and the action needed to be taken by the customerLetters provide steps to mitigate the risk and inform users that a software fix will be installed as soon as it is available. Siemens Service Engineers will contact all locations to arrange installation once fix is available. All personnel who use the ACUSON should be notified of the issue.Update September 17, 2013 : After discussion with CDRH, Siemens revised their letter to state that procedures using the ACUSON SC2000 system with BioSense Webster CARTO 3 equipped with CARTOSOUND Module and SoundSTAR Catheter should not be performed until the users system has been updated with the software upgrade that resolves the issue.For further questions please call (650) 694-5398. |  | 26 | 26 units | Worldwide Distribution - US (nationwide) including the states of IL, PA, AZ, OR, OH, NY, CO, TX and CA., and the countries of France, Germany, Japan, Portugal, Saudi Arabia, Taiwan and Turkey. | Terminated | November 25, 2013 | 138 | 26 | Computer
Z-0884-2013 | 63674 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips DigitalDiagnost X-Ray System with Eleva Software 2.1.3 |  |  | October 25, 2012 | 2012 | Philips Healthcare Inc. | When the operator for a wall stand view selects an "image rotation" different from default, or such image rotation is preset in the examination database, then the resulting image may be rotated in the wrong direction. When the operator manually rotates the image back, electronic side markers (if programmed) are rotated also and as a result may be placed in a wrong position inside the image. | Software design | Philips Medical Systems issued on October 25, 2012 an Urgent Field Safety Notice for the Digital Diagnost System informing consignees of the software upgrade. The UFSN described the problem, the hazard involved, the action to be taken by the consumer/user and the action planned by Philips. For questions contact your local Philips representative: Technical Support Line 1-800-722-9377. |  | 551 | 551 | Worldwide Distribution - USA Nationwide and the Foreign countries of: Austria, Canada, Belgium, China,Denmark, France, germany, Italy, Netherlands,New Zealand, Norway, Philippines, Portugual, SOuth Korea, Spain, Switzerland, Thailand, and UK. | Terminated | March 31, 2014 | 522 | 551 | Computer
Z-0066-2014 | 66341 | No | OXO | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Brivo OEC715/785/865 C-Arm, Models: 715 Prime, 785 Essential, and 865 Plus. |  |  | September 18, 2013 | 2013 | GE Healthcare Beijing , Hangwei Medical Systems, Co. LTD | GE is recalling certain lots of GE Brivo OEC715/785/865 C-Arm due to the potential for the wires in the high voltage cable bundle to fail, resulting in loss of system functionality. | Device Design | GE Healthcare sent an Urgent Medical Device Correction letter dated September 18,2013, to all affected customers to inform them of the issue and provide instructions for immediate mitigation of the issue. Customers were instructed to ensure that all potential users in their facility were made aware of this safety notification and the recommended actions. Customers with questions were instructed to call 1-800-874-7378 option 8.For questions regarding this recall call 801-295-4952. |  | 29 | 29 | Worldwide Distribution - USA (nationwide) and Internationally to China, India, Japan, Malaysia, Australia, Ireland, France, Germany, Spain, Hungary, Belgium, and Switzerland. | Terminated | January 27, 2014 | 131 | 29 | Computer
Z-0082-2014 | 66336 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | MAQUET Getinge Group CARDIOHELP Base Unit REF 70104.7999 REF 70104.8012 Manufactured for: MAQUET Cardiopulmonary AG Made in Germany. |  |  | September 12, 2013 | 2013 | Maquet Medical Systems USA | It has come to the attention of MAQUET that in certain rare instances, the CARDIOHELP-i device may display an error message with audible alarm, indicating "battery defective." The alarm will occur-upon start up or when the main power supply is unplugged from the main power source. The human machine interface touch screen of the CARDIOHELP-i device may gradually become black due to the human machine interface (HMI) back light shutting down. | Software Manufacturing/Software Deployment | MAQUET issued their customers "Urgent-Medical Device Field Correction" notices/Customer Fax Response forms dated September 16, 2013 via FedEx on September 16, 2013. The notification informs the customer of the issue with the product; how to identified affected product; and the action needed to be taken by the customer. The field correction includes service on the affected product. Authorized MAQUET Service Technicians will service the affected product to correct the issue. Customers were requested to complete and fax back (or email a scanned copy) of the enclosed response form. Customers could contact their local field representative or the MAQUET Customer Service line at 1.888.627.8383 (press option 2 followed by option 2) Monday through Friday, 8am to 6pm EST.For questions regarding this recall call 973-709-7000. |  | 63 | 63 | Nationwide Distribution including AZ, FL, CO, TX, OH, VA, MN, WI, SC, IL, ND, UT, CA, OR, NY, IA, CT, SD, MI, HI, AL, MD, KS, MO | Terminated | June 07, 2017 | 1364 | 63 | Computer
Z-0067-2014 | 66205 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Ingenuity TF PET/CT |  |  | August 16, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | Issues with the Ingenuity TF PET/CT Software Version 4.0 could pose a risk for patients in the event that a patient rescan were to require reinjection of PET radiopharmaceutical. | Software design | Philips Healthcare sent an Field Safety Notice dated August 16. 3013, via certified mail to all US customers. Market Representative will distribute the letter internationally. The letter identified the product the problem and the action needed to be taken by the customer.Field Safety Notice 88200460 is intended to inform you about;" what the problem is and under what circumstances it can occur" the actions that should be taken by the customer / user in order to prevent risks for patients or users" the actions planned by Philips to correct the problemAdditionally, a Field Service Engineer will contact each site to schedule a time to implement the Field Change Order.If you need any further information or support concerning these issues, please contact your local Philips representative or local Philips Healthcare office.For North America and Canada contact the Customer Care Solutions Center (1-800.722.9377, option 5: Enter Site ID or follow the prompts). |  | 13 | 13 | Worldwide Distribution - US (nationwide) including the states of NY, OH and Pa, and the countries of Australia, Austria, Egypt, France, Germany, India and Netherlands. | Terminated | March 26, 2014 | 222 | 13 | Computer
Z-0089-2014 | 66472 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | BrightView XCT is a gamma camera for Single Photon Emission Computed Tomography (SPECT), one unit per box. |  |  | September 03, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | Philips received reports from the field that the Flat Panel Detector (FPD) does not always remain securely latched in the stowed position. | Device Design | Philips Healthcare sent an Urgent - Medical Device Correction letter (Field Safety Notice (FSN 88200462), dated August 30, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the information contained in the notice with all members of their staff who need to be aware of the contents of the letter. Customers were asked to retain a copy with the equipment instruction for use. Additionally, a Field Service Engineer will contact each site to schedule time to implement the Field Change Order. A Field Service Engineer will visit the site to install the corrected software. The firm will be tracking the progress.Customers with questions were instructed to contact their local Philips representataive or local Philips Healthcare office. For North America and Canada contact the Customers Care Solutions Center at 1-800-722-9377, option 5. |  | 371 | 371 units | Worldwide Distribution - USA including AZ, CA, CO, DC, DE, FL, GA, HI, IA, IL, KY, MA, MD, MN, MO, MS, MT, NC, NJ, ND, NE, NV, NY, OH, OK, OR, PA, SD, TX, VT, WA, and WI.Internationally to Canada, Algeria, Argentina, Australia, Austria, Belgium, Brazil, China, Czech Republic, Denmark, Finland, France, Gabon, Germany, India, Indonesia, Ireland, Italy, Japan, Korea Republic, Lebanon, Malaysia, Mexico, Netherlands, New Zealand, Norway, Panama, Poland, Portugal, Russia, Saudi Arabia, Slovakia, Spain, Sweden, Switzerland, Taiwan, Turkey, United Arab Emirates, and United Kingdom. | Terminated | August 18, 2014 | 349 | 371 | Computer
Z-0893-2013 | 64313 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker XIA 3 Iliac Screwdriver, 3-PCS. |  |  | December 27, 2012 | 2012 | Stryker Spine | Surgeons have experienced unthreading of the screwdrivers outer shafts during use. | Device Design | Stryker Spine sent a "Urgent Medical Device Removal" letter dated January 14 2013, via FedEx overnight to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to examine your inventory and hospital inventory locations to identify your product. Please fax a copy of the Customer Response form to (855) 632-9049 to Regulatory Compliance. For further questions please call (201) 760-8298. |  | 46 | 46 units US | Nationwide Distribution including the states of AL, AZ, CA, CO, FL, GA, IA, IL,IN, KS, LA, MA, MD, MI, MO, MN, NC, NY. OH, OK, OR, PA, TN, TX, UT and WI. | Terminated | September 30, 2015 | 1007 | 46 | Computer
Z-0022-2014 | 66268 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | Harvest Graft Delivery System; Model Number: GDP-10 |  |  | September 11, 2013 | 2013 | Harvest Technologies Corporation | Potential for leur connector to leak or have cracks. | Device Design | Harvest Technologies sent a notification letter dated September 10, 2013, to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete the attached form to confirm the number of kits received, used, and intending to return. They are asked to return the kits with the form. For questions contact Michelle Sullivan at 508-732-7500. |  | 730 | 730 cases (3 kits per case) | Worldwide Distribution - USA Nationwide: TX, NJ, MI, AZ, GA, LA, ME, MD, AR, MO, VA, MN, GA, MA, NY, PA, CO, CA and in the countries of Australia, Brazil, China, Singapore, Vietnam, Dominican Republic. | Terminated | February 10, 2014 | 152 | 730 | Computer
Z-0894-2013 | 61554 | No | MHY |  |  | Neurology | PMA | Advanced Neuromodulation Systems, Brio DBS System implantable neurostimulator, IPG or R-IPG, Model 6788 Implantable Pulse Generator. |  |  | March 28, 2012 | 2012 | Advanced Neuromodulation Systems Inc. | Unintended Fluid ingress into the device header has been shown. | Device Design | St. Jude Medical sent an Important Product Field Corrective Action Device Information letter dated March 28, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their inventory for any unused product, isolate the affected product for return to St. Jude Medical. Their local representatiave would collect the unused inventory. Any unused product would be replaced with product once inventory becomes available. For patients with implanted product and no degradation of therapy, it is recommended dthat the devices not be unnecesssarily explanted. For patient with implanted product who experience diminished or loss of therapy customers were advised to follow the steps provided in the letter. Customers were asked to forward the notification to any institution in the event one or more patients or products could be potentially affected by this action. Customers wsere asked to maintain a record of thenotice along with the recommendations to ensure the effectiveness of the communication.For questions regarding this recall call 972-309-2154. |  | 388 | 388 units | Worldwide Distribution - 53 consignees out of the US. | Terminated | August 15, 2013 | 505 | 388 | Computer
Z-0896-2013 | 64265 | No | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Firebird Spinal Fixation System, Parallel Rod Connector (52-6800) and Axial Rod Connector (52-6700). |  |  | February 01, 2013 | 2013 | Orthofix, Inc | There is a possibility that the Set Screw (p/n 90-2923) sub-component of the Parallel Rod Connector Assembly (p/n 52-6800) &amp; Axial Rod Connector (p/n 52-6700) may crack during final tightening. Orthofix Inc received six (6) customer complaints related to set screw cracking between November 15, 2010 and December 2012. | Device Design | Orthofix sent an Urgent Medical Device Recall Notification dated February 1, 2013 to all affected customers. The letter identifed the affected product, problem and actions to be taken. Customers were instructed to identify and remove the affected products from inventory immediately and follow the actions listed in the enclosed Product Return Instructions. For questions or concerns contact Regulatory Affairs at 214-937-2100. |  | 565 | 565 total (287 PN 52-6800 and 278 PN 52- 6700) | Worldwidewide Distribution - USA Nationwide including Puerto Rico and the countries of: Australia, Lebanon, Germany, Spain, Sweden and Israel. | Terminated | July 01, 2013 | 150 | 565 | Computer
Z-0024-2014 | 65958 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | GE PET Discovery 610, PET Discovery 710. |  |  | March 26, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue related with the DMPR function and Screen Saved Images on your scanner which involves Optima CT660, Brivo CT385, Discovery PET/CT 610, Discovery PET/CT 710. The DMPR and Screen Save functionality may possibly generate the sameSeries UID. It has been seen that some PACS may combine series under a patient using only the Series UID | Software design (manufacturing process) | Consignees were sent a GE Healthcare "Urgent Medical Device Correction" letter dated March 26, 2013. The letter was addressed to Director of Radiology, Emergency Medicine Physicians and Radiologic Technologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. |  | 28 | 28 (4 US &amp; 24OUS) | Worldwide Distribution-USA (nationwide) including the states of AZ, CA, CO, FL, GA, IL, IA, IS, LA, MA, MO, NH, NM, NY, OH, OK PA, SC, TX, UT, VA, WA, and WI, and the countries of UNITED KINGDOM, DENMARK, CHILE, ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BELGIUM, CANADA, CZECH REPUBLIC, ECUADOR, EGYPT, ETHIOPIA, FINLAND, FRANCE, FRANCE GUIANA, GERMANY, GUADELOUPE, GUATEMALA, GUYANA, INDIA, INDONESIA, IRAQ, ITALY, JAPAN, JORDAN, KAZAHSTAN, KOREA, KUWAIT, LATVIA, MYANMA, NETHERLANDS, NEW ZEALAND, NORWAY, PAKISTAN, POLAND, ROMANIA, RUSSIA, SAUDI ARABIA, SOUTH AFRICA, SPAIN, SWEDEN SWITZERLAND, TAIWAN, THAILAND, TURKEY, UKRAINE, UZBEKISTAN, and VIETNAM. | Terminated | December 05, 2013 | 254 | 408 | Computer
Z-0897-2013 | 64403 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Ingenuity TF PET/CT. Diagnostic imaging system |  |  | January 21, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | The system may not default to the same gating trigger with which the data were acquired. This could cause a risk for patients. | Software design | Philips Medical Systems (Cleveland), Inc. sent an "URGENT-MEDICAL DEVICE CORRECTION" letter dated January 21, 2013 to the affected customer. The letter identified the product, problem, and actions to be taken by the customer. Contact your local Philips representative or Philips Healthcare office for further information concerning this issue. |  | 1 | 1 Unit | Distributed only in Ohio. | Terminated | July 19, 2013 | 179 | 1 | Computer
Z-0011-2014 | 65898 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | ADVIA 2120/2120i Hematology Systems with ADVIA Autoslide connected to it, SMN 10361162 and SMN 10361798 --- CLASSIFICATION NAME: Counter, Differential Cell, Automated (Particle Counter) |  |  | August 12, 2013 | 2013 | Siemens Healthcare Diagnostics | Siemens determined that if an autosampler rack jam error occurs during operation of an ADVIA 2120/2120i connected to an ADVIA Autoslide, it is possible that the next slide processed by the Autoslide could be labeled with the wrong sample identification information. | Software design | Siemens sent an Urgent Medical Device Correction Letter dated August 8, 2013, and Effectiveness Checks Forms were sent to customers on August 12, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were advised if they do not have an ADVIA Autoslide connected to their system, no action is required. To avoid the possibility of mislabeling a slide when an ADVIA 2120/2120i rack jam error occurs: Eject the sample rack. Troubleshoot and correct the reason for the rack jam error. Rerun the sample rack. Field service personnel were sent a support bulletin describing the issue and instructing them how to deal with customer questions.If you need additional assistance, please contact Global Product SupportTarrytown, New York @ osggps.hschematology.healthcare@siemens.com |  | 311 | 311 units total (US: 51 units; Foreign: 260 units) | Worldwide Distribution - USA (nationwide) and the countries of Australia, Austria, Belgium, Canada, Croatia, Denmark, France, Germany, Greece, Italy, Japan, Malaysia, Netherlands, Norway, Philippines, Portugal, South Korea, Singapore, Slovenia, Spain, Switzerland, Thailand, Turkey, United Kingdom and Vietnam. | Terminated | November 18, 2015 | 828 | 311 | Computer
Z-0009-2014 | 66381 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | GE Healthcare, Aisys, Amingo, Avance, and Avance CS2. |  |  | September 16, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue involving the Avance, Avance CS2, and Aisys anesthesia devices.While operating in Pressure Control Ventilation  Volume Guarantee (PCV-VG) mode, the affected products may produce an over delivery tidal volume when given a unique sequence of user inputs and a collapsed bellows (partially filled circuit).Exposure to excessiv | Software design | Consignees were sent a GE Healthcare "Urgent Medical Device Correction" letter dated September 17, 2013. The letter was addressed to Chief of Anesthesia, Health Care Administrator / Risk Manager, and Director of Biomedical / Clinical Engineering. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information (US call 1 800-345-2700, for other countries contact your local GE Healthcare Service Representative). |  | 10263 | 10,263 (3,194 US, 7,069 OUS) | Worldwide Distribution-USA (nationwide) including DC and Puerto Rico except DE, ME, NM, RI, and VT, and the countries of UNITED ARAB EMIRATES, JAPAN, KAZAKHSTAN, KENYA, REPUBLIC OF KOREA, KUWAIT, LATVIA, LEBANON, LIBYAN ARAB JAMAHIRIYA, LITHUANIA, MACEDONIA , MALAYSIA, MEXICO, MOLDOVA, MOROCCO, NAMIBIA, NETHERLANDS, NEW ZEALAND, NORWAY, PAKISTAN, PANAMA, PARAGUAY, PERU, PHILIPPINES, POLAND, PORTUGAL, QATAR, ROMANIA, RUSSIAN FEDERATION,SAUDI ARABIA, SERBIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, SYRIAN ARAB REPUBLIC, TAIWAN, TANZANIA, THAILAND, TURKEY, UNITED KINGDOM, VIETNAM, URUGUAY, and VENEZUELA. | Terminated | October 01, 2015 | 745 | 10263 | Computer
Z-0008-2014 | 66025 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | Draeger Fabius GS Premium, Fabius OS, Fabius Tiro, Fabius Tiro D-M |  |  | August 14, 2013 | 2013 | Draeger Medical, Inc. | During final testing of the Draeger Fabius GS Premium, Fabius OS, and Fabius Tiro anesthesia machines, some units did not pass the high voltage test and the minimum clearance between an electrical component and the unit housing was not maintained for some from a specific batch. | Device Design | The firm, Drager, sent an "Urgent Medical Device Recall" letter dated August 2013 to its customers. The letter described the product, problem and actions to be taken. The letter informed the customers that in the event of a failure, manual ventilation mode should be used and that Draeger reps would be contacting customers to schedule the replacement of the power supply. The customers were instructed to notify personnel accordingly.If you have any questions, contact Director, Quality at 1-800-543-5047 (press 1 at the prompt, then 2, then 32349). For questions regarding the operation and/or servicing of your Draeger anesthesia machine please contact DragerService Technical Support at 1-800-543-5047 (press 4 at the prompt). |  | 1127 | 1127 units | Worldwide distribution: US (nationwide) and countries of: Argentina, Australia, Austria, Azerbaijan, Bangladesh, Brazil, Bulgaria, Canada, Chile, China, Columbia, Ecuador, Egypt, France, Germany, Greece, Hungary, India, Indonesia, Israel, Italy, Ivory Coast, Japan, Kazakhstan, Latvia, Libya, LithuaniaMalaysia, Morocco, Myanmar, Nepal, Philippines, Poland, Rumania, Russian Fed., Saudi Arabia, Senegal, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Turkey, United Kingdom, Uruguay, Vietnam, White RussiaTurkmenistan. | Terminated | September 23, 2014 | 405 | 1127 | Computer
Z-0007-2014 | 66374 | Yes | NIP |  |  | Cardiovascular | PMA | Bard LifeStent Solo Vascular Stent, Item Numbers: EX062001CL, EX072001CL, EX062003CL, EX072003CL. |  |  | September 30, 2013 | 2013 | Bard Peripheral Vascular Inc | Bard Peripheral Vascular is recalling specific lots of Bard Lifestent Solo Vascular Stent because they have determined a higher incidence rate of partial stent deployment associated with their use. | Device Design | The firm, Bard Peripheral Vascular, sent a "Urgent: Medical Device Recall Notification" letter dated September 30, 2013 via FedEx to all the customers who received the specific lots on recall. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to fill out the Recall and Effectiveness Check Form and fax to: BPV at 1-800-994-6772, even if you no longer have possession of the recalled product; call Bard Peripheral Vascular's Recall Coordinator at (800) 321-4254 Option #2 Ext 2727 (M-F 7am to 4pm MST) or email silvia.carrillo@crbard.com. The Recall Coordinator will issue you either a Return Authorization (RCL) Number or Consignment Recall Number (CRC) to facilitate the expedient return of the product; and to return the product to Bard Peripheral Vascular, Inc.,1415 W. 3rd Street,Tempe, AZ 85281. BPV will provide replacement product for your returned product.If you have any questions, call 1-800-321-4254 Option #2 Ext. 2727. |  | 4112 | 4112 units | US Nationwide Distribution. | Terminated | February 05, 2014 | 128 | 4112 | Computer
Z-0900-2013 | 63671 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips Digital Diagnost |  |  | October 30, 2012 | 2012 | Philips Healthcare Inc. | When the operator for a wall stand view selects an "image rotation" different from default, or such image rotation is preset in the examination database, then the resulting image may be rotated in the wrong direction. When the operator manually rotates the image back, electronic side markers (if programmed) are rotated also and as a result may be placed in a wrong position inside the image. | Software design | Philips Healthcare issued an Urgent Field Safety Notice dated October 25, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to review the information contained in the Notice with all member of their staff who need to be aware of the contents. Also, customers were asked to retain a copy with the equipment Instruction for Use. For further information or questions customers were instructed to contact their local Phillips representative Technical Support Line at 1-800-722-9377.For questions regarding this recall call 978-659-4519. |  | 360 | 360 | Worldwide Distribution - USA (nationwide) and Internationally to Argentina, Australia, Austria, Canada, Chile, Czech republic, Estonia,France, Germany, Hungary, Inida, Ireland, Italy, Japan, Libyan Arab Jamahari, Mexico, Netherlands, New Zealand, Norway , Oman, Panama, Peru, Poland, Qatar, Reunion, Russia, Saudi Arabia, South Africa, South Korea,Spain, Sweden, Swizerland, Thailand, UAE, and the UK. | Terminated | July 29, 2013 | 272 | 360 | Computer
Z-0901-2013 | 64071 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Volume Viewer Plus 5.x (Advantage Workstation 4.2) Volume Viewer 3.x (Advantage Workstation 4.1) Volume Analysis (Advantage Workstation 4.0) |  |  | May 25, 2012 | 2012 | GE Healthcare, LLC | Incorrect views and measurements possible. GE Healthcare became aware of multiple potential safety issues involving the Advantage workstation, Volume Viewer Plus and Volume Analysis. | Software design | GE Healthcare sent a "Urgent Medical Device Correction " letter dated May 23, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to discontinue use of the Edit Contour feature until this software functionality is corrected or replaced.GE Healthcare will correct all affected systems by providing a software upgrade at no cost to you. A GE Healthcare service representative will contact you to arrange for this correction. If you have any questions or concerns regarding this notification, please call one of the following phone numbers:United States: 800 437 1171Japan: 0120 - 055 - 919For other countries, please contact your local GE Healthcare Service Representative. Please be assured that maintaining a high level of safety and quality is our highest priority. If you have any questions, please contact us immediately per the contact information above. |  | 363 | 363 | Worldwide Distribution - USA including the states of AR, CA, CO, FL, GA, IL IN, IA, LA, MD, MA, MI, MN, MO, NJ, NM, NY, NC, OH, OK, PA, PR, SC, SD, TN, TX and VA. and the countries of ALGERIA, AUSTRALIA, AUSTRIA, ARGENTINA, BELARUS, CANADA, CHINA, CHILE, EGYPT CZECH REPUBLIC, FRANCE, FINLAND, GERMANY, GABON, HONG KONG, INDIA, IRELAND, ITALY KENYA, KOREA, JAPAN, MEXICO, NIGERIA, PHILIPPINES, PORTUGAL REPUBLIC OF MALTA, RUSSIAN FEDERATION SAUDI ARABIA, SPAIN SWITZERLAND, TURKEY, TAIWAN, TUNISIA, UNTIED KINGDOM and VENEZUELA. | Terminated | January 28, 2014 | 613 | 363 | Computer
Z-0905-2013 | 64195 | No | DZL | Intraosseous fixation screw or wire. | Dental | Dental | 510(k) | Synthes Multi-Vector Distractor - Multi-axial Mandible Distractor (Left), Part Number 487.967. |  |  | January 11, 2013 | 2013 | Synthes USA HQ, Inc. | It was discovered that the sheer pin on the Multi-Vector Distractor may not be made according to product specifications. | Device Design | The firm, Synthes (USA) sent an "URGENT NOTICE: MEDICAL DEVICE RECALL" letter dated January 11, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firmat 610-719-5450 for questions regarding this notice. |  | 18 | 18 | Nationwide Distribution-including the states of CO, LA, MD, MN, NY, PA, RI, TN, TX and UT. | Terminated | August 31, 2015 | 962 | 88 | Computer
Z-2318-2013 | 66301 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | McKesson Horizon Medical Imaging (HMI) Radiological Image Processing System, product versions 4.6.1 to 11.6., is a medical image and information management application intended to receive, transmit, store, retrieve, display, print and process digital medical images, digital medical video, and associated medical information from medical imaging systems. |  |  | September 04, 2013 | 2013 | Mckesson Information Solutions LLC | Software malfunction may occur resulting in a certain combination of server configurations which has the potential to cause image loss. | Software change control | Consignees were notified via Urgent Field Safety Notice dated September 4. MMIC Customer support will be contacting customers known to have a multi-server configuration and will implement the configuration change. |  | 329 | 329 units | Worldwide Distribution - United States including the territories of Guam and Puerto Rico and the countries of Canada and France. | Open, Classified |  |  | 329 | Computer
Z-0892-2013 | 64136 | No | GYC | Cortical electrode. | Neurology | Neurology | 510(k) | Ad-Tech Medical Instrument Corporation Macro Micro Subdural Electrode. For temporary (&lt;30 days) use with recording, monitoring and stimulation equipment, for the recording, monitoring and stimulation of electrical signals on the surface level of the brain. The recording of electrical activity supports definition of the location of epileptogenic foci and brain mapping. |  |  | December 18, 2012 | 2012 | Ad-Tech Medical Instrument Corporation | The microelectrodes are not flush with the silastic substrate, may contain sharp edges capable of abrading the cortical surface, and may have remaining pieces which may fracture and remain in the cortical tissue with the electrode is explanted. Recall expanded on 5/2/2013 to include additional Macro Micro Subdural Electrodes. | Device Design | AD-TECH Medical Instrument Corporation sent an "Urgent Medical Device Recall" letter dated December 18, 2012 to all affected consignees. The letter described the Reason for the Voluntary Recall, Risk to Health, Actions to be taken by the Customer, Product and Distribution Information, Type of Action by Ad-Tech Medical Instrument Corporation and Other Information. Consignees were instructed to quarantine the affected product and contact the firm for a Return Material Authorization number. If the product have been further distributed consignees were advised to notify their customers of the recall. The letter requested consignees to return the Acknowledgement and Receipt Form.For questions contact your Ad-Tech Clinical Specialist.Expanded recall letters dated May 2, 2013 were sent out April 29, 2013. The letters were formatted as the initial ones sent on 12/18/2013. |  | 144 | 144 (103 US; 41 OUS) | Worldwide Distribution - USA Nationwide including the states of CA, GA, LA, MA, MD, MI, MN, NH, PA, RI, TX, UT, WI, NY IA, WI and the countries of: Australia, CANADA, ENGLAND, JAPAN, KOREA, SPAIN, FRANCE, and SWITZERLAND. | Terminated | October 15, 2014 | 666 | 144 | Computer
Z-0907-2013 | 64214 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | SIEMENS brand ARTISTE syngo RT Therapist (RTT) version 4.3, ARTISTE Model, Part No. 08162815, Digital linear accelerators, MOSAIQ OIS system, Distributed by SIEMENS Healthcare IM CR RO, CR Radiation Oncology, Manufactured by SIEMENS AG, MEDICAL SOLUTIONS, GERMANY. |  |  | January 17, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Siemens became aware of a potential safety issue when using the "In-session resumption" feature of the syngo RT Therapist (RTT) 4.3 software. The RTT will send again the previously completed beam or segment for delivery to the control console, thus repeating the already delivered Monitor Units of the same beam or segment. | Software design | The firm issued "Urgent Medical Device Correction" letters via mail to all its direct consignees, informing them of the affected product issue and providing instructions on the recall. The firm stated that the release of this UI TH005/13/S to distribute the Customer Safety Advisory Notice began on January 17, 2013. |  | 5 | 5 units | 5 devices, all located outside of the United States in Germany. | Terminated | March 05, 2013 | 47 | 5 | Computer
Z-0906-2013 | 64259 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Philips Medical System Extended Brilliance Workspace (EBW), Computed Tomography X-Ray System. |  |  | September 28, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Incorrect values may be displayed. Philips is issuing a software update (version 4.5.5) to the Extended Brilliance Workspace (EBW) systems currently running software version(s) 4.0.2, V4.5.2 or V4.5.3, addressing a number of non-conformances. | Software design | Philips Healthcare sent a Urgent Medical Device Correction letter dated September 28, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.This Field Safety notification is to inform the actions that should be taken by the customer/user to prevent risks for patient or users, and the actions planned by Philips to correct the problem.If you need further information or supplies please contact your local Philips Healthcare office. For North American and Canada contact the Customer Care Solutions Center ( 1-800 722-9377) option 5 enter site or follow the prompts. |  | 2580 | 2,580 units | Worldwide Distribution - USA including the states of AL, AT, AZ, BAH, CA, CO, CT, DC, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NJ, NV, NY, OH, OR, PA, PR, PR, SC, TN, TX, VA, VT, WA, WI &amp; WV. and the countries of Albania, Algeria, Argentina, Australia, Austria, Bahrain, Bangladesh, Belgium, Bolivia, Brazil, Brunei Darussalam, Bulgaria, Canada, Chile, China, Columbia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican, Republic, Ecuador, Egypt, Ethiopia, Finland, France, French Polynesia, Gabon, Georgia, Germany, Greece, Guadeloupe, Guinea, Hong Kong, Hungary, Iceland, India, Indonesia, Iran, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kuwait, LAO People's Dem, Rep., Latvia, Lebanon, Libyan Arab Jamahiriya, Lithuania, Malawi, Malaysia, Malta, Martinique, Mexico, Netherlands, Netherlands, Netherlands Antilles, New Caledonia, New Zealand, Nigeria, Norway, Oman, Pakistan, Palestine, Panama, Peru, Philippines, Poland, Portugal, Qatar, Reunion, Romania, Russia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sudan, Sweden, Switzerland, Syrian Arab Republic, Taiwan, Tanzania, Thailand, Tunisia, Turkey, Uganda, Ukraine, United Arab Emirates, United Kingdom, Uzbekistan, Venezuela &amp; Viet Nam. | Terminated | December 26, 2013 | 454 | 2580 | Computer
Z-2131-2013 | 65923 | No | JTC | Microtiter diluting and dispensing device. | Microbiology | Microbiology | 510(K) Exempt | WASP (Walk Away Specimen Process). Automatic Planting and Streaking Instrument |  |  | August 09, 2013 | 2013 | Copan Italia | The recall of the WASP (Walk Away Specimen Process) instrument was initiated because Copan Italia SpA has upgraded the software concerning the broth inoculation table feature. The upgrade has been considered necessary from the manufacturer in order to eliminate the very low risk of rotation of the broth table when the door is open. | Device Design | Copan sent an Urgent WASP recall letter dated July 31, 2013, to all affected customers that received the WASP (Walk Away Specimen Process) instrument to notify them of the software upgrade concerning the broth inoculation table feature. The upgrade has been considered necessary from the manufacturer in order to eliminate the very low risk of rotation of the broth table when the door is open. This event may happen only if damage in the door locking system happens. This could result in a minor, transient injury to the user due to the impact with the rotating table.The WASP instruments involved in this upgrade are all whose serial numbers show the configuration with the broth inoculation feature activated.The machine installed in your laboratory is one of this serial number involved in the SW upgrade. The Software upgrade will be performed on site by personnel trained by Copan and no action is required by yourself.Customers will be contacted by the WASP Technical Support in order to define a date that will fit with their planning. Customers with questions are instructed to call +39 0303666130 or write to regulatory.wasp@copanitalia.com. |  | 63 | 63 units | Worldwide Distribution - USA including Oregon and California., and the countries of France, Belgium, Italy, Denmark, Norway, Switzerland, Romania, Finland, Germany, South Africa, Canada, Ireland, UK, Sweden and Japan | Terminated | December 12, 2013 | 125 | 63 | Computer
Z-2130-2013 | 65902 | No | GDJ | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Stryker Howmedica Osteonics Triathlon tibial alignment ankle clamp extramedullary (EM) referencing instrument (Cat. No. 6541-2-609), for use in Total Knee Arthroplasty (TKA) when the surgeon elects to use Stryker's Triathlon Total Knee System in conjunction with EM referencing. Non-Sterile Instrument. |  |  | July 02, 2013 | 2013 | Stryker Howmedica Osteonics Corp. | Cracks or fracture of the Triathlon Tibial Alignment Ankle Clamp, an instrument. | Device Design | Stryker Orthopaedics sent an Urgent Medical Device Correction letter dated July 2, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by t he customers. Customers are advised of the issue, potential hazards and risk mitigation factors. Customers are asked to complete the attached form and fax to 855-251-3635. Any questions should be directed to (201) 972-2100.All letters included a Product Correction Bulletin. |  | 10537 | 10,537 units | Nationwide Distribution | Open, Classified |  |  | 10537 | Computer
Z-2111-2013 | 65779 | No | MKQ |  |  | Pathology | PMA | PrepStain Tecan US-I. |  |  | July 08, 2013 | 2013 | Becton Dickinson &amp; Co. | PrepStain system rack setting was incorrectly entered on one PrepStain system during a routine preventive maintenance visit by a BD Field Service Representative. | Software design | Ex-US BD sites were notified via email and teleconference, and will communicate directly with customer sites in their region as required by local regulation. All US customers were initially contacted via telephone and letters were mailed. |  | 890 | 890 instruments | Worldwide Distribution-USA (nationwide) and Puerto Rico and the country of Canada. | Terminated | May 01, 2014 | 297 | 890 | Computer
Z-1021-2013 | 64469 | No | LJG | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | Microbiologics KWIK-QC SLIDES ( KWIK-QC Acid Fast Stain Slides and KWIK-QC Mycobacterium Slides). |  |  | February 11, 2013 | 2013 | Microbiologics Inc | Microbiologics Inc is initiating a voluntary recall on products KWIK-QC Acid Fast Stain Slide, SL42-10 and KWIK-QC Mycobacterium, SL43-10. These products are being recalled because the control organisms are not properly fixated on the slide, which may cause the user to unintentionally contaminate the patient area of the slide by flooding the control well with stain. This contamination could pote | Device Design | Microbiologists an "Urgent Medical Device Recall" letter dated February 11, 2013. The Letter described the product and the reason for the recall. Advised consignees to "Do not Distribute" and "Return" the product. Urged consignees to "Notify" end users and to complete the attached "Recall Response Form". Contact the form at 320-229-7057 for questions regarding this notice. |  | 3960 | 396 boxes of 10 slides each. | Worldwide Distribution-USA including the states of CA, IN, MS, NY, OH, PA, and SD, and the countries of AUSTRALIA, AUSTRIA, UNITED ARAB EMIRATES, MALAYSIA, TAIWAN, JAPAN, TURKEY, BELGIUM, COLOMBIA, CANADA, SOUTH AFRICA, and VENEZUELA. | Terminated | March 10, 2016 | 1123 | 3960 | Computer
Z-1927-2013 | 65444 | No | HXX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Invue Split Driver, Assembly with Handle, Model SI70006 Rev G. Used in conjunction with the Invue Anterior Cervical Plate System under 501(k) clearance number K121060. |  |  | May 10, 2013 | 2013 | SpineFrontier, Inc. | Device design. | Device Design | SpineFrontier sent an URGENT Advisory Notice dated June 27, 2013, to all affected consignees regarding the use of Invue Drivers. The notice advises users that the Split Tip Stylet Driver Assembly (11-70099) and Sold Driver, AO Connection (11-70070 or 11-70077) are being shipped in place of the Split Driver Assembly w/Handle (SI70006). The uses of the replacement parts are explained. Consignees are instructed to complete the Acknowledgement Form and return it to SpineFrontier.On June 28, 2013, customers were sent a second Advisory Notice requesting the removal of the Split Driver Assy w/ Handle Part Number SI70006 Lot Numbers 027104 and 027105. Consignees were instructed to complete the Acknowledgement Form and return it to SpineFrontier.Consignees with questions were instructed to call 978-232-3990. |  | 12 | 12 devices | Nationwide Distribution including Texas, Ohio, and Missouri. | Terminated | July 20, 2015 | 801 | 12 | Computer
Z-1943-2013 | 65133 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | SIGNA INFINITY MR SYSTEM WITH EXCITE TECHNOLOGY |  |  | November 21, 2011 | 2011 | GE Healthcare, LLC | Screen save images are included in wrong exam due to duplicate UID creation. This may cause a safety issuedue to a potential patient data mismatch.While the issue has only been observed at one customer site on the Signa HDxT 1.5T system, internal investigation revealed that the issue also impacts multiple GE healthcare MR Systems. | Software design | GE Healthcare communicated internally to their Service Team the directions for completing the correction on the customers' systems. GE Healthcare's recall strategy was to update all affected systems that were distributed. At this time,Week of 02/06/2012, deployment of this recall is considered complete.Should you have any questions or require additional information, please contact Regulatory Affairs Leader by telephone at 262-548-2608 or email paulcorrigan@ge.com. |  | 7016 | 7016 (2919 US; 4097 OUS) | Worldwide distribution: USA (nationwide) including DC and Puerto Rico and countries: ALBANIA, ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA, AZERBAIDJAN, BELARUS, BELGIUM, BOLIVIA, BOSNIA-HERZEGOVINA, BRAZIL, BULGARIA, CANADA, CHILE, CHINA, COLOMBIA, CROATIA, CYPRUS, CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, ECUADOR, EGYPT, FINLAND, FRANCE, GEORGIA, GERMANY, GREAT BRITAIN, GREECE, GUADELOUPE, GUATEMALA, GUYANA, HONDURAS, HONG KONG, HUNGARY, ICELAND, INDIA, INDONESIA, IRAN, IRELAND, ISRAEL, ITALY, JAPAN,JORDAN, KAZAKHSTAN, KENYA, KUWAIT, LATVIA, LEBANON, LIBYA, LITHUANIAN, LUXEMBOURG, MALAYSIA, MALTA, MARTINIQUE, MEXICO, MOROCCO, NETHERLANDS, NEW ZEALAND, NORWAY, PAKISTAN, PANAMA, PARAGUAY, PERU. PHILIPPINES, POLAND, PORTUGAL, REPUBLIC OF MACEDONIA, REUNION, ROMANIA, RUSSIA, SAUDI ARABIA, SINGAPORE, SLOVAK REPUBLIC, SLOVENIA, SOUTH AFRICA, SOUTH KOREA, SPAIN, SUDAN, SWEDEN, SWITZERLAND, SYRIA, TAIWAN, THAILAND, TUNISIA, TUR,EKY, UKRAINE, UNITED ARAB EMIRATES, URUGUAY, VENEZUELA, VIETNAM, YUGOSLAVIA, MALAWI, MAURITANIA, NIGERIA, UKRAINE, IVORY COST, OMAN, ZAMBIA, UZBEKISTAN, ESTONIA, IRAQ, and JAMAICA. | Terminated | October 22, 2014 | 1066 | 7016 | Computer
Z-1940-2013 | 65333 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | GE LightSpeed CT750 HD |  |  | April 05, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare conducted a recall involving 2 events affecting the GE LightSpeed CT750 HD. Event 1 - There was a report regarding CT number variance in time line at posterior fossa region of head perfusion. When beam hardening and helical artifacts in this region are reconstructed into the perfusion map, these artifacts will give subtle differences in the CT numbers.Event 2 - The software does no | Software design | GE Healthcare's recall strategy was to update CT750HD CT Systems with software 10MW25.6 on all systems that were distributed. Updates started on 4/5/2011. At this time, deployment of this recall is considered complete. The recall was completed 2/2/2012.Should you have any questions or require additional information, please contact Regulatory Affairs Leader by telephone at 262-548-2198 or email paulcorrigan@ge.com. |  | 244 | 244 | Worldwide distribution: US (nationwide) including states of: AL, AZ, CA, CT, FL, GA, HI, IL, IN, IA, KY,ME, MD,MA, MI,MN, MS, NV, NH, NJ, NY, NC, PA, TN, TX, VA, WA,WV, and WI; and countries of: AUSTRALIA, BELGIUM, BRAZIL, CANDA, CHINA, DENMARK, FINLAND, FRANCE, GERMANY, GREECE, HONG KONG, INDIA, INDONESIA, IRELAND, ISRAEL, ITALY, JAPAN,JORDAN, REPUBLIC OF KOREA, KUWAIT, LATVIA, NEW ZEALAND, POLAND, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SWEDEN, SWITZERLAND, TUKEY, and UNITED KINGDOM. | Terminated | August 11, 2014 | 1224 | 244 | Computer
Z-1944-2013 | 65583 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | GE Healthcare, LOGIQ S8 Ultrasound System |  |  | August 13, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare became aware of a potential issue related to incorrect image measurement associated with the software of the LOGIQ S8 system. This issue was found internally. Users may notice that the system shows incorrect image measurement results on Pulsed Wave (PW) or TimeIntensity Curve (TIC) or depth measurement when using Virtual Convex mode and putting a cursor (Sample Volume, Depth measur | Software design | GE initiated a field action on 8/13/12. GE plans to update the software on all of the affected GE LOGIQ S8 Ultrasound systems that were distributed, at no cost to the customer.Should you have any questions or require additional information, please contact RA Leader, US Region by telephone at 262-548-2089 or email James.Giles@med.ge.com. |  | 755 | 755 | Worldwide distribution: US (nationwide) including states of: AL, AZ, CA, CO, CT, DC, FL, GA, IL, IN, IA, LA, MA, MI, MS, MO, MT, NE, NV, NJ, NY, NC, OH, OK, PA,RI, SC, TN, TX, VA, WA, and WI; and countries of: ALGERIA, AUSTRALIA, AUSTRIA, BRAZIL, CANADA, COLOMBIA, CZECH REPUBLIC, DENMARK, ESTONIA, FINLAND, FRANCE, GERMANY, GREECE, INDIA, INDONESIA, IRELAND, ITALY, JAPAN, LATVIA, LEBANON, MALAYSIA, MOROCCO, NETHERLANDS, NEW ZEALAND, NORWAY, POLAND, PORTUGAL, REPUBLIC OF MOLDOVIA, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SOUTH KOREA, SPAIN, SWEDEN, SWITZERLAND, THAILAND, TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, and YUGOSLAVIA. | Terminated | September 15, 2014 | 763 | 755 | Computer
Z-1942-2013 | 64992 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Optima XR220amx, Optima XR200amx, and Brivo XR285amx. |  |  | August 07, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare conducted a recall on various mobile x-ray units (Optima XR220amx, Optima XR200amx, and Brivo XR285amx systems) because of an Erratic Drive Performance. | Software design | GE Healthcare recall strategy was to update all of the affected Optima XR220amx, Optima XR200amx, and Brivo XR285amx systems that were distributed. At this time ( 2/14/2013), deployment of this recall is considered complete.Should you have any questions or require additional information, please contact Regulatory Affairs Leader by telephone at 262-548-2198 or email paulcorrigan@ge.com |  | 458 | 458 (306 US; 152 OUS) | Worldwide distribution: USA (nationwide) including states of: AL, AK, AZ, AR, CA, CO, CT, DC, FL, GA, GUAM, IL, IN, IA, KY, LA, ME, MD, MA, MI, MN, MS, MO, NV, NH, NH, NM, NY, NC, ND, OH, OK, OR, PA, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, and WY; and countries of: Albania, Australia, Bahrain, Belgium, Canada, Chile, Czech Republic, Dominican Republic, France, Germany, Great Britain, Honduras, Hungary, Ireland, Israel, Italy, Japan, Korea, Kuwait, Netherlands, Nigeria, Norway, Romania, Saudi Arabia, Spain, United Arab Emirates and Venezuela. | Terminated | February 04, 2014 | 546 | 458 | Computer
Z-1025-2013 | 64075 | No | FZZ | Skin marker. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Dual Tip Marker, Sterile |  |  | August 22, 2012 | 2012 | Viscot Medical LLC | Dual Tip Marker has inconsistent ink flow on the black permanent ink side. | Device Design | Viscot Medical LLC sent a Customer Notification Letter dated August 22, 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to return the affected product for a replacement. To return the product call 1-800-221-0658 or e-mail jeffl@viscot .com for the RGA. |  | 458589 | 458,589 markers | USA Nationwide Distribution | Terminated | July 21, 2014 | 698 | 458589 | Computer
Z-1945-2013 | 65721 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Siemens AXIOM Artis zeego |  |  | May 13, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Siemens issued a customer safety notice about the overlay of system information after a functional check of the brakes on the Artis Zeego system. | Software design | The firm, Siemens, sent an "Important safety notice" by letter on May 13, 2013 to its customers. The notice described the product, problem and actions to be taken. The customers were instructed to immediately notify the service department to clarify what further measures need to be taken, if after carrying out the regular brake test, the system can only be moved at a reduced speed. If this device has been sold and is therefore no longer in your possession, send this safety notice to the new owner and provide the new owner to Siemens.Update instruction AX012/13/S was issued, and a software update will be released as AX006/13/S and AX007/13/S which will modify the system software so that it will not be possible for information to be overlaid.If you have any questions call +49 (9191) 18 0. |  | 132 | 132 | Nationwide distribution: AL, AK, AR, AZ, CA, CO, CT, DC, FL, GA, IA, IL, IN, KY, LA, MA, MD, MN, MO,MS, NC, NE, NH, NJ, NY, OH, PA, PR, SD, TN, TX, UT, VA, WA, and WI. | Terminated | November 06, 2014 | 542 | 132 | Computer
Z-2118-2013 | 49516 | No | KPA | Automated slide stainer. | Pathology | Pathology | 510(K) Exempt | Benchmark XT/LT and Discovery XT |  |  | June 12, 2008 | 2008 | Ventana Medical Systems Inc | Ventana Medical System is initiating the recall of the BenchMark and DISCOVERY series instruments because there is a possibility that the heating element could separate from the underside of the heater cap and cause a non-staining or inconsistent staining event. Additionally, the system may not detect this anomaly and therefore it may not generate an error message for the slide position. | Device Design | Ventana sent an Urgent Medical Device Correction Notification letter dated June 16, 2008 to all affected customers. The letter identifed the affected product, problem and actions to be taken. The letter instructed customers to carefully read and perform all of the actions outlined in the notification letter concerning potential safety issue. Customers with concerns regarding the staining results on any patient specimen are instructed to follow their internal Quality Procedures regarding the review of patient reports. For question contact your local Customer Service Center. |  | 1217 | 1217 units total | Worldwide Distribution - USA Nationwide and the countries of Canada, Europe, Asia, Japan. | Terminated | August 29, 2013 | 1904 | 1217 | Computer
Z-1934-2013 | 65837 | No | MMO | Ear, nose, and throat manual surgical instrument. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(K) Exempt | Tray Surgical, ENT (various instrument trays) |  |  | July 09, 2013 | 2013 | Medtronic Xomed, Inc. | The firm will recall and replace the identified instrument trays in the field, with trays that have improved labeling, including lids that contain the statement indicating 'PREVACUUM ONLY" | Device Design | Medtronic sent a Customer Commincation Recall Letter to all affected customers with the following information: Subject: URGENT PRODUCT Recall  Immediate Attention RequiredMEDTRONIC XOMED INSTRUMENT TRAYS- model numbers provided on Attachment One- all lots distributed prior to July 2013Dear Purchasing Manager, Risk Manager and Central Supply Manager/Supervisor:This letter is to inform you that Medtronic Surgical Technologies (ENT Division) is initiating a voluntary recall to remove certain instrument trays from the field and replace them with trays that have modified labeling. Medtronic Xomed Instrument Trays are intended for the storage and transportation of reusable surgical instruments. These trays are not intended and not designed to maintain sterility. We are aware, however, that some users may also choose to use the trays for holding instruments during steam sterilization. If the user chooses to sterilize instruments within these instrument trays, they must validate the sterilization process. These users should be aware of the following issue:Issue Description:During the validation of a new instrument, testing revealed that use of a GRAVITY steam cycle with one of the listed instrument trays did not produce the Sterility Assurance Level (SAL) of 10-6 recommended by industry standard (ANSI/AAMI ST77:2006). A subsequent investigation of all Medtronic Xomed instrument trays identified additional trays as having certain locations within thetray that would not meet the recommended sterilization level of SAL10-6 during GRAVITY steam sterilization cycles. The trays identified through the testing are the subject trays involved in this field action. Instruments that are sterilized in the worst-case locations of the identified trays have a higher potential to be inadequately sterilized. An instrument that is inadequately sterilized can become the source of cross-contamination or cross-infection when used in multiple surgical procedures. At this time |  | 3562 | 3,562 trays | Worldwide Distribution - USA (nationwide) Nationwide, Hawaii, Puerto Rico and Guam.Internationally to: Canada, Austria, Belgium, Cyprus, France, Germany, Greece, India, Iran, Ireland, Israel, Italy, Kuwait, Lebanon, Luxembourg, Morocco, Netherlands, Nigeria, Portugal, Qatar, Reunion, Saudi Arabia, South Africa, Spain, Sweden, Switzerland, Turkey, United Arab Emirates, United Kingdom, Australia, New Zealand, Singapore, Thailand, Vietnam, China, Hong Kong,Taiwan, Korea, Japan, Brazil, Chile, Colombia, Costa Rica, Ecuador, Guatemala, Mexico, Nicaragua, and Venezuela. | Terminated | September 25, 2014 | 443 | 3562 | Computer
Z-1949-2013 | 64304 | No | NQI | Immunomagnetic circulating cancer cell selection and enumeration system. | Immunology | Pathology | 510(k) | CELLTRACKS AUTOPREP System (Software Version 2.4.1 &amp; Below. |  |  | December 28, 2012 | 2012 | Veridex, LLC | Veridex, LLC has received a report that a patient's demographics were incorrectly associated with another patient's identification (PID) number. The purpose of this notification is to inform you of the issue and provide instructions to help prevent an occurrence on your CELLTRACKS AUTOPREP System and CELLTRACKS ANALYZER II. | Software design | Veridex sent an Important Product Correction Notification letter dated December 28, 2012, to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete and return the Confirmation of Receipt form. For questions contact Customer Technical Services at 1-800-421-3311. |  | 160 | US: 82 Units; Ex-US 78 Units. | Worldwide Distribution - USA Nationwide and the countries of Australia, Canada, China, India, Japan, Singapore, England, France, Germany, Italy and Spain. | Terminated | August 11, 2015 | 956 | 306 | Computer
Z-1954-2013 | 65502 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | REF 00-4309-071-40 Trabecular Metal Reverse Shoulder System Instrumentation, Glenosphere Helmet, 40 mm Diameter Nonsterile, Qty -1 |  |  | June 28, 2013 | 2013 | Zimmer, Inc. | There is potential for one or both tabsof theTrabecular Metal Reverse Glenosphere Helmet to fracture during use and become disassociated from the device. There have been 47 reported complaints of tab fracture. | Device Design | Zimmer Inc sent Urgent Medical Device Recall notifications on June 28, 2013, to all affected customers. The initial letter identified the issue, possible adverse consequences and informed customers that a replacement would be available in the coming months. In addition, the letter included references to both the TM Reverse Surgical Technique and to the video available on Zimmer TV and Zimmer.com. Both resources explain in detail the proper technique for utilization of the Glenosphere Helmet. Proper utilization of the Glenosphere Helmet dramatically reduces the likelihood of fracture. Distributors will receive a copy of the surgeon letter for their reference along with the list of affected surgeons. The affected surgeons will be identified using invoice data and in cases where the surgeon is unknown, the distributor will be instructed to assist with the surgeon list. In addition, information regarding the return process for the damaged devices will be included. Questions and concerns are directed to Zimmer's customer call center (1-877-946-2761).Phase 2: As replacement devices with the new design became available for distribution, a 2nd letter was sent to the distributors. Estimated replacement device availability will begin in July 2013. The distributors will be instructed to order replacement devices as a means of completing the affected instrument cases in circulation within their territories. Instructions for the return of devices were included.For questions regarding this recall call 800-613-6131. |  | 3163 | 3,163 both sizes | Worldwide Distribution - USA (nationwide) and internationally to Australia, Austria, Belgium, Brazil, Canada, China, Czech Republic, Denmark, Finland, France, Germany, Greece, India, Israel, Italy, Korea, Lebanon, Malta, Netherlands, Norway, Portugal, Singapore, Spain, Sweden, Switzerland, Taiwan, Thailand, United Kingdom, Utd.Arab.Emir. | Terminated | March 02, 2015 | 612 | 3163 | Computer
Z-1952-2013 | 65560 | No | HTO | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Synthes Flexible Medullary Reamer. |  |  | June 11, 2013 | 2013 | Synthes USA HQ, Inc. | Due to the coiled design for this product, the product is difficult to clean and the potential for corrosion on the device exists. The potential for harm exists if there is an introduction of foreign material at the operative site originating from the presence of corrosion and/or the retained undefined material that could not be adequately cleaned. | Device Design | Sythes notified direct accounts by letter on 6/11/13 and requested them to check their inventory and remove affected lots from stock. Synthes requested that affected lots be returned and to call Synthes at 1-800-479-6329 for a Return Authorization Number. Completion and return of the Verification section at the end of the letter was requested, including negative responses. |  | 13050 | 13,050 | Nationwide Distribution. | Terminated | August 31, 2015 | 811 | 13050 | Computer
Z-1957-2013 | 64994 | No | HXX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Invue Solid Driver, Tipped, AO. Part Number: 11-70083. |  |  | May 09, 2012 | 2012 | SpineFrontier, Inc. | The nipple of the tipped driver broke off when surgeons were attempting to insert the screw. | Device Design | On May 9, 2012 an email was sent to consignees requesting that the product be returned for servicing and that a replacement part would be sent. |  | 9 | 9 | Distributed in DC and the states of MA, and OK. | Terminated | January 10, 2014 | 611 | 9 | Computer
Z-1956-2013 | 65743 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | StealthStation S7 Framelink Software version 5.4. |  |  | July 15, 2013 | 2013 | Medtronic Navigation, Inc. | Medtronic Navigation, Inc. has initiated a field correction to their Framelink software version 5.4 which can be used on Medtronic Navigation StealthStation 57 and i7 systems or a Medtronic Planning Station due to the potential of discrepancy between the exam image and the displayed measurement or overlay under certain conditions of use. | Software design | A Field Safety Notice was sent on 7/15/13 to users informing them that the Framelink software version 5.4 has been revised to correct the software anomalies described in this letter. A new version will be installed on their system(s) by a Medtronic Navigation representative. Until the software upgrade has been installed on their system, do not use high resolution examinations (greater than 256 x 256). If users are not sure of the exam resolution being used, firm recommend contacting the hospital imaging department or local Medtronic Navigation representative for assistance. Medtronic representatives will visit all identified hospitals and upgrade all systems to version 5.4.1. This will account for all consignees. |  | 572 | 572 | Worldwide Distribution. | Terminated | January 29, 2014 | 198 | 572 | Computer
Z-1931-2013 | 65786 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MHI-TM2000 Linear Accelerator System, Software Version 3.0.0 revision 13 or earlier; Proprietary and Trade Name: MHI-TM2000/ VERO (TM); Common Name: Medical Linear Accelerator. |  |  | June 21, 2013 | 2013 | MITSUBISHI HEAVY INDUSTRIES, LTD., HIROSHIMA MACHINERY WORK | Software anomaly: Mitsubishi Heavy Industries, Ltd. became aware of a potential safety issue associated with the VERO (TM) Linear Accelerator System. The VERO (TM) may save treatment delivery record into wrong location under certain conditions. | Software Design Change | The firm, Mitsubishi Heavy Industries, Ltd. (MHI), sent an "URGENT FIELD SAFETY NOTICE" letter dated June 21, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to do the following: If any minor failure occurs, please always press [ALARM RESET] button on operator console first; and if [EXIT] button is clicked after a minor failure occurs, please click [YES] on the pop up dialog to finish treatment completely. MHI is developing a correction for this issue. Tentatively planned availability date: End of October 2013. The correction will be free of charge and a MHI Customer Service representative will contact affected customers to schedule its installation as the correction becomes available.If you need any further information or support concerning this issue, please feel free to contact our Customer Service representative: VERO GmbH, Kapellenstrasse 12, 8566 Feldkirchen Germany, at +49-89-99-1568-44 or email: franz.gum@vero-sbrt.com. |  | 14 | 14 (2 in US) | Worldwide distribution: US (nationwide) including states of TX and FL; and countries of: Belgium, Italy, Japan and Korea. | Terminated | March 23, 2017 | 1371 | 14 | Computer
Z-1930-2013 | 65843 | No | LPP | Cutaneous oxygen (PcO2) monitor. | Anesthesiology | Anesthesiology | 510(K) Exempt | TCM4 series Base Units, and spare part Power Supply, Radiometer Medical ApS, Akandevej 21, BrÂżnshÂżj, Denmark, www.radiometer.com. The devices are intended for continuous monitoring of carbon dioxide and oxygen partial pressures, oxygen saturation, and pulse rate. |  |  | November 27, 2012 | 2012 | Radiometer America Inc | RADIOMETER became aware that the TCM4 series Base Unit does not comply with IEC 60601-1 Sub Clause 15B. | Device Design | On 11/30/2012 the firm sent an "Urgent Field Safety Notice" to their customers. The notification explained the issue and provided recommended actions. Customers may contact 1-800-736-0600 option 4 with questions. |  | 5874 | 5,874 units | Worldwide distribution: US (nationwide) AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, TX, UT, VA, VT, WA, WI, WV &amp; WY; and country of: Canada. | Terminated | June 16, 2014 | 566 | 5874 | Computer
Z-1933-2013 | 65653 | No | FSM | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Indus Invue Screw Caddy, Model: SI70097 |  |  | June 21, 2013 | 2013 | SpineFrontier, Inc. | Difficulty removing screws from the Invue caddies. | Device Design | SpineFrontier sent a Urgent Invue Caddy Advisory letter dated July 10, 2013, to all affected customers. Customers were advised the removal of the Invue Screw Caddy, SI70097 from the field for modification. Customers were asked to return the parts to SpineFrontier and to complete an Acknowledgement Form indicating that they have read the notice and will initiate the immediate return of the parts noted above.Please feel free to contact us at (978) 232-3990.On 8/21/2013 the recall was expanded. |  | 13 | 13 | Worldwide Distribution - USA, including the states of CT, VA, MA, FL, TX, CO and MS. and the country of Jamaica. | Terminated | January 11, 2016 | 934 | 13 | Computer
Z-1926-2013 | 64984 | No | HXX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Rev E PedFuse Return Screw Inserter, Part Number: 11-80030 |  |  | March 14, 2013 | 2013 | SpineFrontier, Inc. | Inserters have a set screw that fixes a collar to the inserter. The screw set can loosen and fall out of the inserter which allows the collar to slide off the proximal end of the inserter. When the collar is not in place the locking sleeve can also slide off the inserter. | Device Design | Spine Frontier sent an email dated December 10, 2012, to the affected customers that a Pedfuse Inserter with part number PM3979 needed to be returned.Further questions please call (978) 232-3990. |  | 8 | 8 | Distribution only in Texas. | Terminated | June 03, 2015 | 811 | 8 | Computer
Z-1029-2013 | 64393 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Ingenuity TF PET/CT. |  |  | December 20, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Issues have been detected in the Ingenuity TF PET/CT currently running software version 4.0.60605.0 that if they were to re-occur, could cause a risk for patients, users or service personnel, including inaccurate standard uptake values (SUV). Philips Medical Systems (Cleveland), Inc. is issuing a software update to address the issues. | Software design | Philips Healthcare sent an "URGENT-MEDICAL DEVICE CORRECTION" letter dated December 19, 2012 to the affected customer. The letter identified the product, problem, and actions to be taken by the customer. A Philips Service Engineer will contact the customer for implementation of the software update on the affected system. |  | 1 | 1 | Distributed to one customer in Ohio. | Terminated | June 18, 2013 | 180 | 1 | Computer
Z-1907-2013 | 65470 | No | LTW | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | bioMerieux Vitek 2 Systems Software version 6.01 software, used with the VITEK 2 automated system. Kit, V2S 6.01 RP5800-WES7 Media. Part number: 413861. |  |  | June 14, 2013 | 2013 | Biomerieux Inc | A product issue has been identified for customers using software version 6.01 and are interfaced with a Laboratory Information System (LIS). During the update of software from version 5.xx to 6.01, the Bid-directional Computer interface (BCI) antibiotic codes are not updated properly for Cefovecin, Prulifloxacin and Ceftaroline. For the referenced antibiotics, the firm code in the reference table | Software change control | The firm, bioMerieux, sent an "Urgent Product Correction Notice" letter dated June 14, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure the Urgent Product Correction Notice and the Instructions for modifying BCI Translation Codes are distributed to all appropriate personnel in your organization; follow the instructions in the attached file (Instructions for modifying BCI Translation Codes) to resolve the discrepancy associated with Cefovecin, Prulifloxacin and Ceftaroline in the translation table and complete and return the Attachment A: Acknowledgement Form via Fax to: Attn: RA North American ComOps at (919)287-2542.If you require additional assistance or have any questions, please contact your local bioMerieux Clinical Customer Service organization at (800) 682-2666 or Industrial Customer Service organization at (800) 634-7656. |  | 12384 | 12,384 total kits | Worldwide distribution: US (nationwide) and countries of: Argentina, Australia, Austria, Bahrain, Belarus, Benelux, Brazil, Bulgaria, Canada, Chile China, Costa Rica, Croatia, Czech Republic, Dominican Republic, Dutch Antilles, Ecuador, Egypt, France, Germany, Hong Kong, Hungary, India, Indonesia, Israel, Italy, Japan, Kazakhstan, Korea, Kuwait, Libya, Lithuania, Macedonia, Malaysia, Malta, Mexico, Morocco, Myanmar, Netherlands, Nicaragua, Nigeria, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Romania, Russia, Serbia, Singapore, Slovenia, Spain, Sweden, Switzerland, South Africa, Thailand, Tunisia, United Arab Emirates, United Kingdom, Uruguay, and Vietnam. | Terminated | June 01, 2015 | 717 | 12384 | Computer
Z-0834-2013 | 64244 | No | KPA | Automated slide stainer. | Pathology | Pathology | 510(K) Exempt | VENTANA (a member of the Roche Group) DISCOVERY ULTRA; REF/MU 750-601; Research Use Only |  |  | September 28, 2012 | 2012 | Ventana Medical Systems Inc | Potential for a waste fluid overflow condition to accur while running decontamination cycles on the BenchMark ULTRA automated staining platform or the DISCOVERY ULTRA automated staining platform.The waste fluid drainage system may become restricted by build-up of debris in the tubing and in-line filter, thereby inhibiting optimal flow rates. This type of overflow event can lead to migration of flu | Device Design | The firm, Roche sent a "Technical Bulletin" and an letter dated October 3, 2012 to its customers. The letter and bulletin described the product, problem and actions to be taken. The customers were informed that Ventana is aggressively working on solutions to this issue. Until mitigations to the overflow condition are available for implementation; inform all personnel in your organization that work with the BenchMark ULTRA or DISCOVERY ULTRA instrument about the potential for fluid overflow to occur during decontamination cycles; follow the instructions provided in the letter to avoid the documented failure mode and follow the instructions provided in the technical bulletin.Contact your local support center with any questions or concerns or call 520-877-7191. |  | 54 | 54 instruments | Worldwide distribution: USA (nationwide) and country of: Germany. | Terminated | February 24, 2016 | 1244 | 1789 | Computer
Z-1881-2013 | 65162 | No | JLW | Thyroid stimulating hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA Centaur Systems TSH3 Ultra ReadyPack ((REF 06491072 - 100 Tests, REF 06491080 - 500 Tests, REF 04862625 - 2500T). |  |  | May 10, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc | A rare variant of TSH, identified in a small cluster of patients, is not detected. | Device Design | On May 10, 2013, an Important Safety Information notice was sent to all affected Siemens Healthcare Diagnostic customers in the United States. On May 10, 2013, an Urgent Field Safety Notice was sent to all customers outside of the United States. These notices inform the customer that there is a rare variant of TSH, identified in a small cluster of patients that is not detected by some Siemens assays. This notice informs customers that they should continue to use the affected TSH assays as the observed rate of occurrence during a 30-month time period was 0.6 X 10 -7 and there are strong mitigating factors. These mitigating factors include co-interpretation of TSH results in conjunction with thyroxine and TS, patient history and clinical signs and symptoms. Although Siemens believes a look back is not required as the expected frequency of the variant is extremely rare, these notices state that this decision should be made by the medical director. |  | 172913 | 172913 | Worldwide Distribution-USA (nationwide) and the countries of Canada, AE, AR, AT, AU, BA, BD, BE, BG, BO, BR, BY, CH, CL, CM, CN, CO, CR, CY, CZ, DE, DK, DO, DZ, EC, EE, EG, ES, FI, FR, GB, GR, GT, HK, HN, HR, HU, ID, IE, IL, IN, IR, IS, IT, JO, JP, KA, KR, KW, KZ, LB, LK, LT, LV, LY, MA, MK, MT, MX, MY, NI, NL, NO, NP, OM, PA, PE, PH, PK, PL, PT, PY, QA, RO, RU, SA, SE, SG, SK, SV, SY, TH, TM, TN, TR, TW, UA, UY, UZ, VE, VN, XS, and ZA. | Open, Classified |  |  | 330020 | Computer
Z-1033-2013 | 64332 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access 2 Immunoassay System; P/N 107405S 2/12 |  |  | February 12, 2013 | 2013 | Beckman Coulter Inc. | Beckman Coulter is initiating a recall due to a defective newly released aluminum incubator belt pulley that may rub against the light shield, which causes the formation of debris. Debris that is caused by the incubator belt pulley rubbing against the light shield can fall into the reaction vessels (RVs) and present a risk of erroneous test results. | Device Design | Beckman Coulter sent an "Urgent Product Correction" letter dated February 12, 2013 via mail and email to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to place your Access system in the Ready or Not Ready mode. Lift the cover (as you do to access the probes) and look at the left side of the incubator unit under the peripump by the aspirate probe, if you see a silver pulley, immediately discontinue use and contact Technical Support also check for QC failures and system check failures; If you see a white pulley, this is the correct part and you can resume normal activity. Beckman will replace the metal pulley with the plastic pulley if you verified the your instrument has the silver pulley. The customers were also instructed to share this information with staff; and complete and return enclosed Response Form within 10 days.If you have any questions regarding this notice, please contact Technical support from the website: http://www.beckmancoulter.com or by phone at: 1-800-854-3633 in the United States and Canada. |  | 3500 | 3500 units | Worldwide distribution: USA (nationwide) and countries of: Aurstria, Belgium, Bulgaria, Canada, Croatia, Czech Republic, Denmark, Djibouti, Finland, France, Germany, Greece, Hungary, Ireland, Italy, Latvia, Lithuania, Netherlands, Poland, Portugal, Romania, Slovakia, Slovenia, Spain, Sweden, Switzerland, Turkey and United Kingdom. | Terminated | December 23, 2013 | 314 | 10236 | Computer
Z-1851-2013 | 65689 | No | BRZ | Intravascular administration set. | General Hospital | General Hospital | 510(k) | List No. 14203-28; BLOOD SET; 200 Micron Filter, 80 Inch, Non-DEHP; For Blood Product Administration; Made in Costa Rica; Product Inquiries should be directed to Hospira., Inc., Lake Forrest, IL 60045 USA |  |  | April 01, 2013 | 2013 | Hospira Inc. | It was discovered that there was a potential for the piercing pin on certain Hospira blood sets to puncture the outer wall of blood bags during insertion of the pin into the blood bag. | Device Design | Hospira sent an Important Safety Information letter dated April 1, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to exercise EXTREME CAUTION when using Hospira blood bags and to make sure that all instructions for use are included with the blood bag and their facility's protocol for spiking blood bags are completely followed in order to minimize the possibility of puncturing the outer wall of the blood bag. Customers with issues were instructed to call Hospira at 1-800-441-4100 or email ProductComplaintsPP@hospira.com. Customers were asked to complete the attached replay formand return it to the fax number on the form. If customers distributed the product further they were instructed to notify their accounts that received the affected product and ask them to contact Stericycle at 1-866-608-3936 to receive a reply form. Customers with questions were instructed to contact Hospira at 1-800-441-4100.For questions regarding this recall call 224-212-2000. |  | 1152 | 1,152 units | Worldwide Distribution - USA (nationwide) and Puerto Rico and Internationally to Costa Rica. | Terminated | May 24, 2017 | 1514 | 701870 | Computer
Z-1875-2013 | 65605 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Mobiletta Mira |  |  | May 02, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Siemens issued a Field Safety Alert about the potential risk to users when operating the Mobilett Mira. When a Mobilett Mira is regularly exposed to extensive mechanical strain (driven over high thresholds for doors or elevators), it is possible that one of its main wheel shafts could break because of material fatigue, resulting in the chassis falling on the ground on the affected wheel side. This | Device Design | Siemens sent a Field Safety Notice dated May 2. 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.Customers were instructed to :1) Never step or sit on the unit or otherwise put substantial weight on it.2) Avoid moving the unit over high thresholds for doors or elevators where possible, or do so carefully and slowly. 3) Make sure that no person had his/her extremities under the chassis, also when the unit is not moving. Update instructions were also released to alert customers of the potential risk and to provide a hardware change which is intended to improve corrosion resistance.An acknowledgement of receipt form was also sent out for the field safety notice and requested to be returned to Siemens.For further questions please call (888) 826-9702. |  | 39 | 39 | US Distribution including the states of IA, SD, NC, WI, OH, NE, TX, KS, ID, MI, GA, AR and IL. | Terminated | February 24, 2015 | 663 | 39 | Computer
Z-1873-2013 | 63605 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Beckman Coulter AU5800 Clinical Chemistry Analyzer, All Serial Numbers. |  |  | October 26, 2012 | 2012 | Beckman Coulter Inc. | Beckman Coulter is recalling the AU8500 Clinical Chemistry Analyzer because there is no labeling on the diluted detergent tank to show the proper disposition of the on/off valve. | Software design | A recall notification letter was forwarded to all their customers on 11/1/12 who purchased the AU5800 Clinical Chemistry Analyzer. The customers were informed of the problems identified and the actions to be taken. Customers are instructed to complete and return the enclosed Reponse form within 10 days. Customers with questions are instructed to call (800) 854-3633 or contact their local BCI Representative. |  | 338 | 338 units (11 units in the US) | Worldwide Distribution-USA (nationwide) and Puerto Rico and the countries of Germany, China, Austria, Brazil, Italy, Argentina, Korea, Vietnam, Hungary, Ireland, France, Spain, Turkey, Switzerland, Israel, Australia, India, United Kingdom, Netherlands, Malaysia, Russian Federation, Japan, Taiwan, and Canada | Terminated | October 18, 2013 | 357 | 338 | Computer
Z-1871-2013 | 64985 | No | FZX | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Drill Guide - Chameleon FacetFuse Fixation System, Model: 11-00028 Rev. A |  |  | November 09, 2012 | 2012 | SpineFrontier, Inc. | Prior to use in surgery, the Drill Guides were presenting noticeable degrees of denting and/or bending. As a result, drills and taps could not be directed down guides during surgery due to mechanical interference. | Device Design | Spine Frontier sent an Advisory Notice dated November 9, 2012, via email to all affected customers. Customers were advised that their FacetFuse Drill Guides need to be replaced. Customers were instructed to write their number ( included in the email) on the outside of the package and provide us with the tracking information. A replacement part has been sent out to you.If you have any further questions please call (978) 232-3990. |  | 13 | 13 drill guides | US Distribution including Puerto Rico and the states of CO, KS, MA and TX. | Terminated | June 05, 2015 | 938 | 13 | Computer
Z-1874-2013 | 65778 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | MRidium 1145 Dose Reduction System (DERS) drug library kit. (This is an option for the MRidium 3860+MR Infusion Pump System). |  |  | July 01, 2013 | 2013 | Iradimed Corporation | The Dose Error Reduction System (DERS) can indicate an incorrect recommended value for the rate via a specific key sequence during the initial infusion setup. This can result in a risk of over-infusion or under-infusion if the infusion is started with this rate value. | Software Manufacturing/Software Deployment | The firm, Iradimed Corporation, notified by phone and sent an "URGENT MEDICAL DEVICE RECALL NOTICE" via email dated July 1, 2013 all consignees/customers. The notice described the product, problem and action to be taken. The customers were instructed to identify the MRidium 3860+ MR Infusion Pumps which have the 1145 DERS Library card (AM05) installed. If these are not installed, no further action is required at this time; however, if they are installed-remove the 1145 DERS Library Card from the 3860 Pump using the instructions provided. Product will be reconditioned by installing a software upgrade that corrects this problem. The updated software (version 3.5.1) can be seen during the 3860+ Pump's power on cycle.Should your users have any questions following these instructions, please contact Iradimed Customer Service; 866-677-8022 (toll free) or 407-677-8022 ext. 100 or via e-mail - recallsupport@iradimed.com |  | 176 | 176 individuals kits | Worldwide distribution: US (nationwide) including states of: AL, CA, CO, FL, GA, IL, KY, LA, MD, MA, MI, MN, NJ, NY, NC, OH, PA, RI, TN, TX, UT, VA, and WA; and countries of: Australia, Austria, Denmark, Epsom, France, Germany, London, Netherlands, Saudia Arabia, and United Kingdom. | Terminated | July 13, 2015 | 742 | 176 | Computer
Z-1030-2013 | 64665 | No | GYC | Cortical electrode. | Neurology | Neurology | 510(k) | Integra Ojemann Cortical Stimulator; Product / Catalogue No: OCS2. |  |  | March 18, 2013 | 2013 | Integra Burlington MA, Inc. | If the headphone jack is in use during a surgical procedure with the OCS2 and a non-intended voltage (such as static electricity) comes in contact with the outer case of the unit, that voltage could be transmitted to the patient and could be a potential source of injury. | Device Design | Integra initiated a voluntary recall on March 18, 2013 by providing written notification to consignees either by traceable courier service or traceable emails regarding the correction that Integra records indicate have an OCS2 since it was introduced to the market in 2007. The consignee notification:Advised of the nature of the issue and to immediately stop using the headphone jack while using the OCS2. Requested identification by S/N any OCS2's they have, and complete / return the Acknowledgment Return Form.Advised that Integra will contact the consignee to schedule a time / date to correct the OCS2's they have identified. Questions, contact service hotline at 1-888-772-7378. |  | 371 | 371 units | Worldwide Distribution-USA (nationwide) and the countries of Austria, Belgium, Czech Republic, Belgium, Denmark, Spain, Great Britain, Hungary, Ireland, Italy, Netherlands, Poland, Morocco, Netherlands, Poland, Portugal, Saudia Arabia, Sweden, Turkey, and South Africa. | Terminated | February 25, 2014 | 344 | 371 | Computer
Z-1038-2013 | 64675 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | SIEMENS brand COHERENCE RT Therapist versions RTT MR1.5, RTT 2.1/2.1A and RTT 2.2/ 2.2B; Software release for COHERENCE Therapist 2.3 for Primus and Oncor Linear Accelerators; |  |  | March 05, 0013 | 13 | Siemens Medical Solutions USA, Inc | Siemen's issued a Customer Information Letter under Update Instruction TH016/12S to alert all its affected customers of the software issues. The software release is a result of the final system testing addressing multiple performance defects and several safety issues that were not considered critical (for immediate correction) for the safe use of the device. | Software design (manufacturing process) | Siemens Medical Solutions sent a Customer Information letter via mail to all affected customers. The letter identified the affected product, problem and provide instructions on the correction. Customers were instructed to include the Customer Information letter in the Owner Manual chapter "Safety Advisory Letters".9-16-13 UPDATE: - Control Console SW 9.2.18 was released to support the rol-out10/18/2013; Distribution of update instructions for TH017/13S and TH018/13S - Distribtuion of the Control Console SW 9.2.18 released |  | 621 | 621 units (foreign account) in total | Worldwide Distribution | Terminated | August 28, 2014 | 731026 | 621 | Computer
Z-1848-2013 | 65491 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Siemens Luminos dRF Fluoroscopic X-Ray system |  |  | May 07, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Siemens became aware of an unintended behavior when using the Luminos dRF with software version VD10A. The possibility exists that a software error may occur causing the table to tilt into an upright position and may collide with the floor. Four consignees were notified of this recall by Update Instruction XP067/12/S -- all four units have been corrected. | Software design | Siemens initiated their correction of this product in May, 2013 by sending an Safety Update Instruction XP067/12S to resolve the software error to all affected customers. All units have been corrected.Further questions please call (610)-219-6300. |  | 4 | 4 | US Distribution including the states of CA, IN, NJ and OK. | Terminated | December 24, 2013 | 231 | 4 | Computer
Z-1973-2013 | 65381 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | TerumoÂż Advanced Perfusion System 1 (APS1, System 1) |  |  | July 26, 2013 | 2013 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (TCVS) has become aware that an Advanced PerfusionSystem 1 (APS1) Centrifugal Pump 'Service Pump' message may occur due to user inducedmotor movement. In this instance, the APS1 will identify the movement (as it isdesigned) and display a Service Pump message error. The user might interpret this as apump which requires service rather than a message displayed d | Device Design | Terumo Cardiovascular Systems (TCVS) notified their consignees of the Urgent Medical Device Recall letter dated 7/26/13 and was mailed 7/29/13. The notification provided reason for the device correction, affected product, how the device will be corrected, and the potential health hazard associated with a user's inaccurate interpretation of the service pump error message. Instructions for consignee responses to the recall notice communication may be faxed to 734-741-6149 or emailed to the address on the return form. If you have any questions or concerns, call Terumo Cardiovascular Systems (TCVS) Customer Service at 1-800-521-2818 M-F, 8am-6pm ET. Terumo Cardiovascular Systems (TCVS) will provide an addendum to the APS1 Manual to their consignees. The addendum will be placed at the beginning of Chapter 6, Centrifugal Control Unit. |  | 1657 | 1657 total | Worlwide Distribution - US Nationwide: AL,AR,AZ, CA, CO, CT, DC, DE, FL,GA, HI, IA, ID, IL, IN, KY, LA, MA , MD, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WI, WV and the countries of: AUSTRALIA,UNITED ARAB EMIRATES (UAE),Hong Kong, onesia,Singapore,Taiwan,Thailand, Philippines,CHILE, Malaysia,USA,South Korea,BELGIUM,Japan,Mexico,CANADA | Open, Classified |  |  | 1657 | Computer
Z-1015-2013 | 53991 | No | BRY | Anesthetic cabinet, table, or tray. | Anesthesiology | Anesthesiology | 510(K) Exempt | Pyxis Anesthesia System Model 2000 (PAS 2000) |  |  | November 20, 2009 | 2009 | CareFusion Corporation | The recall was initiated because Carefusion has confirmed that when the PAS 2000 drawers due to inaccessability of drawers caused by system down time or lock-up. | Device Design | Carefusion sent a Safety Alert Notification letter dated January 13, 2010 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to ensure back-up procedures were in place in case of system down-time or lock-up, including review of the back-up procedures with staff. The letter included an update to the PAS 2000 User Guide. The letterstates that a CareFusion field representative will be contacting customers to make arrangements to apply warning labels to the PAS 2000 electronics doors. Customers were instructed to complete and return the Response Card included. For questions contact your local field support representative ot the Advocacy department at customer.advocacy@carefusion.com. |  | 2637 | 2637 | Worldwide Distribution - USA (Nationwide) including CANADA. | Terminated | March 28, 2013 | 1224 | 2637 | Computer
Z-1002-2013 | 64461 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Vitrea Enterprise Suite, Vitrea, VitreaAdvanced, VitreaCore (VitalConnect) and Vitrea fX prior to version 6.3.3, except 6.1.6, 6.2.3 and 6.3 |  |  | February 05, 2013 | 2013 | Vital Images, Inc. | Vital Images has found a potential error in measurement generated through the software which affects all versions of Vitrea Enterprise Suite, Vitrea, VitreaAdvanced, VitreaCore (VitalConnect) and Vitrea fX prior to version 6.3.3, except 6.1.6, 6.2.3 and 6.3.This error occurs when images are rotated at the time of reconstruction by the scanner in non-90 degree increments of the transverse/axia | Software design | Vital Images sent an "Urgent Field Safety Notice' letter dated February 5, 2013, to all affected customers. The letter identified the probelm the product and the action needed to be taken by the customer. Please immediately inform all potential users of Vitrea of the information contained in this letter. We recommend that all measurements associated with the error described in this letter be verified against results from other technologies.You will be contacted by Vital Images Customer Support or by the customer support department of your distributor to arrange for installation of this patch on your system. We recommend that you implement the patch. If you have an earlier version, Vital Images Customer Support or your distributor will contact you to discuss your options for replacement.Please sign this Notice in the space below to verify that you have received and understood it, and return your signed Notice by email to fieldnotices@vitalimages.com.If you have questions, direct customers of Vital Images should contact Vital Images Customer Support at 1-800-208-3005 or by e-mail to support@vitalimages.com; customers through a distributor of Vital Images products should contact that distributor.Updated 5/28/2013:A second consignee Vital Images, Inc "Urgent Field safety Notice" letter dated March 26, 2013 was sent to consignees on 3/25/13. The follow-up letter was sent via e-mail and mail to 4,616 already-notified consginees, who had received affected versions 6.3.0, 6.2.3, or 6.1.6. The letter described the problem and the affected product being recalled and provided a number of steps to follow which included to retun the Effectiveness Check form. |  | 38245 | 7649 (Expnaded 5/28/13 to 7654) | Worldwide Distribution - All states in continental USA including DC, PR and the countries of ALGERIA, ARGENTINA, ANGOLA, ARMENIA, AUSTRALIA, AUSTRIA, BENGLADESH, BELARUS, BELGIUM, BOLIVIA, BOTSWANA, BRAZIL BULAGARIA, CAMBODIA, CANADA, CENTRAL AFRICAN REPUBLIC, CHAD, CHILE, CHINA , COLOMBIA, CONGO, CROATIA, CYRPRUS, CZECH REPUBLIC, DENMARK, DJIBOUTI, DOMINICAN REPUBLIC, ECUADOR, EGYPT, ESTONIA, ETHIOPIA, FINLAND, FRANCE, GREAT BRITAIN, GEORGIA, GERMANY, GHANA, GREECE, GUAM, HONDURAS, HONG KONG, HUNGARY, ICELAND, INDIA, INDONESIA, IRAQ, IRELAND, ISRAEL, ITALY, IVORY COAST, JAPAN, JORDAN, KAZAKHASTAN, KOREA, LEBANON, LITHUANIA, MACEDONIA, MALYASIA, MALI, MAURITIUS, MEXICO, MALDOVA, MOSCOW, NYNMAR, NAMBIA, NAMIRGIA, NEPAL, NETHERLANDS, NEW ZEALAND, NIGER, NIGERIA, NORTHERN IRELAND, NORWAY, OMAN, PAKISTAN, PANAMA, PARAGUAY, PERU, PHILLIPPINES, POLAND, PORTUGAL, QATAR, ROMANIA, RUSSIA, RUSSIAN FEDERATION, SAUDI ARABIA, SENEGAL, SERBIA, SEBIA AND MONROBIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOMALIA, SOUTH AFRICA, SPAIN, SRI LANKA, SUDAN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, THE NETHERLANDS, TIBET, TUNISIA, TURKEY, UKRAINE, UNITED ARAB EMIRATES, UNITED KINGDOM, UZBERKISTAN, VENEZUELA, VIETNAM and YEMEN. | Terminated | May 08, 2015 | 822 | 38245 | Computer
Z-2056-2013 | 65824 | No | HBG | Manual cranial drills, burrs, trephines, and their accessories | Neurology | Neurology | 510(k) | Integra Leyla Ball Joint Clamp, also as included in Leyla Table Mounting Hardware. |  |  | July 17, 2013 | 2013 | Integra Limited | As a result of complaint investigations it was identified that there is a potential that the Leyla Ball Joint Clamp (R2383) may not provide the level of stability needed to maintain retraction in all cases. | Device Design | Integra sent an Urgent Medical Device letter dated July 17, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were requested to review your inventory and determine if you have any of the affected products. If so, stop using them immediately and remove them from your service. When your form is received, Customer Service will you an RMA number and directions for returning the products.Should you have any questions regarding these instructions, please contact Integra Customer Service at 1-800-997-4868. |  | 65 | 65 | Worldwide Distribution - USA including Puerto Rico and the states of AZ, CA, FL, GA, IL, IN, KY, MA, MI, MS, NE, NM, NY, OH, PA, TN, TX, UT, VA, WA, WI and WV., and the countries of Spain, Singapore, South Africa, Canada, Libya,Pakistan and Cayman Islands. | Terminated | October 01, 2014 | 441 | 65 | Computer
Z-2114-2013 | 65852 | No | GKL | Automated cell counter. | Hematology | Hematology | 510(k) | ADVIA 120 Hematology Analyzer, Catalog Numbers SMN 10360958 and 10360959 . |  |  | June 26, 2013 | 2013 | Siemens Healthcare Diagnostics | The ADVIA 120 optics cover or hood is located on the very top of the ADVIA 120 analyzer. When the optics cover is raised, it is supported by two gas struts (SMN 10309266) attached to the cover on the left and right sides. Over time, it is possible that these struts may lose their effectiveness and fail to support the optics cover in the open position. This could cause the cover to fall during m | Equipment maintenance | Siemens sent an Urgent Device Correction Letter dated June 26, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed of the ADVIA 120 system displays a problem when raising and supporting the optics cover, immediately contact your service provider to arrange to have the cover struts serviced. If the optics cover will not stay in the up position, as required, do not perform maintenance procedures on the optics area until the hood struts have been replaced. Siemens field support will be proactively replacing the optics cover hood struts on a regular preventive maintenance cycle based on the age of the system.Please forward this notification to whomever you may have distributed this productWe apologize for the inconvenience that this situation has caused. Thank you for your patience and continued support.This response form is to confirm receipt of the enclosed Siemens Healthcare Diagnostics Urgent Device Correction dated June 2013 regarding Optics Cover Struts (SMN 10309266).Please read each question and indicate the appropriate answer. Fax this completed form to Siemens Healthcare diagnostics at the fax number indicated at the bottom of this page.PLEASE FAX THIS COMPLETED FORM TO THE TECHNICAL SOLUTIONS CENTER AT (302) 631-7597. |  | 4319 | Domestic: 950 units; Foreign: 3369 units | Worldwide Distribution - USA (nationwide) and the countries of Austria, Australia, Belgium, Bosnia - Herzegovina, China, Croatia, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, Hungary, Iceland, India, Indonesia, Ireland, Iraq, Israel, Italy, Japan, Kenya, Republic of Korea, Kuwait, Lebanon, Lithuania, Luxembourg, Martinique, Malaysia, Netherlands, Norway, New Zealand, Philippines, Poland, Portugal, Romania, Russian Republic, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom and Vietnam. | Open, Classified |  |  | 4319 | Computer
Z-2109-2013 | 65886 | No | FMK | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ACCU-CHECK FastClix Lancing Device. It is intended for use in conjunction with a blood glucose monitoring system. |  |  | August 19, 2013 | 2013 | Roche Diagnostics Operations, Inc. | Internal Roche inspections have revealed that after the lancet is used for a finger stick, the needle may not fully retract back within the lancet drum or back into the cap. This allows the needle to protrude inside the cap or outside of the cap of the ACCU-CHEK FastClix lancing device. | Equipment maintenance | On 8/19/ 2013, Roche sent Distributor/Retailer URGENT MEDICAL DEVICE CORRECTION notifications which described the recall issue and health consequences, provided product description, and provided actions required of consignees. Roche sent a CUSTOMER BULLETIN to their direct customers. URGENT MEDICAL DEVICE CORRECTION faxback form for consignee acknowledgement of recall is to be faxed to 1-888-240-2804. Consignees may contact the ACCU-CHEK Customer Care Service Center, 24 hours a day, seven days a week at 1-800-440-3638 if they have questions about the information contained in the Distributor/Retailer Letter. |  | 36332 | 36,332 | Nationwide Distribution | Terminated | May 12, 2014 | 266 | 36332 | Computer
Z-2098-2013 | 65880 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Ultrasound ViewPoint 6. |  |  | August 07, 2013 | 2013 | GE Healthcare, LLC | GE has recently become aware of a potential safety issue due to the incorrect calculation of the pulsatility index value (PIV) for venous flow measurements associated with obstetric reporting of the GE ViewPoint 6 system. The ViewPoint calculated Doppler PIV for veins could be incorrect. An incorrect calculation of the ductus venosus PIV could lead to false high risk for trisomy 13, 18, and 21, | Software Design Change | GE Healthcare sent An Urgent Medical Device Correction letter dated August 7, 2013 to all affected customers. The letter described the affected product, problem and instructed customers to discontinue use of venous PIV for making any and all clinical decisions. The letter informed customers that a GE service engineer will install a correction at no cost. For questions contact your local GE Healthcare Representative or call 877-644-3114 (US) and 49 (0)8153-9311-11(Europe/Asia) |  | 225 | 225 | Worldwide Distribution - USA Nationwide including Puerto Rico and the countries of: ANDORRA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BRAZIL, BULGARIA, CANADA, CHILE, COLOMBIA, CZECH REPUBLIC, ESTONIA, FINLAND, GERMANY, HONG KONG, INDIA, INDONESIA, IRAQ, ITALY, JAPAN, KUWAIT, MALAYSIA, Norway, OMAN, PALESTINIAN TERRITORY, OCCUPIED, POLAND, PORTUGAL, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TURKEY, and UNITED KINGDOM. | Terminated | April 25, 2014 | 261 | 225 | Computer
Z-0965-2013 | 64546 | Yes | LWS |  |  | Cardiovascular | PMA | Boston Scientific/Cameron Health SQ-RX Model 1010 subcutaneous pulse generators. The S-ICD System is intended to provide therapy for the treatment of life-threatening ventricular tachyarrhythmias. |  |  | March 01, 2013 | 2013 | Boston Scientific CRM Corp | An internal protective fuse can be unintentionally activated while the device is charging its capacitors for shock delivery or induction. Should this occur, the defibrillator would not be able to deliver therapy or communicate with the Q-TECH Model 2020 programmer, and would be unable to emit tones or otherwise respond to magnet application. | Software design | Consignees were sent a Boston Scientific "Important Medical Device Information" letter dated March 1, 2012. The letter was addressed to "Dear Doctor". The letter described the product involved in the recall, the problem and provided Recommendations/Actions. US customers may contact Technical Services for more information at 800-227-3422. |  | 2090 | 2090 (570 USA; 1520 OUS) | Worldwide Distribution, USA including the states of AZ, CA, CO, DC, GA, IL, MD, MA, MN, NJ, NY, NC, OH, PA, SC, VA, WA and the countries of AUSTRIA, BELGIUM, CANADA, CZECH REPUBLIC, DENMARK, FRANCE, GERMANY, IRELAND, ITALY, NETHERLANDS, NEW ZEALAND, PORTUGAL, SLOVAKIA, SPAIN, SWEDEN, SWITZERLAND, and the UNITED KINGDOM. | Terminated | May 23, 2013 | 83 | 2090 | Computer
Z-2086-2013 | 65672 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Illico Ti Cannulated Polyaxial Screw 7.5mm X 40mm. The Illico MIS Posterior Fixation System is intended to facilitate the surgical correction of noncervical spinal deformities by providing temporary internal fixation and stabilization during bone graft healing and/or fusion mass development. |  |  | June 28, 2013 | 2013 | Alphatec Spine, Inc. | The firm initiated this recall of certain lots of Illico Ti Cannulated Polyaxial Screws 7.5mm x 40mm implants, due to an off-center cannulation hole at the proximal end of the bone screw. There is a potential risk that when the screw is mated to a polyaxial screwdriver, the guide wire may not pass through the screwdriver. | Equipment maintenance | The firm, Alphatec Spine, sent an "URGENT-MEDICAL DEVICE RECALL NOTIFICATION" with response sheet dated June 28, 2013 via a combination of e-mail, Fedex, and telephone contact to its customers. The notification described the product, problem and actions to be taken. The customers were instructed to immediately return any and all affected lots of the illico Ti Cannulated Polyaxial Screw 7.5 mm x 40 mm implants that have been sent to them; if the device has previously been implanted, there is no further or additional risk post-implantation; immediately contact Alphatec Spine Customer Service Department - customerservice@alphatecspine.com on how to replace the implants; and complete and return the last page of the notification via mail to Alphatec Spine, Inc. 5818 El Camino Real, Carlsbad, CA 92008; email: sarrhenius@alphatecspine or fax to: (760) 431-0289.If you have any questions or need any additional information, please contact Director of Compliance at 760-494-9636 or email at dmcgrath@alphatecspine.com. |  | 333 | 333 screws | Worldwide distribution: US (nationwide) including states of:US states: AZ, CA, FL, GA, IL, MD, NJ, NY, OK, PA, TN, TX,and WA ; and countries of: Spain, France, Japan.International: Spain, France, Japan. | Terminated | October 09, 2014 | 468 | 333 | Computer
Z-2095-2013 | 65848 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | PRIMUS, ONCOR, ARTISTE Linac systems with component Tabletop, 550TxT, TT-A s/n 1001-1233 and Tabletop 550TxT, TT-S s/n 1001-1083. |  |  | July 18, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Siemens has become aware of a customer complaint concerning the rail threaded holes of the 550TxT treatment tabletop. Due to a manufacturing error, the screws holding the side rails in position on the tabletop might break off. | Device Design | Siemens released Updated Instruction containing the Medical Device Correction Field Safety Notice of Improvement letter, which was delivered by the Siemens Customer Service Group to all affected customers. The release of this UI TH002/12/S to distribute the update instruction began on July 18, 2013. This document includes completion protocol, which will be completed at each facility and returned to Siemens. A copy of the notification should be kept in the Digital Linear Accelerator System owners manual under Safety Advisory letters.UPDATE: a secnd update intruction, TH003/12/S has been released to begin the replcement of tabletops identified through the inspection carried out through TH002/12/S. |  | 179 | 179 active devices. | Worldwide Distribution - USA Nationwide in the states of NJ, IN, PA, FL, OH, MN, MO, CA, AL, NE MI, LA. and countries of:Argentina, China, South Africa, Pakistan, Spain, Switzerland, Italy, Algeria, Croatia, Turkey, Canada, Germany, France, Poland, Japan, Russian Federation. Lebanon, Egypt, Slovakia, Latvia, Kenya, Brazil, Kuwait, Saudi Arabia, Greece, Belgium, Taiwan, Portugal. | Terminated | November 24, 2014 | 494 | 179 | Computer
Z-2055-2013 | 65677 | No | KWT | Shoulder joint metal/polymer non-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | 00 4309-049-00 Trabecular Metal Reverse Shoulder Glenosphere Distractor Rx, Non sterile (sterilize before use) |  |  | August 01, 2013 | 2013 | Zimmer, Inc. | Zimmer is initiating a lot specific recall of the Trabecular Metal" Reverse Glenosphere Distractor due to the potential of the device not functioning properly should it be utilized without proper lubrication. As a result, there is a potential for the device to fail to provide impact to the Glenosphere when the trigger is pulled. In addition, there is a potential for the foot of the device to fract | Device Design | Zimmer sent a Urgent Medical Device Recall Notification letter August 1, 2013, and via E-mail and courier to all affected customers. The initial recall notice identified the reason for the recall, possible adverse risks and injuries and informed customers that a replacement will be available in the coming months. In addition, this letter will remind that the instrument is to be lubricated before each use and to inspect the foot. Customers are instructed to inspect any recalled product before and after use for cracks at the foot and ensure the foot is intact. Also ensure all moving parts are lubricated prior to each use according to the Recommendations for Care, Cleaning, Maintenance, and Sterilization Manual 97-5000-170-00. If you find a device which does not function properly or has a damaged foot, provide it to your Zimmer sales representative for return to Zimmer on a Product Experience Report (PER). Questions and concerns should be addressed to the Zimmer customer call center at 1-877-946-2761 between the hours of 8 a.m. and 5 p.m. EST. Phase 2: As devices become available for distribution, a 2nd letter will be sent to the distributors and risk managers. Distributors will be instructed to order replacement devices as a means of completing the affected instrument cases in circulation within their territories. Instructions for the return of devices will be included. |  | 771 | 771 | Worldwide Distribution - USA (nationwide) including the countries of Singapore, Brazil, Taiwan, Australia, Canada, Germany(with further distribution to UAE, Austria, Belgium, Switzerland, Czech Republic, Denmark, Spain, Finland, France, Greece, Israel, Italy, Malta, Norway, Portugal, Sweden and South Africa, | Terminated | March 02, 2015 | 578 | 771 | Computer
Z-2059-2013 | 65950 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips Easy Upgrade DR, Code No: 712086 |  |  | August 08, 2013 | 2013 | Philips Healthcare Inc. | The mirror icon on the bottom of each image could be misinterpreted as a RIGHT lead marker for side indication if the image is mirrored within a PACS system. | Software design | Philips Healthcare sent an Urgent Field Safety Notice letter dated July 24, 2013, to all affected customer. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed : When reading an image, a physician should not use the mirror icon for image side interpretation. 2. Clinical staff should follow the Instructions for Use and use lead letters to indicate body orientation of patient or equivalent electronic markers to provide Left/Right indications, during all X-ray image acquisitions. Philips will upgrade the system, which involves software and hardware upgrades of the system. The software provides a replacement icon that reduces the risk of the potential misinterpretation of the original mirror icon. Philips will also provide updated Instructions for Use for the work spot and the system (including an addendum).If you need any further information or support concerning this, please contact your local Philips representative Technical Support Line at 1-800-722-9377.We apologize for any inconvenience this may cause and trust that this information is adequately addressing any concerns you may have. |  | 89 | 89 | Worldwide Distribution - US Nationwide including the states of AZ, CA, CO, DE, FL, GA, MO, MS, NC, NH, NY, PA and TX., and the countries of Canada, Australia, China, Czech Republic, France, Germany, Greece, Italy, Netherlands, Portugal, Russia, Saudi Arabia, Slovenia, Spain, Sweden, Switzerland, Turkey, United Arab Emirates and United Kingdom. | Completed |  |  | 89 | Computer
Z-2082-2013 | 65080 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | 2008 MeDS Pump, Venofer(R) Pump |  |  | April 01, 2013 | 2013 | Fresenius Medical Care Holdings, Inc. dba Renal Solutions | Product is being recalled because the combination of the current 50 mg vial, the 2008 MeDS pump, and the 2008 MeDS pump tubing Set for CombiSet Bloodlines can lead to a situation where Venofer(R) delivery does not meet dosage accuracy requirements. This may lead to a slight over or under delivery of Venofer(R) that does not pose an acceptable risk to the patient. Please note that this recall | Device Design | Renal Solutions sent an Urgent Notification letter dated April 1, 2013 via certified mail to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to cease further use of the device. The letter informed customers that a Fresenius Medical Care Technical Service/Customer Service representative will contact customers to arrange for removal of the device. For questions contact Fresenius Medical Care Customer Service at 800-227-2572. |  | 132 | 132 | USA Nationwide Distribution in the states of CA, MA, and NY. | Terminated | March 02, 2015 | 700 | 132 | Computer
Z-2089-2013 | 65769 | No | BRY | Anesthetic cabinet, table, or tray. | Anesthesiology | Anesthesiology | 510(K) Exempt | Mc Kesson Anesthesia-Rx Model Number 8.0.2 &amp; 8.0.1 with RapiDispense |  |  | July 22, 2013 | 2013 | MCKESSON TECHNOLOGIES INC | The firm issued a clinical alert after users reported an issue where the Anesthesia-Rx system was interrupted ANF became temporarily unavailable during use. | Software design | McKesson sent a Clinical Alert letter dated July 22, 2013 to all affected customers. The letter identified the affected product, symptoms, and actions to be taken. McKesson committed to releasing 8.0.2 Update 1.5 to affected customers which will include enhanced code to prevent the unhandled excpetion errors. Additional resolution steps were included in the letter. For questions call 1-800-700-8737. |  | 106 | 106 | USA Nationwide Distribution in the states AR, MD, NJ, NY, PA, and WV. | Terminated | June 19, 2014 | 332 | 106 | Computer
Z-0971-2013 | 63973 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Specialist 2 IM Rod, Non-Sterile and associated Surgical Techniques: |  |  | February 05, 2013 | 2013 | DePuy Orthopaedics, Inc. | Potential for the IM rod to break, leaving fragments in the patient. Surgical Techniques guidance has been updated until design changes can be made to reduce likelihood of potential failure. | Device Design | DePuy Orthopaedics sent an Urgent Medical Device Correction Notice letter dated February 5, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were asked to complete and return the FAX back form to FAX 574-371-4939 or email kseppa@its.jnj.com Questions about device correction information call Post Market Surveillance, 574-372-7333, Sales force questions call Product Specialist, 574-372-7303 and Surgeon questions DePuy's Scientific Information Office at 1-888-554-2482. |  | 2600 | 4000; 2600 in USA | Worldwide Distribution - USA Nationwide. | Terminated | January 06, 2014 | 335 | 2600 | Computer
Z-2057-2013 | 65556 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes 2.4mm VA-LCP Two Column Volar Distal Radius Plate and Torque limiting attachment, 1.2Nm |  |  | May 31, 2013 | 2013 | Synthes (USA) Products LLC | The firm is initiating a Field Safety Correction for the labeling of the 2.4mm VA-LCP Two Column Volar Distal Radius Plate and 1.2 Nm Torque Limiting Attachment. The US Technique guide lists 1.2Nm TLA as an option for use with the 2.4mm Variable Angle LCP (VA-LCP) Two-Column (Standard &amp; Narrow) Volar Distal Radius Plate. The potential exists for the VA locking screw to be driven through the 2.4mm | Device Design | Synthes sent an Urgent Notice: Medical Device Labeling Correction letter dated May 31, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The firm is initiating a Field Safety Correction for the labeling of the 2.4mm VA-LCP Two Column Volar Distal Radius Plate and 1.2 Nm Torque Limiting Attachment. The US Technique guide lists 1.2Nm TLA as an option for use with the 2.4mm Variable Angle LCP (VA-LCP) Two-Column (Standard &amp; Narrow) Volar Distal Radius Plate. The potential exists for the VA locking screw to be driven through the 2.4mm VA-LCP Two Column Volar Distal Radius Plate when used with a 1.2Nm Torque Limiting Attachment. A screw that passes through the VA-LCP plate is not likely to result in a serious adverse event although should it occur, additional medical intervention during the index procedure could be needed.If you have any questions, please call (610) 719-5450, or contact your Synthes sales Consultant. |  | 119891 | 119, 891 | Nationwide Distribution including the states of AK. AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MD, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI SC, SD, TN, TX, UT, VA, VT, WA WI and WV. | Terminated | September 21, 2015 | 843 | 119891 | Computer
Z-2088-2013 | 65607 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension Vista 500 and Dimension Vista 1500 System Aliquot Well Double Dispense with software versions 3.5.1 and 3.6. |  |  | June 19, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Siemens is conducting a field correction for Dimension Vista 500 and Dimension Vista 1500 systems using software versions 3.5.1 and 3.6. Under specific rare conditions which require a system reset, there is a potential for a sample to be dispensed into an aliquot well that has previously been used for a QC or calibrator product from a Vista vial. If two samples are dispensed into the same aliquot | Software Design Change | The firm, Siemens, sent an "Urgent Medical Device Correction" letter dated June 2013 to its consignees/customers. The letter describes the product, problem and actions to be taken. The customers were instructed to take the following actions: 1) In the event that an instrument Reset is required for any reason, review the Current Activities Screen to determine if there are any pending QC or Calibration tests.2) If there are no pending QC of Calibration tests, Rest the instrument or troubleshoot the instrument per the Operators Guide.3) If there are pending QC or Calibration tests, restarting the software is necessary before resuming patient processing. Go to Operation -&gt; Shutdown, confirm Yes, then on the blue console menu screen select start Vista. Consignees/customers were also asked to complete the Field Effectiveness Check form attached to the letter via fax to Siemens Healthcare Diagnostics at (302) 621-8467, and retain the letter with laboratory records, and forward the letter to those who may have received the product.If you have any questions, contact your Siemens Customer Care Center at (800-441-9250)-Technical Solutions or your local Siemens Technical Support representative. |  | 1937 | 1937 | Worldwide distribution: US (nationwide) including PR and countries of: Australia, Austria, Belgium, Canada, Czech Republic, Denmark, Finland, France, Germany, Italy, Japan, Malaysia, Netherlands, New Zealand, Norway, Portugal, Republic Korea, Saudia Arabia, Slovakia, Slovenia, South Africa, Spain, Switzerland, and United Kingdom. | Terminated | August 27, 2014 | 434 | 1937 | Computer
Z-2083-2013 | 65773 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | BRAINLAB; FRAMELESS SRS QA TARGET POINTER |  |  | July 16, 2013 | 2013 | Brainlab AG | The Frameless SRS QA Target Pointer - Pointer Cap with the engraved cross hairs may become loose even when only a moderate force is applied to the Pointer Cap. | Device Design | BrainLab sent an "FIELD SAFETY NOTICE / PRODUCT NOTIFICATION" dated July 9, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed to discontinue use of the SRS QA Target Pointer. Consignees can instead use the Phantom Pointer feature until Brainlab can provide replacement hardware. Brainlab anticipates being able to provide all impacted facilities with replacement hardware by the end of July 2013.We sincerely apologize for any inconvenience and thank you in advance for your co-operation. If you require further clarification, please feel free to contact your local Brainlab Customer Support Representative. Customer Hotline: +49 89 99 15 68 44 or +1 800 597 5911 (for US customers) or byE-mail: support@brainlab.com. Fax Brainlab AG: + 49 89 99 15 68 33 |  | 6 | 6 units | Worldwide Distribution - US Distribution only to AR., and the countries of Australia, Germany and Japan. | Open, Classified |  |  | 6 | Computer
Z-2078-2013 | 65800 | No | HWT | Template for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | TRIAL STEM for MP(R) RECONSTRUCTION PROTHESIS, REF 99-0155/20, Stem size S-4, prox. Stem dia. 20 mm, Length 180 mm, WALDEMAR LINK GmbH &amp; CoKG, Barkhausenweg 10 . D . 22339 Hamburg |  |  | July 16, 2013 | 2013 | Waldemar Link GmbH &amp; Co. KG (Corp. Hq.) | The trial stems were manufactured using the wrong design. | Device Design | The firm notified their consignees of their product recall by parcel post on 07/16/2013. |  | 8 | 8 units | US Distribution to the state of Tennessee only. | Terminated | June 04, 2014 | 323 | 127 | Computer
Z-0996-2013 | 63621 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Medfusion Model 3500 Syringe Pump, Model #3500-415 and 3500-500. |  |  | May 14, 2010 | 2010 | Smiths Medical ASD, Inc. | An issue with the plunger rod seal was identified where the seal had the potential to migrate inside the pump case top during slow rate/ high back pressure infusions. | Device Design | On May 14th, 2010, Smiths Medical contacted the facility that received the pumps via e-mail and requested that they immediately return the pumps. The pumps were returned per Returned Goods Authorization (RGA) and exchanged for replacement pumps manufactured with plunger rod seal G6000077. |  | 6 | 6 | Distributed in the state of NH. | Terminated | April 11, 2013 | 1063 | 6 | Computer
Z-2084-2013 | 48997 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Continuous Ventilator/Continuous Ventilator with Respiratory Mechanics, Model No. V1000 |  |  | October 29, 2007 | 2007 | Respironics California Inc | The power to the Esprit Display's Backlight is interrupted, causing the GUI to flicker or go dark and making the display unreadable. | Device Design | A Medical Device Field Correction letter dated 10/26/2007 was sent to all customers who received the Esprit Ventilator. The recall letter informed the customers of the problems identified and the actions to be taken. Customers with questions are instructed to contact Respironics' Customer Service Dept. at (800) 345-6443, select Option#5, then select Option #4. |  | 3255 | 3,255 units (2302 units in US) | USA Nationwide and Worldwide to Algeria, Argentina, Australia, Bangladesh, Brazil, Cambodia, Canada, Chile, Colombia, Ecuador, Egypt, France, Germany, Guatemala, Hong Kong, India, Italy, Japan, Libya, Malaysia, Mexico, Pakistan, Peru, Philippines, Puerto Rico, Russia, South Africa, South Korea, Spain, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Venezuela, and Vietnam. | Terminated | August 26, 2013 | 2128 | 3255 | Computer
Z-2046-2013 | 65865 | No | HKI | Ophthalmic camera. | Ophthalmic | Ophthalmic | 510(k) | NAVIS-EX , Image Filing Software for NIDEK AFC Fundus camera, Ver 1.1.0 to 1.3.5. NAVIS-EX is indicated for use as file management software for ophthalmic images. |  |  | August 01, 2013 | 2013 | Nidek Inc | Under certain circumstances, the Left eye image may be saved as the Right eye image. | Software design | "Medical Device Correction Notification" letters dated August 5, 2013 were sent to consignees through their US contracted distributor - Marco Opthalmic. The letter described the product issue and provided information on the correction. Customers can contact 800-223-9044 for questions. |  | 86 | 86 | Nationwide US distribution only. | Terminated | February 25, 2014 | 208 | 86 | Computer
Z-2047-2013 | 65609 | [6 empty] | PSA Assays for use in the IMMULITE/IMMULITE 1000 Systems Analyzers, Catalog Numbers LKPS1, LKPS5 and LKPTS1(D); and in the IMMULITE 2000/2000XPi Systems Analyzers, Catalog Numbers L2KPS2, L2KPS6, L2KPTS2(D) and L2KPTS6(D). IVD Immunoassay kits in the following sizes: 100, 200, 500 and 600 test kits; For in vitro diagnostic use with the IMMULITE/IMMULITE 1000 and the IMMULITE 2000/ 2000 XPi Systems Analyzers. PMA # P930027/ Device Listing # D002374. |  |  | June 26, 2013 | 2013 | Siemens Healthcare Diagnostics | Customers complained of a positive bias in patient results on the IMMULITE/ IMMULITE 1000 and the IMMULITE 2000/ 2000 XPi PSA assay when compared against the IMMULITE /IMMULITE 1000 3rd Generation PSA assay. Siemens has confirmed an overall average positive bias of approximately 20% to 23% across the assay range relative to the WHO 96/670 with the IMMMULITE Systems PSA Assay. This positive bias i | Software design | Seimens sent an Urgent Medical Device Recall Letters to US consignees via FedEx and an Urgent Field Safety Notices were sent to the foreign consignees starting on June 26, 2013. In addition, Support Bulletins were distributed to Siemens Healthcare Diagnostics service personnel describing the issue and instructing them how to deal with customer questions. Customers were informed to discontinue use, discard the kits remaining in inventory and there are no replacement PSA kits available at this time. Siemens is instructing customers to contact their local Siemens representative for assistance with determining appropriate PSA testing solutions for your laboratory. |  | 24361 | 24,361 kits | Worldwide Distribution - USA Nationwide including Puerto Rico and the countries of: Albania, Afghanistan, Algeria, Argentina, Austria, Bangladesh, Belgium, Bolivia, Bosnia- Herzegovinia, Brazil, Bulgaria, Camaroon, Canada, Canary Islands, Chile, China, Colombia, Coast Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, El Salvador, Estonia, Finland, France, Germany, Great Britain, Greece, Guatemala, Honduras, Hungary, India, Iran, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Latvia, Lebanon, Lithuania, Macedonia, Malaysia, Malta, Mexico, Morocco, Nicaragua, Oman, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Romania, Russian Federation, Saudi Arabia, Serbia, Slovakia, Slovenia, Spain, South Africa, Rep. of Korea (S), Sweden, Switzerland, Syria, Taiwan, Thailand, The Netherlands, Tunisia, Turkey, Turkmenistan, Ukraine, United Arab Emirates, Uruguay, White Russia, Venezuela, and Vietnam. | Open, Classified |  |  | 24361 | Computer
Z-0972-2013 | 64285 | No | HTO | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | ReClaim Reamer Extension; Product code 2975-00-500, used with Reclaim Modular Revision Hip Prosthesis System. The ReClaim Reamer Extension is as an attachment to the distal reamers to allow the surgeon to clear soft tissue while reaming the distal femoral canal. |  |  | February 25, 2013 | 2013 | DePuy Orthopaedics, Inc. | The tabs on the reamer extension device have broken. DePuy is notifying surgeons of the issue to increase the awareness and to ask surgeons and users to look for tabs before and after surgery to reduce any potential that the tabs are left in the patient. Once the root cause is determined and a design change is implemented, DePuy will conduct a formal swap-out of the affected devices. | Device Design | DePuy Synthes sent a Urgent Information Medical Device Correction Notice to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed upon visual inspection, if the reamer extension is found to have one or more tabs broken, please do not use the reamer extension. The reamer extension will still engage with one tab intact and no clinical implications have been identified with the use of the reamer extension with one tab, but DePuy recommends it be returned and the alternate reamer extension device be used.If both tabs are broken, the reamer extension will not engage and DePuy recommends it be returned and the alternate reamer extension device be used. A reamer extractor is available in the instrument kit for removal of the reamer if necessary.In order to reduce the risk of breaking the tabs during surgery, the Reamer Extension should only be used to drive the Distal Reamer axially within the patients femoral canal. The instrument is NOT intended to be used for side-cutting or lateralizing the femoral canal, therefore, no off-axis force should be applied during distal reaming.The purpose of this communication is to inform you of this device correction and request acknowledgement of receipt of this letter by signing and returning the Surgeon Response Card within 14 days of receiving this notice. Response cards should be faxed to Fax Number: 574-371-4939.For product-related questions, please contact your local DePuy Orthopaedics Sales Representative. For clinical-related questions from surgeons, please have them contact DePuys Scientific Information Office at 1-888-554-2482 (M-F; 8 am to 5 pm EST.) For questions about device correction information provided, 574-372-7333 (M-F: 8 am to 5 pm EST.) |  | 750 | 750 | Worldwide Distribution - USA including the states of ME, MA, PA, MD, FL, LA, IL, IA, OH, MN, TX, MT, WA, CA, HI, NV, AZ, KY, WI, CO, GA, MI, TN, NY, AR, NC and VA and the countries of Canada, Chile, Germany, Czech Republic, Israel , Norway, Poland, Slovenia, Australia, New Zealand, South Africa, Belgium, Austria, Switzerland, France, England and County Cork, Ireland. | Terminated | December 30, 2013 | 308 | 750 | Computer
Z-0977-2013 | 64112 | No | CFR | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | BS-200 Chemistry Analyzer |  |  | November 28, 2012 | 2012 | Mindray DS USA, Inc. d.b.a. Mindray North America | An issue has been identified with the BS-200 where clinical chemistry results may be inaccurate. | Device Design | Mindray North America /Mindray DS USA Inc. sent Urgent Product Corrective Action letters via certified mail, return receipt requested, signature required on 11/28/2012. The letters identified the affected product as well as the issue, effect on patients, and the corrective action. Customers are to remove the analyzer from service until the affected pump has been replaced. Also, customers are to contact Mindray Customer Service Representative Tony Li at 1-800-288-2121 x2112 for technical assistance, and to arrange for replacement pumps at no charge. |  | 3 | 3 units | Distributed in the states of NC and TX. | Terminated | June 12, 2013 | 196 | 3 | Computer
Z-1989-2013 | 65862 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Precision 500D |  |  | May 02, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue involving the overhead Video Monitor Suspension involving Precision 500D Systems. There is a potential for the locking nut to loosen out of the Video Monitor Suspension yoke which could result in impact. This will not lead to an immediate fall of the suspended monitor since it will be held in place by the remaining bushing and | Device Design | GE Healthcare "Urgent Medical Device Correction" letter dated May 1, 2013 to all affected customers. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. The letter informed customers that GE Healthcare will correct all affected systems at no cost. A GE Healthcare service representative will contact you to arrange for this correction. For questions or concerns regarding this notification call: US 800-437-1171, Japan: 0120-055-919. |  | 445 | 445 (421 US, 24 OUS) | Worldwide Distribution - USA Nationwide in the states of AL, AK, AR, CA, CO, CT, FL, GA, HI, ID, IL IN, IA, KS, KY, LA, MD, MA, NMI, MN, MS, MO, NE, NV,NH,NJ, NM, NY, NC,ND, OH, OK OR, PA,SC, SD, TN, TX, UT, VA, WA, WV, WI, WY and the countries of CANADA, FRANCE GERMANY, GREAT BRITAIN ITALY, PANAMA, SAUDI ARABIA. | Terminated | November 26, 2013 | 208 | 445 | Computer
Z-0994-2013 | 64674 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Toshiba Infinix-i X-Ray System. A radiographic and fluoroscopic x-ray system for use in diagnostic and angiographic procedures. |  |  | November 08, 2012 | 2012 | Toshiba American Medical Systems Inc | Lost and incorrect images. When using the "F-Rec" fluoroscopy image acquisition recording option in the "manual mode" an infrequent software error may cause images from another study to be displayed on the monitor. The acquired images may be lost requiring additional image acquisitions to be performed. | Software design | In November, 2012, Toshiba America Medical Systems, Inc. sent an Urgent: Medical Device Correction letter to customers. The letter contains a return reply form that is to be faxed to TAMS for retention. |  | 95 | 95 units | USA nationwide. No international. | Terminated | March 26, 2013 | 138 | 95 | Computer
Z-0992-2013 | 64684 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Siemens Axiom Artis zeego systems with software version VC14, VC20 and VC21 in conjunction with specific technical configuration |  |  | November 26, 2012 | 2012 | Siemens Medical Solutions USA, Inc | In the course of the firm's product monitoring activities, a potential risk for patient or operators was determined during the operation of an Artis zeego system with software revision VC14, VC20 and VC21 in conjunction with a specific technical configuration which cannot be completely executed. | Software design | Siemens sent an Urgent Customer Information letter dated November 8, 2012 to affected consignees. The letter notified consignees of the issue and informed them the update AX039/12/S would be started and affected systems would be checked and corrected if required. Customers were instructed that they would be contacted to arrange a date for checking their system.For questions regarding this recall call 610-219-6300. |  | 52 | 52 | Nationwide Distribution including AL, AR, AZ, CA, CT, DC, FL, GA, IA, IL, LA, MD, MI, MO, NC, NH, OR, PA, SD, TN, TX, WA and WI and Puerto Rico. | Terminated | December 24, 2013 | 393 | 52 | Computer
Z-0991-2013 | 64394 | No | LTW | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | VITEK 2 Compact computer system when equipped with the following Hewlett-Packard PC models: |  |  | February 18, 2013 | 2013 | Biomerieux Inc | There is a potential to link an isolate result to the wrong patient and then upload the results to the LIS. | Software design | bioMerieux sent an "URGENT PRODUCT CORRECTION NOTICE" dated February 20, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact Emails were issued by the recalling firm on 2/18/2013 and 2/19/2013 to their subsidiaries and to distributors. Letters began issuing to customers via e-mail on 2/20/2013. Contact bioMerieux Clinical Customer Service at (800) 682-2666 for questions regarding this notice. |  | 17293 | For all 3 computers types: 16,736 originally distributed; 554 refurbished and redistributed | Worldwide Distribution-USA (nationwide) and the countries of Algeria, Angola, Argentina, Australia, Austria, Azerbaidjian, Bahrain, Belarus, Belgium, Benin, Bolivia, Bosnia, Brazil, Bulgaria, Burkina Faso, Camaroun, Canada, Chad, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Democratic Republic of Congo, Demark, Dominican Republic, Dutch Antilles, Ecuador, Egypt, El Salvador, Ethiopia, Finland, France, Gabon, Georgia, Germany, Greece, Guatemala, Guinea, Herzegovina, Honduras, Hungary, India, Indonesia, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Korea, Kosovo, Kuwait, Latvia, Lebanon, Libya, Lithuania, Macedonia, Malaysia, Mali, Mexico, Moldavia, Mongolia, Morocco, Myanmar, Netherlands, Netherlands, Nicaragua, Niger Niamey, Nigeria, Norway, Oman, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Puerto Rico, Qatar, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Thailand, Tunisia, Turkey, Uganda, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Venezuela, Vietnam, and Yemen. | Terminated | August 14, 2014 | 542 | 17293 | Computer
Z-0997-2013 | 64662 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Artis Systems with Software Artis VC20x/VC21A/VD10x that have a DSA license installed. |  |  | March 08, 2013 | 2013 | Siemens Medical Solutions USA, Inc | The firm has become aware of a potential issue on Artis systems with the software Artis VC20x/VC21A/VD10x that have a DSA license installed. Under certain preconditions, executing roadmap OGP (organ program) in the DSA overlay mode may lead to imprecise registration of the DSA mask image with the roadmap image displayed on the live monitor. | Software design | Siemens sent a Customer Safety Advisory Notice on March 8, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to pass the information on to all those who need to be aware of it and to instruct their personnel accordingly. Customers were asked to forward the safety information to other organizations that couald be affected by this action.For questions regarding this recall call 610-219-6300. |  | 81 | 81 | Nationwide Distribution including AL, CA, CO, FL, IA, IL, IN, KY, MA, MD, ME, MI, MN, MO, NC, ND, NJ, NY, OH, PA, SC, TX, UT, VA, WA, and WI and to Puerto Rico. | Terminated | November 06, 2014 | 608 | 81 | Computer
Z-0998-2013 | 64609 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | STERRAD 200 System Carriage, P/N 10208. |  |  | March 13, 2012 | 2012 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) is recalling the STERRAD 200 System Carriage because it can potentially become dislodged when an excessive load is applied. | Device Design | Advanced Sterilization Products (ASP) sent an "URGENT MEDICAL DEVICE NOTIFICATION" letter dated March 13, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. The letter informed the customers that ASP recently determined that the push bar located on the STERRAD 200 System Carriage can potentially become dislodged when an excessive load is applied. Customers with questions are instructed to contact Stericycle at (877) 931-0642. |  | 2457 | 2,457 units | Worldwide Distribution-USA (nationwide) and the countries of Egypt, Germany, Italy, France, Saudi Arabia, UAE, Japan, Singapore, China, Thailand, Hong Kong, Hungary, Israel, Slovania, Poland, Czech Republic, Taiwan, Sweden, Greece, Korea, Ireland, and United Kingdom. | Terminated | May 28, 2014 | 806 | 2457 | Computer
Z-1977-2013 | 65380 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | 100/120V AC, Advanced Perfusion System 1 |  |  | July 26, 2013 | 2013 | Terumo Cardiovascular Systems Corporation | TCVS is taking action to update the Operators Manual with the specific instructions to respond to this condition and will issue a labeling addendum for current users. In the event of a false high pressure alarm condition occurs that displays a "999" pressure reading and results in stopping a pump, this pressure display can be cleared and the pump re-started by disconnecting the pressure transduce | Software design | Terumo sent an Urgent Medical Device Recall notification letter dated July 26, 2013 to all affected customers. The letter identified the affected product, reason for correction, potential hazards of the recall issue, the correction to be implemented, and customer instructions for replying to the recall notification. Terumo CVS is providing an addendum to be placed in the Operator's Manual at the beginning of Chapter 7, Modules. Questions or concerns are directed to Terumo CVS Customer Service at 1-800-521-2818, Recall Fax 1-877-275-6860. Customer Service hours are Monday to Friday 8am-6pm ET. |  | 1657 | 1657 | Worldwide Distribution - US Nationwide in the states of :AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, MSMT, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WI, WV and countries of AUSTRALIA, UNITED ARAB EMIRATES (UAE),Hong Kong,Indonesia, Singapore, Taiwan, Thailand,Philippines, CHILE,Malaysia,South Korea, BELGIUM,Japan,Mexico,CANADA | Open, Classified |  |  | 1657 | Computer
Z-1975-2013 | 65719 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Siemens syngo.plaza SW version VA20 with Mammography license. |  |  | March 04, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Siemens issued a Customer Safety Notice to inform customers of a potential problem when using syngo.plaza. In some instances, it can happen that the laterality information in mammography images is not displayed and these images may not be properly aligned to mammography segments. | Software design | The firm, Siemens, notified their customers on March 4, 2013 and sent a "Customer Safety Advisory Notice" letter dated April 03, 2013. The letter described the product, problem and actions to be taken. The customers were instructed to always check whether the laterality information is displayed in the image text of mammography images; if not, the DICOM header is accessible and can be checked for the tag "Laterality" (0020, 0060). If a system is not able to provide the DICOM tag "Image Laterality" (0020, 0062), configuring the image text by adding the tag was proposed. If you have sold the device/equipment and it is no longer in your possession, forward this safety notice to the new owner of the device/equipment and inform us about the new owner of the device/equipment. This issue will be resolved in the syngo.plaza software version VA20D_HF04 which is planned to be released in 02/2013. Customers were also instructed to contact the local Siemens service office to plan the update.If you have any questions, contact the SYNGO CEO or Vice President Quality at +49 (9131) 84-4874 or email to: michael.mlb.braun@siemens.com. |  | 6 | 6 | Nationwide distribution: FL, MA, KS, NY, OH, PA, TX and WI. | Terminated | March 09, 2015 | 735 | 6 | Computer
Z-1969-2013 | 65354 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | TerumoÂż Advanced Perfusion System 1: 220/240V AC, Advanced Perfusion System 1 |  |  | July 26, 2013 | 2013 | Terumo Cardiovascular Systems Corporation | The operators manual does not fully describe the sequence of expected events that will occur when a specific operating condition exists. Specifically, if the user is adjusting the System 1 centrifugal pump speed via the Central Control Monitor at the same time an alarm condition occurs with configuration for a pump coast response, the labeling in the operators manual does not fully describe how th | Software design | Terumo CVS, sent an URGENT MEDICAL DEVICE RECALL notification letter dated July 26, 2013, to all affected customers explaining the correction to the operators manual, potential hazards and what to do in response to the issue.If you have any questions or concern please contact Terumo CVS Customer Service 1-800-521-2818 Recall Fax 1-734-741-6149 Customer Service Hours:Monday  Friday, 8 AM - 6 PM ET |  | 1657 | 1657 | Worldwide Distribution - USA (nationwide) including the countries of Australia, Dubai, Hong Kong, Indonesia, Singapore, China, Mexico, British Columbia, Taiwan, Thailand, Chile, Malaysia, Japan, Belgium, United Arab Emirates (UAE), South Korea, Manila and Canada. | Open, Classified |  |  | 1657 | Computer
Z-0630-2013 | 63616 | Yes | MSX | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | DPM Central Station |  |  | May 12, 2011 | 2011 | Mindray DS USA, Inc. d.b.a. Mindray North America | There is a software anomaly with the DPM Central Monitoring System where the trend data of one patient maybe replaced with the trend data of a second patient. | Software design | Mindray sent a Product Corrective Action letter dated May 12, 2012, to all affected customers via certified mail return receipt requested to two medical facilities. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed that the issue would be corrected through a softweare upgrade of the affected product. A Mindray Service Representative would contact their facility to arrange for the software upgrade at no cost to them.For questions regarding this recall call 1-800-288-2121. |  | 2 | two units | Nationwide Distribution including OH and GA | Terminated | February 19, 2013 | 649 | 2 | Computer