## Sheet1
Recall Number | Recall Event ID | Device Type | Product Code | Regulation Description | Medical Specialty | Review Panel | Submission Type | Trade Name/Product | Main Name | Recall Class | Date Posted | Year Posted | Recalling Manufacturer | Reason for Recall | FDA Determined Cause | Action | Consumer Instructions | Clean Quantity | Quantity in Commerce | Distribution | Recall Status | Termination Date | Time to Terminate | Merged Quantities | Fault_Class
Z-1981-2013 | 65833 | No | JDG | Hip joint femoral (hemi-hip) metallic cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Moore Hip Prosthesis Long Fenestrated Stem, several head Diameter, 203 mm Stem length. |  |  | July 16, 2013 | 2013 | Zimmer, Inc. | Zimmer recently conducted a review of historic packaging validations completed. Based on this review, it was determined that a subset of product packaged at a single Zimmer facility required testing to ensure that previously packaged product from this facility met the established requirements for sterile barrier integrity. This testing determined that a single packaging tray configuration, which | Packaging process control | Zimmer is removing the devices from the field including distributor and hospital inventories. A notification was issued to known users of the product on 07/16/2013. |  | 1710 | 1,710 units are in distribution | Worldwide Distribution-USA (nationwide) including the states of CT, FL, GA, IL, IN, KS, KY, MA, MI, MN, MT, NC, NJ, NY, OH, OR, PA, SC, TN, TX, UT, VI, and WI, and the countries of Singapore, Canada, Mexico, Arab Emirates, Saudi Arabia, and Hong Kong. | Terminated | April 30, 2015 | 653 | 1710 | Not_Computer
Z-1983-2013 | 65785 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | GYRUS ACMI, REF 3740, Everest BiCOAG Bipolar Dissecting Forceps 5mm/45 cm, Sterile EO. |  |  | July 24, 2013 | 2013 | Gyrus Medical, Inc | On April 29, 2013 Olympus Surgical Technologies America R&D; received results from an accelerated aging study conducted as part of a new market submission. The study found that one of 60 Dissector units that underwent 3 year accelerated aging failed the dye penetration test indicating that the seal on sterile packaging was compromised. The dissector products are labeled with a 3 year shelf life. | Packaging process control | Olympus sent a "Medical Device Recall" letter dated July 24, 2013, to all affected customers. The letter identified the problem the product and the action needed to be taken by the customer.. Customers were advised to immediately cease any further use of the product and remove it from inventory and quarantine it until is shipped back to the company. To return the product, call 1-888-524-7266 Olympus customer service representative. Customers were requested to complete and return by fax the enclosed questionnaire to 484-896-7128. If the product was further distributed then they need to identify their customers and notify them at once of the product recall. |  | 1140 | 1140 | Worldwide Distribution - USA (nationwide) including the states of AL, AZ, CA, FL, GA, IL, IN, KS,KY, MD, MA, MI, MN, MS, MO, NY, NC, OH, PA, SD, TN, TX, VA and WI., and the countries of BRAZIL, CHINA, AUSTRALIA, CHINA, GERMANY and INDIA. | Terminated | June 20, 2014 | 331 | 1140 | Not_Computer
Z-1001-2013 | 64585 | No | OJU | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Sepp, Compound Benzoin Tincture Packaged and Tincture 3,000 Bulk. |  |  | March 04, 2013 | 2013 | CareFusion 213, LLC | An internal review of CareFusion's labeling for preoperative skin prep product codes was recently conducted. The labeling of the product codes identified above state that the contents of the package are sterile and does not clarify that this sterility claim is only applicable to the applicator. The antiseptic solution that is applied to the skin is a cutaneous disinfectant, and is not sterilized. | Labeling False and Misleading | CareFusion sent an Urgent Recall Notice dated March 4, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immediately check inventory, quarantine and return the affected product for replacement. For questions call (913) 345-3504. |  | 13749000 | 13,749,000 | USA Nationwide Distribution | Terminated | May 12, 2014 | 434 | 13749000 | Not_Computer
Z-0109-2014 | 66222 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Vented Spike Adapter Product Code 2C0471, Sterile, nopyrogenic. Manufactured by an affiliate of Baxter Healthcare Corporation, Clintec Nutrition Division, Deerfield IL 60015 USA. |  |  | August 30, 2013 | 2013 | Baxter Healthcare Corp. | Potential tears in the pouch of six lots of Vented Spike Adapter Product Code 2C0471 | Packaging | The firm, Baxter, sent an "Urgent Product Recall" letter dated September 10, 2013 to its consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed to locate and remove all affected product from your facility; If you distributed the products to other facilities or departments, forward them a copy of the letter; complete and return the attached Customer Reply Form via fax to 1-224-270-5457 or scan and email to fca@baxter.com; and contact Baxter Healthcare Center for Service at 1-888-229-0001 between hours of 7:00 am and 6:00pm Central Time to return affected product and receiving credit.For additional information contact the Center for One Baxter at 1-800-422-9837, Monday thru Friday during the hours of 8:00am to 5:00pm Central Time. |  | 49344 | 49,344 units | Nationwide distribution. | Terminated | April 19, 2017 | 1328 | 49344 | Not_Computer
Z-1938-2013 | 65698 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | C-QUR V-Patch Mesh (All sizes and shapes). |  |  | July 19, 2013 | 2013 | Atrium Medical Corporation | Additional Instructions for Use and Storage Conditions as Coated mesh can adhere to the inner packaging liner due to exposure to high humidity conditions | Package design/selection | Atrium Medical issued Recall Letter via UPS and e-mail on 7/19/13 to the accounts and field representatives. The notification identifies the problem, product, and risk factors. If the Product is exposed to excessive humidity for an extended period of time, then the increased humidity occurring inside the pouch can potentially cause the coating on the mesh to strongly adhere to the inner handling sleeve. A reply for is requested to be completed to acknowledge receipt to the notification.Additional language to the instructions for use (IFU) to include:Prolonged exposure to high humidity may result in increased rate of adherence of the C-QUR mesh to its handling sleeve. Store in a Controlled Room Temperature (25 C / 77 F ) or less. Brief exposure to up to 40 C (104 F ) is acceptable.Questions please contact Atrium Medical Customer Service at 1- 800-528-7486 Monday through Friday 9:00 am to 5:00 pm EDT. |  | 95286 | 95,286 units | USA (nationwide) including Puerto Rico and the countries of Australia Austria Bahrain Brazil Canada Chile Colombia Dominican Republic Ecuador El Salvador France Germany Great Britain Greece Honduras Hong Kong India Ireland Israel Italy Japan Jordan Korea Malaysia Mexico Netherlands New Zealand Nicaragua Norway Panama Peru Portugal Romania Saudi Arabia Singapore South Africa Spain Sri Lanka Switzerland Taiwan Thailand Turkey and Venezuela. | Terminated | March 18, 2016 | 973 | 145251 | Not_Computer
Z-0987-2013 | 63946 | No | LOM |  |  | Microbiology | PMA | ARCHITECT HBsAg Qualitative Assay, List Number 4P53 and ARCHITECT HBsAg Qualitative Confirmatory Assay, List Number 4P54 |  |  | December 18, 2012 | 2012 | Abbott Laboratories, Inc | Abbott has identified that assay files for the ARCHITECT HBsAg Qualitative and HBsAg Qualitative Confirmatory assays (LNs 4P53 and 4P54) on the ARCHITECT i1000SR System were mistakenly made available on abbottdiagnostics.com in October 2012. Additionally, some i1000SR customers received notification through AbbottLink in November 2012 that assay files for LNs 4P53 and 4P54 were available. The AR | No Marketing Application | The firm decided to correct the issue and and sent out Urgent Field Safety Notices to their customers dated 10/18/2012. |  | 12 | 12 | Nationwide Distribution including AZ, CA, FL, GA, HI, LA, MA, MD, MO, NJ, NY, PA, RI, TN, TX, UT, and WA. | Terminated | July 31, 2013 | 225 | 12 | Not_Computer
Z-0111-2014 | 66524 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Single Site Curved Needed Driver |  |  | October 02, 2013 | 2013 | Intuitive Surgical, Inc. | A limited number of FDA-cleared instruments for da Vinci Single Site use, Bipolar Maryland and curved needle Driver, were shipped without the correct revision of the user manual which provided added instruction for use of these additional Single Site instruments including indication, warnings and precautions. | Labeling Change Control | The firm, Intuitive Surgical, sent an " Urgent Medical Device Correction" letter dated October 2, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure all affected personnel are fully informed of this notice; after receiving the updated manual, locate and discard all previous versions of the Single Site Instruments manual and replace with the new manual; complete and return the attached Acknowledgement Form via FAX to Intuitive Surgical, Inc., ATTN: REGULATORY COMPLIANCE, U.S. (408) 716-3040, or Scan and Email to: isi.compliance@intusurg.com.; Inform affected personnel when the correction has been completed; and Retain a copy of this Notice and the Acknowledgement Form for your files.If you have any questions, contact your Clinical Sales Representative or contact Intuitive Surgical Customer Service at the numbers listed below:" North and South America: 800-876-1310 Option 3 (6 am to 5 pm PST)" Japan: 0120-56-5635 or 003-5575-1362 (9 am to 6 pm JST)" South Korea: 02-3271-3200 (9 am to 6 pm KSTJ)" Europe, Middle East, Asia and Africa: +800 0821 2020 or +41 21 821 2020 (8 am to 6 pm CET). |  | 1118880 | 504 devices for a total of 2,220 affected clinical product. | Worldwide distribution: US (nationwide) and countries of:Australia, Belgium, France, Italy and United Kingdom. . | Terminated | September 04, 2014 | 337 | 1118880 | Not_Computer
Z-0875-2013 | 64322 | No | NUC | Condom. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | KY TINGLING JELLY Personal Lubricant. The product is packaged into 5 oz tubes with the UPC printed on it. Distributed By: McNeil-PPC Inc. |  |  | January 25, 2013 | 2013 | Mc-NEIL-PPC, Inc. | This product is not cleared for marketing in the US. | No Marketing Application | Johnson and Johnson Healthcare Products sent recall notification letters dated 1/25/2013 to its customers (retailers and wholesalers). The letter identified the product, problem, and actions to be taken by the customers. Customers were asked to return all cases of affected product in their possession. The letter included delivery confirmation and a return packing slip, return authorization label and business reply card. Contact J &amp; J Sales &amp; Logistics at 866-912-8414 for questions regarding the return of affected product. Contact your J &amp; J Sales &amp; Logistics Company Sales Representative for questions about the product. |  | 425304 | 425,304 units | Nationwide Distribution. | Terminated | April 11, 2014 | 441 | 2361232 | Not_Computer
Z-0101-2014 | 65920 | No | GKR | Automated hemoglobin system. | Hematology | Hematology | 510(k) | HemoCueÂż Hb 201 Microcuvettes, carton/4 boxes/25 pouches of Microcuvettes, HemoCue AB Angelholm, Sweden |  |  | July 25, 2013 | 2013 | HemoCue AB | HemoCue AB has become aware of a production related problem involving punctured HemoCue Hb 201 Microcuvettes single pack pouches. | Packaging | On 8/13/2013 the firm sent Field Safety Notices to their customers. |  | 169320 | 42,330 packages/4 boxes/25 microcuvetes | Worldwide distribution which includes the following states: AK, AL, AS, AZ, CA, CO, CT, FL, GA, IA, IL, IN, KS, KY,MD, ME, MI, MN, MO, NC, ND, NE, NJ, NM, NY, OH, OK, OR, SC, SD, TN, TX, UT, VA, WA, WI, VT, WV &amp; WYand the countries of Armenia, Australia, Austria, Belgium, Cyprus, Czech Republic, Denmark, El Salvador Finland, France, Germany, Great Britain, Guatemala, Hong Kong, Israel, Ivory Coast, Netherlands, New Zealand, Norway, Peru, Poland, Russia, South Africa, Sweden, Switzerland and Taiwan. | Terminated | August 11, 2014 | 382 | 169320 | Not_Computer
Z-0338-2014 | 66763 | No | DXZ | Intracardiac patch or pledget made of polypropylene, polyethylene terephthalate, or polytetrafluoroethylene. | Cardiovascular | Cardiovascular | 510(k) | Vascu Guard, Peripheral Vascular Patch with Apex Processing. |  |  | October 28, 2013 | 2013 | Synovis Surgical Innovations, Inc. | Synovis, a subsidiary of Baxter International Inc., is recalling certain shipments of VASCUGUARD and DURA-GUARD because Synovis, a subsidiary of Baxter International Inc., is recalling certain shipments of VASCUGUARD and DURA-GUARD because the shipping configurations in which the 6 fluid ounce units containing a specific revision of the Instructions for Use (IFU) shipped in a divided cell shipping | Packaging change control | A customer notification letter dated 10/28/13 was sent to all customers who purchased the Vascu Guard and the Dura Guard. The letter informs the customers of the problem identified and the actions to be taken. Customers with any questions regarding the recall are instructed to contact Customer Service unit at (800) 423-2090. If customers are distributors or if they have have distributed the recalled products to any other health care provider, they are instructed to notify their accounts and provide them with a copy of the recall notification. If customers would like Baxter to perform the recall task for them, then they are instructed to contact Stericycle at 1-866-912-8417. |  | 6014 | 6014 units | Nationwide Distribution. | Terminated | June 03, 2014 | 218 | 6460 | Not_Computer
Z-0098-2014 | 66362 | No | KWI | Elbow joint radial (hemi-elbow) polymer prosthesis. | Orthopedic | Orthopedic | 510(k) | Small Bone Innovations (SBi) rHead Radial Stem, rHead Recon Radial System, rHead Radial Stem Lateral Assembly, rHead Radial Stem 6mm Collar, rHead Recon Radial Stem 6mm Collar, rHead Radial Stem Lateral Assembly 6mm Collar, rHead Standard Extended Stem, rHead Radial Stem Plasma Coated, rHead Recon Radial Stem Plasma Coated, rHead Radial Stem 6mm Collar Plasma Coated, rHead Recon Radial Stem 6mm Collar Plasma Coated, uHead Ulnar Stem Standard, uHead Ulnar Stem 20mm Collar packaged in a sterile pouch inside a shrink-wrapped box. |  |  | September 06, 2013 | 2013 | Small Bone Innovations, Inc. | The recall was initiated due to an increased risk of implants breaching the sterile pouches inside the shrink-wrapped box. | Packaging change control | Email correspondence and a notification letter, dated September 6, 2013, were sent to users requesting the immediate return of affected devices. Another notification letter, dated December 20, 2013 was also sent for recall of the additional devices after the recall was expanded. |  | 2036 | 2036 devices | Worldwide Distribution-USA (nationwide) and the countries of Canada and Mexico. | Terminated | March 30, 2015 | 570 | 2036 | Not_Computer
Z-1987-2013 | 65678 | No | KDQ | Vacuum-powered body fluid suction apparatus. | General Hospital | General Hospital | 510(K) Exempt | PleuraGuide Disposable Chest Tube Kit; Product codes: 17100, 17150, and 17200. |  |  | July 16, 2013 | 2013 | Atrium Medical Corporation | Individual component's IFUs are not included with the kit. | Labeling Change Control | Notice of correction was sent on July 18, 2013 via UPS and e-mail to the field representatives that represent these accounts. The notification identifies the problem, product, and risk factors. It provides instructions how to handle the product that may be affected. Additional labeling will be added to the kit to include instructions for use for the additional components. Atrium Medical will confirm through tracking of the UPS letters that each consignee has received the letter of correction. This shall be 100%. Where a UPS delivery cannot be confirmed, Atrium Medical will make good faith effort to contact the consignee to ensure that the letter was received. |  | 6516 | 6,516 units | Nationwide Distribution. | Terminated | July 15, 2014 | 364 | 6516 | Not_Computer
Z-1928-2013 | 65370 | No | MBL | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | R3 (TM) XLPE ACETABULAR LINER UHMWPE, 20o OVERHANG, 36 MM ID, 62 MM OD, QTY: (1), STERILE H2O2, REF 71335762. |  |  | May 29, 2013 | 2013 | Smith &amp; Nephew Inc | One batch of 60 mm 20o R3 XLPE liner may be laser etched and packaged as a 62 mm 20o liner and one batch of 62 mm 20o R3 XLPE liner may be laser etched and packaged as a 60 mm 20o liner. | Packaging process control | All affected Smith &amp; Nephew Sales Representatives were notified of problem and the recall via email on 5/29. They were instructed to cease use and distribution of the product and quarantine for return. Surgeons were notified by letter beginning on 6/12. International consignees were notified on 05/29 via email. |  | 53 | 53 units | Worldwide Distribution-USA including the states of PA, WA, TX, FL, IA, MI, WI, TN, KY, IN, NY, MO and OR, and the countries of India, Canada, Chile, Venezuela, Dubai, South Africa, Australia, Poland, Romania, Belgium, Germany and United Kingdom. | Terminated | March 22, 2016 | 1028 | 53 | Not_Computer
Z-0028-2014 | 66325 | No | MDB | Microbial growth monitor. | Microbiology | Microbiology | 510(k) | BBL (tm) MGIT (tm) Mycobacteria Growth Indicator Tube, part # 245122, packaged in 100 ct cartons and labeled in part |  |  | September 09, 2013 | 2013 | Becton Dickinson &amp; Co. | An in-vitro diagnostic reagent may be labeled with an incorrect barcode, which when scanned on associated diagnostic instrument may cause a false negative result. | Error in labeling | BD sent an Urgent Product Recall letter dated September 2013 to customers identified as microbiology lab managers. The letter identified the affected product, problem and actions to be taken. The letter advised that referenced lot numbers may contain an overlabel barcode that may begin to lift, exposing the original incorrect barcode pictured in the notification. Customers were instructed to complete the attached form and fax to Regulatory Compliance 410-316-4258. For further assistance regarding replacements contact BD Customer Service Department at 1-800-675-0908 and all other inquiries contact BD Technical Services Department at 1-800-638-8663. |  | 187000 | 187,000 tubes | USA Nationwide Disatribution in the states of: AL, AR, AZ, CA, CO, CT, FL, GA, IL, KY, MA, ME, MN, MO, NC, ND, NJ, NV, NY, OH, PA, TX, UT, VA, WA, WI | Terminated | October 29, 2013 | 50 | 187000 | Not_Computer
Z-0876-2013 | 64303 | No | EJF | Orthodontic appliance and accessories. | Dental | Dental | 510(K) Exempt | Avex CX Suite, Avex CX 0.22 UL4,5 HK REF #224-017, |  |  | February 08, 2013 | 2013 | Ultradent Products, Inc. | Orthodontic brackets distributed was incorrectly labeled. | Packaging | Ultradent Products, Inc. sent an Urgent Medical Device Correction letter dated February 8 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were asked to return affected products for replacement. Customers with questions can call firm representative, Corey Jaseph at (801) 553-4220. |  | 124 | 124 | USA Nationwide Distribution including the states of CA, AZ, CO and WA. | Terminated | June 07, 2013 | 119 | 124 | Not_Computer
Z-1013-2013 | 64327 | No | LTW | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | bioMerieux VITEK 2 AST-ST01 Streptococcus Susceptibility Card, REF 410 028. |  |  | February 06, 2013 | 2013 | Biomerieux Inc | The product package did not contain the correct version of the package insert listing the limitations of use and performance requirements. | Error in labeling | bioMerieux sent an "PRODUCT SAFETY CORRECTION NOTICE" dated February 13, 2013 to the affected customer. The letter identified the product, problem, and actions to be taken by the customer. Contact the firm at (800) 682-2666 for questions regarding this notice. |  | 4 | 4/20-card cartons | Distributed only in HI. | Terminated | October 30, 2013 | 266 | 4 | Not_Computer
Z-0051-2014 | 65969 | No | HMX | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | Synergetics 25 gauge Soft Tip Cannula, packaged in 6 units per box. Catalog number: 30.08.25. |  |  | July 30, 2013 | 2013 | Synergetics Inc | One lot of 25 gauge Soft-Tip Disposable Cannulas are co-mingled with 23 gauge cannulas; the product is labeled as 25 gauge cannulas. | Labeling Change Control | The firm sent recall notification letters to their consignees on 7/30/13. |  | 48 | 48/6-cannula boxes | Nationwide Distribution-USA including the states of MO, CO, KY, TN, NC, PA, IL CA, NY, IN, MI, LA, NJ and OH. | Terminated | April 07, 2014 | 251 | 48 | Not_Computer
Z-0787-2013 | 64028 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Exactech NOVATION CROWN CUP, NEUTRAL LINER, Use with Crown Cup Only, GXL UHMWPE, SN 0123456, REF 130-32-53. |  |  | November 05, 2012 | 2012 | Exactech, Inc. | Recall of Novation Crown Cup Connection GXL Neutral Liners Catalog No. (130-32-53) and Novation Crown Cup Connection GXL + 5mm Lateralized Liners Catalog No. (136-32-53). Exactech discovered the device labeling was inadvertently transposed. | Labeling mix-ups | Exactech, Inc. notified consignees by telephone on 11/05/12 of the affected devices and sent an "IMPORTANT PRODUCT RECALL NOTICE" letter dated 11/06/12. The letter describes the purpose of the recall, the product affected, the problem and actions to be taken by the consignees. A Product Recall Response Form was attached for the consignees to contact the Exactech inventory representative within 5 business days to confirm quantities of product in their inventory and return them to Exactech. Consignees with any questions relating to these devices should call 1-800-392-2832. |  | 26 | 26 Units | Distributed in the states of TN, FL, VA, WA and NV. | Terminated | December 24, 2013 | 414 | 26 | Not_Computer
Z-0086-2014 | 65918 | No | KQO | Automated urinalysis system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | HemoCueÂż Urine Albumin Microcuvettes, 25 Microcuvettes HemoCue AB Angelholm, Sweden |  |  | July 16, 2013 | 2013 | HemoCue AB | HemoCue AB has become aware of a production related problem involving punctured HemoCueÂż Urine Albumin Microcuvettes single pack pouches. | Packaging | HemoCue sent an Urgent Field Safety Notice dated July 24, 2013. to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed to :1. Quarantine any HemoCueÂż Urine Albumin (110608); lot numbers 1208848 to 1304847 you may have in your possession2. Return the Field Safety Notice Verification Form to: fax: 562-668-5794We apologize for the inconvenience this advisory has caused you. Please feel free to contact Technical Support at 800-426-7256 with any questions. |  | 517800 | 10,356 boxes/50 microcuvettes | Worldwide Distribution - Nationwide Distribution including the states of AL, AR, AZ, CA, CA ,CT, FL, GA, HE, ID, IL, IN, LA, MA, ME, MI, MN, MO, MS, NC, NH, NM, NY, OH, PA, SC,TX, UT, VA, VT, WA and WI., and the countries of Australia, Bahrain, Belarus, Canada, Czech Republic, Denmark, Finland, France, Germany, Great Britain, Hong Kong, India, Indonesia, Italy, Kenya, Lesotho, Malawi, Mauritius, Mexico, Myanmar, Nigeria, Norway, Omen, Poland, Qatar, Romania, South Africa, South Africa, South Korea, Spain, Sweden, Taiwan, Tanzania, Ukraine, Yemen and Zambia. | Terminated | April 28, 2014 | 286 | 517800 | Not_Computer
Z-0983-2013 | 63951 | No | FSM | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Needle Counters; 12.2 x 10.5 x 3.5 cm; 40 Count/70 Capacity; Sterile; Removable Numbered Foam Block; Adhesive Area (Cat. No. 3FA70SSA) |  |  | December 18, 2012 | 2012 | Cardinal Health | Small holes (approximately 1 mm or less) found in device unit level packaging potentially compromising product sterility. | Package design/selection | Cardinal Health sent an Urgent Product Recall letter dated December 18, 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immedaitely stop using affected product and discard any product that may be on hand, return the enclosed acknowledgement form, notify any customers to whom you may have distributed product affected by this recall, and contact the appropriate Customer Service group to arrange for credit of any affected product discarded. For questions contact Cardinal Health Professional Services at 800-292-9332. |  | 54912 | Cat. No. 3FA70SSA: 23,808; Cat. No. 3FM70SSA: 54,912 | Worldwide Distribution - USA (Nationwide) and the countried of Australia, Canada and Japan. | Open, Classified |  |  | 316752 | Not_Computer
Z-0890-2013 | 64316 | No | FRO |  |  | General &amp; Plastic Surgery | 510(k) | Atrapro Antipruritic Hydrogel, 1.5 oz gel tube, 4 Net wt/ 1.5 oz.. |  |  | February 12, 2013 | 2013 | Oculus Innovative Sciences Inc | Out of specification result - stability test failure of free available chlorine (FAC) which functions as a preservative for the hydrogel. The 13 month test result obtained was 4 ppm FAC and the stability specification is &gt;10 ppm FAC. | Incorrect or no expiration date | Oculus sent an Urgent Medical Device Recall letter dated February 12, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately discontinue distribution and segregate the affected product. Immediately remove the affected lots of Recall product from the inventory of sales representative, sub contracted distribution and or fulfillment centers. Segregate the affected product in a secure location for return or destruction. If the product was further distributed customers were instructed to notify healthcare professionals of this recall. Customers were also instructed to completed and return the Business Reply Card and fax to 1-415-462-5163. Once the firm receives the affected product the customer would be contacted by an Oculus Representative and provided a Returns Goods Authorization (RGA) number.For questions customers were instructed to call 1-707-559-2445.For questions regarding this recall 707-559-7191, ext 14. |  | 56567.5 | 56,566 1.5 oz tubes | Nationwide Distribution - One distributor in PA | Terminated | November 25, 2013 | 286 | 56567.5 | Not_Computer
Z-0336-2014 | 66670 | No | DSY | Vascular graft prosthesis. | Cardiovascular | Cardiovascular | 510(k) | Vascutek Gelsoft Vascular Graft (peripheral) |  |  | October 23, 2013 | 2013 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (Terumo CVS) became aware of a customer complaint that a 15mm x 8mm Gelsoft graft was received which the customer believed was a 15mm x 6mm Gelsoft graft unit. The manufacturer confirmed a batch of product code 631508 was incorrectly labeled. | Error in labeling | Terumo CVS called affected consignees (via telephone)on 10/23/2013 of this device removal. Terumo CVS advised users to discontinue use and return all affected grafts in inventory. Terumo will replace or issue credit for returned grafts. Terumo provided consignees with the reason of the removal, potential hazard, and description of affected product. Terumo addressed any questions from the consignees. Consignees with questions were instructed to contact Terumo CVS Customer Service at 800-521-2818.For questions regarding this recall call 734-741-6173. |  | 9 | 9 | US Distribution in the states of: NC and CA. | Terminated | February 27, 2014 | 127 | 9 | Not_Computer
Z-0870-2013 | 64329 | No | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | MEDRAD Hand Controller Sheath, Catalog number AVA 500 HCS/3010903, for cardiac catheterization procedures. The Sterile Sheath is specifically intended for use in the x-ray angiography environment to cover the Medrad Avanta Sterile Hand Controller to prevent cross contamination between patients. |  |  | November 05, 2012 | 2012 | Medrad Inc | The seal between the white polyethylene and the clear plastic may be missing on the Hand Controller Sheath package, resulting in a potential breach of the sterility of this accessory. | Packaging | Medrad sent a Recall letter dated November 5, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customer were instructed to review their inventory and quarantine any packages with an expiration date prior to April, 2017. Customers were also given the option of conducting a visual inspection on site to determine if the product should be returned. Alternatively, the customer has the option of returning the affected product. If the customer has affected product, they are instructed to contact Medrad Customer Support to make arrangements for the return and replacement of the product.For further questions please call our Customer Support Team at 1-877-229-3767. |  | 832175 | 832,175 | Worldwide Distribution including USA nationwide. | Terminated | March 10, 2015 | 855 | 832175 | Not_Computer
Z-1958-2013 | 64970 | No | FSM | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Case Lid and Base (Inspan, Invue, SLIFT) Caddy Lid (Inspan, Arena-C); Inspan Lid: 12-60039A, 11-60042B, 11-60043B, 11-60045B, 11-60046C, 11-60045A, 11-60046A, 11-60053A; Invue Lid: SI70095A; S-Lift Lid: SI50070A; S-Lift Base: SI50067A; Arena-C Lid: 11-72013D, 11-72014D; Generics Lid: 12-99005A. |  |  | April 22, 2013 | 2013 | SpineFrontier, Inc. | Non-toxic inks, used on the silk screen graphics of the case bases and lids, and caddy lids, were found to be either delaminating or running off from the substrate after repeated washing and sterilization cycles. | Labeling design | On 2/15/2013, SpineFrontier emailed distributors informing them that they needed specific parts returned to work towards completing the conversion of product graphics. |  | 258 | 258 units (30 in distribution at the time of the recall) | Worldwide Distribution-USA including the states of FL, OH, MA, and TX, and the countries of Mexico, and Dominican Republic. | Terminated | November 06, 2015 | 928 | 258 | Not_Computer
Z-1018-2013 | 64510 | No | HDW | Contraceptive diaphragm and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Omniflex Diaphragm - Size 65 (Milex Wide Seal Silicone Vaginal Diaphragm) Model MXWF65 |  |  | February 26, 2013 | 2013 | Cooper Surgical, Inc. | Incorrect size was printed on the outer product packaging. | Packaging | CSI notified consignees affected by this recall by a Recall Notification letter, dated February 26, 2013, sent via Federal Express with confirmed delivery receipt. The letter identified the affected product, and asked customers to contact the firm's Customer Service Department at 1-800-243-2974, if they had any of the affected lot in their inventory. Customers also have the option of completing the form attached to the letter and faxing it back to 800-262-0105. At CooperSurgical's expense, arrangements will made to replace the affected products that customers have on hand. The firm will provide further instruction and disposition of the product at the time of replacement. |  | 80 | 80 individual diaphragms | Worldwide Distribution -- USA, including the states of NE, MA, FL, TN, CA, ME, SC, CA, NH, OK, TX, NM, MO, NY, PA, and AZ, and countries of Germany, Italy, and Canada. | Terminated | March 04, 2014 | 371 | 80 | Not_Computer
Z-0029-2014 | 66105 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Model BP7507 OPTI LION E-Plus Cassettes 25 Cassettes with Sample Probes, OPTIMedical |  |  | August 09, 2013 | 2013 | OPTI Medical Systems, Inc | OPTI LION E-Plus Cassettes were labeled with an incorrect barcode. The incorrect barcode could perfect the performance of one or more parameters as the product ages. Also a Customer complaint from a distributor reported that a customer was getting QC failures. | Error in labeling | The firm, OPTIMedical, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated August 9, 2013 via email and fax to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately discontinue use of the product; return any unused product to OPTI Medical System; distributors outside the U.S. may destroy unused product and certify the instructions on the enclosed Recall Response Card; contact your distributor or OPTI Medical System Technical Support at 1-800-490-6784 or technicalsupport@optimedical.com to obtain authorization number and shipping instructions; and complete and return the Recall Response Card to technical support via fax to 1-770-510-4447; mail to: OPTI Medical Technical Support, OPTI Medical Systems, Inc., 235 Hembree Park Drive, Roswell, GA 30076 or email: technicalsupport@optimedical.com Please contact OPTI Medical Systems, Technical Support department at 1-800-490-6784 if you have questions related to this notice. |  | 9025 | 361 boxes of 25 cassettes each | Worldwide distribution: US (nationwide) including states of: CA, MA, MN, and VA; and countries of: China, Guatemala, India, Iraq, South Korea, and Untied Arab Emirates. | Terminated | December 12, 2013 | 125 | 9025 | Not_Computer
Z-1960-2013 | 65005 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Biomet 3i Dental Implant, Parallel Walled 6mm (D) x 10mm (L) |  |  | May 10, 2013 | 2013 | Biomet 3i, LLC | A small number of outer implant boxes may be mislabeled. The inner tray, containing the implant, is label correctly. This condition, if present and not recognized could potentially result in the dental implant not being able to be placed into the osteotomy or seated to the desired depth if placement is attempted. | Labeling mix-ups | The firm issued "Urgent Medical Device Recall" notification letters to their customers dated May 10, 2013. The notification described the issue and provided recommendations actions regarding affected product. Customers with questions may contact 1 800-342-5454. |  | 51 | 51 | Distribution to the state of Massachusetts, and the foreign countries of Australia, Belgium, Spain, Denmark, France, Luxembourg, Switzerland, Germany, Bonaire, Sweden, and Italy. | Terminated | August 13, 2014 | 460 | 84 | Not_Computer
Z-0879-2013 | 64250 | No | GXI | Radiofrequency lesion probe. | Neurology | Neurology | 510(k) | BMC RF Cannula Curved Sharp RadiOpaque, Non-Pyrogenic, Active Tip: 10mm, Gauge 16 Ga, Length: 100mm. The BMC RF Cannula is sealed in a Tyvek pouch (primary packaging) and then placed into a white shelf-box (secondary packaging) with the instructions for use. A predetermined number of white shelf-boxes are then placed into shipping cartons (tertiary packaging). The BMC RF Cannula is intended for use in radiofrequency heat lesion procedures for relief of pain. |  |  | October 04, 2012 | 2012 | Baylis Medical Company Inc. | A report of an event was received indicating that upon opening a box containing 10 sterile cannula by a nurse, the pouch package of 5 of the 10 individually packaged cannula had not been sealed along one edge. Thus the pouch was open and the device could fall out of the pouch and on to the floor when retrieved from the shelf box. The cannulas were not used on any patients, as the unsealed packages | Packaging | A recall was initiated on 10/04/2012 and a letter was sent out to the distributor with a Fax-back Acknowledgment Form. The letter described the issue and provided recommendations and a contact number. |  | 1780 | 1780 | International distribution to the country of Australia. | Terminated | May 23, 2013 | 231 | 1780 | Not_Computer
Z-0081-2014 | 66483 | No | DTK | Cardiovascular intravascular filter. | Cardiovascular | Cardiovascular | 510(k) | Cordis OPTEASE Vena Cava Filter |  |  | September 27, 2013 | 2013 | Cordis Corporation | Cordis identified a printing error on one unit of the Optease Vena Cava Filter, in which the orientation arrow for the femoral approach was printed in the incorrect direction. The error resulted in the filter being implanted upside down, requiring and additional percutaneous procedure to retrieve the filter. | Error in labeling | Cordis sent an Urgent Medical Device Recall letter dated October 8, 2013, to all affected customers. The letter identified the product (Cordis OPTEASEÂż Vena Cava Filter Catalog Numbers 466F220A and 466F220B), the problem, and the action to be taken by the customer. Customers were instructed to:1) Read this Urgent Medical Device Recall letter.2) Immediately identify and set aside all product listed below in a manner thatensures the affected product will not be used.3) Review, complete, sign and return the enclosed Acknowledgement Formdirectly to Cordis at the fax number on the form.4) Return any affected product with reference to the RA# on the form. ContactCustomer Service (see Available Assistance section) if assistance is needed for return of the affected product. Credit will be provided or consignment levels reduced.5) Share this letter with others in your facility that need to be made aware of thisrecall.6) Contact any other facility to arrange the return of OPTEASEÂż if any productlisted below has been forwarded to them.7) Maintain awareness of this notice until all affected product has been returnedto Cordis.8) Keep a copy of this notice with the affected product.Please contact the special Customer Service number (800) 551-7683,Monday through Friday from 8:30 AM to 6:30 PM ET.o Return instructions are provided on the Acknowledgment Form.Questions related to the Acknowledgement Form" Please contact Cordis QA at (786) 313-8730.For questions regarding this recall call 1-800-551-7683. |  | 29395 | 29,395 units | Worldwide Distribution - US (nationwide) including Hawaii and Puerto Rico and states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY; and countries of: Canada and Panama. | Terminated | September 23, 2014 | 361 | 29395 | Not_Computer
Z-1016-2013 | 64271 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Synthes Dens Instrument Set (Dens Graphic Case), Part Number 687.030 (also known as the Anterior Column Retractor Set) |  |  | January 11, 2013 | 2013 | Synthes USA HQ, Inc. | The Dens Instrument Set has been designed to facilitate the anterior internal fixation of the Dens (odontoid process). There are currently no screws specifically indicated to treat Dens fractures via anterior screw fixation therefore the affected product will be removed. | No Marketing Application | Synthes sent an Urgent Notice Medical Device Recall letter dated January 11, 2013, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to examine their inventory, remove the recalled product from use and return it to Synthes. Customers were asked to completed the attached verification Section at the end of the letter by checking the appropriate box indicating that no affected product had been located and fax to 610-251-9005.For questions customers were instructed to call 610-719-5450 or email at FieldAction@synthes.com.For questions regarding this recall call 610-719-5000. |  | 148 | 148 | Nationwide Distribution including AL, AR, AZ, CA, CO, CT, DE, FL, GA, IL, KS, KY, LA, MD, ME, MI, MN, MO, NC, NE, NH, NJ, NV, NY, OH, OR, PA, RI, TN, TX, VA, VT, WI, WV, and WY. | Terminated | September 10, 2015 | 972 | 148 | Not_Computer
Z-0784-2013 | 63390 | pack, hot or cold, reuseable | IME | Cold pack. | Physical Medicine | Physical Medicine | 510(K) Exempt | THERA PEARL 3-in-1 BREAST THERAPY, 2 hot and cold packs in plastic box. |  |  | August 14, 2012 | 2012 | Lansinoh Laboratories Inc | The Spanish translation on the Instructions for Use (IFU) insert and product package for hot and cold packs has incorrect heating instructions which would cause product to over heat. | Error in labeling | Lansinoh Laboratories, Inc. sent an Urgent Product Recall letter dated August 14, 2012, to all affected customer. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to stop distributing and quarantine the affected product. Customers were informed that a credit would be issued for the affected product and they were asked to mail the postage paid Business Reply Card even if they do not have the affected product. If necessary call 1-877-366-1182 to arrange pickup. For questions customers were instructed to call the Customer Service Center at 1-800-292-4794.For questions regarding this recall call 703-299-6477. |  | 51472 | 51,472 units | Nationwide Distribution, including the states of TX, CA, NJ, NY, SD, RI, WV, IL PA, NC and MN. | Terminated | March 08, 2013 | 206 | 51472 | Not_Computer
Z-1004-2013 | 64531 | No | HHW | Vaginal pessary. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Cooper Surgical Incontinence Ring Flexible Pessary #4, 2 3/4", 70mm. Milex |  |  | February 26, 2013 | 2013 | Cooper Surgical, Inc. | Pessary mislabeled outer package may not reflect the size of the product contained in the package | Packaging process control | Cooper Surgical initiateda recall by letter on February 26, 2013, advising accounts to examine inventory and report quantity back to the firm for replacement. For questions customers should call Customer Service at 1-800-243-2974.For questions regarding this recall call 203-601-9825. |  | 80 | 80 units | Worldwide Distribution - USA (nationwide) and internationally to Australia, Canada, Colombia, Demarkt, France, Holland, New ZeaJand, andThailand. | Terminated | March 04, 2014 | 371 | 750 | Not_Computer
Z-0342-2014 | 66688 | No | JEB | Manual operating table and accessories and manual operating chair and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | TruSystem Surgical Table (TS7000)  Leg Section One Part Light HU |  |  | October 28, 2013 | 2013 | Trumpf Medical Systems, Inc. | The coupler point (hood) of the Leg Section One Part Light may become detached from the leg section frame. | Packaging | TRUMPF sent an Urgent Safety Information letter dated October 28, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.Customers were instructed to inspect Leg Section One Part Light for damage to the coupler point (hook) between each case. If any damage is noticed to the coupler, they were instructed to stop use immediately and contact TRUMPF Technical Service. If they drop or collide the Leg Section One Part Light into anything then stop use immediately and contact TRUMPF Technical Service. Review the User Instruction for proper use and care of the device. The leg section should be used only as intended as part of the TRUMPF TS7000 surgical table. The intended use is to support and position the patients legs during the pre-operative and post-operative stages of surgery. The leg section should be removed during surgery. TRUMPF will be exchanging the hooks, counter plates and screws of the coupler point to increase the area that the load is distributed over. This will allow the screws securing the hood to the leg frame to withstand external forces. TRUMPF anticipates parts to be available the first week of November. TRUMPF trained technicians will complete the TRUMPF Medical Systems, Inc. for the part exchange. Completion is anticipated by November 11, 2013.Further questions please call (843) 534-0606 |  | 15 | 15 components | US Distribution including the states of CA, IL, NJ, NY, NC, TX, WA and SC. | Terminated | September 16, 2014 | 323 | 15 | Not_Computer
Z-0064-2014 | 66248 | No | MBB | Polymethylmethacrylate (PMMA) bone cement. | Orthopedic | Orthopedic | 510(k) | SmartSetÂż Endurance GMV, Gentamicin, REF 3105-040, Medium Viscosity Antibiotic Bone Cement, Sterile. |  |  | September 16, 2013 | 2013 | DePuy Orthopaedics, Inc. | The SmartSetÂż GMV Endurance Gentamicin Bone Cement, product ref: number 3105-040 is intended for use in indications that have not been cleared by the US FDA. The intended use stated in the IFU for the SmartSetÂż GMV Endurance Gentamicin Bone Cement is for the fixation of prostheses to living bone in arthroplasty procedures of joints in which infection by gentamicin-sensitive organisms is a potentia | No Marketing Application | The firm DePuy Synthes, sent an "URGENT INFORMATION-MEDICAL DEVICE RECALL NOTICE" dated September 16, 2013 to its consignees via E-mail, hand delivery or fax depending on the most appropriate delivery method for each consignees. The Urgent recall notice described the recalled product, explained the reason for the recall, and directed consignees to cease further distribution, complete the attached Reconciliation form and quarantine and return all affected products through the DePuy Orthopaedics Sales consultant, Fax 1-787-287-7681 or Email to vbaez1@its.jnj.com. Consignees were directed to maintain a copy of the notice and forward the notice if product was further distributed. Clinical questions should be addressed to DePuy Orthopaedics, Inc's Scientific Information Office at 1-888-554-2482 (M-F; 8 a.m. - 5 p.m. EDT). Recall questions or concerns should be addressed to the Quality &amp; Compliance Supervisor, at 1-787- 272-1900 x3114 or vbaez1@its.jnj.com (MF; 8 am- 5 pm EST). |  | 16040 | 16,040 units | Distributed in Puerto Rico and US Virgin Islands. | Terminated | February 24, 2015 | 526 | 16040 | Not_Computer
Z-0074-2014 | 65799 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | NexGenÂż Femoral Component (CR-FLEX GSF POR FEM, C-L; CR-FLEX GSF POR FEM, C-R; CR-FLEX GSF POR FEM, C-L MINUS; CR-FLEX GSF POR FEM, C-R MINUS; CR-FLEX GSF POR FEM, D-L; CR-FLEX GSF POR FEM, D-R |  |  | August 30, 2013 | 2013 | Zimmer, Inc. | The low density polyethylene (LDPE) bag used to package implants adheres to the highly polished surface of the devices. | Package design/selection | August 30, 2013 Urgent Recall Notices were sent to Distributors, hospital risk managers and surgeons identifying the issue and directing consignees to make all affected personnel aware of the contents of the notification. The notice identified the products involved, the issue, possible adverse events associated with the issue and recommended that implants found with LDPE bag adhesion not use the product and return it to Zimmer sales representatives. Questions or concerns should be addressed to Zimmer customer call center at 1-877-946-2761 between the hours of 8 a.m. and 5 p.m. EST. |  | 11694 | 11694 | Worldwide Distribution: US (nationwide) and countries of: Argentina, Australia, Bolivia, Brazil, Canada, Chile, China, Colombia, Costa Rica, Dominican Republic, Ecuador, El Salvador, Germany, Guatemala, Honduras, Hong Kong, India, Israel, Jamaica, Japan, Korea, Malaysia, Mexico, New Zealand, Nicaragua, Panama, Peru, Singapore, Taiwan, Thailand, and Venezula. | Terminated | September 25, 2014 | 391 | 204220 | Not_Computer
Z-0360-2014 | 66720 | No | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Salter Labs EtCO2 Sample line 15' (4.6 M) tube, male luer connectors. |  |  | November 01, 2013 | 2013 | SALTER LABS | It has been determined that three lots of model number 4515-15-25 CO2 Sample Line 15' tube, male Luer lock connectors may contain a 30' tube instead of the 15' length. | Labeling Change Control | Salter Labs called the five customers that received product from the three lots, then follow-up immediately by sending a recall notification letter directly to each customer. Each letter will identify all shipments of subject product that they received.All customers letters were sent certified mail, return receipt requested. |  | 91 | 91 cases (25 units per case) | Distributed in the states of MS, TX FL,, WI, MI, VA, NY, MO, NC, and MS. | Terminated | December 17, 2013 | 46 | 91 | Not_Computer
Z-0114-2014 | 65997 | No | JDQ | Bone fixation cerclage. | Orthopedic | Orthopedic | 510(k) | Synthes Application Instrument Sternal Zip Flex System |  |  | August 13, 2013 | 2013 | Synthes USA HQ, Inc. | When using the 1st generation Zip Fix Application Instrument it is possible to cut the implant while the tensioning trigger is being squeezed which could lead to loosening of the implant. | Labeling False and Misleading | Synthes sent a Urgent Medical Device Notification dated August 13, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Action to be taken: This notification is to make you aware of the issue and the importance of completely releasing the tensioning trigger before cutting the implant. Please take the following action:"To ensure your receipt and understanding of this information, Please complete the verification section of this letter and return to Synthes by either of the following methods:" Fax: 1-888-352-9616 or" Scan/email: FieldAction@synthes.comIf you have any questions, please call 610-719-5450 or contact your Synthes Sales Consultant. |  | 319 | 319 | Worldwide Distribution - US (nationwide) and country of Mexico. | Open, Classified |  |  | 319 | Not_Computer
Z-0042-2014 | 66433 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Brand: The ForeCYTE Breast Health Test, Mammary Aspiration Specimen Cytology Test (MASCT), MASCT System Kits, Clarity System Kits, MASCT Patient Sample Kits &amp; Clarity Patient Sample Kits Instrument, Biopsy. |  |  | September 19, 2013 | 2013 | Atossa Genetics, Inc. | Atossa Genetics is recalling ForeCYTE Breast Health Test, Mammary Aspiration Specimen Cytology Test (MASCT) due to concerns raised by the U.S. Food and Drug Administration (FDA) in a warning letter received by Atossa in February 2013. The FDA raised concerns about (1) the current Instructions for Use (IFU); (2) certain promotional claims used to market these devices; and (3) the need for FDA cl | No Marketing Application | Atossa Genetics Inc. issued a press release to announce a voluntary recall on October 4, 2013.The firm started sending out the URGENT - VOLUNTARY PRODUCT RECALL NOTIFICATION ForeCYTE Breast Health Test; Mammary Aspiration Specimen Cytology Test (MASCT) RESPONSE REQUIRED letter to consignees on October 7, 2013. The letter identified the affected product, problem and actions to be taken. The firm initiated a voluntary recall to remove the ForeCYTE Breast Health Test and the Mammary Aspiration Specimen Cytology Test (MASCT) device from the market. The MASCT device has not been cleared by the FDA for the screening or diagnosis of breast cancer. In addition, the ForeCYTE Breast Health Test has not been cleared or approved by the FDA for any indication. The ForeCYTE Breast Health Test and the MASCT device are not a replacement for screening mammograms, diagnostic imaging tests, or biopsies. Patients should follow the recommendations and instructions of their physician with respect to breast cancer screening and diagnosis.Distributors are instructed to notify downstream customers and physicians are instructed to notify patients who have been tested using the MASCT and/or ForeCYTE systems. Consignees should return the Response Form back to Atossa.Distributors and customers should call 800-219-4629 to obtain a Returned Materials Authorization (RMA) number to return affected products. Customers with questions about this voluntary recall or Atossa's products, or would like assistance with the recall documentation may contact Atossa at 1-888-219-4629.Any problems experienced with the use of this product may be reported to the FDA at http://www.fda.gov/Safety/MedWatch/HowToReport/default.htm,Online (form available to fax or mail), or call FDA 1-800-FDA-1088. |  | 58277 | 577 MASCT System kits, 101 Lab Kits, and 2557 Patients Kits (5 test per kit) | USA Nationwide Distribution in the states of: AL, CA, FL, GA, IL, IN, KY, MD, ME, MI, MO, NC, ND, NJ, NV, NY, OK, PA, SC, TN, TX, UT, VA, WA, Washington DC, and Puerto Rico. | Terminated | August 19, 2014 | 334 | 58277 | Not_Computer
Z-1965-2013 | 65829 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Arrow International, Inc. Two-Lumen CVC Kit with Blue Flex Tip Catheter. The multiple lumen catheter permits venous access to the central circulation of patients. |  |  | July 31, 2013 | 2013 | Arrow International Inc | Arrow is recalling the product due to a labeling inconsistency. One portion of the lidstock states the kit contains a 3-Lumen catheter while the rest of the lidstock correctly identifies the catheter as 2-Lumen. | Labeling Change Control | An "Urgent Medical Device Recall Notification" letter dated July 31, 2013 was sent to direct accounts requesting immediate discontinuation of use and quarantine of all affected product on hand. Negative responses were also requested and the immediate return of product to Arrow was authorized. Questions were directed to Customer Service at 800 343-2935. |  | 156 | 156 | AL, PA, CT, NC, VA, KS, RI, AZ, NJ, and IL. | Terminated | April 28, 2015 | 636 | 156 | Not_Computer
Z-0128-2014 | 66423 | No | EWL | Laryngeal prosthesis (Taub design). | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | Blom-Singer Adjustable Tracheostoma Valve II |  |  | October 01, 2013 | 2013 | Helix Medical LLC | InHealth Technologies (IHT) is recalling the Blom-Singer Adjustable Tracheostoma Valve II (ATSV II) because the body component of the device is missing from the packaging. The ATSV II Valve product cannot be used as intended without the product body. | Packaging change control | A phone script and a customer notification letter are provided. The recall was initiated on 10/1/13 and customers will be notified that InHealth Technologies, a division of Helix Medical, LLC has initiated a recall of the following 3 lots of the Blom-Singer Adjustable Tracheostoma Valve (BE8025H and BE8025H-R3): 1057760, 1057919 and 1059222. The phone script provides the problem identified and the actions to be taken. Customers are instructed to complete the Customer response form and return by email to belinda.jackson@helixmedical.com, fax to (805) 566-5381 or by mail to InHealth Technologies, 1110 Mark Avenue, Carpinteria, CA 93013. Customers are also instructed to Contact UPS at 1-800-742-5877 and provide In Health Technologies UPS account no., 846718. UPS will come to your site, package the recalled product(s) and ship it to InHealth Technologies at InHealth Technologies expense. Customers are instructed to respond within 10 days. |  | 47 | 47 units | Worldwide Distribution-USA (nationwide) and the countries of Belgium, Canada, Croatia, Germany, Ireland, and Italy. | Terminated | January 21, 2014 | 112 | 47 | Not_Computer
Z-2243-2013 | 64406 | No | HTY | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Synthes Cannulated Distal Femoral Nail (Trauma Nail), Part Number 450.851S, 420MM-Sterile. Orthppedic implant device. |  |  | April 18, 2013 | 2013 | Synthes USA HQ, Inc. | Firm discovered that a nail that was within the scope of a previous Trauma Nail recall had been distributed. | Packaging process control | A letter dated April 18, 2013 was sent to one consignee who was identified as receiving the affected device. A review of inventory and return of any affected product on hand was requested. Negative responses were also requested. |  | 1 | 1 | US distribution to MA. | Terminated | March 30, 2015 | 711 | 1 | Not_Computer
Z-1995-2013 | 65834 | No | KKX | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Blue Belt Technologies, Inc. Monitor Drape REF: PFSD-V0001, Rx only, Sterile EO. |  |  | July 19, 2013 | 2013 | Blue Belt Technologies MN | Blue Belt Technologies is recalling sterile Monitor Drapes used with their Navio PFS system during orthopedic knee surgery. The drapes are used to cover the Navio PFS system computer monitor and arm assembly and provide a sterile barrier in the surgical suite. The defect is evident when the computer monitor is being draped. The front section of the drape separates from the rest of the drape at | Package design/selection | Blue Belt Technologies contacted consignees by telephone using a script on July 19, 2013. The consignees were informed of the problem and the product involved in the recall. Consignees were advised to immediate stop using the product and to locate and quarantine all unused product. Additionally, an "Urgent Medical Device Recall" letter dated July 24, 2013 was issued on July 24, 2013 and August 12, 2013. The letter includes the description of the Issue, description of safety concerns, affected product, and action to be taken by user. For questions contact Blue Belt Technologies, Lindsey Bring (+1 763.452.4922) or Robin Gonzalez (+1 763.452.4910) and get a Return Authorization number. Lindsey or Robin will arrange a pickup of the product. BBT will send replacement product at no charge. |  | 2400 | 60 cartons (20 pieces/carton) 40 in EU, 20 in US. | Worldwide Distribution and US Nationwide in the states of: CA and PA.United Kingdom. | Terminated | November 14, 2013 | 118 | 2400 | Not_Computer
Z-2158-2013 | 66029 | No | HTY | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Steinmann Pin, packed in a clear tube with end caps, tubes packed in a sterile barrier pouch, 6 pouches/sales unit. Pouches are labeled in part ***MicroAire Surgical Instruments 3590 Grand Forks Blvd Charlottesville, VA LLC 22911*** |  |  | August 12, 2013 | 2013 | MicroAire Surgical Instruments, LLC | A medical device used in surgical procedures was packaged in a pouch in which a defective seal may compromise sterility | Packaging process control | MicroAire Surgical LLC sent an Urgent Medical Device Recall letter dated August 12, 2013, to all affected customers. The letter advised consignees recover and quarantine the affected product so that it does not make it into the surgical process. Customers were instructed to either return affected product to MicroAire for inspection or they could inspect the product for the defect by looking at the seal area on the inside sterile package (short non-chevron side). The recalling firm further advised distributors to contact all of their customers to recover and quarantine the product so that it does not make it into the surgical process. Customers with questions were instructed to call 434-975-8370.For questions regarding this recall call 434-975-8370. |  | 15786 | 15786 total units | Worldwide Distribution - USA (nationwide) and internationally to Australia, Ireland, Canada, Netherlands, Sweden, Japan, and United Kingdom. | Terminated | February 12, 2014 | 184 | 15786 | Not_Computer
Z-0958-2013 | 64162 | No | CGX | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Creatinine, Part Numbers OSR6178, OSR6678 |  |  | December 28, 2012 | 2012 | Beckman Coulter Inc. | Beckman Coulter is including additional information and instructions in the Creatinine (OSR6x78) IFU and Setting Sheet, relating to calibration stability and quality control frequency to assure product performance. | Labeling design | Beckman Coulter sent an Urgent Product Correction letter to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter informed the customers that Beckman Coulter is including additional information and instructions in the Creatinine (OSR6x78) IFU and Setting Sheet, relating to calibration stability and quality control frequency to assure product performance. Customers are instructed to complete and return the enclosed response form. Customers with questions about the notification letter are instructed to contact the call center (hotline) at (800) 854-3633 in the US and Canada. |  | 57962 | 57,962 units (50,452 units in the US) | Worldwide Distribution - USA Nationwide and the countries of: Canada, Brazil, Colombia, Mexico, Puerto Rico, Venezuela, Guam, Honduras, Peru, Ecuador, Costa Rica. | Terminated | October 17, 2013 | 293 | 57962 | Not_Computer
Z-0186-2014 | 65739 | No | ONW | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Combiset Blood Line Hemodialysis use |  |  | August 29, 2013 | 2013 | Fresenius Medical Care Holdings, Inc. | Updated Instructions for Use: Bloodline connections not tightened per instructions in the operator's manual may disconnect and lead to blood loss | Labeling design | Fresenius issued a Field Notification Letter dated 8/29/13 to Patient Clinics/ associated at home patients and provides additional detail instructions to ensure the connection to the transducer protector bloodline female luer interface is secure.This field advisory notice sent by Certified Mail, Return Receipt Requested. Accounts requested to return both the receipt and the completed reply form (enclosed) to confirm receipt of the notice.If you have any questions regarding this field advisory, please contact FMCNA Customer Service Team at 1-800-323-5188. Customer Service will be available Monday-Friday 7am-6pm CST. |  | 508029 | 508,029 cases | Worldwide Distribution-USA (nationwide) and the country of Canada. | Open, Classified |  |  | 540759 | Not_Computer
Z-0917-2013 | 64367 | No | GEY | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | D3000I Advantage Drive System, 230v, surgical instrument motor. |  |  | January 24, 2013 | 2013 | Linvatec Corp. dba ConMed Linvatec | The Membrane Switch Panel (referred to as "front panel" of the component part number A34-003-207, used on both the D3000 and D3000I Advantage Drive System, contained the incorrect symbology per third edition IEC60606-01 Medical Electrical Equipment: General requirements for basic safety and essential performance. | Error in labeling | Linvatec Corp sent an "URGENT: Medical Device Recall Notification" letter dated January 24, 2013 to its customers after initial phone notifications. The letter describes the affected product, problem and actions to be taken. Customers were instructed to immediate check facility's inventory for the affected product, use the enclosed FedEx air-bill when returning the product, and complete and return to the Recall Reply Form. For questions contact your Customer Service department at 800-535-8536 or fax at 727-319-5701 or email Custserv1@linvatec.com. |  | 23 | 23 units | Worldwide Distribution - USA (nationwide) including the states of: CA, FL, GA, IL, KS, MA, MD, MO, NE, NY, and TX and the countries of: Canada, Germany, Romania, Spain, Palestine, South Africa, France, Sweden and Singapore | Terminated | July 05, 2013 | 162 | 54 | Not_Computer
Z-0800-2013 | 63433 | No | MBP | Resorbable calcium salt bone void filler device. | Orthopedic | Physical Medicine | 510(k) | Bacterin International, Inc., Osteoselect Demineralized Bone Matrix DMB Putty; Graft ID# B090056-664 Donor# B090056 2.5cc Expiration Date: 12/25/2011; LBL 210-0010-1 REV: 040109; and Graft ID# B090056-675. Bone void filler for voids or gaps in bone. Indicated for treatment of surgically created osseous defects. |  |  | April 18, 2012 | 2012 | Bacterin International, Inc. | Expired product was shipped to two customers. | Incorrect or no expiration date | Bacterin International phoned the two customers and they were instructed to return any of the unused product.For questions regarding this recall call 406-388-0480.The 3rd expired product was discovered by Bacterin Sales Rep 7/12/2012 at the consignee. |  | 3 | 3 packages each containing 2.5cc of putty | Nationwide Distribution in TX , GA and CA. | Terminated | May 13, 2013 | 390 | 3 | Not_Computer
Z-2165-2013 | 66082 | No | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Baxa Vial Adapter, 14mm - Luer Lock, Order No. REF: 90, Quantity: 50 each, Single Use. Product Code H93890. |  |  | August 22, 2013 | 2013 | Baxa Corporation | Baxter Healthcare Corporation is voluntarily recalling one lot of Baxa Vial Adapters, 14mm - Luer Lock, due to incorrect expiry date. | Labeling Change Control | An Urgent Product Recall notification was mailed to one customer via USPS First Class mail on 8/22/13. The one customer affected by this action was instructed to locate and remove the affected product from inventory and to return to Baxter for credit and/or replacement. Any customer returns received will be discarded per the Corporate Hold. |  | 34 | 34 cases | Distributed to One Customer in NC. | Terminated |  |  | 34 | Not_Computer
Z-1269-2013 | 64696 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | A) Product Code 2N3371: Non-DEHP Y-Type Catheter Extension Set, 5.7", Vol. 0.8 mL, 2 Injection Sites, Male Luer Lock Adapter, Single use only |  |  | March 28, 2013 | 2013 | Baxter Healthcare Corp. | Customers have reported that when separating an individual package from its attached grouping, the adjacent package has opened, compromising its sterile barrier properties. | Packaging process control | Baxter sent an "Urgent Product Notification" dated March 28, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taked. Customers are advised to locate and remove all of the affected products from their inventory. If the products were further distributed, consigneees are requested to notify their customers of the recall. Products can be returned for credit by contacting Baxter. An acknowledgement form is required to be compelted by customers. |  | 228200 | A) Product Code 2N3371: 228,200 units | Worldwide Distribution - USA Nationwide and the countries of Puerto Rico, Brazil, Canada and Columbia. | Terminated | July 31, 2017 | 1586 | 844300 | Not_Computer
Z-0138-2014 | 65782 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Stereotactic Circular Collimator - 3D Line stereotatic Hardware Accessories |  |  | July 23, 2013 | 2013 | Elekta, Inc. | Potential for clinical errors. | No Marketing Application | Elekta sent an Important Field Safety Notice dated 7/2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were advised of the issue and recommendation to resolve either by upgrading their system or to discontinue a clinical use of the device.For further questions please call ( 770 ) 300-9725. |  | 89 | 89 | Nationwide distribution: USA including the states of AZ, CA, FL, GA, MO, NJ, NM, NC, OR, TN and WV. | Terminated | November 18, 2014 | 483 | 89 | Not_Computer
Z-0918-2013 | 64343 | No | GAG | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | PSX PROXIMATE Skin Staple Extractor, |  |  | October 04, 2012 | 2012 | Ethicon Endo-Surgery Inc | Damage to the packaging may have compromised the sterility of the device. | Package design/selection | On 10/28/2012, the firm sent Urgent: Device Recall, Event 5928 notification letters to their customers. The letters identified the affected product, the reason for the recall, the action to be taken by customers, as well as the actions taken by the firm. Customers are to examine their inventory immediately and remove the affected product. The Business Reply Form should be completed and faxed to Stericycle at 1-877-650-0360 within 3 business days, even if there is no affected product on hand. Customers must complete the "Replacement Purchase Order Number Field" on the Business Reply Form to ensure adequate identification of the return product in their facility. Customers should follow the instructions provided in order to return the product. The information should be shared with all appropriate staff at the customer's facility. EES planned to replace any unopened an unexpired recall product that was returned by December 31, 2012. Questions regarding the issue should be directed to Stericycle at 1-877-650-5412. |  | 5894975 | 5,894,975 pieces | Worldwide Distribution -- USA, including the states of AK, AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, &amp; WY; and, the countries of ARGENTINA, AUSTRALIA, AUSTRIA, BAHRAIN, BELGIUM, BRAZIL, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, CZECH REPUBLIC, DOMINICAN REPUBLIC, ECUADOR, EGYPT, GERMANY, GREECE, GUATEMALA, HONG KONG, HUNGARY, INDIA, INDONESIA, ISRAEL, ITALY, JAMAICA, JAPAN, JORDAN, KUWAIT, LIBYAN ARAB JAMAHIRIYA, MALAYSIA, MEXICO, NETHERLANDS ANTILLES, NICARAGUA, NICOSIA, PAKISTAN, PANAMA, PERU, PHILIPPINES, POLAND, PORTUGAL, PROVINCE OF CHINA, PUERTO RICO, REPUBLIC OF KOREA, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SWITZERLAND, THAILAND, TRINIDAD AND TOBAGO, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY, VENEZUELA &amp;, YEMEN. | Terminated | June 11, 2013 | 250 | 5894975 | Not_Computer
Z-0943-2013 | 64342 | No | JDS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | INTERNAL HEX RECON SCREW, TI-6AL-4V, REF 7162380, TRIGEN (R) INTERNAL HEX CAPTURED RECON SCREW, QTY: (1), 6.4 MM, 80 MM, TITANIUM, Smith &amp; Nephew, Inc. Orthopedic. Bone screw for intramedullary rod fixation. |  |  | February 07, 2013 | 2013 | Smith &amp; Nephew Inc | 75 mm screws were packaged and labeled using an 80 mm screw outer box and label, while 80 mm screws were packaged and labeled using a 75 mm screw outer box and label. | Labeling mix-ups | Smith &amp; Nephew sales representatives and Smith &amp; Nephew international distributors were notified of problem and the recall via email on February 7, 2013. They were instructed to cease use and distribution of the product and quarantine for return. Non responding consignees were notified again on February 28, 2013 by email. Hospitals were notified by letter beginning on February 7, 2013. Smith &amp; Nephew also contacted the surgeons who implanted the screws by letter the week of February 11th.For questions regarding this recall call 978-749-1330. |  | 69 | 69 units | Worldwide Distribution - USA including NM, TN, TX, NC, CA, MI, KY, MA, NJ, OH, and HI. Internationally to Canada, United Arab Emirates, Mexico, Argentina, Japan, and Chile. | Terminated | January 11, 2016 | 1068 | 159 | Not_Computer
Z-0803-2013 | 64177 | No | ILO | Water circulating hot or cold pack. | Physical Medicine | Physical Medicine | 510(K) Exempt | BREG Cold Therapy Combination Units containing Sterile Polar Pads. 02356 Rev M, PAD M/U XL STER, Mfg.Date: 112012. Model #s: 10903, 10703, 09101, 09111, 09131, 09731, 09611, 09621, and 09631. Local anesthetic effect. |  |  | January 14, 2013 | 2013 | Breg Inc | The recall was initiated because Breg has determined that some Sterile Polar Pads products manufactured from January 2010 to October 2012 may have sustained damage to the product packaging which may compromise product sterility assurance. | Packaging | The firm, BREG, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated January 14, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately examine your inventory and quarantine the product; complete the Return Response Form by email, Fax or mail; and contact Breg for a Return Authorization for the affected product and request replacement by contacting Breg Customer Care at 800-321-0607. Note: response is required even if you have no affected inventory.Should you have any questions regarding this communication or need to report an adverse event, please contact Breg Customer Care at 800-321-0607. |  | 44883 | 44,883 units | Worldwide Distribution-USA (nationwide) and the countries of Australia, Singapore, Chile, Latvia. | Terminated | July 01, 2013 | 168 | 44883 | Not_Computer
Z-2277-2013 | 65998 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco RTP System Radiation Therapy Treatment Planning System. |  |  | May 22, 2013 | 2013 | Elekta, Inc. | Patient needs to be positioned as indicated by the treatment plan. | Labeling design | Computerized Medical Systems issued an Important Safety Notice dated May 22, 2013 to all affected customers. This letter idenitfies the problems the products and the actions needed to be takem by the custom. |  | 72 | 72 | Distributed in Puerto Rico and the states of AK, AZ, CA, CO, CT, FL, IL, IN, MA, MD, MI, MN, MO, NC, NJ, NM, NY, OH, OK, OR, PA, RI, SC, TN, TX, VA, and WA. | Terminated | December 23, 2013 | 215 | 72 | Not_Computer
Z-2278-2013 | 65009 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes 3.5 MM LCP(R) Distal Humerus System |  |  | March 08, 2013 | 2013 | Synthes USA HQ, Inc. | Recall is being initiated due to the part being mislabeled (Part Number 241.267 was mislabeled with Part Number 241.276). | Labeling mix-ups | Synthes sent an Urgent Notice Medical Device Recall letter dated March 8, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to obtain a Return Authorization Number, complete the Verification Section at the end of the letter by the checking the appropriate box indicating affected product located. Return the Verification Form with the product to: Credit/Returns, Synthes, 1101 Synthes Avenue, Monument, CO 80132. Customers with questions were instructed to call 610-719-5450 or email Fieldaction@synthes.com.For questions regarding this recall call 610-719-5000. |  | 2 | 2 | Nationwide Distribution including CA,and VA.. | Terminated | May 21, 2015 | 804 | 2 | Not_Computer
Z-2189-2013 | 65984 | No | DRS | Extravascular blood pressure transducer. | Cardiovascular | Cardiovascular | 510(k) | "***Flo Trac Sensor Model MHD85***" |  |  | July 31, 2013 | 2013 | Edwards Lifesciences, LLC | The Edwards Lifesciences FloTrac Sensor is being recalled due to a manufacturing non-conformance related to incorrect packaging configuration where the IFU was placed on top of the Tyvek instead of being placed at the bottom of the box. There is no evidence available at this time that could validate the sterility of this incorrect packaging configuration. | Packaging process control | Edwards Lifesciences initiated a US recall via an Urgent Product Recall letter dated July 31, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised that an Edweards Lifesciences representative would come to their facility to retrieve the affected device. A device replacement would be provided to the customer. Customers were asked to complete the enclosed Confirmation Form and give it to the Edwards Lifesciences representative that will come to their facility. For customers with questions were instruct4ed to call Edwards Customers Service at 1-800-424-3278.For questions regarding this recall call 972-410-7100. |  | 35 | 35 units | Nationwide Distribution including MO and MI. | Terminated | April 14, 2017 | 1353 | 35 | Not_Computer
Z-0809-2013 | 64110 | No | IOR | Mechanical wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Chenica Gunnell Rehab Series of Wheelchairs (GRSW); Mechanical Wheelchair, Sold as Rehab Tough and Tilt (TNT), Rehab Recline and Mobility (RAM), Rehab Kidster (Kidster). Provide mobility to individuals limited to a seating position. |  |  | January 18, 2013 | 2013 | CHENICA INC. | The wheelchairs may not meet the specifications of WC-19 (crash testing), therefore the WC-19 verbiage and corresponding logo are removed from the flyers on the Gunnell Rehab Series of Wheelchairs (GRSW). | Labeling False and Misleading | Gunnell Inc. sent a "RE: URGENT MEDICAL DEVICE CORRECTION/RECALL" letter dated January 18, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm 1-800-551-0055 for questions relating to this letter. |  | 53 | 53 | Nationwide Distribution-including the states of Florida, Michigan, South Carolina, California, Texas, Maryland, New York, Tennessee, Virginia, Illinois, Ohio, Massachusetts, Arizona and Washington. | Terminated | April 15, 2014 | 452 | 53 | Not_Computer
Z-0795-2013 | 63858 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | The Lumenis VersaCut Tissue Morcellator |  |  | January 02, 2013 | 2013 | Lumenis Limited | Lumenis Ltd is recalling certain models of VersaCut Morcellator devices to perform corrective labeling as a result of three reported events of air embolism secondary to incorrect reversal tubing system hook-up. | Labeling design | Lumenis sent an Urgent Safety Advisory Notice letter dated January 2, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to re-label the devices with the enclosed labels provided, replace the Operation Manual with the current revised Manual, complete the information requested on the Confirmation Card and return the Confirmation Card using the enclosed pre-paid shipping label by UPS. For question contact VersaCut Morcellator field correction administrator by either Cell: +972-(54)-9599196, Tel: +972-(4)-9599296 or E-mail: VersaCutMorcellator@lumenis.com. |  | 931 | 931 units | Worldwide Distribution - US Nationwide including: CA, WI, NC, FL, PA, IN, NM, NJ, IA, WA, OH, FL, MD, MA, WY, NY, AZ, NH, MN, TN, OK, KY, TX, AK, MI, VT. Internationally: Japan, Nepal, India, UK, South Korea, Netherlands, Germany, Australia, China, Canada, Turkey, Spain, Italy, Egypt, Mexico, Bulgaria, Utd. Arab, Emir., Poland, Kuwait, Hong Kong, Saudi Arabia, Yemen, Qatar, Czech Republic, New Zealand, Malaysia, Portugal, France, Argentina, Singapore, Romania, Chile, Colombia, Switzerland, Bahrain, Brazil, Serbia/Montenero, Taiwan, Israel, Indonesia, Peru, Pakistan, Nepal, Belarus, Latvia, Sri Lanka, Russian Fed, Thailand, Azerbaijan, Belgium, and Denmark. | Terminated | May 20, 2013 | 138 | 931 | Not_Computer
Z-0176-2014 | 66191 | No | NGX | Powered muscle stimulator. | Physical Medicine | Physical Medicine | 510(k) | Pinook Mini Massager |  |  | May 23, 2013 | 2013 | Pinook-Usa | Failure to submit a premarket submission and gain approval of a medical device. | No Marketing Application | On October 18, 2013 the firm sent the following notification to their customers:This is to inform you of a product recall involving: Pinooks Mini-massager and Pinooks Micro Massager model number PT-11 See enclosed product label (for ease in identifying the product at retail/user level). This recall has been initiated due to failure to submit a premarket submission and gain approval.We began shipping this product on 11/27/2012. Immediately examine your inventory and quarantine product subject to recall. In addition, if you may have further distributed this product, please identify your customers and notify them at once of this product recall. Your notification to your customers may be enhanced by including a copy of this recall notification letter.Pinook USA is removing the products (mini and micro massager) from the market until receiving FDA approval to market them again under a new 510(k).If you have stock on your end please return them to our office:901 C FL PKWY STE A6, Orlando FL, 32824.Before returning, please contact Sheva Ben-Shushan: accounting@pinookusa.com in order to receive an RMA number.Once you have RMA number, place that number on the package label.This recall should be carried out to the retail level. Your assistance is appreciated and necessary to prevent patient harm. If you have any questions, call Dvir Lev-Ran, telephone number: 407-854-0055 Ext 304. This recall is being made with the knowledge of the Food and Drug Administration. |  | 28994 | 28,994 units (Mini and Micro) | Worldwide Distribution-USA (nationwide) including Puerto Rico and Guam and the states of AZ, CA, CO, FL, IL, MA, MD, MO, MT, NC, ND, NJ, NV, NY, TX, VA, and WA and the country of Canada. | Open, Classified |  |  | 28994 | Not_Computer
Z-2205-2013 | 66143 | No | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | C2244.4308 CONELOG Titanium Base CC Âż4.3GH2 0.8 inkl HS |  |  | August 28, 2013 | 2013 | Camlog Usa | Product not cleared for sale/use in the United States. | No Marketing Application | Camlog Usa notified four US consignees of the recall by telephone on August 5, 2013. The notification identified the recalled product, which did not have proper marketing clearance for distribution in the United States and asked customers to return all unused product. Cusotmers were sent an overnight UPS label via e-mail and asked customers to return the device to CAM:LOG USA Attn: Debby Callahan, 1822 Aston Avenue Carlsbad, CA 92007. |  | 6 | 6 to US customers | USA Distribution in the states of NH, MD, CO, and IL. | Terminated | January 29, 2014 | 154 | 12 | Not_Computer
Z-2209-2013 | 66155 | No | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | Blanketrol III, Model Number 233, Cincinnati Sub-Zero Products, Inc., 12011 Mosteller Road, Cincinnati, OH 45241 |  |  | September 03, 2013 | 2013 | Cincinnati Sub-Zero Products Inc | During a label review it was discovered that both the Blanketrol III Operations (57201-N) &amp; Operations &amp; Technical (56201-Q) manuals had verbiage added back into the manual during ECN M111-4260 (released 12/29/2011) which was previously removed for clarification of the intended use and subsequent 501(K) approval. | Labeling Change Control | Cincinnati Sub-Zero (CSZ) sent an Urgent Medical Device Field Action letter dated September 5, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to contact CSZ at 513-772-8810 or 1-800-989-7373 to request new versions of the manuals. At that time customers are instructed to provide the quantity and serial numbers of the affected devices that are at their facility. Upon receipt of the new versions of the manual, immediately collect and destroy/discard all the affectred devices Operation and Technical Manuals their facility is currentlyin possession of and replace the discarded manuals with the new versions of the manuals provided by CSZ. After the manuals have been discarded and replaced, customers were asked to complete and return the enclosed response form as soon as possible to acknowledge receipt of this notification and to inform CSZ that they have performed and completed the requested actions. Customers should return the form by fax to 513-772-9119 scan and email the information to B3manual@cszinc.com or mail to Cincinnati Sub-Zero, 12011 Mosteller Road, Cincinnati, Ohio 45241. Customers with questions were instructed to contact their Blanketrol III Model 233 field action coordinator at 1-800-989-7373 or 513-772-8810.For questions regarding this recall call 513-772-8810, ext 5295. |  | 1126 | 115V - 396 units; 230V - 385 units | Worldwide Distribution - USA including AK, AR, AZ, CA, CO, DC, DE, FL, GA, IA, ID, IL, IN, KY, LA, MA, MD, ME, MI, MN, MT, , NC, ND, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, SC, TN, TX, UT, VA, WA, WI &amp; WV. Internationally to ARGENTINA, AUSTRALIA, AZERBAIJAN, BELGIUM, CANADA, CHILE, CHINA, COSTA RICA, CROATIA, GERMANY, HONDURAS, HONG KONG, INDIA, INDONESIA, ISRAEL, KOREA, KUWAIT, MEXICO, NEW ZEALAND, PERU, SAUDI ARABIA, SINGAPORE, SPAIN, TAHITI, THAILAND, THE NETHERLANDS, THE PHILIPPINES, UNITED ARAB EMERITES &amp; UNITED KINGDOM. | Terminated | August 22, 2014 | 353 | 1126 | Not_Computer
Z-2202-2013 | 66027 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | Sonogage EyeScan, A-Scan Biometer, |  |  | August 09, 2013 | 2013 | Sonogage Inc | During an FDA inspection on 7/15-26/2013, it was discovered that the firm lacked registration and 510(k) clearance for the device. | No Marketing Application | Sonogage sent an Urgent: Medical Device Recall letter dated August 9, 2013 to their customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immediately examine inventory and quarantine the affected product for return. Customers were asked to complete and return the enclosed response form. For question call 216-464-1119 or 800-798-1119. |  | 564 | 564 units | Worldwide Distribution - USA Nationwide in the following states: AK, AZ, CA, FL, GA, IL, KY, LA, MA, MD, MN, NE, NJ, NM, NY, OH, OR, PA, TX, VA and WA. Product was also shipped to Canada and Mexico. | Terminated | September 19, 2015 | 771 | 564 | Not_Computer
Z-0166-2014 | 66425 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | OmniPod Insulin Management Product Demonstration Kit; Model Reference Number: (PDK-ZXP490) |  |  | September 27, 2013 | 2013 | Insulet Corporation | Alcohol Prep Pads contained within the kit are not properly labeled. | Labeling design | Insulet sent an Important Corrective Action Notification letter dated September 27, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers are asked to discontinue use and properly dispose of the affected product. For questions contact Customer Care at (866) 941-0294. |  | 5050 | 5,050 kits | Worldwide Distribution - USA Nationwide and the country of Israel. | Terminated | March 04, 2014 | 158 | 6769 | Not_Computer
Z-2260-2013 | 66186 | No | HGB | Contraceptive tubal occlusion device (TOD) and introducer. | Obstetrics/Gynecology | Obstetrics/Gynecology | PMA | Cranial Access (CRAK) Kits, Product / Catalogue No.s: INSHITH, INSHITHND, INSHITHRZN |  |  | September 04, 2013 | 2013 | Integra LifeSciences Corp. d.b.a. Integra Pain Management | Integra is recalling certain lots of Cranial Access Kit due to a possibility that the seal integrity of the sterile package may become compromised during shipping. | Packaging | Integra sent an Urgent Voluntary Medical Device Recall letter dated September 4, 2013, to all affected customers that have been shipped unexpired affected product lots of IntegraÂż Cranial Access Kits, Catalogue No.'s INSHITH, INSHITHND, INSHITHRZN.The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to review their inventory and immediately stop the use or distribution of any affected product. Customers were instructed to identify and report to Integra if they do or do not have any of the affected product lots. Customers were advised how to return any affected product lot numbers in their possession.Effectiveness Checks will be conducted to determine the level of customers who:- Integra has verified, have been notified, and / or acknowledged the recall either by returning the acknowledgement form.- Indicate they will be returning / have returned unexpired affected product lots of kits.Any affected product lot numbers that are returned from consignees will be forwarded to the Integra Salt Lake City, UT facility for final disposition per its site QMS processes.Customers with questions were advised to contact Customers Service at 1-855-532-1723.For questions regarding this recall call 801-886-9505. |  | 8648 | 8,648 | Nationwide Distribution including CA, CO, FL, IN, KS, NC, NV, NY, OK, TN, TX, VA, WA, and WI. | Terminated | January 28, 2014 | 146 | 8648 | Not_Computer
Z-0811-2013 | 64272 | No | ELW | Impression material. | Dental | Dental | 510(k) | Ultradent Products, Chromaclone PVS, Regular Set heavy body, Bubble Gum Flavor, 50 ml, Impression Material, Part No. 5012, Lot No. B7PPC. Used in prosthodontics. |  |  | February 04, 2013 | 2013 | Ultradent Products, Inc. | Incorrect packaging of regular set and fast set cartridges. | Packaging process control | Ultradent Products, Inc. sent an Urgent Medical Device Correction letter dated February 4, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their inventory for the affected products. They may continue to use product or dispose unwanted cartridges and reply back to firm indicating acknowledgement of letter and if replacement is needed. Consumers with questions can call the firm at 801-553-4420.For questions regarding this recall call 801-553-4420. |  | 32 | 32 cartridges | Nationwide Distribution | Terminated | May 30, 2013 | 115 | 132 | Not_Computer
Z-2215-2013 | 66037 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | BIORAPTOR 2.9 mm Suture Anchor with 2 ULTRABRAID Suture; Part Number: 72200775 |  |  | August 06, 2013 | 2013 | Smith &amp; Nephew, Inc. Endoscopy Division | Pinholes detected in a small number of primary pouches that may constitute a breach of the sterile barrier. | Packaging | Smith&Nephew; sent an Urgent Product Recall Notification letter dated August 6, 2013, via Federal Express to all affected customers. Addendum customer letters for 4 additional part numbers were sent via Federal Express on August 19, 2013, The letters identified the product, the problem, and the action to be taken by the customer. Customers were asked to inspect their inventory, locate unused devices and quarantine them. Customers were also asked to complete the last three columns in the inventory Return Certification form indicating the quantities that need to be returned, and include their contact information in the spaces provided. Customers were instructed to contact Smith &amp; Nephew's Demo Service Departmentat 800-343-5717 for a Return Authorization (RA#), and provide the information response form. Replacement product would be provided. For questions regarding this recall call 978-749-1073. |  | 43385 | 43385 US | Worldwide Distribution - USA (nationwide) Canada, Europe: Austria, Belgium, Cyprus, Czech Republic, Denmark, Germany, Greece, Finland, France, Italy, Ireland, Luxembourg, Netherlands, Norway, Poland, Portugal, Russia, Slovenia, Spain, Sweden, Switzerland, Turkey, Ukraine and United Kingdom, Asia-pacific: Australia, China, Dubai, Hong Kong, India, Japan, Malaysia, New Zealand, Philippines, Sri Lanka, Singapore, South Korea, Taiwan, Thailand, Argentina, Bermuda, Brazil, Chile, Colombia, Costa Rica, Israel, Mexico, Puerto Rico, South Africa, and Venezuela. | Terminated | April 20, 2016 | 988 | 224222 | Not_Computer
Z-0939-2013 | 64447 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Draeger Infinity Acute Care System (IACS) Monitoring Solution. Catalog Number MS20783 (Infinity MCable-Dual Hemo) |  |  | December 13, 2012 | 2012 | Draeger Medical, Inc. | The displayed value for the invasive blood pressure measurement of the IACS was different than the actual patient values. Liquid ingress in the cable connection may cause inaccurate measurements. This can happen if the Dual Hemo MCable is used in a non-vertical position, a vertical position but upside down, or laying it on the floor or on the patient bed. | Labeling design | Drager sent an "URGENT MEDICAL DEVICE RECALL" letter dated December 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at 1-800-543-5047 for questions relating to this notice. |  | 260 | 260 | Distributed in the states of CA, IL, and MA. | Terminated | May 20, 2013 | 158 | 260 | Not_Computer
Z-2255-2013 | 66185 | No | CAE | Oropharyngeal airway. | Anesthesiology | Anesthesiology | 510(K) Exempt | King LT-D Oropharyngeal Airways |  |  | September 03, 2013 | 2013 | King Systems Corp. | Report that size 5 KLTD airways were packaged in pouches indicating a size of 2.5 and shipped to your facility. The correct size of the KLTD airway is clearly marked on the actual airway product and the physical size difference between 2.5 and 5 is distinct and obvious. | Labeling mix-ups | King Systems notified one consignee who received the affected product verbally on August 7, 2013, via email and an Urgent Medical Device Recall - Immediate Action Required notification on September 3, 2013. The Urgent Medical Device Recall - Immediate Action Required notification described the product and issue as well as includes a response form for the consignee to complete and send back to King Systems via FAX to 317-776-5175, or email to wnet@ambu.com. Questions regarding the notice are directed to: King Systems customer service representative at 800.642.5464, Monday through Friday 8:00am- 5pm EDT.For questions regarding this recall call 317-776-6823. |  | 54 | 9 cases; 45 products | Distributed in NY. | Terminated | April 15, 2014 | 224 | 54 | Not_Computer
Z-0932-2013 | 62993 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Cynch Lordotic T-PLlF Implant. Intervertebral fusion device with bone graft, for use with supplemental fixation. |  |  | August 24, 2012 | 2012 | Spine Smith Partners LP | Implants were incorrectly laser marked as 23mm instead of 28mm. | Labeling mix-ups | Spine Smith sent a Field Action Notice dated August 24, 2012 to its affected customer. The letter identified the product, problem, and actions to be taken by the customer. Contact the firm at 512-637-2097 for questions concerning this notice. |  | 3 | 3 units | Distributed in Texas. | Terminated | May 17, 2013 | 266 | 3 | Not_Computer
Z-0173-2014 | 66132 | No | CEW | Parathyroid hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA Centaur Intact PTH Calibrator (6 pack); Catalog Number: SMN 10492387 |  |  | August 16, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc | PTH Calibrator packaged with only high calibrator. | Packaging | Packaging error was brought to Siemens' attention through escalations received from (3) customers in the United States. Internal Investigation confirmed the issue was limited to the five (5) kits distributed to these customers. As the customers were aware of the issue and non affected product was provided to the customers to use during Siemens' investigation, no further communication was warranted. |  | 595 | 595 kits total (5 kits with packaging error) | Nationwide Distribution including MD, PA, and CT. | Terminated | February 13, 2014 | 181 | 595 | Not_Computer
Z-2254-2013 | 66223 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | POWERWAND Safety Introducer with Extended Dwell Catheter; |  |  | August 16, 2013 | 2013 | Access Scientific Inc | Access Scientific is recalling the POWERWAND Saftey Introducer with Extended Dwell Catheter because it is mislabeled. | Labeling Change Control | On 8/16/13, Access Scientific visited 2 consignees and initiated a field correction. This is the communication given to the two(2) accounts by our field personnel: "Access Scientific has inadvertently shipped mislabeled devices to your facility. The products affected were 5 Fr. POWERWANDs, Catalog # 92006, lot # I666473."The consignees were instructed to "Quarantine the devices until Access Scientific personnel arrived with the appropriate labeling and performed the corrective action of applying the correct labels to the devices and inserting the DFU into the boxes." |  | 160 | 160 units | Distributed in MI and TX. | Terminated | September 24, 2013 | 39 | 160 | Not_Computer
Z-0174-2014 | 66597 | Yes | OAD |  |  | Cardiovascular | PMA | US Safire Duo Ablation Catheter MediGuide Enabled, REF A700240 and Cool Path Duo Ablation Catheter, MediGuide Enabled REF A700244, Rx only, Sterile EO. |  |  | October 11, 2013 | 2013 | St. Jude Medical | St. Jude Medical Cardiovascular and Ablation Technologies Division is performing a voluntary notification regarding selected batches of Coolpath Duo Ablation Catheter MediGuide Enabled products and Safire Duo Ablation Catheter MediGuide Enabled. These units were inadvertently shipped with the incorrect version of the Instructions for Use (IFU). | Error in labeling | The firm, St. Jude Medical (SJM), sent a "Urgent Medical Device Notice" dated 11 October 2013 to its consignees. The letter described the problem, the product involved in the recall and the actions to be taken. The consignees were instructed to discard the IFU received with the batches and replace it with the IFU received in conjunction with the notification letter. SJM has determined that it is acceptable to continue using these devices as long as the product is used in accordance with the IFU.If you have any questions, contact Vice President, Quality, at 651-756-5402. |  | 59 | 59 | US distribution: KS, MA, UT,TX, KY, OH. | Terminated | March 10, 2016 | 881 | 59 | Not_Computer
Z-2162-2013 | 65905 | No | FQZ | Clinical color change thermometer. | General Hospital | General Hospital | 510(K) Exempt | Tempa DOT Single-Use Clinical Thermometer |  |  | July 10, 2013 | 2013 | Medical Indicators, Inc. | TempaDot products may have been mislabeled. | Labeling mix-ups | The firm, Medical Indicators Inc.sent an "URGENT" Recall letter dated July 10, 2013 and Acknowledgement Form to all distributors/customers via email and Priority Mail with delivery confirmation. The letter described the product, problem and actions to be taken. The customers were instructed to check full &amp; partial cases of the tempadot; open and inspect all 5122NS cases; pull out and quarantine any of the lots found in the 5122NS cases, pending return to medical indicators; if individual Tempadot 100-count boxes of 5122 and 5122NS are stored in different warehouse follow the instructions provided in the letter; if Tempadot 100-count boxes of 5122 were sold to a customer with the affected lots, contact the customer and have them return the product; and complete and return the Recall Return Response Acknowledgement and Receipt Form via mail to: 16 Thomas J Rhodes Industrial Drive, Hamilton NJ, 08619; fax: (609) 587-5636 or email to: www.medicalindicators.com or General Manager at: pbaker@medicalindicators.com.Contact information for any question(s) related to this recall may be addressed to General Manager of Medical Indicators, Inc. at 609-737-1600, extension 204 or email to: pbaker@medicalindicators.com, Mondays through Fridays from 8:00am to 5:00pm EST. |  | 111800 | 111,800 boxes (5,590 full boxes) | Nationwide distribution: USA including states of: CA, GA, FL, IN, NY, NV OH, PA, TN,TX, VA, and WI. | Terminated | January 16, 2014 | 190 | 111800 | Not_Computer
Z-0956-2013 | 64194 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Variable Angle Locking Compression Plate/VA-LCP Dorsal Distal Radius Plate, Part Number 02.115.431. |  |  | January 11, 2013 | 2013 | Synthes USA HQ, Inc. | It was discovered that one of the VA-LCP Dorsal Distal Radius Plates was incorrectly etched as Part Number 02.115.431 (110 degree) instead of the correct Part Number 02.115.231 (90 degree). | Labeling mix-ups | The firm initiated their recall of this product on January 11, 2013 by sending an Urgent Notice: Medical Device Recall letter to consignees. The letter identified the affected product and the issue. Synthes requested customers to examine their inventory for the affected product, remove them from use, and return them to Synthes. Additionally, customers were to follow the steps provided in the event that they did or did not have any affected product on hand. Questions should be directed to 610-719-5450 or FieldAction@synthes.com. |  | 4 | 4 | Distributed in the states of IN, OH, PA, and TX. | Terminated | August 31, 2015 | 962 | 4 | Not_Computer
Z-0027-2014 | 66346 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | P21x/5-1 MHz transducer, Part Number P07698-23 |  |  | September 11, 2013 | 2013 | SonoSite, Inc. | FUJIFILM SonoSite Inc. remove P21x/5-1 Transducer because the firm discovered the Biopsy Bracket Sensor is not working as intended. | Error in labeling | SonoSite started calling their consignees on September 13, 2013. For customers who could not be reached via phone, a letter Urgent -Medical Device Removal will be sent on September 17, 2013. Affected customers will be offered a replacement transducer (P07698-24) free of charge. Consignees in the U.S. can call 1-877-657-8118, Monday-Friday, 6:00 a.m. - 5:00 p.m. (PST) or send E-mail to service@sonosite.com. Customers outside the US can call 425-951-1200. |  | 10 | 10 units | Worldwide Distribution - USA Nationwide in the states of Florida, Georgia, Oklahoma, Pennsylvania, Tennessee, Texas, West Virginia, and in the country Japan and Italy. | Terminated | December 11, 2013 | 91 | 10 | Not_Computer
Z-2272-2013 | 66184 | No | CAI | Anesthesia breathing circuit. | Anesthesiology | Anesthesiology | 510(K) Exempt | Universal Flex2 Breathing Circuit |  |  | August 23, 2012 | 2012 | King Systems Corp. | King Systems received a customer complaint indicating that the customer had ordered a breathing circuit with a Latex breathing bag. The product they received contained a Latex breathing bag, but was labeled as non-latex. | Labeling mix-ups | Ambu, Inc. sent a letter dated June 24, 2013, informing the customer that they have decided to discontinue the manufacture and sales of the King Systems latex breathing bags and to conver all customers to their King Systems latex-free breathing bags. The customer was advised to adjust their inventories accordingly and transition to the latex free part number as their inventory levels reach their minimums. For questions the customer was advised to call 317-403-8677.For questions regarding this recall call 317-776-6823. |  | 41 | 1 case; 40 units | Nationwide Distribution - NE only. | Terminated | November 18, 2013 | 452 | 41 | Not_Computer
Z-0891-2013 | 64434 | No | GDX | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Technocut Plus, Sterile Single Use Safety Scalpel, Distributed by: Myco Medical. |  |  | February 07, 2013 | 2013 | Myco Medical Supplies Inc | Pouch labeling/printing errors and defective packaging resulting in possible compromised sterility. | Packaging | MYCO Medical sent an Advisory Notice dated February 7, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers wsere instructed to review their inventory for the affected product and immediately quarantine the affected product per their internal procedures. Customers were also instructed to complete the attached Material for Return form and return to MYCO via email QA@mycomedical.com or fax to 919-800-3919. Customers were have the returned material packaged and ready for UPS to pick up at their location. If product was shipped to other customers, customers were asked to contact a MYCO Customer Service Representative with the additional customer's contact and ship-to information. Customers with questions were instructed to call 919-460-2535.For questions regarding this recall call 919-460-2535, ext 105. |  | 110768 | 110,768 scalpels in individual pouches | Worldwide Distribution - USA including CA, FL, IL, MI, MO, NY, OH, PA, SC, TN, WI and Internationally to Canada | Terminated | July 09, 2013 | 152 | 110768 | Not_Computer
Z-1998-2013 | 65413 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | 00597001101 Femur CR PRECOAT FEM COMP SIZE AML, Rx, Sterile; |  |  | June 07, 2013 | 2013 | Zimmer, Inc. | Surgeons are implanting components of the NexGen Cruciate Retaining (CR) Complete Knee System in combinations that are not approved (micro with standard). Includes CR Micro Femurs, Articular Surface, Patellae, instruments and provisionals combined with standard components. A March 2012 recall reiterated the proper combination labeling, but the mis match problem continues. Zimmer is removing the d | Labeling design | Beginning June 7, 2013 all distributors will be sent Urgent Device Recall notices via electronic mail, with distributors, hospital risk managers with affected inventory also receiving the notice via courier. The Notice outlined the products recalled; the reason for the recall, possible adverse events related to the recall and directed consignees to locate, quarantine and count all affected product recording the information on the Inventory Return Form acknowledging that recall directions were followed. Further questions should be directed to 1-877-946-2761. Affected product is to be returned to Zimmer Product Service Department , 1777 West Center St. Warsaw, IN 46580 or internation return at Zimmer International Logistics GmbH, Max - Immelmann Allee 12 79427 Eschbach Germany."Communications outside of the United States will occur approximately two weeks after the United States communications. |  | 192355 | 192,355 all devices | Worldwide Distribution-USA (nationwide) and the country of Japan. | Terminated | December 15, 2014 | 556 | 192355 | Not_Computer
Z-0129-2014 | 66196 | No | NDG | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Ti-Magna Fx Washer |  |  | August 28, 2013 | 2013 | Zimmer, Inc. | Zimmer recently conducted a review of historic packaging validations completed. Based on this review, it was determined that a subset of product packaged at a single Zimmer facility required testing to ensure that previously packaged product from this facility met the established requirements for sterile barrier integrity. The number of samples required for packaging testing, to provide statisti | Packaging process control | The firm, Zimmer Inc., sent a letter entitled "Voluntary Recall of Ti-VFxII Condylar Tube Plates and Ti-Washers" dated September 12, 2013 to its customer. The letter described the product, problem and action to be taken. The customer was instructed to immediately discontinue use of the listed products. A Sales representatives from Zimmer K.K. or agencies will be in contact with you. Zimmer Inc. is planning to resume product manufacturing upon the completion of package validation review. Zimmer K.K. will contact you when a detailed schedule has been determined.If you have any questions, contact Zimmer K.K. Safety Office (9:00 to 17:00; closed weekends and public holidays) TEL: 0120-015-180. |  | 2632 | 2,632 devices | International distribution: Japan only. | Terminated | April 30, 2015 | 610 | 2632 | Not_Computer
Z-0973-2013 | 64633 | No | NDP | Endosseous dental implant accessories. | Dental | Dental | 510(K) Exempt | PrimaConnex Tapered Implant TC, SD Internal Connection 3.5 x 10mm |  |  | March 11, 2013 | 2013 | Keystone Dental Inc | Mislabeled: Cover screw that was in the package was a WD cover screw and not a SD cover screw as labeled | Packaging process control | Keystone Dental initially called customers on 3/11/13 and forms were completed to document these calls. In addition, a letter sent via Federal Express/or hand delivered to all distributors and customers. Product requested to be returned. |  | 120 | 120 units | Worldwide Distribution-USA including DC and the states of CA, FL, IN, MA, MN, NJ, NY, NH, PA, WI, and NY, and the countries of Republic of China, Germany, and France. | Terminated | September 26, 2013 | 199 | 120 | Not_Computer
Z-1996-2013 | 65738 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | GENESIS(R) II CRUCIATE RETAINING ARTICULAR INSERT, TRIAL, SIZE 3-4, A/P 48 MM, M/L 68 MM, QTY 1, NON-STERILE, REF 71430490, |  |  | July 08, 2013 | 2013 | Smith &amp; Nephew Inc | The products were produced using the wrong color material, resulting in the devicesbeing red (used for deep dished inserts) instead of blue (used for cruciate retaining inserts). | Labeling mix-ups | Smith &amp; Nephew sent an Urgent Recall Notification letter dated July 8, 2013 to all affected customers. Surgeons were notified by letter beginning on July 16, 2013. The letter identified the affected product, problem and actions to be taken. Customers were instructed to inspect inventory, locate unused devices, quarantine immediately, and return affected product to the address provided in the notification. For questions email to FieldActions@smith-nephew.com. |  | 9 | 9 units | International only: Germany, France, United Arab Emirates, and United Kingdom. | Open, Classified |  |  | 9 | Not_Computer
Z-0974-2013 | 64435 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | INTIO Inc., ClearStart-SVM Segmentation and Volumetric Measurement System, Users Manual. Part #: UM-06-20-0003. ClearStart-SVM is used for patient assessment using CT (computed tomography) images, in order to plan potential therapies or monitor tumor response to treatment. |  |  | February 26, 2013 | 2013 | Intio Inc | INTIO Inc. distributed a SVM User's Manual and it has an editing mistake that could be interpreted as promoting an unapproved use. | Labeling False and Misleading | INTIO sent a Recall Notice to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. INTIO has corrected this version of the SVM User's Manual. An INTIO Inc. employee will contact the customers to replace the current manual with the correct one. Consignees were notified by letter that an INTIO Inc. employee will contact you within the next 30 days to replace the current manual with the correct one. The new SVM User's Manual will be provided to you once INTIO Inc. has retrieved the incorrect version of the SVM User's Manual. Customers with questions were instructed to email support@intio.us or call 303-396-1777.For questions regarding this recall call 303-396-9291. |  | 7 | 7 | Nationwide Distribution including MN, FL, PA, WI, NY, and RI. | Terminated | April 17, 2013 | 50 | 7 | Not_Computer
Z-0020-2014 | 66047 | No | CEO | Phosphorus (inorganic) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Clinical Chemistry Phosphorus Reagent |  |  | August 12, 2013 | 2013 | Abbott Laboratories, Inc | Reagent lots 64736UN12, 03991UN13, 51611UN13, 74805UN12, and 03992UN13 have incorrect expiration dates assigned by Supplier. | Error in labeling | The firm, Abbott Laboratories, sent a "PRODUCT CORRECTION" letter dated August 12, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to update existing inventory and manually track the implanted lots with the corrected expiration dates listed in the letter; retain a copy of letter and if you have forwarded any of the lots listed to another laboratory, provide a copy of the letter to them.If you or any of your healthcare providers you serve have any questions regarding this information, U.S. Customers please contact Customer Service at 1-877-4ABBOTT (422-2688). Customers outside the U.S., please contact your local area customer service. |  | 7213 | 7213 | Worldwide distribution: US (nationwide) including Puerto Rico and countries of: Albania, Algeria, Argentina, Austria, Australia, Azerbaijan, Bangladesh, Bahamas, Belgium, Bosnia &amp; Herzegovina, Brazil, Bulgaria, Burkina, Canada, Cayman Islands, Chile, China, Columbia, Croatia, Curacao, Cyprus, Czech Republic, Denmark, Dominican Republic, El Salvador, Egypt, Estonia, Ethiopia, Finland, France, Germany, Greece, Guatemala, Hong Kong, India, Iraq, Ireland, Israel, Italy, Ivory Coast, Jamaica, Jordan, Kazakhstan, Kenya, Kuwait, Latvia, Lebanon, Libya, Lithuania, Luxembourg, Malaysia, Mali, Mauritania ,Mexico, Moldova, Morocco, Netherlands, New Zealand, Norway, Oman, Pakistan, Panama, Paraguay, Philippines, Poland, Portugal, Qatar, Republic of Yemen, Romania, Russia, Saudi Arabia, Senegal, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Tajikistan, Taiwan, Tanzania, Thailand, Trinidad &amp; Tobago, Tunisia, Turkey, Turks and Caicos, United Arab Emirates, United Kingdom, Uruguay and Vietnam. | Terminated | April 06, 2015 | 602 | 7213 | Not_Computer
Z-0014-2014 | 65955 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Synthes Midfoot Fusion Bolt 6.5 mm. It is indicated for fracture, fixation, osteotomies, nonunions, and fusions of large bones in the foot and ankle. |  |  | March 08, 2013 | 2013 | Synthes USA HQ, Inc. | A post market safety review following several customer complaints indicated additional clarification was necessary in the Midfoot Fusion Bolt Technique Guide to reduce potential for bolt back-out or migration by emphasizing the need for supplemental fixation. | Labeling design | An "Urgent Notice: Medical Device Recall" letter dated March 8, 2013 was sent to customers asking them to check their inventory, remove all affected product technical guides from use, and return them to Synthes along with the Verification form. Customer questions were directed to (610) 719-5450. |  | 62 | 62 | Nationwide Distribution, including the states of MI, PA, SC, IA, MA, CA, WI, IL, and KS. | Terminated | August 12, 2015 | 887 | 62 | Not_Computer
Z-2091-2013 | 65513 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Free Lock Femoral Hip Fixation System Compression Tube/Plate, different length (59, 75, 91, 107, 123 mm), holes (2, 3, 4, 5, 6) and degree angles (130, 135, 140, 145 and 150). Item codes starting with 00-1181-130, 00-1181-135, 00-1181-140, 00-1181-145, 00-1181-150). |  |  | June 19, 2013 | 2013 | Zimmer, Inc. | Single packaging configuration used for the VERSA-FX-II Femoral Fixation System failed the leak test compromising the sterility of the product. | Packaging process control | The firm, Zimmer, sent an "URGENT MEDICAL DEVICE NOTIFICATION" letter dated July 11, 2013 to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to inspect the device according to the insert instructions; not use the device if damage to the corner of the inner tray is observed, and contact the firm for return and questions. Customers should also keep this notification in mind and inform Zimmer of product with this condition using zimmer.per@zimmer.com.If after reviewing this notification you have questions or concerns, please contact Zimmer at 1-877-946-2761. |  | 85 | 85, 794 units in total | Worldwide distribution: US Nationwide (including PR and USVI), and countries of: Israel, Argentina, Brazil, France, Mexico, Saudi Arabia, Beirut, China, Korea, San Salvador, Honduras, Canada, Australia, Switzerland, Germany, Arab Emirates, UK, Belgium, Egypt, Spain, France, Italy, Iran, Lebanon, Jordan, Iraq, Canary Islands. | Terminated | June 16, 2015 | 727 | 85 | Not_Computer
Z-0276-2014 | 66450 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | 3.5MM LOCKING SCREW, T15 HEXALOBE, SELF TAPPING, 34MM |  |  | September 26, 2013 | 2013 | OrthoPediatrics Corp | The firm identified, through complaints that the 3.5MM LOCKING SCREW packaging material was marked as a 3.5mm screw and the part number was correct but that the description indicated a 7.0mm screw. Upon investigation, it has been determined that a lot number 009U2 was packaged and mislabeled internally by OrthoPediatrics. | Labeling mix-ups | OrthoPediatrics sent an Urgent Medical Device Recall letter dated September 26, 2013, to all affected customers.The letter identified the product the problem and the actions needed to be taken by the customerThe sales representatives and distributors who have been identified via shipping records as having received one or more discrepant devices affected by this Recall were requested to quarantine the devices identified in this Recall by telephone on 18-September-2013. If you or any of your customers have any of these devices in the original packaging please take possession of them and quarantine them in preparation for returning them to OrthoPedlatrics. They are misbranded. If the Screw has already been removed from the original packaging and the original packaging has been discarded and the Screw placed in a sterilization tray, it need not be returned.For further questions please call ( 574) 268-6379. |  | 7 | 7 units shipped | Worldwide Distribution - US including Michigan and the countries of Ireland and Australia | Terminated | December 26, 2013 | 91 | 7 | Not_Computer
Z-2096-2013 | 65874 | [6 empty] | OLA Buffer |  |  | July 12, 2013 | 2013 | Celera Corporation | The package insert for FG OLA Buffer, GPR displayed the wrong concentration of ÂżM instead of the correct mM unit. | Error in labeling | Celera sent a customer notification letter dated July 12, 2013 by e-mail and mailed on July 19, 2013 via first class mail to all affected customers. The letter identified the affected product, problem and actions to be taken. The information is being provided for records and assures customer that there were not quality issues with the product and it may be used as usual in laboratory test procedures. For questions call 1-810-749-4200 or contact your local authorized Celera representative. |  | 43 | 43 units | USA Nationwide in the state of CA | Terminated | October 01, 2013 | 81 | 43 | Not_Computer
Z-0902-2013 | 64383 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Century Console Base, Model #043625051b. |  |  | January 31, 2013 | 2013 | Salyer Prn Biomedical | Concerns that the addition of the UPS Battery Backup to the Century Console Base Shelf did not comply with the FDA premarket notification regulations. | No Marketing Application | An Urgent: Medical Device Recall letter, dated 1/31/13, was sent to all the customers that received the Century Console Base Shelf option with the Linear/1.5k 115v UPS Battery Backup for the Century Heart Lung Machine. The letter informed the customers that Salyer PRN Biomedical will have a Salyer Biomedical technician return the customers' Century Console Bases to their original configuration without the Shelf option and restore the Linear/1.5k UPS Battery Backup to its original condition as a stand-alone unit with its own wheel base for ease of use. A refund or credit will be issued for the return of the product. Customers are instructed to contact Jeff Poland at (314) 497-5232. |  | 9 | 9 units | Distributed in the states of MA,CA, and WI. | Terminated | April 15, 2013 | 74 | 9 | Not_Computer
Z-2058-2013 | 65631 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Synthes Trauma Nail System. The devices are indication for bone fixation. |  |  | November 02, 2012 | 2012 | Synthes USA HQ, Inc. | Recall was initiated due to the possibility that the outer pouch was compromised. The product is packaged in two pouches; a sterile, internal pouch within an external pouch. It is the outer, external pouch that may be compromised. | Packaging change control | The firm initiated their recall of this product on November 2, 2012 by sending an "URGENT NOTICE: MEDICAL DEVICE RECALL" letter to all consignees. The letter explained the problem and gave instructions for consignees to examine their inventory and urged consignees to return the identified devices. The letter stated If returning the product would adversely impact their ability to provide necessary medical care to patients, they should re-sterilize the product per Pre-Vacuum Sterilization instructions contained in the Instruction for Use. |  | 1011783 | 1,011,783 | Nationwide distribution. | Open, Classified |  |  | 1011783 | Not_Computer
Z-0794-2013 | 64069 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | SeriScaffold Surgical Scaffold |  |  | January 08, 2013 | 2013 | Allergen Medical | Product is dual packaged in an inner and outer pouch, the outer pouch seal may be compromised and sterility may not be assured | Packaging process control | Allergan sent an "URGENT MEDICAL DEVICE RECALL" letter dated January 8, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers are requested to quarantine and return any unused devices in inventory . In addition, customers are asked to complete and return the acknowledgement form. Customers with questions should contact Allergan at 781-658-2800, Mon- Fri, 9 a.m. to 5 p.m. EST. |  | 225 | 225 units (US) and 81 (Outside US) | Worldwide Distribution-USA including the states of CA, FL, GA, MI, MN, NJ, NY, OH, OR, RI, TX, VA, and WA, and the countries of Germany, Spain, and UK. | Terminated | July 29, 2013 | 202 | 225 | Not_Computer
Z-0277-2014 | 66199 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | VerSysÂż Cemented Revision Calcar (CRC) Hip System Femoral Stems |  |  | August 28, 2013 | 2013 | Zimmer, Inc. | Based on a review of historic packaging validations, it was determined that a subset of product packaged in one of the clean rooms at a single Zimmer facility required testing to ensure that previously packaged product from this facility met the established requirements for sterile barrier integrity. The number of samples required for packaging testing (seal integrity and seal strength) to provid | Packaging process control | Removal of the subject devices from the field including distributor and hospital inventories. Distributors and Risk Managers and Surgeons at hospitals that received a direct shipment from Zimmer or its Distributors received a notification on 09/19/2013. |  | 12866 | 12866 units | Worldwide distribution: US (Nationwide) including Puerto Rico and Virgin Islands; and countries of: Australia, Algeria, Belgium, Bolivia, Canada, Canary Islands, Chile, Croatia, Finland, France, Germany, Greece, Guatemala, Honduras, India, Israel, Italy, Japan, Jordan, Lima-Peru, Malaysia, Morocco, Nicaragua, Portugal, Saudi Arabia, Seoul Korea, Serbia, Singapore, South Africa, Spain, Switzerland, Turkey, Utd.Arab.Emir., United Kingdom, and Venezuela. | Terminated | April 30, 2015 | 610 | 12866 | Not_Computer
Z-2054-2013 | 65787 | No | GXP | Methyl methacrylate for cranioplasty. | Neurology | Neurology | 510(k) | Stryker Part Numbers 79-43903 (3cc), 79-43905 (5cc), 79-43910 (10 cc), 79-43915 (15cc), 397003 (3cc), 397005 (5cc), 397010 (10 cc), 397015 (15 cc) HydroSet Injectable HA Bone Substitute Sterile R/EO, nonpyrogenic, Product of Ireland |  |  | July 30, 2013 | 2013 | Stryker Craniomaxillofacial Division | Stryker received reports of the desiccant bag in the kit leaking silica gel granulate. This could increase exposure to anesthesia during surgical delay up to 60 minutes or inadvertent implantation of silica gel granulate which could cause foreign body reaction including inflammation and might lead to the need for revision surgery. | Package design/selection | Stryker sent an Urgent Medical Device Recall Notices with attached Product Correction Bulletin dated, July 30, 2013, describing the product, problem and possible adverse events associated with the recall were sent to all affected customers. Customers were instructed to pass the notice and bulletin to all individuals who may be using the device, provided instructions on procedures which would help mitigate risk and return the customer response form to our Customer Service Team by email, fax or letter.Email: CMF-custserv@stryker.comFax: 1-877-648-7114Address: 750 Trade Centre Way, Suite 200Portage, MI 49002 Customers responses will be tracked and non-responders will be contacted.Further questions please call (877) 534-2464. |  | 12719 | 12,719 | Worldwide Distribution - USA (nationwide) including the countries through Stryker Distribution: Austria, Belgium, Denmark, Estonia, Finland, France, Germany, Ireland, Italy, Latvia, Luxembourg, Netherlands, Norway, Poland, Portugal, Romania, Serbia, Spain, Sweden, Switzerland, United Kingdom, Argentina, Australia, Brazil, Jordan, Kingdom of Saudi Arabia, Kuwait, Latin America (others), New Zealand, Russia, Singapore, South Africa, and United Arabian Emirates. | Terminated | April 17, 2014 | 261 | 12719 | Not_Computer
Z-0001-2014 | 66365 | No | OFA | Cardiovascular surgical instruments. | Cardiovascular | Cardiovascular | Enforcement Discretion | Ultima OPCAB System, Sterile, Rx Only, |  |  | September 20, 2013 | 2013 | Maquet Cardiovascular, LLC | MAQUET has identified a potential issue with the ULTIMA OPCAB (part numbers OM-2003S &amp; OM-2001D). Labeling for the 2 blade lots had been incorrectly applied. The standard blades from Lot #25052176 were incorrectly labeled as deep blades, and the 20 deep blades from Lot #25052200 were incorrectly labeled as standard blades. | Error in labeling | MAQUET sent an "Urgent Medical Device Recall" letter and "Field Recall Response Forms" dated September 20, 2013 to customers via Federal Express. The notification identifies the affected product; informs the customers of the nature of the issue; and advises of potential adverse patient consequences that may occur. The customer is asked to complete and return the Field Recall Response Form. If the customer has questions they are directed to contact their field representative or MAQUET Technical Support (888.627.8383 between 8 am and 5 pm EST Monday through Friday). |  | 20 | 20 units (Lot 25052176); 12 units (Lot #25052200) | USA Nationwide Distribution in the states of: Florida, Maryland, New Jersey, New York and Texas. | Terminated | June 10, 2014 | 263 | 20 | Not_Computer
Z-0003-2014 | 66026 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Cordis ADROIT 6F Guiding Catheter |  |  | August 12, 2013 | 2013 | Cordis Corporation | Due to a labeling error related to the conversion of the inner diameter of the catheter from inches to millimeters. The English dimension is correct and is the primary dimension. | Error in labeling | Cordis sent an Urgent Medical Device Recall letter dated August 12, 2013 Acknowledgement Form to all affected customers. The letter identifed the affected product, problem and actions to be taken. Customers were instructed to immediately identify and set aside all affected product in a manner that ensures the product will not be used. Customers were asked to review, complete, sign and return the enclosed Acknowledgement form directly to Cordis to the fax number on the form. For questions call the Customer Service number at 800-561-7683. |  | 1238 | 1238 units | Worldwide Distribution - USA Nationwide in the states of AZ, CA, FL, ID, LA, MA, MD, MI, MN, MO, NC, NH, NY, OH, OK, PA, SC, TX, VA, WI, and WV and the countries of: Austria, Belgium, Czech Republic, France, Germany, Greece,Hungary, Israel, Italy, Latvia, Netherlands, Poland, Portugal, Russian Federation, Slovenia, Spain, Sweden, Switzerland, United Arab, and United Kingdom | Terminated | September 30, 2014 | 414 | 1238 | Not_Computer
Z-0004-2014 | 66357 | No | MXE |  |  | General &amp; Plastic Surgery | 510(k) | SHOPKO Waterproof Adhesive Pads Antibacterial. |  |  | August 21, 2013 | 2013 | ASO, LLC | No antibacterial bandages were packed in boxes labeled as antibacterial. | Labeling Change Control | A letter was sent via e-mail and mail to the only customer on August 23, 2013 that a recall of two lots was in progress and to examine their inventory and stop distribution of the product. Product was to be either place in quarantine or destroyed. A credit for all returned product or by returned completed certificate of destruction. |  | 37 | 37 cases (72 retail boxes in each case) | Distributed in the states of ID, WI, IN, and NE. | Terminated | August 04, 2014 | 348 | 37 | Not_Computer
Z-2107-2013 | 65914 | No | GAA | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | SpineNet Bone Marrow Aspiration Needle Kit (which is the Ranfac Bone Marrow Aspiration Needle #74150-01M). |  |  | February 04, 2013 | 2013 | SpineNet | SpineNet was marketing this device to be used with SpineNet ACC device without FDA clearance or approval. | No Marketing Application | SpineNET sent a "VOLUNTARY PRODUCT RECALL" letter dated 02/04/2013 to Halifax Medical Center. The letter identified the product, problem, and actions to be taken by the customer. Contact the firm at 407-539-2483 for questions regarding this notice. |  | 100 | 100 units | Distributed to one customer in Florida. | Terminated | November 15, 2013 | 284 | 100 | Not_Computer
Z-2103-2013 | 65917 | No | ODP | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | SpineNet Anterior Cervical Cage, Lordotic, Peek (originally submitted as Daytona Anterior Cervical Cage) |  |  | February 05, 2013 | 2013 | SpineNet | The recall decision has been made because because it has recently been determined that the instructions for use (IFU) had certain statements that should have been removed prior to initial product marketing. | Error in labeling | SpineNET sent a Customer Notification letter dated 02/05/2013 to their only customer (Halifax Medical Center). The letter identified the affected product, problem and actions to be taken. Customers were informed to remove any copies of the original IFU-ACC-01 form from their records and replace with the IFU-ACC-02 form attached to the letter. For questions contact SpineNET office directly. |  | 400 | 400 | USA Nationwide Distribution in the state of Florida | Terminated | May 02, 2014 | 451 | 400 | Not_Computer
Z-0002-2014 | 66189 | No | CAE | Oropharyngeal airway. | Anesthesiology | Anesthesiology | 510(K) Exempt | King LTS-D Oropharyngeal Airways |  |  | December 13, 0012 | 12 | King Systems Corp. | On December 13, 2012, King Systems initiated a voluntary recall of one (1) case(s) (5) individual products) from one (1) lot [(IV1V3) 62 cases, 310 products]of King LTSD Oropharyngeal Airways because the package label indicated that the airway was a size 5 when in fact the airway in the package was a size 4. When King manufacturers a lot, all of the components are the same size, part number etc. I | Error in labeling | King Systems notified the only one (1) consignee who received the affected product of the problem via a customer complaint. Five (5) of the five (5) affected products were destroyed by consignee. |  | 6 | 1 case; 5 units | USA Nationwide Distribution in the state of NY | Terminated | November 08, 2013 | 730815 | 6 | Not_Computer
Z-2311-2013 | 65697 | No | DSY | Vascular graft prosthesis. | Cardiovascular | Cardiovascular | 510(k) | Advanta SST Grafts &gt;or= 6mm |  |  | July 12, 2013 | 2013 | Atrium Medical Corporation | Inner tray label may adhere to wall of outer tray which make it difficult to remove the tray onto the sterile field | Package design/selection | Atrium Medical Corporation issued an Urgent Medical Device Safety Alert letter on July 16, 2013, via UPS to customers and e-mail to the field representatives that represent these accounts. The notification identifies the problem, product, and risk factors. In addition, it provides instructions how to handle product that may be affected.For questions customers were instructed to contact their Atrium Medical Customer Service at 603-880-1433 Prompt 1 or by e-mail at returns@atriummed.com |  |  |  | Worldwide Distribution - USA (nationwide) and internationally to the following countries:American Samoa AntillesArgentina Australia Austria Bangladesh Belgium Bolivia Brazil Canada Chile Colombia Egypt El Salvador Finland France Germany Great Britain Greece Guatemala Hong Kong India Indonesia Ireland Italy Jamaica Japan Korea Lebanon Libya Malaysia MauritiusMexico Neal New Zealand Nicaragua Norway PakistanPanama Peru Philippines Poland Portugal Puerto Rico Romania Russia Saudi Arabia Singapore Slovenia South Africa Spain Switzerland Taiwan Thailand the Netherlands Trinidad Tunisia | Terminated | July 15, 2014 | 368 |  | Not_Computer
Z-2101-2013 | 65850 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | ZIMMER PERIARTICULAR LOCKING PLATE SYSTEM |  |  | July 30, 2013 | 2013 | Zimmer, Inc. | Zimmer is initiating a lot specific voluntary recall of the Distal Medial Humeral Locking Plate, Short due to the incorrect information printed on the Chinese label. The original label indicates that implants are RIGHT sided, but the Chinese label indicates that the implants are LEFT sided. | Labeling mix-ups | Zimmer sent an Urgent Medical Device Recall Notification letter dated July 30, 2013 to Distributors and Hospitals in China. The letter described the affected product, problem, labeling issue, directions on how to respond if affected product is discovered, and potential health risks associated with the labeling error. An inventory return certification form is included with the recall notification. Questions or concerns are directed to Zimmer Shanhai at 021-2211 5196, Monday to Friday, between the hours of 9 a.m. and 5 p.m. |  | 11 | 11 units | Internationally Distributed in China only. | Terminated | April 15, 2014 | 259 | 11 | Not_Computer
Z-2113-2013 | 65764 | No | NKC | Laparoscopic insufflator. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(K) Exempt | DeRoyal Insufflation Tubing sets: |  |  | July 15, 2013 | 2013 | DeRoyal Industries Inc | DeRoyal Industries is recalling their product Insufflation Tubing Sets due to materials incompatibility which could possibly result in occlusion at or near the filter housing which could prevent the flow of gas through the device. | Packaging process control | The firm notified customers via letters dated 7/15/2013 sent UPS and via DeRoyal Representatives. Customers were asked to notify DeRoyal of any remaining inventory to return for replacement product or credit. Non-responders will receive additional notifications. Effectiveness checks will be performed by contacting non-responders by phone. |  | 524486 | 506,632 sets in US; 17,854 sets International | Worldwide Distribution-USA (nationwide) including DC and the states of AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, and the countries of Dubai, Arab Emirates, Quebec, Canada, Japan, and Ireland. | Terminated | May 20, 2015 | 674 | 524486 | Not_Computer
Z-0144-2014 | 66040 | No | HNJ | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | MLS Torpedo Mini-Light Ophthalmic single-port illumination. |  |  | June 27, 2013 | 2013 | Insight Instruments, Inc. | The product was registered and distributed as a Class I medical device and was subsequently determined to be a Class II medical device which requires FDA market clearance. | No Marketing Application | Insight Instruments, Inc. sent a letter dated April 8, 2013, via USPS priority mail to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue using and dispensing all affected product, perform a physical count of their inventory and complete the Return Response Form and fax to 772-219-9342. Customers were also instructed to call 800-255-8354 to arrange for return and prompt reimbursement.For questions regarding this recall call 727-219-9393. |  | 10212 | 10212 | Worldwide Distribution - USA (nationwide) including AZ, CA, FL, IN, MA, MD, TX, and WA and Internationally to Czech Republic, Germany, Netherlands, Spain, Switzerland, and Turkey. | Open, Classified |  |  | 10212 | Not_Computer
Z-0825-2013 | 63933 | No | GEH | Cryosurgical unit and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | HealthTronics Endocare 3.8mm RenalCryo Cryoprobe, R3.8L. The Endocare CryoProbe is an accessory probe intended for use with the CryoCare Surgical System, for freezing and ablation of tissue. |  |  | December 20, 2012 | 2012 | Healthtronics, Inc. | The Endocare 3.8mm RenalCryo Cryoprobe is being recalled because the label on the outer box has the incorrect expiration date. | Incorrect or no expiration date | Healthtronics sent an Urgent Device Recall letter dated December 20, 2012, to all affected customers that allows tracking. The letter included a suggested format/content for a correction status/acknowledgement letter. The customer was requested to complete the verification section of the recall letter and return all affected products to Healthtronics. A return material authorization (RMA) and pre-paid FedEx return labels was provided for all customers. All affected Healthtronics field personnell were notified of the recall via email.For questions regarding this recall call 512-439-8355. |  | 50 | 50 units | Nationwide Distribution including CA, CO, FL, GA, NC, NJ, OH, TX, and WI | Terminated | April 22, 2013 | 123 | 50 | Not_Computer
Z-0272-2014 | 66689 | No | LFD |  |  | Dental | 510(k) | Biotene Moisturizing Mouth Spray, with xylitol, For dry mouth symptom relief, 44.3ml/1.5 FL. OZ., Sugar Free Sweetened with xylitol and sorbitol, GlaxoSmithKline, Consumer Healthcare, L.P., Moon Township, PA 15108. |  |  | October 17, 2013 | 2013 | GlaxoSmithKline, LLC. | Label indicates, "Sweetened with xylitol and Sorbitol", however product does not contain sorbitol. | Labeling Change Control | GlaxoSmithKline sent an Urgent Medical Device Recall letter dated October 17, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to return all affected product. If product was further distributed customers were instructed notify their customers of the recall and the actions required. Customers were asked to complete and return the enclosed business replay card even if they have no affected product in their inventory. Customers should complete the return goods packing slip, enclose it with their return goods shipment and affix the return label to the package. Customers with questions should call Customers Service at 1-800-743-4014.For questions regarding this recall call 919-269-1194. |  | 483204 | 483,204 bottles | Nationwide Distribution . | Terminated | July 26, 2017 | 1378 | 483204 | Not_Computer
Z-0200-2014 | 65863 | No | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | Product Catalog (part number 20-0020) utilized with the Electric and Pneumatic Motor Systems including: MicroMax Motor Systems, XMax Motor Systems, EMax2 Motor Systems, and Emax2Plus Motor Systems, all models and serial numbers. |  |  | May 06, 2011 | 2011 | The Anspach Effort, Inc. | Anspach Effort, Inc. Palm Beach Gardens, FL initiated a correction of its medical Device labeling for its Pneumatic Systems Operating Manual, Cutting Burr Chart and Product Catalog due to a recent review of their regulations and compliance requirements that determined the claim that the device is capable of dissection of biometals, bioplastics and cement has been verified, but is is not part of th | Labeling Change Control | The firm, Anspach Effort, Inc., sent an "Urgent Medical Device Labeling Correction" letter dated June 27, 2011 to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to do the following: 1) Destroy or delete any items in their inventory immediately and replace with the attached revised Operating Manual, Product Catalog and/or Cutting Bur Chart. 2) Complete and return the attached Customer Reply Form indicating disposition of the labeling materials and confirming their receipt of this letter, and fax to 1-800-327-6661. 3) If they distributed any of the products to other services or facilities, please forward this information as appropriate. For additional copies, contact Anspach Customer Support at (800) 327-6887. |  | 2020 | 2,020 | Worldwide Distribution: USA Nationwide and the countries of: Egypt, India, Germany, Turkey, Canada, Saudi Arabia, Greece, United Kingdom, Cayman Islands, Kazakhstan, Republic of Korea, Taiwan, Hong Kong, China, Lebanon, Argentina, Denmark, Belgium, Netherlands, Mexico, France, Australia, Panama, Italy, Guam, United Arab Emirates, Israel, Polan, Venezuela, Austria, Hungary, Lithuania, Norway, Malaysia, Colombia, Costa Rica, Ecuador, Jamaica, Peru, Thailand, Italy, Czech Republic, Cameroon, South Africa, Portugal, Finland, Sweden, Nigeria, Bahamas, Cambodia, Dominican Republic, Switzerland, Chile, Singapore, Spain, Syrian Arab Republic, Jordan, Japan. | Terminated | January 28, 2016 | 1728 | 2020 | Not_Computer
Z-2190-2013 | 65854 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | HemoCueÂż Glucose 201 Microcuvettes, 25 Microcuvettes Individually Packaged. |  |  | July 15, 2013 | 2013 | HemoCue AB | During a complaint investigation HemoCue AB discovered punctured HemoCue Glucose 201 single pack pouches. | Packaging | HemoCue sent an Urgent Field Safety Notice dated July 12, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to quarantine any affected product and Return the Field Safety Notice Verification Form to e-mail: 2013glufsn@hemocue.com or fax 562-668-5794 Customers would be contacted to coordinate return and replacement of the affected microcuvettes. Customers with questions were instructed to call 800-426-7256.For questions regarding this recall call 440-925-3485. |  | 2221500 | 22,215 boxes of 100 cuvettes distributed in US, 34, 841 distributed outside the US | Worldwide Distribution - USA (nationwide) and internationally to AUSTRALIA, BELGIUM, CANADA, COTE D'IVOIRE, DENMARK, EGYPT, FINLAND, FRANCE, GEORGIA, GERMANY, GREECE, IRELAND, ITALY, KENYA, NETHERLANDS, NEW ZEALAND, NORWAY, OMAN, POLAND, QATAR, RUSSIAN FEDERATION, SPAIN, SWEDEN, SWITZERLAND, UNITED KINGDOM, VIET NAM &amp; YEMEN.. | Terminated | September 02, 2014 | 414 | 2221500 | Not_Computer
Z-1276-2013 | 64988 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | ExoShape Soft Tissue Fastener 12 mm x 30 mm REF 1101-00-1230. |  |  | April 04, 2013 | 2013 | Medshape Solutions, INC. | Expiration date on the patient label and date on the product carton were not the same. | Error in labeling | Medshape sent an "URGENT: MEDICAL DEVICE RECALL" letter to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. |  | 26 | 26 | Nationwide Distribution including the states of CO, FL, GA, IA, MD, MO, OR, TX, VT, VA, and WI. | Terminated | August 07, 2013 | 125 | 26 | Not_Computer
Z-1887-2013 | 65482 | No | DTK | Cardiovascular intravascular filter. | Cardiovascular | Cardiovascular | 510(k) | Cordis OPTEASE Retrieval Inferior Vena Cava (IVC) Filter***REF Cat. No.***466F220A and 466F220B***Vena Cava Filter and Introduction Kit***(BRITE TIP Catheter Sheath Introducer and Angiographic Vessel Dilator)***Cordis***a Johnson and Johnson Company***ENDOVASCULAR*** |  |  | March 29, 2013 | 2013 | Cordis Corporation | Correction to provide clarification and additional labeling to minimize likelihood on implanting filter backwards. | Labeling design | Cordis Corporation sent an "URGENT Medical Device Correction" notification to their consignees in the US by letter dated April 3, 2013 and an e-mail was sent to the Canadian consignees. The letter identified the affected product, problem and actions to be taken. Consignees were instructed to read the description and recommendations sections carefully, then sign and return the enclosed Acknowledgement Form directly to Cordis to their fax number(908) 429-8287 or Scan and e-mail to MiamiQA@crdus.jnj.com. The letter states to pass on the notice to anyone in the consignees' facilities that need to be informed. Additionally, the notification request consignees to maintain awareness of this notification until the information has been incorporated in the Cordis OPTEASE Vena Cava Filter labeling. For questions related to the issue contact Cordis Endovascular Medical Information Line 877-338-4235. |  | 33000 | 33,000 units | Worldwide Distribution - USA Nationwide in the states of AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, Ny, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY, including Hawaii and Puerto Rico and the countries of Canada &amp; Panama | Terminated | October 27, 2014 | 577 | 33000 | Not_Computer
Z-1728-2013 | 65505 | No | GBA | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | 27023WU Balloon Catheter, 3 Fr., 2/pkg, sterile, contains latex. |  |  | May 31, 2013 | 2013 | Karl Storz Endoscopy America Inc | Karl Storz Endoscopy-America has initiated the recall of the 27023WU Balloon Catheter because the packaging may be compromised, and therefore the sterility of the product could be affected. | Packaging | The firm, Karl Storz, sent a "A SAFETY ALERT-ATTENTION NEEDED" letter dated May 31, 2013 to all customers who purchased the 27023WU Balloon Catheter. The letter described the product, problem and actions to be taken. The customers were instructed to check their inventory and to return the recalled product back to Karl Storz.. The customers were also instructed to complete and fax back the Recall Response Form attached with the safety alert letter at Fax #424-218-8559. If you have any questions, please feel free to call 424-218-8289. |  | 40 | 40 units | US distribution to: AZ, FL, IL, IN and TX. | Terminated | July 15, 2013 | 45 | 40 | Not_Computer
Z-1426-2013 | 65072 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Thyroidectomy Indication for the da Vinci Surgical Systems; |  |  | October 13, 2011 | 2011 | Intuitive Surgical, Inc. | Promotional literature for use of the da Vinci system for thyroidectomy, which had not been approved with a change to the 510(k). | No Marketing Application | Intuitive sent a notification letter on October 17, 2011 by email to all consigness whose email address were on file. Also a copy of the letter was sent by registered mail on February 8, 2012 to all surgeons whose addresses were on file. The letters informed consignees that Intuitive Surgical were in the process of reviewing labeling for the da Vinci System in regard to use for thyroidectomy procedures. During this period, Intuitive will not provide or facilitate any promotional, training, case observations, proctoring or in case procedural support for use of the da Vinci System in conjunction with thyroidectomy procedures. The letters recommend consulting the institution policy regarding use of approved or cleared medical devices for procedures that are not specified in the labeled indication for use. Questions may be directed to customer service at 1-877-408-3872. |  | 2423 | 2,423 consignee | USA Nationwide distribution | Terminated | May 30, 2013 | 595 | 2423 | Not_Computer
Z-0464-2014 | 66830 | [6 empty] | Centurion Sterile # 84 Rubber Bands Reorder EB84, Caution: This product contains natural rubber latex which may cause allergic reactions. |  |  | October 28, 2013 | 2013 | Centurion Medical Products Corporation | Package labeling indicates both "latex free" and "contains natural rubber latex" . The rubber bands do contain natural rubber latex. This could cause a significant risk to users with latex allergies. | Labeling False and Misleading | Centurion sent a Urgent Recall Notification letter via Certified Mail October 31, 2013, return receipt to all affected customers. The affected Centurion Medical Products Corporation sales representatives were notified via email on October 28, 2013. Customers were instructed to destroy all implicated product and complete the accountability record included with the notice and fax to 517-546-3356. Customers were asked to forward a copy of the notice if product was further distributed. Additional notices will be mailed to non-responsive customers via Certified Mail ReturnReceipt, and will be documented in the recall file. For further questions please call (517) 546-5400. |  | 500 | 500 lots | US Distribution including the states of GA, LA and NY. | Terminated | January 29, 2014 | 93 | 500 | Not_Computer
Z-0467-2014 | 66761 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | 1) Accolade II 127 Neck Angle Hip Stem, Size # 3 |  |  | October 25, 2013 | 2013 | Stryker Howmedica Osteonics Corp. | Stryker received a report from the field indicating that an Accolade II Hip Stem Size #3 was identified in packaging with an Accolade II Hip Stem Size #5. | Packaging process control | Stryker sent an e-mail to their domestic locations on 10/11/2013. Notification letters and product accountability forms were sent via Fed Ex on 10/25/2013 with return receipt to branches/agencies and hospital risk management, chief of orthopaedics and surgeons. |  | 20 | 20 units | Nationwide Distribution including states: AL, GA, IL, and WA. | Terminated | March 31, 2015 | 522 | 20 | Not_Computer
Z-1424-2013 | 65094 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | User Manual addendum, TORS for da Vinci Surgical System IS 1200, IS 2000 and IS 3000. |  |  | January 24, 2013 | 2013 | Intuitive Surgical, Inc. | Change to User Manual to remove one indication for use and amend warnings to emphasize restrictions for use. | Labeling False and Misleading | Intuitive sent an Important Medical Device Correction letters dated January 24, 2013 via Federal Express to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to locate and discard the existing TORS addendum, place the new TORS addendum included with this letter with your System User Manual, and complete and return the attached Acknowledgement Form to Intuitive Surgical. For questions call Intuitive Surgical Customer Service at 1-800-876-1310. |  | 1374 | 1,374 | US Distribution Only | Terminated | May 30, 2013 | 126 | 1374 | Not_Computer
Z-1660-2013 | 64078 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Custom Procedural Trays: Merit's Custom Angio Kits |  |  | November 21, 2012 | 2012 | Merit Medical Systems, Inc. | Custom Procedural trays contain Medrad Hand Controller Sheaths which were recalled because the seal between the white polyethylene and the clear plastic may be missing on the package resulting in a potential breach of the sterility of this accessory. | Packaging | Merit Medical sent an "URGENT PRODUCT RECALL NOTICE" letter dated November 21, 2012, to affected customer. The letter identified the product, the problem, and the action to be taken by the customer: Ensure all personnel within their institution to which the affected Merit trays were distributed are notified, quarantine all affected trays and return the affected product to Merit - Maastricht. Complete and sign the Customer Response Form attached to this letter and perform the following steps:(1) Immediately fax a copy of the completed Customer Response Form to: 1-804-416-1031. (2) Mail the completed original Customer Response Form from (1) above using the provided postage-paid, self-addressed envelope to Merit. Product Return Instructions: Please return the affected products by shipping them back to Merit via Federal Express 2nd Day Air (Acct #1129-5275-6) to: RGA Department, RGA #NC-54948 Merit Medical Maastricht, Amerikalaan 42, 6199 AE Maastricht - Airport, The Netherlands. NOTE: A copy of the Customer Response Form must accompany the shipment. If you have any questions, please call 1-804-416-1048. |  | 134 | 134 trays | U.S. Military Hospital on base in Landstuhl, Germany | Terminated | July 12, 2013 | 233 | 134 | Not_Computer
Z-1428-2013 | 65083 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Flush Port (used in almost all da Vinci instruments) |  |  | September 19, 2011 | 2011 | Intuitive Surgical, Inc. | Clarification of existing labeling information and operating procedures. | Labeling False and Misleading | Intuitive Surgical sent an "IMPORTANT PRODUCT NOTIFICATION" letter dated October 17, 2011 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact Customer Service at 1-800-876-1310 for questions regarding this notice. |  |  |  | Worldwide Distribution- USA Nationwide and countries of: Argentina, Australia,Canada, Chile, China, Czech Republic, Egypt, India, Israel, Italy,Korea-Republic of, Poland, Qatar, Romania, Saudi Arabia, Singapore, Spain, Thailand, Turkey, and Venezuela. | Terminated | May 30, 2013 | 619 |  | Not_Computer
Z-0473-2014 | 66582 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes TomoFix Plates |  |  | October 09, 2013 | 2013 | Synthes USA HQ, Inc. | A labeling correction was initiated related to the Surgical Technique Guide J5606-C "TomoFix Osteotomy System" due to the potential for a surgical delay if the surgeon inserts the superior screws at a lesser angle (4 or 6 degrees) and is not aware that those angles are not present. | Labeling False and Misleading | Synthes sent an Urgent Field Safety Notification/Medical Device Labeling Correction letter dated October 9, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to exchange the old surgical technique guide version J5606-C with the new version J5606-D provided with this notification. Review the revised surgical technique guide (Pages 25 &amp; 26). Forward this Field Safety Notification to anyone in their facility that need to be informed.If the surgical technique guide has been forwarded to another facility, contact that facility.Maintain awareness of this notice until all surgical technique guides have been exchanged.Maintain a copy of this notice. Customers with questions were instructed to contact their SynthesSales Consultant. For questions regarding this recall call 610-719-5000. |  | 2912 | 2912 | Worldwide Distribution - USA (nationwide) Canada and Mexico | Terminated | December 18, 2015 | 800 | 2912 | Not_Computer
Z-1654-2013 | 64888 | No | LRY | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | HealthLink***BIOPSY PUNCHES*** |  |  | March 18, 2013 | 2013 | Clorox Healthcare Holdings, LLC (dba HealthLink) | Due to lack of sterilization. | Packaging change control | On March 19, 2013, HealthLink sent their customers letters instructing them to do the following: 1) Screen their inventory and quarantine the product. 2) Notify all customers that received the affected products and ask them to return the lots to their firm. 3) Complete the attached form and return the lots to HealthLink Distribution center. For further assistance, customers were directed to call customer service at 1-800-638-2625. |  | 1044600 | 1,044,600 units | Worldwide Distribution-USA (nationwide) including Puerto Rico and the states of AZ, CA, CO, FL, GA, IA, IL, IN, MI, MO, MS, MT, NC, NJ, NY, OH, SC,TX ,VA ,WA, WI, and HI, and the countries of Canada, Mexico, and New Zealand. | Terminated | February 25, 2014 | 344 | 2849900 | Not_Computer
Z-0551-2014 | 66949 | No | JXG | Central nervous system fluid shunt and components. | Neurology | Neurology | 510(k) | Integra Cranial Access Kit, Model Number HITHSP10, Rx Only. |  |  | November 26, 2013 | 2013 | Integra LifeSciences Corp. d.b.a. Integra Pain Management | Integra LifeSciences is voluntarily recalling anyunexpired lots of HITHSP10 Cranial Access Kits due to identified labeling discrepancies. kit contains DuraPrep and Monocryl sutures whose labeling does not align with the intended use and labeling of the Cranial Access Kit. | Labeling mix-ups | Integra sent an Urgent Voluntary Medical Device Recall letter dated November 26, 2013, to all affected consignees. The letter identified the product, the problem, the action to be taken by the consignee and included a Recall Acknowledgement and Return form. Consignees were instructed to review their inventory and determine if they have any affected product, stop using the affected product immediately and remove from service. Consignees were instructed to complete the attached form even if they do not have the affected product in their inventory and return by email or fax as indicated on the form. When the form is received Customer Service will send an RMA number and directions to return the product. Customers with questions were instructed to contact Customer Service at 1-855-532-1723 or their Integra sales representative at 909-201-4099.For questions regarding this recall call 609-936-6822. |  | 731 | 731 | Nationwide Distribution | Terminated | April 09, 2014 | 134 | 731 | Not_Computer
Z-0483-2014 | 66908 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Stryker Orthopaedics Cancellous Bone Screw |  |  | October 01, 2013 | 2013 | Stryker Howmedica Osteonics Corp. | Stryker received a report from the field indicating that a 5.5 mm Cancellous Bone Screw 40 mm was identified in packaging associated with a 6.5 mm Cancellous Bone Screw 25 mm. | Labeling mix-ups | Stryker sent notification letters/accountability forms via Fed Ex on October 1, 2013, with return receipt. to branches/agencies and hospital risk management, chief of orthopedics and surgeons.Domestic branch locations were notified by e-mail on 9/16/2013.The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to fax the Product Recall Acknowledgement Form to 855-251-3635, return the affected product to the attention of Regulatory Compliance, Stryker Orthopaedics, 325 Corporate Drive, Mahwah, New Jersey, 07430. Customers were instructed to attach the fluorescent organge PRODUCT REMEDIATION sticker to their return, indicating the Produc Remediation # RA 2013-145. Mark the outer box with the words "Product Recall."Customers with questions were instructed to call 201-972-2100.For questions regarding this recall call 201-831-5000. |  | 2232619272 | Catalog Number 2030-6525-1- 4 units; Catalog No: 2029-5540-1-7 units | Nationwide Distribution | Terminated | July 15, 2015 | 652 | 2232619272 | Not_Computer
Z-1623-2013 | 65205 | No | EMF | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | -Instratek Carpal Tunnel Release Blades, (1) Triangle Blade, (1) Hook Blade. REF R2056 - LOT 122279 |  |  | May 03, 2013 | 2013 | Surgical Instrument Service And Savings, Inc. | The seal which maintains a sterile barrier for reprocessed medical devices became compromised to the point where product may fall out of the pouch. | Packaging process control | MEDISISS sent a Correction Immediate Action Request letter dated May 2, 2013, to all affected customes. The letter identified the product, the problem, and the action to be taken by the customer. The letter directed the customer to inspect the products prior to use for seal integrity as specified on the product labeling: "Sterile unless package has been damaged or opened. Do not use if package is damaged."This will be followed up by a visit from the customer's sales representatives who are tasked with inspecting all the lots that may be remaining in the customer's inventory. The sales representative is directed to complete a "Verification of Seal Inspection" form and return it to MEDISISS Customer Relations. Devices with unacceptable seals are to be returned to MEDISISS. The sales representative is responsible for obtaining a shipping label from customer service to enable the devices to be returned to the recalling firm via Fed Ex.Customers with questions were instructed to call 866-866-7477. |  | 56 | 56 units | Nationwide Distribution. | Terminated | July 11, 2013 | 69 | 4540 | Not_Computer
Z-0531-2014 | 66930 | No | LZP | Intraocular fluid. | Ophthalmic | Ophthalmic | PMA | Viscocel and Viscocel Plus. |  |  | November 25, 2013 | 2013 | C.L.R. Medicals International, Inc. | CLR Medicals International Inc., initiated a nationwide recall of Viscocel and Viscocel Plus, in response to FDA inspection findings of deficiencies in the Quality System Regulations for Medical Device Manufacturers, 21 Code of Federal Regulation (CFR) Part 820 and lack of Premarket Approval (PMA) for devices sold or distributed in domestic commerce. The Viscoelastic products, intended for human | No Marketing Application | CLR Medicals initiated a nationwide recall on November 25, 2013, of Viscocel and Viscocel Plus, by issuing a press release,in response to FDA inspectional findings of deficiencies in the Quality System regulations for Medical Device Manufacturers, 21 Code of Federal Regulation (CFR) Part 820 and lack of Premarket Approval (PMA) for devices sold or distributed in domestic commerce. The Viscoelastic products, intended for human and animal use, have been determined to be a public health concern due to distribution in the US without FDA approval, which evaluates the safety and efficacy of the product, and violations of the QSR regulations related to sterility, and biocompatibility that may result in the release of contaminated products. CLR Medicals International, Inc. sent a Medical Device Recall letter dated December 6, 2013, to all affected customers. The letter informed customers of the reason for recall, product description with manufacturing date range, possible health hazard, instructions on what to do with recalled product, and contact information for the firm and to report any adverse events to FDA MedWatch. The recall notification letter was accompanied by a response form. Customers with questions were instructed to call 909-398-1880.For questions regarding this recall call 909-398-1880. |  | 249592 | 249,592 units | Worldwide Distribution - USA (nationwide) and Internationally to Egypt, Mexico, and Canada. | Terminated | December 05, 2014 | 375 | 249592 | Not_Computer
Z-1162-2013 | 64508 | No | NAW |  |  | Radiology | PMA | Sirtex SIR-yttrium 90 microspheres (Radionuclide); SIR-Y001. |  |  | February 27, 2013 | 2013 | Sirtex Medical Limited | Incorrect expiration date | Labeling False and Misleading | Sirtex notified medical end users on 2/27/13 via email correspondence with a copy of the recall letter attached. This was done to ensure advice was received at the same time that the goods were arriving. Hard copies of the recall letter, which were stamped "copy", were also sent to the individual facilities via UPS, processed for delivery Thursday, 28 February, 2013. Medical end users were advised in the recall letter to contact the US Customer Service staff at (888) 474-7839 x 702 to arrange for the pick-up if they decide to return the unused product. |  | 21 | 21 doses | Nationwide Distribution-USA (nationwide) including DC and the states of AZ, CA, CT, GA, ID, IL,KY, LA, MD, ME, MO, MT, NJ, NV, WA, and WV. | Terminated | March 30, 2015 | 761 | 21 | Not_Computer
Z-1618-2013 | 65202 | No | KMK | Intravascular catheter securement device. | General Hospital | General Hospital | 510(K) Exempt | StatLock CV Plus w/Pigtail, Catheter Stabilization Device Sterile EO Single Use Only, C.R. Bard, Inc. The StatLockÂż device provides stabilization for compatible medical tubes and catheters. |  |  | June 10, 2013 | 2013 | Bard Access Systems | BARD Access Systems is recalling lot JUWJF283 of StatLock CV Plus Stabilization Butterfly Devices due to incorrect pad assembly used during manufacture. The absence of the hydrocolloid strip on the bottom may pose an incremental risk including localized skin irritation. Missing directional arrow may lead to mis-orientation of the StatLock stabilization which may impact the securement of the cathet | Packaging | On 6/18/13, an "Urgent Kit Recall" notification letter was mailed to all customers informing them of the recall and instructing the consignee to remove and return any affected product lot numbers to Bard Access Systems through return goods authorization. Customers may contact 800-290-1689 if questions. |  | 3875 | 3,875 units. | National and Global distribution (Europe, European Economic Area, Argentina, China, and Brazil). | Terminated | May 21, 2014 | 345 | 3875 | Not_Computer
Z-1111-2013 | 64782 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Edwards Lifesciences Duraflo coated Femoral Cannula. |  |  | March 14, 2013 | 2013 | Edwards Lifesciences, LLC | Two unsealed pouches were found at an Edwards' distribution location in Japan. Incorrectly sealed pouches may result in a sterility barrier breach. | Packaging process control | Edwards Lifesciences sent an Urgent Field Safety Notice letter dated March 28, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to discontinue use of any affected product remaining in inventory and to return any unused devices to Edwards. A completed Acknowledgement Form should be faxed to Edwards at 800-422-9329 within three days of receiving the notice. Questions regarding the recall should be directed to Edwards Customer Services at 800-424-3278, 6:00AM - 4:30PM PST. |  | 379 | 379 units | USA Nationwide Distribution including the states of: FL, MA, MI, NY, PA, TX. | Terminated | March 08, 2016 | 1090 | 379 | Not_Computer
Z-0465-2014 | 66823 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Axillent Debakey Grasping Forceps, Part No. 33821DYD |  |  | November 06, 2013 | 2013 | Karl Storz Endoscopy America Inc | KARL STORZ is recalling the Axillent Debakey Grasping Forceps because the 33821WTD Wavy grasper was mislabeled as 33821DYD Axillent Debakey Grasping Forceps. | Labeling mix-ups | Karl Storz Endosocopy sent a Safety Alert letter dated November 6, 2013, to all customers who received the 33821DYD Axillent Debakey Grasping Forceps. The letter informed the customers of the problems identified and the actions to be taken. Customers were asked to set aside the mislabeled product and return the mislabeled instruments with the RMA (Return Material Authorization) number issued to them. Customers were instructed to fax or email completed recall response form to (424) 218-8519. Customers with any questions were instructed to call (800) 421-0837 ext 8289.For questions regarding this recall call 746-708-8440. |  | 7 | 7 units | Nationwide Distribution including states of: TX, PA, and NY. | Terminated | January 13, 2014 | 68 | 7 | Not_Computer
Z-1667-2013 | 65449 | Yes | BTO | Tracheostomy tube and tube cuff. | Anesthesiology | Anesthesiology | 510(k) | Tracoe Mini 3.0mm Tube, for pediatric use. |  |  | April 17, 2013 | 2013 | Bryan Medical Inc | The pediatric tracheostomy tubes were mislabeled as the smaller infant tracheostomy tubes. | Labeling Change Control | The consignees were contacted by telephone and asked to return the recalled product. |  | 5 | 5 units | Distributed in KY and AL. | Terminated | October 29, 2013 | 195 | 5 | Not_Computer
Z-0461-2014 | 66809 | No | HWT | Template for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | NexGen Complete Knee Solution |  |  | November 19, 2013 | 2013 | Zimmer, Inc. | This is an expansion of the June 2013 recall of NexGen Micro components. Five (5) additional instruments have been identified that are associated only with the NexGen Micro components. There is the potential for these devices to be inadvertently utilized during surgery, resulting in a delay. The associated implants would not be available due to the previous recall. | Labeling design | Zimmer sent an Urgent Medical Device Recall Notice dated November 19, 2013, to all affected customers. The Notice outlined the products recalled; the reason for the recall, possible adverse events related to the recall and directed consignees to locate, quarantine and count all affected product recording the information on the Inventory Return Form acknowledging that recall directions were followed. Further questions should be directed to 1-877-946-2761. Affected product will be removed by a Zimmer Sales representative, Question should be directed to 1-877-946-2761 8a.m - 5 p.m. EST.For questions regarding this recall call 800-613-6131. |  | 638 | 638 units | Worldwide Distribution - USA (nationwide) and Internationally to Australia, Belgium, Canada, Denmark, Ecuador, Germany, India, Italy, Japan, Korea, Malaysia, Mexico, Netherlands, Peru, Singapore, South Africa, Spain, Switzerland, UAE, and UK. | Terminated | May 06, 2015 | 533 | 3027 | Not_Computer
Z-0717-2013 | 63873 | No | HIS | Condom. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Lifestyles Pleasure Collection |  |  | November 05, 2012 | 2012 | Ansell Healthcare Products LLC | One lot of Lifestyles Pleasure Collection Lubricated Condoms is misbranded. Ultra Thin was included rather than Skyn. | Packaging | Ansell Healthcare Products LLC sent an "URGENT: VOLUNTARY CONDOM RECALL" letter/return response form dated November 5, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm's Customer Service Department at 1-800-305-1054 for questions or if additional information is needed. |  | 445 | 445 cases (10680 consumer units) | Nationwide Distribution. | Terminated | July 31, 2013 | 268 | 445 | Not_Computer
Z-0579-2014 | 67003 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | OrthoPediatrics ACL Reconstruction System is a pediatric and small stature adult ACL reconstruction system containing Sleeve Installation Tool Set , Part 01-1008-0040/0047, 6mm; 6.5mm, 7mm, 7.5mm, 8mm, 8.5mm, 9mm, and 10mm. |  |  | November 07, 2013 | 2013 | OrthoPediatrics Corp | Part # 01-1008-0040/0047, ACL Sleeve Installation/ Insertion Tool, shipped November 4, 2013 - November 6, 2013 may be etched with the wrong size. | Labeling design | November 7, 2013 all consignees were notified of the Urgent Device recall via telephone calls and E-mail messages describing the product and reason for the recall,. The notice included directions to identify and quarantine the recalled devices returning product to OrthoPediatrics with completed return form to OrthoPediatrics via mail, email or courier service, or byphotographing with a smart phone and emailing to logistics@orthopediatrics.com or by answering therecall specific questionsin an email to mfox@orthopediatrics.com |  | 40 | 40 | Nationwide Distribution including the states of CA, KY, OH, RI, and TN. | Terminated | May 09, 2014 | 183 | 40 | Not_Computer
Z-1405-2013 | 65119 | No | DQR | Catheter cannula. | Cardiovascular | Cardiovascular | 510(k) | AngioDynamics Micro Introducer Kits under the following labeling: |  |  | April 30, 2013 | 2013 | Angiodynamics, Inc. | AngioDynamics Inc., is recalling Micro Introducer Kits and VenaCure EVLT NeverTouch Procedure Kits, because the potential exists that the micro Introducer (sheath/dilator) packaged within the kit could be the incorrect French size. | Labeling mix-ups | AngioDynamics sent an Urgent Medical Device Recall letter dated April 29, 2013, to all customers by Fed Ex. The letter identified the affected product, problem and actions to be taken. Customers were instructed to lmmediately remove the affected product from inventory and call AngioDynamics, Inc. Customer Service at 1-800-772-6446 to obtain a replacement or credit for your returned product and to obtain UPS Account Number to return the package via second day delivery. Customers are instructed to segregate and return all affected devices to: AngioDynamics 603 Queensbury Avenue Queensbury, NY 12804 Attn: Micro Introducer and VenaCure EVLT NeverTouch Procedure Kits Recall Coordinator. Write the RMA number on the shipping box (Provied on the Recall Verification Tracking Form). Promptly complete, sign and return the enclosed Reply Verification Tracking Form (even if you do not have any product to return); following the directions on the notification form and the Reply Verification Form. Send the Reply Verification form to AngioDynamics either via Fax or Emai. Fax Reply Verification Tracking Form to: Attn: Micro Introducer and VenaCure EVLT NeverTouch Procedure Kits Recall Coordinatory Fax number 1-518-798-1360 or Email Reply Verification Tracking Form to: rdenino@angiodynamics.com. |  | 806 | Domestic(1,760 units): 1)260 units/26 boxes, 2)310 units/31 boxes,3)260 units/26 boxes,4)880 units/88 boxes,5)60 units/6 boxes,Foreign:10 units/1 box | USA Nationwide Distribution including the states of: IL, GA, CA, NY OH, AZ, TX, NE, FL, PA, NC, WI, AL, IN, CO, MO, MT, NE, MA, VA, AK, IN, MI, NJ, MD, TN, CT, WA, MN | Completed |  |  | 1557 | Not_Computer
Z-1401-2013 | 65048 | No | JHI | Human chorionic gonadotropin (HCG) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Alere" hCG Combo Cassette (20/10 mIU/mL) Kit. |  |  | April 16, 2013 | 2013 | Alere San Diego, Inc. | The firm initiated recall of Alere" hCG Combo Cassette (20/10 mIU/mL) kit PN 92215, lots hCG1110133 and hCG1110135, because an investigation has indicated a portion of these Alere" hCG Combo Cassette (20/10 mIU/mL) PN 92215 kit lots were incorrectly packaged with the Alere" hCG Cassette (20 mIU/mL) PN 92210 Product Insert. The Alere" hCG Cassette (20 mIU/mL) Product Insert contains information for | Packaging | Alere initiated this recall by sending recall notification letters, titled "URGENT MEDICAL DEVICE RECALL", dated April 16, 2013 via via fax, email or direct mail. The letter informed customers of the product recalled, reason for recall, customer required action, and contact information, "Alere Technical Services at 866-216-0094 or by e-mail at lateral.flow.support@alere.com.". A verification form was attached to the notification letter. |  | 3975 | 3975 kits | United States Nationwide Distribution | Terminated | October 18, 2013 | 185 | 3975 | Not_Computer
Z-0448-2014 | 66333 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | GENESIS(R) II, PS HIGH FLEXION ARTICULAR INSERT, SIZE 5-6, 11 MM QTY: (1), REF 71421516, STERILE EO, Smith &amp; Nephew, Inc. Memphis, TN 38116 USA |  |  | August 27, 2013 | 2013 | Smith &amp; Nephew Inc | The Tyvek(TM) lid on one, or both, of the two trays that form adouble barrier around the insert may not have been adequately sealed. | Package design/selection | Smith &amp; Nephew sent an Urgent Recall Notification letter dated August 27, 2013, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the customer. They were instructed to cease use and distribution of the product and quarantine for return. Customers were instructed to contact Smith &amp; Nephew via email at FieldActions@smith-nephew.com or fax 901-566-7975 to obtain a return authorization (RA) number. Customers with questions were also instructed use this email address. Customers with affected product were instructed to return it along with the completed inventory Return Certification Form to the address listed on the form. Non-responding consignees were notified again on September 27, 2013, by email.For questions regarding this recall call 901-396-2121. |  | 46 | 46 pieces | Worldwide Distribution - USA including Nevada and Internationally to United Arab Emirates, Italy, Spain, Belgium, Bulgaria, Poland, France, Japan, Netherlands, Australia, Turkey, Russian Federation, Germany and United Kingdom. | Open, Classified |  |  | 227 | Not_Computer
Z-1518-2013 | 65108 | No | LFL |  |  | General &amp; Plastic Surgery | 510(k) | TX1 Tissue Removal System disposable handpieces. |  |  | April 30, 2013 | 2013 | American Optisurgical Inc | American Optisurgical, Inc. is recalling certain lots of TX1 Tissue Removal System disposable handpiece due to inventory being released into distribution without a cleared FDA 510(k) to cover recent product changes. Use of this product may lead to an underperformance of the device. | No Marketing Application | American Optisurgical Inc. (AOI) sent an Urgent Medical Device Recall letter along with a response sheet via USPS registered mail to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immediately examine your inventory and quarantine product subject to recall. Once the product has been identified, return all products by FedEx to AOI using the AOI FedEx account 30655659-2, and referencing the Returned Goods Authorization (RGA) number:RGA# (to be assigned). For questions contact American Optisurgical Inc. Customer Service 949-580-1266. |  | 1436 | 1,436 units | Worldwide Disitribution - USA Nationwide in the states of Alabama, Arkansas, Arizona, California, Colorado, Delaware, Florida, Georgia, Iowa, Illinois, Indiana, Kansas, Kentucky, Massachusetts, Maryland, Michigan, Minnesota, Missouri, North Carolina, Nebraska, New Hampshire, New Jersey, New Mexico, Nevada, New York, Ohio, Oklahoma, Oregon, Pennsylvania, South Carolina, Tennessee, Texas, Utah, Virginia, Vermont, Washington, Wisconsin, West Virginia, Wyoming and the countries of Quezon City, Philippines, Singapore, and Panama City, Panama. | Terminated | March 24, 2014 | 328 | 1436 | Not_Computer
Z-1084-2013 | 64749 | No | DZD | Orthodontic appliance and accessories. | Dental | Dental | 510(K) Exempt | Bracket Buccal Tube U7L DB/UL22 Roth MT HK 0/-14/0 |  |  | March 08, 2013 | 2013 | Ormco/Sybronendo | Ormco Corporation is voluntarily recalling one (1) lot of Bracket Buccal Tubes as a result of a packaging error. More specifically, the BracketBuccal Tubes product package correctly identifies the product as gingival (0Âż) offset brackets, butactually contains distal (10Âż) offset brackets. Use of the affected product may cause unintended rotationof the molars. Although reversible, a delay in | Packaging | Ormco sent an Urgent Medical Device Recall letter dated March 8, 2013 to all affected customers, via USPS 1st class mail. The letter identified the affected product, problem and actions to be taken. The consignees were instructed to complete the Return Form and return any affected product identified in stock. For questions contact Ormco Customer Care. |  | 8210 | 8210 | Worldwide Distribution - USA (nationwide) including the country of EU. | Terminated | September 04, 2013 | 180 | 8210 | Not_Computer
Z-1675-2013 | 65439 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | AngioDynamics RITA Main Cable, 9 Pin to 14 Pin, REF 700-101892. RITA Main Cables are indicated for use with the RITA RF Generators and RITA electrosurgical devices. |  |  | June 11, 2013 | 2013 | Angiodynamics, INC | The RITA Main Cable product contains the incorrect Instructions for Use. | Packaging process control | Consignees were notified via certified mail on 06/11/2013 with an "Urgent Medical Device Field Safety Correction" notification. This included the correct Instructions for Use and the recommended actions. |  | 13 | 13 units | Worldwide Distribution, including the US, Austria, China, Italy, Russia and the UK. | Terminated | June 12, 2015 | 731 | 13 | Not_Computer
Z-0559-2014 | 66535 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | B. Braun Introcan Safety IV Catheter PUR 24G, 0.7X19mm-US, catalog number 4251601-02. |  |  | September 27, 2013 | 2013 | B. Braun Medical, Inc. | One lot of Introcan Safety IV Catheters may not contain the anti-needlestick safety clip which is used for the prevention of needlestick injuries. | Labeling False and Misleading | B. Braun sent an Urgent Medical Device Recall Notification letter dated September 27, 2013 to all affected customers. The letter identifed the affected product, problem and actions to be taken. Customers were instructed to utilize the attached "Product Removal Acknowledgement Form" to record the number of individual units in their inventory and fax the completed form to 610-848-1197. To ensure proper credit for affected product, customers were instructed to return the product to B. Braun Medical Inc. For questions contact Customer Support Department at 800-227-2862. |  | 55200 | 55,200 units | USA Nationwide Distribution including Puerto Rico. | Terminated | October 08, 2014 | 376 | 55200 | Not_Computer
Z-1676-2013 | 65290 | No | HSD | Shoulder joint humeral (hemi-shoulder) metallic uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Bigliani/FlatowÂż The Complete Shoulder Solution, Humeral Stem Provisional 12mm x 170mm |  |  | June 11, 2013 | 2013 | Zimmer, Inc. | The firm is initiating a removal of one lot of the Bigliani/Flatow Humeral Provisional Stem (00-4301-012-17, lot 62283991) as the stems manufactured under the lot are 14mm x 170mm devices incorrectly etched and packaged as 12mm x 170mm devices. | Packaging process control | Zimmer sent an URGENT MEDICAL DEVICE RECALL notification dated June 11, 2013, to all consignees via email and courier describing the affected device and problem and the risks associated with using the affected device. The notifications provided instructions for returning the affected device to Zimmer Product Service Department, 1777 West Center Street, Warsaw, IN 46580. Any questions or concerns concerning the recall, please call the customer call center at 1-877-946-2761. Customers may report adverse events or quality problems experienced with the use of these products to Zimmer by email at zimmer.per@zimmer.com.For questions regarding this recall call 800-613-6131. |  | 1 | 1 | Nationwide Distribution including IL and MO. | Terminated | January 30, 2014 | 233 | 1 | Not_Computer
Z-1683-2013 | 65183 | No | LJH | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Two Lead Arthroscopic Irrigation Set, for use with ARTHROMATIC Plastic Containers, Product Code 2C4030 - Approximate Length 98" (2.5. m) |  |  | May 15, 2013 | 2013 | Baxter Healthcare Corp. | A firm's internal investigation led to the discovery of pinholes in the packaging of the products. The pinholes are not likely to be detected by the end user prior to use and could compromise the sterile barrier properties of the packaging. | Packaging | Baxter sent an Urgent Product Recall letter dated May 15, 2013 to all affected custorners. The letter identified the affected product, problem and actions to be taken. Customers were requested to remove all affected products and return for credit by contacting Baxter Healthcare Center for Service at 888-229-0001. Customers were instructed to complete the attached Customer Reply Form and return it to Baxter by either faxing to 224-270-5457 or scanning and e-mailing the reply form to fca@baxter.com |  | 530488 | 530,488 units | US Nationwide Distribution | Terminated | June 16, 2015 | 762 | 1136792 | Not_Computer
Z-1671-2013 | 65347 | No | KNQ | Esophageal dilator. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | RigiflexTM 11 Single-Use Achalasia Balloon Dilator, Material Number (UPN) M00554510 |  |  | May 28, 2013 | 2013 | Boston Scientific Corporation | Product labeled with incorrect expiration date. | Incorrect or no expiration date | Boston Scientific sent an Urgent: Medical Device Recall letter to all affected customers via Federal Express. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to immediately discontinue use, check their inventory, remove any affected product from their inventory, and return it to Boston Scientific. The letter also requests the completion of a Reply Verification Tracking Form as a method of documenting the presence or absence of affected product from their inventory. Follow-up communications will be made to the accounts that do not respond. For questions contact your local Sales Representative. |  | 91 | 91 | Worldwide Distribution - USA Nationwide including the state of NY, NE, MD, OH, TX, PA, VA, PR, CA, NJ, NY, CA and the countries of: AT, FR, DE, GB, IL, IT, LB, NL, ES, SE, TN, Argentina, and Japan. | Terminated | February 10, 2014 | 258 | 91 | Not_Computer
Z-1409-2013 | 64586 | No | MNI | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | 45 mm Screw, Item #BG7045 |  |  | February 20, 2013 | 2013 | Spinal Solutions, LLC | The recall of the Loaner Bin and Sterilization Trays for implants and instruments was initiated by Spinal Solutions because the directions for use and surgical techniques were not being provided to every physician and for every surgery. | Labeling design | Spinal Solutions sent a recall letter dated March 8, 2013 to all customers who purchased loaner bins and sterilization trays for implants and instruments from Spinal Solutions. The letter identified the affected product, problem identified and actions to be taken. Customers are instructed to call the office at the number (951) 304-9001 if they are requesting paper copies of any of the literature. |  | 24 | 24 units | USA Nationwide Distribution in the states of MS: TX, NV, CA, MD, and WI. | Terminated | March 26, 2014 | 399 | 503 | Not_Computer
Z-0560-2014 | 66934 | No | HTD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Edwards Lifesciences Ergonic Double Acting DeBakey Forceps with Narrow Straight Tip. Model Number ERGF01330. |  |  | November 20, 2013 | 2013 | Edwards Lifesciences, LLC | Edwards Lifesciences is recalling one lot of Ergonic instrument due to mislabeling. Ergonic scissors standard curved labeled as a Double Acting DeBakey forcep. Edwards control confirmed that the laser etching on the device was incorrect. | Labeling mix-ups | Edwards sent an " Urgent Field Safety Notice" to all affected customers dated November 20, 2013. The letter identified the product the problem and the action needed to be taken by the customers.Action to be taken by user:Our records show that you received one of these instruments. Our sales representative will be on-site to remove this instrument and provide a replacement. Please assist our sales representative in reviewing your entire inventory for the instrument listed in the recall letter. The form has been pre-populated with the inventory from our records. If you have additional inventory of this affected lot within your control, please add the quantity to the section on the form provided. Please remove the instrument from use, quarantine the device at your site and support the Edwards Sales Representative to return this product to Edwards. An acknowledgment form is included to assist you in the assessment of your inventory.While our sales representative is on-site, please complete the attached acknowledgment form and fax it back to Edwards Customer Service at 800.422.9329. The return of this form allows us to confirm that you have reviewed this notice and have taken appropriate action. Please contact Customer Service at 800.424.3278 if you have questions about the return and crediting process.The Customer Service organization can answer questions about tracking of the replacement instrument provided by the Edwards Sales Representative. If you have questions that have not been answered by this letter, please call Edwards Customer Service at 800.424.3278 from the hours of 6:00AM - 4:30PM PST . |  | 3 | 3 | Distributed to 3 customers in FL only. | Terminated | February 10, 2014 | 82 | 3 | Not_Computer
Z-1176-2013 | 64706 | No | NEW | Absorbable polydioxanone surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ORTHOCORD One Violet and One Blue Braided Composite Suture with OS-6 Reverse Cutting Needles, Product Code: 223116. |  |  | March 22, 2013 | 2013 | DePuy Mitek, Inc., a Johnson &amp; Johnson Co. | The IFU provided with the free strand ORTHOCORD sutures currently indicates that approximately 100% suture strength remains at six weeks after implantation. While this statement is true for Violet ORTHOCORD sutures, it is an inaccurate statement for Blue ORTHOCORD free strand sutures. Based on in-vivo strength retention study, a Blue ORTHOCORD sutures maintain approximately 80% of strength at six | Labeling Change Control | On March 22, 2013, firm notified consignees of recall via letter. Consignees were informed of the following: The IFU provided with the free strand ORTHOCORD sutures currently indicates that approximately 100% suture strength remains at six weeks after implantation. While this statement is true for Violet ORTHOCORD sutures, it is an inaccurate statement for Blue ORTHOCORD free strand sutures. Based on in-vivo strength retention study, a Blue ORTHOCORD sutures maintain approximately 80% of strength at six weeks after implantation. As Blue and Violet ORTHOCORD free strand sutures share the same IFU, IFU-108147 will be revised. Consignees were requested to confirm the receipt of the letter by March 29, 2013. |  | 46501 | 46501 | Nationwide Distribution. | Terminated | February 24, 2014 | 339 | 449009 | Not_Computer
Z-1115-2013 | 64810 | No | GFA | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Product is distributed by Zimmer (Warsaw, Indiana)/Synvasive Technology, Inc. (El Dorado Hills, CA) &amp; manufactured under Synvasive Technology, Inc. El Dorado Hills, CA |  |  | April 02, 2013 | 2013 | Synvasive Technology Inc | Various reciprocating and oscillating saw blades were incorrectly labeled. The labeling errors include: (1) The pouch says oscillating saw blade when it actually contains a reciprocating saw blade; (2) The box bar code date does not match the human readable expiration date; and (3) The label incorrectly states use with Zimmer/Synthes handset and not the Zimmer Universal Power System. | Packaging process control | Zimmer sent a Urgent Device Removal letter dated April 2, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers that received affected product from both Synvasive and from Zimmer Surgical will receive two customer letters, as each letter sets out specific item and lot numbers dependent upon distribution entity. Zimmer Distributors will also receive an e-mail notification. All notifications include a response form for both Synvasive and Zimmer distributors as well as customer consignees to account for product disposition (disposal in the field per facilities' sharps protocol).(1) Notifications will be mailed via Fed Ex to Risk Managers and Zimmer Distributors with inventory.(2) Notifications will be emailed to international geographies.(3) Notifications will be emailed all Zimmer Distributors.(4) A second Fed Ex mailing will be sent to domestic non-responding consignees.(5) Effectiveness phone calls will be made to any remaining domestic non-responding consignees. (6) If a customer reports they have further distributed the product, the secondary consignee identified will receive a notification from Zimmer.The scope of the recall is limited to certain confirmed lots which were shipped to the affected consignees..E-mails, certified mail receipts, and telephone records will be the firm's course of action in monitoring effectiveness checks. All returns will be stored in a separate area away from regular inventory. Final disposition of returned product is yet to be determined.The firm's recall strategy appears to be adequate.For further questions please call (916) 939-3913. |  | 3442 | 3442 units per the part numbers listed | Worldwide Distribution - USA (nationwide) and countries of: Austria, Australia, Belgium, Czech Republic, Denmark, Finland, France, Germany, Italy, Japan, Netherlands, Norway, Panama, Spain, Sweden, Switzerland, Tunisia, United Kingdom and Yemen. | Terminated | July 15, 2014 | 469 | 3442 | Not_Computer
Z-0692-2013 | 64039 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Eterna Giovinezza System, Powered Laser Surgical Instrument. |  |  | November 07, 2012 | 2012 | Quanta System, S.p.A. | Quanta System S.p.A, Italy, is notifying U.S. consignees that their website contained promotional materials for products distributed in the U.S. which did not specify that some of the claims made therein were not meant to apply to the U.S. distributed devices. | Labeling False and Misleading | Quanta System sent a letter dated November 7, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to perform the following: quarantine and destroy all promotional literature printed from the manufacturer, website or obtained from the manufacturer. Provide any promotional literature independently produced by the initial importer or distributor containing information about Quanta System products. Provide URL of initial importers or distributors website if information about Quanta System products is contained therein. Provide a written certification that all actions have been taken. Quanta System has or will: review all promotional materials and websiteinformation provided by initial importers and distributors; advise initial importers and distributors of any actions that need to be taken to change promotional materials and/or website information.For questions regarding this recall call 303-223-4336. |  | 12 | 12 units | Class II Recall - Worldwide Distribution - USA including TX, AZ, OR, and CO.Only US product is affected by this recall. | Terminated | January 23, 2013 | 77 | 109 | Not_Computer
Z-1611-2013 | 65397 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | PinnacleÂż DestinationÂż Peripheral Guiding Sheath, 8 French, 90cm, Straight Tip, Cross Cut Valve. |  |  | May 30, 2013 | 2013 | Terumo Medical Corporation | Medical devices were incorrectly labeled with extended expiration dates. | Error in labeling | The recalling firm notified consignees by letter on 5/30/2013. The letter was flagged as an urgent product bulletin addressed to risk managers. The notification advised customers to immediately discontinue use and quarantine any available stock for subsequent instructions to be provided for pick-up and replacement. |  | 20 | 20 units | Nationwide Distribution. | Terminated | October 30, 2013 | 153 | 570 | Not_Computer
Z-1543-2013 | 65374 | No | DRS | Extravascular blood pressure transducer. | Cardiovascular | Cardiovascular | 510(k) | Deltran Disposable Pressure Transducer Kit (DPT Kit) |  |  | June 04, 2013 | 2013 | Utah Medical Products, Inc | Utah Medical Products, Inc. is recalling certain Deltran Disposable Pressure Transducer Kits due to complaints related to packaging integrity of the sterile packaged devices. | Packaging | Utah Medical Products (UTMD) sent a Recall Notice letter dated June 3, 2013 via fax and email to all affected customers and will be completing follow up telephone calls by June 7, 2013. The letter identified the affected product, problem, health risk, and actions to be taken. UTMD instructed customers to locate and return affected products for exchange. For questions call Customer Service at 800-533-4984. |  | 1800 | 1,800 kits | Worldwide Distribution - United States Nationwide in the states of (MO, NE, KY, FL, OH, NY), and countries of Columbia, Brazil, Venezuela. West Indies, Latin America, Ecuador are distributed through a US direct consignee exporter. | Terminated | September 19, 2013 | 107 | 1800 | Not_Computer
Z-1486-2013 | 65314 | No | GAM | Absorbable poly(glycolide/l-lactide) surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Digitex Absorbable Polydiaxanone Suture, Size 2-0. |  |  | May 13, 2013 | 2013 | Coloplast Manufacturing US, LLC | Coloplast is conducting a recall on three lots of Digitex Absorbable Polydiaxanone Suture, Size 2-0 because it is mislabeled with a three year expiration date. | Incorrect or no expiration date | The firm, Coloplast, sent a letter dated May 8, 2013 to its customers. The letter described the problem, the product and actions to be taken. The customers will be contacted by Coloplast Customer Service with instructions for the handling and return of any unused product.If you have any additional questions, the dedicated Customer Service number regarding this effort is 800-258-3476. |  | 772 | 772 units | US Distribution to: KY, MI, OH, PR, TX, and VA. | Terminated | April 10, 2014 | 332 | 772 | Not_Computer
Z-0665-2013 | 63941 | No | DKB | Clinical toxicology calibrator. | Toxicology | Toxicology | 510(K) Exempt | Siemens Emit(R) tox(TM) Serum Calibrators, catalog number 7B019UL |  |  | December 06, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Incorrect expiration date was assigned to the D1 lot of Emit(R) tox(TM) Serum Calibrators. The correct expiration date is 2013-12-31. The incorrect assigned expiration date is 2014-12-31. | Error in labeling | Siemens sent a Customer Notification letter dated December 2012 to all affected customers of Emit(R) tox(TM) Serum Calibrators lot D1, informing them that Siemens Healthcare Diagnostics, Inc. has confirmed that the expiration of 2014-12-31 is incorrect. Siemens requested customers to discontinue use and discard any remaining inventory of the D1 lots. Customers were instructed to complete the Effectiveness Check form included with the letter even if no replacement reagents were needed and fax to 302-631-8467. Siemens would replace any unused kits of the affected lots at no charge.Customers with questions were instructed to contact their Siemens Technical Solutions Center or their local Siemens technical support representative.For questions regarding this recall call 302-631-6311. |  | 1998 | 1,998 cartons | Nationwide Distribution to the following states: AL, AR, CA, CO, FL, GA, IL, IN, KY, MA, MD, MI, MO, NC, NY, OH, PA, TN, TX and WI. | Terminated | December 24, 2013 | 383 | 1998 | Not_Computer
Z-0530-2014 | 65477 | No | JXB | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | remel RapID Yeast Plus System, 20 tests, Ref: R8311007, IVD, containing a 15-ml bottle of RapID Yeast Plus Reagent A and a 10-ml bottle of RapID Yeast Plus Reagent B. |  |  | April 04, 2013 | 2013 | Remel Inc | The two reagents supplied in the RapID Yeast Plus System kit (Reagent A and Reagent B) are mislabeled. The labels were reversed. | Labeling mix-ups | Thermo Fisher Scientific sent a Medical Device Recall letter dated April 3, 2013 to affected customers. The letter identified the affected product, problem and actions to be taken. The recalling firm instructed customer to review their inventory for affected lot and discard the product. A Medical Device Recall Return Response was enclosed for the customer to indicate they understood the instructions and to report the amount of product they discarded. For questions call Technical Services Department at 800-255-6730 (US) or 913-888-0939 (International). |  | 150 | 150 kits | Worldwide Distribution - USA Nationwide in the states of AL, AZ, CA, CT, FL, ID, IN, KY, MA, MI, MS, NJ, OH, PA, TX, WI and the countries of Canada, UK, and Japan. | Terminated | March 12, 2014 | 342 | 150 | Not_Computer
Z-1549-2013 | 33728 | microscope, phase contrast | IBM | Microscopes and accessories. | Pathology | Pathology | 510(K) Exempt | Aperio ScanScope System. Aperio ScanScope Models, XT, AT, AT Turbo, CS and CSO. |  |  | February 22, 2013 | 2013 | Aperio Technologies Inc | Aperio initiated this recall because the User Guides incorrectly state the devices had been cleared for specific applications. The User Guides should state the devices are for Research Use Only and are not for use in diagnostic procedures. | No Marketing Application | Aperio sent an Urgent Medical Device Recall letter dated April 22, 2013 to all affected consignees via certified USPS mail. The letter included the reason for the recall, recall product information, instructions, acknowledgement form, contact information ( Aperio Technical Support at 1-866-478-3999), and product labeling. The consignees were requested to acknowledge receipt of the Medical Device Correction letter. In addition, an Acknowledgement Form is required to be signed and e-mailed back to the Manufacturer asevidence of letter receipt.For questions regarding this recall call 760-304-6211, ext 6211. |  | 364 | 364 | Nationwide Distribution - including all states, none internationally. | Terminated | September 25, 2013 | 215 | 364 | Not_Computer
Z-0523-2014 | 66964 | No | HTD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | outer box: Ref 906768, CurvTek Eye Needle 22 MM X-large Stainless Steel, Single Use Gamma Sterilized, Rx |  |  | November 26, 2013 | 2013 | Biomet, Inc. | Inner label of PN: 906768 Lot: 055450 CurvTek Eye Needle 22mm X-Large states incorrectly PN: 906760 7mm Medium. | Labeling mix-ups | Biomet sent an Urgent Medical Device Recall Notice dated November 26, 2013, to all affected customers. The notice identifies the product, problem, and actions to be taken by the customers. The letter included a Customer Response Form for customers to complete and return to the firm by FAX 574-372-1683. Customers with questions were instructed to call 574-372-1570.For questions regarding this recall call 574-372-1570. |  | 21 | 21 | Nationwide Distribution including MI, IA, IL, IN, and TN. | Terminated | July 07, 2015 | 588 | 21 | Not_Computer
Z-1544-2013 | 64216 | No | EKR | Dental hand instrument. | Dental | Dental | 510(K) Exempt | SybronEndo Elements Obturation Unit; REF 973-0322; SN 91-1234; 2005-10; SybronEndo, 1332 S. Lone Hill Ave. Glendora, CA 91740; EU Rep: Kerr Italia, S.r.l, Via Passanti, 332, 1-84018 Scafati (SA), Italy |  |  | January 17, 2013 | 2013 | Ormco/Sybronendo | A recall has been issued by SybronEndo because a Canadian Standards Association (CSA)-directed cautionary statement label was inadvertently omitted from the Elements Obturation Unit and its Replacement A/C Power Cord. | Labeling mix-ups | The firm SybronEndo sent a "FIELD CORRECTION ACTION" letter dated January 30, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were requested to review and determine if they have any of the affected equipment in their inventory; Note: the equipment itself has not been modified in any way, and therefore there is no change in the performance; customers should continue to use the equipment in the same manner they have in the past:; Note: SynbronEndo has enclosed the cautionary label along with instructions for proper placement of the affected equipment; the customers should apply the enclosed cautionary statement label to Elements Obturation Unit and/or Replacement AC Power Cord as per instructions and complete and return the FIELD CORRECTION ACKNOWLEDGEMENT FORM ,whether or not they have any affected equipment in their inventory, via Fax to: 909-962-5605.For questions regarding this recall call 714-516-7602.Note: If you are an authorized SybronEndo international distributor, we request that you identify those customers that may have been shipped the affected lots of equipment and contact these customers to inform them of the issue within forty-eight (48) hours of receipt of this notification.If you have any questions, contact SybronEndo Customer Care at 1-800-346-3636. |  | 3615 | 3615 units total | Worldwide Distribution - USA (nationwide) and the countries of Australia, Azerbaijan, Belgium, Canada, China, Cyprus, Czech Republic, Egypt, Estonia, Finland, France, Germany, Greece, Hong Kong, Iceland, India, Italy, Japan, Kuwait, Latvia, Lebanon, Malaysia, Netherlands, New Zealand, Pakistan, Poland, Romania, Russia, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Thailand, Turkey, Ukraine, United Arab Emirates, and United Kingdom. | Terminated | November 07, 2013 | 294 | 3615 | Not_Computer
Z-1142-2013 | 64884 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Innova 2100IQ, Innova 3100IQ, Innova 4100IQ, |  |  | October 21, 2011 | 2011 | GE Healthcare, LLC | As part of GE Innova IQ table introduction, a label was designed to be affixed on the front Cover of theGantry Pivot section. This safety label is missing on systems manufactured between March 2010 and July 2011. In the situation a staff member (doctor, nurse, operator, anesthetist, etc&amp;) stands in the region between the C-arm gantry pivot and the head section of the table, there is a risk of co | Labeling design | A field action was initiated on October 21, 2011, and completed March 9, 2012, . The issue was corrected to all affected systems by applying the missing label onto the gantry. No recall letter was issued.For further questions please call (262) 513-4122. |  | 48 | 48 | Worldwide Distribution - Nationwide Distribution including the states of FL, IL, IN, LA, MI, MN, NJ, NM, NY, OH, PR, SC, TX and VA and the countries of Vietnam, Latvia, Italy, Israel, Indonesia, Hong Kong, United Kingdom, Finland, Spain, Czech Republic, Canada, Bahrain and Australia. | Terminated | February 04, 2014 | 837 | 48 | Not_Computer
Z-1135-2013 | 64747 | Yes | NPT |  |  | Cardiovascular | PMA | Edwards Lifesciences Crimpers, Model 9100CR26. |  |  | March 26, 2013 | 2013 | Edwards Lifesciences, LLC | Holes in the sterile barrier pouch may be caused by handling of the device during manufacturing or once pouches have been removed from the unit box and stored on the shelf by customers. | Package design/selection | Edwards Lifesciences notified customers via letter and instructed them to maintain the device in its protective unit carton until ready for use and to inspect the pouch for holes or tears prior to use as instructed in the product's labeling. If any holes or tears are found, customers are instructed to contact Edwards Customer Service to arrange for return and replacement of damaged product. |  | 6117 | 6,117 units | Nationwide Distribution-USA (nationwide) including DC and the states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, ME, MD, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, and WI. | Terminated | September 16, 2013 | 174 | 11589 | Not_Computer
Z-1507-2013 | 65182 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | FLEXSTENT Biliary Self-Expanding Stent System. |  |  | April 23, 2013 | 2013 | Cordis Corporation | Holes and/or tears can be introduced in the FLEXSTENTÂż inner pouch under some conditions, and to a lesser degree, in the outer pouch. | Package design/selection | Cordis sent an Urgent Medical Device Correction and Potential Removal letter dated April 24, 2013 e-mail to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immediately identify and set aside all affected products to prevent further use of the affected product, until inspected. Âż Use only product identified with a blue dot next to the lot number on the carton face label, carton end label, and outer pouch label, or product with upgraded sterile barrier system (not yet available, as of the date of this letter), review, complete, sign and return the enclosed Acknowledgement Form and maintain a copy of this notice with the affected product. For questions related to the issue, please contact our FSS/Cordisrepresentative, or please contact Cordis Direct at (800) 781-0282, Mondaythrough Friday from 9:00 AM to 5:00 PM EDT. |  | 23 | 23 units | Worldwide Distribution - USA (nationwide) and the states of: AZ, CA, CT, FL, GA, IA, IL, LA, MD, MI, MS, NC, NJ, NY, OH, PA, RI, SC, TX, WA, WI, WV, and the countries of: Czech Republic, Italy, Netherlands, and New Zealand | Terminated | June 27, 2014 | 430 | 2724 | Not_Computer
Z-1515-2013 | 65224 | No | JIT | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ADVIA CentaurÂż iPTH Calibrator 6 pk; REF: 10492387, 10630628. |  |  | May 14, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc | iPTH Assay- reagent and calibrator lots may have demonstrated higher than anticipated variability | No Marketing Application | Siemens sent an Urgent Medical Device Correction letter dated May 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. For questions, contact your Siemens Technical Solutions Center or your local Siemens technical support. |  | 198 | 198 (6 pk) | Nationwide Distribution. | Terminated | March 05, 2015 | 660 | 3908 | Not_Computer
Z-1127-2013 | 64646 | No | HSD | Shoulder joint humeral (hemi-shoulder) metallic uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | The Oxford Partial Knee; Oxford Knee Resection Procedure 3 Pack CEMENTED |  |  | March 14, 2013 | 2013 | Synvasive Technology Inc | Biomet part # 506076, lot 928182 was received from Synvasive Technology containing the incorrect blade. | Packaging change control | An Urgent Medical Device Recall Notice letter dated March 21, 2013 was sent to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to discontinue use; remove the products from inventory and complete and return the attached Response Form. For questions call 574-372-1570. |  | 100 | 100 kits | USA Nationwide Distributor including the state of IN | Terminated | January 13, 2014 | 305 | 100 | Not_Computer
Z-1517-2013 | 65289 | No | DXE | Embolectomy catheter. | Cardiovascular | Cardiovascular | 510(k) | lnvatec SpA brand Diver C.E., Model/Catalog Numbers: DVR010014145, DVR011014145, |  |  | March 04, 2009 | 2009 | Invatec Llc | Devices labeled for an intended use was not included in the existing 510(k) or PMA, in that, the Indications for use was modified. | Labeling False and Misleading | Invatec sent an Urgent - Device Field Action letter dated March 4, 2009, to all affected customer. The letter identified the product, the problem, and the action to be taken by the customer. Customers were required to return a confirmation form, acknowledging the action taken, and identifying affected units in their inventory for return of the product for replacement. The firm indicated that this recall was fully executed in 2009, and this 806 captures the events that took place at that time.For questions regarding this recall call 707-566-1143. |  | 177 | 177 units | Nationwide Distribution | Terminated | June 10, 2013 | 1559 | 177 | Not_Computer
Z-0546-2014 | 66589 | No | CAZ | Anesthesia conduction kit. | Anesthesiology | Anesthesiology | 510(k) | Continuous Peripheral Nerve Block Catheter Kit and Set |  |  | October 09, 2013 | 2013 | Arrow International Inc | A labeling inconsistency was discovered in which the lidstock states that the kit contains a stimulator cable with 1.5mm socket, however, the kit contains a stimulator cable with 1.5mm pin. | Labeling mix-ups | The firm, Arrow International, Inc., sent an "Urgent Medical Device Recall Notification" letter dated October 14, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately discontinue use and quarantine any of listed affected products/lots you have in stock; return affected products, customer service representative will contact you with RGA# and instructions for the return of the product, and complete enclosed Recall Acknowledgement form and fax to 1-800-343-2935, Attn: Customer Service, even if you have not affected product.If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-800-343-2935. |  | 121 | 121 | Worldwide distribution: US (nationwide) including states of: MA, LA, CA, GA, KS, TX, PA, NY, MO, WI, NJ, MS, CT, ID, AL, CO, and PR; and Singapore. | Terminated | September 21, 2015 | 712 | 121 | Not_Computer
Z-1143-2013 | 64538 | No | MCW | Intraluminal artery stripper. | Cardiovascular | Cardiovascular | 510(k) | Elana Catheter Sizer 2.0*** |  |  | February 14, 2013 | 2013 | Elana, Inc. | Distribution of an unapproved device into interstate commerce. | No Marketing Application | Elana Inc. issued an Urgent: Medical Device Recall letter on February 14, 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to return the Elana Catheter Sizer 2.0 as soon as possible. The Sizer 2.0 was intended as an optional accessory to the Elana Surgical Kit. The Elana Surgical Kit may be used properly without the Elana Catheter Sizer 2.0. Please acknowledge receipt of these instructions immediately by Email: Claartje@elana.com or Fax: 301-215-7038. The return shipping address for the Sizer 2.0 is Elana, Inc., 10480 Little Patuxent Parkway, Suite 400, Columbia, MD 21044 USA. Contact information for questions: Phone: 301-215-7010 or 443-538-4923 or +1 31 30 2537276 www.elana.com. |  | 3 | 3 devices | Distributed in the states of Pennsylvania, Texas, and New York. | Terminated | May 09, 2013 | 84 | 3 | Not_Computer
Z-1550-2013 | 65245 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | IDEV Technologies SUPERA VERITAS Interwoven Self-Expanding Nitinol Stent Transhepatic Biliary System. Indicated for palliative treatment of biliary strictures produced by malignant neoplasms |  |  | May 29, 2012 | 2012 | IDEV Technologies, Inc. | There was an error on the outer labeling of this device. The inner label was correct. | Labeling mix-ups | The firm decided to rework the products that were in their control at the time and the one that was not in their control was corrected by an IDEV employee. That hospital was called and notified of the situation in May 2012. |  | 3 | 3 | Nationwide Distribution to the states of FL, NY, and PA. | Terminated | July 29, 2013 | 426 | 3 | Not_Computer
Z-1465-2013 | 65209 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Active Articulation E1. |  |  | April 26, 2013 | 2013 | Biomet, Inc. | Biomet is recalling Active Articulation E1 Hip Bearing implants. Biomet has initiated this action following an investigation which identified that outer label indicates two conflicting O.D. sizes while the inner labels indicate an incorrect O.D. size.The implant itself is etched with the correct O.D. size. | Labeling mix-ups | Biomet sent an Urgent Medical Device Recall letter dated April 26, 2013 to consignees detailing that the outer label indicates two conflicting O.D. sizes while the inner labels indicate an incorrect O.D. size. This action requires the immediate location and discontinued use of the product and its return to Biomet. Ciustomers were advised to:Immediately locate and remove the identified device(s) listed from circulation .-Carefully follow the instructions on the enclosed "FAX Back Response Form" .- Fax a copy of the Response Form to 574-372-1683 prior to return of product.-Use priority carrier for your shipment.- If you have further distributed this product, you MUST notify hospital personnel of this action via the enclosed "Dear O.R. Manager" notice. This letter MUST be given to hospital personnel responsible for receiving recall notices. However, you are charged with the location and return of these products. |  | 8 | 8 units | Distributed in New York. | Terminated | August 29, 2013 | 125 | 8 | Not_Computer
Z-0656-2013 | 61718 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Accolade TMZF Plus 127 degree Neck Angle V40 Hip Stem #2.5 and Accolade TMZF Plus 127 degree Neck Angle V40 Hip Stem #3.5, Howmedica Osteonics Corp. The intended use is for reconstruction of the head and neck of the femoral joint. |  |  | October 23, 2012 | 2012 | Stryker Howmedica Osteonics Corp. | Stryker has received a report that single size 3.5 stem was packaged as size 2.5 stem resulting in a possible product mix. | Labeling mix-ups | Stryker sent "Urgent Medical Device Recall Notification" letters and product accountability forms via Fed Ex with return receipt on 10/23/12 to branches/agencies and to hospitals and physicians on 10/24/12. The letter described the issue and provided recommended actions to mitigate the risk. Customers can call (201) 972-2100 if questions. |  | 18107 | 6021-2530- four units; 6021-3535- | Nationwide distribution, including the states of NY, NJ, CA and FL. | Terminated | September 11, 2013 | 323 | 18107 | Not_Computer
Z-1477-2013 | 64164 | No | GQO | Herpes simplex virus serological assays. | Microbiology | Microbiology | 510(k) | Plastic bottle containing 10.5 mL of IgG Conjugate |  |  | January 23, 2013 | 2013 | Bion Enterprises Ltd | Affected products have a labeled expiration date of one year too long. | Error in labeling | MBL Bion sent an Urgent Product Recall letter dated January 28, 2013 to all affected customers via email. The letter identified the affected product, problem and actions to be taken. Customers were requested to destroy the affected product and to notify Bion of the number of units consumed and destroyed. For questions call 847-544-5044. |  | 207 | 207 units | Worldwide Distribution - United States (nationwide) in the states of IN, SC and the countries of Canada (Toronto); South Africa (Randburg) and Brazil (Rio de Janeiro) | Terminated | May 05, 2015 | 832 | 912 | Not_Computer
Z-1596-2013 | 65138 | No | MED | Liquid chemical sterilants/high level disinfectants. | General Hospital | General Hospital | 510(k) | Quick Connect QLC1677E, QRC1699E, QPC1721E and QFC1683E. STERIS Corporation, Mentor, OH 44060 |  |  | March 04, 2013 | 2013 | Steris Corporation | During a routine label review, the firm identified four SYSTEM 1E Quick Connects that will be updated to reflect the removal of obsolete device models. This labeling review process also highlighted the opportunity to combine two existing quick connects into a single product while eliminating the obsolete models. | Labeling Change Control | STERIS sent a Customer Notification letter dated February 19, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to replace their existing labels with labels they provided in their customer notification package.We apologize for any inconvenience this may have caused you and as always STERIS Corporation is dedicated to support our products and valued Customers. If you have any questions regarding this matter, please contact your STERIS Capital Account Manager. For further questions please call (440)392-7348 |  |  | QLC1677E: 38; QRC1699E: 128 &amp; QPC1721E/QFC1683E: 0 | Nationwide Distribution including the states of AK, AZ, CA, CO, FL, IL, IN, KY, LA, MA, MD, ME, MI, MO, MS, NC, NE, NH, NY, OH, OK, OR, PA, SD, TN, TX, VA, WA, WI and WY. | Terminated | July 02, 2013 | 120 | 0 | Not_Computer
Z-1080-2013 | 64594 | No | GWQ | Electroencephalograph. | Neurology | Neurology | 510(k) | Harmonie software versions 5.1 uo to 6.2e : |  |  | April 06, 2009 | 2009 | Stellate Systems | The firm is initiating this recall because the implied or stated indications in their labeling associated with Harmonie software have not yet been cleared by the FDA for sale in the USA. | No Marketing Application | Stellate sent an "DEVICE CORRECTION" letter dated April 6, 2009, to all affected customers. The letter identified the product, the problem and the action needed to be taken by the customer. Customers were instructed to check the product code referenced in the service manual to verify if your system is affected and check in "Help" &amp; "About Harmonie" in the menu to determine the software version on these systems. Stellate will shortly release updates to disable the noted features that were present in your particular version of the Harmonie software and revise labeling and Instructions for Use. All customers in possession of an affected system will be contacted by our Customer Support team to set up an appointment for installation of the updated version.If you have any questions or concerns regarding this notification, please do not hesitate to call our customer support line at 1-888-742-1306. |  | 2660 | 2660 Software licenses | Worldwide distribution: USA (nationwide) including states of: AL, AR, AZ, CA, CT, DE, FL, HI, IA, IN, IL, KS, KY, LA, MA, MD, ME, MI, MN, MO, NC, ND, NJ, NY, OH, OR, PA, RI, SC, TX, VA, VT, WI, WV, and WY; and the countries of: Argentina, Australia, Bahrain, Brasil, Canada, China, Denmark, Finland, France, Germany, Greece, Haiti, Hong Kong, Hungary, Indonesia, Iran, Israel, Italy, Japan, Korea, KSA, Mexico, Netherlands, Norway, Pakistan, Portugal, Rep. DE Panama, Saudi Arabia, Spain, Switzerland, Taiwan ROC, Thailand, Turkey, and United Kingdom. | Terminated | April 22, 2013 | 1477 | 2660 | Not_Computer
Z-1119-2013 | 64610 | No | MRG | Vitamin D test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA Centaur Vitamin D Assay (100 tests); Reference Number: 10491994. |  |  | March 13, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc | Negative bias in patient samples for Vitamin D | No Marketing Application | Siemens Healthcare issued Urgent Medical Device Recall was sent to all affected Siemens Healthcare Diagnostics customers in the United States on March 13th, 2013. This notice informs the customer of the recall and the actions that need to be taken. The Urgent Medical Device Recall instructs customers to discontinue the use of ADVIA Centaur VitD Assay Kit Lots (REF 10491994 and 10631021) and ADVIA Centaur Systems VitD Cal (REF 10493589 and 10630911) released prior to the recalibration of the assay and request replacement lots. This notice also directs customers to review the contents of the recall notice with their laboratory or medical director. A confirmation fax-back form is included to ensure that all customers in the United States have been notified and have complied with the mandatory action. |  | 3871 | 3871 | Nationwide Distribution. | Terminated | March 05, 2015 | 722 | 8489 | Not_Computer
Z-1157-2013 | 64578 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | APLIF Implants and Instruments, |  |  | March 08, 2013 | 2013 | Spinal Solutions, LLC | Spinal Solutions is recalling the APLIF system because it is not supported by adequate testing and documentation to demonstrate that it meets performance or safety standards. These inadequacies might result in product performance failures that could cause patient harm due to implant breakage, movement, or inadequate sterilization. | No Marketing Application | Spinal Solutions LLC sent an Urgent Medical Device Recall letter dated April 12, 2013, to all affected customers. The letter informed the customers of the problem identified and the actions to be taken. Customers were instructed to discontinue use of the affected product and return to the firm. Customers were also instructed to complete the bottom of the form and to return it by fax to (858) 764-9739.For questions regarding this recall call 951-304-9001. |  | 220 | 220 units | Nationwide Distribution including NV, WI, MD,and CA. | Terminated | March 26, 2014 | 383 | 220 | Not_Computer
Z-1451-2013 | 65160 | No | FXX | Surgical apparel. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Comfort brand Ear-Loop Face Mask, Item #: 3100, Model Numbers: # 3100B, # 3100P, # 3100G, # 3100Y, # 3100W; Product is manufactured and distributed by Maytex Corporation, CA |  |  | May 15, 2013 | 2013 | Maytex Corp | Recent Lab testing conducted in March 2013 indicates that the product does not meet the standard of "Differential Pressure of only 1.6 Delta P (mm H2O/cm2)" as indicated on the box/package of the product; and also product manufactured between 2008 - 2012 cannot be guaranteed as Fluid Resistant. | Labeling False and Misleading | The firm, Maytex Corporation, sent a "MEDICAL DEVICE SAFETY NOTICE" dated May 15, 2013, to all its customers. The notice described the product, problem and actions to be taken. The customers were instructed to immediately contact your medical professional for treatment and contact Maytex Corporation to file a complaint should any illness or injury occur; discontinue use of the product if you detect any potential health hazards; if this product has been further sold or distributed, notify your customers to the retail level and send copies of safety notice to customer/users; and complete and promptly return the enclosed Medical Device Safety Notice Return Response Acknowledgement and Receipt form via mail to: Maytex Corporation 23521 Foley Street, Hayward CA 94545; fax at; 510-786-0209; email: www.maytex18@aol.com or by delivery to Maytex Corporation.Maytex Corporation has temporarily taken down their Website at www.maytexcorp.com to correct the wrong product specifications as described in the notice.If you have any questions, contact Maytex Corporation at 1-800-462-9839 or email: maytex18@aol.com Monday through Friday, 8:30 AM to 5:00 PM Pacific Time. |  | 43915 | 43,915 Cases | Worldwide distribution: USA (nationwide) including states of: AL, AZ, CA, CO, FL, GA, HI, ID, IL, IN, MA, MD, ME, MI, MN, MO, MT, NE, NH, NJ, NM, NY, OK, OR, PA, PR, TN, TX, UT, WA, and WI; and country of: Mexico. | Terminated | October 29, 2013 | 167 | 43915 | Not_Computer
Z-0493-2014 | 66967 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | G7 Acetabular Screw 6.5mmx35mm |  |  | November 26, 0013 | 13 | Biomet, Inc. | Investigation identified that the outer label, inner label, and patient label all have the incorrect diameter size of 10.5mm. The screw is actually a 6.5mm diameter and correctly corresponds to the items number on the label. | Labeling Change Control | Beginning November 25 Urgent recall electronic notices were sent by Biomet UK to all distributors with hospitals outside the USA notified by ground mail November 27,2013. The notification identified the product and reason for the recall, asking for immediate location and discontinued use of the product and return to Biomet UK, Ltd. Questions can be addressed to Biomet USA at 01 574-372-1570 M-F 8 am to 5 pm. |  | 532 | 532 in USA all under Biomet USA control | Worldwide Distribution - USA Nationwide and in the countries of Europe, Italy, UK, Australia, Japan. all product in USA is under Biomet control. | Terminated | September 29, 2014 | 730792 | 532 | Not_Computer
Z-0516-2014 | 66894 | No | GGN | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(K) Exempt | HemosIL High Abnormal Control 3 UNASSAYED |  |  | November 15, 2013 | 2013 | Instrumentation Laboratory Co. | HemosIL Low, Abnormal, Normal Controls UNASSAYED (no-value assigned) lacks 510(k) approval | No Marketing Application | Instrumentation Laboratory issued Important Correction Notice via USPS to US customers dated November 2013 advising users in the US that unassayed controls are not acceptable for use in their unassayed (non-value assigned) format. IL will no longer sell or distribute controls without assayed values for the US market as of November 15, 2013, and are converting US customers to an assayed control format. Additional contact with customers via phone calls and a reminder email.An enclosed response tracking form provided to be returned. For technical questions, contact the IL Technical Support Center at 1-800-678-0710, Option # 6 (ACL Coagulation Products).For replacement product if value assignments are not available for your product lot, please contact your IL Customer Service Representative at 1-800-955-9525, Option # 6.For questions regarding this recall call 781-861-4467. |  | 84561 | 84561 kits | Nationwide Distribution | Terminated | March 25, 2014 | 130 | 171384 | Not_Computer
Z-0490-2014 | 66735 | [6 empty] | Nephros Dual Stage Ultra Filter (DSU) Brochure |  |  | October 28, 2013 | 2013 | Nephros Inc | Promotional materials for non-medical water filtration products were determined to promote claims which constitute marketing the product as a medical device. | Labeling False and Misleading | Nephros Inc., sent an Urgent Medical Device Recall letter dated October 29, 2013 to all affected customers via UPS 2nd Day Air. The letter identified the affected product, problem and actions to be taken. Customer are asked to provide a customer information list, immediately remove and discard all affected product labeling materials and any independent labeling materials they have created for the affected products. Customers are instructed to complete and sign the attached Response Form to acknowledge receipt, understanding and compliance with the instruction. For questions contact the Nephros representative at 201-343-5202 x 100.The letter was also posted at : http://www.nephros.com/download/news/DSU_Recall_Letters.pdf |  | 3367 | 3,367 untis with affected labeling | USA Nationwide Distribution | Terminated | July 06, 2016 | 982 | 3367 | Not_Computer
Z-1125-2013 | 64568 | Yes | BTL | Powered emergency ventilator. | Anesthesiology | Anesthesiology | 510(k) | Instrumentation Industries BE183-SU and BE 183 SU-12 Manual Jet Ventilators; BE 183-SUR and BE 183 SUR-12 Manual Jet Ventilators with Regulator and Gauge; BE 183-2 4 ft. Small Bore Tubing Assembly w/Inline Filter; and BE 183-R Regulator |  |  | September 07, 2012 | 2012 | Instrumentation Industries Inc | Firm initiated an update to Directions for Use necessitating replacement of previously distributed Directions for Use. | Error in labeling | Instrumentation Industries, Inc. initiated a recall of this product on September 7, 2012 by sending a letter via priority mail to all consignees. The letter identified the product, the problem, and the action to be taken by the customer. Consignees were instructed to replace the package insert of any affected product with the enclosed corrected package insert (now at Rev. H). The Directions for Use can alsbe downloaded from their website if they should need additional copies. If the product was further distributed consignees were asked to notify their consignees about this Field Correction.For questions regarding this recall call 412-854-1133. |  | 2998 | 2998 | Worldwide Distribution - USA including AK, AL, AZ, CA, CO, CT, DC, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NH, NJ, NV, NY, OH, OK, OR, PA, PR, SC, SD, TN, TX, UT, VA, VT, WA, WI and WV. Shipped internationally to Canada, Ireland, Hong Kong, the Netherlands, New Zealand, and Thailand. | Terminated | July 05, 2013 | 301 | 2998 | Not_Computer
Z-1565-2013 | 65254 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Baxter, MiniCap with Povidone-Iodine solution, 5C4466P, Baxter Healthcare Corporation. For dialysis. |  |  | May 21, 2013 | 2013 | Baxter Healthcare Corp. | leaking pouches | Packaging change control | The firm, Baxter Healthcare Corporation sent an "URGENT PRODUCT RECALL" letter dated May 21, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. For general questions regarding this communication, contact The Center for One Baxter at 1-800-422-9837. |  | 629280 | 629,280 units | Nationwide Distribution. | Completed |  |  | 629280 | Not_Computer
Z-1564-2013 | 65458 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Medtronic, Launcher Guiding Catheter, Model # LA8EBU45SH, Lot # 0006722866, 8F and Model # LA6IMAD, Lot # 0006724224, 6F |  |  | June 12, 2013 | 2013 | Medtronic Inc. Cardiac Rhythm Disease Management | Medtronic has identified a labeling issue for 24 units of the Launcher coronary Guide Catheter. The two affected catheter lots are mislabeled and do not accurately reflect the catheters contained within each package. | Labeling mix-ups | Medtronic "Urgent Field Safety Notice" letter dated June 2013 to all affected customers. The letter described the affected product, problem and actions to be taken. The letter instructed customers to remove and quarantine all affected product within their inventory and return the affected product to Medtronic. For questions contact your local Medtronic representative. |  | 15 | 15 | Internationally distributed: to countries of Germany and Poland. | Terminated | November 19, 2013 | 160 | 15 | Not_Computer
Z-1568-2013 | 65362 | No | KYI | Wrist joint carpal trapezium polymer prosthesis. | Orthopedic | Orthopedic | 510(k) | CMC Saddle |  |  | June 03, 2013 | 2013 | Integra LifeSciences Corp. | As a result of a quality review, labelling content discrepancies were identified between the Instruction for Use (IFU) distributed in the US with the Ascension CMC and the content that is cleared by FDA. | Labeling Change Control | Integra sent an Urgent: Medical Device Recall letter to all affected product. The letter identified the affected product, problem and actions to be taken. Customers were instructed to examine inventory, discontinue use, and complete and return the attached form. When the form is received, Customer Service Returns and Repairs will send an RMA number to return the affected inventory. For quwstions contact your regional managers or call 512-520-7697. |  | 467 | 467 devices | USA Nationwide Distribution | Terminated | October 08, 2014 | 492 | 467 | Not_Computer
Z-1572-2013 | 65417 | No | FGY | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | MAQUET GETINGE GROUP Arterial Cannula OD |  |  | May 14, 2013 | 2013 | Maquet Cardiovascular Us Sales, Llc | The specified part numbers and lot numbers of Maquet Pediatric Arterial Cannula Packs may contain an incorrectly sized connector. | Packaging process control | Maquet sent an "Urgent Product Recall"/"Recall - Device Corrective Action Response" forms dated May 14, 2013, via Federal Express to all affected customers. The notification identifies affected product; informs customers of the nature of the issue; advises of potential adverse patient consequences that may occur; and advises customers to follow instructions on the "Recall-Device Corrective Action Response" forms included as part of the notification.If you have any further questions please contact your local MAQUET representative or call Customer Service at (888) 627-8383. |  | 100 | 100 units (60 Domestic; 40 International) | Worldwide Distribution - US Distribution including the states of Delaware, Illinois, Minnesota , New York., Ohio and Pennsylvania and the countries of Germany, Turkey and France. | Terminated | September 18, 2014 | 492 | 100 | Not_Computer
Z-0543-2014 | 66844 | No | MVN | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic | 510(k) | ClearSight(TM) Toric 1 day (ocufilcon D), 30 daily disposable contact lenses for astigmatism, STERILE, Rx only --- CooperVision Hamble, SO31 4RF, UK - Scottsville, NY 14546 USA - Juana Diaz, PR 00795 USA |  |  | November 20, 2013 | 2013 | CooperVision Inc. | Lots were labeled with incorrect expiration dates on secondary package (carton). The primary package (blister) is marked with the correct expiration date. | Incorrect or no expiration date | CooperVision sent an Urgent Medical Device Recall Letters dated November 20, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers. Our records indicate you received some of the affected product. Enclosed you will find a list of affected carton lot numbers, with the correct expiration date, that were shipped to your account." Please immediately examine your inventory," Stop any further distribution of carton(s) subject to the recall," Quarantine carton(s) with incorrect expiration date, and" CooperVision will make arrangements to remove affected, quarantine product.As soon as possible, please complete the attached response form and fax it to 1.866.484.1048 or scan and email to info@coopervision.com. If you have product to return please contact customer service for instructions and to schedule pick-up and removal of affected, quarantined product. Your account will be credited once the product is returned and processed.We appreciate your assistance and apologize for any inconvenience you or your patients experienced by this removal. If you have any further questions, please feel free to contact us at 800-341-2020. |  | 45 | 45 cartons | Worldwide Distribution - USA (nationwide) and the countries of Aruba, Canada, Colombia and Trinidad. | Terminated | July 22, 2015 | 609 | 198 | Not_Computer
Z-0714-2013 | 63959 | No | FIE | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Medisystems 15 Gauge x 3/5 inch Buttonhole Needle Set |  |  | December 20, 2012 | 2012 | Medisystems a NX Stage Company | A total of 5 (ea) single needle pouches are mislabeled as 15 Ga x1" instead of 15 Ga x 3/5". | Labeling mix-ups | Medisystems sent a notification letter dated December 20, 2012 to affected customers. The letter identified the affected products, problem and actions to be taken. Customers were instructed to immediately open and inspect inventory for the affected products. Results of the inspection are to be provided to Medisystems and the product is to be returned to Medisystems. For questions call (978) 332-5973. |  | 26 | 26 casesx 250 needles (6500needles) | US Nationwide Distribution in the states of NY, PA and TN. | Terminated | October 10, 2013 | 294 | 42 | Not_Computer
Z-0716-2013 | 64043 | No | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Abbott Point of Care i-Stat cTnI cartridge. |  |  | November 01, 2012 | 2012 | Abbott Point of Care Canada Limited | APOC has determined that 10 boxes of i-Stat cTnI cartridges from lot P12179 have an incorrect barcode applied to the cartridge portion pack. | Error in labeling | APOC sent an "IMPORTANT PRODUCT INFORMATION" letter dated October 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact Abbott Point Of Care Technical Support at 800-366-8020, Option 1 or your Abbott Point Of Care representative for questions regarding this information. |  | 560 | 560 boxes (25 cartridges in each box) | Worldwide Distribution. | Terminated | June 14, 2013 | 225 | 560 | Not_Computer
Z-0661-2013 | 63877 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | STERRAD NX Cassettes, P/N 10133. |  |  | September 20, 2012 | 2012 | Advanced Sterilization Products | Advanced Sterilization Products is recalling three lots of STERRAD NX System Cassettes because it does not have adequate data to support the entire duration of the product's labeled shelf-life. | Labeling design | Advanced Sterilization Product sent an Urgent Medical Device Recall (Removal) letter dated December 4, 2012, to all affected customers. The letter informed the customers that the Sterrad NX Cassettes do not have adequate data to support the entire duration of the product's labeled shelf-life. Customers are instructed of the problems identifed and the actions to be taken. Customers are instructed to notify their customers if they have further distributed the recalled products and to have them contact Stericycle at (877) 650-7686 to arrange for return of the product. Customers with questions are instructed to call Stericycle at (877) 650-7686 Mon-Fri from 8am to 5pm, ET.For questions regarding this recall call 949-453-6400. |  | 12280 | 12,280 units | Worldwide Distribution - USA (nationwide) and internationally to the following countries:AR, CO, TW, SI, AU, MX, MY, HU, PH, BE, FR, DE, IE, IT, ES, NO, CH, and GB. | Terminated | October 17, 2013 | 392 | 12280 | Not_Computer
Z-1347-2013 | 64618 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | PROGRESSIVE MEDICAL Laparoscopic Electrode PTFE Coated L Hook 33cm. |  |  | February 15, 2013 | 2013 | Unimed Surgical Products, Inc. | Firm is recalling due to concerns regarding changes that were made to the PTFE coating materials. | No Marketing Application | Unimed Surgical Products, Inc. issued a Medical Device Correction and Removal letter dated February 15, 2013 to its customers. Customers were instructed to quarantine any undistributed product, and notify its customers of the recall to retrieve any unused product. For further information, contact Unimed Surgical Products Inc. at 1-800-531-2627. |  | 120416 | 120,416 units | Distributed in the states of MO, TN, and PA. | Terminated | August 07, 2013 | 173 | 120416 | Not_Computer
Z-1304-2013 | 64908 | No | KDD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Curity Staple Remover Kit; Product Code: 66701. |  |  | April 12, 2013 | 2013 | Covidien LLC | On March 26, 2013 Covidien received a recall notice from one of their suppliers CareFusion. CareFusion is recalling a variety of OTC drug products including the 10% Povidone Iodine SEPP Applicator which is utilized in specific convenience kits. | Error in labeling | US Customers were notified of the recall by letter sent by Federal Express on April 12, 2013. The letter requests that they immediately stop using any affected product, quarantine it, and return it to Covidien. Consignees were asked to return a verification form and return it to Covidien via fax. |  | 52363 | 52363 distributed | Worldwide distribution: USA (nationwide) and in the country of: Panama. | Terminated | June 02, 2015 | 781 | 70103 | Not_Computer
Z-1796-2013 | 65635 | No | JPA | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(k) | TEG 5000 Thromboelastograph Coagulation Analyzer |  |  | June 20, 2013 | 2013 | Haemonetics Corporation | Research Use Only (RUO) parameters were included in TEG 5000 User Manual. | Labeling False and Misleading | Haemonetics sent a Medical Device Correction letter dated June 20, 2013 to all affected customers. The letter identified the affected product, problem and action to be taken. The letter instructs customers to discard the current manuals and replace with revised manuals included with the Customer Notification Letter. A Consignee Response Form has been included with the letter, which is to be signed and returned for effectiveness purposes. Follow up will be made to customers who don't return the Response Form. For questions contact TEG Technical Support team at 1-800-438-2834. |  | 1590 | 1590 devices | US Nationwide | Terminated | June 04, 2015 | 714 | 1590 | Not_Computer
Z-0406-2014 | 66413 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes 2.4 mm LC-DCP Plate 6 Holes/51mm |  |  | July 30, 2013 | 2013 | Synthes (USA) Products LLC | A complaint was reported regarding a screw being found in a package labeled for 2.4 mm LC-DCP Plate. | Labeling mix-ups | Synthes sent an Urgent Notice: Medical Device Recall letter dated July 30, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to review their inventory, immediately remove the affected product and return to Credit/Returns, Synthes. For questions call 610-719-5450. |  | 13 | 13 | Worldwide Distribution - USA (nationwide) including IN, NY, WA, NE, MD, TX, AL, MI, and Internationally to Canada. | Terminated | August 17, 2015 | 748 | 13 | Not_Computer
Z-1341-2013 | 64943 | No | OTO | Surgical mesh. | General &amp; Plastic Surgery | Obstetrics/Gynecology | 510(k) | Rostorelle L, Polypropylene Mesh, Catalog number 501440, Manufacturer Coloplast A/S 3050 Denmark. |  |  | April 12, 2013 | 2013 | Coloplast Manufacturing US, LLC | Coloplast is recalling one lot of Restorelle L prosthesis because it is mislabeled. The incorrect Instructions For Use (IFU) pamphlet is included in the packaging. | Labeling mix-ups | Coloplast sent a letter to dated April 12, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The letter included an attachment with the correct IFU to replace the incorrect IFU. Customers with questions were instructed to call Customer Service at 800-258-3476.For questions regarding this recall call 612-302-4992. |  | 111 | 111 | Nationwide Distribution including AL, CA, CO, DC, IL, IN, KY, MD, MA, MO, NH, OH, RI, and VA. | Terminated | January 21, 2014 | 284 | 111 | Not_Computer
Z-0410-2014 | 66414 | No | OAR | Cement dispenser. | Orthopedic | Orthopedic | 510(K) Exempt | Vertebroplasty Needle Kit 10g Diamond Tip |  |  | July 30, 2013 | 2013 | Synthes USA HQ, Inc. | Complaints were received of inconsistencies in the expiration dates shown on the packaging of the Vertebroplasty Needle Kit 10g Diamond Tip device. | Incorrect or no expiration date | Synthes sent an Urgent Notice: Medical Device Notification letter dated October 8, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to returny any of the affect devices, call Synthes at 1-800-479-6329 to obtain a Return Authorization Number, complete the Verification Section indicating the affected product has been located, also indicate the number of devices found and note the Return Authorization Number. Return the Verification Form to:Credit/ReturnsSynthes1101 Synthes AvenueMonument, CO 80132Customers were also instructed to return the Verification Section even if they don't have the affected product. Customers with questions were instructed to call 610-719-5450 or contact their Synthes Sales Consultant.For questions regarding this recall call 610-719-5000. |  | 40 | 40 | Nationwide Distribution including IN, MT, HI, and FL. | Terminated | April 16, 2015 | 625 | 40 | Not_Computer
Z-1810-2013 | 65600 | syringe, irrigating (non dental) | KYZ | Irrigating syringe. | General Hospital | General Hospital | 510(K) Exempt | Medtronic Preimplantation Test Kit, Catalog No. 21047. The kit is indicated for use with Medtronic CSF-Flow Control Valves and Delta Valves, when a simple preimplantation test is desired to verify that the valve conforms to labeled product specifications. |  |  | June 28, 2013 | 2013 | Medtronic Neurosurgery | Medtronic Neurosurgery is recalling the Medtronic Preimplantation Test Kit because the outer carton had expiration dates that were incorrect. The date on the outer carton is later than the expiration date for one or more of the individually packaged sterile components in the kit. | Incorrect or no expiration date | An "Urgent Field Safety Notice" recall letter dated 6/28/13 was sent to customers who purchased the Medtronic Preimplantation Test Kit, catalog number 21047, to inform them of the recall due to the incorrect outer carton expiration dates. The recall letter informed the customers of the problems and actions identified. Customers were instructed to contact the Quality Department at (805) 571-8725 for questions. |  | 915 | 915 units | Nationwide Distribution in the US. | Terminated | September 23, 2013 | 87 | 915 | Not_Computer
Z-1213-2013 | 64859 | No | OIY | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | 3DKnee e+ Tibial Insert. The insert is used along with the other components that comprise the 3DKnee System. |  |  | April 01, 2013 | 2013 | Encore Medical, Lp | U.S. Agent contacted Director Commercial Logistics, notifying him of incorrect color coded labels on two packages of e+tibial inserts. There were 8 e+ tibial inserts of the same lot number in Finished Goods as the two from the agency. These devices were inspected and all were incorrectly labeled. | Labeling mix-ups | The firm sent their US customer a "Urgent Field Safety Notice" letter on April 1, 2013. The letter explained the product problem and provided recommended actions as well as customer contact numbers. |  | 5 | 5 | US Distribution to the state of California and International Distribution to the country of Venezuela. | Terminated | July 01, 2013 | 91 | 5 | Not_Computer
Z-1808-2013 | 65654 | No | NJE |  |  | Neurology | HDE - Humanitarian Device Exemption | Stryker brand Neuroform 3 Microdelivery Stent System, 4.5 mm X 20 mm, REF Catalog Number: E345020, Model/Product Number: M003450200, Product is manufactured by Boston Scientific Corporation, Fremont, CA and distributed by Stryker Neurovascular, Fremont, CA |  |  | June 24, 2013 | 2013 | Stryker Neurovascular | The device is labeled incorrectly, in that, the device Model Number on the outer carton and inner pouch labels did not match, although both referenced the same lot number. The Model Number on the outer box states UPN M003E345020 (4.5mm x 20 mm Neuroform 3 Stent) while the pouch label states UPN M003E3450300 (4.5mm x 30 mm Neuroform 3 Stent). | Packaging process control | Stryker Neurovascular sent an Urgent Field Safety Notice on June 28, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately check their internal inventory and quarantine all subject devices pending return to Stryker. Circulate the Field Safety Notice internally to all interested/affected parties. Maintain awareness of this notice internally until all requred actions have been completed within their facility. Inform Stryker if any of the subject devices were distributed to other organization. Complete the attached customer response form and return to their Stryker Representataive.For questions regarding this recall call 510-413-2500. |  | 23 | 23 units | International Distribution including Austria, China, and Germany. | Terminated | September 30, 2013 | 98 | 23 | Not_Computer
Z-0382-2014 | 66569 | No | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | Tri-Nex Healing Abutment, 0.50 mm x 5 mm height, Catalog No. HA-L-50-5. |  |  | September 13, 2013 | 2013 | Southern Implants, Inc | Southern Implants is recalling the Tri-Nex healing abutment because the package was incorrectly labeled and packaged as the Tri-Nex healing abutment, 05.0 mm x 5 mm height (HA-L-50-5). The package contains the Tri-Nex healing abutment 05.0 mm x 6.0 mm Flare x 5 mm Height (HA-L-50W-5). | Labeling design | Southern Implants, Inc. sent an Urgent Product Action letter dated September 11, 2013, to all affected customers. The letter identiied the product, the problem, and the action to be taken by the customer. Customers were instructed to fill out the customer response form and fax/email the completed form to Customer Support at (866) 903-9272 or customersupport@keystonedental.com.For questions regarding this recall call 866-903-9272. |  | 197 | 197 units | Worldwide Distribution - USA (nationwide) and Internationally to Canada, Guatemala, Taiwan. | Terminated | January 21, 2014 | 130 | 197 | Not_Computer
Z-0384-2014 | 66415 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Synthes Matrix Threaded Persuader |  |  | July 31, 2013 | 2013 | Synthes USA HQ, Inc. | Synthes became aware that during a procedure using the Matrix Threaded Persuader, the Reduction Insert malfunctioned causing pieces to break off. Once there is a breakage to the Reduction Insert it will not allow the instrument to function as required. | Labeling design | The firm, Synthes, sent an "URGENT NOTICE: MEDICAL DEVICE RECALL" letter dated July 31, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to review your inventory and immediately remove the affected lots from stock; If you do have the affected devices, contact Synthes at 1-800-479-6329 to obtain the RA# and return the product to with completed form to Credit/Returns, Synthes, 1101 Synthes Avenue, Monument, CO 80132; complete and sign the Return Receipt Requested form via fax at 1-888-731-7955 or email to: FieldAction@synthes.com, even if you do not have the affected product.If you have questions, please call 610-719-5450. |  | 634 | 634 | Worldwide distribution: US (nationwide) and countries of Canada and Switzerland. | Terminated | August 31, 2015 | 761 | 634 | Not_Computer
Z-1214-2013 | 64820 | Yes | KPO | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | NxStage PureFlow-B Solution, Model/Catalog Number: RFP-401, packaged in a case |  |  | March 18, 2013 | 2013 | NxStage Medical, Inc. | Product may be mislabeled. | Labeling mix-ups | NxStage sent a letter dated March 18, 2013, via email to all consignees instructing them to check their entire inventory and to segregate and quarantine lot F023444 pending further instruction from NxStage.On 4/2/13 NxStage sent a letter via email to all consignees requesting a return of all RFP-400 and RFP-401 labeled with lot F023444.. Consignees were instructed to complete the recall reply form and provide a list of customers (name address, phone number) and fax them to NXStage Medical, Inc.Customers with questions were instructed to call 1-866-697-8243.For questions regarding this recall call 1-866-697-8243. |  | 1480 | 1480 cases (2 bags per case) | Nationwide Distribution including IL, SC, TX, VA, FL, AR, and NY. | Terminated | September 17, 2014 | 548 | 1480 | Not_Computer
Z-1809-2013 | 65391 | No | LXM |  |  | Physical Medicine | 510(k) | Advanced Orthogonal Percussion Adjusting Instrument |  |  | March 06, 2013 | 2013 | Advanced Orthogonal Equipment, Incorporated | Unapproved medical device | No Marketing Application | The firm, Advanced Orthogonal Equipment, sent a "CORRECTION NOTICE" to its consignees/customers. The notice described the product, problem and actions to be taken. The customers were instructed that " none of these devices be used pending final clearance by FDA".Questions and concerns regarding this notice should be directed to the President of Advanced Orthogonal Equipment by phone at: 727-528-8700 or email: Atlasortho@aol.com. |  | 54 | 54 | Nationwide distribution: US including states of: AL, CA, CO, FL, GA, ID, KY, MD, NC, NJ, NY, OH, PA, TN, VA, and WA. | Open, Classified |  |  | 54 | Not_Computer
Z-1395-2013 | 63204 | No | OAR | Cement dispenser. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker Orthobiologics Aliquot Delivery System Manual Surgical Instrument - Aliquot Reamer, 8", Part Number 2110-0525 |  |  | November 10, 2012 | 2012 | Orthovita, Inc., dBA Stryker Orthobiologics. | There have been a few reported cases of the catheter breaking when removing the Aliquot catheter from the needle during vertebroplasty procedures. This has been attributed to bending of the catheter during the retraction from the needle. It was determined to be use errors. The correction is limited to a revision of the Directions for Use. | Labeling design | Stryker sent an Urgent Medical Device Correction letter dated November 10, 2012 to all affected customers. The letter identified the affected product, problem issue, potential hazards and actions to be taken. Customers were instructed to immediately review this notification and the attached IFU. Customers were asked to complete and sign the enclosed Business Reply Form and fax a copy to 1-484-323-8803. For questions contact Stryker Orthobiologics. |  | 5 | 5 | USA Nationwide Distribution in the states of AZ, CA, CO, CT, DE, FL, GA, HI, IL, IN, KS, KY, LA, MD, ME, MI, MO, NC, ND, NH, NJ, NY, OH, OK, PA, RI, SC, TN, TX, UT, VA, WA, and WI. | Terminated | October 29, 2014 | 718 | 1413 | Not_Computer
Z-1057-2013 | 64593 | No | DYG | Flow-directed catheter. | Cardiovascular | Cardiovascular | 510(k) | Arrow/Teleflex Medical Berman Angiographic Balloon Catheter and Reverse Berman Angiographic Balloon Catheter. |  |  | March 04, 2013 | 2013 | Arrow International Inc | The product specifications sheet, part of the Instructions for Use (IFU), for the specified product numbers, reference Renografin-76 contrast medium (viscosity of 8.4 centipoise), which is no longer available. Firm is in the process of updating the labeling for these products. In the meantime, when choosing an alternative contrast medium, health care providers are instructed to select a contrast | Labeling Change Control | Arrow sent an Urgent Medical Device Customer Notification letter dated March 4, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Arrow is in the process of updating the labeling; in the meantime they wanted to instruct all customers to select a contrast medium with a viscosity of 8.4 centipoise or lower for use with the pressure and flow settings specified in the current IFU. In addition, they are placing an insert card in kits for the specified product codes alerting them that Renografin-76 is no longer available and to select a contract medium with a viscosity of 8/4 centipose or lower for use with the pressure and flow setting specified in the current IFU.Arrow remains committed to providing high quality, safe and effective products. If you require additional information or clarification regarding this matter, please contact your Sales Representative or Arrow's Customer Service Department at (866) 396-2111. |  | 99627 | 99627ea | Worldwide Distribution - USA including Puerto Rico and the states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI and WV. and the countries of Australia, Dominican Republic, Uruguay, Ecuador, Chile, Taiwan, Pakistan, Colombia, Argentina, Korea Republic, New Zealand, Guatemala, Thailand, Peru, Bahamas, Philippines, Trinidad and Tobago, Japan, Ireland, India, Singapore, China, Canada and Mexico. | Terminated | December 21, 2013 | 292 | 99627 | Not_Computer
Z-1821-2013 | 65668 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Synthes Synapse SD Drill Guide, Part Number SD389.477, Lot Number 6913900 |  |  | October 12, 2012 | 2012 | Synthes USA HQ, Inc. | A complaint was received from the field, that the drill guide was etched incorrectly, 12mm etch instead of 18mm drill guide. The firm identified that this constituted a field action and retrospectively reported this action. | Labeling Change Control | The firm, Synthes USA, sent an email to its customer. The email identified the product, problem and the actions to be taken. The customer was instructed to do the following: do not use this product; do not remove any product from your account or return any product without further instructions. If you have identified that you have the product, please have the product located in an area that ensures it will not be used. Further instructions will be sent out with regards to any action that needs to be taken next.Please call Synthes if you have any questions at 610-719-5364. |  | 2 | 2 | US distribution: PA only.. | Terminated | June 11, 2015 | 972 | 2 | Not_Computer
Z-0392-2014 | 66389 | No | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Baxter, VIAL-MATE Adaptor, 2B8071, Rx Only, STERILE R, Baxter Healthcare Corporation, Deerfield, IL 30015 USA |  |  | September 25, 2013 | 2013 | Baxter Healthcare Corp. | Incomplete foil seal on one lot of sterile product. | Packaging process control | Baxter Healthcare sent an Urgent Product Recall letter dated September 25, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to locate and remove all affected product from their facility, contact Baxter's Center for Service at 1-888-229-0001 to arrange for return and credit. Complete the attached Customer Reply Form and return it to Baxter by fax or email. Customers with questions were instructed to call 1-800-422-9837. For questions regarding this recall call 224-948-4770. |  | 38400 | 38,400 units | Nationwide Distribution including WV, KY, IA, AR, OH, IL, IN, MI, KS, TN, WI, MN, MO. | Terminated | July 22, 2014 | 300 | 38400 | Not_Computer
Z-1336-2013 | 64468 | No | LPL | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic | 510(k) | ACUVUE ADVANCE Brand Contact Lenses with HYDRACLEAR, Soft, disposable contact lenses. |  |  | December 21, 2012 | 2012 | Johnson &amp; Johnson Vision Care, Inc. | Limited number of individual contact lens packages may not have been completely sealed. | Packaging | The firm, Johnson &amp; Johnson, Vision Care, Inc. sent a letter dated January 9, 2013 to all their consignees stating the following actions:" Please contact your patients that may have received any of the affected product and instruct them to discontinue use immediately and return product to you for replacement " Please return the enclosed Business Reply Card (BRCl confirming for our records that you have received this notice, and indicating whether or not you have in stock, or have distributed to patients,any product from the lots noted above. " Please return any remaining affected product in your possession using the enclosed pre-paid mailer. " Pass this notice on to anyone in your practice that needs to be informed " Customer service will arrange for the return and replacement of any affected product. If you have any questions and/or if replacement lenses are needed, please contact customer Service (1-800-87 4-5278). |  | 3546 | 3546 lenses | Worldwide distribution: USA (nationwide) including states of: AL, AR, AZ, CA, CO, CT, DC, FL, GA, IL, IN, KY, LA, MA, MD, MI, MO, NC, NJ, NY, OH, OK, PA, SC, TN, TX, UT, VA, WA, WI and Puerto Rico; and countries of: Austria, Belgium, Canada, Costa Rica, Denmark, France, Germany, Hungary, India, Ireland, Israel, Italy, Korea, Japan, Luxembourg, Netherlands, Portugal, Norway, South Africa, Slovenia, Spain, Sweden, and United Kingdom. | Terminated | July 05, 2013 | 196 | 39876 | Not_Computer
Z-1331-2013 | 65078 | No | FYE | Surgical apparel. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | DePuy LifeLiner Stick &amp; Cut Resistant Gloves Rt, Lg |  |  | May 08, 2013 | 2013 | DePuy Orthopaedics, Inc. | LifeLiner Stick and Cut Resistant Gloves are being recalled because cases and individual glove packages were mispackaged, with right handed, large sized glove packaging containing left handed, large sized gloves. | Labeling mix-ups | DePuy Synthes Joint Reconstruction, a division of DePuy Orthopaedics, Inc.,sent a " Urgent Information Medical Device Recall Notice" dated May 8, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Please cease further distribution or use of the affected gloves immediately. Our records indicate your hospital purchased the affected lot. The purpose of this communication is to inform you of this recall and requestacknowledgement of receipt of this letter by signing and returning the Hospital/User Facility Reconciliation Form within 10 business days of receiving this notice.Response cards should be faxed to (574) 371-4964. Please contact your DePuy Synthes Joint Reconstruction Sales Representative for assistance returning any unused affected gloves.For product related questions, please contact your local DePuy Synthes Joint Reconstruction Sales Consultant. For clinical questions from surgeons, please contact DePuy Synthes Joint Reconstructions Scientific Information Office at 1-888-554-2482 (M-F; 8 a.m.- 5 p.m. EDT). For questions about device recall information provided, please call (574)371-4917 (M-F; 8 a.m. - 5 p.m. EDT). |  | 180 | 180 gloves (120 gloves distributed in the US) | Worldwide Distribution - US (nationwide) including the states of IL, OH, NC, SC, MN, NV, TN and WA., and the countries of UK, Netherlands and Canada. | Terminated | January 14, 2014 | 251 | 180 | Not_Computer
Z-1054-2013 | 64704 | No | KFX | External assembled lower limb prosthesis. | Physical Medicine | Physical Medicine | 510(K) Exempt | LeMaitre Stent Guide: 1102-00 (100 strips); 1102-50 (50 strips); 1102-20 (20 strips). |  |  | March 18, 2013 | 2013 | LeMaitre Vascular, Inc. | Devices were not sealed correctly during the manufacturing process, and the sterility of these products have been compromised. | Packaging process control | LeMaitre sent an "URGENT:VASCUTAPE RADIOPAQUE TAPE DEVICE FIELD SAFETY NOTICE" dated March 20, 2013 to affected customers. The recall letter contains a form that requests customers to return to LeMaitre Vascular, Inc. as a record of notification and reconciliation. The letter provides instructions on how the customer can inspect the defective tape seals and/or return the affected products for replacement/credit. |  | 6970 | 6970 | Worldwide Distribution-USA (nationwide) and the countries of Canada, EU, and Asia. | Terminated | December 03, 2014 | 625 | 20530 | Not_Computer
Z-0614-2014 | 67128 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | Catheter, Continuous Flush |  |  | December 06, 2013 | 2013 | Merit Medical Systems, Inc. | Merit Medical Systems, Inc. is conducting a recall for 3 lots of Maestro Microcatheters with a graphicdiscrepancy on the carton/box label end panel. The catheter graphic on the front of the cartonbox (correct) and the graphic on the end panel (incorrect) do not match. | Packaging process control | Merit Medical Systems sent an Urgent Product Recall Notice dated December 6, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately discontinue use of the Maestro Microcatheters identified and isolate them in a quarantined area until the product is over-labeled by their Merit Sales Representative Customers were instructed to ensure that all personnel to whom the devices were distributed were made aware of the field action. Customers with questions were instructed to contact their sales representative or call 1-801-826-4062.For questions regarding this recall call 801-253-1600. |  | 59 | 59 | Worldwide Distribution - USA (nationwide) and Internationally to Canada. | Terminated | January 27, 2014 | 52 | 59 | Not_Computer
Z-0403-2014 | 66815 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | STERRAD NX, Product Code: 10033 |  |  | November 11, 2013 | 2013 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) is recalling the STERRAD NX and STERRAD 100NX sterilization systems because ASP has determined that the bronchoscopes may contain a polyurethane material in the suction channel location that is not listed in the current claims for interior channels (lumens) for processing in STERRAD systems. | Labeling design | Advanced Sterilization Products sent an Urgent Medical Device Field Safety Notificaiton letter on November 11, 2013, to all affected customers. The letter informed customers of the problems identified and the actions to be taken. Customers were instructed to provide the notice to anyone in their facility that needs to be informed and post this notice where appropriate. Customers were instructed to contact ASP at 1-888-783-7723 for any suspected problems. ASP is collaborating with KARL STORZ Endoskope to determine the sterilization impact of the suction port lumen material. Customers are instructed to contact KARL STORZ Endoskope Technical Support at (800)421-0837, ext 5350 for alternative reprocessing instructions.For questions regarding this recall call 949-453-6400. |  | 4381 | Total 4,381 units | Nationwide Distribution | Terminated | January 22, 2014 | 72 | 4381 | Not_Computer
Z-0390-2014 | 66666 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | 3.7mm and 5.0mm Dynamic Locking Screw (DLS) |  |  | July 22, 2013 | 2013 | Synthes USA HQ, Inc. | All lots of 3.7mm and 5.0mm Dynamic Locking Screw (DLS) were recalled due to complaints of breakages at the distal tip of the pin identified after successful healing and during planned removal of the device. | Labeling False and Misleading | Synthes sent an Urgent Notice: Medical Device Recall letter to Sales Reps and End Users on July 22, 2013. The letter identified the affect product, problem and actions to be taken. For questions call the Recall center at 610-719-5450. |  | 40248 | 40248 | Worldwide Distribution - USA (nationwide) and internationally to Canada. | Open, Classified |  |  | 40248 | Not_Computer
Z-1053-2013 | 64528 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | LEEP REDIKIT Product No. 6060 |  |  | February 26, 2013 | 2013 | Cooper Surgical, Inc. | LEEP RediKit Products (for use in Loop Electrosurgical Excision Procedures) have Xylocaine (Lidocaine) HCL (2%) listed on the tray lid label but this is not included in the contents of the kit, which contained Polocaine (Mepivacaine) HCL (3%) instead (which was not listed on the tray label). | Labeling mix-ups | CooperSurgical sent a Recall Notification dated February 26, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.CSI intends to track and quarantine returned product in a segregated location at its manufacturing facility to ensure its separation from acceptable product and prevent its misuse. CSI intends to render used product useless by physical destruction and to rework unused product by correcting the labeling for redistribution.If you have any one of these affected lot numbers in your current inventory, please contact our Customer Service Department at 800-243-2974, or simply complete the attached form and fax to 800-262-0105. Arrangements will be made to replace product in your possession at CooperSurgical's expense. Further instruction and disposition of product will be communicated to you at that time.CooperSurgical deeply regrets the inconvenience that this problem has caused. |  | 4395 | 4395 kits | Nationwide Distribution including Puerto Rico | Terminated | June 12, 2014 | 471 | 4395 | Not_Computer
Z-0755-2013 | 63934 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Ortho Clinical Diagnostics VITROSÂż Chemistry Products Vapor Adsorption Cartridge REF/CAT No. 680 0100. VITROSÂż Vapor Adsorption Cartridges are used to remove vapor contaminants that may potentially interfere with Signal Reagent used with MicroWell Assays on the VITROSÂż 3600, VITROSÂż 5600, and VITROSÂż ECi/ECiQ Systems. |  |  | December 13, 2012 | 2012 | Ortho-Clinical Diagnostics | Some VITROSÂż Chemistry Products Vapor Adsorption Cartridge CAT No. 6800100 cartons incorrectly contain a Particulate Cartridge Part No. J19612. In some instances, the Particulate Cartridge has been inadvertently installed on the VITROSÂż System instead of the intended VITROSÂż Vapor Adsorption Cartridge which may bias results. | Packaging process control | Ortho-Clinical Diagnostics sent URGENT PRODUCT CORRECTION Letters (dated 12/13/12) along with Q&A; and Confirmation of Receipt Forms to distributors and consignees on 12/13/12. They were advised of the issue and instructed to discard their current inventory of VITROSÂż Vapor Adsorption Cartridges. On 12/13/12, foreign affiliates were informed by e-mail of the issue and instructed to notify their consignees of the issue and actions. |  | 5237 | Domestic: 15,891; Foreign: 5237 | Worldwide Distribution, including Nationwide (USA) and the countries of AU, GM, FR, IT, SP, BR, BM, AR, CA, JA, MX, UK, CH, CL, CO, IN, EC, PE, PY, PA, PT, NZ, SG, VE and UY. | Open, Classified |  |  | 5237 | Not_Computer
Z-0608-2014 | 66968 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | CPT Hip System, Parts 00-8114-002-10 and 00-8011-004-00 |  |  | December 10, 2013 | 2013 | Zimmer, Inc. | Zimmer is initiating a lot specific recall of the CPT Hip Stem, Versys Build-Up Block, and Trabecular Metal Reverse Glenosphere medical devices due to reports of the products being packaged in the incorrect outer carton. Product labels and patient labels are correct. | Labeling mix-ups | Zimmer sent an Urgent Medical Device Recall letter dated December 10, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the notification and ensure affected personnel were aware of the contents. If affected product was found, it should be quarantines and their Zimmer representative notified. The Zimmer sales representative will remove the recalled product from their facility. Customers with questions were instructed to call 1-877-946-2761. . |  | 36 | 36 | Worldwide Distribution -USA including OH, FL, WI, MI NJ, SC, CA, NC, PA, OR, WA, TN, KY, OK, KS, AZ, TX, CO, NV and Internationally to Germany.There have been 71 distribution events. Note that this product is lot controlled and may be distributed multiple times if it was returned and returned to the shelf. 7 units have been returned, including three through the complaint process, which leaves 64 affected units in the field subject to removal. | Terminated | September 11, 2014 | 275 | 64 | Not_Computer
Z-0404-2014 | 66581 | No | JDS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes 14.5 mm Outer Protection Sleeve for Suprapatellar |  |  | October 14, 2013 | 2013 | Synthes (USA) Products LLC | Complaints were received reporting that a product packaged as a 14.5 mm Outer Protection Sleeve for Suprapatellar contained a 12.0 mm Outer Protection Sleeve for Suprapatellar. | Labeling mix-ups | Synthes sent an Urgent Notice: Medical Device Recall letter dated October 14, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to call Synthes at 1-800-479-6329 if they have any of the affected product to obtain a Return Authorization Number, complete the Verification Section at the end of the letter and return the Verification Form with the affected product to:Credit/ReturnsSynthes1101 Synthes AvenueMonument, CO 80132If customers do not have the affected product they should complete the attached Verification Section and return all documents to Synthes by:Fax: 866-229-7778Scan/email: FieldAction@synthes.comCustomers with questions were instructed to call 610-719-5450.For questions regarding this recall call 610-719-5000. |  | 55 | 55 | Nationwide Distribution | Terminated | August 18, 2015 | 673 | 55 | Not_Computer
Z-1344-2013 | 49026 | No | JHI | Human chorionic gonadotropin (HCG) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Cardinal Health hCG Cassette Rapid Test, Part Number: B1077-22 (Mfg PN: 4582105026) |  |  | January 21, 2013 | 2013 | Alere San Diego, Inc. | Alere San Diego Inc. is recalling the Cardinal Health hCG Cassette Rapid Test because some of them were incorrectly packaged into Cardinal Health hCG Combo Rapid test kit boxes. The kit box incorrectly indicates the product may be used for both urine and serum sample types. | Labeling mix-ups | The firm, Alere, sent an "URGENT MEDICAL DEVICE RECALL" letter dated January 21, 2013 to all their customers. The letter describes the product, problems identified and the actions to be taken. Customers are instructed to discontinue use and/or sale of the product; immediately discard all units of the the product in accordance with your local regulations; share this information with your laboratory staff and retain this letter for your records; if you have forwarded the product to another laboratory, provide a copy of this letter to them and include the name, address and a contact for the laboratory on the enclosed Customer Verification Form; complete and return the Customer Verification Form within 10 days via fax to: Technical Service at 858-805-8457 or e-mail to responses.ts@alere.com. Customers with any questions are instructed to contact Alere Technical Services at (866) 216-0094 or by email at lateral.flow.support@alere.com. |  | 2167 | 2,167 kits | Nationwide distribution: USA including states of: AL, AK, AR, AZ, CA, CO, CT, DE, HI, IA, ID, IL, IN, FL, GA, KY, KS, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, UT, TN, TX, VA, WA, WI, WV AND WY. | Terminated | August 28, 2013 | 219 | 2167 | Not_Computer
Z-0379-2014 | 66663 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Product Brand Name(s): Alere Cholestech LDXÂż Lipid Profile"GLU Test Cassette, Model 10-991. |  |  | October 23, 2013 | 2013 | Alere San Diego, Inc. | Alere initiated a recall of Cholestech LDX Lipid Profile-GLU test cassettes, Catalog Number 10-991, Lot # 320167, becuase an incorrect lot number was printed on the outer 10 pack kit box label. The incorrect Lot # number printed was 321067. | Labeling Change Control | Alere initiated this recall by sending out notification letters to customers via fax, email, or direct mail. The letter titled "URGENT MEDICAL DEVICE NOTICE", dated October 23, 2013 is accompanied by a response form/Verification form. The notification informed customers of the reason for recall, product description with codes, required actions by the customers, and contact information. |  | 3621 | 3,621 kits | Worldwide Distribution: USA (nationwide) and Internationally to Canada. | Terminated | September 25, 2014 | 337 | 3621 | Not_Computer
Z-1778-2013 | 65488 | No | FTC | Ultraviolet lamp for dermatologic disorders. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Dermalight 80 Model Number UVB-080, one unit per box. |  |  | April 09, 2013 | 2013 | National Biological Corp | The incorrect manual may have been packaged with this product. UVB-080 devices may have included the manual for a UVA-083 device. | Labeling mix-ups | The firm, National Biological, sent an "VOLUNTARY PRODUCT RECALL NOTIFICATION" letter dated May 10, 2013 to its consignees/customers by UPS or other delivery service with a signature required. The letter described the product, problem and actions to be taken. Consignees/customers who do not respond to the letter will be contacted via telephone. The customers were instructed to check within your household to determine if you have any of the listed product; immediately destroy the incorrect User Manual, it is not necessary to return the product; record the Customer Name and or the Dermalight-80 serial number and complete and return the enclosed Verification Form, even if you do not have any User Manuals to destroy.Please contact your National Biological Corporation sales representative if you have any questions regarding this recall, any of our products, or would like assistance with the recall call 216-831-0600 or 800-338-5045 email: infor@PhototherapyExperts.com. |  | 124 | 124 units | Worldwide distribution: US (nationwide) including states of: AL, CA, CO, CT, FL, GA, IA, ID, IL, LA, MA, MN, MO, MS, NE, NJ, NY, OH, PA, SC, TN, TX, VA, WA, and WI; and countries of: Canada, Brazil, Malaysia, and Mexico. | Terminated | February 24, 2014 | 321 | 124 | Not_Computer
Z-1181-2013 | 64833 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | VitreaÂż CT Brain Perfusion made available in various configurations under the names Vitrea, Vitrea fX, VitreaAdvanced and Vitrea Enterprise Suite. |  |  | March 15, 2013 | 2013 | Vital Images, Inc. | Vital Images, Inc. (Vital Images) received one report about potential user confusion when viewing a brain perfusion scan on Vitrea CT Brain Perfusion 2D when the scan was performed using a wide coverage shuttle acquisition technique. This may result in patient treatment delays, misdiagnosis or mistreatment. | Labeling design | Consignees were sent on 3/15/2013 a VITAL Images, Inc., "Urgent Field Safety Notice" dated March 15, 2013. The letter described the problem and the product involved in the recall. The letter listed the limitations of the product and the steps to take. Consignees were requested to return the Product Recall Effectiveness Check form. Questions should be directed to Vital Images Customer Support at 1-800-208-3005 or via e-mail, support@vitalimages.com. |  | 1978 | 1,978 | Worldwide Distribution -- USA, excluding RI and SD and including DC and PR, and the countries of ALGERIA, ANGOLA, ARGENTINA, AUSTRALIA, AUSTRIA, BELGIUM, BELARUS, BOLIVIA, BRAZIL, BULGARIA, CAMBODIA, CANADA, CHILE, CHINA, COLOMBIA, CYPRUS, CZECH REPUBLIC, DENMARK, EGYPT, FINLAND, FRANCE, GEORGIA, GERMANY, CHANA, GREECE, GUAM, HONDURAS, HONG KONG, ICELAND, INDIA, INDONESIA, IRAQ, IRELAND, ISRAEL, ITALY, JAPAN,JORDAN, KOREA, KUWAIT, LEBANON, LITHUANIA, MALAYSIA, MEXICO, MOSCOW, MYANMAR, NAMIBIA, NETHERLANDS, NEW ZEALAND, NIGERIA, NORTHERN IRELAND, NORWAY, OMAN, PAKISTAN, PANAMA, PERU, PHILIPPINES, POLAND, PORTUGAL, QATAR, ROMANIA, RUSSIA, SAUDI ARABIA, SERBIA, SERBIA AND MONTENEGRO, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SRI LANKA, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TUNISIA, TURKEY, UKRAINE, UNITED ARAB EMIRATES, UNITED KINGDOM, UZBEKISTAN, VENEZUELA, VIETNAM, and YEMEN. | Terminated | April 08, 2015 | 754 | 1978 | Not_Computer
Z-1045-2013 | 64569 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | Zimmer NexPosure, retractor portals of the NexPosure MIS Access System, Part Numbers or REF numbers: N1830R80-100, N1830R80-090Âż, N1830R80-080Âż, N1830R80-070Âż, N1830R80-060, N1830R50-100, N1830R50-090, N1830R50-080, Âż N1830R50-070, N1830R50-060, N1830L80-100, N1830L80-090, N1830L80-080, N1830L80-070, N1830L80-060, N1830L50-100, N1830L50-090, N1830L50-080, N1830L50-070, N1830L50-060, Âż N182450-085, N182450-075, N182450-065, N182450-055, N182450-045.Âż Retractor access system that allows placement of posterior spinal fixation implants. |  |  | February 08, 2013 | 2013 | Zimmer, Inc. | Possibility that the inner pouch may be compromised. The product is packaged in two pouches - a sterile inner pouch within an external pouch. NexPosure Retractors contained in a breached inner pouch have an increased risk of contamination (i.e.; loss of product sterility) if proper aseptic technique is not followed when transferring the inner package into the sterile field. Patient infection may o | Package design/selection | Consignees were sent on 2/8/13 a Zimmer Spine "Urgent Medical Device Recall" letter dated February 08, 2013. The letter was sent to Surgeons and Facilities using the NexPosure MIS Access System. The letter described the product involved in the recall and the problem. Advised consignees to discontinue the use of the product and to contact their sales representative for removal.A consignee letter dated February 8, 2013 was also sent to Distributors, Sales Representatives, and Distribution Operation Managers Distributing the NexPosureÂż MIS Access System. The letter contained instructions if they had the affected product or not and the steps to follow. |  | 58 | 58 | Nationwide Distribution including the states of AZ, CO, DE, FL, ID, LMA, MI, NH, NJ, NM, PA, RI, TX, and VA. | Terminated | December 09, 2014 | 669 | 58 | Not_Computer
Z-1093-2013 | 64776 | No | LRR |  |  | General &amp; Plastic Surgery | Enforcement Discretion | LHB, Perishable Items Pack Type IV. Kit number: 2114016-100. |  |  | March 22, 2013 | 2013 | Lighthouse For The Blind | The product contains an iodine 10% Ampule (10% Povidone Iodine swab) which is being recalled by the firm's distributor. The labeling of the iodine 10% Ampule labels state that the contents of the package are sterile and does not clarify that this sterility claim is only applicable to the applicator. The antiseptic solution that is applied to the skin is a cutaneous disinfectant, and is not sterili | Labeling False and Misleading | The firm sent "URGENT DEVICE RECALL" letters dated March 22, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at 800-542-3697 for questions regarding this notice. |  | 425 | 425 kits | Nationwide Distribution including DC and the states of AK, AL, AR, AZ, CA, CO, DE, FL, GA, ID, IL, IN, KS, KY, MI, MD, MN, MO, MT, NC, NJ, NM, NY, NV, OH, OK, PA, SC, TN, TX, UT, VA, WA and WY. | Terminated | November 25, 2013 | 248 | 108939 | Not_Computer
Z-1922-2013 | 65781 | No | FPD | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Nestle Nutrition, Compat, Enteral Feeding Tube, sterile, 8 French (REF 08310400), 10 French (REF 08310600), and 12 French (REF 08310500). |  |  | July 24, 2013 | 2013 | Nestle HealthCare Nutrition | Nestle HealthCare Nutrition, Inc. is initiating a recall on some of Compat NG Tubes because they have been delivered in an insufficiently sealed pouch, thus compromising the sterility claimed on the label. | Packaging process control | Nestle Health Science sent a "Medical Device Recall" letter dated July 24, 2013. The letter was addressed to Risk Manger. The letter described the product involved, problem and actions to be taken. Customers were advised that products currently in use should not be returned. Customers were requested to examine their inventory, quarantine the product, carry the recall to the user facility level and to complete and return the Acknowledgment / Response Form.For questions call Nestle HealthCare Nutrition Customer Service at 1-877-463-7853. |  | 37310 | 37,310 | US Nationwide Distribution - including the states of: AK, AZ, CA, CO, CT, FL, GA, HI, IA, IL, IN, KS, KY, MA, MD, MI, MN, MO, NC, NE, NJ, NM, NV, NY, OH, OK, PA, RI, SC, TN, TX, UT, VA, WA. | Terminated | June 13, 2014 | 324 | 37310 | Not_Computer
Z-0577-2014 | 66935 | [6 empty] | ProPlan CMF Guide(s) Model(s) |  |  | November 08, 2013 | 2013 | Materialise USA LLC | Patient Specific Guide contained the incorrect case report. The case report contains patient specific surgical plan information. | Labeling Change Control | Customer/ ordering physician was made aware of the issue11/08/2013. Revised accurate case report was immediately provided. |  | 1 | 1 | Distributed in Canada. | Terminated | January 07, 2014 | 60 | 1 | Not_Computer
Z-1748-2013 | 65303 | No | OOY |  |  | Anesthesiology | PMA | Alair Bronchial Thermoplasty Catheter, Bronchial Thermoplasty System; Material Number: M005ATS25010, Catalog Number: ATS 2-5. Product is manufactured and distributed by Boston Scientific Corporation (BSC), Sunnyvale, CA |  |  | May 28, 2013 | 2013 | Boston Scientific Corporation | Boston Scientific has discovered an inconsistent "Use By" expiration date on a single lot/batch of Alair Bronchial Thermoplasty Catheters. The expiration date indicated on the inner tray, 2012-04, is incorrect. The correct expiration date is 2014-04 as indicated on the outer box label. | Error in labeling | The firm, Boston Scientific, sent a letter dated May 28, 2013, via express mail to all its direct consignees. The letter described the product, problem and actions to be taken. The customers were informed that you can continue to use the device until the expiration date on the outer box lablel, 2014-04; No product is being recalled and the you are not required to return product to Boston Scientific. The customers were instructed to read the enclosed information carefully and post the information in a conspicuous location near the affected product and/or take such other steps, consistent with your internal policies and procedures, to ensure that the information is easily accessible to all users of this device; and to complete and return the Customer Acknowledgment Form even if you no longer have any units from the affected lot/batch via email: MapleGroveFieldAc@bsci.com or Fax to: Field Action Center - 1866-213-1806. If you have any questions, call 763-494-1133 or email: MapleGroveFieldAc@bsci.com . |  | 35 | 35 units | Worldwide distribution: US (nationwide) including states of: CA, FL, IA, ID, MI, OK, TA and WA; and countries of: Canada and Thailand. | Terminated | August 28, 2013 | 92 | 35 | Not_Computer
Z-1888-2013 | 65819 | No | MGO | Hydrophilic wound dressing. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | CVS Pharmacy Wound Closure Adhesive Surgical Tape Strips |  |  | July 26, 2013 | 2013 | ASO, LLC | This recall has been initiated by Aspen Surgical Products, Inc., due to the inability to confirm sterility. | Package design/selection | ASO sent a letter on July 26, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately examine your inventory and stop using/distributing the product. Place in quarantine to be returned or destroyed. ASO customers will receive credit for returned product or a returned completed Certificate of Destruction. Please use Return Authorization Number 1675 when returning the product.Please complete and return the enclosed response form as soon as possible. Return the form to:Address:ASO LLCRegulatory Affairs300 Sarasota Center Blvd.Sarasota, FL 34240Fax:941-378-9040Attn: Regulatory AffairsFor questions regarding this recall call 888-364-7004, ext 146. |  | 2689560 | Total 2,689,560 pouches (2 cards/pouch) | Nationwide Distribution including FL, IA, IL, MI, NC, NY, OH, OR, PA, RI, TX, WA and WI. | Terminated | August 04, 2014 | 374 | 2689560 | Not_Computer
Z-1377-2013 | 64886 | No | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | ***REF S-3D*** 3mm Diamond Ball. |  |  | March 04, 2011 | 2011 | The Anspach Effort, Inc. | Due to three batches of product may have diamond grit size larger than intended applied to the head of the cutting burr. | Labeling mix-ups | Anspach sent an URGENT: Medical Device Correction letter dated March 4, 2011 to its customers identifying the affected product and actions to be taken by the customers. Customers questions or concerns should be directed to Anspach Customer Support at (800) 327-6887. |  | 1827 | 1,827 burrs | Worldwide Distribution-USA (nationwide) including DC and the states of OH, TX, IL, CA, GA, KY, FL, WI, NC, VA, WI, MN, MD, KS, OK, PA, OR ,HI, AZ, WA, TN, AR, IA, MI, CO, LA, NJ, IN, WV, MO, AL, NY, VT, and SC and the countries of France, Canada, Great Britain, Mexico, Australia, Sweden and Netherlands. | Terminated | December 17, 2013 | 1019 | 1827 | Not_Computer
Z-1378-2013 | 64588 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Femoral component, CR, cemented, #1, left2103-1310 |  |  | February 20, 2013 | 2013 | Orthopedic Alliance LLC | The recall of the Loaner Bin and Sterilization Trays for implants and instruments was initiated by Orthopedic Alliance because the directions for use and surgical techniques were not being provided to every physicianand for every surgery. | Labeling design | A recall letter dated 3/8/13 was sent to all customers who purchased loaner bins and sterilization trays for implants and instruments from Orthopedic Alliance. The letter informed the customers of the problem identifiedand the actions to be taken. Customers are instructed to call the office at the number (951) 304-9001 if they are requesting paper copies of any of the literature. |  | 136 | 136 units | Nationwide Distribution including the states of CA, MD, TX, NV, and WI. | Terminated | March 26, 2014 | 399 | 821 | Not_Computer
Z-1206-2013 | 64736 | No | OFP | Anesthesia breathing circuit. | Anesthesiology | General Hospital | Enforcement Discretion | Presource PBDS, Knee Arthroscopy, Kit, Circulator |  |  | March 26, 2013 | 2013 | Cardinal Health, Medical Products &amp; Services | Various Presource Kits contain a pre-assembled Filter and Anesthesia Circuit that may contain outer plastic packaging on one or more components. If the packaging is removed without disassembling the components, remnants of the plastic from the packaging material may become lodged in the filter potentially causing an obstruction of airflow. | Packaging process control | Affected consignees were notified with a letter, dated March 26, 2013, sent via FedEx Priority, of this action. Cardinal Sales Representatives will also be advised and will verbally communicate the action to affected customers. The letter identified the affected product and the problem. Customers are to locate any affected product and quarantine them. In addition, clinicians should be informed of the issue; and, warning labels should be affixed to the front of each kit instructing users to remove and discard the Anesthesia Circuit and Filter Assembly. In order to confirm receipt of this action and completion of labeling the product, the enclosed Acknowledgement Form should be completed and returned to the number provided. If customers do now wish to accept overlabeled product or to overlabel the product in their possession, they are to contact Presource Sales Operations at 800-766-0706, or their sales representative for further instruction. Customers are to contact Cardinal Health Quality Systems at 800-292-9332, if they have any questions or need special assistance. |  | 4571 | 4,571 Kits total | Distributed in the states of AR, MS, SC, and TN. | Terminated | April 12, 2017 | 1478 | 4571 | Not_Computer
Z-0573-2014 | 66914 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Indus Invue Screws: IM71058-XX: 04.2mm, SelfTapping, Tapered. |  |  | November 20, 2013 | 2013 | SpineFrontier, Inc. | Mismarked and unmarked screws | Packaging process control | On 11/20/2013, consignees were notified of the email via letter. Consignees were instructed to return the screws listed in the recall notification letter and return the Recall Acknowledgement form to SpineFrontier via fax or email. |  | 993 | 993 | Nationwide Distribution- including DC and the states of OH, TX, KS, MA, IN, CO, NY, AL, and MI. | Open, Classified |  |  | 1915 | Not_Computer
Z-1305-2013 | 55358 | No | FJL | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Electrode, cutting, 24FR (27050G/6). |  |  | June 01, 2009 | 2009 | Karl Storz Endoscopy America Inc | Possible damage to the sterile packaging (blister tray) of some Karl Storz sterile products, which may result in airborne contamination. | Packaging | A Safety Alert dated May 29, 2009 was sent to affected customers. The letter informed them of possible damage to the sterile packaging (blister tray) of some Karl Storz sterile products which may result in airborne contamination. Karl Storz requested consignees to inspect their inventory of the subject products thoroughly and look for small holes in the blister trays. Any identified units should be returned to Karl Storz, and a replacement package (6 units per package) will be provided to at no charge. The Karl Storz Customer Support number 1-800-421-0837 ext. 7577 was provided to obtain a Return Authorization Number (RMA) and arrange replacement. Questions and comments were forwarded to Technical Support at 1-800-421-0837 ext. 5350. |  | 9432 | 9432 units | Distributed Nationwide and in Canada. | Terminated | May 13, 2013 | 1442 | 9432 | Not_Computer
Z-1366-2013 | 61315 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Pinnacle Hip Solutions ALTRX Polyethylene Acetabular Liner, Ref 1221-36-452 |  |  | March 07, 2012 | 2012 | DePuy Orthopaedics, Inc. | DePuy Orthopaedics, Inc. (DePuy) is issuing a voluntary recall of one lot of the AltrX" liner because the product was incorrectly labeled on the box. The box was labeled for a 36IDX52OD +4 Neutral Liner, but the product enclosed was a 36IDX56OD Neutral Liner. Three ( 3 ) complaints were received in January 2012. | Packaging process control | DePuy Orthopaedics sent an URGENT INFORMATION- MEDICAL DEVICE RECALL NOTICE dated March 7, 2012 via email to all affected distributors and customers. The letter identified the affected product, problem and actions to be taken. Distributors and customers are instructed to cease further distribution or use or of recalled product and to return them to DePuy. The sales representatives are expected to aid customer in returning of recalled product. Distributors and sales representatives are to return the recalled product through the normal DePuy Returns process. Reconciliation form should be completed and returned to your DePuy Sales Representative or faxed to 574-372-7567. For questions about recall information contact the Manager of Customer Quality 574-372-7333. |  | 20 | 20 liners | Worldwide Distribution - USA Nationwide including the state of VA and the country of Ireland. | Terminated | May 24, 2013 | 443 | 20 | Not_Computer
Z-0606-2014 | 66902 | No | CAI | Anesthesia breathing circuit. | Anesthesiology | Anesthesiology | 510(K) Exempt | RTC 25-C Inline Aerosol Tee Adapter. |  |  | November 13, 2013 | 2013 | Instrumentation Industries Inc | Some packages of the RTC 25-C may have been shipped to customers with incomplete Installation and Usage Directions due to a printing error. | Labeling Change Control | A recall notification letter, dated November 13, 2013, was sent to Users. |  | 220 | 220 | Nationwide Distribution. | Terminated | July 09, 2014 | 238 | 220 | Not_Computer
Z-1218-2013 | 64682 | No | GCX | Vacuum-powered body fluid suction apparatus. | General Hospital | General Hospital | 510(K) Exempt | Stryker Irrigator Reusable Tips: |  |  | March 27, 2013 | 2013 | Stryker Endoscopy | The pre-vacuum steam sterilization (wrapped method) temperature in the Instructions for Use incorrectly listed 121-123ÂżC , instead of the correct 132-133ÂżC. | Error in labeling | Stryker sent "Urgent Device Correction" letters dated March 27, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. |  | 147921 | 147,921 - all sizes | Worldwide Distribution: USA (nationwide) and countries of: Mexico, Portugal, Latin America, India, Greece, Taiwan, Singapore, New Zealand, Phillipines, Netherlands, France, Australia, Canada, Japan, Scandinavia, Iberia, Brasil, Romania, South Africa, Chile, Korea, China. | Terminated | September 24, 2013 | 181 | 147921 | Not_Computer
Z-1781-2013 | 65542 | No | GXH | Cryogenic surgical device. | Neurology | Neurology | 510(k) | Myoscience brand iovera Smart Tip, Myoscience Cryo-Touch IV, a Cryogenic Surgical device, REF Catalog Number: STA0311-10. |  |  | June 21, 2013 | 2013 | Myoscience Inc | The outer box label of the product was labeled with a down revision label which did not include the US labeling requirements, specifically the symbol descriptors and also the "Caution: Federal Law (USA) restricts this device to sale by or on the order of a physician." This product was previously labeled for EU distribution only. | Labeling mix-ups | Myoscience representatives visited both affected sites on June 14, 2013 and relabeled the product in person. Neither site had used the product and was unaffected by the relabeling. No formal communication was issued. |  | 3 | 3 units | US distribution including CA and ID. | Terminated | July 15, 2014 | 389 | 3 | Not_Computer
Z-1876-2013 | 65757 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Mullins Transseptal Adult 7F, REF 008551. |  |  | July 08, 2013 | 2013 | Medtronic Inc. Cardiac Rhythm Disease Management | Medtronic identified a specific labeling issue for specific lots of Mullins Transseptal Catheter Introducer Sheath, model 008551, in which packages labeled as Mullins 7F (French) may actually contain an 8F(French) dilator. Medtronic is requesting the return of potentially mis-labeled Mullins Transseptal Catheter Introducer 7 French Sheaths. Due to the differences in size, an 8F dilator will not | Packaging change control | Consignees were sent on 7/8/2013 a Medtronic, Inc " Urgent Medical Device Recall" letter dated 8 July 2013. The letter was addressed to Healthcare Professionals. The letter identified the problem and the product involved in the recall. Advised consignees to remove and quarantine the product from their inventory and return the product. Requested consignees to complete the attached Customer Confirmation Certificate and fax it to Medtronic at 651-367-0612. If the product has been further distributed advised consignees to please notify them of the recall. |  | 269 | 269 | Worldwide Distribution-USA (nationwide) including DC and the states of AZ, AK, CA, DE, GA, IL, KY, MA, MI, MN, MS, NV, NY, NC, OH, PA, TN, and TX, and the countries of AUSTRALIA, BRAZIL, CHILE, INDIA, and SINGAPORE. | Terminated | March 06, 2014 | 241 | 269 | Not_Computer
Z-1779-2013 | 65494 | No | FTC | Ultraviolet lamp for dermatologic disorders. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Dermalight 80 Model Number UVB-083, one unit per box. |  |  | April 08, 2013 | 2013 | National Biological Corp | The incorrect manual may have been packaged with this product. UVB-083 devices may have included themanual for a UVA-080 device. | Labeling mix-ups | The firm, National Biological, sent an "VOLUNTARY PRODUCT RECALL NOTIFICATION" letter dated May 10, 2013 to its consignees/customers by UPS or other delivery service with a signature required. The letter described the product, problem and actions to be taken. Consignees/customers who do not respond to the letter will be contacted via telephone. The customers were instructed to check within your household to determine if you have any of the listed product; immediately destroy the incorrect User Manual, it is not necessary to return the product; record the Customer Name and or the Dermalight-80 serial number and complete and return the enclosed Verification Form, even if you do not have any User Manuals to destroy.Please contact your National Biological Corporation sales representative if you have any questions regarding this recall, any of our products, or would like assistance with the recall call 216-831-0600 or 800-338-5045 email: infor@PhototherapyExperts.com. |  | 6 | 6 units | US distribution including the states of: AL, CA, MN, and NC. | Terminated | February 24, 2014 | 322 | 6 | Not_Computer
Z-1772-2013 | 65504 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Endopath Electrosurgery Probe Plus II, twelve (12) different devices, including hooks, spatulas, right angles, curved dissectors, suction/irrigation devices, accessory ports, pool/sump devices and stone retrieval cannula, packaged and sterilized individually in a transparent flexible blister sealed to a Tyvek lid. 6 flexible packages are placed into a sales unit carton, with 4 cartons per corrugated shipper. |  |  | April 08, 2013 | 2013 | Ethicon Endo-Surgery Inc | The firm is recalling the ENDOPATH Probe Plus II Shafts because the possibility that a tear in the Tyvek packaging may affect the sterility of the device. | Packaging | The firm, Ethicon Endo-Surgery, Inc., sent an "URGENT: MEDICAL DEVICE RECALL" letter dated April 22, 2013 to all consignees/customers via overnight UPS to allow tracking of the receipt. The letter described the product, problem and actions to be taken. The consignees/customers were instructed to do the following: EFFECTIVE IMMEDIATELY-DO NOT USE ANY OF THE FOLLOWING TWELVE (12) PRODUCT CODES within the noted Expiration Dates provide in the letter; examine your inventory immediately to determine if you have affected product on hand; remove the affected product; complete and return the Business Reply Form (BRF) and fax to: 1-888-912-7089 or email to: ees3118@stericycle.com within 3 business days, even if you do not have affected product; return affected product with a photocopy of the completed BRF; and share this information with all of the appropriate staff at your facility. If you need additional shipping labels or a communications package, contact the Stericycle's support center at 1-888-943-4896 and reference Event 3118."Returned ENDOPATH Probe Plus II Shafts will be replaced with corrected ENDOPATH Probe Plus II Shafts once supply becomes available. If replacement product is not available and/or has been discontinued, a credit will be issue."If you need clinical or product support, please contact your local representative or call our Customer Support Center, 1-877-ETHICON (877-384-4266). |  | 327648 | 327,648 instruments | Worldwide distribution: US (nationwide) including states of: AK, AL, AR, AZ, CA, CO, CT, FL, GA, lA, ID, IL, IN, KS, KY, LA, MA, MD, ME, Ml, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, &amp; WY and countries of: Canada, Argentina, Australia, Austria, Bahrain, Belgium, Brazil, Canada, Chile, China, Columbia, Czech Republic, Dominican Republic, Ecuador, Egypt, India, Indonesia, Israel, Italy, Jamaica, Japan, Korea, Kuwait, Libya, Malaysia, Mexico, Pakistan, Panama, Philippines, Poland, Puerto Rico, Russia, Singapore, Slovenia, South Africa, Spain, Saudi Arabia, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Uruguay, &amp; Venezuela. | Terminated | December 24, 2013 | 260 | 327648 | Not_Computer
Z-0183-2014 | 66610 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Applied Medical's InziiÂż 12/15mm Retrieval System. |  |  | October 21, 2013 | 2013 | Applied Medical Resources Corp | Applied Medical is conducting a voluntary recall of the InziiÂż 12/15mm retrieval system. During shipment, the retrieval system packaging has the potential to become punctured with small holes, which could compromise the sterile barrier. The likelihood of this situation to occur and result in patient harm is highly unlikely; however, out of an abundance of caution for patient safety and a commitmen | Package design/selection | Customer notifications of the recall began the week of 10/21/2013 via UPS overnight with letters requesting that all affected products be returned to Applied. The notification letter is dated October 21, 2013 and titled "URGENT: MEDICAL DEVICE RECALL". The notification was accompanied by a Recall Notification Confirmation Form. The letter informed customers of the recall by providing the product description with codes (Model and Lot Numbers). reason for the recall, actions to take and contact information. |  | 29506 | 29,506 units | Worldwide Distribution: US (nationwide) and countries of: Australia, Austria, Belgium, Brazil, Canada, Chile, Colombia, Czech Republic, Denmark, Ecuador, Finland, France, Germany, Hong Kong, India, Ireland, Israel, Italy, Japan, Lebanon, Malaysia, Mexico, New Zealand, Netherlands, Norway, Peru, Poland, Portugal, Saudi Arabia, Singapore, Spain, South Africa, Sweden, Switzerland, United Arab Emirates, and United Kingdom. | Terminated | September 15, 2014 | 329 | 29506 | Not_Computer
Z-0593-2014 | 67052 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart MRx Monitor/Defibrillator |  |  | June 07, 2013 | 2013 | Philips Medical Systems, Inc. | Error in the printing of the therapy energy setting labels on Philips HeartStart M3536A MRx Monitor/Defibrillators. The first energy setting is labeled 10 when it should be labeled 1-10 | Labeling Change Control | Philips sent a Customer Information Letter dated June 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.During an EI of the firm, the firm had not reported the correction to FDA.The Customer Information letter informs the customers what the behavior is and under what circumstances. The correction consisted of a replacement label that was provided free of charge to all units affected by this issue. Philips Healthcare representatives have been contact customers with affected devices to arrange for installation of the replacement label.To correct this issue, Philips will replace the label free of charge. A Philips Healthcare representative will contact customers with affected devices to arrange for installation of the replacement label. This issue does not affect current product shipments.If you need any further information or support concerning this issue, please contact your local Philips representative at 1-800-722-9377.Philips apologizes for any inconvenience caused by this problem. |  | 93 | 93 units | Worldwide Distribution - US Distribution including the states of IN, MS, TX and VA., and the countries of DENMARK, FRANCE, GERMANY, SPAIN and SWITZERLAND. | Terminated | March 22, 2017 | 1384 | 93 | Not_Computer
Z-0380-2014 | 66411 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | 2.4 and 2.7 MM Locking Screw SLF-TPNG with T8 Stardrive Recess 12 MM |  |  | July 31, 2013 | 2013 | Synthes USA HQ, Inc. | Specific lots of the SS 2.4mm and SS 2.7mm Locking Screw (self-tapping) due to a mislabeling issue. The SS 2.7mm Screw was labeled as SS 2.4 mm Screw. | Labeling mix-ups | Synthes sent an URGENT NOTICE: MEDICAL DEVICE RECALL letter dated July 31, 2013, to all affected customer. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to call Synthes at 1-800-479-6329 if they have any of the affected product to obtain a Return Authorization Nuber, complete the Verification Section at the end of the letter by checking the appropriate box indicating affected product has been located, also indicate the number of devices found and note the Return Authorization Number. Fax the documents to Synthes at 1-866-792-6446 or email to FieldAction@synthes.com. Customers with questions were instructed to call 610-719-5450 or contact their Synthes Sales Consultant.For questions regarding this recall call 610-719-5000. |  | 98 | 98 | Nationwide Distribution including PA, CA, FL, MO, KS, IN, IL, WV, KY, LA, NM, SC, IA, TX, AZ, GA, NC, OK, NY, TN, MI, CO, and MT. | Open, Classified |  |  | 98 | Not_Computer
Z-0148-2014 | 66339 | Yes | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius LibertyÂż Cycler |  |  | October 14, 2013 | 2013 | Fresenius Medical Care Holdings, Inc. | Fluid leaking may be observed inside the pump compartment of the Liberty Cycler and may result in dialysate becoming contaminated, potentially causing peritonitis | Device Design | Fresenius Medical Care contacted all affected consignees by telephone on 10/14/13 to inform them of the product safety alert. Follow-up Urgent Safety Alert letter was sent on 10/18/13 and 10/23/13 by Certified Mail, Return Receipt requested. The letter identified the affected product, problem and actions to be taken. Customers were instructed to report any events associated with this safety alert to Customer Service Team at 1-800-323-5188. |  | 16600 | 16,600 | USA Nationwide Distribution | Open, Classified |  |  | 16600 | Computer
Z-0618-2014 | 66468 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | S3 Med/Surg Bed |  |  | November 26, 2013 | 2013 | Stryker Medical Division of Stryker Corporation | Stryker Medical has identified a potential issue associated with Secure II and S3 Med/Surg beds built between April 2nd, 2005 and July 26th, 2011. The affected beds have the potential to experience damage to the Foot End Cover and/or Power Coil Cable due to reduced clearance between the Foot End Cover and Foot End Lift Header. | Device Design | Stryker Medical sent an Urgent Medical Device Recall Notification letter on November 26, 2013, to all affected customers. The notification included the products being recalled, the product description, product issue, potential hazards, and actions needed. Customers were instructed to locate and inspect the beds listed in the notice and ensure that no damage has occurred to the foot-end cover, power coil cable, or sensor coil cable. If the beds were not damaged customers may continue to use them. If damage occurred, customers wsere instructed to remove the beds from service and contact Stryker at 1-800-327-0770, Option 2, then Option 2 to arrange a service call. Continue to inspect the foot-end cover and coil cables as part of their routine preventive maintenance. Customers were instructed to fax the enclosed business reply form to 269-488-8691 or email productfieldaction@stryker.com. Customers were instructed to forward the letter if the product was loaned or sold to other customers. If any beds were disposed of and were no longer in use, customers should advised Stryker by providing them with the serial number. Questions or concerns are directed to Stryker Customer Service at 1-800-327-0770. Consignees are to fax (269-488-8691) or email (productfieldaction@stryker.com) the business reply form to confirm receipt of this notification.For questions regarding this recall call 269-324-6609. |  | 108769 | 108,769 Total | Worldwide Distribution - USA (nationwide) and Internationally to Canada, Chile, China, Argentina, Brazil, Germany, Korea, Mexico, Singapore, UK, Europe, Middle East, Africa, and Latin America. | Terminated | May 29, 2015 | 549 | 108769 | Computer
Z-0872-2013 | 62334 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Liko Overhead Rail System, Ceiling Mounted Rail Systems for all Liko Stationary Lifts including Likorall, Multirall and Masterlift Systems |  |  | October 31, 2012 | 2012 | Hill-Rom, Inc. | Liko/Hill-Rom to date has received two reports from facilities alleging that the overhead rail system has failed to the extent that the overhead rails have fallen down while in use in high humidity (swimming pool) areas. Analysis of the incidents suggest that the components (hardware) used to secure the rail to the ceiling had corroded and may be unsuitable for environments with an atmosphere con | Device Design | Liko, a Hill-Rom company, sent an URGENT MEDICAL DEVICE RECALL notice dated October 31, 2012 to all affected customers. Letter identified the affected product, problem and actions to be taken until a new design is finalized and the field corrective action implemented. The letter instructed customers to complete and return the Response Form immediately even if you do not have any rail systems in a chlorinated environment. For questions concerning this request or the procedure as outlined, please contact Hill-Rom Technical Support at 800-445-3720. |  | 3816 | 3816 in total | Worldwide Distribution: USA (nationwide) including states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, and WY; and CANADA LOCATIONS: AB, BC, MB, NB, NS, ON, and SK. | Terminated | March 09, 2016 | 1225 | 3816 | Computer
Z-0616-2014 | 67002 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | SYSTEM 1E Liquid Chemical Sterilant Processing System, Distributed by STERIS Corporation, Mentor, OH |  |  | November 21, 2013 | 2013 | Steris Corporation | Customers indicated fluctuations in environmental conditions sites and gradual degradation/deformation of check valve o-rings following frequent contact with S40 sterilant use solution result in nuisance aborts. | Device Design | STERIS sent an Urgent Voluntary Field Correction Notice dated November 21, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed that a STERIS Field Service Representative would contact their facility to schedule installation of the new hardware and updated software on their SYSTEM IE Liquid Chemical Sterilant Processing System(s). Customers were instructed that they may continue to operate the SYSTEM IE processor prior to the completion of the update. Customers with questions were instructed to contact STERIS Field Service Dispatch at 1-800-333-8828.For questions regarding this recall call 440-392-7519. |  | 6028 | 6,028 units | Nationwide Distribution - USA (nationwide) and the District of Columbia and Puerto Rico.Product was also shipped to the following countries: Afghanistan, Bahrain, Hong Kong, Italy, Saudi Arabia, Turkey &amp; United Kingdom. | Terminated | June 16, 2014 | 207 | 6028 | Computer
Z-0540-2014 | 66987 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | MedfusionÂż 4000 Syringe Infusion Pumps |  |  | November 22, 2013 | 2013 | Smiths Medical ASD, Inc. | Smiths Medical is conducting a recall involving all MedfusionÂż 4000 Syringe Infusion Pumps.Smiths Medical has identified an issue with electrical shorting at the AC power receptacle. Shorting can occur when saline solution or other conductive fluids (e.g., potassium chloride) enter the AC power receptacle. When this happens, sparking and/ or flames, accompanied by smoke, may appear at the AC po | Device Design | Smiths Medical sent an "Urgent Medical Device Correction" letter dated November 22, 2013, to all affected customers.The letter identified the product the problem and the action needed to be taken by the customers.All customers will be provided with instructions on inspection, cleaning, and prevention regarding this issue, for the continued safe and effective use of these pumps. Additionally, customers will be instructed to:1. Return damaged pumps for servicing or damaged AC power cords for replacement to Smiths Medical.2. Return the completed Confirmation Form to acknowledge receipt and understanding of the information provided in the Urgent Medical Device Correction letter, via fax or email, to Smiths Medical for processing."The Instructions for preventing this issue", "Mitigations", Conclusions", " Action to be taken by the User", and "Transmission of this Urgent Medical Device Correction". The consignees were also asked to complete the "Urgent Medical Device confirmation Form" and fax it to 1-800-237-8033 or by sending an electronic copy via e-mail to medfusion4000nov2013@smiths-medical.com.For questions regarding this information, consignees should contact their local Sales Representative, Smiths Medical's Customer Service Department at 1-800-258-5361, or via e-mail at medfusion4000nov2013@smiths-medical.com.UPDATE: Consignees were sent a follow-up letter dated 13 January 2014 from the one dated November 22, 2013. The letter was addressed to Risk / Safety Managers, Biomedical Professionals, Clinicians who oversee the use of Medfusion Pumps, Distributors, and other users of these devices. The "Urgent Medical Device Recall" letter included Advise on Action to be Taken by the User. This included customers to have their Pumps undergo a free correction. The correction consists of the installation of a new AC power Receptacle seal. Smith Medical provided options for completing the correction process which included: 1) Consignees were a |  | 9560 | 9560 (8601 US, 959 OUS) | Worldwide Distribution - USA (nationwide ) including the states of : AZ, NY, TN, MS, AR, WV, VA, VT, PA, DE, DC, OH, CO, NH, MI, CA, FL, CT, WA, MD, NM, RI, MA, TX, NC, IL, KY, WI and MI., and the countries of CANADA, SAUDI ARABIA and AUSTRALIA. | Terminated | March 25, 2015 | 488 | 9560 | Computer
Z-0545-2014 | 66983 | No | EIA | Dental operative unit and accessories. | Dental | Dental | 510(k) | SEAC Advanced Dental System, Model A591561 |  |  | November 05, 2013 | 2013 | ASI Medical, Inc. | ASI Medical is initiating a field correction for the SEAC Advanced Dental System because the miniature plastic air tank can potentially rupture. | Device Design | ASI Medical sent an Field Correction Action and Safety Notice letter dated November 5, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.REQUIRED FIELD CORRECTIONOur field correction will consist of sending you a warning label to be applied to the back service access panel of the cart and an ASME safety pop off valve that will open and exhaust if pressure should reach over 100 psi. This safety pop off valve is designed be easily installed in line between the air compressor and air tank in your system. We are requesting your cooperation in making this field correction. |  | 337 | 337 | Worldwide Distribution - USA (nationwide) including the countries of Austria, Aruba, Saudi Arabia, Canada, Germany, Spain and West Indies. | Terminated | August 11, 2014 | 279 | 337 | Computer
Z-0544-2014 | 66698 | No | MNI | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | PedFuse Reset Screw; Model: 01-80115-55 |  |  | October 18, 2013 | 2013 | SpineFrontier, Inc. | During routine servicing, it was discovered that one 8mmx55mm Reset screw was assembled with an 8mm tulip head and a 9mm screw. The implant should haveconsisted of an 8mm tulip head and an 8mm screw. A recall was initiated to remove all remaining screws of the same lot from the field to investigate the nonconformity. | Software design (manufacturing process) | SpineFrontier sent an Voluntary Recall Notification dated November 5,2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.All affected implants are to be returned to SpineFrontier for inspection. SpineFrontier is requiring confirmation that you have received this notice and that you will initiate the immediate return of the implants listed above.IMMEDIATE RESPONSE REQUIRED.Please feel free to contact us at 978-232-3990 Monday through Friday8:30AM to 5:30PM EST with any questions |  | 35 | 35 Devices | US Distribution to VA and TX. | Open, Classified |  |  | 35 | Computer
Z-0119-2014 | 66059 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Xario Diagnostic Ultrasound System SSA-660A |  |  | August 05, 2013 | 2013 | Toshiba American Medical Systems Inc | The firm initiated this recall of certain Toshiba Diagnostic Ultrasound Systems because it has been found that under certain conditions, the temperature of the transducer head may increase in 2D+PW mode or 2D+CDI (POWER, ADF) +PW mode. When the system is used under both conditions listed below, the temperature of the transducer head connected to the system may increase in temperature and possibly | Device Design | Toshiba sent an Urgent Medical Device Correction letter dated August 5, 2013. along with a Reply Form via USPS return receipt mail. The letter identified the product, the problem, and the action to be taken by the customer. Until the software is corrected customers were instructed to change the preset settings as described in the letter. Customers were asked to complete and return the attached form and fax it to the toll free number. The form also can be sent via email to raffairs@tams.comCustomers with questions were instructed to call 800-421-1968 or contact their local Toshiba Representative at 800-521-1968.For questions regarding this recall call 714-730-5000. |  | 17 | 17 (1352 Total) | Nationwide Distribution including and US Territory Puerto Rico | Open, Classified |  |  | 17 | Computer
Z-0823-2013 | 64242 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, CARESCAPE Monitor B650. |  |  | September 30, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware of potential issues associated with the CARESCAPE Monitor B650. There have been no reported illnesses or injuries that have occurred with the use of the device. | Software design | GE Healthcare sent an Important Product Notification letteron September 30, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to keep the patient under close observation or use an alternate monitoring device. Customers should not disconnect the affected USB devices from the patient monitor during an active patient case. If the license is enabled, the alarm limit adjustments from the iCentral Client has an L-NETM-00 license should be disabled. GE Healthcare would contact customers to schedule the update. Customers with questions should contact Technical Support or their local representative.For questions regarding this recall call 262-513-4122. |  | 2813 | 2813 | Worldwide Distribution - USA including AL, AZ, AR, CA, CO, CT, FL, GA,IL,IN, IA, KS, KY,LA, MD, MA, MI, M, MO, NE, NE,NJ, NM, NY, NC, OH, OK PA, SC, TN, TX, UT, VA, WI. Internationally to UNITED KINGDOM, UNITED ARAB EMIRATES, TURKEY, THAILAND, SWITZERLAND, SWEDEN, SPAIN, SOUTH AFRICA, SLOVENIA, SLOVAKIA, SINGAPORE, SAUDI ARABIA, ROMANIA, PORTUGAL, POLAND, PERU, NORWAY, NEW ZEALAND, NETHERLAND, MALAYSIA, LITHUANIA, KUWAIT, JAPAN, ITALY, ISRAEL, IRELAND, INDIA, ICELAND, HONG KONG, GERMANY, FRANCE, FINLAND, EGYPT, DENMARK, CZECH REPUBLIC, COLOMBIA, CHILE, CANADA, BELGIUM BAHRAIN, AUSTRIA, AUSTRALIA. | Terminated | January 27, 2014 | 850 | 2813 | Computer
Z-0634-2013 | 63746 | No | JJH | Clinical sample concentrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Abbott m2000sp Instrument; an automated fluid handling system for performing sample preparation for nucleic acid testing; Abbott Molecular Inc., Des Plaines, IL 60018 USA; Abbott m2000sp Instrument; an automated system for performing sample preparation for nucleic acid testing; Abbott Molecular Inc., Des Plaines, IL 60018 USA; list 09K14 |  |  | December 04, 2012 | 2012 | Abbott Molecular | An issue with the m2000sp Instrument software has been identified whereby, after a series of steps, if changes are made to the control or calibrator information on the Sample Extraction Assay Details screen, the changes are not saved. | Software design | Abbott Molecular Inc. sent a Field Correction Recall/Urgent Field Safety Notice letter dated December 4, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The customer was instructed to follow the instructions identified until an upgrade to their m2000sp software can be installed. Abbott Molecular estimates that this software version will be available the first quarter of 2013. They were requested to complete the enclosed reply sheet and fax it to (847) 775-6728 at prior to December 18, 2012 to acknowledge receipt and understanding of the letter.For questions customers were instructed to contact their local Abbott Molecular representative.For questions regarding this recall call 224-361-7274. |  | 215 | 215 units - U.S., 693 units - ROW | Worldwide Distribution - USA (nationwide) and internationally to Algeria, Argentina, Australia, Austria, Azerbaijan, Belgium, Bolivia, Brazil, Burkina Faso, Burundi, Cameroon, Canada, Chad, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Denmark, Ethiopia, Estonia, Finland, France, Gaza &amp; Jericho, Germany, Greece, Guinea, Haiti, Honduras, Hong Kong, India, Indonesia, Ireland, Israel, Italy, Japan, Kazakhstan, Kenya, Libya, Lithuania, Luxembourg, Malawi, Malaysia, Mali, Mexico, Mozambique, Netherlands, New Zealand, Norway, Pakistan, Panama, Paraguay, Peru, Poland, Portugal, Romania, Russia, Saudi Arabia, Singapore, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Turkey, Uganda, Ukraine, United Arab Emirates, United Kingdom, Venezuela, Vietnam and Zambia | Terminated | March 08, 2016 | 1190 | 215 | Computer
Z-0168-2014 | 66239 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | COBAS INTEGRA 800 Analyzer with software 9864.C2 |  |  | September 16, 2013 | 2013 | Roche Diagnostics Operations, Inc. | Roche COBAS INTEGRA 800 system, Software Version 9864.C2 does not perform necessary ISE Service Actions. Whenever the ISE rack is removed, the ISE counters in the software are reset to zero and the corresponding Service Actions are not conducted automatically by the system. Additionally, certain maintenance activities that have to be performed manually by the operator are not requested on the sof | Software Design Change | Roche sent an Urgent Medical Device Correction letter dated September 16, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Roche Diagnostics recommends tracking all ISE counters and manually requesting the corresponding Service Actions listed in the notification be monitored until updated software becomes available during Q4 2013. If the affected product was moved to another site, the notification is to be forwarded. consignees were asked to complete the attached faxback form and fax it to 1-877-266-0997 and file the Urgent Medical Device Correction (UMDC) for future reference. Question should be directed toRoche Support Network Customer Support Center 24 hours a day, seven days a week at 1-800-428-2336For questions regarding thisr ecall call 317-576-3911. |  | 527 | 527 units | Nationwide Distribution and Puerto Rico | Terminated | May 12, 2015 | 603 | 527 | Computer
Z-0127-2014 | 66456 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | Dekompressor Percutaneous Discectomy Probe. |  |  | November 04, 2013 | 2013 | Stryker Instruments Div. of Stryker Corporation | Stryker has identified a potential for the auger and/or cannula to fracture when the Dekompressor is used contrary to the warnings in the Instructions For Use (IFU). Specifically, manually bending the cannula can weaken the device and may result in a fracture of the auger and/or cannula. This could result in the need for additional medical intervention to remove the broken piece, pain or loss of m | Device Design | On 11/4/13, an URGENT MEDICAL DEVICE Recall Notification was sent to all domestic consignees via mail and foreign by email. An advisory was sent to every physician that uses any of the products listed. Every physician thatuses any of these products must print their name and sign below which means you acknowledge that youve received a copy of the attached advisoryand that you understand the information regarding the proper usage of the Decompressor.Once the form has been signed by all physicians, the form mustbe completed and signed by the recall coordinator. Please fax to866-521-2762 or email to kara.spath@stryker.com. |  | 20385 | 20,385 total distributed since 2004 | Nationwide Distribution including the states of AR, AZ, CA, CO, FL, GA, IL, KS, MS, ND, and NE. | Terminated | April 29, 2014 | 176 | 20385 | Computer
Z-0116-2014 | 65471 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | The Centricity PACS Workstation is intended for use as a primary diagnostic and analysis tool for diagnostic images by trained healthcare professionals, including radiologists, physicians, technologists, clinicians and nurses. It is also intended for use as a clinical review workstation throughout the healthcare facility. The workstation interface provides the user with a means to display, manipulate, archive, print and export images when connected with the Centricity PACS infrastructure. |  |  | June 11, 2013 | 2013 | GE Healthcare It | GE has recently become aware of potential issues due to out of context issues associated with specific workflows of the Centricity PACS RA1000 workstation. These issues are as follows: 1) In Centricity RIS-IC, and a RIS-Driven integration, the "close" button closes the exam in RIS-IC, but does not close the exam in Centricity PACS RA1000 Workstation. 2) Interrupted workflow could break the | Software design | URGENT MEDICAL DEVICE CORRECTION letters dated June 11, 2013 was mailed to all customers alerting them to these potential safety issues. The software was modified to correct these defects and a GE Healthcare service representative will contact each customer to make arrangements for installation of the new software version at no charge to the customer. Customers with concerns or questions about the market action can contact their local GE Healtchcare IT service representative or GE Healthcare's Remote Online Center - Customer Care Center at 1-800-437-1171. |  | 623 | 623 units | Worldwide Distribution: US (nationwide) including states of: AK, AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI and WV) and the countries of the UK, Taiwan, Switerzland, Sweden, Spain, Singapore, Saudi Arabia, Netherlands, Malaysia, Korea, Japan, Italy, Israel, India, Hong Kong, Germany, France, Denmark, China, Canada, Bermuda, Beligum and Australia. | Open, Classified |  |  | 623 | Computer
Z-0827-2013 | 64248 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ELECTROLASE(R) HYFRECATOR(R) ELECTRODES - BLUNT, REF/Catalog Number 7-101-8BX, Disposable Hyfrecator(R) Tips, Sterile, Rx Only --- CONMED CORPORATION USA. |  |  | January 28, 2013 | 2013 | ConMed Corporation | Some devices had broken through the seal of the sterile pouch. ConMed confirmed instances where the pouch seal was compromised on the affected products. | Equipment maintenance | ConMed sent Urgent Device Recall Letters (dated 1/28/13) and Business Reply Forms to all domestic customers via USPS Priority Mail. International notices were sent via FedEx International Priority to the foreign consignees. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review their inventory and complete Attachment II response form and return it with unused affected product to: ConMed Corporation 525 French Road, Utica, NY 13502. Customers were asked to return Attachment II even if they don't have any affected product. For questions regarding this recall call 315-624-3533. |  | 104250 | Domestic Distribution: 97,450 units; Foreign Distribution: 6,800 units | Worldwide Distribution-USA (nationwide) and the countries of China, Dubai and Indonesia. | Terminated | April 11, 2017 | 1534 | 104250 | Computer
Z-0132-2014 | 65586 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | DRX Revolution, Catalog #'s 1)1019397, 2) 1023415 (w/ installation warranty, US-only) and 3) 1023423 (w/ parts warranty, US-only). |  |  | June 19, 2013 | 2013 | Carestream Health Inc. | Carestream Health Inc. is conducting a recall for the Column end Cover present on the DRX Revolution Mobile X-Ray System due to the cover dislodging from the column. | Device Design | The firm, Carestream Health Inc., sent a "URGENT: Medical Device Recall" letter dated June 17, 2013 via courier service to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure that there is no opportunity for interference between the End Cover and potential obstructions as the column is descending. Carestream will deploy a field service representative to apply a warning label to the column, and to provide an addendum to the Safety and Regulatory Information manual. Your Carestream service representative will contact you to answer any questions and schedule a visit as soon as possible. If you have any questions or concerns, please contact the Carestream Customer Care Center at 1-800-328-2910 from 8am through 8pm. |  | 331 | 331 units | Worldwide distribution: US (nationwide) including states of: AL, AR, AZ, CA, CO, FL, GA, IA, ID, IL, IN, IO, KS, LA, MA, MD, MI, MN, MO, NC, NY, OH, PA, SC, TN, TX, UT, VA, WI, and WV; and countries of: Australia, Brazil, Canada, Dubai, Germany, Japan, Kuwait, Malaysia, New Zealand, Saudi Arabia, Singapore, South Africa, Taiwan, and Thailand. | Open, Classified |  |  | 331 | Computer
Z-0567-2014 | 66976 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Ziehm Vision R C-Arm Interventional Fluoroscopic X-Ray System. |  |  | November 21, 2013 | 2013 | Ziehm Imaging Inc | Engineering change in 2006 resulted in a minor non-conformity of their devices. Specifically the emergency switch was changed from a standard component to just an option on the devices. | Device Design | Ziehm Imaging sent an "URGENT RECALL NOTICE" letter dated November 22, 2013, to all affected customers. The letter identified the products the problem and the action needed to be taken by the customers. Customers were advised and instructed to:Your Model and Serial number can be found on the product label attached to the side of the mobile C-arm Stand. Our records indicate that your facility has one or more of the products listed above that are being used in your facility.Recall of the product was initiated because according to our manufacturing documentation, the referenced device installed at your facility was not equipped with an emergency stopping device for vertical travel of the c-arm imaging assembly. To fully comply with applicable standards the referenced device must be retrofitted with an emergency stopping device identification and label.The issue:According to our manufacturing documentation, the referenced device installed atyour facility was not equipped with an emergency stopping device.To fully comply with applicable standards the referenced device must be retrofitted with an emergency stopping device. The manufacturer Ziehm Imaging GmbH advised customers they may continue to operate the affected device with caution until the correction has been carried out on-site. During this time Ziehm Imaging GmbH advises the operator to be aware of the following safetymeasures which may resolve the possibility of potentially hazardous situations by:" Operating the power-off switch located either on the C-arm Stand top panel or on the monitor cart" Removing the power cable from the wall outlet or," Activate the push-buttons for reverse vertical movement and keeping the pushbutton pressed until patient /operator is no longer in the risk zone or the device stops Please alert all staff members working with the device to these potential hazardous situations, and familiarize them with all possibilities to resolve the hazardous situation |  | 103 | 103 devices | USA (Nationwide) Distribution including the states of AK, AL, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MT, MA, MD, MI, MN, MN, MO,MS, NC ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY and Hawaii. | Terminated | June 23, 2014 | 214 | 386 | Computer
Z-0570-2014 | 67046 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart MRx Monitor/Defibrillator with CO2/EtCO2 Measurement |  |  | December 10, 2013 | 2013 | Philips Medical Systems, Inc. | MRx defibrillator displays a -?- for EtCO2 and does not display EtCO2 values when patient CO2 level falls below 7.6 mmHg | Software design | Philips Healthcare issued an Customer Information letter dated February 20, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer with the addendum to the Instructions for Use.If you have a ? in the parameter block and no CO2 waveform on the display, the waveform source is invalid. If invalid, check patient, confirm airway status and examine the FilterLine to make sure it is connected to the device and not kinked or pinched. For more information on the EtCO2 Option seeMonitoring Carbon Dioxide chapter in the HeartStart MRx Instructions for Use.Should you have any questions or concerns, please contact your local Philips representative at 1-800-722-9377. |  | 18529 | 18,529 units | US Nationwide and OUS. | Terminated | April 18, 2017 | 1225 | 18529 | Computer
Z-0571-2014 | 66754 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Instrument Manager driver : Driver sysu100i that connects to the Sysmex UF-100 analyzer |  |  | November 06, 2013 | 2013 | Data Innovations, Inc. | When connected to the Sysmex UF-100, theInstrument Manager driver sysu100i driver software versions 7.00.0001, 7.00.0002, 7.00.0003, 7.00.0004, 7.00.0005, 8.00.0001, 8.00.0002, 8.00.0003,and 8.00.0004, and is configured with option "Strip Zeros from Result" set to "Yes" may result in sample error | Software design | Data Innovation notified consignees by Important Medical Device Recall Notification dated 11/5/13 to all list serv recipients of the problem and posted critical alert on Customer Web Portal. A certified letter was issued to impacted customers and consignees if no response was recieved from e-mail. |  | 24 | 24 systems | NationwideCanada and Britain | Terminated | December 05, 2014 | 394 | 24 | Computer
Z-0561-2014 | 65131 | No | JMG | Ferritin immunological test system. | Immunology | Clinical Chemistry | 510(k) | Dimension(R) Chemistry Wash (RD701), Siemens Material Number (SMN) 10445052 |  |  | May 06, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics received complaints of Chemistry Wash causing QC and patient sample result shifts on CTN1, LTN1, FT4, TSH, PBNP and LPBN when using Chemistry Wash lots RD23031, RD23111, RD23241, RD23311, RD23391, and RD23461. These lots shipped between 12/12/2012 and 3/22/2013. Internal investigation has confirmed shifts can be encountered in two scenarios: 1 - When an affected l | Software design | Siemens sent an Urgent Medical Device Recall letter dated May 2013 to all affected customers. The letter identified the product and lots, the reason for the field action, the risk to health and the action(s) to be taken. Customers were instructed to discontinue use and discard inventory of the affected product. Siemens will replace any unused inventory of the affected lots at no charge. For questions contact your Siemens Technical Center or your local Siemens technical support representative. |  | 17719 | 17,719 | US Nationwide and OUS. | Terminated | August 27, 2014 | 478 | 17719 | Computer
Z-0580-2014 | 66980 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | steroEOS Workstation with software versions: 1.3, 1.4 and 1.5 used in processing 3D X-ray images |  |  | November 29, 2013 | 2013 | EOS Imaging | Error copying information to the clinical file and to the report when two 3D models are opened simultaneously on the sterEOS workstation | Software design | EOS imaging issued a safety information letter dated November 29, 2013, to the US customer sites. The letter identifed the affected product, problem and actions to be taken. Users of the sterEOS workstation are asked not to simultaneously open several pairs of images in sterEOS 3D. The letter informed customers that EOS imaging will, at it's own expense, corect all systems affected by means of a software update. An EOS imaging representative will contact affected customers to make an appointment for this update once it is available. Users of the sterEOS workstation are asked not to simultaneously open several pairs of images in sterEOS 3D. For questions call +33 (0) 155 25 6060. |  | 36 | 36 units | USA Nationwide Distribution | Terminated | January 13, 2015 | 410 | 36 | Computer
Z-0645-2013 | 63829 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ABX PENTRA ML, Software Versions: V8.0.x and Below; V.9.0.1 and V9.0.2 |  |  | September 20, 2012 | 2012 | Horiba Instruments, Inc dba Horiba Medical | Horiba Medical is recalling ABX PENTRA ML because they have confirmed an issue with the ABX PENTRA ML data management and validation system. An incorrect patient result is transmitted to the Laboratory Information System. This result is from a previous run of the same patient that is stored in the ML database. This issue occurs in a very rare sequence of specific events. The affected software | Software design | .A recall letter dated 11/09/2012 was sent to all their customers who received the ABX PENTRA ML products. The letter informed the customers of the problem identified and the action to be taken. Customers are instructed to complete and return the enclose response form within 10 days. Customers with questions about the letter are instructed to contact their local HORIBA Medical representative.Recall Letter, as produced per FCD-0377, Field Safety Notice &amp; Acknowledgement will be mailed via FedEx to the ten (10) customers noted on the list the week of November 12, 2012. Effectivity will be evaluated by receipt of the Field Safety Notice Acknowledgment Form. Follow-up communications with customers will be conducted, as needed. |  | 10 | 10 units | Nationwide in the US, including the states of CA, MI, IN, FL, and GA. Center Recall Depth is Retail. | Terminated | March 12, 2013 | 173 | 10 | Computer
Z-0838-2013 | 64246 | No | HIS | Condom. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | LifeStyles ZERO 10 uber-thin lubricated latex condoms. Mfd. for Ansell Healthcare Products LLC, Dothan, AL 96303. |  |  | January 14, 2013 | 2013 | Ansell Healthcare Products LLC | Out of an abundance of caution, Ansell is implementing a voluntary recall of the Zero Latex Condom (regular) product. The condom has met or exceeded test and quality standards required by the Food and Drug Administration. However, we are experiencing a number of complaints concerning the difficulty in unrolling and subsequent breakage, and so we are asking to have all products returned. | Device Design | Ansell HealthCare Product LLC sent an Recall letter dated January 14, 2013, via registered mail to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The recall notification informs customers about the problem with the affected product and provides instructions to the customers on how to handle affected product. Customers are asked to complete the Response Form and return via fax (1-800) 722-8155.If you have further questions, please contact our Customer Service Department at 1-800-952-9916. |  | 469134 | 469,134 units | .Nationwide Distribution including the states of AL, AZ, CA, CO, DE, FL, GA, IA, IN, IL, KS, KY, LA, MI, MN, MO.MS, NC, NH, NJ, NY, OH, OR, PA, RI, SC, TN, TX, VA, WI and UT. | Terminated | August 06, 2015 | 934 | 469134 | Computer
Z-0016-2014 | 65161 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Siemens RAD Fluoro Uro Systems - Luminos dRF, Ysio and Uroskop Omnia with software versions VB10D, VB10F, VB10C and VB10E. |  |  | May 06, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Siemens has discovered through product monitoring a potential malfunction when using system operating with software versions VB10D, VB10F, VB10C and VB10. During a RAD examination the possibility exists that radiation may continue longer than expected and may result in unnecessary dose applied during an acquisition. Additionally, it is not possible for the user to stop the radiation manually; ho | Software design | Siemens initiated this correction on May 6, 2013 by sending Update InstructionXP005/13/S. Firm subsequently issued a letter on June 3, 2013 to all accounts.1. Siemens Field Modification of the Luminos dRF, Ysio &amp; Uroskop Omina running software versions VB10D, VB10F, VB10C and VB10E, which improves the signal transmission for the automatic exposure control, is available as of May 2013, and systems will be updated within one year.2. The revised software fix for models Luminos dRF, Ysio &amp; Uroskop Omina running software versions VB10D, VB10F, VB10C and VB10E, will be provided to all identified customers at no charge.CDRH approves the CAP subject to the following conditions:1. Notification of all dealers and purchasers is to be made within 15 workingdays of receipt of this letter in the manner specified in 21 CFR 1003.21 and1003.22. This office and the Food and Drug Administration (FDA) districtoffice coordinator noted below are to be included in the notification.2. Siemens Medical Solutions USA, Inc., will notify the FDA district office inPhiladelphia, Pennsylvania, with reports of progress regarding this fieldcorrection. |  | 23 | 23 | Nationwide Distribution | Terminated | November 06, 2014 | 549 | 23 | Computer
Z-0125-2014 | 66473 | No | LGW |  |  | Neurology | PMA | Medtronic Activa PC, model 37601, and Activa RC, model 37612, Multi-program Rechargeable Neurostimulator for Deep Brain Stimulation. |  |  | September 30, 2013 | 2013 | Medtronic Neuromodulation | Medtronic notified customers regarding over stimulation or stimulation in the wrong area related to select Medtronic neurostimulators. Upon a specific set of conditions, typically related to device recovery from an over discharge, there is a potential for over stimulation or stimulation directed to a lead electrode other than what was intended. | Software design | The firm, Medtronic, sent an "Urgent: Medical Device Correction" letter dated September 2013 entitled "Loss of Stimulation and Over Stimulation" to customers associated with the affected Medtronic neurostimulators. This letter described the product, problem and actions to be taken. The letter informed customers of the potential for these issues to occur and provided recommendations to significantly reduce the likelihood of the issues. The customers were instructed to avoid over discharge; if a rechargeable device becomes over discharged, contact Medtronic Patient Services at (1-800-510-6735) to resolve the POR and reduce the likelihood of over stimulation.If you have questions, please contact Medtronic Neurostimulation Technical Services at 1-800-707-0933 weekdays 7am-6pm CST. |  | 38294 | 38,294 | Worldwide distribution: US (Nationwide) including Puerto Rico and countries of: Algeria, Aruba, Austria, Belarus, Belgium, Brazil, Bulgaria, Chile, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, El Salvador, Estonia, Faroe Islands, Finland, France, French Polynesia, Georgia, Germany, Greece, Guadeloupe, Hungary, Iceland, India, Iran, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kuwait, Latvia, Lebanon, Libya, Lithuania, Luxembourg, Malta, Martinique, Mexico, Morocco, Netherlands, Netherlands Antilles, New Caledonia, Norway, Oman, Pakistan, Panama, Peru, Poland, Portugal, Qatar, Reunion, Romania, Russian Federation, San Marino, Saudi Arabia, Serbia, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay, and Venezuela. | Terminated | November 20, 2015 | 781 | 102457 | Computer
Z-0836-2013 | 64307 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | Animas(R) 2020, Animas(R) IR 1250 (1200 plus), Animas(R) IR 1200 Insulin Infusion Pumps |  |  | December 17, 2012 | 2012 | Animas Corporation | Due to an internal calendar date limit (December 31, 2015) these pumps will cease to operate on January 1, 2016, and will result in a call service alarm that is displayed on the pump display. | Device Design | The firm initiated their recall of this product on December 17, 2012 by sending a letter to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the Animas Customer Support Department at (866) 796-6373 for questions or concerns. |  | 37998 | 37,998 | Worldwide Distribution-USA (nationwide) including Washington, DC, Puerto Rico, US Virgin Island and the states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY, and the countries of Austria, Australia, Bermuda, Canada, New Zealand, Czech Republic, Italy, France, Finland, Hungary, Israel, United Kingdom, Spain, Sweden, Norway, Denmark, Germany, Mexico and Ireland. | Terminated | December 24, 2013 | 372 | 37998 | Computer
Z-0090-2014 | 65805 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | DRX Revolution Mobile X-Ray System, Model DRXR-1, Verona Street Rochester, NY 14608 |  |  | July 26, 2013 | 2013 | Carestream Health, Inc. | During an audit, Carestream discovered a DRX-1 System that captured an image which displayed an artifact. This issue was discovered in manufacturing. It occurs when the DRX product console software is updated to v5.6b software and the detector is not calibrated before use. | Software design | Carestream sent a Urgent-Medical Device Recall letter dated July 26, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Carestream Health Inc. recently discovered an image artifact issue due to a detector not being recalibrated with the 5.6B software upgrade. As our installation instructions did not require the user to recalibrate during an upgrade, there is a possibility your detector could do the same. When this occurs an image artifact could be evident on images captured in the Extended Exposure mode. These artifacts could potentially contribute to misdiagnosis or the necessity of patient re-exposure.We request that you take a moment to verify that your detector(s) are calibrated. At our next visit, Carestream will also verify that your detector was re-calibrated and, if not, re-calibrate the detector for you. Please note this only affects systems upgraded in the field.For questions or concerns, please contact the Carestream Customer Care Center at 1-800-328-2910, 8 am through 8 pm EST in the US. |  | 23 | 23 units installed total (11 domestically &amp; 12 internationally) | Worldwide Distribution - US (nationwide) including the states of New York, Massachusetts, Pennsylvania, Illinois, Texas, California and Louisiana., and the countries of Australia, Canada, Finland, Italy, Switzerland and United Kingdom. | Open, Classified |  |  | 23 | Computer
Z-0126-2014 | 65469 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Carestream DRX-Revolution Mobile X-ray System. |  |  | June 06, 2013 | 2013 | Carestream Health, Inc. | Carestream Health Inc. is recalling the power plug on their DRX-Revolution Mobile X-Ray System due to the item generating sparks and becoming warm to touch. | Device Design | On June 6, 2013 Carestream Health Inc. initiated a recall for the power plug on their DRX-Revolution Mobile X-Ray System due to the item generating sparks and becoming warm to touch. Carestream Health Inc. will replace the power plug in all potentially impacted DRX-Revolution Mobile X-Ray Systems. Your Carestream service representative will contact you to answer any questions and schedule a visit to replace the plug of your Mobile X-Ray System as soon as possible. If you have any questions or concerns, or if you experience any issue with your plug prior to our visit, please contact the Carestream Customer Care Center at 1-800-328-2910, 8AM through 8pm Eastern Standard Time. |  | 180 | 180 units | Nationwide Distribution including the states of NY, TX, PA, OH, WI, CA, IL, CO, FL, IN, OK, NJ, GA, LA, AL, KS, AZ, MA, SC, MI, NC, IA, UT, VA, TN, and MN. | Open, Classified |  |  | 180 | Computer
Z-0655-2013 | 63938 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | FlowGate Balloon Guide Catheter; |  |  | December 20, 2012 | 2012 | Concentric Medical Inc | Firm received complaints of resistance and possible collapse of the distal tip of the FlowGate Balloon Guide Catheters during use for aspiration. | Device Design | Concentric Medical sent a Urgent Medical Device Recall Notification letter dated December 19, 2012, to all affected customers. The letter identified the product the problme and the action needed to be taken by the customer.Customers were instructed to follow these instructions.1. Immediately locate the subject devices and quarantine them.2. Distribute this notice to all affected Departments in your facility.3. Inform Concentric Medical (Stryker Neurovascular) if any of the subjectdevices have been distributed to other organizations.a. Please provide contact details so that Concentric Medical (StrykerNeurovascular) can inform those recipients appropriately.4. Complete and sign the attached Customer Response Form and: Fax it to: 650-237-5230 Please complete this form even if you do not have any subject devices to return to prevent the need of follow up notices.5. Keep a copy of the completed, signed Customer Response Form for yourrecords. Upon receipt of the returned customer response form, a Stryker Representative will contact you to arrange for credit of returned devices.On behalf of Stryker Neurovascular, we sincerely thank you for your help and support and regret any inconvenience and appreciate your assistance in completing this recall.Should you have any questions concerning this matter please do not hesitate to contact the undersigned. |  | 13 | 13 units | US (nationwide) including the states of CA, FL, GA, OR and TN. | Terminated | February 05, 2013 | 47 | 13 | Computer
Z-0602-2014 | 67094 | No | GWM | Intracranial pressure monitoring device. | Neurology | Neurology | 510(k) | Integra Camino Advanced Monitor with Waveform Display Model CAM01 Integra Neurosciences, 311 Enterprise Dr Plainsboro, NJ 08536 USA |  |  | November 26, 2013 | 2013 | Integra LifeSciences Corp. | Integra became aware through a complaint from a single non-US customer that if the optional SYS/DIAS mode is used on CAM01 monitor that contains Mod 4 Firmware, the displayed values for the systolic and diastolic intracranial pressure (ICP) are reversed. But, if the CAM01 is used with an auxiliary bedside monitor, the bedside monitor displays these values correctly. | Device Design | Integra sent an Urgent Medical Device Correction letter and Correction Acknowledgement form dated November 26, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Our records indicate that you received one or more of the affected serial numbers that will require this firmware update. Integra asks that you do the following:1. Review your inventory and verify that you have Camino Advanced Monitors with a serialnumber formatted as MExxxxxx04. If so, you can continue to use the CAM01monitor(s), being aware that if you choose to use the optional SYS/DIAS mode, thevalues on the CAM01 display will be reversed.2. Complete the attached form. If you do not have the product with the affected serialnumber in service at your facility, check the box:I do not have any product with a serial number of the form MExxxxxx04.3. If you do have the affected serial number, check the box:I do have the product with a serial number of the form MExxxxxx04, and pleaseindicate the serial numbers you have.4. Complete the other information and return it by email or fax as indicated on the form.Complete the form and return it by email or fax as indicated on the form. Should you have any questions regarding these instructions, please contact Customer Service at 1-855-532-1723.We apologize for any inconvenience this may cause and thank you for your cooperation in this effort. |  | 276 | 275 CAM01 Monitors | Worldwide Distribution - USA ( Nationwide ) and the countries of : UK, Saudi Arabia, South Africa, Brunei, China, France, Poland, Turkey, Greece, Sweden, United Arab Emirates, Russia, Spain, Pakistan, Finland, Israel, Italy, Slovenia, Japan, Poland, Malaysia, Germany and Panama. | Terminated | February 13, 2015 | 444 | 276 | Computer
Z-0867-2013 | 64319 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Eclipse Treatment Planning System for Radiotherapy; Model Number: H48. In addition, the Eclipse Proton Eye algorithm is specifically indicated for planning proton treatment of neoplasms of the eye. |  |  | January 28, 2013 | 2013 | Varian Medical Systems, Inc. Oncology Systems | The calculation of dose dynamic plans containing different energy fields is incorrect when the Pencil Beam Convolution 11.0 algorithm (PBC 11.0) is used. | Software design | Varian Medical Systems, Inc. Oncology Systems sent an Urgent Medical Device Correction Letter dated January 30, 2013 to all affected customers. The Letter identified the affected product, problem and actions to be taken. Varian Medical Systems is developing a technical correction for this issue, Its customers will be contacted by a Varian Customer Service representatives when this correction is available to schedule its installation. The letter provides customers with help desk contact information. |  | 1168 | 1168 units | Worldwide Distribution - USA Nationwide | Terminated | April 04, 2014 | 431 | 1168 | Computer
Z-0552-2014 | 66956 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Plan Parameters workspace within |  |  | November 18, 2013 | 2013 | Varian Medical Systems, Inc. | An anomaly has been traced to a software deficiency within the Plan Parameters workspace. When a dose relevant parameter is changed and "do no clear MUs and Reference Point dose" is selected, the Monitor units may change. | Software design | Varian Medical Systems will provide all affected consignees with an Urgent Medical Device Correction Letter commencing on November 18, 2013 by federal express. . The Letter will also be distributed to the Varian Sales, Marketing, and Service organizations, informing them of the reported issue. |  | 6701 | 6701 | Worldwide Distribution. | Terminated | March 13, 2015 | 480 | 6701 | Computer
Z-0830-2013 | 63783 | No | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | LPS Diaphyseal Sleeves |  |  | January 04, 2013 | 2013 | DePuy Orthopaedics, Inc. | DePuy Orthopaedics is initiating a voluntary recall of the LPS Diaphyseal sleeve because the product has the potential for fracture at the taper junction. The LPS Diaphyseal Sleeve to Diaphyseal Sleeve Base taper connection may not be sufficient to accommodate potential physiologic loads that may be transferred to the junction during normal gait activities by some patients. | Device Design | DePuy Othopaedics sent an Urgent Information Medical Device Recall Notice letter dated January 4, 2013 to all affected customers. The letter identified the affected products, reason for recall, clinical implications and actions to be taken. The letter instructed customers to check inventory, discontinue use, cease further distribution and return all inventory on hand for credit. Customers were requested to complete and return the Reconciliation Form. For product-related questions contact your local DePuy Othopaedics Sales Representative, clinical questions contact DePuy's Scientific Information Office at 1-888-554-2482 and recall notice information call 574-372-7333. |  | 162 | 162 total units (US) and 102 units total (OUS) | Worldwide Distribution - US Nationwide in the states of AL, AZ, CA, CO, FL, GA, ID, IL, MA, MD, MI, MN, NC, NM, NV, NY, OH, PA, SC, SD, TN, TX, VA, WA, WI, and WV and the countries of Holland, Australia, Austria, South Africa and Canada including Veteran Administration hospitals. | Terminated | May 20, 2014 | 501 | 162 | Computer
Z-0599-2014 | 66895 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Asante Conset Infusion Set with 60cm tubing &amp; 9mm cannula packaged in 5-packs,singles, or inside a Combination Kit (FG-5229-4); |  |  | November 19, 2013 | 2013 | Asante Solutions, Inc. | Reports of the infusion tubing disconnecting from the connector, resulting in leakage of insulin from the infusion set and providing patient with less insulin than intended. | Device Design | Asante sent urgent medical device recall notifications letters to distributors and customers on November 19, 2013 by US mail, return receipt requested. |  | 1979 | 1979 kits | Nationwide Distribution. | Terminated | April 24, 2014 | 156 | 6977 | Computer
Z-0583-2014 | 66909 | No | HWA | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker Orthopaedics Ball Impactor Tip |  |  | September 30, 2013 | 2013 | Stryker Howmedica Osteonics Corp. | Stryker has received complaints associated with cracks and/or fracture of the Ball Impactor Tip and Rim Impactor Tip Instruments. | Device Design | Stryker sent an Urgent Medical Device Correction letter dated October 1, 2013 to all affected customers via Fed Ex. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete and fax the attached Product Correction Acknowledgement form to 855-251-3635. |  | 1024 | 1,024 units | USA Nationwide Distribution | Open, Classified |  |  | 2714 | Computer
Z-0118-2014 | 64730 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Centricity PACS IW by GE Healthcare Dynamic Imaging Solutions is a device that receives medical images (including mammograms) and data from various imaging sources. Images and data can be stored, communicated, processed and displayed within the system or across computer networks at distributed locations. Lossy compressed mammographic images and digitized film screen images may only be interpreted using an FDA approved monitor that offers at least 5 M pixel resolution and meets other technical specifications reviewed and accepted by the FDA. Typical users of this system are trained professionals, including but not limited to physicians, radiologists, nurses, medical technicians an assistants. |  |  | June 15, 2012 | 2012 | Ge Healthcare It | Multiple issues were identified with the GE Centricity PACS-IW system that could pose a safety risk to patients. These issues include:1) With Auto-Fetch in use another patient's image may appear without user's intention. 2.a) In the case where dose administration time and image acquisition time fall on either side of midnight, the value will be under reported and close to zero for all pixe | Software design | GE Healthcare sent an IMPORTANT PRODUCT INFORMATION letter dated June 15, 2012 to direct accounts informing them of the Auto-Fetch functionality issue and advising them to discontinue the use of the Auto-Fetch feature and to change the system setting to not allow users privileges to use Auto-Fetch. Customers were requested to contact their Field Engineer or GE Service Representative with any questions or to call the GE Healthcare Remote Online Center Customer Care Center at 1-800-437-1171. GE Healthcare sent an URGENT MEDICAL DEVICE CORRECTION letter dated March 12, 2013 to all affected customers informing them of all three of the safety issues. The letter again adivsed customers to discontinue the use of the Auto-Fetch feature and to change the system setting to not allow users privileges to use Auto-Fetch. The letter further infomed customers that all three errors can be resolved through the installation of a new version of the software. Customers were requested to contact their Field Engineer or GE Service Representative with any questions or to call the GE Healthcare Remote Online Center Customer Care Center at 1-800-437-1171. |  | 847 | 847 units | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, WY, District of Columbia, Puerto Rico, and the U.S. Virgin Islands.Internationally to Austria, Azerbaijan, Bahrain, Belgium, Bosnia, Brazil, Canada, Columbia, Egypt, France, Germany, Ghana, Hungary, Ireland, Israel, Italy, Japan, Kuwait, Lebanon, Libya, Lithuania, Mexico, Netherlands, Nigeria, Palestine, Poland, Qatar, Russia, Saudi Arabia, Slovenia, South Africa, Spain, Switzerland, Turkey, United Arab Emirates, United Kingdom, Uzbekistan, and Yemen.Center Recommended Depth - Consumers/User** | Open, Classified |  |  | 847 | Computer
Z-0142-2014 | 66464 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | S5 Heart-lung machine, Stockert S5 System. Catalogue No: 48-30/40/50-00. |  |  | September 24, 2013 | 2013 | Sorin Group Deutschland GmbH | Sorin Group is recalling certain models of S5 Perfusion Systems due to malfunction of cardioplegia control. | Software Manufacturing/Software Deployment | Sorin Group sent an Urgent Field Safety Notice dated September 24, 2013, to all affected customers. The letter informed users of the potential for failing of automatic cardioplegia delivery function and notified them that they would be contacted by the Sorin Service Group to arrange an appointment to update the firmware on site. Customers were asked to complete the Response form to confirm they had received, read and understood the Field Notice. For questions regarding this recall call 303-467-6527. |  | 84 | 84 | Worldwide Distribution - USA including WA, NY, VA, TN, SC, TX, CA, CO, MA, WI, MD, AL, ND, and Puerto Rico and Internationally to Italy, Finland, Germany, Sweden, and Great Britain. | Terminated | July 10, 2014 | 289 | 84 | Computer
Z-0554-2014 | 66873 | No | MNI | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | PedFuse Reset Screw Inserters; PedFuse Screw Inserter 05-7 REspond Part# 11-80111 |  |  | November 20, 2013 | 2013 | SpineFrontier, Inc. | Screw inserters may not mate properly with hex interface of the screws. | Device Design | The firm, SpineFrontier, Inc., sent an "Urgent PedFuse Screw Inserter Voluntary Recall Notification" letter on November 20, 2013 to its consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed to initiate the immediate return of the recalled products and to complete a recall ACKNOWLEDGEMENT FORM via fax or email to: SpineFrontier, Inc., 500 Cummings Center, Suite 3500, Beverly MA 01915, Attn: Customer Service; Fax: 978-232-3991; email: customerservice@spinefrontier.com.Please feel free to contact at 978-232-3990 Monday thru Friday 8:30AM to 5:30PM Eastern with any questions. |  | 25 | 25 | US Distribution to states of:Texas, New York, Illinois, Ohio, Virginia, Washington D.C., and Florida.. | Open, Classified |  |  | 42 | Computer
Z-0633-2013 | 63838 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | PENTRA C200, Part Number: A11A01629. |  |  | October 10, 2012 | 2012 | Horiba Instruments, Inc dba Horiba Medical | Horiba Medical is recalling the ABX PENTRA AST CP because during the AST application validation on the Pentra C200 using an Audit Systems linearity kit, it was determined that two higher values for AST (expected at 800 U/L and 1000 U/L) were recorded at 280 and 199 U/L without alarm. | Software design | Horiba Medical sent a Field Safety Notification and Acknowledgement Letter dated November 20, 2012, to all affected customers. The letter notified the customers of the problems identified and the actions to be taken. Customers were instructed to complete and return the enclosed response form within 10 calendar days via mail, email or fax. Customers with questions were instructed to contact their local representative.For questions regarding this recall call 714-273-9254. |  | 384 | 384 units | Nationwide Distribution including CA, OK, NY, FL, AZ, NY, VA, and TX. | Terminated | September 12, 2013 | 337 | 67270 | Computer
Z-0534-2014 | 66865 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris PC unit model 8000 and 8015 with Alaris PCA module model 8120 and Auto ID barcode scanner model 8600. |  |  | December 05, 2013 | 2013 | CareFusion 303, Inc. | CareFusion is recalling the Alaris PC unit (model 8000 and 8015) because the Alaris PC unit will produce an audible alarm and the Alaris PC unit will display a System Error message. | Software design | CareFusion sent an Urgent Medical Device Recall Notification letter dated December 5, 2013, to all affected customers. The letter identifed the affected product, problem and actions to be taken. Customers were instructed to contact CareFusion Customer Advocacy at (888) 812-3266, if they experienced a SYSTEM ERROR on the Alaris PC unit. CareFusion will contact all affected customers within 60 days to initiate the scheduling process for the remediation.Customers are instructed to complete and return the enclosed Customer Response Card to expedite the corrective action process. Customers with questions were instructed to contact CareFusion Support Center at 888-562-6018, or for Technical Support call 888-812-3229. |  | 4824 | 4824 total units (4242 units in US) | Worldwide Distribution - USA (nationwide) and Internationally to Canada, Netherlands, and Australia. | Completed |  |  | 4824 | Computer
Z-0139-2014 | 66526 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Patient Cart used in conjunction with the da Vinci S surgical and SI System. |  |  | November 30, 0012 | 12 | Intuitive Surgical, Inc. | Increased number of complaints regarding fluid entering the patient card base of the da Vinci S and SI systems. | Device Design | Intuitive Surgical Inc sent an Product Notification to all affected customers via email on November 30, 2012.E-mail notifications advised users of an addendum to the User Manual for Covering the Patient Cart column that provided a recommended method for covering the da Vinci patient cart column to prevent splashing on the cart during a procedure. The addendum was attached to the e-mail and is available online.If you have any questions, please contact Customer Service at 1-800-876-1310. Outside the United States call +41 21 821 2020 or +800.0821.2020 |  | 1754 | 1754 patient carts | Worldwide Distribution - USA including Puerto Rico and the countries of Australia, Austria, Argentina, Belgium, Bulgaria, Brazil, Canada, Chile, China, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, India, Indonesia, Ireland, Israel, Japan, Kuwait, Mexico, Netherlands, Norway, Pakistan, Panama, Philippines, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom, Uruguay and Venezuela. | Terminated | November 04, 2013 | 730824 | 1754 | Computer
Z-0627-2013 | 63878 | No | FWZ | Manual operating table and accessories and manual operating chair and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Maquet Yuno Extension Table Device |  |  | November 02, 2012 | 2012 | Maquet Inc. | There is a potential issue with the screw tension of the extension device 1433.62A1 used with the YUNO table. | Device Design | Maquet sent Urgent Product Field Correction letters/Field Correction Confirmation Form dated November 2, 2012 on November 15, 2012, via Fed Ex to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed that a MAQUET authorized Service Technicians would contact them to arrange an appointment to assess their device and to replace the release lever at no cost to them at their facility. Customers were asked to completed and return the enclosed confirmation form, acknowledging their receipt and understanding of the communication, in the enclosed postage prepaid envelope, by e-mailing a scanned copy to FieldActions@maquet.com or via fax at 973-709-6508.For questions customers were instructed to call MAQUET Technical Support at 800-777-4222, Press option2, followed by option 2.For questions regarding this recall call 973-709-7660. |  | 63 | 63 devices | Worldwide Distribution - USA including AR, GA, SC, TX, TN and internationally to France, United Arab Emirates, Australia, Brazil, Canada, Switzerland, China, Germany, Spain, United Kingdom, Italy, Japan, Sweden, Thailand, Turkey, Serbia, and South Africa. | Terminated | August 21, 2013 | 292 | 63 | Computer
Z-0512-2014 | 66888 | No | MJN | Vascular clamp. | Cardiovascular | Cardiovascular | 510(k) | Stryker brand SUPER COMPLIANT 7MM X 15MM, Catalog Number: |  |  | November 04, 2013 | 2013 | Stryker Neurovascular | Products were shipped labeled with "use by" dates that exceed currently available shelf-life data. All products remain within the earliest expiration date of April 2015. | Software Manufacturing/Software Deployment | Stryker Neurovascular sent an Urgent Medical Device Recall letter dated November 4, 2013, all affected users. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately discontinue use and segregate the affected products. Customers with questions were instructed to contact via NeuroComplaints@stryker.com. Or Fax the response form to 1-800-876-4355.For questions regarding this recall call 510-413-2500. |  |  |  | Worldwide Distribution - USa (nationwide) and Internationally to:Austria, Belgium, Czech Rep, Germany, France, Great Britain, Hungary, Italy, Lithuania, Netherlands, Norway, Poland, Slovakia, Spain, Turkey, Australia, Canada, Chile, Hong Kong, India, Korea, Malaysia, Russia and Japan | Terminated | August 27, 2014 | 296 |  | Computer
Z-0536-2014 | 66963 | No | JJH | Clinical sample concentrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | MagNA Pure 96 System |  |  | November 25, 0013 | 13 | Roche Diagnostics Operations, Inc. | This issue is only relevant for customers importing xml files into the MagNA Pure 96 IVD instrument software for sample identification. There is a potential for sample mismatch when using the MagNA Pure 96 IVD instrument (catalog number 06541089001). Importing an order file in xml format to the MagNA Pure 96 IVD instrument software may result in a wrong sequence of samples in the graphical overvi | Software design | Roche Diagnostics sent an URGENT Medical Device Correction Notice dated November 25, 2013, to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers.Customers were advised to discontinue use of the Pipetting Report and to use the sample list file instead. The MagNA Pure 96 Instrument Operators Guide is being updated with an addendum addressing this issue. The letter included a Customer Response Form for customers to complete and return to the firm by FAX : 1-888-670-4135. Please contact the Roche Support Network Customer Support Center at 1-800-526-1247, 24 hours a day, seven days a week, if you have questions about the information contained in this UMDC. |  | 27 | 27 | US Distribution including the states of AL, AZ, CA, CT, GA, NY, NC, NJ, OH, TX, VA and WA. | Terminated | September 04, 2014 | 730768 | 27 | Computer
Z-0357-2014 | 66439 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Trinica Anterior Lumbar Plate (ALP) system, part 07.01058.001. Non Sterile, Rx only. The Trinica Anterior Lumbar Plate System is a temporary supplemental fixation device consisting of a variety of shapes and sizes of plates and screws. The Trinica Anterior Lumbar Plate System is used as an implant for the correction and stabilization of the spine. This system provides temporary stabilization and augments the development of a solid spinal fusion. Additionally, this system provides the surgeon with the ability to supplement an interbody device with anterior plate fixation. The Trinica Anterior Lumbar Plate System components can be locked into a variety of configurations and each construct may be customized to individual cases. The plates are low profile and anatomically designed to provide optimal fit from either anterior or anterior-lateral approach. This system also features anti-migration locking caps to help secure the fixation screws. All Trinica Anterior Lumbar Plate System implant components are made from titanium alloy (Ti-6AI-4V). |  |  | October 09, 2013 | 2013 | Zimmer, Inc. | Zimmer Spine, Inc. is initiating a correction regarding the Trinica ALP Instrument Tray, part 07.01058.001 (all lots). The Ratchet Handles, part 07.00438.001 may not be effectively sterilized if left in their designated location(s) in the instrument tray. This is due to the current bracketing in the tray potentially interfering with complete steam penetration into the cannulated handle. | Device Design | Zimmer Spine sent a Medical Device Correction letter dated October 9, 2013 via FedEx A Medical Device Correction letter, dated October 09, 2013 via FedEx to distributors, Hospitals and Surgeons who have received and/or used the device. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the notification and ensure affected personnel were aware of its contents. Only sterilize the Ratchet Handles external to the instrument tray. Customers were to be notified when replacement trays were available. Customer with questions were instructed to call the Zimmer Customer Call Center at 866-774-6368.UPDATE: July 17, 2014 Zimmer issued another letter to consignees to inform them of a redesigned tray is available and will replace the current tray. A Zimmer representative will conduct the swap-out of the current Rev A tray for the new Rev B design tray. A Return Confirmation Form is asked to be emailed back to CorporateQuality.Postmarket@zimmer.com. For assistance or questions, contact Zimmer Spine Customer Service at 866-774-6368. |  | 94 | 94 Tray sets | Worldwide Distribution - USA (nationwide) and Internationally to Nationwide and OUS countries include: Australia, Austria, Belgium, Canada, France, Germany, Italy, New Zealand, Spain, Switzerland, and the UK. | Terminated | February 24, 2015 | 503 | 94 | Computer
Z-0743-2013 | 63866 | No | CEO | Phosphorus (inorganic) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Alfa Wassermann System Diluent 750 mL |  |  | October 10, 2012 | 2012 | Alfa Wassermann, Inc. | Inorganic phosphorus assays using System Diluent Lot 155599 may not product acceptable calibration results. | Device Design | Alfa Wassermann sent an important advisory notice/product response form to their distributors on October 10, 2012. The letter identified the product, the problem, and the action to be taken by the customer. The letter informs customers about the affected lot and provide instruction regarding the receipt of replacement product. Any recalled product should be destroyed. If there is further distribution, those customers should provide Alfa with a distribution list so that they may be notified. Any questions can be directed to (866) 419 - 2532. |  | 4104 | 4,104 kits | Worldwide Distribution - USA (nationwide) and the countries of Canada and Cayman Islands | Terminated | October 25, 2013 | 380 | 4104 | Computer
Z-0426-2014 | 65985 | [6 empty] | Custom Ultrasonics System 83 Plus 2 and Plus 9 Endoscope Washer/Disinfector. |  |  | March 12, 2013 | 2013 | Custom Ultrasonics, Inc. | Certain System 83 Plus 2 and System 83 Plus 9 units running the Windows 7 operating system were not fully validated prior to distribution. | Software design | A notification letter, dated February 26, 2013, was sent to direct accounts to inform them of the recall and that technicians would visit facilities to replace the Windows 7 operating system with an appropriately validated Windows XP operating system. |  | 118 | 118 | Worldwide Distribution: USA (nationwide) and Internationally to: Canada, Slovakia Republic, Sydney and Peru. | Terminated | June 26, 2014 | 471 | 118 | Computer
Z-0292-2014 | 65572 | No | LDP | Acetaminophen test system. | Toxicology | Toxicology | 510(k) | VITROS Chemistry Products ACET Slides, 90 slides/pack; catalog number 1037274 |  |  | June 12, 2013 | 2013 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics is recalling Slide Dispense cartridges used with some VITROS Chemistry Product Slides due to slide dispense issues. | Equipment maintenance | Ortho Clinical Diagnostics sent an Important Product Correction Notification letter dated June 12, 2013, to all affected customers via FedEx regarding Slide Dispense issues with various VITROS Chemistry Product Slides. VITROS ALT, CL-, GGT, LDH, K+, Na+ and TP slides use a different ABP, they are not affected by this issue. OCD provided instructions that may allow their customers to reposition the ABP if a jam occurs. OCD will provide you with a credit for any cartridge where the ABP becomes misaligned within the cartridge body in such a way that you are unable to reposition the ABP successfully to resolve the slide dispense issue. In addition to the Confirmation of Receipt form, OCD also enclosed three (3) copies of a Request for Credit form for you to complete when you require credit for affected cartridges. Complete and return the Confirmation of Receipt by June 21, 2013. If you experience slide dispense condition codes more than once a week, call our Customer Technical Support at 1-800-421-3311.For questions regarding this recall call 585-453-4224. |  | 16732 | 16732 units (12208 units domestically &amp; 4524 internationally) | Worldwide Distribution - USA (nationwide) including AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY including Puerto Rico and Internationally to Australia, Brazil, Canada, Chile, China, Colombia, India, Japan, Mexico, Panama, Singapore, Venezuela, England, Germany, France, Italy, and Spain. | Terminated | April 11, 2017 | 1399 | 16732 | Computer
Z-0408-2014 | 66412 | No | HTW | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Slipped Capital Femoral Epiphysis System |  |  | July 31, 2013 | 2013 | Synthes USA HQ, Inc. | There is a possibility for the Guide Wires of the Slipped Capital Femoral Epiphysis System to stick in the cannulation of the drill bit and the drill bits to break during surgery. | Device Design | Synthes sent an Urgent Notice: Medical Device Recall letter dated July 30, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to review their inventory, immediately remove the affected product and return to Credit/Returns, Synthes. For questions call 610-719-5450. |  | 166 | 166 | USA Nationwide Distribution | Terminated | June 01, 2015 | 670 | 166 | Computer
Z-0425-2014 | 65356 | No | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | 13mm addEASE(TM) Binary Connector with 17 Ga. Needle for use with PAB(R) Containers, BC1300 |  |  | April 04, 2013 | 2013 | B. Braun Medical, Inc. | Possible solution leakage at the interface of the addEASE(TM) Binary Connector (Product Codes BC2000 and BC1300) and the PAB(R) bag. Based on extensive root cause investigation, it has been determined that the leakage is the result of a misaligned cannula (needle) on the addEASE(TM) Binary Connector, which pierces the PAB(R) bag stopper to access the solution. The misaligned cannula may result i | Equipment maintenance | B Braun sent an Urgent - Medical Device Removal Notification letter dated April 4, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by he customer. Customers were instructed to review the Product Recall Notification in its entirety and ensure that all users in their organization are informed about this voluntary product recall. Determine their current inventory of the affected lots within their distribution center. Contact their customers that have received the affected lots and inform them of this voluntary product recall. Combine both their inventory, and the inventory reported by their customers, and utilize their specific "Product Removal Acknowledgement" form to record the total number of individual units (within partial cases) and the number of full-unopened cases. If no inventory remaining enter zero (0). Return the completed Acknowledgement Form to BBMI Quality Assurance in the self addressed envelope or fax to 610-849-1197. Customers should call Customer Support Department at 800-227-2862 to arrange for eturn and replacement product or for questions.For questions regarding this recall call 800-227-2862. |  | 4000 | 4,000 units | Nationwide Distribution | Terminated | August 26, 2014 | 509 | 1750079 | Computer
Z-0330-2014 | 64806 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | The Centricity Laboratory System is intended to be an information system designed to support the clinical and administrative activities associated with the provision and utilization of clinical laboratory services and facilities, e.g., the storing and delivering of analytical results. It is a specially designed data program application (software), which is supplied for installation in existing mainframe or decentralized computers or a computer network. This software system will not be the software directly used to run, steer, or control any specific laboratory analyzer or equipment. It is intended for use in laboratory facilities, including a central laboratory or in a multiple laboratory environment servicing satellite laboratories, reference laboratories, clinics, etc. Centricity Laboratory System is not intended for direct patient contact. |  |  | March 12, 2013 | 2013 | GE Healthcare It | GE Healthcare has become aware of a potential safety issue associated with the use of GE Centricity Laboratory where panel reflex does not occur if DIP.Specimen is blank. Panel reflexing fails to occur when ad-hoc re-routing functionality is used and the following are true: 1) The ordered panel is a slave panel, and; 2) The panitem specimen group field of one or more of the ordered items on the ma | Software design | GE Healtchare sent an Urgent Medical Device Correction letter dated March 12, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The letter requested that customers review their workflows to determine whether this series of events could happen at their sites. Customers who self-identified this series of events as a possible workflow are recommended to contact their GE service representative for a software correction. These customers may also elect to identify master panels that may be re-routed ad hoc and then to ensure that the specimen group field is updated to NULL if currently set to blank.Customers with any questions about this issue are advised to contact their GE service representative or the GE Healthcare IT Help Desk: North America at tgg.nahelpdesk@med.ge.com.For questions regarding this recall call 847-277-4468. |  | 7 | 7 units | Worldwide Distribution - USA (nationwide) including GA and Internationally to Canada and the United Kingdom. | Open, Classified |  |  | 7 | Computer
Z-0742-2013 | 63800 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Mindray DPM5 Monitor, a vital signs monitor used on human patients in healthcare facilities. |  |  | September 05, 2012 | 2012 | Mindray DS USA, Inc. d.b.a. Mindray North America | Mindray has identified an issue with the DPM5 monitors where ST and Arrhythmia features may not be available. | Device Design | Mindray DS USA Inc. sent an Urgent Product Corrective Action letter on September 5, 2012 via Fed Ex with delivery confirmation to all users. The letter identifies the affected product, problem and actions to be taken. Customers were instructed to contact their Mindray Service Representative at 1-800-288-2121 Monday through Friday 8:30-5:30 pm, to schedule a software upgrade. |  | 13 | 13 units | US Nationwide Distribution including the states of: VA, TN, FL, &amp; TX | Terminated | October 22, 2013 | 412 | 13 | Computer
Z-0333-2014 | 66635 | No | HJO | AC-powered slitlamp biomicroscope. | Ophthalmic | Ophthalmic | 510(k) | Zeiss brand IOLMaster, Version 1-5, Biomicroscope, Slit-Lamp, AC-Powered, |  |  | October 23, 2013 | 2013 | Carl Zeiss Meditec, Inc. | The user database delivered for download contained an erroneous value. Specifically the "aO" constant of the lens AMO Sensar AR40E is listed as "0.0" in the internal lens database of the IOLMaster. The correct value is -2.420. This could lead to the implantation of lens with a wrong diopter value. | Software in the Use Environment | The firm, Zeiss, sent an "Urgent Medical Device Correction" letter dated October 23, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customers. Customers were instructed to immediately update the Laser Interference Biometry (ULIB) lens constants, provided as an Attachment 1 and complete and return Attachment 2- Confirmation Form via mail to: Stericycle, Inc., 6026 Lakeside Blvd., Indianapolis, IN 46278; Attn: Zeiss IOLMaster.If you have any questions, call the following number for assistance: 1 (866) 891-1922. |  | 4907 | 4,907 units | Nationwide distribution. | Terminated | April 17, 2014 | 176 | 6108 | Computer
Z-0430-2014 | 65408 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension Vista(R) 500 and Dimension Vista(R) 1500 - Reloaded Vista Vials - limited to systems using software versions 3.4 and 3.5.1. |  |  | May 20, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | There is the potential for the vial volume of "in use" Vista calibrators, controls, or sample diluent to incorrectly revert to full volume when removed and reloaded. Dimension Vista(R) 500 and Dimension Vista(R) 1500 - Reloaded Vista Vials - limited to systems using software versions 3.4 and 3.5.1 | Software design | A notification letter was sent to consignees in May 2013. |  | 776 | 776 | Worldwide Distribution: USA (nationwide) and iinternationally to Canada. | Terminated | April 10, 2014 | 325 | 776 | Computer
Z-0431-2014 | 66705 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Easy Diagnost Eleva |  |  | October 30, 2013 | 2013 | Philips Medical Systems, Inc. | The delivered application software was an outdated version to the model of equipment. | Software design | Philips Medical Systems sent an Urgent - Field Safety Notice dated October 30, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Cusromers were informed that the delivered application software was an outdated version to the model of their expipment and that a Philips Field Engineer will be implementing an upgrade of the application software. Customers were instructed to review the information contained in the letter with all memebers of their staff who need to be aware of its contents. Customers with questions were instructed to contact their Philips representataive: Technical Support Line 1-800-722-9377.For questions regarding this recall call 1-800-722-9377. |  | 19 | 19 devices | Nationwide Distribution including MN, NH, DE, OR, MA, TX, OH, FL, CA, WA, and SC. | Terminated | January 09, 2017 | 1167 | 19 | Computer
Z-0791-2013 | 64160 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Medacta Quadra Trial Broach size 0 part # 01.10.10.045, for use with the Medacta Total Hip Prosthesis System. |  |  | December 17, 2012 | 2012 | Medacta Usa Inc | The tip of the Quadra Trial Broach size 0, broke during a surgery. | Device Design | Medacta sent a Urgent Safety Notification letter dated December 13, 2012 to all affected customers. The letter identified the recalled product, problem and actions to be taken. Customers were instructed to check inventory for affected product and contact QA@medacta.us.com for an RMA# to return items to Medacta USA. A new version of the recalled product will be provided to all affected customers. For questions call 805-437-7085 x26. |  | 26 | 26 pieces U.S.A. (226 pieces Worldwide) | US Nationwide Distribution including the states of Arizona, California, Idaho, Illinois, North Carolina, Ohio, Utah, Washington and US Virgin Islands | Terminated | February 07, 2013 | 52 | 26 | Computer
Z-0438-2014 | 66724 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | SIEMENS brand ARTISTE Linear Accelerator Linac systems with component: any photon unflat beam option (Multiple X feature) in combination with the IMRT or mARC option. A family of linear accelerator systems is to deliver X-ray photon and electron radiation for the therapeutic treatment of cancer. The ARTISTE/ONCOR systems are manufactured by: SIEMENS AG, MEDICAL SOLUTIONS, GERMANY and distributed SIEMENS Healthcare IM CR RO, CR Radiation Oncology, Martinez, CA |  |  | October 23, 2013 | 2013 | Siemens Medical Solutions USA, Inc | A safety risk regarding a potentially existing dark current radiation phenomenon on the Linear Accelerator (LINAC) in combination with IMRT or mARC (rotational IMRT) treatments using unflat (flattening filter free) beams. In case there are many long periods of field shape changes in a treatment plan, the dose due to dark current radiation may sum up to a significant value. | Device Design | Siemens sent an Urgent Medical Device Correction letter dated October 23, 2013 including the release of update instruction TH024/13/S to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. |  | 26 | 26 units | Worldwide Distribution: USA (nationwide) including states of: FL, NC, NY, PA, TN, UT, and WI; and countries of: Angola, Brazil, Canada, P.R. China, Germany, India, Japan, Republic Korea, Poland, and Thailand. | Terminated | January 22, 2014 | 91 | 29 | Computer
Z-0335-2014 | 66601 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance iCT Computed Tomography X-Ray System, |  |  | October 11, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | If the couch, with the footrest extension attached, is positioned where the table travel is within the bore of the gantry, and the user begins to lower the couch, the footrest extension can contact the gantry cover or come in contact with the user due to the length of the accessory. | Device Design | Philips sent an Urgent - Medical Device Correction letter dated October. 11, 2013, to their customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to discontinue the use of the footrest extension. Philips also, stated they will replace the old footrest extension with the current one in use. For questions contact your local Philips representative or local Philips Healthcare office. |  | 9 | 9 units | Worldwide distribution: USA (nationwide) Distribution in the states of DC, GA, IL, IN, NY and OH and the countries: Germany, Israel and Netherlands. | Terminated | August 04, 2015 | 662 | 9 | Computer
Z-0439-2014 | 66764 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | C-Series: Clinac, Trilogy, Trilogy Tx. , Novalis high energy linear accelerators and UNIQUE single energy linear accelerator; Versions 7, 8 and 9. The UNIQUE is not sold in the US. |  |  | October 24, 2013 | 2013 | Varian Medical Systems, Inc. | This correction is to notify users that a solution to a previous correction has been developed and Varians reps will be contacting locations to schedule installation. | Device Design | Varian sent an Urgent Medical Device Correction letter, P/N CP-05645 Rev D dated October 24, 2013 to their affected customers. The letter identified the affected product, problem and actions to be taken. The letter stated all affected sites will be contacted by Varian representatives to schedule an upgrade to their devices. Questions may be directed to 888-827-4265 for USA and Canada. Europe +41 41 749 8844 |  | 3809 | 3,809 devices | Worldwide Distribution - USA Nationwide and in the countries of Algeria, Angola, Argentina, Australia, Austria, Azerbaijan, Bangladesh, Belgium, Belorussia, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, Estonia, Finland, France, French Polynesia, Georgia, Germany, Greece, Guadeloupe, Guam, Guatemala, Hong Kong, Hungary, India, Indonesia, Iraq, Ireland, Israel, Italy, Japan, Kazakhstan, Kenya, Kuwait, Latvia, Lebanon, Uthuania, luxembourg, Macau, Macedonia, Malaysia, Martinique, Mauritania, Mexico, Moldova, Monaco, Morocco, Netherlands, New Zealand, Norway, Oman, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Puerto Rico, Qatar, Reunion, Romania, Russia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, SriLanka, Sweden, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Venezuela, Vietnam, and Zimbabwe. | Terminated | January 13, 2016 | 811 | 3809 | Computer
Z-0436-2014 | 66866 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Assembly, Patient Side Manipulator on da Vinci Surgical Systems: |  |  | November 11, 2013 | 2013 | Intuitive Surgical, Inc. | Reports of friction within certain instrument arms can interrupt smooth instrument motion. This can be felt by the surgeon as resistance in the movement of the master.In this situation, the instrument can stall momentarily and then suddenly catch-up to the master position if the surgeon pushes through the resistance. | Device Design | The firm, Intuitive Surgical, sent an "Urgent Medical Device Recall" letter dated November 11, 2013 all affected customers via traceable mail. The letter described the product, problem and actions to be taken. The customers were instructed to take the following actions: 1. Ensure all appropriate personnel are fully informed of this Notification. Forward this letter to all Surgeons and OR staff who perform da Vinci surgical procedures, Risk Manager, OR Director, Purchasing, and Biomedical Engineering staff. 2. Complete the attached Acknowledgement Form and return it to Intuitive Surgical via Fax to: Intuitive Surgical, Inc., ATTN: REGULATORY AFFAIRSU.S. +1 (408) 716-3040, or Scan and Email: isi.compliance@intusurg.com.3. An Intuitive Surgical Representative will contact your site to schedule the inspection of the instrument arms on your system(s) as indicated above. Please work with the Intuitive Surgical Representative to schedule this activity.4. Please inform appropriate personnel when the correction has been completed.5. Please retain a copy of this letter and the Acknowledgement Form for your files.If you need further information or support concerning this medical device correction, then please contact your Clinical Sales Representative or contact Intuitive Surgical Customer Service at the numbers listed below. Âż North and South America: 800-876-1310 Option 3 (6 am to 5pm PST) Âż Japan: 0120-56-5635 or 03-5575-1362 (9am to 6pm JST) Âż South Korea: 02-3271-3200 (9am to 6pm KSTJ) Âż Europe, Middle East, Asia and Africa: +800 0821 2020 or +41 21 821 2020 (8 am to 6 pm CET). |  | 1386 | 1,386 | Worldwide Distribution: US (nationwide) and to countries of: Australia, Austria, Belgium, Brazil, Bulgaria, Canada, China, Columbia, Denmark, France, Germany, India, Israel, Italy, Japan, S. Korea, Luxembourg, Mexico, Monaco, Netherlands, Norway, Romania. Russian Federation, Singapore, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, and United Kingdom. | Terminated | February 18, 2014 | 99 | 1386 | Computer
Z-0434-2014 | 66502 | No | FRG | Sterilization wrap. | General Hospital | General Hospital | 510(k) | STERITITE Containers: STERITITE Container Kit, 1/4 size, 3" high, product code: 13902; STERITITE Mini Mid Size, 3" high, product code: 13903; STERITITE Container Kit, 1/2 size, 4" high, product code: 13904; STERITITE Container Kit, 3/4 size, 4" high, product code: 13905; STERITITE Container Kit, Full size, 4" high, product code: 13906. |  |  | October 03, 2013 | 2013 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) is recalling the SteriTite containers because they do not have consistent data to show that sterility assurance level of 10(-6) is achieved during the half-cycle test runs on new, unprocessed containers. | Device Design | Advanced Sterilization Products (ASP) sent a Urgent Medical Device Field Safety Notification letter dated October 3, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers who received the SteriTite Containers.Customers were instructed to:Provide this notice to anyone in your facility that needs to be informed.Maintain a copy of this notice with the affected product.Maintain awareness of this notice.Customers are instructed to contact Stericycle at (888) 912-7092 for additional questions related to the recall letter. Customers are instructed to report any complaints or suspected problems related to the recall by contacting ASP at (888) STERRAD (888-783-7723). |  | 267 | 267 units total (252 units in US) | Worldwide Distribution - USA (nationwide) and the countries of Australia, Vietnam, Thailand, Malaysia, China, Korea (South), Mexico, Spain, Singapore, Hong Kong and Portugal. | Terminated | January 13, 2014 | 102 | 267 | Computer
Z-0274-2014 | 65760 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | S-LIFT Extension Shims |  |  | August 28, 2013 | 2013 | SpineFrontier, Inc. | Fit variation among all lots of S-LIFT extension shims with corresponding retractor blades. The firm's investigation revealed that extension shims may not fit down all retractor blades. | Device Design | SpineFrontier sent an Urgent Voluntary Recall letter dated August 28, 2013, to all affected customers. The letter identified teh product, the problem, and the action to be taken by the customer. Customers were asked to initiate the immediate return of the parts noted and complete an acknowledgement form and return it to SpineFrontier via fax or email. Customers with questions were instructed to call 978-232-3990. For questions regarding this recall call 978-232-3990. |  | 67 | 67 | Nationwide Distribution including FL, TX, OK, MO, IL, OH, VA, OR, CO, KS, CT | Open, Classified |  |  | 67 | Computer
Z-0433-2014 | 65147 | No | NSX |  |  | General Hospital | Contact ODE | syngo(R) Lab Data Manager - Version VA11B and VA12A Systems |  |  | May 10, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Under certain conditions the system may not perform as intended, causing the release of results to the Laboratory Information System that should have been held for manual review due to auto-verification rules or the delay/omission of result transmission to the LIS. | Software design | Siemens sent an Urgent Medical Device Correction letter in May 2013 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the letter in its entirety and ensure that all operators understand the information presented. Specific actions needed to be implemented by the customer were outlined in the letter. Customers were instructed to retain the letter with their laboratory records, and forward the letter to those who may have received the affected product. Customers with questions were instructed to contact their Siemens Technical Solutions Center or their local Siemens technical support representative.For questions regarding this recall call 800-441-9250. |  | 262 | 262 | Worldwide Distribution: USA (nationwide) and internationally to Canada. | Terminated | August 26, 2014 | 473 | 262 | Computer
Z-0172-2014 | 65795 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Artis zee systems |  |  | January 21, 2013 | 2013 | Siemens Medical Solutions USA, Inc | The firm became aware of a potential issue with Artis zee systems bearing model numbers 10094135, 10094137, and 10094141 equipped with a card collimator from a specific lot where a dislocation of the scatter radiation filter inside the card collimator may occur. | Device Design | Siemens sent an letter to all affected customers in January 2013, notifying them of the potential issue and urging customers to make appointments with Siemens to have software update AX 027/12/S realized in order to resolve the problem. The software update was made available from January 22, 2013, on. Customers were also asked to pass along this information if the device was no longer in their possession.The Center for Devices and Radiological Health (CDRH) understands that the CAP involving 17 unit(s) will consist of the following:1. The firm sent out a customer notification letter in January 2013 notifying them of the issue and highlighting the approved procedures of the Corrective Action Plan.2. Siemens directed in the customer notification letter, for the end-user to make an appointment with the Customer Service organization to have safety update AX 027/12/S realized.3. As stated in the safety update AX 027/12/S and in the e-mail response the update to correct this issue will be provided to the customer at no expense.4. Siemens indicated in the customer notification letter that the AX 027/12/S update will be made available from January 22, 2013 and on.5. The AX 027/12/S update provides general information, work instructions and a completion certificate on how to correct the dislocated scatter radiation filter inside the collimator.You may proceed with implementation of the CAP (subject to the conditions noted above). For further questions please call (610) 219-6300. |  | 17 | 17 | US Distribution including the states of AZ, KY, PA, IA WV, FL, MO, OK, CA, AR, MT, NY, KS, and CT. | Terminated | October 29, 2014 | 646 | 17 | Computer
Z-0140-2014 | 66314 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Siemens RAD Fluoro Uro system (Luminos dRF and Uroskop Omnia with software version VD10A/G |  |  | September 03, 2013 | 2013 | Siemens Medical Solutions USA, Inc | When using systems operating with software versions VD10A/G during a RAD examination, a malfunction can result in radiation continuing longer than required resulting in an overexposed image that is not of diagnostic quality, prompting repetition of the examination. | Software design | Siemens sent an Field Safety Notice dated September 3, 2013, to all affected customers.The letter identified the product the problem and the action to be taken by the customer. Action to be taken: 1. It is not possible for the user to prevent this malfunction. The exposure time with the use of automatic exposure control is limited to 500ms, which is the maximum duration of radiation in this mode. The exposure can be interrupted by the user at anytime by releasing the exposure button.2. There are two possibilities to solve this issue:a.) Switch off the system if the message "wait for FLC" in status bar does not disappear approximately. 30s after the OGP selection. The system gets ready after switching the system on and reselecting the Torno OGP.b.) Press Shift+F10. The pop up message has to be confirmed and you have to wait approximately.40s until! the "i" icon in the status bar stops blinking. After that a RAD OGP has to be selected and then the Torno OGP. Thereafter, the system gets ready.3. It is not possible for the user to prevent this malfunctionThe Center for Devices and Radiological Health (CDRH) understands that yourCorrective Action Plan (CAP) will involve 59 units and consist of the following elements:1. You have notified affected customers with Field Safety Notice XP032/13/S.2. You have prepared a field modification of the Luminos dRF and Uroskop Omniasystems with software version VD10A/G.3. Distribution of this field modification will began in September 2013 as UpdateInstruction XP015/13/S at no cost to the user.For further Customer Support call (610) 219-6300. |  | 59 | 59 | Worldwide Distribution - USA (nationwide) and Internationally to Canada. | Terminated | March 09, 2015 | 552 | 59 | Computer
Z-0346-2014 | 66069 | No | JGS | Sodium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA 2400 System Software V4.01. |  |  | September 03, 2013 | 2013 | Siemens Healthcare Diagnostics | Siemens Healthcare Diagnostics issued a Urgent Medical Device Correction due to Ration Parameters Flagging Behavior for All Software Versions utilized on ADVIA 1200, 1650, 1800 or 2400 Chemistry Systems. | Software design | On September 3, 2013, Siemens Healthcare Diagnostics issued an Urgent Field Safety Notice for the ADVIA 1200, 1650, 1800 and 2400 Chemistry Systems due to a software issue. The software issue has caused the Chemistry Systems to generate reports without error flags. Notices were sent to domestic customers via U.S. and outside the US the distribution is determined at the country level. The customers were instructed to contact Siemens Healthcare Diagnostics Technical Solution Center if they have technical questions regarding this issue at 877-229-3711 (prompt 18). |  | 876 | 876 units | Worldwide Distribution-USA (nationwide) and the countries of Sweden, Switzerland, Taiwan, Turkey, United Kingdom, and Vietnam. | Open, Classified |  |  | 3759 | Computer
Z-0659-2013 | 63559 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | The TM Ardis Interbody System Inserter Stainless Steel, Rubber, Silicone Non Sterile Rx only. |  |  | September 20, 2012 | 2012 | Zimmer, Inc. | Zimmer Spine is has received complaints of difficulty in turning the implant release knob of the TM Ardis Inserter after final positioning of the TM Ardis interbody spacer implant. There has been two reports where the inserter instrument was inadvertently removed from patient with the implant still engaged. The correction includes revised directions explaining the technique to remove the implant | Device Design | Zimmer sent a "MEDICAL DEVICE CORRECTION" letter dated September 26, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. A Risk Manager Confirmation Form was attached for customers to complete and return. Questions and concerns should be addressed to Zimmer Spine Customer Service at 866-774-6368.When re-design is approved, Zimmer will remove and replace all affected devices. |  | 170 | 170 | USA nationwide; Europe, Sweden | Terminated | April 23, 2014 | 580 | 170 | Computer
Z-0337-2014 | 64805 | No | HGM | Perinatal monitoring system and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Centricity Perinatal and Centricity Intensive Care is intended to be used in clinical departments healthcare delivery systems such as a Labor and Delivery, Postpartum Maternal Care, newborn Nursery, critical care units including Neonatal Intensive Care Units (NICU), and may also be used in physicians offices and outpatient clinics. This product is primarily intended to serve the purpose of electronic documentation of clinical data and is designed to accept, transfer, display, calculate, store and manage clinical data. System capabilities provide the user with the ability to acquire data from the medical devices and to document, annotate, display, store, print, export and retrieve patient clinical practice information. This product is intended for professional use only. |  |  | March 11, 2013 | 2013 | GE Healthcare It | 1) Under rare network conditions, a Centricity Perinatal process may maintain an incorrect count of Patient slots in use which may lead to an inability to access the patient roster. The user is temporarily unable to review or edit form and chart information. This could result in a potential delay of care; however active fetal strips and permanent patient data storage are not affected. 2) When | Software design | GE Healthcare sent an URGENT MEDICAL DEVICE CORRECTION letter dated March 11, 2013, to all affected customers. The letters requested that direct accounts ensure that all potential users of the software were made aware of the potential problems and the recommended actions. Direct accounts were asked to return a customer response form and to return it to GE Healthcare via fax at 847-939-1522 or e-mail at QARAFMl1@ge.com.Version 6.94 of the software has been modified to correct these defects.GE Technical Support: 800-433-2009. |  | 1132 | 1,132 units | Worldwide Distribution - USA (nationwide) including Puerto Rico and the countries of: Australia, Canada, Denmark, Germany, Hong Kong, Korea, South Korea, Thailand and United Kingdom. | Open, Classified |  |  | 1132 | Computer
Z-0370-2014 | 66056 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Medcare Products Max-500 Stand, Total Support, patient stand, model #400006. The responsible firm name listed in the Operations Manual is Medcare Products, Burnsville, MN. |  |  | October 30, 2013 | 2013 | Ergosafe Products LLC DBA Prism Medical Services USA | The lift might fail with the labeled 500-lb. load rating. | Device Design | The firm, Medcare Products, sent an "URGENT FIELD CORRECTIVE ACTIONS" letter dated October 21, 2013 via certified mail on October 30, 2013 to its customers. The letter described the product, problem and action to be taken. The customers were instructed to peel off the boom stickers and serial number label and replace with the new ones supplied; review the revised Manual and dispose of the original; and complete the action report form once the products have been relabeled. Please refer to your Operations Manual or visit our website at www.medcarelifts.com to view the manual online.If we can answer any questions that you may have regarding the testing standards or the relabeling of the products, please do not hesitate to contact us at 1-800-695-4479. |  | 97 | 97 units | Nationwide distribution. | Terminated | May 17, 2017 | 1295 | 330 | Computer
Z-0767-2013 | 64127 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Phillips Easy Diagnost Eleva DRF. Model Number: 706034 |  |  | January 16, 2013 | 2013 | Philips Healthcare Inc. | The mirror icon on the bottom of each image could be misinterpreted as a RIGHT lead marker for side indication if the image is mirrored within a PACS system. | Software design (manufacturing process) | Phillips issued an Urgent Field Safety Notice to all customers on January 14, 2013, to notify users that the Phillips EasyDiagnost Eleva systems, release 3.0 has a mirror icon that could be misinterpereted and could potentially lead to misdiagnosis. Philips will upgrade the system to version 3.1 which involves software and hardware upgrades of the system. This software provides a replacement icon that reduces the risk of the potential misinterpretation of the original mirror icon. Additionally, Philips will provide updated Instructions for Use (IFU) for the Eleva workspot and the EasyDiagnost 3.1 system and an addendum to the EasyDiagnost system IFU. Users will be contacted by Phillips to schedule the installation. For questions contact your local Philips representative: 1-800-722-9377. |  | 14 | 14 units | US Nationwide Distribution including the states of FL, CA, WA, AZ, NC, KY, DE, AZ and MO. | Open, Classified |  |  | 14 | Computer
Z-0385-2014 | 66605 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ACCU-CHEK Aviva Combo Blood Glucose Meter (as a part of the ACCU-CHEK Aviva Combo System)\ |  |  | October 25, 2013 | 2013 | Roche Diagnostics Operations, Inc. | There is a software synchronization issue with the ACCU-CHEK Combo system. In rare cases, when the Manual Pump option is chosen on the ACCU-CHEK Aviva Combo meter, there is a possibility of receiving an incorrect bolus advice recommendation that may cause a temporary under delivery of insulin. | Software design | The firm, Roche Diagnostics, sent an "URGENT MEDICAL DEVICE CORRECTION (UMDC)" letter dated October 29, 2013, via UPS to their consignees/customers. The letter described the product, problem and actions to be taken. The UMDC informed them of the software synchronization issue with the ACCU-CHEK Combo system that may cause a temporary under delivery of insulin. The customers were instructed to ensure they properly deliver the correction bolus when utilizing the Manual Pump option, as instructed in the Owners Booklet; and complete and return the Business Reply Card via fax to 1-877-282-0223. The Distributors were instructed to provide the UMDC to their direct customers. The HCPs were instructed to remind their patients to follow the instructions in the Accu-Chek Aviva Combo Meter Advanced Owners Booklet when utilizing Manual Pump option.Questions about the information contained in this UMDC are directed to ACCU-CHEK Pump Support, 24 hours a day, seven days a week at 1-800-688-4578. |  | 6595 | 6,595 units | Nationwide Distribution. | Terminated | January 14, 2015 | 446 | 6595 | Computer
Z-0368-2014 | 66662 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | EPIQ 7 Ultrasound System, Model 795200, 795201, 795202. |  |  | October 17, 2013 | 2013 | Philips Medical Systems North America Inc. | A problem has been detected with Philips Model EPIQ 7 Ultrasound System, software version 1.0 and 1.0.1. Using the International System Date Format may result in calculation errors in Gestational Age Gestational Age (GA) and Established Due Date (EDD). | Software design | Affected customers will be notified via the Urgent-Field Safety Notice (FSN79500286), Philips Model EPIQ 7 Ultrasound system, dated 2013 October. Customers are informed that using the International System Date Format in the Philips Model EPIQ 7 Ultrasound System, software version 1.0 and 1.0.1. may result in calculation errors in Gestational Age and Established Due Date (EDD). Philips will upgrade customer systems to software version 1.1, which corrects the issue, for a free of charge. |  | 306 | 306 units total (206 international; 100 in the US). | Worldwide Distribution-USA (nationwide) including the states of CT, HI, IL, IN, KY, MA, ME, MI, NC, OR, PA, TX, and WA, and the countries of Australia, Austria, Belgium, Chile, Denmark, Finland, France, Germany, India, Ireland, Republic of Korea, Latvia, Netherlands, New Caledonia, New Zealand, Norway, Poland, Portugal, Romania, Saudi Arabia, Slovakia, Spain, Sweden, Switzerland, and United Kingdom. | Terminated | July 16, 2014 | 272 | 306 | Computer
Z-0741-2013 | 64117 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | HAMILTON-T1 Ventilator with software versions 1.1.2 or lower. |  |  | December 19, 2012 | 2012 | Hamilton Medical Inc | Analysis of a customer complaint has shown that during ventilation of small pediatric patients with high airway resistance and low lung compliance, the oxygen consumption of a Hamilton T1 ventilator with software versions 1.1.2 and lower must be calculated using a larger margin than originally expected. | Software design | Hamilton sent a Field Safety alert letter dated January 11, 2013 to all customers. The letter identified the affected product, explained the potential problem and a included Medical Device Safety Alert and Corrective Action document. Hamilton is working on a new software version for this issue. Customers are asked to update their operator manual with the update oxygen consumption formula page which provided revised information for calculating oxygen capacity. When the new software is available, Hamilton will provide loaner devices to each location while their unit is being upgraded.Questions should be addressed to Hamilton Medical at 800-426-6331 ext 215 |  | 6693 | 97 total ventilators; 69 of part number 161005 and 28 of part number 161006 | US Nationwide Distribution including the country of Switzerland. | Terminated | April 22, 2014 | 489 | 6693 | Computer
Z-0359-2014 | 66300 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | BIOMET Spine, 5.5 TI 5.5x50mm Translation Screw Polaris 5.5 TI Spinal System. Rx Only. Non-Sterile. |  |  | August 22, 2013 | 2013 | Ebi, Llc | Polaris Translation Screw Failures were recently reported to Biomet Spine. These failures are generally described as dissociation of the screw head from the shaft of the Polaris Translation Screw. | Device Design | Urgent Product Information notices dated 8/22/2013 were mailed via FedEx to distributors on 8/23/2013. Follow-up emails with a copy of the UPI notice were sent to distributors on 8/23/2013. The UPI notice informed the customers of the issue with the product and how to identify the affected product. Customers were advised that the Polaris Translation Screw System and the SpF Implantable Spine Fusion Stimulator not be used in conjunction with any implants containing cobalt chrome. Firm contacts were provided in case customers had any questions. EBI, LLC had a conference call with their domestic distributors on 8/23/2013 informing them they are required to notify their customers. |  | 7092 | 7,092 units | Nationwide distribution including Puerto Rico. | Terminated | August 31, 2016 | 1105 | 7092 | Computer
Z-0363-2014 | 66137 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Model FL14E Rose |  |  | October 23, 2013 | 2013 | Stryker Medical Division of Stryker Corporation | It was reported that the siderail may appear latched when it is not, due to lack of proper preventive maintenance. This might result in bruises and/or lacerations to a patient who is leaning against the siderails if the siderail releases. Additionally, the siderail may not effectively constrain a patient to prevent them from rolling or falling off the bed. | Equipment maintenance | Stryker Medical sent a Medical Device Notification letter dated October 23, 2013, to all affected consignees. The letter described the recalled product, potential injuries associated with the side rail sticking, as well as providing preventive measures to be taken to mitigate the risk. Consignees were directed to locate the beds listed in this notice, ensure that the siderails have been inspected and adjusted, as required, continue to routinely inspect the siderails and functionality using the enclosed, enhanced product maintenance information. Problems and concerns should be addressed to Stryker at (1-800-327-0770, option 2). Consignees were asked to Fax (269-488-8691) or e-mail (ProductFieldAction@stryker.com ) the enclosed return form and directed to forward the notice if product was further distributed and advise Stryker of the beds new location.For questions regarding this recall call 269-324-6609. |  | 24575 | 4126 units USA; 20449 inits Internationally | Worldwide Distribution - USA (nationwide) and Internationally to Canada; Japan, South America, Latin America and China. | Terminated | May 16, 2014 | 205 | 24575 | Computer
Z-0364-2014 | 66639 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Juggerknotless Drill Bit, REF 110003173, |  |  | October 10, 2013 | 2013 | Biomet, Inc. | Investigation identified that the design of the drill bit may make the suture anchor difficult to insert into the bone. The implant may skip across the hole or miss the hole during insertion. | Device Design | Biomet sent an Urgent Medical Device Recall Notice letter dated October 10, 2013 to all consignees. The letter identified the affected product, reason for recall, and actions to be taken. The letter stated customers will be contacted on an individual basis in an attempt to reconcile the affected product. Customers were directed to locate and remove implicated devices and fax a copy of the response form to 574-372-1683 prior to returning product. For questions call 574-372-1570. |  | 25 | 25 | USA Nationwide Distribution in the states of VA, CA, Ky, and IN | Terminated | July 16, 2014 | 279 | 25 | Computer
Z-0361-2014 | 66534 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Aesculap Arcadius XP L Trial Insertion Instrument (ME020R/ME020R-US) |  |  | September 20, 2013 | 2013 | Aesculap, Inc. | The trial insertion instrument may disassemble during use due to a potentially weak weld on then trial inserter shaft block area located near the thumb wheel. | Device Design | An Important Product Removal Notification dated September 20, 2013 was sent to customers. |  | 30 | 30 | Worldwide distribution: USA (nationwide) GA, TN, CA, WY, NV, TX, OK, OH, and IL; and internationally to Canada. | Terminated | December 16, 2014 | 452 | 30 | Computer
Z-0388-2014 | 66420 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Synthes Tensioning Device |  |  | July 31, 2013 | 2013 | Synthes USA HQ, Inc. | All lots of the Synthes Tensioning Device were recalled due to the potential for the reaction fork to come loose from the main body of the device and not allow the wire to release, nor the device to apply tension. | Device Design | Synthes sent an Urgent: Medical Device Recall letter dated July 31, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer."Please remove this product from your inventory and return it to Synthes. A replacement product will be sent to you when available. If you DO have any of the identified devices, please take the following steps:"Call Synthes at 1-800-479-6329 to obtain a Return Authorization Number."Complete the Verification Section at the end of this letter by checking the appropriate box indicating affected product has been located. Also, please indicate the number of devices found and note the Return Authorization Number. Please include your name, title, telephone number and signature in the spaces provided."Return the Verification Form (page 3 of this letter) with the product to:" Credit/Returns, Synthes, 1101 Synthes Avenue, Monument, CO 80132.If you DO NOT have the identified product, please take the following steps:"Complete the attached Verification Section at the end of this letter by checking the appropriate box indicating that no affected product has been located. Please include your name, title, telephone number and signature in the spaces provided. This return documentation acknowledges your receipt of medical device removal information."Return the document to Synthes by: Fax: 877-272-3517"Email: FieldAction@synthes.comIf you have any questions, please call 610-719-5450 or email Fieldaction@synthes.comWe apologize for any inconvenience this may cause and thank you in advance for your attention and cooperation. |  | 369 | 369 | Nationwide Distribution. | Terminated | June 01, 2015 | 670 | 369 | Computer
Z-0108-2014 | 66166 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Siemens RAPIDLabÂż 1265 system blood gas analyzer |  |  | September 17, 2012 | 2012 | Siemens Healthcare Diagnostics Inc | RAPIDPointÂż 400 Series, RAPIDPointÂż 500, orRAPIDLabÂż 1200 Series blood gas analyzer, users can inadvertently enter a wrong patient ID on the demographics screen, which ultimately could lead to a misidentification of the sample | Software Manufacturing/Software Deployment | Siemens Healthcare sent an Urgent Field Safety Notification letter dated September 2013 via Federal Express, to affected Siemens domestic customers. Effected Siemens customers outside the US was provided a copy of the Urgent Field Safety Notice via hard copy, e-mail, and/or fax as determined by each countries local regulations and procedures. Previously Siemens Healthcare had issued a Customer Notification on August 23, 2012. The letter identified the affected product, issue, information to mitigate the issue and actions to be taken. The firm will issue an upgrades software to users. For questions contact your local Customer Care Center - Technical Support Representative at 1-877-229-3711. |  | 2939 | 2939 units | Worldwide Distribution - USA Nationwide and in the countries of Algeria, Argentina, Australia, Austria, Bahrain, Bangladesh, Belgium, Bosnia Herzeg., Botswana, Brazil, Brunei, Bulgaria, Chile, Colombia, Croatia, Czech Republic, Denmark, Egypt, Estonia, Finland, France, Fren.Polynesia, Germany, Greece, Hong Kong, Hungary, Iceland, India, Indonesia, Iraq, Ireland, Israel, Italy, Ivory Coast, Japan, Kazakhstan, Kenya, Kuwait, Latvia, Lithuania, Luxembourg, Macedonia, Malaysia, Mayotte, Netherlands, New Caledonia, New Zealand, Norway, P.R. China, Philippines, Poland, Portugal, Qatar, Rep. of Yemen, Republic Korea, Romania, Russian Fed., Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, Turkmenistan, U.A.E., United Kingdom, Vatikancity, Venezuela, Vietnam, | Terminated | April 01, 2015 | 926 | 11006 | Computer
Z-0362-2014 | 66609 | component, external, limb, ankle/foot | ISH | External limb prosthetic component. | Physical Medicine | Physical Medicine | 510(K) Exempt | Fixed Offset Adapter, A-642100. |  |  | October 24, 2013 | 2013 | Ossur North America Inc | Ossur inititated a voluntary recall of the A-642100 Fixed Offset Adapter because it can develop fatigue cracks in normal use over an extended period of time and ultimately separate completely. | Device Design | Ossur sent an "Safety Notice Voluntary Product Recall" letter dated October 29, 2013, via regular mail. Follow-up telephones are also made. The letter identified the product the problem and the actions to be taken by the customer.Âżssur is committed to providing superior customer service and we will do our utmost to minimize any inconvenience this may cause your practice. To ensure patient safety we kindly ask you to take the following action:" Please identify the users of each of the A-642100 adapters, see attached list." Also, please check your inventory for any A-642100s you may have on hand but have not fitted on users. A follow-up phone call will be made a few days later to coordinate replacements. A new replacement component will be provided to the customer. For further questions please call (949) 360-3634. |  | 66 | 66 units | US Distribution including the states of AL, TX, NY, CA, VA, GA, WI, NC, ID and NJ. | Terminated | June 17, 2014 | 236 | 66 | Computer
Z-0750-2013 | 63875 | No | OAF |  |  | Ear Nose &amp; Throat | PMA | Esteem Programmer Model 6001, US and ISA, Model 3001 Dell E54XX, part of the Esteem Totally Implantable Hearing System. |  |  | January 09, 2013 | 2013 | Envoy Medical Corporation | Envoy Medical is conducting a voluntary correction of a limited number of Esteem Programmers, part of the Esteem Totally Implantable Hearing System, to reduce their susceptibility to noise interference. | Software design | Envoy Medical Corporation sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated December 7, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at 949-233-1204 for questions relating to this notice. |  | 44 | 44 | Worldwide Distribution-USA including the states of AL, AZ, CA, CT, FL, GA, IL, MA, MI, MN, NC, OK, RI, and TX, and the country of GERMANY. | Terminated | April 09, 2013 | 90 | 44 | Computer
Z-0374-2014 | 66231 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | (Siemens Rad Fluoro Uro Systems) Ysio, Luminos Agile, Luminos dRF, and Uroskop Omnia systems with software version VC10 |  |  | September 03, 2013 | 2013 | Siemens Medical Solutions USA, Inc | A potential malfunction and hazard to patients exists when using the Ysio, Luminos Agile, Luminos dRF, and Uroskop Omnia systems with software version VC10, which may result in radiation continuing longer than required and overexposed images that are not of diagnostic quality. | Software design | Siemens Healthcare sent a Field Safety Notice dated September 3, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to forward the Field Safety Notice to any new users and provide the new owner information to Siemens Healthcare. Customers were asked to send back the attached form confirming receipt of the Field Safety Notice. Customers with questions should call 610-850-4634.For questions regarding this recall call 610-219-6300. |  | 357 | 357 | Worldwide Distribution - USA (nationwide) and internationally to Canada. | Terminated | February 24, 2015 | 539 | 357 | Computer
Z-0752-2013 | 63805 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Extremity Medical, IO FiX System Accessory. Catalog Number 101-00010. |  |  | March 22, 2012 | 2012 | Extremity Medical LLC | Lag screw measurement is short (approximately 5 mm) when using current screw measuring device (p/n 101-00010) per the IO FIX Surgical Technique Rev C. which can lead to possible inappropriate implant size selection (e.g. shorter screw size selected). | Device Design | Extremity Medical LLC sent e-mail notification dated March 22, 2012, to all affected customers. The email identified the product the problem and the action to be taken by the customers.The email includes a buleeting which informes the customer that an upgrade to the existing instrumentation is being released. The upgrade will replace the existing gauge. Updated pdf versions of the technique guide are also included. Customers are instructed to return the recalled device after swapping out the recalled device with the replacement unit. Customers will be contacted by Extremity and Sales Reps should contact their principal.For further questions please call (973) 588-8980. |  | 93 | 93 units | Worldwide Distribution -- Nationwide Distribution including the states of AR, AZ, CA, FL, GA, IL, IN, MD, MT, NC, NE, NH, NJ, NY, OH, PA, TN, TX, WI and WA. | Terminated | June 13, 2013 | 448 | 93 | Computer
Z-0754-2013 | 61611 | No | KPQ | Radiation therapy simulation system. | Radiology | Radiology | 510(k) | AdvantageSim MD versions 7.4 through 7.6, Model 5160092-2 Version 7.4, 5160092-3 Version 7.5, 5160092-4 version 7.6. |  |  | May 24, 2012 | 2012 | GE Healthcare, LLC | It was reported by a customer site that when using GE AdvantageSim MD on Advantage Workstation, the series could be incorrectly labeled in image view when multiple series of an exam are loaded simultaneously in Advantage Sim MD and if their series dates are different. A mismatch of series label for structure sets may lead to under-treatment of a tumor due to too small coverage of volume treated. | Software design | GE Healthcare sent an "Urgent Medical Device Correction" letter dated May 11,2012, to all consignees on May 24, 2012. The letter identified the product, the problem, and the action to be taken by the consignee. The letter addressed the Safety Issue, Safety Instructions, Affected Product,, Product Correction and Contact Information. GE indicated that a GE Healthcare service representative would contact they would correct all affected systems by providing a software upgrade at no cost to the consignee. For questions customers were instructed to call one of the following numers:USA 800-437-1171Japan 0120 - 055 - 919For questions regarding this recall call 262-513-4122. |  | 65 | 65 | Worldwide Distribution - USA including AL, CT, FL, IL, MA, MO, NJ, NY, PA, TX VA, WA, WI and Internationally to TURKEY, TAIWAN, SWITZERLAND, SWEDEN, SOUTH AFRICA, NIGERIA, MALAYSIA, KOREA, JAPAN, INDIA, GERMANY, FRANCE, EGYPT, DENMARK, CHINA, CHILE, CANDA, BRAZIL, AUSTRIA, and GREAT BRITAIN. | Terminated | January 16, 2014 | 602 | 65 | Computer
Z-0341-2014 | 66529 | No | GAD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Synthes Soft Tissue Retractor Small Extendible |  |  | September 24, 2013 | 2013 | Synthes USA HQ, Inc. | The potential exists for the locking screw nut to break on the Soft Tissue Retractor if it is overt-tightened. | Device Design | Synthes sent an Urgent Notice: Medical Device Recall letter dated July 31, 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review their inventory and immediately remove the affected lots from stock, call Synthes at 1-800-479-6329 to obtain a Return Authorization Number. Complete the Verification Section at the letter by checking the appropriate box indicating affected product has been located and return the Verification Form with the product to Credit/Returns, Synthes, 1101 Synthes Avenue, Monument, CO 80132. If customers do not have the identified product, they were instructed to complete the Offset blade facilitates easy preparation of the epipereosteal cavity for percutaneous plate insertion.e attached Verification Section and return to Synthes. |  | 374 | 374 | Worldwide Distribution - USA (Nationwide) and Canada. | Terminated | June 01, 2015 | 615 | 374 | Computer
Z-0348-2014 | 66773 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Candela Laser GentleLase Pro, Powered Laser Surgical Instrument for dermatological use - Model number: 9914-00-9015. |  |  | November 06, 2013 | 2013 | Candela Corporation | Unintended single pulse maybe emitted to handpiece prior to pressing of trigger or footswitch and cause injury. | Software design | Candela notiifed Distributors on 11/8/13 through e-mail, Certified return receipt request, or FedEx package containing distributor letter, customer letter for translation and customer lists. A software update to Version 2.0 will be issued to the user by a field service representative. |  | 1258 | 1258 units | Worldwide distribution: US (nationwide) and Internationally to: Azerbaijan, Bulgaria, Chile, Czech Republic, Greece, Israel, Kuwait, Qatar, Saudi Arabia, Serbia, Turkey, and the United Arab Emirates and to Subsidiaries in Australia, France, Germany, Italy, Japan, Portugal, Spain, and the United Kingdom. | Terminated | June 08, 2015 | 579 | 1258 | Computer
Z-0415-2014 | 66864 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | SIGMAÂż LCSÂż HIGH PERFORMANCE" |  |  | November 18, 2013 | 2013 | DePuy Orthopaedics, Inc. | DePuy Orthopaedics, Inc. is issuing a voluntary device correction of the HP M.B.T. Keel Punch Knee Instrument because of the potential for tabs on the connection end to fracture and be left in the patient. | Device Design | DePuy Orthopaedics sent an Urgent Information - Medical Device Correction letter to all affected customers on November 18, 2013. Immediately after initiation on November 18, the company halted initiation with our US distributors and international affiliates because the surgical technique referenced and images from the surgical technique were not correct. No surgeons received the November 18 surgeon letter. The recall was re-initiated on November 20, 2013. The communication to the surgeons who used the affected devices will be delivered by Distributors either by mail or in person and includes the surgeon letter.The device correction includes awareness, reiteration of surgical techniques and instrument care, and visual check of the device as outlined in the Field Notice and Surgeon letter. The devices are not being removed from the market, but the company is investigating potential design changes to the HP M.B.T. Keel Punch Knee Instrument to reduce the potential for the tabs to fracture. All DePuy United States Distributors and DePuy affiliate contacts for Canada, Latin America and International will be notified via email. The US Distributors will identify surgeons who use the affected instruments on Distributor Response Card #1. The US Distributor will notify surgeons either by mail or in person with the written communication. Once all device correction activities are complete, the US Distributor will complete and return Distributor Card #2 to DePuy Orthopaedics, Inc.The US distributors are the US consignees and are responsible for notifying all new HP M.B.T. Keel Punch Knee Instrument surgeon users of this device correction.All territories are required to ensure the HP M.B.T. Keel Punch Knee Instrument Device Correction is promptly initiated. Customers were asked to review the instructions included, along with the attached documents, whichcontain important information regarding notification of surgeons affected by this Device Correction.Customers |  | 58703880 | 29,235 units in total / 2008 through November 2013 | Worldwide Distribution - USA (nationwide) including AR,AZ,CA,CO,FL,GA,HI,IA,IL,IN,KY,LA,MA,MD,ME,MI,MN,MT,NC,NV,NY,OH,PA, TX,VA,WA,WI and Internationally to: AUSTRALIA AUSTRIA BELGIUM BRAZIL CANADA CZECH REPUBLIC DENMARK EGYPT FRANCE GERMANY GREECE HOLLAND HUNGARY INDIA IRELAND ISRAEL ITALY JAPAN KOREA MALAYSIA NEW ZEALAND NORWAY P R CHINA PAKISTAN PANAMA POLAND PORTUGAL PUERTO RICO RUSSIA SAUDI ARABIA SINGAPORE SLOVENIA SOUTH AFRICA SPAIN SWEDEN SWITZERLAND TURKEY UAE UK | Terminated | February 24, 2015 | 463 | 58703880 | Computer
Z-0279-2014 | 66506 | No | MHY |  |  | Neurology | PMA | Medtronic Activa PC, model 37601, and Activa RC, model 37612, Activa SC, models 37602 and 37603. Multi-program Rechargeable Neurostimulator for Deep Brain Stimulation. |  |  | September 30, 2013 | 2013 | Medtronic Neuromodulation | Medtronic notified customers regarding Loss of Stimulation of specific Implantable Neuro-Stimulators under the following specific conditions: Switching between groups with Multiple programs or creating program groups. Therapy can be restored with the Physician Recharge mode (RPM) of the Implantable Neurostimulator Recharger (INSR). | Software design | In September 2013 Medtronic issued a medical device correction letter titled "Loss of Stimulation and Over Stimulation" to customers associated with the affected Medtronic neurostimulators. This letter informed customers of the potential for these issues to occur and provided recommendations to significantly reduce the likelihood of the issues. Medtronic field representatives will be updating the model 8870 software application card to version AAR/01, which is contained in the physician programmer. A Physician Reply Form was asked to be signed and returned by fax or mail for acknowledging the receipt of the Urgent Medical Device Correction letter. When the Medtronic field rep updates the software application card of the physicains' programmer, they will complete a Rep Reply For: 8870 Card Update to AAR/01. |  | 62634 | 62,634 | Worldwide distribution: US (Nationwide) including Puerto Rico; and countries of: Algeria, Aruba, Austria, Belarus, Belgium, Brazil, Bulgaria, Canada, Chile, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, El Salvador, Estonia, Faroe Islands, Finland, France, French Polynesia, Georgia, Germany, Greece, Guadeloupe, Hungary, Iceland, India, Iran, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kuwait, Latvia, Lebanon, Libya, Lithuania, Luxembourg, Malta, Martinique, Mexico, Morocco, Netherlands, Netherlands Antilles, New Caledonia, Norway, Oman, Pakistan, Panama, Peru, Poland, Portugal, Qatar, Reunion, Romania, Russian Federation, San Marino, Saudi Arabia, Serbia, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay, and Venezuela. | Terminated | November 20, 2015 | 781 | 126797 | Computer
Z-0282-2014 | 66537 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Draeger Movita 603x/DVE 803x Ceiling Supply Unit. It is intended for use in operating theaters and intensive care |  |  | September 30, 2013 | 2013 | Draeger Medical, Inc. | Cases were encountered with the Draeger Movita 603x/DVE 803x Ceiling Supply Units where the drive screw in the motor unit broke and the lift arm fell down approximately 60 cm (23.6 inches). | Equipment maintenance | An "Urgent Medical Device Recall" notification letter was sent to users in September 2013. The issue was described and recommended actions were provided. Customers may contact (800) 543-5047 and select 1, then 2. |  | 139 | 139 | Worldwide distribution: US (Nationwide)and in the country of Canada. (note: the ceiling supply units are not classified as a medical device by Health Canada). | Terminated | September 03, 2014 | 338 | 139 | Computer
Z-0443-2014 | 66632 | No | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | Blanketrol III Data Export Software Version 1.10 |  |  | May 15, 2013 | 2013 | Cincinnati Sub-Zero Products Inc | During a customer inquiry regarding a beta version of the Bill Data Export Software, Model USB-127, it was discovered that the wrong software version was released to a customer. | Software Manufacturing/Software Deployment | Cincinnati Sub-Zero sent a Urgent Medical Device Notification letter dated July 15, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.They instructed their customers to to do the following:1) Upon receipt of the corrected Blanketrol III Data Export Software disc, immediately collect and destroy all previous Blanketrol III Data Export Software disc from this lot # currently in possession and replace all downloaded Blanketrol III Data Export Software with the correct version provided by CSZ.2) After the Blanketrol Data Export Software Version 1.00 have been destroyed and downloaded versions replaced with version 1.10, complete and return the enclosed response form as soon as possible to acknowledge receipt of this notification and to inform CSZ that you have performed and completed the requested actions. Return the form by fax to (513) 772-9119 scan and e-mail the information to BIIIExportSoftware@cszinc.com, or mail to:Cincinnati Sub-Zero12011 Mosteller RoadCincinnati, OH 45241 U.S.A.For further questions please call 1-(800) 989-7373 or (513) 772-8810. |  | 12 | 12 Units | Worldwide Distribution - US Distribution including the following states AK, CO, FL, MA, MI, and NC. and Internationally to Canada, Germany and United Kingdom. | Terminated | May 13, 2014 | 363 | 12 | Computer
Z-0281-2014 | 66600 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Ray Station Radiation Therapy Treatment Planning System, Model Numbers 3.0, 3.5 and 4.0 |  |  | October 23, 2013 | 2013 | RAYSEARCH LABORATORIES AB | Ray Search Laboratories found an issue with the dose calculation for Elekta motorized wedges in Ray Station 3.0, 3.5 and 4.0. The error presents when a wedge is used for a beam and the wedge angle differs from 60 degrees. | Software design | Medical Device Correction Letters dated 10/23/13 were e-mailed to the user accounts on 10/23/13. |  | 195 | 195 units domestically | US Distribution in the states of: California, Colorado, Delaware, Florida, Maryland, Missouri, New Jersey, New York, Ohio, Pennsylvania, Tennessee, and Virginia. | Open, Classified |  |  | 195 | Computer
Z-0489-2014 | 66970 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | HGE Healthcare Patient Monitor B30, Procare Monitor B20, Procare Monitor B40, and Monitor B40 |  |  | October 11, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue due to ECG filter settings associated with the B20/B40 V1, B30 and B40 V2 Patient Monitors.The B20, B30 and B40 Patient Monitors set the ECG filter setting incorrectly if the monitor starts with impedance respiration set in the ON position. Instead of the user selected filter, an additional 0.5 Hz high pass filter is used. The | Software design | GE Healthcare sent an Urgent Medical Device Correction letter dated October 10, 2013, to all affected consignees. The letter described the safety Issue, Safety Instructions, Affected Product, Details, Product Correction and Contact Information (Technical Support at 1-800-558-7044).Consignees were asked to retain a copy of the instructions with the user manual.For questions regarding this recall call 262-513-4122. |  | 16061 | 16,061 (343 US, 15,718 OUS) | Worldwide Distribution - USA (nationwide) to CA, CT, FL, GA, MO, NC, NJ, PA, TX, WI and Internationally to:FRANCE, PORTUGAL, ALGERIA, ANGOLA, ARGENTINA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BANGLADESH, BARBADOS, BELGIUM, BOLIVIA, BOSNIA 7 HERZEGOVINA, BRAZIL, BRUNEI DARUSSALAM, BULGARIA, NEW ZEALAND, CHILE, CHINA, COLOMBIA, CONGO, COSTA RICA, COTE DLVOIRE, CROATIA, CYPRUS,. CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, ECUADOR, EGYPT, EL SALVADOR, ESTONIA, FINLAND, GABON, GERMANY, GREECE, GUATEMALA, HONDURAS, HUNGARIA, INDIA, INDONESIA, IRAQ, IRELAND, ISRAEL, ITALY, JAMAICA, KENYA, REPUBLIC OF KOREA, KUWAIT, LATVIA, LEBANON, LITHUANIA, MALAYSIA, MEXICO, MOLDOVA, MOROCCO, MYAMMARA, NAMBIA, NAPAL, NETHERLANDS, NICARAGUA, NIGERIA, NORWAY, OMAN, PAKISTAN, PANAMA, PARAGUAY, PHILIPPINES, POLAND, PORTUGAL ROMANIA, RUSSIA, SENEGAL SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SIR LANKA, SWEDEN, SWITZERLAND, THAILAND, THE NETHERLANDS, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, UNITED ARAB EMIRATES , UNITED KINGDOM, URUGUAY, UZEBEKISTAN, VENEZUELA, and VIETNAM. | Terminated | February 03, 2016 | 845 | 16061 | Computer
Z-0472-2014 | 66638 | No | BSF | Carbon dioxide absorber. | Anesthesiology | Anesthesiology | 510(K) Exempt | Multi Absorber Original, Disposable |  |  | September 19, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue due to air leakage associated with the CO2 Multi Absorber. | Device Design | GE issued an Urgent Medical Device Correction letter dated September 19, 2013, to all affected customers via Fed Ex overnight mail. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to not use the affected product, isolate all affected product and return the attached form via fax to 800-535-7923. Once the information is received, customers would be contacted with an RMA and replacement information. If product was distributed further the notice should be forwarded to those customers. For questions customers should call 1-800-345-2700, option 2 followed by ooption 2. For foreign customers 1-800-932-0760, option 2.For questions regarding this recall call 262-513-4122. GE issued a press release on 12/30/2013. |  | 444115 | US: 59,721 boxes, 358,326 canisters; OUS: 3,724 boxes, 22,344 canisters | Worldwide Distribution - USA (nationwide) Puerto Rico and Internationally to Canada, Costa Rica, Mexico, and Venezuela. | Terminated | March 10, 2015 | 537 | 444115 | Computer
Z-0349-2014 | 64832 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | The GemStar Pump is a small and lightweight, single channel infusion device designed for use in the home, in the hospital, or anywhere electronic infusion is required. The GemStar Pump can be powered by AC mains adaptor, rechargeable battery pack, docking station, or two disposable AA alkaline batteries. When powered by batteries, The GemStar Pump is ideal for ambulatory patients. |  |  | March 15, 2013 | 2013 | Hospira Inc. | The proximal and distal pressure sensor calibration can drift resulting in the pump failing the Proximal or Distal Occlusion Operational Test, as described in the GemStar Technical Service Manual, or reporting one of the following errors during device setup or infusion: 1) Cassette Check - D; 2) Cassette Check - P; 3) Proximal Occlusion; 4) Distal Occlusion; 5) Pressure Calibration Error; 6) Bad Pressure Sensor Event; 7) Bad Pressure State; 8) Distal Pressure is Out of Range; 9) Proximal Sensor is Out of Range. A pump with this issue may, instead of reporting an error, not detected occlusions or issue false occlusion alarms, which will stop the infusion and invoke visual and audible warnings to the user. | Device Design | Hospira sent an Urgent Device Field Correction letter dated March 15, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The letter asks that customers immediately perform proximal and distal occlusion tests as defined i nthe GemStar Technical Service Manual. If the device fails either of the tests, remove it from clinical service and contact Hospira at 1-800-441-4100 (Monday - Friday, 8:00 AM - 5:00 PM) to report the issue and arrange for the return of the device for recalibration. The letter further reminds customers that the proximal and distal occlusion tests should be performed on an annual basis as a part of the yearly GemStar maintenance schedule. Customers were also requested to complete and return the attached reply form and return it to Hospira by fax to 1-888-714-5077 or by e-mail to Hospira7906@stericycle.com.Customers were directed to notify their consignees if they further distributed the product and confirm they have done so by return of the reply form to Stericycle. For questions regarding this recall call 800-441-4100. |  | 4132 | 1) List Number 13000: 2,162 pumps; 2) List Number 13100: 15 pumps; 3) List Number 13150: 1,955 pumps | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY. Internationally to Australia, New Zealand, China, Hong Kong, Taiwan, Japan, Korea, Philippines, Thailand, Malaysia, Canada, Austria, Bahrain, Belgium, Croatia, Denmark, Egypt, Finland, France, Germany, Gibraltar, Greece, Hungary, Iceland, Ireland, Israel, Italy, Jordan, Kuwait, Lebanon, Libya, Luxembourg, Malta, Netherlands, Norway, Oman, Portugal, Qatar, Romania, Saudi Arabia, Slovenia, South Africa, Spain, Sweden, Switzerland, Turkey, UK, United Arab Emirates, Brazil, Chile and Colombia. | Terminated | March 07, 2017 | 1453 | 47972 | Computer
Z-0685-2013 | 62719 | No | HXC | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | DePuy Spine EXPEDIUM Offset Torque Wrench |  |  | July 24, 2012 | 2012 | DePuy Spine, Inc. | A component within the handle of the Expedium Offset Torque Wrench is prematurely wearing, potentially resulting in the instrument recording torque values outside of the specification. | Device Design | DePuy Spine issued an Urgent - Voluntary Product Recall Notification dated July 24, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to disconstinue the use of the affected product and to remove the EXPEDIUMÂż Offset Torque Wrenches from service and return to DePuy Synthes Spine until direction can be provided by DePuy Synthes Spine regarding a replacement torque limiting device.For questions regarding this recall call 508-828-3769. |  | 140 | 140 units | Worldwide Distribution - USA (nationwide) and internationally to Belgium, Egypt, Ireland, Italy, Portugal, and the UK. | Terminated | December 30, 2013 | 524 | 140 | Computer
Z-0180-2014 | 66067 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | CARESTREAM DRX -Mobile Retrofit Kit, Catalog number: 1019397. Made in U.S.A. by Carestream Health, Inc. 150 Verona Street Rochester, NY 14608 |  |  | August 29, 2013 | 2013 | Carestream Health Inc. | Carestream Health Inc. has conducted a voluntary field action which includes the installation of a software patch on the following: DRX-Revolution Mobile X-Ray, DRX-Mobile and DRX Transportable Systems with DirectView Version 5.5 and 5.6 console software when used with DRX-1 Detector version 69 firmware. Carestream recently received reports of a non-diagnostic image on exams when a Normal Exposure | Software design | The firm, Carestream Health Inc., sent an "Urgent: Medical Device Recall" letter dated August 29, 2013 to its customers. The letter described the product, problem and actions taken. Carestream has developed a software patch that will allow Normal Exposure Mode and Extended Exposure Mode images to be taken in any order and will install the patch in all DRX-Revolution, DRC-Mobile and DRX-Transportable Systems with software version 5.5 or higher and Detector firmware version 69 or higher. No DRX Transportables were affected. Your Carestream service representative will install the patch at your next service call. If you have questions or concerns, or if you experience an issue prior to our visit, please contact the Carestream Customer Care Center at 1-800-328-2910, 8am through 8pm Eastern Standard Time. |  | 364 | 364 units | Nationwide distribution:AL, AZ, AR, CA, CO, CT, DC, FL, GA, ID, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MS, MI, MO, NE, NH, NJ, NM, NY, NC, OH, OK, OR, PA, RI, SC, TN, TX, UT, VT, VA, WA, WI, and WV. | Open, Classified |  |  | 554 | Computer
Z-0520-2014 | 66918 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Mega Needle Driver, 8MM assembly of the da Vinci S IS1200 and da Vinci Si, IS2000. |  |  | November 15, 2013 | 2013 | Intuitive Surgical, Inc. | Firm has become aware of the potential for detachment of the jaw insert for Large Needle Drivers and Mega Needle Drivers. | Software design (manufacturing process) | Intuitive Surgical sent a Field Safety Notice dated November 15, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to ensure all appropriate personnel are fully informed of the Notification. Discontinue use of affected product, Use the attached instructions to identify and segregate affected product. Return affected product to Intuitive Surgical following the established Returned Materials Authorization (RMA) process. Complete the Acknowledgement Form and return to Intuitive Surgical. Customers with questions were instructed to contact Intuitive Surgical Customer Service. Customers were asked to retain a copy of this letter and the acknowledgement form for their files.For questions regarding this recall call 800-876-1310, option 3.Japan 0120-56-5635 or 03-5575-1362South Korea 02-3271-3200Europe, Middle East, Asia and Africa +800 0821 2020 +41 21 821 2020For questions regarding this recall call 408-523-2100.Urgent Medical Device Recall notification letters were sent to all affected sites on November 15, 2013. |  | 3540 | 15, 236 UPDATED 12/19/13 to 11,235 | Worldwide Distribution - USA (nationwide and Puerto Rico) and Internationally to Finland, Belgium, Netherlands, UK, Czech Republic, Sweden, Germany, Austria, Switzerland, Norway, Denmark, South Korea, France, Ireland, and Slovakia. | Terminated | September 04, 2014 | 293 | 95754 | Computer
Z-0675-2013 | 63393 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Computed Tomography X-Ray System Brilliance Workspace Portal, Model #728269 Philips Medical Systems, Highland Heights, OH. |  |  | September 11, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Philips was notified that when performing post processing with the Brilliance Workspace Portal platform with software versions 2.6.1.4, the coronary arteries in the rendered image are segmented automatically. However, the center line is not extracted or highlighted upon hovering and making it impossible to activate the vessel and choose the appropriate name (label) from the drop-down menu. | Software design | Philips Healthcare sent a "CUSTOMER INFORMATION -URGENT-MEDICAL DEVICE CORRECTION" letter dated September 13, 2012. The letter identified the product, problem, and actions to be taken by the customers. Philips Service Engineers will contact customers for implementation of the software update. Contact your local Philips representative or local Philips office for further information concerning this notice. |  | 108 | 108 units | Worldwide Distribution-USA including the states of AZ, CA, CO, IL, IL, IN, MA, ME, MI, MO, NC, NJ, NY, OR, TN, TX, VA and WI and the countries of Austria, Canada, China, France, India, Ireland, Italy, Japan, Netherlands, Norway, Spain, Switzerland, Thailand and United Kingdom. | Terminated | February 14, 2014 | 521 | 108 | Computer
Z-0674-2013 | 63817 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance 64 and Ingenuity CT Computed Tomography X-Ray Scanning Systems: Brilliance 64 - Model #728231; and Ingenuity CT - Model # 728326, Distributed By: Philips Healthcare, Cleveland, OH. The Brilliance 64 and Ingenuity CT are Computed Tomography X-Ray Systems intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. These devices may include signal analysis and display equipment, patient, and equipment supports, components and accessories. |  |  | December 03, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Software anomalies. Philips Medical Systems recalled their Brilliance 64 running software v4.0.0xx516 and their Ingenuity CT running software v4.0.0xx518 due to software issues detected in these CT scanning systems. If these issues were to re-occur they could pose the risk of injury to patients undergoing scans; device users; or device service personnel. | Software design | The firm, Philips Healthcare, sent an 'URGENT- Medical Device Correction" letters' dated November 15, 2012, and December 3, 2012, to their customers. Both letters describe the product, problem and actions to be taken. The letters state that Philips is notifying all affected customers of this information via a 'Field Safety Notice and will be installing a software update on the affected Brilliance 64 and Ingenuity CT system in order to address these issues. A Philips' Field Service Engineer will be contacting each customer site and make arrangements to install the necessary software update on these device systems. The customers should contact their local Philips representative or local Philips Healthcare office. In North America and Canada needing further information or support concerning this update, should contact our Customer Care Solutions Center at 1-800-722-9377. Selection option 5, and enter their site ID or follow the prompts. |  | 6 | A total of nine (9) units were distributed (3 units in the U.S &amp; 6 units overseas) | Worldwide distribution: USA including IN, OH and OR; and the countries: Australia, Belgium, China, Israel, and Sweden. | Terminated | June 18, 2013 | 197 | 6 | Computer
Z-0178-2014 | 66331 | No | CJY | Calcium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Calcium Arsenazo, CALA-0250, For the Quantitative Determination of Total Calcium in Human Serum. |  |  | September 04, 2013 | 2013 | SEPPIM SAS | Calcium Arsenazo Reagent, reference CALA-0250, is recalled due to complaint of plasma specimens report high results than serum. | Device Design | ELITech Group of Bothell, WA sent an Urgent: Product Recall Notice dated September 9, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Actions to be taken:" Laboratories must not use plasma for the determination of total calcium when using this product. The product can still be used to determine total calcium with serum sample." ELITech Clinical Systems SAS has revised the package insert and kit box label for the CALCIUM ARSENAZO reagent reference CALA-0250. All mentions of the use of the plasma will be removed from the documentation. If you have inventory of CALCIUM ARSENAZO, reference CALA- 0250. All lots. in your laboratory, please discard the insert that you have and replace it with the insert enclosed with this notice." If you have transferred any of this product to another location, please forward a copy this bulletin to that location." All users receive CALCIUM ARSENAZO, CALA-0250, from distributors. To be certain everyone who is possibly affected by this change is contacted, we are sending this communication to all registered users. As confirmation of receipt of this notice, please return the attached receipt confirmation to Technical Support by fax (401-642-9003) within 48 hours of receipt.We apologize for the inconvenience that this situation can cause in your laboratories; we remain at your disposal for further information. If you have any questions, please contact Technical Support at 1-855-ELITECH (855-354-324)We thank you for your cooperation and the attention which you will give to this bulletin. |  | 95 | 47 kits in US and 48 kits in Sweden | Worldwide Distribution - US including the states of KY and UT and the country of Sweden. | Terminated | December 12, 2013 | 99 | 95 | Computer
Z-0171-2014 | 63859 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Handset component of the Handicare USA's Minivator Bath Bliss Reclining Bath Lifts (Bath Bliss), Part Number BP300-100/BP300-400RSUS (handset) - Model Numbers BBR0311USAB (blue) and BBR0311USSS (silver) |  |  | November 29, 2012 | 2012 | Handicare Accessibility | Firm became aware that two handsets in the affected batches became inoperable while the user was in the tub. | Device Design | Handicare sent an Urgent Medical Device Recall letter dated November 29, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately advise all of their customers affected by the recall letter to immediately remove all affected product from use. Upon receipt of affected product, replacement handsets would be sent to customers. Customers with questions were instructed to call 866-276-5438 or email Customerservice.usa@handicare.com. For questions regarding this recall call 610-266-5260. |  | 3301 | 3301 handsets | Worldwide distribution: USA (Nationwide) and country of: Canada. | Terminated | December 08, 2014 | 739 | 3301 | Computer
Z-0169-2014 | 66659 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | VitreaCoreÂż software. Versions subject to the limitation: |  |  | October 21, 2013 | 2013 | Vital Images, Inc. | Vital Images, Inc. has found an issue in VitreaCore software though internal testing. There is a potential for incorrect and possibly reversed image orientation for certain snapshots taken in VitreaCore software versions 6.0, 6.1, 6.2, 6.3 (except 6.3.1), 6.4 (except 6.4.4 and 6.4.5), as well as their upgrades, and 6.5. | Software design | Vita Images, Inc. sent an Urgent Medical Device Recall letter on October 21, 2013, to all affected customers via USPS First class mail. The letter described the Reason for Voluntary Recall (Nature of software issue, Potential scenarios which exhibit the issue), Risk to Health (How to confirm if the Snapshots are correct), Actions to be Taken by Customer (Recommended Next Steps which includes to complete and sign the enclosed Effectiveness Check Form and send to fieldnotices@vitalimages.com or fax to 1-952-540-3717), and Actions Being Taken by Vital Images (Vital Images Contact Information: If you have questions, direct Vital Images customers and customers of Toshiba America Medical Systems (TAMS) should contact Vital Images Customer Support at +1-800-208-3005, or by e-mail to support@vitalimages.com; customers through a distributor will be contacted by that distributor).For questions regarding this recall call 952-487-9574. |  | 598 | 598 | Nationwide Distribution - all states in continental USA including DC but not OK or RI.OUS: Australia, Austria, Belgium, Brazil, Canada, China, Denmark, Finland, France, Germany, Honduras, Hong Kong, Iceland, Ireland, Italy, Korea, Netherlands, New Zealand, Norway, Pakistan, Poland, Portugal, Qatar, r Saudi Arabia, Serbia, Singapore, Slovenia, South Africa,Spain , Sweden, Switzerland, Thailand, Turkey, Ukraine. | Terminated | January 16, 2015 | 452 | 598 | Computer
Z-0662-2013 | 62121 | No | DSP | Intra-aortic balloon and control system. | Cardiovascular | Cardiovascular | 510(k) | CARDIOSAVE Intra-Aortic Balloon Pump |  |  | March 27, 2012 | 2012 | Maquet Cardiovascular, LLC | In the Cardiosave Intra-aortic Balloon Pump, there is a possibility that the coiled cord assembly, a prominent connector cord between the Cardiosave IABP housing and the dual display head which secured atop the IABP may fail as a result of mishandling. | Device Design | Maquet Cardiovascular sent an Urgent Product Recall Letter dated March 27, 2012, and Urgent Product Recall Acknowledgement to all affected customers . The letter identified the product, the problem, and the action to be taken by the customer. Datascope Corp determined that the best course of action was to replace the affected product. Customers were instructed that they would be contacted by a service representative to make arrangements for the removal and replacement of their affected product. Customers with questions were instructed to contact the US Call Center at 1-800-777-4222.For questions regarding this recall call 973-709-7498. |  | 29 | 29 units (total) | Worldwide Distribution - USA including ID and Fl and Internationally to United Arab Emirates, France, Luxembourg, Poland, Spain, Sweden, Germany, Italy, Netherlands, Republic of China, Spain, Sweden, and Thailand | Terminated | September 30, 2015 | 1282 | 29 | Computer
Z-0170-2014 | 66408 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | HAMILTON-C2 Ventilator: |  |  | September 26, 2013 | 2013 | Hamilton Medical, Inc. | Manufacturer defined limits for improperly functioning loudspeakers of the HAMILTON C2 alarm system were exceeded. | Device Design | Hamilton Medical sent an Medical Device Safety Alert and Corrective Action letter dated September 27, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The corrective action does not require any action by the operator. In the case of an alarm situation as described in the letter, users should follow the instruction for use (Section 8 - responding to alarms in the operator manual) and have the ventilator serviced by an authorized service technician. An addendum to the Operators Manual as a result of the field correction was provided to customers with the Safety Alert Letter.Manufacture will develop and provision a test protocol for the service tech which test the proper functioning of the loudspeaker during the annual service inspection.We appreciate your support in this matter and sincerely regret any inconveniences that this action may cause you. We consider this action as necessary to ensure that our customers receive only safe and effective products with high quality.Further questions please call (775) 426-6331 Ext. 210 |  | 316 | 316 devices distributed in the US. | Nationwide distribution: USA including Puerto Rico and the states of IL, MT, IN, ND, AZ, MT, KY, NC, OH, MA, MO, VA, CA, NV, TX, WA, NY, GA, SC,MN, WY, PA, ND, UT, NE, CO and MI. | Open, Classified |  |  | 316 | Computer
Z-0524-2014 | 66992 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | InstaTrakÂż 3500 Plus, ENTrakTM Plus, InstaTrakÂż 3500, ENTrakTM Computer-Assisted Navigation Systems (USA), |  |  | November 26, 2013 | 2013 | GE OEC Medical Systems, Inc | GE OEC Medical Systems Inc is recalling InstaTrakÂż 3500 Plus, ENTrakTM Plus, InstaTrakÂż 3500, ENTrakTM Computer-Assisted Navigation Systems (USA) due to several potential safety issues related to the use of these products. | Software Design Change | GE Healthcare sent an An Urgent Medical Device Correction letter dated November 26, 2013, to all affected customers. The letter identified the product, the problem, and provided instructions for immediate mitigation of the issue. No product is being returned. The affected systems in the field will be corrected.Customers with questions were instructed to contact the service team at 800-874-7378.For questions regarding this recall call 801-536-4952. |  | 654 | 654 | Nationwide Distribution | Terminated | April 17, 2017 | 1238 | 654 | Computer
Z-0657-2013 | 63355 | No | FSY | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | (l) Visum Halogen Surgical Light, (2) Visum LED Surgical Light, (3) Stryker Flat Panel and Navigation Arm |  |  | September 25, 2012 | 2012 | Stryker Communications, Inc. | Stryker has become aware that there is a low likelihood of incomplete engagement of the Spring Arm Circlip during service or replacementof the Spring Arm. If the Spring Arm Circlip is not fully seated, the Spring Arm could separate and fall along with the surgical lights or monitors mounted to the end of the Spring Arm. To date, there have been no serious injuries reported from the incomplete e | Equipment maintenance | Stryker sent an Urgent Medical Device Correction letter dated September 25, 2012, to all affected customers. The letter identified the product, the prolem, and the action to be taken by the customer. Customers were instructed to immediately inspect their Spring Arm. If customers experienced any concerns with their Spring Arm they were instructed to contact Stryker immediately at 1-866-841-5663 to schedule a service technician to inspect and correct the Spring Arm circlips.For questions regarding this recall call 972-410-7100. |  | 1281 | United States: 1,082 units and Outside United States: 199 units | Worldwide Distribution - USA (nationwide) and Internationally to Australia, Japan, Taiwan, Canada, United Kingdom, France, Netherlands, Spain, Mexico, India, Korea, Singapore, New Zealand, Hong Kong, Philippines and Brazil. | Terminated | April 22, 2013 | 209 | 1281 | Computer
Z-0487-2014 | 66736 | [6 empty] | Nephros SafeSpout |  |  | October 28, 2013 | 2013 | Nephros Inc | Promotional materials for non-medical water filtration products were determined to promote claims which constitute marketing the product as a medical device. | Device Design | Nephros sent an Urgent Medical Device Recall letter dated October 28, 2013, with a return response form to all affected customers via UPS 2nd Day Air. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to cease shipment of the affected product immediately. Obtain an RGA number to return the new and used filters for credit. In addition, customers were instructed to immediately remove and discard all copies of the labeling materials that pertain to the affected product. Remove all electronic copies of these materials from any publicly accessible locations (e.g. website). Customers with questions were instructed to contact their Nephros representative or call 201-343-5202, ext 100.For questions regarding this recall call 201-343-5202.The letter was also posted at:http://www.nephros.com/download/news/SafeSpout_SafeShower_Recall_Letters.pdfNephros sent a new recall letter dated 2/20/204 as they determined they have additional customers. |  | 3742 | 3742 units | Nationwide Distribution | Terminated | June 28, 2017 | 1339 | 8946 | Computer
Z-0822-2013 | 64111 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Viterion TeleHealth Monitor Model Number V500; Product Codes 08832984 (US) and 02080468 (UK) - Physiological patient monitor (without arrhythmia detection or alarms). |  |  | November 15, 2012 | 2012 | Viterion TeleHealthcare Llc | Viterion telehealth monitors do not transmit "HI" (above the meter's measurable range) and "LO" (below the meter's measurable range) blood glucose readings from Bayer's Contour Elite XL blood glucose meters, or from Roche's Accu-Chek Advantage 03 and 04 meters to healthcare professionals. | Software design | Bayer Healthcare sent "IMPORTANT RECALL INFORMATION" to all affected customers.The letter identified the product, problem, and actions to be taken by the customers. Contact Viterion Support at 1-800-866-0133 option 1 for questions relating to this issue. |  | 369 | Domestic - 296 units, Foreign - 73 units | Worldwide Distribution-USA (nationwide) and the United Kingdom. | Open, Classified |  |  | 27440 | Computer
Z-0824-2013 | 64070 | No | DXN | Noninvasive blood pressure measurement system. | Cardiovascular | Cardiovascular | 510(k) | Philips Healthcare SureSigns VS2+ |  |  | January 14, 2013 | 2013 | Philips Healthcare Inc. | VS2+ screen may lock up and stop monitoring during patient use | Software design | Philips issued an Urgent Medical Device Correction letter on January 14, 2013. The field action will take place in two phases: Phase 1 Philips sent customer an Urgent Medical Device Correction Notification/Field Safety Notice. This notification explains the issues and gives the customers instructions on actions to take while they await the hardware correction.Phase 2 Once the hardware correction Philips will contact customers and perform the correction free of charge. Until the correction is available.Affected products may continue to be used while waiting for the correction from Philips if the unit is not used in NBP interval mode.A Philips Healthcare representative will contact you to arrange for the correction of all affected devices. If you need any further information or support concerning this issue, please contact the Philips Customer Care Center at 1-800-722-9377.On 3/14/13, firm stated Initially the fix would be a hardware upgrade but have since changed that to a software update. Customers within the US will be sent a new Field Safety Notice as well as a USB software update package and will be required to notify us upon the upgrade of their affected units. They will be performing the updates themselves |  | 3248 | 3248 units | Worldwide Distribution - US Nationwide and the countries of: AUSTRALIA, AUSTRIA, BELGIUM,CANADA, CHILE, DENMARK, FRANCEGERMANY, ICELAND, INDIA, IRELAND,ISRAEL,ITALY, JORDAN, KENYAKUWAIT, LEBANON, MALAYSIA, NAMIBIA, NETHERLANDS, NEW ZEALANDNORWAY,PALESTINE, POLAND,PORTUGAL, SAUDI ARABIA, SINGAPORESLOVENIA, SOUTH AFRICA, SOUTH KOREA,SPAIN,SWEDEN, SWITZERLANDTHAILAND, TRINIDAD AND TOBAGO, UNITED ARAB EMIRATESUNITED KINGDOM | Open, Classified |  |  | 3248 | Computer
Z-0522-2014 | 65915 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | KayserBetten model IDA beds |  |  | April 11, 2013 | 2013 | mobility unlimited inc | Some joints on the doors of a limited number of IDA beds may loosen or separate over time which can cause the doors to drop outward and cause injury. | Device Design | Mobility Unlimited sent a notification letters dated April 12, 2013 to affected customers The letter identified the affected product, problem and actions to be taken. The letter informed customers that a Retrofit Kits along with illustrated instructions will be sent. For questions call 1-800-574-7880. |  | 33 | 33 beds | Worldwide Distribution - USA Nationwide in the states of AL, AZ, CA, FL, GA, IL, MN, MO, NY, PA, TN, TX, VA, and Canada. | Terminated | October 08, 2014 | 545 | 33 | Computer
Z-0479-2014 | 66533 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Aesculap GNI161 Bipolar Foot Control |  |  | September 10, 2013 | 2013 | Aesculap, Inc. | Complaints were received for the GN161 Bipolar Foot Control reporting that the bipolar energy did not stop after release of the foot control. | Device Design | Aesculap sent an Important Recall Notification letter dated September 23, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to return affected product and receive replacement product or credit. An Inventory Sheet is attached and must be completed. For questions call 610-984-9265 or 610-984-9414. |  | 85 | 85 | Worldwide Distribution - USA Nationwide and countries of: Austria, Germany, Japan, and Taiwan. | Terminated | October 20, 2014 | 405 | 85 | Computer
Z-0694-2013 | 63931 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Spacelabs Smart Disclosure System, Model 92810, is an integrated component of the Intesys Clinical Suite (ICS) G2, Version 4.03. |  |  | December 05, 2012 | 2012 | Spacelabs Healthcare, Llc | Displaying or printing of an incorrect waveform for the time indicated in the printout or display in the Spacelabs Healthcare Intesys Clinical Suite (ICS) G2, Version 4.03, Clinical Access portion of Smart Disclosure, Model 92810. | Software design | Spacelabs has sent the Urgent -Medical Device Correction Spacelabs Healthcare ICS G2 Version 4.03 Model 92810 Smart Closure letter, dated 19 December 2012, to all U.S. customers. The firm plans to send a customer letter (translated as necessary) via e-mail to all international subsidiaries and distributors of record on December 27, 2012.The letter is addressed to Hospital Administrator/Biomedical Manager/Office Manager.Users are advised on the following actions:1) Advise your staff of this potential failure and advise them to always click the"Refresh" button when scrolling to new waveform data prior to printing or viewing waveform information or storing saved events.2) We will contact you at the earliest possible date to schedule a convenient time for your hospital to upgrade your affected software at no cost.Customers with questions can contact Spacelabs Healthcare at 1-800-522-7025 and select 2 for Technical Support. |  | 44 | 27 in US and 17 International | Worldwide Distribution-USA including Puerto Rico and the states of California, Florida, Iowa, Illinois, Louisiana, Michigan, Minnesota, North Carolina, South Carolina, Texas, and Washington, and the countries of Argentina, Mexico, France, and India. | Terminated | July 08, 2013 | 215 | 44 | Computer
Z-0480-2014 | 66714 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | NexGenÂż Complete Knee Solutions Stemmed Tibial Component Precoat |  |  | October 31, 2013 | 2013 | Zimmer, Inc. | Two complaints were received in which it was reported that the stem would not lock into the tibia. Evaluation of the returned device found that the stemmed portion of the tibial plate would not hold the stem as the connection was too loose. The taper diameter was found to be oversized by .010 . | Device Design | All distributors will be notified via electronic mail. Hospital risk managers and distributors with affected inventory will also be notified via courier. Surgeons that have implanted this manufacturing lot will be identified and notified by the distributor. Distributors will be sent a letter identifying the issue and their responsibilities. These responsibilities include locating and removing the affected product in their territory, as well as identifying and notifying surgeons that have implanted this manufacturing lot. Distributors will return on-hand affected products to Zimmer and ensure all of their products are accounted for using Attachment 1 provided in the letter. Distributors are to complete and return Attachment 2 indicating which surgeons have implanted the product and to which they have provided notice to. Hospitals that received direct shipments from Zimmer will be sent a letter identifying the issue and their responsibilities. These responsibilities include locating the affected product and returning it to their Zimmer sales representative.Communications outside of the United States will occur approximately two weeks after the United States communications.Zimmer will conduct effectiveness checks in the following way:Distributors/Hospitalso 100% of the FedEx notifications will be accounted foro An inventory certification form will need to be returned from each distributor that certifies that all available affected product from their territory has been returned.o Additional notifications via email and FedEx will be sent to those who have not completed the required certification formo Accounts will be deemed unresponsive after 3 attempts.Further questions please call 800-613-6131. |  | 128 | 128 units | Worldwide Distribution - USA (nationwide) including the states of CA, IN, OH, GA, WI, IA, NE, NJ, PA, NY, MD, NC, AZ, NM, TX, CO, IL, AL, MI, WA, FL and KS, and the countries of China, Japan and Taiwan. | Terminated | April 30, 2015 | 546 | 128 | Computer
Z-0270-2014 | 66550 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | AW VolumeShare 4 (ADVANTAGE WORKSTATION 4.3.), AW VolumeShare 5 (AW VOLUMESHARE 5 WITH ANGIOVIZ OPTION), AW Server. |  |  | September 19, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue due to the image orientation within GSI viewer on AW products.The GSI Viewer application running on AW/Volume Share or on the AW Server when making a batch and rotating the batch group clockwise 180 degrees (Inferior to Superior) will display images in the Superior orientation. The orientation annotation on the image is correc | Software design | The firm, GE Healthcare, sent an "Urgent Medical Device Correction" letter dated September 19, 2013 to its consignees/customers. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information (800-437-1171 for USA and Canada or contact local GE Healthcare Applications Specialist). The customers were instructed to "Please ensure that all potential users in your facility are made aware of this safety notification and the action taken". GE Healthcare will provide a software update for all affected systems to address the issue. A GE Healthcare service representative will contact you to arrange for this correction. This activity will be performed at no cost to you.A second updated letter was sent on October 22, 2013 to correct the Affected Product Details communicated in the previous letter dated September 19, 2013. The new statement reads "All versions of GSI Viewer prior to version 2.00-OM on AW and prior to version 2.20- OB on AW Server are affected".If you have any questions, please contact us at 800-437-1171 (USA and Canada) or contact local GE Healthcare Applications Specialist. |  | 1458 | 1458 (336 US, 1122 OUS) | Worldwide distribution: US (nationwide) including states of: AL, AZ, CA, CT,DE, FL,GA, HI, IL, IN, IA, KS, MD,MA, MI,MN, MS, O. NV, NH, NJ, NY, NC, OH, OK, OR, PA, SD, TN, TX , UT, VA, WA, WI. OUS: ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA, BELGIUM, BRAZIL, BYELO RUSSIAN, CANADA, CHINA, DENMARK, ESTONIA, FINLAND, FRANCE, GERMANY, GREECE, HONDURAS, HON KONG, HUNGARY, INDIA, INDONESIA, IRAQ, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KOREA, KUWAIT, LATVIA, LIBYA, LITHUANIA, LUXEMBOURG, MEXICO, NETHERLANDS, NEW ZEALAND, NORWAY, POLAND, PORTUGAL, RUSSIA, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TURKEY, UKRAINE, and UNITED KINGDOM.. | Terminated | July 31, 2014 | 315 | 1458 | Computer
Z-0718-2013 | 63981 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ELECTROLASE(R) HYFRECATOR(R) ELECTRODES - SHARP, Disposable Hyfrecator(R) Tips, REF/Catalog No. 7-100-8CS, Sterile, Rx ONLY -- CONMED CORPORATION Electrosurgical, Cutting and Coagulation and Accessories. |  |  | December 27, 2012 | 2012 | ConMed Corporation | ConMed received complaints of some devices which had broken through the seal of the sterile pouch. ConMed confirmed instances where the pouch seal was compromised on the affected products. In no instance has it been reported to ConMed that a compromise in the sterile barrier has resulted in illness or injury. The compromised seals were discovered prior to use. | Equipment maintenance | ConMed sent an Urgent Field Safety Notice dated January 7, 2013, and Business Reply Form to all affected customers via USPS Priority Mail. International notices were sent via FedEx International Priority on December 27, 2012. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review their inventory and complete Attachment II and return it with unused affected product to:ConMed Corporation525 French Road, Utica, NY 13502Customer were asked to return Attachment II even if they don't have any affected product. For questions regarding this recall call 315-624-3533. |  | 313190 | 313,190 (total for all devices) | Worldwide Distribution - USA (nationwide) Puerto Rico and internationally to Belgium, Colombia, Czech Republic, Dubai, Guatemala, Hong Kong, India, Israel, Italy, Japan, Kenya, Kuwait, Malaysia, Mexico, New Zealand, Norway, Oman, Pakistan, Portugal, Saudi Arabia, Thailand, and Uruguay. | Open, Classified |  |  | 313190 | Computer
Z-0454-2014 | 64762 | No | LMD | Medical image communications device. | Radiology | Radiology | 510(K) Exempt | The GE Centricity Web; Software Version 3.0x, is intended for use under the direct supervision of a licensed healthcare practitioner. The Centricity Enterprise Web is an image and information distribution system for the clinical review of medical images and reports. The Centricity Enterprise Web is not intended for primary diagnosis. The Centricity Enterprise Web is available as a supplemental sub-system to Centricity PACS or as a stand alone Web-based image and information distribution system. The Centricity Enterprise Web provides both the server software and a client application, which utilizes off the shelf browser technology. Centricity Enterprise Web supports several interface protocols which allows the system to be integrated with any image or information systems such as a hospital information system (HIS), radiology information system (RIS), or electronic medical record (EMR), which supports one of the supported integration protocols. |  |  | April 19, 2012 | 2012 | Ge Healthcare It | GE Healthcare has become aware of an issue associated with rejected images using the WADO protocol within GE Centricity Web that may impact patient safety. The WADO (Web Access to DICOM Objects) interface allows for embedding DICOM images as JPEGs on a web page and is available as a licensed operation on Centricity Web. When rejected images are archived by Centricity PACS to Enterprise Archive, an | Software design | Urgent Medical Device Correction letters dated April 19, 2013 were sent to customers. The letter requests that customers ensure that their system is configured to disable the WADO interface using the included instructions until the software revision can be installed at their site. The letter further informs customers that updated software will be sent to them for installation.The software has been modified to correct these defects. The modified software will be released and installed in user sites through GE Healthcare FMI# 85206R. |  | 53 | 53 units | US Distribution incluidng states of: CA, FL, GA, IA, IL, LA, MO, NJ, NY, OK, TN, WA, and WI. | Open, Classified |  |  | 53 | Computer
Z-0219-2014 | 66306 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Carestream Vue PACS; Virtual Enterprise SW StandAlone |  |  | September 16, 2013 | 2013 | Carestream Health Inc. | Carestream Health, Inc. has issued a recall when using the CARESTREAM Client Workstation where under very specific conditions, some images from multi-slice studies such as CT and MR may not load into the viewer application due to a software problem. | Software design | Carestream sent an Urgent Medical Device Recall letter dated September 13, 2013, to all affected customers via Fed Ex on September 16, 2013, informing them that a software problem may have the potential to occur under certain conditions when using the CARESTREAM Client Workstation. When certain conditions are met, it appears that the child background process completes its tasks before the parent process and as a result the parent process considers its work done and terminates. In the event that a software problem does occur, the scan may not be fluent or the image numbers may not be sequential. A software patch is currently being deployed by Carestream Health. Customers were instructed that they would be contacted by a Carestream Service Engineer to coordinate the deployment. Customers with questions were instructed to call 1-866 737 2487 or they may call the Carestream Customer Care Center at 1-800-328-2910, 8AM through 8pm Eastern Standard Time.For questions regarding this recall call 1-800-328-2910. |  |  |  | Worldwide Distribution - USA including AR, AZ, CA, CO, CT, FL, IA, IL, MA, MD, ME, MN, MO, NC, NH, NV, NY, OH, OK, PA, RI, TN, UT, VA WA, and WI and Internationally to Canada, Japan, and Kuwait. | Terminated | December 21, 2015 | 826 |  | Computer
Z-0271-2014 | 65792 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Siemens syngo.plaza image processing system |  |  | August 08, 2011 | 2011 | Siemens Medical Solutions USA, Inc | The firm became aware of an unintended behavior when using the syngo.plaza. The "Rename" functionality can cause an unintended patient merge if the patient is sent to another system. This issue affects customers who transfer patient data between two or more syngo.plaza or syngo Imaging XS systems. | Software design | Siemens sent a Customer Safety Advisory Notice dated August 3, 2011, to all affected customer. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to forward the Customer Safety Advisory Notice to their customers. If necessary, customers should have the letter translated technically into the local language and copy it themselves. A letter was sent to direct accounts to inform them of the issu and provide instructions to avoid encountering the problem.For questions regarding this recall call 610-219-6300. |  | 19 | 19 | Nationwide Distribution including MO, MA, NY, WI, OH, NY, CA, FL, NC, MN, and WA. | Terminated | November 06, 2014 | 1186 | 19 | Computer
Z-0701-2013 | 63949 | No | NJR | Streptococcus spp. serological reagents. | Microbiology | Microbiology | 510(K) Exempt | Becton, Dickinson and Company, BD MAX (tm) (6 channel) Instruments, Catalog number 441916, clinical diagnostic instrument. |  |  | September 18, 2012 | 2012 | Becton Dickinson &amp; Co. | Clinical diagnostic instrument may exhibit minor defects, which could lead to erroneous test results. | Device Design | Becton Dickinson sent an Urgent Field Corrective Action letter September 18, 2012,by fax. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised that BD Field Service Associates will contact them within 7-10 days of the notification to determine if their instrument is impacted. Customers were requested to return a response form indicating the serial number on their instrument. If customers needed further assistance they were instrycted ti contact BD Technical Services Department at 1-800-638-8663.For questions regarding this recall call 410-316-4054. |  | 86 | 86 Instruments | USA including CT, FL, IL, KS, MD, MI, MO, NY, PA, RI, TX and WI. Internationally to Canada, Europe and Japan. | Terminated | February 04, 2013 | 139 | 86 | Computer
Z-0202-2014 | 66485 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Synthes 8.5 MM Medullary Reamer |  |  | October 03, 2013 | 2013 | Synthes USA HQ, Inc. | Complaints were received for intraoperative reamer head breakages which can allow for unretrieved fragments of non-implant grade material. | Device Design | Synthes sent an Urgent Notice: Voluntary Medical Device Recall Notification letters, dated October 3, 2013, to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to review their inventory and immediately remove the affected lots from stock. For questions call 610-719-5450 or contact your Synthes Sales Consultant. |  | 2979 | 2979 | Worldwide Distribution: USA (Nationwide) and countries of: Canada and Mexico. | Open, Classified |  |  | 2979 | Computer
Z-0269-2014 | 66560 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Philips brand IntelliSpace PACS 4.4, Image Management System; The system is a software package used with general purpose computing hardware to acquire, store, distribute, process and display images and associated data throughout a clinical environment. Product is manufactured and distributed by Philips Healthcare Informatics, Inc. Foster City, CA |  |  | October 14, 2013 | 2013 | Philips Healthcare Informatics, Inc. | Philips Healthcare became aware of a software problem with the IntelliSpace PACS system, in that a timing issue has been identified that may result in a condition where the device does not store additional images received before processing is complete. The missing images may contain pathology which may result in a different diagnosis if they had been available at time of interpretation. | Software in the Use Environment | Philips sent an Urgent Field Safety Notice dated October 14, 2013, to all affected customers. The letter identified the product, problem, and actions to be taken by the customers.Please review the following with all radiology technicians who use lntelliSpaceBY CUSTOMER/USER PACS:" Assure you and your colleagues are aware of the problem." Verify that all images sent for each study are present on the iVaultbefore the images are purged from the modality.o Use the image count from the lower right corner of thethumbnails as seen in Enterprise only. Single images (such asCR, DX and MG) show no image count, so the user shouldcount those thumbnails without image counts as one. Add all thenumbers up to come up with the total number of viewableimages, stored to lntelliSpace." Compare number of images sent from the modality and those sent fromother 3rd party source to image count in IntelliSpace PACS." If the image count on the modality or 3 party source indicates there aremore images to the study than are present on the iVault, the studyshould be resent from the modality and/or 3 party source." Resending the studies will typically correct the problem but thetechnician should again verify the image count is correct.If you need any further information or support concerning this issue, pleaseINFORMATION AND contact your Philips representative or contact Care at isitecare@philips.com. Please call (650) 293-2300 for further questions. |  | 71 | 71 units | Worldwide Distribution - USA (nationwide) and country: Canada. | Terminated | June 11, 2014 | 240 | 71 | Computer
Z-0210-2014 | 66449 | No | HBL | Neurosurgical head holder (skull clamp). | Neurology | Neurology | 510(k) | Composite Series Skull Clamp, one clamp per tote. |  |  | August 29, 2013 | 2013 | Integra LifeSciences Corp. | An investigation of an adverse trend of complaints for the Skull Clamp index locking knob not easily unlocking or disengaging and malfunctions of the 80 lbs. torque knob was initiated. None of the complaints resulted in a report of patient injury, adverse health consequence, or a delay / prolongation of the surgical procedure. | Device Design | Integra sent an Urgent: Voluntary Medical Device Recall Notification dated August 29, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Your local Integra NeuroSpecialist will set up a time with you to discuss replacement options but in the interim, Integra is asking that you do the following:1. Review your inventory and determine if you have any MAYFIELDÂż Composite Series Skull Clamps (A3059). If so, stop using them immediately and remove them from service to prevent any inadvertent use of the Composite Series Skull Clamp.Either while the Integra NeuroSpecialist is on-site or independently;2. Complete the attached form. If you do not have affected product, check the box: I do not have Composite Series Skull Clamps.3. If you do have affected product, check the box: I do have Composite Series Skull Clamps. Record the total quantity and Serial No.(s) of all Skull Clamps in your inventory. The Serial No. is located on the clamp upright comprised of two letters (SS) followed by six numerals.4. Complete the other information and return by email or fax as indicated on the form, or give it to the Integra NeuroSpecialist who will return it for you.For assistance returning and exchanging these products or any other questions that you may have, please contact Integra Customer Service at 1-855-532-1723.We sincerely apologize for any inconvenience this may have caused and thank you for your cooperation. |  | 37 | 37 units | US Distribution including the states of AL, AR, CA, CT, FL, GA, IA, IL, MS, NY, OH, SC, TN, and TX. | Terminated | January 14, 2014 | 138 | 37 | Computer
Z-0273-2014 | 66664 | No | EBW | Dental handpiece and accessories. | Dental | Dental | 510(k) | Osseocare Pro Console &amp; Set Osseocare Pro, Model 1600870-001, 1700470-001, and 1700471-001. |  |  | October 24, 2013 | 2013 | Nobel Biocare Usa Llc | The recall was initiated because there is a potential for malfunction of the OsseoCare Pro when used with the OseeoCare Pro iPad application software release 1.2.0.7. | Software Manufacturing/Software Deployment | Bien Air Dental send and Urgent Field Safety Notice dated October 24, 2013, to all affected customers. The letter notified customers of the recall, description of the problem, description of the product (OsseoCare Pro iPad app software release 1.2.0.7), action to be taken by user, an alternative solution, and contact details. The letter was also accompanied by an acknowlegement form. Customers with questions were instructed to call 1-800-433-2436.For questions regarding this recall call 714-282-4800. |  | 1082 | 1082 (391 US) | Worldwide Distribution - USA (nationwide) United States and Internationally to Austria, Belgium, Brazil, Canada, Croatia, Czech Republic, Denmark, Finland, France, Germany, Hong Kong, Hungary, India, Ireland, Italy, Japan, Latvia, Lithuania, Luxembourg, Mexico, Netherlands, Norway, Philippines, Poland, Portugal, Romania, Russia, South Africa, Spain, Sweden, Switzerland, Turkey, and the United Kingdom. | Terminated | April 10, 2014 | 168 | 1082 | Computer
Z-0802-2013 | 60695 | No | DRY | Cardiopulmonary bypass on-line blood gas monitor. | Cardiovascular | Cardiovascular | 510(k) | Terumo CDI 101 Hematocrit/Oxygen Saturation Monitoring System. Monitors blood/oxygen saturation, hematocrit and hemoglobin levels. |  |  | November 01, 2010 | 2010 | Terumo Cardiovascular Systems Corporation | Inaccurate readings. Software contains unused (turned off) accuracy correction factors for slope (gain) and offset (bias) for Hematocrit/Saturation calculations. The accuracy on any cuvette size did not meet the specifications. | Software design | Terumo sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated November 1, 2010 to all affected customers. The letter identified the product, problem, and the actions to be taken by the customers. Contact Terumo CVS Customer Service at 1-800-521-2818 for questions or concerns. |  | 622 | 622 units | Worldwide Distribution. | Terminated | November 08, 2013 | 1103 | 622 | Computer
Z-0763-2013 | 64155 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | St. Jude Medical, AGA Medical corporation, AMPLATZERÂż TorqVueÂż FX Delivery System. Model #'s: 9-ITVFX06F45/60, 9-ITVFX07F45/60, 9-ITVFX007F45/80, 9-ITVFX08F45/60, 9-ITVFX08F45/80, 9-ITVFX09F45/80, 9-ITVFX10F45/80 9-ITVFX12F45/80, 9-ITVFX13F45/80. Cardiac surgical instrument, used to close atrial septal defects between the right and left atriums. |  |  | January 18, 2013 | 2013 | St Jude Medical Inc | The distal end of the core wire of the TorqVue FX Delivery System could potentially fracture when exposed to a combination of certain cardiac anatomies and usage conditions. | Device Design | St. Jude Medical sent an "URGENT MEDICAL DEVICE RECALL NOTICE" dated January 18, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers were advised to discontinue use of the device and remove it from inventory. Customers should contact their SJM sales representative for questions relating to this recall. |  | 635 | 635 | Nationwide Distribution-including DC and the states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NH, NJ, NV, NY, OH, OK, OR, PA, RI, SC, TX, VA, WA, and WI. | Terminated | May 21, 2013 | 123 | 635 | Computer
Z-0456-2014 | 66808 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | iba COMPASS SW Version 3.0, Catalog Number CS10-100 (medical linear accelerator) |  |  | August 20, 2013 | 2013 | Iba Dosimetry Gmbh | The development partner of IBA informed that there is problem in a SW component used in COMPASS when converting Region of Interest (ROI) geometries to dose grid representation. Under certain conditions this problem generates too large ROI voxel volumes causing erroneous dose statistics and DVHs. | Software design | Iba sent an Field Safety Notice on August 20, 2013, by email and letter delivered by UPS, Please distribute this notice to all those who need to be aware within your area of responsibility where the affected software have been installed (e.g. colleagues, satellites).Please maintain awareness of this notice and resulting action until we have provided you with a new software version.We would kindly ask you to acknowledge the receipt of this Field Safety Notice. Please sign and send back the attached Confirmation of Receipt within four weeks after the receipt of this Field Safety Notice.We deeply apologize for the inconvenience. We are working on a new software version with corrected functionality. This should become available during next fall and we will ship it to you as soon as it is released, together with the installation and verification instructions.Should you have questions or require additional information, please contact the Customer Service Team at: Service Department IBA Dosimetry GmbH Bahnhofstrasse 5DE-90592 Schwarzenbruck, GermanyTel.:+49 9128 607 38Fax: + 49 9128 607 26Email: service@iba-group.com. |  | 33 | 33 units | Worldwide Distribution - US Distribution including the states of CT, MA, OH, CA, SC, NY and OK, and the countries of International distribution (from Germany) Poland, Germany, GB, The Netherlands, Spain, Turkey, Canada, Mexico, China, Hong Kong and Singapore. | Terminated | February 24, 2014 | 188 | 33 | Computer
Z-0469-2014 | 66596 | research use only/pathology | OTU |  |  | Pathology | Enforcement Discretion | Digital Pathology System (DPS) Software version 1.1 |  |  | October 01, 2013 | 2013 | Omnyx Llc | There is an issue associated with Digital Pathology System (DPS) Software, in which the annotation tools within the DPS version 1.1 Image Viewer will display incorrect measurements and area calculations for a region of interest on a scanned slide, under certain conditions. | Software design | Omnyx sent a Urgent Field Safety Notice letter dated October 1, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.Please discontinue use of the measurement and area calculation features of the annotation tools (line, rectangle/square, ellipse/circle, and freeform) in the DPS Image Viewer at this time. If any of these tools have been used for measurements or area calculations on clinical cases, verify measurements and calculations on glass slides.Please confirm receipt of these instructions by completing the attached form and scanning/emailing it to FieldAction@ omnyx.com or faxing it to 001-412-894-2137Product Correction: An updated version of the DPS SW will be released and installed on your system as soon as it becomes available.Please note that the appropriate regulatory authorities have been informed of this field action.We apologize for any inconvenience that this may have caused you and are working diligently to address this issue.If you have any additional questions or concerns, please contact you GE Healthcare account representative. Further questions please call (412) 891-2100. |  | 2 | 2 | International Distribution to Canada and Switzerland only. | Terminated | April 16, 2014 | 197 | 2 | Computer
Z-0696-2013 | 63617 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Roche COBAS INTEGRA 400 and 400 plus Analyzer, |  |  | November 19, 2012 | 2012 | Roche Diagnostics Operations, Inc. | A software security issue with Oracles TNS-Listener component has been identified. TNS-Listener is supplied with the Roche COBAS INTEGRA 400/400 plus Analyzers. The Oracle TNS-Listener can be considered as an authentication and redirection component between the client and the database inside the Roche product. This means an attacker could register their own non-Roche database via remote access | Software design | Roche sent an Urgent Medical Device Correction letter on November 19, 2012, to all affected customers via UPS Ground (receipt required). The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to complete the attached form and return by fax and file this Medical Device Correction for future reference. For technical support customers were instructed to contact Diagnostics Technical Support at 1-900-428-2336.For questions regarding this recall call 317-521-3911. |  | 845 | 845 Units | Nationwide Distribution including AL, AR, AZ, CA, CT, DC, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WY and GUAYNABO, PR. | Terminated | January 13, 2014 | 420 | 845 | Computer
Z-0798-2013 | 64114 | Yes | MSX | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Mindray DS USA, Inc., DPM Central Monitoring System with ambulatory Telepack. Part numbers: 300BF-PA00002, and 300BF-PA00003 (DPM Central Monitoring System) 115-006908-01, -03, -04 (DPM software upgrade kit). The DPM Central Monitoring System transfers information between DPM Central station and other networked devices. |  |  | November 28, 2012 | 2012 | Mindray DS USA, Inc. d.b.a. Mindray North America | Mindray has identified a software anomaly that may where, in the presence of an asystole, the Telepack ECG waveform may produce a unique step pattern (instead of a flat line) if Lead 1 or V is being utilized for detecting cardiac arrhythmia. | Software design | Mindray North America/ Mindray DS USA Inc. sent an "URGENT PRODUCT CORRECTION ACTION" letter dated November 28, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact your Mindray Service Representative at 1-800-288-2121 for questions relating to this issue. |  | 51 | 51 units | Nationwide Distribution. | Terminated | August 15, 2013 | 260 | 51 | Computer
Z-0673-2013 | 63319 | No | GZT | Self-retaining retractor for neurosurgery. | Neurology | Neurology | 510(k) | Vycor Medical ViewSite Brain Access System 17mm/ 11mm / 5cm Device***STERILE*** |  |  | August 21, 2012 | 2012 | Vycor Medical, Inc. | Vycor Medical recalled their VBAS (Vycor Viewsite Brain Access System) because an unidentified black fiber was found on the device. | Equipment maintenance | Vycor Medical, Inc. notified all customers of the product recall by telephone and follow-up letter. Customers were instructed to quarantine affected products and hold the shipment(s) in quarantine unti further notice. The units will be replaced with inspected and passed units. All distributors were instructed to notify their customers of the recall and ask to have them return their stock. For question call 561-558-2020. |  | 3000 | 3000 devices | Worldwide Distribution - USA (Nationwide) in the states of TX, IL, NJ, FL, OH, PA and (International) in the country of Japan. | Terminated | May 13, 2013 | 265 | 3000 | Computer
Z-0695-2013 | 63863 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Hoffman LRF Wire Bolt, Catalog number 4933-1-001 through -003. |  |  | October 25, 2012 | 2012 | Stryker Howmedica Osteonics Corp. | Two complaints have been filed where Hoffman LRF Wires have broken during load bearing application by patients. | Device Design | Stryker sent "Urgent Medical Device Removal" notification letters and acknowledgement forms dated 10/25/2012 via Fed Ex to Branches/Agencies with return receipt. Stryker sent notification letters/acknowledgement forms dated 10/26/202 via Fed Ex to Hospital Risk Management, Hospital Chief of Orthopaedics and Surgeons. The letters described the device issue and provided recommended actions. Questions were directed to Stryker Customer Service. |  | 476 | 476 units | Nationwide Distribution, including the states of PA, OH, NY, CA and IL. | Terminated | December 01, 2015 | 1132 | 476 | Computer
Z-0478-2014 | 66874 | No | FLE | Steam sterilizer. | General Hospital | General Hospital | 510(k) | AMSCOÂż EvolutionÂż steam sterilizer Transfer Carriage. |  |  | October 18, 2013 | 2013 | Steris Corporation | The firm discovered that unlatching can occur when overloaded transfer carriages are pulled or maneuvered incorrectly or with excessive force. | Device Design | Steris sent an Urgent Voluntary Field Correction Notice letter dated October 18, 2013 to their customers. The letter identified the affected product, problem, and actions to be taken. The letter informed customers that a Steris Service Representative will contact affected facility to arrange for Steris to upgrade the Transfer Carriage(s). For questions contact Steris Field Service Dispatch at 1-800-333-8828. |  | 936 | 936 units | Worldwide Distribution - USA Nationwide in the states of: AL, AZ, AR, CA, CO, CT, DE, FL, GA, ID, IL, IN, KS, KY, LA, ME, MD, MA, MI, MN, MI, MO, MT, NE, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, WY and the District of Columbia including the countries of: Brazil, Canada, Chile, Guam, Mexico, Saudi Arabia, United Arab Emirates &amp; Venezuela. | Terminated | May 13, 2014 | 207 | 936 | Computer
Z-0871-2013 | 64201 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart XL+ Defibrillator/Monitor |  |  | January 28, 2013 | 2013 | Philips Healthcare Inc. | Philips HeartStart XL+ Defibrillator/Monitor may become locked out of clinical use | Software design | Philips Healthcare sent an Urgent Medical Device Recall letter dated January 2013 on January 28, 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised if the Philips HeartStart XL+ Defibrillator/Monitor is first powered on and then there is disconnection or an interruption in AC power, the device may be locked out of clinical mode until the user enters the service mode password and performs an Operational Check. Users are provided with the Safety Notice that provides information on how to identify affected devices and instructions on actions to be taken. Customers were advised to follow the ACTION TO BE TAKEN BY CUSTOMER / USER section of the notice. To correct this issue Philips will provide a software upgrade to customers with affected units free of charge.If you need any further information or support concerning this issue, please contact your local Philips representative at 1-800-722-9377.For questions regarding this recall call 978-687-1501. |  | 3785 | 3785 units | Worldwide Distribution - USA (nationwide) and Internationally to ALGERIA, ANGOLA, AUSTRALIA, AUSTRIA, BAHRAIN, BELGIUM, BULGARIA, COLOMBIA, CROATIA (LOCAL NAME HYVATSKA), CYPRUS, CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, EGYPT, ESTONIA, FINLAND, FRANCE, GABON, GERMANY, HONG KONG, HUNGARY, ICELAND, INDIA, INDONESIA, IRAN (ISLAMIC REPUBLIC OF), IRAQ, IRELAND, ISRAEL, ITALY, JORDAN, KENYA, REPUBLIC OF KOREA, KUWAIT, LATVIA, LEBANON, LIBYAN ARAB JAMAHIRIYA, LITHUANIA, MALAYSIA, MALTA, MAURITIUS, MOROCCO, NAMIBIA, NETHERLANDS, NEW ZEALAND, NORWAY, OMAN, PALESTINE, POLAND, PORTUGAL, ROMANIA, SAUDI ARABIA, SINGAPORE, SLOVAKIA (SLOVAK REPUBLIC), SLOVENIA, SOUTH AFRICA, SPAIN, SRI LANKA, SWEDEN, SWITZERLAND, SYRIAN ARAB REPUBLIC, TANZANIA, THAILAND, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, AND VIET NAM. | Open, Classified |  |  | 3785 | Computer
Z-0949-2013 | 64580 | No | IOQ | Powered flotation therapy bed. | Physical Medicine | Physical Medicine | 510(K) Exempt | TotalCare SpO2RT2 bed. |  |  | February 26, 2013 | 2013 | Hill-Rom, Inc. | In certain situations a software problem with the TotalCare SpO2RTÂż 2 bed may occur that can lead to inadvertent loss of sufficient surface cushionpressure to adequately support the patient. This may increase the likelihoodof skin irritation for patients at higher risk of developing pressure-related skin conditions. A correction to the software program is currently being developed andtested, | Software design | Hill-Rom sent an Urgent Field Safey Notice letter dated February 26, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter advised customers to follow steps provided until the updated software is available. Customers were requested to forward a copy of this letter to any other facility personnel you deem appropriate. Questions should be directed to Hill-Rom Technical Support at 800-445-3720 |  | 309 | 309 | USA Nationwide Distribution including the state of Canada | Terminated | April 16, 2014 | 414 | 309 | Computer
Z-0869-2013 | 64185 | No | FWZ | Manual operating table and accessories and manual operating chair and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | MAYFIELDÂż Composite Series Base Unit Standard, Model (A3101), Rx ONLY, Integra LifeScience Corporation, 4900 Charlemar Drive, Building A, Cincinnati, OH 45227 |  |  | December 21, 2012 | 2012 | Integra LifeSciences Corp. | Integra determined there is a possibility that the Base Units may break during use. | Device Design | Integra LifeSciences Corporation sent an Urgent Medical Device Recall letter dated December 21, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to stop using the affected product immediately and also remove them from service to prevent any further use of the affected product. Integra would provide an alternative base unit to replace those customers could no longer use. Customers were asked to complete the Recall Acknowledgement and Return Form to ensure they received this information. For assistance customers were instructed to call 1-513-533-7944 or 1-513-533-7946 |  | 96 | 96 units | Worldwide Distribution - USA including AL, AR, AZ, CA, CO, CT, FL, GA, IL, IN, KY, MA, MD, MI, MN, MO, MS, NC, NE, NJ, NM, NY, OH, OK, OR, PA, TN, TX, UT, VA, WA &amp; WI. Internationally to Belgium. | Terminated | July 10, 2013 | 201 | 96 | Computer
Z-1681-2013 | 65320 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | ***REF C7122***QTY 1***APEX Arthroscopy Tubing Set. |  |  | May 30, 2013 | 2013 | Linvatec Corp. dba ConMed Linvatec | Conmed Linvatec received complaints that their products C7120 APEX Arthroscopy Tubing Set, One Connection and C7122 APEX Arthroscopy Tubing Set, Two Connection, contained a scratch/grooved area on the inside diameter of the pressure fitting that may cause the APEX Irrigation (pump) Console to sense a pressure drop and increase fluid flow to the surgical site. | Equipment maintenance | ConMed Linvatec sent an "URGENT: Medical Device Recall Notification" letter by FedEx next day delivery or equivalent on 5/30/2013 to all domestic direct consignees. ConMed Linvatec is translating customer notifications letters into the necessary languages and will send International notification letters no later than 6/4/2013. |  | 4691 | 4,691 units, total of each | Worldwide Distribution-USA (nationwide) including the states of AR, CA, FL, GA, IL, IN, KS, MO, NC, NJ, NM, NY, OH, OK, PA, TN, TX, UT, VA, and WI, and the countries of Austria; Belgium, Canada, China, Czech Republic, Finland, France, Great Britain, Italy, and Japan. | Terminated | September 04, 2013 | 97 | 4691 | Computer
Z-1665-2013 | 65429 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | S-LIFT Inserter Outer Sleeve |  |  | May 24, 2013 | 2013 | SpineFrontier, Inc. | The outer sleeve would not fit over the S-LIFT Inserter assembly because the hole at the top of the sleeve handle was too small. | Device Design | Spine Frontier sent an Urgent Advisory Notice letter dated On June 7, 2013 to all consignees via email. The letter identified the affected product, problem and actions to be taken. The letter informed consignees that a modification has been implemented to the Outer Sleeve component of the S-LIFT Inserter and that SpineFrontier has initiated removal of the affected product from the field.. An acknowledgement form is to be signed and returned from each recipient of the Advisory Notice. Upon receipt of the acknowledgement form and RMA request for part return, replacement parts will be shipped. For questions call 978-232-3990. |  | 12 | 12 devices | USA Nationwide Distribution including the states of: FL, TX, OK, MO, NJ, MA | Terminated | October 08, 2014 | 502 | 12 | Computer
Z-1666-2013 | 65442 | No | HXX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | InSpan Inserter, Right Assembly; InSpan Inserter, Left Assembly |  |  | May 07, 2013 | 2013 | SpineFrontier, Inc. | Potential for set screw to be come deformed. | Device Design | Spine Frontier sent an Urgent Advisory Notice letter dated May 30, 2013 to all affected consignees. The letter identified the affected product. problem and actions to be taken. The letter informed consignees that a correction was implemented for InSpan Inserters. Consignees were instructed to remove parts from the field immediately for modification and that replacement Inserters will be shipped upon receipt of completed acknowledgement form and RMA request for part return. For questions call 978-232-3990. |  | 61 | 61 individual inserter assemblies | Worldwide Distribution - USA Nationwide: in the states of: FL, CO, VA, KS, TX, OH, TN, OK, CA, IL, MO, MN, Wash DC, including PR and the countries of Dominican Republic, Medico, Panama | Terminated | January 11, 2016 | 979 | 61 | Computer
Z-1096-2013 | 64725 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | GE Healthcare LOGIQ S8 (with software revisions R1.1.1 and R1.5.1) and LOGIQ S7 (with software revisions R1.0.1, R1.0.2 and R1.0.3) diagnostic ultrasound system with the S4-10-D Ultrasound Probe. The device is intended for use in ultrasound evaluation. |  |  | February 08, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has become aware of a potential safety issue due to the S4-10-D Ultrasound Probe used in conjunction with the LOGIQ S8 and LOGIQ S7 Ultrasound systems. It was also discovered that this issue also leads to a specification issue of the acoustic output of this probe. There is a potential for a skin burn when the flow model in the cardiac application is activated with the S4-10-D Ultraso | Software design | Consignees were sent on 2/8/2013 a GE Healthcare " Urgent Medical Device Correction" letter dated February 8, 2013. The letter was addressed to Hospital Administrators / Risk Manager, Biomedical Engineering and Managers of General/Diagnostic Imaging. The letter described the Safety Issue, Safety Instructions, Affected Product details, Product correction and Contact Information. US customer contact number is (1)800-437-1171. |  | 2061 | 2061 | Worldwide Distribution, USA, including the states of AL, AZ, AR, CA, CO,K CT, DC, FL, GA, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MS, MO, MT NE, NV N NJ, NM, NY, NC, ND, OH, OK, PA, RI ,SC, SD, TN, TX, VA, WA, WI and the countries of AUSTRALIA, AUSTRIA, BRAZIL, CANDA, CHILE, CHINA, COSTA ARICA, CROATIA, CZECH REPUBLIC, DENMARK, FINLAND, FRANCE, GERMANY, GREECE, HONG KONG, HUNGARY INDIA, INDONESIA ITALY JAPAN, KOREA, LATVIA, LEBANON, MALAYSIA, MEXICO, MOLDOVA, MOROCCO, NETHERLANDS, NEW ZEALAND, NORWAY, PHILIPPINES, POLAND, PORTUGAL, ROMANIA, RUSSIA, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA , SOUTH KOREA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, VIETNAM. | Terminated | November 21, 2013 | 286 | 2061 | Computer
Z-1670-2013 | 65292 | No | HQC | Phacofragmentation system. | Ophthalmic | Ophthalmic | 510(k) | SovereignÂż Compact Phacoemulsification System, v 5.3, Part Number: CMP680300. The system is used for cataract surgery. |  |  | May 23, 2013 | 2013 | Abbott Medical Optics, Inc. | Abbott Medical Optics Inc. (AMO) initiated this Urgent Notice of Medical Device Correction and Advisory Notice to all customers who use the AMO SovereignÂż Compact Phacoemulsification System, Version 5.3, because it has the potential to lose the vacuum tone mode during surgery. | Software design | AMO issued an "Urgent Notice of Medical Device Correction and Advisory Notice" on May 23, 2013 via email, to all customers with a SovereignÂż Compact Phacoemulsification System v 5.3 to inform them of the issue and provide them with actions to take in the event of sound loss during use of the SovereignÂż Compact Phacoemulsification System, v 5.3. |  | 534 | 534 systems | Worldwide Distribution, including Nationwide (US) and countries in Latin America, EMEA, and APAC. | Terminated | October 06, 2014 | 501 | 534 | Computer
Z-1668-2013 | 65430 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Medtronic MiniMed Paradigm Insulin Infusion Sets |  |  | June 07, 2013 | 2013 | Medtronic MiniMed | Medtronic is recalling the Medtronic MiniMed Paradigm Medtronic is recalling the Medtronic MiniMed Paradigm Insulin Infusion sets because of a potential safety issue that can occur if insulin or other fluids come in contact with the inside of Medtronic Paradigm infusion sets. Under certain conditions, the infusion set may malfunction and deliver an incorrect volume of insulin. | Device Design | Medtronic, Inc. sent an Urgent Medical Device Safety Notification letter dated June 10, 2013, to all affected consignees. The letter identified the product and notified healthcare professionals of the potential for over or under delivery of insulin if insulin or other fluids contact the inside of Medtronic Paradigm Tubing Connectors. The letter informed the consignees that it applies to all Medtronic infusion sets designed for use with Medtronic Paradigm family infusion pumps. Paradigm Insulin Pump users, distributors, and healthcare professionals were all instructed to contact Medtronic's 24 hour helpline at (888) 204-7616.For questions regarding this recall call 818-567-4700. |  | 11000000 | 11,000,000 units | Worldwide distribution: US (nationwide) distribution: Algeria, Argentina, Australia, Bahamas, Bahrain, Bermuda, Boliva, Bosnia and Herzegovina, Brazil, Canada, Chile, China, Columbia, Costa Rica, Croatia, Cyprus, Denmark, Dominican Republic, Ecuador, Egypt, El Salvador, Estonia, Finland, France, French Southern Territories, Germany, Great Britain, Greece, Guatemala, Hong Kong, India, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kuwait, Latvia, Lebanon, Macedonia, Malaysia, Mexico, Morocco, Netherlands, Norway, New Zealand, Oman, Panama, Paraguay, Peru, Poland, Province of China, Taiwan, Qatar, Republic of Korea, Romania, Russia Federation, Saint Lucia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Syrian Arab Republic, Thailand, Trinidad and Tobago, Turkey, Ukraine, United Arab Emirates, Uruguay, andÂżVenezuela. | Terminated | June 10, 2015 | 733 | 11000000 | Computer
Z-1099-2013 | 64853 | No | OCX | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Torrent" irrigation tubing. |  |  | March 07, 2013 | 2013 | US Endoscopy Group Inc | The firm was notified by their customers that there was a water leakage from the white tubing connector. | Device Design | US Endoscopy Group sent an Urgent Recall Notification letters to their customers. The letter identified the affected product, problem and actions to be taken. They instructed their customers to remove any of the affected products from their inventory and quarantine them. They also requested each customer complete the attached Customer Response Card and return it to US Endoscopy via fax at 1-440-639-4495 or by email at recall@endoscopy.com as soon as possible. For questions contact Customer Service at 1-800-769-8229. |  | 21189 | 21,189 units | USA Nationwide Distribution - including the states of: AK, AL, AZ, CA, CO, CT, DE, FL, GA, ID, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NE, NH, NM, NV, NY, OH, OR, PA, SC, TN, TX, UT, VA, WA, WI &amp; WV. | Terminated | March 20, 2014 | 378 | 21189 | Computer
Z-1678-2013 | 65577 | No | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter's elastomeric infusion pumps are non-electronic, ambulatory, disposable devices that deliver continuous medication for periods of 30 minutes up to 7 days. They are designed for medication therapies requiring continuous infusion. |  |  | July 31, 2009 | 2009 | Baxter Healthcare Corp. | There are reported incidents of leakage on the INFUSOR and FOLFUSOR pumps at the connection of the blue winged cap and the distal male Luer. Investigation of these reports indicates the incidence of leakage has only occurred after pumps are filled, primed, re-capped (using the blue winged Luer cap) and stored for a period of time, typically overnight. In addition, the reported instances of leakage | Equipment maintenance | Baxter sent an "Important Product Information" notification dated July 31, 2009, with a follow-up "Important Product Information" notification mailed on March 5, 2010. The notifications informed customers of the potential issue, and included instructions to ensure that the winged Luer cap is tightly secured after filling and priming. Customers were also instructed to follow normal clinical practices to reduce contact with any leaking medication, and to store pumps with pre-dispensed medication in a sealed pouch. If customers noticed a leak they were instructed to place the pump in a plastic bag, quarantine the device and call Baxter's Corporate Product Surveillance at 1-800-437-5176 for further instructions. Customers were asked to return the attached reply form via FAX to 1-847-270-5457. If customers further distributed the affected product they were instructed to forward the notification to other facilities. Customers with questions were instructed to call 1-800-933-0303.For questions regarding this recall call 224-270-4667. |  | 112860 | Code 2C1700KP: 7,128 units; Code 2C1702KP: 105,732 units | Worldwide Distribution - USA (nationwide), Puerto Rico and Internationally to Australia, Austria, Belgium, Brazil, Bulgaria, Canada, Chile, Columbia, Czech Republic, Denmark (including Faroe Islands), Ecuador, Finland, France (including Martinique and Reunion Islands), Germany, Greece, Iceland, India, Ireland, Isreal, Italy, Luxemburg, Japan, Malaysia, Mexico, Netherlands, New Zealand, Norway, Panama, Portugal, Qatar, Russia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Korea, Spain, Sweded, Switzerland, Thailand, Turkey, United Arab Emirates, and United Kingdom. | Terminated | March 08, 2016 | 2412 | 559758 | Computer
Z-1083-2013 | 64759 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | VMX, VMX Plus Mobile Radiographic Unit with the following models: |  |  | January 23, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has become aware of an issue associated with the base casting for the articulated arm of your VMX andVMX Plus Mobile X-ray systems that may impact Operator or Patient safety. | Device Design | GE Healthcare sent and "Urgent Medical Device Correction" letter dated January 23, 2012. The letter was addressed to Hospital Administrators / Risk Managers, Radiology Department Managers and Radiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details and Product Correction. For questions contact your local GE Healthcare Service Represenative. |  | 1296 | 1296 | Worldwide Distribution - USA (nationwide) including the states of: AR, CA, GA, IL, IN,IA,LA, MA, MN, MT, NY, OH, PA, TX, VA and the countries of BELGIUM, BURKINA FASO, BRAZIL, CHINA, CANADA, SWITZERLAND, CHILE, COLOMBIA, CYPRUS, CZECH REPUBLIC, GERMANY, DENMARK, DOMINICAN REPUBLIC, EGYPT, SPAIN, FRANCE, UNITED KINGDOM, CHANA, GUADELOUPE, GREECE, GUYANA, HONG KONG, CROATIA, HUNGARY, INDONESIA, ISRAEL, ITALY, KUWAIT, LEBANON LUXEMBURG, LIBYAN, MOROCCO, MARTINIQUE, MEXICO, NETHERLANDS, NORWAY, PANAMA, PERU, POLAND, PORTUGAL , REUNION, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SUDAN, SWEDEN, SINGAPORE, SLOVENIA, SLOVAKIA, SIERRA LEONE, THAILAND, TUNISIA, TURKEY, VENEZUELA, MACEDONIA, PALESTINE, SOUTH AFRICA. | Terminated | February 04, 2014 | 743 | 1296 | Computer
Z-1100-2013 | 64763 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE Healthcare, Discovery MR750w 3.0T System. |  |  | April 23, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential issue due to image flip in frequency encoding direction associated with a specific workflow of your Discovery MR750w system.Image flip issue in frequency encoding direction (including S/I, A/P and R/L) can result in incorrect localization ofanatomy and pathological findings. | Software design | GE Healthcare sent an Important Product Information letter dated April 23, 2012. The letter described the Issue, Instructions, Affected Product Details and Product Correction. For questions contact your local service representative. |  | 46 | 46 | Worldwide Distribution - USA (nationwide) including the states of: AZ, CA, CO, FL, IL, MD, NJ, NY, OH, OK, TX. and the countries of Australia, Belarus, Canada,France, Germany, India, Japan, Lebanon, New Zealand, Saudi Arabia, Singapore, Spain, Sweden, Switzerland and United Kingdom. | Terminated | March 14, 2014 | 690 | 46 | Computer
Z-1695-2013 | 65459 | No | HRY | Knee joint femorotibial metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | REF 154200 Oxford Fixed Bearing Partial Knee |  |  | June 07, 2013 | 2013 | Biomet, Inc. | Biomet investigation found the design of the tibial bearing creates the potential for anterior impingement between the bone and the poly bearing component in terminal extension. | Device Design | The firm, Biomet, sent an "URGENT MEDICAL DEVICE RECALL NOTICE" letter dated June 7, 2013 to its consignees/customers. The letter identified the device being recalled, the reason for the recall, potential adverse events associated with the recall and actions to be taken. The consignees/customers were instructed immediately locate and remove the identified device(s) listed in the letter; carefully follow the instructions on the enclosed "FAX Back Response Form"; complete and return the FAX Back Response Form to 574-372-1683 prior to return of product within three (3) business days; return Goods to Biomet, Inc 56 East Bell Drive Warsaw, IN 46580. To confirm receipt of this notice by call 574-372- 1570. Obtain an RGA # by phone (574)372-6677 email rqarequest@biomet.com or through FAST; use priority carrier for your shipment; and if you have further distributed this product you MUST notify personnel of the letter. Questions should be addressed to 574-372-1570 M-F 8 a.m. to 5 p.m. |  | 26 | 26 | Worldwide distribution: USA including NC, IN, PA, and OH; and countries of: Canada and Netherlands. | Terminated | July 17, 2014 | 405 | 794 | Computer
Z-1686-2013 | 65590 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | da Vinci Si Surgeon Side Cart, model number SS3000; |  |  | June 08, 2013 | 2013 | Intuitive Surgical, Inc. | Factory testing on da Vinci Si Surgical Systems may not be in compliance with UL Standards as one of the testing devices was found to be working incorrectly. | Software design | Intuitive Surgical sent an Urgent Device Correction notice dated July 15, 2013, to all affected customers by Federal Express.The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to ensure that all affected personnel were fully informed of the notification and that the notification should be forwarded to other managers within their facility. Customers were also instructed to complete and return the attached Acknowledgment Form to acknowledge receipt of the notification. Customers with questions were instructed to call Customer Service at 800-876-1310, Option 3.For questions regarding this recall call 408-523-2244. |  | 30 | 30 total, all devices | Worldwide Distribution - USA including UT, CA, AZ, CO, FL, HI. MT, NY,TX and WA and one customer in Russia. | Terminated | January 22, 2014 | 228 | 30 | Computer
Z-1081-2013 | 64855 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Siemens brand ACUSON SC2000 Ultrasound System, Model 10433816, hardware that contains version 4 of the MPI board 10434200; |  |  | March 28, 2013 | 2013 | Siemens Medical Solutions USA, Inc. | When using the V5M transducer, the display of the temperature on the ACUSON SC2000 system can be lower than the actual temperature of the V5M transducer. This occurs only if the system hardware contains revision 4 of the MPI Board. There is a potential to cause esophageal burns in patients. | Device Design | Siemens Medical Solutions USA Inc. sent a CUSTOMER SAFETY ADVISORY NOTIFICATION letter dated March 28, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.A return response form is included for customers to complete and mail to the recalling establishment. Siemens stated additional follow up action will occur, once material is available, where Siemens field service engineers will visit each unit and replace the affected MPI 4 board.For further questions please call (650) 694-5747. |  | 120 | 120 units | Worldwide Distribution - USA and countries of Australia, Austria, Bosnia Herzegovina, Belgium, Finland, France, Germany, India, Italy, Japan, Kuwait, Malaysia, Netherlands, Norway, P.R. China, Portugal, Rep.of Korea, Romania, Saudi Arabia, Spain, Sweden, Switzerland, Taiwan and United Kingdom. | Terminated | February 14, 2014 | 323 | 120 | Computer
Z-1694-2013 | 65487 | No | MUE | Full-field digital mammography system. | Radiology | Radiology | 510(k) | Siemens Mammomat Inspiration - full-field digital mammography |  |  | February 28, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Firm became aware of an unintended behavior when using the Mammomat Inspiration. The possibility exists that a malfunction may occur during Stereo or Tomo acquisitions if the tomo clutch was incorrectly adjusted. Although the possibility is rare, the system could move while a patient's breast is compressed resulting in a possible bruise. Firm released Update Instructions XP031/12/S for this iss | Device Design | Siemens issued an Update Instructions XP031/12/S by letter to the affected customers. For affected sites, a Siemens Service Engineer will inspect the unit and make any adjustments needed. Following the inspection and adjustments, a proper calibration will be completed. For questions call 610-448-3237 and 610-446-4634. |  | 3 | 3 | United States Nationwide Distribution in the states of ND, OH and SC. | Terminated | February 21, 2014 | 358 | 3 | Computer
Z-1727-2013 | 65163 | No | JMM | Hydroxyproline test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | B-CrossLaps/serum (B-CTx in serum) |  |  | April 25, 2013 | 2013 | Roche Diagnostics Operations, Inc. | Roche Diagnostics internal investigations in R&D; have shown that the claims for biotin interference and the sample stability given in the Elecsys Âż-CrossLaps (Âż-CL) assay package insert version 5, could not be reproduced. The Âż-CrossLaps test results contribute to the monitoring of antiresorptive therapy. False low or false normal results may lead to delayed diagnosis and treatment of osteoporosi | Device Design | The firm, Roche, sent an "URGENT MEDICAL DEVICE CORRECTION" (UMDC)letter dated May 2, 2013, to its consignees/customers via UPS Ground. The letter describes the product, problem and actions to be taken. The consignees/customers were instructed to apply the correct sample storage conditions as stated in the UMDC and to inform all relevant parties about the correct biotin interference claim. The consignees/customers were instructed to keep the UMDC on file for reference until the revised package insert is available. The consignees/customers were asked to complete and return a form via fax (1-888-943-4205). Consignees/customers were instructed to call Roche Diagnostics Technical Support, 24 hours a day, seven days a week, at 1-800-428-2336 if they had any questions about the recall. Non-responding accounts are monitored on an ongoing basis. |  | 3056 | 3056 units | Nationwide distribution: USA including states of: AL, CA, GA, FL, IL, IN, KS, KY, MA, MD, MI, MN, MO, NC, NY, OH, PA, TX, VA, VT, and WA. | Terminated | September 27, 2013 | 155 | 3056 | Computer
Z-1052-2013 | 64308 | No | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | ShapeMatch Cutting Guides |  |  | November 08, 2012 | 2012 | Stryker Howmedica Osteonics Corp. | With regards to ShapeMatch Cutting Guides using MRI Imaging, Stryker Orthopaedicshas become aware that there are potential issues associated with internal processesfor planning cases and producing ShapeMatch Cutting Guides. | Device Design | Stryker sent Product Correction Bulletins on 11/8/2012 dated 11/9/2012 and 11/19/2012 to all affected customers. Stryker sent an updated Urgent Medical Device Product Notification letters and Acknowledgement letters dated 1/24/2013 through 1/28/2013 to all affected customers via Fed Ex with return receipt: The letter identified the affect product, problem, potential hazards, risk mitigating factors and actions to be taken. For questions call 201-831-6878.Stryker issued press release on 4/10/2013. |  | 7868 | 7868 units | USA Nationwide Distribution | Open, Classified |  |  | 7868 | Computer
Z-1074-2013 | 64624 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | All OneTouch Verio IQ Blood Glucose Meters sold as: |  |  | March 11, 2013 | 2013 | Lifescan Inc | The Verio IQ meter will shut off and revert to set up mode at glucose values above 1023 instead of displaying EXTREME HIGH GLUCOSE. | Software design | Lifescan sent an Urgent Medical Device Voluntary Recall letter dated March 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to identify and hold all affected product in inventory and discontinue distributing. The letter states that a Lifescan sales representative will contact customers to arrange for pick up and replacement. |  | 710413 | US 710,413 | USA Nationwide Distribution | Terminated | August 03, 2015 | 875 | 710413 | Computer
Z-1661-2013 | 65493 | No | JOW | Compressible limb sleeve. | Cardiovascular | Cardiovascular | 510(k) | VenaFlow Elite 30B-B and 30BI-B |  |  | May 28, 2013 | 2013 | DJO, LLC | DJO, LLC has recently identified a product safety issue with the battery assembly for the VenaFlow 30B-B and 30 BI-B products. There is a potential for battery leakage which could result in thermal meltdown (fire). | Device Design | DJO sent a Customer Notification letter dated May 28, 2013 to Domestic customers via USPS certified return receipt and to International customers via UPS on June 14, 2013. The Customer Notification informed customers of the product safety issue, product description, and instructions to follow. DJO plans to install replacement batteries at no charge. For questions contact your local DJO Customer Service Representative. |  | 2491 | 2491 | Worldwide Distribution - USA Nationwide and countries of: Australia, Belgium, China, Guatemala CentroAmeria, France, Hong Kong, India,Italy, Korea, Lebanon, Malaysia, Morocco, Pakistan, Paraguay, The Netherlands, Romania, Singapore, Spain, Sweden, Thailand, Turkey, United Arab Emirates, United Kingdom and Venezula. | Terminated | February 13, 2014 | 261 | 2491 | Computer
Z-1104-2013 | 64693 | No | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | Aplio 500/400/300 Diagnostic Ultrasound System (TUS-500/400/300). |  |  | December 11, 2012 | 2012 | Toshiba American Medical Systems Inc | Toshiba Medical Systems has received reports of startup issues with the Toshiba TUS-500 andTUS-300 ultrasound systems. The startup issue is predominant when attempting to bring thesystem up from standby mode when the system is used remotely. It has been determined that thismay potentially cause damage to the system. | Device Design | Toshiba sent an "URGENT: MEDICAL DEVICE CORRECTION" letter dated December 11, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at (800) 421-1968 for questions regarding this letter. |  | 233 | 233 units | Nationwide Distribution. | Terminated | April 10, 2014 | 485 | 233 | Computer
Z-1103-2013 | 63856 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Integra XKnifeÂż Stereotactic Radiosurgery and Radiotherapy Treatment Planning, |  |  | December 10, 2012 | 2012 | Integra LifeSciences Corp. | Depending on the system configuration, a software error message in versions 5.0.1 and 5.0.2. occurs if a beam plan is transmitted from XKnife using DICOM-RT. | Software design | Integra initiated a voluntary recall on December 10, 2012 by providing a written Urgent: Medical Device Recall Notification to US consignees who have experienced the software issue due to the configuration of their system. The letter identifies the affected product, as well as, advises customers of the nature of the issue and how it affects treatment records.- Advises customers that they will be contacted when V 5.0.3 software, where the software error has been addressed, is available to schedule an upgrade.- Requests that customer consignees complete and return the Product Recall Acknowledgment Form.If customer have additional questions, they should contact the firm's service hotline at 1-888-772-7378. |  | 12 | 11 XKnife V5.0.1 and 1 XKnife V5.0.2 | Distributed in the states of IL, ID, LA, MA, NJ, NY, OH, PA, and SD. | Terminated | August 01, 2013 | 234 | 12 | Computer
Z-1102-2013 | 64547 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Intellispace Portal (ISP) Philips Medical Systems (Cleveland), Inc., |  |  | February 12, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | Philips was notified of a problem using the IntelliSpace Portal with software version 4.0.2 with the URL PACS integration. Philips was informed that when loading a study to an analysis application while using URL PACS integration configuration, after another study was opened by a review application, the Analysis monitor is updated with the study of the current patient, but the Review monitor conti | Software design | Philips sent a Urgent Medical Device Correction letter dated February 12, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.They advised their customers that when reviewing two disparate image datasets, not to only be focused on the actual image content but to also be alerted to the patient's study name and/or ID prior to interpretation.They recommend to manually exit the Review and Analysis applications before loading a new study to Intellispace Portal v4.0.2 from the PACS.For North America and Canada contact the Customer Care Solutions Center at (1-800) 722-9377. |  | 313 | 313 units | Worldwide Distribution - (USA) Nationwide Distribution including the states of AL, AR, AZ, CA, FL, GA, IA, IN, LA, MI, MN, MO, NC, NJ, NY, OH, OK, OR, PA, TN, TX, WA, WI &amp; WV. and the countries of Australia, Austria, Belgium, Canada, China, Czech Republic, Denmark, Egypt, Finland, France, France, Germany, Greece, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Latvia, Lebanon, Monaco, Netherlands, Norway, Poland, Portugal, Russia, Slovakia, South Africa, South Korea, Spain, Sweden, Switzerland, Thailand &amp; United Kingdom | Terminated | November 04, 2013 | 265 | 313 | Computer
Z-1647-2013 | 65207 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | 1) VITROS 5,1 FS Chemistry System (product code 6801375) and 2) VITROS 5,1 FS Chemistry System Refurbished (product code 6801890) |  |  | May 06, 2013 | 2013 | Ortho-Clinical Diagnostics | There is evidence that an electrical short within the systems connector has caused an accumulation of heat which has radiated back through the wiring harness causing the wires and connectors to melt. | Device Design | Ortho Clinical Diagnostics (OCD) sent an Urgent Product Correction Notification letters dated May 6, 2013 and response forms to all affected customers. Ortho Clinical Diagnostics informed their customers of the issue and advised if they detect the smell of melting plastic or observe smoke originating from your VITROSÂż System, to immediately power down the system and contact OCD for assistance. They were also advised to consider performing the backup procedure more frequently to help prevent the loss of data on their Systems. In the near future, an OCD Field Engineer will contact customers who are affected and perform a Modification (MOD) procedure that will replace their potentially affected Wire Connector(s). Foreign affiliates were informed by email on May 6, 2013 of the issue and instructed to notify their customers of the issue and actions. OCD also found that this wire connector is commonly used in many types of computers (i.e., medical devices, personal computers, etc.). On 29 May 2013, OCD expanded their recall to incllude ORTHO Assay Software (OAS) Servers. As a precaution, OCD sent notifications to encompass all ORTHOÂż Assay Software (OAS) Servers that potentially contain this type of wire connector. The letters (Ref. CL13-154) were sent via US Priority Mail to all customers who were shipped an affected OAS Server. Ortho Clinical Diagnostics has requested that customers with any questions or in need of additional information, please call Customer Technical Services at 1-800-421-3311. |  | 101 | Total 101 units (84 units for # 6801375- 26 units domestic &amp; 58 units foreign) and (17 units for # 6801890-15 units domestic, 2 units foreign) | Worldwide Distribution - USA Nationwide including the states of AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NY, OH, OK, OR, PA, SC, SD, TN, TX, VA, WA, WI &amp; WV, and countries of: Australia, Brazil, Canada, Chile, China, Colombia, India, Japan, Mexico, PR, Panama, Singapore, Venezuela, England, France, Germany, Italy, Spain | Open, Classified |  |  | 264 | Computer
Z-1646-2013 | 65139 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Computed Tomography X-ray systems |  |  | April 15, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | Philips Healthcare received reports from the field that certain Brilliance CT 64, Ingenuity Core, and Ingenuity Core128 systems running software versions 2.6.1 or 3.5.4 exhibited intermittent swirl-like ring artifacts that may appear on reconstructed images, which may impair the diagnostic quality of affected images. | Software design | Philips Healthcare sent an Urgent Medical device Correction letter dated April 15, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter states that Philips Healthcare is implementing software update versions 2.6.2 or 3.5.5 to resolve the issue and a Philips Field Service Engineer will contact each site to schedule the software installation. For questions contact your local Philips representative or local Philips Healthcare office. |  | 2348 | 2348 total units: 2182 Brilliance CT 64 units; 95 Ingenuity Core units; and 71 Ingenuity Core128 units. | USA Nationwide Distribution - including the states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY. | Terminated | January 12, 2015 | 637 | 2348 | Computer
Z-1105-2013 | 64737 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Siemens syngo WorkFlow SLR, Radiological Information System, Version VA31A SP1/SP2/SP3 |  |  | October 05, 2012 | 2012 | Siemens Medical Solutions USA, Inc | Under certain conditions, when using the Search functionality, the potential exists for the wrong patient's Interactive Documents to display when viewing a signed report in the single patient view. This can occur when the Interactive Document is viewed or auto displayed, dismissed by the user, and then a different patient and signed exam is selected in single patient view. The Notes Viewer will | Software design | Siemens Medical sent a Customer Safety Advisory Notice letter dated October 5, 2012 , to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.There will be a Hot Fix available.If you have any further questions please contact Siemens Customer Product Support at (1-888) 767-8326. |  | 37 | 37 | Worldwide Distribution - USA including the states of AL, CA, CO, DC, DE, FL, GA, HI, IL, IN, KS, KY, LA, MA, MD, MO, NC, NE, NH, NJ, NV, NY, OH, OR, PA, TN, TX, VA, WI, WV, and WY. and the countries of Canada and Australia. | Terminated | July 16, 2013 | 284 | 37 | Computer
Z-1620-2013 | 65301 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Bigliani/FlatowÂż The Complete Solution Fukuda Retractors Large Nonsterile 4305-19-40. |  |  | May 23, 2013 | 2013 | Zimmer, Inc. | This is a recall expansion from the 2010 recall which now includes seven additional lots that were not previously included. Zimmer, Inc. is initiating a recall of small and large Fukuda-type retractors due to fractures occurring at the grooves around the perimeter of the blades while being used during surgery. | Device Design | In 2010, Zimmer initiated a recall of the Bigliani/FlatowÂż Fukuda Retractors due to fractures occurring at the grooves around the perimeter of the blades. Zimmer is expanding this recall to include seven lots that were not previously included. On 5/23/13, an Urgent Medical Device Recall notification letter went to all consignees explaining the recall expansion and risk associated with the recall. All consignees are to locate and quarantine the device immediately. Please return the product along with the completed inventory certification form to Zimmer. If you have any questions regarding this recall, call Zimmer at 877-946-2761. |  | 722 | 722 units (includes 2010 market withdrawal) | Worldwide Distribution-USA (nationwide) including the states of NY, MI, MN, OH, IN, IL, VA, MD, PA, WA, KY, CT, RI, NH, TX, MT, and CA and the countries of Chile, China, South Korea, and Australia. | Terminated | December 18, 2014 | 574 | 722 | Computer
Z-1622-2013 | 65159 | No | LGX | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General Hospital | 510(K) Exempt | MAQUET Getinge Group Mobile Operating Table YUNO OTN GA 1433.01 EN 05. |  |  | April 26, 2013 | 2013 | Maquet Cardiovascular Us Sales, Llc | MAQUET has identified a potential issue with the extension device handle (part number 1433.62A1) and will make enhancements to the YUNO table (part number 1433.01F0).Identification of the Extension device issue: By fully actuating the lever of the extension device handle, the operator's skin can enter a gap between the handle lever and a recess in the handle bar. When the lever is released, | Device Design | MAQUET Getinge Group issued an Urgent Medical Device Recall letter/"Field Action Response Form dated April 26, 2013, via Federal Express to all affected customers. The recall notification identifies the product, issue with the product, and measures to be taken by the user. Customers were instructed that the firm would contact them to arrange for a loaner YUNO table and extension device while the table and extension device are upgraded. Upon completion of the upgrades, MAQUET would reinstall their upgraded YUNO table and extension device and remove the loaner device. Customers were instructed to complete and return the enclosed Field Action Response Form, acknowledging their receipt and understanding of the communication by e-mailing a scanned copy to FieldAction@maquet.com or via fax at 973-807-1658. Customers with questions were instructed to contact their MAQUET representative or MAQUET Technical Support at 888-627-8383, option 3, followed by option 1.For questions regarding this recall call 973-709-7691. |  | 37 | 13 Tables; 24 Extension devices | Worldwide Distribution - USA including Arizona, Georgia, North Dakota, South Carolina, Texas and Tennessee, Internationally to United Arab Emirates, Austria, Brazil, Canada, Switzerland, China, Germany, Spain, France, United Kingdom, Italy, Japan, Singapore, Thailand, Turkey, Serbia, and South Africa. | Terminated | April 29, 2014 | 368 | 37 | Computer
Z-1616-2013 | 65022 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Innova 3100 and Innova 4100 systems. |  |  | May 16, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare reported a potential safety issuedue to temporary loss of X-Ray Imaging Modes during Collimator Auto-Calibration which affects Innova 3100 and Innova 4100 Systems. | Software design | A Customer Letter was not originally sent. GE Healthcare has retrospectively re-assessed the reportability of the field action and has determined that the action should be reported to the FDA.GE Healthcare has already updated all of the affected Innova 3100 and Innova 4100 systems that were distributed.As an effectiveness check, they have tracked all service requests for the field action to completion. These service requests have recorded the repair of the affected units. Exceptions to completion of the service request have been recorded as well. Records of field action completion have been retained per their processes and procedures.If you have any questions, please call 262-548-2757.For questions regarding this recall call 262-513-4122. |  | 48 | 48 | Worldwide distribution - USA including AZ, CA, CO, FL, ID, IL, ME, NV,NY, NC, OK, TN, UT, and WI; and Internationally to Canada, Chile, France, Germany, Italy, Japan, Korea, Norway and Spain. | Terminated | March 13, 2014 | 1032 | 48 | Computer
Z-1609-2013 | 65364 | No | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | Brand Name: Coiled Tube Infusors (Portable Elastomeric Infusion Systems). Baxter's Infusor SV Elastomeric Infusion Device |  |  | June 07, 2013 | 2013 | Baxter Healthcare Corp. | Due to an increase in complaints for leaks at the distal male Luer and Luer cap. Baxter implemented enhancements to the blue winged cap to improve the seal between the cap and the Luer. | Device Design | Baxter sent an Urgent Produd Recall notification letter dated June 13, 2013, to all affected consignees via USPS.The letter was accompanied by a Customer Reply Form and Affected Product List (table). The letter stated affected products, problem description, risk to health (or health risks), actions to be taken by customers/users and contact information.The recall notification instructs customers to segregate affected product from current inventory and contact Baxter for instructions to return affected product. Customers with questions were instructed to call Baxter at 1-800-437-5176,For questions regarding this recall call 847-270-5556. |  | 601230 | 601,230 total | Worldwide Distribution - USA (nationwide), Puerto Rico, and Internationally to Brazil, Colombia, Panama, Mexico, Singapore Japan, Australia, New Zealand, Hong Kong, India, Thailand, Taiwan, and Korea. | Terminated | June 04, 2014 | 362 | 601230 | Computer
Z-1124-2013 | 64679 | No | NRA |  |  | Orthopedic | PMA | REF 32-422760 |  |  | March 20, 2013 | 2013 | Biomet U.K., Ltd. | Investigation of complaints found excessive use of the instrument could result in cracking on the weld between the head and the handle, which could lead to part of the head becoming detached from the handle. | Device Design | BIOMET sent an Urgent Medical Device Recall Notice on March 20, 2013, to all affected customers with directions to forward the notice if product was further distributed to Hospital Operating Room managers. The notice identified the recalled device, reason for the recall, and possible adverse events directing consignees to locate and discontinue use of the device. A FAX back Response form was to be returned to 574-372-1683 within 3 days of receipt prior to return of the product to Biomet, USA. Questionsregarding this recall should be addressed to 574 372-1570 M-F, 8a.m. -5 p.m. |  | 430 | 430 | Worldwide Distribution - USA (nationwide), Europe, Japan, Australia, Canada, Singapore, South Africa, and Costa Rica. | Terminated | May 04, 2015 | 775 | 430 | Computer
Z-1112-2013 | 64653 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris PC unit model 8015 with software version 9.12 |  |  | March 06, 2013 | 2013 | Carefusion Corporation | The recall was initiated because Carefusion has identified potential risk associated with bolus programming with the Alaris PC unit (model 8015) software version 9.12. Automating infusion pump parameter input from either the electronic medical record (EMR) or the Alaris Auto-ID module for a continuous infusion with a bolus dose option may result in a loss of information from the Guardrails Data s | Software design | Carefusion sent an Urgent Medical Device Recall Notification letter dated March 6, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers are required to confirm receipt of the notification by returning the Recall response card to Carefusion by postage-paid, self-addressed mail, fax, or email. The letter stated that a Carefusion representatives will contact customers by telephone to schedule field remediation. For questions call Carefusion Support Center at 888-562-6018, Customer Advocacy at 888-812-3266, and/or Technical Support at 888-812-3229. |  | 4090 | 4,090 total units | Worldwide Distribution USA Nationwide and the countries of Australia and Canada. | Terminated | May 08, 2014 | 428 | 4090 | Computer
Z-1113-2013 | 64460 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ACCU-CHEK Inform II Blood Glucose Monitoring System bar code scanner. |  |  | March 06, 2013 | 2013 | Roche Diagnostics Operations, Inc. | Roche investigation into complaints found Accu-Chek Inform II and CoaguChek XS Pro System barcode reading meters had the potential of intermittent erroneous decoding of patient identification if barcodes being read fail to meet size and/or quality standards. | Device Design | Consignees of both devices were notified by Urgent Medical Device Correction letters sent via UPS Ground (receipt required) beginning 3/06/2013. The notice contained recommendation not to use EAN-13 barcode system and ensure barcodes used meet quality standards; use operator and patient lists (downloaded from a data manager) in order to check any decoded identification against the respective list. The notice directed forwarding of the notice if the product was further distributed. Customers were asked to complete the attached fax form. Contact the Customer Care Service Center, 24 hours a day, seven days a week at 1-800-440-3638 for Accu check questions and 1-800-428-4674 for CoaguChek questions about the information contained in this UMDC. |  | 4467 | Both devices 4467 | Nationwide Distribution. | Terminated |  |  | 4467 | Computer
Z-1580-2013 | 65443 | No | LKK |  |  | General Hospital | PMA | Medtronic SynchroMed EL Programmable pumps, Models 8626, 8626L, 8627, 8627L, (The SynchroMed EL Pumps are supplied in 10 ml or 18 ml reservoir size.) Sterilized using ethylene oxide. The pump is part of an infusion system that stores and delivers a prescribed drug to a specific site. The drug is stored in the pump reservoir. Per a programmed prescription, the drug moves from the pump reservoir, through the pump tubing, catheter port, and catheter, to the infusion site. |  |  | June 03, 2013 | 2013 | Medtronic Neuromodulation | Medtronic Neuromodulation is providing Healthcare Providers with safety information and patient management recommendations related to the potential for electrical shorting internal to they SynchroMed infusion pump. An electrical short circuit in a feedthrough may present as a motor stall or low battery reset/alarm and lead to a loss of or reduction in therapy which may result in the return of und | Device Design | Medtronic sent an Urgent Medical Device Correction letter dated May 2013, to all affected customers. The letter described the issue of the pump's potentially shorting, how to identify if a pump is affected, and provided ongoing patient management recommendations. The Medtronic Rep would complete a Physician Notification Confirmation Form. A copy of the recall notification letter and the completed Confirmation Form is to be left with the physician. Customers with questions were instructed to call 1-800-707-0933.For questions regarding this recall call 763-526-9738. |  | 261109 | 261,109 total devices estimated implanted Worldwide | worldwide | Terminated | November 16, 2015 | 896 | 261109 | Computer
Z-1730-2013 | 60480 | No | DRY | Cardiopulmonary bypass on-line blood gas monitor. | Cardiovascular | Cardiovascular | 510(k) | CDI Blood Parameter Monitoring System 500 with Arterial and Venous blood parameter modules and Hematocrit/Saturation probe. |  |  | February 27, 2013 | 2013 | Terumo Cardiovascular Systems Corporation | Terumo discovered through its own testing that CDI 500 Blood Parameter Monitoring System software contains inaccuracies associated with a software algorithm, and a software update has been developed to address this issue and a premarket notification (K130425) has been submitted to the agency. Once clearance is received for this premarket notification from the US FDA, Terumo CVS will implement the | Software design | On 2/27/13, all consignees will receive a recall notification letter via Federal Express. Terumo CVS will correct all affected CDl Monitors. Correction to the devices will be completed concurrently with BPM replacements related to Z-2185/2189-11 RES 57194.Terumo continues to manufacture and ship devices affected by this recall. The long term corrective actions are associated with Recall# Z-2185/2189-11 RES 57194. As committed previously, we will re-notify all customers once we are ready to implement the fix for the BPM in the field.All units produced prior to the change in manufacturing will be included in thescope of this correction. Terumo continues to manufacturer units affected by this correction.All units returned for this correction will be tagged with a service traveler to indicate reason for correction and units will be properly segregated to prevent unintended re-shipment. |  | 15888 | 3972 units since 4/26/12 | Worldwide Distribution-USA (nationwide) including DC and Puerto Rico and the states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, and WV and the countries of Jordan, CANADA, AUSTRALIA, Japan ,Hong Kong ,UNITED ARAB EMIRATES (UAE), BELGIUM, South Korea, Malaysia, Mexico, Philippines, Singapore, Taiwan, and Thailand. | Terminated | February 17, 2016 | 1085 | 15888 | Computer
Z-1739-2013 | 63109 | No | LGW |  |  | Neurology | PMA | Eon Mini - 3721 (Charger). |  |  | July 26, 2012 | 2012 | St. Jude Medical | St. Jude Medical has informed your doctor that a number of cases have been reported in which discomfort associated with heating occurred at the device site while patients were using the Charging System to charge their spinal cord stimulator. In three cases, patients received a burn to the skin (one second-degree and two first-degree burns) at the implant site. As of June 30, 2012, St. Jude Medica | Device Design | ST. Jude Medical sent an Important Medical Device Safety Information dated July 26, 2012 to affected customers. The letter identifed the affected product, problem and actions to be taken. For questions or concerns contact your ST. Jude Medical representative or St. Jude Medical Neuromodulation Division at 1-866-240-6741. |  | 70638 | 70,638 (combined total of the Eon and the Eon Mini systems) | Worldwide Distribution - USA Nationwide Including Puerto Rico and countries of: Argentina, Australia, Belgium, Brazil, Canada, China, Chile, Colombia, Ireland, Israel, India, Malaysia, Mexico, Saudi Arabia, Singapore, South Africa, South Korea, Switzerland and United Kingdom. | Terminated | November 18, 2015 | 1210 | 70638 | Computer
Z-1864-2013 | 64625 | No | KTI | Bronchoscope (flexible or rigid) and accessories. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | Inspira AIR Balloon Dilation system. |  |  | January 11, 2013 | 2013 | Acclarent, Inc. | Labeling correction for all sizes of the Inspira AIR Balloon Dilation System to include additional language and warnings as well as to clarify some instructional steps. | Device Design | Urgent Medical Device Labeling correction letters will be sent to all consignees through a delivery system that can track delivery. After CDRH review and corrections, Acclarent began issuing customer notification on August 7, 2013 |  | 1217 | 1217 | Worldwide Distribution-USA (nationwide) and the countries of Belgium, Canada, and Saudi Arabia. | Terminated | October 25, 2013 | 287 | 28660 | Computer
Z-1834-2013 | 65440 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AU5800 Clinical Chemistry Analyzer with Automation Ready Connection (All Software Versions), Model Numbers AU5811, AU5821, AU5831 and AU5841. This system is for in vitro diagnostic use only. |  |  | June 03, 2013 | 2013 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has identified a software error on AU5800 analyzers with the Automation Ready (AR) Sampler Kit. This software error results in the failure of the concentrated wash solution pump to stop, creating an overflow or leak on each analyzer unit. | Software design | An Urgent Product Correction letter dated 5/29/13 was sent on the week of June 3, 2013 to all the customers who received the AU5800 Clinical Chemistry Analyzer with Automation Ready (AR) Connection (All Software versions). The letter informed the customers of the problem identified and the actions to be taken. Customers were instructed to complete and return the enclosed Response Form within 10 days and to share the information with their laboratory staff. Customers with questions are instructed to contact the 24 hours Call Center (Hotline) at (800) 854-3633 in the US and Canada. |  | 76 | 76 units | Worldwide Distribution, including Nationwide (US) and the countries of China, Italy, Japan, Canada, Turkey, United Kingdom, Singapore, Israel, and the Netherlands. | Terminated | August 23, 2014 | 446 | 76 | Computer
Z-1791-2013 | 65579 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Discovery NM 630, dual detector free-geometry integrated nuclear imaging system. |  |  | June 13, 2013 | 2013 | GE Healthcare, LLC | GE became aware of an incident at a VA Medical Center facility in the US in which a patient died due to injuries sustained while being scanned on an Infinia Hawkeye 4 Nuclear Medicine System. UPDATE: On July 03, 2013 GE notified hospitals that they are recalling several Nuclear Medicine Imaging Systems because serious injuries or deaths could occur due to the failure mode associated with this re | Device Design | GE sent an Urgent Medical Device Correction letter dated June 17, 2013 to all affected customers. The letter identified the affected product, recommended that qualified service personnel maintain the equipment and that Preventative Maintenance procedures are executed according to labeling. Also the Safety Chapter Sections should be re-reviewed with personnel to ensure proper operation of the equipment. GE notified customers again on July 03, 2013, via Urgent Medical Device Recall letter (including confirmation of delivery for US customers) and follow-up telephone call. Customers were instructed to cease using the affected system until GE Healthcare Field Engineer can complete an inspection of the system and perform any necessary repairs at no cost. A GE Healthcare representative will contact the hospitals to arrange for the inspection. |  | 108 | 108 | Worldwide Distribution - US Nationwide including Guam, Puerto Rico and the countries of : Afghanistan, Algeria, Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, Finland, France, Georgia, Germany, Greece, Hong Kong, Hungary, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jordan, Korea (Republic of), Kuwait, Latvia, Lebanon, Lithuania, Malaysia, Malta, Martinique, Mexico, Morocco, Namibia, Netherlands, New Zealand, Norway, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Puerto Rico, Romania, Russia, Saudi Arabia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Ukraine, United Kingdom, Uruguay, Venezuela, Vietnam. | Terminated | June 19, 2014 | 371 | 3454 | Computer
Z-1062-2013 | 64200 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Flexible Grip which is part of the Synthes Universal Nail System. |  |  | January 11, 2013 | 2013 | Synthes USA HQ, Inc. | Synthes is initiating a voluntary recall of the Flexible Grip (Part number 355.28) which is part of the Universal Nail System, due to the potential lack of the ability to thoroughly clean the instrument. | Device Design | Synthes sent an "URGENT NOTICE: MEDICAL DEVICE RECALL" dated January 11, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. The letter requested consignees examine inventory, remove the recalled products from use and return them to Synthes. If you have any questions, contact the firm at 610-719-5450. |  | 328 | 328 | Nationwide Distribution including the states of AZ, CA, CO, FL, IA, ID, GA, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, NC, NH, NJ, NV, NY, OH, PA, RI, TN, TX, UT, VT, WA, WI, and WV. | Terminated | August 31, 2015 | 962 | 328 | Computer
Z-1813-2013 | 65328 | No | DSQ |  |  | Cardiovascular | PMA | Heartware HVAD Pump Implant Kit (Heartware Ventricular Assist Device) |  |  | May 16, 0013 | 13 | HeartWare Inc | Potential to electrostatic discharged (ESD) | Device Design | HeartWare sent an Urgent Medical Device Correction Safety Notification letter on May 16, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to educate their patients about the HeartWare Controller and instructions for equipment handling stated in the IFU and recognized practices to avoid electrostatic discharge. The direct accounts were also told to scan and return the signed Acknowledgment Form via fax to 305-364-2665 or e-mail to quality@heartwareinc.com.Customers with questions were instructed to contact their HeartWare Representative or HeartWare Customers Service.For questions regarding this recall call 305-364-1575. |  | 76 | 76 (US) and 82 (International) | Worldwide Distribution - USA including AL, AZ, CA, DC, DE, FL, GA, IA, IL, IN, KY, LA, MA, MD, MI, MN, MO, NC, NJ, NY, OH, OR, PA, TN, TX, UT, VA, WA, and WI. Internationally to:Australia, Austria, Belguim, Brazil, Canada, Chile, Croatia, Czech Republic, France, Germany, Greece, India, Israel, Italy, Japan, Lebanon, Lithuania, Luxemburg, Malaysia, Netherlands, New Zeland,Norway, PolandSingapore, South Africa, Switzerland, Turkey, and United Kingdom. | Open, Classified |  |  | 76 | Computer
Z-1818-2013 | 65718 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | ViSi Mobile Chest Sensor (3 lead-wire ECG, AAMI), Catalog No. 93-10014 |  |  | May 23, 2013 | 2013 | Sotera Wireless, Inc. | Sotera Wireless, Inc. is recalling the ViSi Mobile Monitoring System due to the potential for the Chest Sensor to cause a blister wound. | Device Design | Sotera sent an Urgent Product Recall letter dated May 23. 2013, to all affected customers. The letter informed the customers of the problems identified and the actions to be taken. Customers were instructed to report any adverse reactions experienced with the use of this product and or quality problems to the FDA's MedWatch Program by: www.fda.gov/medwatch, (800) FDA-1008, fax (800) FDA-0178. Customers were also instructed to contact Sotera Wireless Technical Support at (866) 794-5526 for questions. |  | 75 | 75 units | US distribution including CA and UT. | Terminated | September 09, 2013 | 109 | 165 | Computer
Z-1814-2013 | 65666 | No | NSX |  |  | General Hospital | Contact ODE | Picis ED Pulsecheck - EMR Software Application - 2125, Software Versions: 5.2 and 5.3. The application stores patient information in a database, and it may analyze and/or display the data in different formats for evaluation by healthcare professionals for informational purposes. |  |  | June 21, 2013 | 2013 | Picis Inc. | Notes associated with prescription are not printed to the prescription or to the patient chart. | Software change control | Initial customer notifications were sent via email on June 21, 2013 informing consignees of the recall and providing further instruction regarding the software solution. Consignees were provided with recommended actions until they receive the necessary update. |  | 35 | 35 | Nationwide Distribution, including the states of: AK, AR, AZ, CA, CO, DC, DE, FL, GA, ID, IN, MA, MD, MO, NH, NJ, OH, OR, SC, TN, WA, and WV. | Terminated | December 19, 2014 | 546 | 35 | Computer
Z-1815-2013 | 63935 | No | IEO | Automated tissue processor. | Pathology | Pathology | 510(K) Exempt | Leica ASP6025 Vacuum Tissue Processor |  |  | November 13, 2012 | 2012 | Leica Microsystems, Inc. | The Instruction for Use 1v9 Ref F and all previous versions specify incorrect dimensions for the Ready To Use (RTU) bottles from other suppliers than Leica. In case of a pressure failure of the device in combination with a bottle with incorrect dimensions, overflow can cause contamination of reagents. Additionally, the current instructions for use does not highlight the importance to check the filling levels of all reagent bottles (RTU and system bottles) prior to every process run.Without this important information, risk of tissue damage or loss is possible. Therefore the missing information will be distributed to all customers by a Field Service Notice. | Device Design | Leica Biosystems sent an Urgent Field Safety Notice dated December 12, 2012, to all affected customers. The notification included a description of the problem and a recommendation that the personnel who use the equipment only use Leica Ready To Use (RTU) bottles that are already supplied as a standard delivery item with the ASP6025 machine. The personnel are further reminded that reagent bottle levels need to be checked and corrected when necessary. Customers with questions or questions regarding this recall call 847-405-5413. |  | 128 | 128 machines | Worldwide Distribution - USA (nationwide) and internationally to Australia, Belgium, Brazil, Chile, France, Germany, Ireland, Italy, Japan, Malaysia, Mexico, Netherlands, Philipines, Singapore, South Africa, Spain, Switzerland, Thailand, Turkey, and United Kingdom. | Terminated | January 05, 2017 | 1514 | 128 | Computer
Z-1050-2013 | 64386 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ABX Micros IM2, Part# M60CSB110EN04 and D00A00202 |  |  | February 15, 2013 | 2013 | Horiba Instruments, Inc dba Horiba Medical | Horiba Medical is recalling ABX Micros IM2 Data Management System because it may display incorrect patient test results under unanticipated operating conditions. | Device Design | Horiba Medical sent an Urgent Safety Notice dated February 15, 2013, to all affected customers. The letter informed the customers that Horiba Medical is informing all ABX Micros IM2 Data Management System customers of their intention to remove the ABX Micros IM2 product from the market. The letter informed the customers of the problem identified and the action to be taken. Customers were instructed to complete and return the enclosed response form within 10 days.For questions regarding this recall call 949-453-0500, ext 208. |  | 180 | 180 units | Nationwide Distribution | Terminated | August 11, 2014 | 542 | 180 | Computer
Z-1690-2013 | 65420 | No | LKK |  |  | General Hospital | PMA | MedStream Programmable Infusion Pump- Implantable for the intrathecal delivery of Baclofen or Morphine |  |  | June 07, 2013 | 2013 | Codman &amp; Shurtleff, Inc. | Miscalibrated Fill Level Sensor may affect dosing level | Device Design | Codman issued a Medical Device Correction Notification to Medical Facilities dated June 7, 2013 via FedX to inform them of the potential problem and a Worksheet (instructions) to Identify Pumps with Miscalibrated FLS during each patient's next scheduled refill session or sooner if patient is symptomatic. If a pump is identified as having a miscalibrated FLS, notify Codman Neuro using the same worksheet. In addition a sample letter was provided to Physi cans to provide to patients on the system.Follow the recommendations for managing patients who are identified as having a pump with a miscalibrated FLS.For additional information,contact your Codman Neuro representative, or contact Codman Neuro Clinical Support at 1-800-660-2660. Please report any malfunction or adverse event related to this device to Codman Neuro at 1-866-491-0974, Option 2. |  | 737 | 2 units US and 735 OUS | Worldwide Distribution-USA (nationwide) including the states of FL, NC, NJ, and TX, and the countries of Austria, Belgium, Spain, Great Britain, Germany, Italy, Ireland, Luxemburg, Russia, Saudi Arabia, Portugal, Greece, Egypt, Switzerland, and Turkey. | Terminated | April 21, 2015 | 683 | 1102 | Computer
Z-1836-2013 | 63108 | No | LGW |  |  | Neurology | PMA | Eon Mini - 3788 (IPG), |  |  | July 26, 2012 | 2012 | St. Jude Medical | As part of St. Jude Medical's routine tracking and product monitoring, they have received a total of 214 reports, out of 34,617 Eon Mini IPGs, that lost the ability to communicate or recharge due to an inner battery weld issue resulting in loss of pain relief and subsequent explant. Explant surgery, as with any surgery, presents a risk to patient health. Adverse events associated with an unplanne | Device Design | St. Jude Medical sent a Important Medical Device Recall Advisory dated July 26, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.To further assist in your patient care we are providing you with a list of all serial numbers we show have been distributed to you. ( See Attachment A ).Further questions please call (972) 309-2154. |  | 34617 | 34,617 | Worldwide Distribution--USA (nationwide) including Puerto Rico and the countries of Argentina, Australia, Austria, Belgium, Brazil, canada, Colombia, Cyprus, Czech Republic, Denmark, Finland, France, Germany, Greece, India, Ireland, Israel, Italy, Japan, Mexico, Netherlands, Saudi, South Africa, Spain, Sweden, Switzerland, Turkey and United Kingdom. | Terminated | August 21, 2015 | 1121 | 34617 | Computer
Z-1833-2013 | 65755 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | GE Healthcare, VOLUSON E6, E8, E8 EXPERT, E10 DIAGNOSTIC ULTRASOUND SYSTEM. |  |  | July 13, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has recently become aware that during use of the Voluson E6 or E8 systems, that an error message Out of Memory can be displayed which could cause the system to freeze. This is an intermittent behavior that also requires a certain workflow. This issue can be resolved by rebootingthe system. | Software design | No consignee letter was sent out.Starting on July 13, 2012, Field Engineers visited consignees and carried on instructions as per Field Modification Instruction (FMI 79044A). The Engineer returned the system to the customer and ensure acceptance. Before leaving the customer site have the staff operate the system. Ensure that: all settings are customized, all peripherals are functioning and all communications are working properly. We plan to update the software on all of the affected GE VOLUSON E6, E8, E8 EXPERT, E10 DIAGNOSTIC ULTRASOUND SYSTEM that were distributed.As an effectiveness check, we will track all service requests for the field action to completion. These service requests will record the repair of the affected units. Exceptions to completion of the service request will be recorded as well. Records of field action completion will be retained per our processes and procedures.Further questions please call 262-513-4122. |  | 375 | 375 | Worldwide Distribution - US Distribution including the states AK, AZ, AR, CA, CO, CT, DE, FL, GA, IL, LA, MD, MI, MN, MO, NE, NV, NM, NY, NC, OH, OR, PA, SC, SD, TN,TX, VA, WA and WI., and the countries of ALGERIA, AUSTRALIA, AUSTRIA, BAHRAIN, BELGIUM, BRIL, CANADA, CHINA, COLOMBIA, DENMARK, EGYPT, FINLAND, FRANCE, GERMANY, INDIA, IRAQ, ISRAEL, ITALY, JAPAN, JORDAN, KOREA, REPUBLIC OF MOLDOVA, NETHERLANDS, NORWAY, POLAND, ROMANIA, RUSSIAN FEDERATION, KINGDOM OF SAUDI ARABIA, SLOVAKIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, UZBEKISTAN, VENEZUELA. | Terminated | November 10, 2014 | 850 | 375 | Computer
Z-1837-2013 | 65498 | No | BXB | Powered exercise equipment. | Physical Medicine | Physical Medicine | 510(K) Exempt | Omnicycle Elite; US Importer and Distributor: Accelerated Care Plus, Reno NV. Manufactured by Medica Medizintechnik GmbH, Hochdorf, Germany. The product is indicated for powered exercise equipment intended for medical purposes. |  |  | June 12, 2013 | 2013 | Accelerated Care Plus Corporation | An anomaly in the firmware that may cause Omnicycle Elite to operate at a speed other than what was previously set when in reverse mode. | Software design | Letters were issued on June 12, 2013 that included a memory stick with self loading software so consumers can update the device. Letters were sent by trackable method. Customers with questions or concerns should call Customer Support at 800-350-1100. |  | 1001 | 1001 units | Nationwide Distribution. | Terminated | October 23, 2013 | 133 | 1001 | Computer
Z-1820-2013 | 65591 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | LPS Lower Extremity Dovetail Intercalary component. |  |  | July 11, 2013 | 2013 | DePuy Orthopaedics, Inc. | The LPS Lower Extremity Dovetail Intercalary component has the potential for fracture of the female component when exposed to certain physiological loads. | Device Design | On 7/11/13, URGENT MEDICAL DEVICE RECALL NOTIFICATIONS were sent to consignees: Depuy Distributors, hospitals, and surgeons. End-user Hospital Customers will submit reconciliation forms to DePuy. Possible clinical implications of device failure are stated in these notifications, including poor mechanics and loss of function, pain, component dislocation, adverse tissue reaction, bone damage, potential need for revision surgery and the risks associated with this type of surgery. A patient letter will be included with the surgeon letter to assist in surgeon's notification and discussion with patients. The devices will be returned to DePuy Warsaw through the normal DePuy Returns process, to attention of Returns and marking H13-13 on the outside of the box.If a patient presents with a fractured LPS" Lower Extremity Dovetail Intercalary component with well-fixed proximal and distal stems and the surgeon determines LPS" Lower Extremity Dovetail Intercalary component is the best treatment option, DePuy plans to make the LPS" Lower Extremity Dovetail Intercalary component available through authorized approvals. Directions on how to purchase the component are provided in the Recall Communications.Any questions or concerns about the recall, please contact Kim Earle, Recall Coordinator, at 574-371-4917 (M-F; 8am-5pm EDT) |  | 2262 | 58 U.S. / 39 OUS | Worldwide Distribution-USA including the states of FL, GA, ID, KS, KY, MI, MO, NC, NV, NY, OH, OK, PA, TN, TX, VA, and WA and the countries of AUSTRALIA, CANADA, ENGLAND, GERMANY, INDIA, COUNTY CORK, IRELAND, and TURKEY. | Terminated | July 10, 2014 | 364 | 2262 | Computer
Z-1859-2013 | 65603 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo Dynamics Information System (sDIS) versions 9.5 and VA10A. |  |  | June 10, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Siemens initiate this recall due to a potential issue when using the sDIS matching rules of the syngo Dynamics Information System version 9.5 or VA10A with modality device configured to end study at association close. Images or DICOM SR objects sent from a modality device to syngo Dynamics may not be saved when the modality is configured at syngo Dynamics to end study at association close, and whe | Software design | Siemens sent a Customer Safety Advisory Notice dated June 4, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. 1) With sDIS implemented, use the DICOM Modality List (DMWL), without sDIS matching rules enabled OR 2) Do not implement sDIS. Update instructions were also provided. The reported issue will be resolved in a modification to syngo Dynamics version 9.5 and syngo Dynamics VA10A which will be released in the near future.We regret any inconvenience that this may cause , and we thank you in advance for your understanding.For further questions please call (610) 219-6300. |  | 9 | 9 | US Distribution only including the states of PA, OH, UT, CA, NY, MO, and TX. | Terminated | February 24, 2015 | 624 | 9 | Computer
Z-1869-2013 | 64987 | No | FSM | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Caddy - PedFuse, Screw Caddy, 8MM and 9MM, Part Number: 11-80009 Rev. B |  |  | October 05, 2012 | 2012 | SpineFrontier, Inc. | Two sections that are designated for the 8mm and 9mm screws were switched; the holes that hold the 8mm screws are designed for 9mm screws, and the holes that hold the 9mm screws are designated for 8mm screws. | Device Design | The firm, SpineFrontier, sent an Advisory Notice entitled "URGENT May 1, 2013, PedFuse Screw Caddy Removal" dated May 1, 2013 to their consignees/customers. The notice described the product, problem and action to be taken. Consignees/customers were informed that a removal has been initiated to remove the parts from the field to ensure that all sets contain correctly marked screw caddies. Consignees/customers wre instructed to complete and return the URGENT PedFuse Screw Caddy Removal ACKNOWLEDGEMENT FORM via fax to: SpineFrontier, Inc., 978-232-3991 or email: customerservice@spinefrontier.com. Replacement caddies will be shipped upon receipt of completed acknowledgement form and RMA request for part return.Please feel free to contact SpineFrontier at 978-232-3990 with any questions. |  | 11 | 11 caddies | US distribution in states of: CO, DC, MA and TX. | Terminated | July 29, 2015 | 1027 | 11 | Computer
Z-1797-2013 | 65468 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart MRx Monitor/Defibrillator, |  |  | June 19, 2013 | 2013 | Philips Healthcare Inc. | The MRx may deliver a non-synchronized cardioversion shock resulting in the delivery of incorrect therapy, which may induce ventricular fibrillation. | Software design | Philips sent an Urgent Medical Device Correction letter on June 19, 2013, to all affected customers. The letter informed customers of the issue, identified details of the units affected, gave instructions on actions to be taken by the customer and identified what action Philips planned to take to remedy the issue.Philips will perform a software upgrade and replace the label on the unit that outlines the steps for Synchronized Cardioversion with a revised label for all affected units at no charge to the customer.Philips is asking customers to follow the Action to be Taken by Customer/User section of the Urgent Medical Device Correction notification/Field Safety Notice: You can continue to use your MRx prior to receiving the software upgrade and device label revision, provided that you follow the instruction provided.Customers with questions were instructed to contact their local Philips representataive or call 1-800-722-9377.For questions regarding this recall call 978-687-1501. |  | 66150 | 66,150 units | Worldwide Distribution - USA (nationwide) and countries of: AFGHANISTAN ALBANIA ARGENTINA ARUBA AUSTRALIA AUSTRIA AZERBAIJAN BAHRAIN BANGLADESH BELGIUM BOLIVIA BOSNIA AND HERZEGOWINA BRAZIL BRUNEI DARUSSALAM BULGARIA CAMBODIA CHILE CHINA COLOMBIA COSTA RICA CROATIA (local name: Hrvatska) CYPRUS CZECH REPUBLIC DENMARK ECUADOR EGYPT EL SALVADOR ESTONIA ETHIOPIA FINLAND FRANCE GERMANY GREECE GUADELOUPE HONG KONG ICELAND INDIA INDONESIA IRAQ IRELAND ISRAEL ITALY JAPAN JORDAN KAZAKHSTAN KENYA KOREA, REPUBLIC OF KUWAIT LAOS LATVIA LEBANON LIBYAN ARAB JAMAHIRIYA LITHUANIA LUXEMBOURG MACAU MACEDONIA MALAYSIA MALTA MEXICO MONGOLIA MOROCCO NAMIBIA NEPAL NETHERLANDS NETHERLANDS ANTILLES NEW CALEDONIA NEW ZEALAND NICARAGUA NIGERIA NORWAY OMAN PAKISTAN PALESTINE PANAMA PAPUA NEW GUINEA PARAGUAY PERU PHILIPPINES POLAND PORTUGAL PUERTO RICO QATAR REUNION ROMANIA RUSSIAN FEDERATION SAUDI ARABIA SERBIA SINGAPORE SLOVENIA SOUTH AFRICA SPAIN SRI LANKA SWEDEN SWITZERLAND TAIWAN, PROVINCE OF CHINA TANZANIA, UNITED REPUBLIC OF THAILAND TRINIDAD AND TOBAGO TUNISIA TURKEY UGANDA UKRAINE UNITED ARAB EMIRATES UNITED KINGDOM UZBEKISTAN VENEZUELA VIET NAM and YEMEN. | Open, Classified |  |  | 66150 | Computer
Z-1795-2013 | 65604 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Axiom Aristos MX Solid State Flat Panel Digital Imager, Model number 5895003. The device is used for diagnostic X-ray imaging. |  |  | June 05, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Siemens issued a Field Safety Notice about the potential hazard to patients or operators when using the function "table top float" of the AXIOM Aristos MX. When repositioning a patient lying on the table top and releasing the table top, it is possible that the fingers of the patient or operator are between the table top and the detector tray and if the table top is moved, the fingers can be squee | Device Design | On 6/5/13, consignees were notified by letter of a "Field Safety Notice" and received instructions about how to avoid the potential risk of this issue while using the device. Siemens also rolled out a modification of AXIOM Aristos MX that will avoid these potential injuries in the future. |  | 401 | 401 | Nationwide Distribution. | Terminated | February 24, 2015 | 629 | 401 | Computer
Z-1794-2013 | 65259 | No | IOQ | Powered flotation therapy bed. | Physical Medicine | Physical Medicine | 510(K) Exempt | TotalCare Bed used with AccuMax Quantum VPC mattress |  |  | May 30, 2013 | 2013 | Hill-Rom, Inc. | When the bed is put into the chair egress position, the magnets on the mattress may not hold the bottom of the mattress flat to the bed frame, and the mattress foot section will extend away from the bed at an angle. A patient attempting to get back into the bed may miss the seat and sit on the foot section instead. The patient can slide down to the floor, or fall, with possible injury. | Device Design | The firm, Hill-Rom, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated May 29, 2013 was sent to all consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed to do the following: Until the correction is implemented, after placing the bed into chair position, regularly check the foot section of the mattress. If necessary, push it back against the bed frame until the magnets hold it in place before allowing the patient to get back into or out of bed. In addition, forward a copy of this letter to any other facility personnel you deem appropriate.If you have any questions concerning this notification, please contact your distributor, your Hill-Rom representative, or Hill-Rom Technical Support at 800-445-3720 (Mon-Thurs, 8 am-6:30 pm; Friday, 8 am-6 pm). |  | 4121 | 4,121 Units | Worldwide distribution: USA (nationwide) and countries of: Canada, Colombia, Nigeria, Kuwait and Saudi Arabia. | Terminated | April 17, 2014 | 322 | 4121 | Computer
Z-1782-2013 | 65651 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | ARTIS zee and zeego x-ray, angiographic system with Software Artis VC20x/VC21A that have Patch 7 and a 3D and DSA license installed. |  |  | June 26, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Siemens issued a customer safety advisory notice and field safety corrective action for a potential issue on Artis systems with Artis VC20x/VC21A that have Patch 7 and a 3D and DSA license installed. The issue affects-DSA rotational acquisition. The restriction is related to any rotational 3D-acquisition program with real-time subtraction (e.g. 3D DSA). During the fill phase of the rotational aqui | Software Design Change | Siemens sent a Customer Safety Advisory Notice dated June 26, 2013, to all affected customers. The letter identified the product, the problem, and the action needed to be taken by the customer. Customers were strongly recommend to carefully evaluate one of the following options before running 3D rotational acquisition programs that include real-time subtraction until a patch is installed on their system: 1) position the monitor in a way such that contrast advancement can be monitored during the fill phase on the live monitor in the control room.2) For applications not requiring actions based on the contrast flow during the fill phase, the contrast may not need to be monitored during the fill phase. Although the monitor is not showing the expected subtraction, the acquisition is working correctly. There is no need to interrupt, but continue and finish the workflow.3) Perform 3D rotational programs that include real time subtraction from the exam room. The display on the live monitor in the exam room is not affected by this problem. To resolve the issue, the update AX028/13/P (VC20x/VC21B), will be made available around July 2013, and scheduling an appointment for implementation of the update with Siemens Customer Service organization was recommended. If this device/equipment is no longer in theirr possession, customers were asked to forward this safety notice to the new owner and inform Siemens of the new user's identity. |  | 29 | 29 | Nationwide Distribution and Puerto Rico | Terminated | February 05, 2015 | 589 | 29 | Computer
Z-1066-2013 | 64505 | No | OQG | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stelkast Cross-over Acetabular Shell and Liner Hip System, Part Numbers SC3458-52, SC3474-52, SC3459-52, and SC3460-52. |  |  | January 24, 2013 | 2013 | Stelkast Co | Firm was made aware of incidents relating to the difficulty of properly engaging and securing the Cross-Over 36mm liner into the Cross-Over shell. | Device Design | The firm initiated their recall of this product on January 24, 2013 by notifying their consignees by telephone informing them of the recall. |  | 254 | 254 | Nationwide Distribution including the states of NV, OK and PA. | Terminated | October 08, 2014 | 622 | 254 | Computer
Z-1785-2013 | 63947 | No | LGW |  |  | Neurology | PMA | Eon and Eon Mini charging system, model numbers Eon -3701 (Wall Charger), 3711 (Portable Charger) and Eon Mini -3721 (Charger); This is a continuation of the firm's July 2012 recall. |  |  | November 30, 2012 | 2012 | St. Jude Medical | There is issues of warmth or heating at the implant site during charging for the Eon Mini implantable pulse generators (IPGs). In the July 2012 letter St. Jude Medical informed customers that they would be implementing design improvements to the charger to address possible increased energy dissipation when the charger is misaligned with the IPG or the IPG is implanted too near the surface of the | Device Design | St Jude Medical sent an Important Medical Device Safety Information letter dated November 30, 2012 to affected customers. The letter identified the affected product, problem and actions to be taken. The letter informed customers that a replacement charger for the Eon Mini IPGs is now available and will be provided at no charge. For questions or concerns call St. Jude Medical at 877-650-0367. |  | 75663 | 75,663 | Worldwide Distribution - USA Nationwide including Puerto Rico and countries of: Austria, Australia, Belgium, Brazil, Canada, Chile, China, Colombia, Cyprus, Czech Republic, Denmark, Finland, France, Germany, Greece, Hong Kong, India, Ireland, Israel, Italy, Japan, Luxembourg, Martinique, Mexico, Netherlands, Poland, Russian Federal, Saudi Arabia, Singapore, South Africa, South Korea, Spain, Sweden, Switzerland, Turkey and United Kingdom. | Terminated | November 04, 2015 | 1069 | 75663 | Computer
Z-1070-2013 | 64656 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX CV Reporting Cardiac Catheterization module and IMPAX CV Outbound Reporting. |  |  | March 15, 2013 | 2013 | AGFA Corp. | Baseline pulmonary capillary wedge (PCW) pressure values from IMPAX CV Reporting Cardiac Catheterization module were not correctly transferring into the hospital's electronic health record (EHR) system. | Software design | AGFA sent an "URGENT FIELD SAFETY NOTICE" letter dated March 15, 2013 to the consignees who use the IMPAX CV Reporting tool described for this issue. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at 877-777-2432, referencing HQ_10020006 for questions regarding this notice. |  | 26 | 26 | Nationwide Distribution including the states of AK, GA, IL, KY, MA, NY, NC, RI, SC, TX, and VA. | Terminated | January 09, 2014 | 300 | 26 | Computer
Z-1780-2013 | 65622 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Radiohead Access Point 1.4 GHz Radiohead Access Point |  |  | July 08, 2013 | 2013 | Philips Healthcare Inc. | IntelliVue Smart-hopping 1.4 GHz Access Point may experience unexpected signal loss of severalminutes impacting wireless connectivity | Software Design Change | Philips sent an Urgent Medical Device Correction letter dated July 2013 to affected customers. The letter informs customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer. Philips will perform a firmware upgrade (version C.00..07 for all affected customers at no cost. Philips is asking customers to follow the Action to be Taken by Customer/User section of the Urgent Medical Device Correction notification/Field Safety Notice.If you need any further information or support concerning this issue, please contact Philips Customer Support at 1-800-722-9377 |  | 20105 | 20105 units | USA Nationwide Distribution | Terminated | May 17, 2017 | 1409 | 20105 | Computer
Z-1069-2013 | 64654 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX CV 7.8 SU3 - OCR Service. |  |  | March 15, 2013 | 2013 | AGFA Corp. | Software design error in IMPAX CV 7.8 SU3 | Software design | AGFA Healthcare sent an "URGENT FIELD SAFETY NOTICE" letter dated March 15, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed to distribute this information within your facility to all individuals who need to be aware of this. Please fax back to (864) 421-1664 that the information was received and understood was requested from the customers.If you have any questions about this matter please contact your local AGFA Healthcare organization, or call (864) 421-1754. |  | 2 | 2 | Distribute in the state of FL and KY | Terminated | April 01, 2014 | 382 | 2 | Computer
Z-1774-2013 | 65688 | No | HQF | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | PASCAL and PASCAL Streamline (532 nm.577 nm); |  |  | March 27, 2012 | 2012 | Topcon Medical Laser Systems, Inc | When using the "Favorites" function for saved settings, there is potential for a mismatch between displayed power and the actual power output. This may result in over or under treatment. | Software design | Initial notification was sent on March 27, 2012 to all TMLS Distribution Partners and Sales Representative. This was followed by a follow up notification on June 4, 2012. A final notice was sent on February 11, 2013 to all those who had not sent documentation that the field action had been implemented. |  | 40 | 7 US, 33 foreign | Worldwide Distribution-USA and the countries of China, Hong Kong, Brazil, France, Japan, Netherlands, Pakistan, Malaysia, United Kingdom, South Korea, Thailand, Dubai, and South Africa. | Terminated | July 19, 2013 | 479 | 40 | Computer
Z-1771-2013 | 65630 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Spacelabs ARKON Anesthesia Workstation. |  |  | June 07, 2013 | 2013 | Del Mar Reynolds Medical, Ltd. | Spacelabs ARKON Anesthesia System is recalled due to software defect. The System Status Computer may incorrectly determine that a communication error has occurred. If this situation occurs, a buzzer will activate and a yellow triangle will be displayed on the System Status Computer screen that alerts the user of this error. | Software design | Spacelabs Healthcare sent an Urgent Medical Device Correction" letter, dated June 20, 2013, to all affected consignees. The letter identified the product, th problem, and the action to be taken by the consignee. Consignees were instructed to advise their staff of the situation. Consignees were also advised that Spacelabs Field Service personnel would be contacting them at their facility to schedule installation of a software update to resolve the issue at not cost to the consignee. Consignees with questions were instructed to call Technical Support at 1-800-522-7025 and select option 2.For questions regarding this recall call 800-522-7025. |  | 9 | 9 units (4 in the US and 5 outside US) | Worldwide Distribution - USA including South Carolina and internationally to Canada and Great Britian. | Terminated | August 29, 2013 | 83 | 9 | Computer
Z-1072-2013 | 64664 | No | JAI | Powered radiation therapy patient support assembly. | Radiology | Radiology | 510(k) | Precise Treatment Table. |  |  | March 14, 2013 | 2013 | Elekta, Inc. | The Column rotation can accidentally be moved to a different position which in the locked position. | Device Design | Elekta sent an "IMPORTANT FIELD SAFETY NOTICE" dated March 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact your local Elekta office for questions about this notice. |  | 443 | 443 | Nationwide Distribution and Puerto Rico. | Terminated | March 17, 2015 | 733 | 443 | Computer
Z-1743-2013 | 65652 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | TDHistology/Cytology |  |  | June 20, 2013 | 2013 | Technidata S.A. | A recall was initiated by Technidata because use of symbols (less than or equal to or greater than or equal to or any other symbol) in Microsoft Word reports when using electronic transmissions of results by of HPRIM (including HPRIM Medecin via Kermit) or HL7 must be absolutely prohibited. Such symbols are not supported by the Microsoft Word interface used by the Histology/Cytology software and | Software design | TD Technidata Medical Software sent a Medical Device Recall letter dated June 20, 2013, to one affected customer who received the TD-Synergy Histology/Cytology module. The customer was informed of the problems identified and the actions to be taken. Customers were instructed to complete the attached Reply Form and return it to the firm by fax to 514-270.2770. The customer were also instructed to contact Technidata Canada at (514) 270-7777 or at support-canada@technidata-web.com. |  | 1 | 1 site | US Distribution CA only. | Terminated | July 29, 2013 | 39 | 1 | Computer
Z-1745-2013 | 65561 | No | DZL | Intraosseous fixation screw or wire. | Dental | Dental | 510(k) | Multi Vector Distractor Pin Holding Clamp, limited bone stock |  |  | May 31, 2013 | 2013 | Synthes USA HQ, Inc. | Synthes is initiating a voluntary Medical Device removal for specified part numbers and lot numbers of the Mutis Vector Distractor Pin Holding clamp, limited bone stock. The firm received complaints that the clamp and nut construct would not engage the distractor arm threads and subsequently bind. It is not likely that a serious adverse event will result due to a nonconforming clamp, however there | Device Design | Synthes sent a Urgent Notice: Medical Device Removal letter dated May 31, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers. The firm received complaints that the clamp and nut construct would not engage the distractor arm threads and subsequently bind. It is not likely that a serious adverse event will result due to a nonconforming clamp, however, there is a remote potential that additional intervention would be required due to a non-functioning clamp.Synthes requested the product to be removed from inventory and returned to Synthes by calling 1-800-479-6329 to obtain a Return Authorization Number. Direct Accounts were asked to complete and return the Verification Section at the end of the letter indicating the quantity of affected product was found. Negative responses were also requested.If you have any questions please call (610) 719-5450. |  | 121 | 121 | US Nationwide Distribution including the states of MA, TX, MN, SC, UT, VA, and NY. | Terminated | August 31, 2015 | 822 | 121 | Computer
Z-1750-2013 | 65620 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Siemens AXIOM Artis zee/zeego systems equipped with generator A100 Plus |  |  | January 31, 2013 | 2013 | Siemens Medical Solutions USA, Inc | During the course of product monitoring, Siemens became aware of a tolerance problem in the power supply of the generator control above a specific value, thus, the generator may fail. | Device Design | Siemens sent an Urgent Customer Information letter to their consignees. The letter identified the affected product, problem , update Instructions and actions to be taken. Siemens began to equip affected systems with a modified board assembly in January, 2013. |  | 23 | 23 | United States Nationwide Distribution - including the states of: CA, CO, FL, IA, IL, IN, MA, MD, MI, NC, NY, PA, TX, VA and WA and Puerto Rico. | Terminated | November 12, 2014 | 650 | 23 | Computer
Z-1736-2013 | 65406 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Integra PANTA Arthrodesis Nail System. |  |  | May 09, 2013 | 2013 | Integra LifeSciences Corporation | As a result of the investigation into an adverse trend of complaints for reports of misalignment of the calcaneal or tibial screws during use of the Panta Nail Support Device, Integra has decided to issue this voluntary medical device correction. The misalignment can result in the inability to properly insert the screws into the Panta Nail and create the potential need for revision surgery. | Device Design | Integra initiated a voluntary correction on May 9, 2013 by providing a written notification and copy of the revised Surgical Technique Manual as well as the specific highlighted sections of the manual addressing calcaneal and tibial screw alignment to all consignees that Integra records indicate have been shipped a PANTA Nail. The written notifications and Surgical Technique Manual have been delivered to US Consignees by traceable courier service and by traceable email or facsimile for Canada and Non-US Consignees. The notification informs customers of the nature of the issue; advises of potential adverse patient consequences that may occur; advises customers to review highlighted revisions of the Surgical Technique Manual to minimize the potential for misalignment of the screws and requests customers to complete and return the attached Acknowledgement Return Form. |  | 692 | 692 | Worldwide Distribution-USA (nationwide) and the countries of Austria, Australia, Belgium, Canada, Czech Republic, Denmark, Finland, France, Germany, Great Britain, Ireland, Israel, Italy, Netherlands, Portugal, Russia, Saudia Arabia, South Africa, Spain and United Arab Emirates. | Terminated | November 07, 2013 | 182 | 692 | Computer
Z-1077-2013 | 63974 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ELEKTA MOSAIQ SW version below 2.41 |  |  | December 20, 2012 | 2012 | Siemens Medical Solutions USA, Inc | A potential safety risk with the MOSAIQ OncologyInformation System that has been identified by the manufacturer IMPAC Medical Systems, Inc. | Software design | Siemens sent an Urgent Medical Device Correction letter with update instructions to all affected customers. The letter identified the affected product, problem and actions to be taken. For questions contact your responsible Siemens Service Engineer or Application Specialist. |  | 101 | 101 | Worldwide Distribution | Terminated | July 19, 2013 | 211 | 101 | Computer
Z-1588-2013 | 65137 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance CT 64, Ingenuity CT, Ingenuity Core, and Ingenuity Core128 scanners are whole body Computed Tomography X-ray systems. |  |  | March 21, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | If the operator selects a patient name from the worklist using the mouse and then selects a different patient name using the keyboard (arrow keys then press enter) the patient name that populates the exam information page is the patient name selected with the mouse. Also, it is possible to begin a normal clinical scan while the anti-virus scan is running. However, if the user is running a bol | Software design | Philips Medical Systems sent an "URGENT-MEDICAL DEVICE CORRECTION" letter dated Mar 21, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Additionally, a Field Service Engineer will contact each site to schedule a time to implement the Field Change Order. |  | 50 | 50 units | Nationwide Distribution-including the states of AR, AZ, DE, IN, KS, LA, MA, MN, MS, MT, NJ, NY, NY, OH, OR, PA, TN, TX, VT, WV, and WY. | Terminated | November 27, 2013 | 251 | 50 | Computer
Z-1570-2013 | 65135 | No | LKK |  |  | General Hospital | PMA | Medtronic SynchroMed II, Model 8637, (The SynchroMed II Pump is supplied in 20 ml or 40 ml reservoir size.) Sterilized using ethylene oxide. |  |  | June 03, 2013 | 2013 | Medtronic Neuromodulation | Medtronic is providing Healthcare Professionals with important safety information and patient management recommendations regarding the unintended delivery of drug during the priming bolus function for the SynchroMed implantable infusion pump. The unintended delivery of drug can contribute to patient overdose or underdose symptoms which may be clinically relevant. | Device Design | Medtronic sent an Urgent Medical Device Correction letter dated May 2013, to all affected customers. The letter described the background and nature of the issue and provided recommendations for patient management and monitoring after initiation of Intrathecal Therapy. The letter directed customers to continue to use the priming bolus precedure; to monitor all patients following start or restart of intrathecal therapy; consider priming the pump prior to implant; and educate caregivers and family members to recognize the signs and sypmtoms associated with intrathecal drug therapy complications. For questions contact your local Medtronic field representative or Medtronic Neuromodulation Technical Services at 1-800-707-0933. |  | 261109 | 261,109 total devices estimated implanted Worldwide | Worldwide Distribution - USA Nationwide | Terminated | May 14, 2015 | 710 | 261109 | Computer
Z-1586-2013 | 65411 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | DC power cable (p/n 3010556) supplied with the MEDRAD(R) Veris MR Vital Signs Monitor: |  |  | April 15, 2013 | 2013 | Medrad Inc | The firm is recalling these power cables due to a latent design reliability issue and the potential for shorting which can result in heating/melting of the cable jacket. | Device Design | Medrad Inc sent an Urgent Medical Device Component Recall letter dated April 15, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter states that replacement products will be available in June and shipped in the coming months along with additional instructions for returning the affected product. Customers were instructed to complete the enclosed recall response form and fax it back to Bayer at 412-406-0942. For questions contact customer support team at 1-877-229-3767 and select option 2. |  | 1066 | 1,066 | Worldwide Distribution - US Nationwide including the states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY and the countries of: AE, AU, AZ, BN, BR, CA, CH, CL, CO, DE, DK, DO, DZ, EG, ES, FR, GB, HK, HN, HR, HU, IN, IT, JO, JP, KR, LB, MU, MX, MY, NL, PH, PS, SA, SG, TH, TR, TW, UY, VE, and ZA. | Terminated | April 03, 2014 | 353 | 1066 | Computer
Z-1159-2013 | 64839 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | GemStar Pump - Battery |  |  | March 18, 2013 | 2013 | Hospira Inc. | Battery level lower then 2.4 volts results in corrupt history log and loss of data for preview therapy's parameters and causes an 11/004 error resulting in the pump not being able to be used. | Device Design | Hospira sent an Urgent Device Field Correction letter dated March 18, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to notify their consignees if they further distributed the product and confirm they have done so by return of the reply form to Stericycle. Lithium batteries that are older than three (3) years should be replaced. Customers should contact their Hospira Advanced Knowledge Center at 1-800-241-4002, option 4, to determine if their battery need to be replaced and if necessary, arrange for the return of their device to perform battery replacement. Customers with questions should call 1-877-907-7516. For questions regarding this recall call 224-212-4892. |  | 38801 | List Number 13000: 25,119 pumps; List Number 13100: 185 pumps; List Number 13150: 13,497 pumps | Worldwide Distribution - USA (nationwide) Australia, New Zealand, China, Hong Kong, Taiwan, Japan, Korea, Philippines, Thailand, Malaysia, Canada, Austria, Bahrain, Belgium, Croatia, Denmark, Egypt, Finland, France, Germany, Gibraltar, Greece, Hungary, Iceland, Ireland, Israel, Italy, Jordan, Kuwait, Lebanon, Libya, Luxembourg, Malta, Netherlands, Norway, Oman, Portugal, Qatar, Romania, Saudi Arabia, Slovenia, South Africa, Spain, Sweden, Switzerland, Turkey, UK, United Arab Emirates, Brazil, Chile and Colombia. | Terminated | March 07, 2017 | 1450 | 107955 | Computer
Z-1368-2013 | 64850 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | CobasÂż 8000 modular analyzer series software version 03-01, 02, 03 and 04, impacting the cobas e 602 module only. |  |  | April 04, 2013 | 2013 | Roche Diagnostics Operations, Inc. | When using the cobas e 602 module, Diluent Multiassay may be used for autodilution of certain assays instead of Diluent Universal. This may cause incorrect results under specific circumstances if all of the following conditions are met simultaneously: More than a total of 2000 reagent kits or diluents have been registered on the cobas e 602 modules within one system core, and two different diluent | Software design | Roche sent an Urgent Medical Device Correction (UMDC) letter dated April 4, 2013 via UPS to all affected customers. The letter idenified the affected product, problem and instruction to be followed. Customers were instructed to consult with the physician or pathologist at their facility to determine specific clinical implications for patients. A future software version will correct this issue. Customers were asked to complete and fax back the UMDC attached fax form to 1-88-628-0730. Questions should be addresssed to Roche Diagnostics Technical Support 24 hours a day at 1-800-428-2336. |  | 72 | 72 units | USA Nationwide Distribution | Terminated | February 24, 2015 | 691 | 72 | Computer
Z-1367-2013 | 64718 | No | IFB | Cytocentrifuge. | Pathology | Pathology | 510(K) Exempt | SLIDERPREP PLUS SLIDE PREPARATION CENTRIFUGE. |  |  | March 13, 2013 | 2013 | Separation Technology, Inc. | Separation Technology, Inc. recalled SlidePrep Plus Cytology Centrifuges manufactured prior to January 1, 2013 for having the potential to start up unexpectedly when encountering an electrostatic discharge. | Device Design | On March 13, 2013, Separation Technology, Inc. sent their customers letters instructing them to complete and return the attached Acknowledgement &amp; Receipt form via fax to (407) 788-3677 and requesting the return of the SlidePrep Plus Cytology Centrifuges to STI for correction. For any questions regarding this recall, call STI Customer Service at (800) 777-6668. |  | 102 | 102 | Worldwide Distribution-USA (nationwide) including the states of PA, NC, IL, KS, and FL, and the countries Canada and Turkey. | Terminated | December 24, 2013 | 286 | 102 | Computer
Z-1370-2013 | 64941 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker Orthopaedics Triathlon Femoral Distal Augment- Left; 1 single unit per package |  |  | March 25, 2013 | 2013 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics received a report that a label on the external packaging (box) of a Triathalon Femoral Distal Augment 10mm- Size 5 Left depicted a glyph of an acetabular shell rather than a femoral distal augment. In addition the label was missing a reference to the "10mm" augment thickness which per the label print is required to appear below the size reference. | Software design | The firm, Stryker Orthopaedics, sent an "URGENT MEDICAL DEVICE REMOVAL" letter dated March 25, 2013 and Product Accountability Forms via Fed Ex to its Consignees/ customers. The letter described the product, problem and actions to be taken. The customers were instructed to complete and return the attached Notification Acknowledgment Form within 5 days via fax to 855-251-3635; contact the hospitals in your territory that have the affected product to arrange return of product; and return affected product to the attention of Regulatory Compliance Stryker Orthopaedics, 325 Corporate Drive, Mahwah, NJ 07430, attach flourescent orange PRODUCT REMEDIATION sticker to your return indicating the PR #RA2013-011 and mark outer box with the words "Product Recall".If you have any questions, feel free to contact Manager, Divisional Regulatory Compliance at (201) 972-2100. |  | 11 | 11 units | Nationwide distribution: USA including states of: HI, IN, MS, NC, NJ, OK, TX, and WI. | Terminated | May 06, 2015 | 772 | 11 | Computer
Z-1375-2013 | 65031 | No | IQA | Powered environmental control system. | Physical Medicine | Physical Medicine | 510(K) Exempt | QuietCare. |  |  | April 04, 2013 | 2013 | Intel-GE Care Innovations LLC | Customers at 4 facilities reported receiving duplicate alerts using the QuietCare-Networked product. | Software design | Care Innovations contacted their customers by phone and e-mail regarding the potential problems and fixes on April 4, 17, and 23, 2013. The information provided stated the product, problem, and actions to be taken by the customers. Contact Technical Support at 855-885-2273 for questions regarding this issue. |  | 99 | 1722S; 99 Out of US | US and Great Britain | Terminated | June 04, 2013 | 61 | 145 | Computer
Z-1182-2013 | 64944 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Software correction applies to customers who use the RT Therapist RTT4.2.108 and RT Therapist Connect RTT 4.2.108 |  |  | April 04, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Software update to fix multiple safety related issues. | Software design | Siemens sent updated instructions of TH004/13/S on April 4, 2013. The SW Update will be delivered by the Siemens Customer Service Group.For questions regarding this recall call 925-293-5442. |  | 245 | 245 active devices | Worldwide Distribution - USA (nationwide) and internationally to Poland, India, Canada, Germany, United Kingdom, New Zealand, Belgium, Malaysia, Republic Korea, Ireland, Australia, Norway, Japan, Turkey, Croatia, Saudi Arabia, Colombia, P.R. China, Brazil, Thailand, Romania, Lebanon, Mexico, Italy, Sweden, Czech Republic, Russian Federation, and South Africa. | Terminated | August 12, 2013 | 130 | 245 | Computer
Z-1177-2013 | 64645 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Maquet Getinge Group battery modules for use with Maquet SERVO-i ventilator systems (Part #6487180). |  |  | March 07, 2013 | 2013 | Maquet Cardiovascular Us Sales, Llc | When using the SERVO-i on battery power, in a very small number of cases, some battery modules distributed after January 31, 2010 have a shorter battery run time than expected. | Device Design | The firm, Maquet Getinge Group" sent "Field Safety Notices"/"Field Safety Notice Confirmation Forms" dated March 7 &amp; 11 2013 via Federal Express and an updated Field Safety Notices dated April 15, 2013 to its customers. The Field Safety Notices identifies the problem, product, risk factors and actions to be taken by the customer. The customers were instructed to reconnect the ventilator to mains power immediately or add additional battery modules in case of 'low battery voltage' alarm; ensure that at least two fully charged batteries are installed at all times, and that sufficient resuscitation equipment is accessible to the ventilator; identify and immediately replace battery modules with date codes 1005-1243, that have generated either: a "low battery voltage" alarm or a "replace battery" message, prior to the expiry date; and complete and return the enclosed Field Safety Notice Confirmation Form by e-mailing a scanned copy to FieldActions@maquet.com or via fax at (973) 807-1832.If you have any questions or require replacement batteries, please contact your local MAQUET representative, or MAQUET Customer Service at 888-627-8383 (press option 2, followed by option 3) Monday through Friday, between the hours of 8:00 a.m. and 6:00 p.m. EST. |  | 90000 | 90,000 units (batteries) | Worldwide distribution: USA (nationwide) and countries of: ALGERIA, ANDORRA, ANTILLES (Netherlands), ARGENTINA, ARUBA, AUSTRALIA, AUSTRIA, AZERBADJAN, BANGLADESH, BELARUS, BELGIUM, BOLIVIA, BOTSWANA, BRAZIL, BRUNEI, BULGARIA, CANADA, CHILE, CHINA,COLOMBIA, CONGO, COSTA RICA, CROATIA, CUBA, CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, ECUADOR, EGYPT, EL SALVADOR, ESTONIA, YUGOSLAVIA (Former), FIJI, FINLAND, FRANCE, GEORGIA, GERMANY, GREECE, GUATEMALA, GUYANA, HONDURAS, HONG KONG,HUNGARY, ICELAND, INDIA, INDONESIA, IRAN, IRAQ, IRELAND, ISRAEL, ITALY, JAMAICA, JAPAN, KENYA, KUWAIT, KYRGYZSTAN, LATVIA, LIBYA, LITHUANIA, LUXEMBURG, MACAU, MACEDONIA, MALAYSIA, MALDIVES, MARTINIQUE (French), MEXICO, MOLDAVIA, MOROCCO, NEPAL, NETHERLANDS, NEW CALEDONIA (French), NEW ZEALAND, NICARAGUA, NIGERIA, NORWAY, OMAN, PAKISTAN, PALESTINE, PANAMA, PAPUA NEW GUINEA, PARAGUAY, PERU, PHILIPPINES, POLAND, POLYNESIA (French), PORTUGAL, QATAR, REUNION (French), ROMANIA, RUSSIA, SAUDI ARABIA, SERBIA, SEYCHELLES, SINGAPORE, SLOVAK REPUBLIC, SLOVENIA, SOUTH AFRICA, SOUTH KOREA, SPAIN, SRI LANKA, SUDAN, SURINAM, SWEDEN, SWITZERLAND, SYRIA, TADJIKISTAN, TAIWAN, TANZANIA, THAILAND, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, UKRAINE, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY, VATICAN city state, VENEZUELA, VIETNAM and YEMEN. | Terminated | August 11, 2016 | 1253 | 90000 | Computer
Z-1404-2013 | 65082 | No | HKO | Ophthalmic refractometer. | Ophthalmic | Ophthalmic | 510(K) Exempt | Canon Full Auto Ref-Keratometer RK-F2, Catalog Number 6937B002AA. |  |  | December 05, 2012 | 2012 | Canon Inc. | Canon Inc. had received 17 claims from our foreignsales companies between October to November2012 that measurement value of RK-F2s refractivepower (SPH value) sometimes indicated abnormalvalue which occurs approximately 5% in probability. | Software design | Canon initiated recall by sending letters on November 22, 2012 to all affected customers. The letter identified the affected product, problem, actions to be taken and contact information. |  | 384 | 384 units | Worldwide Distribution - USA Nationwide and countries of: Australia, Japan, Korea, Singapore and United Kingdom. | Terminated | September 20, 2013 | 289 | 384 | Computer
Z-1585-2013 | 64789 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | The Centricity Laboratory System is intended to be an information system designed to support the clinical and administrative activities associated with the provision and utilization of clinical laboratory services and facilities, e.g., the storing and delivering of analytical results. It is a specially designed data program application (software), which is supplied for installation in existing mainframe or decentralized computers or a computer network. This software system will not be the software directly used to run, steer, or control any specific laboratory analyzer or equipment. It is intended for use in laboratory facilities, including a central laboratory or in a multiple laboratory environment servicing satellite laboratories, reference laboratories, clinics, etc. Centricity Laboratory System is not intended for direct patient contact. |  |  | November 20, 2012 | 2012 | GE Healthcare It | If environment variables MICRO_CORRECT_MSG_TYPE and CORRECT_MSG_TYPE are set to different values, in Microbiology result entry when released results (Final Report released) are changed and the Corrected Result messages are edited on the free-text form, some of the originally entered panel free-text is lost. For this issue to occur the original result must be amended and the environmental variabl | Software design | GE Healthcare sent an Urgent Medical Device Correction letter dated November 20, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Software correction is available. Please contact your GE service representative to request or install this update. Please contact your GE Service Representative with any question regarding this issue. For other questions please contact GE Helpdesk.Help Desk: North America: 1-888-778-3375 Help Desk: Europe: + 44 1603 877595 Help Desk: Asia Pacific + 61 3 5327 4000 |  | 9 | 9 Units | Worldwide Distribution - US Distribution including the states of OH, NY, GA, TX, TN and MN and the countries of Australia, Qatar and England | Open, Classified |  |  | 9 | Computer
Z-1407-2013 | 65171 | No | GWF | Evoked response electrical stimulator. | Neurology | Neurology | 510(k) | Natus neurology Dantec KEYPOINT.NET 2.20. |  |  | April 09, 2013 | 2013 | Alpine Biomed Aps | Alpine Biomed ApS is recalling Natus neurology Dantec KEYPOINT.NET ver. 2.20 Software due to the presence of a software bug. | Software design | Alpine Biomed ApS sent a Recall Notification/ Field Safety Notice dated April 8, 2013 to all affected customers via e-mail and courier. Alpine Biomed ApS advised their customers to upgrade all installations of Keypoint.NET v2.20 to Keypoint.NET v2.21. Alpine BioMed ApS has recommended that their customers complete and return the verification form to QA/RA Engineer Michael Jacobsen at e-mail michael.jakobsen@natus.com. These actions and the return of the verification form to Alpine Biomed ApS should be completed within 3 months of receipt of the recall notifications. If any assistance is required to complete these actions, please contact Technical Support Engineer Michael Jensen via e-mail michael.jensen@natus.com, alternatively QA/RA Engineer Michael Jacobsen by e-mail michael.jakobsen@natus.com. |  | 871 | 871 licenses installed | Worldwide Distribution - US Nationwide and the countries of: Germany, France, Bulgaria, India, Colombia, Tunisia, Saudi Arabia, Finland, Denmark, Brasil, Italy, Estonia, South Korea, Russia, Greece, Thailand, Israel, Poland, Turkey, Slovakia, Portugal, Switzerland, Hungary, Great Britain, Netherlands, China, Indonesia, Australia, Sweden, Spain, Iran, Singapore, Norway, Austria and Czech Republic. | Terminated | January 06, 2014 | 272 | 871 | Computer
Z-1419-2013 | 65136 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | RayStation Therapy Treatment Planning System, Model Numbers 2.0, 2.5, and 3.0; CLASSIFICATION NAME: System, Planning, Radiation Therapy Treatment. |  |  | February 15, 2013 | 2013 | RAYSEARCH LABORATORIES AB | Under some circumstances, imported CT, MR and PET images are offset from their true positions by one pixel in one or two directions. Users must be aware of this information to avoid incorrect dose calculations during treatment planning. | Software change control | RaySearch Laboratories sent an Urgent Field Safety Notice, Medical Device Correction letter dated January 28, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to educate their planning staff and all users about the recommended workarounds contained in the letter. Customers were also asked to inspect their product and identify all installed units that were infected and then complete the attached Reply Form and e-mail or scan a copy of the completed form to the firm. Customers were instructed to pass along this information to all those who need to be aware of this Medical Device Correction Customers with questions were instructed to contact the RaySearch support email at support@raysearchlabs.For questions regarding this recall call +46 8 545 06130. |  | 97 | Domestic: 70 units; Foreign: 27 units | Worldwide Distribution - USA including Connecticut, Florida, Massachusetts, Missouri, New York, Ohio, and foreign distribution to(Canada, Italy, Australia, Germany, Netherlands, and Switzerland. | Open, Classified |  |  | 97 | Computer
Z-1421-2013 | 65165 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Leucadia Autolok" Polyaxial Pedicle Screw System. |  |  | May 06, 2013 | 2013 | Alphatec Spine, Inc. | Alphatec Spine initiated this medical device recall for the following reason: due to tolerancing, the flange of the Phygen Leucadia" Autolok set screw may interfere with the top of the screw head prior to fully locking the rod. If this condition occurs, the polyaxialscrew head and rod may fail to provide a rigid construct under loading. To date there have beenno reports of patient injury, adve | Device Design | Alphatec Spine initiated recall by contacting customers via telephone and/or email on 05/06/2013. An Urgent Medical Device Recall Notification letter along with a response form were sent via FedEx to all affected accounts. The letter identified the product, problem, and actions to be taken by the customers. |  | 1814 | 1,814 | Nationwide Distribution. | Terminated | September 25, 2014 | 507 | 1814 | Computer
Z-1425-2013 | 65073 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Tip Cover Accessory, 8 MM, MCS, box of 10 |  |  | September 23, 2011 | 2011 | Intuitive Surgical, Inc. | Clarification of existing labeling and related cautions for the tip cover accessory of the da Vinci Surgical System. | Device Design | Intuitive Surgical sent an Important Product Notification letters dated October 10, 2011 to all affected customers. A second letter was sent to those that had not responded on December 4, 2011. The letters identified the affected product, problem, and actions to be taken, The letter included recommendations and suggestions for proper use and prevention of damage to the Tip cover during surgery. Also attachments include recommendations and further information for the following areas: avoid instrument collisions; carefully install tip cover; straighten wrist prior to removal; inspect cannula prior to use and do not exceed maximum monopolar cautery settings.Attached is a list of validated ESU settings and suggestions and instruction on being aware of critical anatomy in contact with the instrument and to change the tip cover in cases with extended cautery use.Questions may be referred to Customer service at 1-800-876-1310 in the US and at + 41 21 821 2020 or +800-0821-2010. |  | 500000 | 50,000 boxes of 10 tip covers | Worldwide distribution - USA Nationwide and the countries of Australia, Austria, Belgium, China, Czech Republic, Denmark, Finland, France, Germany, Greece, India, Ireland, Israel, Italy, Japan, Malaysia, Mexico, Netherlands, New Zealand, Norway, Pakistan, Philippines, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Slovenia, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom. | Terminated | May 30, 2013 | 615 | 500000 | Computer
Z-1431-2013 | 65194 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | 8mm Precise Bipolar Forceps used in conjunction with the da Vinci Surgical System, Model IS1200. |  |  | May 08, 2013 | 2013 | Intuitive Surgical, Inc. | Reports of dislodging of the Cautery Plug Insert from the Cautery Plug Riser, preventing a bipolar cautery cable from being properly connected which results in loss of cautery function of Bipolar Instruments. | Device Design | Intuitive Surgical sent an "Urgent Medical Device Correction" letter dated May 8, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact Customer Service at 800-876-1310 for questions regarding this notice. |  | 697 | 697 units - all products | Worldwide Distribution-USA (nationwide) including the states of AL, CA, CO, FL, IL, KY, MI, MS, NJ, NY, PA, TX, UT, VA, WA, and WI, and the countries of Austria, Australia, Belgium, Canada, Czech Republic, Denmark, France, Germany, Italy, Netherlands, Norway, Qatar, Romania, Spain, Sweden, Switzerland, Taiwan, United Kingdom, Saudi Arabia, Singapore, and Venezuela. | Terminated | February 18, 2014 | 286 | 697 | Computer
Z-1167-2013 | 64731 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AU5800 Clinical Chemistry Analyzer |  |  | March 20, 2013 | 2013 | Beckman Coulter Inc. | Beckman Coulter, Inc. initiated this recall because there is a software error that has the potential to cause the reporting of erroneous patient results because of misidentification by the AU5800 system. | Software design | Beckman Coulter sent an Urgent Product Correction letter during the week of March 18, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were informed of the software limitation and provides instructions to be taken (temporary workaround) until the software is upgraded and released. Customer contact information provided: "Call Center Hotline at 800 854-3633. New customers will receive a copy of the customer notification at the time of installation by the Field Service Personnel until revised software that corrects this issue is released. |  | 490 | 490 | Worldwide Distribution - USA, including Canada and Puerto Rico and the states of NJ, TX, CA, UT, KS, MA, NY, MI, GA, VA, IN, WI, IL, OH, NV, NC and PA. and the countries of Australia, Belgium, Czech Republic, Japan, Germany, Brazil, China, Italy, Ireland, Israel, India, Hong Kong, Malaysia, Netherlands, Spain, Switzerland, France, Taiwan, Turkey, United Kingdom and Vietnam. | Terminated | September 16, 2014 | 545 | 490 | Computer
Z-1365-2013 | 65158 | No | OUG | Medical device data system. | General Hospital | General Hospital | 510(K) Exempt | MEDITECH Bi-Directional interactive Communication Analyzer Interface; |  |  | May 07, 2013 | 2013 | Medical Information Technology, Inc. | Incorrect transmission of laboratory results. | Device Design | Beginning on May 7, 2013, MEDITECH distributed a notification to MAGIC and Client Server customers who have the MEDITECH Bi-Directional Interactive Communication Analyzer Interface in use with Vitros 3600, 4600, or 5600 Analyzer in TEST and/or LIVE. This notification was made via remailed task updates that can be printed by the customer. Task updates sent via e-mail are immediately transmitted to the customer.This notification informs customers of the potential patient safety issue that could result in reporting inaccurate qualitative results. Customers will receive a corrective change, in which the previous test's results will no longer be associated with the test not performed. Until customers receive the corrective change, they are instructed to monitor all No Result values from their analyzer and confirm that the result in MEDITECH matches. For further questions please call 781-821-3000, ext. 5742. |  | 56 | 56 devices | Worldwide Distribution - USA, including Puerto Rico and the country of Canada. | Terminated | October 01, 2014 | 512 | 56 | Computer
Z-1187-2013 | 64812 | No | DQA | Oximeter. | Cardiovascular | Anesthesiology | 510(k) | Rad-8 Pulse Oximeters |  |  | March 13, 2013 | 2013 | Masimo Corporation | Masimo has identified a very small number (0.02%) of Rad-8 devices with 20-pin patient cable connector that, when moved, can power off without the operator pressing the power button. | Device Design | Masimo Corporation sent an Urgent Recall Notification letter dated March 8, 2013, to all affected customers The letter identified the product the problem and the action needed to be taken by the customer.1.Customers were instructed to please promptly remove all Rad-8 device(s) from use and your inventory and segregate these devices.2. Please call 1-800-326-4890 and select option 2 for Technical Services. Determine with Technical Services whether you will return the affected device(s) to Masimo for repair or if eligible, receive a repair kit. If you will return the device(s) for repair, obtain a Return Material Authorization Number (RMA number).3. Complete the Tracking/Verification Form in Attachment 2 and fax it to Masimo at 1-949-297-7700, even if you do not have or intend to use the affected device(s).4. If you are eligible and intend to repair units onsite fax the completed original form in Attachment 2 to Masimo at 1-949-297-7700.5. If you have elected to return the affected device(s), and have an RMA number include the completed original Form in Attachment 2 with the returned device(s) and indicate the RMA number on the outside of the return-shipping box.We apologize for the inconvenience. Please be assured that Masimo is committed to consistently providing high quality products and services to you, our customers. We thank you for your patience and cooperation while we actively work to resolve this issue... |  | 42908 | 42,908 units | Worldwide Distribution - USA (nationwide) including the countries of Africa, Asia, Australia/New Zealand, Europe, Canada, Latin America, and Middle East / SW Asia. | Terminated | June 10, 2015 | 819 | 42908 | Computer
Z-1189-2013 | 64860 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | HeartStart MRx Monitor/Defribillator; Product Codes: M3536A, M3536J, M3536M, M3536MC, M3536M2, M3536M4, M3536M5, M3536M6. |  |  | April 08, 2013 | 2013 | Philips Healthcare Inc. | Device may fail to deliver defibrillation therapy in either Manual Defib or AED mode. | Software design | The firm, Philips Healthcare, sent an "Field Safety Notice-URGENT - MEDICAL DEVICE RECALL" letter dated April 2013 to its customers. The letter describes the product, problem and actions to be taken. In addition, letter provides instructions for using the MRx prior to receiving the software upgrade and informs the customer that Philips is initiating a correction consisting of a software upgrade for all units affected. The customers were instructed to follow the "ACTION TO BE TAKEN BY CUSTOMER/USER" section of the Field Safety Notice; send a copy of the attached package including customer letter and Field Safety Notice to any customer to whom you have distributed one of the affected device; and provide your local Philips organization with the names and addresses of customers to whom you have sold affected devices, so that arrangements can be made to perform the software updates.Should you have any questions or concerns about this device correction, please contact your local Philips representative at 1-800-722-9377.Philips updated the Field Safety Notice dated May 2013 to users who have not had the software upgrade performed. It provided additional information in the "Problem Description" and "Action to Be Taken By Customer/User". |  | 64972 | 64972 | Worldwide distribution: USA (nationwide) and countries of: Afghanistan, Albania, Argentina, Australia, Austria, Azerbaijan, Bahrain, Bangladesh, Belgium, Bolivia, Bosnia and Herzegowina, Brazil, Brunei Darussalam, Bulgaria, Canada, Cambodia, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, El Salvador, Estonia, Ethiopia, Finland, France, Germany, Greece, Guadeloupe, Hong Kong, Iceland, India, Indonesia, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Korea, Republic Of, Kuwait, Laos, Latvia, Lebanon, Libyan Arab Jamahiriya, Lithuania, Luxembourg, Macau, Macedonia, Malaysia, Malta, Mexico, Mongolia, Morocco, Namibia, Nepal, Netherlands (Antilles), New Caledonia, New Zealand, Nicaragua, Nigera, Norway, Oman, Pakistan, Palestine, Panama, Papua New Guinea, Paraguay, Peru, Philippines, Poland, Portugal, Puerto, Rico, Qatar, Reunion, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovenia, South Africa, Spain, SRI Lanka, Sweden, Switzerland, Taiwan (Province of China), Tanzania, Thailand, Trinidad and Tobago, Tunisia, Turkey, Turkmenistan, Uganda, Ukraine, United Arab Emirates, United Kingdom, Uzbekistan, Venezuela, Veit Nam, and Yemen. | Terminated | January 10, 2017 | 1373 | 64972 | Computer
Z-1346-2013 | 64930 | No | OOI | Instrumentation for clinical multiplex test systems. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | cobasÂż 4800 system z480 instrument. |  |  | April 01, 2013 | 2013 | Roche Molecular Systems, Inc. | In rare instances, channel shifted results have been generated with cobasÂż 4800 assays run on v1.1.1 of the cobasÂż 4800 system. Raw data review indicates that, in these cases, all signals are shifted by one channel:channel 1 results report as channel 2,channel 2 results report as channel 3,channel 3 results report as channel 4, andchannel 4 results report as channel 1.In the reported | Software design | Roche sent an "Urgent Medical Device Correction" (UMDC) notification with response form dated April 1, 2013, to all affected customers via UPS Ground on April 1, 2013. The letter identified the product, the problem, and the action to be taken by the customer. A contact number (Roche Diagnostic Technical Support) is provided for technical support 24 hours a day, seven days a week at phone number: 1-888-912-7090. |  | 425 | 425 instruments | Nationwide Distribution. | Terminated | August 07, 2014 | 493 | 425 | Computer
Z-1345-2013 | 56447 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | TD-Synergy. |  |  | July 29, 2010 | 2010 | Technidata S.A. | Technidata has discovered that using the F8 key in the ERM session to add a combined test can result in the modification of the value of a result that has already been entered for a test of the request. | Software design | The firm, Technidata, sent an "P1 DEFECT INFORMATION" notice dated July 2, 2010 to their consignee/customers. The notice describes the product, problem and actions to be taken. The notice directs them to find the correction to the software issue on the Technidata website where they are able to install the correction kit depending on the version and operating system used. Note "You must back up the database and the installation directory before installing this correction".Should you need any further information, do not hesitate to contact the Support team at +33 (0) 4 76 04 13 00. |  | 10 | 10 | Worldwide Distribution-USA (nationwide) including Washington DC and the states of California, Nevada, Kentucky, Alabama, Rhode Island, Pennsylvania, and Louisiana and the country of Italy. | Terminated | December 05, 2013 | 1225 | 10 | Computer
Z-1339-2013 | 65096 | Yes | LWS |  |  | Cardiovascular | PMA | LATITUDEÂż Patient Management System, Model 6488, Version 7.3. |  |  | April 10, 2013 | 2013 | Boston Scientific CRM Corp | Boston Scientific CRM is conducting a recall on the LATITUDE Patient Management System Model 6488 Version 7.3 because between March 4 and 27, 2013, a Siebel release caused ~5000 distributed PMR clinic faxes to display with illegible characters and not the required, Patient non-compliance information that was intended. | Software change control | Boston Scientific CRM Corp began re-faxing the 4,912 previously illegible PMR notifications on April 10, 2013 and completed on April 15, 2013, successfully fulfilling labeling expectations that these notifications be communicated to the clinic.For questions regarding this recall call 651-582-4000. |  | 1 | 1 | Nationwide Distribution including DC and PR | Terminated | November 26, 2013 | 230 | 1 | Computer
Z-1186-2013 | 64621 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | SL3 SOFT-TISSUE Laser (Model LR2002); Manuf: Discus Dental, LLC, 8550 Higuera Street, Culver City CA 90232. |  |  | July 21, 2011 | 2011 | Discus Dental LLC | The firm is initiating this recall of the SL3 Soft Tissue Laser equipped with a fiber optic Unifiber system which contains a fragile glass filament because the fiber broke near the hand piece. | Device Design | DISCUS Dental, sent an "URGENT MEDICAL DEVICE NOTIFICATION" letter dated July 21, 2011. The letter described the product, problem and action to be taken. Customers were instructed tocontact Discus Dental at 1-888-427-9279 for a no-cost replacement.For questions regarding this recall call 310-845-8200. |  | 150 | 150 units | Nationwide Distribution including AL, AZ, CA, CO, FL, GA, IA, ID, IL, IN, KS, MA, MD, MI, MO, MT NH, NJ, NM, NV, NY, OH, PA, SD, TX, UT, VA, and WA. | Terminated | April 29, 2013 | 648 | 150 | Computer
Z-1340-2013 | 64647 | No | KQM | Surgical camera and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Ref: VP1600 HD Still Capture System ConMed Linvatec . |  |  | March 18, 2013 | 2013 | Linvatec Corp. dba ConMed Linvatec | The VP1600 HD Still Capture Image System is recalled due to an anomaly between the software code in the VP1600 console and the communications interface for various keyboard key strokes (Caps lock and Num Lock) which may lock up and requires the user to reboot to continue. | Software design (manufacturing process) | The firm, ConMed, sent an "URGENT: MEDICAL DEVICE RECALL NOTIFICATION" letter dated 3/8/2013 to its customers after initial phone notifications. The letter identified the product, problem, and actions to be taken by the customers. |  | 267 | 267 units | Worldwide Distribution-USA (nationwide) including the Virgin Islands and the states of CA, CT, FL, GA, IN, KY, KS, LA, MD, MI, MN, MO, MS, MT, NE, NC, NJ, NY, OH, OR, PA, SC, TN, TX, UT, VA, and WI, and the countries of Austria, Belgium, Canada, Denmark, Hong Kong and Zambia. | Terminated | November 06, 2013 | 233 | 267 | Computer
Z-1326-2013 | 65028 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Fluoroscan Mini C-arm InSight, software versions 1.x, 2.x, 3.x |  |  | April 25, 2013 | 2013 | Hologic, Inc. | The audible alarm to alert the operator when cumulative irradiation time reaches five minutes is not accurate and the total cumulative irradiation time is not accurate. | Software design | The firm, Hologic, Inc. sent an "Urgent: Medical Device Recall" letter dated April 25, 2013, via FedEx to consignees/customers notifying them of the software defect and providing them with corrected user installable software version. The customers were instructed to complete and return the Enclosure to Customer Notification, Confirmation of Notification form via fax to Hologic at 866-652-8674 or mail the document to Hologic using the enclosed pre-addressed, pre-paid return envelope.Should you have any questions regarding this recall, require further information or assistance with the software installation call Hologic's Help Desk at 1-800-321-4659. |  | 287 | 287 | Worldwide distribution: USA (nationwide) and countries of: DE, RU, AU, IT, GB, ES, FR, KR, ZA, CA, AR, LT, TW, CN, BE, AT, JP, QA, SA, SG, HK, SI, NL, and SE. | Terminated | August 21, 2014 | 483 | 3067 | Computer
Z-1170-2013 | 64837 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | The GemStar Pump is a small and lightweight, single channel infusion device designed for use in the home, in the hospital, or anywhere electronic infusion is required. The GemStar Pump can be powered by AC mains adaptor, rechargeable battery pack, docking station, or two disposable AA alkaline batteries. When powered by batteries, The GemStar Pump is ideal for ambulatory patients. |  |  | March 18, 2013 | 2013 | Hospira Inc. | Pump shutting off during use without warning. | Device Design | Hospira sent an Important Device Information letter dated March 18, 2013, to all affected customer. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to inspect the affected product for signs of leakage, corrosion or other damage prior to each use. In addition, each time the batteries are replaced the battery compartment should be inspected for damage. If a device exhibits damage caused by leakage, it should immediately be removed from clinical service and Hospira should be contacted at 1-800-441-4100 to arrange for return of the device for repair. If the product has been further distributed they should notify their consignees and confirm they have done so by returning the reply form to Stericycle.For questions regarding this recall call 224-212-4892. |  | 69604 | List Number 13086: 286 pumps; List Number 13087: 45,376 pumps; List Number 13088: 23,942 pumps | Worldwide Distribution - USA (nationwide) and internationally to Australia, New Zealand, China, Hong Kong, Taiwan, Japan, Korea, Philippines, Thailand, Malaysia, Canada, Austria, Bahrain, Belgium, Croatia, Denmark, Egypt, Finland, France, Germany, Gibraltar, Greece, Hungary, Iceland, Ireland, Israel, Italy, Jordan, Kuwait, Lebanon, Libya, Luxembourg, Malta, Netherlands, Norway, Oman, Portugal, Qatar, Romania, Saudi Arabia, Slovenia, South Africa, Spain, Sweden, Switzerland, Turkey, UK, United Arab Emirates, Brazil, Chile and Colombia. | Terminated | March 10, 2017 | 1453 | 108405 | Computer
Z-1208-2013 | 64918 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE Healthcare, GE 1.5T SIGNA HDe MR SYSTEM. |  |  | June 08, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has recently become aware of a Reformat Image Flip Issue associated with GE MR Products.The axial and coronal reformatted images from a sagittal 3D series (such as with applications of FSPGR-3D T1, BRAVO T1, or CUBE T1) of over 126 slices may display a left-to-right image flip .In addition, the acquired sagittal images may be annotated with left-right location annotation revers | Software design | GE Healthcare sent an "Important Product Information" letter dated June 8, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers should contact their local service representative for questions regarding this notice. |  | 822 | 822 (154 US; 668 OUS) | Worldwide Distribution-USA (nationwide) including DC and Puerto Rico, excluding the state of RI and the countries of Yugoslavia YEMEN Viet Nam Venezuela Uzbekistan Uruguay United Arab Emirates Tunisia Thailand Tanzania Taiwan Turkey Syria Switzerland Sweden Spain South Korea South Africa Slovenia Slovakia Singapore Saudi Arabia Russia Romania Reunion Qatar Portugal Poland Philippines Peru Pakistan Paraguay Philippines Oman Norway Nigeria New Zealand Netherlands Morocco Mexico Martinique Malaysia Macedonia Mauritania Luxembourg Lithuania Libya Lebanon Latvia Kuwait Kenya Korea Korea (Republic Of [South] Korea) KUWAIT Kazakhstan Jordan Japan Jamaica Italy Israel Ireland Iran Iraq Indonesia India Hungary Hong Kong Honduras Guadeloupe Greece Germany Georgia, Republic of French Polynesia France Finland Estonia Ecuador Egypt Denmark Dominican Republic Czech Republic Croatia Colombia China Chile Canada Bulgaria Brazil Bosnia And Herzegovina Belgium Belarus Barbados Bahrain Austria Australia Argentina Arab Emirates Algeria and United Kingdom. | Terminated | November 03, 2014 | 878 | 12048 | Computer
Z-1210-2013 | 64942 | No | HTW | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker Howmedica Osteonics Corp. AxSOS System Instrument. |  |  | March 28, 2013 | 2013 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics has received a report indicating that the scale of the AxSOS Calibrated Drill Bit 4.3,,X263mm, flat , manufactured by Stryker Osteosynthesis, is incorrect. The incorrect scale could result in the surgeon selecting a screw that is 10mm longer than intended. | Device Design | Stryker Orthopaedics sent a Urgent Medical Device Removal letter dated March 28, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.1. Please inform users of this Medical Device Removal and pass thisnotice to all those individuals who need to be aware within yourorganization.2. Return all affected products available at your location toStryker Osteosynthesis or Contact Stryker Customer Service.3. Complete and sign the enclosed Business Reply Form and fax a copyto: 1-865-252-3635 .4. Keep a copy of the completed and executed Business Reply Form foryour records.Report any adverse events or product quality problems to StrykerOrthopaedics: 1-866-0R-ASSIST. (1-866-672-7747).As we strive for products that meet your expectations for quality and reliability, please do not hesitate to contact us, in case you have any further questions at (201) 831-5970.. |  | 107 | 107 units | Nationwide Distribution including the states of AL, AZ, IL, KY, LA, NJ, NC, FL, CA, NM, MO, TN, VA, CO, MA, TX, IN and MT. | Terminated | November 17, 2015 | 964 | 107 | Computer
Z-1324-2013 | 64986 | No | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | Segmented Cylinder Applicator Set, Model # GM11004150, when used for Pulsed Dose Rate (PDR) |  |  | April 11, 0013 | 13 | Varian Medical Systems, Inc. | Segmented cylinder applicator set may slip during treatment, causing the delivery of radiation to areas outside the target. | Device Design | Varian Medical Systems, Inc. notified customers of the recall with 'Urgent Medical Device Correction Urgent Field Safety Notice," dated 4/11/13, and delivered by mail to customers. The notice identified the product, problem, and instructions for the customers. Varian Medical Systems, Inc. informed customers of the possible malfunction of the GM11004150 Segmented Cylinder Applicator Set for pulsed dose rate brachytherapy treatment, with codes H640383, H640277, H640384, H640788, H640D126, H640385, H640372, H640303, H640387. The notice instructed customers to cease use of the Segmented Cylinder Applicator Set and return the devices along with a completed proof of notification form (included with the recall notice) to the recalling firm. |  | 21 | 21 devices | Worldwide distribution: USA (nationwide) and in the countries of: Belgium, France, Martinique, Slovenia, Sweden, and the United Kingdom. | Terminated | January 13, 2014 | 730762 | 21 | Computer
Z-1323-2013 | 64390 | No | DSQ |  |  | Cardiovascular | PMA | Heartware HVAD Pump Implant Kit (Heartware Ventricular Assist Device) |  |  | December 27, 2012 | 2012 | Heartware Inc | HeartWare, Inc. of Miami Lakes, FL is recalling their Ventricular Assist System due to the loosening of the HVAD Pump driveline connector. | Device Design | HeartWare, Inc. sent an Urgent Medical Device Correction letter to all their customers on February 7, 2013 instructing that at each routine clinic visit, they should inspect the patient's driveline connector as described in the attached Technical Bulletin. They were asked to sign and return the acknowledgement form to HeartWare. For questions contact your HeartWare representative or Customer Service at 1-877-367-4823. |  | 3544 | 3,544 | Worldwide Distribution-USA (nationwide) including the states of MN, IL, MI, FL, WI, TX, MO, CA, MA, TN, OH, NC, GA, VA, UT, KY, MD, IN, PA, NY, NJ, LA, OR, WA, AL and CO, and the countries of Poland, Australia, Austria, Italy, Germany, New Zealand, South Africa, Czech Republic, France, Brazil, Switzerland, Singapore, United Kingdom, Saudi Arabia, Netherlands, China, Chile, Belgium, Canada, Norway, Lebanon, Israel, Malaysia, Turkey, India, Argentina, Greece, Japan and Lithuania. | Terminated | July 09, 2013 | 194 | 3544 | Computer
Z-1322-2013 | 64221 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Proplege Coronary Sinus Catheter model PR9, 11 French Introducer, kitted with PR9. |  |  | January 25, 2013 | 2013 | Edwards Lifesciences, LLC | Edwards Lifesciences is recalling certain Proplege Coronary Sinus Catheter, model PR9 due to confirmed reports of blood leakage through the hemostasis valve in the introducer that is kitted with the catheter. | Device Design | Edwards Lifesciences sent an Urgent Field Safety Notice letter dated February 1, 2013. The letter identified the affected product, problem and actions to be taken. Customers were instructed to review entire inventory for the affected product, quarantine and return to Edwards. Contact Customer Services at 800-424-3278 for any question, to obtain a Returned Goods Authorization number and replacement product. |  | 807 | 807 | Worldwide Distribution - USA Nationwide including the states of: MD, NY, PA, CA, FL, AZ, AL, TX, VA, OH, MO, NC, GA, TN, CO, OR, SC, ME, IL, HI, NJ, WA, WI, CT, KY, LA and Canada. | Open, Classified |  |  | 807 | Computer
Z-1238-2013 | 64771 | No | KWI | Elbow joint radial (hemi-elbow) polymer prosthesis. | Orthopedic | Orthopedic | 510(k) | Synthes Radial Head Prosthesis: Trial Radial Head |  |  | February 22, 2013 | 2013 | Synthes USA HQ, Inc. | The Trial Head may come loose from the implant stem during manipulation of the arm during surgery. To date, no adverse events have been reported related to this issue and this recall is not being initiated as a result of adverse events. | Device Design | Synthes sent an Urgent Notice: Medical Device Recall letter dated February 22, 2013, to all affected consignees. The letter requested consignees examine their inventory, remove them from use and return them to Synthes. Customers were asked to complete the attached Verification Section at the end of the letter indicating whether or not they have any of the affected product. For questions customers were instructed to call 610-719-5450 or email FieldAction@synthes.com.For questions regarding this recall call 610-719-5000. |  | 417 | 417 | Nationwide Distribution: USA including states of: IL, IN, MA, MI, MO, NJ, NY, OH, OR, PA, SC, TX, VA, WA and WI. | Terminated | August 31, 2015 | 920 | 417 | Computer
Z-1245-2013 | 64876 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | FreeStyle InsuLinx Blood Glucose Monitoring System: |  |  | April 15, 2013 | 2013 | Abbott Diabetes Care, Inc. | AT rare, extremely high glucose levels (1024 mg/dL and above), the FreeStyle InsuLinx Blood Glucose Meter may provide an inaccurate reading. | Software design | Abbott Diabetes Care sent an Urgent Product Recall letter dated April 15, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to take one of the following options immediately to address the issue. 1) Customers can access a software update to resolve the issue at www.freestyleinsulinx.com/swupdate.2) Customers can contact Abbott Diabetes Care Customers Service at 1-866-723-2697 to expedite return and replacement of the affected product at no charge. For questions customers were instructed to call Abbott Diabetes Care Customers Service at 1-866-723-2697.For questions regarding this recall call 510-239-2775. |  | 544609 | 335,273 meters; 209,336 remain at firm. | Nationwide Distribution | Terminated | August 03, 2015 | 840 | 544609 | Computer
Z-1168-2013 | 64694 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Ventri Discovery NM 530c, Ventri with Premium Table Discovery NM 530c, models H3000ZW &amp; H3000SA. The system is used to perform nuclear imaging procedures. |  |  | January 09, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of an incident which caused a patient injury during an unload following an exam on a Ventri scanner. During a patient unload following a scan, the patients fingers may be pinched between the moving cradle edge and the stationary base of the table while the table is moving outward automatically. | Device Design | Consignees were sent on 1/9/2013 a GE Healthcare "Urgent Medical Device Correction" letter dated January 9, 2013. The letter was addressed to Managers of Nuclear Medicine, Hospital Administrators / Risk Managers andManagers of Radiology / Head of Biomedical Engineering. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. |  | 556 | 556 | Worldwide Distribution - US, including the states of AL, AZ, AR, CA, CO, CT, DE, DC, FL, IL, IN, IA, KY, LA, ME, MD, MA, MI, ,N,MS, MO, NE, NV, NH, NJ, NM, NY,NC, OH, OK,OR, PA, RI, SC, TN, TX, UT, VA, WA, WV, WI and the countries ofTAIWAN, TURKEY, SINGAPORE, SWD EN, PORTUGAL, POLAND, NEW ZEALAND, NORWAY, NETHERLANDS, MALAYSIA, MEXICO, SOUTH KOREA, JAPAN, ITALY, INDIA, ISRAEL GREAT BRITAIN, FRANCE, DENMARK, GERMANY, SWITZERLAND, CANADA, BRAZIL BELGIUM, AUSTRALIA, and ARGENTINA. | Terminated | August 06, 2013 | 209 | 556 | Computer
Z-1441-2013 | 65191 | No | HLI | Ophthalmoscope. | Ophthalmic | Ophthalmic | 510(K) Exempt | Optovue iVue with Normative Database (NDB) with Software Version 3.0 and 3.1. |  |  | May 14, 2013 | 2013 | Optovue, Inc. | Colors from the NDB comparison for Ganglion Cell Complex thickness in the singular OU report in iVue version 3.0 (US) and 3.1(international) are different from those in the individual iWellness ID and OS reports. | Software design | Optovue sent a Field Correction Notice letter dated May 16, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. |  | 329 | ~329 users | Distributed Nationwide and in Canada. | Terminated | October 24, 2013 | 163 | 329 | Computer
Z-1163-2013 | 64828 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand RPM Respiratory Gating System, RPM Gating v1.7.5 with 3D Option, Model Number: HS2 |  |  | March 26, 2013 | 2013 | Varian Medical Systems, Inc. Oncology Systems | Varian has identified an anomaly with the RPM Respiratory Gating System version I.7.5 with 3D Option. While operating with the 6-dot marker block, this version of RPM Gating may produce an inverted respiratory waveform in amplitude and phase-gating sessions. | Device Design | Varian sent an Urgent Medical Device Correction - Urgent Field Safety Notice Letter dated March 26, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to retain a copy of this document along with their most current product labeling, advise the appropriate personnel of the content of this letter, and for future reference this document will be posted to the Varian Medical Systems customer support website: http://www.MyVarian.com. The letter stated that a Varian Medical Systems has developed a technical correction and a Customer Service representative will contact customers to schedule installation. For question contact your local Varian medical Systems Customer Support District or Regional Manager. |  | 2508 | 2508 units | Worldwide Distribution - USA (Nationwide) and countries of: Australia, Austria, Bangladesh, Belgium, Brazil, Bulgaria, Canada, China, Columbia, Costa Rica, Cyprus, Czech Republic, Denmark, Dominican Republic, Estonia, Finland, France, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Kuwait, Latvia, Lebanon, Lithuania, Luxembourg, Macau, Malaysia, Mexico Morocco, Netherlands, New Zealand, Norway, Peru, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Singapore, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, Ukraine, United Arab E, and United Kingdom. | Terminated | July 09, 2014 | 470 | 2508 | Computer
Z-1546-2013 | 64424 | apparatus, traction, non-powered | HST | Nonpowered orthopedic traction apparatus and accessories. | Orthopedic | Neurology | 510(K) Exempt | Zimmer Patient Helper Bed Frame Adapter Assembly |  |  | February 14, 2013 | 2013 | Zimmer, Inc. | The bed frame adapter assemblies may have a weak weld where the patient helper adapter mounting bars attach to the tube. | Device Design | Zimmer sent an Urgent Device Removal letter dated February 14, 2013, to all affected consignees. The letter identified the product, the problem,and the action to be taken by the consignee. Consignees were advised to review the notification letter, follow the instructions contained in the letter, remove from use and quarantine the affected product. Consignees were also instructed to complete and return the fax back from to 877-787-0375 or zimmer3632@stericycle.com. Consignees were instructed to return the affected product back using the shipping label provided. If affected product was further distribution, consignees were asked to provide the customer's information to the firm on the back of the form. For questions regarding this recall call 866-201-9081. |  | 344 | 344 Units | Worldwide distribution: USA (nationwide) including states of: AL, AR, AZ, CA, CT, FL, HI, IL, KY, MA, MD, MN, MO, NE, NJ, NY, OH, OK, PA, TN, TX, WA, and WI; and country of Canada. | Terminated | May 01, 2014 | 441 | 344 | Computer
Z-1584-2013 | 65246 | No | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | CoaguChek XS System (HCP) |  |  | May 16, 2013 | 2013 | Roche Diagnostics Operations, Inc. | Roche has confirmed the potential for an undetected elevated INR test result with the CoaguChek XS, CoaguChek XS Plus, and CoaguChek XS Pro meters. In rare cases, instead of a value, an ERROR 6 message is displayed. | Software design | Rochet sent an Urgent Medical Device Correction letter dated May 16, 2013, via UPS Ground to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were issued the following instructions: If you are using Warfarin therapy in combination with antibiotics and/or chemotherapy drugs, which could potentially lead to an extremely high test result (&gt; 10 INR), you may in rare cases receive repeated ERROR 6 messages instead of a test result. If the ERROR 6 message is displayed repeatedly on your CoaguChek XS meter, immediately contact your treating physician.The CoaguChek XS PT Test Strip package insert; the online version of the CoaguChek XS PST system users manual; and the CoaguChek XS systems are being updated to provide additional ERROR 6 message information and actions. Updated versions will be available on the www.poc.roche.com website by the end of June 2013.Health Care Professionals are instructed to use an alternative method to determine INR results if an ERROR 6 message is displayed repeatedly on the CoaguChek XS, CoaguChek XS Plus, or CoaguChek XS Pro meters for the same patient. Self testing patients are instructed to contact their Health Care provider immendiately if the ERROR 6 message is displayed repeatedly on the CoaguChek XS.Please contact your Roche Account Manager or Roche Diagnostics Point of Care Technical Service, 24 hours a day, seven days a week at 1-800-428-4674 if you have questions about the information contained in this UMDC.For questions regarding this recall call 317-521-3711. |  | 86553 | 86553 meters | Nationwide Distribution | Terminated | April 16, 2014 | 335 | 89896 | Computer
Z-1587-2013 | 65371 | No | PAA |  |  | Radiology | PMA | Soma v Automated Breast Ultrasound System (ABUS), Automated Breast Ultrasound. |  |  | May 16, 2013 | 2013 | U-systems Inc | Devices labeled for an intended use not included in the existing 510(k) or PMA, in that, the user manual contains the Indication for Use of the soma v Platinum ABUS [a 510(k) product]. | Software design | On 6/5/13, an "URGENT MEDICAL DEVICE CORRECTION" letter was distributed to all affected users, with a description of the problem and user corrective action steps. One consignee received an Urgent Notice to Customer - Medical Device Correction letter on 5/16/13. Contact the firm at 408-328-7925 for questions regarding this notice. |  | 3 | 3 units | Distributed in California and Illinois. | Terminated | November 04, 2013 | 172 | 3 | Computer
Z-1574-2013 | 65342 | No | LKK |  |  | General Hospital | PMA | Medtronic Intrathecal Catheter Pump Segment Revision Kit, model 8596SC. Contents: 60-cm pump segment with attached sutureless pump connector, Spinal Segment Strain-relief sleeves, Pump segment strain-relief sleeves, Connector pin. Contents of inner package are STERILE. |  |  | June 03, 2013 | 2013 | Medtronic Neuromodulation | The Sutureless Connector (SC) Intrathecal Catheter connector has been redesigned to reduce the potential for occlusion at the catheter to pump interface. Medtronic is removing the unused products from the market that were manufactured with the previous design, and recommend the previous design no longer be used due to greater potential for misalignment and subsequent occlusion. | Device Design | The firm, Medtronic, sent an "Urgent: Medical Device Removal" letter dated May 2013 to its customers. The letter described the product, problem and actions to be taken. Representatives (Rep) are visiting all locations, beginning June 3, 2013, to retrieve devices with a Use By date of 2014 08 14 (August 14, 2014) or sooner. The Rep is leaving a letter with the hospitals to tell them that the Sutureless Connector Intrathecal Catheter connector has been redesigned and that they are removing unused devices. They also do not recommend using any of the old design. An Account Specific Customer Confirmation Form will be completed by the Rep. A copy of the completed form will be left with the Hospital along with the Urgent Medical Device Removal letter.If you have questions, please contact Medtronic Neuromodulation Technical Services at 1-800-707-0933 weekdays 7am - 6pm CST. |  | 115722 | 115,722 total | Worldwide distribution: US (nationwide) and countries of: Aruba, Australia, Austria, Belgium, Canada, Chile, China, Colombia, Costa Rica, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Finland, France, Georgia, Germany, Greece, Guadeloupe, Hong Kong, Hungary, Iceland, India, Ireland, Israel, Italy, Jordan, Korea, Kuwait, Lebanon, Lithuania, Luxembourg, Malta, Martinique, Morocco, Netherlands, Netherlands Antilles, New Zealand, Norway, Panama, Poland, Portugal, Puerto Rico, Qatar, Romania, Russian Federation, San Marino, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Thailand, Trinidad and Tobago, Turkey, United Arab Emirates, United Kingdom, Uruguay, and Venezuela. | Terminated | April 07, 2014 | 308 | 115722 | Computer
Z-1098-2013 | 64728 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris PC unit model 8015 with software version 9.12 |  |  | March 06, 2013 | 2013 | Carefusion Corporation | The recall was initiated because Carefusion has received reports of a communication error on the Alaris PC unit model 8015 with software version 9.12 when the Alaris EtCO2 module or Alaris SpO2 module is attached. | Software design | CareFusion sent an Urgent Medical Device Recall Notification letter dated March 6, 2013 to all affected customers. The letter identified the affected product, problem and action to be taken. Customers are required to confirm receipt of the notification by returning the Recall Response card to Carefusion by postage-paid, self-addressed mail, fax, or email. Carefusion representatives will contact customers by telephone to schedule field remediation. For questions contact CareFusion Support Center at 888-562-6018, Customer Advocacy at 888-812-3266 and/or Technical Support at 888-812-3229. |  | 9241 | 9,241 units | Worldwide Distribution - USA (nationwide) and the countries of Australia and China. | Terminated | September 02, 2014 | 545 | 9241 | Computer
Z-1562-2013 | 65435 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARTISTE with syngo RT Therapist version 4.3 |  |  | May 29, 2013 | 2013 | Siemens Medical Solutions USA, Inc | The potential for a safety issue, for the image quality degradation, that is addressed by the Customer Safety Advisory Notice when verifying the patient position based on 2D images that were acquired as single-exposure MV or In-Line kView images using the syngo RT Therapist 4.3. | Device Design | Siemens sent an Urgent: Medical Device Correction Customer Safety Advisory Notice letter to all affected customers. The letter identified the affected product, problem, what preventive measures can the user take and actions to be taken. Customers were advised to perform the preventive measures provided and inform all affected personnel immediately. The letter states that Siemens is currently working on a new version of the software. Customers are asked to include the Customer Safety Advisory Notice in the System Owner Manual chapter "Safety Advisory Letters". For question call 925-602-8157. |  | 8 | 8 units | Worldwide distribution | Terminated | November 26, 2013 | 181 | 8 | Computer
Z-1567-2013 | 65379 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | CARESCAPE Monitor B850 |  |  | July 20, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare become aware of 4 potential issues associated with the CARESCAPE Monitor B850. These issues include CARESCAPE Monitor B850 uDOM failure, CARESCAPE Monitor B850 unstable USB communication, CARESCAPE Monitor B850 remote printing issue with currently printing PRN 50 recorder and CARESCAPE Monitor B850 reboots due to an X server freeze. | Software design | GE Healthcare sent "Important Product Information" letter dated 12 July 2012 to all affected customers. The letter described the Subject (issues) , Instructions, Affected Product Details, Product Correction and actions to be taken. For questions contact Technical Support 1-800-558-7044 or your local Service Representative. |  | 12779 | 12,779 | Worldwide Distribution - USA Nationwide including the states of: AL, AK, AZ, AR, CA, CO, CT, DC, FL, GA, ID, IL, IN, IA, KS, KY, IA, ME, MD,MA, MI, MN, MS, MO, MT, NE,NV,NH, NJ, NM, NY, NC, OR, OK OH, PA, PR, SC, TN, TX, UT, VA, WA, WY, WI, WY, and the countries of: ARGENTINA, AUSTRALIA, AUSTRIA, BAHRAIN, BELGIUM, BRAZIL, CANADA, CHILE, COLOMBIA, COSTA RICA, CYPRUS, CZECH REPUBLIC, DENMARK,, EGYPT, ESTONIA, FINLAND, FRANCE, GERMANY, HONG KONG, INDIA, IRAQ, ISRAEL, ITALY, JAPAN, KOREA, KUWAIT, LEBANON, LITHUANIA, MALAYSIA, MALTA, MEXICO, NETHERLAND, NEW ZEALAND, NORWAY, PANAMA, PERU, POLAND, PORTUGAL, QATAR, SAUDI ARABIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY. | Terminated | June 30, 2014 | 710 | 12779 | Computer
Z-1558-2013 | 65129 | No | CCW | Rigid laryngoscope. | Anesthesiology | Anesthesiology | 510(K) Exempt | GlideScope Video Laryngoscope (GVL) 3, 4, 5 and GlideScope (AVL) 2, 3, 4 and 5 reusable blades |  |  | May 08, 2013 | 2013 | Verathon, Inc. | Verathon GlideScope Reusable GVL and AVL Blades are being recalled due to the potential risk of premature failure/breakage of the blade tip that may not be readily visible during routine inspection before or after intubation. | Device Design | Verathon issued an Urgent - Safety Alert and Medical Device Recall letter, dated May 10, 2013, to all affected customers The letter identified the affected product, problem and actions to be taken. The letter advises customers to discontinue the use of the specific serial numbers listed in the letter and asked that the affected product be returned to Verathon for replacement. The customers are also directed to fill out a return response form (included in the letter) regarding this recall.Consignees with questions are directed to contact their Verathon representative or Verathon Customer Care at 800-331-2313 (US and Canada) and 1-425-867-1348 for those consignees outside of US and Canada. the support team will be available Monday through Friday from 6 AM - 5 PM PST for the US and from 6 AM to 4:30 pm PST outside the US. Alternately they can e-mail Verathon at customerservice@verathon.com |  | 4784 | 4784 units worldwide (3,381 in the US and 1403 international) | Worldwide distribution: USA (nationwide) including states of : AK, AL, AZ, CA, CO, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY, Washington DC, Guam, and countries of : Argentina, Australia, Brazil, Canada, China, Ecuador, France, Germany, Hong Kong, India, Indonesia, Italy, Japan, Japan, Malaysia, Panama, Peru, Philippines, Singapore, South Korea, Taiwan, Thailand, The Netherlands, United Kingdom, and Venezuela. | Terminated | December 17, 2013 | 223 | 4784 | Computer
Z-1560-2013 | 65331 | No | JQW | Pipetting and diluting system for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Tecan Freedom EVO with Touch Tools Suite Version 3.0 |  |  | February 19, 2013 | 2013 | Tecan US, Inc. | WIth TouchTools 3.0 after the execution of a Rich User Prompt in Display and Wait mode, door lock errors are ignored. If an operator presses "Continue" following a Rich User Prompt and the door has been left open, the script will progress (continue without warning). This could lead to the instrument moving unexpectedly and potentially cause an injury to the operator if they are physically inside t | Software change control | The firm, Tecan, sent an "Urgent Field Corrective Action" letter dated February 15, 2013 to its consignees/customers and also contact them by phone February 19, 2013. The letter describes the product, problem and actions to be taken. The customers were instructed to not use Rich User Prompts in Display and Wait mode until TouchTools 3.0 Patch 1 or a higher version has been installed. Tecan is currently testing and releasing a patch that will correct this issue. The patch will be available by the end of February. Tecan will contact you upon its release as installation of the patch is required. If you have any further questions, please contact your local Tecan Helpdesk for assistance or call 919-361-5200 ext 19519. |  | 12 | 12 units | Worldwide distribution: US (nationwide) including states of: CA, KY, NC, OH, TX, and WI; and countries of: France, Germany and Italy. | Terminated | June 20, 2013 | 121 | 12 | Computer
Z-1559-2013 | 65438 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Echelon, Echelon Oval and Oasis MR imaging device systems; one unit per box. |  |  | January 31, 2013 | 2013 | Hitachi Medical Systems America Inc | Hitachi discovered a software error that can occur when simultaneously scanning a patient while processing images previously acquired from a different patient. The software error can occur randomly. The error causes an image in the processed dataset to have the same unique identifier (UID) as in the scanning dataset. The processed image is stored in the place of the original image, which is lost. | Software design | Hitachi sent a Device Correction Notice beginning on January 31,2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Software correction is targeted to start by March 1, 2013. All sites will be visited by Hitachi Service to perform corrections. All letters are recorded by tracking number and date received by the customers.For further questions please call 800-800-3106 x 3720. |  | 45 | 45 system units | Worldwide Distribution - US Nationwide including the states of AZ, CA, CO, DE, FL, IL, IN, KS, KY, MD, NC, ND, NY, OH, OR, PA, TX, VA, WA, and WI. One (1) system unit was also distributed to Mexico. | Terminated | September 20, 2013 | 232 | 45 | Computer
Z-1551-2013 | 65229 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, PET VCAR. |  |  | May 09, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue with respect to the automated PERCIST response criteria calculation applied in the PET VCAR application on Advantage Workstation or AW Server. | Software design | GE Healthcare sent an "Urgent Medical Device Correction" letter dated May 14, 2013, to all affected consignees. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. Customers with questions were instructed to call one of the following numbers:United Sates 800-437-1171Japan 0120-055-919For other countries they were instructed to contact their local GE Healthcare Service Representative.For questions regarding this recall call 407-475-4717. |  | 231 | 231 (38 US; 193 OUS) | Worldwide Distribution - USA including CA, CO, CT, FL, GA, IL, NY, NY NC OH, TX, VA, WV, WI.Internationally to: AUSTRALIA, AUSTRIA, BAHARIN, CANADA, CHILE, CHINE, CZECH REPUBLIC, DENMARK, FINLAND, FRANCE, GEORGIA, GERMANY, HONG KONG, INDIA, INDONESIA, ITALY, JAPAN, KOREA, MEXICO, MOROCCO, NIGERIA, POLAND, ROMANIA, RUSSIA,SAUDI ARABIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TURKEY, UNITED KINGDOM. | Terminated | January 21, 2014 | 257 | 231 | Computer
Z-1548-2013 | 65212 | No | PAO | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Superior Approach Resection Guide Assembly. |  |  | May 15, 2013 | 2013 | Biomet, Inc. | Biomet has initiated this action following an investigation which identified that the slot on the resection guide is offset in the wrong direction.When the slot on the resection guide is offset in the wrong direction, the Steinman pin that is intended to hold the guide into place may not come into contact with the bone. | Device Design | On 5/15/2013 an URGENT MEDICAL DEVICE RECALL notification was sent to all consignees detailing to immediately locate and discontinue use of the product and return to Biomet.The following actions are REQUIRED:-Immediately locate and remove the identified device(s) listed below from circulation .-Carefully follow the instructions on the enclosed "FAX Back Response Form" .-Fax a copy of the Response Form to 574-372-1683 prior to return of product.-Use priority carrier for your shipment.- If you have further distributed this product. you MUST notify hospital personnel of this action via the enclosed "Dear O.R. Manager" notice. This letter MUST be given to hospital personnel responsible for receiving recall notices. |  | 17 | 17 | Worldwide Distribution-USA (nationwide) including the states of UT, TX, AL, KY, and WI, and the countries of Australia, The Netherlands, Korea, and Canada. | Terminated | June 09, 2014 | 390 | 17 | Computer
Z-1537-2013 | 65376 | Yes | MNT | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Respironics V60 Ventilator, Model #V60 |  |  | June 04, 2013 | 2013 | Respironics California Inc | Respironics California, Inc. has initiated a recall on the V60 ventilator because of an issue with the software on the V60 Power Management Board Assembly. If a component fails on the Power Management Board Assembly, it may cause ventilator support to be lost with potentially no audible alarm from the ventilator. | Software design | The firm, Philips Healthcare, sent a "MEDICAL DEVICE CORRECTION" letter dated June 3, 2013 to all customers who received the V60 Ventilators. The letter describes the product, problem identified and the actions to be taken. The customers were informed that the V60 ventilator may continue to be used in accordance with its directions for use, pending the completion of the update. In addition, the customers were instructed to refer to their manual for additional information on warnings. A Philips Field Service Engineer, Approved Sevice Provider or Distributor will be contacting the customers to schedule a no-cost update and replacement of the Power Management Board Assembly Software on all V60 ventilators shipped from the manufacturer prior to April 1, 2013.If you need any further information or support concerning this issue, contact your local Philips Respironics representative at (800) 722-9377. |  | 19200 | 19,200 units total (12,060 units in US) | Worldwide distribution: US (nationwide) and countries of: Argentina, Australia, Austria, Bahrain, Belgium, Bolivia, Brazil, Canada, Chile, China, Croatia, Czech Republic, Denmark, Dominican Rep., Ecuador, Egypt , Finland, France, Germany, Hong Kong, Iceland, India, Indonesia, Iran, Ireland, Israel, Italy, Kenya, Kuwait, Lebanon, Libya, Malaysia, Malta, Mexico, Morocco, Namibia, Netherlands, New Zealand, Norway, Oman, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Puerto Rico, Qatar, Reunion, Russian Fed., Saudi Arabia, Singapore, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom, Utd.Arab.Emir. and Vietnam. | Terminated | April 24, 2014 | 324 | 19200 | Computer
Z-1133-2013 | 64748 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare Muse v7 cardiology information system. Model number 2026443-001: MUSE 7.0 app cd, Model Number 2026443-015: MUSE 7.0.2 app cd, Model number 2026443-017: MUSE 7.1 app cd, Model number 2026443-029:MUSE 7.1.1 A app cd, and Model Number 2026443-031: MUSE 7.1.1 app cd. |  |  | January 04, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare conducted a recall on the Muse v7 because incorrect medications were displayed on the MUSE, which could result in incorrect treatment. | Software design | Consignee letters were not sent for this recall. GE Healthcare's recall strategy was to update all affected systems that were distributed. Updating of the devices was carried out from February 2012 to August 2012 by GE Field personnel. |  | 963 | 963 | Worldwide Distribution -- USA, including all states except ME, and the U.S. Virgin Islands; and, the countries of AUSTRALIA, AUSTRIA, BELGIUM, CANADA, CHILE, CHINA, CZECH REPUBLIC, DENMARK, ESTONIA, FINLAND, FRANCE, GERMANY, GREECE, INDIA, IRELAND, ISRAEL, ITALY, JAPAN, KOREA, KUWAIT, MALAYSIA, NETHERLAND, NEW ZEALAND, NORWAY, PHILIPPINES,S QATAR, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA SPAIN, SWEDEN, SWITZERLAND, TAIWAN, TURKEY, UNITED ARAB EMIRATES, and UNITED KINGDOM. | Terminated | July 19, 2013 | 562 | 963 | Computer
Z-1547-2013 | 65210 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Biomet Trauma Phoenix Retrograde Femoral Connecting Bolt Trauma Fixation Systems. |  |  | April 29, 2013 | 2013 | Biomet, Inc. | Biomet Trauma ("Biomet") has initiated a recall of Retrograde Femoral Connecting Bolt, which involves Part Number 14-442021. The connecting bolt has an undersized diameter specification that may cause an interference fit with the 4mm hex driver (Part Number: 41024) near the edge of the tolerance. If the 4mm hex driver becomes stuck in the connecting bolt it cannot be taken apart and a delay in su | Device Design | Biomet sent an URGENT MEDICAL DEVICE RECALL notification letter dated April 29, 2013 to all affected customers. The letter identifed the affected product, problem and actions to be taken. The letter instructed customers to: immediately locate and remove the identified affected device(s), carefully follow the instructions on the enclosed "FAX Back Response Form", fax a copy of the Response Form to 574-372-1683 prior to return of product, use priority carrier for your shipment, if you have further distributed this product, you MUST notify hospital personnel of this action via the enclosed "Dear O.R. Manager" notice. For questions call 574-372-1570. |  | 75 | 75 | Worldwide Distribution - USA Nationwide: CA, TX, NY, PA, NJ, FL, OH, MI, CO, SD, IN, MT, GA, and WI. and the countries of Netherlands, Costa Rica, Japan, and San Juan. | Terminated | May 06, 2015 | 737 | 75 | Computer
Z-1536-2013 | 65091 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | 90 Degree Cannulated Infant Blade Plate 35mm x 5mm x 3 Hole |  |  | April 19, 2013 | 2013 | OrthoPediatrics Corp | Post Market Surveillance of the 90 degree Cannulated Infant Blade Plate indicated that there have been five reports concerning bending of the plate intra-operatively. | Device Design | OrthoPediatrics sent an Urgent Voluntary Recall notifiation via telephone and email to all affected consignees. The letter instructed consignees to take possession of any affected devices and quarantine immediately. A Mandatory Reply Form is included with the notification letter to be completed and returned. Please contact Greg Teghtmeyer if you have any questions (574-268-6379). |  | 130 | 130 devices were shipped | USA Nationwide Distribution | Terminated | September 19, 2013 | 153 | 130 | Computer
Z-1164-2013 | 64804 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Healthcare, Xeleris 2 Processing &amp; Review Workstation. |  |  | October 06, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware of a software issue in your Xeleris 2.1 workstation (that is used for processing Discovery 530/570 data), that may impact patient safety.On imported CT scans, Myocardial Perfusion Imaging (MPI) studies may create artifacts in attenuation corrected images. This may lead to a potential for misdiagnosis. | Software design | GE Healthcare sent an Urgent Medical Device Correction letter dated October 5, 2011, to all affected consignees. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. Consignees were instructed to contact their local GE Healthcare Service Representative if they have questions or concerns. They may also contact GE Healthcare Call Center , United States or Canada at 800-437-1171.Other countries 262-896-2890.For questions regarding this recall call 262-513-4122. |  | 11 | 11 | Worldwide Distribution - USA including GA, and NY; and countries of: Canada, Israel, Italy and Switzerland. | Terminated | February 04, 2014 | 852 | 11 | Computer
Z-1137-2013 | 64900 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Healthcare, AW Server. |  |  | June 28, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue with respect to the Results Viewer in the AW Server.After starting the Results Viewer with the selected series of a patient, the pixel and certain annotation data from a different patient is displayed in one of the viewports. This occurs on an occasional basis, when multiple users are simultaneously accessing the Results Viewe | Software design | GE Healthcare sent an "Important Product Information" letter dated June 27, 2012. The letter was addressed to Hospital Administrators , Radiology Managers and Risk Managers. The letter described the Subject, Instructions and Affected Product Details. GE Healthcare will correct all affected product by providing a software upgrade at no cost. For questions call: US 800-437-1171, Japan 0120-055-919, and/or local GE Healthcare Service Representative. |  | 429 | 429 | Worldwide Distribution - USA Nationwide including the states of: AK, AZ, CA, CO, CT, DE, DC, FL, GA, HI, IL, IN, IA, LA, MD, MI, ,N, ,S, MO, MT, NJ, NM, NY, NC, OH, OK, PA, SC, SD, TN,TX, UT, VA, WA, WI. and the countries of : AUSTRALIA, AUSTRIA, BELGIUM, BOSNIA, BRAZIL, CANADA, DENMARK, EGYPT, ESTONIA, FINLAND, FRANCE, GERMANY, GREAT BRITAIN, GREECE, GUADELOUPE, HONG KONG, INDIA, ISRAEL, ITALY, JAPAN, KOREA, LIBYA, MALAYSIA, NETHERLANDS, NEW ZEALAND, NORWAY, PANAMA, POLAND, QATAR, REUNION, ROMANIA, RUSSIA, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, THAILAND, TURKEY and UNITED ARAB EMIRATES. | Terminated | February 04, 2014 | 586 | 429 | Computer
Z-1530-2013 | 65330 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ECHELON 60mm Endoscopic Linear Cutter Reloads Black in Flexible Blister pack with Tyvek lid. |  |  | May 13, 2013 | 2013 | Ethicon Endo-Surgery Inc | Ethicon Endo-Surgery is initiating a voluntary worldwide recall for ECHELON 60mm Black Reload (ECR60T) due to the potential for incomplete staple line formation from reload damage during the firing sequence. | Device Design | Ethicon sent a Urgent: Medical Device Recall letter dated May 29, 2013, to all affected customers. The letter identified the product the problem and the action to be taken by the customer.Customers were instructed1.To examine your inventory immediately to determine if you have affected product manufactured by Ethicon and remove the affected product(s).See Attachment A for assistance in identifying affected product.2. Complete the Business Reply Form (BRF) Attachment C and fax or email it to Stericycle at 1-888-943-4190. Please reference Event 4706. If you have product to be returned, keep a copy of this form for your records.Your JJHCS number has been pre-populated on the BRF. Replacement product will be issued for all returned recalled product.3. All affected product must be returned by August 30, 2013, to receive replacement product. To return affected product, photocopy your completed Business Reply Form, placeit in the box with your product, and affix the pre-printed UPS shipping label to return to Stericycle. Returned ECHELON 60mm Black Reload (ECR60T) will be replaced with corrected ECHELON 60mm Black Reload (ECR60T) as soon as supply becomes available.4. If you do not have affected product in your inventory, you should still complete and return the Business Reply Form, indicating you have no affected product.5. Please share this information with all of the appropriate staff at your facility.6. If you need assistance identifying an alternative product code, please speak directly with your Ethicon Sales Representative. If you need clinical or product support, please contact your local sales representative or call ourCustomer Support Center, 1-877-ETHICON (877-384-4266 OPTION 6 to speak with a registered nurse). If you need additional shipping labels or a communications package, contact the Stericycle support center at 1-888-943-2397 and reference Event 4706. |  | 57540 | 57,540 Units | Worldwide Distribution - US Nationwide all fifty states and including the countries of Canada, Belgium, India, Malaysia, Saudi Arabia, Brazil, Israel, Mexico, Singapore, Jordan, New Zealand, Taiwan, Chile, Korea, Puerto Rico, Turkey, Columbia, Kuwait, Qatar, United Arab Emirates, Egypt, Lebanon and Russia. | Terminated | February 11, 2014 | 274 | 57540 | Computer
Z-1531-2013 | 65341 | No | LZF | Infusion pump. | General Hospital | General Hospital | 510(k) | VIA Medical Collection Bag; The collection bag is an accessory waste bag used in conjunction with the sterile LVM blood gas sensor with the LVM Monitor. The collection bag is packaged in a tyvek pouch which is included in either replacement kit (777-3103) or LVM sensor kit (777-3102). |  |  | May 24, 2013 | 2013 | International Biomedical, Ltd. | International Biomedical has received reports of interference between the collection bag and the extension set connectors. Evaluation of the affected components indicates that the collection bag connector can become occluded due to an over-tightening of the luer lock by the end user. To alleviate the occurrence of this issue in the fixture, the connector on the collection bag has been improved to | Device Design | International Biomedical decided to recall and sent "Product Bulletin" notification letters to their consignees on 05/24/2013 or 05/30/2013. The letter described the product issue and provided customers with recommended actions. Questions were directed to (512) 873-0033. |  | 7000 | 7,000 | Nationwide Distribution. | Terminated | July 29, 2013 | 66 | 7000 | Computer
Z-1149-2013 | 64641 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | MAQUET CARDIOHELP Support System: CARDIOHELP-I Software (up to version 3.3.0.0) |  |  | March 07, 2013 | 2013 | Maquet Cardiovascular, LLC | It has come to the attention of MAQUET that there have been episodes of brief unexpected shutdown, followed by an automatic device restart of the human machine interface (touch screen) of the CARDIOHELP device. The human machine interface (HMI) is the central display of the CARDIOHELP device, where measured values, settings and alarms are displayed and adjusted/confirmed, with the exception of bl | Software change control | MAQUET sent an Urgent Medical Device Field Correction letters dated March 7, 2013 via Federal Express to all affected customers. The letter identifies the affected product, problem and actions to be taken. Customers are asked to complete and submit the included Customer Fax-back Form. Maquet Service will exchange all affected products with the updated version. For questions MAQUET Customer Service at 888-627-8383 (press option 2, followed by option 2). |  | 154 | 154 units affected | Worldwide Distribution - USA Nationwide and the following countries: Australia, Austria, Belarus, Belgium, Canada, Columbia, Croatia, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, Hong Kong, Iceland, Iran, Ireland, Italy, Jordan, Kazakhstan, Korea, Kuwait, Luxembourg, The Netherlands, Norway, Oman, Poland, Portugal, Russia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Thailand, Turkey, United Arab Emirates and the United Kingdom. | Terminated | August 22, 2013 | 168 | 154 | Computer
Z-1144-2013 | 64683 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Vitrea CT Multi-Chamber Cardiac Functional Analysis application on Vitrea version 5.2 (Vitrea Enterprise Suite 1.3 and Vitrea fX 3.1) and subsequent versions of VitreaWorkstation, Vitrea fX, VitreaAdvanced and Vitrea Enterprise Suite, through version 6.4. |  |  | March 15, 2013 | 2013 | Vital Images, Inc. | Vital Images has made a decision to perform a field correction to the Vitrea CT Multi-Chamber Cardiac Functional Analysis application. Vital Images has identified a defect in which numerical values may be incorrect in the Results Table under the Cardiac Analysis portion of the user interface. | Software design | Vital Images sent an Urgent Field Safety Notice dated March 15, 2013, to all affected consignees. The letter described the problem and the product affected by the recall. Letter provided a series of steps to follow which included to cease use of the device and to send the signed enclosed Effectiveness Check form to fieldnotices@vitalimages.com or fax to 1-952-540-3717. For questions customers were instructed to call Customer Support at 1-800-208-3005.For questions regarding this recall call 952-487-9500. |  | 102 | 102 | Worldwide Distribution - USA including AL, AR, CA, CO, IA, MI,MN, MO, NY, NC, OH, OR, PA, TX, and UT.Internationally to AUSTRIA, BELGIUM, CANADA, CHINA, COLOMBIA, GEORGIA, JAPAN, KOREA, MACEDONIA, NETHERLANDS, POLAND, PORTUGAL, RUSSIAN FEDERATION, SINGAPORE, SWEDEN, SWITZERLAND, TURKEY, UKRAINE, UNITED KINGDOM. | Terminated | February 27, 2015 | 714 | 102 | Computer
Z-1522-2013 | 65283 | No | NME | Dental bur. | Dental | Dental | 510(k) | StarDental Classique(TM) Diamond Dental Bur. |  |  | May 28, 2013 | 2013 | Dental EZ Stardental Division | Recall was initiated because a manufacturing issue was found that prevented the diamond coating from adhering properly to the shank of the diamond bur for these lots. | Equipment maintenance | Dental EZ Inc. sent an "URGENT: MEDICAL DEVICE RECALL" letter on May 28, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at 866-383-4636 ext.4350 for questions regarding this recall. |  | 3560 | 3,560 | Worldwide Distribution-USA (nationwide) including the states of CA, FL, IA, IN, NJ, NV, NY, PA, TN, TX, and WI, and the country of Japan. | Terminated | March 18, 2014 | 294 | 3560 | Computer
Z-1141-2013 | 64723 | No | JGS | Sodium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | HORIBA ABX PENTRA 400 I.S.E. Module, Model Numbers: P400ISE110EN02 and P400ISE110US02, All Serial Numbers, Software Versions 5.0.7 and Below. |  |  | March 20, 2013 | 2013 | Horiba Instruments, Inc dba Horiba Medical | HORIBA Medical is informing all ABX PENTRA 400 Analyzer with I.S.E. Module customers that there are noLinearity Limit Flags for the I.S.E. assays currently programmed into the ABX PENTRA 400 Analyzer software. Allother assays run on this analyzer do have Linearity Limit Flags to alert the user to values that fall below or above thevalidated linearity range for each assay. | Software design | A recall letter, Field Safety Notice &amp; Acknowledgement mailed to the customers March 20, 2013. The letter included the product information, reason for recall, instructions on what to do with the recalled product, and contact information:"If you have any questions or concerns, please contact your local HORIBA Medicalrepresentative. We sincerely apologize for any inconvenience this may cause your laboratory. Please contact your Customer Service Representative at 1-888-903-5001 (option 3) for any additional questions related to this customerletter." |  | 905 | 905 units | Nationwide Distribution. | Terminated | September 26, 2014 | 555 | 905 | Computer
Z-1488-2013 | 65062 | No | GDC | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Hand Controls for Operon D750, D752, D760, D820 and D850 |  |  | March 12, 2013 | 2013 | Berchtold Corp. | The firm received complaints for unintended movements. | Device Design | Berchtold sent an Urgent Medical Device Recall letter in March 2013, to all customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were recommended to clean the affected product and use the instructions to reduce the possibility of unintended movements until the firm could replace the affected product. Customers with questions were advised to contact the Berchtold Operon Pendanct Recall Coordinator at 800-243-5135, or 843-569-6100, ext 306.For questions regarding this recall call 843-569-6133. |  | 3763 | 3763 | Nationwide Distribution | Terminated | May 08, 2017 | 1518 | 3763 | Computer
Z-1516-2013 | 65113 | No | LFL |  |  | General &amp; Plastic Surgery | 510(k) | American Optisurgical's TX1 Tissue Removal System Console. The console is a software unit that is indicated for use in surgical procedures. |  |  | March 30, 2013 | 2013 | American Optisurgical Inc | The recall has been initiated due to an overheating protection voltage suppressor on the redundant safety protection circuit board. Secondary to the recall, American Optisurgical, Inc. will be making a correction to add spacers between the valve and cassette to correct a small irrigation leak that can occur after the priming cycle. Use of this product may cause the device to lose power and stop wo | Device Design | "Urgent Medical Device Recall" notifications letters were sent to customers on May 20, 2013 via USPS registered mail. Included in the Recall Notification Letter were the recommended actions and instructions for customers. Questions were referred to (949) 580-1266. |  | 194 | 194 units | Worldwide distribution, US, including the states of United States: Alabama, Arkansas, Arizona, California, Colorado, Florida, Georgia, Iowa, Illinois, Indiana, Kansas, Kentucky, Massachusetts, Maryland, Michigan, Minnesota, Missouri, North Carolina, Nebraska, New Jersey, New York, Ohio, Pennsylvania, South Carolina, Tennessee, Texas, Utah, Virginia, Washington Wisconsin, West Virginia and the countries of Singapore and the Philippines. | Terminated | March 24, 2014 | 359 | 194 | Computer
Z-1480-2013 | 64927 | No | OSI | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | E-Z WIDER BARIATRIC CHAIR BED; Single Motion Expandable Width Technology |  |  | August 02, 2012 | 2012 | Kreg Medical Inc. | An EZ Wider bed (capacity of 1000 lb. safe working load collapsed during patient use, leaving the patient in a forward tilt position. | Device Design | Kreg Medical sent an Urgent Voluntary Medical Device Recall letter dated May 10, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Kreg Medical started replacing potentially defective beds with refurbished beds in August 2012. All of the potentiall defective beds were refurbished by November 20, 2012. Customers were requested to send back the Medical Device Recall Return Response to confirm receipt of the notification letter.For questions regarding this recall call 312-275-7070. |  | 59 | 59 beds | Nationwide Distribution including IL and IN. | Completed |  |  | 59 | Computer
Z-1472-2013 | 65093 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Versaport Bladeless Optical Fixation Cannula - 5mm |  |  | May 06, 2013 | 2013 | Covidien LP | Seals may disengage from the cannula which may result in a component inadvertently disengaging into the patients cavity | Device Design | Covidien sent an Urgent Medical Device Recall letter dated May 6, 2013,via federal express to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use of the product and return all inventory to Covidien Surgical Solutions, 195 McDermott Road, North Haven, CT 06473. Customers were requested to return completed Response Form.Customers with questions were instructed to contact their Covidien representative or Covidien Customer Service, Monday through Friday, Sam - 6:30pm ET, at (800) 962-9888, option 1, and then option 2.For questions regarding this recall call 203-492-8165. |  | 57461 | 57461 untis | Worldwide Distribution - USA (nationwide) and internationally to the following countries:CanadaAustriaBahrainBenoniCenturionChatsworthCzech RepublicDenmarkEgyptFinlandFranceGermanyIrelandIsraelItalyJapanKuwaitNetherlandsNorwayPortugalQATARSaudi ArabiaSerbiaSingaporeSouth AfricaSpainOMANSwedenSwitzerlandUAEUnited KingdomVenezuela | Terminated | August 12, 2014 | 463 | 218221 | Computer
Z-1476-2013 | 64550 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | SlingBar Standard. |  |  | April 30, 2013 | 2013 | Hill-Rom, Inc. | Retrospective review found 4 sling bars that may present a potential for impalement if a patient, caregiver or bystander were to fall onto the hooks based on a simulated test method. In effort to eliminate any risk of injury, Hill-Romis requesting products be removed from field use andreplaced with an updated design. | Device Design | On 4/30/13 an URGENT MEDICAL DEVICE RECALL CORRECTION was issued to all consignees detailing the hazards and actions that are needed. Included with this notification is a Response Form/Receipt and a Replacement Guide, complete and return the form to Hill-Rom as soon as possible. Hill-Rom will send you the applicable number of sling bars to you at no cost. After you receive the new sling bars, we request that you discard the old ones. |  | 726 | 726 | Worldwide Distribution-USA (nationwide) including the states of AZ, CA, CO, FL, IL, IN, KS, LA, MA, MI, MN, MO, NC, NH, NJ, NV, OH, OK, PA, SC, TN, UT, WA, and WI, and the countries of Canada, Australia, Austria, Belgium, Switzerland, Germany Denmark, Spain, Finland, France, United Kingdom, Greece, Hong Kong, Ireland, Israel, Italy, Japan, Netherlands, Norway, New Zealand, Portugal, Sweden, and Turkey. | Terminated | November 28, 2014 | 577 | 4710 | Computer
Z-1156-2013 | 64584 | No | MNI | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Orthopedic Alliance Spine System (alias "Blue &amp; Gold") |  |  | March 07, 2013 | 2013 | Orthopedic Alliance LLC | Orthopedic Alliance is recalling Blue &amp; Gold Implants and Instruments because the Blue &amp; Gold system is not supported by adequate testing and documentation to demonstrate that it meets performance or safety standards. These inadequacies might result in product peformance failures that could cause patient harm due to implant breakage, loosening, or inadequate sterilization. | Device Design | A recall letter dated March 7, 2013 was sent to one customer who purchased the Blue &amp; Gold Implants and instruments. The letter informed the customer that the Blue &amp; Gold system is not supported by adequate testing and documentation to demonstrate that it meets performance or safety standards. These inadequacies might result in product performance failures that could cause patient harm due to implant breakage, loosening, or inadequate sterilization. Customer is informed of the actions to be taken. Customer is instructed to complete the customer response form and return it by FAX to (858) 764-9739. |  | 3300 | 3,300 units | Distributed in Texas. | Terminated | March 26, 2014 | 384 | 3300 | Computer
Z-1446-2013 | 64968 | cassettes, tissue | IDZ | Tissue processing equipment. | Pathology | Pathology | 510(K) Exempt | LPC Fine Mesh Cassette for the Leica LPC Printer. Part No: 14060546848; 14060546846; 14060546845; 14060546933; 14060546932; 14060546852; 14060546850; 14060546847. The intended use of this device is for storage of specimens/tissues during processing. |  |  | April 04, 2013 | 2013 | Leica Microsystems, Inc. | An LPC Fine Mesh Cassette failed to stay closed during processing operations. | Device Design | A Medical Device Correction letter dated April 17, 2013 was sent to customers via USPS 1st class mail. The product issue was described and recommended actions were provided. Customers were requested to acknowledge receipt of the letter by fax-back. Customer questions were directed to (847) 405-5413. |  | 20160 | 20,160 units | Worldwide distribution: US (nationwide) to states of AZ and KY; and the country of New Zealand. | Terminated | July 21, 2017 | 1569 | 20160 | Computer
Z-1049-2013 | 64463 | No | CEO | Phosphorus (inorganic) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens Dimension(R) Phosphorus (PHOS) Flex(R) reagent cartridge (DF61) |  |  | February 22, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has initiated a voluntary field action for Dimension(R) PHOS Flex(R) Reagent due to reagent interaction causing falsely depressed Creatinine results when EZCR is processed from open wells that are in close proximity to open wells of PHOS reagent. | Device Design | Siemens sent an Urgent Medical Device Correction letter dated February 22 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Action to be taken by Customer:1. Customers with two or more Dimension systems should perform testing of PHOS and EZCR on separate Dimension systems.2. Customers with only one Dimension system should discontinue use of either PHOS or EZCR on the Dimension system. Please contact your local sales representative to discuss alternate solutions that will best meet the needs of your laboratory.Please complete the attached Effectiveness Check form and fax it to the Siemens Technical Solutions Center at (302) 631-8467, to indicate you have received the information. If you have any questions, please contact the Siemens Technical Solutions enter or your local Siemens technical support representative. We apologize for the disruption this situation will cause your laboratory. We are working diligently to resolve this issue with a permanent solution.. |  | 60674 | 60,674 | Nationwide Distribution including the states of AK, AL, AR, AZ, CA, CO, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WV. | Terminated | January 29, 2015 | 706 | 60674 | Computer
Z-1870-2013 | 65033 | No | FZY | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | S-LIFT Slap Hammer, Part Number: 11-50204, Revision A |  |  | April 25, 2013 | 2013 | SpineFrontier, Inc. | Malfunction of slap hammer. | Device Design | The firm, SpineFrontier, sent an advisory notice entitled "URGENT April 25, 2013 S-LIFT Slap Hammer Removal" dated April 25, 2013 via email to its customers. The notice described the product, problem and action to be taken. The notice requested the removal of all distributed SLIFT Slap Hammer with listed lot number. The customers were instructed to complete and return the URGENT S-LIFT Slap Hammer Removal ACKNOWLEDGEMENT FORM via fax to: SpineFrontier, Inc., 978-232-3991 or email to customerservice@spinefrontier.com.Please feel free to contact SpineFrontier at 978-232-3990 with any questions. |  | 3 | 3 | US distributions to states of: KS and TX. | Terminated | January 09, 2014 | 259 | 3 | Computer
Z-0877-2013 | 63957 | No | KMI | Bed-patient monitor. | General Hospital | General Hospital | 510(K) Exempt | QuietCare-Networked product, Facility Server/model: QC101601; |  |  | December 16, 2012 | 2012 | Intel-GE Care Innovations LLC | Care Innovations detected a software anomaly: During the engineering investigation into the cause of why a Wander Alert was not generated, a defect in the implementation of the Wander and Night Motion ADLs was discovered. | Software design | Care Innovations sent a "DEVICE CORRECTION NOTIFICATION" letter dated 12/20/12 to all affected customers. The letter identified the product and problem. Correction was performed remotely by the firm. Contact the firm at 917-232-4775 for questions regarding this notice. |  | 19 | 19 units | Nationwide Distribution. | Terminated | February 26, 2013 | 72 | 19 | Computer
Z-2257-2013 | 66126 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Healthcare, Seno Advantage, 5183685, 5183685-2, 5183685-3, 5183685-4, 5183685-5, 5183685-6, 5183685-7, 5183685-8, 5183685-9, 5342425, 5342425-2, 5342425-3, 5342425-4, 5342425-5. |  |  | September 20, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has recently become aware that patient mismatch without a warning message could occur on Seno Advantage (SA) workstations under certain circumstances. This issue may impact patient safety. | Software design | GE Healthcare sent an Important Product Information letter dated September 19, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. GE Healthcare has recently become aware that patient mismatch without a warning message could occur on Seno Advantage (SA) workstations under certain circumstances. This issue may impact patient safety.The Contact Information (US 800-437-1171, Japan 0120-055-919, Other countries contact local GE Healthcare Service Representative). |  | 179 | 179 | Worldwide Distribution - USA (nationwide) including the states of : AL AZ, CA, CO, CT, FL, GA, IL, IN, MD,MI, MO, NJ, NY, NC, OK, PA, TX and WI.,and the countries of : AUSTRIA, BELGIUM, CANADA, FINLAND, FRANCE, GERMANY, GREAT BRITAIN, IRAQ, JAPAN, LEBANON, NORWAY, PORTUGAL, SPAIN, SWEDEN and TAIWAN. | Completed |  |  | 179 | Computer
Z-2214-2013 | 66098 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Cypher MIS Screw System, Biomet Spine Parsippany, NJ 07054. The Cypher MIS Screw System is a non-cervical spinal fixation device. |  |  | August 13, 2013 | 2013 | Ebi, Llc | Several complaints were reported to exhibit screw head splay in the Cypher MIS Screw System. | Device Design | Biomet Spine &amp; Bone Healing Technologies sent "Urgent Medical Device Recall Notice" letters with Fax-Back Response Forms dated 8/13/2013 to distributors and OR managers via Fed Ex. Follow-up e-mails with a copy of the Recall Notice were sent to distributors on 8/13/2013. Customers with questions may contact 973-299-9300 x2162 Monday to Friday from 8:00 am to 5:00 pm EST. |  | 427 | 427 units | Nationwide Distribution (US), including the states of IN, MD, MS, NC, OH, PA, SC, and TX. | Terminated | March 11, 2015 | 575 | 427 | Computer
Z-2241-2013 | 66111 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Healthcare, Advantage Workstation, READY View Application. |  |  | March 27, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has become aware that when analyzing fMRI images using the READY View application software, incorrectfMRI activation maps may be computed and displayed. | Software design | GE provided Urgent Medical Device Correction letters to consignees in the following countries: Canada, Finland, France, Germany, Guadeloupe, Italy, Japan, Latvia, Martinique, Netherlands, Norway, Poland, Reunion, Romania, Saudi Arabia, Singapore, Spain,Sweden, Switzerland, Turkey and the United Kingdom.The letters were dated March 27, 2012 and were addressed to Hospital Administrators, Radiology Mangers and Risk Managers. The letters described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information (US 800-437-1171, Japan 0120-055-919 and for other countries contact your local GE Healthcare Service Representative).*Letters to other consignees is being planned for September 2013. |  | 318 | 318 (28 USm 290 OUS) | Worldwide Distribution-USA including the states of IL, MN, NJ, NY, and VA, and the countries of ALGERIA, AUSTRALIA, BELGIUM, BRAZIL, CANADA, DENMARK, FINLAND, FRANCE, GERMANY, GUADELOUPE, INDIA, INDONESIA, ISRAEL, ITALY, JAPAN,KOREA, LATVIA, MARTINIQUE, NETHERLANDS, NORWAY, PHILIPPINES,S POLAND, REUNION, ROMANIA, RUSSIA, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SPAIN, SWEDEN SWITZERLAND, THAILAND, TUNISIA, TURKEY, UNITED KINGDOM, and VIETNAM. | Terminated | November 21, 2014 | 969 | 318 | Computer
Z-0931-2013 | 64165 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Elekta Synergy XVI. Radiation Therapy Digital Imager. Product XVI R3.5.1, R4.2.1, and R4.5.1 |  |  | December 19, 2012 | 2012 | Elekta, Inc. | The DICOM tags do not contain sufficient information for XVI to identify the different volumes that have the same name. | Software design | Elekta sent an "IMPORTANT FIELD SAFETY NOTICE" dated December 2012 to all affected customers. The letter identifies the product, problem, and the actions to be taken by the customers. The IFSN informs the user in the TPS, to give the different volumes different structure names. After importing the patient data set from the TPS, make sure that the structure contours are correct. Users are requested to return the signed acknowledgement. Contact your local Elekta office for questions regarding this notice. |  | 164 | 164 | Nationwide Distribution-DC, Puerto Rico and the states of AL, AK, AZ, CA, CO, CT, FL, GA, IL, IA, KY, MD, MA, MI, MN, MS, MT, NE, NV, NH, NJ, NM, NY, NC, OH, OK, OR, PA, SC, TN, TX, VT, VA, WA, WV, and WI. | Open, Classified |  |  | 164 | Computer
Z-2248-2013 | 65893 | No | GKP | Coagulation instrument. | Hematology | Hematology | 510(k) | ACL TOP 700 CTS; PN 0000280050 |  |  | July 19, 2013 | 2013 | Instrumentation Laboratory Co. | Potential carryover issue that can can cause shortened APTT clotting times on the instrument. | Device Design | Instrumentation Laboratory sent an Important Product Correction letter on July 19, 2013, to all affected customers. On July 23, 2013, the same communication was circulated by the IL International Group in Milan, Italy to the affected international distributors and dealers. The interim field notification was followed by the release of corrective Software V5.1.0 and ParametersP-16.3.00 (Phase II) on September 30,2013.Consignees were advised of a potential carryover issue and a workaround was provided. Consignees were asked to fax the Mandatory Response Tracking Form to 781-861-4207 or email to ra-usa.ilww.comFor technical questions customers were instructed to contact the IL Technical Support Center at 1-800-678-0710, option 2. For product availability questions, customers were instructed to contact their IL Customer Service Representative at 1-800-955-9525, Option #2.For questions regarding this recall call 781-861-4467. |  | 185 | 185 | Worldwide Distribution - USA including AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WV. Internationally to Canada, Argentina, Australia, Austria, Belarus, Belgium, Brazil, Brunei, Chile, China, Colombia, Costa Rica, Czech Republic, Denmark, Dubai, Finland, France, Germany, Great Britian, Greece, Guatemala, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Korea, Kuwait, Luthuania, Luxembourg, Malaysia, Mexico, Netherlands, New Zealand, Norway, Panama, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Uruguay, Venezuela, and Viet Nam. | Terminated | October 27, 2014 | 465 | 4799 | Computer
Z-2253-2013 | 66168 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Discovery IGS 730 Cardiovascular X-Ray Imaging System. |  |  | June 19, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue with respect to Discovery IGS 730 Imaging Systems.A patient data management error may prevent the Discovery IGS 730 Imaging System from booting on start-up or re-booting during procedure. The boot issue may occur when a largeamount of patient data remains in the browser and is not deleted. The inability to boot the system ma | Software design | The firm, GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated June 19, 2013 to its Consignees/customers. The letter described the Safety Issue, Safely Instructions, Affected Product Detail, Product Correction and Contact Information. The consignees were instructed to please ensure that all potential users in your facility are made aware of this safety notification and the recommended actions. GE Healthcare will correct all affected systems at no cost to you. A GE Healthcare service representative will contact you to arrange for this correction.If you have any questions or concerns regarding this notification, please call : United States: 800 437 1171. For other countries, please contact your local GE Healthcare Service Representative. |  | 4 | 4 (2 US, 2 OUS) | Worldwide distribution: US (nationwide) including states of : PA and TN; countries of: FRANCE and SINGAPORE. | Terminated | October 25, 2013 | 128 | 4 | Computer
Z-2256-2013 | 66148 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, CARESCAPE Monitor B650. |  |  | July 12, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has recently become aware of a number of potential issues associated with the CARESCAPE Monitor B650.No injuries or illnesses have been reported. | Software design | GE Healthcare sent an Important Product Information letter dated July 12, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.GE Healthcare will provide a software and hardware correction at no charge once it is available. We will contact you to arrange for this correction.If you have any questions regarding this notification, please contact Technical Support (1-800- 558-7044) or your local Service Representative |  | 9466 | 9466 (3261 US, 6205 OUS) | Worldwide Distribution - USA (nationwide) including the states of : AL, AZ, AR, CA, CO, CT, FL, GA, ID, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MS, MO, NE, NV,NH, NH, NM, NY, NC, OH, OK OR, PA, PR, SC, TN, TX, VA, WA, WV, WI and WY., and the countries of : ARGENTINA, AUSTRALIA, AUSTRIA, BAHRAIN, BANGLADESH, BELGIUM, BOLIVIA, BRAZIL, BULGARIA, CANADA, CHILE, COLOMBIA, COSTA RICA. CROATIA, CZECH REPUBLIC, DENMARK, EGYPT, FINLAND, FRANCE, GERMANY, HONG KONG, HUNGARIA, ICELAND, INDIA, INDONESIA, IRAQ, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KOREA, KUWAIT, LEBANON, LITHUANIA, MALAYSIA, MOROCCO, NETHERLAND, NEW ZEALAND, NICARAGUA, NORWAY, OMAN, PAKISTAN, PANAMA, PARAGUAY, PERU, POLAND, PORTUGAL, QATAR, ROMANIA, RUSSIA, SAUDI ARABIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, SYRIA, THAILAND, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM and VENEZUELA. | Terminated | April 21, 2014 | 648 | 9466 | Computer
Z-0928-2013 | 64344 | Yes | MRM |  |  | Cardiovascular | PMA | Isoline Implantable Defibrillation Lead, Model Isoline 2CR5, 2CR6 and 2CT6. |  |  | January 28, 2013 | 2013 | Sorin Group USA, Inc. | Isoline defibrillation leads distributed may have internal insulation breach. | Device Design | Sorin C.R.M., srl sent an Urgent Medical Device Field Safety Notice dated January28, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were requested to return any unused products by certified mail with signed acknowledgment form, patients should continue to be monitored by their physician field reps during notification visit, and customers should make sure all personnel involved in the management of patients implanted with affected product are aware of the information described in the Field Safety Notice. Customers with questions should contact their local Sorin representative at 1-610-304-9728.For questions regarding this recall call 303-467-6222. |  | 1376 | 1376 Isoline leads | Worldlwide Distribution - USA (Nationwide) and worldwide to Argentina, Australia, Austria, Belgium, Brazil, Canada, Cuba, Czech Republic, Denmark, Egypt, France, Germany, Greece, Hong Kong, India, Iran, Ireland, Israel, Italy, Japan, Kuwait, Lebanon, Morocco, The Netherlands, Portugal, Russia, Saudi Arabia, Slovenia, Spain, Sweden, Switzerland, Tunisia, Ukraine, Uruguay and the United Kingdom. | Terminated | December 09, 2015 | 1045 | 1376 | Computer
Z-0915-2013 | 64440 | No | FLE | Steam sterilizer. | General Hospital | General Hospital | 510(k) | AMSCO C and AMSCO 400 Steam Sterilizers designed for sterilization of heat and moisture-stable materials used in healthcare facilities. |  |  | February 12, 2013 | 2013 | Steris Corporation | The control boards alarm set points and temperature tolerances allowed for erroneous activation of the AD (analog/digital) alarm as temperature fluctuated during operation. When this alarm occurs, the sterilizer locks up and cannot be used until serviced, causing Customer nuisance. | Software design | Steris Corporation initiated their recall on February 12, 2013. They will notify customers via onsite visits by STERIS Field Service Technicians to each affected Customer site.For further questions, please call (440) 392-7601. |  | 395 | 395 Units | Worldwide Distribution - USA including Canada and the states of AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NM, NY, OH, OK, OR, PA, SC, SD, SK, TN, TX, UT, VA, VT, WA, WI, and WV. | Terminated | December 05, 2013 | 296 | 395 | Computer
Z-2266-2013 | 65992 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE 1.5T Signa HDxt, 3.0T Signa HDxt, Signa Vibrant, HD23 Software and the following model numbers: 2395001, 2395001-2, 2395001-3, 2395001-4, 5269997-2, 5269997-4, 2377062-61, 5127452-3, 5127452, 5337894-3, 5337894-4, 5911000-4, 5911000-5, 53085735308573-2. |  |  | April 09, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue due to non-unique series identifiers embedded in image headers of certain GE MR Products.For ClariView and Screen Save from GVTK (Gems Visualization ToolKit) Viewer, the MR Product can generate image headers with non-unique series identifiers, which when exported to some PACS systems, results in the series being associated wit | Software design | Customers were sent two GE Healthcare "Urgent Medical Device Correction" letters one dated April 9, 2013, and the other May 22, 2013, each addressing one of the 2 groups of MR Products involved in the recall. The letters were addressed to Hospital Administrators / Risk managers, Radiology Department Mangers and Radiologists. The letters described the Safety Issue and provided Safety Instructions, Affected Product Details, Product Correction and Contact Information.GE Healthcare will correct all affected systems by providing a software update at no cost to you. A GE Healthcare service representative will contact you to arrange for this correction.Contact your local service representative if you have any questions concerning this notification.For further questions please call (262) 513-4122. |  | 1246 | 1,246 (521 US and 725 OUS) | Worldwide Distribution - All states in continental USA including DC except AK, AZ, RI and WY. and the countries of : AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BELARUS, BELGIUM, BRAZIL, CANADA, CHILE, CHINA, CZECH REPUBLIC, DENMARK, EGYPT, FINLAND, FRANCE, GERMANY, GREECE, HONG KONG, HUNGARY, INDIA, INDONESIA, ISRAEL, ITALY, JAPAN, KOREA, KUWAIT, LEANON, MARINIQUE, MEXICO, NETHERLANDS, NEW ZEALAND, NORWAY, POLAND, QATAR, ROMANIA, RUSSIA, SAUDI ARABIA, SERBIA, SINGAPORE,SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, LUXEMBOURG, PORTUGAL, VIETNAM and URUGUAY. | Terminated | May 05, 2014 | 391 | 1246 | Computer
Z-2271-2013 | 66211 | cart, emergency, cardiopulmonary (excluding equipment) | BZN | Cardiopulmonary emergency cart. | Anesthesiology | Anesthesiology | 510(K) Exempt | Stryker Vision Mounting Arm, for use with other Stryker Endoscopy medical device products for human patients. Model number 0240-095-200, manufactured by Stryker Endoscopy San Jose |  |  | August 30, 2013 | 2013 | Stryker Endoscopy | When positioning the display located on top of the video cart, it is possible for the user to apply excessive force and push the display up and out of the mounting arm. This can result in the display falling upon the user/patient. | Device Design | Stryker Endiscopy sent an Urgent Device Correction letter dated August 30, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to inform individuals within their organization who needed to be aware of this Medical Device Field Correction, contact Stericycle at 1-877-546-0482 or email stryker4645@stericycle.com to schedule a time for Stericycle to perform the correction, or perform the correction themselves by requesting the materials and following the correction instructions included. Customers were also instructed to report all adverse events or product quality problems to Stryker Endoscopy at 1-800-624-4422, press 3 for Customer Service. --UPDATE--On September 17, 2013, the recall was expanded to include additional customers who may have received the screw and washer included with the design change update. Notification was sent to these consignees on September 17, 2013. . |  | 14271 | 14271 units | Worldwide Distribution - USA including AL, AK, AZ, AR, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, WY and internationally to Australia, Canada, Chile, China, Argentina, Guam, Taiwan, India, Japan, Korea, Latin America, Philippines, Mexico, New Zealand, Hong Kong, Singapore, and Switzerland. | Terminated | April 30, 2014 | 243 | 14271 | Computer
Z-2275-2013 | 66051 | No | OUG | Medical device data system. | General Hospital | General Hospital | 510(K) Exempt | TEGRIS system |  |  | August 19, 2013 | 2013 | Maquet Cardiovascular Us Sales, Llc | MAQUET has identified a potential issue that might occur because of the lack of an emergency stop button on a secondary touch screen display on the TEGRIS System. | Device Design | Maquet sent an Urgent Medical Device Recall Letters/Field Action Response Forms dated August 19, 2013, via Federal Express to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were informed that a MAQUET authorized service technician would contact them to arrange an appointment to equip their system with a second emergency stop button that would be located in close proximity to the secondary touch screen. Customers were instructed to complete and return the enclosed Field Action Response Form acknowledging their receipt and understanding of the communiction by emailing a scanned copy to FieldActions@maquet.com or via fax at 973-807-1658. Customers with questions were asked to contact their MAQUET representative or MAQUET Technical Support at 888-627-8383, option 3.For questions regarding this recall call 973-709-7361. |  | 21 | 21 units | Worldwide Distribution - USA including NY and PA and internationally to Australia, Brazil, Canada, Germany, Hong Kong, Japan, Sweden. | Terminated | April 17, 2014 | 241 | 21 | Computer
Z-2276-2013 | 65996 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco RTP System |  |  | July 02, 2013 | 2013 | Elekta, Inc. | Upgrade installation script is designed to overwrite the same pouch files in the Treatment Couch Library with the original default values. | Device Design | Monaco 3.3 upgrades have been temporarily halted and Important Safety Notice (LINMON0007) has been issued to effected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to ensure that couch files are checked for releases 3.20.00 and 3.30.00 to verify that the correct relative electron density values are being used in the dose calculation. If customers have edited the sample couch files, and previously upgraded to Monaco 3.20.00 or 3.30.00, their couch files have already been overwritten. Open the couch files, make the necessary updates to the couch parameters, and use the "Save As Treatment Couch" option to save the modified sample couch files to a new, unique name.Please note that the couch files can also be edited in Focal and then the patients opened in Monaco. Comparable Focal releases would be Focal4.64.00(Monaco 3.10.00), Focal4.70.00(Monaco3.20.00) and Focal4.80.00 (Monaco 3.30.00).For questions regarding this recall call 404-993-5886. |  | 45 | 45 | Nationwide Distribution including AK, CA, CO, CT, FL, IL, IN, MI, MN, MO, NC, NJ, NY, OH, OK, OR, PA, SC, TN, TX, WA and Puerto Rico. | Terminated | May 08, 2014 | 310 | 45 | Computer
Z-2283-2013 | 65922 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | The FL23SE Bed is an AC-powered adjustable hospital bed with four built-in electric DC motors and remote controls that is intended for medical purposes. It can be operated by the patient or the caregiver to adjust the height and surface contour of the bed. The bed includes movable and latchable siderails. The FL23SE Bed is intended for use with patients for procedures, therapy, recovery in healthcare environments, and to transport patients between bays and procedural rooms. |  |  | August 19, 2013 | 2013 | Stryker Medical Division of Stryker Corporation | The nurse call cable (Part Number: QDF23-0571) on the FL23SE Bed is located right under the opening of the bed head box. If an accessory (i.e. IV pole) is inserted into this slot, the circuit directing the nurse call system can be damaged. | Device Design | Stryker sent an Urgent Medical Device Recall Notification letter dated August 19, 2013, to all affected consignees. The notification described the affected product, reason for recall, potential hazards, and recall customer actions and reply instructions. A Stryker Representative will contact the consignees to arrange a time to repair the unit(s) and eliminate the risk of the potential interference . Urgent questions or concerns are directed to Blaine Burnett (269)389-6954. Stryker Medical business hours are Monday-Friday 8 a.m-5 p.m. (ET).For questions regarding this recall call 269-324-6609. |  | 61 | 61 units | Worldwide Distribution - USA including MI, WI, VA and internationally to Canada. | Terminated | January 29, 2014 | 163 | 61 | Computer
Z-0911-2013 | 64013 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | BD Vacutainer PST Gel and Lithium Heparin (LH); 83 Units; Plus Blood Collection Tubes; Sterile; IVD; REF 367962; 4.5 mL, 13 X 100 mm; BD, Franklin Lakes, NJ USA; Made in USA. |  |  | March 05, 2012 | 2012 | Becton Dickinson &amp; Company | BD is conducting a recall of the BD Vacutainer Plus Plastic Plasma Separator Tube (PST) with Light Green/BD Homogard Closure due to a deformed top of the tube reservoir reported in two customer complaints. | Equipment maintenance | BD sent advisory letters via UPS dated 3/5/2012 to customers. The letter identified the affected product and the reason for recall. BD has taken steps to modify the existing manufacturing and assembly processes to help reduce the potential for this condition. Questions regarding this issue should be directed to BD Global Technical Services at 1-800-631-0174 or email at www.bd.com/Vascutainer/contact. |  | 225000000 | 225 million units | Nationwide Distribution | Terminated | July 20, 2015 | 1232 | 225000000 | Computer
Z-2201-2013 | 66156 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, CARESCAPE Monitor B850. |  |  | July 20, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has recently become aware of an issue associated with the CARESCAPE Monitor B850 when used with the iCentral network. The iCentral must have the alarm management software license (L-NETM-00) and the alarm limit change option enabled to experience this issue. When changing the HR, SpO2 or Invasive pressure alarm limits from the iCentral over the network, the alarm priority on th | Software design | Consignees were sent on 7/20/2011 a GE Healthcare "Important Product Information" letter. The letter was addressed to Healthcare Administrator / Risk Manager, chief of Nursing and Director of Biomedical Engineering. The letter described the Issue, Affected Product Details, Instructions, Product Correction, and Contact Information (Technical Support 1-800-558-7044). |  | 636 | 636 (53 US, 583 OUS) | Worldwide Distribution-USA (nationwide) including the states of CA, CO, CT, HI, IL , IA, KY, LA, MA, MI, NE, NY,OH, OK, TX, and WI, and the countries of UNITED KINGDOM, UNITED ARAB EMIRATES, TURKEY, TUNISIA, THAILAND, SYRIAN ARAB REPUBLIC, SWITZERLAND, SWEDEN, SPAIN, SOUTH AFRICA, SLOVAKIA, SINGAPORE, SAUDI ARABIA, RUSSIAN FEDERATION, REPUBLIC OF SERBIA, PORTUGAL, POLAND, PARAGUAY, NORWAY, NETHERLANDS, NEPAL, MEXICO, MALAYSIA, KUWAIT, REPUBLIC OF KOREA, JORDAN, JAPAN, ITALY, ISRAEL, IRELAND, IRAQ, INDONESIA, INDIA, HONG KONG, GREECE, GERMANY, FRANCE, FINLAND, ESTONIA, EGYPT, DENMARK, CZECH REPUBLIC, COSTA RICA, COLOMBIA, CHINA, CHILE, CANADA, BOSNIA AND HERZEGOVINA, BELGIUM, BAHRAIN, AUSTRIA, and AUSTRALIA. | Terminated | October 25, 2013 | 828 | 636 | Computer
Z-0935-2013 | 64291 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Optima MR450w, Discovery MR450, Discovery MR750, Signa HDx, Signa HDxt, Signa HDi, Signa Vibrant, and GE 1.5T Signa HDe MR System. |  |  | May 10, 2011 | 2011 | GE Healthcare, LLC | An error on the SR Viewer Reporting Tool in which edits can be made on a report and can be saved into incorrect reports. | Software design | GE Healthcare sent an Urgent Field Safety Notice dated May 5, 2011, to all affected customers. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information. Customers were instructed to open only one patient report at a time when editing a report in SR Viewer.US Customers did not receive a customer letter, however all US systems have been updated.For questions customers were instructed to contact their local Sales/Service Representative.For questions regarding this recall call 262-513-4122. |  | 4026 | 4026 (1672 US; 2354 OUS) | Worldwide Distribution - USA (All States in continental USA including PR except RI.). Internationall to:ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BRAZIL, BELARUS, BELGIUM, BULGARIA, BOSNIA &amp; HERZEGOVINA, CANADA, CHILE, CHINA, COLOMBIA, CROATIA, CZECH REPUBLIC, DENMARK, EGYPT, ESTONIA, FINLAND, FRANCE, FRENCH POLYNESIA, GEORGIA, GERMANY, GREECE, GUADELOUPE, HONDURAS, HUNGARY, INDIA, INDONESIA, IRAN, ISLAMIC REPUBLIC OF IRAN, IRAQ, ITALY, IRELAND, ISRAEL, JAPAN, JORDAN, KENYA, KOREA, KUWAIT, KAZAKHSTAN, LATVIA, LEBANON, LIBYA, LITHUANIA, LUXEMBOURG, MALAYSIA, MEXICO, MACEDONIA, MARTINIQUE, MOROCCO, NETHERLANDS, NORWAY, PAKISTAN, PERU, PHILIPPINES, REUNION, NEW ZEALAND, OMAN, POLAND, PORTUGAL, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SERBIA, SOUTH AFRICA, SPAIN, SUDAN, SWEDEN, SINGAPORE , SLOVAKIA, SOUTH KOREA, SWITZERLAND, SYRIA, TAIWAN, THAILAND, TURKEY, TUNISIA, UNITED ARAB EMIRATES, UNITED KINGDOM, UZBEKISTAN, VENEZUELA, URUGUAY, VIET NAAM, and YUGOSLAVIA. | Open, Classified |  |  | 4026 | Computer
Z-2207-2013 | 66204 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Osseotite 2 Certain Prevail 5mm(D)X 4.1mm (P) X X 11.5mm (L) Dental Implant |  |  | June 10, 2013 | 2013 | Biomet 3i, LLC | One of the dental implants do not have the internal hex feature. One of the dental implants do not have the internal thread feature. | Device Design | Biomet 3i sent an Urgent: Medical Device Recall letter dated June 10, 2013 to affected customer via email, facsimile, or postal mail. The letter identified the affected product, problem, and actions to be taking. Customer were instructed to check their respective inventory for the affected product and return to Biomet 3i. Biomet 3i will issue replacement product. For questions or concerns call 1-800-342-5454. |  | 215 | 215 | Worldwide Distribution - USA Nationwide in the states of MA, CT, NJ, DE, NY, FL, OK, TX, PA, DC, IL, KS, LA, MD, CA, RI, TN, MN, and SC and the countries of Canada, Japan, Australia, Italy Germany, France, Switzerland, Apain, Brazil, Uruguay, Singapore, Dominican Republic, Taiwan, Austria, Portugal, and the United Kingdom. | Terminated | November 06, 2014 | 514 | 393 | Computer
Z-2199-2013 | 66057 | No | HTZ | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | AEQUALIS Reversed II or AEQUALIS Reversed Fracture Impactor. These products are contained within the Aequalis reversed II Instrument set tray no.YKAD83 and the Aequalis reversed fracture instrument set tray no.YKAD95. The items affected are MWD21, MWD023, MWD024 and MWD025. |  |  | August 13, 2013 | 2013 | Tornier, Inc | Following the identification of the disassociation of an impactor tip during surgery, Tornier is initiating a voluntary recall of all lots of Aequalis Reversed II and Reversed Fracture impactors. Upon review of the event, Tornier determined that a redesign of the instrument was warranted. These products are contained within the Aequalis Revered II instrument set tray no YKAD83 and the Aequalis R | Device Design | Tornier sent an "Urgent Medical Device Recall" letter dated August 12, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customer were advised to cease further use and distribution of the product and to contact all end users. Also requested the completion and return of the attached questionnaire via fax to 952-236-4007, e-mail to CustomerServices@Tonier.com or using the envelope provided.We apologize for the incovenience and thank you in advance for your cooperation in this matter. If you have any questions please contact our Customer Service Department at 1-888 494-7950. |  | 2298 | 2,298 (1045 US, 1253 OUS) | .Worldwide Distribution - USA including Puerto Rico and the states of AL, AZ, CA, CO,CT, FL, GA, IL, KY, MD, MA, MI, MN, NE, NJ, NC, OH, OR, PA and TX., and the countries of CANADA, MEXICO, HONG KONG, SOUTH AFRICA, ARGENTINA, COLOMBIA, ISRAEL, AUSTRALIA, BELGIUM, FINLAND,FRANCE, GERMANY, ITALY, NETHERLANDS, SLOVENIA, SPAIN, SWITZERLAND and UNITED KINGDOM | Terminated | March 03, 2015 | 567 | 2298 | Computer
Z-2188-2013 | 64152 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Outlook ES pump is intended for use with B. Braun Medical Inc. Horizon Pump IV Sets. Used to regulate the flow of primary and secondary fluids when positive pressure is required. The infusion system is capable of delivering fluid from a negative head height (when the IV fluid container is lower than the pump), and provides clinically accepted volumetric accuracy for all standard IV fluids, including blood, lipid, and Total Parenteral Nutrition (TPN). Model number: 621-300ES and 621-400ES. To deliver a broad range of drugs with use with the B. Braun Medical Inc. Horizon Pump IV Sets. |  |  | November 14, 2012 | 2012 | B Braun Medical, Inc. | This notice is being provided in follow up to the field correction initiated by B. Braun Medical Inc. on August 26, 2011 due to an issue in which the Outlook ES Safety Infusion System may halt infusion but the "RUN" Light Emitting Diodes (LEDs) on the front display continue to advance as if the pump were infusing. The pump emits a backup alarm, but there are no visual indicators that the infusion | Software design | The firm, B Braun Medical Inc, sent an "Urgent Expanded Medical Device Field Correction" letter dated November 14, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed that pumps that have been upgraded (software version 151599 or higher), no additional action is required on their part at this time. For the pumps that have not yet been upgraded, within three weeks of the letter, customers should confirm the serial numbers using the enclosed response form or contact B. Braun Medical, Inc. Infusion Systems Customers Support at 1-800-627-7867 (800-627-PUMP) to schedule their upgrade.For questions regarding this recall call 972-245-2243. |  | 6919 | 6,919 | Nationwide distribution: USA including states of: AR, CA, FL, IL, LA, NC, NJ, NY, MD, ME, MN, SC, TN, TX, VA, WI, and WV. | Terminated | March 06, 2015 | 842 | 6919 | Computer
Z-2194-2013 | 66173 | No | OAT | Endosseous dental implant. | Dental | Dental | 510(k) | Brand Name: Vector TAS 8mm Screw. |  |  | July 12, 2013 | 2013 | Ormco/Sybronendo | Ormco Corporation is voluntarily recalling one lot of 8 MM Vector TAS Screws because some screws of this lot were machined undersized at the top of the threads may potentially cause a weakening of the screw. Use of the undersized product may result in breakage of the screw during insertion or removal when torqued, possibly requiring surgical removal of any remaining pieces. | Device Design | Ormco sent an "Urgent Medical Device Recall letter dated July 12, 2013, via USPS 1st class mail, along with a return form to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were informed of the recall, with product description, product codes, ship dates, contact information and instructions to return any affected product identified in stock, Additionally, doctors are instructed to proceedwith treatment if the screw has already been placed and to take special care not to over-torque the screw upon removal. Periodic effectiveness checks will be conducted in order to follow up with non-responders.Please fax this acknowledgement form to the the following number to confirm receipt of this notification whether or not you have any affected products. (714) 288-4603 |  | 4786 | 4,786 | Worldwide Distribution - USA (nationwide) including the countries of Canada, Australia, New Zealand, EU and ROW consignees. | Terminated | September 11, 2014 | 426 | 4786 | Computer
Z-2195-2013 | 66110 | No | DYJ | Orthodontic appliance and accessories. | Dental | Dental | 510(K) Exempt | AOA Mini RPE Screw Assembly , a component of |  |  | June 06, 2013 | 2013 | Sybron Dental Specialties | The firm initiated the recall of AOA Mini RPE Screw Assembly, because it may be missing a complete weld joining one side of the arm with the screw body. Herbst, MARA and Fixed Metal devices may be affected, because they contain an AOA Mini RPE Screw Assembly. | Device Design | AOA sent a recall notification letter titled "Urgent-Medical Device Correction", on June 6, 2013 via USPS 1st class mail. The letter identified the affected product, reason for recall, images, and actions to be taken. Customers were instructed to complete and return the attached Acknowledgement Form to AOA. For questions contact the AOA Customer Support Department at 1-800-262-2551. |  | 1091 | 1,091 | Worldwide Distribution - USA Nationwide and the countries of Australia, Canada, and EU. | Terminated | September 13, 2013 | 99 | 1091 | Computer
Z-1993-2013 | 66046 | No | MNO |  |  | Cardiovascular | PMA | Cardiogenesis Corporation - Cardiogenesis Transmyocardial Laser. |  |  | May 01, 2011 | 2011 | Cardiogenesis Corporation | Following acquisition of Cardiogenesis, we began receiving complaints of sparks, burns, or smoke incidents related to the handpieces. | Device Design | Cardiogenesis is responsible for the correction of all noncompliant products. CDRH understands that the corrective action plan (CAP), involving 2,437 Sologrip III handpieces and 61 PEARL 5.0 handpieces, consists of the following:1. Notification to purchasers of the handpiece removal and replacement procedures,2. Replace existing hospital inventory of handpieces with a redesigned handpieceCDRH approves the CAP subject to the following conditions:1. Notification of all distributors and purchasers will be successfully implemented, and2. Corrections will be made at no cost to the purchasers by October 1, 2013.For further questions please call (949) 420-1800. |  | 2498 | 2498 | US Distribution | Terminated | September 15, 2014 | 1233 | 2498 | Computer
Z-2191-2013 | 66133 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Vision One Laser System Model GA-0025020, Serial No: 10050, 10051, 10053. |  |  | August 19, 2013 | 2013 | Lumenis, Inc. | Lumenis has initiated a recall on certain models of Vision One System due to a potential for unintended laser exposure to the user. | Software design | Customer's were notified via letter on 8/19/13. Service visits to consignees were scheduled with anticipated completion on 8/26/13 to replace the control board. Consignee monitoring was performed by use of return reply verification tracking cards returned by the Lumenis service engineer after completion of CPU board replacement. |  | 3 | 3 units | Distributed in the states of NJ, PA, and IL. | Terminated | December 19, 2013 | 122 | 3 | Computer
Z-2173-2013 | 66036 | No | MAQ |  |  | Microbiology | PMA | KIT cobas 4800 HPV AMP/DET 240T / 960T US-IVD, CE-IVD |  |  | July 25, 2013 | 2013 | Roche Molecular Systems, Inc. | During an internal investigation, a formulation error that occurred was identified during the manufacture of the cobasÂż 4800 HPV Master Mix reagent that was packaged in four (4) cobasÂż 4800 AMP/DET kits, US- and CE-IVD (240 and 960 test kit configurations). Examination of the batch production record of the material revealed there was an error in the tare weight used for the carboy during the fina | Device Design | Roche Molecular Diagnostics sent an Urgent Medical Device Correction letter dated July 25, 2013, to all affected customers via UPS return receipt. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to:1. Immediately discontinue use of the product lots listed above.2. Discard affected product according to local guidelines.3. To receive replacement inventory, complete the entire attached faxback form and fax it to 1-866-918-8741.4. The faxback form must be completed and returned even if you do not have affected inventory.5. File this Urgent Medical Device Correction (UMDC) for future reference.Please contact Roche Molecular Diagnostics Technical Support, 24 hours a day, seven days a week at 1-800-526-1247, if you have questions about the information contained in this UMDC. |  | 791 | US-IVD: 350 Kits; CE-IVD: 441 KITS | Worldwide Distribution - USA (nationwide) including the countries of Austria, Australia, Belgium, Brazil, Czech Rep, Canada, Croatia, Columbia, Denmark, Dominican Republic, Germany, Ecuador, France, Japan, Greenland, Mexico, Italy, New Zealand, Netherlands, Peru, Norway, Singapore, Romania, South Africa, Spain, South Korea, Sweden, Switzerland and UK. | Terminated | August 06, 2015 | 742 | 791 | Computer
Z-2177-2013 | 64601 | No | FSO | Surgical microscope and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Surgical microscope |  |  | December 18, 2012 | 2012 | Leica Microsystems (Schweiz) Ag | The brake in the Z axis and/or brakes in the AB and C axes failed during installation. | Device Design | The firm, Leica Microsystems, sent an "URGENT MEDICAL DEVICE RECALL" letter dated December 19, 2012 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to DISCONTINUE USE until your system software is updated and to complete and return the attached Acknowledgement Form via email at LMGRA@leica-microsystems.com or fax to: 847-236-3747 or +41 71 726 32 49 within 10 working days after receipt of field safety notice. Your Leica Mycrosystems will contact you to make the necessary arrangements to complete an upgrade at no charge to the customers. If you have any further questions, contact Regulatory Affairs/Quality Assurance Manager at 847-405-5413 or +41 71 726 32 16. |  | 28 | 28 units | Worldwide distribution: USA (Nationwide) including states of: AK, CA, CO, NY, and OR; and countries of: Belarus, China, Hong Kong, India, Indonesia, Italy, Norway, South Africa, Spain, Sweden, Turkey, and Vietnam. | Open, Classified |  |  | 33 | Computer
Z-0948-2013 | 64376 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Siemens Somatom CTs equipped with software syngo CT2010A, syngo CT2010B, syngo CT2011A and CT2012B. |  |  | November 27, 2012 | 2012 | Siemens Medical Solutions USA, Inc | Sporadic hang-up malfunctions concerning 3D reconstructions during activated "Preview Image" functionality. | Software design | Siemens sent a "CUSTOMER SAFETY ADVISORY NOTICE" dated November 27, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. The correction for this problem will be available with Service Pack 1 for syngo CT2011A And Service Pack 4 for syngo CT2012B. Both Service pack are planned to be released starting February 2013. Systems still running on syngo CT2010A or syngo CT2010B are already scheduled to be upgraded to syngo CT2011A or syngo CT2012B respectively. Contact your local Siemens Application Specialist for additional information. |  | 896 | 896 | Nationwide Distribution. | Terminated | January 31, 2014 | 430 | 896 | Computer
Z-0950-2013 | 64436 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance 64 and Ingenuity Computed Tomography X-Ray Systems |  |  | February 15, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | This field change order is being released to update software and customer release notes to the affected installed base. | Software design | Philips Healthcare sent an Urgent Medical Device Correction letter dated February 13, 2013, via certified mail to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to review the information contained in the letter with all members of their staff who need to be aware of the contents of the letter. Customers were instructed to retain a copy with the equipment Instruction for Use. For questions customers were instructed to contact their local Phillips representative or local Philips Healthcare office. For North Americal and Canada contact Customer Care Solutions Center at 1-800-722-9377, option 5: Enter Site ID or follow the promptsFor questions regarding this recall call 1-800-722-9377. |  | 42 | 42 Units | Worldwide Distribution - USA including AR, AZ, IN, KS, LA, MA, MN, MS, MT, NJ, NY, NY, OH, OR, PA, TN, TX, VT, WV, WY, and worldwide including the contries of Australia, Belgium, China, France, Germany, Israel, Japan, Norway, Sweeden, and Switzerland. | Terminated | April 11, 2014 | 420 | 42 | Computer
Z-0951-2013 | 64320 | No | GJT | Factor deficiency test. | Hematology | Hematology | 510(k) | Vial Label: |  |  | January 28, 2013 | 2013 | Diagnostica Stago, Inc. | Diagnostica Stago has confirmed that sometimes the delta OD measured can be changed leading to a lowered percentage result. | Software design (manufacturing process) | Diagnostica Stago Inc. sent e-mail notifications on 1/28/2013 to all consignees (laboratory supervisor). The written field safety notice, entitled Important Field Safety Notice, with its response form was attached. The letter identified the affected product and gave a description of the defect. Customers were asked to confirm all the results obtained in the abnormal range by either of the two methods provided. Users are to complete and return the enclosed response form as soon as possible to the Stago Hotline via fax, 973-644-9348. Questions should be directed to the Stago Hotline at 800-725-0607. |  | 2901 | 2901 kits | Worldwide Distribution - USA and Canada. | Terminated | July 19, 2013 | 172 | 2901 | Computer
Z-2164-2013 | 64146 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Outlook ES pump is intended for use with B. Braun Medical Inc. Horizon Pump IV Sets. Used to regulate the flow of primary and secondary fluids when positive pressure is required. The infusion system is capable of delivering |  |  | October 31, 2012 | 2012 | B Braun Medical, Inc. | This notice is being provided in follow up to the field correction initiated by B. Braun Medical Inc. on 26 August, 2011 due to an issue in the Outlook ES Safety Infusion System where the Care Area can inadvertently be exited thereby removing DoseGuard dose limit protection. This occurs when a sequence of HOLD - HOLD key strokes are performed on the pump key panel while the user is in the Dose Dat | Software design | The firm decided to conduct a field correction to correct the issue and carried it out through routine field services. The firm was notified by FDA that this is a reportable event, so they notified their consignees via letters dated 10/31/2012. |  | 1970 | 1,970 pumps added for this expansion | Nationwide Distribution. | Terminated | February 03, 2016 | 1190 | 1970 | Computer
Z-2157-2013 | 66048 | No | IXI | Radiation therapy beam-shaping block. | Radiology | Radiology | 510(k) | DMLC IV-ERGO - Beam shaping block for Radiation Therapy. |  |  | August 15, 2013 | 2013 | Elekta, Inc. | Clinical mistreatment may occur when there are errors in the clinical workflow, or settings are incorrect. | Device Design | Important Field Safety Notices 200 01 406 030 (Changes to your system to improve the safe operation of DMLC) and 200 01 602 031 (Changes to your system to improve the safe operation of 3DLINE DMLC) letters were sent to customers on 8/15/2013 advising them of the issue and recommendations to resolve either by upgrading their system or to discontinue a clinical use of the device. |  | 31 | 31 | Nationwide distribution: AL, CT, FL, GA, IA, MS, MO, NJ, NY, ND, OH, OK, PA, RI, TX, VA and Puerto Rico. | Terminated | July 28, 2015 | 712 | 31 | Computer
Z-0953-2013 | 64465 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Optetrak Logic Comprehensive Knee Prosthesis System, Proximal Tibial Spacer and screws, size 3.5, 8mm, Catalog # 02-012-42-3508. |  |  | February 01, 2013 | 2013 | Exactech, Inc. | A dimensional mismatch was identified which can potentially allow for interference between devices. | Device Design | Exactech Inc. sent a *** Important *** Product Market Recall Notice dated January 30, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.1. Customers were instructed to immediately cease distribution or use of these products. 2. Extend this information to your accounts that may have this product in their possession.3. Verify if you have any of the subject Optetrak Logic Proxima Tibial Spacer (PTS) in the specified lots.Please complete the attached fax back form and return it to Exactech within the next 5 working days. Thank you for your prompt attention to this matter. For further questions please call (352) 377-1140. |  | 123 | 123 devices | Nationwide Distribution including the states of FL, AL, LA, PA, VA, IL, OH, NY, NC, KS, TN, NJ, CO, GA, CA, MD, MN, IN, &amp; Puerto Rico. | Terminated | September 22, 2014 | 598 | 123 | Computer
Z-2282-2013 | 64236 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Synthes Oracle Spacer System Slap Hammer, Part Number 03.809.690, Lot Number 3723847 |  |  | January 11, 2013 | 2013 | Synthes USA HQ, Inc. | This recall is being initiated in response to complaints received where it was reported that the proximal end of the Slap Hammer broke during use. | Device Design | SYNTHES sent an Urgent Notice Medical Device Recall letter dated January 11, 2013, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. The letter requested consignees examine their inventory, remove the recalled product from use and return them to Synthes. Customers were also instructed to completed the attached Verification Section at the end of the letter indicating whether or not they had any of the affected product and the number of devices found. Customers with questions were instructed to call 610-719-5450 or email Fieldaction@synthes.com.For questions regarding this recall call 610-719-5000. |  | 27 | 27 | Worldwide Distribution - USA including CO, PA, TN, VA, TN, WA, and internationally to Czech Republic. | Terminated | August 31, 2015 | 962 | 27 | Computer
Z-0913-2013 | 64000 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Beckman Coulter AU5800 Clinical Chemistry Analyzer. An automated chemistry analyzer that measures analytes such as Glucose, Magnesium, and Potassium. |  |  | November 28, 2012 | 2012 | Beckman Coulter Inc. | Due to a software data storage limitation which only occurs under certain conditions, measurements of the Reagent Blank (RB) or Calibration (CAL) sample results are not saved and the software reverts to a previously stored, acceptable RB/CAL result. | Software design | Beckman Coulter sent an "URGENT PRODUCT CORRECTION" letter dated November 28, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the Call Center Hotline at (800) 854-3633 for questions regarding the notice. |  | 350 | 28 in U.S.; 4 Canada; 318 to other foreign countries. | Worldwide Distribution-USA (nationwide) and the countries of Argentina, Australia, Austria, Brazil, China, Czech Republic, France, Germany, Hungary, India, Ireland, Israel, Italy, Korea Republic, Malaysia, Netherlands, Russian Federation, Spain, Switzerland, Taiwan, Turkey, United Kingdom, and Viet Nam. | Terminated | February 24, 2014 | 453 | 350 | Computer
Z-1849-2013 | 65759 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | GE Healthcare, LOGIQ e Ultrasound System. |  |  | November 09, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare became aware of a potential issue associated with the LOGIQ e system. Image data from one patient has the potential to be linked to the patient ID of another patient. | Software design | Starting on November 9, 2012, GE Healthcare, will plan to update the software on all of the affected GE LOGIQ e Ultrasound systems that were distributed at no cost to the customer.As an effectiveness check, we will track all service requests for the field action to completion. These service requests will record the repair of the affected units. Exceptions to completion of the service request will be recorded as well. Records of field action completion will be retained per our processes and procedures.No Customer letter was sent.Further questions please call (262) 513-4122. |  | 588 | 588 | Worldwide Distribution - USA (nationwide) including the states of AL, AK, AZ, AR, CA, CO, DE, FL, GA, ID, IL , IA, MD, MA, MI, MS, MO, NJ, NM, NY, NC, OH, PA, RI, SC, TN,TX, UT,VT, VA and WI., and the countries of ARAB EMIRATES, AUSTRALIA, AUSTRIA, BELGIUM, DENMARK, FINLAND, FRANCE, GERMANY, GREAT BRITAIN, HONG KONG, INDIA, ITALY, JAPAN, NETHERLANDS,NEW ZEALAND, NORWAY, POLAND, RUSSIA, SAUDI ARABIA, SINGAPORE,SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND and UNITED EMIRATES. | Terminated | September 15, 2014 | 675 | 588 | Computer
Z-0880-2013 | 63822 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Philips Medical Systems, Brilliance Computed Tomography (CT) X-Ray Scanners, Big Bore, Model #728243; and Brilliance Computed Tomography (CT) X-Ray Scanners, Big Bore Oncology, Model #728244. |  |  | October 24, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Software anomaly. Philips determined that the TumorLOC software program for the referenced PET/CT X-Ray scanning systems has the capability to generate Intensity Projection data sets from respiratory gated data. A circumstance was identified by the firm whereby these generated datasets are 'flipped' and incorrectly labeled left-to-right when the are saved to disk. If the flipped Intensity Project | Software design | Philips Medical issued an URGENT - Medical Device Correction letter to their customers concerning the Tumor LOC IPD issue on October 25, 2012. The letter informs the customer of the nature of this software problem and names the Brilliance and GEMINI scanning systems affected by this issue. The letter goes on to describe the hazard inherent to this software issue and provides 'Workaround procedures which the customers can use in order to help mitigate the effects of this problem. Customers with questions should contact their Customers Care Solutions Center at 1-800-722-9377, option 5.For questions regarding this recall call 1-800-722-9377. |  | 765 | 765 units | Worldwide Distribution - USA including AK, AK, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, IL, IN, KS, KY, OA MA, MD MI, MN, MO, MS, NC, D, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WV; Internationally to Canada and other countries.. | Terminated | July 23, 2015 | 1002 | 871 | Computer
Z-0013-2014 | 66296 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance CT Big Bore Oncology, Brilliance CT Big Bore Radiology and Brilliance CT 16 (Air) |  |  | August 30, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | There are artifacts appearing on the Tracker images during threshold-triggered bolus scans with protocols using a Rotation Time of 0.4 seconds. | Software design | Philips sent an "Urgent - Medical Device Correction" letter dated August 30, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. This Field Safety Notice is intended to inform you about:" what the problem is and under what circumstances it can occur" the actions that should be taken by the customer / user in order to prevent risks for patients or users" the actions planned by Philips to correct the problem If you need any further information or Support concerning this issue, please contact your local Philips representative or local Philips Healthcare office.For North America and Canada, contact the Customer Care Solutions Center (1 -800-722-9377, Option 5: Enter Site ID or follow the prompts).This notice has been reported to the appropriate Regulatory Agency.Philips apologizes for any inconveniences caused by this problem |  | 258 | 258 | Worldwide Distribution - USA (nationwide) including Puerto Rico and the states of AR, AZ, CA, CO, CT, GA, HI, IL, IN, KS, MA, MD, MI, MO, MS, NC, NH, NJ, NV, NY, OH, OK, OR,TN, TX, UT, VA, VT and WV., and the countries of Australia, Austria, Bangladesh, Belgium, Brazil, Chile, China, Denmark, Egypt, France, Germany, Honduras, India, Indonesia, Italy, Kazakhstan, Latvia, Lebanon, Libyan Arab Jamahiriya, Mexico, Netherlands, Norway, Palestine, Philippines, Poland, Russia, Saudi Arabia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, Ukraine, United Kingdom and Vietnam. | Terminated | May 14, 2014 | 257 | 258 | Computer
Z-0704-2013 | 64134 | No | JAC | Radiographic film marking system. | Radiology | Radiology | 510(K) Exempt | GE Healthcare Definium 5000/5220493-2 X-Ray System |  |  | September 26, 2012 | 2012 | GE Healthcare, LLC | When quickly switching the protocols from AEC to Fixed under the same view, if the user selection process time is less than 0.5 seconds, the User Interface (UI) may display different exposure techniques than selected, and may not display AEC setting. The exposure can still be made as commanded, however it may not be properly indicated. | Software design | GE Healthcare will remedy the issue or bring the product into compliance with each applicable Federal standard. A GE Healthcare Service Representative will update the software on the system to address the issue.Call Center phone number: 800-437-1171 (United States). For other countries, please contact your local GE Healthcare Service Representative if you have any questions or concerns regarding this notification.Please be assured that maintaining a high level of safety and quality is our highest priority. If you have any questions, please contact us immediately per the contact information above. |  | 312 | 312 installed in US | Nationwide Distribution | Terminated | December 05, 2013 | 435 | 312 | Computer
Z-0034-2014 | 64787 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Centricity Cardio Imaging 5.0 is a software only Picture Archiving and Communication System (PACS). |  |  | October 02, 2012 | 2012 | GE Healthcare It | GE Healthcare is aware of a potential safety risk where an image labeling situation exists in the Image Viewer software of the Centricity Cardio Imaging product, version 5.0 SP1. Multiple studies may be displayed side by side by selecting "Compare Study" from the Study List, or by opening prior studies in the Patient Folder. In these cases, the Viewer differentiates the studies wityh a label (Curr | Software design | GE Healthcare IT mailed an IMPORTANT PRODUCT INFORMATION letter dated October 2, 2012 to the only consignee of this product. The letter notified the customer of the issue, provided temporary instructions on how to use safely use the software until a new software version can be installed, and advises that a software correction will be made available in the near future. Customers who have any questions can contact their local GEHC-IT service representative or call GE Healthcare's Remote Online Center (ROC) Customer Care Center at 1-800-437-1171. |  | 1 | 1 unit | USA Nationwide Distribution in the state of OK | Open, Classified |  |  | 1 | Computer
Z-0033-2014 | 66312 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Philips MR Systems: Asset 0.5T, Apollo 0.5T, Infinion 1.5T, Eclipse/Polaris 1.5T, Panorama 0.6T |  |  | August 07, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | There is a potential for water to collect in the vent pipe elbow in the magnet venting system. Water may freeze, blocking the venting system. | Device Design | Philips Healthcare sent an Urgent Medical Device Correction letter via certified mail on August 8, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the information with all of their staff who need to be aware of the contents of this communication. Philips Healthcare will perform any necessary corrective actions prior to the customer's next cryogen refill. Customers with questions were instructed to contact their Philips field service representative: Technical Support Line 1-800-722-9377.For questions regarding this recall call 1-800-722-9377. |  | 17885 | Total 511 (Asset 0.5T,: 35 units Apollo 0.5T: 15 units, Infinion 1.5T: 116 units, Eclipse/Polaris 1.5T: 290 units, Panorama 0.6T: 55 units) | Worldwide Distribution - US (nationwide) including FL, CO, KY, MS, PA, NC, NJ, TX and Puerto Rico and internationally to Italy. | Terminated | March 26, 2014 | 231 | 17885 | Computer
Z-0035-2014 | 66424 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX 5.2 DB Server (running Oracle 10.1.0.4.0), medical imaging Picture Archiving and Communication System (PACS). |  |  | September 19, 2013 | 2013 | AGFA Corp. | System downtime and/or slow performance may result if software in use is an old version of IMPAX and is not at current supported levels. | Software in the Use Environment | An "URGENT FIELD SAFETY NOTICE" letter was sent via email to the consignees on 9/19/2013, Acknowledgement, via FAX-Back or email, that the information was received and understood has been requested from the consignees. The letter specifically describes the safety alert for the device, informs consignees of the problem and recommends the upgrades for IMPAX and Oracle. |  | 4 | 4 | Worldwide distrbution: US(nationwide) distribution including NY and OK, and country of Canada. | Terminated | October 24, 2013 | 35 | 4 | Computer
Z-0032-2014 | 66293 | No | LOM |  |  | Microbiology | PMA | FSE Image Upgrade Kit, Catalog # 92019C |  |  | August 20, 2013 | 2013 | Bio-Rad Laboratories | The FSE Image Upgrade kit (part 92019C) is recalled due to issue related to the export of data from the EVOLIS analyzer to Laboratory Information Systems (LIS) when using the MONOLISA Anti-HBs EIAQuantitative Determination in EVOLIS APF version 6.0. | Software change control | BioRad has sent an Urgent Product Correction letter dated August 19, 2013, to customers who had the EVOLIS FSE Upgrade Image Kit installed on the EVOLIS at their site by a Bio-Rad Laboratories Field Service Engineer (FSE).This issue impacts laboratory only if you are running the MONOLISA Anti-HBs EIA Quantitative Determination method with EVOLIS APF version 6.0 and TimeLiner Data Management Software. If you use this Quantitative Determination method with EVOLIS APF version 6.0, please contact Bio-Rad Laboratories for additional information or assistance.Bio-Rad Laboratories is currently developing and qualifying a software update CD consisting of the EVOLIS APF V6.0b and Timeliner Data Management Software to correct the issue. The software update CD will be installed by Bio-Rad Field Service Engineers at each of the customer sites wheninstalling the EVOLIS APF V6.0b.Customers should contact Bio-Rad Laboratories at 1-800-224-6723, Option2 and then Option 3 for any questions about this product correction. |  | 23 | 23 kits | US Distribution including the states of FL, IA, ME, MO, MS, NE, NY, TN, TX, and WA | Terminated | November 13, 2013 | 85 | 23 | Computer
Z-0050-2014 | 66361 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Solea 1.0 Laser Surgical instrument for use in general and plastic surgery and dermatology. |  |  | September 12, 2013 | 2013 | Convergent Dental | Specific error condition could lead to unintended laser emission during laser activation. | Software design | The following Notice of Correction letter was sent out to all customers via email on 9/12/13, and followed up with a direct phone call to confirm receipt and review the contents of the letter: Dear Dr. [Customer name],This email is to inform you of a pending notice of correction regarding your Solea system, serial #[Customer serial number].We have identified a specific error condition that could lead to unintended laser emission during laser activation. The software defect that caused the error condition has been identified and tested.Please do not use your Solea system until further notice. Resolution of this issue will require the installation of an updated release of the Solea software. A representative from Convergent Dental will be in touch with you to schedule the upgrade.Convergent Dental is committed to product quality and patient safety. If you have questions about this issue, please contact us. Software fix and testing:New software version (Release 2.0.4.0) was created by the software development group. Convergent Dental has implemented two software changes to eliminate this problem. The first is a defect fix to correct the handling of fatal errors, specifically laser communication timeout errors, so that these errors are not able to be cleared. This will prevent the user from being able to re-enter the workflow after a fatal error. The second is a defect fix to not set the laser state to disconnected when a laser timeout error condition occurs. This will ensure that the command to stop laser emission is always sent to the laser control module anytime the laser cutting state is exited due to error condition or the normal workflow event of stepping off the foot pedal.The software development team ran full Unit and Integration Tests on the new software build. The software verification activities included execution of the full verification test plan (Doc #790-00001), which has been revised to include specific defect |  | 5 | 5 | US distribution including states of: MA, CT, NY. | Terminated | November 15, 2013 | 64 | 5 | Computer
Z-2301-2013 | 66195 | No | MUE | Full-field digital mammography system. | Radiology | Radiology | 510(k) | Selenia Full Field Digital Mammography System with software versions 5.0.x and 5.1.x |  |  | September 06, 2013 | 2013 | Hologic, Inc. | Software: When the Selenia upon meeting the hard drive threshold of 4,000 images, may delete the most recently dated images | Software design | Hologic sent an Urgent Medical Device Recall letter dated September 5, 2013, by Certified Mail to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were advised how to avoid loss of images until Hologic corrects the problem. A software upgrade to customers who are connected to Hologic Connect who have already been connected remotely and customers who have already been connected via a Service Call will receive notification of the problem and the date the system was remotely corrected notified by Certified Mail.A Dealer notification issued by certified mail to inform Dealers of the need to notify their customers of the problem and to correct the identified Selenia system using Hologic Technical Bulletin TB-00072.Selenia software version 5.2 will implement a double safeguard to both guarantee a 28 day retention period, as well as ensure images are removed in a least recently used order.If there are any questions regarding this field correction, please call Hologic Customer Support at 877 371-4372. |  | 62 | 62 systems | Worldwide Distribution - USA including the states of AL, AR, CA, CO, FL, IL, KY, LA, NE, NC, NJ, NY, OH, PA, TN, TX, UT, WA and WV., and the countries of Canada, Denmark, Iran, Italy, Mexico and Taiwan. | Terminated | December 13, 2013 | 98 | 62 | Computer
Z-0052-2014 | 65812 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Acuson SC2000 volume imaging ultrasound systems at software version 3.0. |  |  | July 10, 2013 | 2013 | Siemens Medical Solutions USA, Inc. | Software issue results in inaccurate 3D image may result in ablation in wrong location. | Device Design | Siemens sent an Customer Safety Advisory Notification letter to all affected customers on July 10, 2013.The letter identified the product the problem and the action needed to be taken by the customerLetters provide steps to mitigate the risk and inform users that a software fix will be installed as soon as it is available. Siemens Service Engineers will contact all locations to arrange installation once fix is available. All personnel who use the ACUSON should be notified of the issue.Update September 17, 2013 : After discussion with CDRH, Siemens revised their letter to state that procedures using the ACUSON SC2000 system with BioSense Webster CARTO 3 equipped with CARTOSOUND Module and SoundSTAR Catheter should not be performed until the users system has been updated with the software upgrade that resolves the issue.For further questions please call (650) 694-5398. |  | 26 | 26 units | Worldwide Distribution - US (nationwide) including the states of IL, PA, AZ, OR, OH, NY, CO, TX and CA., and the countries of France, Germany, Japan, Portugal, Saudi Arabia, Taiwan and Turkey. | Terminated | November 25, 2013 | 138 | 26 | Computer
Z-0884-2013 | 63674 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips DigitalDiagnost X-Ray System with Eleva Software 2.1.3 |  |  | October 25, 2012 | 2012 | Philips Healthcare Inc. | When the operator for a wall stand view selects an "image rotation" different from default, or such image rotation is preset in the examination database, then the resulting image may be rotated in the wrong direction. When the operator manually rotates the image back, electronic side markers (if programmed) are rotated also and as a result may be placed in a wrong position inside the image. | Software design | Philips Medical Systems issued on October 25, 2012 an Urgent Field Safety Notice for the Digital Diagnost System informing consignees of the software upgrade. The UFSN described the problem, the hazard involved, the action to be taken by the consumer/user and the action planned by Philips. For questions contact your local Philips representative: Technical Support Line 1-800-722-9377. |  | 551 | 551 | Worldwide Distribution - USA Nationwide and the Foreign countries of: Austria, Canada, Belgium, China,Denmark, France, germany, Italy, Netherlands,New Zealand, Norway, Philippines, Portugual, SOuth Korea, Spain, Switzerland, Thailand, and UK. | Terminated | March 31, 2014 | 522 | 551 | Computer
Z-0066-2014 | 66341 | No | OXO | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Brivo OEC715/785/865 C-Arm, Models: 715 Prime, 785 Essential, and 865 Plus. |  |  | September 18, 2013 | 2013 | GE Healthcare Beijing , Hangwei Medical Systems, Co. LTD | GE is recalling certain lots of GE Brivo OEC715/785/865 C-Arm due to the potential for the wires in the high voltage cable bundle to fail, resulting in loss of system functionality. | Device Design | GE Healthcare sent an Urgent Medical Device Correction letter dated September 18,2013, to all affected customers to inform them of the issue and provide instructions for immediate mitigation of the issue. Customers were instructed to ensure that all potential users in their facility were made aware of this safety notification and the recommended actions. Customers with questions were instructed to call 1-800-874-7378 option 8.For questions regarding this recall call 801-295-4952. |  | 29 | 29 | Worldwide Distribution - USA (nationwide) and Internationally to China, India, Japan, Malaysia, Australia, Ireland, France, Germany, Spain, Hungary, Belgium, and Switzerland. | Terminated | January 27, 2014 | 131 | 29 | Computer
Z-0082-2014 | 66336 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | MAQUET Getinge Group CARDIOHELP Base Unit REF 70104.7999 REF 70104.8012 Manufactured for: MAQUET Cardiopulmonary AG Made in Germany. |  |  | September 12, 2013 | 2013 | Maquet Medical Systems USA | It has come to the attention of MAQUET that in certain rare instances, the CARDIOHELP-i device may display an error message with audible alarm, indicating "battery defective." The alarm will occur-upon start up or when the main power supply is unplugged from the main power source. The human machine interface touch screen of the CARDIOHELP-i device may gradually become black due to the human machine interface (HMI) back light shutting down. | Software Manufacturing/Software Deployment | MAQUET issued their customers "Urgent-Medical Device Field Correction" notices/Customer Fax Response forms dated September 16, 2013 via FedEx on September 16, 2013. The notification informs the customer of the issue with the product; how to identified affected product; and the action needed to be taken by the customer. The field correction includes service on the affected product. Authorized MAQUET Service Technicians will service the affected product to correct the issue. Customers were requested to complete and fax back (or email a scanned copy) of the enclosed response form. Customers could contact their local field representative or the MAQUET Customer Service line at 1.888.627.8383 (press option 2 followed by option 2) Monday through Friday, 8am to 6pm EST.For questions regarding this recall call 973-709-7000. |  | 63 | 63 | Nationwide Distribution including AZ, FL, CO, TX, OH, VA, MN, WI, SC, IL, ND, UT, CA, OR, NY, IA, CT, SD, MI, HI, AL, MD, KS, MO | Terminated | June 07, 2017 | 1364 | 63 | Computer
Z-0067-2014 | 66205 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Ingenuity TF PET/CT |  |  | August 16, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | Issues with the Ingenuity TF PET/CT Software Version 4.0 could pose a risk for patients in the event that a patient rescan were to require reinjection of PET radiopharmaceutical. | Software design | Philips Healthcare sent an Field Safety Notice dated August 16. 3013, via certified mail to all US customers. Market Representative will distribute the letter internationally. The letter identified the product the problem and the action needed to be taken by the customer.Field Safety Notice 88200460 is intended to inform you about;" what the problem is and under what circumstances it can occur" the actions that should be taken by the customer / user in order to prevent risks for patients or users" the actions planned by Philips to correct the problemAdditionally, a Field Service Engineer will contact each site to schedule a time to implement the Field Change Order.If you need any further information or support concerning these issues, please contact your local Philips representative or local Philips Healthcare office.For North America and Canada contact the Customer Care Solutions Center (1-800.722.9377, option 5: Enter Site ID or follow the prompts). |  | 13 | 13 | Worldwide Distribution - US (nationwide) including the states of NY, OH and Pa, and the countries of Australia, Austria, Egypt, France, Germany, India and Netherlands. | Terminated | March 26, 2014 | 222 | 13 | Computer
Z-0089-2014 | 66472 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | BrightView XCT is a gamma camera for Single Photon Emission Computed Tomography (SPECT), one unit per box. |  |  | September 03, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | Philips received reports from the field that the Flat Panel Detector (FPD) does not always remain securely latched in the stowed position. | Device Design | Philips Healthcare sent an Urgent - Medical Device Correction letter (Field Safety Notice (FSN 88200462), dated August 30, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the information contained in the notice with all members of their staff who need to be aware of the contents of the letter. Customers were asked to retain a copy with the equipment instruction for use. Additionally, a Field Service Engineer will contact each site to schedule time to implement the Field Change Order. A Field Service Engineer will visit the site to install the corrected software. The firm will be tracking the progress.Customers with questions were instructed to contact their local Philips representataive or local Philips Healthcare office. For North America and Canada contact the Customers Care Solutions Center at 1-800-722-9377, option 5. |  | 371 | 371 units | Worldwide Distribution - USA including AZ, CA, CO, DC, DE, FL, GA, HI, IA, IL, KY, MA, MD, MN, MO, MS, MT, NC, NJ, ND, NE, NV, NY, OH, OK, OR, PA, SD, TX, VT, WA, and WI.Internationally to Canada, Algeria, Argentina, Australia, Austria, Belgium, Brazil, China, Czech Republic, Denmark, Finland, France, Gabon, Germany, India, Indonesia, Ireland, Italy, Japan, Korea Republic, Lebanon, Malaysia, Mexico, Netherlands, New Zealand, Norway, Panama, Poland, Portugal, Russia, Saudi Arabia, Slovakia, Spain, Sweden, Switzerland, Taiwan, Turkey, United Arab Emirates, and United Kingdom. | Terminated | August 18, 2014 | 349 | 371 | Computer
Z-0893-2013 | 64313 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker XIA 3 Iliac Screwdriver, 3-PCS. |  |  | December 27, 2012 | 2012 | Stryker Spine | Surgeons have experienced unthreading of the screwdrivers outer shafts during use. | Device Design | Stryker Spine sent a "Urgent Medical Device Removal" letter dated January 14 2013, via FedEx overnight to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to examine your inventory and hospital inventory locations to identify your product. Please fax a copy of the Customer Response form to (855) 632-9049 to Regulatory Compliance. For further questions please call (201) 760-8298. |  | 46 | 46 units US | Nationwide Distribution including the states of AL, AZ, CA, CO, FL, GA, IA, IL,IN, KS, LA, MA, MD, MI, MO, MN, NC, NY. OH, OK, OR, PA, TN, TX, UT and WI. | Terminated | September 30, 2015 | 1007 | 46 | Computer
Z-0022-2014 | 66268 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | Harvest Graft Delivery System; Model Number: GDP-10 |  |  | September 11, 2013 | 2013 | Harvest Technologies Corporation | Potential for leur connector to leak or have cracks. | Device Design | Harvest Technologies sent a notification letter dated September 10, 2013, to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete the attached form to confirm the number of kits received, used, and intending to return. They are asked to return the kits with the form. For questions contact Michelle Sullivan at 508-732-7500. |  | 730 | 730 cases (3 kits per case) | Worldwide Distribution - USA Nationwide: TX, NJ, MI, AZ, GA, LA, ME, MD, AR, MO, VA, MN, GA, MA, NY, PA, CO, CA and in the countries of Australia, Brazil, China, Singapore, Vietnam, Dominican Republic. | Terminated | February 10, 2014 | 152 | 730 | Computer
Z-0894-2013 | 61554 | No | MHY |  |  | Neurology | PMA | Advanced Neuromodulation Systems, Brio DBS System implantable neurostimulator, IPG or R-IPG, Model 6788 Implantable Pulse Generator. |  |  | March 28, 2012 | 2012 | Advanced Neuromodulation Systems Inc. | Unintended Fluid ingress into the device header has been shown. | Device Design | St. Jude Medical sent an Important Product Field Corrective Action Device Information letter dated March 28, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their inventory for any unused product, isolate the affected product for return to St. Jude Medical. Their local representatiave would collect the unused inventory. Any unused product would be replaced with product once inventory becomes available. For patients with implanted product and no degradation of therapy, it is recommended dthat the devices not be unnecesssarily explanted. For patient with implanted product who experience diminished or loss of therapy customers were advised to follow the steps provided in the letter. Customers were asked to forward the notification to any institution in the event one or more patients or products could be potentially affected by this action. Customers wsere asked to maintain a record of thenotice along with the recommendations to ensure the effectiveness of the communication.For questions regarding this recall call 972-309-2154. |  | 388 | 388 units | Worldwide Distribution - 53 consignees out of the US. | Terminated | August 15, 2013 | 505 | 388 | Computer
Z-0896-2013 | 64265 | No | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Firebird Spinal Fixation System, Parallel Rod Connector (52-6800) and Axial Rod Connector (52-6700). |  |  | February 01, 2013 | 2013 | Orthofix, Inc | There is a possibility that the Set Screw (p/n 90-2923) sub-component of the Parallel Rod Connector Assembly (p/n 52-6800) &amp; Axial Rod Connector (p/n 52-6700) may crack during final tightening. Orthofix Inc received six (6) customer complaints related to set screw cracking between November 15, 2010 and December 2012. | Device Design | Orthofix sent an Urgent Medical Device Recall Notification dated February 1, 2013 to all affected customers. The letter identifed the affected product, problem and actions to be taken. Customers were instructed to identify and remove the affected products from inventory immediately and follow the actions listed in the enclosed Product Return Instructions. For questions or concerns contact Regulatory Affairs at 214-937-2100. |  | 565 | 565 total (287 PN 52-6800 and 278 PN 52- 6700) | Worldwidewide Distribution - USA Nationwide including Puerto Rico and the countries of: Australia, Lebanon, Germany, Spain, Sweden and Israel. | Terminated | July 01, 2013 | 150 | 565 | Computer
Z-0024-2014 | 65958 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | GE PET Discovery 610, PET Discovery 710. |  |  | March 26, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue related with the DMPR function and Screen Saved Images on your scanner which involves Optima CT660, Brivo CT385, Discovery PET/CT 610, Discovery PET/CT 710. The DMPR and Screen Save functionality may possibly generate the sameSeries UID. It has been seen that some PACS may combine series under a patient using only the Series UID | Software design (manufacturing process) | Consignees were sent a GE Healthcare "Urgent Medical Device Correction" letter dated March 26, 2013. The letter was addressed to Director of Radiology, Emergency Medicine Physicians and Radiologic Technologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. |  | 28 | 28 (4 US &amp; 24OUS) | Worldwide Distribution-USA (nationwide) including the states of AZ, CA, CO, FL, GA, IL, IA, IS, LA, MA, MO, NH, NM, NY, OH, OK PA, SC, TX, UT, VA, WA, and WI, and the countries of UNITED KINGDOM, DENMARK, CHILE, ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BELGIUM, CANADA, CZECH REPUBLIC, ECUADOR, EGYPT, ETHIOPIA, FINLAND, FRANCE, FRANCE GUIANA, GERMANY, GUADELOUPE, GUATEMALA, GUYANA, INDIA, INDONESIA, IRAQ, ITALY, JAPAN, JORDAN, KAZAHSTAN, KOREA, KUWAIT, LATVIA, MYANMA, NETHERLANDS, NEW ZEALAND, NORWAY, PAKISTAN, POLAND, ROMANIA, RUSSIA, SAUDI ARABIA, SOUTH AFRICA, SPAIN, SWEDEN SWITZERLAND, TAIWAN, THAILAND, TURKEY, UKRAINE, UZBEKISTAN, and VIETNAM. | Terminated | December 05, 2013 | 254 | 408 | Computer
Z-0897-2013 | 64403 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Ingenuity TF PET/CT. Diagnostic imaging system |  |  | January 21, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | The system may not default to the same gating trigger with which the data were acquired. This could cause a risk for patients. | Software design | Philips Medical Systems (Cleveland), Inc. sent an "URGENT-MEDICAL DEVICE CORRECTION" letter dated January 21, 2013 to the affected customer. The letter identified the product, problem, and actions to be taken by the customer. Contact your local Philips representative or Philips Healthcare office for further information concerning this issue. |  | 1 | 1 Unit | Distributed only in Ohio. | Terminated | July 19, 2013 | 179 | 1 | Computer
Z-0011-2014 | 65898 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | ADVIA 2120/2120i Hematology Systems with ADVIA Autoslide connected to it, SMN 10361162 and SMN 10361798 --- CLASSIFICATION NAME: Counter, Differential Cell, Automated (Particle Counter) |  |  | August 12, 2013 | 2013 | Siemens Healthcare Diagnostics | Siemens determined that if an autosampler rack jam error occurs during operation of an ADVIA 2120/2120i connected to an ADVIA Autoslide, it is possible that the next slide processed by the Autoslide could be labeled with the wrong sample identification information. | Software design | Siemens sent an Urgent Medical Device Correction Letter dated August 8, 2013, and Effectiveness Checks Forms were sent to customers on August 12, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were advised if they do not have an ADVIA Autoslide connected to their system, no action is required. To avoid the possibility of mislabeling a slide when an ADVIA 2120/2120i rack jam error occurs: Eject the sample rack. Troubleshoot and correct the reason for the rack jam error. Rerun the sample rack. Field service personnel were sent a support bulletin describing the issue and instructing them how to deal with customer questions.If you need additional assistance, please contact Global Product SupportTarrytown, New York @ osggps.hschematology.healthcare@siemens.com |  | 311 | 311 units total (US: 51 units; Foreign: 260 units) | Worldwide Distribution - USA (nationwide) and the countries of Australia, Austria, Belgium, Canada, Croatia, Denmark, France, Germany, Greece, Italy, Japan, Malaysia, Netherlands, Norway, Philippines, Portugal, South Korea, Singapore, Slovenia, Spain, Switzerland, Thailand, Turkey, United Kingdom and Vietnam. | Terminated | November 18, 2015 | 828 | 311 | Computer
Z-0009-2014 | 66381 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | GE Healthcare, Aisys, Amingo, Avance, and Avance CS2. |  |  | September 16, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue involving the Avance, Avance CS2, and Aisys anesthesia devices.While operating in Pressure Control Ventilation  Volume Guarantee (PCV-VG) mode, the affected products may produce an over delivery tidal volume when given a unique sequence of user inputs and a collapsed bellows (partially filled circuit).Exposure to excessiv | Software design | Consignees were sent a GE Healthcare "Urgent Medical Device Correction" letter dated September 17, 2013. The letter was addressed to Chief of Anesthesia, Health Care Administrator / Risk Manager, and Director of Biomedical / Clinical Engineering. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information (US call 1 800-345-2700, for other countries contact your local GE Healthcare Service Representative). |  | 10263 | 10,263 (3,194 US, 7,069 OUS) | Worldwide Distribution-USA (nationwide) including DC and Puerto Rico except DE, ME, NM, RI, and VT, and the countries of UNITED ARAB EMIRATES, JAPAN, KAZAKHSTAN, KENYA, REPUBLIC OF KOREA, KUWAIT, LATVIA, LEBANON, LIBYAN ARAB JAMAHIRIYA, LITHUANIA, MACEDONIA , MALAYSIA, MEXICO, MOLDOVA, MOROCCO, NAMIBIA, NETHERLANDS, NEW ZEALAND, NORWAY, PAKISTAN, PANAMA, PARAGUAY, PERU, PHILIPPINES, POLAND, PORTUGAL, QATAR, ROMANIA, RUSSIAN FEDERATION,SAUDI ARABIA, SERBIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, SYRIAN ARAB REPUBLIC, TAIWAN, TANZANIA, THAILAND, TURKEY, UNITED KINGDOM, VIETNAM, URUGUAY, and VENEZUELA. | Terminated | October 01, 2015 | 745 | 10263 | Computer
Z-0008-2014 | 66025 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | Draeger Fabius GS Premium, Fabius OS, Fabius Tiro, Fabius Tiro D-M |  |  | August 14, 2013 | 2013 | Draeger Medical, Inc. | During final testing of the Draeger Fabius GS Premium, Fabius OS, and Fabius Tiro anesthesia machines, some units did not pass the high voltage test and the minimum clearance between an electrical component and the unit housing was not maintained for some from a specific batch. | Device Design | The firm, Drager, sent an "Urgent Medical Device Recall" letter dated August 2013 to its customers. The letter described the product, problem and actions to be taken. The letter informed the customers that in the event of a failure, manual ventilation mode should be used and that Draeger reps would be contacting customers to schedule the replacement of the power supply. The customers were instructed to notify personnel accordingly.If you have any questions, contact Director, Quality at 1-800-543-5047 (press 1 at the prompt, then 2, then 32349). For questions regarding the operation and/or servicing of your Draeger anesthesia machine please contact DragerService Technical Support at 1-800-543-5047 (press 4 at the prompt). |  | 1127 | 1127 units | Worldwide distribution: US (nationwide) and countries of: Argentina, Australia, Austria, Azerbaijan, Bangladesh, Brazil, Bulgaria, Canada, Chile, China, Columbia, Ecuador, Egypt, France, Germany, Greece, Hungary, India, Indonesia, Israel, Italy, Ivory Coast, Japan, Kazakhstan, Latvia, Libya, LithuaniaMalaysia, Morocco, Myanmar, Nepal, Philippines, Poland, Rumania, Russian Fed., Saudi Arabia, Senegal, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Turkey, United Kingdom, Uruguay, Vietnam, White RussiaTurkmenistan. | Terminated | September 23, 2014 | 405 | 1127 | Computer
Z-0007-2014 | 66374 | Yes | NIP |  |  | Cardiovascular | PMA | Bard LifeStent Solo Vascular Stent, Item Numbers: EX062001CL, EX072001CL, EX062003CL, EX072003CL. |  |  | September 30, 2013 | 2013 | Bard Peripheral Vascular Inc | Bard Peripheral Vascular is recalling specific lots of Bard Lifestent Solo Vascular Stent because they have determined a higher incidence rate of partial stent deployment associated with their use. | Device Design | The firm, Bard Peripheral Vascular, sent a "Urgent: Medical Device Recall Notification" letter dated September 30, 2013 via FedEx to all the customers who received the specific lots on recall. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to fill out the Recall and Effectiveness Check Form and fax to: BPV at 1-800-994-6772, even if you no longer have possession of the recalled product; call Bard Peripheral Vascular's Recall Coordinator at (800) 321-4254 Option #2 Ext 2727 (M-F 7am to 4pm MST) or email silvia.carrillo@crbard.com. The Recall Coordinator will issue you either a Return Authorization (RCL) Number or Consignment Recall Number (CRC) to facilitate the expedient return of the product; and to return the product to Bard Peripheral Vascular, Inc.,1415 W. 3rd Street,Tempe, AZ 85281. BPV will provide replacement product for your returned product.If you have any questions, call 1-800-321-4254 Option #2 Ext. 2727. |  | 4112 | 4112 units | US Nationwide Distribution. | Terminated | February 05, 2014 | 128 | 4112 | Computer
Z-0900-2013 | 63671 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips Digital Diagnost |  |  | October 30, 2012 | 2012 | Philips Healthcare Inc. | When the operator for a wall stand view selects an "image rotation" different from default, or such image rotation is preset in the examination database, then the resulting image may be rotated in the wrong direction. When the operator manually rotates the image back, electronic side markers (if programmed) are rotated also and as a result may be placed in a wrong position inside the image. | Software design | Philips Healthcare issued an Urgent Field Safety Notice dated October 25, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to review the information contained in the Notice with all member of their staff who need to be aware of the contents. Also, customers were asked to retain a copy with the equipment Instruction for Use. For further information or questions customers were instructed to contact their local Phillips representative Technical Support Line at 1-800-722-9377.For questions regarding this recall call 978-659-4519. |  | 360 | 360 | Worldwide Distribution - USA (nationwide) and Internationally to Argentina, Australia, Austria, Canada, Chile, Czech republic, Estonia,France, Germany, Hungary, Inida, Ireland, Italy, Japan, Libyan Arab Jamahari, Mexico, Netherlands, New Zealand, Norway , Oman, Panama, Peru, Poland, Qatar, Reunion, Russia, Saudi Arabia, South Africa, South Korea,Spain, Sweden, Swizerland, Thailand, UAE, and the UK. | Terminated | July 29, 2013 | 272 | 360 | Computer
Z-0901-2013 | 64071 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Volume Viewer Plus 5.x (Advantage Workstation 4.2) Volume Viewer 3.x (Advantage Workstation 4.1) Volume Analysis (Advantage Workstation 4.0) |  |  | May 25, 2012 | 2012 | GE Healthcare, LLC | Incorrect views and measurements possible. GE Healthcare became aware of multiple potential safety issues involving the Advantage workstation, Volume Viewer Plus and Volume Analysis. | Software design | GE Healthcare sent a "Urgent Medical Device Correction " letter dated May 23, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to discontinue use of the Edit Contour feature until this software functionality is corrected or replaced.GE Healthcare will correct all affected systems by providing a software upgrade at no cost to you. A GE Healthcare service representative will contact you to arrange for this correction. If you have any questions or concerns regarding this notification, please call one of the following phone numbers:United States: 800 437 1171Japan: 0120 - 055 - 919For other countries, please contact your local GE Healthcare Service Representative. Please be assured that maintaining a high level of safety and quality is our highest priority. If you have any questions, please contact us immediately per the contact information above. |  | 363 | 363 | Worldwide Distribution - USA including the states of AR, CA, CO, FL, GA, IL IN, IA, LA, MD, MA, MI, MN, MO, NJ, NM, NY, NC, OH, OK, PA, PR, SC, SD, TN, TX and VA. and the countries of ALGERIA, AUSTRALIA, AUSTRIA, ARGENTINA, BELARUS, CANADA, CHINA, CHILE, EGYPT CZECH REPUBLIC, FRANCE, FINLAND, GERMANY, GABON, HONG KONG, INDIA, IRELAND, ITALY KENYA, KOREA, JAPAN, MEXICO, NIGERIA, PHILIPPINES, PORTUGAL REPUBLIC OF MALTA, RUSSIAN FEDERATION SAUDI ARABIA, SPAIN SWITZERLAND, TURKEY, TAIWAN, TUNISIA, UNTIED KINGDOM and VENEZUELA. | Terminated | January 28, 2014 | 613 | 363 | Computer
Z-0905-2013 | 64195 | No | DZL | Intraosseous fixation screw or wire. | Dental | Dental | 510(k) | Synthes Multi-Vector Distractor - Multi-axial Mandible Distractor (Left), Part Number 487.967. |  |  | January 11, 2013 | 2013 | Synthes USA HQ, Inc. | It was discovered that the sheer pin on the Multi-Vector Distractor may not be made according to product specifications. | Device Design | The firm, Synthes (USA) sent an "URGENT NOTICE: MEDICAL DEVICE RECALL" letter dated January 11, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firmat 610-719-5450 for questions regarding this notice. |  | 18 | 18 | Nationwide Distribution-including the states of CO, LA, MD, MN, NY, PA, RI, TN, TX and UT. | Terminated | August 31, 2015 | 962 | 88 | Computer
Z-2318-2013 | 66301 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | McKesson Horizon Medical Imaging (HMI) Radiological Image Processing System, product versions 4.6.1 to 11.6., is a medical image and information management application intended to receive, transmit, store, retrieve, display, print and process digital medical images, digital medical video, and associated medical information from medical imaging systems. |  |  | September 04, 2013 | 2013 | Mckesson Information Solutions LLC | Software malfunction may occur resulting in a certain combination of server configurations which has the potential to cause image loss. | Software change control | Consignees were notified via Urgent Field Safety Notice dated September 4. MMIC Customer support will be contacting customers known to have a multi-server configuration and will implement the configuration change. |  | 329 | 329 units | Worldwide Distribution - United States including the territories of Guam and Puerto Rico and the countries of Canada and France. | Open, Classified |  |  | 329 | Computer
Z-0892-2013 | 64136 | No | GYC | Cortical electrode. | Neurology | Neurology | 510(k) | Ad-Tech Medical Instrument Corporation Macro Micro Subdural Electrode. For temporary (&lt;30 days) use with recording, monitoring and stimulation equipment, for the recording, monitoring and stimulation of electrical signals on the surface level of the brain. The recording of electrical activity supports definition of the location of epileptogenic foci and brain mapping. |  |  | December 18, 2012 | 2012 | Ad-Tech Medical Instrument Corporation | The microelectrodes are not flush with the silastic substrate, may contain sharp edges capable of abrading the cortical surface, and may have remaining pieces which may fracture and remain in the cortical tissue with the electrode is explanted. Recall expanded on 5/2/2013 to include additional Macro Micro Subdural Electrodes. | Device Design | AD-TECH Medical Instrument Corporation sent an "Urgent Medical Device Recall" letter dated December 18, 2012 to all affected consignees. The letter described the Reason for the Voluntary Recall, Risk to Health, Actions to be taken by the Customer, Product and Distribution Information, Type of Action by Ad-Tech Medical Instrument Corporation and Other Information. Consignees were instructed to quarantine the affected product and contact the firm for a Return Material Authorization number. If the product have been further distributed consignees were advised to notify their customers of the recall. The letter requested consignees to return the Acknowledgement and Receipt Form.For questions contact your Ad-Tech Clinical Specialist.Expanded recall letters dated May 2, 2013 were sent out April 29, 2013. The letters were formatted as the initial ones sent on 12/18/2013. |  | 144 | 144 (103 US; 41 OUS) | Worldwide Distribution - USA Nationwide including the states of CA, GA, LA, MA, MD, MI, MN, NH, PA, RI, TX, UT, WI, NY IA, WI and the countries of: Australia, CANADA, ENGLAND, JAPAN, KOREA, SPAIN, FRANCE, and SWITZERLAND. | Terminated | October 15, 2014 | 666 | 144 | Computer
Z-0907-2013 | 64214 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | SIEMENS brand ARTISTE syngo RT Therapist (RTT) version 4.3, ARTISTE Model, Part No. 08162815, Digital linear accelerators, MOSAIQ OIS system, Distributed by SIEMENS Healthcare IM CR RO, CR Radiation Oncology, Manufactured by SIEMENS AG, MEDICAL SOLUTIONS, GERMANY. |  |  | January 17, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Siemens became aware of a potential safety issue when using the "In-session resumption" feature of the syngo RT Therapist (RTT) 4.3 software. The RTT will send again the previously completed beam or segment for delivery to the control console, thus repeating the already delivered Monitor Units of the same beam or segment. | Software design | The firm issued "Urgent Medical Device Correction" letters via mail to all its direct consignees, informing them of the affected product issue and providing instructions on the recall. The firm stated that the release of this UI TH005/13/S to distribute the Customer Safety Advisory Notice began on January 17, 2013. |  | 5 | 5 units | 5 devices, all located outside of the United States in Germany. | Terminated | March 05, 2013 | 47 | 5 | Computer
Z-0906-2013 | 64259 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Philips Medical System Extended Brilliance Workspace (EBW), Computed Tomography X-Ray System. |  |  | September 28, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Incorrect values may be displayed. Philips is issuing a software update (version 4.5.5) to the Extended Brilliance Workspace (EBW) systems currently running software version(s) 4.0.2, V4.5.2 or V4.5.3, addressing a number of non-conformances. | Software design | Philips Healthcare sent a Urgent Medical Device Correction letter dated September 28, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.This Field Safety notification is to inform the actions that should be taken by the customer/user to prevent risks for patient or users, and the actions planned by Philips to correct the problem.If you need further information or supplies please contact your local Philips Healthcare office. For North American and Canada contact the Customer Care Solutions Center ( 1-800 722-9377) option 5 enter site or follow the prompts. |  | 2580 | 2,580 units | Worldwide Distribution - USA including the states of AL, AT, AZ, BAH, CA, CO, CT, DC, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NJ, NV, NY, OH, OR, PA, PR, PR, SC, TN, TX, VA, VT, WA, WI &amp; WV. and the countries of Albania, Algeria, Argentina, Australia, Austria, Bahrain, Bangladesh, Belgium, Bolivia, Brazil, Brunei Darussalam, Bulgaria, Canada, Chile, China, Columbia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican, Republic, Ecuador, Egypt, Ethiopia, Finland, France, French Polynesia, Gabon, Georgia, Germany, Greece, Guadeloupe, Guinea, Hong Kong, Hungary, Iceland, India, Indonesia, Iran, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kuwait, LAO People's Dem, Rep., Latvia, Lebanon, Libyan Arab Jamahiriya, Lithuania, Malawi, Malaysia, Malta, Martinique, Mexico, Netherlands, Netherlands, Netherlands Antilles, New Caledonia, New Zealand, Nigeria, Norway, Oman, Pakistan, Palestine, Panama, Peru, Philippines, Poland, Portugal, Qatar, Reunion, Romania, Russia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sudan, Sweden, Switzerland, Syrian Arab Republic, Taiwan, Tanzania, Thailand, Tunisia, Turkey, Uganda, Ukraine, United Arab Emirates, United Kingdom, Uzbekistan, Venezuela &amp; Viet Nam. | Terminated | December 26, 2013 | 454 | 2580 | Computer
Z-2131-2013 | 65923 | No | JTC | Microtiter diluting and dispensing device. | Microbiology | Microbiology | 510(K) Exempt | WASP (Walk Away Specimen Process). Automatic Planting and Streaking Instrument |  |  | August 09, 2013 | 2013 | Copan Italia | The recall of the WASP (Walk Away Specimen Process) instrument was initiated because Copan Italia SpA has upgraded the software concerning the broth inoculation table feature. The upgrade has been considered necessary from the manufacturer in order to eliminate the very low risk of rotation of the broth table when the door is open. | Device Design | Copan sent an Urgent WASP recall letter dated July 31, 2013, to all affected customers that received the WASP (Walk Away Specimen Process) instrument to notify them of the software upgrade concerning the broth inoculation table feature. The upgrade has been considered necessary from the manufacturer in order to eliminate the very low risk of rotation of the broth table when the door is open. This event may happen only if damage in the door locking system happens. This could result in a minor, transient injury to the user due to the impact with the rotating table.The WASP instruments involved in this upgrade are all whose serial numbers show the configuration with the broth inoculation feature activated.The machine installed in your laboratory is one of this serial number involved in the SW upgrade. The Software upgrade will be performed on site by personnel trained by Copan and no action is required by yourself.Customers will be contacted by the WASP Technical Support in order to define a date that will fit with their planning. Customers with questions are instructed to call +39 0303666130 or write to regulatory.wasp@copanitalia.com. |  | 63 | 63 units | Worldwide Distribution - USA including Oregon and California., and the countries of France, Belgium, Italy, Denmark, Norway, Switzerland, Romania, Finland, Germany, South Africa, Canada, Ireland, UK, Sweden and Japan | Terminated | December 12, 2013 | 125 | 63 | Computer
Z-2130-2013 | 65902 | No | GDJ | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Stryker Howmedica Osteonics Triathlon tibial alignment ankle clamp extramedullary (EM) referencing instrument (Cat. No. 6541-2-609), for use in Total Knee Arthroplasty (TKA) when the surgeon elects to use Stryker's Triathlon Total Knee System in conjunction with EM referencing. Non-Sterile Instrument. |  |  | July 02, 2013 | 2013 | Stryker Howmedica Osteonics Corp. | Cracks or fracture of the Triathlon Tibial Alignment Ankle Clamp, an instrument. | Device Design | Stryker Orthopaedics sent an Urgent Medical Device Correction letter dated July 2, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by t he customers. Customers are advised of the issue, potential hazards and risk mitigation factors. Customers are asked to complete the attached form and fax to 855-251-3635. Any questions should be directed to (201) 972-2100.All letters included a Product Correction Bulletin. |  | 10537 | 10,537 units | Nationwide Distribution | Open, Classified |  |  | 10537 | Computer
Z-2111-2013 | 65779 | No | MKQ |  |  | Pathology | PMA | PrepStain Tecan US-I. |  |  | July 08, 2013 | 2013 | Becton Dickinson &amp; Co. | PrepStain system rack setting was incorrectly entered on one PrepStain system during a routine preventive maintenance visit by a BD Field Service Representative. | Software design | Ex-US BD sites were notified via email and teleconference, and will communicate directly with customer sites in their region as required by local regulation. All US customers were initially contacted via telephone and letters were mailed. |  | 890 | 890 instruments | Worldwide Distribution-USA (nationwide) and Puerto Rico and the country of Canada. | Terminated | May 01, 2014 | 297 | 890 | Computer
Z-1021-2013 | 64469 | No | LJG | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | Microbiologics KWIK-QC SLIDES ( KWIK-QC Acid Fast Stain Slides and KWIK-QC Mycobacterium Slides). |  |  | February 11, 2013 | 2013 | Microbiologics Inc | Microbiologics Inc is initiating a voluntary recall on products KWIK-QC Acid Fast Stain Slide, SL42-10 and KWIK-QC Mycobacterium, SL43-10. These products are being recalled because the control organisms are not properly fixated on the slide, which may cause the user to unintentionally contaminate the patient area of the slide by flooding the control well with stain. This contamination could pote | Device Design | Microbiologists an "Urgent Medical Device Recall" letter dated February 11, 2013. The Letter described the product and the reason for the recall. Advised consignees to "Do not Distribute" and "Return" the product. Urged consignees to "Notify" end users and to complete the attached "Recall Response Form". Contact the form at 320-229-7057 for questions regarding this notice. |  | 3960 | 396 boxes of 10 slides each. | Worldwide Distribution-USA including the states of CA, IN, MS, NY, OH, PA, and SD, and the countries of AUSTRALIA, AUSTRIA, UNITED ARAB EMIRATES, MALAYSIA, TAIWAN, JAPAN, TURKEY, BELGIUM, COLOMBIA, CANADA, SOUTH AFRICA, and VENEZUELA. | Terminated | March 10, 2016 | 1123 | 3960 | Computer
Z-1927-2013 | 65444 | No | HXX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Invue Split Driver, Assembly with Handle, Model SI70006 Rev G. Used in conjunction with the Invue Anterior Cervical Plate System under 501(k) clearance number K121060. |  |  | May 10, 2013 | 2013 | SpineFrontier, Inc. | Device design. | Device Design | SpineFrontier sent an URGENT Advisory Notice dated June 27, 2013, to all affected consignees regarding the use of Invue Drivers. The notice advises users that the Split Tip Stylet Driver Assembly (11-70099) and Sold Driver, AO Connection (11-70070 or 11-70077) are being shipped in place of the Split Driver Assembly w/Handle (SI70006). The uses of the replacement parts are explained. Consignees are instructed to complete the Acknowledgement Form and return it to SpineFrontier.On June 28, 2013, customers were sent a second Advisory Notice requesting the removal of the Split Driver Assy w/ Handle Part Number SI70006 Lot Numbers 027104 and 027105. Consignees were instructed to complete the Acknowledgement Form and return it to SpineFrontier.Consignees with questions were instructed to call 978-232-3990. |  | 12 | 12 devices | Nationwide Distribution including Texas, Ohio, and Missouri. | Terminated | July 20, 2015 | 801 | 12 | Computer
Z-1943-2013 | 65133 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | SIGNA INFINITY MR SYSTEM WITH EXCITE TECHNOLOGY |  |  | November 21, 2011 | 2011 | GE Healthcare, LLC | Screen save images are included in wrong exam due to duplicate UID creation. This may cause a safety issuedue to a potential patient data mismatch.While the issue has only been observed at one customer site on the Signa HDxT 1.5T system, internal investigation revealed that the issue also impacts multiple GE healthcare MR Systems. | Software design | GE Healthcare communicated internally to their Service Team the directions for completing the correction on the customers' systems. GE Healthcare's recall strategy was to update all affected systems that were distributed. At this time,Week of 02/06/2012, deployment of this recall is considered complete.Should you have any questions or require additional information, please contact Regulatory Affairs Leader by telephone at 262-548-2608 or email paulcorrigan@ge.com. |  | 7016 | 7016 (2919 US; 4097 OUS) | Worldwide distribution: USA (nationwide) including DC and Puerto Rico and countries: ALBANIA, ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA, AZERBAIDJAN, BELARUS, BELGIUM, BOLIVIA, BOSNIA-HERZEGOVINA, BRAZIL, BULGARIA, CANADA, CHILE, CHINA, COLOMBIA, CROATIA, CYPRUS, CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, ECUADOR, EGYPT, FINLAND, FRANCE, GEORGIA, GERMANY, GREAT BRITAIN, GREECE, GUADELOUPE, GUATEMALA, GUYANA, HONDURAS, HONG KONG, HUNGARY, ICELAND, INDIA, INDONESIA, IRAN, IRELAND, ISRAEL, ITALY, JAPAN,JORDAN, KAZAKHSTAN, KENYA, KUWAIT, LATVIA, LEBANON, LIBYA, LITHUANIAN, LUXEMBOURG, MALAYSIA, MALTA, MARTINIQUE, MEXICO, MOROCCO, NETHERLANDS, NEW ZEALAND, NORWAY, PAKISTAN, PANAMA, PARAGUAY, PERU. PHILIPPINES, POLAND, PORTUGAL, REPUBLIC OF MACEDONIA, REUNION, ROMANIA, RUSSIA, SAUDI ARABIA, SINGAPORE, SLOVAK REPUBLIC, SLOVENIA, SOUTH AFRICA, SOUTH KOREA, SPAIN, SUDAN, SWEDEN, SWITZERLAND, SYRIA, TAIWAN, THAILAND, TUNISIA, TUR,EKY, UKRAINE, UNITED ARAB EMIRATES, URUGUAY, VENEZUELA, VIETNAM, YUGOSLAVIA, MALAWI, MAURITANIA, NIGERIA, UKRAINE, IVORY COST, OMAN, ZAMBIA, UZBEKISTAN, ESTONIA, IRAQ, and JAMAICA. | Terminated | October 22, 2014 | 1066 | 7016 | Computer
Z-1940-2013 | 65333 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | GE LightSpeed CT750 HD |  |  | April 05, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare conducted a recall involving 2 events affecting the GE LightSpeed CT750 HD. Event 1 - There was a report regarding CT number variance in time line at posterior fossa region of head perfusion. When beam hardening and helical artifacts in this region are reconstructed into the perfusion map, these artifacts will give subtle differences in the CT numbers.Event 2 - The software does no | Software design | GE Healthcare's recall strategy was to update CT750HD CT Systems with software 10MW25.6 on all systems that were distributed. Updates started on 4/5/2011. At this time, deployment of this recall is considered complete. The recall was completed 2/2/2012.Should you have any questions or require additional information, please contact Regulatory Affairs Leader by telephone at 262-548-2198 or email paulcorrigan@ge.com. |  | 244 | 244 | Worldwide distribution: US (nationwide) including states of: AL, AZ, CA, CT, FL, GA, HI, IL, IN, IA, KY,ME, MD,MA, MI,MN, MS, NV, NH, NJ, NY, NC, PA, TN, TX, VA, WA,WV, and WI; and countries of: AUSTRALIA, BELGIUM, BRAZIL, CANDA, CHINA, DENMARK, FINLAND, FRANCE, GERMANY, GREECE, HONG KONG, INDIA, INDONESIA, IRELAND, ISRAEL, ITALY, JAPAN,JORDAN, REPUBLIC OF KOREA, KUWAIT, LATVIA, NEW ZEALAND, POLAND, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SWEDEN, SWITZERLAND, TUKEY, and UNITED KINGDOM. | Terminated | August 11, 2014 | 1224 | 244 | Computer
Z-1944-2013 | 65583 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | GE Healthcare, LOGIQ S8 Ultrasound System |  |  | August 13, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare became aware of a potential issue related to incorrect image measurement associated with the software of the LOGIQ S8 system. This issue was found internally. Users may notice that the system shows incorrect image measurement results on Pulsed Wave (PW) or TimeIntensity Curve (TIC) or depth measurement when using Virtual Convex mode and putting a cursor (Sample Volume, Depth measur | Software design | GE initiated a field action on 8/13/12. GE plans to update the software on all of the affected GE LOGIQ S8 Ultrasound systems that were distributed, at no cost to the customer.Should you have any questions or require additional information, please contact RA Leader, US Region by telephone at 262-548-2089 or email James.Giles@med.ge.com. |  | 755 | 755 | Worldwide distribution: US (nationwide) including states of: AL, AZ, CA, CO, CT, DC, FL, GA, IL, IN, IA, LA, MA, MI, MS, MO, MT, NE, NV, NJ, NY, NC, OH, OK, PA,RI, SC, TN, TX, VA, WA, and WI; and countries of: ALGERIA, AUSTRALIA, AUSTRIA, BRAZIL, CANADA, COLOMBIA, CZECH REPUBLIC, DENMARK, ESTONIA, FINLAND, FRANCE, GERMANY, GREECE, INDIA, INDONESIA, IRELAND, ITALY, JAPAN, LATVIA, LEBANON, MALAYSIA, MOROCCO, NETHERLANDS, NEW ZEALAND, NORWAY, POLAND, PORTUGAL, REPUBLIC OF MOLDOVIA, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SOUTH KOREA, SPAIN, SWEDEN, SWITZERLAND, THAILAND, TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, and YUGOSLAVIA. | Terminated | September 15, 2014 | 763 | 755 | Computer
Z-1942-2013 | 64992 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Optima XR220amx, Optima XR200amx, and Brivo XR285amx. |  |  | August 07, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare conducted a recall on various mobile x-ray units (Optima XR220amx, Optima XR200amx, and Brivo XR285amx systems) because of an Erratic Drive Performance. | Software design | GE Healthcare recall strategy was to update all of the affected Optima XR220amx, Optima XR200amx, and Brivo XR285amx systems that were distributed. At this time ( 2/14/2013), deployment of this recall is considered complete.Should you have any questions or require additional information, please contact Regulatory Affairs Leader by telephone at 262-548-2198 or email paulcorrigan@ge.com |  | 458 | 458 (306 US; 152 OUS) | Worldwide distribution: USA (nationwide) including states of: AL, AK, AZ, AR, CA, CO, CT, DC, FL, GA, GUAM, IL, IN, IA, KY, LA, ME, MD, MA, MI, MN, MS, MO, NV, NH, NH, NM, NY, NC, ND, OH, OK, OR, PA, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, and WY; and countries of: Albania, Australia, Bahrain, Belgium, Canada, Chile, Czech Republic, Dominican Republic, France, Germany, Great Britain, Honduras, Hungary, Ireland, Israel, Italy, Japan, Korea, Kuwait, Netherlands, Nigeria, Norway, Romania, Saudi Arabia, Spain, United Arab Emirates and Venezuela. | Terminated | February 04, 2014 | 546 | 458 | Computer
Z-1025-2013 | 64075 | No | FZZ | Skin marker. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Dual Tip Marker, Sterile |  |  | August 22, 2012 | 2012 | Viscot Medical LLC | Dual Tip Marker has inconsistent ink flow on the black permanent ink side. | Device Design | Viscot Medical LLC sent a Customer Notification Letter dated August 22, 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to return the affected product for a replacement. To return the product call 1-800-221-0658 or e-mail jeffl@viscot .com for the RGA. |  | 458589 | 458,589 markers | USA Nationwide Distribution | Terminated | July 21, 2014 | 698 | 458589 | Computer
Z-1945-2013 | 65721 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Siemens AXIOM Artis zeego |  |  | May 13, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Siemens issued a customer safety notice about the overlay of system information after a functional check of the brakes on the Artis Zeego system. | Software design | The firm, Siemens, sent an "Important safety notice" by letter on May 13, 2013 to its customers. The notice described the product, problem and actions to be taken. The customers were instructed to immediately notify the service department to clarify what further measures need to be taken, if after carrying out the regular brake test, the system can only be moved at a reduced speed. If this device has been sold and is therefore no longer in your possession, send this safety notice to the new owner and provide the new owner to Siemens.Update instruction AX012/13/S was issued, and a software update will be released as AX006/13/S and AX007/13/S which will modify the system software so that it will not be possible for information to be overlaid.If you have any questions call +49 (9191) 18 0. |  | 132 | 132 | Nationwide distribution: AL, AK, AR, AZ, CA, CO, CT, DC, FL, GA, IA, IL, IN, KY, LA, MA, MD, MN, MO,MS, NC, NE, NH, NJ, NY, OH, PA, PR, SD, TN, TX, UT, VA, WA, and WI. | Terminated | November 06, 2014 | 542 | 132 | Computer
Z-2118-2013 | 49516 | No | KPA | Automated slide stainer. | Pathology | Pathology | 510(K) Exempt | Benchmark XT/LT and Discovery XT |  |  | June 12, 2008 | 2008 | Ventana Medical Systems Inc | Ventana Medical System is initiating the recall of the BenchMark and DISCOVERY series instruments because there is a possibility that the heating element could separate from the underside of the heater cap and cause a non-staining or inconsistent staining event. Additionally, the system may not detect this anomaly and therefore it may not generate an error message for the slide position. | Device Design | Ventana sent an Urgent Medical Device Correction Notification letter dated June 16, 2008 to all affected customers. The letter identifed the affected product, problem and actions to be taken. The letter instructed customers to carefully read and perform all of the actions outlined in the notification letter concerning potential safety issue. Customers with concerns regarding the staining results on any patient specimen are instructed to follow their internal Quality Procedures regarding the review of patient reports. For question contact your local Customer Service Center. |  | 1217 | 1217 units total | Worldwide Distribution - USA Nationwide and the countries of Canada, Europe, Asia, Japan. | Terminated | August 29, 2013 | 1904 | 1217 | Computer
Z-1934-2013 | 65837 | No | MMO | Ear, nose, and throat manual surgical instrument. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(K) Exempt | Tray Surgical, ENT (various instrument trays) |  |  | July 09, 2013 | 2013 | Medtronic Xomed, Inc. | The firm will recall and replace the identified instrument trays in the field, with trays that have improved labeling, including lids that contain the statement indicating 'PREVACUUM ONLY" | Device Design | Medtronic sent a Customer Commincation Recall Letter to all affected customers with the following information: Subject: URGENT PRODUCT Recall  Immediate Attention RequiredMEDTRONIC XOMED INSTRUMENT TRAYS- model numbers provided on Attachment One- all lots distributed prior to July 2013Dear Purchasing Manager, Risk Manager and Central Supply Manager/Supervisor:This letter is to inform you that Medtronic Surgical Technologies (ENT Division) is initiating a voluntary recall to remove certain instrument trays from the field and replace them with trays that have modified labeling. Medtronic Xomed Instrument Trays are intended for the storage and transportation of reusable surgical instruments. These trays are not intended and not designed to maintain sterility. We are aware, however, that some users may also choose to use the trays for holding instruments during steam sterilization. If the user chooses to sterilize instruments within these instrument trays, they must validate the sterilization process. These users should be aware of the following issue:Issue Description:During the validation of a new instrument, testing revealed that use of a GRAVITY steam cycle with one of the listed instrument trays did not produce the Sterility Assurance Level (SAL) of 10-6 recommended by industry standard (ANSI/AAMI ST77:2006). A subsequent investigation of all Medtronic Xomed instrument trays identified additional trays as having certain locations within thetray that would not meet the recommended sterilization level of SAL10-6 during GRAVITY steam sterilization cycles. The trays identified through the testing are the subject trays involved in this field action. Instruments that are sterilized in the worst-case locations of the identified trays have a higher potential to be inadequately sterilized. An instrument that is inadequately sterilized can become the source of cross-contamination or cross-infection when used in multiple surgical procedures. At this time |  | 3562 | 3,562 trays | Worldwide Distribution - USA (nationwide) Nationwide, Hawaii, Puerto Rico and Guam.Internationally to: Canada, Austria, Belgium, Cyprus, France, Germany, Greece, India, Iran, Ireland, Israel, Italy, Kuwait, Lebanon, Luxembourg, Morocco, Netherlands, Nigeria, Portugal, Qatar, Reunion, Saudi Arabia, South Africa, Spain, Sweden, Switzerland, Turkey, United Arab Emirates, United Kingdom, Australia, New Zealand, Singapore, Thailand, Vietnam, China, Hong Kong,Taiwan, Korea, Japan, Brazil, Chile, Colombia, Costa Rica, Ecuador, Guatemala, Mexico, Nicaragua, and Venezuela. | Terminated | September 25, 2014 | 443 | 3562 | Computer
Z-1949-2013 | 64304 | No | NQI | Immunomagnetic circulating cancer cell selection and enumeration system. | Immunology | Pathology | 510(k) | CELLTRACKS AUTOPREP System (Software Version 2.4.1 &amp; Below. |  |  | December 28, 2012 | 2012 | Veridex, LLC | Veridex, LLC has received a report that a patient's demographics were incorrectly associated with another patient's identification (PID) number. The purpose of this notification is to inform you of the issue and provide instructions to help prevent an occurrence on your CELLTRACKS AUTOPREP System and CELLTRACKS ANALYZER II. | Software design | Veridex sent an Important Product Correction Notification letter dated December 28, 2012, to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete and return the Confirmation of Receipt form. For questions contact Customer Technical Services at 1-800-421-3311. |  | 160 | US: 82 Units; Ex-US 78 Units. | Worldwide Distribution - USA Nationwide and the countries of Australia, Canada, China, India, Japan, Singapore, England, France, Germany, Italy and Spain. | Terminated | August 11, 2015 | 956 | 306 | Computer
Z-1954-2013 | 65502 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | REF 00-4309-071-40 Trabecular Metal Reverse Shoulder System Instrumentation, Glenosphere Helmet, 40 mm Diameter Nonsterile, Qty -1 |  |  | June 28, 2013 | 2013 | Zimmer, Inc. | There is potential for one or both tabsof theTrabecular Metal Reverse Glenosphere Helmet to fracture during use and become disassociated from the device. There have been 47 reported complaints of tab fracture. | Device Design | Zimmer Inc sent Urgent Medical Device Recall notifications on June 28, 2013, to all affected customers. The initial letter identified the issue, possible adverse consequences and informed customers that a replacement would be available in the coming months. In addition, the letter included references to both the TM Reverse Surgical Technique and to the video available on Zimmer TV and Zimmer.com. Both resources explain in detail the proper technique for utilization of the Glenosphere Helmet. Proper utilization of the Glenosphere Helmet dramatically reduces the likelihood of fracture. Distributors will receive a copy of the surgeon letter for their reference along with the list of affected surgeons. The affected surgeons will be identified using invoice data and in cases where the surgeon is unknown, the distributor will be instructed to assist with the surgeon list. In addition, information regarding the return process for the damaged devices will be included. Questions and concerns are directed to Zimmer's customer call center (1-877-946-2761).Phase 2: As replacement devices with the new design became available for distribution, a 2nd letter was sent to the distributors. Estimated replacement device availability will begin in July 2013. The distributors will be instructed to order replacement devices as a means of completing the affected instrument cases in circulation within their territories. Instructions for the return of devices were included.For questions regarding this recall call 800-613-6131. |  | 3163 | 3,163 both sizes | Worldwide Distribution - USA (nationwide) and internationally to Australia, Austria, Belgium, Brazil, Canada, China, Czech Republic, Denmark, Finland, France, Germany, Greece, India, Israel, Italy, Korea, Lebanon, Malta, Netherlands, Norway, Portugal, Singapore, Spain, Sweden, Switzerland, Taiwan, Thailand, United Kingdom, Utd.Arab.Emir. | Terminated | March 02, 2015 | 612 | 3163 | Computer
Z-1952-2013 | 65560 | No | HTO | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Synthes Flexible Medullary Reamer. |  |  | June 11, 2013 | 2013 | Synthes USA HQ, Inc. | Due to the coiled design for this product, the product is difficult to clean and the potential for corrosion on the device exists. The potential for harm exists if there is an introduction of foreign material at the operative site originating from the presence of corrosion and/or the retained undefined material that could not be adequately cleaned. | Device Design | Sythes notified direct accounts by letter on 6/11/13 and requested them to check their inventory and remove affected lots from stock. Synthes requested that affected lots be returned and to call Synthes at 1-800-479-6329 for a Return Authorization Number. Completion and return of the Verification section at the end of the letter was requested, including negative responses. |  | 13050 | 13,050 | Nationwide Distribution. | Terminated | August 31, 2015 | 811 | 13050 | Computer
Z-1957-2013 | 64994 | No | HXX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Invue Solid Driver, Tipped, AO. Part Number: 11-70083. |  |  | May 09, 2012 | 2012 | SpineFrontier, Inc. | The nipple of the tipped driver broke off when surgeons were attempting to insert the screw. | Device Design | On May 9, 2012 an email was sent to consignees requesting that the product be returned for servicing and that a replacement part would be sent. |  | 9 | 9 | Distributed in DC and the states of MA, and OK. | Terminated | January 10, 2014 | 611 | 9 | Computer
Z-1956-2013 | 65743 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | StealthStation S7 Framelink Software version 5.4. |  |  | July 15, 2013 | 2013 | Medtronic Navigation, Inc. | Medtronic Navigation, Inc. has initiated a field correction to their Framelink software version 5.4 which can be used on Medtronic Navigation StealthStation 57 and i7 systems or a Medtronic Planning Station due to the potential of discrepancy between the exam image and the displayed measurement or overlay under certain conditions of use. | Software design | A Field Safety Notice was sent on 7/15/13 to users informing them that the Framelink software version 5.4 has been revised to correct the software anomalies described in this letter. A new version will be installed on their system(s) by a Medtronic Navigation representative. Until the software upgrade has been installed on their system, do not use high resolution examinations (greater than 256 x 256). If users are not sure of the exam resolution being used, firm recommend contacting the hospital imaging department or local Medtronic Navigation representative for assistance. Medtronic representatives will visit all identified hospitals and upgrade all systems to version 5.4.1. This will account for all consignees. |  | 572 | 572 | Worldwide Distribution. | Terminated | January 29, 2014 | 198 | 572 | Computer
Z-1931-2013 | 65786 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MHI-TM2000 Linear Accelerator System, Software Version 3.0.0 revision 13 or earlier; Proprietary and Trade Name: MHI-TM2000/ VERO (TM); Common Name: Medical Linear Accelerator. |  |  | June 21, 2013 | 2013 | MITSUBISHI HEAVY INDUSTRIES, LTD., HIROSHIMA MACHINERY WORK | Software anomaly: Mitsubishi Heavy Industries, Ltd. became aware of a potential safety issue associated with the VERO (TM) Linear Accelerator System. The VERO (TM) may save treatment delivery record into wrong location under certain conditions. | Software Design Change | The firm, Mitsubishi Heavy Industries, Ltd. (MHI), sent an "URGENT FIELD SAFETY NOTICE" letter dated June 21, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to do the following: If any minor failure occurs, please always press [ALARM RESET] button on operator console first; and if [EXIT] button is clicked after a minor failure occurs, please click [YES] on the pop up dialog to finish treatment completely. MHI is developing a correction for this issue. Tentatively planned availability date: End of October 2013. The correction will be free of charge and a MHI Customer Service representative will contact affected customers to schedule its installation as the correction becomes available.If you need any further information or support concerning this issue, please feel free to contact our Customer Service representative: VERO GmbH, Kapellenstrasse 12, 8566 Feldkirchen Germany, at +49-89-99-1568-44 or email: franz.gum@vero-sbrt.com. |  | 14 | 14 (2 in US) | Worldwide distribution: US (nationwide) including states of TX and FL; and countries of: Belgium, Italy, Japan and Korea. | Terminated | March 23, 2017 | 1371 | 14 | Computer
Z-1930-2013 | 65843 | No | LPP | Cutaneous oxygen (PcO2) monitor. | Anesthesiology | Anesthesiology | 510(K) Exempt | TCM4 series Base Units, and spare part Power Supply, Radiometer Medical ApS, Akandevej 21, BrÂżnshÂżj, Denmark, www.radiometer.com. The devices are intended for continuous monitoring of carbon dioxide and oxygen partial pressures, oxygen saturation, and pulse rate. |  |  | November 27, 2012 | 2012 | Radiometer America Inc | RADIOMETER became aware that the TCM4 series Base Unit does not comply with IEC 60601-1 Sub Clause 15B. | Device Design | On 11/30/2012 the firm sent an "Urgent Field Safety Notice" to their customers. The notification explained the issue and provided recommended actions. Customers may contact 1-800-736-0600 option 4 with questions. |  | 5874 | 5,874 units | Worldwide distribution: US (nationwide) AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, TX, UT, VA, VT, WA, WI, WV &amp; WY; and country of: Canada. | Terminated | June 16, 2014 | 566 | 5874 | Computer
Z-1933-2013 | 65653 | No | FSM | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Indus Invue Screw Caddy, Model: SI70097 |  |  | June 21, 2013 | 2013 | SpineFrontier, Inc. | Difficulty removing screws from the Invue caddies. | Device Design | SpineFrontier sent a Urgent Invue Caddy Advisory letter dated July 10, 2013, to all affected customers. Customers were advised the removal of the Invue Screw Caddy, SI70097 from the field for modification. Customers were asked to return the parts to SpineFrontier and to complete an Acknowledgement Form indicating that they have read the notice and will initiate the immediate return of the parts noted above.Please feel free to contact us at (978) 232-3990.On 8/21/2013 the recall was expanded. |  | 13 | 13 | Worldwide Distribution - USA, including the states of CT, VA, MA, FL, TX, CO and MS. and the country of Jamaica. | Terminated | January 11, 2016 | 934 | 13 | Computer
Z-1926-2013 | 64984 | No | HXX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Rev E PedFuse Return Screw Inserter, Part Number: 11-80030 |  |  | March 14, 2013 | 2013 | SpineFrontier, Inc. | Inserters have a set screw that fixes a collar to the inserter. The screw set can loosen and fall out of the inserter which allows the collar to slide off the proximal end of the inserter. When the collar is not in place the locking sleeve can also slide off the inserter. | Device Design | Spine Frontier sent an email dated December 10, 2012, to the affected customers that a Pedfuse Inserter with part number PM3979 needed to be returned.Further questions please call (978) 232-3990. |  | 8 | 8 | Distribution only in Texas. | Terminated | June 03, 2015 | 811 | 8 | Computer
Z-1029-2013 | 64393 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Ingenuity TF PET/CT. |  |  | December 20, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Issues have been detected in the Ingenuity TF PET/CT currently running software version 4.0.60605.0 that if they were to re-occur, could cause a risk for patients, users or service personnel, including inaccurate standard uptake values (SUV). Philips Medical Systems (Cleveland), Inc. is issuing a software update to address the issues. | Software design | Philips Healthcare sent an "URGENT-MEDICAL DEVICE CORRECTION" letter dated December 19, 2012 to the affected customer. The letter identified the product, problem, and actions to be taken by the customer. A Philips Service Engineer will contact the customer for implementation of the software update on the affected system. |  | 1 | 1 | Distributed to one customer in Ohio. | Terminated | June 18, 2013 | 180 | 1 | Computer
Z-1907-2013 | 65470 | No | LTW | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | bioMerieux Vitek 2 Systems Software version 6.01 software, used with the VITEK 2 automated system. Kit, V2S 6.01 RP5800-WES7 Media. Part number: 413861. |  |  | June 14, 2013 | 2013 | Biomerieux Inc | A product issue has been identified for customers using software version 6.01 and are interfaced with a Laboratory Information System (LIS). During the update of software from version 5.xx to 6.01, the Bid-directional Computer interface (BCI) antibiotic codes are not updated properly for Cefovecin, Prulifloxacin and Ceftaroline. For the referenced antibiotics, the firm code in the reference table | Software change control | The firm, bioMerieux, sent an "Urgent Product Correction Notice" letter dated June 14, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure the Urgent Product Correction Notice and the Instructions for modifying BCI Translation Codes are distributed to all appropriate personnel in your organization; follow the instructions in the attached file (Instructions for modifying BCI Translation Codes) to resolve the discrepancy associated with Cefovecin, Prulifloxacin and Ceftaroline in the translation table and complete and return the Attachment A: Acknowledgement Form via Fax to: Attn: RA North American ComOps at (919)287-2542.If you require additional assistance or have any questions, please contact your local bioMerieux Clinical Customer Service organization at (800) 682-2666 or Industrial Customer Service organization at (800) 634-7656. |  | 12384 | 12,384 total kits | Worldwide distribution: US (nationwide) and countries of: Argentina, Australia, Austria, Bahrain, Belarus, Benelux, Brazil, Bulgaria, Canada, Chile China, Costa Rica, Croatia, Czech Republic, Dominican Republic, Dutch Antilles, Ecuador, Egypt, France, Germany, Hong Kong, Hungary, India, Indonesia, Israel, Italy, Japan, Kazakhstan, Korea, Kuwait, Libya, Lithuania, Macedonia, Malaysia, Malta, Mexico, Morocco, Myanmar, Netherlands, Nicaragua, Nigeria, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Romania, Russia, Serbia, Singapore, Slovenia, Spain, Sweden, Switzerland, South Africa, Thailand, Tunisia, United Arab Emirates, United Kingdom, Uruguay, and Vietnam. | Terminated | June 01, 2015 | 717 | 12384 | Computer
Z-0834-2013 | 64244 | No | KPA | Automated slide stainer. | Pathology | Pathology | 510(K) Exempt | VENTANA (a member of the Roche Group) DISCOVERY ULTRA; REF/MU 750-601; Research Use Only |  |  | September 28, 2012 | 2012 | Ventana Medical Systems Inc | Potential for a waste fluid overflow condition to accur while running decontamination cycles on the BenchMark ULTRA automated staining platform or the DISCOVERY ULTRA automated staining platform.The waste fluid drainage system may become restricted by build-up of debris in the tubing and in-line filter, thereby inhibiting optimal flow rates. This type of overflow event can lead to migration of flu | Device Design | The firm, Roche sent a "Technical Bulletin" and an letter dated October 3, 2012 to its customers. The letter and bulletin described the product, problem and actions to be taken. The customers were informed that Ventana is aggressively working on solutions to this issue. Until mitigations to the overflow condition are available for implementation; inform all personnel in your organization that work with the BenchMark ULTRA or DISCOVERY ULTRA instrument about the potential for fluid overflow to occur during decontamination cycles; follow the instructions provided in the letter to avoid the documented failure mode and follow the instructions provided in the technical bulletin.Contact your local support center with any questions or concerns or call 520-877-7191. |  | 54 | 54 instruments | Worldwide distribution: USA (nationwide) and country of: Germany. | Terminated | February 24, 2016 | 1244 | 1789 | Computer
Z-1881-2013 | 65162 | No | JLW | Thyroid stimulating hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA Centaur Systems TSH3 Ultra ReadyPack ((REF 06491072 - 100 Tests, REF 06491080 - 500 Tests, REF 04862625 - 2500T). |  |  | May 10, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc | A rare variant of TSH, identified in a small cluster of patients, is not detected. | Device Design | On May 10, 2013, an Important Safety Information notice was sent to all affected Siemens Healthcare Diagnostic customers in the United States. On May 10, 2013, an Urgent Field Safety Notice was sent to all customers outside of the United States. These notices inform the customer that there is a rare variant of TSH, identified in a small cluster of patients that is not detected by some Siemens assays. This notice informs customers that they should continue to use the affected TSH assays as the observed rate of occurrence during a 30-month time period was 0.6 X 10 -7 and there are strong mitigating factors. These mitigating factors include co-interpretation of TSH results in conjunction with thyroxine and TS, patient history and clinical signs and symptoms. Although Siemens believes a look back is not required as the expected frequency of the variant is extremely rare, these notices state that this decision should be made by the medical director. |  | 172913 | 172913 | Worldwide Distribution-USA (nationwide) and the countries of Canada, AE, AR, AT, AU, BA, BD, BE, BG, BO, BR, BY, CH, CL, CM, CN, CO, CR, CY, CZ, DE, DK, DO, DZ, EC, EE, EG, ES, FI, FR, GB, GR, GT, HK, HN, HR, HU, ID, IE, IL, IN, IR, IS, IT, JO, JP, KA, KR, KW, KZ, LB, LK, LT, LV, LY, MA, MK, MT, MX, MY, NI, NL, NO, NP, OM, PA, PE, PH, PK, PL, PT, PY, QA, RO, RU, SA, SE, SG, SK, SV, SY, TH, TM, TN, TR, TW, UA, UY, UZ, VE, VN, XS, and ZA. | Open, Classified |  |  | 330020 | Computer
Z-1033-2013 | 64332 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access 2 Immunoassay System; P/N 107405S 2/12 |  |  | February 12, 2013 | 2013 | Beckman Coulter Inc. | Beckman Coulter is initiating a recall due to a defective newly released aluminum incubator belt pulley that may rub against the light shield, which causes the formation of debris. Debris that is caused by the incubator belt pulley rubbing against the light shield can fall into the reaction vessels (RVs) and present a risk of erroneous test results. | Device Design | Beckman Coulter sent an "Urgent Product Correction" letter dated February 12, 2013 via mail and email to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to place your Access system in the Ready or Not Ready mode. Lift the cover (as you do to access the probes) and look at the left side of the incubator unit under the peripump by the aspirate probe, if you see a silver pulley, immediately discontinue use and contact Technical Support also check for QC failures and system check failures; If you see a white pulley, this is the correct part and you can resume normal activity. Beckman will replace the metal pulley with the plastic pulley if you verified the your instrument has the silver pulley. The customers were also instructed to share this information with staff; and complete and return enclosed Response Form within 10 days.If you have any questions regarding this notice, please contact Technical support from the website: http://www.beckmancoulter.com or by phone at: 1-800-854-3633 in the United States and Canada. |  | 3500 | 3500 units | Worldwide distribution: USA (nationwide) and countries of: Aurstria, Belgium, Bulgaria, Canada, Croatia, Czech Republic, Denmark, Djibouti, Finland, France, Germany, Greece, Hungary, Ireland, Italy, Latvia, Lithuania, Netherlands, Poland, Portugal, Romania, Slovakia, Slovenia, Spain, Sweden, Switzerland, Turkey and United Kingdom. | Terminated | December 23, 2013 | 314 | 10236 | Computer
Z-1851-2013 | 65689 | No | BRZ | Intravascular administration set. | General Hospital | General Hospital | 510(k) | List No. 14203-28; BLOOD SET; 200 Micron Filter, 80 Inch, Non-DEHP; For Blood Product Administration; Made in Costa Rica; Product Inquiries should be directed to Hospira., Inc., Lake Forrest, IL 60045 USA |  |  | April 01, 2013 | 2013 | Hospira Inc. | It was discovered that there was a potential for the piercing pin on certain Hospira blood sets to puncture the outer wall of blood bags during insertion of the pin into the blood bag. | Device Design | Hospira sent an Important Safety Information letter dated April 1, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to exercise EXTREME CAUTION when using Hospira blood bags and to make sure that all instructions for use are included with the blood bag and their facility's protocol for spiking blood bags are completely followed in order to minimize the possibility of puncturing the outer wall of the blood bag. Customers with issues were instructed to call Hospira at 1-800-441-4100 or email ProductComplaintsPP@hospira.com. Customers were asked to complete the attached replay formand return it to the fax number on the form. If customers distributed the product further they were instructed to notify their accounts that received the affected product and ask them to contact Stericycle at 1-866-608-3936 to receive a reply form. Customers with questions were instructed to contact Hospira at 1-800-441-4100.For questions regarding this recall call 224-212-2000. |  | 1152 | 1,152 units | Worldwide Distribution - USA (nationwide) and Puerto Rico and Internationally to Costa Rica. | Terminated | May 24, 2017 | 1514 | 701870 | Computer
Z-1875-2013 | 65605 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Mobiletta Mira |  |  | May 02, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Siemens issued a Field Safety Alert about the potential risk to users when operating the Mobilett Mira. When a Mobilett Mira is regularly exposed to extensive mechanical strain (driven over high thresholds for doors or elevators), it is possible that one of its main wheel shafts could break because of material fatigue, resulting in the chassis falling on the ground on the affected wheel side. This | Device Design | Siemens sent a Field Safety Notice dated May 2. 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.Customers were instructed to :1) Never step or sit on the unit or otherwise put substantial weight on it.2) Avoid moving the unit over high thresholds for doors or elevators where possible, or do so carefully and slowly. 3) Make sure that no person had his/her extremities under the chassis, also when the unit is not moving. Update instructions were also released to alert customers of the potential risk and to provide a hardware change which is intended to improve corrosion resistance.An acknowledgement of receipt form was also sent out for the field safety notice and requested to be returned to Siemens.For further questions please call (888) 826-9702. |  | 39 | 39 | US Distribution including the states of IA, SD, NC, WI, OH, NE, TX, KS, ID, MI, GA, AR and IL. | Terminated | February 24, 2015 | 663 | 39 | Computer
Z-1873-2013 | 63605 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Beckman Coulter AU5800 Clinical Chemistry Analyzer, All Serial Numbers. |  |  | October 26, 2012 | 2012 | Beckman Coulter Inc. | Beckman Coulter is recalling the AU8500 Clinical Chemistry Analyzer because there is no labeling on the diluted detergent tank to show the proper disposition of the on/off valve. | Software design | A recall notification letter was forwarded to all their customers on 11/1/12 who purchased the AU5800 Clinical Chemistry Analyzer. The customers were informed of the problems identified and the actions to be taken. Customers are instructed to complete and return the enclosed Reponse form within 10 days. Customers with questions are instructed to call (800) 854-3633 or contact their local BCI Representative. |  | 338 | 338 units (11 units in the US) | Worldwide Distribution-USA (nationwide) and Puerto Rico and the countries of Germany, China, Austria, Brazil, Italy, Argentina, Korea, Vietnam, Hungary, Ireland, France, Spain, Turkey, Switzerland, Israel, Australia, India, United Kingdom, Netherlands, Malaysia, Russian Federation, Japan, Taiwan, and Canada | Terminated | October 18, 2013 | 357 | 338 | Computer
Z-1871-2013 | 64985 | No | FZX | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Drill Guide - Chameleon FacetFuse Fixation System, Model: 11-00028 Rev. A |  |  | November 09, 2012 | 2012 | SpineFrontier, Inc. | Prior to use in surgery, the Drill Guides were presenting noticeable degrees of denting and/or bending. As a result, drills and taps could not be directed down guides during surgery due to mechanical interference. | Device Design | Spine Frontier sent an Advisory Notice dated November 9, 2012, via email to all affected customers. Customers were advised that their FacetFuse Drill Guides need to be replaced. Customers were instructed to write their number ( included in the email) on the outside of the package and provide us with the tracking information. A replacement part has been sent out to you.If you have any further questions please call (978) 232-3990. |  | 13 | 13 drill guides | US Distribution including Puerto Rico and the states of CO, KS, MA and TX. | Terminated | June 05, 2015 | 938 | 13 | Computer
Z-1874-2013 | 65778 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | MRidium 1145 Dose Reduction System (DERS) drug library kit. (This is an option for the MRidium 3860+MR Infusion Pump System). |  |  | July 01, 2013 | 2013 | Iradimed Corporation | The Dose Error Reduction System (DERS) can indicate an incorrect recommended value for the rate via a specific key sequence during the initial infusion setup. This can result in a risk of over-infusion or under-infusion if the infusion is started with this rate value. | Software Manufacturing/Software Deployment | The firm, Iradimed Corporation, notified by phone and sent an "URGENT MEDICAL DEVICE RECALL NOTICE" via email dated July 1, 2013 all consignees/customers. The notice described the product, problem and action to be taken. The customers were instructed to identify the MRidium 3860+ MR Infusion Pumps which have the 1145 DERS Library card (AM05) installed. If these are not installed, no further action is required at this time; however, if they are installed-remove the 1145 DERS Library Card from the 3860 Pump using the instructions provided. Product will be reconditioned by installing a software upgrade that corrects this problem. The updated software (version 3.5.1) can be seen during the 3860+ Pump's power on cycle.Should your users have any questions following these instructions, please contact Iradimed Customer Service; 866-677-8022 (toll free) or 407-677-8022 ext. 100 or via e-mail - recallsupport@iradimed.com |  | 176 | 176 individuals kits | Worldwide distribution: US (nationwide) including states of: AL, CA, CO, FL, GA, IL, KY, LA, MD, MA, MI, MN, NJ, NY, NC, OH, PA, RI, TN, TX, UT, VA, and WA; and countries of: Australia, Austria, Denmark, Epsom, France, Germany, London, Netherlands, Saudia Arabia, and United Kingdom. | Terminated | July 13, 2015 | 742 | 176 | Computer
Z-1030-2013 | 64665 | No | GYC | Cortical electrode. | Neurology | Neurology | 510(k) | Integra Ojemann Cortical Stimulator; Product / Catalogue No: OCS2. |  |  | March 18, 2013 | 2013 | Integra Burlington MA, Inc. | If the headphone jack is in use during a surgical procedure with the OCS2 and a non-intended voltage (such as static electricity) comes in contact with the outer case of the unit, that voltage could be transmitted to the patient and could be a potential source of injury. | Device Design | Integra initiated a voluntary recall on March 18, 2013 by providing written notification to consignees either by traceable courier service or traceable emails regarding the correction that Integra records indicate have an OCS2 since it was introduced to the market in 2007. The consignee notification:Advised of the nature of the issue and to immediately stop using the headphone jack while using the OCS2. Requested identification by S/N any OCS2's they have, and complete / return the Acknowledgment Return Form.Advised that Integra will contact the consignee to schedule a time / date to correct the OCS2's they have identified. Questions, contact service hotline at 1-888-772-7378. |  | 371 | 371 units | Worldwide Distribution-USA (nationwide) and the countries of Austria, Belgium, Czech Republic, Belgium, Denmark, Spain, Great Britain, Hungary, Ireland, Italy, Netherlands, Poland, Morocco, Netherlands, Poland, Portugal, Saudia Arabia, Sweden, Turkey, and South Africa. | Terminated | February 25, 2014 | 344 | 371 | Computer
Z-1038-2013 | 64675 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | SIEMENS brand COHERENCE RT Therapist versions RTT MR1.5, RTT 2.1/2.1A and RTT 2.2/ 2.2B; Software release for COHERENCE Therapist 2.3 for Primus and Oncor Linear Accelerators; |  |  | March 05, 0013 | 13 | Siemens Medical Solutions USA, Inc | Siemen's issued a Customer Information Letter under Update Instruction TH016/12S to alert all its affected customers of the software issues. The software release is a result of the final system testing addressing multiple performance defects and several safety issues that were not considered critical (for immediate correction) for the safe use of the device. | Software design (manufacturing process) | Siemens Medical Solutions sent a Customer Information letter via mail to all affected customers. The letter identified the affected product, problem and provide instructions on the correction. Customers were instructed to include the Customer Information letter in the Owner Manual chapter "Safety Advisory Letters".9-16-13 UPDATE: - Control Console SW 9.2.18 was released to support the rol-out10/18/2013; Distribution of update instructions for TH017/13S and TH018/13S - Distribtuion of the Control Console SW 9.2.18 released |  | 621 | 621 units (foreign account) in total | Worldwide Distribution | Terminated | August 28, 2014 | 731026 | 621 | Computer
Z-1848-2013 | 65491 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Siemens Luminos dRF Fluoroscopic X-Ray system |  |  | May 07, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Siemens became aware of an unintended behavior when using the Luminos dRF with software version VD10A. The possibility exists that a software error may occur causing the table to tilt into an upright position and may collide with the floor. Four consignees were notified of this recall by Update Instruction XP067/12/S -- all four units have been corrected. | Software design | Siemens initiated their correction of this product in May, 2013 by sending an Safety Update Instruction XP067/12S to resolve the software error to all affected customers. All units have been corrected.Further questions please call (610)-219-6300. |  | 4 | 4 | US Distribution including the states of CA, IN, NJ and OK. | Terminated | December 24, 2013 | 231 | 4 | Computer
Z-1973-2013 | 65381 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | TerumoÂż Advanced Perfusion System 1 (APS1, System 1) |  |  | July 26, 2013 | 2013 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (TCVS) has become aware that an Advanced PerfusionSystem 1 (APS1) Centrifugal Pump 'Service Pump' message may occur due to user inducedmotor movement. In this instance, the APS1 will identify the movement (as it isdesigned) and display a Service Pump message error. The user might interpret this as apump which requires service rather than a message displayed d | Device Design | Terumo Cardiovascular Systems (TCVS) notified their consignees of the Urgent Medical Device Recall letter dated 7/26/13 and was mailed 7/29/13. The notification provided reason for the device correction, affected product, how the device will be corrected, and the potential health hazard associated with a user's inaccurate interpretation of the service pump error message. Instructions for consignee responses to the recall notice communication may be faxed to 734-741-6149 or emailed to the address on the return form. If you have any questions or concerns, call Terumo Cardiovascular Systems (TCVS) Customer Service at 1-800-521-2818 M-F, 8am-6pm ET. Terumo Cardiovascular Systems (TCVS) will provide an addendum to the APS1 Manual to their consignees. The addendum will be placed at the beginning of Chapter 6, Centrifugal Control Unit. |  | 1657 | 1657 total | Worlwide Distribution - US Nationwide: AL,AR,AZ, CA, CO, CT, DC, DE, FL,GA, HI, IA, ID, IL, IN, KY, LA, MA , MD, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WI, WV and the countries of: AUSTRALIA,UNITED ARAB EMIRATES (UAE),Hong Kong, onesia,Singapore,Taiwan,Thailand, Philippines,CHILE, Malaysia,USA,South Korea,BELGIUM,Japan,Mexico,CANADA | Open, Classified |  |  | 1657 | Computer
Z-1015-2013 | 53991 | No | BRY | Anesthetic cabinet, table, or tray. | Anesthesiology | Anesthesiology | 510(K) Exempt | Pyxis Anesthesia System Model 2000 (PAS 2000) |  |  | November 20, 2009 | 2009 | CareFusion Corporation | The recall was initiated because Carefusion has confirmed that when the PAS 2000 drawers due to inaccessability of drawers caused by system down time or lock-up. | Device Design | Carefusion sent a Safety Alert Notification letter dated January 13, 2010 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to ensure back-up procedures were in place in case of system down-time or lock-up, including review of the back-up procedures with staff. The letter included an update to the PAS 2000 User Guide. The letterstates that a CareFusion field representative will be contacting customers to make arrangements to apply warning labels to the PAS 2000 electronics doors. Customers were instructed to complete and return the Response Card included. For questions contact your local field support representative ot the Advocacy department at customer.advocacy@carefusion.com. |  | 2637 | 2637 | Worldwide Distribution - USA (Nationwide) including CANADA. | Terminated | March 28, 2013 | 1224 | 2637 | Computer
Z-1002-2013 | 64461 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Vitrea Enterprise Suite, Vitrea, VitreaAdvanced, VitreaCore (VitalConnect) and Vitrea fX prior to version 6.3.3, except 6.1.6, 6.2.3 and 6.3 |  |  | February 05, 2013 | 2013 | Vital Images, Inc. | Vital Images has found a potential error in measurement generated through the software which affects all versions of Vitrea Enterprise Suite, Vitrea, VitreaAdvanced, VitreaCore (VitalConnect) and Vitrea fX prior to version 6.3.3, except 6.1.6, 6.2.3 and 6.3.This error occurs when images are rotated at the time of reconstruction by the scanner in non-90 degree increments of the transverse/axia | Software design | Vital Images sent an "Urgent Field Safety Notice' letter dated February 5, 2013, to all affected customers. The letter identified the probelm the product and the action needed to be taken by the customer. Please immediately inform all potential users of Vitrea of the information contained in this letter. We recommend that all measurements associated with the error described in this letter be verified against results from other technologies.You will be contacted by Vital Images Customer Support or by the customer support department of your distributor to arrange for installation of this patch on your system. We recommend that you implement the patch. If you have an earlier version, Vital Images Customer Support or your distributor will contact you to discuss your options for replacement.Please sign this Notice in the space below to verify that you have received and understood it, and return your signed Notice by email to fieldnotices@vitalimages.com.If you have questions, direct customers of Vital Images should contact Vital Images Customer Support at 1-800-208-3005 or by e-mail to support@vitalimages.com; customers through a distributor of Vital Images products should contact that distributor.Updated 5/28/2013:A second consignee Vital Images, Inc "Urgent Field safety Notice" letter dated March 26, 2013 was sent to consignees on 3/25/13. The follow-up letter was sent via e-mail and mail to 4,616 already-notified consginees, who had received affected versions 6.3.0, 6.2.3, or 6.1.6. The letter described the problem and the affected product being recalled and provided a number of steps to follow which included to retun the Effectiveness Check form. |  | 38245 | 7649 (Expnaded 5/28/13 to 7654) | Worldwide Distribution - All states in continental USA including DC, PR and the countries of ALGERIA, ARGENTINA, ANGOLA, ARMENIA, AUSTRALIA, AUSTRIA, BENGLADESH, BELARUS, BELGIUM, BOLIVIA, BOTSWANA, BRAZIL BULAGARIA, CAMBODIA, CANADA, CENTRAL AFRICAN REPUBLIC, CHAD, CHILE, CHINA , COLOMBIA, CONGO, CROATIA, CYRPRUS, CZECH REPUBLIC, DENMARK, DJIBOUTI, DOMINICAN REPUBLIC, ECUADOR, EGYPT, ESTONIA, ETHIOPIA, FINLAND, FRANCE, GREAT BRITAIN, GEORGIA, GERMANY, GHANA, GREECE, GUAM, HONDURAS, HONG KONG, HUNGARY, ICELAND, INDIA, INDONESIA, IRAQ, IRELAND, ISRAEL, ITALY, IVORY COAST, JAPAN, JORDAN, KAZAKHASTAN, KOREA, LEBANON, LITHUANIA, MACEDONIA, MALYASIA, MALI, MAURITIUS, MEXICO, MALDOVA, MOSCOW, NYNMAR, NAMBIA, NAMIRGIA, NEPAL, NETHERLANDS, NEW ZEALAND, NIGER, NIGERIA, NORTHERN IRELAND, NORWAY, OMAN, PAKISTAN, PANAMA, PARAGUAY, PERU, PHILLIPPINES, POLAND, PORTUGAL, QATAR, ROMANIA, RUSSIA, RUSSIAN FEDERATION, SAUDI ARABIA, SENEGAL, SERBIA, SEBIA AND MONROBIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOMALIA, SOUTH AFRICA, SPAIN, SRI LANKA, SUDAN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, THE NETHERLANDS, TIBET, TUNISIA, TURKEY, UKRAINE, UNITED ARAB EMIRATES, UNITED KINGDOM, UZBERKISTAN, VENEZUELA, VIETNAM and YEMEN. | Terminated | May 08, 2015 | 822 | 38245 | Computer
Z-2056-2013 | 65824 | No | HBG | Manual cranial drills, burrs, trephines, and their accessories | Neurology | Neurology | 510(k) | Integra Leyla Ball Joint Clamp, also as included in Leyla Table Mounting Hardware. |  |  | July 17, 2013 | 2013 | Integra Limited | As a result of complaint investigations it was identified that there is a potential that the Leyla Ball Joint Clamp (R2383) may not provide the level of stability needed to maintain retraction in all cases. | Device Design | Integra sent an Urgent Medical Device letter dated July 17, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were requested to review your inventory and determine if you have any of the affected products. If so, stop using them immediately and remove them from your service. When your form is received, Customer Service will you an RMA number and directions for returning the products.Should you have any questions regarding these instructions, please contact Integra Customer Service at 1-800-997-4868. |  | 65 | 65 | Worldwide Distribution - USA including Puerto Rico and the states of AZ, CA, FL, GA, IL, IN, KY, MA, MI, MS, NE, NM, NY, OH, PA, TN, TX, UT, VA, WA, WI and WV., and the countries of Spain, Singapore, South Africa, Canada, Libya,Pakistan and Cayman Islands. | Terminated | October 01, 2014 | 441 | 65 | Computer
Z-2114-2013 | 65852 | No | GKL | Automated cell counter. | Hematology | Hematology | 510(k) | ADVIA 120 Hematology Analyzer, Catalog Numbers SMN 10360958 and 10360959 . |  |  | June 26, 2013 | 2013 | Siemens Healthcare Diagnostics | The ADVIA 120 optics cover or hood is located on the very top of the ADVIA 120 analyzer. When the optics cover is raised, it is supported by two gas struts (SMN 10309266) attached to the cover on the left and right sides. Over time, it is possible that these struts may lose their effectiveness and fail to support the optics cover in the open position. This could cause the cover to fall during m | Equipment maintenance | Siemens sent an Urgent Device Correction Letter dated June 26, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed of the ADVIA 120 system displays a problem when raising and supporting the optics cover, immediately contact your service provider to arrange to have the cover struts serviced. If the optics cover will not stay in the up position, as required, do not perform maintenance procedures on the optics area until the hood struts have been replaced. Siemens field support will be proactively replacing the optics cover hood struts on a regular preventive maintenance cycle based on the age of the system.Please forward this notification to whomever you may have distributed this productWe apologize for the inconvenience that this situation has caused. Thank you for your patience and continued support.This response form is to confirm receipt of the enclosed Siemens Healthcare Diagnostics Urgent Device Correction dated June 2013 regarding Optics Cover Struts (SMN 10309266).Please read each question and indicate the appropriate answer. Fax this completed form to Siemens Healthcare diagnostics at the fax number indicated at the bottom of this page.PLEASE FAX THIS COMPLETED FORM TO THE TECHNICAL SOLUTIONS CENTER AT (302) 631-7597. |  | 4319 | Domestic: 950 units; Foreign: 3369 units | Worldwide Distribution - USA (nationwide) and the countries of Austria, Australia, Belgium, Bosnia - Herzegovina, China, Croatia, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, Hungary, Iceland, India, Indonesia, Ireland, Iraq, Israel, Italy, Japan, Kenya, Republic of Korea, Kuwait, Lebanon, Lithuania, Luxembourg, Martinique, Malaysia, Netherlands, Norway, New Zealand, Philippines, Poland, Portugal, Romania, Russian Republic, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom and Vietnam. | Open, Classified |  |  | 4319 | Computer
Z-2109-2013 | 65886 | No | FMK | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ACCU-CHECK FastClix Lancing Device. It is intended for use in conjunction with a blood glucose monitoring system. |  |  | August 19, 2013 | 2013 | Roche Diagnostics Operations, Inc. | Internal Roche inspections have revealed that after the lancet is used for a finger stick, the needle may not fully retract back within the lancet drum or back into the cap. This allows the needle to protrude inside the cap or outside of the cap of the ACCU-CHEK FastClix lancing device. | Equipment maintenance | On 8/19/ 2013, Roche sent Distributor/Retailer URGENT MEDICAL DEVICE CORRECTION notifications which described the recall issue and health consequences, provided product description, and provided actions required of consignees. Roche sent a CUSTOMER BULLETIN to their direct customers. URGENT MEDICAL DEVICE CORRECTION faxback form for consignee acknowledgement of recall is to be faxed to 1-888-240-2804. Consignees may contact the ACCU-CHEK Customer Care Service Center, 24 hours a day, seven days a week at 1-800-440-3638 if they have questions about the information contained in the Distributor/Retailer Letter. |  | 36332 | 36,332 | Nationwide Distribution | Terminated | May 12, 2014 | 266 | 36332 | Computer
Z-2098-2013 | 65880 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Ultrasound ViewPoint 6. |  |  | August 07, 2013 | 2013 | GE Healthcare, LLC | GE has recently become aware of a potential safety issue due to the incorrect calculation of the pulsatility index value (PIV) for venous flow measurements associated with obstetric reporting of the GE ViewPoint 6 system. The ViewPoint calculated Doppler PIV for veins could be incorrect. An incorrect calculation of the ductus venosus PIV could lead to false high risk for trisomy 13, 18, and 21, | Software Design Change | GE Healthcare sent An Urgent Medical Device Correction letter dated August 7, 2013 to all affected customers. The letter described the affected product, problem and instructed customers to discontinue use of venous PIV for making any and all clinical decisions. The letter informed customers that a GE service engineer will install a correction at no cost. For questions contact your local GE Healthcare Representative or call 877-644-3114 (US) and 49 (0)8153-9311-11(Europe/Asia) |  | 225 | 225 | Worldwide Distribution - USA Nationwide including Puerto Rico and the countries of: ANDORRA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BRAZIL, BULGARIA, CANADA, CHILE, COLOMBIA, CZECH REPUBLIC, ESTONIA, FINLAND, GERMANY, HONG KONG, INDIA, INDONESIA, IRAQ, ITALY, JAPAN, KUWAIT, MALAYSIA, Norway, OMAN, PALESTINIAN TERRITORY, OCCUPIED, POLAND, PORTUGAL, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TURKEY, and UNITED KINGDOM. | Terminated | April 25, 2014 | 261 | 225 | Computer
Z-0965-2013 | 64546 | Yes | LWS |  |  | Cardiovascular | PMA | Boston Scientific/Cameron Health SQ-RX Model 1010 subcutaneous pulse generators. The S-ICD System is intended to provide therapy for the treatment of life-threatening ventricular tachyarrhythmias. |  |  | March 01, 2013 | 2013 | Boston Scientific CRM Corp | An internal protective fuse can be unintentionally activated while the device is charging its capacitors for shock delivery or induction. Should this occur, the defibrillator would not be able to deliver therapy or communicate with the Q-TECH Model 2020 programmer, and would be unable to emit tones or otherwise respond to magnet application. | Software design | Consignees were sent a Boston Scientific "Important Medical Device Information" letter dated March 1, 2012. The letter was addressed to "Dear Doctor". The letter described the product involved in the recall, the problem and provided Recommendations/Actions. US customers may contact Technical Services for more information at 800-227-3422. |  | 2090 | 2090 (570 USA; 1520 OUS) | Worldwide Distribution, USA including the states of AZ, CA, CO, DC, GA, IL, MD, MA, MN, NJ, NY, NC, OH, PA, SC, VA, WA and the countries of AUSTRIA, BELGIUM, CANADA, CZECH REPUBLIC, DENMARK, FRANCE, GERMANY, IRELAND, ITALY, NETHERLANDS, NEW ZEALAND, PORTUGAL, SLOVAKIA, SPAIN, SWEDEN, SWITZERLAND, and the UNITED KINGDOM. | Terminated | May 23, 2013 | 83 | 2090 | Computer
Z-2086-2013 | 65672 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Illico Ti Cannulated Polyaxial Screw 7.5mm X 40mm. The Illico MIS Posterior Fixation System is intended to facilitate the surgical correction of noncervical spinal deformities by providing temporary internal fixation and stabilization during bone graft healing and/or fusion mass development. |  |  | June 28, 2013 | 2013 | Alphatec Spine, Inc. | The firm initiated this recall of certain lots of Illico Ti Cannulated Polyaxial Screws 7.5mm x 40mm implants, due to an off-center cannulation hole at the proximal end of the bone screw. There is a potential risk that when the screw is mated to a polyaxial screwdriver, the guide wire may not pass through the screwdriver. | Equipment maintenance | The firm, Alphatec Spine, sent an "URGENT-MEDICAL DEVICE RECALL NOTIFICATION" with response sheet dated June 28, 2013 via a combination of e-mail, Fedex, and telephone contact to its customers. The notification described the product, problem and actions to be taken. The customers were instructed to immediately return any and all affected lots of the illico Ti Cannulated Polyaxial Screw 7.5 mm x 40 mm implants that have been sent to them; if the device has previously been implanted, there is no further or additional risk post-implantation; immediately contact Alphatec Spine Customer Service Department - customerservice@alphatecspine.com on how to replace the implants; and complete and return the last page of the notification via mail to Alphatec Spine, Inc. 5818 El Camino Real, Carlsbad, CA 92008; email: sarrhenius@alphatecspine or fax to: (760) 431-0289.If you have any questions or need any additional information, please contact Director of Compliance at 760-494-9636 or email at dmcgrath@alphatecspine.com. |  | 333 | 333 screws | Worldwide distribution: US (nationwide) including states of:US states: AZ, CA, FL, GA, IL, MD, NJ, NY, OK, PA, TN, TX,and WA ; and countries of: Spain, France, Japan.International: Spain, France, Japan. | Terminated | October 09, 2014 | 468 | 333 | Computer
Z-2095-2013 | 65848 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | PRIMUS, ONCOR, ARTISTE Linac systems with component Tabletop, 550TxT, TT-A s/n 1001-1233 and Tabletop 550TxT, TT-S s/n 1001-1083. |  |  | July 18, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Siemens has become aware of a customer complaint concerning the rail threaded holes of the 550TxT treatment tabletop. Due to a manufacturing error, the screws holding the side rails in position on the tabletop might break off. | Device Design | Siemens released Updated Instruction containing the Medical Device Correction Field Safety Notice of Improvement letter, which was delivered by the Siemens Customer Service Group to all affected customers. The release of this UI TH002/12/S to distribute the update instruction began on July 18, 2013. This document includes completion protocol, which will be completed at each facility and returned to Siemens. A copy of the notification should be kept in the Digital Linear Accelerator System owners manual under Safety Advisory letters.UPDATE: a secnd update intruction, TH003/12/S has been released to begin the replcement of tabletops identified through the inspection carried out through TH002/12/S. |  | 179 | 179 active devices. | Worldwide Distribution - USA Nationwide in the states of NJ, IN, PA, FL, OH, MN, MO, CA, AL, NE MI, LA. and countries of:Argentina, China, South Africa, Pakistan, Spain, Switzerland, Italy, Algeria, Croatia, Turkey, Canada, Germany, France, Poland, Japan, Russian Federation. Lebanon, Egypt, Slovakia, Latvia, Kenya, Brazil, Kuwait, Saudi Arabia, Greece, Belgium, Taiwan, Portugal. | Terminated | November 24, 2014 | 494 | 179 | Computer
Z-2055-2013 | 65677 | No | KWT | Shoulder joint metal/polymer non-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | 00 4309-049-00 Trabecular Metal Reverse Shoulder Glenosphere Distractor Rx, Non sterile (sterilize before use) |  |  | August 01, 2013 | 2013 | Zimmer, Inc. | Zimmer is initiating a lot specific recall of the Trabecular Metal" Reverse Glenosphere Distractor due to the potential of the device not functioning properly should it be utilized without proper lubrication. As a result, there is a potential for the device to fail to provide impact to the Glenosphere when the trigger is pulled. In addition, there is a potential for the foot of the device to fract | Device Design | Zimmer sent a Urgent Medical Device Recall Notification letter August 1, 2013, and via E-mail and courier to all affected customers. The initial recall notice identified the reason for the recall, possible adverse risks and injuries and informed customers that a replacement will be available in the coming months. In addition, this letter will remind that the instrument is to be lubricated before each use and to inspect the foot. Customers are instructed to inspect any recalled product before and after use for cracks at the foot and ensure the foot is intact. Also ensure all moving parts are lubricated prior to each use according to the Recommendations for Care, Cleaning, Maintenance, and Sterilization Manual 97-5000-170-00. If you find a device which does not function properly or has a damaged foot, provide it to your Zimmer sales representative for return to Zimmer on a Product Experience Report (PER). Questions and concerns should be addressed to the Zimmer customer call center at 1-877-946-2761 between the hours of 8 a.m. and 5 p.m. EST. Phase 2: As devices become available for distribution, a 2nd letter will be sent to the distributors and risk managers. Distributors will be instructed to order replacement devices as a means of completing the affected instrument cases in circulation within their territories. Instructions for the return of devices will be included. |  | 771 | 771 | Worldwide Distribution - USA (nationwide) including the countries of Singapore, Brazil, Taiwan, Australia, Canada, Germany(with further distribution to UAE, Austria, Belgium, Switzerland, Czech Republic, Denmark, Spain, Finland, France, Greece, Israel, Italy, Malta, Norway, Portugal, Sweden and South Africa, | Terminated | March 02, 2015 | 578 | 771 | Computer
Z-2059-2013 | 65950 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips Easy Upgrade DR, Code No: 712086 |  |  | August 08, 2013 | 2013 | Philips Healthcare Inc. | The mirror icon on the bottom of each image could be misinterpreted as a RIGHT lead marker for side indication if the image is mirrored within a PACS system. | Software design | Philips Healthcare sent an Urgent Field Safety Notice letter dated July 24, 2013, to all affected customer. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed : When reading an image, a physician should not use the mirror icon for image side interpretation. 2. Clinical staff should follow the Instructions for Use and use lead letters to indicate body orientation of patient or equivalent electronic markers to provide Left/Right indications, during all X-ray image acquisitions. Philips will upgrade the system, which involves software and hardware upgrades of the system. The software provides a replacement icon that reduces the risk of the potential misinterpretation of the original mirror icon. Philips will also provide updated Instructions for Use for the work spot and the system (including an addendum).If you need any further information or support concerning this, please contact your local Philips representative Technical Support Line at 1-800-722-9377.We apologize for any inconvenience this may cause and trust that this information is adequately addressing any concerns you may have. |  | 89 | 89 | Worldwide Distribution - US Nationwide including the states of AZ, CA, CO, DE, FL, GA, MO, MS, NC, NH, NY, PA and TX., and the countries of Canada, Australia, China, Czech Republic, France, Germany, Greece, Italy, Netherlands, Portugal, Russia, Saudi Arabia, Slovenia, Spain, Sweden, Switzerland, Turkey, United Arab Emirates and United Kingdom. | Completed |  |  | 89 | Computer
Z-2082-2013 | 65080 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | 2008 MeDS Pump, Venofer(R) Pump |  |  | April 01, 2013 | 2013 | Fresenius Medical Care Holdings, Inc. dba Renal Solutions | Product is being recalled because the combination of the current 50 mg vial, the 2008 MeDS pump, and the 2008 MeDS pump tubing Set for CombiSet Bloodlines can lead to a situation where Venofer(R) delivery does not meet dosage accuracy requirements. This may lead to a slight over or under delivery of Venofer(R) that does not pose an acceptable risk to the patient. Please note that this recall | Device Design | Renal Solutions sent an Urgent Notification letter dated April 1, 2013 via certified mail to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to cease further use of the device. The letter informed customers that a Fresenius Medical Care Technical Service/Customer Service representative will contact customers to arrange for removal of the device. For questions contact Fresenius Medical Care Customer Service at 800-227-2572. |  | 132 | 132 | USA Nationwide Distribution in the states of CA, MA, and NY. | Terminated | March 02, 2015 | 700 | 132 | Computer
Z-2089-2013 | 65769 | No | BRY | Anesthetic cabinet, table, or tray. | Anesthesiology | Anesthesiology | 510(K) Exempt | Mc Kesson Anesthesia-Rx Model Number 8.0.2 &amp; 8.0.1 with RapiDispense |  |  | July 22, 2013 | 2013 | MCKESSON TECHNOLOGIES INC | The firm issued a clinical alert after users reported an issue where the Anesthesia-Rx system was interrupted ANF became temporarily unavailable during use. | Software design | McKesson sent a Clinical Alert letter dated July 22, 2013 to all affected customers. The letter identified the affected product, symptoms, and actions to be taken. McKesson committed to releasing 8.0.2 Update 1.5 to affected customers which will include enhanced code to prevent the unhandled excpetion errors. Additional resolution steps were included in the letter. For questions call 1-800-700-8737. |  | 106 | 106 | USA Nationwide Distribution in the states AR, MD, NJ, NY, PA, and WV. | Terminated | June 19, 2014 | 332 | 106 | Computer
Z-0971-2013 | 63973 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Specialist 2 IM Rod, Non-Sterile and associated Surgical Techniques: |  |  | February 05, 2013 | 2013 | DePuy Orthopaedics, Inc. | Potential for the IM rod to break, leaving fragments in the patient. Surgical Techniques guidance has been updated until design changes can be made to reduce likelihood of potential failure. | Device Design | DePuy Orthopaedics sent an Urgent Medical Device Correction Notice letter dated February 5, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were asked to complete and return the FAX back form to FAX 574-371-4939 or email kseppa@its.jnj.com Questions about device correction information call Post Market Surveillance, 574-372-7333, Sales force questions call Product Specialist, 574-372-7303 and Surgeon questions DePuy's Scientific Information Office at 1-888-554-2482. |  | 2600 | 4000; 2600 in USA | Worldwide Distribution - USA Nationwide. | Terminated | January 06, 2014 | 335 | 2600 | Computer
Z-2057-2013 | 65556 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes 2.4mm VA-LCP Two Column Volar Distal Radius Plate and Torque limiting attachment, 1.2Nm |  |  | May 31, 2013 | 2013 | Synthes (USA) Products LLC | The firm is initiating a Field Safety Correction for the labeling of the 2.4mm VA-LCP Two Column Volar Distal Radius Plate and 1.2 Nm Torque Limiting Attachment. The US Technique guide lists 1.2Nm TLA as an option for use with the 2.4mm Variable Angle LCP (VA-LCP) Two-Column (Standard &amp; Narrow) Volar Distal Radius Plate. The potential exists for the VA locking screw to be driven through the 2.4mm | Device Design | Synthes sent an Urgent Notice: Medical Device Labeling Correction letter dated May 31, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The firm is initiating a Field Safety Correction for the labeling of the 2.4mm VA-LCP Two Column Volar Distal Radius Plate and 1.2 Nm Torque Limiting Attachment. The US Technique guide lists 1.2Nm TLA as an option for use with the 2.4mm Variable Angle LCP (VA-LCP) Two-Column (Standard &amp; Narrow) Volar Distal Radius Plate. The potential exists for the VA locking screw to be driven through the 2.4mm VA-LCP Two Column Volar Distal Radius Plate when used with a 1.2Nm Torque Limiting Attachment. A screw that passes through the VA-LCP plate is not likely to result in a serious adverse event although should it occur, additional medical intervention during the index procedure could be needed.If you have any questions, please call (610) 719-5450, or contact your Synthes sales Consultant. |  | 119891 | 119, 891 | Nationwide Distribution including the states of AK. AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MD, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI SC, SD, TN, TX, UT, VA, VT, WA WI and WV. | Terminated | September 21, 2015 | 843 | 119891 | Computer
Z-2088-2013 | 65607 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension Vista 500 and Dimension Vista 1500 System Aliquot Well Double Dispense with software versions 3.5.1 and 3.6. |  |  | June 19, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Siemens is conducting a field correction for Dimension Vista 500 and Dimension Vista 1500 systems using software versions 3.5.1 and 3.6. Under specific rare conditions which require a system reset, there is a potential for a sample to be dispensed into an aliquot well that has previously been used for a QC or calibrator product from a Vista vial. If two samples are dispensed into the same aliquot | Software Design Change | The firm, Siemens, sent an "Urgent Medical Device Correction" letter dated June 2013 to its consignees/customers. The letter describes the product, problem and actions to be taken. The customers were instructed to take the following actions: 1) In the event that an instrument Reset is required for any reason, review the Current Activities Screen to determine if there are any pending QC or Calibration tests.2) If there are no pending QC of Calibration tests, Rest the instrument or troubleshoot the instrument per the Operators Guide.3) If there are pending QC or Calibration tests, restarting the software is necessary before resuming patient processing. Go to Operation -&gt; Shutdown, confirm Yes, then on the blue console menu screen select start Vista. Consignees/customers were also asked to complete the Field Effectiveness Check form attached to the letter via fax to Siemens Healthcare Diagnostics at (302) 621-8467, and retain the letter with laboratory records, and forward the letter to those who may have received the product.If you have any questions, contact your Siemens Customer Care Center at (800-441-9250)-Technical Solutions or your local Siemens Technical Support representative. |  | 1937 | 1937 | Worldwide distribution: US (nationwide) including PR and countries of: Australia, Austria, Belgium, Canada, Czech Republic, Denmark, Finland, France, Germany, Italy, Japan, Malaysia, Netherlands, New Zealand, Norway, Portugal, Republic Korea, Saudia Arabia, Slovakia, Slovenia, South Africa, Spain, Switzerland, and United Kingdom. | Terminated | August 27, 2014 | 434 | 1937 | Computer
Z-2083-2013 | 65773 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | BRAINLAB; FRAMELESS SRS QA TARGET POINTER |  |  | July 16, 2013 | 2013 | Brainlab AG | The Frameless SRS QA Target Pointer - Pointer Cap with the engraved cross hairs may become loose even when only a moderate force is applied to the Pointer Cap. | Device Design | BrainLab sent an "FIELD SAFETY NOTICE / PRODUCT NOTIFICATION" dated July 9, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed to discontinue use of the SRS QA Target Pointer. Consignees can instead use the Phantom Pointer feature until Brainlab can provide replacement hardware. Brainlab anticipates being able to provide all impacted facilities with replacement hardware by the end of July 2013.We sincerely apologize for any inconvenience and thank you in advance for your co-operation. If you require further clarification, please feel free to contact your local Brainlab Customer Support Representative. Customer Hotline: +49 89 99 15 68 44 or +1 800 597 5911 (for US customers) or byE-mail: support@brainlab.com. Fax Brainlab AG: + 49 89 99 15 68 33 |  | 6 | 6 units | Worldwide Distribution - US Distribution only to AR., and the countries of Australia, Germany and Japan. | Open, Classified |  |  | 6 | Computer
Z-2078-2013 | 65800 | No | HWT | Template for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | TRIAL STEM for MP(R) RECONSTRUCTION PROTHESIS, REF 99-0155/20, Stem size S-4, prox. Stem dia. 20 mm, Length 180 mm, WALDEMAR LINK GmbH &amp; CoKG, Barkhausenweg 10 . D . 22339 Hamburg |  |  | July 16, 2013 | 2013 | Waldemar Link GmbH &amp; Co. KG (Corp. Hq.) | The trial stems were manufactured using the wrong design. | Device Design | The firm notified their consignees of their product recall by parcel post on 07/16/2013. |  | 8 | 8 units | US Distribution to the state of Tennessee only. | Terminated | June 04, 2014 | 323 | 127 | Computer
Z-0996-2013 | 63621 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Medfusion Model 3500 Syringe Pump, Model #3500-415 and 3500-500. |  |  | May 14, 2010 | 2010 | Smiths Medical ASD, Inc. | An issue with the plunger rod seal was identified where the seal had the potential to migrate inside the pump case top during slow rate/ high back pressure infusions. | Device Design | On May 14th, 2010, Smiths Medical contacted the facility that received the pumps via e-mail and requested that they immediately return the pumps. The pumps were returned per Returned Goods Authorization (RGA) and exchanged for replacement pumps manufactured with plunger rod seal G6000077. |  | 6 | 6 | Distributed in the state of NH. | Terminated | April 11, 2013 | 1063 | 6 | Computer
Z-2084-2013 | 48997 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Continuous Ventilator/Continuous Ventilator with Respiratory Mechanics, Model No. V1000 |  |  | October 29, 2007 | 2007 | Respironics California Inc | The power to the Esprit Display's Backlight is interrupted, causing the GUI to flicker or go dark and making the display unreadable. | Device Design | A Medical Device Field Correction letter dated 10/26/2007 was sent to all customers who received the Esprit Ventilator. The recall letter informed the customers of the problems identified and the actions to be taken. Customers with questions are instructed to contact Respironics' Customer Service Dept. at (800) 345-6443, select Option#5, then select Option #4. |  | 3255 | 3,255 units (2302 units in US) | USA Nationwide and Worldwide to Algeria, Argentina, Australia, Bangladesh, Brazil, Cambodia, Canada, Chile, Colombia, Ecuador, Egypt, France, Germany, Guatemala, Hong Kong, India, Italy, Japan, Libya, Malaysia, Mexico, Pakistan, Peru, Philippines, Puerto Rico, Russia, South Africa, South Korea, Spain, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Venezuela, and Vietnam. | Terminated | August 26, 2013 | 2128 | 3255 | Computer
Z-2046-2013 | 65865 | No | HKI | Ophthalmic camera. | Ophthalmic | Ophthalmic | 510(k) | NAVIS-EX , Image Filing Software for NIDEK AFC Fundus camera, Ver 1.1.0 to 1.3.5. NAVIS-EX is indicated for use as file management software for ophthalmic images. |  |  | August 01, 2013 | 2013 | Nidek Inc | Under certain circumstances, the Left eye image may be saved as the Right eye image. | Software design | "Medical Device Correction Notification" letters dated August 5, 2013 were sent to consignees through their US contracted distributor - Marco Opthalmic. The letter described the product issue and provided information on the correction. Customers can contact 800-223-9044 for questions. |  | 86 | 86 | Nationwide US distribution only. | Terminated | February 25, 2014 | 208 | 86 | Computer
Z-2047-2013 | 65609 | [6 empty] | PSA Assays for use in the IMMULITE/IMMULITE 1000 Systems Analyzers, Catalog Numbers LKPS1, LKPS5 and LKPTS1(D); and in the IMMULITE 2000/2000XPi Systems Analyzers, Catalog Numbers L2KPS2, L2KPS6, L2KPTS2(D) and L2KPTS6(D). IVD Immunoassay kits in the following sizes: 100, 200, 500 and 600 test kits; For in vitro diagnostic use with the IMMULITE/IMMULITE 1000 and the IMMULITE 2000/ 2000 XPi Systems Analyzers. PMA # P930027/ Device Listing # D002374. |  |  | June 26, 2013 | 2013 | Siemens Healthcare Diagnostics | Customers complained of a positive bias in patient results on the IMMULITE/ IMMULITE 1000 and the IMMULITE 2000/ 2000 XPi PSA assay when compared against the IMMULITE /IMMULITE 1000 3rd Generation PSA assay. Siemens has confirmed an overall average positive bias of approximately 20% to 23% across the assay range relative to the WHO 96/670 with the IMMMULITE Systems PSA Assay. This positive bias i | Software design | Seimens sent an Urgent Medical Device Recall Letters to US consignees via FedEx and an Urgent Field Safety Notices were sent to the foreign consignees starting on June 26, 2013. In addition, Support Bulletins were distributed to Siemens Healthcare Diagnostics service personnel describing the issue and instructing them how to deal with customer questions. Customers were informed to discontinue use, discard the kits remaining in inventory and there are no replacement PSA kits available at this time. Siemens is instructing customers to contact their local Siemens representative for assistance with determining appropriate PSA testing solutions for your laboratory. |  | 24361 | 24,361 kits | Worldwide Distribution - USA Nationwide including Puerto Rico and the countries of: Albania, Afghanistan, Algeria, Argentina, Austria, Bangladesh, Belgium, Bolivia, Bosnia- Herzegovinia, Brazil, Bulgaria, Camaroon, Canada, Canary Islands, Chile, China, Colombia, Coast Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, El Salvador, Estonia, Finland, France, Germany, Great Britain, Greece, Guatemala, Honduras, Hungary, India, Iran, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Latvia, Lebanon, Lithuania, Macedonia, Malaysia, Malta, Mexico, Morocco, Nicaragua, Oman, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Romania, Russian Federation, Saudi Arabia, Serbia, Slovakia, Slovenia, Spain, South Africa, Rep. of Korea (S), Sweden, Switzerland, Syria, Taiwan, Thailand, The Netherlands, Tunisia, Turkey, Turkmenistan, Ukraine, United Arab Emirates, Uruguay, White Russia, Venezuela, and Vietnam. | Open, Classified |  |  | 24361 | Computer
Z-0972-2013 | 64285 | No | HTO | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | ReClaim Reamer Extension; Product code 2975-00-500, used with Reclaim Modular Revision Hip Prosthesis System. The ReClaim Reamer Extension is as an attachment to the distal reamers to allow the surgeon to clear soft tissue while reaming the distal femoral canal. |  |  | February 25, 2013 | 2013 | DePuy Orthopaedics, Inc. | The tabs on the reamer extension device have broken. DePuy is notifying surgeons of the issue to increase the awareness and to ask surgeons and users to look for tabs before and after surgery to reduce any potential that the tabs are left in the patient. Once the root cause is determined and a design change is implemented, DePuy will conduct a formal swap-out of the affected devices. | Device Design | DePuy Synthes sent a Urgent Information Medical Device Correction Notice to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed upon visual inspection, if the reamer extension is found to have one or more tabs broken, please do not use the reamer extension. The reamer extension will still engage with one tab intact and no clinical implications have been identified with the use of the reamer extension with one tab, but DePuy recommends it be returned and the alternate reamer extension device be used.If both tabs are broken, the reamer extension will not engage and DePuy recommends it be returned and the alternate reamer extension device be used. A reamer extractor is available in the instrument kit for removal of the reamer if necessary.In order to reduce the risk of breaking the tabs during surgery, the Reamer Extension should only be used to drive the Distal Reamer axially within the patients femoral canal. The instrument is NOT intended to be used for side-cutting or lateralizing the femoral canal, therefore, no off-axis force should be applied during distal reaming.The purpose of this communication is to inform you of this device correction and request acknowledgement of receipt of this letter by signing and returning the Surgeon Response Card within 14 days of receiving this notice. Response cards should be faxed to Fax Number: 574-371-4939.For product-related questions, please contact your local DePuy Orthopaedics Sales Representative. For clinical-related questions from surgeons, please have them contact DePuys Scientific Information Office at 1-888-554-2482 (M-F; 8 am to 5 pm EST.) For questions about device correction information provided, 574-372-7333 (M-F: 8 am to 5 pm EST.) |  | 750 | 750 | Worldwide Distribution - USA including the states of ME, MA, PA, MD, FL, LA, IL, IA, OH, MN, TX, MT, WA, CA, HI, NV, AZ, KY, WI, CO, GA, MI, TN, NY, AR, NC and VA and the countries of Canada, Chile, Germany, Czech Republic, Israel , Norway, Poland, Slovenia, Australia, New Zealand, South Africa, Belgium, Austria, Switzerland, France, England and County Cork, Ireland. | Terminated | December 30, 2013 | 308 | 750 | Computer
Z-0977-2013 | 64112 | No | CFR | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | BS-200 Chemistry Analyzer |  |  | November 28, 2012 | 2012 | Mindray DS USA, Inc. d.b.a. Mindray North America | An issue has been identified with the BS-200 where clinical chemistry results may be inaccurate. | Device Design | Mindray North America /Mindray DS USA Inc. sent Urgent Product Corrective Action letters via certified mail, return receipt requested, signature required on 11/28/2012. The letters identified the affected product as well as the issue, effect on patients, and the corrective action. Customers are to remove the analyzer from service until the affected pump has been replaced. Also, customers are to contact Mindray Customer Service Representative Tony Li at 1-800-288-2121 x2112 for technical assistance, and to arrange for replacement pumps at no charge. |  | 3 | 3 units | Distributed in the states of NC and TX. | Terminated | June 12, 2013 | 196 | 3 | Computer
Z-1989-2013 | 65862 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Precision 500D |  |  | May 02, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue involving the overhead Video Monitor Suspension involving Precision 500D Systems. There is a potential for the locking nut to loosen out of the Video Monitor Suspension yoke which could result in impact. This will not lead to an immediate fall of the suspended monitor since it will be held in place by the remaining bushing and | Device Design | GE Healthcare "Urgent Medical Device Correction" letter dated May 1, 2013 to all affected customers. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. The letter informed customers that GE Healthcare will correct all affected systems at no cost. A GE Healthcare service representative will contact you to arrange for this correction. For questions or concerns regarding this notification call: US 800-437-1171, Japan: 0120-055-919. |  | 445 | 445 (421 US, 24 OUS) | Worldwide Distribution - USA Nationwide in the states of AL, AK, AR, CA, CO, CT, FL, GA, HI, ID, IL IN, IA, KS, KY, LA, MD, MA, NMI, MN, MS, MO, NE, NV,NH,NJ, NM, NY, NC,ND, OH, OK OR, PA,SC, SD, TN, TX, UT, VA, WA, WV, WI, WY and the countries of CANADA, FRANCE GERMANY, GREAT BRITAIN ITALY, PANAMA, SAUDI ARABIA. | Terminated | November 26, 2013 | 208 | 445 | Computer
Z-0994-2013 | 64674 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Toshiba Infinix-i X-Ray System. A radiographic and fluoroscopic x-ray system for use in diagnostic and angiographic procedures. |  |  | November 08, 2012 | 2012 | Toshiba American Medical Systems Inc | Lost and incorrect images. When using the "F-Rec" fluoroscopy image acquisition recording option in the "manual mode" an infrequent software error may cause images from another study to be displayed on the monitor. The acquired images may be lost requiring additional image acquisitions to be performed. | Software design | In November, 2012, Toshiba America Medical Systems, Inc. sent an Urgent: Medical Device Correction letter to customers. The letter contains a return reply form that is to be faxed to TAMS for retention. |  | 95 | 95 units | USA nationwide. No international. | Terminated | March 26, 2013 | 138 | 95 | Computer
Z-0992-2013 | 64684 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Siemens Axiom Artis zeego systems with software version VC14, VC20 and VC21 in conjunction with specific technical configuration |  |  | November 26, 2012 | 2012 | Siemens Medical Solutions USA, Inc | In the course of the firm's product monitoring activities, a potential risk for patient or operators was determined during the operation of an Artis zeego system with software revision VC14, VC20 and VC21 in conjunction with a specific technical configuration which cannot be completely executed. | Software design | Siemens sent an Urgent Customer Information letter dated November 8, 2012 to affected consignees. The letter notified consignees of the issue and informed them the update AX039/12/S would be started and affected systems would be checked and corrected if required. Customers were instructed that they would be contacted to arrange a date for checking their system.For questions regarding this recall call 610-219-6300. |  | 52 | 52 | Nationwide Distribution including AL, AR, AZ, CA, CT, DC, FL, GA, IA, IL, LA, MD, MI, MO, NC, NH, OR, PA, SD, TN, TX, WA and WI and Puerto Rico. | Terminated | December 24, 2013 | 393 | 52 | Computer
Z-0991-2013 | 64394 | No | LTW | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | VITEK 2 Compact computer system when equipped with the following Hewlett-Packard PC models: |  |  | February 18, 2013 | 2013 | Biomerieux Inc | There is a potential to link an isolate result to the wrong patient and then upload the results to the LIS. | Software design | bioMerieux sent an "URGENT PRODUCT CORRECTION NOTICE" dated February 20, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact Emails were issued by the recalling firm on 2/18/2013 and 2/19/2013 to their subsidiaries and to distributors. Letters began issuing to customers via e-mail on 2/20/2013. Contact bioMerieux Clinical Customer Service at (800) 682-2666 for questions regarding this notice. |  | 17293 | For all 3 computers types: 16,736 originally distributed; 554 refurbished and redistributed | Worldwide Distribution-USA (nationwide) and the countries of Algeria, Angola, Argentina, Australia, Austria, Azerbaidjian, Bahrain, Belarus, Belgium, Benin, Bolivia, Bosnia, Brazil, Bulgaria, Burkina Faso, Camaroun, Canada, Chad, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Democratic Republic of Congo, Demark, Dominican Republic, Dutch Antilles, Ecuador, Egypt, El Salvador, Ethiopia, Finland, France, Gabon, Georgia, Germany, Greece, Guatemala, Guinea, Herzegovina, Honduras, Hungary, India, Indonesia, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Korea, Kosovo, Kuwait, Latvia, Lebanon, Libya, Lithuania, Macedonia, Malaysia, Mali, Mexico, Moldavia, Mongolia, Morocco, Myanmar, Netherlands, Netherlands, Nicaragua, Niger Niamey, Nigeria, Norway, Oman, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Puerto Rico, Qatar, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Thailand, Tunisia, Turkey, Uganda, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Venezuela, Vietnam, and Yemen. | Terminated | August 14, 2014 | 542 | 17293 | Computer
Z-0997-2013 | 64662 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Artis Systems with Software Artis VC20x/VC21A/VD10x that have a DSA license installed. |  |  | March 08, 2013 | 2013 | Siemens Medical Solutions USA, Inc | The firm has become aware of a potential issue on Artis systems with the software Artis VC20x/VC21A/VD10x that have a DSA license installed. Under certain preconditions, executing roadmap OGP (organ program) in the DSA overlay mode may lead to imprecise registration of the DSA mask image with the roadmap image displayed on the live monitor. | Software design | Siemens sent a Customer Safety Advisory Notice on March 8, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to pass the information on to all those who need to be aware of it and to instruct their personnel accordingly. Customers were asked to forward the safety information to other organizations that couald be affected by this action.For questions regarding this recall call 610-219-6300. |  | 81 | 81 | Nationwide Distribution including AL, CA, CO, FL, IA, IL, IN, KY, MA, MD, ME, MI, MN, MO, NC, ND, NJ, NY, OH, PA, SC, TX, UT, VA, WA, and WI and to Puerto Rico. | Terminated | November 06, 2014 | 608 | 81 | Computer
Z-0998-2013 | 64609 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | STERRAD 200 System Carriage, P/N 10208. |  |  | March 13, 2012 | 2012 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) is recalling the STERRAD 200 System Carriage because it can potentially become dislodged when an excessive load is applied. | Device Design | Advanced Sterilization Products (ASP) sent an "URGENT MEDICAL DEVICE NOTIFICATION" letter dated March 13, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. The letter informed the customers that ASP recently determined that the push bar located on the STERRAD 200 System Carriage can potentially become dislodged when an excessive load is applied. Customers with questions are instructed to contact Stericycle at (877) 931-0642. |  | 2457 | 2,457 units | Worldwide Distribution-USA (nationwide) and the countries of Egypt, Germany, Italy, France, Saudi Arabia, UAE, Japan, Singapore, China, Thailand, Hong Kong, Hungary, Israel, Slovania, Poland, Czech Republic, Taiwan, Sweden, Greece, Korea, Ireland, and United Kingdom. | Terminated | May 28, 2014 | 806 | 2457 | Computer
Z-1977-2013 | 65380 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | 100/120V AC, Advanced Perfusion System 1 |  |  | July 26, 2013 | 2013 | Terumo Cardiovascular Systems Corporation | TCVS is taking action to update the Operators Manual with the specific instructions to respond to this condition and will issue a labeling addendum for current users. In the event of a false high pressure alarm condition occurs that displays a "999" pressure reading and results in stopping a pump, this pressure display can be cleared and the pump re-started by disconnecting the pressure transduce | Software design | Terumo sent an Urgent Medical Device Recall notification letter dated July 26, 2013 to all affected customers. The letter identified the affected product, reason for correction, potential hazards of the recall issue, the correction to be implemented, and customer instructions for replying to the recall notification. Terumo CVS is providing an addendum to be placed in the Operator's Manual at the beginning of Chapter 7, Modules. Questions or concerns are directed to Terumo CVS Customer Service at 1-800-521-2818, Recall Fax 1-877-275-6860. Customer Service hours are Monday to Friday 8am-6pm ET. |  | 1657 | 1657 | Worldwide Distribution - US Nationwide in the states of :AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, MSMT, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WI, WV and countries of AUSTRALIA, UNITED ARAB EMIRATES (UAE),Hong Kong,Indonesia, Singapore, Taiwan, Thailand,Philippines, CHILE,Malaysia,South Korea, BELGIUM,Japan,Mexico,CANADA | Open, Classified |  |  | 1657 | Computer
Z-1975-2013 | 65719 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Siemens syngo.plaza SW version VA20 with Mammography license. |  |  | March 04, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Siemens issued a Customer Safety Notice to inform customers of a potential problem when using syngo.plaza. In some instances, it can happen that the laterality information in mammography images is not displayed and these images may not be properly aligned to mammography segments. | Software design | The firm, Siemens, notified their customers on March 4, 2013 and sent a "Customer Safety Advisory Notice" letter dated April 03, 2013. The letter described the product, problem and actions to be taken. The customers were instructed to always check whether the laterality information is displayed in the image text of mammography images; if not, the DICOM header is accessible and can be checked for the tag "Laterality" (0020, 0060). If a system is not able to provide the DICOM tag "Image Laterality" (0020, 0062), configuring the image text by adding the tag was proposed. If you have sold the device/equipment and it is no longer in your possession, forward this safety notice to the new owner of the device/equipment and inform us about the new owner of the device/equipment. This issue will be resolved in the syngo.plaza software version VA20D_HF04 which is planned to be released in 02/2013. Customers were also instructed to contact the local Siemens service office to plan the update.If you have any questions, contact the SYNGO CEO or Vice President Quality at +49 (9131) 84-4874 or email to: michael.mlb.braun@siemens.com. |  | 6 | 6 | Nationwide distribution: FL, MA, KS, NY, OH, PA, TX and WI. | Terminated | March 09, 2015 | 735 | 6 | Computer
Z-1969-2013 | 65354 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | TerumoÂż Advanced Perfusion System 1: 220/240V AC, Advanced Perfusion System 1 |  |  | July 26, 2013 | 2013 | Terumo Cardiovascular Systems Corporation | The operators manual does not fully describe the sequence of expected events that will occur when a specific operating condition exists. Specifically, if the user is adjusting the System 1 centrifugal pump speed via the Central Control Monitor at the same time an alarm condition occurs with configuration for a pump coast response, the labeling in the operators manual does not fully describe how th | Software design | Terumo CVS, sent an URGENT MEDICAL DEVICE RECALL notification letter dated July 26, 2013, to all affected customers explaining the correction to the operators manual, potential hazards and what to do in response to the issue.If you have any questions or concern please contact Terumo CVS Customer Service 1-800-521-2818 Recall Fax 1-734-741-6149 Customer Service Hours:Monday  Friday, 8 AM - 6 PM ET |  | 1657 | 1657 | Worldwide Distribution - USA (nationwide) including the countries of Australia, Dubai, Hong Kong, Indonesia, Singapore, China, Mexico, British Columbia, Taiwan, Thailand, Chile, Malaysia, Japan, Belgium, United Arab Emirates (UAE), South Korea, Manila and Canada. | Open, Classified |  |  | 1657 | Computer
Z-0630-2013 | 63616 | Yes | MSX | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | DPM Central Station |  |  | May 12, 2011 | 2011 | Mindray DS USA, Inc. d.b.a. Mindray North America | There is a software anomaly with the DPM Central Monitoring System where the trend data of one patient maybe replaced with the trend data of a second patient. | Software design | Mindray sent a Product Corrective Action letter dated May 12, 2012, to all affected customers via certified mail return receipt requested to two medical facilities. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed that the issue would be corrected through a softweare upgrade of the affected product. A Mindray Service Representative would contact their facility to arrange for the software upgrade at no cost to them.For questions regarding this recall call 1-800-288-2121. |  | 2 | two units | Nationwide Distribution including OH and GA | Terminated | February 19, 2013 | 649 | 2 | Computer