## Sheet1
Recall Number | Recall Event ID | Device Type | Product Code | Regulation Description | Medical Specialty | Review Panel | Submission Type | Trade Name/Product | Main Name | Recall Class | Date Posted | Year Posted | Recalling Manufacturer | Reason for Recall | FDA Determined Cause | Action | Consumer Instructions | Clean Quantity | Quantity in Commerce | Distribution | Recall Status | Termination Date | Time to Terminate | Merged Quantities | Fault_Class | 0
Z-2132-2014 | 68766 | No | HQF | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | MC-500 Multicolor Laser Photocoagulator; Distributed by NIDEK Inc. Manufactured by NIDEK CO., Ltd, Japan. For use in ophthalmic and retinal procedures. |  |  | June 09, 2014 | 2014 | Nidek Inc | Laser aperture label was not affixed to device prior to shipment in the US. | Error in labeling | Nidek sent a Recall/Field Correction letter dated June 9, 2014, to all affected customers. The letter advised customers of the issues and informed customers that Nidek Representatives will visit each location to affix the proper labeling to the device. Customers with questions were instructed to contact Nidek Customer Service at 1-800-223-9044 M-F 8 am to 5 pm PST.For questions regarding this recall call 510-353-7785. |  | 29 | 29 units (26 already re-labeled) | Nationwide Distribution including IL NE NY, NH, CD, MI, TX, CA, PA, FL, KS, UT, MN, WI,. MO. and OR. | Terminated | November 24, 2014 | 168 | 29 | Not_Computer | Not_Computer
Z-0019-2015 | 67029 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | ArjoHuntleigh Contoura C1000 series beds are designed to accommodate patients weighing up to 990 lbs. (450 kg). Model C1000 is the basic bed. Model C1080 includes a weighing system and a power outlet for an ArjoHuntleigh approved mattress pump. |  |  | March 17, 2014 | 2014 | Arjo Hospital Equipment AB | There is a low but existing entrapment risk for a patient being treated on the Contoura 1000 and Contoura 1080 bariatric beds if they have a BMI &lt; 40. | Labeling design | ArjoHuntleigh sent an URGENT FIELD SAFETY NOTICE - IMMEDIATE ATTENTION REQUIRED letters dated March 17, 2014 to all affected cusomers. The letters included instructions for customers to: 1) ensure that the Contoura Bed Field Safety Notice brought to the attention of all caregivers and clinical staff and they are aware that the Contoura 1000 or 1080 medical bed is intended to be used as a bariatric medical bed, designed for the patients with BMI greater than 40; 2) ensure that all users of the Contoura 1000/1080 bed, applicable staff and caregivers are trained to ensure awareness of the revised Instructions for Use Manual and new bed label; 3) discard all previous versions of Contoura 1000/1080 Bed Instructions for Use Manuals and ensure that a copy of this Contoura Bed Field Safety Notice is placed in the enclosed revised Contoura 1000/1080 bed Instructions for Use Manual (746-173-AM-3 - 10/2013); 4) where applicable, include this requirement in subsequent training of applicable staff and caregivers; and, 5) complete and sign the enclosed Customer Response Form and return this form to the ArjoHuntleigh Regulatory Department. An ArjoHuntleigh Service Technician will contact the designated facility representative listed on the Customer Response Form to schedule service to apply the two labels shown above, free of charge. For question contact ArjoHuntleigh at 1-800-323-1245, ext. 57985. |  | 1695 | 1,695 beds (worldwide) | Worldwide Distribution - US: Nationwide in the states of (AZ, CA, CO, CT, GA, ID, IL, MA, MD, MN, NC, NE, OH, PA, RI, SD, TX, VA, WI); and the countries of Australia, Austria, Bahrain, Belgium, Canada, Denmark, France, Germany, Greece, Iceland, India, Ireland, Israel, Italy, Lebanon, Mexico, Netherlands, New Zealand, Norway, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, United Arab Emirates | Open, Classified |  |  | 1695 | Not_Computer | Not_Computer
Z-2143-2014 | 68647 | No | IWZ | Radiographic film. | Radiology | Radiology | 510(K) Exempt | Ultra-speed Dental Film, Catalog Number/REF 834 8658, KODAK and Carestream DENTAL labels --- Carestream Health Inc. |  |  | July 02, 2014 | 2014 | Carestream Health Inc. | The Indicator Dot on the film packet's outer printed paper labeling was incorrectly positioned relative to the location of the raised dot on the dental intraoral film. This may cause confusion for the dental practitioner. If the processed film is placed on the view box by referencing the film's raised dot, the image orientation will be rotated 180 degrees. | Package design/selection | Recall notices were mailed to US Customers via overnight carrier on 07/02/2014. Customers are asked to complete an inventory of goods to be returned which will be ultimately destroyed by recalling company. If you have any questions or concerns, please contact the Carestream Customer Care Center in the U.S. at 1-800-933-8031. |  | 1633 | 1633 boxes | Worldwide Distribution. | Open, Classified |  |  | 1633 | Not_Computer | Not_Computer
Z-1317-2014 | 67734 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Citation TMZF HA 132 degrees Neck Angle, V40 Hip Stem - Right Size #3: catalog # 6265-5113 , Left Size #4: catalog #6265-5104 , and Right Size #4: catalog #6265-5114; Stryker Howmedica Osteonics Hip Implants. |  |  | February 03, 2014 | 2014 | Stryker Howmedica Osteonics Corp. | Stryker received a report from the field indicating that a citation TMZF HA 132 degrees neck angle V40 Hip Stem - Left Size # 4 was identified in a packaging associated with a citation TMZF HA 132 degrees Neck Angle V40 Hip Stem - Right Size #3. | Error in labeling | Stryker sent an Urgent Medical Device Recall Notification letters and Product Accountability Forms to Hospital Risk Management, Chief of Orthopaedics, and Surgeons via Fed Ex with return receipt on February 26, 2014. The letter identified the affected product, problem and actions to be taken. Customer were instructed to return the affected product to the attention of Regulatory Compliance. Please attach the fluorescent orange PRODUCT REMEDIATION sticker to your return. For questions call 201-972-2100. |  | 48 | 48 units | USA Nationwide Distribution | Terminated | December 02, 2014 | 302 | 48 | Not_Computer | Not_Computer
Z-2047-2014 | 68565 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | SUTUREFIX Ultra S with 1 #2 Ultrabraid (blue) |  |  | June 02, 2014 | 2014 | Smith &amp; Nephew, Inc. Endoscopy Division | Sterility of device is compromised due to breach in sterile barrier | Package design/selection | Smith &amp; Nephew issued an Urgent - Product Recall Notification letter dated June 2, 2014 via Federal Express to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to inspect their inventory and locate all affected product and quarantine them immediately.Complete the Return Certification Form .Contact the Smith &amp; Nephew Returns Group at 800-343-5717 (option 3) or send an email to ProductRecoverv@smith-nephew.com. They will provide customers with instructions for returning their product and receiving a replacement.For questions regardin g this recall call 978-749-1073. |  | 907 | 907 units | Worldwide Distribution - USA (nationwide) and Internationally to Australia, Belgium, Denmark, France, Germany, Great Britain, Italy, Netherlands, Norway, Portugal, Spain, and Sweden. | Open, Classified |  |  | 3151 | Not_Computer | Not_Computer
Z-1337-2014 | 67665 | [6 empty] | Syphilis Test Strip (Serum) |  |  | March 04, 2014 | 2014 | LumiQuick Diagnostics Inc. | Products shipped to US locations that do not have approved applications for PMA , IDE or 510(k) clearance. | No Marketing Application | Urgent Device Recall notification letters were sent to all US consignees on March 4, 2014 by e-mail and US mail as well as a phone call. The letter identified the products and actions to be taken by the customers. |  |  | ALL | Distributed in the states of CA, FL, NJ, TX, and NY. | Terminated | May 09, 2014 | 66 |  | Not_Computer | Not_Computer
Z-1417-2014 | 67757 | No | GCY | Nonpowered, single patient, portable suction apparatus. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | 7mm Drain Kit, Flat Full Perforated, Saber 100ml Bulb and 3C System Sterile (10/case) |  |  | March 20, 2014 | 2014 | Aspen Surgical Products, Inc. | The recall has been initiated due to a potential compromise of the sterile barrier of the Surgidyne Wound Drainage and Surgidyne Bulb Evacuator &amp; Collection System medical devices. | Packaging process control | Aspen Surgical sent an Urgent Medical Device Recall letter dated March 20, 2014, to all affected customers and will be carried out at the distributor and retailer level, which can include doctor's offices, hospitals, and clinics. The distributor was instructed to contact their local FDA District office, conduct a sub-recall and to forward this Recall Notification and Response form to their retailers.A written notification to the consignees was sent by certified mail.Aspen has given the customers the option to either 1.) Return the product 2.) Destroy the product with a Certificate of Destruction to be provided to Aspen. Any returned product to Aspen will be put on MRB and dispositioned by Quality. The returned product will not be reprocessed.In order to determine effectiveness of the recall, any consignee that does not return the response form and/or product, will be contacted within three weeks of receipt of the Recall letter. Non-responders will be contacted a total of three (3) times in order to verify the effectiveness of the recall activity.Customers with questions were instructed to call 888-364-7004, ext 177 or 1-888-364-7004, ext. 156.For questions regarding this recall call 616-698-7100. |  |  |  | Worldwide Distribution - USA including LA, MA, NH, TX, FL, OH, NC, PA, CA, UT, TN, MI, CO, WY, OK, WA, MN, PA, IN, VA, IA, IL, MD, CT, DE, SC, NY, AR, MA, AZ, NJ, MO, GA, KS, WV, NE, NC, Puerto Rico, and Internationally to Canada, UAE, England, Italy, Belgium, Costa Rica, and Ecuador. | Terminated | November 25, 2014 | 250 |  | Not_Computer | Not_Computer
Z-0488-2015 | 69642 | No | HHK | Endometrial suction curette and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Cooper Surgical CANNULA-CURETTE 12MM; Model Number: MX562 |  |  | October 27, 2014 | 2014 | CooperSurgical, Inc. | The product has been identified to contain an incorrect curette type. | Packaging process control | CooperSurgical sent an Recall letter dated October 27, 2014, via UPS to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The subject device does not meet the required release specification as it contains an incorrect curette type. No adverse events or injuries have been reported to date in connection with the use of the CANNULA CURETTE. The event was discovered during internal investigation. Products that do not meet the above criteria are safe to use alongside the DFU guideline.The letter requests that customers return the recalled product for a full refund or exchange. If you have any further questions please feel free to call (203) 601.5200. |  | 30 | 30 | US Distribution including the states of NY, MO and NV. | Terminated | May 28, 2015 | 213 | 30 | Not_Computer | Not_Computer
Z-2152-2014 | 68749 | No | MTF |  |  | Immunology | PMA | VIDAS TPSA, in vitro diagnostic, for use as an automated quantitative measurement of prostate specific antigen levels in human serum or plasma. |  |  | October 14, 2013 | 2013 | BioMerieux SA | Product is not approved or authorized for distribution in the US. | Packaging process control | A letter was issued on October 14, 2013 to consignees requesting the consignee to check their stock and destroy all kits of the VIDAS TPSA ref. 30428. The letter also requested that the consignee notify their accounts of the recall. An acknowledgment form was also included with the letter which is to be returned to BioMerieux. |  | 1902 | 1902 | FL, LA, MI, MS, NV, OH, OK, TN, TX | Completed |  |  | 1902 | Not_Computer | Not_Computer
Z-2041-2014 | 68468 | No | JSS | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | BBL CrystaLEAD Enteric/Nonfermentor (E/NF) Identification (ID) System, catalog number 245000. in vitro diagnostic for the identification of aerobic gram-negative bacteria that belong to the family Enterobacteriaceae as well as some of the more frequently isolated glucose fermenting and nonfermenting gram-negative bacilli. |  |  | May 21, 2014 | 2014 | Becton Dickinson &amp; Co. | A portion of the BBL Crystal Enteric/Nonfermentor (E/NF) Identification kits from lot numbers 3224281 and 3290379 are labeled with incorrect expiration dates. A portion of these kits contain Crystal Inoculum fluid bottles that will expire prior to the date that was printed on the kit carton. The expiration date printed on the individual Crystal Inoculum fluid bottles is correct. | Labeling mix-ups | BD Diagnostic sent an Urgent Product Recall letter dated May 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed:Please discard any outer kit boxes that are labeled with the incorrect expiration date. All components of the kit will perform as expected with the expiration dates on each component.Our records indicate you may have been shipped a portion of the above-referenced product between August 23, 2013, and the present day.Please complete the attached form whether or not you have any inventory remaining so that we may acknowledge your receipt of this notification. It is not necessary to contact BD by telephone. Simply complete and fax the form to: Attention: Regulatory Compliance, Fax: 410-316-4258.If further assistance is needed, please contact the BD Customer Service Department at 1-800-675-0908. For all other inquiries, please contact the BD Technical Services Department at 1-800-638-8663.Please accept our apology for any inconvenience this may cause. BD is committed to providing you with the highest quality products. Thank you for your continued support. : -"/ 06 . ;/ i )!.U |  | 18.21 | 18.210 | Worldwide Distribution - US (nationwide), Australia, Belgium, Canada, Colombia, China, Indonesia, India, Japan, Korea, Mexico, Malaysia, Peru, Philippines, Singapore, Thailand, and Taiwan. | Terminated | August 12, 2016 | 814 | 18.21 | Not_Computer | Not_Computer
Z-0484-2015 | 69636 | No | HGI | Gynecologic electrocautery and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | CooperSurgical LEEP RADIUS LOOP ELECTRODE (W 2.0 cm x D 1.0 cm D, 12 cm shaft ). Box of 5 |  |  | October 27, 2014 | 2014 | CooperSurgical, Inc. | Outer box is mislabeled as P/N R2010- Large Radius Loop Electrode instead of P/N R1010- Medium Radius Loop Electrode | Packaging process control | CooperSurgical issued recall letter dated October 27, 2014, to all affected customers via certified mail delivery (UPS). Acccounts were requested to return product for replacement and complete the return response form. Customers with questions were instructed to call 203-601-5200. |  | 595 | 595 units | Worldwide Distribution - USA (Nationwide) and Internationally to Canada, Portugal, and Hong Kong. | Terminated | May 29, 2015 | 214 | 595 | Not_Computer | Not_Computer
Z-2151-2014 | 67640 | dam, rubber | EIE | Rubber dam and accessories. | Dental | Dental | 510(K) Exempt | Latex Dental Dams labeled as Non-Latex Dental Dams |  |  | December 11, 2013 | 2013 | Coltene Whaledent Inc | ColtÂżne/Whaledent Inc. received a complaint about the dental dams being the incorrect color. During an internal investigation, the firm found that Latex dental dams had been packaged in Non-Latex packaging. | Packaging change control | The firm issued a press release, called the distributor and sent a certified letter. The firm obtained the distributor's distribution list and sent recall notifications to all of these consignees.The firm included envelopes and return/reply cards, via response form mailed, certified mailing with return receipt. |  | 480 | 32 boxes, 15 individual dams per box | Distributed in Memphis, TN. | Open, Classified |  |  | 480 | Not_Computer | Not_Computer
Z-0076-2015 | 69290 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Multi-Lumen Central Venous Catheterization |  |  | September 18, 2014 | 2014 | Arrow International Inc | Arrow International, Inc. has initiated a voluntary recall for product code AK-12703, lot number RF3039028 because the label on the outer corrugate is labeled with the incorrect expiration date. All kits packaged within the corrugate are correctly labeled. There is no risk to the patient as the kits within the corrugate are labeled with the correctexpiration date. | Labeling False and Misleading | Arrow International sent an Medical Device Advisory Notifications letter dated September 18, 2014, sent to direct accounts to notify them about the product, problem, and actions to be taken.Arrow International, Inc. ( Arrow ) is notifying customers that the label on the outer corrugate for the above mentioned product code and lot number is labeled with the incorrect expiration date. All kits packaged within thecorrugate are correctly labeled. There is no risk to the patient as the kits within the corrugate are labeled with the correct expiration date.Arrow International, Inc. is committed to providing high quality, safe and effective products. We sincerely apologize for any inconvenience this action may cause your operations. If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-866-246-6990. |  | 65 | 65 | US Distribution including the states of MD, NC , LA , IN , OH, KY, CA , PA , NY and NV. | Terminated | December 02, 2014 | 75 | 65 | Not_Computer | Not_Computer
Z-1005-2014 | 66500 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | VERSA-FX II KEYLESS TUBE, VERSA-FX II KEYLESS SHORT, VERSA-FX II SHORT TUBE PL, VERSA-FX II STD TUBE PLAT and VERSA-FX II SUPRACONDYLR |  |  | October 11, 2013 | 2013 | Zimmer, Inc. | Zimmer recently conducted a review of all historic packaging validations at its Zimmer Manufacturing B.V. (ZMBV) facility in Ponce, Puerto Rico. Based on this review, it was determined that the packaging operations conducted in the Building 1 packaging operation at this facility were not properly validated. As a result, Zimmer conducted seal integrity testing to evaluate if previously packaged product from this facility met the established requirements for sterile barrier integrity. The acceptance criteria established for this testing was 99% reliability with a 90% confidence interval. These packaging tray configurations met the acceptance criteria for sterile barrier integrity of both the inner and outer trays, however, Zimmer has decided to recall all manufacturing lots that have not expired and that were sterile packaged in the ZMBV facility Building 1 packaging operations. | Packaging | Distributors received a letter identifying the issue and their responsibilities dated November 4, 2013. These responsibilities include quarantining the affected fielded inventory in their possession, providing hospital risk managers with the recall notification letter and retrieving affected fielded inventory and direct sales from the hospitals. |  | 7044680 | 7,044,680 total | Worldwide Distribution-USA (nationwide) and the countries of Belgium, Canada, Canary Islands, China, Czech Republic, Denmark, Finland, France, Germany, Greece, India, Italy, Japan, Korea, Luxemburg, Netherlands, New Zealand, Norway, Portugal Reunion, Russian Federation, Singapore, Slovakia, Spain, Sweden, Switzerland, Turkey, Taiwan, Thailand, and United Kingdom. | Terminated | April 14, 2017 | 1281 | 7044680 | Not_Computer | Not_Computer
Z-1121-2014 | 67490 | No | MAV | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Everest 20cc Inflation Device; Model Number and Catalog Number: AC2205P |  |  | February 06, 2014 | 2014 | Medtronic Vascular | Incorrect master label. | Labeling mix-ups | Medtronic sent an Urgent Medical Device Recall letter on February 6, 2014, via UPS 3-day delivery. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to remove the affected product that remains in their inventory. Return the affected product to Medtronic: Contact Customer Service at 888-283-7868 and reference RADAR # 170276 to initiate a return and credit of product. Complete the attached Customers Confirmation Certificate and fax it to Medtronic at 651-367-0612 to the attention of Customer Focused Quality. If any affected product has been forwarded to another facility, customers were asked to notify that facility accordingly and facilitate the retrieval of the affected product. Customers with questions were instructed to contact their Medtronic representative.For questions regarding this recall call 888-283-7868. |  | 191 | 191 devices | Nationwide Distribution including TX, ME, IA, PA, OR, and NH. | Terminated | June 02, 2014 | 116 | 191 | Not_Computer | Not_Computer
Z-2145-2014 | 68745 | No | HET | Gynecologic laparoscope and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Instructions for Use for the blue endo MOREsolution Electromechanical Laparoscopic Morcellator System, Model 55-000011, consisting of the control unit, hand-grip, flexible shaft, trocars/obturators, sleeves, valves, and blades. |  |  | July 16, 2014 | 2014 | Market-Tiers Inc | Some pages of the Instructions for Use were omitted, including pages regarding the Indications for Use, Warning, and Contraindications. | Labeling Change Control | The recalling firm issued a letter dated 7/16/14 via Federal Express requesting the customer destroy the incorrect IFU and replace it with the corrected version. |  | 100 | 100 units | Nationwide Distribution. | Terminated | May 11, 2015 | 299 | 100 | Not_Computer | Not_Computer
Z-1576-2014 | 68143 | No | EIA | Dental operative unit and accessories. | Dental | Dental | 510(k) | Enaly 1000 BT-12 Ozone Generator, Model # 1000BT-12 |  |  | April 23, 2014 | 2014 | William C. Domb, D.M.D., A Professional Corporation | William Domb is recalling the Enaly 1000 BT-12 Ozone Generator because it is not approved or cleared by the FDA for medical use. | No Marketing Application | William Domb, DMD sent an Urgent Medical Device Recall letter on April 23, 2014, to all affected customers who purchased the Enaly 1000BT-12 Ozone Generator. The letter informed the customers of the problems identified, risk to health, and the actions to be taken. Customers were instructed to discontinue use of the Enaly 1000 BT-12 Ozone Generator and return their units to the following address:William Domb, DMD190 N Mountain AvenueUpland, CA 91786Customers were also instructed to complete the Medical Device Field Action Return Response form. For questions customers should contact William Domb at (909) 949-2968 or by email at bill@smilestudio.com from Monday-Friday at 8am-5pm Pacific Time.For questions regarding this recall call (909) 949-2968. |  | 16 | 16 units | Worldwide Distribution - USA including IN, HI, WA, TN, CA, TX, FL, NC, and Internationally to Canada, Portugal, and Ireland. | Terminated | January 08, 2015 | 260 | 16 | Not_Computer | Not_Computer
Z-2124-2014 | 68744 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | DeRoyal KNEE ARTHROSCOPY PACK, REF 89-6192.04, 1 Per Pack, Distributed by: DeRoyal Industries, Inc.; Manufacturer: DeRoyal Industries, Inc. |  |  | June 04, 2014 | 2014 | DeRoyal Industries Inc | The firm distributed surgical kits which contained Irrigation Sets which were subsequently recalled by Hospira. | Packaging change control | The recall was initiated by letter delivered via UPS on 6/4/2014. The firm requested that the consignee contact Stericycle to arrange for return and replacement of the product. A second notice was sent on 7/22/2014 to those accounts who did not respond to the initial notice. Distributors were requested to notify their end-users or provide a list of end-users to DeRoyal for direct notification. Effectiveness checks will be sent via email for those customers who have provided an email contact and via UPS for those customers who have not responded. |  | 128 | 128 units | Distributed in the states of MI, GA, OH, IL, MO, and IN. | Terminated | May 05, 2015 | 335 | 358 | Not_Computer | Not_Computer
Z-0470-2015 | 69839 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Non-sterile drape |  |  | November 04, 2014 | 2014 | Novadaq Technologies, Inc. | Non sterile-drape LUNA drape was inadvertently mislabeled with an additional contents sterile label. | Error in labeling | Novadaq initiated their recall by telephone and followed by an email and a Medical Device Recall Notice dated November 11, 2014, to all affected customers. Distributors were instructed to notify their customers. The firm is removing the recalled products and replacing them with correctly labeled product. Customers were instructed to confirm receipt of the Recall Notice and their understanding of the actions to be taken by completing the attached confirmation form and faxing it to Novadaq Quality Department at 1-800-886-2419. For questions regarding this recall call 604-432-9861. |  | 324 | 324 units | Nationwide Distribution to IA, LA, GA, MD, ID, FL, TN, OH, TX, NC, MA, KY, KS, WI, IN, SC, OR, CT, CA, WA, PA, OK, and PR. | Terminated | November 06, 2015 | 367 | 1152 | Not_Computer | Not_Computer
Z-2021-2014 | 68196 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVOCATE Redi-Code+ BMB-BA006A Blood Glucose test strips |  |  | May 05, 2014 | 2014 | DIABETIC SUPPLY OF SUNCOAST, INC. | "Advocate Redi-Code + BMB-BA006A Blood Glucose test strips" are inappropriately labeled and marketed as "Advocate Redi-Code + Blood Glucose test strips". | Labeling False and Misleading | Diabetic Supply of Suncoast, Inc. sent an Urgent Voluntary Medical Device Recall letter dated June 6, 2014, to all affected consignees. The firm is asking the distributors to examine their inventory immediately and quarantine any product subject to recall. In addition, if they may have further distributed this product, to please identify their customers and notify them at once of this product recall. Customers who have the BMB-BA006A Advocate Redi-Code+ test strips can contact Diabetic Supply of Suncoast, Inc. directly for further information to determine if the products are being used in the proper manner or if they need to be returned to Suncoast for replacement. Test strips will be replaced with relabeled (with new labels) boxes of test strips to avoid any chance of confusion. Consumers should stop using the above mentioned product and contact Suncoast to find out if this recall pertains to the products they have in their possession and if so how to have them replaced for the relabeled product manufactured by BroadMaster Bio-Tech. Customers with questions should call Diabetic Supply of Suncoast, Inc. at 561-296-0488.For questions regarding this recall call 561-254-1007. |  | 1852800 | 1,852,800 cases (50 strips ea) | Worldwide Distribution - USA (nationwide) and Internationally to Venezuela, Ecuador, Canada &amp; Vietnam | Terminated | June 13, 2017 | 1135 | 1852800 | Not_Computer | Not_Computer
Z-0064-2015 | 69173 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Omni Micro-electrode/reference electrode for cobas b221 analyzer. |  |  | September 12, 2014 | 2014 | Roche Diagnostics Operations, Inc. | The default references for normal values are inconsistent between cobas b 221 and other blood gas instruments. In addition, the reference ranges are different when comparing with corresponding ranges listed in the Instructions for Use for the analyzer and the reference literature source. | Labeling Change Control | The Urgent Medical Device Correction will be sent via UPS Ground (receipt required) to the consignee list. The Urgent Medical Device Correction will be sent via UPS Ground (receipt required) to the consignee list.Consignees are instructed as follows:" Check the current set of reference values on your cobas b 221 blood gas system and, if appropriate, change the reference ranges accordingly." If you have not established your own reference ranges, please follow the instructions in the Workaround section of the UMDC or contact the Roche Support Network Customer Support Center. |  | 492 | 492 devicees | Nationwide Distribution including Puerto Rico and Guam, and the states of AK, AL, AZ, CA, CO, CT, FL, GA, HI, IA, IL, IN, KS, KY, LA, ME, MD, MI, MN, MO, MS, MT, NC, NH, NJ, NM, NY, OH, OK, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA and WI. | Open, Classified |  |  | 492 | Not_Computer | Not_Computer
Z-0821-2014 | 67163 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Puritan Bennett 840 Series Ventilator PB840 Operators Manual Addendum PN 066009A 09/02 |  |  | January 25, 2013 | 2013 | Nellcor Puritan Bennett Inc. (dba Covidien LP) | Covidien is correcting labeling to clarify the operational life of the oxygen sensor (O2 sensor), PN 4-072214-00 used in the Puritan Bennett" 840, and PN, G-062010-00, used in Puritan Bennett" 740, and760 ventilator systems. The Operators Manual for PB840 states a nominal life of one year and the Addendum incorrectly states 2 years. | Labeling Change Control | The firm, Covidien, sent an "URGENT MEDICAL DEVICE VOLUNTARY FIELD CORRECTION" letter dated January 24, 2014 to its Consignees/customers via certified mail. The letter described the product, problem and actions to be taken. The customers were instructed to read and understand the enclosed information; communicate this information to those who need to be aware of this labeling update; for the Puritan Bennett 840 ventilator only: If you have addendum P/N 066009A 09/02, remove and discard the addendum and replace it with the new information in Attachment I; for the Puritan Bennett 740 and 760 ventilators only: Add the new information in Attachment I; Ensure all O2 sensors in use or in inventory conform to the instructions in this notification. to your current operators manual and service manual, and read and understand the warnings, cautions and notes related to safety, ventilator use and operation that are contained in the Preface of the Operators Manuals. are asked to remove incorrect labeling and replace with the revised labeling.Covidien enclosed a labeling update (Attachment I) which provides detailed information related to the life of oxygen sensors used with the Puritan Bennett 840, 740 and 760 ventilators. Covidien has validated another oxygen sensor, P/N 10097559, for use in the Puritan Bennett 840, 740 and 760 ventilators. Information about that sensor is also included in the update in Attachment I.Should you have any questions regarding this letter or to report any issues with the Puritan Bennett" 840, 740 or 760 ventilators contact our Technical Support Department @ 1-800-255-6774, option 4, then option 1, to ensure proper device reporting procedures are followed. |  | 70160 | 70,160 | Worldwide Distribution: US (nationwide) and Internationally to: Albania, Argentina, Australia, Bahrain, Belgium, Belize, Bermuda Bolivia, Bosnia and Herzegovina, Brazil, Cambodia, Canada, Chile, China, Colombia, Costa Rica, Croatia, Ecuador, Egypt, El Salvador, Finland, France, French Polynesia, Georgia, Germany, Greece, Guam, Guatemala, Honduras, Hong Kong, Hungary, India, Indonesia, Iran, Ireland, Israel, Italy, Jamaica, Japan, Jordan, Kazakhstan, Kenya, Kuwait, Lebanon, Libya, Lithuania, Macedonia, The former Yugoslav, Malaysia, Mexico, Moldova, Republic of, Montenegro, Morocco, Netherlands, Nicaragua, Northern Mariana Islands, Oman, Panama, Paraguay, Peru, Philippines, Poland, Puerto Rico, Qatar, Romania, Russia, Saudi Arabia, Singapore, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Sudan, Switzerland, Syria, Taiwan, Thailand, Tunisia, Turkey, Uganda, Ukrainian SSR, United Arab Emirates (UAE) United Kingdom, Uruguay, Venezuela and Yugoslavia. | Terminated | November 24, 2014 | 668 | 78119 | Not_Computer | Not_Computer
Z-0049-2015 | 69404 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | StarMedTec LightTrailÂż Reusable Fibers, 365 um; Material/Part Number: 6453 |  |  | September 30, 2014 | 2014 | Boston Scientific Corporation | Product is not cleared for use with lasers other than the Auriga XL system | No Marketing Application | Boston Scientific sent an " Urgent Medical Device Recall Removal Immediate Action Required" letter dated September 30, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The notification instructs the account to check their inventory, remove any affected product from their inventory, and contact Boston Scientific for further instructions on returns. The notification also requests the completion of a Reply Verification Tracking Form as a method of documenting the presence or absence of affected product in their inventory.We ask that you complete and return the Reply Verification Tracking form, included with this letter, according to the instructions on page 3, even if you no longer have inventory of the recalled lots. If you identify any of the affected lots within your inventory, please segregate it immediately and you will be contacted by Boston Scientific with additional instructions after you return your form. You will receive credit, as appropriate, for any affected product in your inventory. Fax to : Field Action Center at 1-866-213-1806.For further questions call (508) 382-9555 |  | 5 | 5 units | US Distribution including the states of NY and TX. | Terminated | November 26, 2014 | 57 | 45 | Not_Computer | Not_Computer
Z-2093-2014 | 68701 | No | BYG | Oxygen mask. | Anesthesiology | Anesthesiology | 510(K) Exempt | Oxygen Mask, Model No. 8120-7 |  |  | June 25, 2014 | 2014 | Salter Labs | The case label correctly states partial rebreather but the insert label incorrectly states non-rebreather. | Error in labeling | Salter Labs sent a Medical Device Recall letter dated June 25, 2014 to all affected customers. The letter identified the affected product, problems and actions to be taken. Customers are instructed to complete the attached Return Response Form and return it to Salter as soon as possible. Customers are instructed to send completed form to Salter by either Fax (800) 628-4690, attn: Sarah Velasquez, email to SVelasquez@us.salterlabs.com, or mail to Salter Labs, attn: Sarah Velasquez, 100 W. Sycamore Rd. Arvin, CA 93203. |  | 584 | 584 units | Worldwide Distribution - US (nationwide) and the countries of Brazil, United Kingdom, Indonesia, Morocco, Mexico, Norway, Venezuela, and South Africa. | Terminated | September 10, 2014 | 77 | 584 | Not_Computer | Not_Computer
Z-1073-2014 | 67405 | Yes | KPO | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | NxStage PureFlow-B Solution; Model/Catalog Number: RFP-401. Product is packaged in a case containing two individual 5L bags. |  |  | January 24, 2014 | 2014 | NxStage Medical, Inc. | Mislabeled product. | Packaging process control | NxStage sent a Voluntary Medical Device Recall letter dated January 24, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. They were asked to check all boxes in inventory for the recalled lot and segregate and quarantine them. In addition, they were asked to complete the recall reply form and return it by fax to NxStage Medical as well as to contact NxStage Customer service to arrange for return of all affected product and for replacement to be sent. Distributor was asked to notify all downstream customers and to provide NxStage Medical with a list of customers. Please complete and fax this form to 978-687-4810 even if you do not have any of the affected product on hand.If you have any questions or comments, please feel free to contact NxStage Customer Service at 1-866- NXSTAGE (1-866-697-8243). |  | 969 | 969 cases (2 bags per case) | US Distribution including the states of NY, TX, ID, AR, CA and AL | Terminated | August 29, 2014 | 217 | 969 | Not_Computer | Not_Computer
Z-2074-2014 | 68590 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Baxter MiniCap with Povidone-Iodine Solution, 5C4466P, Baxter Healthcare Corporation. For dialysis. |  |  | June 17, 2014 | 2014 | Baxter Healthcare Corp. | Inadequate iodine and packaging related defects. | Packaging process control | The firm notified their direct consignees on 06/17/2014 by letter requesting customers to locate and remove all affected product from their facility, and contact Baxter Healthcare Center for Service to arrange for return and credit. |  | 585360 | 585,360 units | Distributed USA (nationwide) and Canada. | Terminated | April 03, 2015 | 290 | 585360 | Not_Computer | Not_Computer
Z-0066-2015 | 68987 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Arrow Select Kits (ASK) and Maximal Barrier Precautions Kits (CDC) with BD Medical Posiflush 10mL Saline Flush Syringe. |  |  | August 06, 2014 | 2014 | Arrow International Inc | The component supplier recalled their 0.9% Sodium Chloride Injection USP BD Posiflush SF Saline Flush Syringe 10mL because some unit packages may exhibit open seals which impact package integrity and potentially product sterility. | Packaging | Arrow International sent an Urgent Medical Device Recall Notification letter, dated August 6, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately discontinue use and quarantine any affected product. Customers were instructed to remove the BD pre-filled saline syringe from the exterior of the packaging. To return the BD syringes, customers were instructed to complete the enclosed Recall Acknowledgement Form and fax it to 1-855-419-8507, Attn: Customer Service. A customer service representative will contact the customer with a Return Good Authorization (RGA) Number and will provide instructions for the return of producet to Arrow International. All customers were asked to complete the enclosed Recall Acknowledgement Form and fax it to 1-855-419-8507. Attn: Customer Service. Customers with questions were instructed to contact their local sales representative or Customer Service at 1-866-246-6990. |  | 71255 | 71255 | Nationwide Distribution | Open, Classified |  |  | 71255 | Not_Computer | Not_Computer
Z-2115-2014 | 68703 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Pad-Pak and Pedi-Pak accessories for the HeartSine Samaritan PAD 300P and 350P device (SAM 300P and 350P). |  |  | June 12, 2014 | 2014 | Heartsine Technologies, Limited | A small number of sealed foil pouches containing the electrodes were found to be difficult to open, resulting in potential delay in therapy. | Packaging process control | HeartSine notified distributors in the United States by e-mail on June 12, 2014 and by follow-up letter on June 13, 2014. Distributors were given a letter to use in communications with customers on June 17, 2014. Customers were instructed to return affected product to there distributors. Distributors were asked to return the affected product to HeartSine. In order to effect an immediate return, replacement Pad-Paks and Pedi-Paks have already been delivered to all respective distributors. A seconde letter, dated June 30th, was sent to end users to clarify where on the label to find lot specific information.For question contact Quality Manager at +44 (0)28 90 93 94 19. |  | 3245 | 3245 | Worldwide Distribution - US (nationwide) and in Canada, Mexico and Latin America. | Terminated | April 16, 2015 | 308 | 3245 | Not_Computer | Not_Computer
Z-2113-2014 | 68709 | No | MVC |  |  | Pathology | PMA | InSite HER-2 Detection Kit, an immunohistochemistry assay for in vitro diagnostic use. |  |  | July 01, 2014 | 2014 | Biogenex Laboratories, Inc. | Product was distributed without 510(k) approval. | No Marketing Application | Each customer contacted by phone - neither is in possession of the product. |  | 4 | 4 kits | Brazil and Ecuador. No US distribution | Terminated | July 24, 2014 | 23 | 4 | Not_Computer | Not_Computer
Z-2072-2014 | 68742 | No | GWN | Nystagmograph. | Neurology | Neurology | 510(k) | Integra Camino 110-4 Intracranial Pressure (ICP) Monitoring Catheters, |  |  | July 09, 2014 | 2014 | Integra LifeSciences Corp. | Integra LifeSciences is recalling the Camino Intracranial Pressure Monitoring Kits because of revised labelling to indicate they are unsafe in an MR environment. | No Marketing Application | A customer notification letter dated 7/9/14 was sent to customers to inform them that Integra LifeSciences has recently revised the labeling of certain Camino Intracranial Pressure Monitoring Kits to indicate they are UNSAFE in an MR environment. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to complete the form attached the letter and return it by email or fax. Customers with questions are instructed to contact Customer Service at (855) 532-1723. |  | 28752 | 28,752 units | Algeria, Austria, Belgium, Bulgaria, Cyprus, Egypt, Finland, France, Germany, Great Britain, Greece, Hungary, Greece, Hungary, Ireland, Israel, Italy, Jordan, Lebanon, Lithuania, Luxemborg, Netherlands, Poland, Portugal, Romania, Saudi Arabia, Slovenia, South Africa, Spain, Sweden, Switzerland, Tunisia, Turkey, United Arab Emirates, Argentina, Brazil, Brunei, Chile, China, India, Japan, Korea, Mexico. | Terminated | October 14, 2014 | 97 | 28752 | Not_Computer | Not_Computer
Z-0476-2015 | 69386 | No | LBH | Cavity varnish. | Dental | Dental | 510(k) | Dharma Cavity Varnish 15mL |  |  | September 29, 2014 | 2014 | Dharma Research, Inc. | During an FDA inspection it was found that the products are marketed without a cleared 510k. | No Marketing Application | On September 29, 2014 the firm sent a notification to its consignees informing them of the product being discountued and removed. The firm asked for the product to be returned directly to Dharma's office at 5220 NW 72 Avenue, Suite 15, Miami FL 33166. For any questions please do not hesitate to call (305) 482-9669. |  | 916 | 916 | Worldwide Distribution: US Distribution in states of: CA, PA, FL, IL, including Puerto Rico; and countries of: Honduras, Dominican Republic, Turkey, Nicaragua, Colombia, Pakistan, Trinidad &amp; Tobago, Guatemala, Phillipines, Saudi Arabia, and United Arab Emirates. | Open, Classified |  |  | 3758 | Not_Computer | Not_Computer
Z-1552-2014 | 68074 | No | LFL |  |  | General &amp; Plastic Surgery | 510(k) | TX1 Tissue Removal System. TX1 Procedure Pack P/N 554-1003-001 (outer box), TX1 MicroTip P/N 554-1002-001 (inner tray). For use with the TX System Console as an ultrasonic surgical aspirator of soft tissue. |  |  | April 14, 2014 | 2014 | Tenex Health Inc | The sterile barrier in the packaging may be compromised due to cracks in the thermoform tray that may occur during transport of the product in extreme cold weather conditions, in conjunction with mechanical impact. | Packaging | Tenex Health sent an Urgent Medical Device Recall notification letter dated April 15, 2014, to all consigness. The letter informed consignees of the recall, manufacturer, product name, product code, reason for recall, recall carried out to the user level, instructions to customers, contact information, label examples, and response form was provided. Consignees were instructed to immediately examine their inventory per the enclosed instructions, if they identify any of the affected product within their inventory, they should segregate the product and immediately return only MicroTips with damaged trays to the address below, using the Tenex Health, Inc. FEDEX account number and referencing the Returned Goods Authorization (RGA) number.RGA#Tenex Health, Inc.26902 Vista TerraceLake Forest, CA 92630Customers were asked to complete and return the enclosed response form as soon as possible.Customers with questions should contact Tenex at (949) 454-7500, ext. 209or email tanaka@tenexhealth.com.For questions regarding this recall call (949) 454-7500. |  | 9927 | 9,927 units | Worldwide Distribution - USA (Nationwide) and Internationally to Panama and Singapore. | Terminated | July 16, 2014 | 93 | 9927 | Not_Computer | Not_Computer
Z-0715-2014 | 67177 | No | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | Anspach MCA Cutting Burr, DFU Rx Only |  |  | March 29, 2012 | 2012 | The Anspach Effort, Inc. | Current directions for use for the MCA cutting burs (item 18-0001) do not provide enough details regarding the intended use specifically for the MCA burrs. | Error in labeling | Anspach sent an Urgent Medical Device Labeling Correction letter dated Marach 29,, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to include the item 18-0079 with their inventory immediately, complete the attached reply form indicating their receipt of the letter. Return the completed form by fax to the number or email provided on the form. If product was further distributed to other facilities, the correction letter should be forwarded to them. Customers with quesitons should call 1-800-327-6887 or email customer.support@synthes.com or ProductSupportTeam@Synthes.com.For questions regarding this recall call 561-627-1020. |  | 492 | 492 (321 US) | Worldwide Distribution - USA (nationwide) and Internationally to Great Britain, Australia, Switzerland, Sweden, France, Japan, Spain, Italy, Canada, and Netherlands. | Terminated | April 15, 2016 | 1478 | 492 | Not_Computer | Not_Computer
Z-0819-2014 | 67149 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Composix L/P with Echo PS 10" X 13" Product Code 0144113 |  |  | January 03, 2014 | 2014 | Davol, Inc., Subs. C. R. Bard, Inc. | Product labeling does not match product configuration. | Labeling Change Control | Davol, Inc., Subs. C. R. Bard, Inc.sent a customer notification letter on January 3, 2013, to all affected customers via Fedex to inform consignees that the Instructions for Use (IFU) and graphic contained onthe inside of the product carton bottom references two features added to the inflation tube that are not present on the product they received. Please contact the Davol Customer Service Department at 1-800-556-6275 or C. R. Bards Medical Services &amp; Support Department at 1-800-562-0027 if you have any questions. Consignees are asked to complete an effectiveness check form and return it via fax to 401-825-8753.For questions regarding this recall call 800-556-6275, |  | 19 | 19 | Nationwide Distribution including CA, PA, IL, VA, FL, GA, NY, TX, MO, NE, IN, NM, OH, IA. | Terminated | December 23, 2015 | 719 | 65 | Not_Computer | Not_Computer
Z-0070-2015 | 69368 | No | OMP | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | V.A.C. GranuFoam Silver Large Dressing is barrier to bacterial penetration and may help reduce infection. Featuring SensaT.R.A.C. Technology REF M8275099/10 STERILE Single Use only Latex Free Rx Only. |  |  | September 18, 2014 | 2014 | KCI USA, Inc. | A portion of the V.A.C. GranuFoam Silver Dressing M8275099/10 lot 50775977 was incorrectly manufactured with an incorrect expiry date on the sterile pouch. The correct expiry date of 2015-11 is on the case label, however the pouch label identifies the expiry date of 2016-11. | Error in labeling | KCI will send a customer letter to each consignee notifying them of the issue, and to return the affected units. |  | 1090 | 1090 pouches (109 boxes) | The affected product was distributed to Mexico, Turkey, and United Arab Emirates only. The affected product was not distributed to the US. | Terminated | July 24, 2015 | 309 | 1090 | Not_Computer | Not_Computer
Z-2069-2014 | 68480 | No | EJK | Calcium hydroxide cavity liner. | Dental | Dental | 510(k) | Calcium Hydroxide Powder, USP |  |  | June 05, 2014 | 2014 | Axcentria Pharmaceuticals, LLC | One lot of Calcium Hydroxide Powder is being recalled because it was manufactured and distributed without an approved 510k, and the current label does not include instructions for use. | No Marketing Application | Axcentria Pharmaceuticals sent an urgent medical device recall notice, dated June 5, 2014, was sent to customers and described the product, problem, and actions to be taken. Customers were asked to fill out and return the form included.For questions customers were instructed to call 215-453-5055, ext 317. |  | 528 | 528 | US distribution to NY and CA | Open, Classified |  |  | 528 | Not_Computer | Not_Computer
Z-0183-2014 | 66610 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Applied Medical's InziiÂż 12/15mm Retrieval System. |  |  | October 21, 2013 | 2013 | Applied Medical Resources Corp | Applied Medical is conducting a voluntary recall of the InziiÂż 12/15mm retrieval system. During shipment, the retrieval system packaging has the potential to become punctured with small holes, which could compromise the sterile barrier. The likelihood of this situation to occur and result in patient harm is highly unlikely; however, out of an abundance of caution for patient safety and a commitmen | Package design/selection | Customer notifications of the recall began the week of 10/21/2013 via UPS overnight with letters requesting that all affected products be returned to Applied. The notification letter is dated October 21, 2013 and titled "URGENT: MEDICAL DEVICE RECALL". The notification was accompanied by a Recall Notification Confirmation Form. The letter informed customers of the recall by providing the product description with codes (Model and Lot Numbers). reason for the recall, actions to take and contact information. |  | 29506 | 29,506 units | Worldwide Distribution: US (nationwide) and countries of: Australia, Austria, Belgium, Brazil, Canada, Chile, Colombia, Czech Republic, Denmark, Ecuador, Finland, France, Germany, Hong Kong, India, Ireland, Israel, Italy, Japan, Lebanon, Malaysia, Mexico, New Zealand, Netherlands, Norway, Peru, Poland, Portugal, Saudi Arabia, Singapore, Spain, South Africa, Sweden, Switzerland, United Arab Emirates, and United Kingdom. | Terminated | September 15, 2014 | 329 | 29506 | Not_Computer | Not_Computer
Z-1574-2014 | 68109 | No | HTT | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | ANSPACH Cutting bur, 3 mm Diamond Ball, Standard Length. REF L-3SD Use with LONG Attachments. Used for cutting and shaping bone, including the spine and cranium. |  |  | January 08, 2014 | 2014 | The Anspach Effort, Inc. | Anspach product code L-3SD contained product code S-3SD which is shorter in length than the labeled product. | Labeling mix-ups | Anspach sent an "URGENT: MEDICAL DEVICE REMOVAL" letter dated January 13, 2014, to all affected customers. The letter described the product, the problem and actions to be taken. If they have distributed any of the products to other services or facilities, customers were instructed to please forward this information as appropriate. Customers with questions you contact Anspach Customer Support at (800) 327-6887, option 3 or email customer.support@synthes.com.. Customers were asked to complete the attached reply form confirm receipt of the letter, return the completed form by fax to 800-327-6661.For questions regarding this recall call 561-494-3706. |  | 30 | 30 cutting burs affected | Nationwide Distribution including FL, TX, NC, IL, CO, and PA. | Terminated | January 20, 2015 | 377 | 30 | Not_Computer | Not_Computer
Z-2123-2014 | 68706 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | PHAROS Excimer Laser, Model No. EX-308 |  |  | June 30, 2014 | 2014 | Ra Medical Systems Inc | RA Medical Systems is recalling the PHAROS EX-308 Excimer Laser because there is a possibility that the laser may calibrate with the iris closed. | Error in labeling | Ra Medical Systems sent a customer letter dated June 30, 2014, to who purchased the PHAROS EX-308 Excimer Laser Phototherapy System Hand Piece. The letter informed the customers of the problems identified and the actions to be taken. Customers were also provided with the Instructions for Use (IFU) for the PHAROS EX-308 Excimer Laser Phototherapy System Hand Piece and a replacement iris indication label for the handpiece. |  | 435 | 435 units | Worldwide Distribution - USA (Nationwide) and Internationally to: Korea, Kuwait, and Saudi Arabia. | Terminated | September 09, 2014 | 71 | 435 | Not_Computer | Not_Computer
Z-2090-2014 | 68718 | No | DRM | External cardiac compressor. | Cardiovascular | Cardiovascular | 510(k) | Zoll Circulation LifeBand Single Pack (accessory), and Nickel-metal hydride (NIMH) Battery. |  |  | July 16, 2013 | 2013 | Zoll Circulation, Inc. | Typographical error in the Japanese legal label of the AutoPulse Resuscitation System Model 100. The label is specific to the Japanese version of the product. | Labeling Change Control | Zoll Japan handled the recall. |  | 201 | 161 Nickel Metal batteries, 40 LifeBand | Distributed in Japan. | Terminated | July 22, 2014 | 371 | 201 | Not_Computer | Not_Computer
Z-2156-2014 | 68807 | No | LFL |  |  | General &amp; Plastic Surgery | 510(k) | Sonopet Ultrasonic Surgical System Console, Ultrasonic Aspirator, Stryker Medtech. |  |  | June 24, 2014 | 2014 | Stryker Instruments Div. of Stryker Corporation | The firm received complaints regarding insufficient irrigation to the tip of the Sonopet system. Lack of irrigation can potentially lead to excessive heat at the tip. | Labeling design | The firm sent Medical Device Upgrade Notification letters dated July 11, 2014 to their customers. |  | 2460 | 2,460 units | Nationwide Distribution-including DC and the states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, VT, WA, WI &amp; WV. | Terminated | October 27, 2015 | 490 | 2460 | Not_Computer | Not_Computer
Z-0709-2014 | 66869 | No | CAE | Oropharyngeal airway. | Anesthesiology | Anesthesiology | 510(K) Exempt | King Systems |  |  | December 13, 0013 | 13 | King Systems Corp. | KLTD214 airways found inside KLTSD414 packages. The KLTSD414 airway has a lumen for use of a gastric suction tube; the KLTD214 airway does not include a lumen for a gastric suction tube. | Labeling Change Control | December 13, 2013 Urgent Recall Notices were sent to all Direct Accounts identifying the recalled item, risks associated with the incorrectly labeled product and asking customers to stop use, quarantine and destroy any recalled product on hand. Customers were asked to froward the notifcation if the product was further distributed and FAX the completed recall acknowledgement form to 317-776-5175 or email wnethercutt@kingsystems.com. Call the customer service representative (800.262.8673) to arrange for credit of your affected product(s). |  | 108 | 108 cases( 540 products) | Distributed in the states of CA, NY | Terminated | May 15, 2014 | 730638 | 108 | Not_Computer | Not_Computer
Z-0487-2015 | 69648 | No | BZD | Noncontinuous ventilator (IPPB). | Anesthesiology | Anesthesiology | 510(k) | Cannulaide, Model 101. The Cannulaide product is composed of die-cut hydrocolloid coated polyurethane film with an integral hook strip (i.e., Velcro) to allow for the use of a loop material to help secure a nasal interface. |  |  | January 18, 2007 | 2007 | Beevers Manufacturing &amp; Supply, Inc. | Cannulaide model CA 101 was recalled because the firm discovered this model was packed and labeled as CA 102. | Labeling Change Control | Beevers Manufacturing and Supply, Inc. (BMS) of McMinnville, Oregon, began notifying consignees about this recall on 1/18/2007.Customers were advised to inspect blue box that say size 2 (CA 102) and check if size 1 (CA 101) was packed inside. Salter Labs provided BMS' recall report to FDA in October 2014. |  | 393 | 393 units | Worldwide Distribution: US (nationwide) and country of: ME-Musaffah. | Terminated | December 01, 2014 | 2874 | 393 | Not_Computer | Not_Computer
Z-2592-2014 | 68993 | Yes | MSD | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Split Cath III Trays, 16F x 28CM |  |  | August 04, 2014 | 2014 | Medical Components, Inc dba MedComp | The Split Cath III Trays were incorrectly labeled as 16F x 32CM Split Cath III instead of 28CM. | Labeling False and Misleading | Medcomp sent a Recall Notification letter, dated August 4, 2014, to all affected customers. The letter identified the affected product, problem, and action to be taken. Customers were requested to respond and return all affected product. For questions call contact Susan Smith at ssmith@medcompnet.com. |  | 110 | 110 | USA Nationwide Distribution | Terminated | December 10, 2014 | 128 | 110 | Not_Computer | Not_Computer
Z-2281-2014 | 68346 | No | HQL | Intraocular lens. | Ophthalmic | Ophthalmic | PMA | AMO TECNIS CL Foldable Silicone IOL with OptiEdge Design, Model Z9002. Indicated for primary implantation for the visual correction of aphakia in adult patients. Tecnis CL Mono 13 mm with the following diopter sizes: 19.5-20 |  |  | May 13, 2014 | 2014 | AMO Puerto Rico Manufacturing, Inc. | Units may be labeled with the incorrect diopter power. | Error in labeling | AMO sent an Urgent Medical Device Recall dated May 13, 2014 to the affected customers. The letter identified the affected product, problem and actions to be taken. This recall only affects the specific Silicone Tecnis IOLs Model Z9002. Implanted IOLs with appropriate post-op refractive outcomes are not affected by this recall. Please underatake the following actions:1. Stop using and remove from your inventory the affected product.2.Complete and return the included Facsimile Form as AMO. If you have inventory to return, please complete the Facsimile Form, and returned to AMO. For questions contact AMO Customer Service Representative 877-266-4543. |  | 77055 | 77,055 IOLs in total | Worldwide Distribution - Nationwide (USA and Puerto Rico), and foreign: Latin America, Asia Pacific, Canada, Europe, Africa and Middle East, and Japan. | Open, Classified |  |  | 77055 | Not_Computer | Not_Computer
Z-1501-2014 | 67976 | [6 empty] | Hot Ice, 1 fl. oz., Meta Labs LLC, UPC 80705-2 |  |  | February 03, 2014 | 2014 | Altasource LLC dba Meta Labs LLC | Firm is not registered as a drug manufacturer. | No Marketing Application | Meta Labs notified customers of the Voluntary Recall Notification letter on February 4, 2014. The letter requested that the products be returned. For questions call 770-645-0602. |  | 155 | 155 | USA Nationwide in the states of AZ, AR, CA, CO, FL, GA, ID, IL,KY, LA, MI, NE, NV, NY, NC, OH, PA, TN, TX, UT, and VA. | Open, Classified |  |  | 274 | Not_Computer | Not_Computer
Z-0934-2014 | 67302 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | 3.5mm Locking Cortical Bone Screw, T15 Hexalobe, Self Tapping, 20mm |  |  | January 08, 2014 | 2014 | OrthoPediatrics Corp | OrthoPediatrics has initiated a Voluntary RECALL of Part # 00-0903-2620, 3.5MM Locking CorticalScrew, Lot # 009SV because the device within the package may be incorrect. These packages incorrectly contain a 3.5mm Non Locking Cortical Screw # 00-1050-3524. | Labeling mix-ups | Orthopediatrics Corporation has requested sales representatives and distributors who have been identified via shipping records as having the possibility of having received one or more discrepant devices affected by this Recall were requested to quarantine the devices identified in this Recall by telephone on January 8, 2014. The decision to Recall these devices was made and a formal written Recall Notice sent to the consignees via email on January 13, 2014.Notification will not be sent to hospitals at this time provided that all devices areretrieved or accounted for within 60 days of the initiation of recall. Hospitals will be notified on an as needed basis if OrthoPediatrics determines that all recalled devices have not been recovered or accounted for, and that contacting the hospitals involved will expedite the recall efforts. All discrepant devices subject to this voluntary Recall are planned to be returned from distribution to OrthoPediatrics Corp. in Warsaw, Indiana and maintained in quarantine. All remaining unshipped devices from these lots have been placed in quarantine at OrthoPediatrics Corp.All quarantined devices at OrthoPediatrics will be inspected and repackaged to correct the mislabeling per normal documented procedures. Objective evidence will be maintained of the rework of the devices.For questions regarding this recall call 574-268-6379. |  | 12 | 12 units | Worldwide Distribution - USA including OK, KY, FL, LA, OH, CO, MO, MN and Internationally to Ireland, Saudi Arabia, Italy, Turkey, South Africa, and Germany. | Terminated | April 23, 2014 | 105 | 12 | Not_Computer | Not_Computer
Z-1456-2014 | 67499 | No | GXL | Skullplate screwdriver. | Neurology | Neurology | 510(K) Exempt | Synthes CMF Battery Powered Driver, Part no. 305.840. Used to place screws in the cranialfacial bones, and to drill and/or tap holes to prepare cranialfacial bones for screw insertion. |  |  | January 30, 2014 | 2014 | Synthes, Inc. | The CMF Battery Powered Driver graphic case contains outlines of the Battery Powered Driver, taps, screwdriver blades, and a screw holding sleeve. However, the device is not intended to be used in conjunction with a tap for tapping holes in a patient's skull, upper and/or lower jaw. Therefore some outlines in the graphic case suggest a potential use of this device that is considered off-label. | Labeling design | Synthes sent an Urgent Notice: Field Safety Notification letter dated January 30, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Please take the following actions;Please see the Battery Powered Driver Users Manual for on-label procedures. You can access an electronic version of the Battery Powered Driver Users Manual by accessing the following : linlc http://www.synthes.com/pages/default.aspx. Click on North America Tab at the top of the screen, and then type Battery Powered Driver Users Manual in the search bar to download the Users Manual Forward this Field Safety Notification to anyone in your facility that needs to be informed. If the graphic case has been forwarded to another facility, contact that facility. " Maintain a copy of this notice.If you have any questions, please call 610-719-5450 or contact your Synthes Sales Consultant. Thank you for your attention to this issue. |  | 20081 | 20081 | Nationwide Distribution | Open, Classified |  |  | 20081 | Not_Computer | Not_Computer
Z-2558-2014 | 68951 | No | KQT | Gastroenterology-urology evacuator. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | CollectEVAC, Sterile device with tubing and adaptor for use with ACMI and STORZ/Wolf resectoscopes. Part number 610-954. Rx only. Sterile EO. |  |  | August 01, 2014 | 2014 | Civco Medical Instruments Inc | A pinhole leak identified in the packaging of the recalled product potentially could compromise the sterility of the product. | Packaging | Civco Medical Instrucments contacted customers by phone and sent an Urgent Medical Device Recall letter dated August 1, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were requested to return all the affected product. Customers were instructed to contact Customer Service at 319-248-6520 with any concerns or questions. |  | 52 | 52 kits | Worldwide Distribution - USA including AZ, IL, MA, MD, MI, MN, MO, MS, LA, NC, NJ, NY, SD, TN, VA, WI, and Internationally to Brazil. | Terminated | January 28, 2015 | 180 | 52 | Not_Computer | Not_Computer
Z-1494-2014 | 67831 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | Arrow Pressure Injectable PICC Powered by Arrow VPS Stylet Catalogue number CDC-35052-VPS, peripherally inserted central catheter. |  |  | March 24, 2014 | 2014 | Arrow International Inc | Kits are packaged with the incorrect catheter. The kits contain a 40cm catheter and should contain a 50cm catheter. | Labeling mix-ups | A recall notification letter, dated March 24, 2014 was sent to end users." Urgent Medical Device Recall Notification " |  | 789 | 789 | Nationwide Distribution- Including the states of Az, Tn, Fl, Ga, Ca | Terminated | September 21, 2015 | 546 | 789 | Not_Computer | Not_Computer
Z-2302-2014 | 68536 | No | OEQ | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | Enforcement Discretion | Customed Angiodrape Pack |  |  | May 20, 2014 | 2014 | Customed, Inc | Customed has determined that there is the possibility that packaging integrity may be compromised on different catalog numbers (multiple lots). This could result in an injury to the patient due to product contamination or loss of sterility condition. | Package design/selection | Customed Inc. sent a Recall Notification letter dated June 3, 2014 via email and followed by letter to all affected customers. The letter identified the affected products, problem and actions to be taken. The letter informed customers if the affected products are in their inventory to: 1. Remove products from sale (quarantine them) and cease distribution. 2. Complete and return the enclosed acknowledgement form to Customed Inc .Recall Coordinator. Affected products should be returned to Customed Inc for credit.Customers with questions were instructed to contact Customed Inc. Recall Coordinator at 1-787-801-0100 Ext. 7540 |  | 39936 | 26 lots, 1536 units (multiple units per lot) | Worldwide Distribution - US (nationwide) in the state of Florida including Puerto Rico (USA); &amp; USVI (St John) and the countries of Tortola, Trinidad, Dominican Republic, Costa Rica and Mexico. | Terminated | August 14, 2017 | 1182 | 15299586 | Not_Computer | Not_Computer
Z-2678-2014 | 69102 | No | JAX | Pneumotachometer. | Anesthesiology | Anesthesiology | 510(K) Exempt | No packaging. The device is a computed tomography scanner. |  |  | January 09, 2007 | 2007 | Philips Medical Systems (Cleveland) Inc | When the Extended Display Firled ov View reconstruction is used, the CT numbers and anatomy in teh extended region are not accurate. | Labeling design | Philips sent a Product Safety Notification letter dated December 21, 2006 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to distribute the Product Safety Notification to all affected users and keep a copy of the notification with the Instructions for Use until further notice. Customers with questions were instructed to contact Customer Care Center at 1-800-722-9377, option 5 - Diagnostic Imaging, option 1 - CT.For questions regarding this recall call 440-483-7600. |  | 61 | 61 | Worldwide Distribution - US including MO, TX, MD, AR, VA, KY, RI, NJ, MA, NY, LA, PA, VT, SC, CA, OH, FL, MN, MI and Internationally to Netherlands, Canada, Finland, Austria, France, Malaysia, Denmark, Norway, UK, Australia, Belgium, New Zealand. | Terminated | November 06, 2015 | 3223 | 61 | Not_Computer | Not_Computer
Z-1457-2014 | 67560 | No | GAD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | The Synthes Hohmann Retractor, 43 mm Long Shank, Angled 235 mm |  |  | February 11, 2014 | 2014 | Synthes, Inc. | The Synthes Hohmann Retractor was mis-etched on the product and package as part number 399.24 instead of part number 399.22. | Error in labeling | Synthes sent an Urgent Notice: Medical Device Recall letter dated February 11, 2014, to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to Complete the attached Verification Section form. For questions call 610-719-5450. |  | 20 | 20 | Worldwide Distribution - USA Nationwide Distribution in the states of: TN, MD, NY, IN, NJ, WA, VA,and the country of Canada. | Terminated | September 10, 2015 | 576 | 20 | Not_Computer | Not_Computer
Z-2670-2014 | 68989 | No | NGT | Intravascular catheter. | General Hospital | General Hospital | 510(k) | 0.9% Sodium Chloride Injection, USP BD PosiFlush(TM) SF Saline Flush Syringe 10 mL REF#306553 used in Centurion Convenience Kits. Recalled product is attached to the outside of the Convenience Kits. |  |  | July 25, 2014 | 2014 | Centurion Medical Products Corporation | Unit packages may exhibit open seals which impacts package integrity and potentially product sterility. | Packaging | Centurion sent an Urgent Product Recall Notice dated July 25, 2014, to all affected customers via Certified Mail. The notification discussed the affected product issue and provided instructions for initiating product return. Customers were instructed to identify and remove all inventory of these kits that they may have within their possession. Customers were asked to forward a copy of the notice to any customers to whom this product may have been further distributed. Customers were also asked to complete the enclosed accountability record, whether or not they have the product, and fax or email at their earliest convenience to 517-546-3356 or lcarpenter@centurionmp.com.Customers with questions were instructed to call (517) 546-5400, Ext 1135.For questions regarding this recall call 517-546-5400, ext 1156. |  | 21388 | 21388 | The product was distributed domestically to the following states: UT, WI, TX, NE, AZ, CA, IA, MD, NY, OK, PA, TN, TX, VA, NC, IL, MA, OR, WA, IN, SD. | Terminated | November 25, 2014 | 123 | 21388 | Not_Computer | Not_Computer
Z-1491-2014 | 64791 | No | MVU | Analyte specific reagents. | Pathology | Hematology | 510(K) Exempt | The CEP 8 SpectrumGreen (SG) ASR Probe Kit, 20ul (list 06J37-018) is an Analyte Specific Reagent (ASR) containing a single vial of CEP 8 spectrum green probe and a Certificate of Analysis (COA). Chromosome enumeration probes (CEP) are chromosome-specific FISH probes that hybridize to highly repetitive human satellite DNA sequences, usually located near centromeres. CEP signals enable the identification and enumeration of human chromosomes in interphase and metaphase cells from fresh and archived samples. The label for CEP 8 SG ASR Probe Kit indicates it is an Analyte Specific Reagent, and analytical and performance characteristics are not established. |  |  | March 28, 2013 | 2013 | Abbott Molecular | 20 kits of CEP 8 SG ASR Probe (list 06J37-018, lot 440023), contained Vysis CEP 4 (Alpha Satellite) SG Probe, 20u, (part 32-112004 lot 438814), instead of the expected CEP 8 probe. | Packaging change control | Customers were sent Field Correction Notices dated 3/28/13 instructing them to inspect their CEP 8 SpectrumGreen Probe kits for the presence of CEP4 probe vials. If any CEP4 vials are discovered, customers should contact their local Abbott Molecular Technical Support Represtative for coordination of reagent replacement with the correct CEP8 vial. If customers have any specific questions they can contact Abbott Molecular at 1-800-553-7042, Option 2. |  | 20 | 20 kits | Nationwide (CA, CO, CT, FL, HI, IL, IN, MA, MD, MI, MN, MO, MT, NC, NE, NJ, NM, NV, NY, OH, OR, PA, TN, TX, VA, WA, WI); Canada | Terminated | March 08, 2016 | 1076 | 20 | Not_Computer | Not_Computer
Z-2671-2014 | 68974 | No | MZO |  |  | Microbiology | PMA | VITROS Immunodiagnostic Products Anti-HCV Reagent Pack for in vitro qualitative detection of immunoglobulin G antibody to hepatitis C virus (anti-HCV) in human serum and plasma. REF 680 2454, IVD -- Ortho-Clinical Diagnostics High Wycombe UK. Note that one sales unit has 5 Reagent Packs contained within it. Each Reagent Pack can test for 100 samples (patient, Quality Control, etc.). |  |  | August 06, 2014 | 2014 | ORTHO-CLINICAL DIAGNOSTICS | VITROS Immunodiagnostic Products Anti-HCV kit lot 9090 has been observed producing sporadic lower than expected VITROS Anti-HCV test results due to an approximate signal loss of up to 66.2% which can result in unexpected negative test results being obtained for reactive sample fluids (quality control and patient samples). | Packaging process control | All US consignees (users) who were shipped the affected lot VITROS HCV Reagent Pack were sent communications via FedEx and/or US Priority Mail (PO Boxes only) on 8/06/14, and advised of the issue and to immediately discontinue using all remaining inventory of VITROS Anti-HCV Reagent Packs Lot 9090. For additional information, please contact OCD's Customer Technical Service representatives at 1-800-421-3311. |  | 463 | 463 sales units | NY, NJ, CT, MA, PA, VA, MD, FL and GA | Terminated | April 14, 2017 | 982 | 463 | Not_Computer | Not_Computer
Z-2568-2014 | 69072 | No | DSP | Intra-aortic balloon and control system. | Cardiovascular | Cardiovascular | 510(k) | LINEAR 7.5 FR. 40cc IAB Catheter |  |  | August 11, 2014 | 2014 | Datascope Corporation | During label reconciliation a duplicate serial number label was detected to be printed for the LINEAR 7.5 FR 40 cc Intra-Aortic Balloon Catheter. | Labeling False and Misleading | Maquet Inc.is sued an Urgent Medical Device Recall Letter/ Fax Back Form dated August 11, 2014. Customers are asked to quarantine the affected product and upon return they will be provided with an exchanged unaffected product. The attached form should be completed an returned to Maquet Inc. Questions can be directed to Maquet Technical Support at (888) 627-8383 - Press option "3" followed by option "1" Monday through Friday between the hours of 8 am and 5 pm (EST). |  | 1 | one unit | US Distribution to TX. | Terminated | February 05, 2015 | 178 | 1 | Not_Computer | Not_Computer
Z-1489-2014 | 68034 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | ev3 EverCross 0.035" OTW PTA Dilatation Catheter, 6mm x 120mm x 135cm, REF: AB35W06120135. and ev3 EverCross 0.035" OTW PTA Dilatation Catheter, 6mm x 40mm x, 80cm REF: AB35W06040080. Peripheral Dilatation Catheter |  |  | April 09, 2014 | 2014 | Ev3, Inc. | Covidien is recalling 2 production lots of EverCross 0.035 PTA Balloon Catheter because of a product labeling error. There is a discrepancy in the PTA Balloon length and the delivery catheter length printed on the product labeling. | Labeling Change Control | Consignees were sent a Covidien "Urgent Product Recall" letter dated April 09, 2014. The letter described the problem and the product involved in the recall. The letter described the required actions which included to immediately quarantine and discontinue use of the affected devices and to return affected product. Requested consignees to complete the Verification Form and return it to their Covidien Sales Representative. For questions they can contact Covidien Customer Service at 1-800-716-6000. |  | 39 | 39 units (31 OUS, 8 US) | Worldwide Distribution-USA including the states of WV, TX, OK, IN, NY, TX, and NC, and the countries of AUSTRALIA, ISRAEL, TURKEY, UNITED KINGDOM, ITALY, FRANCE, and CHINA. | Terminated | October 06, 2014 | 180 | 39 | Not_Computer | Not_Computer
Z-2591-2014 | 69022 | No | MRD |  |  | General &amp; Plastic Surgery | 510(k) | EN MENTOR Round High Profile Single Use Saline Breast Implant Sizer Single Use Saline Breast Implant Sizer |  |  | July 02, 2014 | 2014 | Mentor Texas, LP. | The device labeling is stamped with 330cc and the device is actually 380cc device. | Error in labeling | Mentor sent an Important Customer Notification letter dated July 17, 2014 to all affected customers and follow-up with a visit by the local sales representative. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immediately identify all unsused product from lot list provided and maintain a copy of this notice with the affected product. For question contact your local Sales Representative or Mentor Customer Service at 1-800-235-5731. |  | 34 | 34 units | U.S Nationwide Distribution in the states of Michigan, Maryland, Texas, New Jersey, Indiana, Nevada, Florida, California, New York, Tennessee, Pennsylvania, and Colorado. | Terminated | April 06, 2015 | 278 | 34 | Not_Computer | Not_Computer
Z-1441-2014 | 67752 | No | DSA | Patient transducer and electrode cable (including connector). | Cardiovascular | Cardiovascular | 510(k) | FAS-LOC 6 Foot Disposable Extension Cable with Safe Connect/Screw Down. |  |  | March 05, 2014 | 2014 | Remington Medical Inc. | Remington Medical discovered an error on the FL-601-97 IFU; part number IFU-RM-0154. Specifically, the polarity for the black chuck in the picture is wrong. It shows (+), but should show (-). | Labeling Change Control | Remington Medical Inc sent an Field Safety Notice dated March 5, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Please be aware that the typographical error described above is included in the FL-601-97 Instructions for Use (RM-0154 Rev 05) provided with the lots that you have received as identified below. RMI has corrected the typographical error in the revised FL-601-97 Instructions for Use (RM-0154 Rev 5.1)which is attached to this communication. Please incorporate the information contained in the attachment with the existing FL-601-97 labeling (Instructions for Use) provided with the product. If you have distributed these products further, please forward copies of this notice and the revised Instructions for Use to anyrecipients as appropriate.If you have any questions regarding this communication, please contact RMIs sales representatives or customer service team at 800-989-0057, Monday - Friday, between the hours of 8 am - 4pm EST.Patient safety is the highest priority for Remington Medical, Inc. We appreciate your assistance with this matter and regret any inconvenience this may cause. |  | 41499 | 41,499 cables | USA (nationwide) Distribution including Canada | Open, Classified |  |  | 41499 | Not_Computer | Not_Computer
Z-0920-2014 | 66953 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius Bicarbonate Jug Adapter to the 2008 Series Hemodialysis machine |  |  | December 17, 2013 | 2013 | Fresenius Medical Care Holdings, Inc. | Bicarbonate Jug Adapter distributed prior to510(k) approval | No Marketing Application | Fresenius notified customers by telephone and Formal notification on 12/17/13 via the Urgent Bicarbonate Jug Adapter Recall letter sent certified mail with signature confirmation and fax-back form.Customers were instructed to examine their stock immediately to determine whether they have any of the affected parts on hand. If customers have the affected parts, they are instructed to discontinue use and place all Bicarbonate Jug Adapters in a secure area for return to Fresenius Medical Care Renal Therapies Group, LLC (FMC-RTG). Customers instructed to contact their Fresenius Medical Care Technical Services Team for instructions on how to return the recalled product. |  | 703 | 703 units | Nationwide Distribution. | Open, Classified |  |  | 703 | Not_Computer | Not_Computer
Z-2661-2014 | 69182 | No | KOD | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Rusch, Soft Rubber (Nelaton and Nelaton Robinson), Bladder Catheter, Sterile. |  |  | September 04, 2014 | 2014 | Teleflex Medical | Misbranding: Although the product labeling identifies the catheters as &quot;soft rubber,&quot; the Natural Rubber Latex caution statement was not printed on the product label. | Labeling Change Control | Consignees were notified via letter on/about 09/04/2014. Subsequently, consignees were notified of the expanded recall on 01/13/2015. |  | 193363 | 193,363 ea. + 1,140 units (total 194,503 units) | Worldwide Distribution-USA (nationwide) and the countries of Austria, Belgium, Bahrain, Belgrad, Canada, Croatia, Cypress, Czeck Republic, Denmark, Finland, Germany, Great Britain, Hungary, Ireland, Italy, Libya, Netherlands, Norway, Romania, Russia, Saudi Arabia, Slavonia, Slovakia, Spain, Switzerland, Turkey, Portugal and Zambia. | Open, Classified |  |  | 193363 | Not_Computer | Not_Computer
Z-2570-2014 | 68754 | No | LKF | Obstetric-gynecologic specialized manual instrument. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | CooperSurgical H/S CATHETER (PART # 61-5007) |  |  | July 03, 2014 | 2014 | CooperSurgical, Inc. | Unsealed pouches. | Packaging process control | Customer notification letter was sent on July 1, 2014 via UPS with confirmed delivery receipt. The letter identified the product, the problem, and the action needed to be taken by the customer. At CooperSurgical's expense, arrangements were made to replace any of the affected products that customers may have on hand. For questions contact James Keller, Vice President, Regulatory Affairs/Quality Assurance at 203-601-5200. |  | 1400 | 1,400 units | Worldwide Distribution -- USA, in the states of CA, TX, IL, MI, MS, ID, NY, FL, MD, PA, OR, GA, CO, WA, NJ, HI, and NC, including DC; and the countries of the countries of Spain, England, Holland, and United Kingdom. | Terminated | May 28, 2015 | 329 | 1400 | Not_Computer | Not_Computer
Z-2595-2014 | 69066 | No | LKB |  |  | General Hospital | Enforcement Discretion | DualCap IV Pole Strips Disinfectant Caps for Luer Access Valves, REF 500-LB, Sterile. |  |  | August 05, 2014 | 2014 | Catheter Connections, Inc. | Catheter Connections is recalling DualCap IV Pole Strips disinfectant caps for luer access valves due to potential improper packaging cap seal which may render the sealed cap non-sterile. | Packaging | The firm, Catheter Connections, sent an "Urgent: Medical Device Field Action" letter dated August 5, 2014 to its consignees/customers. The letter described the product, problem and actions to be taken. The consignees/customers were instructed to identify any customers who received this product having lot no. 1406509; make all potential users in your facility are made aware of this notification and the recommended actions; forward the attached URGENT MEDICAL DEVICE CORRECTION LETTER to customers immediately via fax or email; search your inventory and instruct all users to check each cap coming from the specific affected lot; quarantine and discard any defective product; and complete and return the attached URGENT MEDICAL DEVICE FIELD ACTION REPLY FORM via fax to 1-888-862-2693 or email to: info@cathconn.com.Catheter Connections does not require the return of devices having the affected lot number.If you have questions, please contact Catheter Connections at 1-801-906-0820 (M-F, 8:00 am - 5:00 pm MST) and ask for Quality Department. |  | 24910 | 24,910 units | US Distribution to states of: CT, GA, MD, and PA.. | Terminated | October 08, 2014 | 64 | 24910 | Not_Computer | Not_Computer
Z-2668-2014 | 68874 | No | BRZ | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Pedi-Pak Pedi-Syringe Filter 60 mL Becton Dickinson/Monoject Syringe |  |  | July 17, 2014 | 2014 | Genesis BPS, LLC. | Genesis BPS is recalling numerous devices since they did not notify the agency of its intent to introduce material changes. | No Marketing Application | Genesis BPS sent an Urgent Medical Device Recall letter with a Recall Letter/Confirmation Form on July 17, 2014 to all affected customers. Customers were asked to immediately examine their inventory and quarantine product subject to recall. In addition, if there is further distribution of the product, customers should be identified and be notified of the recall. Product use should be discontinued and disposed of according to applicable federal regulations and internal protocols.The attached confirmation form should be completed and returned to Genesis by either fax or (201) 708-1104 or e-mail (info@genesisbps.com). Questions can be directed to (201) 708-1400 ext 2004 (9:00 and to 5:00 pm EST).For questions regarding this recall call 201-488-1174. |  | 7272 | 7272 units | Nationwide Distribution &amp; Switzerland | Terminated | March 04, 2016 | 596 | 19440 | Not_Computer | Not_Computer
Z-2645-2014 | 69117 | No | OLI | Low level laser system for aesthetic use | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Ultraslim. These include all products labeled as UltraSlim, Luxurian, BioSlim Lipo, Cellulize, Slim 'n Trim, UltraRed, UltraBlue, Slim-PRO, Dazzle Light, UltraSlim Cold Light, SlenderRay, and U-Slim It, as well as those products with private labelling branding. |  |  | May 06, 2014 | 2014 | Advanced Photonic Systems, LLC | Marketing without marketing clearance or approval. | No Marketing Application | Advanced Photonic Systems sent an Urgent Medical Device Recall letter dated July 12, 2014 via UPS to all those who are known by APS to have possession, custody, or control of the affected device. Customers were instructed to return the recalled product at the sole expense of APS. Each 90 days for the following year, APS plans to provide a progress report to the FDA that identifies which recipients have complied. For more information , customers may write to Advanced Photonic Systems, LLC at the address provided or call 1-800-210-5060 (9am to 5pm, Monday through Friday).For questions regarding this recall call 800-210-5060. |  | 235 | 235 Total | Worldwide Distribution - US- AL, AZ, CA, CO, FL, GA, ID, IL, IN, LA, MA, MD, MI, MN, MO, NV, NJ, NY, NC, OH, PA, TX, UT, and VA.Internationally to Canada, France, and Switzerland | Terminated | June 24, 2015 | 414 | 235 | Not_Computer | Not_Computer
Z-1472-2014 | 67803 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Fassier-Duval IM Telescopic System. Orthopedic intramedullary rod implant. Catalogue number FD-040(L)-SS. |  |  | March 05, 2014 | 2014 | Pega Medical Inc. | The outer package label contained the incorrect length of the device. There is a label on the outer package and an additional label on the product contained in the package. | Error in labeling | There is an error on the outside label of the implant. The length was mistaken on the label. It was indicated 245mm instead of 338mm. It has to be clarified that the Fassier-Duval implant is provided with two labels, one that is outside the tube used for packaging and one inside the tube. |  | 48 | 48 | Worldwide distribution, USA nationwide, Greece, United Arab Emirates, Peru, Canada, Israel, United Kingdom, France, Germany, Colombia, Mexico, Singapore, Lithuania, Denmark, and Spain. | Terminated | May 19, 2014 | 75 | 48 | Not_Computer | Not_Computer
Z-2625-2014 | 69043 | No | LGM | Patient care reverse isolation chamber. | General Hospital | General Hospital | 510(k) | Life Force Chamber |  |  | June 27, 2014 | 2014 | Life Force of Tampa, LLC | The Life Force chamber is marketed and promoted as a medical device, intended for use in the diagnosis of disease or other conditions or in the cure, mitigation, treatment, or prevention of disease, or is intended to affect the structure or function of the body. Life Force of Tampa does not have an approved application for premarket approval in effect, or an approved application for an investigat | Labeling False and Misleading | Life Force of Tampa LLC sent an URGENT: MEDICAL DEVICE RECALL letter dated June 27, 2014 to all affected customers. The letter instructed the customers to: 1) Immediately remove and discard all copies of the labeling material. 2) Remove all electronic copies of the materials from any publicly accessible locations (e.g. website). 3) Complete and return the Acknowledgment form to Life Force of Tampa. For questions or concerns contact Douglas G. Cornell at (813) 961-4655. |  | 10 | 10 | Nationwide Distribution in the states of AZ, FL, MD, NV, OK, PA, and TX. | Terminated | July 30, 2015 | 398 | 10 | Not_Computer | Not_Computer
Z-2655-2014 | 69054 | No | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | B.Braun Accuflo 100ML/H 100ML Elastomeric Infusion Device. |  |  | August 06, 2014 | 2014 | B. Braun Medical, Inc. | The primary packaging label (inner pouch label) indicates that the package contents are Catalog Item: CT-1000-250, while the secondary packaging label (outer carton label) indicates the contents are CT-1000-100. The actual contents are a CT-1000-100. | Labeling Change Control | B. Braun Medical Inc. provided the recall notification received from B. Braun Melsungen AG to Progressive Medical Inc. who is the initial importer of the product in the U.S. market. The notification was sent via FedEx and e-mail on/about August 6, 2014 and identified the product, problem, and action to be taken. The notification indicated that if the affected units were further distributed that any distributors are required to forward the letter on to their consignees. |  | 8640 | 8,640 units | Distributed in the state of Missouri. | Terminated | January 29, 2015 | 176 | 8640 | Not_Computer | Not_Computer
Z-1463-2014 | 67835 | No | ONG | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | FRAXEL DUAL 1550/1927 Laser System, (FRAXEL SR 1500 DUAL Laser System), for dermatological procedures. |  |  | March 20, 2014 | 2014 | Solta Medical, Inc. | Obsolete versions of the Laser Treatment Settings card for the Fraxel DUAL 1550/1927 Laser system may have contained incorrect indication for use for treatment of Melasma. | Labeling Change Control | Solta Medical sent an Correction Important Product letter dated March 20, 2014 to affected customer. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete and return the enclosed response form. |  | 1077 | 1077 units | Worldwide distribution | Terminated | April 27, 2015 | 403 | 1077 | Not_Computer | Not_Computer
Z-0901-2014 | 67206 | No | BWN | Manual operating table and accessories and manual operating chair and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Radial Assist RAD BOARD, used to support the |  |  | December 18, 2013 | 2013 | Merit Medical Systems, Inc. | Merit Medical Systems, Inc. is voluntarily recalling one lot (B507171) of RAD BOARD RB 100 devices. The affected devices are missing the main label which graphically depicts patient orientation on the board. The issue may allow the board to be oriented in the wrong direction where the semi-radiopaque material in the board (embedded Xenolite) could obscure patient anatomy during fluoroscopy. There | Labeling Change Control | Merit Medical Systems, Inc.sent a letter voluntarily recalling one lot (B507171) of RAD BOARD" RB100 devices. The affected devices are missing the main label which graphically depicts patient orientation on the board. The issue may allow the board to be oriented in the wrong direction where the semi-radiopaque material in the board (embedded Xenolite) could obscure patient anatomy during fluoroscopy. There have been no reports of patient harm or injury from Merit customers as a result of this issue. Merit has chosento remove this lot from the field by replacing the unlabeled RAD BOARDs. Merit requests that customers stop using these devices immediately until the product has been replaced by their Merit sales representative.Customers with questions were instructed to call 1-801-208-4344 or 1-801-208-4365.For questions regarding this recall call 801-208-4301. |  | 11 | 11 | Worldwide Distribution - USA including NJ, TX, WV, MN, GA, and Internationally to Mexico and Japan. | Terminated | February 13, 2014 | 57 | 11 | Not_Computer | Computer
Z-1465-2014 | 67925 | No | OEZ | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | Enforcement Discretion | Terumo Cardiovascular Procedure Kit- X-Coated Pump and Table Pack with FX25RE |  |  | March 28, 2014 | 2014 | Terumo Cardiovascular Systems Corp | Incorrect customer notification label applied to the Tyvek lid of this specific lot of TerumoÂż Cardiovascular Procedure Kits | Labeling mix-ups | Terumo CVS contacted the affected user on 3/28/14 by phone via phone script to discontinue use and return all affected units in inventory. Terumo CVS will replace or issue credit for returned product. |  | 37 | 37 units | CA | Terminated | December 12, 2014 | 259 | 37 | Not_Computer | Not_Computer
Z-2640-2014 | 69083 | No | MGQ |  |  | General &amp; Plastic Surgery | 510(k) | Amerigel Post-Op Surgical Kits . Kit includes: Amerigel Wound Dressing, Amerigel Wound Wash, 22x2 Gauze Pads, and 1x3 Fabric Bandages |  |  | July 11, 2014 | 2014 | Amerx Health Care Corp. | Not approved labeling claims (antimicrobial and autolytic debridement) | Labeling design | On July 21, 2014, Amerx Health Care Corporation telephoned all affected customers advising them of the recalled products. During the call the customers were instructed to: 1) Discontinue dispensing the lot(s) and promptly return via parcel post to the Clearwater facility. 2) If you dispose of the product on your own, provide proof of disposal. 3) You will be reimbursed by check or credit memo for returned or disposed goods and postage. |  | 116633 | 116,633 units | Worldwide Distribution-USA (nationwide) including the District of Columbia, Virgin Islands, Puerto Rico, and Guam, and the states of AK, AL, AZ, CA, CO, CT, FL, GA, HI, IL, IN, KS, KY, LA, MA, ME, MD, MI, MT, MO, NC, NH, NE, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WY, and WV, and the countries of Canada, England, Africa, and Australia. | Terminated | July 22, 2015 | 376 | 522101 | Not_Computer | Not_Computer
Z-2636-2014 | 69049 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Volcano TrakBack II Disposable Pullback Device; |  |  | August 13, 2014 | 2014 | Volcano Corporation | Internal testing had identified a potential breach to the sterile barrier that may impact the sterility of the product. | Packaging | Volcano sent an Urgent Product Recall Notification letter dated August 13, 2014, to all customers who received the device in the last 25 months. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately cease use of the affected product and return unused devices to the Volcano distribution center. Customers were asked to prepare the attached inventory log and then contact Volcano Customer Service at 800-228-4728, option 4 to arrange for return of any affected product in their inventory. A credit would be issued for any returned product. Customers with questions were asked to call 800-228-4728.For questions regarding this recall call 978-439-3586. |  | 8251 | 8251 total devices (3777 US) | Worldwide Distribution - USA (nationwide) and Internationally to Korea, Thailand, Canada, Australia, Chili, Mexico, Argentina. | Terminated | June 30, 2015 | 321 | 8251 | Not_Computer | Not_Computer
Z-2673-2014 | 69047 | No | NGT | Intravascular catheter. | General Hospital | General Hospital | 510(k) | BD PosiFlush SF Saline Flush Syringe 0.9 Sodium Chloride Injection, USP 10 mL REF 306553. |  |  | July 22, 2014 | 2014 | Becton Dickinson &amp; Company | BD received some reports of open seals found on the BD Posiflush SF Flush Syringe 10 mL (Ref #306553). | Packaging | BD sent recall letters/recall response cards dated 7/22/2014 via e-mail or UPS 2nd day mail. Customers are asked to examine their inventory and quarantine any affected product subject to recall. Customers should complete the Recall Response Card and fax it back to BD at 1-201-847-6990. Any questions can be directed to 1-888-237-2762 option 3. |  | 3088320 | 3,088,320 units | Nationwide Distribution. | Terminated | April 21, 2015 | 273 | 3088320 | Not_Computer | Not_Computer
Z-2550-2014 | 68818 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Trauma Tubes and Plates: |  |  | July 22, 2014 | 2014 | Zimmer, Inc. | Zimmer is recalling sterile-packaged implants and instruments packaged in six different package configurations due to packaging design verification test failures. Specifically, multiple test samples fromeach of the six configurations failed simulated distribution and shipping testing when the devices either breached inner or outer package trays or seal integrity was lost. | Package design/selection | Zimmer sent an URGENT MEDICAL DEVICE RECALL letter dated July 2014 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Please call the Customer Call Center at 1-800-348-2759 for questions or concerns involving this notice. |  | 972125 | 972,125 units | Worldwide Distribution. | Terminated | September 15, 2015 | 420 | 972125 | Not_Computer | Not_Computer
Z-2290-2014 | 68753 | [6 empty] | PENTAX Upper/Lower G.I. Endoscope OF-B130 and OF-B194 Gas/Water Feeding Valve |  |  | July 17, 2014 | 2014 | Pentax Medical Company | The Instructions for Use for the Pentax OF-B130 and OF-B194 Gas/Water Valve IFU if not followed carefully may cause a potential risk to health. | Labeling design | Pentax Medical sent an Urgent Medical Device Recall Field Safety Corrective Action letter dated July 16, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.If the Instruction for Use are not followed carefully by the user there is a potential risk to health for the patient, such as a perforation.Customers are asked to complete the attached response form and fax to (201) 799-4063. Customers should see the attached revised Instructions for Use Z632 R09. If you have any questions regarding this action, please feel free to contact your Sales Representative at : 800-431-5880 ext 2064 |  | 1206 | 98 units US; 1108 units OUS | Worldwide Distribution - US Distribution including the states of CA, CO, IN, KS and ID., and the countries of Asia, Canada, Europe and Japan. | Terminated | August 20, 2015 | 399 | 1206 | Not_Computer | Not_Computer
Z-2689-2014 | 68955 | No | BZE | Breathing system heater. | Anesthesiology | Anesthesiology | 510(k) | CareFusion AirLife Heated Infant Breathing Circuit |  |  | July 31, 2014 | 2014 | Carefusion 2200 Inc | The CareFusion AirLife Heated Infant Breathing Circuit is being recalled due to a regulatory compliance risk involving a material change. The changes to the gas pathway may potentially impact the safety of the device and quality of the gas condensates to the patient. | No Marketing Application | CareFusion sent an URGENT: PRODUCT RECALL letters dated July 31, 2014 to all affected customers. The letters instructed customers to: 1) perform a 100% physical inventory to verify if any of the affected product codes are in stock; 2) complete and return the attached Recall Response Form within 15 days; 3) notify any customers/parties that the recalled products were distributed to of the recall; 4) distributors should have all their customers return the products and a completed Recall Response Form to the distributor for credit; and, 5) customers should destroy the recalled products according to their disposal procedure. If a customers procedure(s) does not allow for the destruction of the recalled products, the customers was advised to contact CareFusion AirLife Customer Support at 800-323-9088 (Monday - Friday, 8:00 AM - 5:00 PM CST) for product return options. Customers will receive credit for all destroyed products. Direct accounts with questions about this recall should contact CareFusion AirLife Customer Support at 800-323-9088 (Monday - Friday, 8:00 AM - 5:00 PM CST) for additional information. |  | 318280 | Approximately 318,280 devices | Worldwide Distribution - US Nationwide in the states of AZ, CA, FL, IL, KY, LA, MI, MO, NE, NY, OH, OK, PA, RI, TN, TX, UT including Puerto Rico and the countries of Canada, China, Kuwait, Mexico, Saudi Arabia, and Switzerland. | Terminated | February 08, 2016 | 557 | 318280 | Not_Computer | Not_Computer
Z-2171-2014 | 68814 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | EndoWrist Stapler System User Manual Addendum and Wall Chart; |  |  | June 02, 2014 | 2014 | Intuitive Surgical, Inc. | Changes were made to the Stapler EndoWrist Stapler 45 System Instruments and Accessories User Manual Addendum (550986-03) and the EndoWrist Stapler 45 Wall Chart (PN 551524-01) in which the pH range specification for the enzymatic cleaner used throughout the cleaning process was modified from pH-neutral to pH-neutral to mild alkaline (mild alkaline is defined as pHd11). Mild Alkaline enzymatic | Labeling Change Control | Intuitive Surgical representatives hand delivered corrected manuals to the two affected sites.. |  | 2 | 2 | Distribution to Florida and South Carolina | Terminated | August 08, 2014 | 67 | 2 | Not_Computer | Not_Computer
Z-1525-2014 | 66841 | No | MHY |  |  | Neurology | PMA | Medtronic Activa RC (37612), Activa PC (37601), Activa SC (37602), Activa SC (37603), Activa PC+S (37604), Multi-program Rechargeable Neurostimulator for Deep Brain Stimulation for movement disorders. |  |  | February 27, 2014 | 2014 | Medtronic Neuromodulation | Medtronic neurostimulators (used with pain and deep brain stimulation therapy) have a cycling feature that, when enabled, allows stimulation output to be toggled on and off based on the programmed settings. The use of this feature is determined by the physician and set using the clinician programmer. Current labeling indicates the use of cycling improves device longevity and recharge interval | Labeling False and Misleading | Medtronic sent a Medical Device Correction letter dated February 2014 to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to do the following:If you have patients with cycling enabled or wish to use cycling, use the updated calculations and graphs available online at professional.medtronic.com/cycling to assess the impact of cycling on device longevity. Medtronic will update the product labeling in the second half of 2014." Remind your patients to continue to check their battery status. Instructions on how to check battery status can be found in the patient programmer or recharger system manual." Pain Stimulation Trialing Only: The cycling feature should not be enabled on the External Neurostimulator (ENS) when estimating device longevity because the resulting longevity estimate for non-rechargeable devices may not provide accurate information.Consignees were also instructed to acknowledge receipt of the notice by completing the attached reply form and returning it to Medtronic using the contact details on the reply form. For questions consignees can contact Medtronic Neuromodulation Technical services at 1-800-707-0933 weekdays 7am - 6pm CST.For questions regarding this recall call 800-707-0933. |  | 244207 | 170,286 US, 73,921 OUS for all devices | Nationwide Distribution - all states including DC and Puerto Rico. | Terminated | November 07, 2014 | 253 | 244207 | Not_Computer | Computer
Z-2734-2014 | 69053 | Yes | CBF | Hyperbaric chamber. | Anesthesiology | Anesthesiology | 510(k) | The Rejuvenator Chamber manufactured by Baro-Therapies, Inc. |  |  | August 08, 2014 | 2014 | Baro-Therapies, Inc | The firm is distributing the Rejuvenator device without an approved 510(k). | No Marketing Application | Baro-Therapies issued a Correction Notice letter on August 11, 2014, to all affected customers. The letter stated they have suspended marketing of the device and plans to shut down their website until they receive clearance from the FDA. For questions or concerns contact the firm at (850) 654-0051. |  | 10 | 10 | US Nationwide Distribution in the states of FL, GA, MS, TN and CO. | Terminated | August 06, 2015 | 363 | 10 | Not_Computer | Not_Computer
Z-2176-2014 | 68798 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Agfa Digital Radiography X-Ray System DX-D 100 |  |  | May 13, 2014 | 2014 | AGFA Healthcare Corp. | Unit was mislabeled with a factory label showing 32 KW output power. | Labeling mix-ups | Only one customer was affected no customer letter was issued field correction was conducted.. |  | 1 | 1 | US in the state of WV | Terminated | August 13, 2014 | 92 | 1 | Not_Computer | Not_Computer
Z-2181-2014 | 68803 | No | PDH | Neonatal phototherapy unit. | General Hospital | General Hospital | 510(k) | Luma Wrap Phototherapy Swaddler is designed to provide centered and comfortable boundaries as a benefit for restive babies in the Neonatal Intensive Care Unit. |  |  | July 11, 2014 | 2014 | Beevers Manufacturing &amp; Supply, Inc. | Beevers Manufacturing &amp; Supply Inc. is recalling Luma Wrap neonatal phototherapy blankets because the product was inadvertently misclassified and marketed without a cleared 510(k) submission from the US FDA. | No Marketing Application | BMS is contacting their customers and requesting that distributors report the amount of affected product they have in stock, prior to returning affected product. BMS is also asking its distributors to forward the recall notification to its customers to whom affected product was sold. BMS expects those customers to also report to BMS the amount of affected product they have. |  | 22080 | 240 total (112/ lot 66812; 92/ lot 66813; 17/ lot 30660; and 19/ lot 67879). | Product was distributed to US, Germany, Austria, France, Italy, Spain, Australia. | Terminated | July 17, 2015 | 371 | 22080 | Not_Computer | Not_Computer
Z-2729-2014 | 68725 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Reusable laparoscopic instruments are packaged as non-sterile. Cleaning and sterilization of devices must occur prior to use. |  |  | May 30, 2014 | 2014 | Instrumed International, Inc. | Additional caution statements have been added to the Instructions For Use for monopolar electrodes. 1) To avoid tissue carbonation, the operation voltage of the HF generator must not exceed 650 peak voltage (Vp) for all monopolar electrodes. and 2) For all monopolar electrodes, the useful life of this product is less than or equal to 50 cycles and less than or equal to 2 years. | Labeling design | Instrumed sent an URGENT: MEDICAL DEVICE CORRECTION letters dated May 30, 2014 to all consignees. The letters advised the consignees of the new warning statements on the product Instructions For Use and an instruction to complete and return the attached Response Request Form to Instrumed International. Consignees were further instructed to extend the recall to the retail level. For questions contact Instrumed Quality Manager at 847-908-6119. |  | 15 | 15 instruments | Distribution US Nationwide in the states of: AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, and WY. | Open, Classified |  |  | 4504 | Not_Computer | Not_Computer
Z-0622-2014 | 67025 | No | JXG | Central nervous system fluid shunt and components. | Neurology | Neurology | 510(k) | Integra Large Rickham -Style Reservoir |  |  | November 13, 2013 | 2013 | Integra LifeSciences Corp. | Integra became aware through an adverse complaint trend that Rickham Style Reservoirs with bent or broken connectors that were shipped in sleeve-type packages were found prior to use but after the sterile package was open. | Packaging | Integra sent an Urgent Voluntary Medical Device Recall letter dated November 13, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to inspect their inventory for the affected product and if located stop using immediately. Customers were asked to complete the attached form and return it to Integra as indicated and keep a copy of the form for their records. Once the form was received by Integra they will send an RMA number, directions to return the product, and input an order to replace the quantity they indicated on the form.Customers with questions were instructed to call 1-855-532-1723.For questions regarding this recall call 855-532-1723. |  | 78 | 78 packages | Worldwide Distribution - USA including FL, NE, WI, WA, CT, ND, AL, NY, NH, OR, OK, MN, and TX. Internationally to Belgium, Czech Republic, Germany, France, Great Britian, Italy, Poland, and Portugal. | Terminated | February 13, 2015 | 457 | 272 | Not_Computer | Not_Computer
Z-0627-2014 | 67093 | No | ODP | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | C-Plus" 10mm x 12mm Webless implants |  |  | October 07, 2013 | 2013 | RTI Surgical, Inc. | As the result of a recent internal review of regulatory documents, we have determined that the 10 mm x 12 mm Webless C-Plus implants were incorrectly documented as meeting the requirement for labeling as a partial Vertebral Body Replacement Device (VBR). While the use of this device as a partial VBR is unlikely due to the size of the device, use of this device as a partial VBR could result in a s | Labeling False and Misleading | RTI Surgical, Inc. sent an Important Notice letter dated October 8, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed remove fall affected product from their inventory and send back to Customers Service at 375 River Park Circle, Marquette, MI 49855 by October 16, 2013. Customers with questions were instructed to call 906-226-4489.For questions regarding this recall call 906-226-4812. |  | 1906 | 1906 | Nationwide Distribution - USA including FL, MO, NY, CA, MI,and IL. | Terminated | February 26, 2014 | 142 | 1906 | Not_Computer | Not_Computer
Z-0071-2015 | 69333 | No | DKZ | Amphetamine test system. | Toxicology | Toxicology | 510(k) | ArcPoint Labs 5 Panel Dip Drug Screen, Part No. APD-5M. |  |  | September 23, 2014 | 2014 | Ameditech Inc | Ameditech is recalling the ArcPoint Labs Dip Drug Screen Dip Card because the positive interpretation illustration was incorrect on the cap labels as well as on the card cap graphic in the package insert. | Labeling Change Control | A customer notification letter dated 9/23/14 was sent to all customers who purchased the ArcPoint Labs Dip Drug Screen Dip Card. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to complete and return the Verification Form. Customers with questions are instructed to contact Ameditech at (858) 805-8714 or email at amd.complaint@alere.com. |  | 400 | 400 kits | Distributed in AL. | Terminated | February 27, 2015 | 157 | 800 | Not_Computer | Not_Computer
Z-0986-2014 | 67390 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Synthes 3.2 mm Guide Wire 400mm. |  |  | January 27, 2014 | 2014 | Synthes, Inc. | Synthes 3.2 mm Guide Wire 400mm included an incorrect raw material listed on the label. | Labeling Change Control | A recall notification letter, dated January 27, 2013, was sent to End Users and sales consultants. |  | 1132096 | 1,132,096 | Worldwide Distribution-USA (nationwide) and the countries of Canada and Switzerland. | Terminated | December 18, 2015 | 690 | 1132096 | Not_Computer | Not_Computer
Z-2215-2014 | 68654 | No | JDG | Hip joint femoral (hemi-hip) metallic cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Moore Hip Prosthesis Long Fenestrated Stem 38 mm (1-1/2 Inch) Head Diameter 203 mm Length. |  |  | July 01, 2014 | 2014 | Zimmer, Inc. | Zimmer is initiating a voluntary recall of 20 production lots of Moore Hip Prosthesis Long Fenestrated Stems due to packaging design verification test failures. Specifically, 9 of 16 test samples failed simulated distribution and shipping testing when the Moore hip stem breached either the inner or outer package tray. Zimmer has not received any complaints of damaged package trays for these 20 lot | Packaging process control | On July 1, 2014 Zimmer, Inc. issued recall notification to their consignees via letter. This notice was issued to affected distributors and hospitals with instructions for returning the affected product. Distributors were sent a letter identifying the issue and their responsibilities. These responsibilities include locating and removing the affected product in their territory, as well as identifying hospitals and surgeons who have previously implanted the affected product and providing related contact information (name, address, and phone number). Distributors are to return on-hand affected products to Zimmer and ensure all of their products are accounted for using the form provided in the letter. Hospital risk managers and surgeons were provided with a letter identifying the issue and their responsibilities. These responsibilities include locating the affected product and returning it to their Zimmer sales representative. |  | 68 | 68 | Worldwide Distribution - US (nationwide) in the states of IN, WI, MN, OR, and TX, and one foreign consignee in Canada. | Terminated | January 09, 2015 | 192 | 68 | Not_Computer | Not_Computer
Z-2233-2014 | 68892 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | BD Multi Check CD4 Low Control Catalog No. 340916 is intended as a complete process control for immunophenotyping by flow cytometry. |  |  | July 18, 2014 | 2014 | BD Biosciences, Systems &amp; Reagents | The outer box label contains the incorrect distribution date: Box states 2017-09-02 ; should be 2014-09-02 | Labeling mix-ups | Global hold initiated on July 18, 2014. Customer letters sent by certified mail on July 18, 2014 to advised users of the problem. |  | 85 | 85 | IN, MA, TX, VA, Ca, KY, MD, PA, GA, and BD locations in Singapore and Belgium | Terminated | May 19, 2015 | 305 | 85 | Not_Computer | Not_Computer
Z-0975-2014 | 67384 | No | KKY | Polytetrafluoroethylene with carbon fibers composite implant material. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Stryker CMF Customized Implant Kit X-Large; 5444-0-401 X-Large Implant (2) |  |  | December 06, 2013 | 2013 | Stryker Howmedica Osteonics Corp. | Stryker received a report from the field that an incorrect patient-specific Simplex P CMF Customized Implant Kit was delivered to a hospital. | Labeling Change Control | Stryker Orthopaedics sent Urgent Medical Device Recall Notification letters/Product Recall Acknowledgement Forms dated 12/6/2013 to US branch/agency. surgeons, hospital risk managers, and sales representatives. |  | 1 | one unit | Distributed in PA. | Terminated | May 06, 2015 | 516 | 1 | Not_Computer | Not_Computer
Z-2719-2014 | 69025 | No | GZB | Implanted spinal cord stimulator for pain relief. | Neurology | Neurology | 510(k) | Percutaneous Trial Lead Kit Model 3086 (8 ch percutaneous lead). |  |  | July 29, 2014 | 2014 | St. Jude Medical, Inc. | The Directions for Use (DFU) manual contains information for devices that are not approved in the US. The correct US approved versions of the manuals were sent to the customers to replace the incorrect international manuals. | Packaging process control | A field action was initiated with customer letters, dated July 29, 2014, and attached replacement DFUs. The letter identified the affected product and the reason for the recall. The recalling firm requested the consignees refer to the replacement DFU for guidance while using the device, and to discard the international manuals. Questions regarding the letter should be directed to customers' St. Jude Neuromodulation Representative or call the Technical Services Department at 1-800-727-7846. |  | 4 | 4 units | Distributed in the states of Texas and Illinois. | Terminated | August 21, 2015 | 388 | 6 | Not_Computer | Not_Computer
Z-0977-2014 | 67267 | No | MBP | Resorbable calcium salt bone void filler device. | Orthopedic | Physical Medicine | 510(k) | Medtronic GRAFTON (R) DBM, Flex 2.5cm x 10cm, REF T42100, Rx only, STERILE A, QTY: 1 EA, Medtronic Sofamor Danek, USA, Inc., 1800 Pyramid Place, Memphis, TN 38132 |  |  | January 06, 2014 | 2014 | Medtronic Sofamor Danek USA Inc | Product was labeled as being measured as 2.5 cm x 10 cm, but contained product that measured as 2.5 cm x 5 cm. | Packaging process control | The firm, Medtronic, sent an "URGENT-V voluntary Market Removal (Recall)" communication letter and questionnaire packet (via FedEx) dated January 3, 2014 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to remove the product from normal storage locations; Do not use the product; and as applicable, contact your Medtronic Sales Representative to coordinate the return and replenishment of affected product or call our Global Quality Department at 800-876-3133 ext 3197. In addition, the customers were instructed to disseminate this information to additional personnel within your facility as appropriate. A Medtronic Sales Representative will contact you to facilitate the return of any impacted Set Screw Retaining Drivers you may have in your possession.If you need assistance or if you have any related questions or concerns, please do not hesitate to contact your Medtronic Sales Representative or our Global Quality Department at 800-876-3133 extension 3197. |  | 30 | 30 units | US Distribution in the states of: MO, FL, MN, OH, MI, TX, and NJ. | Terminated | June 24, 2014 | 169 | 30 | Not_Computer | Not_Computer
Z-2244-2014 | 68819 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Magna-Fx Cannulated Screw Fixation System |  |  | July 22, 2014 | 2014 | Zimmer, Inc. | Zimmer is recalling various trauma screws that failed a leak test due to damage of the inner cavity only. | Package design/selection | On 7/22/14, Zimmer sent recall notification to the consignees explaining the issue, identifying affected product, providing the health risks, and instructing the consignees what to do for the recall. Questions or concerns, please contact Zimmer at 1-877-946-2761. |  | 47682 | 47,682 units total | National Distribution: AK, AL, AZ, CA, CT, DC, FL, GA, US VIRGIN ISLANDS, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NJ, NM, NV, NY, OH, OK, OR, PA, PR, SC, SD, TN, TX, VA, WA, WI, WV, WYForeign Distribution: Canada, Mexico, ARGENTINA, AUSTRALIA, BELGIUM,BRAZIL, CHINA, EL SALVADOR, ENGLAND, GERMANY, INDIA, ITALY, JAPAN, JORDAN, KOREA, NEW ZEALAND, SAUDI ARABIA, SINGAPORE, SPAIN, SWITZERLAND, TAIWAN, UAE | Terminated | June 18, 2015 | 331 | 47682 | Not_Computer | Not_Computer
Z-0215-2015 | 69619 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | AxoTrack I Sterile Procedure Kit 18G, K1001-18XTW, (24/case labeled for US, 5/case labeled for US, 24/case labeled for EU, 5/case labeled for EU). |  |  | March 24, 2014 | 2014 | Soma Access Systems LLC | There is a possibility the sterile packaging may fail before the expiration date specified on the package. | Packaging | Consignees were notified via USPS Certified Letter, email and FedEx on 3/24/2014. |  | 3530 | 3530 units | Worldwide Distribution-US (nationwide) including the states of AL, AZ, CA, CT, FL, GA, IN, KY, LA, MA, MI, MO, NC, NJ, NY OH, OR, PA, SC, TN, TX, VA, and WA and the countries of Belgium, Hong Kong, India, Japan, Panama and Singapore. | Terminated | August 06, 2015 | 500 | 3530 | Not_Computer | Not_Computer
Z-1508-2014 | 67818 | No | DIO | Cocaine and cocaine metabolite test system. | Toxicology | Toxicology | 510(k) | ToxCup Drug Screen Cup, Catalogue Numbers: HT15, HT15A, PT11, PT11A, PT15A, PT23A, a screening device for the detection of drugs and drug metabolites in human urine. For in vitro diagnostic use only. |  |  | March 26, 2014 | 2014 | Branan Medical Corporation | May be labeled with incorrect result interpretation graphics on the ToxCup lid label. | Labeling Change Control | The firm, Branan Medical Corp (BMC), sent an "URGENT MEDICAL DEVICE RECALL" letter dated March 26, 2014 to all the customers who received the ToxCup Drug Screen Cup. The letter describes the product, problems identified and the actions to be taken. Customers are instructed to discontinue use and/or sale of product; quarantine all units of these product lots; destroy unused product(s) in compliance with your laboratory and state requirements; share this information with your staff and retain notification as part of your QS documentation; if distributed to your customers, provide a copy of this letter to them; complete and return the enclosed Customer Verification Form via Fax to: Branan Medical Corporation at 949-598-7167 or by e-mail at: branan.complaints@brananmedical.com within 10 days. Should you have any questions about the information contained in this notification, contact Branan Medical Corporation at (949) 598-7166 or email branan.complaints@brananmedical.com. |  | 9800 | 9,800 units | Worldwide Distribution: US (nationwide) and country of Canada. | Terminated | July 23, 2014 | 119 | 9800 | Not_Computer | Not_Computer
Z-2253-2014 | 68898 | No | KWY | Hip joint femoral (hemi-hip) metal/polymer cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Pivot Bipolar Cup 22x42/43-51/53mm, REF 133-YYYY |  |  | July 11, 2014 | 2014 | Ortho Development Corporation | Ortho Development is recalling various Pivot Bipolar Cup, Patella, PS Tibial Insert, UC Tibila Insert, and CK Tibil Insert because the foil pouches did not properly seal, allowing air to seep back into the pouch after sealing which increases the risk that the material properties of the implant would be substantially degraded due to oxidation. | Packaging process control | Ortho Development notified all affected consignees/distributors of the recall via email on July 11, 2014 which included a list of the affected product (part numbers, lot numbers and quantities) under their control. An RMA number and return shipping label were also included with the email. Emails were followed up with phone calls to verify receipt of the notification. |  | 188 | 188 | Worldwide Distribution - US Nationwide and the country of Japan. No military/govt/VA consignees. | Terminated | October 06, 2014 | 87 | 227 | Not_Computer | Not_Computer
Z-2264-2014 | 67814 | No | PDZ | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | All brochures printed before November 20, 2013 for the HyperBlue 1530 medical laser, including the product brochure and the online information regarding the product at www.hyperionMed.com before December 3, 2013. |  |  | November 20, 2013 | 2013 | Hyperion Medical | Brochures and web site information for the HyperBlue 1530 diode laser promoted features which had not been cleared in the product's premarket notification to the FDA. The web site and brochures had claims that it could use 940 nm, 808 nm, and 532 nm wavelengths as single wavelengths or in combination with other wavelengths, and a claim that the product could be used with a 15 mm hand piece. | Labeling False and Misleading | Hyperion Medical notified all sales representatives by private courier letter to stop distributing the the sales brochure and return or destroy existing stocks of the brochure. Sales representatives were instructed not to demonstrate the 15 mm hand piece with the HyberBlue 1530 and to remove from any of their individual internet web sites any information about multiple wavelength use or the 15 mm handpiece accessory. The letter to the sales representatives included delivery confirmation. The information on the firm's web site for this product regarding multiple wavelengths and the 15 mm handpiece was removed.For further information please call (214) 914-3133. |  | 3346 | 3,346 brochures. Quantity not applicable to web site. | US Distribution including the states of PA, MD, UT, TX, OH, GA, MO, FL, CO, and WA. | Terminated | May 15, 2015 | 541 | 3346 | Not_Computer | Not_Computer
Z-2694-2014 | 68487 | No | HMT | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | eye-pak 7407 Tray Support Cover REF 8065740745 |  |  | June 03, 2014 | 2014 | Alcon Research, Ltd. | The peel pouches of the Tray Support Covers may be insufficiently sealed so that sterility cannot be assured. | Packaging process control | Alcon Research sent an Urgent Medical Device Recall letter dated June 3, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Please follow these important steps:1) Review your inventory of Tray Support Covers to determine if you have any units from the affected Lots in stock.2) Place the product aside to ensure it is not used.3) Return the attached Response Form via fax or email to Alcon.Please return the Response Form even if you do not have any inventory from the affected lots.4) Alcon will contact you to arrange the return of your inventory.If you have any further questions, concerns, or need assistance with obtaining replacement product, please contact Alcon Customer Service at 1-800-862-5266. |  | 2462 | 2462 | Worldwide Distribution - Nationwide Distribution and the countries of Australia, Belgium, and Canada. | Terminated | July 27, 2015 | 419 | 2462 | Not_Computer | Not_Computer
Z-2266-2014 | 68952 | No | DSY | Vascular graft prosthesis. | Cardiovascular | Cardiovascular | 510(k) | Gelweave (TM) Vascular Prostheses |  |  | July 15, 2014 | 2014 | Vascutek, Ltd. | The cautery that is provided with the Gelweave graft in the box has a shorter shelf life indicated on its label, versus that of the graft (4 years and 5 years, respectively). | Incorrect or no expiration date | Vascutek sent an Urgent Field Safety Notice dated July 15, 2014, to all affect customers. The letter identified the product the problem and the action needed to be taken by the customer. User Instructions:A. Vascutek advises that you ensure suitable equipment is available prior to implantation to trim the graft if required.B. For product received with a cautery in the interim, Vascutek request that you check the shelf life of each product before every operation. If you find that the cautery is out of date, remove and discard the cautery. If you require an appropriate cautery supplied as a separate item, please contact your local sales representative for a referral.C.Vascutek are taking this opportunity to reinforce the cautions from the Gelweave TM Vascular Prostheses IFU regarding soaking of the graft in saline to prevent focal burning if a cautery is to be used. This statement is already included in the IFU and now appears in bold, red font to ensure appropriate awareness.This notice needs to be passed on to all persons who need to be aware within your organization or to any organization where the devices have been transferred or distributed.Please maintain awareness of this Field Safety Notice while all actions are taken in your organization and until these products are provided without the cautery.Please return the User Return Slip by e-mail or fax to the Distributor's address given on page 3.This action by Vascutek Ltd. is being taken with the knowledge of the National Competent AuthorityÂż Medicines and Healthcare Products Regulatory Agency (MHRA).If you have any further questions or comments, please do not hesitate to contact us at tcvs.recall@terumomedical.com. Return Completed Form immediately fax to (734) 741-6149. If required, your Vascutek Ltd. representative can discuss and provide more information on the use of cautery with Vascutek products. For further questions call (800) 262-3304 ext. 6056 |  | 63162 | 63,162 | Nationwide Distribution including the states of : AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI and WV. | Terminated | March 19, 2015 | 247 | 63162 | Not_Computer | Not_Computer
Z-0966-2014 | 66995 | No | GWN | Nystagmograph. | Neurology | Neurology | 510(k) | Normative Data Template CD for use with the I-Portal devices. |  |  | November 21, 2013 | 2013 | Neuro Kinetics, Inc. | The data provided on the Normative Data Template CD for use with the I-Portal devices has not received clearance by FDA through the premarket 510(k) notification process. | No Marketing Application | A recall notification letter, dated November 2013, was sent to Users. A follow up letter regarding scheduling removal of the software is planned for December 2013. |  | 148 | 148 | Distributed USA (nationwide) and the country of Canada. | Terminated | July 20, 2015 | 606 | 148 | Not_Computer | Not_Computer
Z-1452-2014 | 67808 | No | BZO | Ventilator tubing. | Anesthesiology | Anesthesiology | 510(K) Exempt | Reusable latex breathing bags with multiple part numbers and sold as part of test lung part number and reusable anaesthesia sets. |  |  | March 25, 2014 | 2014 | Draeger Medical, Inc. | Reusable latex breathing bags with various part numbers and sold as part of test lung and reusable anaesthesia sets were recalled when Draeger became aware of an incident where a user suffered an acute allergic reaction after using a DrÂżger latex test lung. The user did not suffer permanent injury. The breathing bag that belongs to the test lung and other latex breathing bags are labelled NR .A | Labeling design | A recall notification letter, dated March 2014, was sent to wholesale distributors and end users. |  | 433 | 433 | USA Nationwide Distribution in the states of: MI, SC, CO, OH, NE, WI, NC, SD, AZ, TX, MA, WA, PA, KY, MN, GA, OK, NY, IN, VA, UT, MT, ME,FL, CA, MD,MO, IA, TN, MD, WV, IL,NJ, CT, and AR. | Terminated | October 08, 2014 | 197 | 433 | Not_Computer | Not_Computer
Z-0237-2015 | 69448 | No | MCZ | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Surgical packs (includes trays, bags and kits) |  |  | October 08, 2014 | 2014 | Customed, Inc | Customed products are being recalled because of compromised sterility, not limited to potentially damaged packaging. Serious deficiencies in the manufacturing processes (QSR) and uncontrolled and inadequate storage conditions increase the risk of contamination and may result in patient infection. | Packaging | Customed initiated a recall on October 8th, 2014 of sterile convenience surgical Packs-trays, bags (known as Convenience Packs) and new lots of products added to the initial June 3, 2014 recall of 233 sterile surgical packs. The products are being recalled because of significant compromised sterility, not limited to potentially damaged packaging. Customers were instructed to quarantine and cease use of the affected product identified. Customers are to return the response form included with the letter, to the contact information provided. Upon receipt, the firm will contact customers with further instructions to obtain credit, replacement product, and for return of the recalled product or any other disposition action. If customers have any questions, they should contact the Customed Recall Coordinator at 407-850-5558 extensions 8001 or 8005, or via e-mail: pablo.diaz@customedhealing.com. |  | 185000 | 185000 | US Distribution to Florida, New York and Puerto Rico.. | Terminated | August 14, 2017 | 1041 | 1079228 | Not_Computer | Not_Computer
Z-2269-2014 | 68768 | No | KNH | Contraceptive tubal occlusion device (TOD) and introducer. | Obstetrics/Gynecology | Obstetrics/Gynecology | PMA | Gyrus ACMI Falope-Ring Dilator, item 000878-4. The Falope-Ring Dilator is a conical plastic device that is a component of the Falope-Ring Band Contraceptive Tubal Occlusion Device procedure pack (FRB-30). The Falope-Ring Band Contraceptive Tubal Occlusion Device (TOD) consists of a cylindrical radiopaque silicone band that occludes the lumen of the Fallopian tube when placed over a looped segment of the tube. The FRB-30 (used with the reusable applicators) consists of one reusable Falope-Ring Guide, an instruction manual and 30 procedure kits. Each kit is provided sterile and includes one dilator and two Falope Bands. |  |  | July 08, 2014 | 2014 | Gyrus Medical, Inc | Falope-Ring Dilator for Reusable Falope-Ring Band Applicators have been sold as a stand-alone non-sterile component. The dilator is intended to assist placement of the Falope-Ring band on the reusable Falope applicator. | Packaging process control | Olympus sent a "Medical Device Recall" letter dated July 7, 2014 to affected customers. The letter was addressed to Surgery Department &amp; Risk Management Department. The letter described the product being recalled, problem and actions to be take. Consignees were advised to immediately cease further use of any affected product and to remove and discard it. Customers were to fax the completed questionnaire to 484-896-7128. For questions contact Laura Storms at 484-896-5688 or at laura.storms@olympus.com |  | 507 | 507 units | Worldwide - US Nationwide Distribution in the states of AL, AR, CA, CO, CT, DE, FL, GA, IA, IL, IN, LA, MD, MI, ,MN, MS, NH, HY, OH,OR, PA, TN TX, UT, VA, WA, WI,and WV. and the country of CANADA. | Terminated | May 07, 2015 | 303 | 507 | Not_Computer | Not_Computer
Z-2688-2014 | 69137 | No | OMP | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | RENASYS EZ Negative Pressure Wound Therapy (NPWT) System (RENASYS EZ) for patients who would benefit from a suction device particulary as the device may promote wound healing. |  |  | August 13, 2014 | 2014 | Smith &amp; Nephew, Inc. | Modification were made to the RENASYS EZ NPWT Systems without a 510(k) or PreMarket Approval notification. | No Marketing Application | All RENASYS EZ Products will be removed from the field and customers will be transitioned to other product. On August 13, 2014 the firm sent an "Urgent- Medical Device Removal RENASYS EZ" notification via FedEx overnight asking that the use of the RENASYS EZ should be immediately discontinued and the device should be returned to Smith &amp; Nephew. If you have questions, please reach out to your Customer Care representative or call 1-800-876-1261 Monday through Friday, 8:00 AM to 6:00 PM, Eastern Time. |  | 2142 | 2,142 units | Worldwide Distribution- US (nationwide) including MA, MN, NY and Puerto Rico; Canada, Mexico, Dubai, Australia, Chile, Colombia, Germany, France, England, Spain, Italy, Norway, Hong Kong, Shanghai, New Zealand, South Africa, Singapore, Japan, Thailand and the Caribbean Islands. | Terminated | August 06, 2015 | 358 | 2142 | Not_Computer | Not_Computer
Z-1507-2014 | 67834 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Arrowgard Blue Plus Antimicrobial Multi-lumen Central Venous Catheter, catalogue number AK-45854-CDC. |  |  | July 17, 2013 | 2013 | Arrow International Inc | Arrow is recalling the Arrowgard Blue Plus Antimicrobial Multi-lumen Central Venous Catheter due to a labeling inconsistency. One portion ofthe lidstock states that the kit contains 8.5FR x 16CM catheter, while the rest of the lidstockcorrectly identifies the catheter as 8.5FR x 20CM | Labeling False and Misleading | Arrow International Inc, ("Arrow") sent a Urgent Medical Device Recall Notification dated July 17, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. 1. If you have affected stock, immediately discontinue use and quarantine any products with the catalog and lot number listed above. 2. To return product, complete the enclosed Recall Acknowledgement Form and fax it to : 1-866-804-9881. Attn: Customer Service. This will allow us to document the amount of product you have on hand for return Goods Authorization (RGA) number and will provide instructions for the return of product to Arrow International.3.If you have no affected stock please complete the enclosed Recall Acknowledgement Form and fax it to : 1-866-804-9881, Attn: Customer Service. This will allow us to document your receipt of this letter.For further questions please contact Arrow Customer Service at 1-800-343-2935. |  | 170 | 170 | US Distribution including the states of IA, OR, WV, MI, AZ, IL, PA, GA, DE, NC, LA, and VA. | Terminated | June 02, 2015 | 685 | 170 | Not_Computer | Not_Computer
Z-0847-2014 | 67207 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Tapered Screw-Vent Implant, MTX, 4.7mm x 11.5mm, 4.5mm Platform. Catalog Number TSVWB11, Lot 62470008. |  |  | January 06, 2014 | 2014 | Zimmer Dental Inc | Zimmer Dental is conducting a voluntary recall of a single lot of the Tapered Screw-Vent Implant, Catalog TSVWB11, Lot 62470008, because some of the packages of this lot may have the Cap Label state Âż3.7 x 10mm instead of Âż4.7 x 11.5mm. | Labeling Change Control | Zimmer Dental initiated the recall on 01/06/2013 by sending out recall notifications via FedEx to domestic customers and via electric mail to international customers.Customers are sent a letter identifying the issue, product description with codes, risk, instructions, and contact information. The letter's instructions include locating the affected product and return unused product to Zimmer Dental and notifying the amount of product that has been consumed. |  | 307 | 307 | Worldwide Distribution-USA (nationwide) and the countries of Thailand, Japan, Hong Kong, Morocco, and UAE. | Terminated | July 07, 2014 | 182 | 307 | Not_Computer | Not_Computer
Z-1603-2014 | 68017 | No | DWL | Medical support stocking. | General Hospital | General Hospital | 510(K) Exempt | DROP LOK" Knee Brace***LATEX FREE" |  |  | March 19, 2014 | 2014 | Zimmer, Inc. | During a transfer of products from a recently shutdown facility, the firm discovered raw material labeled as latex free actually contained latex. | Labeling mix-ups | Zimmer sent an Urgent: Device Removal Letters dated April 8, 2014 to their customers. The letter identified the affected product, probllem and actions to be taken. Customers were instructed to review this notification, iIdentify and quarantine the affected product to prevent further distribution or use, complete the attached Response Form and return it via email. orporateQuality_PostMarket@zimmer.com. Return affected product to: Zimmer SurgicalAttn: QA/RA Dept.- Recall200 West Ohio AvenueDover, Ohio 44622 USAPlease include a copy of the Response Form with the shipment.For returns outside the US, please email Rhonda.duncan@zimmer.com obtain an IRA (international return authorization) number. The IRA request should include the part number(s) being returned and the quantity. Please write the associated IRA number on the outside of the box.4. Zimmer will credit your account for returned Drop-Lok" Knee Braces, Cartilage Knee Braces, Hinged Knee Supports, or Neoprene Tennis Elbow Supports. Please return a copy of the completed response form along with your returned product to ensure proper credit.Important: Please distribute this notification to all personnel within your organization who need to be aware. If you have further transferred affected product(s), please provide the customer's information on the Business Response Form to Zimmer.For questions call 330-354-0989 |  | 131 | 131 units | Worldwide Distribution - US Nationwide in the states of: AK, AZ, CA, FL, GA, IA, IL, IN, LA, MD, MI, MO NC, NE, NV, NY, OH, OK, OR, PA, RI, SD, TX, UT, VA, VT, WA, WI &amp; WV and countries of: AUSTRALIA, CANADA, GERMANY, IRAQ, ITALY, SAUDI ARABIA, TUNISIA &amp; UNITES ARAB EMIRATES. | Terminated | April 14, 2015 | 391 | 3548 | Not_Computer | Not_Computer
Z-2073-2014 | 68674 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | LP,Surgical Fibers, Bare Laser Fiber 600 micron, DBLF-60-2, Sterile EO. |  |  | June 20, 2014 | 2014 | Laser Peripherals LLC | Laser Peripherals LLC is recalling DBLF-60-2 Bare Laser due to a mislabeling that indicated a sterility expiration date as '2016-16' as opposed to the intended date of '2016-06'. | Labeling False and Misleading | Laser Peripherals contacted all consignees via telephone on June 20, 2014, in regard to the product being recalled by following the firm's Standard Operating Procedures. The SOP's included Instructions to immediately stop the use of the recalled product and provided other communications related to the recall.For questions regarding this recall call 612-280-1318. |  | 25 | 25 | Nationwide Distribution to IL, IN, and MD | Terminated | October 22, 2014 | 124 | 25 | Not_Computer | Not_Computer
Z-1166-2014 | 67308 | No | JOJ | Sterilization process indicator. | General Hospital | General Hospital | 510(k) | VERIFYÂż SixCess 270FP Challenge Pack, STERIS Corporation, 5960 Heisley Road, Mentor, OH 44060, 440-354-2600. |  |  | December 19, 2013 | 2013 | Steris Corporation | STERIS has identified that the outer shipping package for one lot (Lot #024552A) of VERIFYÂż SixCess Challenge Packs incorrectly identifies the lot as Lot #024522A. | Labeling False and Misleading | The firm, Steris, sent an "URGENT VOLUNTARY RECALL NOTICE" dated December 20, 2013 to their customers. The letter describes the product, problem and actions to be taken. The customers were instructed to verify if the outer packaging of your on-hand inventory contains lot number 024522A; destroy any remaining inventory of the affected product in your possession and contact STERIS Customer Service at 1-800-548-4873 to obtain no cost replacement product and if you do not have the affected product in your inventory, no further action is required.If you have any questions regarding this matter, please contact your STERIS Account Manager or STERIS Customer Service at 1-800-548-4873. |  | 1880 | 94 boxes/20 units | Worldwide Distribution: US (nationwide) including states of: IL, LA, ME, MI, NC, OH, OK, VA, WI &amp; WV; and to the country of: Canada.Product was also shipped to Canada. | Terminated | March 26, 2014 | 97 | 1880 | Not_Computer | Not_Computer
Z-1793-2014 | 68229 | No | OIE | Specimen transport and storage container. | Pathology | Clinical Chemistry | Enforcement Discretion | APTIMA Urine Collection Kit; Catalog Number |  |  | April 17, 2014 | 2014 | Gen-Probe Inc | Incorrect labeling was used on each pouch of the affected lots of APTIMA Urine Specimen Collection Kits. The outer box of the kits are identified correctly as "APTIMA Urine Collection Kits"; however, the inner pouches are incorrectly labeled as "APTIMA Unisex Swab Collection Kit for Endocervical and Male Urethral Swab Specimens." | Labeling mix-ups | Hologic sent a Recall Notice letter dated April 14, 2014 via overnight mail (FedEx) to affected customers. The letter identified the affected product, problem and actions to be taken. The notification letter requests customers to: 1. segregate the entire contents of the affected kit lots2. reconcile and destroy all partial and/or full kits remaining in inventory and to document; the amount destroyed on the Customer Response Form provided with the Recall Notice; 3. if the urine collection kits were further distributed to other collection sites, then to notify them of this issue; and 4. complete and return the Customer Response Form to Technical Support (contact information is provided in the notification letter).For additional information, questions, or concerns, contact Hologic Molecular (Technical) Support.Toll Free: 888-484-4747Direct: 858-410-8511Fax: 858-410-8250Email: molecularsupport@hologic.com |  | 4740 | 4,740 units | Worldwide Distribution - US Nationwide and the countries of Canada (CA), Denmark (DK), France (FR), Germany (DE), and the United States (US), Belgium (BE), United Kingdom (GB). | Terminated | August 26, 2014 | 131 | 4740 | Not_Computer | Not_Computer
Z-0145-2015 | 69224 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | Parker Eclipse Probe Cover, Latex Free, REF 38-01, 10 Eclipse Probe Covers, Size: 2.37/1.75"W x 9.5"L (60/44 mm x 241 mm), Lot #D424, Expiry date 04/2016, Pre-gelled inside with Aquasonic 100 Ultrasound Transmission Gel |  |  | August 06, 2014 | 2014 | Parker Laboratories, Inc. | Parker Laboratories, Inc. received a report of an incorrect expiration date on the inner boxes and on the plastic bags of the Eclipse Probe Covers. | Incorrect or no expiration date | Parker Laboratories Inc. sent an "Urgent Medical Device Product Recall" Letter/Response Form via Fedex to their customers. The letter indentified the affected product, problem and actions to be taken. For questions contact Parker Laboratories Quality Assurance at 973-276-9500, ext. 7142. |  | 367 | 367 boxes | Worldwide Distribution: US (Nationwide) and country of: Canada. | Terminated | April 29, 2015 | 266 | 367 | Not_Computer | Not_Computer
Z-0212-2015 | 69588 | No | JJY | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | IMMULITE Calcitonin Control Module (IMMULITE, IMMULITE 1000, IMMULITE 2000, IMMULITE 2000 XPi) |  |  | October 20, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc. | A typographical error regarding the expiration date was identified in the control module instructions for use. | Error in labeling | Siemens Healthcare Diagnostics issued an Urgent Medical Device Correction letter October 17, 2014, to US customers, and an Urgent Field Safety Notice issued October 16, 2014, outside the US. The letter identified the product the problem and the action needed to be taken by the customer.Customers are advised to review the notice with their medical directors. Customers should discard the control module IFU listed in table 1 and refer to table 2 for the list of revised IFU versions as noted on the notice. Finally, the customers should check the system for the expiration date and if correct, no further action is required. Field Correction Effectiveness Check should be returned via fax (312) 275-7795 within 30 days. For further questions please call 914-631-8000. |  | 2212 | 2212 units (92 units domestically &amp; 2120 units internationally) | Worldwide Distribution - US Distribution including the states of FL, KY, MN, MA, PA, UT, NY, VA and CA., and the countries of Algeria, Australia, Argentina, Austria, Belgium, Brazil, Bulgaria, Cameron, Chile, China, Colombia, Costa Rica, Cyprus, Czech Republic, Denmark, Estonia, Finland, France, Germany, Great Britain, Greece, Hong Kong, Hungary, India, Iraq, Ireland, Israel, Italy, Jordan, Kazakhstan, Latvia, Macedonia, Mexico, Netherlands,Pakistan, Peru, Poland, Portugal, Rep. Korea, Romania, Russian Fed, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Tadjikistan, Taiwan,Turkey, Unit. Arab Emir and Vietnam. | Terminated | March 31, 2017 | 893 | 2212 | Not_Computer | Not_Computer
Z-1174-2014 | 67391 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes Small Notch Titanium Reconstructive Plate |  |  | January 27, 2014 | 2014 | Synthes, Inc. | Certain modules containing the Synthes Small Notch Titanium Reconstructive Plate set was distributed by Spine sales consultants although it was contraindicated for certain spine applications. | Labeling design | Synthes sent an Urgent Medical Medical Device Recall letter, dated January 27, 2013, to Users and Sales Consultants. The letter identified the affected product, reason for recall, potential hazard, and actions to be taken. Customers were instructed to complete the verification section. For questions call 610-719-5450 or contact your Synthes Sales Consultant. . |  | 18747 | 18747 | USA Nationwide Distribution | Open, Classified |  |  | 18747 | Not_Computer | Not_Computer
Z-0772-2014 | 67098 | No | FGY | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | BD Interlink threaded lock cannula |  |  | November 26, 2013 | 2013 | Becton Dickinson &amp; Company | BD Interlink Threaded Lock Cannula REF# 303369 and lot #3021392 is being recalled due to the potential that a small percentage of the lot may exhibit open seals which adversely impacts unit package integrity and product sterility. | Packaging | BD Diagnostics sent an Urgent: Voluntary Product Recall letters/Notice of Return dated November 26, 2013, to all affected customers. Customers were asked to identify and segregate stock of the affected number, inform the appropriate personnel in their organization to discontinue use of the affected product, return the afected product, following the instructions on the enclosed packing slip, BD would issue a credit for product returned. If product was further distributed customers were instructed to identify their customers and notify them at once of the product recall using the provided customer communication. Customers were asked to complete the attached notice and fax to 855-731-0157. Customers with questions were instructed to call 1-855-731-0157.For questions regarding this recall call 755-731-0157. |  | 268500 | 268,500 products | Worldwide Distribution - USA (nationwide) and Internationally to Canada and Japan. | Open, Classified |  |  | 268500 | Not_Computer | Not_Computer
Z-0708-2014 | 67124 | No | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | Multiples Anspach Micro Curved Attachments (MCA) and Curved Burr Support S leves (Cutters and Attachments) |  |  | October 26, 2012 | 2012 | The Anspach Effort, Inc. | During a recent review of documentation Anspach determined that the current directions for use for the MCA cutting burrs and MCA Micro Curved Attachment do not provide clear details regarding the intended use specifically for the MCA burrs. The MCA burrs are intended for use in delicate bone in otologic procedures such as cochleostomy. The use of the device in dense bone removal may result in cu | Labeling False and Misleading | Anspach sent an Urgent Medical Removal letter dated October 26, 2012, to all affected customers. The letter stated that they were initiating a Medical Device removal regarding the MCA Cutting Burrs. During a recent review of documentation it was determined that the current directions for use for the MCA cutting burrs (item 18-0079) and MCA Micro Curved Attachment (item 18-0060) did not provide clear details regarding the intended use specifically for the MCA burrs. The MCA burrs are intended for use in delicate bone in otologic procedures such as cochleostomy. The use of the device in dense bone removal may result in cutter/ burr fracture. If a fracture occurs, contact Customer Support at (800) 327-6887 or email ProductSupportTeam@synthes.com. If the procedure was done successfully, no additional patient follow up is necessary.Customer Immediate Actions:1. Please remove and return any MCA items from your inventory immediately.2. Complete the attached reply form indicating your receipt of this letter. Return the completed form by fax or email to the number or email address provided on the form. If you distribute any of the products to other services or facilities, please forward this information as appropriate. Customers with questions were instructed to contact Anspach Customer Support at (800) 327-6887 or email ProductSupportTeam@synthes.com.For questions regarding this recall call 1-800-327-6887. |  | 14330400 | 6,824 imits (4,7324 in USA / 2,100 Foreign) | Worldwide Distribution - USA (nationwide) and Internationally to Canada, Netherlands, Great Britain, France, Sweeden, Italy, Spain, Japan, South Africa, Australia, Switzerland, | Terminated | June 23, 2014 | 605 | 14330400 | Not_Computer | Not_Computer
Z-0147-2015 | 69446 | No | DRC | Trocar. | Cardiovascular | Cardiovascular | 510(K) Exempt | Blunt Hurd Cannula Blunt Tip 10 mm with 5 mm Reducer simplifies open laparoscopy. |  |  | September 13, 2013 | 2013 | CooperSurgical, Inc. D.B.A. Lone Star Medical Products | CooperSurgical is recalling two lots of 900-863 because the wrong package insert (IFU) was placed into the shipper. | Error in labeling | CooperSurgical sent an Urgent Field Safety Notice ( Correction ) dated September 27, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The recalling firm provided correct package inserts (IFU) to affected consignees. The recalling firm intends to conduct effectiveness checks by reviewing return receipts from the initial recall communication.We apologize for any inconvenience this communication may cause. If you have any questions or concerns, please contact CooperSurgical at (203) 601-9818. |  | 90 | 90 units | Worldwide Distribution: US (nationwide) to North Carolina; and Internationally to Australia and Italy. | Terminated | September 23, 2015 | 740 | 90 | Not_Computer | Not_Computer
Z-0131-2015 | 67987 | No | CHI | Lipase test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Lipase Reagent, 7D80. |  |  | April 07, 2014 | 2014 | Abbott Laboratories, Inc | Package inserts have incorrect SmartWash parameters for Triglyceride. | Labeling Change Control | Firm issued a Product Correction Letter to all Lipase reagent customers who received lots 43827UQ12, 43819UQ03, 44449UQ05, 44127UQ06, 44334UQ07, 43994UQ07, 44596UQ08, 44775UQ09, and 44930UQ11. The letter identified the affected product and provided an explanation of the problem, as well as, the patient impact and actions to be taken. The Product Correction Letter provided guidance for the laboratory to correctly configure the Architect cSystem's Lipase assay parameters. Customers are to retain a copy of the letter for their lab records, and forward a copy to any laboratories that the product was further distributed to. Questions should be directed to Customer Service at 1-877-4ABBOTT or your local customer service for customers outside of the US. |  | 11916 | 11916 units | Worldwide Distribution: US (nationwide) and countries of: Algeria, Angola, Argentina, Armenia, Australia, Austria, Bangladesh, Belgium, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, Chile, China, Columbia, Croatia, Czech Republic, Dominican Republic, Egypt, Estonia, France, Germany, Greece, Guatemala, Hungary, India, Indonesia, Iraq, Ireland, Israel, Italy, Jamaica, Jordan, Kenya, Kuwait, Kyrgyzstan, Latvia, Lebanon, Libya, Lithuania, Luxembourg, Moldova, Montenegro, Morocco, Netherlands, New Zealand, Norway, Pakistan, Panama, Paraguay, Philippines, Poland, Portugal, Qatar, Republic of Yemen, Romania, Saudi Arabia, Senegal, Serbia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Tunisia, Turks and Caicos, Turkey, Uganda, United Arab Emirates, United Kingdom, and Uruguay. | Terminated | April 01, 2015 | 359 | 11916 | Not_Computer | Not_Computer
Z-1161-2014 | 67477 | No | OMP | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Smith &amp; Nephew RENASYS-G Sterile, Gauze Dressing Kit with Soft Port REF #'s: 66800961 (The RENASYS Soft Port Dressing Kits are used in conjunction with negative pressure wound therapy pumps) |  |  | January 09, 2014 | 2014 | Smith &amp; Nephew Inc. | Reports of wound fluid and /or blood that were not evacuated from beneath the wound dressing and the pump blockage alarm did not activate has been associated with inadequate Soft Port sealing at the wound site or improper positioning of the Soft Port orifice over the wound contact dressing drape. Inadequate sealing can result in air leaks and contribute to inefficient fluid evacuation from the wo | Packaging change control | A product bulletin titled "Soft Port Enhancements" dated March 14, 2013, was released for external distribution describing the changes to the Soft Port products, e.g. increasing the Soft Port orifices size and adding a second slit to the release paper covering the port adhesive, and stating the purpose of the modifications to the device. These product bulletin was provided to sales representatives for distribution to customers requesting information regarding the design enhancement made to the Soft Port product range. The bulletin was not distributed to all the customers.To ensure that the product modification and its purpose was appropriately communicated to consignees of affected product a Medical Device Correction Notice, dated 01/22/2014, retrospectively advising customers of the modification to the Soft Port products has been prepared and was sent to consignees of the affected products on the same date. |  | 7353 | 7,353 individuals units | Worldwide Distribution-USA (nationwide) including Puerto Rico and the states of AK, AL, AR, AZ, CA, CO, CT, FL, GA, ID, IL, KY, LA, MA, MS, MT, NC, NJ, NV, NY, OH, OR, PA, TN, TX, VA, WA, WI, WV, HI and the countries Dubai, Curacao, Argentina, Australia, Chile, Germany, France, England, Hong Kong, Kuwait, Lebanon, Libya, Sri Lanka, Malaysia, New Zealand, Manila, Saudi Arabia, South Africa, Singapore, Turkey, Taipei and Vietnam. | Terminated | September 26, 2014 | 260 | 2423739 | Not_Computer | Not_Computer
Z-1833-2014 | 68072 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes External Fixation Systems (Small, Medium, Large and DO), bone fixation appliance components. |  |  | April 10, 2014 | 2014 | Synthes, Inc. | Labeling changes have been made related to MR (Magnetic Resonance imaging) conditions as a result of changes in required protocols to designate a product MR Safe, MR Conditional, or MR Unsafe. | Labeling Change Control | Synthes sent an Urgent Notice dated April 10, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to:All Synthes External Fixation devices should be treated as MR Conditional."Synthes asks that you review the information contained in this Labeling Notification and complete the Verification Section located on Page 3."Discard outdated revisions of the Technique Guides noted in the table below."Update your records with updated Labeling Infonnation."Forward this Medical Device Labeling Update to anyone in your facility that needs to be infonned, especially those personnel that conduct MR testing."If the Verification Fonn is answered on behalf of more than one facility and/or individual, please clearly indicate the name and address of the facility and/or individual in pg. 3."Updated product literature can be located on the Synthes website athttp://www.syntheskyo.com/trauma kyo/home/home.htm."Please see the attached inseii for current conditions for use in the MR environment."Maintain a copy of this notice.Customers with questions were instructed to call 610-719-5450.For questions regarding this recall call 610-719-5000. |  | 5114052 | 5,114,052 | Worldwide Distribution - USA (nationwide) and Canada. | Open, Classified |  |  | 5114052 | Not_Computer | Not_Computer
Z-1164-2014 | 67658 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific, RunWay Guide Catheter, UPN H749389696640, Catalog Number 38969-664. |  |  | March 03, 2014 | 2014 | Boston Scientific Corporation | The sterile pouch of the devices from this lot may not be completely sealed and sterility may be compromised. The most common adverse health consequence expected from an unsealed pouch is a transient bacteremia, which may require antibiotic therapy. | Packaging | Consignees were sent a Boston Scientific "Urgent Medical Device Recall Removal " letter dated March 3, 2014. The letter described the problem and the product involved in the recall. The letter advised consignees to return the affected product and to complete and return the Account Reply Verification Tracking Form. The local Sales Representative can answer any questions that you may have regarding this Recall Removal. |  | 7 | 7 | TX and Russia | Terminated | June 09, 2014 | 98 | 7 | Not_Computer | Not_Computer
Z-1843-2014 | 68352 | No | PAM | Multiplex nucleic acid assay for identification of microorganisms and resistance markers from positive blood cultures. | Microbiology | Microbiology | 510(k) | FirmArray Blood Culture Identification (BCID) Panel, Model 2.0 |  |  | May 27, 2014 | 2014 | BioFire Diagnostics, Inc. | BioFire has identified an increased risk of false positive results when the FilmArray Blood Culture Identification (BCID) Panel is used with bioMÂżrieux BacT/ALERT Standard Anaerobic (SN) blood culture bottles. False positive results have been observed for Pseudomonas aeruginosa and Enterococcus. | Labeling design | BioFire Diagnostics sent an Urgent Recall/Correction letter on May 28, 2014. Customers were instructed If the BCID Panel is used to test BacT/ALERT SN bottles, positive results for Pseudomonas aeruginosa and Enterococcus should be confirmed by another method prior to reporting the test results.Customer with questions were instructed contact customer support at Support@biofiredx.com or via telephone by dialing 1-800-735-6544 and selecting option 5 for customer support and then option 1 for FilmArray.For questions regarding this recall call 801-736-6354, ext 360. |  | 1084 | 1084 | Worldwide Distribution - USA (nationwide) European Union, Hong Kong, and VA. | Terminated | November 07, 2014 | 164 | 1084 | Not_Computer | Not_Computer
Z-1152-2014 | 67418 | No | NOT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(k) | HER2 Image Analysis |  |  | February 07, 2014 | 2014 | Leica Biosystems Imaging, Inc. | In an abundance of caution, Aperio (now known as Lecia Biosystems) is providing additional information to clarify that the use of Digital IHC Breast/ER/Dako Image Analysis, Digital IHC Breast/HER2 Digital Read, Digital IHC Breast/HER2/Dako Image Analysis, Digital IHC Breast/PR Digital Read and Digital IHC Breast/PR/Dako Image Analysis with the RUO products identified in that earlier notification i | Labeling False and Misleading | Aperio (now known as Leica Biosystems) sent an recall notification letter on February 7, 2014, via USPS mail. The letter dated January 31, 2014, informed customers of the reason for the notification, product description, provided additional information to clarify that the use of Digital IHC Breast/ER/Dako Image Analysis, Digital IHC Breast/HER2 Digital Read, Digital IHC Breast/HER2/Dako Image Analysis, Digital IHC Breast/PR Digital Read and Digital IHC Breast/PR/Dako Image Analysis with the RUO products identified in that earlier notification is also for Research Use Only and not for use in diagnostic procedures. The notice also included updated manuals for the image analysis algorthithms so that the intended use in the hardware and software manuals is consistent. An acknowledgement form and contact information was also provided.We sincerely apologize for any inconvenience this may have caused. If you have any questions, please do not hesitate to contact Aperio Technical Support at (866) 478-3999. |  | 78 | 78 | US Distribution including the states of FL, WA, PA, MS, TX, MN CA, NY, NJ, NC, CO and OH. | Terminated | April 15, 2014 | 67 | 78 | Not_Computer | Computer
Z-0790-2014 | 66933 | No | LLH | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | BD GeneOhm Cdiff Assay, Part #441401, containing Lysis Kit sample buffer tubes, 200 tests/kit; labeled as ***GeneOhm Sciences Canada, Inc. 2555 boul. du Parc-Technologique Quebec, Qc, Canada, G1P 4S5*** |  |  | October 11, 2013 | 2013 | Becton Dickinson &amp; Co. | An incorrect reagent was packed with an in vitro diagnostic kit, which may cause an increase in unresolved results or false negative results for Clostridium difficile in patient samples. | Error in labeling | Becton Dickinson sent an Urgent Product Recall Letter dated October 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.They were advised that a portion of the BD GeneOhm Cdiff assay test kits from one lot, 08T13108, may contain incorrect bags of sample buffer tube. Customers were requested to return response forms indicating acknowledgement of the recall and quantity of remaining product that was discarded. Please fax to 410-316-4258. |  | 33 | 33 kits | US Distribution including the states of NY, IN, UT, MN, OH, MA, SD, CA, NJ, and PA. | Terminated | February 19, 2014 | 131 | 33 | Not_Computer | Not_Computer
Z-1851-2014 | 68387 | Yes | BTR | Tracheal tube. | Anesthesiology | Anesthesiology | 510(k) | RUSCH, LaserTube (Rubber), Laser Resistant Tracheal Tube, cuffed, Teleflex Medical, Willy-Rusch, GmbH. |  |  | May 23, 2014 | 2014 | Teleflex Medical | Although the Natural Rubber Latex caution symbol is clearlydisplayed on the product label, the accompanying caution statement "Caution: This product contains Natural Rubber Latex which may cause Allergic Reactions" was not printed on the product label. | Error in labeling | Teleflex Medical sent a Urgent Medical Device Recall Notification dated June 6, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Our records indicate that you have received products that are subject to this recall. We are now notifying our customers to take the following actions:1. If you have affected stock, immediately discontinue use and quarantine any products with the catalog numbers listed above from the lot numbers listed on the attachment.2. To return product, complete the enclosed Recall Acknowledgement Form and fax it to 1-866- 804-9881, Attn: Customer Service. This will allow us to document the amount of product you have on hand for return. A customer service representative will contact you with a Return Goods Authorization (RGA) Number and will provide instructions for the return of product to Teleflex Medical.3. If you have no affected stock, please complete the enclosed Recall Acknowledgment Form and fax it to 1-866-804-9881, Attn: Customer Service. This will allow us to document your receipt of this letter.Teleflex Medical is committed to providing high quality, safe and effective products. We sincerely apologize for any inconvenience this action may cause your operations. If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-866-246-6990. |  | 1658 | 1658 ea | Nationwide Distribution and the country of Colombia. | Terminated | January 13, 2016 | 600 | 1658 | Not_Computer | Not_Computer
Z-0138-2015 | 69279 | No | NAD | Occlusive wound dressing. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Biatain Super Non-Adhesive, Catalog number 46300, 10 pcs, REF 0463001400. Country of Origin:HU. |  |  | September 08, 2014 | 2014 | Coloplast Manufacturing US, LLC | Coloplast is conducting a recall for one lot (4201931) of the Biatain Super Non-Adhesive wound dressing which was mislabeled. | Labeling mix-ups | Coloplast sent an "Urgent Voluntary Recall Notification" dated September 9, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. If they have further distributed the product they need to contact their customers immediately. For questions they can contact Coloplast Customer Care Representative at 800-533-0464.A second follow up "Urgent Voluntary Recall Notification" letter dated September 23, 2014 was also sent to consignees. |  | 6 | 6 | US Distribution to MO and RI. | Terminated | January 21, 2015 | 135 | 6 | Not_Computer | Not_Computer
Z-1273-2014 | 67638 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Spacelabs Healthcare G2 Clinical Access Mislabeled CDs for PN: 063-1829-11 Rev A and Update Kit, PN: 040-1565-10. |  |  | February 14, 2014 | 2014 | Spacelabs Healthcare, Llc | Spacelabs has voluntarily recalled G2 Clinical Access Software CD because the programs installed on the Clinical Access CD are incorrect and it will prevent the customer frominstalling Clinical Access. | Labeling Change Control | Spacelabs Healthcare began sending out the Medical Device Correction letter, dated 28 February 2014, to domestic customers on 28 February 2014 and to international customers by 7 March 2014. The letter identified the product, the product, the problem, and the action to be taken by the customer. Customers are also directed to call 1-800-522-7025 and select 2 for Technical Support.Customers are advised to dispose the CD, PN: 063-1829-11 Rev A previously supplied and replace it with the replacement CD (Rev B) provided. |  | 36 | 36 units total: 17 units were distributed in the US and 19 units distributed outside the US. | Worldwide Distribution- USA (nationwide) including the states of Alabama, Florida, Georgia, Louisiana, Missouri, New York, South Carolina, Tennessee, Texas and Wyoming, and the countries of Australia, Canada, Czechoslovakia, China, Colombia, France, Great Britain, Italy, Netherlands, Poland, Saudi Arabia, and Turkey. | Terminated | May 20, 2014 | 95 | 36 | Not_Computer | Computer
Z-0132-2015 | 69101 | No | MWS | Cardiovascular surgical instruments. | Cardiovascular | Cardiovascular | 510(K) Exempt | Acrobat V Vacuum Off-Pump System |  |  | August 13, 2014 | 2014 | Maquet Cardiovascular, LLC | An internal finding identified the presence of pinholes/cuts in some of the Tyvek pouches of the vacuum tubing sets included as an accessory to Acrobat V Vacuum Off-Pump System, Acrobat SUV Vacuum Off-Pump System, Axius XPOSE 3 Access Device and XPOSE 4 Access Device. | Package design/selection | Maquet Inc. issued Recall letters/Acknowledgement forms dated via Fed-Ex on 8/13/2014. |  | 16532 | 16,532 units (19,270 units international) | Worldwide Distribution: US (nationwide) and countries of: ARGENTINA, AUSTRIA, BANGLADESH, BELARUS, BRAZIL, BRUNEI, CANADA, CHILE, CHINA, COLOMBIA, CROATIA, CZECH REPUBLIC, DENMARK, ECUADOR, GERMANY, FINLAND, FRANCE, GREECE, HONG KONG, INDICA, INDONESIA, ISRAEL, ITALY, JAPAN, MEXICO, MALAYSIA, NETHERLANDS, PANAMA, PERU, PHILIPPINES, POLAND, PORTUGAL, RUSSIA, SAUDI ARABIA, SINGAPORE, SLOVAK REPUBLIC, SOUTH KOREA, SPAIN, SWITZERLAND, TAIWAN, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM AND VIETNAM. | Terminated | March 30, 2016 | 595 | 16532 | Not_Computer | Not_Computer
Z-0124-2015 | 69417 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Spirotome" Soft-Tissue Biopsy Needle Set. |  |  | September 26, 2014 | 2014 | Cook Medical Incorporated | MEDINVENTS, the manufacturer of the Spirotome&quot; Soft-Tissue Biopsy Needle Set distributed by Cook Medical, has initiated a voluntary recall of these products. The recall hasbeen initiated due to customer complaints that the protective caps were dislodged in the pouch. Furtherinvestigation disclosed potential dislodging of protective caps on the needle tips on other distributedproducts. | Package design/selection | On 9/26/2014, URGENT: MEDICAL DEVICE RECALL notifications were sent to the affected distributors and institutions Risk Managers/Recall Administration with instructions for returning the affected product. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. For questions, call Customer Relations at 1-800-457-4500 or 1-812-339-2235. |  | 1239 | 1239 units | Worldwide Distribution-USA (nationwide) including the states of Arizona, California, Colorado, Hawaii, Iowa, Illinois, Massachusetts, Maryland, Michigan, Missouri, New Mexico, Pennsylvania, Texas, Virginia, Vermont, and Wisconsin, and the countries of AT; BE; CH; CY; CZ; DE; DK; ES; FI; FR; GB; HU; IC; IE; IS; IT; LB; LT; LU; NL; NO; PT; QA; SE; TR; and ZA. | Terminated | February 02, 2016 | 494 | 1239 | Not_Computer | Not_Computer
Z-1868-2014 | 68330 | No | MBP | Resorbable calcium salt bone void filler device. | Orthopedic | Physical Medicine | 510(k) | Accell Evo3C Demineralized Bone Matrix, 5 cc, Model Number: 02-6000-050 , Lot Number: 132160. |  |  | May 19, 2014 | 2014 | Integra LifeSciences Corp. | This lot of Accell Evo3C Demineralized Bone Matrix, 5cc was incorrectly packaged in boxes that describe the product as Accell Evo3 Demineralized Bone Matrix. Evo3C is contraindicated for treatment of vertebral compression and not indicated for use in the posterolateral spine. | Packaging | Integra sent an Urgent Voluntary Medical Device Recall dated May 19, 2014 to customers via mail (traceable courier service). The recall notification letter identified affected product, reason for recall and actions to be taken Customers were instructed to complete the "RECALL ACKNOWLEDGEMENT AND RETURN FORM". and return to:FCA1@integralife.com or FAX to 1-609-275-9445. For questions contact Customer Service at 1-800-550-7155. |  | 61 | 61 | US Nationwide Distribution | Terminated | September 05, 2014 | 109 | 61 | Not_Computer | Not_Computer
Z-0117-2015 | 69379 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Straumann Bone Level Tapered Implant 04.1 mm RC, SLActiveÂż 8mm, RoxolidÂż, Loxim |  |  | September 19, 2014 | 2014 | Straumann USA, LLC | Assembled with Narrow Connect (NC) transfer piece instead of Regular Connect (RC) | Packaging process control | Strauman notified consignees of this Field Corrective Action by letter dated September 19, 2014 and delivered via hand by Straumann Representative. Accounts are requested to return inventory. Contact (978) 747-2514 for assistance. |  | 6 | 6 units | US Distribution in states of: FL and MA. | Terminated | March 13, 2015 | 175 | 6 | Not_Computer | Not_Computer
Z-1928-2014 | 68469 | No | NBC | B-type natriuretic peptide test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Alere Triage BNP Calibrators for the Beckman Coulter Access Family of Immunoassay Systems, Model No. 98202 |  |  | May 30, 2014 | 2014 | Alere San Diego, Inc. | Alere San Diego is recalling Alere Triage BNP Calibrators for the Beckman Coulter Access Family of Immunoassay Systems because the bottle labeled as Calibrator Level 5 actually contains Calibrator Level 4. This results in a calibration failure when attempting to calibrate using a calibrator kit which contains the affected Level 5 Calibrator. | Error in labeling | Alere sent an Urgent Medical Device Recall letter dated May 30, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The primary consignees were instructed to block all stock of the affected lot in their warehouses; discontinue use and return all stock on hand, complete the provided verification form to document receipt of the notice and quantity to be returned, contact their primary consignees and advise them of the requireddiscontinuation and return of the affected lot. For customers (or their sub-distributors) who have inventory on hand, Alere San Diego, Inc. will provide replacement product to the customer for the quantity of units that they must return.Customers are instructed to complete the attached reply form with the amount of unused product that will be returned for replacement. Customers with questions are instructed to contact Alere San Diego at 9975 Summers Ridge Rd, San Diego, CA 92121, (877) 308-8287, triage.support@alere.com |  | 3121 | 3,121 kits | Worldwide Distribution - Nationwide and the countries of Australia, Canada, China, EU, Mexico, Russia, Singapore, South Africa, South Korea, Taiwan and United Arab Emirates. | Terminated | February 11, 2015 | 257 | 3121 | Not_Computer | Not_Computer
Z-1930-2014 | 67011 | No | JJY | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | MAS Omni Immune PRO Controls |  |  | December 02, 2013 | 2013 | Microgenics Corporation | The constituent, Inhibin A, was removed from MAS Omni IMMUNE and MAS Omni IMMUNE PRO controls; however the package insert was not updated to reflect the change. | Labeling Change Control | Urgent Medical Device Correction letters were sent on December 2, 2013 to all affected customers and distributors via trackable mail. |  | 1393 | 1393 | Worldwide Distribution-USA (nationwide) and the countries of Australia, China, Germany, Hong Kong and Malaysia. | Terminated | July 08, 2014 | 218 | 4578 | Not_Computer | Not_Computer
Z-1149-2014 | 67562 | No | NZW | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Heparin I.V. Flush Syringe is available is pre-filled polypropylene syringes, 100 units/mL, 3 mL fill in a 12 mL syringe. |  |  | February 19, 2014 | 2014 | Medefil Incorporated | The bar code that is printed on the Master Carton label does not correspond to the product code fully displayed on the Master Carton label. The non-matching information can only be detected by using a bar code reader. The defective bar codes will be read as Product No. MIH-3335 ( Heparin I.V. Flush Syringe 100 units/mL; 5 mL fill in 12 mL Syringe) whereas this product is actually Product No. MIH-3 | Labeling Change Control | URGENT MEDICAL DEVICE CORRECTION letters dated February 19, 2014 were sent to all direct accounts. The letters included instructions to: quarantine the recalled products; and,complete and return the attached customer response form (via Fax to 630-681-9100 or via mail to Medefil Inc., 250 Windy Point Drive, Glendale Heights, IL 60139) so that the correct number of replacement labels can be sent from Medefil. |  | 1608 | 1,608 Master Cartons | USA (nationwide) including the states of AZ, CA, CO, FL, GA, IL, LA, MA, MD, MI, MN, MO, MS, NC, NJ, NY, OH, OK, PA, SC, TX, UT, and WI, and the country of Canada. | Open, Classified |  |  | 1608 | Not_Computer | Not_Computer
Z-1942-2014 | 68482 | No | HSX | Knee joint femorotibial metal/polymer non-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stride Femoral Component Size 2, PFSI-00053, Right Lateral /Left Medial. Unicondylar Knee prosthesis. |  |  | June 04, 2014 | 2014 | Blue Belt Technologies MN | Mislabeled. Packages labeled PFSI-00053, Right Lateral/Left Medial contains PFSI-00054, Right Medial/Left lateral implants. Packages labeled PFSI-00054, Stride Right Medial/Left Lateral contain PFSI-00053, Left Medial/Right Lateral. | Labeling mix-ups | Firm initiated recall on 6/4/2014 by Blue Belt Technologies representatives hand delivering an Urgent Medeical Device Recall letter. The letter was addressed to the attention of Safety Officer/Regulatory Compliance Officer. The letter described the issue of the labeling mix-up, asked for devices to be quarantined, and stated that a Blue Belt Technologies Clinical Specialist will be contacting the consignee or return and replacement. A Receipt Acknowledgement form was asked to be returned by fax or mail. |  | 2 | 2 | US Nationwide Distribution in the states of CA, PA, and TX | Terminated | August 04, 2014 | 61 | 2 | Not_Computer | Not_Computer
Z-1633-2014 | 67833 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Arrowgard Blue and Arrowgard Blue Plus Pressure Injectable Central Venous Catheter, a multiple-lumen catheter that permits venous access to central circulation. Catalog number ASK-45703-PHF. |  |  | July 17, 2013 | 2013 | Arrow International Inc | The labeling states that the kit contains Introducer Needle: Echogenic 18 Ga. x 2.5 (6.35cm) XTW and 5mL ArroW Raulerson Spring-Wire Introduction Syringe; however, the kit contains, and should state, Introducer Needle: 18 Ga. x 2.5 (6.35cm) XTW and 5mL ArroW Raulerson Spring-Wire Introduction Syringe. As a result of this labeling inconsistency, if the clinician were to place the needle under x-r | Labeling False and Misleading | Arrow sent an Urgent Medical Device Recall Notification letter dated July 17, 2013, was sent to the customer requesting they discontinue use and quarantine any remaining affected product for return to Arrow.Our records indicate that you have received products that are subject to this voluntary recall. We are now asking that you take the following actions:1. If you have affected stock, immediately discontinue use and quarantine any products with the product and lot numbers listed above.2. To return product, complete the enclosed Recall Acknowledgement Form and fax it to 1-800-343-2935,Attn: Customer Service. This will allow us to document the amount of product you have on hand for return. Your sales representative will contact you with a Return Goods Authorization (RGA) Number and will provide instructions and assistance for the return of product to Arrow International.3. If you have no affected stock, please complete the enclosed Recall Acknowledgment Form and fax it to 1-800-343-2935, Attn: Customer Service. This will allow us to document your receipt of this letter. |  | 4285 | 4285 | US Distribution only to MI. | Terminated | March 30, 2015 | 621 | 4285 | Not_Computer | Not_Computer
Z-1792-2014 | 68294 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | cobas 8000 analyzer series Operators Manual Roche cobas |  |  | May 21, 2014 | 2014 | Roche Diagnostics Operations, Inc. | Under certain conditions loose ProCell/CleanCell aspiration tube filters may cause: consecutive discrepant results; sipper alarm; foam on reservoirs and system reagent alarms. | Labeling design | Roche sent an Urgent Medical Device Correction (UMDC) letter dated May 20, 2014 via UPS Grount (receipt required) to affected customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to: 1. Follow the instructions in the Precautions Regarding Handling the Aspiration Tube Filter section in this UMDC; 2. Discontinue performing ProCell/Clean Cell aspiration tube filter check as this maintenance item willbe performed by your Roche FSR as a part of service maintenance; 3. If you have distributed the products listed within this UMDC to other sites or facilities, please provide them with a copy of the UMDC; and 4. Complete the faxback form enclosed with this UMDC and fax it to 1-888-410-6620. For questions contact the Roche Support Network Customer Support Center, 24 hours a day, seven days a week at 1-800-428-2336 if you have questions about the information contained in this UMDC. |  | 149 | 149 | US Nationwide Distribution | Terminated | June 04, 2015 | 379 | 1758 | Not_Computer | Not_Computer
Z-1787-2014 | 65483 | No | MUE | Full-field digital mammography system. | Radiology | Radiology | 510(k) | Beekley Spots Light Image O-Spots for Mammography |  |  | June 19, 2013 | 2013 | Beekley Corporation | Light Image O-Spots Dispenser mislabeled as Soft n' Stretchy Light Image S-Spots | Error in labeling | Beekley sent an Urgent Medical Device Recall letters dated June 19, 2013 via FedEx with signature and tracking confirmation. The letter identified the affected product, problem and actions to be taken. Firm requests users to examine inventory, remove and return recalled product. For questions contact Chelsea Fuller, Regulatory Manager at 1-800-233-5539 x 491 |  | 81 | 81 boxes | Worldwide Distribution - USA in the states of CA, FL, MA, NC, NY, and PA. | Terminated | June 16, 2014 | 362 | 81 | Not_Computer | Not_Computer
Z-1945-2014 | 68640 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Baxter, MiniCap with Povidone-Iodine solution, 5C4466P. For dialysis. |  |  | July 08, 2013 | 2013 | Baxter Healthcare Corp. | the white printed paper layer on the pouch packagingmay not be fully adhered to the pouch foil layer. The upper paper layer may serve as a protective barrier and help support overall packaging. | Packaging change control | An Urgent Product Recall notification was mailed via U.S.P.S. first-class mail to facilities, (dialysis centers and distributors) on July 8, 2013 and mailed by first-class mail to home patients on July 11, 2013. |  | 20940 | 20,940 units | Distributed in the states of AR, TN MO, KS, TX, MS, and AL. | Terminated | June 30, 2014 | 357 | 20940 | Not_Computer | Not_Computer
Z-0749-2014 | 63624 | No | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | Micro Curved Attachment (MCA) Cutting Burr, DFU 18-0001. All lots are included in recall. Part Number MCA |  |  | March 29, 2012 | 2012 | The Anspach Effort, Inc. | The Anspach Effort, Inc. is recalling all lots of the Anspach Micro Curved Attachment (MCA) Cutting Burrs and Dissection Tools. The recall was initiated after the indications for use were found to be broader than the validation of the MCA Cutting Burrs and its directions for use did not clearly state the correct intended use of the device. After a labeling correction, Anspach Effort, Inc. decide | Labeling Change Control | Firm is performing a Medical Device Recall by sending customers a recall letter and requesting the return of Anspach Micro Curved Attachments and Dissection Tools currently on hand. |  | 163 | 163 | Worldwide Distribution: US (nationwide) and Internationally to: AU, BE, CA, CH, ES, FR, GB, IT, JP, NL, SE, and ZA. | Terminated | March 11, 2015 | 1077 | 1412 | Not_Computer | Not_Computer
Z-1698-2014 | 68065 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | Novocastra Liquid Mouse Monoclonal Antibody Glutathione S-Transferase pi (NCL-L-GSTpi-438), a liquid tissue culture supernatant. For in vitro diagnostic use. |  |  | February 26, 2014 | 2014 | Leica Microsystems, Inc. | Novocastra Liquid Mouse Monoclonal AntibodyGlutathione S-Transferase pi (Product Code: NCL-L-GSTpi-438) does not function as intended up to the expiry date on the product labeling. This was detected through an ongoing stability program. There is a linkage between the age of the product and staining intensity. | Incorrect or no expiration date | Leica Biosystems sent a Medical Device Recall Notification letter dated February 26, 2014 to all direct accounts. The letter advised that clinicians may want to review results obtained from the affected products. The letter also included instructions for customers to: 1) appropriately destroy any unused or partially used NCL-L-GSTpi-438 (lots 6018777, 6023260,6026297, 6026980) reagent or confirm that they are not in stock; 2) complete and return the attached Medical Device Recall Notification Acknowledgement Form to 847-236-3747; and, 3) contact the local Leica representative immediately if replacement is required. Customers with questions or concerns about this recall can contact Leica via e-mail at LMGRA@leica-microsystems.com.For questions regarding this recall call 847-405-5413. |  | 33 | 33 units | Nationwide Distribution including AZ, CA, FL, and NY. | Terminated | August 14, 2017 | 1265 | 33 | Not_Computer | Not_Computer
Z-1697-2014 | 68284 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Integra Newdeal SURFIX Fixed Angle Locking Screw 3.5mm x 22mm (catalog number 285322S) and 3.5mm x 30mm (catalog number 285330S). Screws are provided in different diameters and lengths and are used to anchor various sizes and shapes of bone fixation plates when treating bone fractures or reconstruction of the ankle, foot or hand. |  |  | March 28, 2014 | 2014 | Integra LifeSciences Corporation | New Deal has identified through a complaint report that some 3.5 mm Diameter x 22 mm Length Surfix screws in Lot no. F626 (Catalogue No. 285322S) have been packaged and labeled as 3.5mm Diameter x 30 mm Length Surfix screws (Catalogue No. 285330S). | Package design/selection | Integra issued an Urgent Voluntary Medical Device Recall letter dated March 28, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review their inventory for affected product, immediately stopo using or distributing them. Customers were asked to complete the Acknowledgement Form and return by email or fax as indicated on the form. When the form is received customer service will contact customers to provide an RMA number and directions to return the affected products. Customers with questions were instructed to call 1-855-532-1723.For questions regarding this recall call 609-275-2700. |  | 417 | 417 units | Worldwide Distribution - USA including AR, CA, NC, TX and Internationally to Austria, Chile, Denmark, Finland, France, Germany, Greece, Ireland, Italy ,Lebanon, Portugal, Russia, South Africa, Spain, Switzerland, United Arab Emirates, and United Kingdom. | Terminated | April 18, 2016 | 752 | 417 | Not_Computer | Not_Computer
Z-0718-2014 | 63071 | No | HTW | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Anspach Single Use Sterile Bone Cutting Burs and Irrigation Clips. Rx Only |  |  | September 09, 2011 | 2011 | The Anspach Effort, Inc. | Recall was initiated due to the observation of separation between clear layers of the pouch during real time aging studies. As a result, it is possible to have a breach in the sterile barrier. | Packaging change control | Recalling firm notified direct accounts by letter on 09/09/11. Direct accounts were instructed to screen their inventory immediately for product packaged with the Tyvek Pouch/Tube Configuration, complete the reply form, and return the products to the firm. |  | 1860048 | 1,805,022 cutting burs and 55,026 irrigation clips | Worldwide Distribution - USA (nationwide) and the countries of Japan, Canada, Nigeria, Guam, China, Czech Republic, Thialand, Ecuador, Lithuania, Argentina, Turkey, Cayman Island, Jamaica, India, Jordan, Germany, Brazil, Mexico, Israel, Cameroon, Switzerland, South Africa, Greece, Finland, Taiwan, Hong Kong, France, Spain, Cyprus, Australia, United Kingdom, Sweden, Bahama, Korea, Malaysia, Iceland, Portugal, Polan, Austria, Belgium, Egypt, Singapore, Netherland, Arab Emirates, Kenya, Italy, Venezula, Columbia, Dominican Republic, Chile, Panama, Honduras, Costa Rica, Philippines. | Terminated | April 15, 2016 | 1680 | 1860048 | Not_Computer | Not_Computer
Z-1242-2014 | 67618 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Part 475620 Intramedullary Bone Saw CAM Assembly |  |  | February 26, 2014 | 2014 | Biomet, Inc. | Incorrect labeling on specific lots of the Intramedullary Bone Saw.The Saw Blade Assembly was identified as CAM Assembly. | Labeling False and Misleading | The firm, BIOMET, sent an "URGENT MEDICAL DEVICE SAFETY ALERT" letter dated February 26, 2014 via Fed Ex overnight to domestic consignees. Australia and GSCC were notified by email and also received a hard copy letter via mail. The letter described the product, problem and actions to be taken. The letter stated that this action requires the immediate location and discontinued use of the product and its return to Biomet. The Consignees were were instructed to immediately locate, discontinue use of the product and remove the identified device(s) listed; carefully follow the instructions on the FAX Back Response Form and reply within three (3) business days; complete and return Form via Fax to: 574-372-1683 prior to return of product, and use priority carrier for your shipment. If you have further distributed this product, you MUST notify hospital personnel of this action via enclosed "Dear Risk/Recall Manager Notice", and note that you are charged with the location and return of these products. Shipping address is Return Goods, 56 E. Bell Drive Warsaw, IN 46582. Questions related to the notice should be directed to 574-372-1570, Monday through Friday, 8 a.m. to 5 p.m. |  | 1 | 1 | Worldwide Distribution: US (nationwide) to states of: NY, TX, MI, and M; to countries of: Australia and The Netherlands. | Terminated | July 07, 2015 | 496 | 17 | Not_Computer | Not_Computer
Z-0191-2015 | 69464 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | BardÂż PerFix Light Plug |  |  | October 08, 2014 | 2014 | Davol, Inc., Subs. C. R. Bard, Inc. | The patient record peel-off label product code and size did not match the other labeling and actual product in the packaging for the Bard Small PerFix Light Plug. | Error in labeling | Davol Inc. (Subsidiary of C.R. Bard, Inc)sent Customer Notification letters on Friday, October 10, 2014, via FedEx next day delivery with signature confirmation to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Action required:Examine your inventory and identify any product subject to this communication. A sample of the product labeling for the lot has been enclosed to assist in product code and lot number identification. The enclosure also depicts the supplied label as well as an image of the label with the correct information.If you have any remaining inventory, a copy of this communication can be kept with the units for future medical record documentation purposes.If you have already used some of the product lot, you may amend any patient record(s) based upon the content of this communication.Please complete and return the accompanying Effectiveness Check Form attached to this letter regardless of whether or not you have any remaining units of the affected product lot.If you may have further distributed or transferred this product, please identify the respective organizations and notify them of this product communication. Your notification to these organizations may be enhanced by including a copy of this product communication letter.This product communication is being made with the knowledge of the Food and Drug Administration. We apologize for any inconvenience this may have caused your organization. Please contact the Davol Customer Service Department at 1-800-556-6275 or C.R. Bard Medical Services &amp; Support Department at 1-800-562-0027 if you have any questions. |  | 78 | 78 units | US Distribution including the states of CA, NM, OH, DE, MN, UT, MD and MO. | Terminated | December 23, 2015 | 441 | 78 | Not_Computer | Not_Computer
Z-1696-2014 | 68124 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | BardÂż Ventralight" ST Mesh with Echo PS" 8"x10" Reorder Number: 5955810 |  |  | April 24, 2014 | 2014 | Davol, Inc., Subs. C. R. Bard, Inc. | Pouch holding sterile inflation assembly and inflation adapter may be open or have a weak seal. | Packaging process control | Davol, Inc., Subs. C. R. Bard, Inc.sent an Urgent Medical Device Recall letter dated April 28, 2014, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to examine their inventory and quarantine product subject to the recall, contact Davol Customer Service at 1-800-556-6275 or C.R. Bard's Medical Services and Support Department at 1-800-562-0027 for instructions on how to return the product to Davol, Consignees were also instructed to complete and return the recall effectiveness check form, and notify their customers if the product was further distributed.For questions regarding this recall call 1-800-556-6275. |  |  |  | Worldwide Distribution - USA (nationwide) and Internationally to AUSTRIA, BELGIUM, CANADA, CYPRUS, ESTONIA, FINLAND, FRANCE,GERMANY, GREECE, IRELAND, ITALY, LUXEMBOURG, NETHERLANDS, NORWAY, POLAND, PORTUGAL, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, and UNITED KINGDOM | Terminated | December 23, 2015 | 608 |  | Not_Computer | Not_Computer
Z-0719-2014 | 66907 | No | FGX | Manual gastroenterology-urology surgical instrument and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | TWISTER PLUS 22mm Rotatable |  |  | November 20, 2013 | 2013 | Boston Scientific Corporation | Sterility of device may be compromised due to lack of pouch seal integrity | Packaging process control | Boston Scientific issued Customer Notification letters to US accounts via Federal Express Priority mail on November 20, 2013. The notification instructs the account to check their inventory, remove any affected product from their inventory, and return it to Boston Scientific. The notification also requests the completion of a Reply Verification Tracking Form. |  |  | 0 eaches | Worldwide Distribution-USA (nationwide) and the countries of Canada, Austria, Australia, New Zealand, France, Spain, Portugal, Italy, Switzerland, Denmark, and Netherlands. | Terminated | February 24, 2015 | 461 | 5338 | Not_Computer | Not_Computer
Z-1710-2014 | 68090 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | Liquid Mouse Monoclonal Antibody CD15 (NCL-L-CD15), a liquid tissue culture supernatant containing 15 mM sodium azide as a preservative. For in vitro diagnostic use. |  |  | January 10, 2014 | 2014 | Leica Microsystems, Inc. | Product does not function as intended up to the expiry date on the product labeling. There is a link between the age of the product and staining intensity. | Incorrect or no expiration date | Leica sent an Urgent Field Safety Correction Notice letter dated January 10, 2014 was sent ot all affected customers. The notice included instructions for recipients to: 1) destroy any unused or partially used NCL-L-CD15 (lots 6015073 and 6017936) reagent or confirm that the products are no longer in stock; 2) indicate on the attached Field Safety Correction Notice Acknowledgement Fom which action was taken; and, 3) contact the local Leica represetnative immediately if a replacement is required. Customers with any concerns about this notice can contact Leica via e-mail at LMGRA@leica-microsystems.com. |  | 145 | 145 units | Worldwide Distribution - US Nationwide in the states of CA, CT FL, IL, MO, MS, NC, NY, PA, and the countries of Argentina, Australia, Brazil, Colombia, Italy, Malaysia, New Zealand, Nigeria, Poland, Romania, Russia, Saudi Arabia, Serbia, South Africa, Sweden, Taiwan, Tunisia, Turkey, and United Kingdom | Open, Classified |  |  | 145 | Not_Computer | Not_Computer
Z-1233-2014 | 67194 | No | LYC | Bone grafting material. | Dental | Dental | 510(k) | ENDOBON XENOGRAFT GRANULES 0.5 ML (Bovine bone substitute) |  |  | October 03, 2013 | 2013 | Biomet 3i, LLC | Individual pieces of a batch of Biomet's Endobon Xenograft Granules 0.5 ML (Bovine Bone substitute) may have an incorrect label with the wrong batch number. | Error in labeling | Biomet 3i sent a Medical Advisory Notice dated October 3, 2013 via mail, facsimile or telephone to all affected customers. The notice identified the product, the problem, and the action to be taken by the customers. The customers were instructed to do the following: 1) Review their inventor for any incorrectly labeled products and return them to their local Biomet distributor. 2) Provide this information to persons in their organization that uses or order these products. 3) Ensure that a copy of this letter is provided to any other organization to which the products maybe have been transferred. Customer were also instructed to respond to the attached Response Form via fax at 561-514-6316 or email to: domesticcomplaints@biomet.com. For any questions or concerns, call Biomet 3i Customer Service at 1-800-342-5454.For questions regarding this recall call 561-776-6706. |  | 704 | 704 | Worldwide Distribution - USA including CA, DC, FL, LA, MA, MO, NC, NJ, NY, TX, VA, WA and Hawaii and Internatially to Austria, Belgium, Canada, Switzerland, Germany, Spain, France, Great Britain (UK), Ireland, Mexico, Poland, Portugal, and Singapore. | Terminated | December 11, 2014 | 434 | 704 | Not_Computer | Not_Computer
Z-0174-2015 | 69454 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Straumann Bone Level Implant 04.1 mm RC, SLActiveÂż 8mm, RoxolidÂż, Loxim |  |  | September 24, 2014 | 2014 | Straumann USA, LLC | Product transfer piece was fitted with a Narrow Connect (NC) instead of a Regular Connect (RC) | Packaging process control | Straumann USA sent a Field Corrective Action Notice letter dated September 24, 2014 and delivered via United Parcel Service. The letter identified the affected product, problem and actions to be taken. Straumann is requesting the return of unused implants. Returns and Customer Confirmation form is to be returned. Contact firm 781-747- 2514. |  | 111 | 111 units | US Nationwide Distribution | Terminated | June 26, 2017 | 1006 | 111 | Not_Computer | Not_Computer
Z-1720-2014 | 68039 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | SBi RingFix Olive Wires |  |  | April 07, 2014 | 2014 | Small Bone Innovations, Inc. | Small Bone Innovations (SBI) is recalling non-sterile RingFIX Olive Wires labeled as item #451-3004 Drill tip 1.8 mm diameter 400mm olive k-wire packaged with item #451-3005 Bayonet tip 1.8mm diameter 400mm olive k-wire inside. | Packaging process control | SBi sent an Urgent Medical Device Recall dated April 7, 2014, to affected customers (sales consultants, distributors, and end users). The letter identified the affected product, problem and actions to be taken. For questions call 215-428-1791 ext 248. |  | 44 | 44 | US Nationwide Distribution in the states of IN, NC, WA, NH, and CT. | Terminated | October 20, 2014 | 196 | 44 | Not_Computer | Not_Computer
Z-1226-2014 | 67282 | No | CIG | Bilirubin (total or direct) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | VITROS Chemistry Products White Reference Slides, White Correction Factor Slide (WCF), DT Slides, 25 Slides, Part Number J02315; Made in USA by Ortho- Clinical Diagnostics, Inc. 100 Indigo Creek Drive, Rochester, NY 14626 --- This product is a part/ tool used when performing the correction factors adjustment on the VITROS Chemistry Systems. It does not have a 510(k) number. |  |  | January 09, 2014 | 2014 | Ortho-Clinical Diagnostics | An incorrect assay value (0.3785) was printed on the product carton for the 460 nm wavelength (filter) that may cause a delay in the ability to obtain and report TBIL, Bu, Bc, and derived tests results. The correct assay value is 0.8658. | Error in labeling | Ortho Clinical Diagnostics sent an Urgent Distributor Product Correction Notification letter dated January 9, 2014, via FedEx overnight courier to all affected customers. The letter notified users of the issue and the correct 460 nm value. Foreign affitliates were notified on January 9, 2014 via e-mail. The letter advised users to inspect their current inventory of VITROS White Correction Slides to determine if Lot 9052-0045-6358 is present. If there is remaining inventory of VITROS White correction Slides, Lot 9052-0045-6358, the remaining inventory may be discarded or the user may continue to use the product with the updated 460nm value provided. If VITROS White Correction Slides, Lot 9052-0045-6358 have been previously used to perform the reflectometer correction factors adjustment , then this procedure should be repeated using the corrected 460 nm values or an alternate lot of VITROS White Correction Slides. Customers were instructed to complete and return the Confirmation of Receipt Form and identify all customers and any other branches who were shipped the affected lot from their facility and send a copy of the customer communication. Customers were also instructed to forward the information if they have distributed the affected product outside of their facility. Customers with questions were instructed to call Customer Technical Services.For questions regarding this recall call 908-218-8776. |  | 64 | 64 units worldwide | Worldwide Distribution - USA (nationwide) Puerto Rico, and Internationally to Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Republic of Panama, Singapore, Spain, and Venezuela. | Terminated | July 07, 2015 | 544 | 64 | Not_Computer | Not_Computer
Z-0177-2015 | 69512 | No | DRF | Electrode recording catheter or electrode recording probe. | Cardiovascular | Cardiovascular | 510(k) | St. Jude Medical, Fixed Curve Response 6F Electrophysiology Catheter, CSL, Model 401353 |  |  | October 14, 2014 | 2014 | St Jude Medical | St. Jude Medical is performing a voluntary recall of batch 4671318, model number 401353 of the Response Fixed Curve DECA 6F 65cm CSL 2/8 Catheter. Specifically, the firm has identified that every device packaged within batch 4671318 contains the incorrect catheter (Supreme FC QUAD 5F catheter) within the packaging. To date, no serious injuries have occurred as a result of this issue | Labeling mix-ups | St. Jude Medical sent an "Urgent Medical Device Recall Notice" letter dated October 7, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Advised customers to discontinue the use of the recalled device and inform them that a St. Jude Medical sales representative will contact them in the near future to facilitate removal and reconciliation of the recalled device. For questions they can contact their St. Jude Medical Sales Representative. For further questions please call (651) 756-2295. |  | 26 | 26 | US Distribution including the states of CA, GA, NJ, NY, OR and PA. | Terminated | December 10, 2014 | 57 | 26 | Not_Computer | Not_Computer
Z-1268-2014 | 67599 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | AMS URIC ACID For the In vitro quantitative determination of Uric Acid in serum |  |  | February 11, 2014 | 2014 | AMS Diagnostics, LLC | Marketing the devices outside 510(k) requirements | No Marketing Application | AMS Diagnostics sent a Recall Notification letter dated February 11, 2014 to all affected customers. The letter instructed the customers to: 1) Review their inventory for the lots listed and if they currently have them, to discard and notify the firm for a replacement. 2) If they don't have the lots listed, please check the box, sign, date and return the letter via fax at (954) 217-0046 or email to dj@amsdiagnostics.com. Customers with questions were instructed to call 954-217-0040 ext. 302.For questions regarding this recall call 954-217-0040. |  | 13 | 13 kits | Nationwide Distribution including FL, GA, MD, CA, NC, MA, SC, WY, and TX. | Open, Classified |  |  | 296 | Not_Computer | Not_Computer
Z-0207-2015 | 69500 | No | HHW | Vaginal pessary. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | CooperSurgical MILEX PESSARY KITS-Ring w/Support Folding Pessary- #1 |  |  | September 30, 2014 | 2014 | CooperSurgical, Inc. | Incorrect expiration date printed on the MILEX PESSARY KIT packaging | Incorrect or no expiration date | CooperSurgical notified customers a Recall Notification dated September 29, 2014, via UPS . The letter identified the product problem and the action needed to be taken by the customer. CooperSurgical's will replace any of the affected and provided further instruction on the disposition of the product at the time of replacement. If you have any further questions, contact 203.601.5200.. |  | 17 | 17 units | US Distibution in states of:OH, CA, WI, WA, FL, MN, NH, RI, MA, IL, MI, IA, and AL | Terminated | May 28, 2015 | 240 | 23 | Not_Computer | Not_Computer
Z-1270-2014 | 67607 | No | DJG | Opiate test system. | Toxicology | Toxicology | 510(k) | Single Analyte DAU Calibrator (5 mL) |  |  | February 18, 2014 | 2014 | Lin-Zhi International Inc | Customer reported that the Synchron NBUP cutoff calibrator (10 ng/mL) could not be qualified through quality control testing. The cutoff calibrator was recovering at a concentration close to the NBUP Control Level 1 (7 ng/mL). | Packaging | Lin-Zhi International sent a Recall letter dated February 18, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Letters acknowledge that Lin Zhi was notified by the customer that the product could not be qualified and that this was confirmed by Lin-Zhi internal testing.Letters request that the distributor notify all customers of the situation and ask them to destroy any remaining product from the affected lots. the root cause is identified and letters also state that a corrective action has been implemented to prevent future problems.Questions may be directed to Lin Zhi International at 408-732-3856. |  | 7200 | 120 units, 60 of each lot number | US Distribution including CA. | Terminated | April 25, 2014 | 66 | 7200 | Not_Computer | Not_Computer
Z-1210-2014 | 67686 | No | FMQ | Protective restraint. | General Hospital | General Hospital | 510(K) Exempt | DeRoyal(R) LIMB HOLDERS, COTTON ADJUSTABLE, UNIVERSAL, REF M2017 |  |  | February 21, 2014 | 2014 | DeRoyal Industries Inc | Moldy smell in product | Packaging | DeRoyal initiated an Urgent Voluntary Recall dated February 21, 2014 to all affected customers. The letter identified the affected product, problem and actions to be taken. DeRoyal has asked that products be destroyed, and that notice of destruction be sent back. For questions contact your DeRoyal Sales Representative or DeRoyal Customer Service at 1-800-251-9864. |  | 8962 | 8,962 pairs | Worldwide Distribution - USA Nationwide in the states of FL, NC, IA, AL, IL, TX and the countries of Argentina. | Terminated | May 08, 2015 | 441 | 8962 | Not_Computer | Not_Computer
Z-1664-2014 | 68281 | No | HMX | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | Visitec Anterior Chamber Cannula, .40x22mm (27Gx7/8in) |  |  | May 06, 2014 | 2014 | Beaver-Visitec International Inc. | Contains a sharp tip needle when the product should contain a blunt tip cannula | Packaging process control | Beaver Visitec (BVI) sent a Product Recall Notice dated May 6, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately return any un-opened product of the affected lot so that BVI can replace the product or issue credit to their company. Customers were asked to fax or email the customer response form to 866-906-4304 or claimsUS@beaver-visitec.com as acknowledgement of the letter. Customers with questions should call 781-906-7986.For questions regarding this recall call 781-906-7950. |  | 9970 | 9970 units | Nationwide Distribution including AL, CT, IN, NC, NY, TX, and VA. | Terminated | September 05, 2014 | 122 | 9970 | Not_Computer | Not_Computer
Z-1663-2014 | 68221 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | MAGNETOM Aera magnetic resonance system. |  |  | April 24, 2014 | 2014 | Siemens Medical Solutions USA, Inc | Siemens discovered a processing error for the MAGNETOM Aera magnetic resonance system that may have resulted in some users not receiving Operator Manuals. | Labeling design | A customer safety advisory notice, dated April 24, 2014 was sent to notify direct accounts of the delivery of user manuals. |  | 5 | 5 | Distributed in the states of MN, CA, CO, TX, and IA. | Terminated | January 08, 2015 | 259 | 5 | Not_Computer | Computer
Z-0716-2014 | 67142 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Locking Cortical Screw |  |  | December 19, 2013 | 2013 | Biomet, Inc. | 3.5mmx38mm Locking Cortical Screw, implants are being recalled since an investigation identifiedthat a 3.5mmx28mm Cortical Locking Screw is labeled as a 3.5mmx38mm Cortical Locking Screw. | Labeling mix-ups | URGENT MEDICAL DEVICE RECALL notices were sent to all consignees on December 19, 2013. Distributors who have affected hospitals in their regions were notified and provided with a copy of the notice that went to each hospital. Consignees are being contacted on an individual basis in an attempt to reconcile the product 100%. Fax back forms were provided to all consignees in order for Biomet to track returns. Non-responders will be contacted on an individual basis. |  | 52 | 52 | US Distribution including states of: AZ, CA, GA, IL, KY, MI, NC, NH, NY, OH, PA, SD, TX, UT, and VA. | Terminated | July 07, 2015 | 565 | 52 | Not_Computer | Not_Computer
Z-1208-2014 | 67627 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | REF 623421 RS-ROM Noiles Rotating Hinge, Femoral Rotating Hinge Cemented (with Pin), Right , X-Small, Sterile R |  |  | February 24, 2014 | 2014 | DePuy Orthopaedics, Inc. | DePuy Orthopaedics has identified the potential for holes to develop in the inner and outer flexible pouches that form the sterile barrier for S-ROM Noiles Rotating Hinge Femur with Pin devices. The outer carton and shrink wrap are intact. A package redesign is underway to resolve this issue. S-ROM Noiles Rotating Hinge Femur with Pin devices will remain in US Distributors inventory during the pa | Packaging | DePuy Orthopaedics sent an Urgent Information - Recall Notice dated February 24, 2014, to all affected customers. The notice identified the product, reason for the recall and potential risks associated with using the recalled product. The notice includes response cards for the US Distributors to verify they have read, understood, and are adhering to the directions given in the field communication and for hospitals where consignment inventory is removed a hospital letter is available, as needed. No hospital signatures are required. Distributors are asked to send completed distributor cards to DePuy. Product related questions should be addressed to DePuy Orthopaedics, .Sales Consultant. Clinical questions should be addressed to DePuy Scientific Information Office 1-888-554-2482; recall questions to 574-371-4917. |  | 132 | 132 | Worldwide Distribution - USA (nationwide) and Internationally to Canada, Australia, Austria, China, Czech Republic, Denmark, Dubai / UAE, Estonia, France, Germany, Hungary, Iberia, India, Ireland, Israel, Italy, Malaysia, Netherlands, New Zealand, Norway, Pakistan, Poland, Portugal, Russia, Singapore, Slovenia, Spain, South Africa, Sweden, Switzerland, Turkey, United Kingdom, and Vietnam. | Terminated | March 03, 2015 | 372 | 790 | Not_Computer | Not_Computer
Z-1748-2014 | 68081 | formulations, paraffin, all | KEO | General purpose reagent. | Pathology | Pathology | 510(K) Exempt | Tissue-Tek PARAFORM PROC/EMB/MEDIUM, FORMULA 3 - 8x1Kg. |  |  | April 16, 2014 | 2014 | Sakura Finetek USA Inc | This Lot has variations in melting temperature. | Packaging process control | Customer notifications sent via e-mail letter.The letter titled "PRODUCT RECALL" dated 04/17/2014 informed customers (of Lot 263922), informed customers of the recall by providing the following information: product name, code, lot number, reason for recall, requested actions, contact information and an "FIELD CORRECTION EFFECTIVENESS CHECK" form. |  | 115 | 115 | Nationwide Distribution. | Terminated | September 02, 2014 | 139 | 115 | Not_Computer | Not_Computer
Z-0153-2015 | 69392 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Spline Twist Implant, HA 3.25mmD X 15mmL; Part No. 2135. |  |  | October 01, 2014 | 2014 | Zimmer Dental Inc | Zimmer Dental is recalling Tapered Screw-Vent Implants because the incorrect diameter size was printed on the inner vial. | Labeling Change Control | Customer Notification letters dated 10/6/14 were sent to all customers who purchased the Tapered Screw Vent Implant and the Spline Twist Implant because the incorrect diameter size was printed on the inner vial. The letters informs the customers of the problems identified and the actions to be taken. Customers are instructed to complete the attached Business Reply Form and fax it to (574) 372-4265 or email to corporatequality.postmarket@zimmer.com. Customers with questions are instructed to contact Zimmer Dental Customer Service at (800) 854-7019 from 7am-5pm, Monday-Friday. |  | 1044 | 1,044 units | Worldwide Distribution-US (nationwide) and the countries of Poland, Canada, Belgium, Italy, France, Israel, Morocco, Spain, Costa Rica, and Japan. | Terminated | April 28, 2015 | 209 | 1044 | Not_Computer | Not_Computer
Z-1641-2014 | 67991 | Yes | BTR | Tracheal tube. | Anesthesiology | Anesthesiology | 510(k) | Portex Low Dead Space Connector with Sideport, 3.5mm |  |  | April 10, 2014 | 2014 | Smiths Medical ASD, Inc. | One lot of 3.0mm sized Connectors, Product Reorder No. 10035-05PS, Lot No. 2553426 were placed in packages labeled as size 3.5mm. | Packaging process control | Smiths Medical sent an Urgent Medical Device Recall Notice dated April 10, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to inspect their inventory and quarantine all affected products and to complete the attached Response Form and return it to Smith Medical within 7 days of receipt of the notice. Customers with questions were instructed to contact Smiths Medical's Customer Service Department at +1-800-258-5361. |  | 100 | 100 devices | US distribution in the state of North Carolina. | Terminated | November 29, 2016 | 964 | 100 | Not_Computer | Not_Computer
Z-1652-2014 | 67956 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Siemens Linear Accelerator (LINAC) models. |  |  | April 01, 2014 | 2014 | Siemens Medical Solutions USA, Inc | Siemens Radiation Oncology became aware that customers may be using the Siemens Linear Accelerator in combination with stereotactic accessories which have not been validated as being compatible with Siemens LINAC models. | Labeling design | Siemen sent an Urgent Medical Device Correction letter to all affected customers. The letter identified the affected product, problem and actions to be taken. For questions contact your local Elekta representative. |  | 9 | 9 | US Nationwide Distribution in the states of AL, ME, PA, IL, and CA. | Terminated | March 16, 2015 | 349 | 9 | Not_Computer | Computer
Z-1300-2014 | 67567 | No | LJT | Subcutaneous, implanted, intravascular infusion port and catheter. | General Hospital | General Hospital | 510(k) | PowerPort Slim Implantable Port, indicated for patient therapies requiring repeated access to the vascular system. Product Codes: 1716000, 1716070 |  |  | February 27, 2014 | 2014 | Bard Access Systems | Labeling discrepancy.The label states that the implantable port contains silicone suture plugs, however approximately 23% of the units contained in the affected lots have a PowerPort Slim titanium implantable port without silicone suture plugs. | Packaging | Notified by Urgent Customer Notification Letter date February 13, 2014, customers were asked to immediately examine their inventory and identify any product subject to this notification. In addition, if they have further distributed this product, please identify customers and notify them at once of this notification. Customers may include a copy of this letter in your notification to customers. Customers may choose to use the involved product with the understanding that the kit may contain a PowerPortÂż Slim implantable port without suture plugs, or may return the applicable product for replacement or credit. |  | 777 | 777 | US Nationwide distribution including WA, KY, IL, and IA. | Terminated | September 05, 2014 | 190 | 777 | Not_Computer | Not_Computer
Z-0114-2015 | 69256 | No | CAZ | Anesthesia conduction kit. | Anesthesiology | Anesthesiology | 510(k) | ASK-05501-AFH1, Lot Number 23F13K0969 Anesthesia Conduction Catheter, Epidural Anesthesia Kit |  |  | September 15, 2014 | 2014 | Arrow International Inc | Arrow International, Inc. is notifying customers that lot 23F13K0969 of product code ASK-05501-AFH1 contains an incorrect component. The lidstock indicates that the kit contains one Syringe, 10mL, Luer-Slip Loss of Resistance, Glass but contains one Syringe, 10mL, Luer-Lock Loss of Resistance, Glass. | Labeling mix-ups | Arrow International sent a Medical Device Advisory Notification dated September 15, 2014, to affected customers notifying them that lot 23F13K0969 of product code ASK-05501-AFH1 contains an incorrect component, and requesting they provide this Advisory Notice to all those who need to be aware of it within your organization. Questions were directed to local sales representative or Customer Service at 1-866-246-6990. |  | 300 | 300 | US Distribution in Utah only. | Terminated | December 18, 2015 | 459 | 300 | Not_Computer | Not_Computer
Z-0232-2015 | 69579 | No | JWY | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | EtestÂż Ceftaroline (CPT32) US B30. |  |  | October 16, 2014 | 2014 | Biomerieux | The products have a wrong expiration date on their labeling, 5 years instead of 2 years. | Labeling Change Control | Consignees of bioMÂżrieux, impacted by the identified issue, were notified of the FSCA and corrective action (FSCA notification) electronically on October 16th, 2014. Additionally, the bioMÂżrieux United States subsidiary notified US consignees of the FSCA notification via FedEx on October 23rd, 2014. |  | 652 | 652 devices | Worldwide Distribution-US (nationwide) including DC and the states of AL, AK, AR, AZ, CA, CO, DE, FL, GA, IA, IN, KY, MA, MD, MI, MO, NC, NE, NY, OH, OK, PA, SC, TN, TX, UT, VA, WA, WI, and the countries of Australia, Austria, China, Czech Republic, France, Germany, India, Italy, Japan, Netherlands, Poland, Portugal, Switzerland, Sweden, Turkey, United Kingdom, Egypt, Pakistan, Slovenia and Uruguay. | Open, Classified |  |  | 652 | Not_Computer | Not_Computer
Z-2024-2014 | 68607 | No | GAR | Nonabsorbable polyamide surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Dafilon Suture Black 10/0 (0.2) 15 cm DRM4 NS, non-sterile, non-absorbable polyamide Surgical Suture, additional label affixed on the box states the product is sterile, packaged in a cassette holding 12 sutures |  |  | June 13, 2014 | 2014 | Aesculap, Inc. | Aesculap Inc. (AIC (USA)) has initiated a voluntary recall of Dafilon Black 10/0 (0.2) 15 cm DRm4 non-sterile non-absorbable polyamide Surgical Suture due to a labeling issue. Box labels do state non-sterile; however there is an additional label that is affixed on the box that states the product is sterile. | Labeling design | A correction and removal notification, dated June 13, 2014, was sent to end users, sales reps, and distributors which described the product, problem, and actions to be taken. |  | 11 | 11 | Distributed in the states of IL, NY, GA, and RI. | Terminated | October 07, 2014 | 116 | 11 | Not_Computer | Not_Computer
Z-2025-2014 | 68608 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | SmartSite Low Sorbing Infusion Set, Model Codes: 10014855A, 10015861A, 10061661A |  |  | June 17, 2014 | 2014 | CareFusion 303, Inc. | CareFusion is recalling the SmartSite Low Sorbing Infusion Set because they were assembled with a drip chamber that consist of a 15 micron filter. The Directions For Use did not indicate that the drip chamber has a 15 micron filter and it does not provide the instruction that the set should not be used with blood administration due to the potential for occluding the 15 micron filter in the drip c | Labeling Change Control | A recall letter is dated 6/25/14 and will be sent out to customers who purchased the SmartSite Low Sorbing Infusion Set because they were assembled with a drip chamber that consist of a 15 micron filter. The Directions For Use did not indicate that the drip chamber has a 15 micron filter and it does not provide the instruction that the set should not be used with blood administration due to the potential for occluding the 15 micron filter in the drip chamber. The letter informs the customers of the problems identified, potential risk, and actions to be taken. Customers are instructed to promptly complete and return the enclosed mandatory Customer Response Card to acknowledge the receipt of this communication and to expedite the corrective action process. |  | 5600 | 5,600 units | Distributed in the states of FL, NM, WA, WI, and TX. | Open, Classified |  |  | 5600 | Not_Computer | Not_Computer
Z-1130-2014 | 67468 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | FMP X-alt Acetabular Liner |  |  | February 07, 2014 | 2014 | Djo Surgical | Packaging error -- two different types and sizes of acetabular liners used during hip replacement surgery were mistakenly packaged into each other's outer boxes after sterilization, resulting in the products being mislabeled. The units labeled as Model/Catalog Number 931-28-248 Lot/Serial Number 685F1034 on the outside and containing Model/Catalog Number 932-36-252, Lot# 728F1089 on the inside (1 | Packaging process control | DJO Global sent an Urgent Field Safety Notice dated February 7, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. All 10 devices with the outer label of 931-28-248, Lot#685F1034 were found in-house and were quarantined. Theagencies who have the affected product (devices labeled as932-36-252, Lot# 728F1089 containing 931-28-248, Lot#685F1034) were contacted to quarantine their devices. Of the affected devices, 2 of the lot of 10 have already been returned via product complaint.Customers were instructed to pass on the notice to all those who need to be aware within their organization or to any organization where the potentially affected devices have been transferred. Customers were instructed to place a replacement order and receive an RMA number to return all affected devices.. Customers with questions were instructed to call 512-834-6255. |  | 931282480 | 10 units of Model/Catalog Number 931-28-248 Lot/Serial Number(s) 685F1034 labeled asModel/Catalog Number 932-36-252, Lot# 728F1089 | Nationwide Distribution - USA including TX, LA, FL, CA, and AZ. | Terminated | October 01, 2014 | 236 | 931282480 | Not_Computer | Not_Computer
Z-1122-2014 | 67545 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Triton FP Infusion Pump, Part Number 400000, packaged in boxes. |  |  | February 17, 2014 | 2014 | WalkMed Infusion, LLC | WalkMed Infusion is initiaing a recall on their Triton FP Infusion Pump, Part Number 400000 due to the label on the right side of the pump containing an incorrectinstruction. | Labeling design | WalkMed Infusion sent an Important Medical Device Recall letter dated February 14, 2014, to all affected customers. The letter identified the product the problem and the action customers need to take.Our records Indicate that you have purchased product affected by this recall (see the attached list).Our sales representative will contact you to determine a plan for label replacement for your units. Replacement of the label will include a simple process of removing the old label (PN 300063), replacing it with a corrected label(PN 400033) and documenting completion of the replacement. This action can be completed by you,your customers,our sales representative, or by returning the infusion pump to WalkMed Infusion.For further questions please call (720) 351-4934 or (720) 351-4965.Product will be corrected in the field when possible and supported by the customer. Otherwise, WalkMed Infusion employees with travel to the customer and complete the correction activities. |  | 797 | 797 | US Distribution including the states of IL, MA, KS, and, NJ. | Terminated | October 06, 2014 | 231 | 797 | Not_Computer | Not_Computer
Z-0229-2015 | 69461 | No | NBC | B-type natriuretic peptide test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA Centaur BNP Assay for in vitro diagnostic use in the quantitative determination of B-type Natriuretic Peptide (BNP) in human plasma using the ADVIA Centaur, Centaur XP, Centaur CP Systems; Catalog Numbers: |  |  | October 06, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc | Siemens Healthcare Diagnostics has confirmed that approximately 2000 ReadyPacks from the ADVIA CentaurÂż Systems BNP reagent contains an ADVIA Centaur Systems FT4 human-readable side label. | Labeling mix-ups | A Customer Notification (CN) was sent to all affected Siemens Healthcare Diagnostics customers in the United States on October 03, 2014 to be delivered to customers on October 06, 2014 via Federal Express. An Urgent Field Safety Notice (UFSN) was sent out for distribution to all customers outside the United States on October 06, 2014. These notices inform the customer of the potential to receive an ADVIA Centaur BNP ReadyPack with a FT4 human-readable side label. Siemens instructs customers to inspect their inventory and discard any BNP Readypacks with a FT4 side label. In addition, Siemens informs customers the end labels of the affected ReadyPacks correctly identify the pack as BNP and Siemens has confirmed the reagent within these packs is BNP Reagent. Therefore valid BNP results will be generated when calibrating and running quality control or patient samples. Fax-back forms will be used to confirm that the customers received the notification. For those customers that do not return a fax-back form, a Siemens representative will follow up with those customers to ensure they received the Customer Notification and completed the required actions. |  | 7078 | 100 test - 4,600 kits; 500 test - 1,878 kits | Worldwide Distribution - US Nationwide including Puerto Rico; and countries of: Canada, Austria, Belgium, Czech Republic, Denmark, Estonia, Finland, France, Fren.Polynesia, Germany, Great Britain, Greece, Guadeloupe, Italy, Latvia, Lithuania, Netherlands, Norway, Poland, Portugal, Reunion, Russian Fed., Slovakia, Slovenia, Spain, Switzerland, Turkey, Unit.Arab Emir., AR, AU, BR, CL, CN, CR, IN, JP, KR, LK, MY, SG, TW, UY, VN, and ZA. | Terminated | July 22, 2015 | 289 | 7078 | Not_Computer | Not_Computer
Z-1984-2014 | 68383 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Synthes B37 Replacement Screws. M3.5 Screw for insertion handle guide/aiming arm, part numbers 3.113.025 and 03.113.026. Surgical instrument for use with Low Bend Medial Distal Tibia Plate. Orthopedic. |  |  | May 15, 2014 | 2014 | Synthes, Inc. | All lots of B37 Replacement Screws are being recalled due to potential mislabeling where the package may not contain the screws identified on the label. | Labeling design | An urgent medical device recall notice, dated May 29, 2014, was sent to end users and sales consultants to advise them of the issue and provide instructions for returning affected devices. |  | 44 | 44 | Distributed US nationwide, in DC and the states of KY, PA, IL, WA, AZ, CA, OK, FL, and KS. | Terminated | December 18, 2015 | 582 | 44 | Not_Computer | Not_Computer
Z-0424-2015 | 69366 | No | HWA | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Trabecular Metal Tibial Impactor Replacement Pads. |  |  | October 14, 2014 | 2014 | Zimmer, Inc. | To provide clarifying instructions relating to the assembly/disassembly of the Trabecular Metal Tibial Impactor with/from TM tibialImpactor Pads due to complaints regarding Tibal Impactor pad fracture. | Labeling design | On 10/14/2014, URGENT MEDICAL DEVICE CORRECTION notifications were sent to the affected distributors and hospital Risk Managers. All distributors were notified via electronic mail, and all hospital Risk Managers and distributors with affected inventory were also notified via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. For questions, call the customer call center at 1-877-946-2761. |  | 4961 | 4961 units | Worldwide Distribution - US Nationwide in the state of AL, CA, CO, FL, GA, HI, IA, ID, IL, IN, KS, KY, MA, MD, ME, MI, MN, MO, ND, NH, NJ, NV, NY, OH, OR, PA, SC, TX, VA, WI, NC, TN, OK, AZ, UT, AK, PR, DC. and the countries of: Germany, China, Canada, Australia, Japan, United Arab Emirates, Austria, Belgium, Switzerland, Germany, Denmark, Spain, Finland, France, United Kingdom, Greece, Iceland, Italy, Malta, Netherlands, Norway, Poland, Russian Fed., Saudi Arabia, Sweden, Slovakia, and South Africa. | Terminated | June 09, 2015 | 238 | 6126 | Not_Computer | Not_Computer
Z-1614-2014 | 68161 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Coaxial Micro-Introducer Kit. |  |  | April 28, 2014 | 2014 | Greatbatch Medical | Greatbatch identified that the documentation does not support the five-year shelf life of the Bard Access Division Coaxial Micro-Introducer Set. | Incorrect or no expiration date | Greatbatch sent an "Urgent Medical Device Recall" letter dated April 25, 2014 to affected customers via email and Fedex. The letter described the Reason for Recall, Risk to Health, and Instructions to Customers. Requested consignees to immediately place affected devices in quarantine and to complete and return the Field Correction Action Response Form. For questions they should contact their local Greatbatch representative or 1-763-951-8235, or e-mail FieldActionCenter@Greatbatch.com. |  | 6642 | 6,642 sets (66,420 devices) | US Nationwide Distribution in the state of UT. | Terminated | May 24, 2016 | 757 | 6642 | Not_Computer | Not_Computer
Z-2002-2014 | 68415 | No | FRC | Sterilization process indicator. | General Hospital | General Hospital | 510(k) | 3M Attest Auto-reader. Model numbers 390, 390G,and 490. |  |  | June 09, 2014 | 2014 | 3M Company - Health Care Business | Labeling on units shipped prior to May 16th, 2014 did not contain the statement "This product contains dry natural rubber" as required by the United States Food and Drug Administration (FDA) for medical devices. The non-slip pads at the bottom of the unit contain dry natural rubber. One incident of an allergic reaction in a sensitized individual has been reported. | Labeling Change Control | Consignees were sent on 6/9/2014 a 3M "Urgent Medical Devices Correction" letter dated May 21, 2014. The letter described the problem and the product involved in the recall. The letter described what action is 3M taking and what action was required of the consignees. Consignees were sent an adhesive label and a response card. They requested consignees to adhere the label to a portion of the plastic covering each unit in a location where it is visible to the end-user. After that, they are requested to fill out and return the response card. |  | 5343 | 5343 (2621 USA, 2722 OUS) | Worldwide Distribution-USA (nationwide) including DC and Puerto Rico and the countries of AUSTRALIA, CANADA, CHILE , COLOMBIA, COSTA RICA, SOUTH AFRICA, BRAZIL, GERMANY, ECUADOR, DUBAI, JAPAN, HONG KONG , CHINA, KOREA, MALAYSIA, PANAMA, SINGAPORE, VATICAN, and THAILAND. | Terminated | June 03, 2015 | 359 | 5343 | Not_Computer | Not_Computer
Z-1132-2014 | 67444 | No | NPJ | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Restoris Multicompartmental Knee (MCK) System is an implant system designed to be used with MAKO's Robotic Arm Interactive Orthopedic System (RIO). It is composed of a unicompartmental implant system (Restoris MCK Uni) and a patellofemoral implant system (Restoris MCK PF). The Restoris MCK Uni is designed for use when load bearing ROM is expected to be less than or equal to 155 degrees. In Restoris MCK combination where multi-compartmental areas are being treated, the Restoris MCK components were designed with 3 mm of gap between the components to ensure that the components do not interfere. |  |  | January 09, 2014 | 2014 | Mako Surgical Corporation | The product may be mislabeled and could result in the incorrect implant being used. | Error in labeling | MAKO Surgical Corporation sent an Important Medical Device Recall Notice letter dated January 9, 2014 to all affected customers. The letter identified the product, the problem, and the actions to be taken by the customer. The letter informed the customers that their MAKOplasty Sales Specialist will coordinate with them the return of the product that will be replaced at no charge. Customers were asked to sign and return the Acknowledgement form to MAKO Surgical using one of the following methods: Email, Fax or Mail. For questions, customers were instructed to call (954) 628-1721. For questions regarding this recall call 954-628-0607. |  | 41 | 41 units | Nationwide distribution including AZ, FL, HI, OH, MI, NC NH, WA, TX, WI, OK, CA, RI, LA, ID, IL, IN, and TN. | Terminated | August 14, 2014 | 217 | 41 | Not_Computer | Not_Computer
Z-2018-2014 | 68244 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | TerumoÂż Advanced Perfusion System 1 |  |  | June 17, 2014 | 2014 | Terumo Cardiovascular Systems Corporation | An evaluation of the APS1 Operator Manual found that existing instructions lack clarity and accuracy related to specific items. | Labeling Change Control | On June 17, 2014, TCVS sent URGENT MEDICAL DEVICE RECALL notification to consignees. Notification included reason for recall, potential hazard, correction, affected population, and customer instructions. Questions concerning the recall are directed to Terumo CVS Customer Service @ 800-521-2818 (M-F, 8am-6pm) |  | 1713 | 1713 units | Worldwide Distribution-USA (nationwide) including DC and the states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WI, and WV, and the countries of AUSTRALIA, BELGIUM, CANADA, CHILE, Hong Kong, Indonesia, Japan, Malaysia, Mexico, Philippines, Singapore, South Korea, Taiwan, Thailand, UNITED ARAB EMIRATES (UAE), and Vietnam. | Terminated | September 22, 2014 | 97 | 1713 | Not_Computer | Not_Computer
Z-0433-2015 | 69638 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | 8MM,MARYLAND BIPOLAR FORCEPS; |  |  | November 03, 2014 | 2014 | Intuitive Surgical, Inc. | Deviations in reprocessing steps from those stated in the reprocessing instructions can cause surface degradation of the housing and/or accelerate mechanical wear of the instrument. | Labeling design | Urgent Medical Device Recall letters were sent to affected sites on November 3, 2014. Letters inform customers of the reason for recall, risk to health and the affected instrument. A table of all of the affected IS4000 instruments affected is attached to the letter indicating the part number and product name for each. Letters also inform customers that all instruments shipped since July 2014 are manufactured with a new material and not subject to recall.Customers are to take the following actions:1. Ensure that all affected personnel are aware of the information.2. Identify affected product in their possession3. Contact ISI rep for replacement.4. Complete and return the attached Acknowledgement From.Until replacement instruments are received, the old instruments may be used provided they do not show signs of degradation and that the reprocessing instructions are followed. Questions to 800-876-1310 Option 3 ( 6 am to 5 pm PST) or ics@intusurg.com. |  | 59 | 5-9 units - total all devices | US Distribution -- FL, TX, NY, NV, OR, IN, UT, and MA. | Terminated | February 03, 2015 | 92 | 59 | Not_Computer | Not_Computer
Z-1611-2014 | 67932 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | DPM 6 Monitor in use with a Multi Parameter Module (MPM). Patient monitor used for monitoring, displaying, reviewing, storing, and alarming of multiple physiological parameters including ECG, arrhythmia detection, ST segment analysis, heart rate (HR), respiration rate (RESP), temperature, pulse oxygen saturation (SpO2), pulse rate (PR), non-invasive blood pressure (NIBP), invasive blood pressure (IBP), carbon dioxide (CO2), anesthetic gas (AG), impedance cardiograph (ICG), bispectral index (BIS), cardiac output (C.O.) and respiration mechanics (RM) |  |  | March 05, 2014 | 2014 | Mindray DS USA, Inc. dba Mindray North America | The MPM included a cardiac ST/Arrhythmia analysis feature which is not cleared for marketing. | No Marketing Application | Mindray DS Usa Inc. contacted their sole customer via phone on March 5, 2014. The affected products will receive software upgrade. |  | 4 | 4 units | US Nationwide Distribution in the state of OH | Terminated | April 27, 2015 | 418 | 4 | Not_Computer | Not_Computer
Z-2017-2014 | 68493 | No | BTT | Respiratory gas humidifier. | Anesthesiology | Anesthesiology | 510(k) | Humidifier Adaptor, 040 Shelf Pak, French, Teleflex Medical, respiratory gas humidifier adaptor. |  |  | June 05, 2014 | 2014 | Teleflex Medical | The seals of the adaptor packaging may be creased which may potentially affect packaging integrity. | Packaging change control | Consignees were notified by letter on 06/06/2014. |  | 7758650 | 7,758,650 ea. (total for all types listed) | Worldwide Distribution-USA (nationwide) and the countries of Belgium, Canada, China, Germany, Great Britain, France, India, Malaysia, Mexico, Philippines, Singapore, Thailand and Zambia. | Terminated | July 02, 2015 | 392 | 7758650 | Not_Computer | Not_Computer
Z-0100-2015 | 69226 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | GYRUS ACMI, REF 3345 -Everest BiCOAG MOLly Forceps 5mm/45cm |  |  | August 27, 2014 | 2014 | Gyrus Medical, Inc | Due to an anomaly in the packaging process, it is possible that a defective seal could be present that may allow a breach of the package's sterile barrier and may compromise the sterility of the product. | Packaging process control | Olympus sent a Medical Device Recall letter dated August 26, 2014, via FEDEX to all affected consignees. The letter described the issue and the product involved in the recall. It also described the action that the consignee needs to take. Consignees were instructed to cease further use of the device, quarantine it and return it to the manufacturer.Consigness were instructed to complete and return by fax (484-896-7128) the questionnaire enclosed. If the product was further distributed consignees should notify their customers of the product recall. For additional information they can call 484-896-5688.For questions regarding this recall call 484-896-5688. |  | 5560 | 5520 US, 40 OUS | Worldwide Distribution - USA including US: AL, AZ, CA, FL, GA, HI, IL, IN, IA, KY, MA,MI,MS, NJ, NY, NC, OH, PA, SC, TN, UT, and WA.Internationally to ITALY, SPAIN, UNITED KINGDOM, and CHINA. | Terminated | September 15, 2015 | 384 | 5560 | Not_Computer | Not_Computer
Z-0111-2015 | 69420 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Symmetric Total Knee Femoral Components |  |  | September 30, 2014 | 2014 | Signal Medical Corporation | Packaging of a device was compromised | Packaging | Signal Medical sent a Notice letter dated September 30, 2014 to customers. Customers were instructed to carefully inspect devices to ensure package integrity before using. If packaging is compromised it should not be used. Return the affected device and packaging to Signal Medical Corporation. Replacement devices with new packaging configuration is shipped to each location with the old packaging configuration in the event that the packaging is found to be compromised. For questions contact Signal Medical Corp. at 1-800-246-6324. |  | 41 | 41 | Worldwide Distribution - US (Nationwide) in the state of: MO and internationally to: Panama. | Terminated | March 14, 2016 | 531 | 41 | Not_Computer | Not_Computer
Z-1946-2014 | 68592 | No | MQK | Assisted reproduction labware. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Name: Pasteur Pipets, used by the embryologist in the IVF Laboratory. |  |  | June 20, 2014 | 2014 | ORIGIO, Inc | Identification of packaging breach of the nylon pouch sterilebarrier for the 3-pack configuration | Package design/selection | Field Safety Corrective Action / Urgent Product Recall letter dated June, 2014. |  | 134086680 | Total number of Boxes Distributed 32,760. US 4,093 / Canada 1,156 / EU &amp; Other 27,511 | Worldwide Distribution: US Nationwide, Canada, Australia, Barbados, Chile, China, Colombia, Denmark, Ecuador, France, Germany, Great Britain, Greece, India, Italy, Japan, Jordan, Malaysia, Netherlands, Peru, Russia, Singapore, South Africa, Spain, Taiwan, Turkey, and United Arab Emirates. | Terminated | July 20, 2015 | 395 | 134086680 | Not_Computer | Not_Computer
Z-1129-2014 | 67421 | No | GEY | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Punch Hair Matic |  |  | February 18, 2013 | 2013 | MEDICAMAT S.A. | Medicamat is initiating a field correction due to misbranded marketing literature for the Punch Hair Matic SAFER. | Labeling Change Control | MEDICAMAT sent an Advisory Notice dated May 7, 2013, to the affected customer. The letter identified the product the problem and the action needed to be taken by the customer.The notification advises customers of the issue (misbranded brochure and website). The letter advises the distributor to delete the misbranded language from their communication, marketing, or other information provided to clients. They are also advised to ask their clients change their marketing if necessary and provided corrective labeling.For further questions please call +33146562404.. |  | 71 | 71 units | US Distribution including TX. | Terminated | March 10, 2014 | 385 | 71 | Not_Computer | Not_Computer
Z-1133-2014 | 67475 | No | KZP | Bracket adhesive resin and tooth conditioner. | Dental | Dental | 510(k) | BencoDental tartar &amp; stain remover ultrasonic cleaning solution |  |  | February 10, 2014 | 2014 | Heraeus Kulzer, LLC. | As the result of a complaint to Heraeus Kulzer's customer service department, Heraues Kulzer is recalling Benco Tartar and Stain Solution part #1010-451, lot #13121434 because it has incorrect labeling on the immediate container. While the outer carton correctly states "Benco Tartar and Stain", the immediate gallon container in which the solution is contained incorrectly states "Benco Temporary C | Labeling mix-ups | Heraeus Kulzer, notified their one consignee of the recall February 4, 2014 via phone and then by official electronic notification. On February 7, 2014, an Important Communication Product Recall letter was also mailed to Benco direct accounts. On 2/10/14, Heraeus Kulzer also called Benco customers notifying them of the recall. The notifications explained the issue, identified the affected product and actions to be taken. For questions call 574-293-5444 |  | 812 | 812 units (gallons) | USA Nationwide Distribution in the states AL, AR, AZ, CA, CO, FL, GA, IL, IN, KS, KY, LA, MI, MN, MO, MS, NC, ND, NE, NM, NV, OH, OK, PA, SC, TN, TX, UT, VA, WA, WI, WV | Terminated | June 18, 2014 | 128 | 812 | Not_Computer | Not_Computer
Z-0093-2015 | 69380 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | McKesson Cardiology Hemo. |  |  | June 25, 2014 | 2014 | Mckesson Information Solutions LLC | McKesson Cardiology Hemo front-end parameter boxes must undergo annualmaintenance testing to establish the proper functioning of the parameter boxes,and to identify in a timely manner whether a parameter box is out of calibration.McKesson has identified that its communication to customers regarding therequired maintenance tests was insufficient. While each unit has an attachedsticker i | Labeling design | Consignees were notified via Field Safety Notice on 06/25/2014. |  | 625 | 625 devices | Worldwide Distribution-USA (nationwide) and the countries of Australia, Botswana, Canada, Germany, New Zealand, Netherlands, Pakistan, Portugal, South Africa, Turkey and The United Kingdom. | Open, Classified |  |  | 625 | Not_Computer | Computer
Z-2034-2014 | 68390 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer M/L Taper Reduced Neck Standard &amp; Extended Offset, HIP JOINT PROSTHESIS, Sizes:4, 5, 6, 7.5, 9, 10, 11 &amp; 12.5. |  |  | May 23, 2014 | 2014 | Zimmer, Inc. | Black residue observed in the sterile pouch of the M/L Taper and Kinectiv stems. | Packaging | Zimmer sent an Medical Device Recall Letter on June 4, 2014, to their Distributors. On June 9, 2014, Zimmer sent Medical Device Recall Letters to the physicians. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed:1. Locate all products from the affected lots and quarantine them immediately, if not previously done so per the May 23, 2014 Notice of Suspension of Loaner Pool. 2. Confirm the physical count of all affected product in loaner pool in your territory and complete the Confirmation of Loaner Pool Inventory and Return Certification Form (Attachment 2), and email a completed copy of Attachment 2 to corporatequality.postmarket@zimmer.com.3. Return the recalled product along with the completed Confirmation of Loaner Pool Inventory and Return Certification Form (Attachment 2). 4. Identify additional customers that may have used this product in the past 12 months by completing the spreadsheet template provided to you via email or confirm that Zimmer has identified all current customers in your territory.5. If after reviewing this notification you have further questions or concerns please call the customer call center at 1-877-946-2761 between 8:00 am and 5:00pm EST. |  | 852 | 852 units | Worldwide Distribution - USA (nationwide) and the countries of Australia, Canada, China, Europe, India, Japan, Korea and Singapore. | Terminated | March 06, 2015 | 287 | 2654 | Not_Computer | Not_Computer
Z-0865-2014 | 64795 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | The Centricity Laboratory System is intended to be an information system designed to support the clinical and administrative activities associated with the provision and utilization of clinical laboratory services and facilities, e.g., the storing and delivering of analytical results. It is a specially designed data program application (software), which is supplied for installation in existing mainframe or decentralized computers or a computer network. This software system will not be the software directly used to run, steer, or control any specific laboratory analyzer or equipment. It is intended for use in laboratory facilities, including a central laboratory or in a multiple laboratory environment servicing satellite laboratories, reference laboratories, clinics, etc. Centricity Laboratory System is not intended for direct patient contact. |  |  | November 20, 2012 | 2012 | GE Healthcare It | GE Healthcare is aware of a potential safety issue associated with the use of GE Centricity Laboratory Core Lab calculated results in two scenarios. In Scenario 1: Calculated results greater than 6 digits plus a decimal are reported as 0 instead of TOO BIG when significant figures are not defined for the item. In Scenario 2: Calculated results greater than 6 digits plus a decimal are trunc | Software design | An URGENT MEDICAL DEVICE CORRECTION letter date March 12, 2013 was sent to customers. The letter requests that customers review all calculated results to determine if either of the scenarios could occur at their site. If a customer determines that either of the scenarios proposed in the letter could occur at their site, the customer can contact their GE service representative to request or install the software update. Customers with questions can contact the GE Healthcare IT Helpdesk at 1-847-589-8496 or tgg.nahelpdesk@med.ge.com. |  | 7 | 7 units | Worldwide Distribution - US (Nationwide) to states of: MN, GA, and CA; and International to: England. | Open, Classified |  |  | 7 | Computer | Computer
Z-0810-2014 | 66996 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Medrad Veris MR Monitor units |  |  | November 21, 2013 | 2013 | Medrad Mr Inc | The main board, P/N 301641, installed in some Medrad Veris MR Monitor units may be faulty and could lead to unexpected shutdown of the system while in use, resulting in the loss of vital signs information from the monitor. | Device Design | The firm, Bayer HealthCare, sent a "MEDICAL DEVICE FIELD CORRECTIVE ACTION" letter dated November 21, 2013 via certified mail to its users/customers. The letter described the product, problem and actions to be taken. The firm stated that, "You may choose to continue to use your Veris units but you MUST be vigilant about this potential issue. Should a shutdown occur, the system can be quickly restarted by pressing the POWER button. If thisproblem does occur, we ask that you report this immediately to Bayer Customer Support at 1-877-229-3767." The customers were instructed to verify the serial number on your Medrad Veris unit(s). If this problem does occur, report this immediately to Bayer Customer Support at 1-877-229-3767. The customers were also instructed to fill out and fax back the attached form. Bayer Service Representative will then be in touch to schedule an appointment to come and replace your main board, at no cost to you.Bayer also informed their customers that based on the maximum inventory that can be produced by the supplier, they anticipate that this Field Corrective Action will take up to six months.If you have questions, please contact our Customer Support team at 1-877-229-3767. |  | 922 | 922 units | Worldwide Distribution: US (nationwide) and Internationally to: Canada. | Terminated | March 30, 2015 | 494 | 922 | Computer | Computer
Z-0085-2015 | 69181 | No | KPE | I.V. container. | General Hospital | General Hospital | 510(k) | Baxter 2B8013 Empty INTRA VIA Container with PVC Ports (Sterile, nonpyrogenic fluid path) ,500 ml. Admixture Products and Accessories |  |  | September 16, 2014 | 2014 | Baxter Healthcare Corp | Particulate matter found inside the fluid path. | Equipment maintenance | Baxter issued a letter to their consignees stating the following:Baxter is requesting that you take the following actions:Action to be taken if product was purchased directly from Baxter.1. Locate and remove all affected product from all locations in your facility(the product code and lot numbers can be found on the individualproduct package or shipping carton).2. Contact Baxter Healthcare Center for Service to arrange for return andcredit. The Center for Service can be reached at 888-229-0001between the hours of 7:00 am and 6:00 pm Central Time. Please haveyour Baxter eight-digit ship-to account number ready when calling.3. Complete the enclosed customer reply form , and return it to Baxter byeither tax or scanned e-mail.4. If you distributed any of the affected products to other facilities, pleaseconduct a recall with your end-user customers in accordance with yourcustomary procedures. Baxter distributed product 2B8011 lotU R 13D 15112 to customers between 4/26/2013 and 6/20/2013. Baxterdistributed product 2B8013 lot UR13K1 4095 to customers between 11 /27/2013 and 3/ 10/2014.Action to be taken if product was received from an alternate source.1. Locate and remove all of the affected products from all locations inyour facility (the product code and lot numbers can be found onthe individual product package or shipping carton).2. Contact Baxter Healthcare Center for Service to arrange for returnand credit. The Center for Service can be reached at 888-229-0001 between the hours of 7:00 am and 6:00 pm Central Time.Please have your ship-to account number ready when calling.3. Follow your supplier's reply and recall process. Please do not return the customer reply form to Baxter. |  | 20448 | 20448 | Nationwide Distribution including Puerto Rico, and the country of Canada. | Terminated | June 21, 2017 | 1009 | 114672 | Computer | Not_Computer
Z-0089-2015 | 69262 | No | JXG | Central nervous system fluid shunt and components. | Neurology | Neurology | 510(k) | Codman EDS 3 Drainage System with Ventricular Catheter (External Drainage System) |  |  | September 15, 2014 | 2014 | Codman &amp; Shurtleff, Inc. | tubing within the system that drains CSF may leak or disconnect from the joints and may result in over- or under-drainage of CSF from the ventricular system or introduction of air into the ventricular system (pneumocephalus) | Device Design | Codman Neuro sent an Urgent Notice dated September 12, 2014, to all affected customers via overnight mailing and OUS customers notified via email the week of September 15, 2014.Accounts were requested to isolate all inventory of the affected product and return to Stericycle at the following address:Stericycle 2670 Executive Drive, Suite A Attn: Event #8894 Indianapolis, IN 46241For patients currently being managed with the EDS 3 System, the system should be replaced immediately. In the case where no substitute drainage system is immediately available, the EDS 3 System may continue to be used until an alternative product can be obtained.If you have any questions or concerns regarding this notification, please contact your local Codman Neuro Representative or Scientific and Medical Affairs at SciMedAffairs@its.jnj.com or (866) 685-7325.For questions regarding this recall call 508-880-8000. |  | 66120 | 66,120 units | Worldwide Distribution - USA (nationwide) and Internationally to:United Arab EmiratesArgentinaAustriaAustraliaArubaBelgiumBahrainBoliviaBrazilBahamasCanadaSwitzerlandChilePeople's Republic of ChinaColumbiaCosta RicaCyprusCzech RepublicGermanyAlgeriaEstoniaEgyptSpainFinlandFranceGabonGreat BritainFrench GuianaGreeceGuatemalaHong KongCroatiaHungaryIrelandIsraelIndiaIranItalyJordanKenyaKuwaitKazakhstanLebanonLithuaniaLuxembourgLatviaMauritiusMexicoMalaysiaMozambiqueNew CaledoniaNigeriaNetherlandsNorwayNew ZealandOmanPanamaPeruPolandPalestinePortugalReunionRomaniaSerbiaRussian FederationRwandaSaudi ArabiaSwedenSingaporeSloveniaSlovakiaTunisiaTurkeyTrinadad And TobagoTaiwanUkraineUruguayVenezuelaMayotteSouth Africa | Terminated | August 11, 2017 | 1061 | 160823 | Computer | Not_Computer
Z-0733-2014 | 66202 | No | CJY | Calcium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Consult DiagnosticsÂż Eon Calcium Reagent For In-Vitro Diagnostic Use |  |  | August 29, 2013 | 2013 | Vital Diagnostics, Inc. | High positive calcium bias on plasma sample versus results on serum | Device Design | Vital Diagnostics issued a Recall Notice dated August 29, 2013, to all affected customers via certified mail. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed not to use plasma for the determination of total calcium when using the affected product. The use of serum is acceptable, thus the product can still be used to determine total calcium with this sample type. Vital Diagnostics released a corection to the package insert for the Envoy Calcium Reagent. Customers with the affected product in their laboratory, were instructed to discard the insert packed with the product and replace it with the insert enclosed with this bulletin. If product was transferred to another location, the bulletin should be forwarded to that location. The package insert was revised indicating to only use Serum samples when running Calcium. Vital Diagnostics issued a revised (expanded) notice dated September 19, 2013, to include the Consult DiagnosticsÂż Eon Calcium Reagent (PN: 1429) requesting users to discontinue using the single vial Eon Calcium kit (PN: 1429). Customers with questions were instructed to call Vital Diagnostics Technical Support at 1-855-354-8324..For questions regarding this recall call 401-642-8400. |  | 213 | 213 units | Nationwide Distribution | Terminated | August 26, 2015 | 727 | 608 | Computer | Not_Computer
Z-0045-2015 | 69257 | Yes | DRO | External transcutaneous cardiac pacemaker (noninvasive). | Cardiovascular | Cardiovascular | 510(k) | Covidien Kendall Adult Multi-Function Defibrillation Electrodes |  |  | September 18, 2014 | 2014 | Covidien LLC | Covidien Medi-Trace Cadence and Kendall Defibrillation Electrodes Not Compatible for use with Philips model FR3 or FRx AEDs | Device Design | Covidien notified customers by letter via Federal Express on 9/18/14. The letter requests that they immediately verify that they are not stocking Covidien electrodes for use with Philips FR3 or FRx AED units. A customer reply acknowledgement form is provided to allow Covidien to track the effectiveness of the notification. Covidien is adding a label to all levels of packaging that alerts the user that the product is not compatible with the FR3 and FRx AED units. Distribution partners notified to address product in the distribution chain that is available to be relabeled. For questions customers should call (800) 962-9888, option 8, then extension 2500.For questions regarding this recall call 508-261-8000. |  | 31590 | 31,590 pairs | Worldwide Distribtion-USA (nationwide) including Puerto Rico and the countries of Canada Greece Israel Belgium Finland France Germany Great Britain Ireland Italy Norway Poland Portugal Spain Sweden Switzerland Brazil Chile Philippines Taiwan Australia New Zealand. | Terminated | May 24, 2017 | 979 | 644460 | Computer | Not_Computer
Z-0633-2014 | 66919 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | The Propaq LT Series monitors. Model numbers 802LTAN, 802LT0N, and 802LTRN. |  |  | November 21, 2013 | 2013 | Welch Allyn Protocol, Inc | Welch Allyn will update Propaq 802 Series Vital Signs Monitor, models are 802LTAN, 802LTON and 802LTR with software because the fault conditions relating to hardware or software failure or electrical noise caused by proximal defibrillation activity could, in rare instances, potentially interrupt device operations. | Device Design | Welch Allyn began sending out the Urgent Medical Device Correction letter, dated December 03, 2013, to consignees on December 10, 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customer Actions: In the unlikely event that operation of your Propaq LT is interrupted and the display becomes blank or white, pressing the power (on/off) button will restart the device and quickly bring it back to full operational capacity.This notice should be passed on to all those who need to be aware within your organization, or to any organization where the potentially affected devices have been transferred.Welch Allyn is in the process of developing a software installation service tool that will enable installation of the software update by your biomed, IT staff or other technically qualified personnel.Upon release, the service tool will be made available to you via a download from Welch Allyn's website at www.welchallyn.com/PropaqltUpgrade or other media.In the meantime and on an on-going basis, the software update will be included if you schedule routine or other service with Welch Allyn or an authorized service provider.Customers with questions were instructed to call Welch Allyn Technical Support at 1-800-535-6663.For questions regarding this recall call 503-530-7935. |  | 6003 | 6003 units total (1715 units in the US and 4288 outside the US) | Worldwide Distribution - USA (nationwide) and Internationally to United Arab Emirates, Argentina, Australia, Belgium, Brazil, Canada, Switzerland, Chile, China, Columbia, Costa Rica, Czech Republic, Germany, Denmark, Ecuador, Spain, France, United Kingdom, Greece, Hong Kong,Taiwan, Ireland, Italy, Japan, Korea, Republic of, Kuwait, Lithuania, Luxemborg, Mexico, Malaysia, Mozambique, Namibia, Netherlands, Poland, Portugal, Qatar, Russia, Saudi Arabia, Singapore, Sierra Leone, El Salvador, Thailand, Turkey, Uganda, Vietnam, Kosovo, and South Africa. | Terminated | February 17, 2015 | 453 | 6003 | Computer | Computer
Z-2569-2014 | 68704 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Artis zee and Artis zeego systems. |  |  | June 05, 2014 | 2014 | Siemens Medical Solutions USA, Inc | There is a potential problem with Artis zee and Artis zeego systems running software version VC21B and being used in conjunction with the Large Display, in that under certain circumstances, the release of radiation can become blocked unnecessarily. | Software design | Siemens sent an "Important Customer Safety Notice" dated August 21, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Please forward this Safety information to other organizations that could be affected by this measure. If the device has been sold and therefore no longer in your possession, please forward this safety notice to the new owner. We would also request that you inform us of the identity of the device's new owner where possible.Please observe this Customer Safety Notice and comply with the corresponding measures until the update has been fully completed. For further questions please call (610) 219-6300. |  | 206 | 206 | Nationwide Distribution | Terminated | April 21, 2015 | 320 | 206 | Computer | Computer
Z-0164-2015 | 69430 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens Calibrator E; SMN 10309079 (2 pack), SMN 10321075 (6 pack), SMN 10335532 (6 pack Ref) |  |  | September 26, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc | Positive bias for Calibrator E lot kits ending in 38 and 39 with the ADVIA Centaur Systems Cortisol and Progesterone assays compared to the respective master curves. | Device Design | An Urgent Medical Device Correction (UMDC) was sent to all affected Siemens Healthcare Diagnostics customers in the United States on September 26, 2014. An Urgent Field Safety Notice (UFSN) was sent out for distribution to all customers outside the United States on September 26, 2014. These notices inform the customer of the positive bias with ADVIA Centaur Systems Cortisol and Progesterone when calibrated with Calibrator E kit lots ending in 38 and 39.Customers may continue use of the affected lots as the clinical utility of the assay is not impacted. The communications also provide information on the performance of the assays when they use Calibrator E kit lots ending in 40. The notices should be reviewed with the facility's laboratory or Medical director.All affected domestic customers were sent a hard copy of the Urgent Medical Device Correction via Federal Express. Fax-back forms will be used to confirm that the customers received the UMDC. For those customers that do not return a fax-back form, a Siemens representative will follow up with those customers to ensure they received the Urgent Medical Device Correction and completed the required actions.All affected customers outside the US will be provided a copy of the Urgent Field Safety Notice via hard copy, e-mail, and/or fax as determined by each country's local regulations and procedures. Each customer will be required to return a completed confirmation fax-back to Siemens. For those customers that do not return a fax-back form, a Siemens representative will follow up with those customers to ensure they received the Urgent Field Safety Notice and completed the required actions. |  | 6 | Total = 19,352 (13,881 - 2 pack, 4878 - 6 pack, 593 - 6 pack REF) | Nationwide including Puerto Rico. Worldwide: Afghanistan, Armenia, Austria, Bahrain, Belgium, Bulgaria, Burkina Faso, Canary Islands, Croatia, Cyprus, Czech Republic, Denmark, Estonia, Finland, France, Georgia, Germany, Great Britain, Greece, Hungary, Iran, Iraq, Ireland, Italy, Jordan, Kazakhstan, Latvia, Lebanon, Libya, Lithuania, Luxembourg, Macedonia, Malta, Morocco, Netherlands, Norway, Oman, Pakistan, Poland, Portugal, Qatar, Romania, Russian Fed., Saudi Arabia, Serbia, Slovakia, Spain, Sweden, Switzerland, Turkey, Unit.Arab Emir., Vatican, Yemen, AR, AU, BD, BR, CA , CH, CL, CN, CO, DO, EC, EG, GT, HK, ID, IE, IL, IN, JP, KR, LY, MM, MX, MY, NI, NP, NZ, PE, PH, PY, SA, SG, SV, TH, TW, UY, VN, ZA | Terminated | January 29, 2016 | 490 | 6 | Computer | Not_Computer
Z-2644-2014 | 69032 | No | NHT | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(k) | JBAIDS Instrument System Software version 3.5.0.72. |  |  | August 13, 2014 | 2014 | Biofire Defense | BioFire Defense, LLC is initiating a field action for JBAIDS Software version 3.5.0.72 because the wrong version software may have been installed. | Software Manufacturing/Software Deployment | Consignees were notified on 8/13/14 via letter and telephone and were provider with instructions on how to install the correct software. |  | 43 | 43 Units w/ laptops | Distributed USA (nationwide) and Guam. | Terminated | October 24, 2016 | 803 | 43 | Computer | Computer
Z-0748-2014 | 65366 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge Mammo is a multi-modality, vendor-neutral digital mammography workstation that enables imaging centers to use a single workstation to display and read images from different vendors and different acquisition devices, including digital mammography, ultrasound, MRI and CT. |  |  | August 26, 2013 | 2013 | GE Healthcare It | There is a potential safety issue with the 8.0.2 version software of the Merge (Cedara) Mammo viewer. The Merge Mammo viewer can load and display images not in synch with the selected or highlighted images in the thumbnail navigator. This incorrect image display issue could result in an inaccurate reading on the Merge Mammo workstation. This could result in a missed intervention or mis-diagnosis. | Software design | GE Healthcare sent an URGENT MEDICAL DEVICE CORRECTION letter dated August 26, 2013, to all affected customers. The letter includes important safety instructions for healthcare providers to refer to that can minimize the chance that an incorrect image will be used in the diagnosis or treatment of a patient. The letter further states that a GE Healthcare service representative will contact the customer to schedule a time to install a software update to version 8.0.2 that will correct the software defect.. Customers with any concerns about this defect or recall are encouraged to contact their local GEHC-IT service representative or the GE Remote Online Center (ROC) Customer Care Center at 1-800-437-1171.For questions regarding this recall call 800-437-1171. |  | 28 | 28 units | Worldwide Distribution - USA including AK, AR, CA, IA, ID, IL, MI, MS, MT, NJ, NV, NY, OR, PA, SC, TN, WI and Internationally to Israel, Mexico, and Spain. | Open, Classified |  |  | 28 | Computer | Computer
Z-0473-2015 | 69793 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | IntelliSpace Portal DX/HX/EX model number: 881001 |  |  | November 09, 2014 | 2014 | Philips Medical Systems (Cleveland) Inc | Software defect. In certain circumstances, the application may display incorrect measurements of cardiac and aortic anatomy used to select and size the Transcatheter heart valve (THV) to be implanted . | Software design | Consignees were notified of the Field Safety Notice via certified mail. |  | 32 | 32 | Worldwide Distribution: US Distribution in states of: FL, NC, and VT; and the countries of Argentina, Austria, Belgium, Canada, Denmark, France, India, Italy, Japan, Netherlands, United Kingdom, and Vietnam. | Terminated | September 14, 2015 | 309 | 32 | Computer | Computer
Z-0731-2014 | 64724 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Centricity PACS System versions 3.X and higher and 4.X and higher |  |  | September 15, 2012 | 2012 | Ge Healthcare It | There is a potential loss of image(s) associated with the Centricity to Centricity (C2C) exam transfer module of Centricity PACS System versions 3.X and higher and 4.X and higher when transferring from the source to a destination server. When another process in the destination server attempts to access the same object or table, the transfer process of a particular image may be terminated. Once ter | Software design | GE Healthcare sent an Important Product Information letter dated September 15, 2012 to affected customers. The letter identified the affected product problem and actions to be taken. GE Healthcare will provide a patch to 3.2.X and 4.0 Systems to remediate the issue. For questions contact your local GEHC-IT service representative. |  | 748 | 748 units | Worldwide Distribution - US: Nationwide (AK, AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WV, WY); FOREIGN: Australia, Austria, Belgium, Bermuda, China, Denmark, Egypt, France, Germany, Hong Kong, India, Ireland, Israel, Italy, Japan, Korea, Kuwait, Malaysia, Malta, Netherlands, Portugal, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, United Kingdom. | Open, Classified |  |  | 748 | Computer | Computer
Z-0791-2014 | 66311 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance iCT, Brilliance 64 and Brilliance Big Bore |  |  | March 07, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | The DoseRight feature suggest a mAs based on the measured patient size, a reference size and a reference mAs. When scanning large children, the suggested mAs may be higher than clinicians would expect. | Software design | Philips Medical Systems sent a Field Safety Notice letter dated February 25, 2011, to all affected customers.The letter identified the product the problem and the action needed to be taken by the customers.Your Corrective Action Plan (CAP) is approved and consists of:1. Notifying all affected users with Field Safety Notice 72800527/72800528 starting March 8, 2011. This Field Safety Notice has been reviewed and is approved.2. Following up with all responses by dispatching a Field Service engineer to each site to install the software update at no cost to the user.3. Tracking the responses from users by certified mail.For further questions please call (440) 483-7600. |  | 23 | 23 | Worldwide Distribution - US Distribution including the states of DC, IL, IN, MI, MO, NY, OH, OR, PA, TX and VT., and the countries of Australia, Belgium, China, Germany, India, Israel, Italy and The Netherlands. | Terminated | April 28, 2014 | 1148 | 23 | Computer | Not_Computer
Z-0165-2015 | 69424 | No | HAB | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | UNIVERSAL Modular Electric/Battery Handpiece for surgery |  |  | October 06, 2014 | 2014 | Zimmer Surgical Inc | Non-injury complaints from Europe identified a malfunction that made it possible for a handpieceto start by itself when the power source is connected. | Device Design | Zimmer sent an Urgent Medical Device Correction letter via email to US distributors with traceable Fed Ex mailings to all consignees dated October, 2014. The letter identified the affected product, problem and actions to be taken. Customers were advised of the potential malfunction and the operating procedure for the handpiece per the Instructions for Use (IFU). Zimmer Surgical, as the importer, will provide notice to consignees in the US accordingly. In addition, units will receive an upgrade as part of the annual preventive maintenance per the Instructions for Use. |  | 828 | 828 | US Nationwide Distribution in the states of: AK, AL,AR, AZ,CA,CO,CT, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA,ME, MI, MN, MO, NC, NE, NJ, NY, OH, OR, PA, RI, SC, TN, TX, VA, WA, WI, WV, and WY. | Terminated | April 15, 2015 | 191 | 828 | Computer | Computer
Z-0211-2015 | 69557 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Cardiohelp-I System |  |  | October 17, 2014 | 2014 | Maquet Cardiovascular Us Sales, Llc | Cardiohelp- i System may have a software issue that can potentially result in an erroneous display of a &quot;Battery Needs Service&quot; message after startup of the units when using either AC or DC power | Software Design Change | Maquet Inc. sent a Product Recall Urgent Medicall Device Field Correction letter to all of their customers via Federal Express. The letter identified the affected product, problem and actions to be taken. The letter stated that a software upgrade will be performed by an authorized Maquet Service Technician. A Maquet representative will contact customer to schedule an on-site service of the device. Any questions can be directed to the local Maquet representative or Maquet Customer Service at 1-888-627-8383 (option 2, option 2) Monday through Friday, between the hours of 8:00 am and 6:00 pm (EST). |  | 331 | 331 units | Worldwide Distribution - US Nationwide | Terminated | April 06, 2016 | 537 | 331 | Computer | Computer
Z-0861-2014 | 65179 | No | JXG | Central nervous system fluid shunt and components. | Neurology | Neurology | 510(k) | Codman Certas Programmable Valve In Line Valve only with SiphonGuard Device |  |  | May 16, 2013 | 2013 | Codman &amp; Shurtleff, Inc. | Codman Certas Programmable Valves used for hydrocephalus may not operate properly | Device Design | Codman issued an Urgent Notice dated May 16, 2013, to all affected customers. The letter identifed the affected product, problem and actions to be taken. Mitigation instructions are provided and if not followed or are not successful, either issue could lead to symptoms associated with under- or over-drainage. Units not implanted are to be returned. A U.S Acknowledgement Form and fax are requested to be completed and faxed to: 1-888-943-4193, or e-mail the form to codman8894@stericycle.com. If you have any questions on concerns in regards to this recall, please contact your Codman Representative or Stericycle at 1-888-943-2395. |  | 1481 | 1481 units | Worldwide Distribution - US Nationwide, Australia, Austria, Belgium, Brazil, Switzerland, Czech Republic, Germany, Spain, Finland, France, Great Britain, Israel, Italy, Netherlands, Norway, Portugal, Russia, Saudi Arabia, Sweden, Slovenia, Slovakia, Turkey, South Africa, and UK. | Terminated | June 02, 2015 | 747 | 5791 | Computer | Not_Computer
Z-0853-2014 | 67325 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Innova IGS 630. Biplane Cardiovascular and Interventional Imaging System. |  |  | November 15, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue with respect to IGS 630 Imaging Systems.While performing fluoroscopy on a biplane fluoroscopy unit, there is a potential for loss of the x-ray imaging function when the user changes the size of the lateral FOV (Field of View) and releases the biplane footswitch pedal simultaneously. This could lead to a system lock up requiring | Software design | GE Healthcare sent an "Urgent Medical Device Correction" letter dated November 15. 2013. The letter was addressed to Hospital Administrators / Risk Mangers, Mangers of Radiology / Cardiology / Surgery Department and Radiologists / Cardiologists / Surgeons. The letter described the Safety Issue, Safety Instructions, Affected Product Details, and Product Correction. The letter stated that GE Healthcare will correct all affected systems at no cost. For questions call (In US 800-437-1171; for other countries contact your local GE Healthcare Service Representative). |  | 28 | 28 (12 US, 16 OUS). | Worldwide Distribution - US Nationwide in the states of CA, CO, FL, NV, NJ, PA, TX and the countries of CZECH REPUBLIC, FRANCE, JAPAN, REPUBLIC OF KOREA, SAUDI ARABIA, and TURKEY. | Terminated | March 28, 2014 | 133 | 28 | Computer | Computer
Z-0854-2014 | 66982 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare CARESCAPE Monitor B650 and CARESCAPE Monitor B850. |  |  | December 27, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of potential safety issues associated with the CARESCAPE Monitor B850 and CARESCAPE Monitor B650. These included 25 issues related to NIBP (Non-InvasiVentilation Alarms:ve Blood Pressure) , ECG, Central Monitor, General, Tram, and Bed to Bed Issues. | Software design | GE Healthcare sent an Urgent Medical Device Correction letter between December 27, 2013 and June 27, 2013, to all affected customers. The letter described the Safety Issues and Instructions, Affected Product Details, Product Correction and Contact Information. Customers with questions were instructed to contact Technical Support at 1-800-558-7044 or their local Service Representative.For questions regarding this recall call 800-558-7044. |  | 40239 | 40,239 (17,182 units US, 23,147 units OUS). | Nationwide Distribution including DC and PR, except DE, SD, ND, and VT. OUS: Algeria, Argentina, Australia, Austria, Azerbaijan, Bahrain, Bangladesh, Belgium, Bermuda, Bolivia,Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic ,Denmark, Ecuador, Egypt, El Salvador, Estonia ,Finland, France, Germany, Greece,Guam, Guatemala, Honduras, Hong Kong, Hungary, Iceland, India, Indonesia, Iraq, Ireland, Israel, Italy, Jamaica, Japan, Jordan, Kazakhstan, Kuwait,Latvia, Lebanon, Libyan Arab Jamahiriya, Lithuania, Malaysia, Malta, Mexico, Morocco,Namibia, Nepal, Netherlands, New Zealand, Nicaragua, Norway, Oman, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Puerto Rico, Qatar, Romania, Russian Federation, Saudi Arabia,Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Sweden, Switzerland, Syrian Arab Republic, Taiwan, Tanzania, United Republic of Tunisia, Turkey, United Arab Emirates, United Kingdom, Venezuela, Uruguay. | Open, Classified |  |  | 40239 | Computer | Computer
Z-0161-2015 | 58185 | Yes | DTR | Cardiopulmonary bypass heat exchanger. | Cardiovascular | Cardiovascular | 510(k) | Terumo Custom Cardioplegia Delivery Sets |  |  | March 01, 2011 | 2011 | Terumo Cardiovascular Systems Corp | Defect in the luer connector can compromise the seal on the female side of the connector and the compromised seal can introduce air in the table line | Device Design | Terumo Cardiovascualr notified accounts by letter dated 3/1/2011 advising users users that the luer connector to the table line in certain lots of myocardial protection sets may leak. Users should not continue to use affected myocardial protection sets in a manner that exposes the connector to negative pressure without taking correction or replacement activities as described. Firm provided mitigating instructions:Clamp the table line distal to the luer connector when ending an infusion of cardioplegia thus isolating the leak from negative pressure. Contact Terumo CVS Customer Service to report the defect and arrange return of the unit: 800-521-2818. |  | 24 | 24 units | Worldwide Distribution: US (nationwide) and the countries of Canada, Brazil, Hong Kong, Korea, and Thailand. | Terminated | November 06, 2014 | 1346 | 46190 | Computer | Not_Computer
Z-0094-2015 | 69301 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Continuum Acetabular System Trabecular Metal Shell with Multi Holes, Porous 68mm, Prosthesis, Hip, Semi-Constrained, Metal/Polymer |  |  | October 01, 2014 | 2014 | Zimmer, Inc. | The affected products are missing polar boss threads. | Equipment maintenance | Zimmer sent an "URGENT MEDICAL DEVICE RECALL- LOT SPECIFIC" notifications dated October 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Your Responsibilities1. Review the notification and ensure affected personnel are aware of the contents.2. If you find any product from the affected lot , quarantine the product and notify your Zimmer sales representative.3. Your Zimmer sales representative will remove the recalled product from your facility.4. For patients that previously had this product implanted, it is recommended that you continue your normal post operative follow up routine.5. If after reviewing this notification you have further questions or concerns please call the customer call center at1-800-348-2759 between 8:00 am and 5:00pm EST. |  | 19 | 19 | Worldwide Distribution: US (nationwide) in the states of : IA, CA, NC, AL, WA, FL, TN, MA, AZ, TX, UT, and NV; and Internationally to Japan. | Terminated | March 02, 2015 | 152 | 19 | Computer | Not_Computer
Z-0851-2014 | 66585 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | User Defined Method Flex Assignment/Siemens Dimension Vista 500 or Dimension Vista 1500 System with software version 3.5.1 or lower |  |  | October 07, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | When utilizing the Routine Inventory screen to enter a User Defined Method (EMPTY) Flex the system may assign the User Defined Method Flex to a different Flex that is currently in inventory on the system, and then use the incorrect Flex Cartridge to process the user defined method. | Software design | Siemens sent a Urgent Medical Device Correction letter dated October 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Please use the following actions if using Vista SW Versions 3.5.1 and below to prevent the issue from occurring:1. When assigning a User Defined (EMPTY) Flex, use the Advanced Inventory Screen per the Operators Guide.2. Verify that EMPTY Flex is no longer listed in the inventory before proceeding with any testing.3. If the EMPTY Flex is still listed in the inventory, contact your Siemens Customer Care Center - Technical Solutions for assistance.Please retain this letter with your laboratory records and forward this letter to those who may have received this product. Discuss the content of this letter with your Medical Director.We apologize for the inconvenience this situation has caused. If you have any questions, please contact your Siemens Customer Care Center - Technical Solutions at (800-441-9250) or your local Siemens technical support representative. |  | 2038 | 2038 | USA (nationwide) Distribution including Puerto Rico. | Terminated | August 27, 2014 | 324 | 2038 | Computer | Computer
Z-0152-2015 | 69425 | No | IWB | Radionuclide radiation therapy system. | Radiology | Radiology | 510(k) | Laksell Gamma Knife Perfexion |  |  | October 06, 2014 | 2014 | Elekta, Inc. | Elekta has become aware that the latches may be locked before they have been fully turned resulting in a poorly locked frame. | Device Design | Elekta sent an" Important Field Safety Notice" distributed on October 6, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. This Important Field Safety Notice describes the problem and possible work arounds for customers to follow until a permanent solution can be developed. They are:" Always assure that the plastic lever is only operated when at a right angle with the frame adapter." Do not force the plastic lever in place if it meets significant resistance when turned." Ensure that the plastic lever is completely flush with the frame adapter and that no angulations are present.Corrective Action #2 - Elekta will check and replace as necessary affected frame adapters in the field. This will be released at the end of October 2014, and Elekta will have six months to complete the checks and replacements for those products in the field. If you have any questions please contact your local Elekta office, or call (770) 300-9725. |  | 181 | 181 | Worldwide Distribution - USA (nationwide) including the states of AK, AZ, AR, CA, CO, CT, FL, GA, HI, IL, KY, LA, MA, MI, MN, MS, MO, MT, NJ, NM, NY, NC, OH, OR, PA, RI, SC, TX, TN, UT, VA, WA and WV, and the countries of Australia, Austria, Belgium, Brazil, Canada, China, Colombia, Czech Republic, Egypt, France, Germany, Greece, Hong Kong, India, Indonesia, Italy, Japan, Mexico, Morocco, Netherland, Norway, Poland, Portugal, Russia, Saudi Arabia, South Korea, Spain, Sweden, Switzerland, Taiwan, Turkey and United Kingdom. | Open, Classified |  |  | 181 | Computer | Not_Computer
Z-0677-2014 | 67166 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | 14" (36 cm ) Appx 0.50 ml, Smallbore Ext Set w/MicroClaveÂż Clear, |  |  | December 24, 2013 | 2013 | ICU Medical, Inc. | ICU Medical is recalling certain lots of the NanoClave Needlefree Connector product line because of a small percentage of devices that may leak. | Equipment maintenance | icumedical sent an Urgent Medical Device Recall letter dated December 24, 2013, to all affected customers. The letter informed the customers of the problems identified and the actions to be taken. Customers were instructed to return completed verification forms. Customers with questions were instructed to contact ICU Customer Service Monday through Friday between the hours of 8:30AM and 4:00PM Pacific Time: (866) 829-9025 and select option 8, or email: customerservice@icumed.com.ICU Medical reported an update on January 6, 2014, an expansion of the recall to additional lots. Updated recall letter was sent out to all customers on additional lots on January 6, 2014 via UPS with tracking.For questions regarding this recall call 949-366-2183. |  | 100 | 100 units | Worldwide Distribution - USA including UT, TX, GA, WI, VT, NH, AL, OH, TN, IL, CA, KY,HI, NC, and Internationally to CA, BE, IN, KW, and AU. | Terminated | April 01, 2014 | 98 | 147975 | Computer | Not_Computer
Z-0204-2015 | 69367 | walker, mechanical | ITJ | Mechanical walker. | Physical Medicine | Physical Medicine | 510(K) Exempt | Transport Rollator with Padded Seat; Large 8-inch wheels with wheel locks (indoor and outdoor use), fold down footrests, ergonomic handles, curved padded backrest and handy wire basket for storage. Item# 30192 Burgundy each/Item# 30194 Blue each.. |  |  | October 01, 2014 | 2014 | Roscoe Medical Inc | The caster may separate completely from the device frame causing injury to user. | Device Design | The firm, Roscoe Medical, issued a "URGENT RECALL NOTICE-Timely Response Required" Quarantine Communication to customers between 09/08/2014 and 09/10/2014. The notice identified affected products as 30194, serial number range MH140407092MH140407291, the problem and actions to be taken. The customers were instructed NOT to use this product and to destroy any affected product that may be in their inventory/control; immediately complete and return the attached Recall Response Form/Certificate of Destruction to Roscoe Medical- "Transport Rollator Recall" within 15 days of receipt of the recall notification via fax to (440) 572-4261 or email: corporate@roscoemedical.com, even if you do not have any affected product; and notify your customers immediately of this product recall.Should you have any questions, contact Roscoe Medical Customer Support at (800) 376-7263 Monday-Friday 8:00 am EST - 5:00 pm EST. |  | 831 | 831 units | Nationwide Distribution: US including states of: AZ, CA, CO, FL, GA, IA, ID, IL, IN, KY, LA&lt; MA, MD, MI, MN, MO, NC, NJ, NM, NV, NY, OH, PA, SC, TN, TX, VA, WA, WI,and WY. | Open, Classified |  |  | 831 | Computer | Not_Computer
Z-0714-2014 | 67056 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Draeger Infinity Acute Care System Monitoring Solution |  |  | December 10, 2013 | 2013 | Draeger Medical Systems, Inc. | After a test in which a low SpO2 (yellow) medium grade alarm initiated in neo-natal mode and was subsequently paused, the high grade life threatening SpO2 alarm (red) did not activate when the SpO2 values were decreased to reach the red alarm limit. | Software design | An Urgent Medical Device Recall Letter, dated December 2013, was sent to Users. Once available Drager will update components of the IACS Monitoring Solution System. |  | 439 | 439 | US Distribution including the states of: MA, NH, IN, CA, TX, NY, OK, and IL. | Terminated | May 20, 2015 | 526 | 439 | Computer | Computer
Z-0747-2014 | 67115 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare Carescape Patient Data Module |  |  | November 27, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue due to intermittent failure of the spring loadedmounting latch of PDMs produced from April 2007 through December 2012. | Device Design | The firm, GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated November 26, 2013 to its consignees/customers. The letter was addressed to Healthcare Administrator / Risk Manger, Chief of Nursing and Director of Biomedical Engineering. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. The consignees/customers were instructed to please ensure that all potential users in your facility are made aware of this safety notification and the recommended actions. GE Healthcare will provide a new latch at no charge once it is available. GE will contact you to arrange for installation of the latch. If you have any questions regarding this notification, please contact Technical Support at 1-800-558-7044, or your local Service Representative. |  | 43033 | 43,033 (25,681 USA; 17,352 OUS) | Worldwide Distribution: US (nationwide) including states of: AL, AZ, AR, CA, CO, CT, DC, FL, GA, ID, IL, IN, IA, KS,KY, LA, ME, MD, MA, MI, MN, MS, NO, NE, NV, NJ, NM, NY, NC, OH, OK,OR, PA, RI, SC, TN, TX, UT,VT, VA, WA, WI, and WY. Internationally to: Argentina, Australia, Austria, Bangladesh, Belgium, Brazil, Canada, Chile, China, Colombia, Costa Rica, Czech Republic, Denmark, Egypt,England, Estonia, Finland,France, Germany,Greece, Guatemala, Hungary,India, Indonesia, Ireland, Israel, Italy, Jamaica, Japan, Jordan, Korea, Kuwait, Lebanon, Lithuania, Malaysia, Mauritius, Mexico, Morocco, Nepal, Netherlands, New Zealand, Norway, Oman, Pakistan, Panama, Paraguay, Peru, Poland, Portugal, Qatar, Russia, Saudi Arabia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Republic of China, Thailand, Trinidad and Tobago, Tunisia, Turkey, United Arab Emirates and United Kingdom. | Open, Classified |  |  | 43033 | Computer | Not_Computer
Z-0900-2014 | 67298 | No | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | Atlantis PV, 8.5F, 15 MHz Peripheral Imaging Catheter: |  |  | January 06, 2014 | 2014 | Boston Scientific Corporation | Reports of physicians experiencing inability to pass a 0.035" guidewire through the guidewire lumen and exit port at the (proximal) Y-manifold of the imaging catheter. | Equipment maintenance | Boston Scientific sent an Urgent Medical Device Recall Removal letter dated January 14, 2014 to all affected customers by overnight mail. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete an Account Reply verification Tracking Form in order to received a Return Goods Authorization number for product return. For questions contact your local Sales Representative. |  | 1068 | 1068 single units | US Nationwide Distribution and Mexico | Terminated | October 29, 2014 | 296 | 1068 | Computer | Not_Computer
Z-0905-2014 | 66835 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Medtronic Advanced Energy Aquamantys3 Pump Generator; Product Catalog Number: 10-1357 (Generator assembly), 40-404-1(as shipped, including accessories) |  |  | November 12, 2013 | 2013 | Medtronic Advanced Energy, LLC | The internal protection circuitry on the electrosurgical output that prevents a patient from exposure to DC voltage when the device is activated is being compromised during normal use. | Device Design | Medtronic sent an Urgent Product Removal Notification on November 12, 2013, via Next Day FEDEX. The communication advises users to immediately stop using the affected AQM3 electrosurgical generators and quarantine them until they can be returned to Medtronic Advanced Energy. Medtronic field personnel will collect them in order to remove them from service and return them to Medtronic Advanced Energy. Customers with questions were instructed to contact Customer Service at 866-777-9400.For questions regarding this recall call 866-777-9400. |  | 242 | 242 devices | Nationwide Distribution including CA, MO, PA, NJ, NY, OH, VA, TX, NC, DE, KS, TN, LA, IL, MI, SC, WI, and FL. | Terminated | April 03, 2014 | 142 | 242 | Computer | Computer
Z-2653-2014 | 68909 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | VERO Linear Accelerator System, Model Number MHI-TM2000; Common Name: Medical Linear Accelerator. |  |  | July 05, 2014 | 2014 | MITSUBISHI HEAVY INDUSTRIES, LTD., HIROSHIMA MACHINERY WORK | Software Anomaly: If a user changes the calendar setting from Workday to Holiday or vice versa, that would erroneously change the status of treatment/fractions, which are completed and to become Treated (completed) status on the exact day when such change is made, to Untreated status. Furthermore, such treatment would be cloned and mistakenly added to the schedule as Untreated treatment | Software design | A Preliminary Customer Information Letter (dated 7/4/14) was emailed to customers on 7/05/14 from Mitsubishi's sales and service representative in Germany. The letter informed customers of the potential safety issue associated with the VERO Linear Accelerator and the recommended actions to be taken. A formal URGENT FIELD SAFETY NOTICE Letter (dated 7/15/14) was emailed to customers on 8/07/14 from Mitsubishi's sales and service representative in Germany. The letter instructed customers to ensure that all potential users in the facility are made aware of the safety notification and the recommended actions to be taken. |  | 3 | 3 units in use in the US | Distributed in the states of NY, FL, and TX. | Terminated | April 01, 2015 | 270 | 3 | Computer | Computer
Z-0463-2015 | 69542 | No | JJH | Clinical sample concentrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Abbott m2000sp, an automated system for performing sample preparation for nucleic acid testing. |  |  | October 23, 2014 | 2014 | Abbott Molecular | Some versions of Application Specifications (App Spec) are incompatible with m2000sp system software version 6.0 and 7.0. This issue will cause Error Code 9000 (An unexpected error has occurred: Software error) to be generated when the Available Reagent Packs screen is selected. | Software design (manufacturing process) | Field Correction Recall / Urgent Field Safety Notice letters dated September 23, 2014 were sent to all direct accounts (customers). The letters included instructions for customers to: 1) review the app spec files that are installed on the m2000sp system and delete any incompatible app specs; 2) if one or more app spec files were deleted, installation of new app spec file(s) is needed; and, 3) review this information with laboratory personnel and retain the letter for future reference. Customers with questions about this market action can contact Technical Support at 800-553-7042, Option 2, or the local Abbott Molecular representative. |  | 9 | 9 CD ROMs (US); 1,273 CD ROMs (Foreign) | Worldwide distribution. US nationwide (CA, CT, FL, NC, PA, SC, and DC), Algeria,Australia,Austria,Belgium,Bosnia and Herzegovina,Cameroon,Canada,Chad,Colombia,Costa Rica,Croatia,Denmark,Egypt,Ethiopia,Finland,France,Germany,Greece,Guinea,Honduras,Hong Kong,India,Indonesia,Ireland,Israel,Italy,Japan,Kenya,Korea,Lithuania,Mali,Mexico ,Netherlands,New Zealand,Nigeria,Norway,Pakistan,Paraguay,Poland,Portugal,Romania,Russia,Saudi Arabia,Senegal,Slovenia,South Africa,Spain,Sweden,Switzerland,Taiwan,Tanzania,Thailand,Turkey,Uganda,Ukraine,United Arab Emirates, United Kingdom,Vietnam | Terminated | March 08, 2016 | 502 | 9 | Computer | Computer
Z-0725-2014 | 67136 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | GE Optima XR220amx and Optima XR200amx with Digital Upgrade Mobile Radiographic Systems. Model numbers 5555000-5, 5555000-6, and 5421698. |  |  | November 07, 2013 | 2013 | GE Healthcare, LLC | There exists a software issue associated with patient selection from the worklist on the Optima XR200 with Digital Upgrade, and Optima XR220amx Mobile X-ray Systems. Intermittently when a large number of patient exams are retrieved from the HIS/RIS system, the exams when displayed on the mobile x-ray system are not properly displayed. If the user continues to attempt to select one of the exams w | Software design | GE sent an Urgent Medical Device Correction letter, dated 11/7/2013 to consignees. The letter described the issue, identified devices affected and stated that users should pay close attention to the patient exam/entry selected from the worklist prior to selecting the Start Exam button to ensure that the system did not update and select an incorrect exam. The user may also confirm during the exam that the appropriate patient information was selected via the display on the upper left portion. GE will fix the units with a software fix at no cost. |  | 1027 | 1027 | Worldwide Distribution-USA (nationwide) including Washington D.C., and Guam and the countries of ALBANIA, AUSTRALIA, BAHRAIN, BELGIUM, BRAZIL, CANADA, CHILE, CHINA , COLOMBIA, DENMARK, FINLAND, FRANCE, GERMANY, GREECE, HUNGARY, INDIA, IRAQ, IRELAND, ITALY, JAPAN, JORDAN, KOREA, REPUBLIC OF, KUWAIT, MALAYSIA, MEXICO, NEW ZEALAND, NIGERIA, NORWAY, POLAND, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SPAIN, SWEDEN, THAILAND, TRINIDAD AND TOBAGO, TURKEY, UNITED ARAB EMIRATES, and UNITED KINGDOM. | Terminated | May 06, 2014 | 180 | 1027 | Computer | Computer
Z-2652-2014 | 68113 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Tenor is a mobile passive lift, intended to be used on horizontal surfaces for lifting and transfer in hospitals, nursing homes or other healthcare facilities. Transfers with the Tenor are made indoors in the patient's room, in communal areas or in a bathroom. Tenor is meant to be used with the patient type who: sits in a wheelchair, has no capability to support themselves at all, cannot stand unsupported and is not able to bear weight (not even partially), is dependent on caregiver in most situations, is a passive resident, might be almost completely bed ridden, is often stiff and has contracted joints, and/or is totally dependent. The Tenor has a safe working load of 320 KG (705 lbs). |  |  | June 10, 2014 | 2014 | ARJOHUNTLEIGH POLSKA SP. ZO.O. | ArjoHuntleigh received reports where the Tenor Lift without Scale (spreader bar) bolt broke during use. There is a remote risk concerning the breakage of the spreader bar bolt connection, part # TEN.018. During use, the bolt is subjected to mechanical forces, which may lead to breakage. | Device Design | ArjoHuntleigh sent an URGENT - FIELD SAFETY NOTICE letter dated June 10, 2014, to all affected customers. The letter included instructions for direct accounts to: 1) ensure that caregivers and users of the Tenor patient transfer lift referenced above are made aware of the field safety notice; 2) ensure that a copy of the field safety notice is placed in the Tenor Lift Instructions for Usemanual for future reference and training; 3) complete and sign the enclosed Customer Response Form and return the form to the ArjoHuntleighRegulatory Department; 4) and to be advised that an ArjoHuntleigh Service Technician will contact the facility representative listed on the Customer Response Form to schedule service to replace your spreader bar, free of charge. Customers with questions about this recall can contact ArjoHuntleigh at 800-323-1245 x 57985 or via e-mail at AHNA-RA@arjohuntleigh.com. |  | 382 | 382 lifts total worldwide (160 US) | Worldwide Distribution - USA (nationwide) and Internationally to Australia, Bahrain, Canada, France, Germany, India, Ireland, Israel, Italy, Mexico, Netherlands, New Zealand, Poland, Romania, Saudi Arabia, South Africa, Spain, Sweden, United Kingdom, and Uruguay. | Open, Classified |  |  | 382 | Computer | Not_Computer
Z-0471-2015 | 69748 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Cardiac Function Analysis Software |  |  | November 05, 2014 | 2014 | Toshiba American Medical Systems Inc | Potential problem with the cardiac function analysis software (CFA). Incorrect analysis results may be displayed in the Function Parameters for the Entire Heart displayed as analysis results of the CFA and in a Left-Ventricular Volume Curve generated based on some of those parameters. | Software design | Toshiba sent an Urgent Medical Device Correction letter dated November 5, 2014, to all affected customers. Customers were asked to refrain from using the function parameters for the entire heart and the left-ventricular volume curve until corrective actions have been taken.Customers were instructed to share this information with all users and reviewing radiologist as well as clinical engineering or Biomedical group at facility.Please complete and return the attached form and fax it to the toll free number at the top of the form 877-349-3054. This form can also be sent via email to raffairs@tams.com.Revised software to prevent occurrence of this problem will be installed on the system. When the new software becomes available the Toshiba service representative will contact for an appointment to install it on the system.Customers with questions were instructed to call 800-421-1968 or contact their local Toshiba Representative at 800-521-1968. |  | 1 | 1 unit within the U.S. only | US Nationwide Distribution to CA, MA, MD | Terminated | October 09, 2015 | 338 | 3 | Computer | Computer
Z-0196-2015 | 69472 | No | GDC | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | The 5085 is a general surgical table with high patient weight capacity, extended width capability, and lower minimal table top elevation. |  |  | October 10, 2014 | 2014 | Steris Corporation | The 5085 and 5085SRT stainless steel column shrouds may become damaged as a result of misalignment during extreme table articulation, excessive side force on the table column, or pressure on the base of the table. Severe damage to the shrouds may prevent the user from raising or lowering the surgical table. | Device Design | STERIS will mail Consignee Notification Letters to the affected Consignees via FedEx with tracking numbers for delivery confirmation. The Consignee Notification mailing will initiate the week of October 13, 2014. |  | 999 | 999 | US distribution: AL, AZ, AR, CA, CO, CN, DC, FL, GA, HI, ID, IL, IN, IA, KN, KY, LA, ME, MD, MI, MN, MS, MO, NB, NV, NY, NC, ND, OH, OK, OR, PN, RI, SC, SD, TN, TX, VA, WV, WI. Foriegn: Austrialia, Brazil, Canada, China, Ecuador, India, Korea, Mexico, Pnama, and Thailand. | Terminated | June 26, 2015 | 259 | 999 | Computer | Not_Computer
Z-0077-2015 | 69234 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS 4600 Chemistry System (VITROS 5,1 FS System family member), System Product Code 6802445. |  |  | September 08, 2014 | 2014 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics (OCD) identified an anomaly with Software Version 3.0 and below on the VITROS 4600 Chemistry Systems and VITROS Integrated Systems. Internal testing confirmed that when using calibrator barcode labels supplied with VITROS Chemistry Products Calibrator Kits 1, 2, 3, 4, 6, or 9, an unexpected assay calibration may occur if assay targets are unassigned (i.e., hidden). | Software design | On 9/08/14, URGENT PRODUCT CORRECTION NOTIFICATION Letter (CL2014-232, dated 9/8/14) and Confirmation of Receipt Form were sent via FedEx overnight courier or ORTHO PLUS e-Communications and/or US Postal Service Priority Mail (for PO Boxes only) to US consignees (OCD Direct, Drop-Ship and US Federal Government only) to inform them of the issue and advise them of the required actions. Foreign affiliates were informed of the issue via e-mail on 9/8/14 and instructed to notify their consignees of the issue and the required actions. |  | 232 | USA: 97; Foreign: 232 | Worldwide Distribution-USA (nationwide) including Puerto Rico and the countries of Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Singapore, Spain and Venezuela. | Open, Classified |  |  | 1177 | Computer | Not_Computer
Z-0181-2015 | 68723 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Philips Medical Systems Allura Xper FD20C Radiological Imaging with software version R8.2.O; System Code: 722028. |  |  | July 08, 2014 | 2014 | Philips Medical Systems, Inc. | Faulty Automatic Motion Controller (AMC), a problem in the Power On Self Test (POST) error handling was detected, can result in a hazardous movement of the C-arc. system. | Software design | Philips Medical Systems issued certified letters on 7/8/14 and receipt returned for tracking purposes. A mandatory Field Change Order with reference FC072200253 is being released that requires Philips field service engineers to install Software release R8.2.0.3. If you need any further information or support concerning this issue, contact your local Philips representative: SUPPORT 1-800-722-9377. |  | 22 | 22 systems | Distributed in the states of IL, and TN, and the countries of Argentina , Australia, Belgium, Denmark, Egypt, France, Germany, Hong Kong, Italy, Japan, Korea, Republic of Netherlands, Norway, Saudi Arabia, Switzerland, United Kingdom, and Viet Nam. | Open, Classified |  |  | 24 | Computer | Computer
Z-2649-2014 | 69067 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | AXIOM Luminos dRF system with SW VD10 or on Luminos dRF Max |  |  | August 12, 2014 | 2014 | Siemens Medical Solutions USA, Inc | It was determined that a potential malfunction may occur on the Luminos dRF system with SW VD10 or on Luminos dRF Max when the table side control is exposed to fluids. Fluid may infiltrate the tableside control panel and result in malfunction. In the very unlikely case of fluid infiltration a permanent short circuit within the control panel of the Luminos dRF / Luminos dRF Max will automatically d | Equipment maintenance | Siemens sent a Customer Safety Advisory Notice XP045/14/S, dated August 12, 2014, to direct accounts, which identified the product, problem, and action to be taken to avoid potential risk to patients and users. Customers were instructed to avoid extensive exposure of fluids to the tableside control panel. When cleaning the system, customers were instructed not to appoly excessive fluids, only use damp towels, in accordance with the product manual. Customers were asked to forward the safety notice to any of their customers and to inform Siemens of the identify of the device's new owner where possible.For questions regarding this recall call 610-219-6300. |  | 33 | 33 | Nationwide Distribution including IL VA TN PA TX IN MN NM GA MI FL CA NY ME OK MO NJ and AZ. | Terminated | May 21, 2015 | 282 | 33 | Computer | Not_Computer
Z-0183-2015 | 69480 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Artis zeego systems, model number 10280959, with serial numbers 1000-1599 : Artis zeego is a dedicated angiography system developed for single and biplane diagnostic imaging and interventional procedures including, but not limited to, pediatric and obese patients. Procedures that can be performed include cardiac angiography, neuro angiography, general angiography, rotational angiography, multipurpose angiography and whole body radiographic/fluoroscopic procedures. |  |  | September 23, 2014 | 2014 | Siemens Medical Solutions USA, Inc | A cable inside the C-arm systems may have been routed in an inappropriate manner. Improper routing of the cable may result in increased wear over time. Without additional measures, the potential exists for a cable inside the C-arm to break which could result in restricted functionality or system failure and an ongoing procedure could be terminated. Additionally, a subgroup of affected C-arm systems will receive a replacement of the cable harness. | Device Design | Siemens sent an " Important Safety Information " letter dated September 23, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Siemens will, without charge, remedy the defect or bring the product into compliance with each applicable Federal standard in accordance with a plan to be approved by the Secretary of Health and Human Services.Our service organization will contact with you to schedule an appointment to perform the corrective action. Please feel free to contact our service organization for an earlier appointment.Please observe this safety notice and comply with the corresponding measure until the update has been fully completed. For further questions, please call (610) 219-6300. |  | 164 | 164 | Nationwide Distribution including Puerto Rico. | Terminated | May 08, 2017 | 958 | 164 | Computer | Not_Computer
Z-2654-2014 | 69126 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | 4D Integrated Treatment Console (4DITC), versions 10 and 11. |  |  | August 18, 2014 | 2014 | Varian Medical Systems, Inc. | Due to user error, the 4DITC can improperly allow users to clear a safety interlock imposed by the 4DITC on behalf of Varian's Bar Code Conical Verification system from the BrainLab's ExacTract console when it is used with BCCV. | Software design | Varian issued an Urgent Medical Device Correction Notification letter dated August 19, 2014, to all affected sites. The letter informed users of the specific product and versions numbers as well as the issue. A list of third party systems that perform verification and communicate authorizations via ADI are included in Table 1 of the letter. The issue is only now to occur with the ExacTract positioning system in combination with a C-Series HE linac BCCV.User recommended actions are provided for users using auxiliary devices requiring ADI interface and for users performing SRS/SBRT.Varian has developed a fix for the issue and customer service representatives will be contacting all affected sites to schedule installation.A copy of the correction letter is to be kept with the most current labeling and all personnel who work in the radiotherapy department should be made aware of the content of the letter.The recall Return response form is to be completed and returned.Questions may be directed to 1-888-827-4265 for US and Canada. Contact information is also provided for other areas of the world. |  | 512 | 512 affected devices | Worldwide Distribution - USA (nationwide) and Internationally to Argentina, Brazil, Canada, Chili, China, Columbia, Czech Republic, France, Georgia, Germany, hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Latvia, Lithuania, Macau, Malaysia, Mexico, Netherland,s Portugal, Qutar, Russia, Singapore, Slovenia, South Korea, Spain, Sri Lanka, Switzerland, Taiwan, Thailand, Ukraine and United Kingdom. | Terminated | July 30, 2015 | 346 | 512 | Computer | Not_Computer
Z-0871-2014 | 67243 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | ARIATELE TELEMETRY TRANSMITTER, Model 96281, with SpO2 Option C. |  |  | December 19, 2013 | 2013 | Spacelabs Healthcare, Llc | The AriaTele Model 96281 with SpO2 monitoring Option C is recalled because the AriaTele (transmitter) display and the Telemetry Central Station display may show a SpO2 value when the sensor is not connected to the patient or in certain sensor failed conditions. | Software design | On January 6, 2014, Spacelabs sent an "URGENT - MEDICAL DEVICE CORRECTION Spacelabs Healthcare AriaTele Model 96281 with SpO2 Option C" letter, dated 03 January 2014 to US consignees.On January 9, 2014, Spacelabs e-mailed the "URGENT - MEDICAL DEVICE CORRECTION Spacelabs Healthcare AriaTele Model 96281 with SpO2 Option C" letter (translated as necessary), dated 03 January 2014, to foreign consignees.The letter described product, problem and actions to be taken. The Consignees where instructed to follow the "Advise on what action the clinician takes" on the letter.Spacelabs Healthcare will contact you at the earliest possible date to schedule a convenient tie for Spacelabs to correct all of your facility's affected transmitters (software update) at no cost.If you have any questions about this corrective action program, please contact Spacelabs Healthcare at 1-800-522-7025 and select 2 for Technical Support. |  | 72933 | 1057 units (988 in the US and 69 outside of the US). | Worldwide Distribution: US (nationwide) including states of: AZ, CA, FL, GA, IA, IL, LA, MA, ME, MI, MO, MS, NY, PA, SC, TN, TX, VA, and Puerto Rico; and Internationally to: Canada, Chile, and Panama. | Terminated | June 12, 2014 | 175 | 72933 | Computer | Computer
Z-0623-2014 | 67050 | No | GET | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Aesculap Hi-Line XXS handpiece |  |  | December 09, 2013 | 2013 | Aesculap, Inc. | Certain lots of the Aesculap Hi Line XXS handpiece (part # GB790R) were recalled due to a possible failure of the handpiece. | Device Design | Aesculap sent an Important Correction &amp; Removal Notification letter dated December 9, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to quarantine all affected product and returned to AIC (USA). Customers were instructed that a Aesculap Sales Representative would contact them to schedule return of any affected inventory and replace their product. Customers were asked to complete the inventory sheet with the quantity being returned. If customers could not locate the product, they were to provide an explanation as to why the inventory would not be returned (discarded, etc.). Customers with questions were instructed to call 610-984-9414,For questions regarding this recall call 800-258-1946, ext 5067. |  | 14 | 14 | Worldwide Distribution - USA including TX, NY, PA, and TN. Internationally to Austria, Switzerland, Germany, Denmark, Spain, Finland, France, Hong Kong, Israel, Italy, Japan, Netherlands, Poland, Russian Federation, and Taiwan. | Terminated | December 02, 2014 | 358 | 14 | Computer | Not_Computer
Z-0074-2015 | 69283 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | EndoWrist Stapler 45 instrument for the da Vinci Si System |  |  | September 18, 2014 | 2014 | Intuitive Surgical, Inc. | Intuitive is issuing an urgent product notice to all users of the EndoWrist Stapler 45 to discontinue use due to reports of inability to remove the stapler from tissue. | Device Design | Intuitive Surgical sent an Urgent Product Notice letter dated September 19, 2014 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to:-Immediately stop using the EndoWrist Stapler 45 instrument.-Quarantine the EndoWrist Stapler 45 instruments within your hospital. Keep the instruments quarantined until Intuitive Surgical sends you a subsequent communication with instructions regarding the quarantined product. -Ensure all affected personnel are fully informed of this notice. Forward this notice to your Risk Manager, OR Director, Purchasing Manager, Biomedical Engineering staff, and members of your medical staff who perform da Vinci Si surgery procedures using the EndoWrist Stapler 45 instrument.-Please retain a copy of this notice for your records with your da Vinci Si User Manual.If Âżyou Âżneed Âżfurther Âżinformation Âżor Âżsupport Âżconcerning Âżthis ÂżProduct ÂżNotice, Âżplease contact your Clinical Sales Representative or Intuitive Surgical Customer Service at 800-876-1310 Option Âż3 Âż(6 Âżam Âżto Âż5 Âżpm ÂżPST). Âż |  | 1034 | 1,034 Devices | US Nationwide Distribution. | Terminated | March 03, 2015 | 166 | 1034 | Computer | Not_Computer
Z-0201-2015 | 69251 | No | GWM | Intracranial pressure monitoring device. | Neurology | Neurology | 510(k) | Integra Licox Pt02 Monitor, Rx Only, Manufacturer: Integra LifeSciences (Ireland) Limited, IDA Business and Technology Park, Sragh, Tullamore, County Offaly, Ireland, Distributed by: Integra LifeSciences Corporation, 311 Enterprise Drive, Plainsboro, NJ 08536 |  |  | August 18, 2014 | 2014 | Integra LifeSciences Corp. | Integra LifeSciences Corporation received 2 complaints that the USB port on the Licox Pt02 monitor does not consistently provide the user the ability to extract the Pt02 trend data according to the User's Manual for the device. | Device Design | Integra LifeSciences sent an "Urgent Voluntary Medical Device Correction" letter/Return Response Form dated August 18, 2014, to their affected customers via traceable courier service. The letter identified the product the problem and the action needed to be taken by the customer.Complete the attached form. If you do not have an affected monitor, check the box: I do not have any monitors with an affected Serial Number listed below.If you do have affected monitors, check the box: I do have monitors with an affected Serial Number listed below. Circle all the Serial Numbersfor the affected monitors you have and record the total quantity.Complete the other information on the form and return the form by email or fax as indicated. Keep a copy of the form for your records.When your form is received, Integra will schedule this service with your facility and provide a temporary loaner monitor to minimize any disruption this action may cause.Should you have any questions, please contact Neuro Customer Service at 1-855-532-1723. We apologize for any inconvenience this may cause and thank you for your cooperation in this effort. |  | 43 | 43 (LCX02) and 17 (LCX02R) | US Distribution including the states of : CA, DC, FL, IL, IN, MI, MN, NC, NY, PA and TX | Open, Classified |  |  | 43 | Computer | Computer
Z-2657-2014 | 69113 | No | ITI | Powered wheelchair. | Physical Medicine | Physical Medicine | 510(k) | TRx Vertical Positioner, Classification Name: Wheelchair Component - Power Positioning System. |  |  | August 28, 2014 | 2014 | Motion Concepts | Motion Concepts identified that there is a possibility of knee support failure occurring through fatigue of two attachment bolts in the Vertical Position System (TRx Vertical Positioner). | Device Design | URGENT RECALL NOTIFICATION Letters (dated 8/26/2014) were sent to customers on 8/26/14 via Federal Express. Motion Concepts is requesting that all customers that have purchased a Motion Concepts TRx VerticalPositioner system in between June 2011 - May 2014 to make arrangements for replacement of the knee support mount in accordance with the instructions enclosed in the letter. For questions, please contact us at 905-695-0134. |  | 59 | US: 59 units | US Distribution in states of: AR, CA, CT, DE, IA, ID, IL, IN, MI, MD, MO, NE, NY, OH, PA, TN, TX, and WA. | Terminated | June 03, 2015 | 279 | 59 | Computer | Not_Computer
Z-0067-2015 | 69357 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | AC Powered adjustable bed frame. |  |  | September 15, 2014 | 2014 | Primus Medical LLC | Firm received several complaints that the head deck actuator bracket had separated from the bed frame. | Device Design | Primus Medical will send a letter notifying all customers with PCB900 beds at their facility of the potential head bracket separation. |  | 1726 | 1726 | Worldwide Distribution-USA (nationwide) including the states of CA, DE, FL, GA, IL, IN, KS, KY, MA, MN,MS, NC, NM, NV, NY, OH, PA, TN, TX, and VA, and the country of Canada. | Terminated | August 10, 2017 | 1060 | 1726 | Computer | Not_Computer
Z-0220-2015 | 69615 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | LANX, REF, PCR8XXX-XXX, LOT, RXONLY. Packaged within aluminum sterilization case. |  |  | March 09, 2011 | 2011 | Biomet Spine LLC. | Biomet Spine, LLC. announces voluntary recall of Durango Anchored ALIF Plate due to the set screw component potentially becoming dislodged from the plate assembly. | Device Design | Biomet notified customers by phone on March 9, 2011 to return the affected product. |  | 120 | 120 | US Nationwide Distribution in the states of California, Indiana, Louisiana, and Texas, | Terminated | January 07, 2015 | 1400 | 120 | Computer | Not_Computer
Z-2663-2014 | 69074 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ADVIA Centaur XP Immunoassay System, Catalog Numbers/Siemens Material Numbers (SMN) 10285219, 10316507, 10317207, 10317284, 10319668, 10320757, 10323213, 10324519, 10327135, 10327836, 10328940, 10329339, |  |  | August 13, 2014 | 2014 | Siemens Healthcare Diagnostics | The firm is conducting a field correction for the ADVIA Centaur and ADVIA Centaur XP Immunoassay systems, regarding manually loaded sample racks that are improperly loaded or pushed too far on the sample entry queue during normal operation of the ADVIA Centaur and ADVIA Centaur XP Immunoassay systems. Pushing the sample racks may cause misreads of Sample IDs (SID) if the sample rack loading instructions are not followed exactly as described in the ADVIA Centaur or ADVIA Centaur XP Operators Guide. | Software design (manufacturing process) | The firm, Siemens, sent "Urgent Medical Device Correction" Letters (dated 8/13/2014) to their consignees/customers informing them of the issue and actions to be taken. Field service personnel were sent a Support Bulletin describing the issue and instructing them how to deal with customer questions. The UFSN must be used by customers as an addendum to the system operating instructions until further notice. The consignees/customers were also instructed to review this letter with your Medical Director; complete and return the attached FIELD CORRECTION EFFECTIVENESS CHECK within 30 days via fax to the Customer Care Center at (302) 631-7597; retain the letter with your laboratory records and forward the letter to those who may have received this product.If you have any questions, please contact your Siemens Customer Care Center, your local Siemens technical support representative and/or call 914-524-2868. |  | 10515 | 10,515 (Domestic: 2771; Foreign: 7744) in total | Worldwide Distribution: US (nationwide) and countries of: Argentina, Austria, Australia, Bahrain, Bangladesh, Belgium, Belgrade, Brazil, Botswana, Canada, Chile, China, Colombia, Croatia, Czech Republic, Cyprus, Denmark, Egypt, Estonia, Finland, France, Germany, Georgia, Great Britain, Hungary, Indonesia, India, Ireland, Israel, Italy, Japan, Jordan, Korea, Kazakstan, Lebanon, Libya, Lithuania, Luxembourg, Macedonia, Malaysia, Malta, Martinique, Mexico, Myanmar, New Caledonia, Netherlands, Norway, New Zealand, Pakistan, French Polynesia, Philippines, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovenia, Slovakia, South Africa, Spain, Switzerland, Thailand, Turkey, United Arab Emirates, Vatican, Venezuela, Vietnam, Yemen, and Zimbabwe. | Terminated | May 02, 2017 | 993 | 10515 | Computer | Computer
Z-2664-2014 | 69045 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | The Brilliance 64 and Ingenuity CT are Computed Tomography X-Ray Systems intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. These devices may include signal analysis and display equipment, patient, and equipment supports, components and accessories. |  |  | September 01, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | During a Field Test (a customer external evaluation period), the customer detected a problem in the CIRS 4.0 Beta 2 software. When using the CIRS 4.0 Beta 2 software, during the recon during ready stage of reconstruction, images may be overlapped with, or superimposed on other images. Philips has disabled this feature at all sites evaluating 4.0 Beta 2 software. Philips will issue a software up | Software design | The firm, Philips, sent a Field Safety Notice (FSN 72800534) entitled "URGENT-Medical Device Correction" dated September 1, 2011 via mail on September 6, 2011 to its customers. The notice described the product, problem and actions to be taken. The customers were instructed to review the following information with all members of your staff who need to be aware of the contents of this communication and retain a copy with the equipment Instructions for Use. A Field Service Engineer will visit each site to install the corrected software.If you need any further information or support concerning this issue, please contact your local Philips representative or local Philips Healthcare office. For North America and Canada contact the Customer Care Solutions Center (1-800-722-9377, option 5: Enter Site ID or follow the prompts). |  | 6 | 6 | Worldwide Distribution: US distribution in OH and countries of: Australia, China, Belgium, Israel, and Sweden. | Terminated | July 01, 2015 | 1399 | 6 | Computer | Computer
Z-0190-2015 | 69473 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | MX 16-slice CT Scanner System and MX 16-slice CT Scanner System Phase 11 can be used as a whole body computed tomography X-ray system featuring a continuously rotating X-ray tube and detector array with multi-slice capability up to 16 slices simultaneously. The acquired X-RAY transmission data is reconstructed by computer into cross-sectional images of the body from the same axial plane taken at different angles. The system is suitable for all patients. Access Dual CT system can be used as a Head and Whole Body Computed Tomography X-ray System featuring a continuously rotating X-ray tube and detector array with multi-slice capability. The acquired X-ray transmission data is reconstructed by computer into cross-sectional images of the body from the same axial plane taken at different angles. |  |  | October 09, 2014 | 2014 | Philips Medical Systems (Cleveland) Inc | &quot;The orientation of images is displayed incorrectly. When operator chooses gantry on the right side from operation console, and the patient position is HFDL, HFDR, FFDL or FFDR, the patient position picture and text is not consistent. &quot;L&quot; and &quot;R&quot; of scan images are reversed and this is not consistent with the patient actual position. &quot;When a surview which is longe | Software Design Change | Philips is informing the consignees of this issue through a Field Safety Notice (FSN) 72800615 on October 09, 2014. The US consignees will receive this notice via certified mail and Philips Key Market representatives will distribute the FSN internationally. Additionally, A Field Service Engineer will contact each site to schedule a time to install the software update. |  | 708 | 708 | Domestic Distribution: AR, CA, FL, GA, IL, IN, KY, MI, MS, NC, ND, NJ, NV, NY, OH, PA, PR, SC, TN, and VA. Foreign distribution: Albania, Algeria, Argentina, Australia, Austria, Bangladesh, Belgium, Brazil, Chile, China, Colombia, Congo, Ecuador, Egypt, France, Germany, Ghana, Greece, Gutamala, Hungray, India, Indonesia, Ireland, Italy, Japan, Jordan, Latvia, Lybyan Arab JHamahiri, Lithuania, Malawi, Mexico, Myanmar, Namibia, Nepal, Netherlands, Oman, Palestine, Paraguay, Philippines, Poland, Portugal, Romania,Russia, Rwanda, Saudi Arabia, Senegal, South Agrica, South ZKorea, Spain, Sri Lanka, Switzerland, Tunisia, Turkey, Uganda, UAE, Venezuela, Vietnam, Zambia. | Terminated | July 02, 2015 | 266 | 708 | Computer | Computer
Z-0068-2015 | 69295 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | eFilm Workstation 4.0 and 4.0.1. eFilm Workstation is one of the components of a PACS (Picture Archiving and Communications System). |  |  | August 28, 2014 | 2014 | Merge Healthcare, Inc. | There is an issue related to eFilm Workstation versions 4.0 and 4.0.1 when having multiple studies open and utilizing the thumbnail panel to select multiple series from multiple studies may lead to the selection of an incorrect study. | Device Design | Merge sent an Important Product Information Notice letter, dated August 22, 2014 to consignees. The letter stated the issue, identified affected product, provided instructions, and told of a product update expected to be available by September 15, 2014. Customers with questions should contact Customer Service at 877-741-5369 |  | 2630 | 2,630 | Worldwide Distribution - US Nationwide and the countries of : Argentina, Australia, Canada, Chile, France, Germany, Hungary, Iceland, Italy, Lithuania, Mexico, New Zealand, Romania, Russia, South Africa, Sweden, and Yemen. | Terminated | June 01, 2017 | 1008 | 2630 | Computer | Computer
Z-2666-2014 | 69051 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | The Brilliance CT Big Bore is a whole-body Computed Tomography X-ray system intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data from the same axial plane taken at different angles. Model Numbers 728244 and 728243. |  |  | August 31, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Philips Healthcare received a complaint stating the system opens e-stop while sitting idle causing all movements and scan to stop. The problem has only occurred one time when the scan was out of Ready for Scan or scanning mode. | Software design | A Customer Letter (FSN 72800580), Attachment #1, will be sent to all US and Canadian consignees via certified mail. Philips Key Market representatives will distribute the letter internationally. Additionally, a Field Service Engineer will contact each site to schedule time to implement the Field Change Order. |  | 257 | 257 | Nationwide Distribution incuding states of: AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, IL, IN, KS, KY, MA, MD, MI, MN, MO, MS, NC, ND, NH, NJ, NM, NY, OH, OK, OR, PA, TN, TX, VA, WA, and WI. | Terminated | July 01, 2015 | 1034 | 257 | Computer | Computer
Z-0151-2015 | 69400 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Joerns EasyCare Bed |  |  | September 24, 2014 | 2014 | Joerns Healthcare | The Easy-Care Care-Lock feature installed at the foot of the bed may potentially fail to stay engaged (locked) after the Care-Lock System has been engaged by stepping on the red (locked) pedal. | Device Design | Joerns sent an Urgent: Medical Device Correction letter dated September 24, 2014 to affected customers. The letter identified the affected product, problem and actions to be taken. The firm will provide the customers with a temporary solution kit to safegard the bed until a replacement Care-Lock assembly can be shipped and installed. Customers were instructed to complete the response form. For questions call 800-826-0270. |  | 3135 | 3135 | Worldwide Distribution: US (nationwide) and country of: Canada. | Terminated | February 01, 2016 | 495 | 3135 | Computer | Not_Computer
Z-0737-2014 | 66788 | Yes | CAC | Autotransfusion apparatus. | Anesthesiology | Anesthesiology | 510(k) | Atrium Ocean Water Seal Chest Drain ATS (Autotransfusion) Blood Recovery System with stopcock Blood Recovery System; Product Code: 2050-300. |  |  | November 19, 2013 | 2013 | Atrium Medical Corporation | Chest Drain tubing of of the ATS Blood recovery may leak or disconnect | Device Design | Atrium Medical issued notification letters to direct consignees by UPS next day on 19 November 2013, and indirect consignees (distributors customers) by UPS next day or USPS priority (for customers with a PO Box) on 27 November 2013. The Letter explain the issue and the steps that should be taken. The customer will be offered replacement with unaffected product or credit. The firm describes Procedures for Emergency Use. The firm identifies potential product substitutions. These products are the same drainage product without the auto-transfusion line.International Distributors. International distributors provided with the Recall Authorization' package and the customer letter.Accounts are requested to complete the Recall Reply Form acknowledging receipt of the notification.Contact Atrium Medical Customer Service at 1-800-370-7899, Monday through Friday between 9:00 am to 5:00 pm. |  | 6684 | 6,684 units | Worldwide Distribution-USA (nationwide) and the countries of Australia, Canada, Austria, Ireland, Romania, South Africa, Colombia, Lebanon Netherlands, Jordan, Russia, and Taiwan. | Terminated | June 11, 2015 | 569 | 71990 | Computer | Not_Computer
Z-2665-2014 | 69129 | [6 empty] | Siemens CentraLink Data Management System, software version 14. |  |  | August 19, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has identified a task scheduler issue with the CentraLink&quot; Data Management System, software version 14. The issue may cause the software to stop executing commands, including uploading validated results to the LIS. The issue is related to an internal software timer that overflows after 24 days. Automated rules triggered by task schedulers or by actions in the CentraLink user interface may also be affected. As each customer configuration is unique, in addition to the delay in uploading of validated results, the specific consequences of discontinued command execution will vary. | Software design | An Urgent Medical Device Correction (UMDC #10819364) letter was issued on August 19, 2014 to all affected customers. The letter instructed customers as follows:Siemens Healthcare Diagnostics is conducting a field correction for the CentraLink" Data Management System. To prevent the issue from occurring, log out of user sessions and restart the CentraLink services and translators as soon as it is convenient, then periodically restart going forward. Any restart frequency that occurs more often than every 24 days is acceptable and will prevent the issue from occurring. Refer to the Additional Information section for detailed instructions. |  | 858 | 858 | Worldwide Distribution-USA (nationwide) including Puerto Rico and the countries of Argentina, Australia, Austria, Bangladesh, Belgium, Brazil, Canada, Chile, China, Colombia, Denmark, Egypt, Finland, France, Germany, Greece, Hungary, India, Ireland, Israel, Italy, Lebanon, Lithuania, Luxemburg, Malaysia, Mexico, Netherland, New Zealand Norway, Pakistan, Portugal, Peru, Qatar, Russia, Saudi Arabia, Singapore Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand United Arab Emirates, United Kingdom, Uruguay, Zambia. | Terminated | May 12, 2017 | 997 | 858 | Computer | Computer
Z-0168-2015 | 69583 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | COMPASS, Model No. SW Version 3.1, catalogue number CS10-100 (medical Linear accelerator) radiological |  |  | September 25, 2014 | 2014 | Iba Dosimetry Gmbh | Error in the software. A deviation between reconstructed and planned dose distribution may not be detected prior to treatment and this can result in an over or under-estimation of the pretreatment delivered dose. | Software design | The field safety notice including a confirmation of receipt was sent to the affected hospitals; by email and UPS letter on 09/25/2014. |  | 124 | 124 units | Worldwide Distribution-Distributed in the US including Guam and the countries of Germany, Austria, France, UK, Poland, Finland, The Netherlands, Spain, Turkey, Russia, and South Africa, Algeria, Mexico, Brazil, Venezuela, Columbia, Panama, China, Japan Philippines, Korea, India, Singapore, Thailand, Taiwan, Hong Kong, and Malaysia. | Terminated | January 26, 2015 | 123 | 124 | Computer | Computer
Z-0169-2015 | 69463 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Elite XPL (Pulsed Light) Handpiece for the Cynosure Elite MPX Laser System |  |  | October 09, 2014 | 2014 | Cynosure, Inc. | If the laser system is turned on and the user adjusts, installs or removes the XPL pulsed light handpiece, there is a potential risk of an electrical shock | Device Design | Cynosure sent an Urgent Medical Device Recall Notification dated October 9, 2014, to all affected customers. They were also notified by phone. The letter identified the product the problem and the action needed to be taken by the customer.There are two methods for returning the Elite XPL handpieces. If the handpiece is NOT currently connected to the laser console, the customer will be asked to return it to the Cynosure office. If the handpiece is not connected there is no risk.If the handpiece is currently connected to the laser console, the customer is instructed to immediately discontinue use of the laser system and to turn off the system immediately. They are also being instructed to leave the handpiece attached to the console and to not touch it. A Cynosure Field service engineer will then come to the office and remove the handpiece and return it to the Cynosure factory. In addition, the Cynosure Field Service engineer will install an interlock into the connector receptacle on the laser console. This interlock plug is an existing component that is used on systems that do not have an IPL handpiece in place.If your XPL pulsed light handpiece is not attached to the laser system please call Cynosure customer service (1-877-222-6822) to arrange for the handpiece to be returned and repaired. |  | 322 | 322 units | Worldwide Distribution - Nationwide Distribution and the countries of : Australia, Canada, France, Germany, Italy, Netherlands, Ukraine and UK. | Terminated | March 26, 2015 | 168 | 322 | Computer | Computer
Z-0105-2015 | 69348 | No | HGM | Perinatal monitoring system and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Philips Avalon Monitors with software revision J.30.58: |  |  | October 03, 2014 | 2014 | Philips Medical Systems, Inc. | Philips Intellivue and Avalon Fetal Monitors in time-synchronized automatic/sequence mode, the NBP automatic measurement series is stopped | Software Design Change | Philips Healthcare issued on 10/3/14 the Urgent Medical Device Correction notification/Field Safety Notice. The letter informs customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue.The correction will consist of a software upgrade free of charge.A Philips Healthcare representative will contact customers with affected devices to arrange for service. Philips is asking customers to follow the Action to be Taken by Customer/User section of the Urgent Medical Device Correction notification/Field Safety Notice:During the interim period until the SW is upgraded please make sure, if the NBP repetition interval needs to be increased (e.g. from 5 to 15 minutes), to restart the automatic NBP measurement.Please review this information with all members of your staff who need to be aware of the contents of this communication. It is important to understand the implications of this communication. |  | 118 | 118 units | Worldwide Distribution-USA (nationwide) and the countries of AUSTRALIA AUSTRIA BAHRAIN BELGIUM CHINA COLOMBIA CZECH REPUBLIC DENMARK FINLAND FRANCE GERMANY HONG KONG INDIA IRELAND ISRAEL ITALY JAPAN KAZAKHSTAN NETHERLANDS NEW ZEALAND NICARAGUA NIGERIA NORWAY OMAN POLAND ROMANIA RUSSIA SINGAPORE SLOVENIA SOUTH AFRICA SOUTH KOREA SPAIN SWEDEN SWITZERLAND TAIWAN THAILAND TURKEY UKRAINE and UNITED KINGDOM. | Open, Classified |  |  | 7705 | Computer | Computer
Z-0773-2014 | 64740 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Imagecast PACS is an image management system intended to be used by trained professionals, including but not limited to physicians, nurses and medical technicians. The system is used with general purpose computing hardware to acquire, transmit, process and store images and data throughout a clinical environment. Data and images are acquired through DICOM compliant imaging devices and modalities. Lossy compressed mammographic images and digitized film screen images must not be reviewed for primary image and interpretations. Mammongraphic may only be interpreted using an FDA approved monitor that offers at least 5 Mpixel resolution and meets other technical specifications reviewed and accepted by the FDA. |  |  | March 04, 2011 | 2011 | Ge Healthcare It | In Imagecast PACS versions prior to 3.6.51.01, the estimated magnification factor (0018,1114) was not accounted for. These versions of Imagecast PACS look for values to be present in the following tags to calculate a measurement from pixels to millimeters (mm): (0020,0030) Pixel Spacing and, (0018,1164) Imager Pixel Spacing. Because the estimated magnification factor is not used, distance meas | Software design | Important Product Information Letters dated March 4, 2011 were sent to all customers that received the defective software. The letters advised customers to upgrade their PACS software version to Imagecast 3.6. The software upgrade can be requested from GE Support via the eService portal or by telephone at 1-800-437-1171. If the customer elects to not upgrade the software, instructions were provided for a new system calibration procedure. |  | 43 | 43 units | Nationwide Distribution: US including states of: AL, AZ, CT, FL, GA, IA, IN, LA, MA, MI, MN, MO, NC, NH, NJ, NY, OH, PA, TX, VA, WI, WV + District of Columbia. | Open, Classified |  |  | 43 | Computer | Computer
Z-2587-2014 | 68618 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | AXIOM Artis or Artis zee system |  |  | June 05, 2014 | 2014 | Siemens Medical Solutions USA, Inc | There is a potential problem with the AXIOM Artis or Artis zee system in connection with the hand switch operated at the patient table which may result in uncontrolled release of an x-ray image. | Device Design | Siemens sent an Important Safety Information letter dated June 5, 2014, to affected customers. The letter described the affected product, problem, and actions to be taken. Customers were instructed to observe this safety notice and comply with the corresponding measures until the update has been fully completed. |  | 1067 | 1067 | US Nationwide Distribution. | Terminated | April 06, 2015 | 305 | 1067 | Computer | Not_Computer
Z-0805-2014 | 67013 | No | LNX |  |  | General Hospital | 510(k) | ORCHESTRA/ORCHESTRA PLUS Programmer |  |  | November 20, 2013 | 2013 | Sorin Group Italia S.r.l. | Sorin has voluntarily issued a notification to physicians related to the overestimation of the residual longevity displayed by its programmer (ORCHESTRA or ORCHESTRA PLUS) during a follow-up exam of patients implanted with a REPLY or ESPRIT pacemaker. | Software design | On Nov 20 and Nov 25, 2013, the Dear DoctorLetter was sent by certified mail to 1,310 physicians who are following 13,003 patientsimplanted with a REPLU or ESPRIT pacemakers in the US. Sorin provided the following recommendations to the physicians: You should consider checking the battery impedance of the last follow-up exam. In casethe battery impedance is greater than or equal to 3.5kÂż, a follow-up visit must bescheduled within a maximum of 6 months from the last follow-up visit. When pacemaker operation is checked by the simple application of a magnet, a magnetrate less than 95 min-1 should trigger a follow-up exam in the pacemaker centre. As a general rule, a maximum of 6 month follow-up interval when the battery impedancebecomes greater than or equal to 3.5 kÂż. This recommendation should also be followedsubsequent to the installation of the new programmer software version. |  | 1718 | 1,718 | Nationwide. | Terminated | September 08, 2014 | 292 | 1718 | Computer | Computer
Z-0787-2014 | 67126 | No | JAI | Powered radiation therapy patient support assembly. | Radiology | Radiology | 510(k) | PRIMUS, ONCOR, ARTISTE Linac systems with component 550 TxT Treatment Table. |  |  | December 05, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Update to correct the possibility for patients or operators to have their fingers trapped or pinched between the table top and the bottom frame of the 550 TxT Treatment Table. Combined with this update instruction is a second complaint issue related to an incident in which a cable inside of the TxT" Treatment Table has been damaged due to contacting a moving part. The resulting short-circuit cau | Device Design | The firm, Siemens, sent a "Medical Device Correction" letter to affected users/customers starting on December 5, 2013 via Federal express. The letter described the product, problem and actions to be taken. Siemens has arranged for modification of the affected 500 TxT Treatment Table. A service representative will modify the mounts of the TT-A, TT-D, TT-M, or TT-s table tops ( Pleases note that this modification only applies to the Siemens provided table tops TT-A, TT-D, TT-M, and TT-S.). In addition, the Siemens service representative will modify the cable routing inside of the patient table base. The customers were instructed to include this Field Safety Notice in your Digital Linear Accelerator System Owner manual chapter 'Safety Advisory Letters' where it should remain.If you have any questions, please contact the Senior Manager, Regulatory Affairs at 1-925-293-5442 or email: christine.dunbar@siemens.com. |  | 305 | 305 active devices. | Worldwide Distribution: US (nationwide) and internationally to: Canada, Italy, P.R. China, Argentina, Pakistan, South Africa, Egypt, Iran, Vietnam, Indonesia, Brazil, Hungary, Bulgaria, India, Spain, Switzerland, Trinidad, Tobago, Italy, Algeria, Croatia, Turkey, Japan, Germany, Portugal, Poland, France, Sri Lanka, Russian Federation, France, Lebanon, Slovakia, Ukraine, Pakistan, and Korean Republic. | Terminated | January 30, 2015 | 421 | 305 | Computer | Not_Computer
Z-0107-2015 | 69260 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Mindray V21, Size: 515 mm X 335 mm X 685 mm, N.W.: 8 kg, G.W.: 11 kg, Qty:1 |  |  | July 31, 2014 | 2014 | Mindray DS USA, Inc. dba Mindray North America | Mindray DS USA Inc. initiated a voluntary field corrective action for V-Series Patient Monitor because there is an issue with the V-Series Drug Calculator function. | Software design | Mindray DS USA, Inc. sent a Recall Correction Action Letter to their customers via certified mail with return receipt on July 31, 2014. |  | 2247 | 2,247 units | Worldwide Distribution: US (nationwide) and internationally to: Austalia, Canada, Columbia, Germany, India, Indonesia, Ireland, Israel, Kuwait, Malaysia, New Zealand, Pakistan, Russian Federation, Saudi Arabia, South Africa, Sri Lanka, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom, and Venezula. | Open, Classified |  |  | 2247 | Computer | Not_Computer
Z-0786-2014 | 67153 | No | KPQ | Radiation therapy simulation system. | Radiology | Radiology | 510(k) | SIMVIEW NT and SIMVIEW NT with IDI: |  |  | December 11, 2013 | 2013 | Siemens Medical Solutions USA, Inc | In a very rare occurrence an intended movement of the SimView NT simulation system components could not be stopped with usual hand control or the collision detection functionality. | Device Design | Siemens sent an Urgent Field Safety Notice to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. What preventive measures can the user take?To stop the movement in this case, push the emergency button / stop button at control console or push the in-room emergency button. The emergency button/ stop button functionality is described in the User Manual.There is no preventive measure to avoid this situation. This letter is intended to bring this potential safety issue to your awareness and to ensure that your personnel are trained about the use of the emergency button / stop button inside the simulation room and on the control console. Always monitor any movement of the simulator components.The Update Instruction containing the Urgent Field Safety Notice will be delivered by the Siemens Customer Service Group, as per their procedureSiemens is working on a solution to mitigate the described situation. You will be informed as soon as the update distributing this solution is available.Please include this Field Safety Notice in your SimView NT Simulator System Owner Manual chapter Safety Advisory Letters where it should remain.In the interests of safety, we ask that you perform the described preventive measures and inform all affected personnel immediately.The relevant National Competent Authority has been informed of this Customer Safety Advisory Notice.We regret any inconvenience that this may cause, and we thank you in advance for your understanding. For further questions please call (925) 602-8157. |  | 65 | 65 active devices | Class II Recall - Worldwide Distribution - US Distribution including the states of California, Louisiana, Massachusetts, Pennsylvania, New York, Kentucky and Michigan., and the countries of Bulgaria, Germany, Belarus, Russian Fed., P.R. China, India, Republic Korea, Germany, Costa Rica, Argentina, Portugal, Greece,Bulgaria, Vietnam, Hungary, Italy, Palestinian Ter, Hungary, Jordan, EgyptPoland, Greece, Curacao, St Eus, Malaysia, Egypt, Malaysia, South AfricaTurkey, Thailand, Qatar and Iraq. | Terminated | March 31, 2014 | 110 | 65 | Computer | Computer
Z-2738-2014 | 69122 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | SPECTRUM Pump, Model No. 35700BAX. |  |  | September 03, 2014 | 2014 | Baxter Healthcare Corp. | One Service Technician may not have correctly serviced specific Sigma Spectrum Infusion Pumps according to established procedures during the time period of 5/5/2014 through 6/3/2014. | Equipment maintenance | Initial notification was initiated via telephone call by the Baxter Medina Service Center on 8/28/14 to all affected Sigma SPECTRUM Infusion Pump consignees. URGENT DEVICE CORRECTION Letters (dated 9/03/2014) were sent to the consignees via USPS first class mail on 9/03/14 formally informing them of the recall. The letters identified the affected product, the description of the issue, the hazard involved, as well as, the actions to be taken by customer. Customers are being instructed to immediately remove the pump from use and return the pump to Baxter for inspection. The firm will provide replacements. Customers are to contact Baxter Healthcare Medina at 800-356-3454 for technical questions regarding the letter. |  | 58 | USA: 56 units, Canada: 2 units | Worldwide Distribution -- USA and Canada. | Terminated | March 16, 2017 | 925 | 58 | Computer | Not_Computer
Z-0222-2015 | 69531 | No | HQF | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | MC-500 Multicolor Laser Photocoagulator using software version 2.20. |  |  | May 24, 2013 | 2013 | Nidek Inc | Retrospective review found that a Engineering Change Order initiated in May 2013 for a memory function problem was not reported as required. | Software design | Nidek Engineers visited all affected sites to install new software version.For questions regarding this recall call 510-353-7785. |  | 5 | 5 | US Distribution in the states of: CA, NE, and NY. | Terminated | November 17, 2014 | 542 | 5 | Computer | Not_Computer
Z-0491-2015 | 69541 | No | LFL |  |  | General &amp; Plastic Surgery | 510(k) | Integra Selector Kit |  |  | September 29, 2014 | 2014 | Integra LifeSciences Corp. | Some specific lots of Selector Tubing may potentially leak irrigation fluid during use if not placed on the Selector handpiece in a specific manner. | Device Design | Integra sent an 'Urgent- Medical Device Correction' letters/acknowledgement forms dated 9/29/2014 by traceable courier service. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review their inventory for the affected product, complete the attached form. Complete the other information on the form and return the form by email or fax.Customers with questions should contact Customer Service at 1-855-532-1723. |  | 6358 | 6,358 tubing kits | Worldwide Distribution - USA (nationwide) and Internationally to Algeria, Australia, Bulgaria, Canada, Germany, Great Britain, Greece, Hungary, Indonesia, Israel, Italy, Japan, Mexico, Portugal, Spain, Sweden, Thailand, Turkey, and Wales. | Terminated | May 27, 2015 | 240 | 6358 | Computer | Not_Computer
Z-2732-2014 | 69132 | No | LTI |  |  | Gastroenterology/Urology | PMA | LAP-BAND AP SYSTEM AP Large with Access Port, Allergan, LAP-BAND AP SYSTEM Access Port I. Product was distributed for both the AP Large and AP Standard sizes. |  |  | August 15, 2014 | 2014 | Apollo Endosurgery Inc | Product was distributed past its expiration date. | Software Manufacturing/Software Deployment | Firm called its consignees and advised them to either return or destroy the devices which had not already been implanted. |  | 10 | 10 | Distributed in the states of California, Delaware, Indiana, Missouri, New Jersey, New York, and Oregon. | Terminated | February 01, 2016 | 535 | 10 | Computer | Not_Computer
Z-0217-2015 | 69565 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Edwards Lifesciences Vent Catheters, . Packaged sterile and non-pyrogenic in a sealed, peel-type pouch. Model Numbers: E061, E060, PE062, EM012 |  |  | October 08, 2014 | 2014 | Edwards Lifesciences, LLC | Edwards Lifesciences is conducting a field action on the Vent Catheters due to the potential for disruption of the white plastic coating on the stylet. | Device Design | Edward Lifesciences sent an Urgent Medical Device Safety Notice date October 8, 2014, via Fed-Ex to all affected customers. On 3/4/15, letters were sent to surgeons. As part of the customer instructions in the safely notice, customers were requested to forward copies of the safely alert to any organization where the potentially affected devices have been transferred. No product is being returned. The distribution of the device has not been stopped as a result of this field correction. No additional production lots will be released for distribution until the IFU has been revised. Customers with questions were instructed to call (800) 424-3278.For questions regarding this recall call 801-553-7531. |  | 157305 | 157,305 | Worldwide Distribution - USA including Alabama, Arkansas, Arizona, California, Colorado, Connecticut, Florida, Georgia, Hawaii, Iowa, Idaho, Illinois, Indiana, Kansas, Kentucky, Louisiana, Massachusetts, Maryland, Maine, Michigan, Missouri, Mississippi, Montana, North Carolina, Nebraska, New Hampshire, New Jersey, Nevada, New York, Ohio, Oklahoma, Oregon, Pennsylvania, Rhode Island, South Carolina, South Dakota, Tennessee, Texas, Utah, Virginia, Vermont, Washington, Wisconsin, West Virginia. US Territory: Puerto Rico and Internationally to Australia, Bangladesh, Brazil, Canada, Chile, China, Colombia, Hong Kong, India, Japan, Malaysia, Mexico, Netherlands, Singapore, South Korea, Sri Lanka, Taiwan, and Venezuela. | Open, Classified |  |  | 157305 | Computer | Not_Computer
Z-0042-2015 | 69377 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | McKesson Cardiology" Hemo intended for complete physiological/hemodynamic monitoring, clinical data acquisition, medical image and data processing, and analytical assessment. |  |  | August 27, 2014 | 2014 | Mckesson Information Solutions LLC | Software Error: The McKesson Cardiology Hemo calculation section incorrectly converts the Hemoglobin value before it is utilized in the applicable formula calculations. | Software design (manufacturing process) | On August 27, 2014, McKesson initiated a correction via a Field Safety Notice to all affected domestic customers. McKesson Israel will be responsible for contacting foreign consignees. |  | 272 | 272 devices | Worldwide distribution: US (nationwide) including states of: AR, CA, FL, DE, GA, HI, IL, IN, KS, LA, MA, MD, MN, MO, MS, MT, NJ, NV, OK, PA, SC, TNTX, , UT, WA; and countries of: Australia, Canada, Israel and the United Kingdom. | Terminated | May 18, 2016 | 630 | 272 | Computer | Computer
Z-0137-2015 | 69296 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | Animas Vibe Insulin Infusion Pump and System. |  |  | October 25, 2012 | 2012 | Animas Corporation | Animas determined that keypad wear over time was contributing to an increased complaint rate associated with damaged keypads on the Vibe Insulin Infusion Pump and System. Animas has implemented an improvement to the keypad related to the material used to reduce susceptibility to damage and increase durability. | Device Design | Animas initiated the notification to Distributors by letter, dated July 2012, and Patients and Health Care Professionals were notified by letter in September 2012. The delay in notifying patients and health care professionals was driven by the need to have sufficient inventory of replacement product on hand in order to meet anticipated demand before sending the letters. |  | 1938 | 1938 | Worldwide Distribution in the countries of France, Italy, Sweden, and UK.. | Terminated | June 29, 2015 | 977 | 1938 | Computer | Not_Computer
Z-0221-2015 | 69656 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MOSAIQ |  |  | November 03, 2014 | 2014 | Elekta, Inc. | A problem can exist in MOSAIQ resulting in the display of incorrect numeric data due to a dose rounding error on printed reports. | Software design | Elekta sent an Urgent Important Field Safety Notice letter to alert customers of the issue. This notice was distributed on November 3, 2014. A fix to the defect has been developed and has been released by Product Creation. For question call 408-830-8000. |  | 15 | 15 | US Nationwide Distribution in the states of AZ. CO, IL, MD, MO, NY, OK, OR, PA, RI, TX, VA | Terminated | July 27, 2015 | 266 | 15 | Computer | Computer
Z-0450-2015 | 69783 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | BrightView product code: 882478 |  |  | November 04, 2014 | 2014 | Philips Medical Systems (Cleveland) Inc | Software issues | Software design | Philips Healthcare sent an URGENT - Field Safety Notice Medical Device Correction letter dated November 4, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.This Field Safety Notice 88200485 is intended to inform you about:" what the issues are and under what circumstances they can occur" the actions that should be taken by the customer / user in order to prevent risks for patients or users" the actions planned by Philips to correct these issues.If you need any further information or support concerning these issues, please contact your local Philips representative:For North America contact the Customer Care Solutions Center (1-800-722-9377: follow the prompts).Philips apologizes for any inconveniences caused by this problem. |  | 85 | 85 | Worldwide Distribution - USA including the states of AR, CA, FL, GA, ID, LA, ME, MI, MN, MO, MS, NJ, NY, OH, OK, PA, and TX. Government Facilities: AR, FL, MN, OH, and OK. and the countries of: Canada, China, Denmark, Germany, Italy, South Korea, Spain, Sweden, Switzerland, Taiwan, and United Kingdom. | Open, Classified |  |  | 85 | Computer | Computer
Z-0784-2014 | 66846 | No | KRO | Knee joint femorotibial metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer Segmental Systems; Polyethylene Inserts, Size C. |  |  | December 19, 2013 | 2013 | Zimmer, Inc. | After evaluation of 13 complaints(0.4% rate) for hyper-extension of the Zimmer Segmental System, Zimmer is updating the related surgical technique and instructions for use to provide additional instructions and patient conditions that may place excessive loading on the polyethylene insert. | Device Design | Depth of recall" Notifications were sent to customers, surgeons and risk managers, which use the affected product. Distributors were also informed of the notifications sent to surgeons and risk managers. The customers were identified by performing an invoice search and distribution history for the Polyethylene Inserts (00-5850-012-95 and 00-5850-013-95) since product launch. Recall Communication" All distributors were notified via electronic mail. Distributors that have received affected inventory were also notified via courier. Hospital risk managers and surgeons were also notified via courier.On 12/19/13, hospital risk managers and surgeons were provided with a notification identifying the issue and their responsibilities. These responsibilities include ensuring the affected personnel are aware of the contents of the notification. On 12/19/13, Distributors were sent a notification identifying the issue and to provide a copy of the notification sent to surgeons and risk managers. The distributors will also be instructed as needed to assist with identifying a comprehensive list of surgeons. " Communications outside of the United States will occur approximately two weeks after the United States communications.Effectiveness ChecksZimmer will conduct effectiveness checks in the following way:"Surgeons/Hospitalso100% of the FedEx notifications will be accounted for.oSecond and third notification attempts will be made.oFailed delivery attempts following the third attempt will result in the account being deemed as nonresponsive. |  | 5305 | 5305 distribution events | Worldwide Distribution. USA nationwide and Germany, with further distribution into EU countries. | Terminated | February 24, 2015 | 432 | 5305 | Computer | Not_Computer
Z-2733-2014 | 69139 | No | HQF | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | Array LaserLink, Manufactured by Lumenis, |  |  | August 15, 2014 | 2014 | Lumenis, Inc. | Lumenis initiated a field-correction for the Array Laser Link&quot;, GA-0006700 (SN XXYYZZ) because the system may project an erroneous pattern display on the retina, which is different than the desired pattern. | Software design | Lumenis sent a Safety Advisory Notice dated July 15, 2014, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed that Lumenis would contact consignees to schedule a product update, to be completed at a convenient time. For questions consignees should call 801-656-2549.For questions regarding this recall call 801-656-2690. |  | 16 | 16 | Worldwide Distribution - USA (nationwide) and Internationally to JAPAN, CHINA, and CANADA. | Terminated | October 21, 2014 | 67 | 16 | Computer | Computer
Z-0039-2015 | 69146 | No | OMP | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | RENASYS EZ/ RENASYS EZ Plus 250 mL canister with Solidifier |  |  | July 18, 2014 | 2014 | Smith &amp; Nephew, Inc. | RENASYS EZ Canisters exhibiting visible evidence of deformation of the inlet port on the canister lid. | Device Design | Smith &amp; Nephew sent an Urgent Medical Device Correction Notice dated July 15, 2014 to all US customers via Federal Express delivery. The letter identified the affected product, problem and actions to be taken. The firm instructed customers to discard the affected product and replace with a new one from an unaffected lot. For questions contact 727-399-3785. |  | 8650 | 19,499 = (10,849 (Item 66800058) &amp; 8,650 (Item 66800913)) | Worldwide Distribution - US Nationwide in the states of FL, NC, CA, MO, MA, TX, NY, OH, CO, VA, IL, NJ, TN, NV, MN including Puerto Rico and countries ofCanada, Mexico, Dubai, Argentina, Australia, Colombia, Germany, Hong Kong, Jordon, Malaysia, Thailand, Tunisia, New Zealand, South Africa, Shanghai, Brazil, Saudi Arabia and Lebanon. | Terminated | August 05, 2015 | 383 | 55661 | Computer | Not_Computer
Z-0780-2014 | 67078 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO RTP System |  |  | November 25, 2013 | 2013 | Elekta, Inc. | When the number of fractions is changed for a proton spot beam, the MU are recomputed and rounded if necessary, but the dose is not recalculated as it should be even after saving and re-opening the plan. | Software Design Change | An Important Safety Notice, LINXIO0017 Proton Sport Beams, Changes in Fractionation not Updating Dose, is in the process of being distributed to affected customers. The release date for this notice is November 25, 2013. The notice included a Confirmation of Receipt which was to be returned to Elekta.For further questions please call (404) 993-5886. |  | 13 | 13 | Worldwide Distribution - US Distribution including the states of IL, IN, LA, NJ, TX, WA and WI., and the countries of Belgium, Germany, Italy and Netherlands. | Terminated | May 13, 2014 | 169 | 13 | Computer | Not_Computer
Z-0002-2015 | 69258 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Siemens Ysio Max system |  |  | September 02, 2014 | 2014 | Siemens Medical Solutions USA, Inc | Two malfunctions posing potential risk to patients were identified with Siemens Ysio Max, AXIOM Luminos dRF Max, Uroskop Omnia Max, and Luminos Agile Max systems with a specific set of serial numbers that will cause lost images due to an automatic Radiology Information System (RIS) worklist update, thus causing examinations to be repeated. Image label may also be displayed incorrectly after images | Software design | Siemens sent a Safety Advisory Notice, dated September 2, 2014, to all affected customers. The Notice identified the product, the problem, and the action to be taken by the customer. Customers were instructed to disable the automatic RIS worklist update and perform this procedure manually. Customers should disable the Measurement Field Display and Collimator Display by setting the display time for both setting to 0 sec (if editable). Customers should forward the safety notice to the new owner and inform Siemens of the new owner where possible. Customers with questions should call 610-219-2119 or 610-219-4834.For questions regarding this recall call 610-219-6300. |  | 10 | 10 | Nationwide Distribution - US including UT, MO, CO, MA, MN, MS, OH, SC, MI, and NY. | Terminated | February 24, 2015 | 175 | 18 | Computer | Computer
Z-0007-2015 | 69094 | No | DIP | Diphenylhydantoin test system. | Toxicology | Toxicology | 510(k) | VITROS Chemistry Products PHYT (Phenytoin) Slides, REF/Product Code 829 8671, IVD. |  |  | August 20, 2014 | 2014 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics identified a potential for biased results to be generated when using VITROS PHYT Slides, Lot 2613-0150-4913. Some positively biased results were identified during testing of a different VITROS PHYT Slide lot; the affected slides were not released for distribution. Since Lot 2613-0150-4913 was manufactured in the same event, there is the potential for slides from this lot to also be affected. | Software Manufacturing/Software Deployment | On 8/20/14, an URGENT PRODUCT CORRECTION NOTIFICATION Letter (CL2014-222, dated 8/20/14) was sent via FedEx overnight courier and/or ORTHO PLUS e-Communications and/or US Postal Service Priority Mail (for PO Boxes only) to US consignees (OCD Direct, Drop-Ship and US Federal Government) to inform them of the issue and advise them to discontinue use of the affected lot(s) and to inform their Laboratory Medical Director if they suspect that a previously reported result may have been affected. One foreign affiliate was notified on 8/20/14. One distributor was notified on 8/20/14. The customers were advised to: - Immediately discontinue using and discard VITROS Chemistry Products PHYT slide Lot 2613-0150-4913. Review VITROS PHYT results obtained using Lot 2613-0150-4913 and discuss any concerns they may have regarding previously reported results with their Laboratory Medical Director or with the requesting physician. Complete the attached Confirmation of Receipt form no later than 8/28/14. Forward the information in this notification if they have distributed this product outside of their facility facility. |  | 349 | Domestic: 339 sales units; Foreign: 10 sales units | Distributed USA (nationwide) and in Singapore. | Terminated | April 11, 2017 | 965 | 349 | Computer | Not_Computer
Z-0529-2014 | 66833 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | EOS System X- ray beam |  |  | November 07, 2013 | 2013 | EOS Imaging | EOS imaging discovered during production internal testing that the X-ray beam collimation set up by the operator at the user interface of the acquisition workstation is different from the actual collimation ofthe X-ray beam operated by the EOS system during the X-ray acquisition. | Device Design | CDRH approves the following conditions for EOS Imaging SA:1. Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR 1003.21 and 1003.22. This office and the Food and Drug Administration (FDA) district office coordinator notes below what are to be included in the notification. The draft letter provided on November 28, 2013, is deficient for the following reasons:a. It does not provide a statement that all repairs will be made withoutcharge (21 CFR 1003.21(3)).b. It does not provide instructions, with respect to the use of the product pending the correction of the defect (21 CFR 1003.21(1). Specifically, the letter should include directions for the user to determine if the collimation defect is currently present in their system, which would initiate immediate correction, rather than at the next semi-annual service.2. The customer notification letter will be sent in the manner required by 21CFR 1003.21(b).3. The repairs outlined are to be completed by July 1, 2014. All systems under service contracts (which includes semi-annual service) would be visited between receipt of this letter and July 1, 2014. All repairs, including those to systems which are not under service contracts, are also expected to be completed by July 1, 2014.EOS may proceed with implementation of the CAP (subject to the conditions noted above. The customer notification letter may be sent out immediately upon addressing the above deficiencies. However, it is suggested that you provide a copy of the revised letter to FDA for approval prior to issuance.For further questions please call (678) 564-5400. |  | 22 | 22 EOS systems are installed in US | Nationwide Distribution | Terminated | January 27, 2015 | 446 | 22 | Computer | Not_Computer
Z-0839-2014 | 66742 | No | ILO | Water circulating hot or cold pack. | Physical Medicine | Physical Medicine | 510(K) Exempt | DeRoyal (R) Aquarius Hot/Cold Therapy Combo Unit, w/ Foot/Ankle Blanket NS, REF T656NS, Rx only, Manufacturer DeRoyal Industries, Inc. 200 DeBusk Lane, Powell, TN 37849. |  |  | September 17, 2013 | 2013 | DeRoyal Industries Inc | The unit displays an error code of hose kinking when no hose kink is present; the unit fails to monitor the temperature of the water in the blanket as intended and controlling it through the automatic temperature control. This failure could result in temperature injuries to the treatment site; and there have been two reports of electric shock. | Device Design | DeRoyal sent a "URGENT VOLUNTARY RECALL DEROYAL HOT/COLD THERAPY UNIT" letter dated September 17, 2013. The letter identified the product, problem, and actions to be taken by the customers. |  | 25 | 25 units | Nationwide Distribution. | Terminated | June 05, 2017 | 1357 | 1343 | Computer | Not_Computer
Z-0832-2014 | 67306 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance iCT &amp; Brilliance iCT SP Computed Tomography X-Ray Systems, Philips Healthcare, Cleveland, OH |  |  | December 09, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | Patient images exhibited ring artifacts. | Software design | The firm, Philips Healthcare, sent an "URGENT - Medical Device Correction" letter dated December 9, 2013 to their customers. The letter described the product, problem and actions to be taken. The customers were instructed to do the following to avoid the occurrence of these specific ring artifacts:1. Disable Automatic Scan Time for all Exam Cards2. Leave Automatic Scan Time on during planning and prior to clicking Goto initiate the scan, click the acquisition series of the scan and disableAutomatic Scan Time, then proceed with the scan.3. Disable Edit Before Final Recon (Reconstruction) for all Exam CardsIf a ring artifact is present, clinical judgment should be used by thecustomer/user to determine if the images acquired in the dataset contain enough information to make a diagnosis.Philips will be releasing a Field Change Order (FCO) to correct the issue with software version release 4.1.1.19261. In addition, a Philips Field Service Engineer will contact you to schedule the software installation at your site.If you need any further information or support concerning this issue, pleasecontact your local Philips representative or local Philips Healthcare office. ForNorth America and Canada, contact the Customer Care Solutions Center (1-800-722-9377, Option 5: Enter Site ID or follow the prompts). |  | 7 | 7 units | Worldwide Distribution: US (nationwide) including states of: FL, IN, NY, OR and PA; and internationally to: Australia and Israel. | Terminated | May 29, 2014 | 171 | 7 | Computer | Not_Computer
Z-0224-2015 | 69537 | No | HQF | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | Nidek MC-500 Multicolor Laser Photocoagulator using software version 2.20. |  |  | October 24, 2012 | 2012 | Nidek Inc | When the MC-500 is turned on and coagulation is performed without changing the laser color, the delivery unit occasionally fails to emit the laser, or the power output is lower than the set program. | Device Design | Nidek Field Service engineers visited each affected location to upgrade software. |  | 21 | 21affected devices | US Distribution in the states of: NJ, SC, IL, MI, NY, CA, TX, OH, IN, UT, and MN. | Terminated | November 17, 2014 | 754 | 21 | Computer | Computer
Z-0038-2015 | 69104 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | RECLAIM Distal Reamer Extension |  |  | September 03, 2014 | 2014 | DePuy Orthopaedics, Inc. | Specific lots are being recalled due to the potential for the tabs to break, leaving fragments in the patient. | Device Design | On February 25, 2013, DePuy Orthopaedics, Inc. issued a device correction for the specific lots of the RECLAIM Distal Reamer Extension impacted by this recall. The letter stated the reason for the recall and swap out. A subsequent letter outlined a formal swap-out process of the affected reamer extensions as part of the device correction strategy. Newly designed devices are now available. Questions about device swap-out information may be directed to Kim Earle, Recall Coordinator, 574-371-4917; and Surgeon questions may be directed to DePuy's Scientific Information Office at 1-888-554-2482. |  | 875 | 875 units | Worldwide Distribution -- USA, including the states of ME, MD, PA, MA, FL, LA, IL, IA, OH, MN, TX, MT, WA, CA, HI, NV, AZ, KY, WI, CO, GA, MI, NC, VA, NY, TN, and IN; and, the countries of Canada, Australia, Austria, Belgium, Czech Republic, United Kingdom, France, Germany, Ireland, Israel, New Zealand, Norway, Poland, Slovenia, South Africa, Switzerland, and Chile. | Terminated | November 09, 2015 | 432 | 875 | Computer | Not_Computer
Z-0801-2014 | 67168 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Computed Tomography X-Ray System (Brilliance CT Big Bore Oncology &amp; Brilliance CT Big Bore Radiology) Distributed by Philips Healthcare System, Cleveland, OH. |  |  | December 04, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | Machine Name in TumorLOC Basic Mode is obscured. When TumorLOC is in Basic Mode and Set Vertical Laser equals Yes, the Machine Name label and field are covered by the controls above it. The Machine Name cannot be viewed or changed. | Software design | On 12/4/2013 the firm sent Urgent Medical Device Corrections letters to their customers. |  | 32 | 32 units | Worldwide Distribution-USA (nationwide) including the states of AK, MD, NH, NV, OH, and OR, and the countries of Australia, China, France, India, Indonesia, Italy, Korea, Mexico, Poland, Romania, Russia, Spain, Switzerland, Taiwan, Thailand &amp; United Arab Emirates. | Terminated | July 01, 2015 | 574 | 32 | Computer | Computer
Z-0101-2015 | 69242 | No | BSB | Blood and plasma warming device. | Hematology | Hematology | 510(k) | Ultra Battery 1 Model: TA-UB1 NSN: 6515-01-592-5803 |  |  | July 25, 2014 | 2014 | Estill Medical Technologies, Inc | Estill has discovered that a fully discharged Ultra Battery 1 (TA-UB1) will reach 96.9% of its capacity during a 60 minute charging period. If you have a fully discharged Ultra Battery 1 (TA-UB1), in order to achieve a 100% of charge capacity you should charge it for a period of 90 minutes. In addition, the charging voltage range of 10-36V specified on the Ultra Battery 1 (TA-UB1) label and in t | Device Design | Thermal Angel sent a Notice of Corrections letter dated July 25, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to communicate the contents of the notice to all employees, consignees, consultants, or other parties affected by this correction. The firm also updated the labeling and instructions for use. For additional information cusotmers were instructed to contact Medical Technologies at (877)354-0286.The recalling firm issued a Notice of Corrections to its distributors and customers.For questions regarding this recall call 214-561-6001. |  | 6886 | 6,886 | Worldwide Distribution - USA (nationwide) and Internationally to Serbia. | Terminated | August 21, 2015 | 392 | 6886 | Computer | Computer
Z-0828-2014 | 67278 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Accuray CyberKnife Robotic Radiosurgery System; Accuray Incorporated Sunnyvale, CA. |  |  | January 13, 2014 | 2014 | Accuray Incorporated | Potential Safety issue with Synchrony Boom Arm Mounting Assembly - one complaint of mounting assembly detaching. | Device Design | Urgent Medical Device Correction notification letter were sent out beginning January 14, 2014 by trackable method. Letters describe potential problem and provide inspection instructions to be performed prior to use. |  | 44998 | 302 devices; 149 in US, 153 out of US. | Worldwide Distribution. - US (Nationwide) and Mexico; and Internationally to: Australia, Belgium, Canada, China, Colombia, Czech Republic, England, Finland, France, Germany, Greece, India, Ireland, Italy, Japan, Korea, Malaysia, Luxembourg, Netherlands, Poland, Russia, Saudia Arabia, Spain, Switzerland, Taiwan, Thailand, Turkey, Ukraine, United Kingdom, Venezuela, and Vietnam | Terminated | November 25, 2014 | 316 | 44998 | Computer | Not_Computer
Z-0844-2014 | 67188 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | KODAK DirectView DR 7500 Dual Detector System, MODEL DR 7500, Manufactured by Carestream Health, Inc. Made in U.S.A. |  |  | December 26, 2013 | 2013 | Carestream Health, Inc. | Carestream Health, Inc. has recalled DR 7500 Dual Detector System utilizing Version 5.X Software due to a possible patient image display error. | Software Manufacturing/Software Deployment | On December 26, 2013 Carestream Health, Inc. distributed Urgent Medical Device Recall notification letters dated December 19, 2013 to their customers via Fed Ex. Carestream Health, Inc. recommends that once the daily detector calibration process is initiated on the DR 7500 System running Version 5.X software it should not be interrupted. The calibration process should be completed only when patients are not being imaged. A Carestream Health Service representative will contact all customers to install a software modiciation to Version 5.X software which will not allow the cancellation of detector calibration. Customers will questions can call 1-800-328-2910 7 days per week on a 24 hour basis. |  | 164 | 164 units Total (42 domestically &amp; 122 internationally) | worldwide Distribution - US (Nationwide) including states of: IN, PA, MD, MI, NY, GA, MN, OK, VA, TX, WI, CA, IL, TN, WA, NC, SC, MA, NJ, and KY; and Internationally to: Canada, New Zealand, South Korea, Australia, Hong Kong, Saudia Arabia, Guinea, Kuwait, Israel, Oman, South Africa, China, UK, Sweden, Poland, Denmark, Germany, Finland, Czech Republic, Italy, France, Portugal and Spain. | Open, Classified |  |  | 164 | Computer | Computer
Z-0011-2015 | 69092 | No | INM | Mechanical chair. | Physical Medicine | Physical Medicine | 510(K) Exempt | Nemschooff, Model Numbers - Serenity II:790-12, Serenity III: 791-12, Serenity III Exam: 791-12E, and Serenity IIS: 792-12. Designed to provide patient comfort during patient treatment sessions. The chair may include a manual control reclining mechanism. Customer configurable models are also available. The Serenity III Exam 791-12E was discontinued 11/1/2010. |  |  | January 18, 2011 | 2011 | Nemschoff Chairs, Inc. | Nemschoff Chairs, Inc. is recalling various model of the Serenity Treatment Chair because of potential elevated temperatures in the seat and back lumbar areas of the chair. | Device Design | The corrections have been performed via heat service kits being installed "on demand" to address heating element malfunctions starting on 1/18/2011. Consignee letter to all customers is pending HHE assessment. |  | 1868 | 1868 | Worldwide Distribution: USA (nationwide) including DC and the states of AK, AL, AR, AZ, CA, CO, CT, FL, GA, IA, IL ,IN, KS, KY, MA, MD, MI, MN, MO, MN, MT, NC, ND,NE, BH, NJ NM, NV, NY, OH, OK, OR, PA, PE, SC, SD, TN, TX, UT, VA, WA, WI, WV, WY, and the country of CANADA. | Terminated | July 24, 2015 | 1648 | 1868 | Computer | Not_Computer
Z-0110-2015 | 69467 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Mortara Surveyor S12 and Surveyor S19 Patient Monitors. |  |  | October 07, 2014 | 2014 | Mortara Instrument, Inc | Martara Instrument, Inc. has recently became aware of a potential safety hazard involving our Surveyor rand S12 S19 Patient Monitors. A hazardous situation may be created when mounting these monitors using the 75mm VESA mounting hole pattern provided in the battery cover of these monitors. If the mounting screws provided by Mortara are not used and if the mounting screws used are longer than the p | Device Design | Mortar Instruments, Inc. sent an "Urgent Medical Device Correction" letter dated October 3, 2014. The letter was addressed to Health Care Administrator / Risk Manger and Director of Biomedical / Clinical Engineering.The letter described the problem and the product involved in the recall. Customers were advised to contact Mortara Technical Support at 1-888-667-8272 or TechSupport @mortara.com to schedule a service representative to update their Surveyor S12 and Surveyor S19 units or for additional questions they can contact their local Service Representative. The update will include replacement of the battery cover and replacement of the battery pack if mounting screw damage is evident. International consignees received similar dated letter except it provided instructions detailing the correction steps and requested customers to complete and return the Response Acknowledgement and Receipt Form. For questions they can contact Technical Support at +1.414.354.1600 or TechSupport@mortara.com, or contact your local Service Representative. |  | 624 | 624 | Worldwide Distribution: US (Nationwide) in the states of: AZ, AR, CA, DC, FL, ID, IL IN, KY, LA, ME, MD, MA, MI, MN, MO, MT, NV, NJ, NY, NC, OK PA, SC, TN, TX, VA, and WI; and internationally to: AUSTRALIA, BANGLADESH, CANADA, CHILE, CROATIA, EGYPT, GERMANY, ITALY, JORDAN, LEBANON, MALAYSIA, MOROCCO, OMAN, PHILIPPINES, POLAND, ROMANIA, SAUDI ARABIA, THAILAND, THE NETHERLANDS, UNITED KINGDOM. | Terminated | February 13, 2015 | 129 | 624 | Computer | Not_Computer
Z-0800-2014 | 67030 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens ADVIA 1200 and ADVIA 1650 Chemistry Systems. |  |  | December 17, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | An overflow flag //// is displayed in the result column for a test when the algorithm for the test cannot calculate a result or when the concentration cannot be determined because it is outside the absorbance limits, and it is possible for the user to misinterpret or edit the result. | Software design | The distribution of the Urgent Medical Device Correction letter scheduled to begin on December 17, 2013.Âżx Field service personnel will be notified of the Urgent Medical Device Correction Letter via the Siemens Knowledge Base, as well as information regarding placing labels on the customer system to remind the customer that ///// is not a reportable result.Âżx The Customers were instructed to contact Siemens Healthcare Diagnostics Technical SolutionCenter if they have technical questions regarding this issue.Âżx Customers are requested to complete and return the Effectiveness Check questionnaire attached to the Urgent Medical Device Correction. |  | 1563 | 1563 | Worldwide Distribution. USA nationwide and the country of Canada. | Terminated | August 26, 2014 | 252 | 1563 | Computer | Computer
Z-0802-2014 | 66591 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Civco Belly Board, Reusable non-sterile carbon fiber belly board with cushion, REF 125012, packaged 1 per box. |  |  | October 25, 2013 | 2013 | Civco Medical Instruments Inc | Scales adhered to the product may be misaligned from the zero reference such that one side will not match the other. | Device Design | Civco Medical Instruments Inc sent a FIELD SAFETY NOTIFICATION letter dated October 28, 2013 via UPS to their customers. The notification identified the affected device, problem, and field safety corrective action. The letter describes how customers can check to see whether the scales on their device are out of alignment. The letter states CIVCO will rework the system to correct the difference. For questions contact sales representative at CIVCO Medical Solutions, Radiation Oncology at 800-842-8688 or 712-737-8688. |  | 138 | 138 belly boards | Worldwide Distribution - USA Nationwide in the states of AL, AR, AZ, CA, CO, CT, FL, GA, IL, KS, MA, MD, MI, MN, MT, NC, NJ, NY, OH, OK, PA, SC, TN, TX, VA, WA, WI, and WV. Foreign distribution was made to Brazil, Canada, Chile, Czech Republic, France, Germany, Iraq, Israel, Italy, Japan, Malaysia, Saudi Arabia, Slovenia, South Korea, Spain, Turkey, and United Arab Emirates. | Open, Classified |  |  | 140 | Computer | Not_Computer
Z-0831-2014 | 66197 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance iCT and Brilliance iCT SP |  |  | July 16, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | An artifact that resembles thrombus may appear on the image. | Device Design | Philips sent an Field Safety Notice dated July 15, 2013, to all US customers via certified mail. Market representatives will distribute the letter internationally. The letter identified the product the problem and the action needed to be taken by the customers. Additionally, a Field Service Engineer will contact each site to schedule a time to implement the Field Change Order.If you need any further information or support concerning this issue, please contact your local Philips representative or local Philips Healthcare office.For North America and Canada, contact the Customer Care Solutions Center (1"800-722"9377,Option 5: Enter Site 10 or follow the prompts).This notice has been reported to the appropriate Regulatory AgencyPhilips apologizes for any inconveniences caused by this problem |  | 506 | 506 | Worldwide Distribution - USA (nationwide) including Puerto Rico and the states AL, AR, AZ, CA, DC, DE, FL, GA, IA, IL, IN, KY, MA, MD, MN, MO, MS, NC, NV, NY, OH, OK, OR, PA, SC, TX, VT, WA, and WV., and the countries of Argentina, Australia, Austria, Azerbaijan, Bangladesh, Belgium, Brazil, Chile, China, Czech Republic, Denmark, Egypt, France, Germany, Greece, Hungary, Iceland, India, Indonesia, Iraq, Ireland, Israel, Italy, Japan, Jordan, Lebanon, Netherlands, New Zealand, Norway, Oman, Panama, Philippines, Poland, Qatar, Romania, Russia, Saudi Arabia, Singapore, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, and Uzbekistan. | Terminated | September 19, 2014 | 430 | 506 | Computer | Not_Computer
Z-0127-2015 | 69334 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | Dimension VistaÂż HbA1c |  |  | August 14, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc. | In March, Siemens Healthcare Diagnostics issued an Urgent Medical Device Correction for decreased recovery of patient samples with the Dimension Vista&reg; CSA &amp; CSAE Flex&reg; reagent cartridges when whole blood samples are run from Small Sample Containers (SSC). Siemens has confirmed that when whole blood samples are run from SSC, HbA1c values may exhibit a percent bias ranging from - 8.2% t | Device Design | An Urgent Medical Device Correction, dated August 2014, was provided to all Dimension VistaÂż HbA1c customers who received the affected lots to notify them of the issue, the potential risk to health, and actions to be taken by customers. |  | 20043 | 20043 | Worldwide Distribution: US (nationwide) and country of: Canada. | Terminated | August 10, 2015 | 361 | 20043 | Computer | Not_Computer
Z-0235-2015 | 69601 | No | KPQ | Radiation therapy simulation system. | Radiology | Radiology | 510(k) | Siemens SimView NT and SimView NT with IDI simulation system components: |  |  | October 10, 2014 | 2014 | Siemens Medical Solutions USA, Inc | Siemens became aware of a very rare occurrence where an intended movement of the SimView NT simulation system components could not be stopped with usual hand control commands or by use of the collision detection functionality.It is possible that simulation system components could collide with patient and/or system equipment. | Device Design | Customers were initially notified about the issue by letter on January 7, 2014 (this action was reported as Z-0786-2014 which has been terminated). Siemens promised to notify them when a fix became available. Customer letter, dated November 7, 2014, was sent to notify end users that a fix has become available and Siemens would contact them to schedule an appointment for the correction. |  | 6 | 6 | Worldwide Distribution - US Nationwide in the states of PA, KY, LA, MA, MI, and CA; and countries of: Argentina, Belarus, Bulgaria, Costa Rico, P.R. China, Curacao, Egypt, Repulic Korea, Germany, Greece, Hungary, Italy, Iraq, Jordan, Malaysia, Palestanian Ter, Poland, Thailand, Turkey, and Vietnam. | Terminated | June 29, 2017 | 993 | 6 | Computer | Computer
Z-0020-2015 | 68846 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Proteus 235 |  |  | March 27, 2014 | 2014 | Ion Beam Applications S.A. | While the Proteus 235 is in the pencil beam scanning (PBS) treatment mode, on rare occasions, Low Level Radio Frequency (LLRF) pulsed dee voltage (i.e., voltage applied between two electrodes in the cyclotron to accelerate protons) may generate a hot spot on a pencil beam position within a tumor, generating a dose that might exceed the prescribed dose on a single voxel. | Device Design | A User Notice was sent to direct accounts in/about March/April 2014 to notify them of the issue and provide instructions to avoid the identified potential risk to patients. |  | 8 | 8 | Worldwide Distribution-USA (nationwide) including the states of PA, NJ, IL, WA, and TN, and the countries of Germany, Italy, and Czech Republic. | Terminated | December 10, 2014 | 258 | 8 | Computer | Computer
Z-0216-2015 | 69487 | No | IOR | Mechanical wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Transport Chair and Rollator in 1, Model HE21630011 Burgundy and HE21630011 Blue. |  |  | October 10, 2014 | 2014 | Silver Star Brands | Silver Star Brands is recalling the Transport Chair and Rollator in One because they have received several reports of the wheel assembly failing resulting in damaged wheels.. | Device Design | Consignees were sent on 10/10/2014 an Easy Comforts "Urgent: Rollator Recall" letter dated October 10, 2014. The letter described the problem and the product involved in the recall. Advised consignees to stop using the product immediately. The consignees can contact the Product specialist at Easy Comforts at 1-888-327-2390 Monday though Friday from 7:30 am to 3:30 pm central time. |  | 129 | 129 | Nationwide Distribution including the states of AL, AZ, AR, CA, CO, CT, FL, GA, HI, IL, IN, IA, KS, LA, MD, MI, MO, MT, NE, NV,NH, NM, NY, NC, OH,OK, and PA. | Open, Classified |  |  | 129 | Computer | Not_Computer
Z-2708-2014 | 69191 | No | KPA | Automated slide stainer. | Pathology | Pathology | 510(K) Exempt | Dako Autostainer Link 48 with software version Dako Link 4.0.3, Model No. AS480 |  |  | August 29, 2014 | 2014 | Dako North America Inc. | Dako is recalling the Dako Link 4.0.3 software on the Dako Autostainer Link 48 because Dako has determined that an upgrade to this version of the software may cause premature rejection of reagent bottles or inadequate aspiration of reagent if the instrument has also had a preventative maintenance or waste pump replacement, and the baseplate is not completely aligned. | Software change control | Dako North America sent an Urgent Medical Device Correction letter to all customers who were upgraded to the Dako Link 4.0.3 software that is used on the Dako Autostainer Link 48 and who also had preventative maintenance or a waste pump replacement. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed that Dako will contact customers to set up an appointment for a service visit in order to ensure that the instruments are running safely or reliably. Customers with questions are instructed to contact Dako. |  | 191 | 191 units total (57 units in US) | Worldwide Distribution - US Nationwide in the states of AT, AU, BE, CA, CN, DE, DK, ES, Fl, FR, GB, IE, IT, NL, NO, SE, US, DZ, GR, IS, OM, SG, UY. | Terminated | March 03, 2015 | 186 | 191 | Computer | Not_Computer
Z-0857-2014 | 67199 | Yes | DWA | Cardiopulmonary bypass pump speed control. | Cardiovascular | Cardiovascular | 510(k) | CP5 centrifugal pump system |  |  | December 04, 2013 | 2013 | Sorin Group Deutschland GmbH | Sorin is initiating a field correction on S5/C5 Heart-lung machine and CP5 centrifugal pump system due to reports of pump speed control knob failures resulting in no response. | Device Design | Sarin Group US distributed a Field Safety Notice dated December 5, 2013 to all affected customers and distributors. The letter informed users of the potential for shaft encoder failures and notified them that they would be contacted by the Sorin Service Group to arrange an appointment to repair the units onsite. Customers and distributors were also provided with a Response form to confirm they had received, read and understood the Field Notice. They were instructed to return the completed form to assist in monitoring the effectiveness of the communication. For questions call 303-467-6527 or 303-467-6503. |  | 75 | 75 | Worldwide Distribution - US (Nationwide) and PR; and Internationally to:BE, ES, FI, GB, NL, SE, AR, AT, AU, CA, CH, CN, CR, EG, HK, ID, JO, JP, KR, KW, LB, LY, MM, PL, RO, RS, RU, SG. TH, TN, TR, TW, UA, VN, and ZA. | Terminated | December 28, 2015 | 754 | 398 | Computer | Not_Computer
Z-0611-2014 | 67108 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Puritan Bennett 840 ventilator systems software part number 4-070212-85, revision AB-AG. |  |  | December 16, 2013 | 2013 | Covidien | Covidien is conducting a voluntary field corrective action on Puritan Bennett 840 ventilator systems with certain software revisions in response to customer reports of ventilator diagnostic code XB0069 in which the device stops mechanical ventilation due to a software error. | Software design | Covidien sent an Urgent Medical Device Voluntary Field Correction letter to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to continue using their Puritan Bennett 840 ventilators until it is feasible to install the software update discussed in the letter. Covidien has developed a software solution to address this issue. Upon availability, there will be several ways to install the software update. To initiate the process for updating your software and to select the method that best suits your facility, please visit our software update management portal at www.PB840technicalupdate.com and follow the instructions to register your ventilators. If you need further assistance contact our Technical Support Department @ 1-800-255-6774, option 4, then option 1.For questions regarding this recall call 800-255-6774. |  | 56295 | 16,923 US, 39,372 OUS | Worldwide Distribution - USA (nationwide) and Internationally to Austria, Estonia, Belgium, Spain, Italy, Poland, Bulgaria, Finland, Portugal, Switzerland, France, Lithuania, Romania, Cyprus, Great Britain, Czech Republic, Greece, Latvia, Slovenia, Germany, Hungary, Slovak Republic, Ireland, Netherlands, Turkey, Croatia, Andorra, Albania, Algeria, Bahrain, Bosnia and Herzegovina, Egypt, Georgia, Iran, Iraq, Israel, Jordan, Kenya, Kuwait, Lebanon, Libya, Macedonia, Montenegro, Morocco, Russia, Saudi Arabia, Serbia, South Africa, Syria, Sudan, Tunisia, UAE, Uganda, Puerto Rico, Mexico, Columbia, Chile, Brazil, Argentina, Uruguay, Panama, Costa Rica, Peru, China, Korea, Japan, Taiwan, Hong Kong, Malaysia, Singapore, Sri Lanka, Australia, New Zealand. | Terminated | November 29, 2016 | 1079 | 56295 | Computer | Computer
Z-0060-2015 | 69207 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare Carescape Patient Data Module |  |  | September 19, 2014 | 2014 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue to the ECG calculations following a disconnect/reconnect cycle with the Patient Data Module, when used with the Carescape Bx50 monitors.If the PDM is quickly disconnected and reconnected (a less than 10 second cycle) from a Carescape Bx50 monitor following a Patient Discharge or during active monitoring, the ECG waveform an | Software design | Consignees were sent on 9/19/2014 a GE Healthcare "Urgent Medical Device Correction" letter Ref# 336107 dated September 18, 2014.The letter was addressed to Director of Biomedical Engineering, Chief of Nursing and Healthcare Administrator / Risk Manager. The letter described the Safety Issue, Safety Instructions , Affected Product Details, Product Correction and Contact Information. For questions they can contact Technical Support at 1-800-558-70444 or your local Service Representative. |  | 1691517372 | 68,452 (39,505 units US, 24,711 units OUS, 4,236 units unidentified/unknown) | Worldwide Distribution: US (nationwide) including DC, PR, Guantanamo except DE, HI, ND, SD &amp; WY; and countries of: SOUTH AFRICA, KOREA, UNITED KINGDOM, UNITED ARAB EMIRATES, TURKEY, TUNISIA, TRINIDAD TOBAGO, THAILAND, TANZANIA, TAIWAN, SWITZERLAND, SWEDEN, SPAIN, SLOVAKIA, SINGAPORE, SAUDI ARABIA, RUSSIA, QATAR, BRASIL, PORTUGAL, POLAND, PHILIPPINES, PERU, PANAMA, PARAGUAY, OMAN, NEW ZEALAND, NORWAY, NETHERLAND, NEPAL, MYANMAR, MEXICO, MOROCCO, MALDOVA, MAURITIUS, MALAYSIA, LITHUANIA, LEBANON, BAHRAIN, KUWAIT, REPUBLIC OF KOREA, JAPAN, JORDAN, JAMAICA, ITALY, ISRAEL, IRELAND, INDONESIA, INDIA, ICELAND, HUNGARY, HONG KONG, GUATEMALA, GREECE, GERMANY, FRANCE, FINLAND, ESTONIA, EGYPT,ECUADOR, DENMARK, CECH REPUBLIC, COSTA RICA, COLOMBIA,, CHINA, CANADA, BULGARIA, BELGIUM BANGLADESH, AUSTRIA, AUSTRALIA, and ARGENTINA. | Open, Classified |  |  | 1691517372 | Computer | Computer
Z-0823-2014 | 66694 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens CentraLink" Data Management System V14.0.4, V14.0.5 or V14.0.8 |  |  | October 30, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Under extremely rare circumstances, a patient result that had been previously disabled may be released without proper review. (Disabled Test in CentraLink System or Aptio System May Become Enabled in CentraLink System.) | Software design | The firm, Siemens, sent an "Urgent Medical Device Correction" letter dated October 23, 2013 to its customers. The firm also sent a follow-up letter dated October 30, 2013. The letter describes the product, problem and actions to be taken. The customers were instructed Until a service visit is scheduled, do the following: Instruct the laboratory staff to recognize this sequence of events and Ensure that whenever the operator disables a method on an Aptio system or a CentraLinksystem and replenishes reagent for that method, the operator then confirms that the method remains disabled in the CentraLink system or disables it again. The operator must allow up to five minutes to verify that the test remains disabled. In addition, the customers were instructed tocomplete and return the FIELD CORRECTION EFFECTIVENESS CHECK response form via fax to Siemens Healthcare Diagnostics Customer Care Center at 800-441-9250, and to retain this letter with your laboratory records, and forward this letter to those who may have received this product. After a service visit is performed and an interim solution is implemented, the laboratory staff will be trained on this update to understand the minor but necessary changes in the workflow.If you have any questions, please contact your Siemens Customer Care Center at 800-441-9250 or your local Siemens representative. |  | 40 | 40 | US Distribution in the states of: MN, NJ, and NV. | Terminated | December 15, 2014 | 411 | 40 | Computer | Computer
Z-0485-2015 | 69728 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Esprit V1000 Ventilator and Esprit/V200 Conversion Ventilator |  |  | October 29, 2014 | 2014 | Respironics California Inc | If the power supply fan mounting screws are installed incorrectly, the ends of the screws may touch the power supply printed circuit board assembly (PCBA) and short to ground. This may prevent the ventilator to switch to battery power if AC power is disrupted. | Device Design | Philips Healthcare sent an Urgent - Field Safety Notice to all affected consignees to inform them that Respironics is recalling the Esprit V1000 and Esprit/V200 Conversion ventilators because if the power supply fan mounting screws are installed incorrectly, the ends of the screws may touch the power supply printed circuit board assembly (PCBA) and short to ground. This may prevent the ventilator to switch to battery power if AC power is disrupted. A high urgency audible alarm will sound alerting medical personnel of the loss of power. In the absence of AC power, failure of battery power will cause loss of ventilator support, which may result in hypercarbia or hypoxemia. The potential issue affects ventilators manufactured prior to 12/31/04 due to design of the fan/shroud assembly. Esprit V1000 and Esprit/V200 Conversion Ventilators manufactured after that affected date are not affected. Customers with concerns were instructed to contact a Philips Respironics representative at (800) 722-9377. For questions regarding this recall call 760-918-7300. |  | 3943 | 3,943 units total (3,071 units in US) | Worldwide Distribution - USA (nationwide) and Internationally to Argentina, Australia, Brazil, Bulgaria, Canada, Chile, China, Colombia, Ecuador, Egypt, Greece, Hong Kong, India, Italy, Korea, Lebanon, Malaysia, Mexico, Pakistan, Panama, Peru, Philippines, Saudi Arabia, Singapore, South Africa, Spain, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Venezuela, Vietnam, and Yemen. | Open, Classified |  |  | 3943 | Computer | Computer
Z-0782-2014 | 67123 | No | JEQ | Powered toothbrush. | Dental | Dental | 510(K) Exempt | UP &amp; UP kids' power toothbrush |  |  | December 18, 2013 | 2013 | Ranir Corporation | Ranir, LLC is recalling all lots of UP &amp; UP brand Kid's Power Toothbrush and TopCare brand Clifford Kids Power Toothbrush because the tufted disk may become detached from the power brush handle. If thetuft retention disk detaches from the handle, it is a potential choking hazard for children three years old and younger. | Device Design | Ranir, LLC phoned their two direct accounts notifying them of this recall on December 12, 2013. On December 18, 2013, Ranir, LLC sent an URGENT: MEDICAL DEVICE RECALL notification letter was sent to their direct accounts. The notification stated the affected product, issue, and instructions for responding to the notification, including instructions for return of the affected product. Questions are directed to Paula Bojsen (616)698-8880. |  | 601616 | 601,616 total | Distributed in MN and IL. | Terminated | August 20, 2015 | 610 | 601616 | Computer | Not_Computer
Z-0483-2015 | 69627 | No | BZD | Noncontinuous ventilator (IPPB). | Anesthesiology | Anesthesiology | 510(k) | Cannulaide, Model 101, 102, 103. Respirator. |  |  | May 13, 2008 | 2008 | Beevers Manufacturing &amp; Supply, Inc. | Cannulaide recalled May 2008 due to improper seal of the pouch. The improper seal of the Cannulaide's package could potentially affect the device shelf life. | Equipment maintenance | Beevers Manufacturing &amp; Supply, Inc. (BMS) did contact the customers via letter and phone call in May 2008. Customers that didn't reply were further contacted via email on June 2008. Salter Labs provided BMS' recall report to FDA in October 2014.Salter Labs was unable to locate a recall letter issued by BMS to the affected customers. For further questions please call (503) 472-9055 |  | 16575 | 16,575 units total (15,625 units in US and 950 units outside US) | Worldwide Distribution - US Distribution including the states of LA, GA, OR, MA, NJ, FL and CA., and the countries of Canada, France, Japan, Netherlands and United Kingdom. | Terminated | November 28, 2014 | 2390 | 16575 | Computer | Not_Computer
Z-2693-2014 | 68594 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Minstrel (with scale) |  |  | July 28, 2014 | 2014 | ARJOHUNTLEIGH POLSKA Sp. z.o.o. | ArjoHuntleigh received three reports where the bolt under the scale unscrewed from the scale. There is a remote risk concerning the bolt unscrewing from the scale. When the scale and bolt parts are not correctly serviced, or when they receive a sharp blow during or outside of use, it is possible the bolt holding the spreader bar to the scale could unscrew during subsequent use. As a result, the sp | Device Design | ARJOHUNTLEIGH sent an URGENT - FIELD SAFETY NOTICE letter dated July 28, 2014 to all customers. The letter identified the affected product, problem and actions to be taken. The letters provided customers with two options with how to address this recall: 1) because the devices are close to their 10-year lifetime, customers can return the devices for a trade-in discount that will be applied to the purchase of a new ArjoHuntleigh patient transfer lift; or 2) an ArjoHuntleigh Service Technician will install a new scale with added locking features to the existing unit free of charge. Customers were further instructed to: 1) ensure that all caregivers and users of the devices are given a copy of the letter; 2) place a copy of this letter with the device's instruction for use; and, 3) to complete and return the enclosed Customer Service Response form. For questions or assistance completing the Customer Response Form, please contact ArjoHuntleigh at 1-800-323-1245, ext. 57985 during office hours; Monday to Friday, 8:00am to 5:00pm (CST) or via email at AHNA-RA@arjohuntleigh.com. |  | 18 | 18 Lifts | Worldwide Distribution - US Nationwide in the state of MA and the countries of Australia, Germany, and Italy. | Open, Classified |  |  | 18 | Computer | Not_Computer
Z-0063-2015 | 69270 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | EndoWrist One Vessel Sealer used in conjunction with the da Vinci Si Surgical System IS3000. |  |  | September 17, 2014 | 2014 | Intuitive Surgical, Inc. | Intuitive Surgical is initiating a voluntary correction relating to the labeling and software associated with the EndoWrist. One Vessel Sealer for the da Vinci Si (IS3000) Surgical System. | Device Design | Urgent medical device correction letter was sent to all consignees on 9/22/14 by trackable mail system.Letters identified the issue and provides information for users should a error code be displayed. |  | 1483 | 1,483 devices | Worldwide Distribution-USA (nationwide) and the countries of Australia, Austria, Belgium, China, Chile, Colombia, Czech Republic, Denmark, Finland, France, Germany, Greece, Italy, Japan, Netherlands, Norway, South Korea, Spain, Sweden, Switzerland, Turkey, and United Kingdom. | Terminated | June 01, 2015 | 257 | 1483 | Computer | Not_Computer
Z-0146-2015 | 69323 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | ACCU-CHEK Spirit Combo insulin infusion pump sold as part of Accu-Chek Combo System |  |  | September 25, 2014 | 2014 | Roche Diagnostics Operations, Inc. | The insulin pump may lose time and date settings during a power interruption (such as a battery change) due to a faulty capacitor. If the capacitor fails, the time and date will return to default setting. Though the pump prompts customers to confirm the time and date, if the user does not recognize the time and date have returned to default, a shift of their basal rate time block could occur. | Equipment maintenance | Roche sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated September 25, 2014 via UPS Ground (receipt required) to the consignees. The letter included a description of the reason for recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. No product returns were expected. For questions, contact the ACCU-CHEK Customer Care Service Center at 1-800-688-4578. |  | 8116 | 8116 | US Nationwide Distribution in the states of CA, CT, IA, NV, OH, IN, WI, GA, MO, NY, VA, LA, ME, TX, NC, ND, SD, NJ, CO, AL, SC, IL, KY, FL, OR, MN, MD, MS, AR, WV, WA, TN, CT, ID, KS, MI, PA, UT, MA, RI, NM, OK, DE, AZ, NE, MT, NV, WY, AK, and NH. | Open, Classified |  |  | 8116 | Computer | Computer
Z-0062-2015 | 69294 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | MaxLock Extreme Mod-Foot Concave Reamer; Individually packaged in a 3 x 4 bag and heat sealed with label on the outside. |  |  | September 19, 2014 | 2014 | Orthohelix Surgical Designs Inc | Incorrect sizing was being laser etched on the MaxLock Extreme Mod-Foot Concave Reamer | Device Design | An Urgent - Medical Device Recall customer notification, dated September 19, 2014, was sent out via email and FEDEx to all distributors with instructions on returning the product. The letter identified the affected product and the reason for the recall. Customers are to immediately cease distribution of the affected product, and complete and return the questionnaire. The letter also states that the recalled product will be replaced. Questions are to be directed to the Customer Service Department 1-888-494-7950. |  | 38 | 38 | USA Distribution -- in the states of WI, NM, TX, WA, MO, OH, UT, TN, MI, VA, NC, VA, and PA. | Terminated | October 31, 2016 | 773 | 38 | Computer | Not_Computer
Z-0099-2015 | 66541 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | CRN Crono Syringe CRN20 with 19G x 1 "Steel Needle" |  |  | October 18, 2013 | 2013 | Intra Pump Infusion Systems | Product has the potential to have a clear debris on the tip of the needle. | Device Design | Cane S.p.A. Medical Technology sent an Urgent: Correction and Removal Action Required letter dated October 21, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Actions to be taken:a. Stop using the 19G x 11/2" 1.10 x 40mm steelneedle included in the CRONO Syringe blister pack. B. Safely discard the 19G x 11/2" 1.10 x 40mm steel needle into a Sharps container.c. Use a TRANSFER DEVICE (e.g.,Needleless adapter with female-to-female luer connector) or alternative device to transfer medications into the CRONO syringe.d.For intravenous infusions,we are recommending that a filter also be used during infusions.e. Complete the required 'Acknowledgment and Receipt Form'and fax back to our Distributor;Intra Pump Infusion Systems at 630-845-2768.Contact Information for our US Distributor: INTRA PUMP INFUSION SYSTEMS " Monday through Friday 8:00am to 5:00pm CST,TollFree:1-866-211-7867 " Direct:630-845-7867 " Fax:630-845-2768 www.lntrapump.com |  | 25440 | 25,440 pieces | Worldwide Distribution: US (nationwide) in states of CA, PA, MA, IL, OH and NY; and Internationally to Canada. | Terminated | February 01, 2016 | 836 | 25440 | Computer | Not_Computer
Z-0142-2015 | 64774 | walker, mechanical | ITJ | Mechanical walker. | Physical Medicine | Physical Medicine | 510(K) Exempt | DMI Duro-Med Industries Rollator Transport Chair; Wheel Size: 7 1/4"; Overall Width: 23 1/2"; Seat Height: 21 1/4"; Seat Size: 14" x 12 3/4"; Handle Height: 31" - 36"; Width Between Handles: 18"; Weight: 19 Lbs; |  |  | June 05, 2014 | 2014 | Briggs Medical Service Company d.b.a. Mabis | There have been 5 customer reports of a 501-1018 series DMI Rollator Transport Chair collapsing during use. | Device Design | The firm, Briggs Healthcare, sent "URGENT PRODUCT RECALL" letters dated July 28, 2014 to all distributors, retailers and consumers. The letters included instructions for distributors and dealers to: 1) examine their inventory; 2) quarantine any affected products; 3) cease distribution of the affected products; and, 4) complete and return the attached Customer Recall Acknowledgement Form. Distributors and dealers who further distributed the products were also instructed to notify their customers of this recall by forwarding a copy of the enclosed consumer notification letter and its associated documents. Distributors and dealers with questions about this recall can contact Briggs at 800-342-4426 (Monday - Friday, 8:00 AM - 6:00 PM EDT) or by e-mail at RecallCoord@briggscorp.com. The consumer-directed letter dated July 28, 2014 included instructions for consumers to examine their rollator transport chair to determine whether or not it is subject to the recall. If the consumer has the affected product, they should discontinue using it immediately and contact Briggs at 800--342-4426 (Monday - Friday, 8:00 AM - 6:00 PM EDT) or by e-mail at RecallCoord@briggscorp.com to arrange for the return of the product. Consumers will be given the option of receiving a cash voucher or a replacement product in exchange for the recalled product. |  | 1869 | 1,869 units | Worldwide Distribution: US (nationwide) including states of:US: Nationwide (AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY) + District of Columbia + Puerto Rico; and country of : Canada.); FOREIGN: Canada | Open, Classified |  |  | 1869 | Computer | Not_Computer
Z-2709-2014 | 69123 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance CT System 6/10/16/16P/40 configuration. No packaging, sold as single unit. |  |  | December 27, 2005 | 2005 | Philips Medical Systems (Cleveland) Inc | The system has a gap that can appear between the upper and lower tilt stand covers. The fingers may be squeezed if a person was to place them in this region while the gantry was tilting from an angle to zero tilt angle. | Device Design | Customers will be contacted by a field service technician with the field correction order. The firm will install the new labeling to the device. |  | 364 | 364 | Worldwide Distribution -- AZ, CO, GA, IA, IL, IN, KS, KY, MD, MI, MN, MS, NC, NY, OH, OK, TX, VT, and WI; and, countries of Singapore, New Zealand, India, Australia, Korea, Japan, Korea, China, India, Taiwan, Australia, Turkey, France, South Africa, Spain, Italy, Austria, Belgium, Switzerland, Israel, Netherlands, England, and Brazil. | Terminated | November 06, 2015 | 3601 | 364 | Computer | Not_Computer
Z-0112-2015 | 69320 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Covidien Puritan Bennett 980 Ventilator, Rx ONLY. |  |  | October 01, 2014 | 2014 | Nellcor Puritan Bennett Inc. (dba Covidien LP) | A software issue may lead to ventilator inoperative situations. | Software design | The firm sent Urgent: Field Correction Action letters, dated October 3, 2014, to consignees. The letter identified the affected product and the reason for the recall. Customers were provided actions to be taken, as well as, important safety reminders. The letter informed customers that they may continue to use their ventilators pending the software correction as long as two gas sources are connected to the ventilator at all times. Customers are to complete the attached acknowledgement and receipt form and fax it to Covidien at the number provided. For further assistance, customers are to contact the Technical Support Department at 1-800-255-6774. |  | 324 | 324 units | Worldwide Distribution -- USA, including the states of CA, CO, FL, GA, KY, MA, MN, NC, NY, OH, OK, PA, SC, TN, TX, UT, WI; and, the countries of Canada, Mexico, Saudi Arabia, South Africa, and United Arab Emirates. | Terminated | May 02, 2016 | 579 | 324 | Computer | Computer
Z-2711-2014 | 69235 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | The Ingenuity CT, Ingenuity Core, and Ingenuity Core128 scanners are Computed Tomography X-ray systems intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. This device may include signal analysis and display equipment, patient and equipment supports, components and accessories |  |  | September 11, 2014 | 2014 | Philips Medical Systems (Cleveland) Inc | Ingenuity CT system can mislabel the exposure that results in an image in the incorrect order that it was taken by the user. | Software design | Philips sent an Urgent - Field Safety Notice Medical Device Correction letter dated September 5, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Please review the following information with all members of your staff who need to be aware of the contents of this communication. It is important to understand the implications of this communication. Please retain a copy with the equipment Instruction for Use.Additionally, a Field Service Engineer will contact each site to schedule a time to install the software update.If you need any further information or support concerning this issue, please contact your local Philips representative or local Philips Healthcare office. For North America, contact the Customer Care Solutions Center(1.800.722-9377 and follow the prompts). |  | 79 | 79 | Worldwide Distribution - US including the states of : AZ, CA, CT, FL, GA,MA, MS,OH, PA, TX, and WY., and the countries of : Australia, Denmark, Finland, France, Germany, India, Ireland, Japan, Malaysia, Netherlands, New Zealand, Norway, Russia, Slovakia, South Africa, South Korea, Sweden, Switzerland and United Kingdom. | Open, Classified |  |  | 79 | Computer | Computer
Z-0486-2015 | 69602 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | The Artis zee / zeego Angiography System is designed as a set of components that |  |  | September 11, 2013 | 2013 | Siemens Medical Solutions USA, Inc | There is a potential issue on running Artis systems running software VC1x software if a network problem arises, the function cannot be deactivated again by pressing the &quot;Block Radiation&quot; key on the touch screen control. If this behavior occurs, the system image generation function is not available for patient examinations without switching the system off and back on again manually. This | Device Design | Siemens sent an Important Customer Safety Notice in September 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. If the fault described above occurs, we urgently recommend switching the system OFF manually ("System shutdown") and then back ON again. The fault cannot generally be rectified sufficiently, either by executing the "System Restart" or "Application Restart" functions. A corresponding supplement to the Operator Manual is enclosed with this letter.The problem will be solved when the next Service Update is available (currently scheduled for 2014).We thank you for your cooperation in dealing with this safety notice and request that you promptly notify and instruct accordingly all the staff at your organization who have to be aware of this problem. Please also forward this safety information to other organizations that could be affected by this measure. Please observe this safety notice, and comply with the corresponding measures until the update has been fully completed.If this device has been sold and is therefore no longer in your possession, please forward this safety notice to the new owner. We would also request that you inform us of the identity of the device's new owner where possible.For further questions please call (610) 219-6300. |  | 1131 | 1131 | Worldwide Distribution - USA (nationwide) including Puerto Rico., and the countries of: United Arab Emirates, Armenia, Austria, Australia, Belgium, Bahrein ,Belarus, Switzerland, Chile, China, Colombia, Czech Republic,Germany, Denmark, Egypt, Spain, Finland, France, United Kingdom, Hong Kong, Hungary, Indonesia, Ireland, Israel, India, Italy, Jordan, Japan, Kenya, Korea, Kazakhstan, Malaysia, Netherlands, Norway, New Zealand, Philippines, Pakistan, Poland, Palestina, Portugal, Serbia, Russian Federation, Saudi Arabia, Sweden, Singapore, Slovakia, Thailand, Turkey, Taiwan, Ukraine, Vietnam and South Africa. | Terminated | March 30, 2015 | 565 | 1131 | Computer | Computer
Z-0825-2014 | 67226 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON SC2000 |  |  | December 22, 2013 | 2013 | Siemens Medical Solutions USA, Inc. | The locking mechanism that is intended to hold the control panel in a fixed position can become loose and fail to lock the rotation of the control panel, making it unable to steer or control during transport. | Device Design | Customer Safety Advisory Letters were sent to all affected users on December 22, 2013. Letters advised of the issues and the actions they should take until a Siemens Field Engineer will visit and upgrade the control panel. |  | 271 | 271 units | Worldwide Distribution-USA (nationwide) and the countries of Australia, Austria, Canada, Ecuador, Finland, France, Germany, India, Ireland, Israel, Italy, Japan, Mexico, P.R. China, Philippines, Portugal, Republic Korea, Russian Fed. Saudi Arabia, Serbia, South Africa, Spain, Suriname, Sweden, Taiwan, Turkey, U.A.E., and United Kingdom. | Terminated | March 05, 2015 | 438 | 271 | Computer | Not_Computer
Z-2717-2014 | 68329 | No | GFD | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | CelluTome Epidermal Harvesting System Part number CT-H25 or Part Number CT-H50 |  |  | May 19, 2014 | 2014 | KCI USA, Inc. | Blade can shift during shipping which makes the product ineffective. | Device Design | KCI sent an Urgent - Voluntary Medical Device Field Notification letter dated June 6, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The firm is sending Field Notifications to affected customers instructing them to complete appropriate actions.Option I: Contact your local KCI Representative or KCI Customer Technical Support at 1-800-275-4524 and select Option 3 to arrange for return and replacement of the affected harvester(s).Option 2: If you prefer, you may re-set the blade after removal from the packaging by gently tapping the curved end of the harvester against the palm of your hand, as shown in Figure 3, until the blade is no longer visible in the harvester top plate, as shown in Figure 4. |  | 1331 | 1,331 units | US nationwide Distribution | Terminated | September 18, 2015 | 487 | 1331 | Computer | Not_Computer
Z-0143-2015 | 69406 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Merge Healthcare, Merge Hemo Programmable diagnostic computer. |  |  | September 29, 2014 | 2014 | Merge Healthcare, Inc. | It has been reported that during use, the SpO2 value displayed on the Hemo Monitor may not update to reflect changes in the patient's oxygen value. It is also possible that if using the pulse rate of the SpO2 finger clip to calculate the patient's heart rate, it too may be subject to displaying a stale value. | Software design | Consignees were e mailed on 5/29/14 a Merge "Important Alert" letter. The letter described the problem and the product involved in the recall. Informed consignees that they had the opportunity to replace the Nellcor SpO2 module with a Masimo SpO2 module free of charge. The letter also provided Clinical Information and the Next Steps to follow. This included to contact Merge Technical Support if they wanted to replace their Nellcor module. For questions, they can contact Lee Harrop, Director Solutions Management, Cardiology, 386-239-2224, Lee.Harrop@Merge.com.A second letter was sent Sept. 11, 2014 via e-mail. Consignees were made aware of a firmware patch available. |  | 52 | 52 | Nationwide Distribution including the states of AZ, CA, FL, IL, IA, MA, MO, OH, SC, TN, and WI. | Open, Classified |  |  | 52 | Computer | Computer
Z-0464-2015 | 69515 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | AlignRT- Intended for prescription use. The system is indicated for use during simulation, setup and stereotactic radiosurgery and precision radiotherapy for lesions, tumors and conditions anywhere in the body where radiation is indicated. |  |  | October 15, 2014 | 2014 | Vision Rt Inc | Potential failure of AlignRT to assert interlock. | Software in the Use Environment | The firm, visionrt. sent an "URGENT MEDICAL DEVICE CORRECTION NOTIFICATION-URGENT FIELD SAFETY NOTICE" dated October 15, 2014 to all affected customers on 21-Oct-2014. The notice describes the product, problem and actions to be taken. The customers were informed that they did not need to return their devices; they should pass the letter to all those who need to be aware of it within their organization; promptly inform Vision RT if they believe that patient harm occurred due this issue, and complete and return the acknowledgement form via email to service@visionrt.com.Vision RT is designing an upgrade that will resolve the issue, and will subsequently contact the customer in order to install the upgrade on their device.Should you have any queries on this letter, please do not hesitate to contact Vision RT by telephone on +44 20 83464300 (866 778-2379 from the US) or as per http://www.visionrt.com/contact/details. |  | 7 | 7 (Six Units distribited in the US and One in OUS.) | Worldwide Distribution: US distribution in states of: CA, MN, TX, and UT; and internationally to: Republic of Ireland. | Open, Classified |  |  | 7 | Computer | Not_Computer
Z-0218-2015 | 69645 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | CyberKnife Robotic Radiosurgery System with the first generation IRIS Variable Aperture Collimator. |  |  | October 31, 2014 | 2014 | Accuray Incorporated | Software upgrade to correct potential safety issue related to CyberKnife System that occurs when upgrading the Treatment Delivery Software for the first generation Iris Variable Aperture Collimator. | Software design | Accuray sent an Urgent Safety Advisory Notification letter dated October 17, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check the radiaiton field size (beam width) for each Iris setting to ensure that it matches the beam commissioning data in the MultiPlan System to within the stated tolerance of 0.2 mm. Customers with questions were instructed to contact Accuray Customer Support at 1-877-668-8667 or visit the Accuray Technical Solutions Center at www.accuray.com or call North America: +1.877.668.8667, Europe: +800.4141.9595, Hong Kong +852.2247.8688, Japan +81.3.6269.9556.For questions regarding this recall call 408-716-4600. |  | 84 | 84 devices subject to correction. | Worldwide Distribution - USA including CA, MS, FL, AZ, DC, TX, CO, CT, IL, WI, RI, NV, OK, MA, AL, NY, MO, NC, TN, PA, DE, NJ, MI, IN, KS, OH. Foreign distribution to Turkey, Switzerland, Spain, Saudi Arabia, Russia, Poland, Mexico and Korea, Italy, India, Greece, Germany, France, England, Canada, and Belgium. | Terminated | February 20, 2015 | 112 | 84 | Computer | Computer
Z-2715-2014 | 69084 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ADVIA Centaur Immunoassay System (including refurbished units), |  |  | August 19, 2014 | 2014 | Siemens Healthcare Diagnostics | The firm has identified an issue with patient demographic information sent to the LIS from the ADVIA Centaur/ADVIA Centaur XP Immunoassay systems. Siemens has confirmed that under extremely rare circumstances patient demographic data from the previous order received from the LIS is merged with the next order. This issue can occur when the LIS data buffer on the ADVIA Centaur system becomes full | Software design | Siemens sent an Urgent Medical Device Correction Letter dated August 19, 2014, via Fed Ex to US customers and via email to regional countries for implementation outside the US (OUS). The letter identified the product the problem and the action needed to be taken by the customer.The letters inform the customers of the issue and specific actions to be taken if the instrument is interfaced to an LIS system that transmits patient demographics with each order. Field service personnel were sent a support bulletin describing the issue and instructing them how to deal with customer questions. Please complete and return the Field Correction Effectiveness Check attached to this letter within 30 days. If you have received any complaints of illness or adverse events associated with the products listed in Table 1, immediately contact your local Siemens Customer Care Center or your local Siemens technical support representative.Please fax this completed form to the Customer Care Center at (302) 631-7597. If you have any questions, contact your local Siemens technical support representative |  | 10515 | Domestic: 2771 units; Foreign: 7744 units | Worldwide Distribution including the United States., and the countries of Argentina, Austria, Australia, Bahrain, Bangladesh, Belgium, Belgrade, Brazil, Botswana, Canada, Chile, China, Colombia, Croatia, Czech Republic, Cyprus, Denmark, Egypt, Estonia, Finland, France, Germany, Georgia, Great Britain, Hungary, Indonesia, India, Ireland, Israel, Italy, Japan, Jordan, Korea, Kazakhstan, Lebanon, Libya, Lithuania, Luxembourg, Macedonia, Malaysia, Malta, Martinique, Mexico, Myanmar, New Caledonia, Netherlands, Norway, New Zealand, Pakistan, French Polynesia, Philippines, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovenia, Slovakia, South Africa, Spain, Switzerland, Thailand, Turkey, United Arab Emirates, Vatican, Venezuela, Vietnam, Yemen and Zimbabwe. | Open, Classified |  |  | 10515 | Computer | Computer
Z-2714-2014 | 69194 | No | CAI | Anesthesia breathing circuit. | Anesthesiology | Anesthesiology | 510(K) Exempt | Hudson RCIÂż Pediatric Anesthesia Breathing Circuits, Single Patient Use, Rx Only, |  |  | September 08, 2014 | 2014 | Teleflex Medical | The pediatric breathing circuits can crack prior to and during use. | Device Design | Teleflex Medical sent an Urgent Medical Device Recall dated September 8, 2014, to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to immediately discontinue use, quarantine any products with the product codes and lot numbers provided in the notification, return any affected product, complete and return the enclosed Recall Acknowledgement Form and fax it to the attention of Customer Service. For questions contact your local sales representative or Customer Service at 1-866-246-6990. |  | 27176 | 27,176 units | Worldwide Distribution - US Nationwide in the states of AZ, CA, FL, IL, KY, OH, NJ, NY, TX, LA, OR, SD, WA, including Puerto Rico, and the countries of Belgium, Canada, Chile, China, Colombia, Korea, Mexico, Peru, Venezuela and Zaire. | Terminated | August 16, 2016 | 708 | 27176 | Computer | Not_Computer
Z-0651-2014 | 66603 | No | MUH | Extraoral source x-ray system. | Dental | Radiology | 510(k) | The Vantage Digital Panoramic System delivers high-value features standard on every system: - Five pre-programmed exam settings, including bitewing, minimize exam set-up time and allow routine diagnostic exams to be performed extraorally for improved clinical efficiency and patient experience. - Progeny Clarity Enhanced Filter Set provides superior diagnostic imaging quality - Exclusive Focal Trough Alignment reminder assists in consistent patient positioning - Progeny Imaging software included for download on an unlimited number of operatory workstations with no additional costs - Exclusive VantageTrust remote support and training service offered at no charge for the first 60 days of use - Ceph upgradeable for expanded extraoral exam capabilities - The Progeny Vantage Panoramic System s proudly designed and manufactured in the USA; 2013 Midmark Corporation; 675 Heathrow Drive Lincolnshire, IL 60069; 847-415-9800 Toll-free 888-924-3800 |  |  | September 20, 2013 | 2013 | Midmark Corp dba Progeny Inc | An error was identified in the Real Time Controller (RTC) firmware versions 3.0, 3.1, 3.2, and 3.3 used in Vantage Panoramic X-Ray systems. The error may cause the column of the Vantage Panoramic X-Ray unit to continue moving vertically despite the release of the movement control button (i.e. failure to stop"). | Software design | An URGENT MEDICAL DEVICE FIELD CORRECTION letter dated September 20, 2013 was sent to all consignees. Letters sent to distributors include instructions on how to identify the defective units, to quarantine the machine(s), and to contact Midmark at 847-415-9739 or via- e-mail at abavuso@midmark.com. Midmark will either schedule an on-site field correction or arrange for the replacement of the affected device(s) depending on the customer's preference. Midmark further requested the distributors provide Midmark with the names, addresses, phone numbers, and e-mail addresses of any customers that purchased any defective units. Field corrections include the upgrading of RTC firmware to revision 3.4 on the defective units and will be performed as no additional charge by Midmark personnel. Letter sent to practitioners include instructions on how to identify the defective units, instructions on how to temporarily reduce the likelihood of the unintentional movement, and to contact Midmark at 847-415-9739 or via- e-mail at abavuso@midmark.com. Midmark will schedule an on-site field correction by upgrading the firmware to revision 3.4. |  | 58 | 58 units | Worldwide Distribution-USA (nationwide) including Puerto Rico and the states of AK, AL, CA, CO, CT, FL, GA, IL, KS, MA, NC, NH, NJ, OH, OK, PA, SC, TN, TX, VA, and WI, and the countries of Australia, Canada, and Russia. | Open, Classified |  |  | 58 | Computer | Computer
Z-0647-2014 | 67132 | No | CAW | Portable oxygen generator. | Anesthesiology | Anesthesiology | 510(k) | On Site Gas Systems POGS 33C Portable Oxygen Generation System |  |  | December 19, 2013 | 2013 | On Site Gas Systems Inc. | Heat exchanger hose becomes weak and may fail. | Device Design | On Site Gas Systems issued letter on 12/19/13 and stated replacement parts to retrofit the blue Federal Heater hose with stainless steel will be provided and the letter included the Technical Bulletin TB 12-19-13-1 . |  | 198 | 198 units | US Distribution: PA only. | Terminated | October 02, 2014 | 287 | 198 | Computer | Not_Computer
Z-0840-2014 | 66917 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | GE Optima CT520 and Optima CT540 Computed Tomography (CT) systems. |  |  | October 09, 2013 | 2013 | GE Healthcare, LLC | There is an issue with the Manual Film Composer feature on some CT products. There is an opportunity, while following a specific workflow, to create a film image with one patient's images and another patient's name in the footer. This is also an issue when within Manual Film Composer, if a color image is selected for printing, another patient's black and white image appears in the preview screen | Software design | Consignees were sent on 10/9/2013 a GE Healthcare "Urgent Medical Device Correction" letter dated October 9, 2013. The letter was addressed to Hospital Administrators / Risk Managers, Radiology Directors and Radiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information (Contact your local GE Healthcare Service Representative or the GE Healthcare Call Center : US or Canada : 800-437-1171 and Other countries: 262-896-2890). |  | 272 | 272 (29 US; 243 OUS) | Worldwide Distribution-USA (nationwide) including the states of AL, CO, FL, IN, KS, KY, MN, NC, OH, TX, and VA, and the countries of ALGERIA, AUSTRALIA, BELGIUM, BULGARIA, CANADA, CHINA, ABIDJAN COTEDIVOIRE EGYPT, FRANCE, GERMANY, GREAT BRITAIN, INDIA, IRAQ, ITALY, JAPAN, KENYA, LATVIA, LEBANON, MALAYSIA, MARTINIQUE,MOROCCO, NEW ZEALAND, NORWAY, PAKISTAN, POLAND, PORTUGAL ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SOUTH AFRICA, SPAIN, SWITZERLAND, THAILAND, TUNISIA, TURKEY, UNITED ARAB EMIRATES, and VIETNAM. | Terminated | June 19, 2014 | 253 | 272 | Computer | Computer
Z-2737-2014 | 69039 | No | GFC | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Persona The Personalized Knee System 2.5 mm Female Screw 48 mm Length Sterile Qty-2. Sterilized using irradiation, single use. |  |  | September 09, 2014 | 2014 | Zimmer, Inc. | The Persona 48mm X 2.5mm Female Screw is being recalled due to potential use of the screw in cortical bone, and corresponding stripping of the 2.5mm Persona Hex Driver. | Device Design | Zimmer sent an URGENT MEDICAL DEVICE RECALL notifications dated September 9, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.1. Review the notification and the updated portion of the Persona surgical technique provided in Attachment 1, and ensure affected personnel are aware of the contents.2. Locate and quarantine any affected product listed above and notify your Zimmer sales representative.3. Your Zimmer sales representative will remove the recalled product from your facility.4. If after reviewing this notification you have further questions or concerns please call the customer call center at1-877-946-2761 between 8:00 am and 5:00pm EST. |  | 16380 | 16,380 units | Nationwide Distribution including the states of : HI, CA, UT, NY, OH, IN, AL, FL, GA, WI, IA, MN, NE, SD, NJ, PA, ND, IL, MO, VA, MD, NC, MI, OR, WA, KS, KY, TN, WV, OK, CT, MA, ME, AZ, NM, TX, LA, ID, CO, MT, UT, NV and AK., and Internationally to Singapore, Japan, Germany and Canada. | Terminated | May 05, 2015 | 238 | 16380 | Computer | Not_Computer
Z-0141-2015 | 69426 | No | JEA | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | TITAN Operating Room Table |  |  | October 09, 2014 | 2014 | Trumpf Medical Systems, Inc. | Unintended movement of the operating table during surgical cases due to the use of damaged TITAN remotes. | Device Design | TRUMPF sent an Urgent Medical Device Correction letter dated October 3, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Actions to be taken:Complete and return the attached Confirmation of Receipt form with serial numbers of the TITAN table(s) and the serial number(s) for any remote controls you have. We will then arrange to send replacement remotes, which have been redesigned to eliminate this mode of failure. You will then be required to dispose of your recalled remotes.If you have any questions concerning this request or the procedure as outlined, please contact TRUMPF Medical Systems, Inc. at 888-474-9359. |  | 184 | 184 units | Worldwide Disttribution: US (nationwide) including the states of AL, AZ, CT, FL, GA, IA, IL, IN, KS, MD, MI, MN, MO, NC, ND, NJ, NM, NV, NY, OH, OR, PA, SC, TX, VA, WA and WY: and Internationally to Canada. | Terminated | March 23, 2016 | 531 | 184 | Computer | Not_Computer
Z-0079-2015 | 63897 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Medium 2-Hooks Spreader Bar for use with Loop Slings, sizes small, medium and large on the Maxi Sky 2 Ceiling Lift; Model 700-19425. |  |  | January 21, 2013 | 2013 | Arjo, Inc. dba ArjoHuntleigh | The upper portion of the hooks on the spreader bar are too flexible for holding the loops of the sling. The loop of the sling could potentially come out of the hooks if positioned in the upper portion of the hooks. | Device Design | URGENT FIELD CORRECTION RECALL letters dated January 21, 2013 were sent to all direct accounts (customers). The letters identified the potential safety related issue; included instructions on how to use the Medium Two Hooks Spreader Bars, and to complete and return the enclosed Customer Response form. Once the Customer Response form has been returned customers will be contacted by an ArjoHuntleigh service representative to make arrangements for the replacement of the existing spreader bars with spreader bars of a different design. Customers with any questions about this recall can contact ArjoHuntleigh by telephone at 800-323-1245, extension 57985, or by e-mail at usa.quality@arjohuntleigh.com. |  | 363 | 363 spreader bars | USA Distribution - AR, GA, MI, PA, and the District of Columbia. | Terminated | October 27, 2015 | 1009 | 363 | Computer | Not_Computer
Z-1469-2014 | 67569 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Brainlab, Spine &amp; Trauma 3D 2.0, Navigation Software. An intraoperative image-guided localization system to enable minimally invasive surgery. Catalog Number: 22264 |  |  | November 18, 2013 | 2013 | Brainlab AG | The "fluoro match registration" function of the Navigation Software Spine &amp; Trauma 3D 2.0 allows the user to intraoperatively match e.g., CT data sets to the current patient anatomy visible on fluoroscopic images, used by the navigation software during spinal surgeries to display the position and orientation of instruments. For this registration function in combination with a digitally integrated | Software design | Brainlab sent a FIELD SAFETY NOTICE / PRODUCT NOTIFICATION letter dated November 18, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The letter included instructions for the modification of C-arm configurations: 1) If possible, Use analog image transfer between C-arm and Brainlab Navigation System and to consult with the local Brainlab Customer Support Representative for this modification, and 2) If possible, adjust the image size of the C-arm to 8 bit and to follow the instructions provided with the C-arm and if required to consult the C-arm manufacturer. For measures when performing a Fluoro match registration, the following instructions are provided to reduce the probability that the algorithm results in an inaccurate match: 1) perform the preoperative CT scans according to the descriptions in the Brainlab scan protocol, 2) set the threshold so that a smooth bone surface is displayed and the quality of the chosen bone threshold influences the accuracy of the match result (refer to section "Bone Threshold" in the Software User Guide"), 3) use the Radiolucent Spine Reference Clamp instead of the Spine Reference X-Clamp, 4) make sure that no additional objects (e.g. wires, retractors, instruments, etc.) are within the view of the C-arm. Additional reminders in the letter include: 1) always verify registration accuracy by holding the pointer or instrument tip to at least three anatomical landmarks and verifying their position in the software, 2) verify that the registration is at the correct level on the patient and data set, and 3) accuracy must be checked on the treated bone surface. Brainlab will provide a software update with the issue solved to affected customers tentatively by January 2014. Customers with questions can contact the Customer Hotline at 800-597-5911 or via e-mail at us.support@brainlab.com. |  | 133 | 133 units | Worldwide Distribution - US Distribution including the states of AZ, CA,GA, IL, IN, LA, MA, MI, NY, NC, OH, PA, TN, TX and WI and the countries of Australia, Austria, Belgium, Brazil, Chile, Czech Republic, France, Germany, Italy, Japan, Malaysia, Netherlands, Norway, Philippines, Poland, Russian Federation, Singapore, South Africa, Spain, Sweden, Switzerland, Thailand, Turkey and United Kingdom. | Open, Classified |  |  | 133 | Computer | Computer
Z-0902-2014 | 67191 | No | JCX | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Neptune 2 Rover Ultra Waste Management System. |  |  | December 30, 2013 | 2013 | Stryker Instruments Div. of Stryker Corporation | The purpose of this correction is to upgrade the Neptune 2 Waste Management Systems in the field to the recently obtained 510(k) clearance specifications. | Device Design | On 12/30/13, Stryker sent an URGENT MEDICAL DEVICE UPDATE notification to consignees including government accounts and Sales Representatives, Stryker Canada and 3 in-house accounts and 1 account for lost rovers. Domestic and government consignees were sent the notification by mail FedEx overnight, and Sales Representatives and Stryker Canada were sent the notification by email. Notification described the affected product, instructions for the consignees, and upgrade information. Stryker also stated that a training kit and in-servicing will be provided to consignees. For questions regarding this recall, please contact Stryker Instruments: Monday-Friday 8am-5pm (EST)Neptune Customer Care Center 855-458-7441 or 269-389-2316strykerinstrumentsrecalls@stryker.com |  |  |  | Nationwide and Canada, Germany, Australia, New Zealand | Terminated | January 12, 2016 | 743 |  | Computer | Not_Computer
Z-1221-2014 | 67598 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Sarns Malleable Dual-stage Venous Return Cannulae, 34/46 Fr with 1/2" connector, 15" (38 cm) long |  |  | March 03, 2014 | 2014 | Terumo Cardiovascular Systems Corporation | During in-process inspection, Terumo Cardiovascular Systems (Terumo CVS) identified exposed basket wire on certain product lots of Sarns Malleable Dual-stage Venous Return Cannulae and Sarns Dual-stage Venous Return Cannulae. | Device Design | Terumo sent an Urgent Medical Device Recall letter dated March 3, 2014, to all affected consignees. The letter included the reason for removal, potential hazards, and customer instructions. Terumo CVS advised users to stop using the affected devices and return any unused product on hand. Questions or concerns were directed to Terumo CVS Customer Service: 1-800-521-2818 M-F, 8am-6pm.For questions regarding this recall call 734-741-6173. |  |  |  | Worldwide Distribution - USA including CA, CT, FL, GA, IL, IN, LA, MI, MN, MO, NC, NY, PA, TN, TX, and Internationallly to United Arab Emirates (UAE), Taiwan, Thailand, Malaysia, Belgium, Japan, Mexico, and Canada. | Terminated | August 12, 2014 | 162 |  | Computer | Not_Computer
Z-1195-2014 | 67447 | No | NQI | Immunomagnetic circulating cancer cell selection and enumeration system. | Immunology | Pathology | 510(k) | CELLTRACKS ANALYZER IIÂż (Product Code 9555) |  |  | January 08, 2014 | 2014 | Veridex, LLC | Janssen Diagnostics, LLC received complaints of camera failures on the CELLTRACKS ANALYZER IIÂż. | Device Design | Janssen Diagnostics, LLC sent an Important Product Correction Notification letter dated January 8, 2014 to all affected customers via FedEx overnight courier letter and Foreign Affiliate Consignees by email on January 8, 2014. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete and return the Confirmation of Receipt Response form attached. For questions contact Customer Technical Service at 1-877-837-4339. |  | 19 | 19 units | Worldwide Distribution - USA Nationwide | Terminated | July 06, 2017 | 1275 | 19 | Computer | Not_Computer
Z-1746-2014 | 68349 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | IntelliVue MX40 Wearable Patient Monitor in use with the Philips IntelliVue Information Center iX (PIIC iX); 865350, 865351, 865352 |  |  | May 22, 2014 | 2014 | Philips Medical Systems, Inc. | IntelliVue MX40 (part numbers 865350, 865352) with software revisions B.00 or B.01 may not resume alarming when the alarms are paused and the pause timer ends. | Software design | On May 22, 2014, each affected customer was sent an Urgent Medical Device Correction notification/Field Safety Notice. The letter explains the problem, explains the action to be taken by the consumer, and the action planned by Philips. The Urgent Medical Device Correction notification/Field Safety Notice informs customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue.The correction will consist of a free of charge software upgrade. A Philips Healthcare representative will contact customers with affected devices to arrange for service. Philips is asking customers to follow the Action to be Taken by Customer/User section of the Urgent Medical Device Correction notification/Field Safety Notice:The following two items must be configured when using MX40 with the IntelliVue Information Center iX: 1.Utilize the default state of Pause Alarms at Yellow only, which disables the ability to Pause/Suspend all Red level alarmsUtilize the default state of Remote Suspend Alarms as Off, which disables the use of the Pause Alarms feature at the MX40 device |  | 42600 | 42,600 | United States: NationwideForeign Countries:Germany, Austria, Finland, Sweden, Australia, Qatar, Canada, Netherlands, New Zealand, Portugal, United Kingdom, Spain, Belgium, France, Switzerland, Norway, Czech Republic, Poland, Italy, Singapore, Denmark, Iceland, India, Saudi Arabia, Ireland, Oman, Bahrain, Japan, Israel, Malaysia, South Africa, Turkey, Thailand, Latvia, Bulgaria, Indonesia, Slovenia, United Arab Emirates, Hong Kong, Aruba, Chile, Lebanon, Korea, Kuwait, Taiwan, Argentina, Hungary, Luxembourg, Mexico, Gabon | Open, Classified |  |  | 42600 | Computer | Computer
Z-1197-2014 | 67625 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO RPT System |  |  | February 17, 2014 | 2014 | Elekta, Inc. | Dose and Monitor Unit values are not computed correctly when Elekta Motorized Wedges are used with Dynamic Conformal arcs. | Software design | Elekta, Inc. issued an Important Safety Notice (LINXi000019) issued February 14, 29, 2014. The Safety Notice identified the product, the problem, and the action to be taken by the customer. All affected customers of the issue globally by February 28, 2014. Provide a fix to the defect in an expeditious timeframe. It is mandatory for all customers to apply the fix. Version 5.00.01 will be released for the fix in March 2014. Follow up with the affected customers for confirmation receipt of the LIN, and confirmation of installation of patch for the fix by November 30, 2014.For questions regarding this recall call 770-670-2422. |  | 1776 | 1776 | USA including AL, AK, AZ, CA, CO, CT, DE, DC, FL, GA, IL, IN, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NJ, NY, NC, ND, OH, OK, PA, RI, SC, TN, TX, VA, WA, WV, WI, WY, Puerto Rico and Internationally to Albania, Algeria, Argentina, Australia, Austria, Bahamas, Barbados, Bermuda, British Guiana, Boliva, Brazil, British West Indies, Burundi, Canada, Chile, China, Columbia, Costa Rica, Cuba, Curacao, Cyprus, Czechoslovakia, Dominica, Equator, Estonia, Egypt, Finland, France, Gambia, Germany, Great Britain, Georgia, Greece, Guatemala, Hong Kong, India, Indonesia, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kuwait, Lativa, Lithunia, Madagascar, Malaysia, Malta, Mexico, Mongolia, Myanmar, Netherlands, New Zealand, Nigeria, Pakistan, Peru, Philippines, Poland, Portugal, Romania, Russia, San Salvador, Saudia Arabia, Serbia, Singapore, Sir Lanka, South Korea, Spain, Sudan Slovenia, Suriname, Tajikistan, Thailand, Turkey, Turkmenistan, Taiwan, Ukraine, Uruguay, Venezuela, Viet Nam, Zambia. | Terminated | July 21, 2015 | 519 | 1776 | Computer | Not_Computer
Z-1741-2014 | 68269 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | enGen (TM) Laboratory Automation System Configured with TCAutomation" Software v3.5 -- For in vitro quantitative measurement of a variety of analytes of clinical interest --- |  |  | April 17, 2014 | 2014 | Ortho-Clinical Diagnostics | Software anomaly: potential buffer timeout issues when using enGen (TM) Laboratory Automation Systems Configured with TCAutomation (TM) Software v3.5. | Software Design Change | Ortho Clinical Diagnostics sent an Important Product Correction Notification dated April 17, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.On April 17, 2014, foreign affiliates were informed by e-mail of the issue and instructed to notify their customers of the issue and actions.The letters instruct customers: to ensure that samples prematurely routed to Check are detected and processed appropriately, please monitor completion of all pending tests following standard laboratory practice. Customers were informed of the issue and advised to: -- Routinely monitor pending result reports from your Laboratory Information System (LIS) to determine if results are pending for samples. -- Investigate to determine the next action identified for a sample if the sample is routed to the Check location prematurely. If you have any additional questions, please call Customer Technical Services at 1-800-421-3311. |  | 9 | Domestic: 6 units, Foreign: 3 units | Worldwide Distribution - US including the states of Arizona, Delaware, Ohio, Tennessee, and North Carolina; and the countries of Mexico, England and Italy. | Terminated | April 11, 2017 | 1090 | 9 | Computer | Computer
Z-1739-2014 | 68135 | Yes | KPO | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | NaturaLyteÂż Liquid Bicarbonate Concentrate, 6.4 liter bottle to be used with a hemodialysis unit; Part Number: 08-4000-LB. |  |  | May 01, 2014 | 2014 | Fresenius Medical Care Holdings, Inc. | Container cap may contain foreign material- 0.60% - 0.80% Manganese | Equipment maintenance | Fresenius Medical notified customer by letter dated 5/1/14 to examine stock immediately and instructed to discontinue use and place all Naturalyte Liquid Bicarbonate Concentrate in a secure area for return to FMC-RTG. The customer instructed to contact the Fresenius Medical Care Customer Service Team for instructions on how to return the recalled product. |  | 12 | 12 bottles | Distributed in the state of MO. | Open, Classified |  |  | 12 | Computer | Not_Computer
Z-1064-2014 | 65450 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | SL &amp; SLi Series (Mod 15/18/20/25) Linear Accelerator |  |  | May 28, 2013 | 2013 | Elekta, Inc. | Customers are resetting the default values outside recommended factory settings. | Device Design | The firm, Elekta, sent an "Important Field Safety Notice (IFSN 200 01 103 073 - Electron Applicator Auto-Tracking Defaults)" released on May 28, 2013 and is in the process of being distributed to all customers. The Notice requests users to please adhere to the instructions and advice given in manuals and system dialogues, which explicitly state that a change to auto-tracking values could compromise compliance with safety standards and can cause increased unwanted radiation dose to the patient. All machines should be checked against the default settings attached. If the actual settings are greater than the default settings, then Elekta's recommendation is that the default settings should b re-applied and all relevant physics checks (e.g. beam uniformity, output factor) performed. If, however, auto-tracking values are increased beyond their factory-set defaults, then electron applicator radiation leakage measurements should also be made to ensue that the system remains compliant to IEC safety. All users of Desktop Pro and Integrity 1.1 software will receive a mandatory upgrade to Integrity 1.2 (released 28-Jan-14. The upgrade path to Integrity 1.2 differs depending on system configurations at each customers site, therefore Elekta has released three separate Important Field Safety Modifications in order to upgrade all customers successfully. Important Field Safety Notice 200 01 501 036  Important Field Safety Modifications 200 02 501 037  Applicable System Configuration Desktop Pro with Fast Tuning Magnetron, New Patient Support System; Important Field Safety Notice 200 01 501 038  Important Field Safety Modification 200 02 501 039  Applicable System Configuration Desktop Pro with Linac Control System MK3i, Fast Tuning Magnetron, New Patient Support System; Important Field Safety Notice 200 01 502 035  Important Field Safety Modification 200 02 502 036  Applicable System Configuration All Integrity 1.1 sites. All users of Integrity 3.1 will receive a man |  | 449 | 449 | Worldwide Distribution: US (Nationwide) including states of: AK, CA, CT, FL, IL, IN, IA, LA, MD, MA, MI, MN, NJ, NY, OH, OK, OR, PA, TN, TX, VA, WA, WV, and Puerto Rico; and countries of: Argentina, Australia, Austria, Belarus, Belgium, Bosnia &amp; Herzegovinia, Brazil, Canada, China, Czech Republic, Denmark, France, Germany, Greece, Hong Kong, Hungary, India, Indonesia, Iran, Ireland, Italy, Japan, Jordan, Libya, Mexico, Netherlands, New Zealand, Norway, Panama, Philippines, Poland, Portugal, Russia, Senegal, Serbia, Singapore, Slovenia, South Africa, Spain, Sudan, Sweden, Switzerland, Taiwan, Turkey, and United Kingdom. | Open, Classified |  |  | 449 | Computer | Computer
Z-1736-2014 | 68271 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS 5,1 FS Chemistry System and VITROS 5,1 FS Refurbished, Product Codes/Catalog Numbers 6801375, 6801890 for in vitro diagnostic use. |  |  | April 10, 2014 | 2014 | Ortho-Clinical Diagnostics | Potential for biased results when a Multiple Window (MW) code is associated with results generated by the VITROS 4600 Chemistry System, the VITROS 5600 Integrated System and the VITROS 5,1 Chemistry System. MW codes may be associated with results from any of the following VITROS Chemistry Products using multiple point (multipoint) rate measurements. The MicroSlide assays that could be affected | Software design | Important Product Correction Letters (CL2014-116, dated 4/10/2014) were sent on 4/10/2014 via FedEx overnight courier or ORTHO PLUS e-Communications to all customers with VITROS 5,1 FS, 5,1 FS Refurbished, 4600,and/or 5600 Systems. They were informed of the issue and advised to: -- Review VITROS ALT, ALTJ, ALKP, ALKPJ, AST, ASTJ, CHE, CHEJ, CK, CKJ, GGT, GGTJ, LDH and LDHJ results generated on VITROS 4600/5600 and 5,1 FS Systems. Do not report any results associated with an MW code. -- Perform repeat testing in singlicate on all samples with results having an MW code. Only results without an MW code are acceptable. -- Review their current Laboratory Information System (LIS) configuration to determine if changes are required in order to restrict reporting of results with an MW code. Customers with questions should contact Customer Technical Services at 1-800-421-3311. --- Foreign affiliates were informed of the issue by e-mail on 4/10/2014 and instructed to notify their consignees of the issue and actions. |  | 1978 | Domestic - 597 units, Foreign - 1381 units | Worldwide distribution including US Nationwide, Puerto Rico, Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Singapore, Spain, and Venezuela. | Open, Classified |  |  | 4029 | Computer | Not_Computer
Z-1738-2014 | 68270 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | enGen Laboratory Automation Systems (enGen) |  |  | April 10, 2014 | 2014 | Ortho-Clinical Diagnostics | Software anomaly: potential for delay in the reporting of patient sample test results when using enGen with TCAutomation Software v2.6 , v3.2 and v3.5 Configured with Bypass Modules for ADVIA Centaur, Abbott Architect, VITROS and/or Tosoh AIA Systems. | Software design | Urgent Product Correction Letters (CL2014-114, dated 4/10/2014) were sent to all US customers and Foreign affiliates starting on 4/10/2014. The letter informed customers of the issue and provided required actions to be taken: Routinely monitor your Bypass module(s) -- If you detect a stoppage due to the software timing issue, follow the steps in the How to Detect the Bypass Module Time Issue and Release the Samples section on Page 2 of this notification. -- Complete the Confirmation of Receipt Form and return by 4/23/2014. -- Place this notification near each system or with your enGen (TM) Laboratory Automation System user documentation. |  | 108 | Domestic - 31 units; Foreign - 77 units | Worldwide Distribution: US Nationwide including Puerto Rico, and countries of: Australia, Bermuda, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Singapore, Spain, and Venezuela | Open, Classified |  |  | 108 | Computer | Computer
Z-1215-2014 | 66787 | No | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | Dental Abutment labeled in part: "4555 Riverside Drive Palm Beach Gardens, FL 33410 USA 1-800-342-5454***EDAZ , EDAZx and ILDExxx***BellaTek Zirconia Abutment Position Number: XX Abutment; Pilier; Abutment; Abutment; Pilastro; Pilar***Rx only***Non-sterile, Sterilize prior to use.***Biomet 3i Dental Iberica S.L.*** |  |  | August 26, 2013 | 2013 | Biomet 3i, LLC | Reports relating to Biomet 3i's BellaTek Zirconia Dental Abutment fractures were received. | Device Design | The firm, BIOMET 3i, sent an "URGENT: Medical Device Removal - BellaTek Zirconia Abutments" letter dated August 2013 via email, facsimile or postal mail to its customers. The letter described the product, problem and actions to be taken. All customers were instructed to check their inventory for the affected product; return the abutments to BIOMET 3i for replacement and full credit; complete and return the attached Recall Return Response Form, and to advise their customers and forward a copy of the Medical Device Removal Notice and the contained recommendations.BIOMET 3i has decided to initiate a design change to mitigate the fracture risk, and consequently, will cease distribution of BellaTek Zirconia Abutments unit l the corrective actions are completed.If you have any further questions, please contact Customer Service at 1-800-342-5454 for assistance. |  | 10454 | 10,454 | Worldwide Distribution: US (nationwide) including states of: AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, KS, KY, LA, MA, MD, MI, MN, MO, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, Hawaii and Puerto Rico; and countries of: Spain, Australia, Canada, Colombia, Japan, Uruguay, Mexico, Estonia, United Kingdom, Denmark, Germany, Spain, Italy, Poland, Ireland Netherlands, France, Sweden, Portugal, Australia, France, Belgium, Switzerland, Norway, Czech Republic, France, Holland, and Luxembourg. | Terminated | September 26, 2014 | 396 | 10454 | Computer | Not_Computer
Z-1729-2014 | 68237 | No | LDS | Medical washer. | General Hospital | General Hospital | 510(K) Exempt | HLD Systems Model 610 Washer/High Level Disinfector. |  |  | April 24, 2014 | 2014 | Cenorin, LLC | Cenorin LLC is voluntarily recalling HLD Systems Model 610 Washer/High Level Disinfector due to a potential risk of the system not injecting the proper amount of cleaning fluid during the wash cycle and not achieving the desired level of debris reduction. | Equipment maintenance | Cenorin sent a letter on April 24, 2014 informing their customers of the potential failure of the pump. This letter also explains the customer's option of upgrading their entire HLD System to the new cleaning fluid delivery system or to continue purchasing a six month PM kit that has a replacement pump added. Customers are instructed to complete and return an enclosed response form within 15 days and to email Jamie Wright at jwright@cenorin.com or call 800-426-1042, extension 224 for questions. |  | 25 | 25 HLD systems | Distribution USA Nationwide | Terminated | June 30, 2015 | 432 | 25 | Computer | Not_Computer
Z-1728-2014 | 68431 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Horizon Medical Imaging |  |  | May 29, 2014 | 2014 | Mckesson Medical Immaging | Firm received a complaint that an unreported study performed that day could not be opened from the archive. | Software design | McKesson sent an Urgent Field Safety Notice dated May 20, 2014, to all affected customers. The notice identified the product, the problem, and the action to be taken by the customer. Customers were informed that the potentially affected software versions will be corrected without removing the systems from their location. Customers were instructed to distribute the notice to all those within their organization who need to be aware and alert all those affilliated with their organization that are affected by this field safety notice. All customers receiving the Field Safety Notice should contact McKesson Support as soon as possible at 1-800-663-2533 to arrange installation of the software update. Customers with questions should contact the Customer Support Department and ask to speak to their Support Manager. |  | 659 | 659 | Worldwide Distribution - USA including AL, AK, AZ, AR, CA, CO, CT, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OR, PA, RI, SC, TN, TX, VT, VA, WA, WV, WI, WY, and Puerto Rico. Internationally to Australia, Canada, France, Great Britain,and Guam. | Terminated | March 31, 2017 | 1037 | 659 | Computer | Not_Computer
Z-1216-2014 | 67559 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo Dynamics Picture Archiving and Communication System (PACS). Model numbers 10091804, 10091805, 10091806, 10091807, 10091673 |  |  | January 27, 2014 | 2014 | Siemens Medical Solutions USA, Inc | Potential patient data mixup. Improper error handling could allow for two different patients to be displayed in syngo Dynamics. Images shown for one patient and reporting system shown for a different patient, leading to the potential for misdiagnosis. | Software design | A recall notification letter, dated January 27, 2014, was sent to end Users. Customer Safety Advisory Letter SY016/14/S. Siemens is preparing a Service Patch modification of syngo Dynamics v9.5 which will be available in March 2014. |  | 50 | 50 | Nationwide distribution to MN, AR, MA, TX, NJ, WV, GA, CA, PA, OH, MI, WI, TN, NY, KS, CT, NY, MO, IL, IN, OR, IA, and WA. | Terminated | March 16, 2015 | 413 | 50 | Computer | Computer
Z-1712-2014 | 68326 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | PATH(R) THREAD CUP ADAPTOR, REF 2007-0013, Wright Medical Technology, Inc. Orthopaedic surgical instrument. |  |  | April 30, 2014 | 2014 | Microport Orthopedics INC. | Difficulty removing the cup adaptor from the impaction handle during surgery | Device Design | Firm has initiated the recall by email on 04/28/2014 to all distributor's offices, including international, then followed with written communication on 05/02/2014 via Fed Ex P1 delivery to all United States consignees. Follow-up to non-responding consignees will commence 30 days after initiation. |  | 140 | 140 units | Nationwide Distribution. | Terminated | December 01, 2014 | 215 | 140 | Computer | Not_Computer
Z-1234-2014 | 67584 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | MEDITECH Microbiology; Client Server Releases 5.6, 5.65, 5.66, 5.67, 6.0, 6.05, 6.07, 6.1, 6.13 and 6.14. |  |  | February 21, 2014 | 2014 | Medical Information Technology, Inc. | Potential for erroneous result reporting. | Device Design | Beginning on February 3, 2014 MEDITECH distributed a notification to Client Server customers who have the Microbiology product in TEST and/or LIVE only for the releases noted in this correction report. This notification was made via e-mailed task updates that can be printed by the customer. Task updates sent via e-mail are immediately transmitted to the customer. The issue will be corrected by CS LAB DTS 12279. |  | 438 | 438 | Distributed USA (nationwide) and Canada. | Terminated | November 12, 2015 | 629 | 438 | Computer | Computer
Z-1783-2014 | 64507 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Philips IntellVue Info Center iX 866023; PIIC iX Upgrade 866024; IntellVue Server iX Upgrade 866026; PIIC Classic Upgrade 866117 Patient Physiological Monitor. |  |  | March 05, 2013 | 2013 | Philips Healthcare Inc. | If a customer creates customized trend scales in the trend review tile and the iX or primary server reboots for any reason, the iX(s) will enter a reboot loop. If one iX reboots, the issue will impact only those patients monitored on that device. If the primary server reboots, all patients across all iX devices will be impacted. | Software design | On 3/5/13, Philips sent each known affected customer an Urgent Medical Device Correction notification/Field Safety Notice to inform them that Philips has received a customer report that the PIC iX (release A.01.XX) central stations that are connected to a database server (DBS) may not be able to complete a restart/reboot resulting in the loss of central monitoring. In order to prevent this from occurring, customers were advised to not modify/customize the default trend scales for reviewing stored patient data in the retrospective review applications. Recall letter also states that a Philips Healthcare representative will contact customers with affected devices to arrange for the installation of updated PIIX iX software resolving this issue on affected units. Philips will conduct these updates for all affected devices at no charge. |  | 899 | 899 | Worldwide Distribution: US (nationwide) and country of: Canada. | Terminated | May 17, 2017 | 1534 | 899 | Computer | Computer
Z-1711-2014 | 68110 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | AXIOM Luminos dRFsystem |  |  | April 11, 2014 | 2014 | Siemens Medical Solutions USA, Inc | It was determined that under rare environmental conditions (i.e., extremely dry air in combination with a highly isolated floor) electrostatic discharges (ESD) &gt;8 kV may trigger an unintended movement of the AXIOM Luminos dRFsystem which may lead to an emergency situation and pose danger to the patients, operating personnel or to the unit. | Device Design | Siemens sent an Safety Advisory Notice dated April 11, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.What steps can the user take to avoid the potential risk of this issue?Press one of the red emergency STOP buttons immediately in case of an unintended movement. The system needs to be restarted afterwards to be fully functional again.How will the issue finally be resolved?Siemens is preparing a modification of the Luminos dRF system control console that will resolve this potential malfunction. The field modification will be available by end of May 2014.We appreciate your understanding and cooperation with this safety advisory notice and ask you to immediately instruct your personnel accordingly. Please ensure that this safety advisory notice is placed in the system's instructions for use until the update has been installed.If you have sold this device/equipment and it is no longer in your possession, we kindly ask that you forward this safety notice to the new owner of this device/equipment. Please inform us about the new owner of the device/equipment.For further questions please call (610) 219-6300. |  | 37 | 37 | Class II Recall - US Distribution including the states of IL,VA, TN, PA, TX, MN, GA, MI, IN, OH, FL, CA, OR, NY, ME, MO, OK, CO and NJ. | Terminated | December 15, 2014 | 248 | 37 | Computer | Not_Computer
Z-1231-2014 | 67630 | Yes | BTS | Bronchial tube. | Anesthesiology | Anesthesiology | 510(k) | Rusch Brochopart and Rusch Brochopart White |  |  | February 28, 2014 | 2014 | Teleflex Medical | The product may fail to achieve seal of right lung due to the cuff inflating to one side | Device Design | Teleflex sent a Urgent Medical Device Recall Notification dated February 28, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed:1. If you have affected stock, immediately discontinue use and quarantine any products with the catalog numbers listed above.2. To return product, complete the enclosed Recall Acknowledgement Form and fax it to 1-866-804-9881, Attn: Customer Service. This will allow us to document the amount of product you have on hand for return. A customer service representative will contact you with a Return Goods Authorization (RGA) Number and will provide instructions for the return of product to Teleflex Medical.3. If you have no affected stock, please complete the enclosed Recall Acknowledgment Form and fax it to 1-866-804-9881, Attn: Customer Service. This will allow us to document your receipt of this letter.Teleflex Medical is committed to providing high quality, safe and effective products. We sincerely apologize for any inconvenience this action may cause your operations. If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-866-246-6990. |  | 63568 | 63,568 ea. | Worldwide Distribution - USA (nationwide) Distribution including Puerto Rico and the states of AL, AZ, CA, CO, CT, DC, FL, GA, ID, IL, IN, KS, KY, LA, MD, MA, MI, MN, MS, MO, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, VA, WA, WV and WI,and the countries of United Arab Emirates, Austria, Australia, Azerbaijan, Belgium, Belgrade, Bahrain, Canada, Switzerland, Cameroon, China, Costa Rica, Cyprus, Czech Republic, Germany, Denmark, Dominican Republic, Algeria, Ecuador, Estonia, Egypt, Spain, Finland, France, Great Britain, Greece, Guyana, Croatia, Hong Kong, Hungary, Ireland, Israel, India, Iraq, Italy, Japan, Jordan, Kenya, Kuwait, Lebanon, Latvia, Libya, Morocco, Serbia, Malaysia, Martinique, Nigeria, Netherlands, Norway, Oman, Panama, French Polynesia, Philippines, Poland, Portugal, Qatar, Reunion, Russia, San Salvador, Saudi Arabia, Sweden, Singapore, Slovenia, Slovakia, San Marino, Thailand, Trinidad, Tunisia, Turkey, Venezuela, Yemen and Zambia | Open, Classified |  |  | 63568 | Computer | Not_Computer
Z-1706-2014 | 68219 | No | KPY | Personnel protective shield. | Radiology | Radiology | 510(K) Exempt | Kenex Radiation Shield and surgical lamps installed with GE Healthcare Interventional, Surgery, CT and X-Ray Imaging systems. |  |  | May 02, 2014 | 2014 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue due to the improper installation of a Kenex ceiling suspended radiation shield and surgical lamp. This issue is not related to the design of the Kenex product. The affected products are limited to installations completed or contracted by GE Healthcare.There has been a reported incident of a radiation shield falling from the ceil | Equipment maintenance | GE Healthcare sent an "Urgent Medical Device Correction" letter # 12209 dated May 2, 2014. The letter was addressed to Director of Biomedical Engineering, Director of Radiology and Chief of Cardiology. The letter described the Safety Issue, Safety Instructions, Affected product Details, Product Correction and Contact Information. For questions they should call 1-800-345-2700. For other countries, they should contact their local GE Healthcare Service Representative. . |  | 3221 | 3221 (108 US, 3113 OUS). | Worldwide Distribution - US Nationwide, ALBANIA, ALGERIA, ANGOLA, ARGENTINA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BELARUS, BELGIUM, BOLIVIA, BOSNIA &amp; HERZEGOVINA, BRAZIL, BULGARIA, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, CROATIA, CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, ECUADOR, EGYPT, FINLAND, FRANCE, GEORGIA, GERMANY, GREECE, GUADELOUPE, GUATEMALA, GUYANA, HONDURAS, HONG KONG, HUNGARY, ICELAND, INDIA, INDONESIA, IRAN, IRAQ, IRELAND, ISRAEL, ITALY, JAMAICA, JAPAN, JORDAN, KAZAKHSTAN, REPUBLIC OF KOREA, KUWAIT, LATVIA, LEBANON, LIBYAN ARAB JAMAHIRIYA, LITHUANIA, MACEDONIA, MALAYSIA, MALTA, MARTINIQUE, MAURITIUS, MEXICO, MOROCCO, NETHERLANDS, NEW ZEALAND, NICARAGUA, NIGERIA, NORWAY, OMAN, PAKISTAN, PANAMA, PARAGUAY, PHILIPPINES, POLAND, PORTUGAL, REUNION, ROMANIA, RUSSIA, SAUDI ARABIA, SENEGAL, SERBIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SUDAN, SWEDEN, SWITZERLAND, SYRIA, TAIWAN, THAILAND, TUNISIA, TURKEY, UKRAINE, UNITED KINGDOM, VENEZUELA, VIETNAM, and YEMEN. | Terminated | February 25, 2015 | 299 | 3221 | Computer | Not_Computer
Z-1707-2014 | 67926 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ExacTrac 5.5, an Image Processing System used for patient positioning for radiation therapy or radiosurgery. |  |  | April 15, 2014 | 2014 | Brainlab AG | When using a specific plan update workflow in combination with a non-Brainlab treatment planning system (TPS), there is a possibility in ExacTrac 5.5.5 to generate an unintended blend of two initially separate plans of the same patient. | Software design | Brainlab sent a "FIELD SAFETY NOTICE / PRODUCT NOTIFICATION" letter dated March 13, 2014 was sent to all affected customers. The letter identifies the affected product, problem, and actions to be taken. The letter includes specific instructions to users that are intended to reduce the likelihood of similar events occurring when using ExacTrac software version 5.5.5. The letter further states that Brainlab will be providing a software update to all customers. Brainlab will contact all customers to arrange for the software update within the following 6 months. Customers with any questions about this issue can contact Brainlab at 800-597-5911 or via e-mail at us.support@brainlab.com. |  | 40 | 40 systems (worldwide) | Worldwide Distribution - US Nationwide in the states of AZ, AR, CA, FL, KY, MA, NE, NY, TN, TX and the countries of Argentina, Australia, Brazil, Finland, India, Japan, Netherlands, Oman, Russia, and Taiwan | Open, Classified |  |  | 40 | Computer | Computer
Z-1709-2014 | 68266 | [6 empty] | CARINAsim MAN-1236 Rev 2 and CARINAiso MAN-1237 Rev 2 laser software. LAP isocenter marking control software. Laser systems for patient alignment. |  |  | October 07, 2013 | 2013 | LAP of America Laser Applications, LLC | Potential for patient to be marked incorrectly. Customers currently using CARINAsim or CARINAiso systems can encounter two safety relevant errors in the operating instructions (CARINAsim MAN-1236 Rev 2 / CARINAiso MAN-1237 Rev 2) | Software design | On 10/16/2013 Customers were sent an Urgent field safety notice for Corrective action concerning the operating instructions of CARINAsim and CARINAiso. The letter identified the affected product, as well as the issues involved. Users were to review their current operating steps and validate their operating sequence. Users should have received a modified software version with software update in February 2014. All users should have been notified of this letter and a copy should be maintained in the users' records. Questions should be directed to Thomas Gaudet on 561-416-9250 at the Service Center. |  | 280 | 280 Manuals | Worldwide Distribution -- USA, including the states of AK, AZ, CA, CT, DC, FL, GA, IA, ID, IL, KS, LA, MA, MD, MI, MN, MO, MS, NC, NH, NJ, NM, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WY and the territory of Puerto Rico; and, the countries of Brazil, Canada, Chile, Colombia, Cuba, Honduras, and Mexico. | Open, Classified |  |  | 280 | Computer | Computer
Z-1705-2014 | 68357 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Spacelabs Medical Ultraview SL Command Module, Model 91496. |  |  | May 08, 2014 | 2014 | Spacelabs Healthcare Inc | A software anomaly causes temperature alarm limits to change to limits which are not physiologically possible and result in nuisance false positive alarms. | Software Manufacturing/Software Deployment | Spacelabs Healthcare identified a software anomaly in the Ultraview SL" Command Module, Model 91496. During an internal audit, Spacelabs identified a reportable recall that was not reported. We have a product called the Spacelabs Medical Ultraview SL" Command Module, Model 91496. A software error was discovered at a customer site which was not a safety issue, but in the interest ofcustomer service it was decided to downgrade all their modules with an older software version. |  | 64 | 64 devices | US Distribution to GA. | Terminated | June 03, 2014 | 26 | 64 | Computer | Computer
Z-1676-2014 | 68091 | No | JXG | Central nervous system fluid shunt and components. | Neurology | Neurology | 510(k) | Aesculap Inc.(AIC) Miethke Ventricular Catheter with Deflector/Miethke Shunt System |  |  | April 21, 2014 | 2014 | Aesculap, Inc. | AIC (USA) received information regarding 3 complaints in which the deflector did not move freely on the ventricular catheter. | Device Design | Aesculap sent an Important Correction &amp; Removal recall letter dated April 21, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. An Aesculap Sales Representative will contact you to schedule return of any affected inventory and replace your product. Please complete the attached Inventory Sheet which is necessary to comply with FDA regulations. When completing the inventory sheet, please be sure to fill in the quantitybeing returned. If you cannot locate the product, please provide an explanation as to why the inventory will not be returned (discarded, etc.).AIC (USA) appreciates your cooperation on this matter and apologizes for the inconvenience this may cause. Thank you for your patience and continued support of this product. Please contact (610) 984-9414 with any questions. |  | 48 | 48 | US Distribution including the states of PA, KY, NY, RI, SC, KS, TX, CT, TN and WA. | Terminated | October 20, 2014 | 182 | 48 | Computer | Not_Computer
Z-1235-2014 | 67416 | No | INI | Motorized three-wheeled vehicle. | Physical Medicine | Physical Medicine | 510(k) | Buzzaround XL scooters' |  |  | January 21, 2014 | 2014 | Golden Technologies, Inc. | Buzzaround XL scooters' front to rear lockup can become unintentionally disengaged due to possible improper fit of the front alignment cup to the rear alignment screw. | Device Design | Golden Technologies sent a recall notification letter, dated January 14, 2014 to providers/dealers. A second letter, dated January 21, 2014 was sent to these same customers. Based on a slow initial response, the firm intends to send out the correction kit with the inspection acknowledgement form as soon as possible. Based on the overall response from dealers, a third consumer letter was drafted (notification3) for mailing directly to known consumers, if necessary. For help contact a special technical service hotline at 1-800-624-6374 ext. 392. |  | 5044 | 5044 | Worldwide Distribution - USA Nationwide | Terminated | July 23, 2015 | 548 | 5044 | Computer | Not_Computer
Z-1245-2014 | 67592 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Optima CT580/Discovery CT590RT scanners running software version 13BW16.3. |  |  | February 14, 2014 | 2014 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue due to a software issue associated with the Head Scan Protocols used on your Optima CT580 or Discovery CT590RT scanner.A potential hazardous situation may occur during a CT head scan causing a novel and rare artifact that may hide pathology. When using the head and large head SFOV, which applies the Iterative Bone Option (IBO) | Software design | The firm, GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated February 14, 2014 to its Consignees. The letter was addressed to Hospital Administrators / Risk Managers, Radiology Directors and Radiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. The consignees were instructed to ensure that all potential users in your facility are made aware of this safety notification and recommended actions. A GE Healthcare Service representative will perform the required software update on each affected system. This activity will be performed at no cost to consignees.Please contact your local GE Healthcare Service Representative if you have any questions or concerns regarding this notification. You may also contact GE Healthcare Call Center: United States or Canada: 800-437-1171 and in other countries, please contact your local GE Healthcare Service Representative. |  | 33 | 33 (21 US, 12 OUS) | Worldwide Distribution: US (nationwide) to states of: CO, IN, MA, NJ, NC, OH, PA, TX, and VA; and countries of: Algeria,France, Germany, Poland, Russia, and Taiwan. | Terminated | April 29, 2015 | 439 | 33 | Computer | Computer
Z-1677-2014 | 68263 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | enGen (TM) Laboratory Automation System configured with Data Innovations, LLC Instrument Manager (TM) (IM) v8.06 --- For in vitro quantitative measurement of a variety of analytes of clinical interest --- |  |  | March 20, 2014 | 2014 | Ortho-Clinical Diagnostics | Software Anomaly: Ortho Clinical Diagnostics (OCD) has become aware that if a sample requires follow-up action, such as dilution or retesting due to a mechanical issue, it is possible for that sample to be routed to the sample storage drawer (Storage) prior to the completion of the necessary pending action. A delay in patient test result(s) reporting could occur as operator is not alerted that results have not been produced and that further actions are required. | Software design | US consignees were sent Urgent Product Correction Letters (Cl2014-102, dated 3/20/2014) on 3/20/2014 by FedEx or USPS (for PO Box addresses only) overnight mail, informing them of the issue and instructing them to ensure that all required data elements are entered and in the correct format when using the auto verification feature or custom rules. J&J; Foreign affiliate consignees were notified by e-mail on 3/20/2014 informing them of the issue. |  | 81 | Domestic: 20 units, Foreign: 61 units | Worldwide Distribution-USA (nationwide) including Puerto Rico, and the countries of Australia, Bermuda, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Singapore, Spain, and Venezuela. | Terminated | April 11, 2017 | 1118 | 81 | Computer | Computer
Z-1678-2014 | 68179 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | IntelliSpace Portal (ISP), (ISP DX/HX/EX, Model number 881001; ISP IX, Model number 881030; ISP LX SPECT, Model number 881035), Philips Medical Systems. A medical software system for diagnostic imaging data viewing. |  |  | March 11, 2014 | 2014 | Philips Medical Systems (Cleveland) Inc | During SPECT reconstructions using Attenuation Correction and Scatter Correction, no Scatter Correction is being applied in the AutoSPECT Pro application on IntelliSpace Portal. In addition, resolution recovery is not applied correctly in SPECT reconstructions using Astonish in AutoSPECT Pro on Intellispace Portal. | Software design | On 03/10/2014 the firm sent Urgent Medical Device Correction Letters to their customers. |  | 157 | 157 units | Worldwide distribution, including US nationwide, Argentina, Australia, Austria, Canada, Cayman Islands, Chile, Denmark, Finland, France, Germany, India, Iraq, Israel, Italy, Latvia, Malaysia, Mexico, Netherlands, Panama, Peru, Poland, Saudi Arabia, Singapore, Slovakia, Spain, Switzerland, Turkey and United Kingdom. | Terminated | August 04, 2015 | 511 | 157 | Computer | Computer
Z-1263-2014 | 67685 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Philips Model EPIQ 5 Ultrasound System. Intended for diagnostic ultrasound imaging and fluid flow analysis. |  |  | February 28, 2014 | 2014 | Philips Medical Systems North America Inc. | A problem has been detected with Philips Model EPIQ 5 and 7 Ultrasound System. The system will report incorrect Doppler velocities if you use High Pulse Repetition Frequency (HPRF) with multiple sample volume gates. While in High Pulse Repetition Frequency (HPRF) in Display Zoom Velocity Scale using manual spectral Doppler trace, the velocities in the velocity results table may be overstated by | Software design | The firm, Philips Healthcare Ultrasound, sent an "Medical Device Correction" letters dated 2014 Feb 28 via Certified Mail via the USPS by March 5, 2014. The letters described the product, problem and actions to be taken. The customers were instructed to not use HPRF in PW Doppler mode, to avoid the noted issues. Philips Field Service Engineer will be contacting them to schedule software upgrade to the systems free of charge.If you need any further information or support concerning this issue, please contact the Quality Specialist II at 978-659-4519 or your local Philips representative. |  | 262 | 489 total =227 units in US and 262 units Internationally (no Canadian consignees) | Worldwide Distribution: US (nationwide) including states of: AL, AK, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NM, NJ, NM, NY, OH, OK, PA, SC, TN, TX, VA, WA, and WV; and countries of: Australia, Austria, Belgium, Canada, Chili, Ecuador, Egypt, Finland, France, Germany, Greece, India, Ireland, Italy, Japan, Korea Rep, Malaysia, Netherlands, New Zealand, Norway, Peru, Philippines, Portugal, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Thailand, Turkey and United Kingdom. | Open, Classified |  |  | 262 | Computer | Computer
Z-1238-2014 | 67588 | No | KOD | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | SLC 2-way standard Lubricious Coated |  |  | February 07, 2014 | 2014 | Convatec Inc. | The lumen of the Medline Sterile Foley Catheter was found to be blocked by excess lubricious coating material. | Device Design | Convatec sent Urgent: Voluntary Medical Device Recall letters dated 2/7/2014 via UPS and email. Foreign customers were provided the notice via DHL. The letter identified the affected product, problem and actions to be taken. Customers are asked to perform a count of recalled product in inventory. All consignees should be contacted and be instructed to return product based on their procedures. Customers were instructed to complete and return the enclosed respoonse form. |  | 5540 | 5,540 catheters | Worldwide Distribution and in the state of TX | Terminated | December 18, 2014 | 314 | 100580 | Computer | Not_Computer
Z-1198-2014 | 67464 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | GE Healthcare Infinia and Infinia Hawkeye Nuclear Medicine Systems. For Dual-Head Variable Angle Gamma Camera, Hawkeye Option for Dual-Head Variable Angle Gamma Camera (this includes GE Quasar Nuclear Medicine Systems). Nuclear Medicine Systems Infinia devices Product Usage: The intended use of the Quasar system is to perform general Nuclear Medicine imaging procedures for detection of radioisotope tracer uptake in the patient body, using a variety of scanning modes supported by various acquisition types and optional imaging features designed to enhance image quality in Oncology, Cardiology, Neurology and other clinical diagnostic imaging applications. |  |  | February 07, 2014 | 2014 | GE Healthcare, LLC | Potential safety issue related to unintended radial detector motion, which may occur during patient setup or at the end of a patient scan. This failure mode has been observed on a single system, in which the detector moved to the home position and then slowly starting drifting inward toward the patient. This generated both a visual and audio alert condition on both the console and the gantry. N | Device Design | GE sent an Urgent Medical Device Correction letter, dated 2/7/2014 to medical facilities beginning 2/7/2014. The letter was addressed to Manager of Nuclear Medicine, Hospital Administrators/Risk Managers, and Managers of Radiology/Cardiology. The letter described the issue and provided safety instructions. For Infinia and Infinia Hawkeye, Millennium VG, VG Hawkeye and VariCam nuclear Medicine Systems, the letter stated that a software upgrade would correct the issue. For the Helix Nuclear Medicine Systems, GE stated that a software upgrade can not be applied to the older systems and that additional information will be provided at a later date. UPDATE ON THE HELIX DEVICES: An Urgent Medical Device Correction letter, dated 4/7/2014 was sent to customers who still have a GE Helix Nuclear Medicine System. The letter requested to immediately stop using the device, and reiterated the March 2008 letter in which GE stated that they would no longer be able to offer full service support. The main reason for this was due to unavailability of parts required for maintenance and repair of the device because outside suppliers cease to produce the older parts. GE offered assistance in removal and disposal of the Helix system. |  | 3062 | 3062 units (1182 US, 1880 OUS) | Worldwide Distribution-USA (nationwide) including Puerto Rico and Guam and the countries of Algeria, Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, Finland, France, Georgia, Germany, Greece, Hong Kong, Hungary, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jordan, Korea (Republic of), Kuwait, Latvia, Lebanon, Lithuania, Malaysia, Malta, Martinique, Mexico, Morocco, Namibia, New Zealand, Norway, Pakistan, Panama, Philippines, Poland, Portugal, Russia, Saudi Arabia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Ukraine, and United Kingdom. | Terminated | November 03, 2014 | 269 | 3062 | Computer | Computer
Z-1781-2014 | 68111 | No | HTZ | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Synthes Matrix Mandible Short Cut Plate Cutter |  |  | April 22, 2014 | 2014 | Synthes, Inc. | It was discovered internally that the face of the Synthes Matrix Mandible Short Cut Plate Cutter has the potential for discoloration/corroded material in the affected lot. | Device Design | Synthes Inc. sent an Urgent Medical Device Recall Notice dated April 22, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.If you DO have any of the identified devices, please take the following steps:Contact DePuy Synthes at 1-800-479-6329 to obtain a Return Authorization (RA) Number . Complete the Verification Section at the end of this letter by checking the appropriate box indicating affected product has been located. Also, please indicate the number of devices found. Please include your name, title, telephone number and signature in the spaces provided.Return the Verification Section (page 3 of this letter) with the product to: Credit/Returns, DePuy Synthes, 1101 Synthes Avenue, Monument, CO 80132. Return a copy of the Verification Section (page 3 of this letter) to DePuy Synthes by: Fax: 866-324-3731 or Scan/email: Synthes7726@stericycle.comIf you DO NOT have the identified product, please take the following steps:Complete the attached Verification Section at the end of this letter by checking the appropriate box indication that no affecteded product has been located. Please include your name, title, telephone number and signature in the spaces provided this return documentation acknowledges your receipt of medical device removal information.Return the Verification Section (page 3 of this letter) to DePuy Synthes by:Fax: 866-324-3731 or Scan/email: Synthes7726@stericycle.comNote: If the Verification Section is answered on behalf of more than one facility and/or individual, please clearly indicate the name and address of the facility and/or individual on page 3 of the notification.If you have any questions, please call 610-719-5450 or contact your DePuy Synthes sales consultant. |  | 14 | 14 | US Distribution including the states of CA, CT, FL, GA, NY, SD and TX. | Terminated | September 10, 2015 | 506 | 14 | Computer | Not_Computer
Z-1147-2014 | 67430 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare Solar 8000M and Solar 8000i Patient Monitor software version 5.4, 5.5, or 5.6 with the Automatic-View-On- Alarm (AVOA) feature enabled. |  |  | February 13, 2014 | 2014 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue with the Solar 8000M/i Patient Monitor Automatic-View-On-Alarm (AVOA) feature.The Solar 8000M/i AVOA feature will stop functioning if the patient monitor is discharged from the CIC Central Station while a remote automatic bedside alarm view is currently being displayed on the Solar. Once this issue occurs, subsequent remote AV | Software design | GE Healthcare sent a "Urgent Medical Device Correction" letter dated February 12, 2014. The letter was sent to all affected customers. It identified the product the problem and the action needed to be taken by the customers. This issue affects all Solar 8000M/i Patient Monitors running software version 5.4, 5.5, or 5.6 with the Automatic-View-On-Alarm (AVOA) feature enabled.GE Healthcare will provide a correction at no charge once it is available. We will contact you to arrange for this correction.Please contact Technical Support at 1-800-558-7044 or your local Service Representative. |  | 10211 | 10,211 (7,549 units US, 2,662 units OUS) | Worldwide Distribution - USA (nationwide) including the states of AL, AZ, AR, CA, CO, CT, DE, DC, FL, GA, ID, IL, IN, KS, KY, LA, ME, MD, MA, MI, MN, NS, NO, NE, NH, NH, NM, NY, NC, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WV, WI and WY., and the countries of : AUSTRALIA, CANADA, AUSTRIA, BANGLADESH, BELGIUM, BOSNIA AND HERZEGOVINA, BRAZIL, CHILE, CHINA, COSTA RICA, CYPRUS, EGYPT, FINLAND, FRANCE, GERMANY, GREECE, INDIA, INDONESIA, JAPAN, MALAYSIA, MEXICO, NAMBIA, NETHERLANDS, NORWAY, POLAND, RUSSIA, SAUDI ARABIA, SINGAPORE, SOUTH KOREA, SPAIN, SWEDEN, TAIWAN, THAILAND, TURKEY, UNITED KINGDOM, SERBIA AND YUGOSLAVIA. | Terminated | July 10, 2015 | 512 | 10211 | Computer | Computer
Z-1839-2014 | 67610 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Altrus Thermal Tissue Fusion Handpieces under the following specifications: 1) 10mm x 16 cm Catalog Number: 60-9510-001, 2) 10 mm x 23 cm Catalog Number: 60-9511-001, 3) 10 mm x 36 cm Catalog Number: 60-9512-001, 4) 5 mm x 16 cm Catalog Number 60-9520-001, 5) 5 mm x 23 cm Catalog Number 60-9521-001 and 6) 5 mm x 36 cm Catalog Number 60-9522-001. Products are packaged within cartons, 6 units per carton. |  |  | February 20, 2014 | 2014 | ConMed Corporation | ConMed Corporation initiated an Urgent Medical Device Correction for Altrus Thermal Tissue Fusion Handpieces due to risk of burns. | Device Design | On February 20, 2014 and February 28, 2014 ConMed Corporation distributed Urgent Medical Device Correction notices to their domestic and foreign customers via FedEx Priority. The notice was generated by ConMed for information purposes only. Customers are to ensure that the jaws of the energy vessel sealing device does not come in contact with unintended tissue, surfaces, or objects until the jaws have cooled sufficiently. Customers who have further distributed this product, please forward a copy of the correction notice to the recipient as appropriate. Customers will questions should contact ConMed Regulatory Affairs at (315) 624-3237 or by email at altrus@conmed.com. |  | 12653 | 12,653 units (10,769 domestically &amp; 1,884 internationally) | Worldwide Distribution. | Terminated | April 07, 2017 | 1142 | 12653 | Computer | Not_Computer
Z-1812-2014 | 68419 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | HydroFinity Hydrophilic guidewire; Nitinol Guidewire with Hydrophilic Coating, Sterile. |  |  | May 28, 2014 | 2014 | Nitinol Devices and Components, Inc. | Reports of outer polymer jacket to the core wire being damaged during use on the HydroFinity Guidewire. | Device Design | Covidien sent an Urgent Product Recall letter dated May 30, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to return any unused devices to Covidien, their account would be credited. Customers were also instructed to::1) Stop using product listed in this letter immediately2) Segregate this product from other inventory3) Fill out the reply (verification) form at the end of this letterÂż If you do not have any product identified in this letter, please fax oremail the completed form to Covidien at (877) 523-9109 orcovidien8770@stericycle.com.Âż If you do have product, your sales representative will assist you incompleting the verification form and arranging for return of theproduct.Âż Please fax the completed form to Covidien (877) 523-9109 or email tocovidien8770@stericycle.com.4) Your sales representative will be available to answer any questions regarding this recall and assist you in completing the verification form, returning product and addressing any account credits.This action is being conducted with the knowledge of the United States FDA and other regulatory authorities. Âż Customers with questions were instructed to call (202) 310-5120.For questions regarding this recall call 202-310-5120. _ |  | 8485 | 8485 in US, 9595 - ROW - total, all models | Worldwide Distribution - USA (nationwide) and Internationally to Italy, France, Norway, Sweden, Germany, Belgium, United Kingdom and Canada. | Terminated | January 13, 2015 | 230 | 8485 | Computer | Not_Computer
Z-1082-2014 | 67223 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Desktop Pro TM linear accelerator control software intended to assist a licensed practitioner in the delivery of radiation to defined target volumes. |  |  | November 25, 2013 | 2013 | Elekta, Inc. | During the treatment table, gantry, and collimator calibration procedures, it is possible to type incorrect values or invert a value. | Software design | Corrective Action #1 - Release of Important Field Safety Notice 200 01 502 040 - "Recommended quality assurance after a calibration procedure" Released 25 November, 2013. All software versions (r7.01) detailed in this notice are affected by this recall. 200 01 501 041 - Recommended quality assurance after a calibration procedure Released November 25, 2013. Only software versions R6.0 and R6.1 are affected by this recall. Important Field Safety Notices are in the process of being distributed to customers advising them of the recommended Quality Assurance checks to avoid/detect the problem. The permanent solution to this problem is currently been developed. Corrective Action #2 - Permanent Solution - All users of Desktop Pro software will be informed of a mandatory upgrade to Integrity 1.2 upon its release which will prevent calibration from being inverted. |  | 1309 | 1309 | Worldwide Distribution: US (nationwide): AL, AK, AZ, AR, CA, CO, CT, FL, GA, ID, IL, IA, KY, LA, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, including Puerto Rico, DC; and countries of: Albania, Argentina, Australia, Austria, Bangladesh, Belarus, Belgium, Bosnia &amp; Herzego, Brazil, Bulgaria, Canada, Chile, China, Colombia, Croatia, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, Estonia, Finland, France, Gabon, Germany, Ghana, Greece, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Jordan, Libya, Lithuania, Malaysia, Malta, Mexico, Morocco, Netherlands, New Zealand, Norway, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Serbia, Singapore Slovakia, South Africa, South Korea, Spain, Sudan, Suriname, Sweden, Switzerland, Syria, Taiwan, Thailand, Turkey, United Kingdom, Ukraine, Venezuela, and Vietnam. | Open, Classified |  |  | 1309 | Computer | Computer
Z-1869-2014 | 68386 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MOSAIQ |  |  | May 29, 2014 | 2014 | Elekta, Inc. | MOSAIQ does not re-calculate the agent volume when the user changes the ordering dose (to other than what was calculated) prior to saving the order. | Software design | Elekta issued an Important Safety Notice dated May 29, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Corrective action that shall be taken:Notify the affected customers of the issue - To be completed 06/22/2014Follow up with the affected customers for:Confirmation receipt of the LIN, ANDConfirmation of installation of patch for the fix- To be completed 11/30/2014The notice includes a confirmation which is to be returned.For further questions call (770) 670-2422. |  | 101 | 101 | Worldwide Distribution - USA (nationwide) including Puerto Rico and the states of AL, AK, AZ, CA, CO, CT, DE, FL, GA, HI, IL, IN, IA, KY, LA, ME, MD, MA, MI, MO, MT, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, SC, SD, TX, VT, VA and WA,. and the countries of Australia, Canada, New Zealand and United Kingdom. | Open, Classified |  |  | 101 | Computer | Computer
Z-1867-2014 | 63486 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens StreamLAB(R) Analytical Workcell; Catalog number 746002.901. |  |  | October 11, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Firm has received reports from customers that the small pins at the bottom of the center door panel of the Input/Output Module, which protrude approximately 5/8 inch (16 mm), have contributed to trip and fall incidents. Customers are advised that when working in the area of the Input/Output Module to be aware of the two pins that protrude and use caution. | Device Design | The firm initiated their recall/correction of this product on October 11, 2012 by sending an Urgent Customer Notification letter to all StreamLAB(R) Automation Solutions customers. The letter informed customers that the firm received reports from customers that the small pins at the bottom of the center door panel of the Input/Output Module have contributed to two trip and fall incidents. Customers were advised that when working in the area of the Input/Output Module to be aware of the two pins that protrude and use caution. Also, StreamLAB operators were to be notified of the potential safety issue and the notice was to be posted on or near the system. Additionally, customers were to complete and return the attached Field Correction Effectiveness Check form. |  | 317 | 317 | Nationwide Distribution -- AL, AR, AZ, CA, CO, CT, FL, GA, ID, IL, IN, KS, LA, MA, MD, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NY, OH, OK, PA, SC, SD, TN, TX, UT, VA, WA, WI, and WV. | Terminated | August 26, 2014 | 684 | 317 | Computer | Not_Computer
Z-1866-2014 | 68259 | [6 empty] | LAP LASER PICTOR 3D Multi Color Laser System |  |  | March 21, 2014 | 2014 | LAP of America Laser Applications, LLC | While projecting yellow lines it was notice that the yellow projection has the potential to splits into two separate red and green projections. In this case the difference between the red and green projection was up to 2.5mm. | Software design | LAP of America Laser sent an Urgent safety Information FFST 2014-2-3 via UPS Next Day Air to the customer. |  | 1 | 1 system | North Carolina | Open, Classified |  |  | 1 | Computer | Not_Computer
Z-1065-2014 | 65451 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | All Elekta digital accelerators with electrons and beam modulator, MLCi/MLCi2, Agility or Asymmetric Heads |  |  | May 28, 2013 | 2013 | Elekta, Inc. | Customers are resetting the default values outside recommended factory settings. | Device Design | Elekta sent an Important Field Safety Notice dated May 2013 to all affected customers. The Notice identified the affected product, problem, clinical impact and actions to be taken. Customers were instructed to adhere to the instructions and advice given in manuals and system dialogues. For question contact your local Elekta office. |  | 472 | 472 units | USA Nationwide Distribution in the states of: AL, AK, AZ, AR, CA, CO, CT, FL, GA, ID, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NY, NC, ND, OH, OK, OR, PA, RI, SC, TN, TX, UT, VT, VA, WA, WV, WI, WY, DC and including Puerto Rico. | Open, Classified |  |  | 472 | Computer | Computer
Z-1850-2014 | 68128 | No | FSY | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Vertical Spring Arm for Single or Dual Monitor Mount (Harmony LL 500 and Harmony LL 700 surgical lighting systems), Steris Corporation. Model numbers YG19051 and YG19262. |  |  | April 04, 2014 | 2014 | Steris Corporation | Potential for the welded seam in affected monitor spring arms to crack and/or break. | Device Design | On April 11, 2014 the firm sent Urgent Voluntary Field Correction Notices to their customers. |  | 68 | 68 units | Distribution to US (states: MA, NY, OH, PA &amp; TX) and Canada. | Terminated | October 15, 2014 | 194 | 68 | Computer | Not_Computer
Z-1849-2014 | 68466 | No | CAN | Pressure regulator. | Anesthesiology | Anesthesiology | 510(K) Exempt | Model RG3000 and RG4000 Series Valve Integrated Pressure Regulator (VIPR) |  |  | April 07, 2014 | 2014 | Western / Scott Fetzer Company | Separation events have occurred, including the VIPR separating from the Aluminum cylinder. | Device Design | All initial consignees of subject product were notified by letter dated June 5, 2014 via U.S. mail. Initial consignees were directed to forward copies of the recall notification letters to their contacts in the event the subject product has been further distributed. |  | 118325 | 118,325 units | Worldwide Distribution-USA (nationwide) including Puerto Rico and the states of Alaska, Arkansas, Arizona, California, Colorado, Connecticut, Delaware, Florida, Georgia, Hawaii, Idaho, Illinois, Indiana, Kansas, Kentucky, Louisiana, Massachusetts, Maryland, Michigan, Minnesota, Missouri, Mississippi, Montana, North Carolina, North Dakota, Nebraska, New Hampshire, New Jersey, New Mexico, Nevada, New York, Ohio, Oklahoma, Oregon, Pennsylvania, Rhode Island, South Carolina, South Dakota, Tennessee, Texas, Utah, Virginia, Washington, Wisconsin, and West Virginia and the countries of Canada,China, Costa Rica, France, Honduras, India, Ireland, Italy, Japan, Korea (South), Malaysia, Panama, Spain and Taiwan. | Open, Classified |  |  | 118325 | Computer | Not_Computer
Z-1848-2014 | 68456 | No | FCN | Urine collector and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | SteriGear Urinary Drain Bag with Fig Leaf Cover, 2000 ml, REF 10270, The Fig Leaf, SteriGear 362 S. University Avenue, provo, UT 84601, 1-800-398-3259, SteriGear.com, Sterile, EO |  |  | January 11, 2011 | 2011 | SteriGear, LLC | SteriGear notifed customers of circumstances in which The Fig Leaf urinary drain bag valve can close creating a negative pressure that can suck the non-return valve closed. | Device Design | SteriGear, LLC sent a special note to all affected customers between January 12, 2011 and September 28, 2012. The note identified the product, the problem, and the action to be taken by the customer. Customers were instructed on how to reposition the inlet tube. Customers with questions were instructed to contact Director of Product Development at 801-607-3211. No correction or removal was performed.For questions regarding this recall call 801-607-3209. |  | 110560 | 110,560 | Nationwide Distribution | Terminated | January 02, 2015 | 1452 | 110560 | Computer | Not_Computer
Z-1847-2014 | 68256 | No | NEP | Intravascular administration set. | General Hospital | General Hospital | 510(K) Exempt | Abacus 3.1, Pharmacy Calculator |  |  | May 28, 2014 | 2014 | Baxter Corporation Englewood | Baxter Corporation is initiating a field correction following an upgrade to ABACUS Software v3.1, when Clinimix 4.25110 is used in an order, the Nutritional Summary and any labels containing energy information may display 1,000 times less protein than is actually present in the solution. | Software design | Baxter sent an Urgent Device Correction letter dated May 27, 2014 via certified mail to affected customers. The letter identified the affected product, problem and actios to be taken. Customers were instructed to: 1. Review and distribute this communication to pharmacy staff and compounding staff that use ABACUS software. This notice is intended only for customers that have upgraded to version 3.1 of the software.2. Please acknowledge receipt of this notification by completing the enclosed customer reply form and return it to Baxter by either faxing it to 1-224-270-5457 or scanning and e-mailing it to fca@baxter.com.3. If you have distributed ABACUS to other facilities, please notify the end users in accordance with your customary procedures. For questions call Baxter Techenical Support line at 1-800-678-2292. |  | 42 | 42 | US Nationwide Distribution and the state of VA. | Open, Classified |  |  | 42 | Computer | Not_Computer
Z-1845-2014 | 66315 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Siemens RAD Fluoro Uro System (Ysio, Luminos dRF, and Uroskop Omnia) with software versions VB10G. |  |  | August 28, 2013 | 2013 | Siemens Medical Solutions USA, Inc | When using systems operating with software versions VD10A/G during a RAD examination, a malfunction can result in the image appearing very dark, which may prompt unnecessary repetition of the examination. | Software design | Siemens sent a notification letter, dated August 28, 2013, to Affected customers. The letter identified the affected product, problem and actions to be taken. The letter provided work-around instructions and announced that a field modification was prepared to resolve the issue. |  | 76 | 76 | USA Nationwide Distribution in the states of OH, NJ, MA, CO, NY, CA, MD, OR, KY, PA, IL, VA, TX, MN, KS, GA, and WY. | Terminated | February 24, 2015 | 545 | 76 | Computer | Computer
Z-1156-2014 | 67495 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Panaroma Central Station |  |  | January 08, 2014 | 2014 | Mindray DS USA, Inc. dba Mindray North America | Mindray has identified two issues with the Panorama Central Station. Full Disclosure data may lag behind the Panorama system time clock when the Panorama CentralStation is in use with MindrayÂżs V-Series Monitor. Additionally, event indicators lag behind the Panorama system time clock when the Panorama Central Station is in use with a Panorama Telepack. | Software design | Mindray DS USA Inc. sent an Urgent Corrective Action letter dated January 8, 2014 via certified mail to their customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to contact Mindray Service Representative, at 1-800-288-2121 (Monday  Friday 8:30 am-5:30 pm), to arrange for the software update of the Panorama System(s) in their facility. This work will be performed at no cost. |  | 732 | 732 units in the US | Worldwide Distribution - USA Nationwide Distribution | Terminated | January 21, 2015 | 378 | 732 | Computer | Computer
Z-1154-2014 | 67470 | No | HKI | Ophthalmic camera. | Ophthalmic | Ophthalmic | 510(k) | RetCam 3, RetCam Portable and RetCam Shuttle; |  |  | January 24, 2014 | 2014 | Clarity Medical Systems Inc | Software anomaly for RetCam 3, RetCam Shuttle and RetCam Portable with software versions 6.0.x, 6.1. x and 6.2x where the patients age is calculated incorrectly. | Software Manufacturing/Software Deployment | Urgent Safety Alert letters were sent to Risk Manager and to NICU US consignees via FedEx on 1/24/2014.Template of letters to OUS distributors to be sent on distributors letterhead and sent to device recipients via electronic communication and./or hard copy on 1/24/2014. |  | 480 | 166 Ret Cam Shuttles, 288 RetCam 3 and 23 RetCam Portables. | Worldwide Distribution-USA (nationwide) including DC and the states of NM, LA, CA, MI, PR, GA, MD, MT, AL, MA, NY, IL, OH, TX, VA, MI, FL, MO, KY, TN, CT, NJ, PA, UT, SD, OK, MS, and NC and the countries of Algeria, Australia, Austria, Bangladesh, Belarus, The Netherlands, Brazil, Bulgaria, Chile, China, Columbia, Czech Republic, Denmark, Ecuador, Egypt, Finland, France,India, Indonesia, Iran, United Kingdom, Israel, Italy, Kuwait, Lebanon, Malaysia, Mexico, Sweden, Pakistan, Peru, Philippines, Poland, Portugal, Saudi Arabia, Taiwan, Thailand, Turkey, UAE, Argentina, Uzbekistan, Oman, Canada, Germany, and Switzerland. | Terminated | August 11, 2014 | 199 | 480 | Computer | Computer
Z-1163-2014 | 67257 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Innova 2100IQ, Innova 3100, Innova 3100IQ, Innova 4100, Innova 4100IQ Innova 2121IQ, Innova 3131IQ; Cardiovascular X-ray imaging systems (X-ray generator installed with Rotor V3 version and Heater V4 version boards). |  |  | December 11, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has become aware of a potential safety issue involving the innova systems. The Innova system can unexpectedly stop delivering X-rays after a power-on or a reset cycle which may result in the loss of real-time interventional imaging. No injuries have been reported due to this issue. | Software design | GE Healthcare sent an "Urgent Medical Device Correction " letter dated December 10, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letter described the Safety Issue, The long term solution is to provide a software correction for all of the affected systems that were distributed. The plan for public warning is to contact all customers with an affected system by means of an Urgent Medical Device Correction letter within the US. In the US this will be sent by a traceable means (either US Postal Service with receipt acknowledgement (registered mail) or equivalent courier service such as FEDEX). This letter will be sent to the following titles within the affected accounts:GE Healthcare will correct all affected systems at no cost to you. A GE Healthcare service representative will contact you to arrange for this correction. If you have any questions or concerns regarding this notification, please call one of the following phone numbers: United States: 800 437 1171For other countries, please contact your local GE Healthcare Service Representative. |  | 98 | 257(159 US, 98 OUS) | Worldwide Distribution - USA (nationwide) including the states of AL, AK, AZ, AR, CA, CO, CT, FL, GA, IL, IN, KS, KY, LA, ME, MD, MA, MI, MS, MO, NV, NJ, NY, NC, ND, OH, OK OR, PA, PR, RI, SC, TN, TX, UT, VA,WA,WY and WI., and the countries of ARGENTINA, AUSTRALIA, CANADA, CHINA, CZECH REPUBLIC, DENMARK, EGYPT, FINLAND, FRANCE, GERMANY, HUNGARY, INDIA, IRELAND, ITALY, JAPAN, JORDAN, REPUBLIC OF KOREA, MALAYSIA, MEXICO, NEW ZEALAND, NORWAY, PHILIPPINES, POLAND, RUSSIA, SAUDI ARABIA, SPAIN, SRI LANKA, TURKEY and UNITED KINGDOM. | Terminated | July 29, 2014 | 230 | 98 | Computer | Not_Computer
Z-1177-2014 | 67463 | No | LIB | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | bioMerieux PREVI Isola System (General purpose,microbiology, diagnostic Medical Device), (REF 29500/29500R), bioMerieux, Hazelwood, MO. |  |  | February 10, 2014 | 2014 | Biomerieux Inc | The firm has determined the product may fail to dispense the sample to the agar plate resulting in a "failure to dispense" in conjunction with urine specimens. This could lead to a "false" negative growth result. | Software design | The firm sent a recall notification letter to all of their consignees on 2/10/2014. |  | 471 | 471 instruments | Distribution was made to AK, CA, CO, CT, DC, FL, GA, IL, IN, KY, LA, MA, MO, NC, NJ, NY, OH, OR, PA, RI, SC, TN, TX, VA and WI. | Terminated | March 16, 2017 | 1130 | 471 | Computer | Not_Computer
Z-1832-2014 | 68320 | equipment, laboratory, general purpose, labeled or promoted for a specific medical use | LXG | General purpose laboratory equipment labeled or promoted for a specific medical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VersaCell X3 Solution and VersaCell X3 Expansion Pack, a robotic laboratory sample-management system. It connects up to 3 instruments and provides a user interface that routes samples, reduces manual tasks and consolidates results. |  |  | April 25, 2014 | 2014 | Siemens Medical Solutions Diagnostics | The keyboard and monitor may not be properly secured to the monitor arm. | Device Design | Siemen's Healthcare Diagnostics sent an Urgent Medical Device Correction letter dated April 25, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to complete and return teh Field Correction Effectiveness Check form attached to the letter within 30 days by fax to 302-631-7597.Inspect the monitor, keyboard and bracket assembly. If the screw and washer assembly is secure as shown in Figure 1, the monitor and keyboard assembly is safe to use. After making adjustments or repositioning the monitor, keyboard and bracket assembly, inspect the screw adn washer to ensure it is ssecure as shown in Figure 1.If there is a gap between the underside of the monitor support bracket and cylindricalmounting feature as shown in Figure 2, contact your Siemens technical supportrepresentative to schedule a service visit as soon as possible. Front load and manage samples directly on the connected instruments until your Siemens technical support representative addresses the issue.Customers were instructed to retain the letter with their laboratory records, and forward the letter to those who may have received the affected product. Customers with questions were instructed to contact their Siemens Customer Care Center or their local Siemens technical support representative.For questions regarding this recall call 877-229-3711. |  | 19 | 19 units | Worldwide Distribution - USA including Iowa and New York and Internationally to Canada, Chile, France, India, Italy, Malaysia, Peru, Korea, and the United Kingdom. | Terminated | April 16, 2015 | 356 | 19 | Computer | Computer
Z-1165-2014 | 67515 | Yes | DTR | Cardiopulmonary bypass heat exchanger. | Cardiovascular | Cardiovascular | 510(k) | MAQUET PLEGIOX Cardioplegia Heat Exchanger |  |  | January 31, 2014 | 2014 | Maquet Cardiovascular Us Sales, Llc | The MAQUET post-marketing vigilance program has identified a discrepancy that thedesired cardioplegia temperature may not be obtained when the PLEGIOX CardioplegiaHeat Exchanger (PLEGIOX) is used in conjunction with a heater/cooler unit. | Device Design | Maquet Inc.sent an" Field Safety Notice" dated January 31, 2014, to all affected customers via Fed Ex. The letter identified the product the problem and the action needed to be taken by the customer.Any questions or additional information, can be obtained from the MAQUET representative, or MAQUET Customer Service at 1-888-627-8383 (press option 2, followed by option 3), Monday through Friday, between the hours of 8:00 a.m. and 6:00 p.m. EST.Please complete and return the enclosed confirmation form, acknowledging your receipt and understanding of this communication, by emailing a scanned copy to Fieldactions@maquet.com or via fax at (973) 807-9290. |  | 1424 | 1,424 units | Worldwide Distribution - USA (nationwide) and the countries of Brazil, Germany and Turkey | Terminated | April 24, 2015 | 448 | 1424 | Computer | Not_Computer
Z-1827-2014 | 68311 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Philips MR systems: Trolley Variable Height IRF (for Ingenia 1.5T &amp; 3.0T products) |  |  | May 15, 2014 | 2014 | Philips Medical Systems, Inc. | Ingenia customers have experienced clamping of the foot under the central column of the Height Adjustable Flex Trak trolley and the Trolley Variable Height IRF. | Device Design | Philips sent an Urgent Field Safety Notice (FSN78100414) dated May 2014 via certified mail to US and Canadian consignees. Philips Key Market representatives will distribute the letter internationally. The letter identified the affected product, explains the problem, actions that should be taken in order to prevent risks for users, and the actions planned by Philips to correct the problem. On May 31, 2014, Philips released a Field Change Order that requires service engineers to replace the front and rear covers of the Ingenia MR FlexTrak Trolley with new covers with protection bumper that prevents foot clamping. For further information and support concerning this issue, consignees should contact their local Philips representative at 1-800-722-9377. |  | 5 | 5 | Worldwide Distribution - US (nationwide), Canada, Mexico Australia, Austria, Germany, Belarus, Bulgaria, Brazil, China, Denmark, Estonia, Finland, France, Germany, Greece, Italy, Jordan, Korea, Republic of Korea, Netherlands, New Zealand, Norway, Oman, Pakistan, Russian Federation, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, and Vietnam. | Terminated | August 07, 2015 | 449 | 425 | Computer | Not_Computer
Z-1825-2014 | 68535 | No | ISW | External assembled lower limb prosthesis. | Physical Medicine | Physical Medicine | 510(K) Exempt | RUSH81 Prosthetic Foot |  |  | June 02, 2014 | 2014 | Ability Dynamics LLC | Ability Dynamics is recalling RUSH81 Prosthetic Foot because the existing bolts may be bottomed out in the threaded hole and this could lead to bolt failure. | Device Design | Ability Dynamics sent an Urgent Medical Device Correction Notice letter dated June 2, 2014 to all affected customers (Prosthetist) who purchased the RUSH Prosthetic Foot. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to fill out the Bracket Bolt Replacement Form and return to Ability Dynamics. Customers with questions are instructed to contact Rich Sainz, National Clinical Manager, 2082 E. University Dr. Tempe AZ 85281 (480) 717-1412, rsainz@abilitydynamics.com. |  | 175 | 175 units | Worldwide Distribution - US (nationwide) and in the countries of Australia and Canada. | Terminated | September 22, 2014 | 112 | 175 | Computer | Not_Computer
Z-1173-2014 | 67528 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Device Name: ABX PENTRA ML |  |  | February 14, 2014 | 2014 | Horiba Instruments, Inc dba Horiba Medical | HORIBA Medical initiated this recall of all software versions of ABX PENTRA ML (Model A11A06152) due to a software bug with the potential of transmitting incorrect results to the Laboratory Information System. | Software design | Horiba sent an "URGENT SAFETY NOTICE letter datedFebruary 13, 2014 via by mail to customers. The letter identified the affected product, issue, impact, actions to be taken. Customers were instructed to share this information with laboratory staff and retain this notification as part of their laboratorydocumentation. Customers were requested to complete and return the enclosed response form via one of the following methods:Fax: 949-453-0600Email: jim.knowles@horiba.comMail:HORIBA MedicalAttention: Jim Knowles, Recall Coordinator34 Bunsen DriveIrvine, CA 92618.For questions contact your local Horiba Medical representative. |  | 15 | 15 | USA Nationwide Distribution | Terminated | August 22, 2014 | 189 | 15 | Computer | Computer
Z-1807-2014 | 68430 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | iGUIDE System |  |  | May 29, 2014 | 2014 | Elekta, Inc. | A bug in the software prevents the iGUIDE software from logging off. | Software design | Elekta sent an Important Field Safety Notice 618 01 303 024 dated May 29, 2014 to affected customers. The letter identified the affected product, problem, potential clinical impact and actions to be taken. It also gives a recommendation how the risk can be mitigated - Users should only use the iGUIDE software when the key is in the on position and the Position Guard is active. Users are advised to restart the workstation if any error occurs to ensure the software is running properly. The notice also advises the customer that a patch will be released that corrects the behavior of the iGUIDE software. This solution will be implemented on all affected products when available. For questions contact your local Elekta office. |  | 224 | 224 | Worldwide Distribution - USA (nationwide) in the states of AZ, CA, CT, HI, IL, IA, KY, MD, MA, MI, MO, NH, NJ, NC, ND, OH, OK, OR, PA, TX, VA, WA, WI, and the countries of Algeria, Australia, Austria, Brazil, Canada, China, Colombia, Denmark, Finland, France, Germany, Hong Kong, India, Italy, Japan, Kazakhstan, Mexico, Netherlands, New Zealand, Peru, Poland, Russia, Saudi Arabia, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, United Kingdom and Vietnam | Open, Classified |  |  | 224 | Computer | Computer
Z-1175-2014 | 67157 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare CARESCAPE Monitor B650 and CARESCAPE Monitor B850. |  |  | February 28, 2014 | 2014 | GE Healthcare, LLC | GE Healthcare has recently become aware of potential safety issues associated with the CARESCAPE Monitor B850 and CARESCAPE Monitor B650. These issues include both NIBP (Non-Invasive Blood Pressure) and Bed to Bed Issues. | Software design | GE Healthcare sent an Urgent Field Safety Notice dated February 26, 2014, to all International consignees. The letter was addressed to Healthcare Administrator / Risk Manager, Chief of Nursing and Director of Biomedical Engineering. The letter described the Safety Issues and Instructions, Affected Product Details Product Correction and Contact Information ( Contact Technical Support at 1-800-558-7044 or their local Service Representative).GE Healthcare determined that there were no US sites affected since all units were under GE control therefore, no customer letter was deployed in US.For questions regarding this recall call 847-277-4472. |  | 676 | 676 units OUS | Worldwide Distribution - USA (nationwide) and Internationally to AUSTRIA, BELGIUM, CZECH REPUBLIC, FINLAND, FRANCE GERMANY, GREECE, ITALY, LITHUANIA, POLAND, SPAIN, SWEDEN, SWITZERLAND, UNITED KINGDOM. | Terminated | April 29, 2015 | 425 | 676 | Computer | Computer
Z-1806-2014 | 68162 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | RayStation software version 4.0.3.4 |  |  | April 25, 2014 | 2014 | RAYSEARCH LABORATORIES AB | This notice concerns a problem with the computation of ROI voxel volumes for cases where the CT has variable slice spacing. The dose grid volumetric representation of an ROI can be wrong. This affects all dose-volume properties for such ROIs, including DVH, dose statistics, clinical goals and constraints or objective functions. Furthermore, if material override ROIs are used, or if the external is not represented by contours, the bug can trigger an error in the 3D dose distribution. | Software design | On April 25, 2014, RaySearch emailed all consignees of the affected product a copy of the recall notice as well as the response forms. Back up notification methods include phone calls and site visits. |  | 6125 | 25 sites/245 units (US), 25 sites/124 units (foreign) | Worldwide Distribution-USA (nationwide) including the states of AZ, CA, FL, IL, MA, MD, MO, MS, NJ, NV, NY, OH, PA, TN, VA, and WA, and the countries of Australia, Belgium, Germany, France, Italy, New Zealand, Republic of Korea, Switzerland, Spain, and United Kingdom. | Terminated | June 26, 2017 | 1158 | 6125 | Computer | Computer
Z-1194-2014 | 67459 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart MRx Monitor/Defibrillator |  |  | February 18, 2014 | 2014 | Philips Medical Systems, Inc. | When the HeartStart MRx.is used with the Q-CPR Meter in defibrillation mode, the Q-CPR Meter may incorrectly display the Do Not Touch the Patient icon. | Software Design Change | The firm, Philips Healthcare, sent an "URGENT -Medical Device Correction/Field Safety Notice" letter dated February 2014 on February 18, 2014 to its domestic and foreign customers. The letter describes the product, problem and actions to be taken. The customers were instructed to follow the Action to be Taken by Customer/User section of the attached Field Safety Notice; to continue to use your HeartStart MRx without the Q-CPR Meter attached. If you choose to use the HeartStart MRx with the Q-CPR Meter attached and you observe that the Q CPR Meter is displaying the Do Not Touch the Patient icon at an unexpected time, follow the MRx Instructions for Use and remove the meter from the patient and continue CPR according to your organization's protocol.Philips plans to take to remedy the issue. Philips will provide a software upgrade free of charge.If you have questions regarding this notification or need any further information or support, please contact your local Philips representative or call at 1-800-722-9377. |  | 6583 | 6583 units | Worldwide distribution: US (nationwide) and to countries of: ARUBA AUSTRALIA AUSTRIA BELGIUM BRAZIL CANADA ECUADOR EGYPT FINLAND FRANCE GERMANY HONG KONG INDIA INDONESIA IRELAND ISRAEL ITALY KENYA KOREA, REPUBLIC OF KUWAIT LATVIA LIBYAN ARAB JAMAHIRIYA MALAYSIA MEXICO MOROCCO NETHERLANDS NORWAY OMAN PAKISTAN PHILIPPINES POLAND QATAR RUSSIAN FEDERATION SAUDI ARABIA SINGAPORE SOUTH AFRICA SPAIN SWEDEN SWITZERLAND TAIWAN, PROVINCE OF CHINA THAILAND TURKEY UKRAINE UNITED ARAB EMIRATES and UNITED KINGDOM. | Open, Classified |  |  | 6583 | Computer | Computer
Z-1800-2014 | 68355 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MOSAIQ, an oncology information system used to manage workflows for treatment planning and delivery. It supports information flow among healthcare facility personnel and can be used wherever radiotherapy and/or chemotherapy are prescribed. |  |  | May 23, 2014 | 2014 | Elekta, Inc. | When using CMA, if the user hovers the mouse pointer exactly above the direction dropdowns and then uses the mouse thumbwheel or touchpad scroll bar, the direction selection might change when the intention was only to move the mouse to a different area of the screen. | Software design | Elekta sent an Important Safety Notice on May 23, 2014, to all affected customers. The notice identified the product the problem and the action needed to be taken by the customer."Notify the affected customers of the issue - To be completed 06/22/2014"Provide a fix to the defect in an expeditious timeframe. It is mandatory for all customers to apply the fix - Version 2.50.05 SP8 &amp; Version 2.60 - To be released 05/16/2014"Follow up with the affected customers for: -Confirmation receipt of the LIN, AND -Confirmation of installation of patch for the fix- To be completed 11/30/2014. For further questions please call (404) 993-5886. |  | 723 | 723 | Worldwide Distribution - USA (nationwide) including Puerto Rico, and the countries of Albania, Algeria, Argentina, Australia, Austria, Bahrain, Bangladesh, Belarus, Belgium, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, Chile, China, Colombia, Croatia, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, Estonia, Finland, France, Gabon, Germany, Ghana, Greece, Honduras, Hungary, India, Iran, Iraq, Ireland, Israel, Italy, Japan, Jordan, Libya, Lithuania, Malaysia, Malta, Mexico, Morocco, Myanmar, Netherlands, New Zealand, Norway, Pakistan, Panama, Peru, Poland, Portugal, Romania, Russia, Saudi Arabia, Serbia, Slovakia, Slovenia, South Africa, Spain, Sudan, Suriname, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom, Ukraine and Venezuela. | Open, Classified |  |  | 723 | Computer | Computer
Z-1724-2014 | 68255 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Synergy XVI |  |  | April 23, 2014 | 2014 | Elekta, Inc. | XVI can incorrectly calculate the target position of the treatment table. | Software design | Elekta is responsible for the correction of all products with this defect.1. Notifications regarding the possibility of this software error and subsequent table location error and potential hazards have been sent out. The notifications include steps to take to reduce the possibility of the occurrence of an injury.2. A software correction has been released for some versions of the XVI software. All versions of software with this error will be corrected by the end of July 2014.3. The software correction will enable verification that values have been assigned to table position parameters; failure to pass verification will inhibit the system. The software correction will also detect internal communication errors (incomplete registrations) and inhibit the system when those errors are detected.4. The update will be mandatory and will be made at no cost to the users.CDRH approves this CAP. You may proceed with implementation of the CAP.For further questions please call (770) 670-2422. |  | 380 | 380 US | Worldwide Distribution - USA including the states of AZ, CT, IA, OH, MD, MI, MN, NV, TX and VA., and the countries of India, Ireland, Malta, Netherlands, South Africa and UK. | Terminated | July 17, 2017 | 1181 | 380 | Computer | Not_Computer
Z-1801-2014 | 68216 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | EZPass Suture Retriever, Nitinol, Single Pack, Part number 10007380. Orthopedic Manual Surgical Instrument, Sterile, single use arthroscopic instrument utilized in conjunction with a nitinol wire to aid in passing suture through soft tissue. |  |  | April 15, 2014 | 2014 | Biomet, Inc. | The nitinol wire or monofilament may buckle within the EZPass Suture Passer and become kinked making the wire/monofilament difficult to use. | Device Design | On 2/28/14, customers received a letter stating these devices have been recalled following an investigation which identified that the nitinol wire or monofilament may buckle within the EzPass Suture Passer and become kinked making the wire/monofilament difficult to use. The letter also stated, there is no adverse health outcome expected for the patient and this action requires the immediate location and discontinued use of the product and its return to Biomet. |  | 100 | 100 units | Nationwide Distribution. | Terminated | July 07, 2015 | 448 | 202 | Computer | Not_Computer
Z-1785-2014 | 68292 | No | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Convatec Flexi Seal Control Fecal Management System Rx Only |  |  | April 29, 2014 | 2014 | Convatec Inc. | Convatec Inc. is recalling Flexi-seal FMS Control Kit since it does not have a 510K. In addition, an internal assessment of product performance, including a review of customerfeedback, has confirmed that Flexi-Seal Control is not meeting expectations or those of customers. Specifically, the Auto-Valve feature that is unique to the Flexi-Seal Controlhas not been consistently performing relative | Device Design | Convatec Inc. sent an Urgent Voluntary Medical Device Recall letter dated April 29, 2014 via UPS. The letter identified the affected product, problem and actions to be taken. Customers are asked to immediately stop distributing and quarantine all recalled lots. Affected products must be returned to Distributor. Contact your Distributor for shipping instruction. Customers should completed the attached return response form and return it to the address on the response form. Questions can be directed to Gregory Kowzun at (336) 547-3710. Convatec issued a press release on 7/3/2014. |  | 97887 | 73,021 units US; 24,866 units OUS | Worldwide Distribution - USA Nationwide | Terminated | August 28, 2015 | 486 | 97887 | Computer | Not_Computer
Z-1668-2014 | 68197 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON SC2000 Ultrasound Systems with software version 3.5 (VA35) and using Stress Echo Feature. Ultrasound imaging system. |  |  | April 23, 2014 | 2014 | Siemens Medical Solutions USA, Inc. | Potential malfunction due to a software issue for ACUSON SC2000 ultrasound systems using software version 3.5 (VA35) when using the stress echo feature which could cause the need for the study to be repeated. | Software design | Siemens sent a Consumer Safety Advisory Notification letter on April 23, 2014, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use of the stress echo feature with software version .5. For studies acquired with systolic-only captures: review all patient studies on exams performed after software revision 3.5 was installed on your system. If you are unsure when this software was installed on your system,please contact your local Customer Service representative. Foreign distributors will implement distribution of the field correction to their customers.For questions regarding this recall call 650-694-5398. |  | 177 | 177 units | Worldwide Distribution - USA (nationwide) and Internationally to Australia, Belarus, Dominican Rep., France, Germany, India, Indonesia, Israel, Japan, Mexico, New Zealand, Norway, Philippines, Poland, Republic of Korea, Romania, Russian Fed., Saudi Arabia, Singapore, Sweden, Switzerland, Turkey, U.A.E., Ukraine, and United Kingdom. | Terminated | November 12, 2015 | 568 | 177 | Computer | Computer
Z-1667-2014 | 68195 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON SC2000 Ultrasound Systems with software version 3.5 (VA35); ultrasound imaging system. |  |  | April 23, 2014 | 2014 | Siemens Medical Solutions USA, Inc. | Software issues with ACUSON SC2000 ultrasound with software version 3/5 (VA35) - may incorrectly display the color flow data in the wrong position and when Doppler Frequency is adjusted, the system does not restart the pulsed wave Doppler or continuous wave Doppler waveform. | Software design | Siemens sent an Customer Safety Advisory Notification letter on April 23, 2014, to all affected customers. Distributors in foreign countries will implement delivery of notification to their customers." Use the factory default presets with software version 3.5 instead of user-defined presets restored from software version 2.0or 3.0. Review all patient studies on exams performed after software revision 3.5 was installed on your system. If you are unsure when this software was installed on your system, please contact your local Customer Service representative." After adjusting the Doppler frequency, freeze and unfreeze the system. This action refreshes the display of the Doppler spectrum resulting in accurate measurements. Additional measurements performed after system freeze results in correct measurement values.Your local Customer Service Engineer will contact you to schedule an appointment to replace your existing version of software, free of charge. If you have any questions, please contact your local service support person forinformation regarding timelines and status.Until the new software has been installed on your system, please share this information with all personnel within your organization who need to be aware of this issue. As always, patient safety issues are a high priority, and we alert customers of product concerns which may occur under various conditions. To date, no patient injury has been reported as a result of these issues.We sincerely regret any inconvenience this condition may cause in your daily operations. For further questions call (650) 694-5398. |  | 257 | 257 units | Worldwide Distribution - USA including the states of AL, AZ, AR, CA, CO, CT, FL, GA, ID, IL, IN, KS, LS, MD, MI, MN, MO, NH, NY, NC, OH, OR, PA, SC, SD, TN, TX, UT, VA, WA and WI., and the countries of Australia, Austria, Dominican Rep, France, Germany, Iceland, India, Indonesia, Ireland, Israel, Italy, Japan, Mexico, New Zealand, Norway, Philippines, Poland, Portugal, Republic Korea, Romania, Russian Fed, Saudi Arabia, Sweden, Switzerland, Turkey, U.A.E. and United Kingdom. | Terminated | November 12, 2015 | 568 | 257 | Computer | Computer
Z-1665-2014 | 68092 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | The Centricity PACS Workstation is intended for use as a primary diagnostic and analysis tool for diagnostic images by trained healthcare professional, including radiologists, physicians, technologists, clinicians, and nurses. It is also intended for use as a clinical review workstation throughout the healthcare facility. The workstation interface provides the user with a means to display, manipulate, archive, print, and export images when connected wit the Centricity PACS infrastructure. |  |  | April 03, 2014 | 2014 | GE Healthcare | It has been identified that in a reject image workflow where multiple studies are simultaneously opened on the Centricity PACS RA1000 workstation and images are rejected from the active displayed study, images for the previously displayed study may also be rejected. The issue can occur with a previously displayed study that remains open in the background. If one or more, but not all, images in | Software design | An URGENT MEDICAL DEVICE CORRECTION letter dated April 3, 2014 was sent to all direct accounts. The letter included instruction to customers to use until the permanent fix for this issue is deployed: 1) Check to see if the DelRejImg task is active and running. a) Log into CA Tool. b) From the Navigation menu, select System Monitor. c) From the pull down menu, select Dbase Maintenance. d) Scroll to the DelRejImg task (this list is alphabetical). e) Review the Task Status. Possible statuses are IDLE or SUSPEND f) If the task status is set to SUSPEND The task is not running. g) If the task is set to IDLE, move to Step 2. 2) Check the Minimum Rejected Image Deletion Time. a) From the Navigation menu, select Sys. Configuration. b) From the pull down menu, select PACS Configuration. c) Scroll to Minimum Rejected Image Deletion Time (this list is alphabetical). d) Review the Value field. e) If the value field is a negative integer (such as the default setting of -1), the Delete Rejected Image process is NOT running and there is no patient safety hazard. If the value field is a positive integer (such as 1 or higher), the Delete Rejected Image process IS running and there is a potential patient safety hazard. 3) A GE Healthcare Service Representative will remove the DelRejImg task from the task table in the Centricity PACS database. This will prevent the DelRejImg task from running, and prevent any rejected images from being deleted. This feature is being removed entirely from the 3.2 release family of Centricity PACS and will be reinstated in a future 4.X workstation release of Centricity PACS. 4) Perform the reject image workflow with only one patient study open at a time. A future release of Centricity PACS will address the unintentional rejected image issue. Until a correction for the unintentional rejected image issue is available, a GE Healthcare service representative will remove the DelRejImg task table in the Centricity PACS database. This will remove the re |  | 297 | 297 units | Worldwide Distribution-USA (nationwide) including the states of AK, AL, AZ, CA, CO, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NE, NH, NJ, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, VA, VT, WA, WI, and WV, and the countries of Australia, Canada, Chile, France, Germany, India, Ireland, Israel, Italy, Japan, Korea (South), Kuwait, Lebanon, Netherlands, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Taiwan, Turkey, and United Kingdom. | Terminated | June 05, 2015 | 428 | 297 | Computer | Computer
Z-1445-2014 | 67755 | No | OOE | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | Alcon LenSx Ophthalmic Laser, for use in cataract surgery for the creation of corneal cuts/incisions, anterior capsulotomy and laser phacofragmentation. |  |  | March 14, 2014 | 2014 | Alcon LenSx, Inc. | Alcon LenSx received reports of unexpected downward motion of the gantry. | Device Design | Alcon LenSx Inc advised customers of a Medical Device Correction being initiated for specific serial numbers of the LenSx Laser System for use in cataract surgery for the creation of corneal cuts/incisions, anterior capsulotomy and laser phacofragmentation.Alcon sent customer notification letters via mail during the week of March 17, 2014. The letter titled "Medical Device Correction", advised customers of the correction and provided: the product's description with serial number(s), description of the potential condition, advise on action to be taken by the user, transmission of the notice, and a contact reference person. The letter also contained a section for customers to sign and return the page of the Notice to Alcon. |  | 75 | 75 | Worldwide Distribution. | Terminated | October 16, 2014 | 216 | 75 | Computer | Not_Computer
Z-1315-2014 | 67549 | [6 empty] | Omni Micro-electrode/reference electrode for cobas b221 analyzer, Model/Catalog/Part Number: |  |  | February 28, 2014 | 2014 | Roche Diagnostics Operations, Inc. | Reference electrode used beyond the guaranteed in-use 52 week lifetime,may leak and potentially cause erroneous pH and/or sodium results. Other Ion parameters are not affected. | Device Design | Roche Diagnostics Operations sent an Urgent Medical Device Correction letter dated February 27, 2014 via UPS Ground (receipt required) to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete the enclosed fax form and fax it to 1-888-345-0253. For questions contact Roche Customer Business Support Center at 1-800-428-5076. |  |  |  | Worldwide Distribution - USA nationwide including PR and the country of Guam. | Open, Classified |  |  |  | Computer | Not_Computer
Z-1569-2014 | 67927 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Shimadzu Digital Radiography X-Ray System, Catalog No. DAR-8000f. |  |  | April 03, 2014 | 2014 | Shimadzu Medical Systems | Shimadzu Corporation is recalling the Shimadzu Digital X-Ray System because they have received a report of a lost image under certain circumstances. | Software design | A Service Bulletin and a customer notification letter were sent on 4/3/2014 for all the customers who purchased the Shimadzu Digital X-Ray System (Catalog No. DAR-8000f). The notification informs the customers of the problems identified and the actions to be taken. Customers with questions or concerns about the device correction, are instructed to contact their local Shimadzu representative or contact Shimadzu National Technical Support at (800) 228-1429. |  | 685 | 685 units total (14 units in US) | Worldwide Distribution-USA (nationwide) and the countries of Japan, China, Germany, Australia, Thailand, Sri Lanka, South Africa, UAE, Saudi Arabia, Korea, and Taiwan. | Terminated | June 25, 2015 | 448 | 685 | Computer | Computer
Z-1565-2014 | 63145 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | SpineFrontier S-Lift Lock Shaft Inserter, an instrument used with S-LIFT Intervertebral Body Fusion Device in spinal fusion surgery. Product Usage: The S-Lift Lock Shaft SI50001 mates with SI50002, Implant Inserter, to aid in securing the implant to the whole implant inserter assembly, SI50000. |  |  | December 23, 2011 | 2011 | SpineFrontier, Inc. | It has been verified through field use that there are two possible breakages that may occur when using SI50001 Revision D Lock Shaft Inserter, possibly resulting in instrument fragments entering the sterile field:Impact Cap: The impact cap on the Lock Shaft Inserter can shear off under high-impact force during insertion and removal.Pin: The pin used to guide the Lock Shaft Inserter into the ma | Device Design | Spine Frontier sent an Urgent Notification dated December 23, 2011, consignees. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete the attached acknowledgement form. For questions call 866-914-7717 |  | 30 | 30 devices | US Nationwide - OR, FL, MA, TX, CO, OK, IL, and MO | Terminated | May 20, 2014 | 879 | 30 | Computer | Not_Computer
Z-1318-2014 | 67716 | No | FLF | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | 3M" Steri-Vac" Gas Sterilizer/Aerators, Model #'s 4XL,5XL, 8XL |  |  | March 17, 2014 | 2014 | 3M Company - Health Care Business | Recently, 3M received a single report of a potential exposure to Ethylene Oxide (EO) at a hospital installation involving 3M" Steri-Vac" Gas Sterilizers. The cause was traced to a rare situation involving a failure of the EO cartridge puncture assembly. This assembly was only in production during the period November 10, 2006 through December 14, 2007 and in service during the period February 7, 20 | Device Design | 3M sent an "Urgent Medical Device Field Correction" letter dated March 6, 2014. The letter was sent to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. You are being notified because our records indicate you received a unit that contains the cartridge puncture assembly that is the subject of this field correction. 3M will contact you to schedule an appointment for a field correction on any unit in the range of serial numbers provided in the table. 3M's goal is to service the affected units with minimal disruption to your facility. Units manufactured outside of the dates November 10, 2006 to December 14, 2007 are not affected. .For Further Information contact 3M Health Care Service Line at 1-800-228-3957 or your local 3M representative. |  | 419 | 73 US, 346 OUS | Worldwide Distribution - USA (nationwide) Distribution including the states of AL, AZ, CA,CO, FL, GA, HI, IL, IN, IA, KS, KY, MD, MA, MI, MN, MS, NE, NJ, NY, NC, ND, OH, OK,OA, SC, TX, WA, WV and WI. and the countries of COLOMBIA, VENEZUELA, UAE, THAILAND, GERMANY, CHINA, KOREA, CANADA, JAPAN, RUSSIA, INDIA, POLAND, PERU, AUSTRIA, ISRAEL, INDONESIA, SPAIN, CHILE, MALAYSIA, SWITZERLAND, EL SALVADOR, ECUADOR, LEBANON, MEXICO, ARGENTINA, BULGARIA, TAIWAN, TURKEY,ROMANIA, GUATEMALA and NEW ZEALAND. | Terminated | January 22, 2015 | 311 | 419 | Computer | Not_Computer
Z-1316-2014 | 67586 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | IMMULITE 2000; IMMULITE 2000XPi; VersaCell |  |  | February 03, 2014 | 2014 | Siemens Medical Solutions Diagnostics | There is a potential for incorrect results to be reported from processing of a sample tube. | Device Design | Siemens sent an Urgent Medical Device Correction letter dated February 3, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The customers was told how to correct the issue if it has occurred, and how to prevent the issue from occurring. The completed form should be faxed to (302) 631-7597. Any questions can be directed to Siemens Technical Support Representative. Siemens plans to implement a correction at a later time. |  | 1006 | 1006 units | Nationwide and Foreign Distribution. | Open, Classified |  |  | 1006 | Computer | Not_Computer
Z-1549-2014 | 68067 | No | FMT | Infant radiant warmer. | General Hospital | General Hospital | 510(k) | PandaÂż Freestanding Infant Warmers, integrated with Nellcor SpO2. |  |  | April 02, 2014 | 2014 | GE Healthcare | Medical device software may be defective and may cause user-set alarm limits to result in false positive or false negative Oxygen saturation and pulse rate alarm notifications to the caregiver. | Software design | GE Healthcare notified Biomedical/Clinical Engineering managers, Labor &amp; Delivery/ NICU Nurse managers, Respiratory Therapy Managers and Directors of Risk Management by letter 04/02/2014. End users were advised that they may continue to use the system provided they follow the GE recommended actions:1) If in use with a patient, turn off the Nellcor SpO2 function in the control panel and switch to an approved alternate form of SpO2 measurement.2) Prior to using the affected GE warmer with the Nellcor SpO2 function, or resuming the use of this function with any patient, contact your hospital Bio-Med or qualified technician to check the Nellcor SpO2 alarm functionality using either an adult Nellcor SpO2 sensor or a Nellcor SpO2 simulator following the applicable instructions in the Appendix 3) If the test activates a Low Pulse Rate alarm, the user-adjusted Nellcor SpO2 alarm limits are functioning properly and the warmer can be placed back in clinical service.End users were also advised that alarm functionality must be re-checked following every service event requiring access to the internal electronics and during annual preventative maintenance, until revised software is installed.4) If the test does NOT activate a Low Pulse Rate alarm, remove the warmer from clinical service and contact your local GE Healthcare Service Representative. |  | 21 | 21 units | Worldwide Distribution. | Open, Classified |  |  | 1110 | Computer | Computer
Z-1322-2014 | 67574 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Integrity 1.1. |  |  | March 28, 2013 | 2013 | Elekta, Inc. | The problem is that the "static tolerances" from the calibration files (loaded database) are used for beam delivery instead of the machine calculated values. | Software design | An Important Field Safety Notice (IFSN) has been released and distributed to all customers. Integrity 1.1 Incorrect Static Tolerances for Linac Geometric Axes, 200 01 502 032. The notice explains the issue and that the user should apply the static Tolerance Updater tool to check tolerances and adjust as required. Important Field Safety Modification Digital Accelerator FCO Ref: 200 02 502 031 Integrity R1.1 Static Tolerance Updater Tool was released 03-2013. Elekta has released Integrity version 1.2 which contains the permanent fix for this issue. The permanent solution fixes the error that made the digital accelerator protection functions use the value loaded from the calibration database. The solution now ensures the protection functions use the values the digital accelerator calculates. Important Field Safety Notice 200 01 502 035 has been released to inform customers the solution is available. Important Field Safety Modification 200 02 502 036 will track the installation of the software by Elekta service engineers. |  | 498 | 498 | Worldwide Distribution-USA (nationwide) including the states of AL, AZ, CT, FL, GA, KY, LA, MD, MO, NM, NY, NC, OH, OK, PA, TN, TX, and WA and the countries of Argentina, Australia, Austria, Belarus, Belgium, Brazil, Canada, China, Colombia, Croatia, Czech Republic, Denmark, Ecuador, Egypt, Estonia, France, Germany, Greece, Honduras, Hong Kong, India, Indonesia, Iraq, Ireland, Italy, Japan, Jordan, Malaysia, Mexico, Morocco, Myanmar, Netherlands, New Zealand, Nigeria, Norway, Pakistan, Panama, Peru, Poland, Romania, Russia, Saudi Arabia, Senegal, Singapore, Slovakia, South Korea, Spain, Sudan, Sweden, Switzerland, Taiwan, Thailand, Turkey, and United Kingdom. | Terminated | July 26, 2017 | 1581 | 498 | Computer | Computer
Z-1331-2014 | 67741 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Siemens Medical Solutions, ACUSON P300 Ultrasound system with P300 PA 230E cardiac probe (Esaote product 7348). |  |  | March 07, 2014 | 2014 | Siemens Medical Solutions USA, Inc. | ACUSON P300 ultrasound systems using the PA230 transducer with a user-defined preset which includes the 100% setting for output power could cause the transducer surface temperature to exceed the maximum allowable limit. | Device Design | Siemens sent an Customer Safety Advisory Notifications on March 7, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.A new version of software resolving this error will be provided free of charge. Your local Customer Service Engineer will contact you to schedule an appointment to make the necessary repairs on your system. If you have any questions, please contact your local service support person for information regarding timelines and status. Until the necessary repairs have been completed on your system, please share this information with all personnel within your organization who need to be aware of this issue. As always, patient safety issues are a high priority. To date, no patient injury has been reported. This problem was discovered as part of our ongoing quality process.For further questions please call (650) 694-5398.We sincerely regret any inconvenience this condition may cause in your daily operations. |  | 675 | 675 | Class II Recall - Worldwide Distribution - US Distribution and the countries of Australia, Austria, Belgium, Bhutan, Bolivia, Botswana, Bulgaria, Canada, Chile, Colombia, Croatia, Dem. Rep. Congo, Denmark, Ecuador, France, Germany, Greece, India, Indonesia, Ireland, Israel, Italy, Ivory Coast, Japan, Kuwait, Malaysia, Mauritius, Mongolia, Morocco, Mozambique, Myanmar, Namibia, Netherlands, Norway, Peru, Philippines, Poland, Portugal, Republic Korea, Romania, Saudi Arabia, Serbia, Seychelles, Singapore, South Africa, Spain, Switzerland, Taiwan, Trinidad,Tobago, Turkey, U.A.E. and United Kingdom. | Terminated | November 12, 2015 | 615 | 675 | Computer | Computer
Z-1335-2014 | 67758 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | 5X250 MM Curved Cannula, Arm 2 |  |  | March 13, 2014 | 2014 | Intuitive Surgical, Inc. | Intuitive Surgical has become aware of the potential for certain lots of the Single-Site 5mm Curved Cannulae to become damaged during use. The second part of this action is to update Single-Site labeling. These labeling changes instruct users how to inspect cannulae prior to use and to appropriately handle cannulae during use. | Device Design | Intuitive sent an Urgent Medical Device letter dated March 25, 2014, to all users of cannulas that have the potential to be damaged during use as well as a second letter to all single site cannula users that provided clarity to instruction for use for inspections and handling of the product.Certain lots manufactured prior to August of 2012, and users were asked to use Attachment A to help identify affected lots numbers. Phase 1 - Cannulas should then be inspected with information provided in Attachment B. Any damaged devices should be quarantined and returned to intuitive. If inspection does not find damage, the device may continue to be used. Any other cannulae and any non damaged product should be used according to Attachment C's instructions for use.Phase 2 - Once replacement cannulae are available, the undamaged cannulae should be returned.The Acknowledgement form should be completed and returned to Intuitive. Intuitive will contact consignees when replacement cannulae are available.Questions should be directed to Customer Service at 800-876-1310.Those consignees that use the 5 mm canulae but do not have affected product were provided with a letter that only includes the Single Site Cannula inspection and Use addendum. |  | 5217 | 5217 total cannulae | Worldwide Distribution - USA and the countries of Australia, Belgium, Brazil, Chile, China, Columbia, Denmark, France, Germany, India, Italy, Lebanon, Monaco, Netherlands, Norway, Romania, Saudi Arabia, South Korea, Spain, Sweden, Switzerland, Turkey and United Kingdom. | Terminated | September 03, 2014 | 174 | 5217 | Computer | Not_Computer
Z-1528-2014 | 67947 | Yes | MRN | Nitric oxide administration apparatus. | Anesthesiology | Anesthesiology | 510(k) | INOmax DSIR, model 10007, Nitric Oxide Delivery System, with Software version 3.0.0. The INOmax DS delivery system delivers INOMAX (nitric oxide for inhalation) therapy gas and provides continuous monitoring of inspired O2, NO2, and NO, and a comprehensive alarm system. The primary targeted clinical setting is the Neonatal Intensive Care Unit (NICU) and secondary targeted clinical setting is the transport of neonates. |  |  | March 27, 2014 | 2014 | INO Therapeutics (dba Ikaria) | An anomaly impacts the monitoring system of INOmax DSIR Plus (software 3.0.0). When the integrated pneumatic backup switch is activated with low or no gas flow through the injector module, an intended NO cell failure alarm will occur. In addition, an unintended NO2 cell failure alarm may also occur. The NO/NO2 cell failure alarms are caused by an average NO/NO2 circuit concentration greater tha | Software Design Change | Ikaria sent an Urgent Field Corrective Action letter via FedEx on March 27, 2014. The letter identified the issue, and requested that all INOmax DSIR with software version 3.0.0 be returned to Ikaria. Customers with questions were instructed to contact Ikaria Customer Care at 1-877-566-9466.For questions regarding this recall call 877-566-9466. |  | 22 | 22 | Nationwide Distribution including DC and VA. | Terminated | November 21, 2014 | 239 | 22 | Computer | Computer
Z-1512-2014 | 67802 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | GemStar Docking Station, used with GemStar infusion pump, in combination with other GemStar accessories such as the GemStar Lock Box and patient bolus cord. The docking station is also used for hosting the GemStar pump for general infusions. List Numbers: 13075-XX-XX. |  |  | February 18, 2014 | 2014 | Hospira Inc. | There are two situations that may occur when using the GemStar Docking Station, List Number 13075-XX-XX, in conjunction with the GemStar infusion pump: 1) when the Docking Station is used in conjunction with a GemStar Phase 3 pump (List 13000-XX, 13100-XX, or 13150-XX) the potential exists for the GemStar Phase 3 pump to fail to power up while connected to the Docking Station, and 2) when either a GemStar Phase 3 (List 13000-XX, 13100-XX, or 13150-XX)or GemStar Phase 4 pump (List 13086-XX, 13087-XX, or 13088-XX) is used in conjunction with both a Docking Station and an External Battery Pack accessory (List 13073-XX) there exists a possibility that the GemStar pump will display error code 11/003 and give an audible alarm indicating excessive input voltage from the external sources. If the GemStar pump detects what is perceived to be more than 3.6 Volts measured on the external voltage input, the pump will stop the infusion and alarm both with an audible sound as well as a visual alarm. | Device Design | Hospira sent an URGENT MEDICAL DEVICE CORRECTION letter dated February 18, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The letter included the following instructions: 1) To avoid failure to power up, turn the pump on first, before connecting the pump with the docking station. This will prevent the failure to power up. 2) To mitigate the potential for an 11/003 error code, remove the External Battery Pack accessory (List 13073) from the Docking Station in conjunction with an External Battery Pack accessory (List 13073), this practice should not continue. Please contact Hospira to discuss an appropriate alternative. Additional instructions included: 1) Inform potential users within the recipient's organization of the notification. 2) Should the facility experience a failure to power up or an 11/003 error code, report the issue to Hospira by calling 1-800-441-4100 (Monday - Friday, 8:00 AM - 5:00 PM CT) or e-mail ProductComplaintsPP@hospira.com. 3) Compelte the attached reply form and return it to the fax number or e-mail address on the firm, even if there is no affected product in inventory. 4) If th products were further distributed, notify the accounts that received the products identified in the notification and ask them to contact Steriycle at 866-792-5451 (Monday - Friday, 8:00 AM - 5:00 PM ET) to receive a reply form.For questions regarding this recall call 800-441-4100. |  | 26765 | 26,765 docking stations total (6866 in US) | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY, District of Columbia, Guam, Puerto Rico, and Internationally to Australia, Austria, Bahrain, Belgium, Canada, Chile, Egypt, France, Germany, Greece, Hong Kong, Italy, Jordan, Kuwait, Lebanon, Libya, Luxemburg, Netherlands, New Zealand, Oman, Portugal, Saudi Arabia, Slovenia, South Africa, Spain, Switzerland, Turkey, United Arab Emirates, and United Kingdom. | Terminated | January 23, 2017 | 1070 | 26765 | Computer | Computer
Z-1437-2014 | 67691 | No | HWT | Template for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | GLOBALÂż Anchor Peg Glenoid Plus |  |  | March 28, 2014 | 2014 | DePuy Orthopaedics, Inc. | Urgent Device Correction notifications are being sent to USA distributors with follow-up to surgeons concerning updated surgical techniques stressing the possibility of the shoulder pin guide breaking and potentially being left in the patient. | Device Design | Beginning 3/28/12014 DePuy United States Distributors and DePuy affiliate contacts for Canada and International will be notified via email. The US Distributors will be responsible for adhering to the directions outlined in the field communication, including identifying the Surgeon users and notifying them either by mail or in person with the Surgeon Letter and providing access to revised surgical techniques. The notice emphasized technical points regarding the use of the GLOBAL APG+ 2.5mm Breakaway Guide Pin (singleuse shoulder instrument) that have been added to the surgical techniques. Risks were explained and surgical techniques revised were identified and available through the distributor. Distributors were asked to identify surgeons using the device, filling out the response form. For clinical questions from surgeons, please contact DePuy Orthopaedics, Inc.s Scientific InformationOffice at 18885542482 (MF; 8 a.m.  5 p.m. EDT).For product related questions, please contact your local DePuy Orthopaedics, Inc., Sales Consultant." For questions about recall information provided: Kim Earle, Recall Coordinator, 5743714917 (MF; 8 a.m.  5 p.m. EDT) |  | 21675 | 21,675 | USA Nationwide Distribution including Canada | Terminated | December 04, 2014 | 251 | 21675 | Computer | Not_Computer
Z-1438-2014 | 67862 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart XL+ Defibrillator/Monitor with system |  |  | April 02, 2014 | 2014 | Philips Medical Systems, Inc. | Software communication failure may occur on the HeartStart XL+ locking the user out of clinical mode and possibly causing a delay in therapy. | Software design | Philips Healthcare sent an Urgent Medical Device Correction letter dated March 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. This Field Safety Notice is intended to inform you about:" what the problem is and under what circumstances it can occur" the actions that should be taken by the customer/user in order to prevent risks for patients or users" the actions planned by Philips to correct the problem.A software communication failure may occur on the HeartStart XL+ locking the user out of clinical mode and possibly causing a delay in therapy.Please see the attached Field Safety Notice that provides information on how to identify affected devices and instructions on actions to be taken. Please follow the ACTION TO BE TAKEN BY CUSTOMER / USER section of the notice. To correct this issue Philips will provide a software upgrade to customers with affected units free of charge.This voluntary correction has been reported to the appropriate regulatory agency.If you need any further information or support concerning this issue, please contact your local Philips representative at 1-800-722-9377.Philips apologizes for any inconveniences caused by this problem. |  | 9872 | 9872 units | Worldwide Distribution - USA (nationwide) including the countries of: ARGENTINA AUSTRALIA AUSTRIA BELGIUM CANADA CHINA CZECH REPUBLIC DENMARK FRANCE GERMANY ITALY LATVIA NETHERLANDS NEW ZEALAND NORWAY PHILIPPINES PORTUGAL SOUTH KOREA SPAIN SWITZERLAND THAILAND TURKEY and UNITED KINGDOM. | Open, Classified |  |  | 9872 | Computer | Computer
Z-1449-2014 | 67762 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Cobas c8000 Modular Analyzer Series Cobas c502, Part Number 5964067001, a clinical chemistry analyzer intended for the in vitro quantitative / qualitative determination of analytes in body fluids. |  |  | March 13, 2014 | 2014 | Roche Diagnostics Operations, Inc. | The Initial Cassette Volume Check (ICVC) is not activated on the cobas c 502 module of the cobas 8000 modular analyzer series when manually filled cobas c packs are loaded. This issue occurs with tests that use cobas c pack MULTI or empty pre-labeled cobas c packs. The ICVC feature allows the reagent probe to dive into the reagent cassette and confirm the reagent level in the bottles. This ensures | Software Manufacturing/Software Deployment | March 17, 2014 all consignees will be notified by Urgent Medical Device Correction sent via UPS Ground (receipt required). The notice identified the issue, root cause, affected assays with clinical significance. Consignees were instructed to follow the steps listed in the Urgent Medical Device Correction to correctly activate the ICVC and verify the correction. If your system is not connected to the Internet through the cobas link, and you do not have the option to run the affected reagents on an alternate system (i.e., cobas c 311 or cobas c 501 analyzer), please contact the Roche Support Network Customer Support Center at 1-800-428-2336 to have an update scheduled via a Roche Field Representative. If your system is connected to the Internet through the cobas link, follow the steps below:" Remove all reagent cobas c packs on board your cobas c 502 module for the assays listed on pages 1-2." Dispose of these reagent cobas c packs according to your local waste management ordinances." Load new cobas c packs for the affected assays, and visually verify the ICVC is performed." Complete the attached fax form (5326-00-0314) and fax it to 1-888-276-6166." Please consult with the physician or pathologist at your facility to determine specific clinical implications for your patients." File this Urgent Medical Device Correction (UMDC) for future reference. |  | 76 | 76 units | USA Nationwide Distribution including San Juan PR | Terminated |  |  | 76 | Computer | Not_Computer
Z-1662-2014 | 68211 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX CV Pediatric Echo Reporting, IMPAX CV Echo Measurement Import via Optical Character Recognition (OCR) Service IMPAX. For the creation of digital structured reports for cardiovascular diagnostic studies such as echocardiography. |  |  | April 30, 2014 | 2014 | AGFA Corp. | Failure to effectively install and test a hotfix to correct the known software defect within the IMPAX CV 7.8 SU3-OCR Service. | Software design | AGFA Healthcare sent an URGENT FIELD SAFETY NOTICE LETTER dated April 30, 2014, to the consignee describing the problem that had occurred, the actions being taken by Agfa and the actions Agfa has required the site to perform. As part of the correction Agfa requires the consignee to review the affected reports to confirm that the data is correct. If errors are found the consignee will need to correct the data. The consignee was asked to acknowledge, via FAX-Back to 864-421-1664 or email, that the information was received and understood by the consignee. For questions the consignee was instructed to contact their local Agfa HealthCare organization at 877-777-2432.For questions regarding this recall call 864-421-1754. |  | 1 | One | Distribution to SC only | Terminated | September 18, 2014 | 141 | 1 | Computer | Computer
Z-1510-2014 | 67901 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Nuclear gamma cameras (ADAC Vertex Plus, ADAC Solus, ADAC Vertex V60 &amp; ADAC Vertex Classic) Philips Medical System, Cleveland, OH 44143. |  |  | March 10, 2014 | 2014 | Philips Medical Systems (Cleveland) Inc | Slippage of the radius drive belt in the relative 180 degree orientation of the superior positioned detector head will allow the detector head to drift down and potentially come in direct contact with the patient. | Device Design | On 03/10/2014, the firm sent Urgent - Medical Device Correction letters to their customers. The letter identified the affected product, the reason for recall, the hazard involved, and the actions that should be taken by customers. Customers are to cease use of the identified procedures until implementation of the Field Change Order. Additionally, customers should sign and return the attached Customer Response Form within 10 days of receipt, acknowledging that they have read and understood the Field Safety Notice and Addendum. The Field Safety Notice must be placed in customers' User Documentation. Customers are to contact their local Philips Healthcare representative or office, if they need further information or support. |  | 1043 | 1,043 cameras | Worldwide Distribution -- USA, and the countries of Australia, Austria, Belgium, Brazil, Canada, China, Denmark, Egypt, France, Germany, Hong Kong, Ireland, Italy, Korea, Lebanon, Malaysia, Netherlands, Norway, Pakistan, Portugal, Russia, Singapore, Spain, Sweden, Thailand, &amp; United Kingdom. | Terminated | August 04, 2015 | 512 | 1043 | Computer | Not_Computer
Z-1513-2014 | 67719 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ExacTrac 6.0 is a patient positioning and monitoring system. |  |  | February 24, 2014 | 2014 | Brainlab AG | When using multiple radiation treatment targets within a single plan, the ExacTrac v.6.0.3 might position the patient to an unintended radiation treatment point. As a result, there is a possibility that the radiation treatment dose may be delivered to the unintended target position. | Software design | A FIELD SAFETY NOTICE / PRODUCT INFORMATION letters, dated February 24, 2014, were sent to all direct accounts. The letter identified the affected product, the reason for the recall, and included instructions for direct accounts to take while a permanent fix can be deployed. The letter also advised that a software update will be made available starting the end of April 2014. Customers with questions or concerns can contact Brainlab at 800-597-5911 or via e-mail at us.support@brainlab.com. |  | 41 | 41 systems (13 systems in US; 28 systems to foreign countries) | Worldwide Distribution -- USA, including the states of CA, IN, LA, MA, OH, TX, and WI, and the countries of Australia, China, France, Germany, Japan, Netherlands, and the United Kingdom. | Open, Classified |  |  | 41 | Computer | Computer
Z-1484-2014 | 67523 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | SIGMA SPECTRUM Volumetric Infusion Pump with Master Drug Library |  |  | February 07, 2014 | 2014 | Baxter Healthcare Corp. | Baxter Healthcare Corporation has issued an Urgent Device Correction for the SIGMA SPECTRUM Volumetric Infusion Pump with Master Drug Library due to repeated System Error 322 occurrences. | Software design | On February 7, 2014, Baxter Healthcare Corporation distributed Urgent Device Correction notices and Customer Reply Forms to their customers via first-class mail. Affected customers are instructed to contact Baxter Healthcare Corporation at 1-800-356-3454 (choose option 1) Monday through Friday during the hours of 7:00 am to 7:00 pm Eastern Time. Baxter technicians would determine if an adjustment or repair is necessary. Customers are asked to complete the attached Customer Reply Form and return it to Baxter either by faxing to 1-224-270-5457 or e-mailing a scanned copy to fca@baxter.com. |  | 277096 | 277,096 units (revised) | Worldwide Distribution - USA Nationwide including PR and the country of Canada. | Terminated | May 25, 2017 | 1203 | 277096 | Computer | Not_Computer
Z-1511-2014 | 68032 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Healthcare, CardiacVX [CardiacVX (1.1.0 Ext. d) and CardiacVX Flow (1.1.0 Ext. d) analysis software running on Advantage Workstation]. Model #'s 5434389-2 and 5434384-2. |  |  | April 07, 2014 | 2014 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue due to higher than actual flow values reported on CardiacVX and CardiacVX Flow running on Advantage Workstation. | Software design | GE Healthcare sent an Urgent Medical Device Correction" letter dated August 21,2013, to all affected consignees. The letter was addressed to Hospital Administrators /Risk Managers, Radiology Department Managers and Radiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. Customers were instructed until the issue is corrected on their system with the affected CardiacVX or CardiacVX Flow software for phase-contrast protocols customers use the fast CINE phase-contrast sequence with the magnitude weighting user CV option disabled. For directions on how to update user CV, please refer to the MR operator manual. If customers choose to perform imaging using magnitude weighting, do not utilize CardiacVX or CardiacVX flow for flow value calculations.with the affected product. Customers with questions were instructed to contact their local service representative.For questions regarding this recall call 262-513-4122. |  | 101 | 101 (1 US , 100 OUS) | Worldwide Distribution - USA including IN and Internationally to AUSTRALIA, CANADA, CHILE, FRANCE, ITALY, JAPAN, KOREA, KUWAIT, NEW ZEALAND, PANAMA, POLAND, SAUDI ARABIA, SPAIN, SWEDEN, SWITZERLAND, UNITED KINGDOM, | Terminated | May 29, 2014 | 52 | 101 | Computer | Computer
Z-1504-2014 | 68027 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Siemens ACUSON X700 Ultrasound Systems with software version 1.0.00, 1.0.01, 1.0.02, 1.1.00 and 1.1.02. ultrasound imaging system. |  |  | April 02, 2014 | 2014 | Siemens Medical Solutions USA, Inc. | When fourSight 4D ultrasound imaging, 3-Scrape real time 3D imaging, or Anatomical M-mode are used during an Obstetric or Gynecology study, data from the previous patient could replace the current patient's date in the report. | Software Design Change | Siemens sent a Customer Safety Advisory Notification letter on April 2, 2014 to all affected customers via certified mail. The letter identified the affected product, problem and what actions to take until software is upgraded. For questions call 1-888-826-9702. |  | 490 | 490 | Worldwide Distribution - US Nationwide and the countries of Angola, Argentina, Armenia, Australia, Austria, Bangladesh, Belgium, Bosnia Herzeg., Brunei, Chile, Colombia, Croatia, Ecuador, Egypt, France, Germany, Greece, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Kenya, Kuwait, Latvia, Libya, Liechtenstein, Malaysia, Mauritius, Montenegro, Namibia, Nepal, Nigeria, Norway, Palestinian Ter, Philippines, Poland, Portugal, Qatar, Rep. of Yemen, Republic Korea, Romania, Saudi Arabia, Slovakia, South Africa, Spain, Switzerland, Taiwan, Turkey, Turkmenistan, U.A.E., United Kingdom. | Terminated | November 12, 2015 | 589 | 490 | Computer | Computer
Z-1495-2014 | 67868 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Philips MDC PACS R2.X / IntelliSpace PACS DCX R3.1 software, software |  |  | April 02, 2014 | 2014 | Philips Medical Systems, Inc. | Philips MDC PACS Release R2.x and Philips IntelliSpace PACS DCX R3.x, may have incorrect density measurement on Enhanced CT/MR examinations | Software design | Philips Healthcare sent a Urgent Medical Device Correction Notification/Field Safety Notice dated March 2014, to inform customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue.Philips will replace the software free of charge.Philips is asking customers to follow the Action to be Taken by Customer/User section of the Urgent Medical Device Correction notification/Field Safety Notice:Until you receive a software upgrade from Philips, the user must configure the MRI and CT devices to send regular CT or MR DICOM image objects (not enhanced) to PACS.. This issue does not affect images that are not enhanced image objects.Further questions please call (978) 687-1501. |  | 373 | 373 units | Worldwide Distribution - US Distribution including Puerto Rico and the states of CA, FL, IL, NJ, OK and WA., and the countries of : Argentina Australia, Austria, Azerbaijan, Bahrain, Bangladesh, Bosnia, Brazil, Bulgaria, Chile, China, Egypt, Ethiopia, France, Germany, Ghana, Greece, Guinea, India, Iraq, Israel, Italy, Latvia, Libya, Lithuania, Malawi, Malta, Mexico, Mozambique, Nepal, New Zealand, Papua, New Guinea, Paraguay, Philippines, Poland, Portugal, Romania, Russia, Slovakia, Slovenia, Solomon Islands, South Africa, Spain, Thailand, Turkey, United Arab Emirates and United Kingdom. | Open, Classified |  |  | 373 | Computer | Not_Computer
Z-1502-2014 | 67896 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | On-Board Imager - Exact Arm Elbow Motor Shaft (kVD, kVS, and MV) with ''Type 01" elbow motors. |  |  | March 20, 2014 | 2014 | Varian Medical Systems, Inc. | Varian has identified a failure with the shaft of the Type 01 elbow motor for Exact arms used in the On-Board Imager (OBI) device due to belt over-tensioning. | Equipment maintenance | Varian sent an Urgent Medical Device Correction letters March 20, 2014 to affected custoers via trackable mail. The letter was also distributed to Varian Sales, Marketing and Service organizations to inform them of the issues. The letter identified the affected product, description of problem, details, and recommended user actions to be taken. A Varian Customer Support Service representative will arrange a site visit to check the bel tension of the Type 01 elbow motors. If over tensioned belts are detected, the elbow motor will be replace with a Type O2 elbow motor. Questions may be directed to 888-827-4265. |  | 2622 | 2622 | Worldwide Distribution. | Terminated | May 01, 2015 | 407 | 2622 | Computer | Not_Computer
Z-1492-2014 | 67678 | No | LHN | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Proteus 235 Proton Therapy System Legacy Software version 7.2.7 and higher, models 110, 112, 113, and 116. |  |  | December 04, 2013 | 2013 | Ion Beam Applications S.A. | Internal testing of the Proteus 235 Proton Therapy System revealed incorrect dose rate calculation in case of beam line option other than Zero (0). | Software design | Ion Beam Applications ( IBA) sent a recall notification letter dated on/about December 5, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Short term action: A user notice (see annex 1) has been sent to impacted sites where the affected legacy PTS software version was delivered. This user notice explains the workaround required to be put in place in order to fully mitigate the risk and the 1BA action plan to permanently fix this issue.Corrective action:A patch on the PTS Legacy SW was requested and in implementation. In the same time, the solution was put in main code of PTS legacy software version 7 to be integrated in the next release.The customer representative signing underneath confirms that the user notice above is communicated to the appropriate employees.We apologize in advance for any inconvenience this may cause and we would like to thank you for your cooperation.The IBA Site Staff or the Operations manager is available to provide you with additional information and/or guidelines if necessary.Further questions please call : 215-972-7777 |  | 4 | 4 | US Distribution including the states of OK, IL, NJ, and WA. | Terminated | June 12, 2014 | 190 | 4 | Computer | Not_Computer
Z-1442-2014 | 67781 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | SpaceLabs Arkon Anesthesia Delivery System with Version 2.0 Software installed. This system includes a ventilator. |  |  | March 10, 2014 | 2014 | Del Mar Reynolds Medical, Ltd. | The System may go into a controlled failed state if a series of parameters on the user interface are activated in a specific sequence. | Software in the Use Environment | Spacelabs Healthcare sent an Urgent Medical Device Correction letter, dated 11 March 2014, to consignees. The letter identified the product, the problem, and the actions to be taken by the customer. The letter stated that Spacelabs Field Service personnel will be contacting consignees to schedule a convenient time to install, at no cost, a software upgrade.Consignees with questions about this corrective action program, please contact Spacelabs Healthcare at 1-800-522-7025 and select 2 for Technical Support. |  | 16 | 16 units distributed in the US | USA Nationwide Distribution in the states of North Carolina and South Carolina | Terminated | August 20, 2014 | 163 | 16 | Computer | Computer
Z-1488-2014 | 67905 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | GE Vivid E9 Ultrasound System, a general-purpose ultrasound system, specialized for use in cardiac imaging. |  |  | March 27, 2014 | 2014 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue where the probe power surveillance settings for 4 probes (C1-5-D, C2-9-D, IC5-9-D and M5Sc-D) are incorrect. This may result in a situation where probes can overheat. | Software design | GE Healthcare sent an "Urgent Medical Device Correction" letter dated March 27, 2014. The letter was addressed to Hospital Administrators / Risk Managers, Managers of Radiology / Managers of Cardiology. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information (Contact your local GE Healthcare service representative). |  | 447 | 447 (81US, 366 OUS) | USA Nationwide Distribution in the states of: CA, CO EL, GA, IN, KS, KY , LA, MA, MI, MO, MT, NY, NM, NV, US: OH, TX, VA, WI.OUS: ARABIC EMIRATES, AUSTRALIA, AUSTRIA, BELGIUM, COLOMBIA, CZECH REPUBLIC, DENMARK, FINLAND, FRANCE, GERMANY, GREAT BRITAIN, INDIA,IRELAND, ITALY, JAPAN, KOREA, NETHERLAND, NORWAY, POLAND, PORTUGAL, RUSSIA, SINGAPORE, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN. | Terminated | January 14, 2015 | 293 | 447 | Computer | Computer
Z-1464-2014 | 67763 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Ingenuity Core 128, Computed Tomography X-ray systems, Philips Medical Systems (Cleveland), Inc. |  |  | March 06, 2014 | 2014 | Philips Medical Systems (Cleveland) Inc | A customer reported that after upgrading to software version 3.5.5 from 3.5.4, reconstructions would intermittently not complete. | Software in the Use Environment | .The firm, Phillips, sent an " URGENT ÂżÂż Medical Device Correction" letter dated2013 March 06, to its costumers . The letter disc ribes the product, problem and actions to be taken. The costumers were instructed to Verify that all planned image reconstructions are completed, and Perform off-line reconstructions if expected images are missing. The firm will release A Field Change Order (FCO) to correct the issue with software version release 3.5.7.and A Philips Field Service Engineer will contact you to schedule the software installation at your site.If you need any further information or support concerning this issue, please contact your local Philips representative or local Philips Healthcare office. For North America and Canada, contact the Customer Care Solutions Center (1-800-722-9377, Option 5: Enter Site ID or follow the prompts). |  | 116 | 116 units | Worldwide Distribution:US ( Nationwide) Including stages of MI, OR, PA, TX &amp; WV; and Countries of: Australia, Bangladesh, Belgium, Brazil, Bulgaria, China, France, Germany, India, Indonesia, Ireland, Japan, Nicaragua, Oman, Palestine, Russia, South Korea, Spain, Switzerland, Thailand, Turkey &amp; United Kingdom. | Terminated | November 05, 2014 | 244 | 116 | Computer | Computer
Z-1462-2014 | 67867 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | The OEC 9900 Elite mobile fluoroscopy system (image intensified fluoroscopic x-ray) is designed to provide fluoroscopic and spot-film images of the patient during diagnostic surgical and interventional procedure. |  |  | March 31, 2014 | 2014 | GE OEC Medical Systems, Inc | GE Healthcare has identified a potential safety issue related to the screws that attach the workstation monitor spring arm to the extension arm on certain OEC 9900 systems, causing the spring arm and monitor to fall forward. | Device Design | GE Healthcare sent an Urgent Medical Device Correction letter dated March 31, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Use caution when positioning or extending the workstation articulating arm. In addition, refrain from positioning staff or patients directly underneath or in front of a workstation monitor.Your local GE Healthcare Field Engineer will contact you to coordinate the replacement of the screws, and ensure that your system is operating safely and meets all specifications. These activities will be performed at no cost to you.Should you have any questions or concerns regarding these corrections, please do not hesitate to contact the service team for further information at 800-874-7378 option 8.Please be assured that maintaining a high level of safety and quality is our highest priority. If you have any questions, please contact us immediately.UPDATE: On June 20, 2014. GE Healthcare Surgery began distributing an Urgent Medical Device Correction Letter to customers with affected units. The letter makes clear that this field correction applies to all user sites with the listed products, regardless of whether service work was performed on the monitor extension arm since 3/31/2014. |  | 453 | 453 units | Worldwide Distribution - USA, and the countries of Australia, Argentina, Canada, China, Cambodia, Columbia, Japan, Ecuador, Japan, Korea, Hong Kong, India, Israel, Libya, Mexico, Nigeria, Peru, Russia, Thailand, Taiwan, Saudi Arabia and Singapore. | Terminated | October 26, 2015 | 574 | 453 | Computer | Not_Computer
Z-1481-2014 | 67535 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | RaySearch Ray Station Radiation Therapy Treatment Planning System, Stand-alone software. Versions 2.0, 2.5, 3.0, 3.5 and 4.0. |  |  | February 14, 2014 | 2014 | RAYSEARCH LABORATORIES AB | RaySearch Laboratories has recalled "RaySearch Ray Station 2.0, 2.5, 3.0, 3.5 and 4.0" software due to dose miscalculations for bolus/external/fixation support structures on CT-image. | Software design | RaySearch Laboratories sent a Field Safety Notice, Medical Device Correction letter dated February 14, 2014 to their customers via email. The letter identified the affected product, problem and actions to be taken by the user. Customers with questions may contact Freddie Cardel, Director of Customer Support at 1-877-778-3849 or via email at freddie.cardel@raysearchlabs.com. |  | 251 | 251 units (148 units domestically &amp; 103 units internationally) | Worldwide Distribution - USA Nationwide in the states of: AZ, CA, CT, CO, DE, FL, IL, NJ, NY, MA, MD, MS, OH, PA, TN, VA and WA. and the countries of Australia, New Zealand, Belgium, New South Wales, Republic of Korea, Italy, Germany, China, France, Switzerland, Netherlands, Spain, Austria, United Kingdom and Norway. | Open, Classified |  |  | 251 | Computer | Not_Computer
Z-1572-2014 | 67921 | No | LKK |  |  | General Hospital | PMA | MedStream Programmable Infusion Pump, 40 mL, Product Code: 91-4201 |  |  | November 10, 2015 | 2015 | Codman &amp; Shurtleff, Inc. | A small number of specific pumps may trigger a low battery alarm at the beginning of the implanted life. | Device Design | Codman Neuro sent an Urgent Field Safety Notice to all affected customers on April 18, 2014. The letter identified the product the problem and the action needed to be taken by the customer. The letter states the reason for recall, information about managing patients with affected pumps, and returning affected non-implanted pumps.You are receiving this voluntary recall notification because our records indicate that you are the recipient of one or more MedStream Pumps affected by this recall. Please do not implant any identified products included in this recall. Please complete the enclosed Acknowledgement Form and fax the completed form to: 001-508-977-6665. It is important that we receive this acknowledgement form, even if you have no affected product in your inventory.You may also e-mail the form to us at Codmanproductcomplaints@dpyus.jnj.comWe request that you immediately check all inventory to determine if you have affected product. Please contact your Codman Neuro sales consultant to assist you in the return process. Please use the included instructions to report your inventory status and return affected product. Note this recall applies only to |  | 34 | 34 | Worldwide Distribution including the countries of Belgium Germany, Spain, United Kingdom, Ireland, Italy, Russia and Turkey. | Terminated | November 10, 2015 | 0 | 124 | Computer | Computer
Z-1573-2014 | 67970 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | SCC SA HIS version 4.0.7.0 when used with SoftMic version 4.0.7.0. |  |  | January 29, 2014 | 2014 | SCC Soft Computer | When an isolate is resulted without a SNOMED code, isolate information in the downstream system may be incomplete or missing. As a result, there is the potential for the delay or omission of patient treatment updates. | Software design | On 2/10/2014 SCC Soft Computer used a proprietary task management system to send a Correction Communication to the consignee with the affected software. The letter identifies the product, problem and corrective actions. |  | 1 | 1 | Distributed in IL. | Terminated | November 03, 2014 | 278 | 1 | Computer | Computer
Z-1575-2014 | 68015 | No | OLO | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | NavioPFSTM System, Orthopedic Sterotaxic Instrument, model #'s NPFS-02000, NPFS-02010, NPFS 02020. |  |  | April 11, 2014 | 2014 | Blue Belt Technologies MN | Blue Belt Technologies is conducting a voluntary recall for their Navio PFS system which is used during orthopedic knee surgery. One of the electronic modules inside the computer cart assembly may have a defect in one of its components that could result in a latent failure. | Software design | Beginning April 3, 2014 consignees were hand delivered a Blue Belt Technologies, Inc. "Urgent Medical Device Recall" letter. The letter was addressed to Safety Officer / Regulatory Compliance Officer. The letter included the Description of the Issue, Description of Safety Concerns, Affected Product, Action to be taken by user, Transmission of this field safety notice and Contacts at Blue Belt Technologies. For questions in the USA contact Customer Support Blue Belt Technologies 2900 Northwest Blvd, Suite 40, Plymouth, MN 55441, USA, 1-763-452-4922 (Direct) lbrink@bluebelttech.com. In the European Union, Ross Lethbridge, European Product Manager, Blue Belt Technologies, Office number 207, 3000 Aviator Way. Manchester Business Park, Manchester, UK M22 5TG, Mobile: +44 7787 415069, rlethbridge@bluebelttech.com. |  | 11 | 11 | USA: CA, PA, OH, TX.OUS: BELGIUM, SCOTLAND, UNITED KINGDOM. | Terminated | June 10, 2014 | 60 | 11 | Computer | Not_Computer
Z-1312-2014 | 67623 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Pinnacle3 Radiation Therapy Planning (RTP) System, software version 9.0. |  |  | February 11, 2014 | 2014 | Philips Medical Systems (Cleveland), Inc. | Philips Medical Systems have recently determined that a software nonconformance can cause incorrect beam geometry. This issue affects Pinnacle3 Radiation Therapy Planning (RTP) System, software version 9.0 manufactured from 09/29/09 to 12/01/2013. | Software design | Philips Healthcare sent an Urgent Field Safety Notice dated February 10, 2014, to all affected consignees. The letter included the Problem Description, Hazard Involved, How to Identify Affected Products, Action to be taken by Customer /User, Actions Planned by Philips and Further Information and Support. Consignees were instructed to review the following information with all members of their staff who need to be aware of the contents of this communication including clinics within their institution that are remotely connected to their Pinnacle Server. Customers were questions were instructed to contact their Philips representative. For customers in North America please contact Customer Care Solutions Center at 1-800-722-9377. For all other countries their local Philips Healthcare Office should be contacted.For questions regarding this recall call 1-800-722-9377. |  | 1016 | 1016 | Worldwide Distribution - USA including US: AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WV and Internationally to AUSTRALIA, AUSTRIA. BELGIUM, CANDA, CHINA, DENMARK, FRANCE, GERMANY, HONG KONG, HUNGARY, INDIA, IRELAND, ITALY, JAPAN, LUXEMBOURG, MALAYSIA, MEXICO, MONTENEGRO, NETHERLANDS, NEW ZEALAND, OMAN, RUSSIA, SOUTH ARICA, SOUTH KOREA. SPAIN,. SWITZERLAND, TAIWAN, THAILAND, TURKEY, EMIRATES, UNITED KINGDOM.. | Terminated | November 21, 2014 | 283 | 1016 | Computer | Computer
Z-1661-2014 | 68165 | Yes | LWP |  |  | Cardiovascular | PMA | Medtronic Temporary Pacemaker, Model 5348. Single Chamber External Pulse Generator (EPG) Temporary Pacemaker (AAI/VVI). |  |  | April 29, 2014 | 2014 | Medtronic Inc. Cardiac Rhythm Disease Management | Events related to a pacing rate outside of the intended setting, including events of sudden increased pacing rate up to the maximum setting of 180 pulses per minute (ppm), have been reported. This is caused by the development of high resistance on internal electrical connector contacts due to oxidation over time. | Device Design | Consignees were sent a Medtronic "Urgent Medical Device Correction" letter dated April 2014. The letter was addressed to Risk Manger, Health Care Professional. The letter included the Issue Description, Root Cause, Potentially Affected 5348 EPGs within the Serial Number Ranges, Malfunction Indications, and Recommended Actions. For questions related to EPG service policy or purchasing a replacement EPG, customers should contact your Medtronic representative. For technical questions, please call Instruments Technical Service at (1-800)638-1991. Medtronic asks that customers complete and sign the attached Customer Confirmation Certificate and fax it to Medtronic at 651-367-0612 to the attention of Customer Focused Quality or e-mail to RS.CFOFCA@Medtrnic.com. |  | 16337 | 16,337 (8,989 US; 7,348 OUS) | Worldwide Distribution -- USA, including the territories of Guam and Puerto Rico; and, the countries of Albania, Algeria, Australia, Austria, Bahamas, Bangladesh, Belarus, Belgium, Benin, Bolivia, Bosnia And Herzegovina, Brazil, Bulgaria, Burkina Faso, Chile, Colombia, Costa Rica, Cote D'Ivoire, Croatia, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, El Salvador, Estonia, Faroe Islands, Finland, France, Georgia, Germany, Greece, Guadeloupe, Honduras, Hong Kong, Hungary, India, Iran, Ireland, Israel, Italy, Jamaica, Japan, Jordan, Kazakhstan, Kenya, Kuwait, Latvia, Lebanon, Lithuania, Luxembourg, Macau, Malaysia, Malta, Martinique, Mauritius, Mexico, Morocco, Nepal, Netherlands, Netherlands Antilles, New Caledonia, Norway, Oman, Pakistan, Palestine, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Qatar, Reunion, Saudi Arabia, Senegal, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Sudan, Sweden, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Venezuela, and Yemen. | Terminated | October 08, 2014 | 162 | 16337 | Computer | Computer
Z-1660-2014 | 67963 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | LIFEPAK 1000 defibrillator. Affected Part Numbers: 320371500XXX |  |  | May 09, 2014 | 2014 | Physio-Control, Inc. | Physio-Control has become aware of incidents where customers have attempted to use theirLIFEPAK 1000 defibrillator and the device has shut down unexpectedly due to a very low battery.A defibrillator in this scenario has the potential to fail to deliver a shock, with the potential resultthat therapy is not delivered and a patient is not resuscitated. A software malfunction in the LIFEPAK 100 | Software Manufacturing/Software Deployment | On May 9, 2014 Physio-Control sent an Urgent Medical Device Recall Letter (dated May 2014) to all affected customers. The letter identified the product, the problem and the action needed to be taken by the customer. The customer notification letter will provide new instructions on how to check the functional and power readiness of the device and advise the customer on the actions to take based on the symbol displayed.Customers are asked to record the readiness indicator(s) shown on each device and fax the confirmation sheet back to Physio-Control at 1-866-448-9567. Customers are also directed to contact Physio-Control at 1-800-442-1142 to obtain replacement batteries, if needed. Customers are provided a Device Readiness Guide for use in interpreting the Readiness Display on the device, and are directed to refer to the Operating Instructions which were provided upon purchase.In addition to notifying Physio-Control of any potential quality problems or adverse reactions or events experienced associated to the use of these products, Physio-Control asks customers to report directly to the FDA through the MedWatch Adverse Reporting Program online, by regular mail or by fax.Physio-Control reminds customers that it is critically important in understanding what the device and battery indicators mean on yourdefibrillator and what actions you need to take as a result. At any time the battery charge can be verified by following the instruction provided on page 2-5 of the Operating Instructions. It is also important that you always carry a spare fully-charged battery, as stated in the Operating Instructions.Customers are directed to call Physio-Control at 1-800-442-1142, 6:00 a.m. to 4:00p.m. (Pacific), Monday - Friday for any further questions.*********************************************************************************************On May 18, 2015, Physio Control sent the Urgent Medical Correction-Action Required LIFEPAK 1000 DEFIBRILLATOR to U |  | 96611 | 96,611 units total (40,974 in the US and 55,637 outside the US) | Worldwide Distribution-USA (nationwide) and the countries of ALBANIA, ARGENTINA, AUSTRALIA, AUSTRIA, BAHRAIN, BANGLADESH, BARBADOS, BELGIUM, BOLIVIA, BOSNIA AND HERZOGOVENIA, BRAZIL, BRUNEI DARUSSALAM, CANADA, CANARY ISLANDS, CAYMAN ISLANDS, CHILE, China, COLOMBIA, COSTA RICA, CROATIA, Curacao, CYPRUS, CZECH REPUBLIC, DENMARK, ECUADOR, EGYPT, EspaÂża, FAROE ISLANDS, FINLAND, FRANCE, FRENCH GUIANA, FRENCH POLYNESIA, GABON, GEORGIA, GERMANY, GREECE, GUAM, GUATEMALA, HONG KONG, HUNGARY, ICELAND, INDIA, INDONESIA, IRAN, ISLAMIC REPUBLIC OF, IRAQ, IRELAND, ISRAEL, ITALY, Japan, JORDAN, KAZAKHSTAN, KENYA, KOREA, REPUBLIC OF, KUWAIT, LATVIA, LEBANON, LIBYA, LITHUANIA, MACAU, MACEDONIA, THE FORMER YUGOSL, MALAYSIA, MALDIVES, MALTA, MEXICO, MOROCCO, NAMIBIA, NETHERLANDS, NEW CALEDONIA, NEW ZEALAND, NORWAY, OMAN, PAKISTAN, PANAMA, PERU, PHILIPPINES, POLAND, PORTUGAL, QATAR, REUNION, ROMANIA, RUSSIA FEDERATION, SAN MARINO, SAUDI ARABIA, SERBIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SRI LANKA, Suriname, South America, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TUNISIA, TURKEY, TURKMENISTAN, UKRAINE, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY, and VENEZUELA. | Open, Classified |  |  | 96611 | Computer | Computer
Z-1655-2014 | 67953 | No | JLW | Thyroid stimulating hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA Centaur Systems TSH3 Ultra 2500 tests |  |  | April 10, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc | Samples containing fluorescein may show interference with the ADVIA CentaurÂż Systems TSH3 Ultra, Vitamin D and the ADVIA Centaur BRAHMS Procalcitonin assays. | Device Design | Siemens Healthcare Urgent Medical Device Correction (UMDC) was mailed to all affected Siemens Healthcare Diagnostics customers in the United States on April 11, 2014 for customer receipt on April 14, 2014. This notice informed the customers of fluorescein interference with the ADVIA Centaur Systems TSH3-Ultra and Vitamin D Total assays. This notice instructs customers that the TSH3-Ultra and Vitamin D assays can continue to be used for testing any sample that does not contain fluorescein. However, testing blood drawn from a patient, without allowing clearance of fluorescein is not recommended. This notice should be reviewed with the facility's laboratory or medical director.B. An Urgent Field Safety Notice (UFSN) was sent out for distribution to all customers outside the United States on April 10, 2014. This notice informed the customer of fluorescein interference with the ADVIA Centaur Systems TSH3-Ultra , Vitamin D Total and ADVIA Centaur BRAHMS Procalcitonin assays. This notice instructs customers that the TSH3-Ultra, Vitamin D andProcalcitonin assays can continue to be used for testing any sample that does not contain fluorescein. However, testing blood drawn from a patient, without allowing clearance of fluorescein is not recommended. This notice should be reviewed with the facilitys laboratory or medical director. (Note: The ADVIA Centaur BRAHMS Procalcitonin assay is not sold in the US.) |  | 8052 | 8,052 kits | Worldwide Distribution-USA (nationwide) and the countries of Afghanistan Algeria Armenia Austria Bahrain Belgium Bulgaria Burkina Faso Canary Islands Croatia Cyprus Czech Republic Denmark Estonia Finland France Germany Great Britain Greece Hungary Iceland Iran Iraq Ireland Israel Italy Jordan Kazakhstan Latvia Lebanon Lithuania Luxembourg Macedonia Malta Netherlands Norway Oman Pakistan Poland Portugal Qatar Romania Russian Fed. Saudi Arabia Serbia Slovakia Slovenia Spain Sweden Switzerland Turkey Unit.Arab Emir. Vatican Yemen. | Terminated |  |  | 273253 | Computer | Not_Computer
Z-229-2014 | 67639 | No | DSQ |  |  | Cardiovascular | PMA | HeartMate II LVAD Pump and Pock Controller Kit; |  |  | March 04, 2014 | 2014 | Thoratec Corporation | Discovery of serious injuries and deaths associated with the process of changing from a primary System controller to their back-up System controller in patients using the Pocket System controller model. | Device Design | Thoratec sent an Urgent Medical Device Correction letter dated March 4, 2014 to all affected customers. The letter was sent via e-mail. All distributors outside of the USA and Canada to whom affected product has been distributed have been notified of this field correction notice and have been provided copies of the Customer Letter to forward to their customers. The letter identified the affected product , problem and actions to be taken. Customers were requested to complete and sign the attached Acknowledgement Form and return it to Thoratec via fax (925-847-8571) or e-mail a scanned copy tocustomer.service@thoratec.com. For questions call 1-800-528-2577. |  | 36 | 36 | Worldwide distribution: US (nationwide) including states of: AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NJ, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, and WI; and countries of: Austria, Belgium, Czech Republic, France, Germany, Ireland, Israel, Italy, Lebanon, Lithuania, Poland, Saudi Arabia, Singapore, Slovenia, Sweden, Switzerland, The Netherlands, Turkey, and United Kingdom. | Terminated | August 27, 2015 | 541 | 218533 | Computer | Not_Computer
Z-1271-2014 | 67657 | No | LHN | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MEVION S250, System, Radiation Therapy, Charged-Particle, Medical used for Proton Radiation Therapy |  |  | March 05, 2014 | 2014 | Mevion Medical Systems, Inc. | Geometric positioning error can occur after a rotational correction has been applied to one or more treatment fields | Software design | Mervion issued an Important Safety Notice letter on March 5, 2014 by visit to sites to the single customer using the product to prevent the use of multiple isocenters treatment modalities. The letter identified the affected product, problem and actions to be taken. For questions call 978-540-1555. |  | 1 | 1 unit | USA Nationwide Distribution in the state of MO | Terminated | April 03, 2014 | 29 | 1 | Computer | Not_Computer
Z-1640-2014 | 67620 | No | KOI | Electrical peripheral nerve stimulator. | Anesthesiology | Anesthesiology | 510(k) | GE Healthcare, NeuroMuscular Transmission Module, E-NMT-01 |  |  | April 09, 2014 | 2014 | GE Healthcare, LLC | When the E-NMT-01 module is used in conjunction with the ElectroSensor, the Neuromuscular Transmission (NMT) values may indicate a deeper level of muscle relaxation than the actual level of muscle relaxation. In the clinical situation visual movements of the hand are seen after TOF (Train of Four) stimulation, but the patient monitor shows no counts, or counts are not corresponding to the actual a | Software design | GE Healthcare sent an "Urgent Medical Device Correction" letter (Ref 36099) dated April 9, 2014 to affected customers. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction, and Contact Information (contact Technical Support at 1-800-558-7044 or your local Service Representative).UPDATE: On May 11, 2015 a second GE Healthcare letter FMI 36099-2 was sent to OUS consignees. The updated report was for additional affected modules due to FRU installation. |  | 1794 | 1,794 (471 units US, 1,323 units OUS) | US Nationwide Distribution in the states of: AR, CA, CO,DC, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MI, MN, MO, NC, NV, NY, OH, OK, TN, TX, UT, WA, WI, WY.OUS: ARGENTINA. AUSTRALIA, AUSTRIA, BELGIUM, BOLIVIA, CANADA, CHILE, COLOMBIA, COSTA RICA, CYPRUS, CZECH REPUBLIC, DENMARK, ECUADOR, EGYPT, ESTONIA, FINLAND, FRANCE, GERMANY, GREECE, HONG KONG, HUNGRY, INDIA, IRA, ISRAEL, ITALY, JAPAN, KAZAKTSTAN, KUWAIT, LATVIA, LITHUANIA, MALAYSIA, MEXICO, NETHERLANDS, NEW ZEALAND, NORWAY, OMAN, PANAMA, POLAND, QATAR, ROMANIA, RUSSIAN, FÂżDÂżRATION, SAUDI ARABIA, SINGAPORE, SLOVAKIA, SPAIN, SWEDEN, SWITZERLAND, TURKEY, TAIWAN, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY, VENEZUELA. | Terminated | September 28, 2015 | 537 | 1794 | Computer | Not_Computer
Z-1645-2014 | 67957 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | 550 TxT Treatment Table of the Digital Linear Accelerator (LINAC) |  |  | April 01, 2014 | 2014 | Siemens Medical Solutions USA, Inc | Table may lose calibration during patient treatment, which may cause the possibility of mistreatment. | Software design | Siemens sent an Medical Device Correction, Field Safety Notice of Improvement to End users in March 2014, informing them that a Siemens representative would install a software update.By installation of this software update the issues were solved by the following actions:1. A plausibility check was introduced for comparing the two independent encoder systems of each table axis.2. The time between two encoder comparisons was reduced to 50ms and an error code was introduced.Please include this Field Safety Notice in your Digital Linear Accelerator System Owner Manual chapter "Safety Advisory Letters" where it should remain.The relevant National Competent Authority has been informed of this update.We regret any inconvenience that this may cause, and we thank you in advance for your understanding. Further questions please call (610) 219-6300. |  | 120 | 120 | USA ( nationwide) Distribution. | Terminated | December 11, 2014 | 254 | 120 | Computer | Not_Computer
Z-1272-2014 | 67556 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | AXIOM Artis Systems |  |  | February 07, 2014 | 2014 | Siemens Medical Solutions USA, Inc | Certain units of the AXIOM Artis Systems are using a flat detector cooling unit with inappropriate connectors that may possibly leak, causing a system failure and possible drop out of the gantry system. | Device Design | Siemens sent a recall notification letter, dated February 7, 2014, to all affected customers. The letter identified the affected product, problem and actions to be taken. For questions call +49 9191 18 0. |  | 37 | 37 | USA Nationwide Distribution in the states of WI, KY, OH, NJ, NC, FL, PA, MA, IA, TX, NY, IL, MI, AZ, CA, and LA. | Terminated | December 30, 2014 | 326 | 37 | Computer | Computer
Z-1646-2014 | 68262 | No | OUG | Medical device data system. | General Hospital | General Hospital | 510(K) Exempt | Philips XRE, the dictation editor system included in the Philips XIRIS 8.1.50 |  |  | May 05, 2014 | 2014 | Philips Healthcare Informatics, Inc. | When using XIRIS Rio version (8.1.50.0) in conjuncture with iSite PACS a synchronization error can occur and if not noticed, information for the incorrect patient may be displayed | Software design | Philips Healthcare sent an Urgent Field Safety Notice dated May 5, 2014. to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were advised of the synchronization problem and request that a copy of this notification be kept with the equipment instructions for use. The document goes on to outline the affected product, the problems and conditions when it can occur. A photo of the label for easy identity the actions to be taken are provided. Philips is informing users of the issue via the notification and will deploy a software correction to all customers when available.Questions may be directed to Philips reps or Care at 1-866-946-4701 or OneEMSProd@philips.com |  | 5 | 5 units | Worldwide Distribution - US including the states of FL and TX and the countries Canada and South Africa. | Terminated | April 28, 2015 | 358 | 5 | Computer | Computer
Z-1647-2014 | 67787 | No | MKW | Cyclosporine test system. | Clinical Chemistry | Toxicology | 510(k) | Dimension VistaÂż CSA and CSAE |  |  | March 19, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc. | A complaint was received of Dimension VistaÂż CSA and CSAE under-recovery of cyclosporine when patient samples are run from the Small Sample Container. The under-recovery that has been observed is 18 to 44%. In March 2014, Siemens recalled due to the complaint received of Dimension VistaÂż CSA and CSAE reporting under-recovery of cyclosporine when patient samples are run from theSmall Sample C | Device Design | Siemens sent an Urgent Medical Device Correction letter dated March 2014. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed not to use the SSC with patient samples when running CSE or CSAE.Siemens also recommended a look back at previously generated CSA or CSAE results if the patient samples have been run from the SSC and that the content of this letter should be discussed with the Medical Director. Customers were asked to complete the attached form and fax it to (302) 631-8467.For questions regarding this recall call 800-441-9250. |  | 4175 | 4175 | Worldwide Distribution - USA (nationwide) and Internationally to Canada, Belgium, Switzerland, Germany, Spain, France, Great Britain, Italy, Saudi Arabia, Australia, and Japan. | Terminated | August 10, 2015 | 509 | 4175 | Computer | Not_Computer
Z-1275-2014 | 67615 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | DICOM Export Format for Sum Dose with Velocity v3.0.0. |  |  | July 25, 2013 | 2013 | Velocity Medical Solutions, LLC | An anomaly was detected such that, under certain conditions, a summed dose created by Velocity replaces the calculated dose from an approved plan in Eclipse when importing the dose from Velocity. | Software design | Velocity Medical Solutions sent a Urgent Field Safety Notice dated July 25th, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Velocity advises customers creating Dose volumes in Velocity to refrain from exporting these dose files to Eclipse until the customers install has been upgraded to Velocity 3.0.1, which will be available August 5th, 2013.A new version of Velocity that corrects this error, Velocity v3.0.1, will be available on August 5th, 2013 for upgrade at no cost. Please contact your Velocity Medical representative to obtain this upgrade.Service representative will upgrade the application at your earliest convenience.Please advise the appropriate personnel working in your medical department of the content of this letter.We sincerely apologize for any inconvenience and thank you in advance for your co-operation.Contact Information. Helpdesk: (404) 920-1966 or (855) 857-0802 and press 2. |  | 54 | 54 installs | Worldwide Distribution - US (nationwide) including the states of AL, CA, CO, CT, DC, FL, GA, IL, IN, LA, MA, MD, MI, MS, NC, NY, PA, SC, TN, TX, WI, and the countries of Belarus, Italy and Switzerland. | Terminated | September 11, 2014 | 413 | 54 | Computer | Computer
Z-1539-2014 | 68011 | No | BZO | Ventilator tubing. | Anesthesiology | Anesthesiology | 510(K) Exempt | Infant Nasal CPAP (continuous positive airway pressure) Prong, Nasal. |  |  | April 15, 2014 | 2014 | Fisher &amp; Paykel Healthcare, Ltd. | The Infant Nasal CPAP Prong has the potential to detach from the Nasal Tubing and therapy is likely to be interrupted. | Device Design | Fisher &amp; Paykel Healthcare sent an Urgent Medical Device Recall letter on April 11, 2014, to all affected customers. The letter informed them of the problems identified and the actions to be taken. Customers were instructed to complete Section A 'Inspection of Stock' on the Urgent Medical Device Recall Response Form and return it to their FPH Representative. Destroy the Affected Prongs by cutting them in half. Customers were also instructed to contact their FRH Representative for replacement product. For questions regarding this recall call 949-453-4000, ext 1423. |  | 62063 | 62,063 units total (14,420 units USA) | Worldwide Distribution - USA (nationwide) and Internationally to:ARGENTINAAUSTRALIABANGLADESHBELGIUMCANADACHILECHINAFRANCEGERMANYGREAT BRITAINGREECEHUNGARYINDIAINDONESIAIRAQISRAELITALYJAPANLIBYAMALAYSIANEPALNETHERLANDSNEW ZEALANDPERUPOLANDPORTUGALQATARROMANIARUSSIARWANDASAUDI ARABIASOUTH AFRICASPAINSRI LANKASWITZERLANDTAIWANTHAILANDTRINIDAD &amp; TOBAGOTURKEYUNITED ARAB EMIRATES | Terminated | September 30, 2014 | 168 | 62063 | Computer | Not_Computer
Z-1297-2014 | 67672 | No | MIH |  |  | Cardiovascular | PMA | Ovation Prime Abdominal Stent Graft System; |  |  | February 27, 2014 | 2014 | Trivascular, Inc | Communication to European physician users of the Ovation Prime Abdominal Stent Graft system to notify them of changes to Instructions for Use. | Device Design | Trivascular sent an Important Medical Device Information letter dated February 27, 2014, to all European physicians who have scheduled cases through March 31, 2014 or until IFU;s are translated and available for European physicians. The letter provided notification of an update that will appear in a future revision of the product Instructions for Use (IFU). The IFU will be available in their country when the required translations and product registration/notification have been completed. This IFU does not require rework or return of the affected product. Customers with questions were instructed to contact their local TriVascular representative.For questions regarding this recall call 707-543-8805. |  | 45 | 45 | Worldwide Distribution including Austria, Belgium, Czech Republic, France, Germany, Greece, Hungary, Italy, Luxembourg, Norway, Poland, Spain, and Sweden. | Terminated | April 11, 2014 | 43 | 45 | Computer | Not_Computer
Z-1638-2014 | 68224 | No | MUE | Full-field digital mammography system. | Radiology | Radiology | 510(k) | MAMMOMAT Inspiration. |  |  | April 25, 2014 | 2014 | Siemens Medical Solutions USA, Inc | There is a potential and possible hazard to the user when using the MAMMOMAT Inspiration PC monitor at the control desk, in that the holder of the PC monitor can break causing an unstable monitor to fall causing possible serious injury. | Device Design | A safety advisory notice, dated April 25, 2014, was sent to direct accounts to notify them of the issue and provide instructions to avoid the potential risk of injury. |  | 5 | 5 | Distributed in the states of GA, FL, NY, NV, and MA. | Terminated | December 15, 2014 | 234 | 5 | Computer | Computer
Z-1202-2014 | 67016 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Siemens Artis Q and Artis Q biplane systems |  |  | December 05, 2013 | 2013 | Siemens Medical Solutions USA, Inc | There exists the possibility that during clinical operation, a breakdown of the x-ray tube could result in loss of x-ray imaging for the Artis Q and Artis Q biplane systems with a flat detector as 40 HDR (large 30x40). | Device Design | Siemens sent an Important Customer Safety Notice letter dated November 22, 2013, to affected customers. The letter identifed the affected product, problem and actions to be taken. Customers were instructed to discontinue use until the affected product has been successfully replaced by a Siemen's representative.at no cost. |  | 19 | 19 | USA Nationwide Distribution in the states of TN, MA, CA, IL, OH, FL, UT, WA, GA, IA, NY, MO, IN, and CO. | Terminated | November 06, 2014 | 336 | 19 | Computer | Not_Computer
Z-1307-2014 | 67504 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | RayStation Radiation Therapy Treatment Planning System, models 2.5, 3.0, 3.5 and 4.0; a software system designed for treatment planning and analysis of radiation therapy. |  |  | February 06, 2014 | 2014 | RAYSEARCH LABORATORIES AB | Potential for dose errors due to software program errors. | Software Design Change | RaySearch Laboratories sent an Field Safety Notice, Medical Device Correction letter dated February 4, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.ACTIONS TO BE TAKEN BY THE USERA workaround that eliminates the described error is to select the discard rotations option in the Image Registration Module before creating a rigid registration intended to be used for hybrid deformable registration. Another option is to use the ROI based deformable registration when applicable, as this functionality is not affected by the error.SOLUTIONThis issue will be resolved in a patch release of RayStation, scheduled for market release February 28, 2014. Release schedule is subject to regulatory approval for some markets. In the meantime, this field safety notice is being distributed to all customers.CONTACT INFORMATIONFor further information, please contactFor the Americas:Customer Support, RaySearch Americas, at +1 877 778 3849 For Europe, Asia and rest of the world:RaySearch Laboratories AB, at +46 8 5450 6130 Thank you for your cooperation, and we apologize for any inconvenience |  | 151 | 151 units Total (88 units domestically &amp; 63 units internationally) | Worldwide Distribution - US including the states of NY, OH, MD, CT, CO, TN, IL, AZ, MS, NJ, PA, DE, FL, CA and WA., and the countries of New Zealand, Australia, Republic of Korea, China, Italy, Germany, Norway, Belgium, Switzerland, United Kingdom and Spain. | Completed |  |  | 151 | Computer | Computer
Z-1310-2014 | 67628 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | TDWorkstation version 11.01.A or higher. TDWorkstation Evolution. |  |  | February 17, 2014 | 2014 | Technidata S.A. | If a conflict occurs between a Demography Query and a Result Transmission by the instrument, the TDWorkstation could send an incorrect patient demography to an instrument. | Software design | An email dated 2/17/14 was sent to their customers to inform them about concerns only if their site uses TDWorkstation and TD-Workstation Evolution products that meet all the listed criterias. The criterias are listed in the customer notification letter and it informs the customers of the problems identified and the actions to be taken. Customers are instructed to complete the reply form and return it by fax (514) 270-2770 with in 15 days from receiving the customer notification form. Customers that do not respond the reply form, Technidata will consider those customers do not wish to install the correction. |  | 85 | 85 units total (1 unit in US) | Distributed in the state of PA. | Terminated | March 31, 2014 | 42 | 85 | Computer | Not_Computer
Z-1567-2014 | 68066 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris Pump model 8100 with software version 9.1.18 and software upgrade kits with "Delay Until" Option and "Multidose" Feature |  |  | April 23, 2014 | 2014 | CareFusion 303, Inc. | CareFusion is recalling the Alaris Pump model 8100 version 9.1.18 because it may have a software issue that results in situation where the pump module will not properly delay an infusion when the "Delay Until" option or "Multidose" feature is used. | Software design | CareFusion sent an Urgent Medical Device Recall Notification letter dated April 23, 2014, to all affected customers. The letter informed the customers of the problems identified and the actions to be taken. Customers were informed that if the infusion starts earlier or later than intended and is not immediately detected and stopped by the clinician this could result in serious injury or death. Customers are informed although no adverse events or deaths have been reported there is a potential for this risk. If customers experience the issue while using the Alaris Pump module then they are instructed to contact CareFusion Customer Advocacy at (888) 812-3266. Customers are instructed to promptly complete and return the enclosed Customer Response Card to expedite the corrective action process.For questions regarding this recall call 858-617-4000. |  | 57219 | 57,219 units total (55,756 units in US) | Worldwide Distribution-USA (nationwide) and the countries of Canada, AE, AU, and KW. | Terminated | March 01, 2017 | 1043 | 57219 | Computer | Computer
Z-1617-2014 | 68208 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | McKesson Radiology-PACS. |  |  | April 17, 2014 | 2014 | Mckesson Medical Immaging | High priority alert icon was not displayed in the Study List when there are no Stat studies present. However, the High priority alert icon is visible when studies marked as Stat priority exist in the Study List. | Software Design Change | On April 17, 2014, McKesson Medical Imaging Company initiated a process for applying a software solution to all McKesson Radiology customer systems that may be affected to prevent future occurrences of this issue. The systems with the potentially affected software versions will be corrected without removing the systems from their location. |  | 45 | 45 | Worldwide Distribution-USA (nationwide) including the states of AL, CA, GA, ID, IL, IA, MD, MS, MO, NE, NV, NH, NY, NC, PA, TN, TX, VT, WA, WV, WY, and the countries of Canada and France. | Terminated | February 02, 2016 | 656 | 45 | Computer | Computer
Z-1309-2014 | 67558 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Artis Zee Angiographic x-ray systems. |  |  | February 07, 2014 | 2014 | Siemens Medical Solutions USA, Inc | An inappropriate cable connection was assembled for certain Artis zee systems. which may cause a system failure and possibly lead to a restricted maneuverability of the gantry system. | Device Design | A recall notification letter dated February 7, 2014, was sent to End Users. |  | 9 | 9 | MN, WI, PA, MI, AL, OH, NC, AR, and NJ. | Terminated | March 16, 2015 | 402 | 9 | Computer | Not_Computer
Z-1619-2014 | 68042 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Revision K Actuator Test Boards sold as replacement parts for use with Fresenius 2008-Series Hemodialysis Machines (2008K, 2008K2, 2008K@home and 2008T). |  |  | April 25, 2014 | 2014 | Fresenius Medical Care Holdings, Inc. | 2008-Series Hemodialysis Machines with Revision K Actuator Test Boards may fail to detect a disconnected Ultrafiltration (UF) Pump | Device Design | Fresenius Medical issued Urgent Recall letter issued on 4/25/14 by certified mail with signature confirmation and fax-back form. Customers instructed to examine their inventory to determine whether they have any of the affected 2008T machines or if they have installed any of the affected replacement boards. If customers have the affected machines or replacement boards, they will be instructed to contact FMC-RTG Technical Service to have the Actuator Test Boards replaced. Contact the FMCNA Technical Service Team at 1-800-227-2572. |  | 39 | 39 units | USA (nationwide) and the country of Canada.* | Open, Classified |  |  | 118 | Computer | Not_Computer
Z-1308-2014 | 67511 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | RaySearch Ray Station, Radiation Therapy Treatment Planning System, Stand-alone software |  |  | February 06, 2014 | 2014 | RAYSEARCH LABORATORIES AB | RaySearch Laboratories has recalled RaySearch Ray Station 4.0 software due to clinical dose calculation errors during radiation therapy. | Software Design Change | On February 6, 2014, RaySearch Laboratories distributed Urgent Medical Device Correction notices to their customers via email regarding the affected product. To determine if the version you are using is affected, bring up the About box in the RayStation application and check if the number reported is "4.0.2.9". If so, the notice applies to your version. Customers are advised to inform all planning staff and all users about the instructions listed on the notice. Customers are advised to inspect their system and identify all installed units with the referenced software version numbers listed on the notice. Corrective labeling is distributed to the customers in the form of the safety notification attached to this report and the software problem will be resolved in RayStation patch 4.0.3 scheduled for market release end of February 2014 and in all future releases. For further information and return of the reply form, please contact: Freddie Cardel, Director of Customer Support, RaySearch Americas, (877) 778-3849 or freddie.cardel@raysearchlabs.com. |  | 19 | 19 units (16 units domestically &amp; 3 units internationally) | OH, MD, NJ, WA, IL, MA, FL, Spain, United Kingdom, Italy, | Open, Classified |  |  | 19 | Computer | Computer
Z-1613-2014 | 67960 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo Dynamics v9.5 system. A Picture Archiving and Communication System (PACS) intended for acceptance, transfer, display, storage, archive and manipulation of digital medical images, including quantification and report generation. |  |  | April 03, 2014 | 2014 | Siemens Medical Solutions USA, Inc | Systems with the affected serial number are experiencing a software error which may lead to incorrect patient demographics display. | Software design | Siemens sent a Recall Notification letter dated March 26, 2014 to affected customers. The letter identified the affected product, problem and actions to be taken. Siemens is preparing a Service Patch modification that would be performed to resolve the software issue. |  | 50 | 50 | US Nationwide Distribution in the states of AR, CA, CT, GA, IA, IL, KS, MA, MI, MN, NJ, NM, NY, OH, OR, PA, TX, TN, TX, WA, WI, and WV. | Terminated | December 15, 2014 | 256 | 50 | Computer | Computer
Z-1553-2014 | 67603 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ADVIA Centaur XP Immunoassay system, an automated in vitro diagnostic analyzer. Catalog # 078-A011-03 - ADVIA Centaur XP, Catalog # 078-A011R03 - ADVIA Centaur XP refurb. |  |  | February 18, 2014 | 2014 | Siemens Healthcare Diagnostics | Instrument Cover Gas Spring failures. Over time, the gas spring may lose its effectiveness and fail to support the cover in its full or partially open position. This may lead to the cover falling. | Device Design | On February 18, 2014 Siemens Healthcare Diagnostics distributed Urgent Medical Device Correction notices to their customers via courier service (Urgent Medical Device Correction/). Upon receipt of the communication, customers are asked to follow "Actions to be Taken by the Customer". If the cover is unable to stay partially raised, immediately contact your Siemens technical support representative to arrange for a service. Complete and return the Field Correction Effectiveness Check attached to this letter within thirty (30) days by faxing to (302) 631-7597. Customers with questions should contact their Siemens technical support representative. |  | 5900 | 5,900 units total (1,679 domestically &amp; 4,221 internationally) | Nationwide and internationally | Open, Classified |  |  | 5900 | Computer | Not_Computer
Z-1311-2014 | 67749 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Zoll Medical X Series, PropaqMD, PropaqM Defibrillator/Pacemaker/Monitor With System Software Version 02.10.02.00 or Higher, automated external defibrillator. |  |  | March 05, 2014 | 2014 | ZOLL Medical Corporation | Defibrillator/Pacemaker/Monitor May Latch in a Continuous Device Reset Loop and cannot be used | Software design | Zoll Medical issued a Notification letter daetd March 5, 2014, to all affected consignees. The notification letter provides information on how to avoid the issue and instructs customers to contact ZOLL to get affected devices updated. Customers will be given the option to have a software upgrade kit sent to them to update their devices or return devices to ZOLL Service department in Chelmsford Massachusetts to have the device upgraded with the latest software release (Software version 02.16.04.00). Customers with questions were instructed to contact the 24/7 technical support numbers 1 (800) 348 9011 or +1 (978) 421 9460 to assist users with any aspect of this notice. |  | 3169 | 3,169 units | Worldwide Distribution - USA (nationwide) and Internationally to Canada, AE, Afganistian, Australia, Belgium, Czech Republic , China, Chile, Croatia, Denmark, France, Great Britain, Hungary, Israel, Italy, Japan, Korea, Kuwait, Russia, Saudi Arabia, Singapore, South Africa,Thailand, and Turkey. | Open, Classified |  |  | 3169 | Computer | Computer
Z-1616-2014 | 67961 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ACCU-CHEK Compact Plus Test Strips |  |  | April 14, 2014 | 2014 | Roche Diagnostics Operations, Inc. | Roche Diabetes Care has become aware the ACCU-CHEK Compact Plus test strips may produce erroneously low blood glucose readings in patients undergoing Ceftriaxone therapy. The interference is not described in the product labeling. | Device Design | On 4/14/14, Roche Diagnostics Operations issued URGENT MEDICAL DEVICE CORRECTION notification to consignees which included description of the affected product and issue, clinical significance, and actions required. Consignees are directed to call the ACCU-CHEKÂż Customer Care Service Center at 1-800-858-8072 for any questions about the information contained in the notification. |  | 590136 | 590,136 boxes | Nationwide Distribution. | Terminated | September 18, 2014 | 157 | 590136 | Computer | Not_Computer
Z-1566-2014 | 67884 | No | NEP | Intravascular administration set. | General Hospital | General Hospital | 510(K) Exempt | ABACUS TPN Calculation Software, Product Codes: 8300-0045 ABACUS CE (Calculator Edition), 8300-0046 ABACUS SE (Single-Workstation Edition), 8300-0047 ABACUS ME (Multi-Workstation Edition), 8300-0156 ABACUS 3.0 CE, 8300-0157 ABACUS 3.0 SE, 8300-0158 ABACUS 3.0 ME, 8300-0166 ABACUS 3.1 DE, 8300-0167 ABACUS 3.1 CE, 8300-0168 ABACUS 3.1 SE, 8300-0169 ABACUS 3.1 ME. ABACUS is a Windows - based order entry TPN software application for comprehensive calculations and label printing. The primary use of Abacus is for Total Parenteral Nutrition (TPN) order calculations. |  |  | April 02, 2014 | 2014 | Baxter Corporation Englewood | 1. ABACUS v3.1 may calculate quantities of electrolytes that are double the expected values during the creation of TPN orders.2. ABACUS v3.1 may automatically add additional sterile water to a formula equal to the volume of a premixresulting in an over dilution.3. All software versions of ABACUS software display the calcium phosphate curve points for Premasolincorrectly.4. All software v | Software design | Baxter sent an Urgent Device Correction letter dated April 21, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. . 1. Follow safe compounding practice, which includes 1) using of the Summary button to verify the order against the calculated amounts prior to completing the order, 2) verifying that the ordered ingredients and quantities displayed in the software and printed on the Bag label and the Solution Formula label match the PN prescription prior to preparation, and 2) using a filter for administration of a PN bag. Use of the ABACUS Software must comply with your facility's policies, procedures, and training, including verification that all information entered into the ABACUS Software is accurate and clinically appropriate.2. To determine if any of these Issues are present in your current ABACUSconfiguration, please note the following immediate actions should be taken:"Double Salts: Baxter contacted all facilities installed by Baxter with ABACUS v3.1 to correct the Double Salts Issue via telephone between April 4, and April 16, 2014.To avoid the Double Salts Issue, a pharmacy installed with ABACUS v3.1 should 1) ensure the "Include non-electrolyte ion contributions in balance" checkbox is checked for all salt-based templates, and 2) compare the ordered amount of salts to the delivered amount on the Ions tab in ABACUS."Sterile Water Added to TPN with Pre-Mix: Customer should ensure that all premix not intended for use on the compounder are marked as "Pre-Mix PN Solution" type in the Formulary. Customer should then check the template, under the "Infusion" tab, to ensure the "Infusion Mode" is set to "Pre-Mix (Cyclic)"."Premasol CaP04 Curve Data Points: Review Attachment 1 for information relevant to the use of calcium phosphate curves for Premasol 3% Dex 25 |  | 2034 | 2034 | Worldwide Distribution - USA nationwide, Canada, Colombia, Dominican Republic, Panama, Austria, Bosnia Herzegovina, Bulgaria, Bahrain, Estonia, France, Germany, Georgia, Hungary, Ireland, Israel, Jordan, Kuwait, Latvia, Morocco, Oman, Qatar, Saudi Arabia, Slovenia, Slovakia, South Africa, Switzerland, United Arab Emirates, United Kingdom, Hong Kong, and Taiwan. | Terminated | January 20, 2016 | 658 | 2034 | Computer | Not_Computer
Z-1870-2014 | 68531 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | IS4000 da Vinci Xi - Endoscope accessory; |  |  | June 11, 2014 | 2014 | Intuitive Surgical, Inc. | In some procedures when using the endoscope in conjunction with the IS4000 (da Vinci Xi) System, the surgical view has an orange-yellow tint which can make it hard to differentiate between tissue | Device Design | Medical Device Recall notifications were sent to the 4 affected sites on June 11, 2014. Letters describe the problem and the potential health risk. The two model numbers of the endoscope are listed and the statement that ISI will perform a software upgrade to the da Vinci System and replace the endoscopes.Letters request that the following actions take place:- Ensure that all appropriate personnel are fully informed of the information provided in the notification letter.- Complete and return the acknowledgement form using instructions provided.- Retain a copy of The letter for your records.Intuitive representatives will schedule a site visit to provide the re-calibrated endoscope, update the software and ensure that it is properly integrated with the da Vinci Xi Surgical System.Questions or concerns may be directed to Intuitive Customer Service at 800-876-1310. |  | 20 | 20 units | Distributed in the states of FL, TX, and NY. | Terminated | July 02, 2014 | 21 | 39 | Computer | Not_Computer
Z-1148-2014 | 66119 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Model Number/ Description: |  |  | March 25, 2011 | 2011 | GE Healthcare, LLC | When performing a secondary capture generated from GSI Viewer, the header information is correcton the images. However, when changing the ROI (region of interest) and zooming in, the DICOMheader does not change to reflect the new image, thus the line measurements may be inaccurate. | Software design | No consignee letter was sent out. FMI No. 25431 shows where the Field Engineer had to communicate with the hospital (Instead of them sending a letter).GE Healthcare has updated all of the affected Advantage Workstation systems affected by the GSI viewer issue. At this time, deployment of the recall is considered complete. As an effectiveness check, they have tracked all service requests for the field action to completion.These service requests record the repair of the affected units. Exceptions to completion of the service request have been recorded as well. Records of field action completion will be retained per their processes and procedures. |  | 283 | 283 (59 US, 224 OUS) | Worldwide Distribution: US (nationwide): AL, AZ, CA, CT, FL, IL, IN, IA, KY, MD, MA, MI, MS, NV, NH, NJ, NY, PA, WA, and WI; and countries of: AUSTRALIA, BELGIUM, BRAZIL, CANADA, CHINA, DENMARK, FINLAND, FRANCE, GERMANY, GREECE, HONG KONG, INDIA, INDONESIA, ISRAEL, ITALY, JAPAN, KUWAIT, LATVIA, NEW ZEALAND, POLAND, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SOUTH KOREA, SWEDEN,, SWITZERLAND, TAIWAN, and UNITED KINGDOM. | Terminated | March 13, 2014 | 1084 | 283 | Computer | Computer
Z-0910-2014 | 67152 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | MooreMedical Supporting Health &amp; Care...MooreBrand Blood Glucose Monitor for Self-Testing &amp; Multi-Patient Use TRUEtrack. Manufactured by: NIPRO DIAGNOSTICS, Fort Lauderdale, FL 33309 U.S.A. |  |  | October 02, 2013 | 2013 | Nipro Diagnostics, Inc. | Manufactured with an incorrect factory set unit of measure. | Software design (manufacturing process) | Nipro Diagnostics sent an Urgent Medical Device Voluntary Recall letter dated December 31, 2013 to all affected customers. The customers were instructed to do the following: 1) Screen their inventory immediately for any of the serial numbers listed on the back of the products. 2) If they have any of the affected products, return the product by contacting Stericycle at 1-866-236-4518. 3) Complete and return the enclosed response form via fax at 1-888-358-7294. Customers with questions were instructed to call Stericycle at 1-866-236-4518.For questions regarding this recall call 954-677-9201. |  | 105 | 105 | Worldwide Distribution - USA (nationwide, HI and Puerto Rico and Internationally to Canada, European Union, Jamaica, Suriname, Trinidad and Tobago, Kazakhastan, Russian, Turkey, Australia, Bangladesh, Pakistan, Cambodia, Laos, Myanmar, Thailand, Vietnam, and India. | Terminated | January 28, 2016 | 848 | 606 | Computer | Not_Computer
Z-2234-2014 | 68802 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Vital Sync" VPMP &amp; IM, software, V2.4. (DVD), packaged on a CD or other electronically controlled device. CDs are packaged in clam shell jeweled cases within a carton box. |  |  | July 17, 2014 | 2014 | Nellcor Puritan Bennett Inc. (dba Covidien LP) | Covidien is conducting a field correction of the Vital Sync Virtual Patient Monitoring Platform and Informatics Manager due to a mismatch between the alarm priority on the medical devices and the Vital Sync system. | Software design | The firm, Covidien, sent an "URGENT FIELD CORRECTIVE ACTION NOTICE" letter dated 7/16/14 to its customers on 7/17/14 via FedEx. The letter informs users that a representative will be in contact to schedule a software update.The letter also described the product, problem and actions to be taken. The customers were instructed to disregard the audible and visual alarm priority indicators and only use the alarm description text on the patient tile and on the patient detail page for alarm priority information, until the mismatch is corrected. The Vital Sync" Virtual Patient Monitoring Platform and Informatics Manager or Vital Sync" Remote Monitoring System may still be used until the software update is available.If you have questions about this letter or to report any issues with the Vital Sync" Virtual Patient Monitoring Platform and Informatics Manager or Vital Sync" Remote Monitoring System, contact our Technical Support Department at (800)635-5267, option 3, then option 1. |  | 3 | 3 | US distribution to states of: LA, IL, MS, CO, IN, GA, CA and TX. | Terminated | October 21, 2014 | 96 | 12 | Computer | Computer
Z-2298-2014 | 68817 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Persona (TM) The Personalized Knee System |  |  | June 10, 2013 | 2013 | Zimmer, Inc. | Zimmer had received complaints that reported breakage of the Persona Knee Trial Articular Surface Provisionals (TASPs). This is a retrospective recall which covers Zimmer's corrective actions taken earlier this year concerning the TASPs. | Device Design | On February 10, 2014, Zimmer sent letter to surgeons who had filed complaints regarding the affected TASP fractures. The letter provided the surgeons with Zimmer's investigational findings of the issue. Surgeons were directed to send questions or any additional informtion to Zimmer.PER@Zimmer.com. The Zimmer Weekly Wrap-Up, dated February 14, 2014, was issued electronically to sales force. They were notified to engage in mandatory training to improve the usage of the TASPs and prevent potential breakage. Zimmer has taken additional corrective measures by initiating another recall in July 2014 whereby they notify the distributors as well as all users of the enhanced surgical technique. |  | 70986 | 70,986 distribution events | Worldwide Distribution-USA (nationwide) including DC and Puerto Rico and the states of AK, AL, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, MA, MD, ME, MI, MN, MO, MS, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SD, TN, TX, UT, VA, WA, WI, WY, and the countries of Australia, Canada, Germany, India, Japan, Korea, and Singapore. | Terminated | September 11, 2014 | 458 | 70986 | Computer | Not_Computer
Z-0949-2014 | 65214 | No | LKM | Automated cell counter. | Hematology | Pathology | 510(k) | The UF-1000i is a fully automated urine particle analyzer intended for vitro diagnostic use in urinalysis. The instrument can perform screening for abnormal samples with a high degree of accuracy, so it increases automation and efficiency in laboratories. |  |  | March 31, 2013 | 2013 | Sysmex America, Inc. | The current Century Break patch did not contain the same setting for the Century Break year for both the Red Hat Linux and GUI files. Due to the mismatch in file settings, the software could not properly calculate a purge date for the database files based on the customer data retention setting in the application. This resulted in the database files meeting its maximum capacity and causing system p | Software design | A Product Notification letter dated July 2012, was sent to Siemens Healthcare Diagnostics. The letter suggested that an attached procedure should be followed for an server that requires the Century Break parameter change. If Siemens Healthcare Diagnostics has any questions about the procedure, the firm's representatives can contact the Sysmex Technical Assistance Center at 1-888-879-7639. A response form was also included with the letter requesting the Siemens Healthcare Diagnostics complete and return it. A second Product Notification letter dated September 2012, was sent to Siemens Healthcare Diagnostics. The letter contained largely the same instructions as the July 2012, letter. A third Product Notification letter dated March 2013, was sent to Siemens Healthcare Diagnostics. The letter include substantially similar instructions as the pervious letters; however, the procedure for changing the Century Break parameter was expanded. |  | 214 | 214 units | Worldwide Distribution - USA (nationwide) Distribution including Puerto Rico and the states of AL, AR, CA, CO, CT, FL, GA, HI, IA, IS, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, PA, SC, SD, TN, TX, VA, WA and WI and the country of Canada. | Open, Classified |  |  | 214 | Computer | Computer
Z-2148-2014 | 68071 | No | MQN | Bone plate. | Dental | Dental | 510(k) | DePuy Synthes Craniomaxillofacial Distraction System (AB Distractor Bodies and BC Distractor Bodies) |  |  | April 16, 2014 | 2014 | Synthes, Inc. | DePuy Synthes is initiating a recall of certain lots of the Craniomaxillofacial Distraction System (AB Distractor Bodies and BC Distractor Bodies) because they may reverse post-operatively. | Device Design | The firm, DePuy Synthes, sent "Urgent Notice: Medical Device Recall" letters dated April 16, 2014 to affected customers. The letter identified the affected product, reason for recall, potential patient impact, mitigation steps for patients with distractors currently implanted and actions to be taken. Customers were instructed to review inventory, immediately remove the affected lots from stock, complete verification section of the letter and return the affected product. For questions call 610-719-5450 or contact your DePuy Synthes Sales consultant. |  | 1499 | 1499 | US Nationwide Distribution | Terminated | September 09, 2015 | 511 | 1499 | Computer | Not_Computer
Z-2288-2014 | 68653 | No | FSM | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | PERSONA (TM) The Personalized Knee System Tibial General Instrument Tray, non-sterile |  |  | July 01, 2014 | 2014 | Zimmer, Inc. | The firm discovered that the current Personal Tibial/General Fixed Bracket Trays may not have the specific geometry to accommodate the tibial drill and stop guide with proper position/orientation. The drill bracket and drill guide are not the correct geometry to allow for a proper fit in the Persona Fixed Case Kit which could lead to improper sterilization of the tibial drill and stop guide. | Device Design | On July 1, 2014, Zimmer, Inc. issued recall notification to their consignees via e-mail and/or letter. Notification included affected product issue and instructions for returning the affected product. All distributors were notified via electronic mail. Hospital risk managers, as well as distributors with affected inventory were also notified via courier. Distributors were sent a letter identifying the issue and their responsibilities. These responsibilities include locating and removing the affected product in their territory. Distributors are responsible for returning on-hand affected products to Zimmer and ensuring all of their products are accounted for using the form provided in the letter. Hospital risk managers were provided with a letter identifying the issue and their responsibilities. These responsibilities include locating the affected product, quarantining the product, and returning it to their Zimmer sales representative. |  | 190 | 190 units | Worldwide Distribution - USA including AZ, CA, CO, FL, IL, IN, KS, MI, MN, NC, NJ, NY, OH, OR, PA, TN, TX &amp; VA and Internationally to AUSTRALIA, FRANCE, INDIA, ITALY, and GERMANY. | Terminated | June 09, 2015 | 343 | 190 | Computer | Not_Computer
Z-0827-2014 | 67264 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | Product trade name t:slim" Insulin Delivery System |  |  | January 10, 2014 | 2014 | Tandem Diabetes Care Inc | Tandem Diabetes CareÂż, Inc. announced that it is initiating a voluntary recall of specific lots of insulin cartridges that are used with the t:slimÂż Insulin Pump. The affected cartridges may be at risk for leaking. The cause of the recall was identified during Tandems internal product testing, and has not been associated with any complaints or adverse events reported by customers. However, a c | Equipment maintenance | Tandem Diabetes Care sent an Urgent Medical Device Voluntary Recall letter dated January 10, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed check the lot numbers of their cartridge supply against the llist of affected lot numbers. Do Not use any cartridges from the lots affected by this recall. If customers don't have useable cartridges, discontinue using the pump and revert to their backup plan until new supplies arrive. Customers were instructed to call Tandem Technical Support at 1-877-801-6901 to receive replacement cartridges at no charge. Customers were asked to return the enclosed return response form via email, fax, or mail. UPDATED: Tandem Diabetes Care expanded the recall om January 20, 2014. Tandem Diabetes CareÂż, Inc. (NASDAQ: TNDM), or Tandem, today announced that it is expanding a voluntary recall of specific lots of insulin cartridges that are used with the t:slimÂż Insulin Pump.For questions regarding this recall calll 858-366-6900, ext 6963. |  | 12 | 12, 807 boxes (10packs) | USA (nationwide) | Terminated | September 23, 2014 | 256 | 12 | Computer | Not_Computer
Z-2267-2014 | 68719 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IntelePACS software versions 3-8-1 to 4-9-1, |  |  | July 03, 2014 | 2014 | Intelerad Medical Systems | In specific situation a partially truncated file may be written to the IntelePACS server. This partially truncated file still can be viewed by InteleViewer and InteleViewer will report to the user that the file in question is complete. This situation only occurs if a file is partially truncated. | Software design | Intelerad sent an Urgent Recall for Product Correction letter to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were informed that a new software release will be installed on the system. Afterwards, an audit will be run on the the entire PACS to identify all image data that may have been truncated. Customers were requested to complete and return the attached Client Response Form. |  | 230 | 230 | Worldwide Distribution - USA Nationwide and the countries of Scotland, Ireland, United Kingdom, Canada, Australia, and New Zealand. | Terminated | January 13, 2016 | 559 | 230 | Computer | Computer
Z-0965-2014 | 66947 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | D-10 Dual Program, which contains Update Kit D-10 A1c program diskette. |  |  | November 22, 2013 | 2013 | Bio-Rad Laboratories, Inc. | D-10 A1c Dual Program Floppy was programmed incorrectly and cartridge injections counts are not being updated correctly when switching between the D-10 Dual Short Program and Extended Program on the D-10 Hemoglobin Testing System. | Software design | BIO-RAD sent an Urgent Medical Device Correction letter dated November 22, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to continue to use the affected update kit disk(s) included with the reorder pack(s), but manually track the number of injections on the cartridge, reviewing QC results closely throughout the life of the reorder pack or destroy the affected disk(s) included with these reorder pack(s) and replace with the new Update Kit disk(s) The new Update Kit disks will elminate the problem and may be used with all reorder packs. The replacement disks do not contain reagent or calibrator information; after installing teh new Update Kit disk, they ust manually update the component information using instructions provided with the disk.For questions regarding this recall call 510-741-6265. |  | 2906 | 2906 kits (318 Kits in US; 2588 kits internationally) | Worldwide Distribution - USA (nationwide) and Internationally to Latin America, Asia Pacific, Canada, Australia, New Zealand and India. | Terminated | March 20, 2014 | 118 | 2906 | Computer | Not_Computer
Z-0971-2014 | 67437 | No | LIT | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific, Coyote" MONORAIL" PTA Balloon Dilatation Catheter, UPN # H74939185201510, Catalogue # 39185-20151. |  |  | February 04, 2014 | 2014 | Boston Scientific Corporation | Boston Scientific is initiating this Recall Removal of one lot of the Coyote" OVER-THE-WIRE PTA Balloon Dilatation Catheter and one lot of the Coyote" MONORAIL" PTA Balloon Dilatation Catheter. Boston Scientific has become aware that the balloons of the devices from these lots may not deflate as expected during a procedure. To date, Boston Scientific has received three complaints for this issue. T | Device Design | Consignees were sent a Boston Scientific " Urgent Medical Device Recall Removal" letter on 2/4/2014. The letter was addressed to Materials Manager and Field Action Contact. The letter described the product involved in the recall and the problem. Advised consignees to segregate the affected lots and return to Boston Scientific in accordance to the enclosed recall instructions. The instructions also requested consignees to complete and return the Account Reply Verification Tracking Form (RVTF). |  | 35 | 35 for both products | US: AL, CA, FL, IL, IN, MI, MO, NY, OH, PA, TX.OUS: ITALY, GERMANY, FRANCE, GREAT BRITAIN, NETHERLANDS, WHITE RUSSIA, INDIA, JAPAN. | Terminated | August 04, 2014 | 181 | 35 | Computer | Not_Computer
Z-0979-2014 | 67305 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | GEMINI TF Big Bore PET/CT, Computer Tomography System, Philips Healthcare System, Cleveland, OH |  |  | December 05, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | The TF Big Bore patient table experienced mechanical binding during horizontal table motion resulting in an automatic Emergency Stop (E-stop). An E-Stop will interrupt a current scan, resulting in an incomplete study. | Software design | ON 12/5/2013 the firm sent Medical Device Correction Letters to their customers. |  | 47 | 47 Units | Product was shipped to the following states: AR, CA, CO, FL, IL, LA, MA, MD, MI, MN, MO, NC, NE, NJ, NY, OH, PA, TX, VA, VT &amp; WI.Product was also shipped to the following countries: Belgium, Canada, China, Denmark, France, Germany, Italy, Japan, Netherlands, Saudi Arabia &amp; Venezuela. | Terminated | August 25, 2015 | 628 | 47 | Computer | Not_Computer
Z-2252-2014 | 68979 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Philips IntelliSpace PACS 4.4 image management systems performs digital image processing, measurement and communication and storage. |  |  | July 24, 2014 | 2014 | Philips Healthcare Informatics, Inc. | A problem was detected in the Philips IntelliSpace PAC 4.4 software where if patient information is changed in the DICOM header and the study is exported before the image is updated in the PACS, the unchanged data may export causing the potential for diagnosis on incomplete data. | Software design | Urgent Medical Device Correction was sent to all affected sites on July 24, 2014 by certified mail. |  | 109 | 109 units | US and Australia, Belgium, Canada, Egypt, France, Germany, Netherlands, Oman, Saudi Arabia, Spain, Switzerland, United Arab Emirates, United Kingdom. | Terminated | February 11, 2016 | 567 | 109 | Computer | Computer
Z-0976-2014 | 67374 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Horizon Medical Imaging, Picture Archive Communication System (PACS). |  |  | December 23, 2013 | 2013 | Mckesson Information Solutions LLC | Images may be archived to an incorrect or invalid archive location resulting in incorrect images or no images being displayed for a study following retrieval from archive. | Software design | McKesson sent a field notice to customers with the potentially affected software versions on 12/23/2013. |  | 618 | 618 units | U.S. Nationwide (Puerto Rico), Australia, Canada, France, GU, IE, and The United Kingdom. | Terminated | April 25, 2017 | 1219 | 618 | Computer | Computer
Z-2239-2014 | 68948 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | TomoTherapy Treatment System, Model: Hi-Art, Catalog/Part Number: H-0000-0003, software versions 1.2.0, 1.2.1 and 1.2.2 (Hi-ArtÂż 4.2.0, 4.2.1 and 4.2.2). |  |  | July 17, 2014 | 2014 | TomoTherapy Incorporated | Accuray is voluntarily recalling TomoTherapy H Series software versions 1.2.0, 1.2.1 and 1.2.2 (Hi-ArtÂż 4.2.0, 4.2.1 and 4.2.2). Accuray has identified potential safety issues (anomalies) with these software versions. | Software design | Consignees were sent an Accuray "Urgent Device Correction" letter dated July 16, 2014. The letter was addressed to TomoTherapy System Medical Physicist. The letter described the Purpose of this letter, Description of the Potential Problems, Safety Instructions and software upgrade information. For questions or concerns they can contact Accuray Customer Support at 1-877-668-8667 (USA) or 1-408-716-4700(non-USA) or customersupport@accuray.com |  | 322 | 322 | Worldwide Distribution-USA (nationwide) including DC and Puerto Rico and the states of US:; AL, AK, AZ, CA, CO, CT, FL, GA, HI, IL IN, IA, KS,KY, LA, MD, MA, MI, MO, MT, NE, NV, NJ, NM, NY, NC,ND, OH,OK, OR, PA, SC, SD, TX, UT, VA, WV, and WI, and the countries of AUSTRALIA, BELGIUM CANDA, CHINA, EL SALVADOR, ENGLAND, FRANCE, GERMANY, HONG KONG, INDIA, ITALY, JAPAN, KOREA, MALAYSIA, MEXICO, NETHERLANDS, PHILIPPINE, POLAND, RUSSIA, SAUDI ARABIA, SINGAPORE, SPAIN, SWEDEN, TAIWAN, THAILAND, TURKEY, and UNITED ARAB EMIRATES. | Terminated | September 08, 2015 | 418 | 322 | Computer | Computer
Z-2240-2014 | 68960 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | TomoTherapy Treatment System, Model: Hi-Art, Catalog/Part Number: H-0000-0003, software versions 2.0.0 and 2.0.1 (Hi-ArtÂż 5.0.0 and 5.0.1). |  |  | July 16, 2014 | 2014 | TomoTherapy Incorporated | Accuray is voluntarily recalling TomoTherapy H Series software versions 2.0.0 and 2.0.1 (Hi-ArtÂż 5.0.0 and 5.0.1). Accuray has identified potential safety issues (anomalies) with these software versions. | Software design | Accuray sent an "Urgent Device Correction" letter dated July 16, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers. These anomalies will be corrected with a new software version.You will be contacted by Accuray Customer Support to schedule a software upgrade to your system.If you have any questions or concerns regarding this issue, please do not hesitate to contact Accuray Customer Support at +1.877.668.8667 (USA) or +1.408.716.4700 (non-USA) or customersupport@accuray.com. |  | 168 | 168 (75 US; 93 OUS) | Worldwide Distribution - Nationwide including the states of AZ, AR, CA, FL, GA, IL, IN, KS, KY, MD, MN, NM, NY,NC, OK, OR, SD, TN, TX, UT, WA, WY, WI and WV., and the countries of AUSTRALIA, BELGIUM, CANADA, CHINA, ENGLAND, FRANCE, GERMANY, HONG KONG, ITALY, JAPAN, KOREA, MEXICO, NETHERLANDS, PHILIPPINES, RUSSIA, SPAIN, TAIWAN, TURKEY, MARTINIQUE, ROMANIA, SRI LANKA and SWITZERLAND. | Terminated | July 21, 2015 | 370 | 168 | Computer | Computer
Z-2237-2014 | 68953 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | IntelliSpace Portal (ISP) DX/HX/EX, IntelliSpace Portal (ISP) IX, and IntelliSpace Portal (ISP) LX SPECT |  |  | July 27, 2014 | 2014 | Philips Medical Systems (Cleveland) Inc | The following software issues have been identified in the affected products.Problem 1: When reopening a bookmark generated from processing a MUGA (Multi-Gated Acquisition) scan within the NM Cardiac Application, the ejection fraction (EF) displayed should be the same as the ejection fraction (EF) originally displayed when the bookmark was first created. In some instances, the ejection fraction ( | Software design | Philips Healthcare issued an Urgent Medical Device Correction letter dated July 28, 2014, via certified mail to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Philips Key Market representative would distribute the FSN internationally. Additionally, a Field Service Engineer would contact each site to schedule a time to install the software update.Customers with questions were instructed to contact their local Philips representataive or local Philips Healthcare office. For North America and Canada customers were instructed to contact the Customer Care Solutions Center at 1-800-722-9377, option 1.For questions regarding this recall call 978-659-7703. |  | 1099 | 1099 Units | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO, FL, GA, IL, IN, KS, LA, MA, MD, MI, MN, MO, MS, NC, ND, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, VT, WA, WI, WV, WY,Washington D.C. and Internationally to the following countries: Algeria, Argentina, Australia, Austria, Bangladesh, Belgium, Brazil, Bulgaria, Canada, Cayman Islands, Chile, China, Colombia, Congo, Costa Rica, Czech Republic, Denmark, Ecuador, Egypt, Ethiopia, Finland, France, Gabon, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Indonesia, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Latvia, Lebanon, Libyan Arab Jamahiriya, Lithuania, Malaysia, Mauritius, Mexico, Mongolia, Morocco, Myanmar, Nepal, Netherlands, New Zealand, Nicaragua, Norway, Oman, Pakistan, Palestine, Panama, Peru, Philippines, Poland, Portugal, Puerto Rico, Qatar, Russia, Saudi Arabia, Singapore, Slovakia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Viet Nam, and Yemen. | Terminated | February 02, 2016 | 555 | 1099 | Computer | Computer
Z-0941-2014 | 66467 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | enGen Track System with enGen Select v5.0 with autoverification enabled |  |  | October 08, 2013 | 2013 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics issued an Urgent Product Correction Notification involving the enGen Select V5.0 &amp; enGen Custom v3.2.2 due to an anomaly that affects the autoverification feature and custom rules that depend upon data elements used by the referenced enGen Systems. | Software design | Ortho Clinical Diagnostics (OCD), issued an Urgent Product Correction Notification letter dated October 7, 2013 and on December 17, 2013 to their customers via Federal Express or via US Postal Service. The letter identified the affected product, problem and actions. The notices advise the customers to ensure that all required data elements are entered and in the correct format if using the autoverification or custom rules for the referenced enGen Laboratory Automation Systems. OCD trained service representatives will contact all affected customers. Customers with questions may call the Customer Technical Service at 1-800-421-3311. |  | 8 | 8 units | Worldwide Distribution - USA Nationwide in the states and countries of: AZ, FL, WV, TX, PA, DE, MI, OH, TN, IL, NC, Chile, India, Italy, Portugal, Spain, Sweden, UK | Terminated | April 11, 2017 | 1281 | 28 | Computer | Computer
Z-2222-2014 | 68656 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Double Offset Rasp Handle, Right Hand |  |  | July 23, 2014 | 2014 | Zimmer, Inc. | The firm received complaints for failure of the handle to remain assembled to the rasp during use due to various failure modes, including fracture of the leaf spring. | Device Design | On July 23, 2014, Zimmer issued recall notification to their consignees. Distributors were notified via electronic mail and courier. Recall notification included a description of the issue, affected product, possible risks to patient, and instructions for responding to the recall. Questions or concerns regarding the recall are directed to Zimmer customer call center @ 1-800-348-2759 M-F, 8am-8pm ET. |  | 412 | 412 units | Worldwide Distribution-USA (nationwide) including the states of AK, AL, AZ, CA, CO, FL, GA, IN, KS, KY, MI, MN, NC, NJ, NY, OH, OK, OR, PA, SD, TN, TX, UT, and VA, and the countries of Australia, France, Germany, Italy, Switzerland, and Canada. | Terminated | September 15, 2015 | 419 | 412 | Computer | Not_Computer
Z-0981-2014 | 67343 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | GE Healthcare, PET Discovery 610, Discovery 710, Discovery 600 and Discovery 690. |  |  | November 20, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue due to scatter overcorrection associated with the List Mode Replay feature of your Discovery 600, 610, 710 and Discovery 690. | Software Manufacturing/Software Deployment | GE Healthcare sent an Urgent Medical Device Correction letter dated November 20, 2013, and December 4, 2013, to all affected consignees. The letters described the Safety Issue, Safety Instructions, Affected Products Details, Product Correction and Contact Information ( In the US Call 1-800-437-1171. For other countries contact your local GE Healthcare Service Representative). Customers were given the following instructions:To avoid potential issues, replay the list mode data prior to another acquisition.In the event another acquisition is made before replaying the list data, or following reboot, do the following steps:Âż Perform a PET only acquisition using the same tracer that is intended for the List Mode Replay.Âż Then, proceed to perform the List Mode Replay and subsequent reconstructions.Âż Make sure to perform a quality check on any images created after any replay of data.AffectedProduct DetailsPET Discovery 610 and PET Discovery 710, with software versions pet_coreload.44 and pet_mfk.44. The software version is visible on the homepage of the service desktop.To determine the software version currently on your system, do the following:1. On the right monitor upper left corner select the Service Icon.2. Verify that the Service window is under Home.3. Locate the SW Build Version listed in the center of the page.For questions regarding this recall call 800-437-1171. |  | 376 | 376 (101 US, 275 OUS) | Worldwide Distribution - USA including AL, AZ, AR, CA, CO, FL, GA, IL, IN, IA, KY, LA, MN, MO, NE, NJ, NM, NY, NC, OH, OK, PA, SD, TN, TX, UT, VA, WA, WV, WI and Internationally to:AUSTRALIA, AUSTRIA BELGIUM BOSNIA AND HERZEGOVINA, BRAZIL, CANADA, CHILE, COLOMBIA, CHINA, CROATIA, CZECH REPUBLIC, DENMARK, EGYPT, FRANCE, FINLAND, GEORGIA, GERMANY, HONG KONG, IRAN, ISRAEL INDIA,M ITALY, JAPAN, JORDAN, KUWAIT, LEBANON, MALAYSIA, MEXICO, MOROCCO, NEW ZEALAND, NORWAY, RUSSIA, REPUBLIC OF KOREA, POLAND, ROMANIA, SAUDI ARABIA, SPAIN, SWEDEN SWITZERLAND, TAIWAN, TUNISIA, THAILAND, TURKEY, UNITED KINGDOM. | Terminated | July 29, 2014 | 251 | 376 | Computer | Not_Computer
Z-2213-2014 | 68845 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Navigation Spine &amp; Trauma 3D Version 2.0 and 2.1 |  |  | May 21, 2014 | 2014 | Brainlab AG | Brainlab Navigation Software Spine &amp; Trauma 3D 2.0/2.1 offers automatic registration of intraoperatively acquired CT image data sets. To enable automatic registration, the software requires the gantry position of the scanner. The gantry position can either be entered manually or submitted automatically from the CT scanner. If automatic gantry communication is available, the so-called &quot; | Software in the Use Environment | The firm, Brainlab AG, sent "FIELD SAFETY NOTICE / PRODUCT NOTIFICATION" letters dated May 21, 2014 to all consignees. The letters included instructions for consignees to: 1) not use Automatic Registration of iCT Data in combination with Siemens SOMATOM Definition AS CT scan (with Software Somaris SOM 7 or higher and Automatic Gantry Communication (available for Spine &amp; Trauma 3D 2.0 and 2.1 only); 2) always make sure to perform careful accuracy verification according to the instructions in the Software User Guide; 3) Brainlab will perform corresponding service visits to change the gantry communication from automatic to manual to enable a safe and effective usage of this device combination, and 4) advise the appropriate personnel working in your department of the content of the letter.Customers with any questions can contact Brainlab via telephone at 800-597-5911 or e-mail at us.support@brainlab.com. |  | 17 | 17 systems | Worldwide distribution: US (nationwide) including states of: NC, NY, OH, WI;and countries of: Austria, China, Germany, Italy, Japan, Malaysia, and Saudi Arabia. | Open, Classified |  |  | 17 | Computer | Computer
Z-0982-2014 | 67370 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, Dash 3000, Dash 4000 and Dash 5000 Patient Monitors, Dash Port2 Docking Station. |  |  | November 27, 2013 | 2013 | GE Healthcare, LLC | GE has become aware of multiple issues affecting the Dash 3000/4000/5000 Patient Monitor and Dash Port 2 Docking Station:DASH: 1. Dash stops communicating on the Unity Network resulting in Loss of Monitoring at the CARESCAPE" CIC Pro Central Station*.2. Dash displays distorted ECG waveforms in combination monitoring mode. 3. Dash incorrectly becomes Time Master on the Unity NetworkÂż.4. | Software design | Consignees were sent a GE Healthcare "Urgent Medical Device Correction" letter dated November 27,2013. The letter was addressed to Healthcare Administrator / Risk Manager, Chief of Nursing and Director of Biomedical Engineering. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information (Contact Technical Support at 1-800-558-7044 or your local Service Representative). |  | 178711 | 178,711 (92,123 US, 86,588 OUS) | Worldwide Distribution: USA (nationwide) including DC and PR; and Internationally to: ALGERIA, ANGOLA, ARGENTINA, AUSTRALIA, AUSTRIA, BAHAMAS, BAHRAIN, BANGLADESH, BELARUS, BELGIUM, BOLIVIA, BOSNIA-HERZEGOVINA, BRAZIL, BRUNEI, BULGARIA, CAMBODIA, CAMEROON, CANDA, CARIBBEAN, CHILE, CHINA, COLOMBIA, COSTA RICA, CROATIA, CYPRUS, CZECH REPUBLIC, DENMARK, DJIBOUTI, DOMINICAN REPUBLIC, ECUADOR, EGYPT, EL SALVADOR, ESTONIA, ETHIOPIA, FINLAND, FRANCE, GABON, GEORGIA, GERMANY, GHANA, GREECE, GUATEMALA, HONDURAS, HONG KONG, HUNGARY, ICELAND, INDIA, INDONESIA, IRAN, IRAQ, IRELAND, ISRAEL, ITALY, IVORY COAST, JAMAICA, JAPAN, JORDAN, KAZAKHSTAN, KENYA, KOREA, KUWAIT, LATVIA, LEBANON, LIBYA, LITHUANIA, MACEDONIA, MALAYSIA, MALTA, MAURITIUS, MEXICO, MOLDOVIA, MOROCCO, NEPAL, NETHERLANDS, NEW ZEALAND, NICARAGUA, NIGERIA, NORWAY, OMAN, PAKISTAN, PALESTINE, PANAMA, PARAGUAY, PERU, PHILIPPINES, POLAND, PORTUGAL, QATAR, ROMANIA, RUSSIA, SAUDI ARABIA, SENEGAL, SERBIA MONTENEGRO, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SRI LANKA, SUDAN, SWEDEN, SWITZERLAND, SYRIA, TAIWAN, TAJIKISTAN, TANZANIA, THAILAND, TRINIDAD TOBAGO, TUNISIA, TURKEY, TURKMENISTAN, UGANDA, UKRAINE, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY, VENEZUELA, VIETNAM, YEMEN, and ZIMBABWE. | Open, Classified |  |  | 178711 | Computer | Computer
Z-2199-2014 | 68788 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Prisma M100 PRE pump infusion set. Product number 107791. Each set is packaged in a plastic pouch, 4 pouches per carton box |  |  | July 18, 2014 | 2014 | Gambro Renal Products, Incorporated | Gambro initiated a field action on the Prisma and Prismaflex sets due to breakages of connectors resulting in leakage of liquid or blood. | Device Design | Gambro sent an Urgent Field Safety Notice dated July 18, 2014 to all end-users. A Customer Reply Form was attached to the notice for customers to complete and return. |  | 401451 | 401,451 | Worldwide Distribution. | Terminated | February 05, 2015 | 202 | 401451 | Computer | Not_Computer
Z-2184-2014 | 68837 | No | JXG | Central nervous system fluid shunt and components. | Neurology | Neurology | 510(k) | Codman Certas - In Line Valve Only; Product Code: 82-8800 |  |  | July 03, 2014 | 2014 | Codman &amp; Shurtleff, Inc. | To clarify the CODMAN CERTAS Valve virtual off (setting 8) pressure specification outlined in the product's Instructions for Use (IFU). | Device Design | Codman Neuro sent an Urgent Medical Device Safety Notification dated July 17, 2014, to clinicians who use the device. The notification clarifies the CODMAN CERTAS Valve virtual off (setting 8) pressure specification outlined in the product's Instructions for Use (IFU). Consignees are also asked to complete the enclosed Acknowledgement Form and fax the completed form to 1-888-239-1305. For questions or concerns regarding this notification, consignees are instructed to contact their local Codman Neuro Representative or Scientific and Medical Affairs at SciMedAffairs@its.jnj.com or (866) 685-7325.For questions regarding this recall call 508-828-2726. |  | 9498 | 9,498 devices total | Worldwide Distribution - USA including AL, AZ, CA, CO, CT, DC, DE, FL, GA, IA, ID, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, MT, NC, ND, NE, NJ, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, VT, WA, WI, WV, WY; and Internationally to: Canada, Australia, Austria, Belgium, Brazil, Czech Republic, Denmark, Egypt, Finland, France, Germany, Great Britain, Greece, Hungary, Ireland, Israel, Italy, Kuwait, Lebanon, Luxembourg, Netherlands, Norway, Portugal, Russia, Saudi Arabia, Slovak Republic, Slovenia, South Africa, Spain, Sweden, Switzerland, Turkey. | Open, Classified |  |  | 9498 | Computer | Not_Computer
Z-0988-2014 | 67241 | No | LEZ |  |  | Ear Nose &amp; Throat | Enforcement Discretion | Siemens Prisma 2K Hearing Aid |  |  | July 01, 2005 | 2005 | Siemens Hearing Instruments, Inc | Siemens became aware of an issue where an infant or child could open the tamper resistant battery compartment of the Siemens Prisma 2K hearing aid under certain unintended use conditions. | Device Design | Siemens sent a notice letter to affected customers in 7/2005 and 12/2005. The letter identified the affected product, problem and actions to be taken. Siemens will install an additional security measure to affected product at no charge. For questions contact your Sales Representative at 800-766-4500. |  | 738 | 738 units | USA Nationwide Distribution | Terminated | July 06, 2017 | 4388 | 738 | Computer | Not_Computer
Z-2180-2014 | 68634 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Siemens Syngo RT Therapist and Syngo RT Oncologist intended use of the SIEMENS branded Syngo RT Therapist and Syngo RT Oncologist software applications used with SIEMENS family of linear accelerator systems is to deliver X-Ray photon and electron radiation for the therapeutic treatment of cancer. |  |  | June 17, 2014 | 2014 | Siemens Medical Solutions USA, Inc | Siemens became aware of a potential issue with the syngo RT Therapist 4.3.SP1 and syngo RT Oncologist 4.3SP1 software applications running on Siemens brand LINAC systems that may experience sporadic system crashes that will cause the inability to unload patient data. | Software design | In June, 2014, an urgent medical device correction field safety notice of improvement was sent to end users informing them of the product, problem, and action to be taken to avoid the potential safety issue. |  | 56 | 56 | UT, WI, NY | Terminated | December 15, 2014 | 181 | 56 | Computer | Computer
Z-2179-2014 | 68888 | No | MDS |  |  | Clinical Chemistry | PMA | Medtronic MiniMed Guardian Monitor, Model No. CSS7100, CSS7100K |  |  | July 21, 2014 | 2014 | Medtronic MiniMed Inc. | Medtronic MiniMed is recalling the Guardian REAL-Time monitors because of a manufacturing issue that has the potential to compromise the water resistance of the affected Guardian monitors. | Device Design | Medical Device Recall letters dated July 21, 2014 andJuly 22, 2014 were sent to distributors, patients, and healthcare providers who purchased the Guardian REAL time monitors. The letters inform the customers of the problems identified and the actions to be taken. Distributors with questions are instructed to email rs.nrdlrinbox@medtronic.com. If healthcare professionals want to speak to Medtronic's medical officers then they are instructed to call (818) 576-4211, ext. 64511. Customers and healthcare professionals' patients are instructed to call Guardian Recall Hotline at (800) 646-4633, ext. 21090, 8am-6pm, CST. |  | 86 | 86 units | Worldwide:Distribution - US Nationwide and the countries of Croatia, Denmark, Finland, Germany, Netherlands, Norway, Sweden, Switzerland. | Terminated | October 08, 2014 | 79 | 86 | Computer | Computer
Z-2177-2014 | 68727 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Kodak DirectView DR 7500 Diagnostic X-Ray System |  |  | July 09, 2014 | 2014 | Carestream Health, Inc. | An on-site evaluation of the equipment parts that attach the Beta Assembly (x-ray tube and collimator) to the Over Head Tube Crane (OTC) Telescope was completed by Carestream Health (CSH) Service Engineering. Inspection of these parts revealed that the tab on the lock washer was not bent into the lock nut recess and because the Beta Assembly detached from the OTC telescope, the lock nut was not sufficiently tightened. | Device Design | Carestream will notify all DR 7500 customers of the issue and schedule a Field Engineer to inspect and repair the system if required. Specific instructions (Service Modification) for the inspection and repair will followed by the Carestream Field Engineer in order to complete the correction. Documented evidence of the Field Engineer visit will be noted in the Carestream service engineering record for each affected system. Carestream estimates that the plan of correction will be completed within 2-4 weeks |  | 589 | 589 systems | Worldwide Distribution - US Nationwide and the countries of (United Kingdom, United Arab Emirates, Turkey, Taiwan, Switzerland, Sweden, Spain, South Africa, Portugal, Poland, Oman, New Zealand, Netherlands, Kuwait, Italy, Israel, Hong Kong, Germany, France, Finland, Denmark, Dubai, Czech Republic, China, Korea, Belgium, Australia, and Amsterdam. | Terminated | May 12, 2017 | 1038 | 589 | Computer | Not_Computer
Z-2175-2014 | 68752 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | DX-D 600 - DXD Imaging Package |  |  | July 02, 2014 | 2014 | AGFA Healthcare Corp. | Some customers reported that while using the DX-D600 in room or exam preparation, the Overhead Tube Crane (OTC) showed increased velocity and uncontrolled longitudinal movement. | Device Design | AGFA HealthCare sent an URGENT FIELD SAFETY NOTICE letter dated July 2, 2014 to affected customers. The letter described the affected product, problem and actions to be taken. Customers were instructed to complete and return the Acknowledgment form, via FAX-Back or email. A Mandatory Service Bulletin DIS012.14E/No. 32, was released June 30, 2014, to implement the mandatory upgrade of full automatic DX-D600 systems to version 3.6 to prevent unexpected system movements. This Mandatory Service Bulletin was issued to Agfa DR Specialists to perform the service. A 2nd release will be implemented to upgrade semi-automatic systems. For questions contact your local AGFA HealthCare organization 1-877-777-2432. |  | 17 | 17 units | Worldwide Distribution - US Natinwide in the state of GA, IN, KS, KY, MA, NJ, SC, VA, DC and the country of Canada | Terminated | June 04, 2015 | 337 | 17 | Computer | Not_Computer
Z-2167-2014 | 68648 | No | JOP | Ultrasonic transducer. | Cardiovascular | Cardiovascular | 510(k) | Esaote PA230E Transducer |  |  | June 26, 2014 | 2014 | Biosound Esaote, Inc. | Using the PA230E transducer with a user-defined setting with the output power to 100% (Max) in 2D-TEI mode focusing at maximum depth, could cause the transducer surface temperature to exceed the maximum allowable limit. | Software design | On 06/26/2014, Esaote North America sent an advisory letter via FEDEX (with signature required) to each customer or consignee affected by this voluntary recall. The packet contained the advisory letter, including actions to be taken by the user, information about a pending software upgrade, and a business reply card to confirm they received the letter and have the equipment noted on the advisory letter. Customers with questions were instructed to call 800-428-7378.For questions regarding this recall call 1-317-813-6005 |  | 1301 | 1,301 units | Nationwide-Product was shipped to the following states: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, SC, SD, TN, TX, UT, VA, WA, WI &amp; WV. International**Product was also shipped to consignees in Canada | Terminated | July 17, 2015 | 386 | 1301 | Computer | Computer
Z-2135-2014 | 68596 | No | LMH |  |  | General &amp; Plastic Surgery | PMA | Enhancement Medical expression. model # EXP-150. Injectable Gel 1.5cc Syringe with hyaluronic acid concentration of 20 mg/ml. |  |  | August 07, 2013 | 2013 | Enhancement Medical, LLC | Manufacturer made a change in the production process that resulted in a change in final gel weight. RECALL EXPANDED 7/8/2014 Firm expanded their recall to include all lots of product. | Device Design | Enhancement Medical, LLC called all Independent Sales Representatives and then provided each representative a list of consignees with impacted product and instructed the Representatives to disclose low gel weight issue and potential for common treatment responses. Representatives were instructed to exchange the affected product. EXPANDED Enhancement Medical sent an Urgent Medical Device Recall letter beginning July 8, 2014 to all consignees whom received product. The letter identified all lots as being recalled, the reason for the recall, recommended that Expression not be injected subcutaneously, and asked for product to no longer be used and be returned to the recalling firm. An acknowledgement Form was asked to be mailed, e-mailed or faxed back to the firm.Customers with questions were instructed to call 414-918-4280. |  | 17875 | 17,875 | Worldwide Distribution - USA (nationwide including Puerto Rico) and Internationally to Canada. | Open, Classified |  |  | 17875 | Computer | Not_Computer
Z-2535-2014 | 68983 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Philips IntelliSpace ECG Management System, with software option C61 needed to import ECGs from the Quinton Stress ECG System. |  |  | August 07, 2014 | 2014 | Philips Medical Systems, Inc. | Philips IntelliSpace ECG (IECG) Management Systems that are importing Stress ECG records from a Quinton Stress ECG system have the potential for the patient record to contain multiple and therefore inaccurate patient identifiers such as: patient name, patient ID (PID), Medical record number (MRN). This creates the potential for misdiagnosis and incorrect therapy. | Software design | Philips Healthcare sent an Urgent Medical Device Correction Notification / Field Safety notice dated August 7, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. This Field Safety Notice is intended to inform you about:" what the issue is and under what conditions it can occur" the actions that should be taken by the customer/user in order to prevent risks for patients" the corrective action planned by Philips to address the following issue1)Previously Imported Stress Records:Institutions that have imported Quinton Stress ECG records into IECG should review each previously imported record to make sure that multiple patient ID data is not present. Institutions that imported stress records at the end of each stress exam or did not batch process the records should not find any records with multiple Patient IDs in the record. However, we do recommend that you perform the review to confirm that multiple Patient IDs do not exist.2) Continued Use: Software is now available to address the issue described in this letter. If you have not already been contacted by Philips please call to arrange for immediate installation of the software upgrade.The correction will consist of free of charge software upgrade to A.01.06.. A Philips Healthcare representative will contact customers with affected devices to arrange for service. Further information or support concerning this issue, please contact your local Philips representative or call us at 1-800-722-9377 |  | 11 | 11 units | Worldwide Distribution - US including the states of CA, GA, NC, NH, NY and OR., and the countries of Canada and Spain. | Terminated | October 21, 2015 | 440 | 11 | Computer | Computer
Z-2543-2014 | 68801 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Zimmer PERSONA Distal Valgus Alignment Guide, Nonsterile. orthopedic surgical instrument. |  |  | July 30, 2014 | 2014 | Zimmer, Inc. | Persona Distal Valgus Alignment guide collet locks are seizing in the 'unlocked' position. | Device Design | An Urgent Medical Device Correction letter dated, July 31, 2014, was issued to consignees. The letter identified the affected product, the problem, the risks, and responsibilities. Risk managers and surgeons are to review the notification and the updated surgical technique provided in Attachment 1; as well as, ensure that all affected personnel are aware of the contents. Questions or concerns should be directed to Zimmer customer call center at 1-877-946-2761. |  | 2248 | 2,248 units | Worldwide Distribution -- United States, including Puerto Rico; and, to the countries of Australia, Austria, Japan, Belgium, Italy, Switzerland, India, Germany, Spain, France, Israel, Netherlands, Saudi Arabia, South Africa &amp; Canada. | Terminated | March 06, 2015 | 219 | 2248 | Computer | Not_Computer
Z-1084-2014 | 67184 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Agility. |  |  | November 25, 2013 | 2013 | Elekta, Inc. | During the treatment table, gantry, and collimator calibration procedures, it is possible to type incorrect values or invert a value. | Software design | Corrective Action #1 - Release of Important Field Safety Notice 200 01 502 043 - "Recommended quality assurance after a calibration procedure" Released 25 November, 2013. Corrective Action #2 - Permanent Solution - All users of Integrity 3.1 software will be informed or a mandatory upgrade to Integrity 3.2 upon its release which will prevent calibration from being inverted. Target Release for Integrity 3.2 will be available at the end of February 2014. |  | 300 | 300 | Worldwide Distribution-USA (nationwide) including DC and Puerto Rico and the states of AL, AK, AZ, CA, CT, FL, GA, IL, KY, LA, MD, MA, MI, MO, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SD, TN, TX, VT, VA, WA, WV, WI, and WY, and the countries of Algeria, Argentina, Australia, Austria, Bahrain, Belarus, Belgium, Brazil, Bulgaria, Canada, China, Colombia, Czech Republic, Cuba, Denmark, Ecuador, Estonia, Finland, France, Germany, Greece, Honduras, Hong Kong, Hungary, India, Indonesia, Iraq, Ireland, Israel, Italy, Japan, Jordan, Malaysia, Mexico, Morocco, Netherlands, New Zealand, Nigeria, Norway, Pakistan , Panama, Peru, Poland, Portugal, Romania, Russia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sudan, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom, Ukraine, Venezuela, and Vietnam. | Open, Classified |  |  | 889 | Computer | Computer
Z-2586-2014 | 69125 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | ARIATELE TELEMETRY TRANSMITTER, Model 96281, with SpO2 Option C. Physiological patient monitor. Used in a hospital environment. |  |  | July 02, 2014 | 2014 | Spacelabs Healthcare Inc | The transmitter display and the Telemetry Central Station display may show a SpO2 value when the sensor is not connected to the patient or in certain sensor failed conditions.This is an expansion of recall Z-0871-2014 because additional serial numbers were identified as being part of this recall action as part of the firm's retrospective analysis of all recalls since 2011. | Software design | Spacelabs notified both consignees via email of the expanded recall July 2, 2014. A f/u letter was sent by overnight delivery to the single new consignee the same day. Spacelabs sent an announcement of the expanded recall to both consignees identified in the expanded recall August 25, 2014. A field service technician from Spacelabs will visit each consignee and upgrade the software as necessary.If you have any questions about this corrective action program, please contact SpaceLabs Healthcare at 1-800- 522-7025, and select 2 for Technical Support. |  | 8 | 8 units | US Distribution to TX. | Terminated | October 16, 2014 | 106 | 8 | Computer | Computer
Z-2637-2014 | 69086 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ProVue Instrument Driver for JResultNet (ProVue.jrm) v. 1.5.3 and below |  |  | August 15, 2014 | 2014 | Data Innovations, LLC | Potential issues with the ProVue driver have been identified which may cause data to be misassociated with patient results under a given set of circumstances. | Software design | Data Innovations LLC sent an Important Medical Device Information letter on August 15, 2014, to consignees for the JResultNet using ProVue Instrument Driver JResultNet (provue.jrm) v. 1.5.3 and below. Potential issues with the ProVue driver have been identified which may cause data to be mis-associated with patient results under a given set of circumstances. The user on the ProVue transmits the same results twice back to back by clicking the Send to Host button twice. The ProVue sends two ASTM messages back to back before ending transmission (sending an EOT character) The solution was to disconnect JResultNet from the ProVue instrument and apply the patch to driver provue.jrm when it is available. If you have further questions regarding this notice, please feel free to contact us at northamerica-support@datainnovations.com or 239-931-6004 Toll-Free (from the US and Canada): 1-800-332-0499. |  | 84 | 84 | Nationwide Distribution including AL, AZ, CA, CT, DC, FL, GA, IN, LA, MI, MO, NC, NH, NJ, NY, OH, OK, OR, PA, TN, TX, UT, VA, and WA. | Terminated | March 04, 2015 | 201 | 84 | Computer | Computer
Z-0907-2014 | 67138 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Siemens SOMATOM Definition, Definition AS, Definition Flash, and Definition Edge systems. |  |  | November 15, 2013 | 2013 | Siemens Medical Solutions USA, Inc | During clinical operation, a malfunction may occur causing a system display freeze with SOMATOM Definition, Definition AS, and definition Edge while running software version syngo CT2012 B and syngo CT2013B. | Software design | A recall notification letter, dated November 15, 2013, was sent to Users. |  | 147 | 147 | Nationwide Distribution. | Terminated | March 10, 2015 | 480 | 147 | Computer | Computer
Z-2635-2014 | 69038 | No | LEJ | Sunlamp products and ultraviolet lamps intended for use in sunlamp products. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | SUN DASH RADIUS 252 PN 8000809705. |  |  | July 25, 2014 | 2014 | JK Products &amp; Services, Inc | The integrity of the wires inside main electrical chord become compromised from excessive flexing (during the door open/close process). | Device Design | The owners of the affected device will be notified via certified letter. After the recalling firm confirms ownership, a technician will be dispatched to install a door clamp and door stop, and inspect the condition of the main electrical chord. |  | 1948 | 1948 units | Nationwide Distribution. | Open, Classified |  |  | 1948 | Computer | Not_Computer
Z-2626-2014 | 68908 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker Aria Implant Inserter, REF 48758500, |  |  | July 07, 2014 | 2014 | Stryker Spine | Stryker has initiated the recall of its Aria Standard and Specialty Implant Inserters due to complaints of product fracturing. | Device Design | Stryker Spine sent an "Urgent Medical Device Removal Notification/Customer Response" form dated July 7, 2014 to all their affected customers via Fedex. Customers were instructed to examine their inventory and hospital locations to identify the product. Fax a copy of the Customer Response Form to 855-632-9049 or email it to Spine-RegulatoryAction@Stryker.com.For questions regarding this recall call 201-760-8298. |  | 64 | 64 units | Worldwide Distribution - USA including OR, AZ, TN, MA, MD, GA, FL, NY, IA, IN, TX, MI, PA, UT, CA, MN, MO, OH, KS, OK, MT, FL, AL, CO, TN, and WI and Internationally to Australia. | Open, Classified |  |  | 121 | Computer | Not_Computer
Z-2292-2014 | 68785 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Siemens Luminos dRF, Ysio or Uroskop Omnia systems. |  |  | July 07, 2014 | 2014 | Siemens Medical Solutions USA, Inc | It was discovered that during a RAD examination using Siemens Luminos dRF, Ysio or Uroskop Omnia systems with software version VB10C to VB10F and automatic exposure control, a highly unlikely malfunction can result in longer radiation exposure than required. This may also result in an overexposed image that is not of diagnostic quality. As a result the taken examination needs to be repeated. Thi | Device Design | Siemens had informed all affected customers about this issue via Customer Safety Advisory Notice XP023/13/S (dated June 3, 2013) reported under C&R; # 2240869-05/16/13-0013-C. A corresponding software update to fix the issue was originally planned to be rolled out by May 2014. The originally scheduled roll out date could not be met and to ensure that this new software version works accurately Siemens incorporated a comprehensive monitoring phase. Siemens has now scheduled the roll out for the software version VB10J to be completed by end of October 2014. A letter, dated July 7, 2014, with important safety information was sent to end users to notify them that the new software version VB10J will be introduced to solve the potential malfunction of the generator. |  | 192 | 192 | Nationwide Distribution. | Terminated | February 24, 2015 | 232 | 192 | Computer | Computer
Z-0908-2014 | 67242 | No | LWR |  |  | Cardiovascular | PMA | Product Name: Model 7305- Mosaic Aortic Obturator |  |  | January 09, 2014 | 2014 | Medtronic Cardiovascular Surgery-the Heart Valve Division | Medtronic, Inc. initiated a recall of Medtronic Mosaic Aortic Obturator Model 7305 used with the Mosaic Porcine Aortic Bioprosthesis, Model Number 305,because of higher-than-expected transvalvular gradients occurring post implant. | Device Design | Medtronic sent an Customer Notification letter dated January 9, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers. Medtronic recommends surgeons continue to remain diligent in selecting the appropriate size of Mosaic Aortic Bioprosthesis. Adoption of the updated Mosaic Aortic Bioprosthesis sizing chart (iEOA chart) and Mosaic Obturators/Sizers will facilitate optimal aortic valve size selection.Your Medtronic Field Representative will be contacting you about receiving a new Mosaic Aortic Obturator/Sizer set and sizing chart. If you would prefer to order the new Mosaic Aortic Obturator/Sizer set directly through Medtronic Customer Service, contact them at 1-800-848-9300.Medtronic will notify all applicable regulatory agencies, as required, about this matter. Please share this notification with others in your organization as appropriate. We appreciate your review of this notification and apologize for the inconvenience that it may cause. If you have any questions, please contact your Medtronic Field Representative or Lifeline Technical Services at 1-877-526-7890. |  | 8276 | 8276 units | Worldwide Distribution - USA (nationwide) | Terminated | May 30, 2014 | 141 | 8276 | Computer | Not_Computer
Z-2616-2014 | 69036 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Streamline MIS Spinal Fixation System Fixed Rod Holder |  |  | June 14, 2013 | 2013 | PIONEER SURGICAL TECHNOLOGY, INC. | The affected fixed rod holders did not allow the instrument to lock down on the spinal rod. | Device Design | Pioneer Surgical sent a Product Field Action letter on July 7, 2013, to all affected customers. The letter described the affected product and issue, and provided instructions for return of the recalled product. All 9 affected instruments were returned to Pioneer Surgical, reconditioned and sent back to the field. Customers with questions were instructed to call 906-226-4489.For questions regarding this recall call 906-226-4812. |  | 9 | 9 | Nationwide Distribution including TX, CA, IL, and FL. | Terminated | September 30, 2014 | 473 | 9 | Computer | Not_Computer
Z-2611-2014 | 68959 | No | PFF | High throughput genomic sequence analyzer for clinical use. | Immunology | Toxicology | 510(K) Exempt | Illumina Worklist Manager (IWM) (software v1.0.15), a component of Illumina MiSeqDx Platform. |  |  | July 28, 2014 | 2014 | Illumina Inc | A software component of the MiSeqDx instrument, called the Illumina Worklist Manager(IWM) v1.0.15, fails to perform as intended | Device Design | Illumina notified affected customers via phone, email and Product Quality Notification letter dated July 28, 2014 to describe the software anomaly. The letter identified the affected product, problem and actions to be taken. For questions contact your local support team or Illumina Technical Support at techsupport@illumina.com. |  | 35 | 35 | Worldwide Distribution - US (nationwide) and the countries of: France, Germany, Italy, Norway, UK and United Arab Emirates. | Terminated | November 12, 2014 | 107 | 35 | Computer | Computer
Z-0911-2014 | 67358 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Precision 500D, Precision MPI &amp; Precision RXI X-ray imaging systems. |  |  | December 12, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue involving the Precision 500D, Precision MPI &amp; Precision RXI X-ray imaging systems overhead LCD Video Monitor Suspension (Single monitor &amp; Dual monitor).There is a potential for the locking nut to loosen out of the LCD Video Monitor Suspension yoke. This will not lead to an immediate fall of the suspended monitor since it will | Device Design | The firm, GE Healthcare, sent an "Urgent Medical Device Correction" letter dated December 11, 2013 to its consignees. The letter was addressed to Hospital Administrators / Risk Managers, Managers of Radiology/ Cardiology and Radiologists / Cardiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. The consignees were instructed to immediately contact your local GE Healthcare Service Representative if you observe any inappropriate performance or excessive movement of the monitor suspension yoke.GE Healthcare will correct all affected systems at no cost to you. A GE Healthcare service representative will contact you to arrange for this correction.If you have any questions or concerns regarding this notification, please call one of the following numbers: United States: 800-437-1171, Japan: 0120-055-919, for other countries contact your local GE Healthcare Service Representative. |  | 816 | 816 (722 US, 94 OUS). | Worldwide distribution: US (nationwide) including Guam and DC but excluding CT, DE, HI and RI; and internationally to: AUSTRALIA, BULGARIA, CANADA, CHILE, DOMINICAN REPUBLIC, GERMANY, HONDURAS, IRELAND, ITALY, KUWAIT, LEBANON, MALTA, NAMIBIA, PORTUGAL, ROMANIA, RUSSIA, SAUDI ARABIA, SLOVAKIA, SPAIN, SWITZERLAND, TAIWAN, TURKEY, UNITED KINGDOM, VENEZUELA, and YEMEN. | Terminated | September 11, 2014 | 273 | 816 | Computer | Computer
Z-2609-2014 | 68821 | No | DOP | Mass spectrometer for clinical use. | Clinical Chemistry | Toxicology | 510(K) Exempt | AB SCIEX API 3200MD" LC/MS/MS System with software: |  |  | July 15, 2014 | 2014 | Ab Sciex | Software- AB SCIEX API 3200MD LC/MS/MS SYSTEM with MultiQuantMD 3.0 software may provide incorrect data for patient diagnosis | Software design | AB SIEX issued letter on 7/15/14, notifying customers to implement this field correction, a temporary fix to be implemented immediately.Additionally, a software update is in development and is expected to be available in approximately 8 weeks. Upon availability, AB SCIEX will send a DVD with instructions on how to install the new software update. At that time, outlined temporary actions will no longer be necessary.If you have any questions regarding this recall, potential hazards or actions to be taken please feel free to contact AB SCI EX at +1 289 982 2531.Please confirm receipt of this letter by signing and faxing back the attached Acknowledgement Form. |  | 6 | 6 units | FL, NC, OK, PA, TX | Terminated | October 01, 2014 | 78 | 6 | Computer | Computer
Z-2599-2014 | 69058 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | CLAW PLATE 20mm INTERAXIS CHARLOTTE F&A; SYSTEM, REF 40S20020, 1 EACH, Rx ONLY, NON STERILE |  |  | July 30, 2014 | 2014 | Wright Medical Technology, Inc. | Plates manufactured from an incorrect raw material. | Device Design | Wright Medical sent an Urgent Medical Device Recall letter dated July 25, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their inventory for the affected product and immediately quarantine all affected product and return to Wright Medical as soon as possible. For return instructions customers were instructed to contact Customer Service at 800-238-7117. Customers were also asked to complete the attached response form, confirming receipt of the notice and the status of the product(s) on hand and return by fax to 901-867-7401.Customers with questions were instructed to call 901-290-5662.For questions regarding this recall call 800-238-7117. |  |  |  | Worldwide Distribution - USA (Nationwide) and Internationally to Canada Australia, Great Britain, Germany, Spain, Netherlands, France. | Open, Classified |  |  |  | Computer | Not_Computer
Z-2597-2014 | 68982 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | RADIANCE v. 2.4X, 2.5X and 2.6X |  |  | July 28, 2014 | 2014 | Radiometer America Inc |  | Software design | Radiometer sent an Product Notification letter dated July 31, 2014 to affected customers via FedEx. The letter identified the affected product, problem and actions to be taken. The letter informed customers that local RADIOMETER IT specialist will contact customers in order to schedule the upgrade to v. 2.74 or 3.5.2. All RADIANCE versions below 2.7X will be upgraded to v. 2.74 or 3.5.2.For questions regarding this letter or the Recall, please contact Radiometer America Technical Support at 1-800-736-0600 opt.4. |  |  |  | Worldwide Distribution - US (nationwide) in the states of AZ, CA, MO, MS, TX and UT. and the countries of: Austrialia, Austria, France, UK, Italy, Japan, Spain, India, and the Netherlands. | Terminated | October 15, 2014 | 79 |  | Computer | Computer
Z-2583-2014 | 68655 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Natural-Knee System Patella Bushings. Orthopedic surgical instrument. Part Number: 6290-00-702. |  |  | July 10, 2014 | 2014 | Zimmer, Inc. | Zimmer received a trend of complaints indicating corrosion of product. | Device Design | Zimmer issued an Urgent Medical Device Recall-Lot Specific notification via e-mail/letter dated July 19, 2014, to all affected customers. The notification identified the product the problem and the action needed to be taken by the customer. All distributors were notified via electronic mail. Hospital risk managers and surgeons, as well as distributors with affected inventory were notified via courier. Distributors were sent a letter identifying the issue and their responsibilities. These responsibilities include locating and removing the affected product in their territory as well as identifying hospitals and surgeons who have previously used the affected product. Distributors are to return on-hand affected products to Zimmer, and ensure all of their products are accounted for using the form provided in the letter. Hospital risk managers and surgeons were provided with a letter identifying the issue and their responsibilities. These responsibilities include locating the affected product, quarantining the product, and returning it to their Zimmer sales representative. If you have questions or concerns regarding this recall please call the customer call center at 1-877-946-2761. |  | 8638 | 8638 distribution events | Worldwide Distribution - US Distribution and the countries of : Argentina, Australia, Brazil, France, Germany, India, Japan, Spain, Switzerland, and Venezuela. | Terminated | August 20, 2015 | 406 | 8638 | Computer | Not_Computer
Z-2576-2014 | 68878 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | PERSONA The Personalized Knee System Constrained Posterior Stabilized Tibial Articular Surface Provisional TOP with the following: L 3-5 CD; L 6-9 CD; L 3-5 EF; L 6-9 EF; L 10-11 EF; L 6-9 GH; L 10-12 GH; L 10-12 J; R 3-5 CD; R 6-9 CD; R 3-5 EF; R 6-9 EF; R 10-11 EF; R 6-9 GH; R 10-12 GH; and R 10-12 J. |  |  | August 07, 2014 | 2014 | Zimmer, Inc. | Reports of breakage of the Persona Knee Trial Articular Surface Provisionals (TASPs). The Persona primary knee surgical technique has been updated. This field action correction is to ensure that all distributors and users of this instrument are aware of the enhanced surgical techniques. | Device Design | On 8/7/2014, Zimmer notified all distributors via electronic mail. Distributors that have received affected inventory are also notified via courier. Hospital risk managers and surgeons are also notified via courier.Hospital risk managers and surgeons were provided with a notification identifying the issue and their responsibilities. These responsibilities include ensuring the affected personnel are aware of the contents of the notification. Distributors were sent a notification identifying the issue and to provide a copy of the notification sent to surgeons and risk managers. The distributors will also be instructed as needed to assist with identifying a comprehensive list of surgeons. |  | 77737 | 77,737 units total | Worldwide Distribution: US (nationwide) including states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, and WY; and countries of: AUSTRALIA, INDIA, JAPAN, KOREA, SINGAPORE, AUSTRIA, BELGIUM, FRANCE, GERMANY, ISRAEL, ITALY, LUXEMBOURG, NETHERLANDS, SAUDI ARABIA, SOUTH AFRICA, SPAIN, SWITZERLAND, UNITED KINGDOM, and UNITED ARAB EMIRATES. | Terminated | January 27, 2015 | 173 | 77737 | Computer | Not_Computer
Z-2544-2014 | 68277 | No | OYC |  |  | Clinical Chemistry | PMA | Medtronic MiniMed Paradigm Insulin Infusion Pumps, |  |  | March 13, 2014 | 2014 | Medtronic MiniMed Inc. | Medtronic is recalling the Medtronic MiniMed Paradigm Insulin Infusion Pumps because they have received reports that users have accidentally programmed the pump to deliver the maximum bolus amount. | Device Design | The firm, Medtronic, sent an "IMPORTANT MEDICAL DEVICE SAFETY INFORMATION" letter dated March 13, 2014 to its customers. The letter described the product, problem and actions to be taken. The letter informs the customers that because accidental button pressing errors may occur it is important that customers always confirm the insulin dose flashing on the display is correct before pressing ACT to start delivery. The letter informs the customers of what safety settings will minimize the risk of an accidental programming error and what actions to be taken. Should you have any questions about this letter or concerns with your keypad, please do not hesitate to contact the 24 hour helpline at (866) 222-7304 for a replacement. |  | 559374 | 559,374 units total (444,374 units in US) | Worldwide Distribution-Nationwide in US | Terminated | April 07, 2015 | 390 | 559374 | Computer | Not_Computer
Z-2581-2014 | 68985 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco |  |  | August 18, 2014 | 2014 | Elekta, Inc. | Incorrect Delivery of Composite VMAT Fields - When DICOM is exporting an arc plan (VMAT, Dynamic Conformal or 3D Static Arcs) and using the Composite Field Sequencing (CFS) feature, the DICOM exported plan will not match the approved treatment plan. | Software design | The firm, Elekta, issued an "URGENT IMPORTANT FIELD SAFETY NOTICE (IFSN 382-01-MON-001)" to all affected customers on August 18, 2014. The notice describes the product, problems and actions taken. Confirmation of receipt is to be completed by October 31, 2014. The problem will be resolved in patches to the following releases - Monaco 3.30 and Monaco 5.00. Development is underway with a release date of September 2014 with confirmation of installation of the patch no later than 6 months from release of the patch. An updated notice from the manufacturer is also being sent to affected customers (FCO: 382-01-MON-001 VID:2.0). If you have any queries about this Notice, please contact your local Elekta office or call 770-300-9725 or email: Elekta_Global_PMS@elekta.com. |  | 928 | 928 units | Worldwide distribution: US (nationwide) including states of: AL, AK, AZ, CA, CO, CT, FL, IL, IN, MD, MA, MI, MN, MO, MT, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, SD, TN, TX, VA, WA, WI, Puerto Rico; and countries of: Algeria, Angola, Antigua &amp; Barbuda, Australia, Austria, Bahrain, Belarus, Belgium, Brazil, Canada, Chile, China, Colombia, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, Estonia, Finland, France, Germany, Greece, Hong Kong, India, Iraq, Ireland, Israel, Italy, Japan, Kazakhstan, Libya, Lithuania, Malaysia, Malta, Mexico, Morocco, Myanmar, Namibia, Netherlands, New Zealand, Norway, Pakistan, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Singapore, Slovenia, South Africa, South Korea, Spain, Sir Lanka, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom, and Ukraine. | Open, Classified |  |  | 928 | Computer | Not_Computer
Z-0926-2014 | 67248 | No | IKZ | Powered patient rotation bed. | Physical Medicine | Physical Medicine | 510(K) Exempt | The Rotoprone therapy System is an advanced patient care system for the treatment and prevention of pulmonary complications. The Rotoprone Therapy System provides up to 62 degrees in both prone and supine positions as well as Trendelenburg/Reverse Trendelenburg positioning. |  |  | September 11, 2013 | 2013 | ArjoHuntleigh Inc. | The lock pin can become stuck in the "in" position during the manual rotation, resulting in the inability to prone a patient | Device Design | ArjoHuntleigh Inc. sent a Technical Bulletin package to all affected service centers and an Urgent Field Safety Notice package to all affected customers. For rental units, a training plan was established for employees. For sold units, the notice contain a revised user manual and label explaining what to do if the lock pin becomes stuck in the "in" position. Customers that purchased the affected product were required to complete and return a Customer Response form to ArjoHuntleigh to confirm receipt and understanding of the field safety notice, and that the required actions were completed. For questions call 630-785-4885. |  | 231 | 231 units | US Nationwide and Middle East | Terminated | January 08, 2015 | 484 | 231 | Computer | Not_Computer
Z-0932-2014 | 67167 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Diagnostic imaging systems (GEMINI LXL Model #882412, TruFlight Select Model #882438, GEMINI TF PET/CT16 Model #882470, GEMINI TF PET/CT 64 Model #882471, GEMINI TF Ready Model #882473 &amp; GEMINI TF Big Bore Model #882476) |  |  | December 04, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | The PET Reconstruction Server (PRS) database may randomly lock up after the low dose CT acquisition is complete and will not allow the PET acquisition to begin, resulting in an unexpected incomplete study. | Software design | On 12/4/2013, the firm sent Field Safety Noticies to their customers. |  | 386 | 386 units | Product was shipped to the following states: AL, AR, AZ, CA, CO, DC, FL, GA, HI, IA, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, NC, ND, NE, NJ, NV, NY, NY , OH, OR, PA, TN, TX, VA, VT, WA &amp; WI.Product was also shipped to the following countries: Australia, Belgium, Brazil, Bulgaria, Canada, China, Croatia, Cuba, Denmark, Egypt, France, Germany, Hong Kong, Hungary, India, Indonesia, Italy, Japan, Kuwait, Lebanon, Malta, Mexico, Netherlands, Pakistan, Philippines, Poland, Portugal, Republic of Korea, Russia, Russian Federation, Saudi Arabia, Singapore, Slovakia, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, Ukraine, Ukraine, United Kingdom, Venezuela &amp; Vietnam. | Terminated | February 02, 2016 | 790 | 386 | Computer | Not_Computer
Z-2574-2014 | 68984 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | FOCAL Sim Workstation, Radiation Treatment Planning System. |  |  | August 18, 2014 | 2014 | Elekta, Inc. | When DICOM is exporting an arc plan (VMAT, Dynamic Conformal of 3D Static Arcs) and using the Composite Field Sequencing (CFS) feature, the DICOM exported plan will not match the approved treatment plan. | Software design | Elekta, Inc. issued an Urgent Important Field Safety Notice letter (IFSN 382-01-MON-001) to all affected customers.. Confirmation of receipt is to be completed by October 31, 2014. The problem will be resolved in a patch to Focal 4.80. Development is underway with a release date of September 2014 with confirmation of installation of the patch for the fix no later than 6 months from release of the patch. An updated notice from the manufacturer is also being sent to affected customers (FCO: 382-01-MON-001 VID:2.0). For questions contact your local Elekta office. |  | 554 | 554 units | Worldwide Distribution - US (nationwide)in the states of AK, CA, FL, MI, MN, MO, NJ, NY, NC, OH, PA, TN, VA,WA and Puerto Rico, Algeria, Antigua &amp; Barbuda, Australia, Austria, Bahrain, Belarus, Brazil, Bulgaria, Canada, Chile, China, Colombia, Cyprus, Czech republic, Ecuador, Egypt, Estonia, France, Germany, Greece, India, Iraq, Israel, Italy, Japan, Kazakhstan, Malaysia, Malta, Mexico, Morocco, Netherlands, New Zealand, Pakistan Portugal, Romania, Russia, Saudi Arabia, Singapore, Solvenia, South Africa, South Korea, Spain, Switzerland, Thailand, Turkey, and United Kingdom | Terminated | August 08, 2017 | 1086 | 554 | Computer | Not_Computer
Z-0928-2014 | 67224 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO RTP System. |  |  | November 25, 2013 | 2013 | Elekta, Inc. | When bolus is present, the Effective Depth (with bolus) to the weight point returned on the Source Data Report is incorrect. | Software design | Corrective Action 1 - An important safety notice, LINXIO0018 - Incorrect Effective Depth for Bolused Beams is in the process of being distributed to affected customers. (Release date November 25, 2013). Corrective Action 2 - Develop a fix to the defect in an expeditious timeframe and make available in Patch Release 4.0.03 and Release 5.00.01. |  | 1424 | 1424 | Worldwide Distribution-USA (nationwide) including DC and Puerto Rico and the states of AL, AK, AZ, CA, CO, CT, DE, FL, GA, IL, IN, KS, KY, LA, ME, MD, MI, MN, MS, MO, MT, NE, NV, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, QA, WA, WV, WI, and WY, and the countries of Albania, Algeria, Argentina, Austria, Australia, Bosnia, Bulgaria, Belgium, Bahrain, Brazil, Belarus, Botswana, Bangladesh, Canada, China, Columbia, Costa Rica, Croatia, Cuba, Cyprus, Chile, Czech Republic, Ecuador, Egypt, Estonia, France, Finland, Gabon, Germany, Great Britain, Greece, Guatemala, Hungary, India, Indonesia, Italy, Ireland, Israel, Iraq, Jordan, Japan, Kosovo, Kuwait, Lithuania, Libya, Latvia, Morocco, Malaysia, Malta, Mexico, Myanmar, Netherland, New Zealand, Pakistan, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovenia, Slovakia, South Africa, South Korea, Spain, Sudan, Syria, Taiwan, Thailand, Turkey, Turkmenistan, Ukraine, Venezuela, and Vietnam. | Terminated | December 02, 2015 | 737 | 1424 | Computer | Not_Computer
Z-2564-2014 | 68758 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Hitachi Echelon Oval MRI system is a diagnostic imaging device (one unit per package) and is intended to provide the physician with physiological and clinical information, obtained non-invasively and without the use of ionizing radiation. |  |  | August 18, 2014 | 2014 | Hitachi Medical Systems America Inc | The Gradient Coil was found to have a failure mode that allowed it to overheat and become a burn hazard. | Device Design | Letters will be sent to customers. Hatachi Serivce will be sent to each site to exchange the Gradient Coil Assembly. |  | 39 | 39 systems | Worldwide Distribution: US (nationwide) including states of: CA, DE, FL, ID, KS, KY, MD, MS, NY, OH, and WY; countries of: Japan, Brazil, France, Germany, and Kyrgyzstan. | Terminated | September 14, 2015 | 392 | 39 | Computer | Not_Computer
Z-2563-2014 | 68912 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Roche cobas IT 1000 |  |  | July 31, 2014 | 2014 | Roche Diagnostics Operations, Inc. | Potential that test results would be assigned to the wrong Patient ID if cobas IT 1000 receives an ADT ADMIT, TRANSFER or MERGE event message that contains a genuine Patient ID (e.g. 999999), which is the same as what has been configured as the Generic Patient ID (999999 by default). | Software design | Roche sent an Urgent Medical Device Correction letter dated July 31, 2014 via UPS Ground (receipt required) to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to follow the steps provided in the letter to prevent assigning patient results to an incorrect generic patient ID. Please contact the Roche Support Network Customer Support Center at 1-800-440-3638, prompt 5, 24 hours a day, seven days a week if you have questions about the information contained in this UMDC. |  | 148 | 148 Units | US Nationwide Distribution in the states of: AK, AL, AR, AZ, CA, CT, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MD, ME, MI, MO, NC, NE, NH, NJ, NY, OH, OR, PA, PR, SC, TN, TX, VA, VT, WA, WI, and WV. | Terminated | June 16, 2015 | 320 | 148 | Computer | Computer
Z-2561-2014 | 68965 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens EasyLink Informatics System V4.0/V5 |  |  | August 13, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | May not perform as intended under certain conditions, causing the release of results to the Laboratory Information System (LIS) that should have been held for manual review due to auto-verification rules or the delay/omission of resulttransmission to the LIS. | Software design | Siemens issued an Urgent Medical Device Correction (UMDC) letter on August 13, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to review this letter in its entirety and ensure that all operators understand the information presented. The issues described are specific to the features listed in the letter.. If the laboratory is not using a particular feature, then they are not affected by the issues described for that feature. For questions contact your Siemens Customer Care Center or your local Siemens technical supportrepresentative. |  | 1347 | 1347 | US Nationwide Distribution | Terminated | December 02, 2014 | 476 | 1347 | Computer | Computer
Z-0933-2014 | 66861 | No | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | The Gaymar Medi-Therm device provides a means of regulating patient temperature by supplying temperature-controlled water through a connector hose to accessory Gaymar Hyper/Hypothermia blanket(s)/body wrap(s). The blanket/body wrap provides an interface for heating or cooling the patient. Accessory YSI 400 series patient probe interfaces between the |  |  | December 27, 2013 | 2013 | Stryker Medical Division of Stryker Corporation | Stryker Medical has identified that on some units of the MTA7900, there is a lack of a water temperature display while the unit is in Auto mode. While in Auto mode, only the patient temperature is displayed along with the set point. The water display will be blank during Auto mode. If the user wants to see the water temperature displayed and switches out of Auto mode to Manual mode, then back to | Device Design | On Dec. 27, 2013, Stryker Medical notified their consignees, identifying the product, reason for the recall and associated risks. Consignees were asked . Locate the units listed attach enclosed warning tags to the units listed. For units that do not have the water temperature displayed, advise staff not to switch to manual mode in order to display the water temperature.Stryker will make arrangements to update the software so that the water temperature will be displayed in both Auto and Manual modes. Fax (269-488-8691) or e-mail (ProductFieldAction@stryker.com ) the enclosedacknowledgment to confirm receipt of this notification. Please forward a copy of this letter to the new users and advise us of their new location and let Stryker know if any units have been discontinued. Questions or concerns, should be addressed to Stryker Customer Service at (1-800-327-0770). Monday-Friday 8 a.m-5 p.m. (ET). |  | 1002 | 1002 units | nationwide and worldwide | Terminated | November 04, 2014 | 312 | 1002 | Computer | Computer
Z-0937-2014 | 67237 | No | GFO | Partial thromboplastin time tests. | Hematology | Hematology | 510(k) | Hemochron Jr. Citrate APTT Cuvette |  |  | December 19, 2013 | 2013 | ITC-Nexus Dx | ITC has determined that some Citrate APTT cuvettes may recover higher than expected results in normal individuals. | Device Design | The firm, ITC The Point of Care, sent an "URGENT MEDICAL DEVICE RECALL" letter dated December 19, 2013 a long with an attached Health Professional Account Tracking Form via UPS to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to forward this communication to all those within your organization who need to be aware of this matter including any facility that you forwarded any of the affected product to; check all inventory to determine if you have any of the affected lots listed; if you have any affected lots, stop using them, remove from inventory and return to ITC; complete the attached Health Professional Account Tracking Form and return it within 10 days via fax to: 866-429-3132, email: RA@itcmed.com or mail to: ITC Technical Support, 20 Corporate Place South - Piscataway, NJ 08854 . It is important that you complete the form whether or not you have remaining inventory of these lots. In addition, contact ITC Technical Support immediately if you become aware of an unreported adverse event related to the use of affected J103C cuvettes.ITC will review your response form and will contact you to arrange for anyinventory to be returned to ITC and ITC will replace all unused, returned product.If you have questions, please contact ITC Technical Support: Telephone 800-579-2255 or 732-548-5700, Extension 4011 or E-mail Techsupport@itcmed.com. |  | 25965 | 25,965 curvettes (7,425 cuvettes US; 18,540 cuvettes OUS) | Worldwide Distribution: US (nationwide) and countries of: Austria, Belgium, Canada, Serbia, Germany, Estonia, Italy, Spain and Saudia Arabia. | Terminated | October 30, 2014 | 315 | 25965 | Computer | Not_Computer
Z-2243-2014 | 68914 | No | MMX | Embolectomy catheter. | Cardiovascular | Cardiovascular | 510(k) | CloverSnare 4-Loop Vascular Retrieval Snare. |  |  | July 17, 2014 | 2014 | Cook Inc. | This remedial action is the result of six product complaints associated with separation of the snare from the distal tip of the wire. | Device Design | Cook Medical sent an Urgent Medical Device Recall letter dated July 17, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the attached list of affected products and lot numbers shipped to their account, and quarantine any affected product that remains unused. Immediately collect and return all unused affected products to Cook Medical as soon as possible for credit. Complete the attached Recalled Product Reply Form and return to Cook Medical either with the product or separately. Customers were asked to report any Adverse Event to Cook Medical Customer Relations at 800-457-4500 or 1-812-339-2235. |  | 696 | 696 devices total distribution (671 nationwide) | Worldwide Distribution - USA including AK, AL, AZ, CA, CT, DC, FL, GA, HI, IA, ID, IL, IN, KY, LA, MD, ME, MI, MN, MO, MS, NC, NE, NH, NJ, NV, NY, OH, OK, PA, SC, TX, UT, VA, WA, WI, WV and Internationally to:Canada, Austria, Belgium, Denmark, Germany, Great Britain, Ireland, Italy, Spain, Sweden, and Switzerland. | Terminated | June 16, 2015 | 334 | 696 | Computer | Not_Computer
Z-0940-2014 | 67027 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens ADVIA Chemistry Systems models 1200, 1650, 1800 and 2400 |  |  | November 06, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | The ADVIA Chemistry software systems using certain software versions allows the system to automatically run assay calibrations and quality control (QC). | Software design | Siemens sent an Urgent Medical Device Correct letter, dated November 6, 2013, to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to retain the recall notification letter with their laboratory records and forward to those who may have received the affected product. |  | 3759 | 3759 | Worldwide Distribution and USA Nationwide. | Terminated | August 10, 2015 | 642 | 3759 | Computer | Computer
Z-2553-2014 | 68854 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens Dimension EXL Integrated Chemistry System in vitro diagnostic devices intended to measure a variety of analytes in human body fluids. All Serial #s of material #s: 10636929, 10636928, 10486890, 10472176, 10639337, and 10639338. |  |  | July 16, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc. | The Reagent Lid hinge may lose its effectiveness and slowly shift downward during maintenance procedures, potentially resulting in physical contact and injury to the operator. | Device Design | Siemens notified end users of a potential safety issue with the Reagent Lid hinge on the Dimension EXL instrument via an Urgent Medical Device Correction letter, dated July 2014. The letter indicated Customer Service Engineers will inspect and adjust the tension on the reagent lid hinges during every visit and replace hinges which no longer properly adjust. Users were urged to use caution when the Reagent Lid is open and notify Siemens the Reagent Lid is shifting down. |  | 2757 | 2757 | Nationwide and internationally, including Canada and Mexico | Terminated | December 16, 2014 | 153 | 2757 | Computer | Not_Computer
Z-2166-2014 | 68670 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | PerkinElmer Specimen Gate Screening Center software, version 1.4 and 1.6; Product Code: 5002-0500 |  |  | June 27, 2014 | 2014 | PerkinElmer Health Sciences, Inc. | Potential for demographic information and test result to be mismatch. | Software design | PerkinElmer sent an Urgent Medical Device Recall letter on July 1, 2014, to all affected customers. The letter informed customers of the recall and provided instructions. Customers were asked to complete the enclosed response form and fax it to 1-330-825-8520. For additional information, customers were instructed to contact their local PerkinElmer representative or email specimen.gate.support@perkinelmer.com.Customers with questions were instructed to call 1-800-321-9632. option 1. |  | 13 | 13 devices | Worldwide Distribution - USA including MA, GA, FL, NV, and Internationally to Canada, Ireland, Italy, Saudi Arabia, Belgium, and UK | Terminated | April 06, 2015 | 283 | 13 | Computer | Not_Computer
Z-1061-2014 | 67230 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Synergy XVI |  |  | November 25, 2013 | 2013 | Elekta, Inc. | Potential collision risk when using XVI and external beam shaping devices. | Software design | Elekta sent an Important Field Safety Notice dated November 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers. Corrective Action #1 - Potential collision risk when using XVI and external beam shaping devices. (Only software version R4.2.1 from this notice are applicable to this recall). Recommendations to Users - Monitor the Actual column of the Table Move Assistant dialog box. If the Actual is blank, do not press the ASU button on the FKP. To correct the error - (1) Restart XVI. (2) Put the patient in their initial setup position. (3) Scan the patient again, and continue the usual workflow. Always monitor the patient when do RATM. When you do Volume View acquisitions with the Apex collimator attached, make sure that the stop angle of the Volume View parameters is in the lower half of the gantry rotation. This decreases the risk of a collision with the patient. Read and follow the advice in the Important Field Safety Notice (200 01 507 081) and acknowledge receipt to your Elekta representative within 30 days. Corrective Action #2 -Permanent Solution - XVI Risk control measure - changes to the software to prevent table movement if the start position has not been recorded correctly. Recommendations for changes to product are - (1) Table zero should default to NIHII not zero. (2) Check that the table zero is valid because calculating absolute table positions. (3) Check prescribed table position is not more than 2.5cm from actual table position when permitting ASU from outside room. Target Release XVI R.4.2.2 estimated for release April 2014.For further questions please call (770) 300-9725. |  | 312 | 312 | Worldwide Distribution: US (nationwide) including Puerto Rico and the states of: AL, AK, AZ, CA, CO, CT, FL, GA, IL, IA, KY, MD, MA, MI, MN, MS, MT, NE, NV, NJ, NM, NY, NC, OH, OK, OR, PA, SC, TN, TX, VT, VA, WA, WV, WI, DC, and Internationally to: Algeria, Australia, Austria, Belarus, Belgium, Brazil, Canada, Chile, China, Cyprus, Czech Republic, France, Germany, Greece, Hong Kong, India, Israel, Italy, Japan, Malaysia, Mexico, Pakistan, Portugal, Russia, Singapore, Slovakia, South Korea, Spain, Taiwan, Thailand, and United Kingdom. | Open, Classified |  |  | 312 | Computer | Not_Computer
Z-2147-2014 | 68700 | Yes | LWS |  |  | Cardiovascular | PMA | Boston Scientific ENDOTAK RELIANCEÂż SG, transvenous defibrillation lead, Models 0180 &amp; 0292, Sterile EO. |  |  | June 19, 2014 | 2014 | Boston Scientific CRM Corp | A review of manufacturing test records for Boston Scientific ENDOTAK RELIANCE implantable leads revealed a suspected test data recording error. Specifically, some test results were recorded as "failed" without any other indication of failure. There are no reported injuries from the devices. | Software Manufacturing/Software Deployment | Sales representatives hand delivered a Boston Scientific "Medical Device Retrieval" Letter dated June 19, 2014 or June 20, 2014 when they were retrieving the devices from the hospital shelves. The letter was addressed to Hospital Administrator. The letter described the problem, product being recalled and the retrieval of the device by Boston scientific sales representative. For further questions they can contact their local sales representative or Boston Scientific International Technical Services.Physicians were contacted via telephone starting on July 6, 2014. A Boston Scientific letter dated July 8, 2014 followed the telephone conversation. For further information physicians can contact their local Boston Scientific representative or Technical Services at 1-800-227-3422. |  | 21 | 21 | Worldwide Distribution -US (nationwide) in the states of OH, MN and country of France | Terminated | December 03, 2014 | 167 | 21 | Computer | Not_Computer
Z-1980-2014 | 68234 | No | HKI | Ophthalmic camera. | Ophthalmic | Ophthalmic | 510(k) | Non-Mydriatic Auto Fundus Camera. Model AFC-330. An ophthalmic camera for use in capturing images of the retina and anterior segment of the eye. |  |  | June 16, 2014 | 2014 | Nidek Inc | Image taken by AFC-330 has a white spot which may affect diagnosis or evaluation of image. | Device Design | Field Correction notifications were sent to all affected users on June 16, 2014 by trackable method. |  | 185 | 185 cameras | Nationwide Distribution. | Terminated | March 03, 2015 | 260 | 185 | Computer | Not_Computer
Z-2029-2014 | 68024 | No | KFM | Nonroller-type blood pump. | Cardiovascular | Cardiovascular | 510(k) | Sarns Disposable Centrifugal Pump with X-Coating, 164275X, is a sterile, single use device. |  |  | April 14, 2014 | 2014 | Terumo Cardiovascular Systems Corporation | During set-up and priming of the bypass circuit, leaks were detected on the centrifugal pumphead. | Device Design | Urgent Medical Device Recall letters, dated April 11, 2014, were sent to customers. The letters identified the affected product, as well as the reason for the correction, the potential hazard, the correction, the affected population, and instructions for customers. Customers are to review the notice and ensure that all users have received a copy. Additionally, customers are to confirm receipt of the communication by returning the attached Customer Response Form as instructed; and, return all effected centrifugal pumps. Terumo CVS will issue a Returned Goods Authorization upon receipt of the Response Form. Questions should be directed to Terumo CVS Customer Service at 800-521-2818. |  |  |  | Nationwide Distribution | Open, Classified |  |  |  | Computer | Not_Computer
Z-2037-2014 | 68643 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | GEMINI TF PET/CT16: 882470, GEMINI TF PET/CT 64: 882471, GEMINI TF Big Bore: 882476. A diagnostic imaging system for fixed or mobile installations that combines Positron Emission Tomography (PET) and X-ray Computed Tomography (CT) systems. |  |  | June 18, 2014 | 2014 | Philips Medical Systems (Cleveland) Inc | When beginning a procedure to start the gantry and CT Host, a Close Estop message may appear to the operator, but disappears before action can be taken. If this occurs it results in an inability to close E-Stop and the system is not operational for clinical use. This problem occurs randomly. | Software design | Philips Healthcare is informing the consignee of this issue through a Field Safety Notice (FSN) 88200487-488. The US consignees will receive this notice via certified mail and Philips Key Market representatives will distribute the FSN internationally. Additionally, a Field Service Engineer will contact each site to schedule a time to install the software update. |  | 33 | 33 | Distributed in the states of CA, CO, IL, OK, OR and WI. | Terminated | September 16, 2015 | 455 | 33 | Computer | Computer
Z-2036-2014 | 68696 | No | MEC | Medical washer-disinfector. | General Hospital | General Hospital | 510(K) Exempt | Reliance 1227 Cart and Utensil Washer/Disinfector |  |  | June 30, 2014 | 2014 | Steris Corporation | STERIS has identified that deformation of the washers door handle pin may occur when the door handle is rotated with sufficient force. Should deformation and/or breakage of the door handle pin occur there is a potential for the washer door to become inoperable. | Device Design | Customers will be notified via onsite visits by STERIS Field Service Technicians to each affected Customer site. These visits will be documented and monitored until completion. |  | 37 | 37 | US Nationwide Distribution - in the states of AL, AR, CA, DE, FL, GA, IL, IA, LA, MA, MI, NJ, NY, OH, OK, OR, SC, TX, VA. and in the country of Canada. | Terminated | July 01, 2015 | 366 | 37 | Computer | Not_Computer
Z-1110-2014 | 67427 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart MRx Monitor/Defibrillator used |  |  | February 04, 2014 | 2014 | Philips Medical Systems, Inc. | HeartStart MRx Defibrillator Monitor therapy connection could experience accelerated wear causing risk for patients and/or caregivers | Device Design | Philips Medical North America issued on 2/4/14 the Urgent Medical Device Correction Notification/Field Safety Notice informs customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue.The correction will consist of free of charge service support for the MRx therapy connection, including inspection of therapy connection, and installation of a stabilizing collar if not already present. If wear is detected, therapy components (including therapy pads/cables, CPR therapy cable, external paddles cable, and internal paddles cable/adapter) within their service life will be replaced.A Philips Healthcare representative will contact customers with affected devices to arrange for service. Philips is asking customers to follow the Action to be Taken by Customer/User section of the Urgent Medical Device Correction notification/Field Safety Notice:You can continue to use your MRx prior to inspection and installation of upgrade hardware by a Philips Authorized Service Provider provided that upon receipt of this notification: Ensure you are performing ongoing therapy connection inspections on all of your MRx devices to detect wear. The HeartStart MRx Instructions for Use, 453564396411 Ed 1, 453564307761 Ed 2, or 453564174011, a previously circulated Therapy Connection Maintenance and Inspection When Used in Transport Addendum, Edition 6, describe how users can identify wear. If wear is detected, remove affected devices immediately from use and contact local Philips representative to arrange service or call 1-800-722-9377. |  | 58466 | 58466 units | Worldwide Distribution: USA (nationwide) and to countries of: AFGHANISTAN ALBANIA ARGENTINA ARUBA AUSTRALIA AUSTRIA AZERBAIJAN BAHRAIN BANGLADESH BELGIUM BOLIVIA BOSNIA AND HERZEGOWINA BRAZIL BRUNEI DARUSSALAM BULGARIA CAMBODIA CHILE CHINA COLOMBIA COSTA RICA CROATIA (local name: Hrvatska) CYPRUS CZECH REPUBLIC DENMARK ECUADOR EGYPT EL SALVADOR ESTONIA ETHIOPIA FINLAND France GERMANY GHANA GREECE GUADELOUPE HONG KONG ICELAND INDIA INDONESIA IRAQ IRELAND ISRAEL ITALY JAPAN JORDAN KAZAKHSTAN KENYA KOREA, REPUBLIC OF KUWAIT LAOSÂż LATVIA LEBANON LIBYAN ARAB JAMAHIRIYA LITHUANIA LUXEMBOURG MACEDONIA MALAYSIA MEXICO MONGOLIA MOROCCO MOZAMBIQUE MYANMAR (Burma)Âż NAMIBIA NEPAL NETHERLANDS NETHERLANDS ANTILLES NEW CALEDONIA NEW ZEALAND NICARAGUA NORWAY OMAN PAKISTAN PALESTINE PANAMA PAPUA NEW GUINEA PARAGUAY PERU PHILIPPINES POLAND PORTUGAL PUERTO RICO QATAR REUNION ROMANIA RUSSIAN FEDERATION SAUDI ARABIA SERBIA SINGAPORE SLOVENIA SOUTH AFRICA SPAIN SRI LANKA SWEDEN SWITZERLAND TAIWAN, PROVINCE OF CHINA TANZANIA, UNITED REPUBLIC OF THAILAND TRINIDAD AND TOBAGO TUNISIA TURKEY TURKMENISTAN UGANDA UKRAINE UNITED ARAB EMIRATES UNITED KINGDOM UZBEKISTAN VENEZUELA VIET NAM YEMEN | Open, Classified |  |  | 58466 | Computer | Computer
Z-1120-2014 | 67438 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | HeartStart MRx Defibrillator/Monitor M3538A Lithium Ion Batteries |  |  | February 18, 2014 | 2014 | Philips Medical Systems, Inc. | Philips HeartStart MRx Monitor/Defibrillator, when operating on battery power only, may experience an unexpected shutdown if exposed to elevated levels of electromagnetic interference from RF energy source | Device Design | The firm, Philips Healthcare, sent an "URGENT- Medical Device Correction" letter dated February 2014 and issued a Field Safety Notice on February 18/ 2014 to its customers. The letter and field safety notice informs customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue. Philips will replace the affected batteries free of charge.Philips instructed customers to follow the Action to be Taken by Customer/User section of the Field Safety Notice. Fixed, portable, and mobile radio frequency communications equipment can affect the performance of medical equipment. See the HeartStart MRx Instructions for Use for the minimum recommended separation distance between RF communications equipment and the HeartStart MRx. The condition can be avoided by operating the device on AC or external DC power or by the following actions when operating on battery.If you need any further information or support concerning this issue, please contact your local Philips representative or call us at 1-800-722-9377. |  | 43792 | 43,792 batteries | Worldwide Distribution: US (nationwide) and Internationally to: ALBANIA, ARGENTINA, ARUBA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BANGLADESH, BELGUIM, BOLIVIA, BRAZIL, CANADA, CHINA, COLOMBIA, CONGO, CROATIA, CZECH REPUBLIC, DENMARK, EGYPT, EL SALVADOR, ESTONIA, ETHIOPIA, FINLAND, FRANCE, GERMANY, HONG KONG, ICELAND, INDIA, INDONESIA, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KAZAKHSTAN, KENYA, KOREA REPUBLIC, KUWAIT, LATIVA, LEBANON, LIBYAN ARAB, LITHUANIA, MALAYSIA, MALTA, MEXICO, MONGOLIA, MOROCCO, MOZAMBIQUE, MYANMAR, NEPAL, NETHERLANDS, NEW CALEDONIA, NEW ZEALAND, NICARAGUA, NIGERIA, NORWAY, OMAN, PAKISTAN, PARAGUAY, PERU, PHILIPPINES, POLAND, PORTUGAL, QATAR, ROMANIA, RUSSIAN FEDERAL, SAUDIA ARABIA, SERBIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN PROVIDENCE, THAILAND, TRINIDAND AND TABAGO, TUNISIA, TURKEY, TURKMENISTANIA, UKRAINE, UNITED ARAB EMIRATES, UNITED KINGDOM, UZBEKISTAN, VENEZUELA, and VIETNAM. | Open, Classified |  |  | 43792 | Computer | Computer
Z-2023-2014 | 68399 | No | INN | Mechanical chair. | Physical Medicine | Physical Medicine | 510(K) Exempt | Aquatec Fixed Backrest - accessory for Mobile Shower and Toilet commode Aquatec Ocean, Ocean VIP and Ocean E-VIP family. |  |  | June 25, 2014 | 2014 | Invacare Corporation | The AquatecÂż Fixed Backrest withstands a load of only 250 newtons at uneven load. This may result in the Fixed Backrest becoming detached from the frame. | Device Design | The recall notification will be sent to the two consignees informing them of the recall and to notify their customers. The letter will include a provider letter, customer letter, provider response card and impacted units listing. |  | 465 | 465 units | Distributed in PA and Canada. | Terminated | April 15, 2015 | 294 | 465 | Computer | Not_Computer
Z-1093-2014 | 67380 | No | NYQ |  |  | Immunology | PMA | Product Name: HER2 CISH pharmDx Kit |  |  | January 21, 2014 | 2014 | Dako North America Inc. | Dako has initiated a recall of HER2 CISH pharmDxTM Kit because weak red signals have been observed. | Device Design | Dako sent a Recall Notification letter dated January 21, 2014to affected customers via mail. The letter identified the affected product, description of the problem, actions to be taken by the user, communication, and Dako contact information. The letter was accompanied by a "Recall Form (Product Recall Form to customers)". For questions call 805-566-5464. |  | 49 | 49 | Worldwide Distribution - USA Nationwide and in the country of Chile. | Terminated | June 04, 2014 | 134 | 49 | Computer | Not_Computer
Z-2026-2014 | 68678 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | The STERIS 4085 General Surgical Table |  |  | June 27, 2014 | 2014 | Steris Corporation | STERIS engineering analysis has determined the D1 pressure switch in the hydraulic column manifold does not adequately detect pressure level. If one or more of the table floor lock feet become unlocked due to the loss of hydraulic pressure, there is a potential for the table hand control to incorrectly indicate that the floor lock feet are locked. | Device Design | STERIS sent an Urgent Voluntary Field Correction Noticel letter dated June 27, 2014 to affected Customers.via FedEx with tracking numbers for delivery confirmation. The letter identified the description of the product, description of the problem and the actions to be taken. For questions contact Steris Field Service Dispatch at 1-800-333-8828. |  | 3689 | 3689 | Worldwide Distribution - US (nationwide in the states of AL, AK, AZ, AR, CA,CO, CT, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, WY in the countries of Canada, India, Vietnam, China, India, Korea, Mexico, costa Rica, Brazil, Taiwan, Australia, New Zealand, El Salvador, Indonesia, Hong Kong, Singapore, Nepal. | Terminated | September 16, 2015 | 446 | 3689 | Computer | Not_Computer
Z-1096-2014 | 67424 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Invivo Corporation Expression IP5 Information Portal |  |  | October 03, 2013 | 2013 | Invivo Corporation | Invivo Corporation located at 12151 Research Parkway, Orlando, FL 32826 is recalling their Expression Information Portal (Model IP5). Three different IP5 devices lost audio during use. | Software design | Philips Healthcare sent an "URGENT - MEDICAL DEVICE RECALL" letter dated October 3, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers were instructed to refer to the "Action to be Taken by Customer/User section on the attached pages. A Philips representative will contact the customer regarding their affected devices. All affected devices will have a software update installed in order to correct the problem free of charge. Contact your local Philips representative for further information and support. |  | 271 | 271 | Worldwide Distribution - USA and the countries of Australia, Austria, Denmark, France, Germany, Netherlands, Poland, Singapore, Sweden, Switzerland, United Kingdom. | Open, Classified |  |  | 271 | Computer | Computer
Z-1995-2014 | 68580 | No | LDS | Medical washer. | General Hospital | General Hospital | 510(K) Exempt | HLD System, Model 610, Medical Device Cleaning and High Level Disinfection Washer/Pasteurizer. |  |  | January 24, 2013 | 2013 | Cenorin, LLC | The temperature sensor/control system in the HLD Systems Model 610 may provide a false temperature reading. This caused the water to remain at ambient temperature throughout the pasteurization cycle rather than using hot water. | Device Design | Cenorin sent a Notification letter dated December 18, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter stated: Cenorin has developed an upgrade to mitigate this issue as well as provide additional process control and system monitoring benefits. Cenorin included a Return Response form with the recall letter to be completed and return. For questions contact HLD Systems Sales Representative for questions. |  | 25 | 25 HLD systems | Worldwide Distribution - US (nationwide): CA, DE, FL, MA, MI, NY, PA, TX, UT, WA, and WI and the countries of Uruguay and Taiwan. | Terminated | February 10, 2015 | 747 | 25 | Computer | Not_Computer
Z-1994-2014 | 68593 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Patient Side Cart Battery Box used in conjunction with the da Vinci A surgical system. |  |  | April 27, 2012 | 2012 | Intuitive Surgical, Inc. | Patient Side Cart battery boxes may overheat in rare charging conditions, causing battery bulging. | Device Design | No notification was sent. Issue is being addressed during routine maintenance or through complaint process. |  | 2606 | 2606 | Worldwide Distribution. | Terminated | July 07, 2014 | 801 | 2606 | Computer | Not_Computer
Z-1113-2014 | 66527 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | TSX-101A Aquilion 32/64 ONE Computed Tomography X-Ray System; TXS-101/R Aquilion RXL Computed Tomography X-Ray System; TSX-301A Aquilion One Computed Tomography X-Ray System; TSX-301B Aquilion Premium Computed Tomography X-Ray System; TSX-302A Aquilion Prime Computed Tomography X-Ray System. |  |  | October 01, 2013 | 2013 | Toshiba American Medical Systems Inc | Toshiba intitiated this recall because their investigation revealed that due to a problem of the X-ray output control software, the X-ray output conditions may not be set correctly. As a result, X-ray exposure under the scan conditions used for contrast study may not start. | Software design | Toshiba sent an Urgent Medical Device Correction letter on October 1, 2013, via USPS return receipt mail and accompanied by a response form. The recall notification letter provided the Device Model, serial number, SID, problem, corrective action, request to customers, and contact information (800) 421-1968 or (800) 521-1968. Toshiba designed and test an upgraded ROM for the SC2 PWB to prevent recurrence of the fault. Service technicians will replace the ROM on site at no cost to the system owners. All repairs are expected to be completed by April 2014 and progress reports will be provided on the effectiveness of the plan.For questions regarding this recall call 714-730-5000. |  | 338 | 338 units | Nationwide Distribution. | Terminated | March 03, 2015 | 518 | 338 | Computer | Computer
Z-1124-2014 | 67458 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | MEDITECH Microbiology - Calculator/data processing module, for clinical use. |  |  | January 17, 2014 | 2014 | Medical Information Technology, Inc. | Potential for patient results being removed from LAB/EMR. | Device Design | Beginning on January 17, 2014 MEDITECH distributed a notification to MAGIC and Client Server customers who have the Microbiology product in TEST and/or LIVE only for the releases noted in this correction report. This notification was made via e-mailed task updates that can be printed by the customer. Task updates sent via e-mail are immediately transmitted to the customer.Instructions provided to customers: Task Description: CS LAB DTS 11696/MG LAB DTS 16703 - MIC: Inactive procedure prompt not displaying previous results. We are setting up this task to inform you of a potential patient safety issue that may result in patient results being removed from LAB/EMR. This task will continue to be updated to provide you with the status on this issue and details on the deployment of changes to address. This issue has been determined to impact both Client Server and MAGIC 5.6 and higher. Until the changes to address above are delivered, please refrain from inactivating procedure prompts in the MIC Procedure Dictionary. Although patient results are not permanently deleted from the record, because of the potential impact on a patient should this issue occur and go undetected, we felt it necessary to inform you of the issue. We strongly recommend delivery of the changesto address the issue.Âż We have engaged all necessary resources within MEDITECH to identify and resolve this issue as expediently as possible.Âż CS LAB DTS 11696 - MIC: Inactive procedure prompt not displaying previous resultsÂż MG LAB DTS 16703 - MIC: Inactive procedure prompt not displaying previous resultsÂż Inactivating Procedure Prompts in the MIC Procedure Dictionary causes data to be removed from existing specimens. These results are also removed from the EMR. However, once the prompts are re-activated, the results are again present on the specimens.Âż We strongly recommend delivery of this change to your system to alleviate this potential patient safety concern. Please confirm via this task th |  | 737 | 737 | Worldwide Distribution: US (nationwide) and to countries of: Canada, Ireland, and United Kingdom. | Open, Classified |  |  | 737 | Computer | Not_Computer
Z-1960-2014 | 65562 | No | LPS | Urinary homocystine (nonquantitative) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA Centaur HCY Diluent - 10mL Bottle Catalog number 09877493 (124533), SMN 10310370, and Ancillary Pack Catalog number 03302138 (124581), SMN 10378677. Used on the ADVIA Centaur, ADVIA Centaur XP and ADVIA Centaur CP platforms. For in vitro diagnostic use in the quantitative determination of total homocysteine (HCY) in serum or plasma. |  |  | June 25, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc | ADVIA Centaur Homocysteine dilution recovery on the diluent identified the percent recovery for the 1:10 patient sample dilutions lower than reported in the Instructions for Use (IFU). | Device Design | Siemens Healthcare sent an Urgent Field Safety Notice that was sent to all affected Siemens Healthcare Diagnostics customers in the United States on June 25, 2013. The Urgent Field Safety Notice was sent out for distribution to all customers outside the United States on June 24, 2013. These notices inform the customer that the 1:10 diluted patient recovery is lower than reported in the Instructions for Use (IFU). Customers should discontinue use of both manual and on board 1:10 dilutions of samples with the ADVIA Centaur HCY assay. Customers can continue to use a manual or on board 1 :2 dilution for samples with results outside of the assay range. Siemens has confirmed theperformance at the 1 :2 dilution level meets the IFU ranges.Please retain this letter with your laboratory records, and forward this letter to those who may have received this product.We apologize for the inconvenience this situation has caused. If you have any questions, please contact your Siemens Technical Solutions Center or your local Siemens technical support representative.For further questions please call (508) 668-5000. |  | 51 | 51 units | Worldwide Distribution - USA (nationwide) and the countries of : Austria, Belgium, Bulgaria, Croatia, Cyprus, Czech Republic, Denmark, Finland, France, Germany, Great Britain, Greece, Hungary, Ireland, Italy Kazakhstan, Latvia, Liechtenstein, Luxembourg, Netherlands, Norway, Poland, Portugal, Romania, Saudi Arabia, Serbia, Slovakia, Spain, Sweden, Switzerland ,Turkey, Unit.Arab Emir and Vatican | Open, Classified |  |  | 51 | Computer | Not_Computer
Z-2146-2014 | 68741 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ExacTrac is intended to be used to place patients at an accurately defined point within the treatment beam of a medical accelerator for stereotactic radiosurgery or radiotherapy procedures, in order to treat lesions, tumors and conditions anywhere in the body when radiation treatment is indicated. ExacTrac may also be used to monitor the patient position during the treatment. |  |  | June 24, 2014 | 2014 | Brainlab AG | When using multiple isocenters (radiation treatment targets) within a single plan, in certain workflow conditions ExacTrac v.6.0.4 might move the patient to an unintended isocenter position, despite displaying the green "OK" icon. If this anomaly occurs and is not detected by the user, the radiation treatment dose at the linear accelerator may be delivered to the unintended target position. | Software design | Brainlab sent a FIELD SAFETY NOTICE / PRODUCT NOTIFICATION letter dated June 24, 2014, to all affected customers. The letter included instructions for customers to: 1) Avoid using plans with multiple isocenters in ExacTrac 6.0.4 whenever possible, and 2) If plans with multiple isocenters cannot be avoided and such a plan is used in ExacTrac v.6.0.4, then always select "Yes" in the dialog displayed to "keep the previous correction shift" when switching between isocenters and always verify all final treatment positions. Brainlab will provide a software update with this issue solved to affected customers. Brainlab will actively contact you to schedule the update within the next 6 months. Customers with questions can contact Brainlab at 800-597-5911 or via e-mail at us.support@brainlab.com. |  | 15 | 15 systems | Worldwide Distribution - USA including CA, CT, and IL; Internationally to France, Germany, Japan, Netherlands, Russia, and Slovenia. | Open, Classified |  |  | 15 | Computer | Computer
Z-1989-2014 | 68528 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Revolution XR/d. |  |  | June 04, 2014 | 2014 | GE Healthcare, LLC | GE Healthcare has become aware of a potential safety issue due to a collimator installation error during a service maintenance activity involving the Proteus XR/a and Revolution XR/d X-ray imaging systems. . There was a reported incident of a collimator fall which caused a serious patient injury. | Equipment maintenance | GE Healthcare sent an Urgent Medical Device Correction letter dated June 4, 2014, to all affected consignees. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. For questions in regard to this recall the consignees in the United States were advised to call 800-437-1171. For other countries contact your local GE Healthcare Service Representative.For questions regarding this recall call 800-437-1171. |  | 1053 | 1053 | Worldwide Distribution - USA including PR. DC, GUAM, all states except VT. Internationally to the following countries:ALBANIA, ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA, BAHRAIN, BENIN, BOSNIA AND HERZEGOVINA, BOTSWANA, BRAZIL, BULGARIA, BURKINA FASO, CANADA, CHILE, CHINA, COLOMBIA, CZECH REPUBLIC, DOMINICAN REPUBLIC, ECUADOR, EGYPT, ESTONIA, FINLAND, FRANCE, FRENCH POLYNESIA, GERMANY, GREECE, GUATEMALA, HONDURAS, HONG KONG, HUNGRY, ICELAND, INDONESIA, IRAQ, IRELAND, ISRAEL, ITALY, JAMAICA, JAPAN, JORDAN, KAZAKHSTAN, REPUBLIC OF KOREA, KUWAIT, LATVIA, LEBANON, LIBYAN ARAB JAMHIRIYA, LITHUANIA, MALAYSIA, MALTA, MARTINIQUE, MEXICO, MOROCCO, NETHERLANDS, NEW ZEALAND, NICARAGUA, NIGERIA, NORWAY, PAKISTAN, PANAMA, PHILIPPINES, POLAND, PORTUGAL, ROMANIA, RUSSIA, SAUDI ARABIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SUDAN, SWEDEN, SWITZERLAND, TAIWAN, TANZANIA, THAILAND, TUNISIA, TURKEY, UGANDA, UKRAINE, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY, VENEZUELA, VIETNAM YEMEN. | Terminated | May 06, 2015 | 336 | 4513 | Computer | Not_Computer
Z-1979-2014 | 68459 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | CNS-6200 Series Central Nurse Station and accessories. Model Number: CNS-6201 |  |  | June 03, 2014 | 2014 | Nihon Kohden America Inc | Software Version 02-26, when used with the Central Nursing Station (CNS) 6201, (PU-621 RA) may unexpectedly and without warning reboot, resulting in a period of approximately 3 minutes during which time the system does not display patient data and does not emit alarms based on that data should they be appropriate. | Software design | Nihon Kohden sent an Urgent Medical Device Safety Notification and Field Correction letter dated June 3, 2014 to affected customers. Recall notification letter, dated June 3, 2014 was hand-delivered to site by Nihon personnel. The letter provided the following information: product description with codes and software version, reason for recall, field correction information and actions to be taken, MedWatch reporting information, and contact information for recall. |  | 430 | 430 | Worldwide Distribution - US in the state of Oklahoma and in the countries of Bolivia, Chile, Germany, Japan, Korea, Taiwan, Yemen, andSouth Africa. | Terminated | April 28, 2015 | 329 | 430 | Computer | Computer
Z-1982-2014 | 68644 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | BrightView XCT: product code 882482, BrightView X- upgraded with the XCT Flat Panel Detector (FPD): product code 882454. Gamma camera for Single Photon Emission Computed Tomography (SPECT) Medical Device for imaging. |  |  | June 23, 2014 | 2014 | Philips Medical Systems (Cleveland) Inc | Flat Panel Display (FPD) failed to remain securely locked in the deployed position. | Device Design | Philips sent an Field Safety Notice dated June 17, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.This Field Safety Notice (FSN) 88200490 and Addendum is intended to inform you about the following:" what the problem is and under what circumstances it can occur" the actions that should be taken by the customer / user in order to prevent risks for patients or users" the actions planned by Philips to correct the problem.If you need any further information or support concerning this issue, please contact your local Philips Healthcare representative:For North America and Canada contact the Customer Care Solutions Center (1-800-722-9377: follow the prompts).thcare is informing the consignee of this issue through a Field Safety Notice 88200490. The US consignees will receive this notice via certified mail and Philips Healthcare Key Market representatives will distribute the FSN internationally. Additionally, a Field Service Engineer will contact each site to schedule a time to conduct the appropriate field safety correction through Field Change Order (FCO) 88200490 that will address the above identified issue. |  | 407 | 407 | Worldwide Distribution - USA including the states of : AZ, CA, CO, DC, FL, GA, IL, MA, MD, MN, MO, MS, MT, NC, NJ, NY,OH, OK, OR, PA, TX, VT, WA,WI. Government Facilities: AZ, CA, CO, FL, HI, IA, KY, MA, MN, MO, NC, ND, NE, NV, OH, OK, SD, TX, and WV. and the countries of : Algeria, Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, China, Czech Republic, Denmark, Finland, France, Germany, Greece, Indonesia, Ireland, Israel, Italy, Japan, Lebanon, Malaysia, Mexico, Netherlands, New Zealand, Norway, Panama, Poland, Portugal, Russia, Saudi Arabia, Slovakia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Turkey, UAE and UK. | Terminated | September 24, 2015 | 458 | 407 | Computer | Not_Computer
Z-1966-2014 | 68472 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Vision Troubleshooting Guide for the da Vinci Standard Surgical System; da Vinci S System and da Vinci Si System. |  |  | May 21, 2014 | 2014 | Intuitive Surgical, Inc. | Multiple updates to User Manuals, Instructions for Use (IFU), Quick Reference Guides (QRG), and Instrument Release Kits (IRK) for use with the da Vinci Surgical System, instruments and accessories. Removal of the Emergency Grip Release Wrench, part number 710142. | Device Design | Intuitive Surgical sent a Field Safety Notice dated June 16, 2014, to all consignees affected by changes in labeling, IFU User manuals and QRG's. Consignees were instructed to take the following actions:Ensure all appropriate personnel are fully informed of this Notification. Forward this letter to your Risk Manager, OR Director, Purchasing Manager, Biomedical Engineering staff and members of your medical staff who perform da Vinci surgery procedures.2. Use Attachment C to identify the labeling that has been updated.3. No manuals or products should be returned. Discard old versions of the User Manuals, IFUs, and QRGs for which you have received updated versions.4. Discard all versions of the Emergency Grip Release Wrench (Part Number 710142). Order additional IRKs from Intuitive Surgical as needed.5. Complete and return the attached Acknowledgement Form to acknowledge that you have received the updated Labeling, IRKs, in-service and have discarded all previous versions of Labeling and the emergency grip release wrench. Please return the completed Acknowledgment Form within two days of the in-service.6. Retain a copy of this Notice and the Acknowledgement Form for your records.If you need further information or support concerning this recall, please contact your Intuitive Surgical Representative or Intuitive Surgical Customer Service at the numbers listed below:North America: 800-876-1310 Option 3 (6 am to 5pm PST)For questions regarding this recall call 408-523-2602. |  | 2014 | All distributed prior to June 2014 | Worldwide Distribution - USA (nationwide) and Internationally to Argentina, Australia, Austria, Belgium, Brazil, Bulgaria, Canada, Chile, China, Colombia, Cyprus, Czech Republic, Denmark, Dominican Republic, Egypt, Finland, France, Germany, Great Britain, Greece, Guadeloupe, India, Indonesia, Ireland, Israel, Italy, Japan, Kuwait, Lebanon, Malaysia, Mexico, Monaco, Netherlands, New Zealand, Norway, Pakistan, Philippines, Indonesia, Portugal, Puerto Rico, Qatar, Romania, Russia, Saudi Arabia, Singapore, Slovak Republic, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, Uruguay, Venezuela, and Vietnam. | Terminated | May 05, 2015 | 349 | 2014 | Computer | Not_Computer
Z-1962-2014 | 68529 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON S Family Ultrasound Systems at software versions VC25D, VC30A and VC30B; (S2000, S3000 and S1000 systems) Siemens Medical Solutions USA, Inc. Ultrasound imaging systems. |  |  | June 03, 2014 | 2014 | Siemens Medical Solutions USA, Inc. | There is a potential measurement error with the Auto-Stats (auto statistics) measurement tool on the ACUSON S Family ultrasound system. | Software design | Siemens sent a Customer Safety Advisory Notification letter to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to do the following: To avoid the potential risk of this issue, do not use the Auto-Stats tool; use alternate measurement tools. Only use the Velocity tool on frozen images.Customer Service Engineer will contact the customer to schedule an appointment to replace their existing version of software, free of charge. Customers with questions should contact their local service support person forinformation regarding timelines and status.Customers were instructed to share this information with all personnel withintheir organization who need to be aware of this issue.For questions regarding this recall call 650-694-5135. |  | 790 | 790 units | Worldwide Distribution - USA including CA, MO, VA, PA, WV, and Internationally to Angola, Argentina, Australia, Austria, Belarus, Belgium, Bolivia, Brazil, Canada, Chile, Colombia, Czech Republic, Denmark, Dominican Rep., Ecuador, Egypt, France, Germany, Greece, India, Iraq, Israel, Italy, Japan, Kazakhstan, Malaysia, Mexico, Morocco, Netherlands, Norway, P.R. China, Philippines, Poland, Portugal, Republic Korea, Romania, Russian Fed., Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, U.A.E., Ukraine, United Kingdom. | Open, Classified |  |  | 790 | Computer | Not_Computer
Z-1949-2014 | 68393 | No | BZD | Noncontinuous ventilator (IPPB). | Anesthesiology | Anesthesiology | 510(k) | DeVilbiss Model DV5x Series CPAP units. For use in treating Obstructive Sleep Apnea (OSA) by means of application of positive air pressure. Used in home and clinical environments. |  |  | May 13, 2014 | 2014 | DeVilbiss Healthcare LLC | A limited number of DV5x Series CPAP units were manufactured and shipped without the serial number programmed into the unit, causing the unit serial number on LCD display to appear as "RUN-TEST". | Software Manufacturing/Software Deployment | DeVilbiss Healthcare sent an Urgent: Medical Device Correction notice dated May 13, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers. What to do if you have an affected DV5x Series CPAP unit:1. Verify your inventory of DV5x Series CPAPs and isolate any units that are on the included serial number list, these units will need to have the serial number corrected.2. Please contact DeVilbiss Healthcare Customer Service to either request a Sales Representative to correct the issue on site or return the unit back to DeVilbiss to be corrected.3. If you have already distributed any DV5x Series CPAPs included in this serial number list to end users, a patient communication has been included with this notification - a. Insert your contact information at the bottom of the "User Medical Device Correction Letter" included in this notice.b. Forward a copy of the "User Medical Device Correction Letter'' to each user towhom you have provided an affected DV5x Series CPAP unit.If you have one or more of the affected DV5x Series CPAP unit in your possession (as identified in the serial number list) and have any questions regarding the information provided in this communication, please contact DeVilbiss Healthcare Customer Service at 1-800- 338-1988 from 8:00AM-5:00PM EST. [Note: If you have already spoken with a DeVilbiss Healthcare Representative about this issue, there is no need to call unless you have further questions.] |  | 1809 | 1809 | Worldwide Distribution - USA (nationwide) and the countries of Colombia, Honduras, Iran, Thailand and Uruguay. | Terminated | December 02, 2014 | 203 | 1809 | Computer | Not_Computer
Z-1958-2014 | 68347 | No | OCC | Respiratory viral panel multiplex nucleic acid assay. | Microbiology | Microbiology | 510(k) | Simplexa" Flu A/B &amp; RSV Direct, Model Number: MOL2650. |  |  | May 15, 2014 | 2014 | Focus Diagnostics Inc | Focus Diagnostics is providing an urgent safety notice for a correction to the labeling for Simplexa" Flu A/B &amp; RSV Direct (MOL2650). Focus Diagnostics received some customer complaints of Simplexa Flu A/B &amp; RSV Direct assays with sporadic false signals, which may result in a higher false results rate due to a potential software spectral matrix and Direct Amplification Disc storage. | Device Design | Focus Diagnostics sent an Urgent Safety Notice Correction letter dated May 15, 2014, to all affected customers. The notification, informed customers of the product, the problem, and the action to be taken by the customer. To decrease the likelihood of obtaining any false results caused by either manually entering a spectral matrix or storing the DAD improperly after each use, please follow the recommended warnings and precautions provided in the revised and attached Simplexa Flu A/B &amp; RSV Direct (PI.MOL2650 Rev. D) package insert.1. The spectral matrix must be installed in each 3M Integrated Cycler and should not be changed unless an updated QR code for the instrument is provided by Focus Diagnostics. The spectral matrix is unique to each 3M Integrated Cycler. The spectral matrix was provided with the 3M Integrated Cycler instrument on the cover of the 3M Integrated Cycler Hardware Manual. If the matrix label will not scan or cannot be found, contact Focus Diagnostics. The contact information is on the last page of this document. 2. Not installing or changing the spectral matrix can result in false results.3. After each use store DAD discs flat with the numbered foil side up.Customers were instructed to acknowledge that they have received this notification by signing the enclosed acknowledgement form and email the form to DxTS@focusdx.com or fax back to Focus Diagnostics Technical Services at 562-240-6526 within 10 business days. The Notice of Correction and Acknowledgement Form was sent to customers on 05/15/2014 via email and USPS mail. Additional customers were notified on 05/29/2014 via email and USPS mail.For questions regarding this recall call 714-220-1900. |  | 614 | 614 kits | Nationwide Distribution | Terminated | December 31, 2014 | 230 | 614 | Computer | Computer
Z-1963-2014 | 68506 | No | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | EXACTAMIX 1200 Compounder, model numbers 1200-DY, 1200-DX, 1200-DYR and 1200-DXR; EM1200 DY Display, Replacement EM 1200 DY Display. EXACTAMIX 2400 Compounder model numbers 2400-DY, 2400-DX, 2400-DYR and 2400-DYX. Baxa Corporation. An automated pumping system that compounds multiple sterile ingredients into a finished solution in a single patient IV administration bag. |  |  | June 10, 2014 | 2014 | Baxter Healthcare Corp. | If the universal ingredient (UI) in an active configuration is changed using the Configuration Editor, a flush of the outlet pump tube will not be initiated by the software. It could result in the original UI remaining in the tube which could be delivered into the next bag. | Software design | An Urgent Device Correction communication via 1st class mail was mailed on July 14, 2014, informing customers of a potential issue if the universal ingredient in an active configuration is changed using the Configuration Editor rather than following the normally prescribed and supported methods as outlined in the Operator Manual. |  | 1758 | 1758 | US Nationwide | Open, Classified |  |  | 1758 | Computer | Computer
Z-1060-2014 | 66287 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | McKesson Anesthesia Care |  |  | March 15, 2013 | 2013 | MCKESSON TECHNOLOGIES INC. | There was an occurrence where the patient case data did not match the patient data when the case was recalled in the Anesthesia Care Record (ACR) in that it included data from another case. | Software design | McKesson sent a Clinical Alert dated March 15, 2013 to all affected customers. Additionally, phone calls were placed to each customer, followed up by an email to provide them written copies of the communication and Clinical Alert and obtaining acknowledgement that they have read and understand the issue and preventive actions to take. Customers with questions were instructed to contact McKesson Customer Support at 800-442-6767 (option 3).For questions regarding this recall call 404-338-3556. |  | 10 | 10 | Nationwide Distribution including CA, GA, IN, KY, MD, MI, OH, OK, and WA. | Terminated | May 01, 2014 | 412 | 10 | Computer | Computer
Z-1944-2014 | 68394 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | syngo Lab Data Manager, SMN 10800057, 10803189, and 10804573. |  |  | May 20, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc. | Certain configuration parameters may cause the reference ranges to not hold results as expected and release them to the Laboratory Information System (LIS) that should have been held for manual review. | Device Design | An urgent medical device correction notice, dated May 20, 2014, was sent to end users to provide instructions to avoid potential risks associated with the software defect, until changes can be made by a Siemens technical representative. |  | 336 | 336 | Worldwide Distribution. US Nationwide, Canada, Australia, Belgium, Czech Republic, Germany, Denmark, Finland, France, United Kingdom, Italy, Netherlands, Portugal, Slovakia, Singapore, Japan, and Spain. | Terminated | August 10, 2015 | 447 | 336 | Computer | Computer
Z-1081-2014 | 67217 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | RT Desktop |  |  | November 25, 2013 | 2013 | Elekta, Inc. | During the treatment table, gantry, and collimator calibration procedures, it is possible to type incorrect values or invert a value. | Software design | Elekta sent an Important Field Safety Notice November 2013 to all affected customers. Corrective Action #1 - Release of Important Field Safety Notice 200 01 502 041 - "Recommended quality assurance after a calibration procedure" Released 25 November, 2013. Important Field Safety Notice is in the process of being distributed to customers advising them of the recommended Quality Assurance checks to avoid/detect the problem. The permanent solution to this problem is currently been developed. Corrective Action #2 - Permanent Solution - All users of Desktop Pro software will be informed of a mandatory upgrade to Integrity 1.2 upon its release which will prevent calibration from being inverted.For questions regarding this recall call 770-300-9725. |  | 489 | 489 | Worldwide Distribution - USA including AL, CA, CT, GA, FL, IL, IN, KS, KY, LA, MD, MA, MI, MN, MO, NJ, NY, OH, PA, TN, TX, VA, and Internationally to Argentina, Austria, Bangladesh, Belgium, Botswana, Brazil, Bulgaria, Canada, Chile, China, Colombia, Cuba, Czech Republic, Egypt, Finland, France, Germany, Greece, Hungary, India, Indonesia, Iran, Israel, Italy, Japan, Lebanon, Mexico, Morocco, Netherland, Nigeria, Pakistan, Panama, Peru, Philippines, Portugal, Russia, Serbia, Solvenia, South Africa, South Korea, Spain, Sweden, Switzerland, Syria, Taiwan, Thailand, Turkey, United Kingdom, Venezuela, Vietnam, and Yemen. | Terminated | August 13, 2015 | 626 | 489 | Computer | Computer
Z-1920-2014 | 68503 | No | JXG | Central nervous system fluid shunt and components. | Neurology | Neurology | 510(k) | Medtronic Duet External Drainage and Monitoring System, Interlink Injection Sites, Catalog No. 46913. |  |  | June 09, 2014 | 2014 | Medtronic Neurosurgery | Medtronic Neurosurgery is recalling the Medtronic Duet External Drainage and Monitoring Systems because the patient line tubing in the Duet External Drainage and Monitoring Systems may become disconnected from the patient line stopcock. | Device Design | An Urgent Medical Device Recall letter, dated 6/9/14, was sent to all their customers who purchased the Medtronic Duet External Drainage and Monitoring Systems. The letter informs the customers of the problem identified and the actions to be taken. Customers are instructed to cease use of the affected product, return any used units, account for any product used and verify that connections are in the condition as stated in the Instructions for Use for product currently in use. Additionally, customers are to complete the Customer Product Accountability Form. Customers with questions about the recall letter are instructed to contact Bob Shokoohi at (805) 571-8725. |  | 3026 | 3,026 units | Worldwide Distribution -- USA, Australia, Belgium, Canada, Croatia, Cyprus, Denmark, Egypt, Finland, France, Greece, Iran, Israel, Italy, Kuwait, Netherlands, Norway, Poland, Puerto Rico, Romania, Russian Federation, Saudi Arabia, Singapore, South Africa, Switzerland, Turkey, and United Kingdom. | Terminated | January 14, 2015 | 219 | 10752 | Computer | Computer
Z-1934-2014 | 68457 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | NexGen Complete Knee Solution, MIS Total Knee Procedure, Stemmed Tibial Component, Precoat, Fixed Bearing, Size 4. Part of the NexGen systems of semi-constrained, non-linked condylar knee prosthesis. |  |  | May 19, 2014 | 2014 | Zimmer, Inc. | Zimmer is initiating a voluntary recall of specific identified lots due to the potential that the threads may be out of specification. | Equipment maintenance | Zimmer sent an Urgent Medical Device Recall letter June 2014 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the notification and ensure affected personnel are aware of the contents. Locate and quarantine any affected product and notify their Zimmer sales representative. Their Zimmer sales representative will remove the recalled product from their facility. For patients that previously had this product implanted, it is recommended that they continue their normal operative follow up routine. Customers with questions were instructed to call 1-877-946-2761. |  | 41180 | 41,180 total | Worldwide Distribution - USA (nationwide) and Internationally to Australia, China, Hong Kong, India, Japan, Korea, Malaysia, Nicaragua, Singapore, and Taiwan. | Terminated | October 29, 2015 | 528 | 41180 | Computer | Not_Computer
Z-2043-2014 | 68632 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Symbia E camera system. Emission Computed Tomography System used to detect or image the distribution of radionuclides in the body or organ, using the following techniques: planar imaging, whole body imaging, tomographic imaging for isotopes with energies up to 588 keV. Nuclear medicine diagnostic device. |  |  | April 23, 2014 | 2014 | Siemens Medical Solutions USA, Inc. | Siemens Medical Solutions USA, Inc. has received reports indicating the collimator cart rear casters may become loose. | Device Design | A Customer Safety Advisory Notice dated April 23, 2014 was sent to every direct account. The letter included instructions for direct accounts to discontinue use of the collimator cart and notify the local service representative if they notice or suspect tha one or more casters are loose (which would be indicated by shaking while the cart is in motion). Otherwise, a service representative will be contacting the direct account within the following 3 months to schedule a retention bracket installation. Direct accounts with questions can contact Siemens Medical Solutions USA, Inc. at 800-888-7436. |  | 531 | 531 systems | Worldwide Distribution- US: Nationwide (AK ; AL ; AR ; AZ ; CA ; CO ; CT ; DE ; FL ; GA ; HI ; IA ; ID ; IL ; IN ; KS ; KY ; LA ; MA ; MD ; ME ; MI ; MN ; MO ; MS ; MT ; NC ; ND ; NE ; NH ; NJ ; NM ; NV ; NY ; OH ; OK ; OR ; PA ; RI ; SC ; SD ; TN ; TX ; UT ; VA ; VT ; WA ; WI ; WV ; WY) + District of Columbia + Puerto Rico; FOREIGN: Albania ; Algeria ; Angola ; Argentina ; Australia ; Austria ; Azerbaijan ; Bahrain ; Bangladesh ; Belarus ; Belgium ; Botswana ; Brazil ; Bulgaria ; Canada ; Chile ; China ; Colombia ; Costa Rica ; Croatia ; Cuba ; Czech Republic ; Dominican Rep. ; Ecuador ; Egypt ; El Salvador ; Finland ; France ; Georgia ; Germany ; Ghana ; Greece ; Hong Kong ; India ; Indonesia ; Iran ; Iraq ; Ireland ; Israel ; Italy ; Jamaica ; Japan ; Jordan ; Kazakhstan ; Kuwait ; Macedonia ; Malaysia ; Mexico ; Morocco ; Myanmar ; Netherlands ; New Zealand ; Niger ; Nigeria ; Norway ; Oman ; Pakistan ; Palestinian Territory ; Peru ; Philippines ; Poland ; Portugal ; Qatar ; Romania ; Russia ; Saudi Arabia ; Serbia ; Singapore ; Slovakia ; Slovenia ; South Africa ; South Korea ; Spain ; Sri Lanka ; Sweden ; Switzerland ; Syria ; Taiwan ; Thailand ; Trinidad and Tobago ; Tunisia ; Turkey ; Ukraine ; United Arab Emirates ; United Kingdom ; Uruguay ; Vietnam ; Yemen. | Open, Classified |  |  | 2989 | Computer | Not_Computer
Z-2044-2014 | 68570 | No | EFB | Dental handpiece and accessories. | Dental | Dental | 510(k) | Lynx TM20 Torquemaster Low Speed Handpiece MTI Dental Products |  |  | May 07, 2012 | 2012 | MTI Precision Products LLC. | Sharp edge on the body/housing of the Lynx TM20 TorqueMaster Low Speed Handpiece. | Software design (manufacturing process) | MTI notified their customers via e-mail on 5/712 and 5/8/12. MTI sent end-users recall letters on or about 5/9/12. |  | 134 | 134 devices | US Nationwide Distribution in the states of CA, FL, IL, KY, MA, MD NJ, NY, and RI. | Terminated | March 12, 2015 | 1039 | 134 | Computer | Not_Computer
Z-1108-2014 | 67481 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | M3536A HeartStart MRx with Wireless Link Upgrade (861289) |  |  | February 18, 2014 | 2014 | Philips Medical Systems, Inc. | Wireless Link may be unable to transmit data when configured for a specific data flow. | Software design | Philips Healthcare sent an Urgent Medical Device Correction letter to all consignees on February 18, 2014. The Letter informs customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue.Philips is initiating a correction to affected devices. Philips will release a new version of Wireless Link software V2.0.0.3 sub1 that corrects the issue. The correction will be provided free of charge to all units affected by this issue. A Philips Healthcare representative will contact customers with affected devices to arrange for installation of the software upgrade.Customers with questons were instructed to call 1-800-722-9377.For questions regarding this recall call 800-722-9377. |  | 220 | 220 devices | Worldwide Distribution - USA including AZ, CA, FL, IL, IN, KS, KY, MA, MD, MI, MN, MO, NC, NY, OH, OK, OR, PA, SC, TN, TX,and VA and Internationally to Australia, Norway, and Spain. | Terminated | November 12, 2015 | 632 | 220 | Computer | Computer
Z-1111-2014 | 67441 | No | OOI | Instrumentation for clinical multiplex test systems. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | cobasÂż PCR Urine Kit 100 PKT IVD |  |  | January 06, 2014 | 2014 | Roche Molecular Systems, Inc. | Numerous complaints have been filed that leaking cobas Âż PCR media 4.3 mL IVD have been received from different lots of thecobas Âż PCR Urine kits, cobas Âż PCR Female Swab kits, and cobas Âż PCR media kits. | Device Design | Roche Molecular Diagnostics Inc. sent an Urgent Medical Device Correction letter/Fax-back form dated January 6, 2014. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to follow the test-skpecific Instruction For Use for handling of cobas PCR media 4.3 ml IVD. The use of protective equipment should be worn as indicated in the labeling. Inspection each tube and/or blister packet for evidence of leakage prior to use. If leakage of the cobas PCR media 4.3 mL IVD is observed, do not use the tubes. Discard the affected tubes per local guidelines and contact Roche Support Network Customer Support Center at 1-800-526-1247. Complete the attached fax form and fax it to 1-855-695-8564. File this Urgent Medical Device Correction (UMDC) for future reference. Customers with questions were instructed to call 1-800-526-1247.For questions regarding this recall call 908-253-7200. |  | 84 | 84 kit batches | Nationwide Distribution | Terminated | August 07, 2014 | 213 | 84 | Computer | Not_Computer
Z-1059-2014 | 67231 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Synergy XVI |  |  | November 25, 2013 | 2013 | Elekta, Inc. | XVI can incorrectly calculate the target position of the treatment table. | Software design | Elekta sent an Important Field Safety Notice dated November 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers. Potential collision risk when using XVI and external beam shaping devices released November 25, 2013 (Only software versions R4.5, R4.5.1, R4.6 and R5.0 are applicable to this recall). (All software versions R4.6 in this notices are applicable to this recall). Recommendations to Users - Monitor the Actual column of the Table Move Assistant dialog box. If the Actual is blank, do not press the ASU button the FKP. To correct the error - (1) Restart XVI. (2) Put the patient in their initial setup position. (3) Scan the patient again, and continue the usual workflow. Always monitor the patient when you do RATM. When you do Volume View acquisitions with the Apex collimator attached, make sure that the stop angle of the Volume View parameters is in the lower half of the gantry rotation. This decreases the risk of a collision with the patient. Read and follow the advice int he Important Field Safety Notices and acknowledge receipt to your Elekta representative within 30 days. Corrective Action #2 - Permanent Solution - XVI Risk control measure - changes to the software to prevent table movement if the start position has not been recorded correctly. Recommendations for changes to product are - (1) Table zero should default to NIHII not zero. (2) Check that the table zero is valid because calculating absolute table positions. (3) Check prescribed table position is not more than 2.5cm from actual table position when permitting ASU from outside room. Users of software version 4.5, 4.5.1 and 4.6 will receive a mandatory upgrade to XVI 5.0. The estimated release date for the fix for the XVI 5.0 is December 2014.For further questions call ( 770 ) 300-9725. |  | 1028 | 1028 | Worldwide Distribution - USA (nationwide) including Puerto Rico and the states of AL, AK, AZ, AR, CA, CT, FL, GA, ID, IL, KY, LA, MD, MA, MI, MN, MO, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SD, TN, TX, VT, VA, WA, WV, WI, WY and DC., and the countries of Algeria, Australia, Austria, Bahrain, Belarus, Belgium, Brazil, Bulgaria, Canada, China, Columbia, Croatia, Cuba, Cyprus, Czech Republic, Denmark, Ecuador, Estonia, Finland, France, Germany, Greece, Hong Kong, Hungary, India, Iraq, Ireland, Israel, Italy, Japan, Jordan, Malaysia, Mexico, Morocco, Netherlands, New Zealand, Norway, Pakistan, Poland, Portugal, Romania, Russia, Saudi Arabia, Singapore, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom and Venezuela | Open, Classified |  |  | 1028 | Computer | Not_Computer
Z-2139-2014 | 68790 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS 3600 Immunodiagnostic System --- For use in the in vitro quantitative, semi-quantitative and qualitative measurement of a variety of analytes of clinical interest, using VITROS Immunodiagnostic Products Reagents. The VITROS Software is a part of the VITROS 3600, 4600, &amp; 5600 Systems. The VITROS 3600, 4600, &amp; 5600 Systems share the same software base. VITROS Sample Programming consists of Sample Identification (SID), test requests, and patient demographics. Patient demographics include patient name, address, date of birth, gender, physician, and patient identification (PID). |  |  | July 07, 2014 | 2014 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics has identified an anomaly with VITROS System Software Version 3.0 and below. Internal testing confirmed that under very specific conditions, incorrect patient demographics were associated with a sample identification (ID) number. | Software in the Use Environment | On 7/07/2014, Customer Letter (Ref. CL2014-191) was sent via FedEx overnight courier or ORTHO PLUS e-Communications and/or US Postal Service Priority Mail (for PO Boxes only) to US consignees (OCD Direct, Drop-Ship and US Federal Government) to inform them of the issue and advise them of the required actions. On 7/07/2014, Distributor Letter (Ref. DL2014-191 was sent via US Postal Service Priority Mail and/or FedEx overnight courier to US Distributor consignees instructing them to notify their consignees of the issue and required actions. Foreign affiliates were informed by e-mail on 7/07/2014 of the issue and instructed them to notify their consignees of the issue and required actions. For questions, please contact Customer Technical Services at 1-800-421-3311. |  | 678 | Domestic: 142, Foreign: 536 | Worldwide Distribution-USA (nationwide) including Puerto Rico and the countries of Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Singapore, Spain and Venezuela. | Open, Classified |  |  | 5568 | Computer | Computer
Z-2131-2014 | 68657 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Zimmer Periarticular Locking Plate System Distal Radial Volar Plate Lateral Column (1st generation stainless steel only). Orthopedic fixation plate. For temporary internal fixation and stabilization of osteotomies and fractures. |  |  | June 10, 2014 | 2014 | Zimmer, Inc. | A complaint was received in which it was reported that an 8-hole plate from lot 62276042 had an enlarged hole that is unable to lock the screw, which resulted in the screw head pulling through plate hole. | Device Design | Zimmer sent an Urgent Medical Device Recall letter dated June 17, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.They instructed their distributors to do the following: 1. Please locate, quarantine, and return the affected product (part number 00-2358-014-04; lot 62276042) from all inventories.2. Please note that any hospitals that received direct shipments of this product from Zimmer will be sent a copy of the Risk Manager and Surgeon letter directly. Additionally, Risk Managers and Surgeon that were identified via invoice data will also receive a letter directly. 3. Please review the list of hospital and surgeons to be provided to your territory via email.4. Provide the name, address, and phone number of any additional hospitals or surgeons that may have received or used the affected product, or confirm that the list provided is complete.5. Complete and return Attachment 1.6. Ensure that a copy of the Risk Managers and Surgeons letter is left with the facilities from which product is removed by providing them with a copy of the letter, if they are not identified as a facility receiving a letter directly from Zimmer.7.If after reviewing this notification you have questions or concerns, please call 1-877-946-2761.They instructed the Surgeons and Risk Managers to do the following: 1. Review the notification and ensure affected personnel are aware of the contents.2. Inspect inventory for part number 00-2358-014-04; lot number 62276042. Quarantine the device and contact your Zimmer sales representative to remove and return the affected device. 3. If after reviewing this notification you have questions concerning this notification, please call 1-877-946-2761. |  | 28 | 28 units | Worldwide Distribution - US including the states of CA and OH., and the countries of BRAZIL, CHINA, KOREA and TAIWAN. | Terminated | January 09, 2015 | 213 | 28 | Computer | Not_Computer
Z-2134-2014 | 68737 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Patient Data Manager 2.0 (Content manager 2.0, Patient Browser 4.0, DICOM Viewer 2.0) Brainlab Digital Lightbox. A system for the display of medical images. Model numbers 25100-05 and 25100-06. The software can transfer images to and from picture archiving and communication systems (PACS), file servers, or removable storage media. |  |  | June 27, 2014 | 2014 | Brainlab AG | The following defects have been identified for data loaded with Patient Data Manager version 2.0 (including subversions 2.0.0, 2.0.1 and 2.0.2): 1) When loading non-square pixel image data with PDM 2.0, the data set is not correctly transferred and displayed in the corresponding Brainlab navigation or planning software. When a data set is loaded, the pixels of the data set are sorted in order to | Software design | FIELD SAFETY NOTICE / PRODUCT NOTIFICATION letters dated June 27, 2014 were sent to all direct accounts. The letters instruct product users to not upload image data sets with non-square pixels and to verify with their radiology department the modalities of the different imaging devices (e.g. CT, MRT, PET, ultrasound) from which data may be loaded onto a Brainlab device. Brain lab will provide a software update with this issue solved to affected customers with a tentative planned time line of availability being July 2014. Customers with questions can contact Brainlab by telephone at 800-597-5911 or via e-mail at us.support@brainlab.com. |  | 154 | 154 systems | Worldwide Distribution-USA (nationwide) including the states of AL, AK, AZ, CA, FL, GA, IA, IL, KS, MD, MO, MS, NC, NY, OH, PA, SD, TX, VA, VT, WA, and WI, and the countries of Australia, Austria, Belgium, Brazil, Chile, China, Ecuador, France, Germany, Japan, Netherlands, New Zealand, Norway, Qatar, Russia, Spain, Switzerland, Taiwan, United Arab Emirates, and United Kingdom. | Open, Classified |  |  | 154 | Computer | Computer
Z-1068-2014 | 66893 | No | OMP | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Renasys EZ PLUS, a Canister Component of of Renasys EZ Negative Pressure Wound Therapy System |  |  | November 14, 2013 | 2013 | Smith &amp; Nephew Inc. | Modification of the bacterial overflow guard (filter) and related changes to product labeling. | Device Design | The correction strategy is as follows:" An Urgent Medical Device Correction Notice has been prepared to be distributed to all US consignees via Federal Express delivery for confirmation of the delivery." A copy of the Urgent Medical Device Correction Notice is to be posted on Smith &amp; Nephew's websites, URL corporate, http://www.smith-nephew.com/news-and-media/, and myrenasys.com, concurrent with the distribution of the Notice to consignees." US distribution of the original design of Canister Kits codes (66800912, 66800913, 66801066, 66800423 and 66800058) subject to this correction ceased November 1, 2013." The correction has been implemented on a change being affected basis. The modified design is being distributed to US customers effective November 1, 2013. No product will be removed from the market." Periodic progress reports to FDA, as required per 21 CFR 806 will be submitted monthly until completion of the correction actions.. |  | 6871375 | 6,871,375 canisters | Worldwide Distribution: USA (nationwide) including states of: AL, AZ, CA, CO, CT, FL, GA, HI, IA, IL, IN, KY, LA, MA, MI, MN, MO, MS, MT, NC, NH, NV, NY, OH, OR, PA, RI, SC, SD, TN, TX, VA, WA, WV, and Puerto Rico; and the countries of: Austria, Australia, Belgium, Switzerland, China, Cyprus, Germany, Denmark, Spain, Finland, France, UK, Greece, Israel, Italy, Malta, Netherlands, Norway, Paraguay, Russia, Saudi Arabia, Sweden, Thailand, Turkey, South Africa, Mexico, and Canada. | Terminated | August 04, 2014 | 263 | 13670515 | Computer | Not_Computer
Z-2130-2014 | 68697 | No | CCW | Rigid laryngoscope. | Anesthesiology | Anesthesiology | 510(K) Exempt | RUSCH, TruLite, One-Piece Disposable Laryngoscope Handle Blade Combination, Non-Sterile, Rx Only, Distributed by: Teleflex Medical. |  |  | June 30, 2014 | 2014 | Teleflex Medical | There is a possibility the handle of the laryngoscope may heat up and melt the end cap. | Device Design | Consignees were notified via letter on/about 06/30/2014. |  | 6246 | 6,246 units | U.S. Nationwide | Terminated | March 21, 2017 | 995 | 6246 | Computer | Not_Computer
Z-2121-2014 | 68799 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Horizon Medical Imaging |  |  | May 05, 2014 | 2014 | Mckesson Medical Immaging | The firm has identified a software issue which may cause a discrepancy between the index of images expected by the application, and the actual content of the folder storing the images. | Software design | McKesson sent an Urgent Field Safety Notice dated May 5, 2014, to all affected customers informing them of the issue and advising them of the immediate actions to take to reduce the risk to the patient. These actions include - Users ensure that all expected images are display in the Image Viewer and PACS Administrators should engage McKesson Support for all issues with studies. McKesson is also providing a software solution to HMI systems that may be potentially affected by the problem to prevent future recurrence. Customers were instructed to contact McKesson Support as soon as possible at 1-800-663-2533 to acknowledge that they have read and understood the Field Safety Notice, and to arrange installation of the software update. Customers with questions were instructed to contact Customer Support and ask to speak to their Support Manager.For questions regarding this recall call 404-338-3556. |  | 659 | 659 units | Worldwide Distribution - USA including AL, AK, AZ, AR, CA, CO, CT, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NM, NJ, NY, NC, ND, OH, OR, PA, RI, SC, TN, TX, VT, VA, WA, WV, WI, WY,and France, Canada, Puerto Rico, Guam, Great Britain, Ireland, Australia. | Terminated | April 13, 2017 | 1074 | 659 | Computer | Computer
Z-2122-2014 | 68470 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | InTouch Critical Care Bed, Stryker Medical. Model Numbers 2130, 2131, 2140, 2141, 2151, 2152, 2155, and 2156. |  |  | June 04, 2014 | 2014 | Stryker Medical Division of Stryker Corporation | Stryker has received complaints from the field alleging instances where the In Touch beds are hoisted or lifted from the floor and the casters have fallen off. The investigations into these casters have shown that in these cases the caster has broken on the top of the caster stem where the brake rod inserts into the caster. If more than one caster on a unit is damaged, brake functions can be affec | Device Design | On 06/30/1014 the firm sent "Urgent Medical Device Notification" letters to their customers. |  | 11632 | 11,632 units | Worldwide Distribution -USA (nationwide) including DC and the states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI &amp; WV, and the countries of Argentina, Brazil, Canada, Chile, China, EMEA, Greece, Italy, Japan, Korea, Latin America, Mexico, Poland, Singapore, South Africa, Taiwan, and UK. | Terminated | March 14, 2016 | 649 | 11632 | Computer | Not_Computer
Z-1078-2014 | 67401 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Precision 500D and Proteus XR/a equipped with Wireless Digital Radiography Option (WDR1). |  |  | January 20, 2014 | 2014 | GE Healthcare, LLC | GE Healthcare has recently become aware of a software issue associated with patient selection from the worklist on your WDR1(Precision 500D and Proteus XR/a equipped with Wireless Digital Radiography Option) that may impact patient safety.Intermittently when a large number of patient exams are retrieved from the HIS/RIS system, the exams when displayed on the WDR1 are not properly displayed. I | Software design | GE Healthcare sent an "Urgent Medical Device Correction" letter dated January 20, 2014. The letter was addressed to Hospital Administrators / Risk Mangers, Radiology Department Mangers and Radiologists. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information ( In USA call 800-437-1171, For other countries contact your local GE Healthcare Service Representative). |  | 359 | 359 (333 US, 26 OUS). | Worldwide Distribution - USA (nationwide) including the states of AL, AK, AZ, AR, CA, CO, DC, FL. GA, HI. IA, ID, IL, IN, LA, KY, ME, MD, MA, MI, MN, MS, MO, NE, NV, NH, NJ, NM, NY, NC, OH, OK, OR, PA, TN ,KY, TX, UT, VA, WA, WV, WI and WY., and the countries of ARGENTINA, AUSTRALIA, CANADA, FRANCE, IRELAND, ITALY, JORDAN, ROMANIA, RUSSIA, UNITED ARAB EMIRATES and UNITED KINGDOM. | Terminated | November 03, 2014 | 287 | 359 | Computer | Computer
Z-2116-2014 | 68586 | No | FDS | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | EG-3630U Ultrasound Video Gastroscope |  |  | June 12, 2014 | 2014 | Pentax Medical Company | Aspiration needles used in combination with Ultrasound Gastroscopes can derail from the accessory elevator at the distal end of the endoscope. This may result in patient injury if the Instruction for Use is not followed carefully by the user. | Device Design | Pentax Medical sent an Urgent Medical Device Field Safety Corrective Action dated June 12, 2014 to affected customers. The letter identified the affected product, problem and actions to be taken. Pentax also posted their notification on their website. For questions regarding this action please feel free to contact Paul Silva, RecallCoordinator atTel: 800-431-5880 ext 2064 (8:30am  5:00pm Mon-Fri, EST)Fax: 201-799-4063Email : paul.silva@pentaxmedical.com |  | 531 | 531 units | US Nationwide Distribution | Terminated | November 17, 2015 | 523 | 531 | Computer | Not_Computer
Z-2114-2014 | 68378 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, CARESCAPE Monitor B650. A multi-parameter patient monitor intended for use in multiple areas and intrahospital transport within a professional healthcare facility. |  |  | July 07, 2014 | 2014 | GE Healthcare, LLC | Possible sound loss associated with the CARESCAPE Monitor B650. When an external display is connected to the CARESCAPE Monitor B650 using a USB extender cable with an active USB hub, the system may experience a sound loss including alarm sounds from the patient monitor. This may occur if the external display is powered off and on during normal operation of the monitor. | Software design | GE Healthcare "Urgent Medical Device Correction" letter #36102 dated July 7, 2014 to affected customers. The letter was addressed to Healthcare Administrator / Risk Manger, Chief of Nursing and Director of Biomedical Engineering. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. For questions they can contact Technical Support at 1-800-558-7044 or your local Service Representative. |  | 21799 | 21,799 (7,635 units US, 14,164 units OUS) | Worldwide distribution. US nationwide including Puerto Rico, ZAMBIA, YEMEN, VIETNAM, VENEZUELA, URUGUAY, UNITED KINGDOM, UNITED ARAB EMIRATES, TURKEY, THAILAND, UNITED REPUBLIC OF TANZANIA, TAJIKISTAN, TAIWAN, SYRIA, SWITZERLAND, SWEDEN, SRILANKA, SPAIN, SOUTH AFRICA, SLOVENIA, SLOVAKIA, SINGAPORE, SAUDI ARABIA, RUSSIAN FÂżDÂżRATION, ROMANIA, QATAR, PORTUGAL, POLAND, PHILIPPINES, PERU, PARAGUAY, PANAMA, PAKISTAN, OMAN, NORWAY, NICARAGUA, NEW ZEALAND, NETHERLAND, NAMIBIA, MYANMAR, MOROCCO, MEXICO, MAURITIUS, MALAYSIA, LITHUANIA, LIBYAN ARAB JAMAHIRIYA, LEBANON, LATVIA, KUWAIT, REPUBLIC OF KOREA, KAZAKHSTAN, JORDAN, Italy, JAPAN, JAMAICA, ISRAEL, IRELAND, IRAQ, INDONESIA, INDIA, ICELAND, HUNGARY, HONG KONG, HONDURAS, GUATEMALA, GREECE, GERMANY, GEORGIA, FRANCE, FINLAND, ESTONIA, EL SALVADOR, Egypt, ECUADOR, DENMARK, CZECH REPUBLIC, CYPRUS, CROATIA, COSTA RICA, COLOMBIA, CHINA, CHILE, CANADA, BULGARIA, BRUNEI, BRAZIL, BOLIVIA, BERMUDA, BELGIUM, BANGLADESH, BAHRAIN, BAHAMAS, AZERBAIJAN, AUSTRIA, AUSTRALIA, ARGENTINA. | Open, Classified |  |  | 21799 | Computer | Computer
Z-2110-2014 | 68756 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | QLAB 10.1 Core Module PN 453561728001. QLAB Quantification Software is a software application package designed to view and quantify image data acquired on Philips Medical Systems ultrasound products. |  |  | July 09, 2014 | 2014 | Philips Ultrasound, Inc. | The QLAB a2DQ and aCMQ features have a defect that may use incorrect measurement values under specific conditions. | Software change control | Philips Medical Systems notified customers of the voluntary recall by sending a certified URGENT-Medical Device Correction letter, dated 2014 JUL 09, to all US and Canadian consignees. Philips representatives will distribute the letter internationally. Follow-up letters to wholesale customers will be provided and wholesalers are requested to forward copies of recall letters and response forms to their customers. Philips Medical Systems will also have Customer Reply Cards returned for tracking purposes. Customers are directed to call local Philips representative or Philips Customer Service at 1-866-767-7822. |  | 59 | 59 units total (12 units in the US and 47 units outside the US) | Distributed in the states of CA, IA, KY, MD, NY, PA, UT, VA, WA and WI.Also distributed in Australia, Belgium, Canada, India, Italy, Japan, South Korea, Netherlands, New Zealand, Spain, Sweden, and Switzerland. | Terminated | November 06, 2015 | 485 | 59 | Computer | Not_Computer
Z-2111-2014 | 68769 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Siemens syngo Imaging XS, version VA70A or higher |  |  | June 30, 2014 | 2014 | Siemens Medical Solutions USA, Inc | Siemens has become aware of the potential malfunction when using syngo Imaging XS, version VA70A or higher. In some cases a miscalculation may occur for the grey scale values in the functions Region of Interest , Pixel Lens , Edge enhancement and Histograms . This potential malfunction may occur for images with a pixel depth greater than 12 bit. All values calculated for images with a pixel d | Software design | Siemens sent a Customer letter dated June 30th, to all affected customers to inform them of a new software version VA70B_0114 that will be introduced to all syngo Imaging XS customers with version VA70A or higher to solve a potential issue regarding miscalculations done for the grey scale values in the functions "Region of Interest", "Pixel Lens", "Edge Enhancement", and "Histograms".For questions regarding this recall call 610-219-6300. |  | 141 | 141 | Nationwide Distribution | Terminated | February 24, 2015 | 239 | 141 | Computer | Computer
Z-1074-2014 | 67307 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Single and Dual use Emission Computer Tomography gamma cameras (BrightView XCT, BrightView X &amp; BrightView SPECT) Philips Healthcare Systems, Cleveland, OH. |  |  | December 19, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | Three problems were reported to the firm:When positioning for a relative 180 degree non circular acquisition the second (bottom) detector was not positioned correctly causing an inappropriate rotation orbit calculation. When the acquisition was started and the detectors began to move into position, the incorrect calculation resulted in the detectors coming in contact with the patient.As | Software design | On 12/19/2013 the firm sent Field Safety Notices to their customers. |  | 1143 | 1,143 units | Worldwide Distribution-USA (nationwide) including DC and Puerto Rico and the states of AK, AR, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, VA, VT, WA, WI, WV &amp; WY, and the countries of Algeria, Argentina, Australia, Austria, Bahrain, Belgium, Brazil, Canada, China, Czech Republic, Denmark, Dominican Republic, Egypt, Finland, France, Gabon, Germany, Greece, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Korea, Lebanon, Lithuania, Malaysia, Mexico, Netherlands, New Zealand, Norway, Oman, Palestine, Panama, Philippines, Poland, Portugal, Russia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Syrian Arab Republic, Taiwan, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom &amp; Viet Nam. | Terminated | January 12, 2015 | 389 | 1143 | Computer | Not_Computer
Z-1079-2014 | 66430 | No | KWT | Shoulder joint metal/polymer non-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Trabecular Metal" Reverse Shoulder System |  |  | October 17, 2013 | 2013 | Zimmer, Inc. | As a result of complaints for disassociation of the TM Reverse Glenosphere from the Baseplate, Zimmer updated the related surgical technique and released a video to provide additional instructions to users on how to properly assemble the Glenosphere to the Baseplate. | Device Design | In November and December 2012 Sales staff were provided updated surgical technique and associated video to share with current customers, surgeons, and risk managers. October 17, 2013 Distributors ( via E-mail),and Risk Managers and Surgeons (via traceable courier) were sent written Urgent Recall notifications of these labeling changes. Current customers are identified through an invoice search and distribution history for the affected devices for the past 12 months, since 9/1/12. |  | 73002 | 73,002 units | Worldwide Distribution-USA (nationwide) and the countries of Germany, India, Malaysia, China, Singapore, Korea, Taiwan, and Brazil. | Terminated | September 12, 2014 | 330 | 73002 | Computer | Not_Computer
Z-1080-2014 | 67175 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Siemens Artis zee Ceiling Systems. |  |  | December 16, 2013 | 2013 | Siemens Medical Solutions USA, Inc | During the lifetime of Artis zee Ceiling systems, there is an increased wear due to potential sub-optimal cable guidance that can lead to cable damage. | Device Design | In December 2013, a recall notification letter was sent to Users. |  | 284 | 284 | Nationwide Distribution. | Terminated | March 09, 2015 | 448 | 284 | Computer | Not_Computer
Z-2084-2014 | 68440 | No | KWA | Hip joint metal/metal semi-constrained, with an uncemented acetabular component, prosthesis. | Orthopedic | Orthopedic | PMA | RECLAIM DISTAL TAPERED Distal Stem, various sizes. Taper Protector Sleeve. RECLAIM Revision Hip System Orthopedic Manual Surgical Instrument, used to ensure the locking taper surface of the Distal Stem is protected during proximal preparation. |  |  | June 03, 2014 | 2014 | DePuy Orthopaedics, Inc. | The product can be difficult to remove from the Distal Stem both out of the package andafter proximal reaming. | Device Design | DePuy Synthes sent an Urgent Information - Device Correction Notice letter dated June 24, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to:1. Immediate Remedy: Discard the RECLAIMÂż Taper Sleeve Protector instrument before stem implantation.2. Changes to the literature and IFU: Documents are revised as follows:" RECLAIMÂż Revision Hip Systems Surgical Technique # DSUS/JRC/0614/0200  pages 11, 12, 13, 16,30, 34, 35 (figure numbering from pages 11 to 27 were also revised): Removed information regardingtaper sleeve protector." RECLAIMÂż Revision Hip Systems Surgical Technique FAQ # DSUS/JRC/0614/0202  pages 1 and 2:Removed information regarding taper sleeve protector." RECLAIMÂż Revision Hip Systems IFU #090200799 REV E  Page 5: Verbiage will be revised andimplemented as soon as possible, depending on regulatory approvals.3. Change: The company consulted with design surgeons and developed a strategy to eliminate the taper sleeve protector from the system. As a result of removing the taper sleeve protector, the company isevaluating whether there is a need for a protective component to be used during the distributionCustomers with questions were instructed to call 574-371-4917.For questions regarding this recall call 574-371-4577. |  | 3623 | 3,623 | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI &amp; WV and Internationally to AUSTRALIA, AUSTRIA, CANADA, CHILE, CHINA, CZECH REPUBLIC, DENMARK, FRANCE, GERMANY, IBERIA, IRELAND, ISRAEL, ITALY, NEW ZEALAND, NORWAY, POLAND, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SWITZERLAND &amp; UK. | Terminated | May 17, 2017 | 1079 | 14904 | Computer | Not_Computer
Z-1090-2014 | 67429 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | HeartStart MRx Monitor/Defibrillator; Model(s), Catalogue, or Code Number: M3535A, M3536A, M3536J, M3536MC |  |  | February 18, 2014 | 2014 | Philips Medical Systems, Inc. | A component of the MRx Processor Board may be susceptible to damage from electrostatic discharge (ESD), which can disrupt ECG and SpO2 functionality leading to an inability to perform. | Device Design | The firm, Philips, sent an "URGENT - Medical Device Correction" letter dated February, 2014 with an attached Field Safety Notice to each known affected customer on 2/18/14. The letter describes the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue.The customers were instructed to continue to use your MRx prior to receiving the hardware upgrade; if a failure has been detected that may disable some ECG and SpO2 functionality during patient care, treat your patient per existing protocols and if the condition persists, take the device out of use and call for service from Philips authorized service personnel.Philips is initiating a correction to affected devices. A hardware upgrade will be provided free of charge for all units affected by this issue. A Philips Healthcare representative will contact customers with affected devices to arrange for installation of the hardware upgrade.If you have questions regarding this notification or need any further information or support, please contact your local Philips representative or call us at 1-800-722-9377. |  | 39128 | 39,128 | Worldwide Distribution: US (Nationwide) including states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, GU, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY; and Internationally to: Canada, AFGHANISTAN, ALBANIA, ARGENTINA, ARUBA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BELGIUM, BOSNIA AND HERZEGOWINA, BRAZIL, BRUNEI DARUSSALAM, BULGARIA, CHILE, CHINA, COLOMBIA, CROATIA (local name: Hrvatska), CYPRUS, CZECH REPUBLIC, DENMARK, ECUADOR, EGYPT, ESTONIA, FINLAND, FRANCE, GERMANY, GREECE, GUADELOUPE, HONG KONG, ICELAND, INDIA, INDONESIA, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KENYA, KOREA, REPUBLIC OF, KUWAIT, LATVIA, LEBANON, LIBYAN ARAB JAMAHIRIYA, LITHUANIA, LUXEMBOURG, MACEDONIA, MALAYSIA, MEXICO, MOROCCO, NAMIBIA, NETHERLANDS, NETHERLANDS ANTILLES, NEW CALEDONIA, NEW ZEALAND, NICARAGUA, NORWAY, OMAN, PAKISTAN, PALESTINE, PANAMA, PAPUA NEW GUINEA, PERU, PHILIPPINES, POLAND, PORTUGAL, PUERTO RICO, QATAR, REUNION, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SERBIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SRI LANKA, SWEDEN, SWITZERLAND, TAIWAN, PROVINCE OF CHINA, THAILAND, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, UGANDA, UKRAINE, UNITED ARAB EMIRATES, UNITED KINGDOM, VENEZUELA, VIET NAM, and YEMEN. | Open, Classified |  |  | 39128 | Computer | Computer
Z-2085-2014 | 68569 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, CARESCAPE Monitor B850, B650 or B450. Patient Monitor. |  |  | June 13, 2014 | 2014 | GE Healthcare, LLC | Partial loss of monitoring with the CARESCAPE Monitor B850, B650 or B450. When connecting the PDM (Patient Data Module) to the CARESCAPE B850, B650 or B450 monitor with software version 2 the PDM may not start communicating to the monitor and loss of the PDM parameters could occur. | Software design | Consignees were sent on 6/13/2014 a GE Healthcare "Urgent Medical Device Correction" letter #36103, 36104 and 36105 dated June 6, 2014. The letter was addressed to Healthcare Administrator / Risk Manger, Chief of Nursing and Director of Biomedical Engineering. The letter described the Safety Issue, Safety Instructions, Affected Products Details, Product Correction and Contact Information. Consignees were advised to contact Technical Support at 800-558-7044 in regard to questions on this recall. |  | 41992 | 41,992 (20,509 US; 21,483 OUS) | Worldwide Distribution: US nationwide, UNITED KINGDOM, UNITED ARAB EMIRATES, TURKEY, TAIWAN, SWITZERLAND, SWEDEN, SPAIN, SOUTH AFRICA, SLOVENIA, SLOVAKIA, SINGAPORE, SERBIA, SAUDI ARABIA, RUSSIAN FEDERATION, ROMANIA, REUNION, QATAR, PORTUGAL, POLAND, PHILIPPINEs, PARAGUAY, PANAMA, NORWAY, NEW ZEALAND, NETHERLANDS, MYANMAR, MOROCCO, MEXICO, MALAYSIA, LITHUANIA, LEBANON, LATVIA, JORDAN, JAPAN, JAMAICA, ITALY, ISRAEL, IRELAND, IRAQ, INDONESIA, INDIA, ICELAND, HUNGARY, GREECE, GIBRALTAR, GERMANY, FRANCE, FINLAND, ESTONIA, EGYPT, ECUADOR, DENMARK, CZECH REPUBLIC, CYPRUS, CROATIA, CHILE, BRUNEI, BERMUDA, BELGIUM, BANGLADESH, AUSTRIA, AUSTRALIA. | Open, Classified |  |  | 41992 | Computer | Computer
Z-2088-2014 | 68659 | No | JPA | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(k) | Hemochron Jr. Whole Blood Microcoagulation System APTT Cuvette. For in vitro diagnostic coagulation test. Catalog number J103. |  |  | June 20, 2014 | 2014 | ITC-Nexus Dx | Test results may be erroneously elevated or report an out-of-range high (OOR-H) error code in non heparinized blood samples. | Device Design | ITC issued an Urgent Medical Device Recall Letters/Response Forms dated June 20, 2014, to End-users and US Distributors were notified. ITC asks that this communication be forwarded to all those within your organization who need to be aware of this matter. If any affected products have been forwarded to another facility, a copy of this communication should be forwarded to them immediately.Customers should check their inventory for the affected lot. The affected lot should not be used and quarantined until return to ITC. The attached form should be completed and returned to ITC within 10 days. ITC will review the form and contact the customer to coordinate any returns. For questions, please contact Technical Support: Telephone 732-548-5700, Extension 4011 E-mail: Techsupport@itcmed.comNormal business hours are Monday through Friday, 8:00 AM to 6:30 PM, Eastern Time.For urgent matters, Technical Support is available 24 hours per day, 7 days a week. |  | 3150 | US- 315 cuvettes; OUS - 2835 cuvettes | Worldwide Distribution - US including the states of Georgia, Virginia and New York., and the countries of Canada, France, Italy, Liechtenstein and Romania. | Terminated | December 11, 2014 | 174 | 3150 | Computer | Not_Computer
Z-1091-2014 | 67342 | No | EHD | Extraoral source x-ray system. | Dental | Radiology | 510(k) | The FOCUS is a microprocessor controlled intra-oral x-ray unit with a Load Matching Frequency (LMF) DC generator, which produces dental images on digital and film media. For the digital images the FOCUS is connected to the Sigma USB terminal which is connected to a PC, which reads the image data from a sensor and displays the image on the PC monitor. In order to make installation and service faster and easier the electronics is integrated into the horizontal arm. |  |  | January 14, 2014 | 2014 | Instrumentarium Dental, PaloDEx Group Oy | Instrumentarium Dental has delivered a number of FOCUS 3.0 Intraoral X-ray units in November 2013 with a potential flaw in the mounting hardware. A number of wall mount units were delivered with FOCUS X-ray machines that had not been properly welded at our supplier. As a result of the incomplete welding operation, the wall mount unit on the FOCUS X-ray unit may not properly support the unit as exp | Device Design | The firm, Instrumentarium, sent an "URGENT MEDICAL DEVICE RECALL" letter dated January 14, 2014 was sent to all PaloDEx direct accounts (including subsidiaries). The letter inested recipients to perform a field inspection on potentially defective FOCUS 3.0 X-ray units (units with serial numbers F21221 - F21641 and F21643). If a defective wall mount is discovered during the inspection, the FOCUS X-ray unit must be taken out of service and the wall mount replaced before continuing use of the device. Replacement wall mounts will be made available free of charge for the field repair. A Technical Bulletin including instructions on how to perform the inspection was included with the letter. Direct accounts were further instructed to return the enclosed Acknowledgement Letter confirming receipt of the URGENT MEDICAL DEVICE RECALL letter by mail to Instrumentarium Dental, Jussi Tanskanen, Nahkelantie 160, Fl-04300 Finland; or, via e-mail to Jussi.tanskanen@instrudental.com. The direct accounts are to complete and return the enclosed Confirmation Letter to document that all of the potentially affected units have been inspected and repaired if necessary to Instrumentarium Dental, Jussi Tanskanen, Nahkelantie 160, Fl-04300 Finland; or, via e-mail to Jussi.tanskanen@instrudental.com. A nearly identical URGENT MEDICAL DEVICE RECALL letter dated January 22, 2014 was sent to all US customers. The instructions were the same as noted above in the January 14, 2014 letter to PaloDEx direct accounts.If there are any questions, do not hesitate to contact us at the following: Marketing Director, General dentistryMika.anttila@instrudental.comTel. +358 40 768 3340For more information you can contact also:Director, Quality and Regulatory AffairsTarja.paivinen@instrudental.comTel. +358 40 719 0513After sales ManagerMika.Kyllonen@instrudental.comTel. + 358 40 080 5305 |  | 101 | 101 units (US Distribution) | Worldwide Distribution: US (Nationwide) including states of: AK, AZ, CA, CO, GA, IL, KY, MD, MT, NC, NV, OR, PA, TN, TX, UT, VA, WA,and WI; and Internationally to: Belgium Brazil, Chile, China, Czech Republic, France, Germany, Italy, Spain, Sweden, Switzerland, and United Kingdom. | Terminated | May 22, 2015 | 493 | 101 | Computer | Not_Computer
Z-1086-2014 | 67388 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Specialty Straight Static Inserter Assembly |  |  | December 18, 2013 | 2013 | Stryker Spine | Stryker has received complaints from customers relating to fracture along the weld that binds the handle to the shaft of the static inserter. | Device Design | Stryker sent Urgent Medical Device Removal Notification letter and to their customers via email on 12/18/2013. The letter identified the affected product, problem and actions to be taken. Customers were instructed to examine their inventory, return affected product using enclosed pre-paid mailing label, complete Customer Response Form and fax to Regulatory Compliance at 855-632-9049 or mail it to Spine-RegulatoryActions@Stryker.com. For questions call 201-760-8287, 201-760-8298 or 201-760-8041. |  | 21 | 21 units | US Nationwide Distribution | Terminated | August 25, 2014 | 250 | 21 | Computer | Not_Computer
Z-2068-2014 | 68344 | No | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | DATEX-OHMEDA S/5 SINGLE-WIDTH AIRWAY MODULE, E-MINIC AND ACCESSORIES. |  |  | June 11, 2014 | 2014 | GE Healthcare, LLC | Failure of the CO2 detector in Single-width Airway and Extension modules, which may cause a slow continuous decrease of measured Et-/FiCO2 values. Incorrect Et CO2/Fi CO2 value may impair clinical decision making for both mechanically and spontaneously ventilated patients because of incorrect low CO2 values. | Device Design | GE Healthcare sent an Urgent Medical Device Correction letter dated June 11, 2014, to all affected consignees. The letter described the Safety Issue, Safety Instructions, Affected Product Details, and Product Correction and Contact Information. For questions customers were instructed to contact Technical Support at 1-800-558-7044 or their local Service Representative. |  | 2948 | 2,948 (761 units US, 2,187 units OUS) | Worldwide Distribution - USA nationwide, ARGENTINA, AUSTRALIA, AUSTRIA, BAHRAIN, BRAZIL, CANADA, CHILE, CHINA, Croatia, CZECH REPUBLIC, DENMARK, EGYPT, EL SALVADOR, FINLAND, FRANCE, GERMANY, GREECE, GUATEMALA, HONG KONG, INDIA, IRAQ, ISRAEL, ITALY, JAPAN, KAZAKHSTAN, REPUBLIC OF KOREA, MALAYSIA, NETHERLANDS, NEW ZEALAND, OMAN, PERU, POLAND, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SPAIN, SWEDEN, SWITZERLAND, SYRIAN ARAB REPUBLIC, TAIWAN, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, VENEZUELA, LEBANON, REPUBLIC OF MALDOVIA, COSTA RICA, GABON, INDONESIA, JAMAICA, NIGERIA, NORWAY, NEPAL, PAKISTAN, PANAMA, PHILIPPINES, SENEGAL, SOUTH AFRICA, THAILAND, TUNISIA, and CAMBODIA. | Terminated | March 29, 2016 | 657 | 2948 | Computer | Not_Computer
Z-2075-2014 | 68395 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Siemens syngo RT Dosimetrist 2.7 system |  |  | May 15, 2014 | 2014 | Siemens Medical Solutions USA, Inc | An update to the Virtual Simulation software of the syngo RT Dosimetrist is necessary to address a safety issue. In special scenarios within the Structure Operations function the system did not store the selected name from the list of predefined names. After reloading of the data, this structure name had been changed by the system. Additionally, a performance issue was addressed where import of DI | Software design | Siemens sent a Medical Device Correction Field Safety Notice of Improvement letter to affected customers via USPS Certified Mail. The letter described the affected product, problem, and actions to be taken. |  | 7 | 7 | US Nationwide Distribution in the states of NH, TN, OH, VA, and TX. | Terminated | December 15, 2014 | 214 | 7 | Computer | Computer
Z-2071-2014 | 68353 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens RAPIDCommÂż Data Management System. |  |  | May 22, 2014 | 2014 | Siemens Healthcare Diagnostics Inc | RAPIDCommÂż 5.0 Screen Layout will display incorrect test names | Software design | Siemens Healthcare issued an Urgent Field Safety Notice by e-mail to all affected Siemens Healthcare Diagnostics Regional offices on May 22, 2014, both in the United States and elsewhere for communication with affected customers. This notice informs customers of the issue with the RAPIDCommÂż Data Management System (Version 5.0.A software correction will be issued to address the problem.Please retain this letter with your laboratory records. Complete and return the Field Correction Effectiveness Check attached to this letter within 30 days.We apologize for the inconvenience this situation has caused. If you have any questions, please contact your Siemens Customer Care Center or your local Siemens technical support representative.For further questions please call (781) 269-3000. |  | 6 | 6 systems | Worldwide Distribution - US including the states of OR and PA., and the countries of Denmark and UK. | Terminated | May 26, 2015 | 369 | 6 | Computer | Computer
Z-2059-2014 | 68582 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | V-Series Patient Monitoring System, Manufactured for Mindray DS USA, Inc. Used for monitoring of human physiological parameters. The V-Series Monitor also has the capability of performing IV Drug and Hemodynamic Calculations and interfacing with network devices. |  |  | May 19, 2014 | 2014 | Mindray DS USA, Inc. dba Mindray North America | ECG interpretation feature issues. False measurements reported when the algorithm cannot determine P wave or QRS axis. The Median Complex Report displays the representative beat 25% larger than actual signal input. When a patient date of birth is imported from an external source, the ECG interpretation algorithm uses the default age of 50, and interpretive statements may be incorrect. | Software design | Mindray issued "Urgent: V-Series Monitoring System Corrective Action" notices dated May 19, 2014, to customers that received the affected product via certified mail with return receipt verification. The notification informs the customer of the issues with the affected product and how to identify the affected product. The notification recommends that customers discontinue use of the 12-lead ECG interpretation feature until a software correction is available (anticipate August 2014). A point of contact is provided to customers in case they have questions.For further questions please call 1-800-288-2121. |  | 113 | 113 units (57 United States; 56 International) | Worldwide Distribution - US including the states of Louisiana, New Jersey, Montana, South Carolina and Wyoming. | Open, Classified |  |  | 113 | Computer | Computer
Z-2060-2014 | 68578 | No | GKP | Coagulation instrument. | Hematology | Hematology | 510(k) | ACL TOP (Base), automated coagulation laboratory instrument. |  |  | June 03, 2014 | 2014 | Instrumentation Laboratory Co. | Potential for sample misidentification. | Device Design | Consignees were notified via phone starting on June 3, 2014. On June 4, a USPS mailing was sent to all US/Canadian customers. On June 4, 2014, email notifications were sent to the IL Affiliates and International Distributors/Dealers (through the International Group) requesting them to translate the Field Safety Notice and provide notification to their customer base and secondary distributors.The Urgent Product Correction letter informs consignees of the issue and provides workaround instructions. Consignees are asked to complete a Mandatory Response Tracking Record and return it via Fax (781)861-4207 or email: ra-usa@ilww.com. For technical questions, consignees should contact the IL Technical Support Center at 1-800-678-0710. |  | 1074 | 1074 | Worldwide Distribution-USA (nationwide) including the states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WV, and the countries of Canada, Mexico, Argentina, Australia, Austria, Belarus, Belgium, Bolivia, Brazil, Brunei, Chile, China, Costa Rica, Croatia, Czech Republic, Denmark, Dominican Republic, El Salvador, Finland, France, Germany, Gibraltar, Great Britain, Greece, Guatemala, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Korea, Kuwait, Libya, Lithuania, Luxembourg, Malaysia, Mexico, Netherlands, New Zealand, Norway, Panama, Peru, Philippines, Poland, Portugal, Quatar, Romania, Russia, Saudi Arabia, Serbia, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, UAE, Uruguay, Venezuela, Vietnam. | Terminated | June 11, 2015 | 373 | 6647 | Computer | Not_Computer
Z-0620-2014 | 64807 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | The Centricity Laboratory System is intended to be an information system designed to support the clinical and administrative activities associated with the provision and utilization of clinical laboratory services and facilities, e.g., the storing and delivering of analytical results. It is a specially designed data program application (software), which is supplied for installation in existing mainframe or decentralized computers or a computer network. This software system will not be the software directly used to run, steer, or control any specific laboratory analyzer or equipment. It is intended for use in laboratory facilities, including a central laboratory or in a multiple laboratory environment servicing satellite laboratories, reference laboratories, clinics, etc. Centricity Laboratory System is not intended for direct patient contact. |  |  | March 12, 2013 | 2013 | GE Healthcare It | GE Healthcare is aware of a potential safety issue with the use of GE Centricity Laboratory Instrument Interface (IF) where free-text sent from IM is not being transcribed into free-text in iNET. The instrument may upload results to IF/iNET when order/request is not yet in IF/iNET. In such a situation, test level supporting free-text sent from the instrument is discarded (not captured). | Device Design | GE Healthcare sent an URGENT MEDICAL DEVICE CORRECTION letter dated March 12, 2013 to all affected customers. The letter notified customers of the problem, suggested that customers experiencing this problem should retransmit the result from the instrument once the order exists in iNET and verify that the item level free-text is present in iNET, and to contact their GE service representative to request or install the updated software. Customers with additional questions can contact the GE Healthcare IT Help Desk at 1-847-589-8496 or tgg.nahelpdesk@med.ge.com.For questions regarding this recall call 847-589-8496. |  | 5 | 5 units | Worldwide Distribution - US including MN and Internationally to India, Australia, Canada, and England. | Open, Classified |  |  | 5 | Computer | Computer