## Sheet1
Recall Number | Recall Event ID | Device Type | Product Code | Regulation Description | Medical Specialty | Review Panel | Submission Type | Trade Name/Product | Main Name | Recall Class | Date Posted | Year Posted | Recalling Manufacturer | Reason for Recall | FDA Determined Cause | Action | Consumer Instructions | Clean Quantity | Quantity in Commerce | Distribution | Recall Status | Termination Date | Time to Terminate | Merged Quantities | Fault_Class
Z-2316-2015 | 71300 | No | HXD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | SYMMETRY SURGICAL DISPOSABLE TOWEL CLAMP, REF 19-4137, STERILE EO, Rx only. |  |  | June 16, 2015 | 2015 | Symmetry Surgical, Inc. | Lack of sterility assurance. | Package design/selection | The firm initiated their recall by via US certified mail, with return receipt requested, beginning on 06/18/2015. The Urgent Medical Device Recall letter, dated June 16, 2015, identified the affected product, as well as the reason for the recall. The letter discussed the potential harm associated with the issue, and requested that the consignees examine their inventory and discontinue use immediately. Customers should promptly call (800) 251-3000 for a return material authorization (RMA) number to support the return of the impacted lots. Symmetry Surgical Customer Service should be contacted with questions or any other needs. |  | 2708 | 2,708 units | Worldwide Distribution -- US, including the states of CA, VA, TX, FL, OH, NY, MS, KY, UT, GA, PA, MO, IN, NV, AL, MI, MA, and LA; and, the countries of Germany and Mexico. | Open, Classified |  |  | 2708 | Not_Computer
Z-2331-2015 | 71712 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Biograph mCT S(40)-4R, Model Number 10248671 |  |  | July 14, 2015 | 2015 | Siemens Medical Solutions USA Inc. | Emission computed tomography systems and computed tomography xray systems with optional ceiling mounts may be missing a locking ring. | Packaging change control | The firm initiated their recall on 07/14/2015 by telephone to the sole US consignee. The update is scheduled for 07/15/2015. |  | 1 | 1 unit | Distributed in the state of Texas and the countries of Germany, France, Russia, Denmark, and Switzerland. | Terminated | March 03, 2016 | 233 | 12 | Not_Computer
Z-0556-2016 | 72760 | No | NJR | Streptococcus spp. serological reagents. | Microbiology | Microbiology | 510(K) Exempt | Granada Biphasic Broth |  |  | November 23, 2015 | 2015 | Biomerieux Inc | FDA has determined that the product requires 510(k) clearance to continue marketing/distribution in the United States. | No Marketing Application | The firm, Biomerieux, sent an "Urgent Product Removal Notice" dated November 23, 2015 to its consignees. The letter described the product, problem and actions to be taken. The consignees were instructed to confirm that the letter has been distributed and reviewed by all appropriate personnel, discontinue and discard any remaining inventory, store letter with your bioMerieux documentation and complete the attached Acknowledgement Form via fax to: (919) 620-3306 attention of: bioMerieux NA ComOps QA/RA.If you have any questions or concerns, please contact your local bioMÂżrieux Clinical Customer Service organization at (800) 682-2666. |  | 24 | 24 cartons | US Distribution to states of: CA, GA, IN, MA, MN, MO, NY, OH and PA. | Terminated | February 24, 2016 | 93 | 24 | Not_Computer
Z-2279-2015 | 71602 | No | JSO | Culture medium for antimicrobial susceptibility tests. | Microbiology | Microbiology | 510(k) | Hardy Diagnostics |  |  | June 26, 2015 | 2015 | Hardy Diagnostics | Hardy Diagnostics is recalling MacConkey Agar with Ciprofloxacin, Brain Heart Infusion Broth with Ciprofloxacin 1 ug, MDR Acinetobacter, BEA with Azide and Vancomycin, BEA Agar with Vancomycin, Gentamicin, and Amphotericin B, BEA Broth with Vancomycin, and VRE Broth due to lack of 510(k) clearance. | No Marketing Application | The firm initially called customers to inform them of the recall. The firm's phone script stated that they were discontinung the product and as part of the process were recalling anything that still might be within expiration date. The instructions provided via phone were to discard any remaining plates that are in stock in order for credit. The phone script stated that the firm was going to be sending something in writing. On 06/25-29/15 the firm sent out notification letters to customers to inform them that they were discontuing the product, and were requested anything remaining within expiration to be discarded. The firm requests that customers complete, sign, and fax back the endlosed form stating their compliance with the above action. The notification states if the laboratory was unable to be reached via telephone and would like a credit to contact the Customer Service Department at 800-266-2222, option 1. For any questions the firm states to contact Technical Services Department at 800-266-222, option 2 or via email at TechService@HardyDiagnostics.com.Follow-up calls will be made to customers who do not respond. |  | 900 | 900 each (90 pk/10) | U.S. Nationwide Distribution and the country of: Canada | Terminated | September 11, 2015 | 77 | 194840 | Not_Computer
Z-0543-2016 | 72681 | No | IMD | Hot or cold disposable pack. | Physical Medicine | Physical Medicine | 510(K) Exempt | Cardinal Health Tiny Toes; 11470-010T Infant heel warmer |  |  | November 20, 2015 | 2015 | Philips Electronics North America Corporation | Heel warmers may cause first and second degree burns. | Labeling design | Cardial Health was notified of the recall via letter on 11/20/2015. Cardinal Health notified all of their customers as of 12/1/2015. Philips instructed Cardinal Health do take the following steps: 1.) Check for and destroy any inventory of Cardinal Health Tiny Toes. 2.) Complete the Distributor Reply Form included in this letter to document this action. Fax to Philips Healthcare at +1 404-855-4900. Please do not return any product to Philips. 3.) Determine which of your customers may have affected product. 4.) Ask your customers to follow the instructions in the section entitled: ACTIONS FOR CARDINAL HEALTH CUSTOMERS/USERS TO TAKE. 5.) As you receive signed Customer Reply Forms from your customers, please forward a running monthly summary to us. Please email the summary to Jennifer Maglio at jennifer.maglio@philips.com. To maintain complete documentation, please retain your customers completed reply forms. Philips may request access to particular reply forms to respond to an audit inquiry.Actions for Cardinal Health Customers to Take: 1.) Discontinue use of Cardinal Health Tiny Toes. 2.) Check for any inventory of Cardinal Health Tiny Toes. 3.) Destroy all products on hand (eg. by activating or puncturing each unit). 4.) Complete and sign the Customer Reply Form provided in this letter or Cardinal Healths version thereof, to notify Cardinal Health of how many Cardinal Health Tiny Toes were destroyed. Submit your signed reply form only after you have completed all steps |  | 7327100 | 7,327,100 individual units | Moberly, MO | Terminated | August 07, 2017 | 626 | 7327100 | Not_Computer
Z-2322-2015 | 71550 | No | DJZ | Breath-alcohol test system. | Toxicology | Toxicology | 510(K) Exempt | Mission |  |  | June 17, 2015 | 2015 | Acon Laboratories, Inc. | Acon Laboratories, Inc. is recalling Mission Breath Alcohol Detector due to a labeling amendment. | Labeling False and Misleading | On 06/17/15 the firm sent out customer notification letters to their distributors. The letter identified the affected product, problem and actions to be taken. In the letter the firm requests that Mission Breath Alcohol Detectors be quarantined. The firm asks that distributor complete the "fax back" form to allow for the return of unexpired product for a replacement. Once the form is received the firm is going to issue a return authorization number as well as provide a return label.The contact noted in the letter is Qiyi Xie MD, MPH, Sr. Officer Regulatory and Clinical Affairs at 858-857-8011. |  | 1283 | 1,283 kits | U.S. Nationwide Distribution in the states of: NY, CA, FL, VA, LA, NJ, MN, PA, MI, OH, TX, VA, UT, AL, and GA | Terminated | December 08, 2015 | 174 | 1283 | Not_Computer
Z-2271-2015 | 71592 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | Fabius MRI Anesthesia Machine; Fabius MRI is an inhalation anesthesia machine for use in MRI environments in operating, induction and recovery rooms. |  |  | June 22, 2015 | 2015 | Draeger Medical, Inc. | the Fabius MRI or parts of the system were attracted by the magnetic field of the MRI. If the Fabius MRI is positioned too close to the MRI, its parts can be loosened as a result of the magnetic force of the MRI. | Labeling False and Misleading | Draeger Medical sent an Urgent Medical Device Recall Letter, dated June 2015, and Instructions for Use Supplement (one supplement for each Fabius MRI Anesthesia Machine at the facility) to end users. The letter identified the affected product, problem and actions to be taken. For questions regarding the recall letter contact Michael Kelhart at 1-800-543-5047 (press 1 at the prompt, then 2, then 32349). For questions regarding the operation and/or servicing of the DrÂżger Fabius MRI anesthesia machine contact DrÂżgerService Technical Support at 1-800-543-5047 (press 4 at the prompt). |  | 373 | 373 | US Nationwide Distribution | Terminated | May 05, 2017 | 683 | 373 | Not_Computer
Z-0476-2016 | 72726 | No | GDY | Nonabsorbable gauze for internal use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Dressol-X, Radiopaque D-502, NDC 11004-502-40, Dental Use Only, Active Ingredients: A.S.A. (Acetylsalicylic Acid) + Eugenol |  |  | November 18, 2015 | 2015 | Rainbow Specialty &amp; Health Products | Rainbow Specialty &amp; Health Products initiated a field action for Dressol-X and Dressit-X because the labeling does not currently list the percentage of drug components. | Labeling design | Consignees have been made aware of the recall by FAX, e-mail and via telephone calls on 11/18/2015. Requested that distributors contact their clients, by telephone initially and then by email with written correspondence requesting them to recall the product from their clients. |  | 2775 | 2775 | Nationwide. No foreign, govt/VA/military consignees. | Open, Classified |  |  | 2925 | Not_Computer
Z-2348-2015 | 71575 | No | KCY | Pneumatic tourniquet. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | 18inch Dual Port/Single Bladder Disposable Tourniquet Cuff, Zimmer Surgical, Inc. Tourniquets are intended to be used by qualified medical professionals to exert enough pressure on the arterial blood flow in a limb to produce a bloodless operating field. Single bladder tourniquet cuffs are generally used for operations lasting less than 90 minutes. |  |  | June 24, 2015 | 2015 | Zimmer Surgical Inc | The inside of the folded Instructions For Use (IFU) pamphlet was missing the usage instructions and warnings. | Labeling False and Misleading | Zimmer sent an Urgent Medical Device Correction letter dated June 30, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the notification, identify the affected product and replace the IFU with the enclosed corrected IFU.Customers were also instructed to complete the attached Response Form and return it via email to CorporateQuality.PostMarket@zimmer.com. Customers with questions were instructed to call 330-364-0989. |  | 61 | 61 units | Worldwide Distribution - US including AL, AR, AZ, CA, CO, CT, DC, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV &amp; WY and Internationally to Canada. | Open, Classified |  |  | 391 | Not_Computer
Z-0466-2016 | 72811 | No | GFA | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Zimmer Oscillating Flat - TS ST Hub; |  |  | December 08, 2015 | 2015 | Synvasive Technology Inc | Two lots of oscillating saw blades incorrectly identify the blade cutting width as 25mm instead of the correct specification of 19mm. | Labeling False and Misleading | Zimmer sent an Urgent Medical Device Removal letter on December 10, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed: I. Review this notification.2. You will receive copies of notification sent directly to hospital risk managers and surgeons in your territory. As necessary, review and facilitate understanding of this notification by those entities.3. Further identify hospital risk managers and/or surgeons in your territory who should also receive notification of this update, and supply the information to the entities you have identified.4. If after reviewing this notification you have further questions or concerns, please call 330-364-0989 between the hours of 8 a.m. and 5 p.m. EST, Monday through Friday.Letters to Rick Managers are similar and the actions requested by them are:1. Reviewing this notification and forwarding to all people within your organization that may utilize this product.2. Identifying and quarantining the affected product to prevent further distribution or use.3. Completing the attached Response Form and returning it via e-mail toCorporateQua1ity.PostMarket@zimmerbiomet.com. Returning affected product to:Zimmer SurgicalAttn: QA!RA Dept. - RecallPlease include a copy of the Response Form with the shipment.4. Upon receipt of affected product, Zimmer will credit your account for unused Zimmer Oscillating Flat Blade - TS ST Hub, 75mm x 19mm, xl.OO/l.27mm, lots 43054 and 45104 only. Please return a copy of the completed response form along with your returned product to ensure proper credit.Important: Please distribute this notification to all personnel within your organization who need to be aware. If you have further distributed affected product(s), please provide the customer 's information on the Business Response Form to Zimmer.Letters also included a FAQ sheet regarding the labeling error. |  | 14 | 14 | US Distribution to one location each in CT, MA and CA. | Open, Classified |  |  | 14 | Not_Computer
Z-2326-2015 | 71709 | No | HXC | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | S4C Occiput Torque Wrench F/Set Screw; used to tighten the S4 Cervical Spinal System clamping screws with a defined torque 2.8 Nm for tightening the clamping screw in the screw head/hook head/occipital plate/transverse connector and other connectors. |  |  | July 07, 2015 | 2015 | Aesculap, Inc. | A FW103R torque wrench that was improperly reassembled after reprocessing caused a screw to break during implantation due to too much torque. The firm did not include instructions for reassembly after reprocessing. | Labeling design | Aesculap sent Field Safety Notification Letters, dated July 7, 2015, to end users, sales consultants, and distributors to advise them about the issue. Enhancements to the instructions for use were made available on the firm's website. |  | 10 | 10 | Distributed US (nationwide) and to Canada. | Terminated | April 17, 2017 | 650 | 10 | Not_Computer
Z-0149-2016 | 72017 | No | LBI | Neonatal phototherapy unit. | General Hospital | General Hospital | 510(k) | Natus neoBLUE 3 Phototherapy system or Neonatal Phototherapy, Catalogue Part Number: 010066; Top Assembly Part Number: 001103 |  |  | August 05, 2015 | 2015 | Natus Medical Incorporated | The neoBLUE Instruction For Use (IFU) and Service Manuals for the neoBLUE 3 were revised to show the cleared intensity statement settings so the output intensity aligns precisely with the information listed in the cleared 510(k) submission. | No Marketing Application | Natus Medical starting sending the Product Labeling Change NeoBlue LED Phototherapy System letter, dated August 2015, to all affected customers. The firm also attached the electronic copy of the revised User Manual with the letter.Customers are informed that the user and service manuals related to intensity for the neoBLUE 3 system are revised. Also, while the neoBLUE system does allow the light output to be adjusted with the potentiometers to accommodate different treatment heights, the firm removed the phrase "adjustable to &gt;50 ÂżW/cmÂż/nm" from the specifications.If you would like a CD disk containing the updated User Manual and the Service Manual at no charge, please call Natus Technical Service at 888-496-2887. |  | 4560 | 4560 units | Nationwide Distribution. | Terminated | March 17, 2016 | 225 | 4560 | Not_Computer
Z-0449-2016 | 72729 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Label reading in part as: "NexGen System, Complete Knee Solution; Reference Number 5980-37-01; Tibial Component, Precoat, Stemmed Size 3; |  |  | December 02, 2015 | 2015 | Zimmer Manufacturing B.V. | Complaints were received from Thailand reporting that there was no white paper (Tyvek) seal on the inner package of P/N:00-5980-037-01 unit from lot #62460264 | Packaging | An Urgent Medical Device Recall Letter, dated 12/2/15, was sent to the firm's consignees. All distributors will be notified via electronic mail. Hospital risk managers, surgeons, and distributors with product will be notified via courier. |  | 179 | 179 | US Distribution and the country of Thailand. | Open, Classified |  |  | 179 | Not_Computer
Z-2811-2015 | 72139 | No | GXN | Preformed nonalterable cranioplasty plate. | Neurology | Neurology | 510(k) | Synthes Cranial Flap Tube Clamp and Crimping Device |  |  | August 31, 2015 | 2015 | Synthes (USA) Products LLC | The Synthes Cranial Flap Tube Clamp and Crimping Device for Cranial Tube Clamp had been labeled MR Safe although they do not have the testing protocols currently required to designate them as MR Safe. | Labeling False and Misleading | DePuy Synthes sent an URGENT:FIELD SAFETY NOTIFICATION / MEDICAL DEVICE LABELING CORRECTION letter, dated August 31, 2015 to affected customers. The letter identified affected product, problem and actions to be taken. Customers were instructed to:"Update and review the package insert (GP1345-D)." Discard outdated package insert noted in the table above." Forward this Field Safety Notification to anyone in your facility that needs to be informed, especially personnel that conduct MR testing." If the package insert has been forwarded to another facility, contact that facility." Complete the attached Verification Section (page 3 of this notification). Please include your name, title, address, telephone number and signature in the spaces provided. Send a copy of the completed Verification Section by, Fax: 877-884-7821 or Scan/email: Synthes3911@stericycle.com" If the Verification Form is answered on behalf of more than one facility and/or individual, please clearly indicate the name and address of the facility and/or individual in page three (3) of this notification." Maintain a copy of this notice.For questions call 610-719-5450. |  | 105882 | 105882 | US Nationwide and internationally, including Canada. | Open, Classified |  |  | 105882 | Not_Computer
Z-2740-2015 | 71149 | [6 empty] | DermaPen Pro Model, Micro Needling System, a motorized AC powered surgical instrument for dermabrasive use. Each motorized handpiece is packaged within a small box. Each disposable needle cartridge is packaged individually and then placed into boxes of 30 quantities per box. |  |  | February 05, 2015 | 2015 | Derma Pen, LLC | Derma Pen initiated a field communication for Dermapen after receiving an FDA warning letter for distributing a device without FDA clearance or approval. | No Marketing Application | Derma Pen sent a notification letter to its customers on February 5, 2015. The purpose of the notification was to inform customers that the FDA issued a warning letter and the steps Derma Pen was taking to address the FDA' s concerns.For further questions please call (800) 981-5402. |  | 1466 | 1,466 units | Worldwide Distribution - US (nationwide) Distribution including Puerto Rico and The Virgin Islands and to the countries of : Canada, Brazil, Colombia, Ecuador, Costa Rica, Australia, Hong Kong, Netherlands, Mexico, Romania, Sweden, Peru, Greece, United Kingdom and United Arab Emirates. | Open, Classified |  |  | 1466 | Not_Computer
Z-2744-2015 | 71940 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Biomet Parsippany, NJ. 4 0 MM X 14 MM Variable Screw C-Tek Maxan Anterior Cervical Plate System. Lot 375440. |  |  | August 05, 2015 | 2015 | Biomet Spine, LLC | Some of P/N 14-521614B Lot 375440 was mislabeled as P/N 14-521614 Lot 375440. | Labeling mix-ups | Zimmer Biomet sent an Urgent Medical Device Recall Notice dated August 5, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by customer. Customers were instructed to immediately locate and discontinue use of the affected product. This action requires the immediate location and discontinued use of the item and its immediate return to Zimmer Biomet Spine. Specifically, customers were REQUIRED to take the following steps:" Immediately locate and remove from circulation the items consigned/loaned to youraccount as identified on page two of this notice." Carefully follow the instructions on the enclosed "Response Form"." If you have further distributed these items to medical facilities, you MUST notify them of this action. This letter MUST be given to the person responsible for receiving recall notices. However, you are charged with the location and return of these items.Health care professionals and consumers may report serious adverse events or product quality problems with the use of this product to the FDA's MedWatch Adverse Event Reporting program either online, by regular mail, fax or phone:Online:Fax:Phone:Regular Mail:www. fda .gov /MedWatch/report.htm(800) FDA-0178(800} FDA-1088Use postage-paid FDA form 3500 available at www.fda.gov/Medwatch/getforms.htmand mail to Medwatch, 5600 Fishers Lane, Rockville, MD 20852-9787.Questions related to this notice should be directed to 303-501-8400 Monday through Friday, 8 a.m.to 5 p.m. MST.Zimmer Blomet |  | 7 | 7 parts | Nationwide Distribution including Distributed nationwide to OH, LA, TX, and GA. | Terminated | February 10, 2016 | 189 | 7 | Not_Computer
Z-2748-2015 | 71880 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | Elcam Medical Minimal Residual Volume Luer activated Swabbable stopcock (MRVLS) for pressure monitoring and intravenous fluid administration; Manufactured by: Elcam Medical ACAL, Israel. Sterile, Quantity 20. |  |  | July 19, 2015 | 2015 | Elcam Medical, Inc. | Punctured blister packages were detected during the packaging process at the manufacturing site. Potential for device contamination. | Packaging | Elcam Medical sent "Urgent: Stopcock Recall" notices and "Recall Return Response Forms" dated 7/19/2015 to their customers. The recall notification informed the customers of the reason for the recall and the actions they must take. Customers were instructed to cease distribution and quarantine any affected products; return to Elcam Medical the enclosed Recall Return Response Form; and notify their customers as per their recall procedures. |  | 7721580 | 7,721,580 units | Texas, New Mexico and Pennsylvania | Open, Classified |  |  | 7969680 | Not_Computer
Z-2762-2015 | 72019 | No | NAB | Nonresorbable gauze/sponge for external use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | XGAUZE Trauma Dressing, catalog number FIN-0002-01. |  |  | August 12, 2015 | 2015 | Revmedx Inc | The XGAUZE Trauma Dressing was recalled because the packages may acquire small tears or pinholes during shipping due to use of shipping containers that lacked adequate protective materials. If present, these small tears or pinholes may disrupt package integrity and render the device unsterile. | Package design/selection | RevMedx, Inc. sent the Urgent Medical Device Correction and Removal XGAUZE Trauma Dressing letter to US Customers on August 12, 2015.Customers are instructed to inspect packages for signs of damages or loss of vacuum. The affected packages will appear puffy and the dressing will be loose inside. The non-affected packages will have a good vacuum intact seal and the packaging material should contour snugly against the dressing.Customers are advised to return any packages appear to have damaged or loss of vacuum. Any undamaged packages may be used as indicated.Customers with question can call the firm at (503) 218-2172 and fax the Return Response Form via fax to (503) 218-2274 or send via email to info@revmedx.com. |  | 5004 | 5004 units | US Distribution to the states of : Arizona, Idaho, North Carolina, Tennessee and Virginia. | Terminated | September 17, 2015 | 36 | 5004 | Not_Computer
Z-0237-2016 | 72434 | No | HQL | Intraocular lens. | Ophthalmic | Ophthalmic | PMA | TEKIA Single Piece Hydrophilic Acrylic Intraocular Lenses, Model 811 |  |  | September 29, 2015 | 2015 | Tekia, Inc. | Tekia is recalling the Single Piece Hydrophillic Acrylic Intraocular Lenses (IOLs) because they may be labeled with the incorrect diopter power. | Error in labeling | Tekia sent an Field Safety Notice dated September 29, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers. Customers were advised that Tekia is recalling a limited a number of Single Piece Hydrophilic Acrylic IOLs. If the customer has any questions, the customer is instructed to contact a TEKIA Customer Service or QA Representative at (949) 699-1300. Customers with product complaints or adverse events regarding the use of Single Piece Hydrophilic Acrylic IOLs are instructed to inform TEKIA by phone, email, or fax. |  | 50 | 50 units | Internationally to Vietnam. | Terminated | December 03, 2015 | 65 | 50 | Not_Computer
Z-2782-2015 | 72158 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Fluorescence Imaging Procedure Kits used with the da Vinci Si and Xi Surgical Systems; |  |  | September 09, 2015 | 2015 | Intuitive Surgical, Inc. | B. Braun 2 gang 4-way stopcocks in the FluorescenceImaging Procedure Kit may experience damage to the sterile packaging during shipping. The damage canresult in tears to the packaging and has the potential to cause a breach in sterility of the supplied stopcocks. | Package design/selection | Intuitive sent an Urgent Medical Device Correction letters dated September 9, 2015 to all customers via traceable method. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to :1. Ensure that all affected personnel are fully informed of this notice. Forward this notice to your Risk Manager, OR Director, Purchasing Manager, BiomedicalEngineering staff, anesthesiologist and members of your medical staff who perform da Vinci Surgery procedures.2. Prior to use, open and inspect current supply of Fluorescence Imaging Procedure Kits. If kits include B. Braun stopcocks, discard stopcocks and replace with other stopcocks3. Continue to perform procedures using other stopcocks4. Continue to follow directions provided in Florescence Imaging Procedure Kit IFU (p/n 550976), If a sterile item in a Fluorescence Imaging Procedure Kit becomes compromised in any way, do not use the compromised item; replace it with a new item.5. Complete and return the attached Acknowledgment Form to Intuitive Surgicalusing the instructions provided.6. Please retain a copy of this notice near your Fluorescence Imaging Procedure Kit inventory.Questions or concerns should be directed to ClinicalSales Representative or contact Intuitive Surgical Customer Service at the numbers listed below:" (800) 876-1310, Option 3 (6 AM to 5 PM PST) or mail: customersupportservicesupport@intusurg.com |  | 1200 | 1200 kits | US Nationwide Distribution | Terminated | November 02, 2015 | 54 | 1200 | Not_Computer
Z-2796-2015 | 71933 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | 3.5mm Non Locking Cortical Screw, in the PediLoc Tibial Plate System |  |  | July 28, 2015 | 2015 | OrthoPediatrics Corp | The 00-0903-2510, 3.5mm x 10mm Non Locking Cortical Screw, Lot # 150192-J package may contain a different screw. The screw in the package is a 00-0903-2526, 3.5mm x 26mm, and is etched correctly. | Packaging process control | OrthoPediatrics has initiated a Voluntary RECALL of Part # 00-0903-2510, 3.5mm Non Locking Cortical Screw, Lot # 150192-J, initially by contacting customers via telephone, and then by e-mailing notification letters on 7/28/2015. Customers returned the 5 mislabled screws, which were repacked with correct labeling. All recall activities were complete as of 8/14/2015. |  | 5 | 5 | Domestic: TN, IAForeign: EnglandVA/DOD: None | Terminated | March 17, 2016 | 233 | 5 | Not_Computer
Z-0203-2016 | 72312 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | MaxAn 4.0mm and 4.5mm fixed screws, used in the MaxAn Anterior Cervical Plate System. The MaxAn Anterior Cervical Plate System is a decompression-based technique for cervical spine stabilization. |  |  | September 29, 2015 | 2015 | Biomet Spine, LLC | The packaging of 4.5 mm diameter MaxAn screws identifies/labels the screws as 4.0 mm diameter screws. | Labeling mix-ups | Zimmer Biomet Spine sent an Urgent Medical Device Recall Notice dated September 29, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.If the device has already been implanted, written confirmation of this will be required and the necessary parties will be contacted. They will store returned product in quarantine in an isolated location until they can undergo inspection to verify if they are conforming or not and then will be destroyed or retained. Necessary parts will be destroyed so that it is visually obvious they are no longer functional. For further questions call (303) 443-7500 ext.244. |  | 17 | 17 units | US Distribution to the states of : NY, MO, and FL | Open, Classified |  |  | 17 | Not_Computer
Z-2794-2015 | 71976 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Dynamic Locking Screw System (DLS) 3.7mm &amp; Dynamic Locking Screw (DLS) 5.0mm; intended for use in long bone fractures, the fixation of osteopenic bone, the fixation of osteotomies, and for the fixation of non-unions and malunions. |  |  | August 11, 2015 | 2015 | Synthes (USA) Products LLC | DLS pin breakage during planned implant removal, after uneventful and successful healing of the fracture. This notice to additional consignees identified after recall Z-0390-2014, related to complaints of breakage located at the bottom of the screws and under the locking screw head. | Labeling False and Misleading | Synthes sent an Urgent Notices of Medical Device Recall letter, dated August 11, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed to return affected devices and response forms.For further questions please call (610) 719-6500. |  | 40038 | 40038 | Nationwide Distribution | Open, Classified |  |  | 40038 | Not_Computer
Z-2793-2015 | 71957 | No | LHC | Powered heating unit. | Physical Medicine | General Hospital | 510(K) Exempt | Microtek Medical, Inc. ORS Fluid Warming and Slush Drapes, Model #'s ORS-100, ORS-110, ORS-130, ORS-300, ORS-301, ORS-320, ORS-321, ORS-325, ORS-330, ORS-331 |  |  | August 18, 2015 | 2015 | Ecolab Inc | The sterile packaging may contain small (less than 250 microns wide) channels in the seal of the pouch. If channels are present, the package integrity is compromised and there is a remote chance that a non-sterile drape may be used during a surgical procedure. | Package design/selection | Consignees were sent on 8/18/2015 an Ecolab "Urgent Medical Device Recall" letter dated August 18, 2015. The letter was addressed to Recall Coordinator or Risk Manager. The letter described the problem and the product involved in the recall. Advised consignees to quarantine all affected lots, and notify their customers of the recall which is to be carried out to users level. Requested consignees to complete and return the enclosed Recall Response Form. For questions they can contact Customer Service at 1-800-824-3027 or e-mail customerservice@microtekmed.com. |  | 5172816 | 5,172,816 units | Worldwide Distribution -- US, including DC and PR; and, the countries of SAUDI ARABIA, LEBANON, GERMANY, HONG KONG, SINGAPORE, CANADA, and JAPAN. | Terminated | August 09, 2017 | 722 | 5172816 | Not_Computer
Z-2800-2015 | 72038 | No | GDX | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | PenBlade Safety Scalpel, for the incision of human skin and tissue, and to cut or trim sutures and catheters up to 6Fr. Models PB-M-10, PB-M-11P, PB-M-15 |  |  | August 18, 2015 | 2015 | Zien Medical Technologies, Inc | Cracks were discovered in the PET tray of a small population of sterile product. These cracks create a risk of loss of sterile barrier for the product. Sampling and root cause analysis indicated that the risk could extend throughout the lots manufactured from September 2014-July 2015. This led to the decision to recall the product. | Packaging process control | All consignees have been notified via email with an attached copy of the recall letter. Telephone calls with be made to all customers and distributors that do not respond directly to email notification to ensure compliance to the recall. |  | 28284 | 28,284 | Worldwide distribution to U.S. nationwide, Germany and England. | Terminated | December 09, 2015 | 113 | 28284 | Not_Computer
Z-0198-2016 | 72398 | No | MJC | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Antibacterial Hydro Personal Catheter, Male 12 Fr, Product Code 63512, sold in boxes of 30 units |  |  | October 02, 2015 | 2015 | C.R. Bard, Inc. | Incorrect unit labeling. The case label and shelf box are correct, Product Code 63512 and Catheter Size 12 French. The unit label (catheter package) may state Product Code 63516 and Catheter Size 16 French. | Error in labeling | Bard sent a Customer Notification Letter dated September 29 to customers via FedEx with proof of delivery notice. The letter identified the product, the problem, and the action to be taken by the customer. The consignees were instructed to carry out the notification to the user level. Customers with questions were instructed to contact BMD at 1-770-784-6471. |  | 4170 | 4,170 each | Nationwide Distribution including CA, FL, IL, MO, NJ, OR, PA, TX &amp; WA. | Open, Classified |  |  | 4170 | Not_Computer
Z-2807-2015 | 70858 | No | FDT | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | EVIS EXERA II Duodenovideoscope OLYMPUS TJF TypeQ180V. |  |  | March 26, 2015 | 2015 | Olympus America Inc. | New and validated reprocessing procedures have been issued for the Olympus TJF-Q180V duodenoscope ("TJF-Q180V"), consisting of revised manual cleaning and high level disinfection procedures. | Labeling design | The firm, Olympus, sent an "Urgent Safety Notification" dated March 26, 2015, to direct consignees/customers. The notification described the product, problem and actions to be taken. The notification contained new reprocessing procedures. The customers were instructed to immediately implement new precleaning and high level disinfection procedure; ensure that all personnel involved in the reprocessing are completely knowledgeable and thoroughly trained on the new reprocessing instructions; complete and return the enclosed questionnaire via fax to (484) 896-7128 and implement the new cleaning instructions when the new brushes arrive on/about May 8, 2015.If you have question, contact our Technical Assistance Center (TAC) at 1-800-848-9024, option 1, Monday-Friday between 7 AM EST - 8 PM EST. For additional information on this matter, contact V.P., Regulatory/Clinical Affairs &amp; Quality Assurance at (484) 896-5688 or by email at laura.storms@olympus.com. |  | 6109 | 6109 | Worldwide Distribution: US (nationwide) and countries of: Canada and Mexico. | Open, Classified |  |  | 6109 | Not_Computer
Z-2808-2015 | 71757 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | Marketing brochures for the PowerPICC and PowerPICC SOLO catheters |  |  | July 20, 2015 | 2015 | Bard Access Systems | Bard Access Systems is recalling marketing brochures for the PowerPICCÂż and PowerPICCÂż SOLOÂż catheters due to incorrect indication for use contained in the brochure. | Labeling False and Misleading | Customer letter notification with reply form will be sent via Federal Express with delivery confirmation for US consignees on 7/20/15 informing them of the incorrect indication for use contained in the brochure. |  | 1811 | 1811 | Distributed to domestic and foreign conference attendees. | Open, Classified |  |  | 1811 | Not_Computer
Z-2389-2015 | 71678 | No | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker Orthopaedics Triathlon X3 Tibial Bearing Insert - CR |  |  | May 29, 2015 | 2015 | Stryker Howmedica Osteonics Corp. | Stryker has received a report which identified that a Triathlon Tibial Bearing Insert CS X3 Size #2 9mm (5531-G-209) was reported to be in a pack that was labelled as a CR insert (5530-G-209). | Labeling mix-ups | Stryker Orthopaedics sent a recall notification via email on 5/29/2015 to their local branches. Notification letters and Product Accountability Forms were sent to Stryker branches/agencies with return receipts on 7/2/2015. |  | 10 | 10 units | OH, SC, RI, WV | Open, Classified |  |  | 10 | Not_Computer
Z-2819-2015 | 72035 | No | KCY | Pneumatic tourniquet. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Zimmer, Reprocessed Zimmer Tourniquet Cuffs,34" Single Hose Single Bladder Brown (PLC), 34", Model (Reprocessor) Number: ZIM60-7075-106, Reprocessed by Sterilmed. |  |  | August 24, 2015 | 2015 | STERILMED, INC. | Sterilmed, Inc. is recalling its reprocessed Zimmer Tourniquet Cuff, Part #ZIM60-7075-106, lot number 1846885 due to mislabeling. The affected reprocessed product is labeled as single hose devices but the devices inside the packaging are dual hose devices. | Labeling False and Misleading | Sterilmed sent an "Urgent Medical Device Recall" letter dated August 24 2015. The letter was addressed to Materials Manger, Risk Manger and Operating Room Supervisor. The letter identified the affected product, problem and actions to be taken. Customers are requested to return all affected product, sign and return the enclosed Customer Response Questionnaire and the Customer Response Form per the instructions on the Forms. For questions contact Sterilmed's Customer Care Team at 888-541-0078. |  | 43 | 43 | US Nationwide Distribution in the states of CA, KY, MI, OH, UT and WI. | Terminated | March 11, 2016 | 200 | 43 | Not_Computer
Z-2820-2015 | 72132 | No | FZX | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Zimmer Guide Wires for orthopedic procedures. |  |  | August 27, 2015 | 2015 | Zimmer, Inc. | Product labeling of these sterile guide wires (external carton label and patient record label) states a 10-year expiration date. Package testing supports a 5-year expiration date for these same guide wires. | Package design/selection | Zimmer sent an "URGENT MEDICAL DEVICE RECALL" notifications dated August 27, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Your Responsibilities:1.Review the notification and ensure affected personnel are aware of the contents.2.Locate part number and lot number combinations listed in attachment 1 and quarantine the devices immediately.3.Carry out a physical count of all affected product in your territory, including hospitals in your territory with directsales as well as consignment, and complete the Inventory Return Certification of Acknowledgement (Attachment 2). Email a completed copy of Attachment 2 to corporatequality.postmarket@zimmerbiomet.com.4.Return the recalled product along with the completed Attachment 2. Clearly mark the outside carton of each product return shipment made as Recall. 5.Please notify Zimmer Biomet of any hospitals and surgeons that have consumed the product in the past 12 months that are not included in either the distribution history or inventory system data information. Supply the information forany hospitals and surgeons that you have additionally identified by using the provided spreadsheet Additional Account template. The template, and the list of known consignees from the distribution history and inventory systems data, will be emailed to you for completion and return to: corporatequality.postmarket@zimmerbiomet.com.6.If after reviewing this notification you have further questions or concerns please call the customer call center at 1-877-946-2761 between 8:00 am and 5:00pm ESTFor further questions about this recall, please call ( 1-800-447-5633). |  | 112568 | 112,568 units | Worldwide Distribution - US (nationwide) including Puerto Rico and Virgin Islands., and to the countries of Canada, Brazil, Algeria, Mexico, Austria, Taiwan, Belgium, Hong Kong, Croatia, Singapore, Denmark, Bahamas, Egypt, Panama, Finland, Dominican Republic, France, Guatemala, Germany, Australia, Ireland, Japan,Italy, Haiti, Jordan, Malaysia, Kenya, India, Lebanon, Luxembourg, Malta, Mauritius, Morocco, Netherlands, Norway, Poland, Portugal, Russia, Saudi Arabia, Serbia, South Africa, Spain, Sweden, Switzerland, United Kingdom and United Arab Emirates. | Terminated | March 15, 2016 | 201 | 112568 | Not_Computer
Z-2823-2015 | 72183 | No | OLO | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | P/N: 9735023, DRIVER 9735023 SOLERA 5.5/6.0 MAS; P/N: 9735024, DRIVER 9735024 SOLERA 5.5/6.0 MAS CAN; |  |  | September 14, 2015 | 2015 | Medtronic Navigation, Inc. | Specific use conditions can increase the possibility of a screwdriver tip deforming or breaking during surgery. | Labeling design | Medtronic sent an Urgent Medical Device Correction Notification letter dated September 14, 2015, to all affected consignees via FedEx. Revised Instructions for Use were provided to add additional warnings related to the careful inspection of the instruments and the importance of knowledge of operating procedures, patient selection, and product information. Consignees with questions were instructed to to call the local Medtronic Navigation Technical Services Representative or call 1-800-995-9709. |  | 2346 | 2,346 | Worldwide Distribution - US (nationwide) and Internationally to Canada; United Arab Emirates; Australia; Belgium; Switzerland; China; Germany; Denmark; Spain; Finland; France; Great Britain; Greece; India; Israel; Italy; Japan; Netherlands; New Zealand; Poland; Portugal; Sweden; Singapore; Thailand; Slovakia; Turkey; and South Africa. | Terminated | December 03, 2015 | 80 | 2346 | Not_Computer
Z-0105-2016 | 72107 | No | HEE | Obstetric anesthesia set. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | CooperSurgical LEEP REDIKIT-Used as a surgical preparatory kit for Loop Electrosurgical Excision Procedure |  |  | September 01, 2015 | 2015 | CooperSurgical, Inc. | Kit Label incorrect :Labeled with Lidocaine HCI drug cartridges Lidocaine HCl (2%) but contains vials Lidocaine HCl (1%) | Labeling Change Control | CooperSurgical issued an "Recall Notification" dated September 1, 2015, to all affected customers via FedEX with confirmed delivery receipt. The letter identified the product, the problem, and the action needed to be taken by the customer.Customers asked to discontinue use of affected products and complete the attached Acknowledgement and Receipt Form for a free replacement.Contact Product Surveillance at 203.601.5200 ext. 3300 |  | 305 | 305 units | US Distribution to the states of : AL, AZ, CA, GA, IA, IL, KS, LA, MA, MD, MN, MO, MS, NY, OH, PA, TN, TX, VA, and VT. | Terminated | January 29, 2016 | 150 | 305 | Not_Computer
Z-0167-2016 | 72345 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | 2.7 mm Locking Screw SLF-TPNG with T8 Stardrive Recess 12mm. The Synthes modular Mini Fragment LCP System is intended for fixation of fractures, osteotomies, nonunions, replantations, and fusions of small bones and small bone fragments. |  |  | October 06, 2015 | 2015 | Synthes (USA) Products LLC | Specific lots of the SS 2.7mm locking screw have been mislabeled as SS 2.4mm locking screw. | Packaging | Synthes sent an URGENT NOTICE: MEDICAL DEVICE NOTIFICATION UPDATED INFORMATION: CORRECTED LOT NUMBER SS 2.7mm Locking Screw, self-tapping letter dated January 13, 2014, to the affected customer.Please take the following actions:If you DO have any of the identified devices, please take the following steps:Call Synthes at 1-800-479-6329 to obtain a Return Authorization Number.Complete the Verification Section at the end of this letter by checking the appropriate box indicating affected product has been located. Also, please indicate the number of devices found and note the Return Authorization Number. Please include your name, title, telephone number and signature in the spaces provided.Return the Verification Form (page 3 of this letter) with the product to:Credit/Returns, Synthes.If you DO NOT have the identified product, please take the following steps:Complete the attached Verification Section at the end of this letter by checking the appropriate box indication that no affected product has been located. Please include your name, title, telephone number and signature in the spaces provided this return documentation acknowledges your receipt of medical device removal information.Return the documents to Synthes by: Fax: 1-866-792-5446 orScan/email: FieldAction@synthes.comIf you have any questions, please call 610-719-5450 or contact your Synthes Sales Consultant. For further questions please call 610-719-6500. |  | 98 | 98 | US Distribution to the state of : TN | Open, Classified |  |  | 98 | Not_Computer
Z-0189-2016 | 72052 | No | LHC | Powered heating unit. | Physical Medicine | General Hospital | 510(K) Exempt | Regard, Item Number 830500003, Sterile, CV0578C - CV PACK - Custom surgical kit for surgery. |  |  | August 25, 2015 | 2015 | Resource Optimization &amp; Innovation Llc | Custom surgical kits contain a component (ORS fluid warming and slush drapes) that was recalled due to compromised package integrity potentially affecting sterility assurance. | Packaging | On August 25, 2015, the recalling firm sent an email to their direct customers informing them of the recall. The email instructed their customers to examine their inventory and determine whether they had any of the affected devices. The recalling firm attached the recall notice they received from their supplier.On September 4, 2015, the recalling firm sent a second email clarifying that they were instructing their customer to discard the supplier component . The notice also asked the customer to respond with how many affected devices they had in their inventory at the time they received the notice. |  | 17 | 17 kits | Distributed in Texas and Louisiana. | Open, Classified |  |  | 506 | Not_Computer
Z-0165-2016 | 72216 | No | OIY | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Exprt Precision System: Revision Knee, Model Number 160-010-726/738 |  |  | September 16, 2015 | 2015 | Encore Medical, Lp | The labeling is missing the size/diameter information. | Labeling design | DJO Surgical sent an Urgent Field Safety Notice dated September 16, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The letter requested the consignees complete and return a response card and return affected devices for replacement. For further questions call (512) 832-6302. |  | 209 | 209 units | US Distribution to the states of : SC, WI, RI, NY, MO, AR, OK, GA, FL, AL, CA, AZ and CO. | Terminated | January 04, 2016 | 110 | 209 | Not_Computer
Z-0156-2016 | 72205 | No | HWE | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | Orthopedic | 510(K) Exempt | Smith &amp; Nephew MOSAICPLASTY DP- Disposable Harvesting System 4.5 mm, Sterile |  |  | September 10, 2015 | 2015 | Smith &amp; Nephew, Inc., Endoscopy Div. | Sterility of device is compromised due to breach in the packaging | Package design/selection | Smith &amp; Nephew issued Recall Notification Letter dated September 10, 2015 via overnight Federal Express. Customers were requested to inspect inventory and locate any unused devices from the above listed product and batch numbers, and quarantine them immediately. Complete the Inventory Return Certification Form on the following page, indicating the lot numbers and quantities that need to be returned and include your phone and fax number in the spaces provided. If you do not have product to return, please place an x in the column No product to return. Send the completed form to the Smith &amp; Nephew's Field Action Department via email at FieldActions@smithnephew. com or fax +1-901-566-7975 to obtain a return authorization (RA) number.Return any affected product to the address listed on the Inventory Return Certification Form. Please indicate the RA number on your return shipment. |  | 1287 | 1287 units | Worldwide Distribution - US Nationwide, Canada, Austria, Australia, Belgium, China, Spain, Finland, France, Great Britain, Greece, Germany, Hungary, India, Italy, Japan, Korea, Lithuania, Latvia, Netherlands, Poland, Portugal, Sweden, Thailand, Turkey, Ukraine and Venezuela | Open, Classified |  |  | 4483 | Not_Computer
Z-0119-2016 | 72123 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Zimmer Natural Nail CM Long. Intramedullary Fixation Rod. |  |  | September 03, 2015 | 2015 | Zimmer Gmbh | In about 5% of tested samples, the outer tray was found to be compromised. Risks: A damaged outer tray detected during surgery may cause a delay in surgery to obtain another implant. If a compromised packaging is not detected prior or during a surgery by the user, it may lead to periprosthetic infection. This may result in revision or multi-stage revision to treat the infection. | Package design/selection | On 9/3/2015, URGENT MEDICAL DEVICE RECALL notifications were sent via courier to the affected distributors and Hospital Risk Managers with instructions for locating and inspecting the affected devices and contacting Zimmer if damage is observed. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. If after reviewing this notification you have questions concerning this notification, please call 1-877-946-2761. |  | 57603 | 57,603 units | US Nationwide distribution. | Terminated | March 15, 2016 | 194 | 57603 | Not_Computer
Z-0122-2016 | 72120 | No | JJY | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Alere Cholestech LDX Multianalyte Controls, PN 88769 and 88773. |  |  | August 27, 2015 | 2015 | Alere San Diego, Inc. | Alere San Diego is recalling the Alere Cholestech LDX Multianalyte Control because the assigned control ranges for the Total Cholesterol (TC) and Triglyceride (TRG) analytes are incorrect. | Error in labeling | Customer notification letters dated 8/25/15 and 9/3/15 were sent to customers to inform them that Alere San Diego is recalling the Alere Cholestech LDX Multianalyte Control because the assigned control ranges for the Total Cholesterol (TC) and Triglyceride (TRG) analytes are incorrect. Customers are informed that the Alere Cholestech LDX Total Cholesterol and Triglyceride assays continue to perform as expected. However, if customers' facility has used the control lot to test Total Cholesterol or Triglycerides as part of their Quality Control program, Alere San Diego recommends that they verify previous results or perform control testing using the replacement Expected Value Card or replacement control materials. The letters also provide customers with actions to be taken. Customers with questions are instructed to contact: Alere San Diego, Inc. 9975 Summers Ridge Road San Diego, CA 92121 U.S.A. Phone: 877-308-8289, FAX: 866-333-9839, E-mail: Verifications.ts@alere.com. |  | 190 | 190 kits | Nationwide Distribution. | Terminated | December 22, 2015 | 117 | 190 | Not_Computer
Z-0124-2016 | 72330 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes 2.7MM/3.5MM VA-LCP Medial Distal Humerus Plates 10H/RT/186mm- Sterile and Synthes 2.7MM/3.5MM VA-LCP Extended Medial Distal Humerus Plates 10H/RT/189mm-Sterile |  |  | September 28, 2015 | 2015 | Synthes (USA) Products LLC | The 2.7MM/3.5MM VA-LCP Humerus Plates were found to be labeled incorrectly. Part number 02.117.410S was incorrectly etched with part number 02.117.610S | Labeling mix-ups | The firm, DePuy Synthes, sent an "URGENT NOTICE: MEDICAL DEVICE RECALL" letter dated 9/28/2015 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to"Call DePuy Synthes at 1-800-479-6329 to obtain a Return Authorization (RA) Number and replacement product(s)."Complete the Verification Section (page 3 of this letter), (even if you do not have the product) by checking the appropriate box indicating affected product has been located. Also, please indicate the number of devices found and note the Return Authorization Number. Please include your name, title, address, telephone number and signature in the spaces provided."Return the Verification Section (page 3 of this letter) with the product to:"Credit/Returns, DePuy Synthes, 1101 Synthes Avenue, Monument, CO 80132."Send a copy of the completed Verification Section by:"Fax: 844-294-7175 or"Scan/email: Synthes4672@stericycle.com"Note: If the Verification Section is answered on behalf of more than one facility and/or individual, please clearly indicate the name and address of the facility and/or individual on page 3 of the notification.If you have any questions, please call 610-719-5450. |  | 36 | 36 | US Distribution to states of:OH, IL, PA, WV, FL, AZ, CA, AL, ND, KY, MN, NJ, MI, GA, IN, TN, VA, NC, and LA. | Open, Classified |  |  | 36 | Not_Computer
Z-0132-2016 | 71260 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | ONE-LINK NON-DEHP Y-TYPE MICROBORE CATHETER EXTENSION SET NEEDLE-FREE IV CONNECTOR WITH NEUTRAL FLUID DISPLACEMENT |  |  | May 21, 2015 | 2015 | Baxter Healthcare Corp | Potential for a leak at the tubing to luer bond. | Packaging | Baxter sent their consignees a recall letter dated 05/21/2015 with the following instructions:Actions to betaken if product was purchased directly fromBaxter :1. Locate and remove all affected product from your facility. The product code and lot number can be found on the individual product or shipping carton.2. Contact Baxter Healthcare Center for Service to arrange for return and credit.Baxter Healthcare Center for Service can be reached at 888-229-0001 betweenthe hours of 7:00 am and 6:00 pm Central Time, Monday through Friday. Pleasehave your Baxter 8-digit ship-to account number, product code, lot number(s) and quantity of product to be returned ready when calling.3. Complete the Âżenclosed Baxter customer reply form and return it to Baxter byfaxing it to 224-270-5457 or scanning and e-mailing it to fca@baxter.com .Returning the Baxter customer reply form promptly will prevent you from receiving repeat notices.4. If you distribute this product to other facilities or departm ents within yourinstitution, please forward a copy of this communication to them.5. If you are a dealer, wholesaler, distributor/reseller, or Original EquipmentManufacturer (OEM) that distributed any affected product to other facilities,please conduct a consumer-level recall of the affected product that you distributed to customers.Action to be taken if product was purchased from a distributor:1. Locate and remove all affected product from your facility. The product code and lot number can be found on the individual product or shipping carton.2. Contact Baxter Healthcare Center for Service to arrange for return and credit.The Baxter Healthcare Center for Service can be reached at 888-229-0001between the hours of 7:00 am and 6:00pm central time, Monday through Friday.3. Pl ease note that the Baxter customer reply form is not applicable. If a reply form is provided by your distributor or wholesaler, please return it to the supplier according to the |  | 23600 | 23,600 | Domestic Only | Terminated | May 10, 2017 | 720 | 23600 | Not_Computer
Z-0152-2016 | 72112 | No | LJT | Subcutaneous, implanted, intravascular infusion port and catheter. | General Hospital | General Hospital | 510(k) | PORT-A-CATH Plastic Hub Needles, 0.9mm (20G) x 19mm (3/4 in.) 90-degree. REF 21-2287-24 |  |  | September 14, 2015 | 2015 | Smiths Medical ASD, Inc. | The label on the lid stock of the individual needle packages for a portion of Lot number 46262 contains the incorrect part number, part number description and product image. The part number should be 21-2287-24 Port-A-Cath Plastic Hub Bent Needle. | Labeling mix-ups | The firm, Smiths Medical, sent an "URGENT MEDICAL DEVICE RECALL" letter dated 9/14/2015 to customers. The letter described the product, problem, and actions to be taken. The customers were instructed to Inspect your inventory for and quarantine the affected product, complete and return attached Confirmation form via fax to: 1-651-747-4959 or email to: recall.response@smiths-medical.com within 10 days; upon receipt of completed form, customer representative will contact you and arrange for exchange of your unused affected product; send a copy of letter and confirmation form to your customers identified, replacement product will be arranged for your customers; provide email confirmation of your customer notification to Smiths Medical at recall.response@smiths-medical.com, and notify downstream customers who may have received the product. If you should have any questions regarding this information, please contact Smiths Medical's Customer Service Department at 1-800-258-5361. |  | 130 | 130 Cartons (1560 needles) | Worldwide Distribution: US (nationwide) to CA only and countries of: Australia, Belgium, Germany, Hungary, Italy, Sweden, and Switzerland. | Open, Classified |  |  | 130 | Not_Computer
Z-0234-2016 | 72386 | No | OUB | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | miraDry System consisting of the MD4000-MC Console, the MD4000-HP miraDry Handpiece, and the MD4000-BT miraDry bioTip and accessories. |  |  | October 09, 2015 | 2015 | Miramar Labs, Inc. | Update to labeling for the miraDry Console to remind users to only use approved lubricants in a thin layer to prevent potential burns and blisters. | Labeling design | Miramar Labs sent Medical Device Correction letters to all users on October 9, 2015 to inform them of the labeling clarification by UAPS First Class Certified mail or courier such as Federal Express. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed that the firm added a new warning to their labeling. The firm included an updated copy of the MD4000-MC User Manual with the new warning included and asked that customers destroy their older copy. Customers with questions were instructed to call Customer SService at 1-855-647-2379 or email info@miraDry.com |  | 277 | 277 units | Nationwide Distribution | Terminated | December 23, 2015 | 75 | 277 | Not_Computer
Z-0236-2016 | 72407 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes Small Notch Titanium Reconstructive Plate set (SNTRP). Ti PLATE/ROD/PLATE 231mm 6HOLES/5.0mm ROD/17HOLES. Orthopedic implant. |  |  | October 08, 2015 | 2015 | Synthes (USA) Products LLC | The Synthes SNTRP set was released for sale as a Trauma product in 1996 with specific trauma indications and contraindicated for certain spine applications. In 1999, a module was released containing a small offering of the SNTRP plates. Synthes determined that the product in the SNTRP module is being used in spine surgeries, which are a contraindicated use. | No Marketing Application | Synthes sent an Urgent Notice Medical Device Letter dated October 8, 2015. to all affected customer. The letter identified the product the problem and the action needed to be taken by the customer.Please take the following actions:If you DO have any of the identified devices, please take the following steps:"Call Synthes at 1-800-479-6329 to obtain a Return Authorization Number."Complete the Verification Section at the end of this letter by checking the appropriate box indicating affected product has been located. Also, please indicate the number of devices found and note the Return Authorization Number. Please include your name, title, telephone number, and signature in the spaces provided."Return the Verification Form (page 3 of this letter)If you DO NOT have the identified product, please take the following steps:"Complete the attached Verification Section at the end of this letter by checking the appropriate box indicating that no affected product has been located. Please include your name, title, telephone number, and signature in the spaces provided. This return documentation acknowledges your receipt of medical device removal information."Return the documents to Synthes by:"Fax: (866) 240-5364 or"Scan/email: Synthes4642@stericycle.comIf you have any questions, please call 610-719-5450 or contact your Synthes Sales Consultant. Thank you for your attention to this issue. |  | 20494 | 20,494 | US Distribution to the states of :MO, NJ, MN, MS, AZ, CA, LA, OK, VA, OH, AR, FL, PA, VT, IN, AL, WA and NY. | Open, Classified |  |  | 20494 | Not_Computer
Z-0244-2016 | 72518 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | Cook Medical CXI Support Catheter |  |  | October 21, 2015 | 2015 | Cook Inc. | Cook Medical has received six product complaints associated with reports of the 2.6FR curved catheter being mixed with the 4.0FR straight catheter and vice versa. Investigation revealed the products were mixedduring processing. | Labeling mix-ups | Cook Medical sent an Urgent: Medical Device Recall letter dated October 20, 2015, to all affected customers via courier. The recall notification included a description of the reason for the recall, affected product, customer responsibilities, and instructions for responding to the formal recall notification. Should you have any questions or concerns, please do not hesitate to contact Cook Medical Customer Relations at 1-800-457-4500 or 1-812-339-2235. |  | 38 | 38 | Worldwide Distribution - US Distribution to the states of : CA, FL, ID, IL, MA, NY, OK, TX, UT and VA., and to the countries of Canada, Austria, Germany, Spain, France, Hungary, Sweden and UK. | Open, Classified |  |  | 38 | Not_Computer
Z-2730-2015 | 71832 | No | GWN | Nystagmograph. | Neurology | Neurology | 510(k) | Neuro Kinetics, Inc. (NKI) I-PortalÂż NOTC system with the OVAR research test; a rotary chair system used in vestibular and neuro-otologic diagnostic testing. This device provides no diagnoses nor does it provide diagnostic recommendations.The OVAR test has not yet been cleared by the FDA, and until then can only be used as an investigational device. |  |  | July 21, 2015 | 2015 | Neuro Kinetics, Inc. | The devices that included the OVAR research test were not cleared for marketing and mTBI research package was sold with the incorrect investigational label. | No Marketing Application | Neuro Kinetics sent notification letters to end users on/about July 24, 2015, to alert them to the reason for recall and inform them of the corrective action planned by Neurokinetics.Action to be taken by you the Customer/User:If you have already initiated OVAR research (an IRB approved study with non-significant risk designation is ongoing), please reply to this email and include a copy of the IRB approval letter or the IRB approval information (e.g., investigators name, IRB name, IRB location, and study reference number). In return, NKI will mail you a field correction package with the necessary and appropriate Investigational Device labeling.If you have not initiated OVAR research (an IRB approved study is not ongoing), please discontinue use of the OVAR test and reply to this email to confirm that you are not using the OVAR test as part of an IRB approved study. In return, NKI will mail you a field correction package to disable the OVAR test functionality, which can be restored at the time you have an IRB approved study. Disabling the OVAR research test will have no effects on the functionality of the clinical tests included with your NOTC system. There will be no perceivable changes to the system outside of not being able to run the OVAR test.For questions regarding this recall call 412-963-6649. |  | 45 | 45 | Worldwide Distribution - US (nationwide) | Open, Classified |  |  | 45 | Not_Computer
Z-2405-2015 | 70741 | Yes | MPB | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | RenLane Renal Denervation Catheter; Cat No.: D135601. |  |  | April 15, 2014 | 2014 | Cordis Corporation | Potential Damage may occur to the helical tip of the RENLANE Renal Denervation Catheter, if the hemostasis valve of the arterial access device is not fully open before insertion or withdrawal of the catheter. | Labeling Change Control | Cordis sent an Urgent Field Safety Notice dated April 15, 2014 providing additional labeling to their consignees in Germany only, with a follow-up letter on August 4, 2014. Cordis explained When deploying the catheter via a compatible arterial access device, if a Tuohy-Borst type adjustable hemostasis valve is used, ensure the valve is fully open before insertion or withdrawal of the catheter. This is especially important, since the helical tip at the distal end of the RENLANE Renal Denervation Catheter has a larger diameter compared to the proximal catheter shaft. Although the event was not a Removal, all lots involved have expired. The product is no longer commercialized, and was never commercialized beyond the 10 customers in Germany. |  | 98 | 98 units | Distributed in Germany only. | Terminated | August 26, 2015 | 498 | 98 | Not_Computer
Z-2422-2015 | 71911 | No | DTS | Cardiopulmonary bypass cardiotomy return sucker. | Cardiovascular | Cardiovascular | 510(K) Exempt | IBC Suction Wand Catalog Number 1990S |  |  | August 07, 2015 | 2015 | International Biophysics Corp. | It was discovered that the heat seal placed onto sterile pouches may not meet strength acceptance criteria. | Packaging process control | The recalling firm began notifying their affacted customers on 7/2/2015 |  | 100 | 100 units | Worldwide - United States Nationwide and the countries of: Germany, Norway, Turkey, Chile, Pakistan, and Republic of South Africa | Terminated | March 01, 2016 | 207 | 11201 | Not_Computer
Z-2420-2015 | 71897 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Medtronic EOPA Arterial Cannula, Model # 77620. Lot: 2015060544 intended to deliver oxygenated blood to the heart during cardiopulmonary bypass procedures for up to 6 hours in duration. |  |  | July 24, 2015 | 2015 | Medtronic Perfusion Systems | EOPA Arterial Cannula Devices in this lot were shipped without the guidewire. | Packaging process control | On July 24, 2015 Medtronic sales representatives began verbally notifying the customers of the issue, requesting quarantine and return of un-used, affected productConsignees were sent a Medtronic "Urgent Medical device Recall" letter dated August 2015. The letter was addressed to Risk Manger and Health Care Professional. The letter described the problem and the product involved in the recall. Advised consignees to quarantine and remove the product from their inventory and return it to Medtronic. Requested consignees to complete and return the Customer Confirmation Certificate via email to RS.CFQFCA@Medtronic com or fax to Medtronic 651-367-0612 to the attention of Customer Focused Quality. For questions contact your Medtronic Sales Representative or Lifeline Technical Services at 1-877-526-7890. |  | 221 | 221 (111 US, 110 OUS) | Worldwide distribution. US nationwide including: FL, GA, IL, MI, NY, NC, PA, WV; Germany, Hong Kong, and Italy. | Terminated | April 04, 2016 | 255 | 221 | Not_Computer
Z-2431-2015 | 71841 | No | LSL | Neisseria spp. direct serological test reagents. | Microbiology | Microbiology | 510(k) | Xpert CT /NG Vaginal/Endocervical/ Specimen Collection Kit: |  |  | July 27, 2015 | 2015 | Cepheid | Firm determined that some lots of CT/NG Swab 50 and CT/NG Urine collection kits contain a sub component with an expiry date that is earlier than the kit expiration date. | Incorrect or no expiration date | Notification letters dated July 28 and 29th to all affected customers via Federal Express.. |  | 776 | 729 kits US, 47 kits ROW. | Worldwide distribution. US nationwide, including North Mariana Islands, Aruba, Canada, Chile, Colombia, Costa Rica, Cyprus, Guam, Guatemala, Italy, Malawi, Micronesia Federated States, Panama, San Marino, South Africa, Spain, Switzerland, and Zimbabwe. | Terminated | September 28, 2015 | 63 | 2607 | Not_Computer
Z-2427-2015 | 71676 | No | KXE | Wrist joint ulnar (hemi-wrist) polymer prosthesis. | Orthopedic | Orthopedic | 510(k) | uHead, Ulnar Implant, Standard Stem, Size 1, Sterile, Rx Only. Intended for replacement of the distal radioulnar joint following ulnar head resection arthroplasty. |  |  | June 24, 2015 | 2015 | Stryker Howmedica Osteonics Corp. | Stryker is recalling rHead, uHead, Sigmoid Notch, Remotion, Radio Capitellum because packaging integrity (sterile barrier) of the packaging type KIT I may potentially be compromised by transportation. | Packaging change control | Branches/Agencies were notified by Stryker of this action by e-mails on June 24, 2015 and an Urgent Medical Device Removal Letter/Acknowledgement Response Form dated June 25, 2015 was sent to the attention of the Risk Manager. |  | 16992 | 16992 total | Worldwide Distribution. | Open, Classified |  |  | 16992 | Not_Computer
Z-2443-2015 | 71784 | No | MPO | Hot or cold disposable pack. | Physical Medicine | Physical Medicine | 510(k) | Covidien Argyle Infant Heel Warmers (no tab) |  |  | July 31, 2015 | 2015 | Covidien LLC | Infant Heel warmers may break during activation resulting in spillage of the contents on patients or healthcare worker causing potential skin/eye irritations | Package design/selection | Medtronic/Covidien issued an Urgent Medical Device recall letter dated July 31, 2015 via mailed Certified. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to quarantine and discontinue use of the affected products.. For questions or concerns,contact Medtronic representative or Customer Service at (800)-882-5878. |  | 68800 | 68,800 units | Worldwide Distribution - US Nationwide and the countries: Canada, Mexico, Panama, and Japan. | Open, Classified |  |  | 233549 | Not_Computer
Z-0395-2016 | 72436 | No | HCC | Biofeedback device. | Neurology | Neurology | 510(K) Exempt | The "Pacifier Activated Lullaby" (PAL) System composed of three major components: 1) the PAL Player Module (aka Player Unit), 2) the PAL Pacifier Sensor, and 3) power supply. The system is designed to encourage and reinforce effective non-nutritive sucking of premature infants. |  |  | October 01, 2015 | 2015 | Powers Medical Devices, LLC | PAL unit has undergone significant technological changes since the previous 510(k) and labeling is not consistent with the cleared indications of the 510(k). | No Marketing Application | Powers Medical Devices, LLC intends to notify customers via certified U.S. Mail and supply the updated Instruction Manual via U.S. Mail or other tracked service. Instructions to customers include: 1) destroy the previous Instruction Manual, and replace it with the new one provided, 2) attach the new product label over the existing labeling without covering up the Model # and serial # of the device, 3) certify the destruction and replacement of the Instruction Manual nd replacement of the product label to the company using the form and self-addressed stamped envelopw provided. |  | 128 | Total of 128 units | Nationwide Distribution. | Open, Classified |  |  | 128 | Not_Computer
Z-0389-2016 | 72534 | No | DXE | Embolectomy catheter. | Cardiovascular | Cardiovascular | 510(k) | LeMaitre NovaSil Silicone Single Lumen Embolectomy Catheter 6F, 80 cm, Model 1601-68. Indicated for the removal of arterial emboli and thrombi. |  |  | October 21, 2015 | 2015 | LeMaitre Vascular, Inc. | Manufacturing issue that could cause a pinhole in the formed end of the packaging, which may lead to compromised sterility. | Packaging process control | Consignees were notified of the recall on October 21, 2015 via letter. Consignees were asked to do the following: Identify all catheters with the affected LOT and model numbers in inventory; quarantine all unused catheters from the affected LOTs and record this action in the attached form; send the form via mail, email or fax to Customer Service who will then issue an RMA-number (Return Merchandise Authorization number) for the return shipment and replacement of the products; pass this notice on to all those who need to be aware within your organization and to any organization to which the potentially affected devices have been transferred. For questions concerning this safety notice, please contact Salva Kozin, Quality Assurance Engineer, at 781-221-2266 ext. 183. |  | 4315 | 4315 total | Worldwide Distribution. US nationwide including PUERTO RICO, and Australia, AUSTRIA, BELGIUM, COLOMBIA, COSTA RICA, Czech Republic, DENMARK, EGYPT, FRANCE, GERMANY, ICELAND, Italy, JAPAN, JORDAN, Latvia, LUXEMBOURG, MOROCCO, NETHERLANDS, Norway, Oman, PERU, PORTUGAL, Poland, Slovenia, Spain, Switzerland, Sweden, THAILAND, United Kingdom, United Arab Emirates, and VIETNAM. | Open, Classified |  |  | 4315 | Not_Computer
Z-2451-2015 | 71876 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes Trauma Ex-Fix System (Small, Medium, Distraction Osteogenesis (DO) and Large) ; (S) intended to stabilize and provide treatment for fractures of the small bones, such as the hand, wrist, forearm, foot, and ankle; (M) for the construction of an external fixation frame for the treatment of pediatric and adult fractures; )L) for use to provide treatment for long bone and pelvic fractures that require external fixation; and (DO) for fracture fixation (open and closed),pseudoarthrosis or non-unions of long bones, limb lengthening by epiphyseal or metaphysical distraction, correction of bony or soft tissue deformities and correction of segmental bony or soft tissue defects in adult and pediatric patients. |  |  | July 29, 2015 | 2015 | Synthes (USA) Products LLC | Labeling changes were made related to MR conditions that align with standard definitions for MR Safe, MR Conditional, and MR Unsafe. This recall identifies additional Trauma External Fixation System part numbers that should be considered MR Conditional. One serious injury reported. | Labeling design | Urgent recall letters, dated July 31, 2015, were sent to end users and sales consultants to notifiy them about the issue, the associated risks, and to provide instructions for returning the affected devices and response form. |  | 5306080 | 5306080 | Worldwide Distribution-including US (nationwide) and Canada. | Open, Classified |  |  | 5306080 | Not_Computer
Z-2458-2015 | 71889 | No | OFF | Venous blood pressure manometer. | Cardiovascular | Cardiovascular | Enforcement Discretion | Central Line &gt;34 Weeks Kit . Cardiology department. |  |  | August 03, 2015 | 2015 | Medline Industries Inc | The print plate artwork for this lot was found to state: Central Line &lt;34 Weeks Kit displaying an incorrect sign in front of the 34. The description for this device should have stated Central Line &gt;34 Weeks Kit. The defect may lead to the product to be used on patient population that is &lt;34 weeks rather than intended population of &gt;34 weeks. | Error in labeling | Medline Industries contacted the account directly via phone on August 4, 2015, they confirmed no affected product is in their facility. The lot numbers are easily identified on the outside of each kit. Second and third follow up communications will not be sent as we have confirmed the facility does not have any affected product.For questions regarding this recall call 847-643-4129. |  | 12823 | 320 units for lot 14DB9277 &amp; 320 units for lot 14PB2878 | US including PA | Terminated | January 05, 2017 | 521 | 12823 | Not_Computer
Z-2471-2015 | 64928 | Yes | KDJ | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | MiniCap Extended Life PD Transfer Set Extra Short (4") with Twist Clamp. Product Code: 5C4483. |  |  | February 25, 2013 | 2013 | Baxter Healthcare Corp. | Labeling Revision: Additional instructions for the use of peritoneal dialysis transfer sets including the sanitization the connection, healthcare provider monitoring of thyroid function, caution against reuse or reprocessing of the devices, warning against use if tip protectors are not in place, and advisement that the product does not contain natural rubber latex. | Labeling design | Baxter sent an IMPORTANT PRODUCT INFORMATION letters dated February 25, 2013 to affected customers via USPS. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete the attached reply form and return it to Baxter by faxing it to 224-270-5457 of scanning the form and e-mailing it to fca@baxter.com. For general questions about this communication can be answered by contacting The Center for One Baxter at 800-422-9837 (Monday - Friday, 8:00 AM - 5:00 PM CST). Clinical questions can be answered by calling Baxter's Renal Division Clinic Help line at 888-736-2543 x 2 (Monday - Friday, 8:00 AM - 5:00 PM CST). |  | 260244 | 260,244 units | Worldwide Distribution - US: Nationwide + American Samoa + District of Columbia + Guam + Northern Mariana Islands + Puerto Rico; *** FOREIGN: Antigua, Argentina, Australia, Bahamas, Barbados, Bermuda, Brazil, Canada, Chile, Colombia, Cuba, Curacao, Dominican Republic, Ecuador, El Salvador, France (New Caledonia), Guatemala, Guyana, Honduras, Hong Kong, India, Indonesia, Jamaica, Japan, Macau, Malaysia, Mexico, New Zealand, Nicaragua, Panama, Peru, Singapore, Suriname, Taiwan, Thailand, Trinidad and Tobago, United Kingdom (Cayman Islands), Uruguay, Venezuela | Terminated | March 08, 2016 | 1107 | 1543011 | Not_Computer
Z-2468-2015 | 71805 | No | NLM | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Stryker |  |  | July 10, 2015 | 2015 | Stryker Sustainability Solutions | Stryker Sustainability Solutions has received reports indicating that devices from one lot (lot# 3574704) were labeled as 100mm models (CB5LT) while the devices inside of the packaging were 75mm models (CB5ST). | Labeling mix-ups | On 07/10/15 the firm sent out notification letters to all customers. The letter states for the customer to discontinue use of the affected lot, and to use other trocar lots reporcessed by the firm. The firm requests that the Recall Effectiveness Check Form be completed. The firm is going to be sending a sales representative to the facilities to check inventory for the affected trocars. A Recall Effectiveness Check will also be completed. If the form indicates that affected devices remain in inventory a prepaid shipping label will be issued for the return of the product. Customers will receive credit for all affected devices returned. The form must be completed even if no affected devices are found in inventory. Stryker Sustainability Solutions Complaint Hotline: 1-888-888-3433 x5555 |  | 90 | 90 devices | U.S. Nationwide Distribution to the following states of:: NV, NM, TX, and MI. | Terminated | December 04, 2015 | 147 | 90 | Not_Computer
Z-2476-2015 | 71735 | No | NGT | Intravascular catheter. | General Hospital | General Hospital | 510(k) | 0.9% Sodium Chloride Injection, USP, 10mL in 12mL., Single use, Rx only, Sterile Solution. |  |  | July 02, 2015 | 2015 | MRP, LLC dba AMUSA | Procedures for the acceptance and control of in-process product have not been adequately established. | Packaging process control | The firm notified their consignees of the recall by Fed Ex on 07/02/2015. The firm requested the return of the product. Distributors were instructed to notify their customers. |  | 551 | 551 cases | Distributed in the states of CA. PA, VA, TX, MA, MD, WV, UT, FL, and TN. | Terminated | April 03, 2017 | 641 | 551 | Not_Computer
Z-2478-2015 | 71970 | No | MTD | Electrode recording catheter or electrode recording probe. | Cardiovascular | Cardiovascular | 510(k) | FIRMap 60mm Catheter; |  |  | August 12, 2015 | 2015 | Abbott Electrophysiology | The expiration date was incorrectly printed on the pouch/box labels of 19 products that were shipped to customer sites. The manufacturing date was accidentally used (copied) for the expiration date. | Incorrect or no expiration date | The firm, Abbott Electrophysiology, sent an "URGENT PRODUCT RECALL" letter dated August 11, 2015. A customer recall notification letter was distributed to firm representatives who will visit each consignee and inspect inventory for affected product. Letters were provided to reps on August 12, 2015. The letter describes the product, problem and actions to be taken. The customers were informed that the product they have in stock will be return to the firm and will be provided a replacement product free of charge.If you have any questions, contact Quality Director at 650-681-1744. |  | 86 | 86 units | US Distribution in states of: KY, IN, MO, NY, OH, TX, LA, IL, PA, NJ, WA, VA, FL, GA, CT, CA, AZ, CO, and TN. | Terminated | September 23, 2015 | 42 | 86 | Not_Computer
Z-0332-2016 | 72416 | Yes | NYU | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Palindrome Precision Chronic Catheter Kit Symmetrical Tip, Tal VenaTrac" Stylet 14.5 Fr/Ch (4.8mm) X 33 cm |  |  | October 14, 2015 | 2015 | Covidien LLC | Incorrectly packaged with a 90 degree bend at the distal tip | Packaging process control | The firm, Medtronic (Covidien), sent an "URGENT MEDICAL DEVICE RECALL" letter dated 10/16/2015 to their customers. The letter described the product, problem, and actions to be taken. Customers were instructed to quarantine and discontinue use of the product, return affected product and complete and return the RECALLED PRODUCT RETURN FORM via Fax to Medtronic to (800)-895-6140 or email to: feedback.customerservice@Covidien.com/Distributor to (203) 492-7719 or email to: PalCathFCA@Covidien.com.Questions or concerns contact your Medtronic representative or Customer Service at (800) 882-5878. |  | 150 | 150 units | Worldwide distribution: US (nationwide) and countries of: Canada, Belgium, Chile, Panama, Denmark, France, Germany , Netherlands,Spain, Switzerland, and United Kingdom. | Terminated | June 02, 2017 | 597 | 965 | Not_Computer
Z-0328-2016 | 72197 | No | JJX | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | CONSULT(TM) diagnostics hCG CONTROLS Kit, Part/Reorder Number 5011, IVD |  |  | September 14, 2015 | 2015 | Biochemical Diagnostics Inc | CONSULT diagnostics hCG Controls have discrepant storage temperatures listed on the kit labels when compared with the individual vial labels and the package insert. | Labeling False and Misleading | Biochemical Diagnostics sent a Notice of Correction Letter (dated August 26, 2015) to customers via USPS 1st class mail. The letter identified the affected product, problem and actions to be taken. Customers were instructed to return the verification form, per the notification. For questions contact Alere Technical Service at 866-216-0094 or by email at verifications.at verifications.ts@alere.com. |  | 4473 | Domestic: 4,473 kits | US Nationwide Distribution | Terminated | January 26, 2016 | 134 | 4473 | Not_Computer
Z-0322-2016 | 72537 | No | NTR | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | Thermo Scientific p21 WF1 Ab-3 (DCS-60.2) |  |  | October 27, 2015 | 2015 | Lab Vision Corporation | It was discovered that the labeling of MS-230 antibody incorrectly describes the concentration - label reads 0.2mg/ml but it is actually 0.4 mg/ml. | Labeling design | The firm, ThermoFisher Scientific, sent an "URGENT MEDICAL DEVICE RECALL-Lab Vision Corporation" letter dated 10/29/2015 to its customers on 10/30/2015. The letter describes the product, problem and actions to be taken. The customers were instructed to Upon identification of the affected lot, notify Lab Vision of the quantity previously used and the quantity still in inventory; review any associated test results to ensure an accurate diagnosis was able to be made; If routine laboratory use of this product follows the Lab Vision Data Sheet, the antibody may safely be used as is; hold any affected material in quarantine until the new labels are received-affected product may also be returned to Lab Vision, and complete and return the attached Recall Return Response Acknowledgement &amp; Receipt Form via email to: sarah.rickert@thermofisher.com or Fax to: 269-372-2674, Attn: Quality Assurance &amp; Regulatory Affairs Manager. A new label with the correct concentration can be sent upon request, for relabeling of the product at the customers facility.If you have any questions, please contact Quality Assurance &amp; Regulatory Affairs Manager at 269-544-5628 or sarah.rickert@thermofisher.com. |  | 5 | 5 vials | Worldwide Distribution: US distribution to states of: CA and WA, and countries of: Serbia, Germany and Netherlands. | Open, Classified |  |  | 5 | Not_Computer
Z-0293-2016 | 72451 | No | FRO |  |  | General &amp; Plastic Surgery | 510(k) | Neptune Plus, HEMOSTASIS PAD WITH ANTIMICROBIAL BARRIER, For hemostasis following needle/sheath removal, Part number 8870-03 |  |  | October 14, 2015 | 2015 | TZ Medical Inc. | The IFU of the Neptune Plus Pads, does not match that submitted in the original 510k. Need to remove: for haemostasis following needle/sheath removal | Labeling False and Misleading | TZ Medical Inc. began sending the Urgent: Medical Device Correction 8870-03 Neptune Plus letter, dated 10-7-2015, to consignees in October 2015. The firm is notifying all users of Neptune Plus wound dressing to a correction to the Instructions for Use (IFU). Consignees should discard the IFU in your current inventory and replace with new IFU. Consignees with questions should call TZ Medical Inc. at (503) 639-0282. |  | 32170 | 32,170 units | US nationwide distribution. | Open, Classified |  |  | 32170 | Not_Computer
Z-0279-2016 | 72552 | No | PAH | Surgical mesh. | General &amp; Plastic Surgery | Obstetrics/Gynecology | 510(k) | MiniArc Pro Single incision Sling System, Model number / Product Reference 9000261, Lot 936405xxx. Designed to treat female stress urinary incontinence. |  |  | October 26, 2015 | 2015 | Astora | The lot number on the outer carton label did not match the lot number printed on the tray label or identification labels inside the box. The identification labels inside the box are placed on medical records and other documents associated with the procedure. Use of incorrectly identified labels may prevent direct traceability of the implanted unit in the future. | Labeling mix-ups | Consignees were hand delivered by sales representatives a Astora Women's Health "Urgent Recall Notice". The letter was addressed to Physician, Health Care Professional, Nurse &amp; Risk Manager. The letter described the problem and the product involved in the recall. Consignees were advised to check their inventory and return all affected product, notify those that might be affected and to complete and return the Acknowledgement Form. For questions contact ASTORA Women's Health Customer Service at 1-844-879-0787. |  | 52 | 52 | Distributed in the states of AZ, CA, DE, IL, KS, NE, NJ, NY, OH, PA, TX | Terminated | March 09, 2016 | 135 | 52 | Not_Computer
Z-0272-2016 | 72449 | No | CAN | Pressure regulator. | Anesthesiology | Anesthesiology | 510(K) Exempt | Easy Dial Oxygen Regulator, 168708D and 168715D. |  |  | October 08, 2015 | 2015 | Precision Medical, Inc. | These lots of Oxygen Dial Regulators could have the settings label off by one position. When the indicator arrow points to OFF position, the regulator is actually on the half liter setting (0.50), for the 15 liter. For the 8 liter, instead of being in the OFF position, the regulator will be in the quarter liter setting (0.25). | Error in labeling | The firm sent "Attention Field Corrective Action" letters, dated Oct 8, 2015, to customers. The letter identified the affected device and explained the reason for the recall. The letter provided instructions on how to check the regulator; and asked customers to call the number provided for a replacement if a regulator is labeled incorrectly. |  | 315 | 315 | Distributed in the US to the states of PA, MA, MO, MT, ID, MI, MN, ND, KY, NY, NC, CA, AL, AZ, LA, OH, TX, GA, IN, and OK. | Open, Classified |  |  | 315 | Not_Computer
Z-0267-2016 | 72368 | No | NDP | Endosseous dental implant accessories. | Dental | Dental | 510(K) Exempt | Biomet 3i Narrow Right Angle Square Driver Tip. Driver tips are attached to a surgical drilling unit to carry components such as implant mounts, healing abutments and screws for placement and removal. For dental and restorative procedures. |  |  | August 10, 2015 | 2015 | Biomet 3i, LLC | RASQ3N, a square, abutment screw driver tip has been labeled and distributed as RASH3N, a hexed, abutment screw driver tip. Additionally, the RASQ3N received the RASH3N's laser marking "RASH3N" during manufacturing. | Labeling mix-ups | Beginning on or about 08/10/2015 customers will be notified in writing via FEDEX and instructed to to check their inventory for the identified product and to return any unused quantiles along with the recall response. Biomet 3i will promptly issue replacement product. |  | 103 | 103 pieces | AK, AZ, CA, CO, CT, FL, GA, IA, ID, IL, KY, LA, MA, MD, MI, MS, NC, NE, NJ, NY, OH, OR, PA, SD, TN, TX, UT, VA, WA, and District of ColumbiaArgentina, Australia, Belgium, Canada, Colombia, Greece, Hong Kong, Ireland, Italy, South Korea, and Kuwait. | Open, Classified |  |  | 199 | Not_Computer
Z-0268-2016 | 72550 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Perfusion Tubing Packs, Models TL7M64R3, TL7P17R1 &amp; HY2W85R9, Sterilized using Ethylene Oxide. |  |  | October 28, 2015 | 2015 | Medtronic Perfusion Systems | This recall is being initiated due to a potential breach of the Tubing Pack Header Bags manufactured in these select lots. The seal in question is between the Tyvek and the Poly side of the header bag. | Packaging process control | Medtronic sent an Urgent Medical Device Recall letter dated October 2015 to all affected consignees. The letter described the problem and the product involved in the recall. Consignees were advised to quarantine and return all affected product and to complete and return the Customer Confirmation certificate. For questions contact your Medtronic representative or contact Medtronic's national Answering service at 800-633-8766. |  | 496 | 496 | Nationwide Distribution including US: AL, IL, IN, and WI. | Open, Classified |  |  | 496 | Not_Computer
Z-0270-2016 | 72413 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Stryker Radius 5.5 to 3.5 Rod to Red Connector, REF 48664530. Intended to provide additional support during fusion using autograft or allograft in skeletally mature patients. |  |  | September 25, 2015 | 2015 | Stryker Spine | Stryker Spine is recalling Radius Rod to Rod Connector 5.5 to 3.5 mm (Catalog #486614530) because it does not have an associated 510(k) submisstion and subsequent clearance, which is required as this is a class 2 device. | No Marketing Application | Stryker Spine initiated this recall by sending an "Urgent Product Removal Letter and Product Accountability Form dated September 25, 2015 to the Branch Manager, Agency Manager, Quality Contact via FedEx priority overnight. |  | 42 | 42 units | Nationwide Distribution including the states of AK, AL, AR, AZ, CA, CO, FL, GA, IA, IL, IN, KS, KY, MA, MD, MI, MN, MO, MT, NC, NM, NY, OH, OK, OR, PA, TN, TX, UT, WA and WI. | Open, Classified |  |  | 42 | Not_Computer
Z-0253-2016 | 72204 | No | HHW | Vaginal pessary. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | MILEX PESSARY KIT; Model Numbers: MXPES00; MXKPES00; MXKPREG03; MXKPGSK07 |  |  | September 14, 2015 | 2015 | CooperSurgical, Inc. | The products have been identified to have incorrect size prints on the MILEX PESSARY KIT packaging. While the Pessary pouch is marked with correct size, an incorrect label onthe packaging does not meet the required release specification. | Labeling mix-ups | Cooper Surgical sent an Urgent Medical Device Recall letter dated September 14, 2015 via FedEx with confirmed delivery receipt. The letter identified the affected product, problem and actions to be taken. Customers were asked to discontinue use of the affected products and complete the Acknowledgement and Receipt form for a free replacement. For questions contact Product Surveillance at 203.601.5200 ext. 3300. |  | 144 | 144 units | Worldwide Distribution - US Nationwide in the states of AR, CA, CO, GA, IL, KY, LA, MA, MD, MI, MT, NC, NY, OH, OK, PA, RI, SD, VA, WA, WI,WV and in the countries of FRANCE and VIETNAM. | Terminated | January 08, 2016 | 116 | 144 | Not_Computer
Z-2529-2015 | 71893 | No | NGT | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Mahurkar Elite Acute Dual Lumen Catheter Tray, High Flow Pre-Curved with IC* Safety Components, 13.5 Fr x 20 cm |  |  | August 06, 2015 | 2015 | Covidien LLC | Sterility may be compromised | Packaging process control | Medtronic/Covidien sent an Urgent Medical Device Recall letter August 7, 2015 to affected customers via Federal Express. The letter identified the affected product, problem and actions to be taken. The letter instructs customers to quarantine, discontinue use and return the affected product as instructed in the letter. For questions contact your Medtronic representative or Customer Service at (800)-882-5878 |  | 4 | 4 kits | US Nationwide Distribution | Terminated | June 29, 2017 | 693 | 420002 | Not_Computer
Z-0233-2016 | 72462 | No | KDD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Myelo-Nate CSF Sampling Kit, Neonatal/Pediatric cerebralspinal fluid sampling sterile tray, including a lumbar puncture needle, four collection vials, an iodine swabstick package, a fenestrated drape, a polylined drape, three gauze pads, and one spot bandage. |  |  | October 16, 2015 | 2015 | Utah Medical Products, Inc | Utah Medical is conducting a recall of Myelo-Nate CSF Sampling Kit due to an internal post-sterilization inspection that found a sterile finished tray where the seal between the Tyvek lid and the PETG tray on oneside of the tray had been compromised. | Package design/selection | Consignees notified on 10/15/15 via letter and email, instructed to return product to distributor. |  | 21403 | 21,403 trays | Worldwide Distribution-US (nationwide) including Puerto Rico and the states of CA, CO, MD, WI, FL, GA, IL, IA, ID, IN, KS, KY, LA, MI, MN, MO, MS, NC, NE, NJ, NV, OH, OK, OR, PA, SC, TN, TX, VA, WA, and WI, and the countries of Canada, United Kingdom, Saudi Arabia, Jordan, United Arab Emirates, Kuwait, and Costa Rica. | Open, Classified |  |  | 21403 | Not_Computer
Z-2548-2015 | 71721 | No | OEQ | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | Enforcement Discretion | ANGIODRAPE PK 4/CS |  |  | July 10, 2015 | 2015 | Customed, Inc | Recall initiated due to inadequate validation/qualification of sterilization processes, and package integrity for products released for distribution as sterile. | Packaging | Customed issued an "URGENT VOLUNTARY MEDICAL DEVICE RECALL" notification letter dated July 10, 2015 to all affected customers. The letter identified the product, problem, and actions to be taken by the customer. |  | 192 | 192 packs | Distributed Only in Puerto Rico. | Terminated | April 12, 2017 | 642 | 358405 | Not_Computer
Z-0902-2015 | 69773 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ABX PENTRA Reagent Container, Model No. B1037307 (15 mL), B1034626 (10 mL), and B1034634 (4 mL) |  |  | November 13, 2014 | 2014 | Horiba Instruments, Inc dba Horiba Medical | HORIBA Medical is recalling the ABX PENTRA 400 Reagent Container because they do not always sit properly into the reagent rack. In low reagent level scenarios, this could lead to insufficient reagent volume being pipetted and could lead to incorrect results being reported without an alarm. The defect would result in a false high or false low results for HbA1c and false high results for magnes | Package design/selection | A recall letter was sent on 11/13/14 to all customers who purchased the ABX PENTRA 400 Clinical Chemistry Analyzer. The letter informs the customer of the problems identified and the actions to be taken. Customers are instructed to complete and return the enclosed response form within 10 days. Customers with questions are instructed to contact their local HORIBA Medical representative. |  | 443 | 443 units | Nationwide in US | Terminated | March 19, 2015 | 126 | 443 | Not_Computer
Z-2754-2011 | 59111 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Medtronic 6F Zuma Guide Catheter with AL 1.5 Curve with side holes |  |  | June 08, 2011 | 2011 | Medtronic Vascular | Mislabeled; Product labeled as 6F Zuma Guide with AL 1.5 curve with side holes may contain a 5F Zuma Catheter with SAL 1.0 curve | Labeling Change Control | The firm, Medtronic, sent an "Urgent Medical Device Recall Notice(for US)" letter dated June 7, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to segregate the product so that it will not be used. A Medtronic Sales Representative will visit the customers to account for all units they received and initiate a credit as appropriate.If you have any questions, please call (978) 739-0371 or (978) 739-3259. |  | 5 | 5 units | Nationwide distribution: USA including states of: CA, FL, and MA. | Terminated | April 17, 2012 | 314 | 5 | Not_Computer
Z-0565-2016 | 72810 | No | MJN | Vascular clamp. | Cardiovascular | Cardiovascular | 510(k) | RenovoCath RC 120; |  |  | November 25, 2015 | 2015 | RenovoRx, Inc. | Incorrect version of Instructions for Use were shipped to two US customers. | Labeling Change Control | RenovoRx notified customers by phone and by mail of this issue. This notification requests that the customers remove the affected product from their inventory and remove and replace the incorrect Instructions for Use with the correct Instructions for Use provided by RenovoRx. The customers are requested to return the incorrect Instruction for Use to RenovoRx for disposition and are requested to complete a Recall Confirmation Form enclosed to document the completion of the requested Correction. For further questions, please call (650) 284-4433. |  | 11 | 11 units | US Distribution to the states of : CA and FL. | Terminated | February 29, 2016 | 96 | 11 | Not_Computer
Z-2583-2010 | 56343 | No | MQP | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | SynFix-LR Spinal Implant System (Implant Holder for SynFix-LR and surgical technique brochure) |  |  | July 23, 2010 | 2010 | Synthes USA (HQ), Inc. | A potential for the implant holder tip to break off and lodge in the SynFix-LR implant, which pose a potential risk to the patient. | Error in labeling | Synthes Spine Hospital sent an ' URGENT: MEDICAL DEVICE LABELING CORRECTION" letter dated August 20, 2010 to all customers. The letter described the labeling issue and actions to be taken by the customers. The customers were instructed to review the supplementary labeling provided; contact their Synthes Spine Sales Consultant to arrange for provision of the Holding Forceps (p/n 388.401) prior to each SynFix-LR procedure; and complete and return the enclosed Verification Section form to Synthes via fax at 610-251-9005 or email: FieldAction@Synthes.com.If you have any questions, please call 1-800-620-7025 x 3500 or contact your Synthes Spine Sales Consultant. |  |  |  | Worldwide Distribution: USA and the countries of Canada, Japan, Switzerland, Australia, Mexico, Malaysia, New Zealand, and Singapore. | Terminated | December 18, 2014 | 1609 |  | Not_Computer
Z-0627-2009 | 50139 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Excelsior SL-10 Microcatheter, placed in sterile pouch. Manufactured by Boston Scientific Corporation, Fremont, CA. |  |  | October 07, 2008 | 2008 | Boston Scientific Corporation | Packaging issue may result in failure to maintain sterile barrier. | Package design/selection | Consignees were notified by letter dated October 7, 2008. The letter identified the issue and affected product. Customers were instructed to: 1.) Discontinue use of and segregate recalled product, 2.) Complete and return the enclosed Reply Verification Tracking Form, and 3.) Package/ship the recalled product and send to the Boston Scientific Corporation (BSC).Return the completed Reply Verification Tracking Form via fax (800-836-4804) to the BSC Customer Service Call Center.A follow-up letter dated October 8, 2008 was also issued to customers. The letter notified customers of the correct date range of expiration for product within the scope of the recall which is September 2008 through October 2010. There were no other changes identified in the follow-up letter. |  | 21423 | 21,423 units | Worldwide Distribution. | Terminated | March 02, 2010 | 511 | 21423 | Not_Computer
Z-0906-2009 | 50737 | No | EMG | Dental hand instrument. | Dental | Dental | 510(K) Exempt | Obwegeser Ramus Clamp w/ wide shield, 20 cm, 8", Right. ORAL-MAXILLO-CRANIO-FACIAL SURGERY, Original OBWEGESER Instruments. Walter Lorenz Surgical, Inc., Subsidiary of Biomet Microfixation, Inc. NOTE: None distributed domestically. |  |  | May 21, 2008 | 2008 | Biomet Microfixation, Inc. | The Biomet Microfixation Obwegeser Ramus Clamp, Wide Right 01-7996 were incorrectly etched as the left part number 01-7997. | Labeling mix-ups | Customers were contacted by e-mail notification with read receipt. Customers were notified in the e-mail to immediately discontinue use of the recalled product and return the affected items to Biomet Microfixation, Inc. The e-mail also included an Inventory Reconciliation form for customers to complete and return via fax (904-741-1801) to Biomet Microfixation, Inc. For questions related to this notice, please call 1-800-874-7711 or 904-741-4400 extension 448, Monday through Friday, 8 a.m.-4 p.m. Eastern Standard Time (EST). |  | 4 | 4 | International Distribution including countries of Ireland and Japan. | Terminated | September 11, 2009 | 478 | 4 | Not_Computer
Z-1555-2009 | 52007 | No | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Siemens Stratus CS Acute Care Troponin I CalPak. |  |  | April 28, 2008 | 2008 | Siemens Healthcare Diagnostics, Inc. | The firm received and confirmed that some lots of Cardiac Troponin I (cTnl) CalPaks (CCTNI-CR) may exhibit calibration issues that cause lower than typical signal during calibration with the current population of cTnl TestPaks (CCTNI). This issue may result in failed calibrations due to "Slope Error" messages or falsely elevated patient sample and quality control (QC) results on calibrations that | Labeling False and Misleading | Siemens Healthcare Diagnostics, Inc. issued an Urgent Field Safety Notice dated April 29, 2008 to all affected customers instructing them to discontinue use of 3 CalPak lots and reduce the shelf-life for the remaining lots.For further information, contact Siemens Healthcare Diagnostics, Inc. at 1-302-631-7672. |  | 4415 | 4415 cartons | Worldwide Distribution - including United States, United Arab Emirates, Austria, Bermuda, Belize, Canada, Germany, Egypt, India, Japan, Saudi Arabia and Bengal. | Terminated | July 14, 2009 | 442 | 4415 | Not_Computer
Z-1716-2009 | 52217 | No | JDL | Hip joint metal/metal semi-constrained, with a cemented acetabular component, prosthesis. | Orthopedic | Orthopedic | PMA | CONSERVE, TOTAL A-CLASS HEAD W/BFH TECH., REF 38AM-3635, O.D. Size: 36mm, Neck Length: +3.5mm, Taper: SLT, Sterile R, 1 each, Rx only, Wright Medical Technology, Inc., 5677 Airline, Road, Arlington, TN 38002. |  |  | May 14, 2009 | 2009 | Wright Medical Technology Inc | Incorrect packaging, swapped sizes | Packaging process control | Wright Medical Technology, Inc. notified hospitals, surgeons and their distributors of the recall situation via an "Urgent: Medical Device Voluntary Recall" notice dated May 14, 2009 explaining the problem, the hazard involved and requesting confirmation that the product was removed from their inventory. The letter included a "Fax Back" response page to ensure the receipt of the notice.For further questions, contact Wright Medical Technologies, Inc. at 1-800-874-5630. |  | 15 | 15 units | Worldwide Distribution -- Including states of NC, TN, IA, IN, MD, WI, WV, OK, WA, PA, TN, TX, FL and CA and country of Canada. | Terminated | December 09, 2009 | 209 | 29 | Not_Computer
Z-1291-2015 | 70551 | No | MMP | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General Hospital | 510(k) | Outer cardboard box labeling: |  |  | February 11, 2015 | 2015 | Sybron Dental Specialties | Cases of Cure Sleeve labeled with part number 4513, which fits 13mm tips, have inner boxes labeled with part number 4500 which fit 7mm and 8mm tips. There is a low risk associated with the use of the product as the sleeve may tear causing a potential for cross contamination. | Packaging | The firm sent out letters to customers on 02/11/15.The firm's letter states that they are willing to replace or credit back customers if they have any of the affected product, and instruct them to contact Metrex Research Customer Care at 1-800-841-1428 in order to receive an RMA number. The firm also requests that customer complete and fax back the enclosed reply form in order to confirm receipt of this advisory notification. The firm also states that if the customer is a distributor that they identify the customers that might have received the affected product, and contact them within 48 hours. |  | 172 | 172 | Worldwide distribution. US Nationwide, Canada, Australia, Netherlands, Switzerland, Japan, | Terminated | May 21, 2015 | 99 | 172 | Not_Computer
Z-1599-2015 | 71138 | No | PJQ |  |  | Cardiovascular | PMA | VenaSeal, Sapheon Closure System, Model # SP-101, VS-301 and VS-402. For vascular use. |  |  | April 28, 2015 | 2015 | Covidien LLC | Possible sterility breach of the outer packaging (pouch) material which may compromise the outside surface sterility of the sealed inner tray. | Package design/selection | Consignees were contacted by letter via email on April 28, 2015. |  | 1661 | 1661 units | Worldwide distribution. No U.S. accounts; Australia, Canada, France, Germany, Hong Kong, Ireland, Italy, Latvia, Netherlands, Norway, Portugal, Switzerland, UAE and the United Kingdom. | Terminated | March 29, 2016 | 336 | 1661 | Not_Computer
Z-1301-2015 | 70286 | No | IMD | Hot or cold disposable pack. | Physical Medicine | Physical Medicine | 510(K) Exempt | NovaPlus Brand Warm Gel Infant Heel Gel Warmer, Item # V6391 |  |  | January 26, 2015 | 2015 | CooperSurgical, Inc. | Products without expiration date have the remote potential for minor burn | Package design/selection | CooperSurgical issued Recall Notification letter on January 26, 2015, via Fedex confirmed receipt delivery, to all affected customers. Customers were requested to review inventory and return/destroy product that does not have a expiration and lot number. A fax back form to be returned. Customers with questions should call 203.601.5200.CooperSurgical's expense, arrangements were made to replace any of the affected products that customers had on hand. CooperSurgical provided further instruction on the disposition of the product at the time of replacement. |  | 4938266 | 4938266 units | Worldwide Distribution - USA (nationwide) and Internationally to AUSTRALIA, BAHRAIN, BERMUDA, CANADA, FRANCE, LATVIA, LIBYA, MEXICO, NEW ZEALAND, QATAR, SAUDI ARABIA, SINGAPORE, and UNITEDKINGDOM. | Open, Classified |  |  | 52899493 | Not_Computer
Z-1048-2015 | 70181 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | da Vinci Si Surgical System, IS3000; - Wall chart for EndoWrist Stapler 45 System. |  |  | January 15, 2015 | 2015 | Intuitive Surgical, Inc. | The Wall chart has been updated because it was noted that Wall Chart (PN 551524-03 Rev A) was inconsistent with the User Manual Adendum (PN 550986-07_B) with regards to neceassary sterilization dry time. | Error in labeling | Medical Device Recall Replacement letters were sent to all customers on January 15, 2015. The letter identified the reason for the recall, the risk to health, affected products, and actions to be taken by the customer or user and the firm. Customers were asked to discard the existing EndoWrist Stapler Documentation and replace it with the new Documentation provided by the firm. Intuitive Surgical will ship replacement instruments at no charge to customers. Then customers were asked to reunite the replacement instruments with the additional components. The firm will retrain customers on the updated instructions for use. A Clinical Sales Representative or Intuitive Surgical Customer Service at 1-800-876-1310, should be contacted if further information or support is needed concerning the recall notification. |  | 1034 | 1,034 | Nationwide Distribution | Terminated | March 03, 2015 | 47 | 1034 | Not_Computer
Z-1839-2015 | 71373 | No | GAL | Absorbable surgical gut suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Ethicon Fast Absorbing Plain Gut - MultiPass Needles, Sterile, Absorbable Surgical Suture, 5-0 (1.5 metric), 18" (45 cm), Conventional Cutting, Lot #JD6163, Exp: July 2019. |  |  | May 08, 2015 | 2015 | Ethicon, Inc. | Ethicon Fast Absorbing Surgical Gut (Plain) Suture is being recalled because the labeling on the box and foil are incomplete. | Labeling mix-ups | Ethicon initiated this voluntary recall by sending an Urgent: Medical Device Recall (Removal) Letter/Business Reply Form (BRF) dated May 8, 2015 to their affected customers via UPS. Stericycle will be handling returns. The letters identified the affected device and the reason for the recall. Customers were asked to examine their inventory immediately to determine if they had affected devices on hand. If so, they are to remove the affected product and communicate the issue to any necessary personnel. Customers with questions were advised to contact their local sales representative or call the Customer Support Center at 1-877-384-4266.For questions regarding this recall call 908-218-0707. |  | 1922 | 1,922 units | Worldwide Distribution -- US (nationwide) and Canada. | Terminated | June 06, 2017 | 760 | 1922 | Not_Computer
Z-1845-2015 | 71537 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Prestige Plus 185cm J-Tip: |  |  | June 04, 2015 | 2015 | Volcano Corporation | During a relabeling procedure at the manufacturing facility, the products were labeled with a shelf life in accordance with the date of the relabeling and not in accordance with the original shelf life date at the time of manufacture | Labeling mix-ups | Volcano sent an Urgent Medical Device Recalls lettersdated June 4, 2015 via Federal Express to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to examine product inventory, quarantine the affected product, complete the attached Customer Return form, and contact Volcano Customer Service at 800-228-4728, option 4 to arrange for the return of any product in inventory. |  | 162 | 162 total (both devices) | US Nationwide Distribution and one account in Canada. | Terminated | October 27, 2015 | 145 | 162 | Not_Computer
Z-1840-2015 | 71406 | No | JJX | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | The spare component Anti-HBs Cutoff Calibrator, 10 mLU/mL, Catalog number 26154. |  |  | May 13, 2015 | 2015 | Bio-Rad Laboratories, Inc. | The spare component Anti-HBs Cutoff Calibrator, Catalog number 26154, is recalled because the expiration date of 2016-04-30 printed on label in error. | Incorrect or no expiration date | Bio-Rad Laboratories Inc, sent an Urgent Product Correction letter dated May 13, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The recalled calibrator is a spare component which is sold separately from the MONOLISA Anti-HBs EIA Kit and the kit is not affected by this recall.Customers were informed that testing was performed and the Anti-HBs 10 mLU/mL Calibrator was acceptable to use through the expiration date of the kit dated 2015-11-30.Customers with questions can call the firm at 1-800-224-6723, option #2, then #3. |  | 20 | 20 vials | US Distribution to the states of : AZ, CA, ID, MO, and NH. | Terminated | November 05, 2015 | 176 | 20 | Not_Computer
Z-1585-2015 | 70854 | No | MUE | Full-field digital mammography system. | Radiology | Radiology | 510(k) | Aspire HD, Full Field Digital, System, x-Ray, Mammographic; |  |  | February 03, 2015 | 2015 | Fujifilm Medical Systems U.S.A., Inc. | User Manual lacks a description for handlingsome errors even though the system can be recovered from such errors | Labeling Change Control | Fuji Film on 2/3/15, issued a User Notification letter and the recovery procedure were hand carried to the user. A copy of Notification Letter is user's acknowledgement of receipt of the manual . |  | 1 | 1 unit | Distributed in HI. | Terminated | May 20, 2015 | 106 | 1 | Not_Computer
Z-1309-2015 | 70557 | No | HQL | Intraocular lens. | Ophthalmic | Ophthalmic | PMA | Intraocular Lens (IOL), Part No. EC-3 PAL, Diopter: 17.5, 18.0, 18.5, 19.0, 19.5, 21.5, 23.5, 9.0, 22.5, 23.0, 21.0, 24.0. |  |  | February 13, 2015 | 2015 | Aaren Scientific, Incorporated | Aaren Scientific is recalling intraocular lenses because the instructions for use (IFU) are not fully compliant with US requirements. | Labeling design | A customer notification letter dated 2/13/15 was sent to all customers who received the Intraocular Lens model EC-3 PAL. The letter provides the customers with the problems identified and the actions to be taken. Customers are instructed to confirm receipt of the recall letter and to contact Mark Scheib at (909) 906-5119 or via email at mark.scheib@zeiss.com with any questions or comments. |  | 28 | 28 units | Distributed in the states of OK, KS, FL, and AL. | Terminated | June 29, 2015 | 136 | 28 | Not_Computer
Z-1030-2015 | 70313 | No | OBJ | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Arrow Select Kits (ASK), Maximal Barrier Precautions Kits (CDC) and Arterial Catheterization Kit (NA) with Biopatch Dressing |  |  | January 09, 2015 | 2015 | Arrow International Inc | Certain product codes and lot numbers of Kits with Biopatch Dressing were packaged with the incorrect size Biopatch Dressing. No risk to patients is anticipated. | Packaging process control | Arrow International, Inc. (Arrow) sent notification letters, dated January 9, 2015, to alert their direct account customers that it may be necessary to obtain an alternate Biopatch Dressing if a kit is used that contains an incorrectly sized Biopatch Dressing. Due to the inconvenience, Arrow will reimburse customers $7.50 per kit received. |  | 4484 | 4484 | US (nationwide) and PR. | Open, Classified |  |  | 4484 | Not_Computer
Z-1315-2015 | 70090 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | Liquid Mouse Monoclonal Antibody CD15 used is laboratories to perform immunohistochemical (IHC) staining. |  |  | December 16, 2014 | 2014 | Leica Microsystems, Inc. | The staining intensity decreases over the shelf life. | Labeling design | Leica sent an Urgent Medical Device Recall Notification letters dated December 16, 2014 to all customers of record. The letters included instructions for customers to: 1) destory any usused and partially used recalled materials or confirm that the materials are no longer in stock; 2) complete and return the attached Field Safety Correction Notice Acknowledgement Form; 3) contact your local Leica representative immediately if a replacement is required; and, 4) pass this notice primarily to the end users where the product has been sold and to all those within your organisation who need to be aware of this manufacturing issue. Customers with questions can contact Leica Microsystems via e-mail at LMGRA@leica-microsystems.com. |  | 309 | 309 units | Worldwide Distribution - US Nationwide in the states of CA, CT, FL, IA, ID, IL, MT, NC, NY, TX, WA and in the countries of: Argentina, Australia, Bosnia and Herzegovina, Brazil, Canada, Chile, Denmark, El Salvador, France, Germany, Hungary, India, Italy, Japan, Netherlands, Nigeria, Portugal, Romania, Russia, Saudi Arabia, Serbia, South Africa, Spain, Sweden, Switzerland, Taiwan, Tunisia, Turkey, United Arab Emirates, United Kingdom. | Open, Classified |  |  | 309 | Not_Computer
Z-1824-2015 | 71313 | No | HIB | Obstetric-gynecologic specialized manual instrument. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(K) Exempt | McKesson Medi-Pak Vaginal Speculum, 10/bag |  |  | April 13, 2015 | 2015 | CooperSurgical, Inc. | Small sized specula were incorrectly packaged in printed bags for the medium sized speculum | Packaging change control | CooperSurgical sent an Urget Medical Device Recall letter dated April 13, 2015, to all affected customers. The letter identified the product, the problem, and the actions to be taken by the customer. Customers were instructed to discontinue use, discard, or return the affected to product to Cooper Surgical. Customers were asked to complete the attached Acknowledgment and Receipt Form for replacement. Customers with questions were instructed to call 203-601-5200. |  | 9400 | 9400 units | Nationwide Distribution including IL, FL, GA, MA, MO, TX, and WA. | Terminated | February 05, 2016 | 298 | 9400 | Not_Computer
Z-1741-2011 | 57705 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | CR Tibial Insert, Size 5, 7mm, REF 163-2507, Sterile R, Ortho Development, Draper, Utah 84020. |  |  | January 07, 2011 | 2011 | Ortho Development Corporation | Components of a replacement knee system could be mislabeled. | Error in labeling | The firm, Ortho Development, sent a "Field Notification Memo" date January 7, 2011 to its customers. The memo described the product, problem and action to be taken. The customers were instructed to promptly return any remaining parts of the listed product; locate all the parts in their inventory; identify any of the listed parts that may have been used in surgery; verify the quantity they have on hand and the quantity used in surgery and reconcile them with the quantity sent; if discrepancy exists, notify customer service resolution; prepare items for return and contact customer service to obtain RMA number. Once reconciled, arrangements would be made for sending new inventory and the return of unused parts.If you have any questions regarding this issue, please notify Regulatory Affairs Manager at (801) 619-3450. |  | 14 | 14 units (US), 2 (Intnl) | Worldwide distribution: USA including states of: AZ, CA, FL, MI, TX, and UT; and country of: Japan. | Terminated | December 27, 2011 | 354 | 72 | Not_Computer
Z-2046-2011 | 58240 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Clinac High Energy Medical Linear Accelerator, a Trilogy Radiotherapy Delivery System; Model #s: H14, H18, H27, H29, HCX; Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA |  |  | February 24, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | The wedge angle labeling on the wedge tray may not correctly match the wedge body. However, the wedge body is labeled with the correct wedge angle. | Labeling mix-ups | The firm, Varian Medical, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated February 24, 2011, to all affected users. The letter described the product, problem and actions to be taken. The customers were instructed to IMMEDIATELY inspect all wedges to verify that the wedge angle label on the wedge body matches the wedge angle on the wedge tray; if any wedge is identified with mismatched labeling follow the written instructions provided in the letter; always measure wedge data at installation of new or replacement wedges; verify all treatment accessories of "matched" linear accelerators, and please advise the appropriate personnel working in your Radiotherapy Department of the content of this letter.Varian has enhanced their manufacturing quality control process to further ensure correct labeling of wedges. Also, Varian will replace any wedge that is identified as mislabeled, at the request of the customer. If you require further clarification, please feel free to contact your local Varian Customer Support Service Manager at: 1-888-VARIAN5 (888-827-4265) for USA and Canada; +14 41 749 8844 Europe; North American: support-americas@varian.com; Australia/New Zealand:support-anz@varian.com; Europe: support-emea@varian.com; South East Asia: seasia.apps.helpdesk@varian.com; China/Asia: china.apps.helpdesk@varian.com; Japan: Japan.apps.helpdesk@varian.com; Latin America:soporte.al@varian.com. |  | 5541 | 5541 units | Worldwide distribution. | Terminated | November 05, 2012 | 620 | 5541 | Not_Computer
Z-0916-2012 | 60762 | No | IWZ | Radiographic film. | Radiology | Radiology | 510(K) Exempt | Drystar Mammo. |  |  | March 02, 2010 | 2010 | AGFA Corp. | Labeling problem | Error in labeling | The firm, AGFA Corporation spoke to the customer directly via telephone concerning the mislabeled affected products. The firm also contacted the customer via e-mail on April 13, 2011 explaining the outcome of the customer complaint. Replacement product was sent to the customer. |  | 2 | 2 boxes | Product distributed in Florida. | Terminated | February 14, 2012 | 714 | 2 | Not_Computer
Z-1012-2015 | 70086 | No | JXG | Central nervous system fluid shunt and components. | Neurology | Neurology | 510(k) | Miethke Shunt System accessories |  |  | December 17, 2014 | 2014 | Aesculap, Inc. | Some Miethke Shunt System accessories may have an incorrect MR safety symbol on the shelf box label. The label should have NO symbol but instead is marked with the MR conditional symbol. The product is MR unsafe. This information is only required to be in the Instructions for Use (IFU). | Labeling False and Misleading | Aesculap sent an Important Correction and Removal and Field Safety Notification letter, dated December 17, 2014, to all affected customers. The letter identiifed the affected product, problem, and actions to be taken. Customers were requested to remove from inventory and contact Aesculap regarding return of the affected devices and acknowledgement form. |  | 772 | 772 | Worldwide Distribution - US Nationwide | Open, Classified |  |  | 772 | Not_Computer
Z-0694-2012 | 60851 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | MERZ AESTHETICS, 27G X 1.0 " (25MM) Cannula, Sterile R, QTY 20, Model 9015M1X20, Distributed by Merz Aesthetics, In., Manufactured by TSK Laboratory. |  |  | December 16, 2011 | 2011 | Merz Aesthetics, Inc. | Merz Aesthetics is initiating a medical device recall on TSK Closed Tip Cannula (27GA x 25mm) due to a labeling error on the secondary (outer) packaging. The outer package label states that the carton contains Cannulas that are 27GA x 25 mm in size instead of 27G x 38 mm. Length of the cannula is a matter of user preference and is unlikely to impact patient safety. | Packaging process control | Merz Aesthetics Inc, initially contacted customers by telephone on December 16, 2011. They were informed that the wrong size cannula were sent with their last order and would need to be returned and a replacement order would be sent. . Merz Aesthetics sent a "Urgent Medical Device Recall" letter dated January 6, 2012. The letter identified the product the problem and the action needed to be taken by the customer. In addition, if you have further distributed this product, please identify your customers and notify them of this product recall. Your notification to your customers may be enhanced by including a copy of this recall notification letter.You may contact Merz Aesthetics Customer Service at (866) 862-1211 or (262) 835- 3333 ,for instructions regarding the return and immediate replacement of the product, or if you have any further questions or concerns.. |  | 5 | 5 boxes (20 units per box) | Nationwide Distribution including the states of CA, NY, OK, and TN. | Terminated | July 27, 2012 | 224 | 5 | Not_Computer
Z-1010-2015 | 70102 | No | OCX | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Product Code 2C4013, Four Lead TUR Irrigation Set, for use with UROMATIC |  |  | January 14, 2015 | 2015 | Baxter Healthcare Corp. | Potential for failure of the pouch packaging seal at highaltitudes. | Package design/selection | 1. Locate and remove all affected product from your facility. The product code and lot number can be found on the individual product or shipping carton.2. Contact Baxter Healthcare Center for Service to arrange for return and credit.Baxter Healthcare Center for Service can be reached at 888-229-0001 betweenthe hours of 7:00 AM and 6:00 PM Central Time, Monday through Friday. Pleasehave your Baxter 8-digit ship-to account number, product code, lot number(s) and quantity of product to be returned ready when calling.3. Complete the enclosed Customer Reply Form and return it to Baxter by faxing it to 224-270-5457 or scanning and e-mailing to fca@baxter.com . Returning the Customer Reply Form promptly will prevent you from receiving repeat notices.4. If you distribute this product to other facilities or departments within yourinstitution, please forward a copy of this communication to them.5. If you are a dealer, wholesaler, or distributor/reseller that distributed any product to other facilities, please conduct a consumer-level recall of the affected product that you distributed to customers.If product was purchased from a distributor: 1. Locate and remove all affected product from your facility. The product code and lot number can be found on the individual product or shipping carton.2. Contact Baxter Healthcare Center for Service to arrange for return and credit.The Baxter Healthcare Center for Service can be reached at 888-229-0001between the hours of 7:00 AM and 6:00 PM Central Time, Monday throughFriday.3. Follow your supplier's recall and response process. Please do not return thecustomer reply form to Baxter. |  | 26028 | 26,028 units | US Distribution. | Terminated | April 10, 2017 | 817 | 26028 | Not_Computer
Z-1912-2015 | 71305 | No | MAL | Vascular graft prosthesis. | Cardiovascular | Cardiovascular | 510(k) | Gelseal Straight Diameter 8mm Length 15cm. |  |  | May 15, 2015 | 2015 | Vascutek, Ltd. | Product was labeled as a 30cm graft, when in fact the graft in the box was a 15cm graft. | Error in labeling | On 5/15/2015, Vascutek issued a URGENT FIELD SAFETY NOTICE to their consignee informing them of the corrective action of the affected product. The information included details of the affected devices, including potential risks and actions to be taken by medical staff and distributors. |  | 12 | 12 | Distributed in the state of Michigan (USA), and the countries of United Kingdom, Australia, Poland, and Germany. | Terminated | October 20, 2015 | 158 | 12 | Not_Computer
Z-1395-2015 | 70701 | No | IWE | Light beam patient position indicator. | Radiology | Radiology | 510(K) Exempt | GE Discovery MR950 MRI system |  |  | October 24, 2014 | 2014 | GE Healthcare | The alignment lasers are missing the labels required by radiation safety regulations. | Error in labeling | GE Planned action(s) to repair defect or to bring product into compliance:The Center for Devices and Radiological Health (CDRH) understands that the CAP involving 5 units consists of the following:1. New labels will be mailed to all purchasers, with instructions to apply the label in the appropriate location on the unit, or a GE field engineer will be sent to each user site to complete the correction.2. All mailings and service calls will be made at no cost to the purchaser.3. The corrections will be completed by August 1, 2015.If you have any questions, please contact your local Service Representative or call 1 (800) 437-1171. |  | 5 | 5 units (2 in US) | US Distribution to the state of CA., and Internationally to Italy and Japan. | Terminated | May 06, 2015 | 194 | 5 | Not_Computer
Z-1323-2015 | 70594 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Coblator II System, 120V. |  |  | February 20, 2015 | 2015 | ArthroCare Medical Corporation | The recalling firm was notified by the contract manufacturer that they had reversed the L2, L3, L4, or L5 components on the P/N09010 RoHS boards of the Coblator II Controller (Model RF8000E). | Packaging process control | ArthroCare sent an Urgent Medical Device Recall letter dated February 2015 to customers. A response form will be included in the letter and a follow up notifcation letter will be sent to non-responding consignees. The letter identified the affected product, problem and actions to be taken. For questions contact your local Returns Department Representative at 800-343-5717, option 3. |  | 15 | 15 units | US Nationwide Distribution in the states of TN, IA, AZ, NE, SD, NY, and OK. | Terminated | February 26, 2016 | 371 | 15 | Not_Computer
Z-1009-2015 | 69533 | No | MQN | Bone plate. | Dental | Dental | 510(k) | The Synthes Mandible External Fixator; all lots of part nos.: |  |  | October 17, 2014 | 2014 | Synthes, Inc. | Under new testing protocols for MR Environment safety and compatibility, metal devices are no longer designated MR Safe. Some components in the CMF Mandible External Fixator Systems were labeled and/or etched with information indicating they are MR Safe. They can no longer be labeled MR Safe. | Labeling design | Urgent notices of medical device labeling updates, dated October 17, 2014 and October 24, 2014 were sent to end users and sales consultants, which described the product, problem, and action to be taken. These consignees were instructed to inspect inventory and remove out-dated labeling inserts and replace with updated labeling information. |  | 43159 | 43159 | Worldwide Distribution. | Open, Classified |  |  | 43159 | Not_Computer
Z-1847-2015 | 71342 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Medtronic Navigated Cannulated Taps |  |  | May 26, 2015 | 2015 | Medtronic Navigation, Inc. | Medtronic Navigation is initiating a field correction due to the potential for injury which could occur as a result of cleaning the Medtronic Navigated Cannulated Taps. | Error in labeling | Consignees were notified via Urgent Medical Device Correction Notification on 5/26/15. The letter identified the affected product and described the issue. This field correction does not require the return of the subject products. Revised instructions for use are being provided to the identified consignees. |  | 15798 | 15,798 | Worldwide Distribution. | Terminated | October 23, 2015 | 150 | 15798 | Not_Computer
Z-1916-2015 | 71424 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Bullet-Tip PEEK VBR/IBF System, 32mm(l) 13mm(H) Implant 32-13-32 UDI (01)008464680341 08(17)191218(10)200827; T-Plus PEEK VBR/IBF System, 10mm(W) x 27mm(l) x 13mm(H) Implant, 6degree lordotic 30-T-13-6 UDI (01) 00846468032920(17) 200116(10) 202726; and T-Plus PEEK VBR/IBF System, 10mm(W) x 36 mm(l) x 13mm(H) Implant, 6degree lordotic 30-T-1036-13-6 UDI (01)00846468032760(17)200115(10)202582. These products are part of the lnterbody Fusion Vertebral |  |  | May 21, 2015 | 2015 | RTI Surgical, Inc. (dba Pioneer Surgical Technology, Inc.) | Three lots of the Bullet -Tip VBR 13mm x 32 mm implants are mislabeled with a light green colored label, used for 16 mm trial instead of a lime green label, used for a 13 mm trial. | Error in labeling | RTI Surgical sent an" Urgent Medical Device Voluntary Recall " Notification dated May 21, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Immediate Actions: We are aware that one of this product was distributed to you. This one device has already been returned to RTI Surgical so no further action is needed. This notice is for your records.RTI takes this error seriously and has initiated a corrective and preventative action investigation to prevent recurrence. We apologize for any inconvenience this may have caused. For further questions please call (906) 226-4489. |  | 18 | 18 | US Distribution to the states of : FL and NY | Terminated | October 29, 2015 | 161 | 18 | Not_Computer
Z-1005-2015 | 70017 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Multi-Lumen PI CVC Kit and PSI Kit: 3-L 7 FR X 8 product code: ASK-45703-POSU lot #23F13L0166, exp. 31-Jul-2015: |  |  | December 16, 2014 | 2014 | Arrow International Inc | the Instructions for Use (IFU) were inadvertently removed from the kit. | Labeling False and Misleading | An Urgent Medical Device Corrective Action Notification letter, dated December 16, 2014, was sent to Owens &amp; Minor, 1605 Worldwide Blvd., Hebron, KY 41048-8653, instructing them to immediately discontinue distribution and quarantine any products, use customer letter and Corrective Action Acknowledgement Form templates to communicate this corrective action to any of thier customers that have received product included within the scope of the corrective action, and to have their customers return a completed Corrective Action Acknowledgement Form outlining their remaining inventory of the affected product codes and lot numbers for which replacement IFUs is needed. |  | 420 | 420 | In the state of KY | Terminated | August 10, 2015 | 237 | 420 | Not_Computer
Z-1562-2015 | 70841 | No | GCY | Nonpowered, single patient, portable suction apparatus. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | 3/16 in. Trocar with Single Drain |  |  | March 26, 2015 | 2015 | Stryker Instruments Div. of Stryker Corporation | Stryker Instruments is voluntarily recalling the CBCII Wound Drains and Ambulating Bulb for a potential sterility breach in the packaging | Package design/selection | An URGENT MEDICAL DEVICE NOTIFICATION letter was sent to all customers on 03/31/2015. The letter identified the affected product, the reason for the recall, the risk to health, and actions to be taken by the customer or user. All are asked to sign the Business Reply Form stating they have read the Notification and followed all instructions within. Customers were asked to immediately check all inventory and quarantine any affected product, return affected product to Stericycle for destruction using a return label provided by Stericycle. Upon receipt of affected product a credit will be issued to customer account. |  | 2711 | 2,711 boxes (27,110 each) | Worldwide Distribution. US nationwide, Australia, Great Britain, The Netherlands, Spain, Romania, Greece, Switzerland, South Africa, and Japan. | Terminated | December 31, 2015 | 280 | 42806 | Not_Computer
Z-1830-2015 | 71287 | No | KGY | Organ bag. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | LapSac Surgical Tissue Pouch. |  |  | May 14, 2015 | 2015 | Cook Inc. | Includes Information not stated in the intended use. | Labeling Change Control | Cook Medical sent an "Recall Communication Updated Instructions for Use" letter dated May 15, 2015, to all affected customers . The letter identified the product the problem and the action needed to be taken by the customer. Cook has requested that customers do the following: The intended use for the device is being modified from "Used to isolate tissue" during, or prior to, surgical removal and/or morcellation to "Used to isolate tissue" during, or prior to, surgical removal .A contraindication being added is as follows:" Not for use with power morcellators.The warnings are being modified to read as follows:" The LapSac Surgical Tissue Pouch is not intended for use with sharp surgical instruments." Care should be taken to avoid damaging, puncturing or tearing the LapSac Surgical Tissue Pouch." Use caution when removing potentially malignant or harmful tissues, to isolate contamination from surrounding tissue.The instructions for use are being modified to read as follows:1. Following aseptic technique, introduce the LapSac Surgical Tissue Pouch into the abdominal cavity.2. Place tissue to be removed from the body into the pouch. 3. Using the drawstring, carefully close the LapSac Surgical Tissue Pouch. Please note that if excess force is used, the drawstring may break or tear the LapSac.4. Remove the LapSac Surgical Tissue Pouch from the body, taking care not to spill contents of the bag.Please report any Adverse Event to Cook Medical Customer Relations (800) 457-4500 or 1-812-339-2235. Monday through Friday between 7:30 a.m. and 5:00 p.m. Eastern Daylight Time or email at CustomerrelationsNA@cookmedical.com; and to FDA. The customer notice must be passed on to appropriate personnel within each organization or to any organization where the potentially affected devices have been transferred.If customers have further questions regarding labelling changes for this device, please contact Cook Sales Representative or Cook M |  | 123184 | 123,184 devices were distributed in commerce | Worldwide Distribution - US (nationwide) including PR and to the states of: AZ, AK, AL, AR, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MS, MO, ME, MT, NC, ND, NE, NH, NJ, NY, NM, OH,OK,OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV and WY., and to the countries of : China, Korea, Denmark, Mexico, New Zealand, India and Colombia. | Open, Classified |  |  | 123184 | Not_Computer
Z-1262-2015 | 70544 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | - NTL-30 30ml Nutrio Enteral Feeding Syringe (sterile) |  |  | February 13, 2015 | 2015 | Medela Inc. EF Division | The firm is recalling sterile enteral feeding syringes due to determining that the sterility of the affected products may be compromised as a result of a potential breach in packaging. | Packaging | Medela sent a Device Recall letter dated February 13, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Requested Action of Distributors:1. Stop sale and distribution of product with affected lot numbers.2. Conduct a physical count and record the data on the Product Inventory Form included with the letter.3. Return the Product Inventory Form via fax to 800-995-7867 or e-mail it to syringe@medela.com. The firm should contain a contact name and signature.4. If have the affected product in stock, Medela Customer Service will issue a RMA and schedule a pick-up of the affected product. Contact will be made to ensure product is ready. For further questions please call (888) 399-2455. |  | 1959187 | 1,959,187 | Nationwide Distribution. | Terminated | November 09, 2015 | 269 | 1959187 | Not_Computer
Z-1003-2015 | 70136 | apparatus, traction, non-powered | HST | Nonpowered orthopedic traction apparatus and accessories. | Orthopedic | Neurology | 510(K) Exempt | Fitness Gear Inversion table (Style STE00059FG or Style STE00118FG), Distributed by ASLI, Wilmington, DE (on carton), Paradigm Health &amp; Wellness, Inc. (in user manuals); The products are inversion tables, which allow the user to secure his or her ankles to the table and invert. The products were intended to be fitness equipment; i.e. consumer products. |  |  | December 29, 2014 | 2014 | Dick's Sporting Goods | Reports of customers falling from the Inversion Table and sustaining injuries. The Inversion Table is a fitness product, not a medical device. DSG removed medical claims from product labels and advertising. | Labeling False and Misleading | DSG sent emails and medical device correction notices, dated December 2014, to consumers who made purchases online and retail store consumers for whom DSG has contact information via its Scorecard customer loyalty program. DSG also posted the recall on their web site and notified their retail stores with instructions to post notices in DICKs stores. DSG plans to remind stores approximately monthly for three months to display the recall poster, beginning one month after the recall announcement. Consumers were asked to return the product to the stores where they were purchased for a full refund. A Press Release, dated January 8, 2015, was also issued about the recall. |  | 32581 | 32581 | Nationwide. | Terminated | September 02, 2015 | 247 | 32581 | Not_Computer
Z-1122-2015 | 70438 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Mallory Head 4 Finned Acetabular Shell 50 mm, PN 13-104150, hip prosthesis component (Ringloc) |  |  | January 09, 2015 | 2015 | Biomet, Inc. | Biomet recalled the Mallory Head 4 Finned Acetabular Shell 50mm (Lot 364600) and Ringloc + Acetabular Shell Limited Hole Finned 52 mm (Lot 427840) following a mix up in packaging with the two. | Packaging process control | Biomet notified all domestic and foreign customers via FEDEx overnight on 1/9/2015. Response forms were provided, customers were instructed to notify hospital personnel if product was further distributed, and remove recalled products from circulation. Replacement parts were made available, customers may contact Audrey Daenzer, Field Action Specialist, Regulatory ComplianceBiomet, Inc., audrev.daenzer@biomet.com, or 574-372-1570 Monday- Friday 8 AM to 5 PM. |  | 17 | 17 | Distributed in the states of VA and TX and the country of Japan. | Terminated | October 22, 2015 | 286 | 17 | Not_Computer
Z-1233-2015 | 70436 | No | DSP | Intra-aortic balloon and control system. | Cardiovascular | Cardiovascular | 510(k) | ARROW FiberOptix Intra-Aortic Balloon Catheter Kit, 40 cc 7.5 Fr |  |  | February 05, 2015 | 2015 | Arrow International Inc | Outer package product label incorrectly identifies the product as FiberOptix Flex IAB 7.5Fr rather than the correct FiberOptix IAB 8Fr | Labeling mix-ups | Arrow International notified Users, Distributors, Risk Managers by letter dated February 5, 2015. The letter explains the labeling error and requests immediately discontinue use and quarantine product. Customers were instructed to complete the Recall Acknowledgement Form and return by fax to 1-855-419-8507 whether they have or do not have the affected product. Customers with questions were instructed to contact their local sales representataive or Customers Service at 1-866-246-6990.For questions regarding this recall call 610-378-0131. |  | 165 | 165 units | Worldwide Distribution - USA (nationwide) and Internationally to Belgium | Open, Classified |  |  | 165 | Not_Computer
Z-1697-2015 | 71279 | No | FRO |  |  | General &amp; Plastic Surgery | 510(k) | ActCel |  |  | May 18, 2015 | 2015 | Coreva Health Science LLC | Coreva Health Science, LLC. is recalling ActCel Hemostatic Dressing due to customer literature and web sites promoting that it can help control bleeding from open wounds and in body cavities. ActCel is indicated as a prescription device for management of topical wounds and to temporarily control external surface bleeding. | Labeling False and Misleading | On 05/18/15 the firm sent out customer notification letters to those distributors that were identified as marketing ActCel Hemostatic Dressing for use to help control bleeding from open wounds in body cavities.The firm's notification letter provided the following actions to be taken by the customer:- Review their literature and labeling (including internet websites) to determine whether there are any marketing statements for usage in body cavities; -Destroy all hardcopy literature and labeling with the statement for usage in body cavities - Correct all Internet websites eliminating his statement; and- Complete the attached Medical Device Labeling Clarification Return Response form and return it to Coreva Health Science.The firm states to direct any questions to: Ms. Casssie Inglis, Vice President of Marketing.email: cassie@coreva.netphone: 818-685-0951 ext. 119toll free: 800-808-9094 ext. 119Monday through Friday, 9am-4pm PST |  | 180900 | 180,900 | U.S. distribution to the following; TX, NY, FL, VA, WA, and AZ. No foreign distribution.NOTE: The distribution list the firm provided does not include all consignees that purchased the product; rather target the consignees (distributors) who were observed to be using an unapproved marketing use. | Terminated | August 26, 2015 | 100 | 180900 | Not_Computer
Z-1239-2015 | 70502 | No | HHW | Vaginal pessary. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | CooperSurgical PESSARY KIT-CUBE#2, 1-3/8 in , 35 mm |  |  | February 09, 2015 | 2015 | CooperSurgical, Inc. | Incorrect size printed on the Milex Pessary Kit packaging label | Error in labeling | CooperSurgical sent an Recall Letter dated February 9, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letter asked customers to contact CooperSurgical Product Surveillance Department if they had any questions.At CooperSurgical's expense, arrangements were made to replace any of the affected products that customers had on hand. CooperSurgical provided further instruction on the disposition of the product at the time of replacement.Questions contact CooperSurgical 203.601.5200. |  | 95 | 95 units | Nationwide Distribution | Terminated | January 15, 2016 | 340 | 95 | Not_Computer
Z-1222-2015 | 70397 | No | BSY | Tracheobronchial suction catheter. | Anesthesiology | Anesthesiology | 510(K) Exempt | KimVent Wet Pak, Closed Suction System for Adults, Double Swivel Elbow, 14 Fr. Sterile, Distributed by Kimberly Clark Global Systems, LLC. |  |  | January 16, 2015 | 2015 | Halyard Health | Defect: A hole or crack was found which may cause air to leak. | Packaging process control | Consignees were notified by letter via FedEx on 01/16/2015. |  | 83 | 83 cases (20 tubes per case) | CA, FL, IL, LA, MA, MK, MN, MS, NM, NY, OK and TX and Japan. | Terminated | July 07, 2015 | 172 | 93 | Not_Computer
Z-1710-2015 | 71098 | No | LZS |  |  | Ophthalmic | PMA | EC-5000 Excimer Laser Corneal Surgery System (EC-5000, EC 5000 CXIII, EC-5000 Quest. |  |  | March 24, 2015 | 2015 | Nidek Inc | Multi-stage treatment option for the EC-5000 Operator's Manual was not approved by the FDA. | No Marketing Application | Nidek sent an Device Correction letter dated April 30th, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Nidek will contact customers to schedule an appointment to replace the operators manual and implementing a software lockout device on the device so that the multi stage feature is no longer available.If a Nidek rep does not contact you before April 30, 2015, please call 800-223-9044 to schedule an appointment, All non responding customers.will be contacted three time via Fed Ex letter |  | 114 | 114 units | Us distribution only. | Terminated | April 05, 2016 | 378 | 114 | Not_Computer
Z-1707-2015 | 71176 | No | CDP | Total triiodothyronine test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS Immunodiagnostic Products Signal Reagent, Product Code 1072693, Unique Device Identifier No. 10758750008360 VITROS Immunodiagnostics Products Signal Reagent (SR) is supplied to users as a sales unit box containing two Signal Reagent Packs. Each Signal Reagent Pack is labeled with a barcode label. The label should include the name of the reagent, a barcode, lot information and expiry dating. |  |  | May 06, 2015 | 2015 | ORTHO-CLINICAL DIAGNOSTICS | The Immunodiagnostic Products Signal Reagent, Lot 3082 as the reagent packs may contain blank labels. | Error in labeling | On 5/06/2015, Customer Letter and Verification Form (Ref. CL15-095) were sent notifying the users of the issue and advised them of the required actions. The notification was sent via FedEx overnight courier and/ or ORTHO PLUS e-Communications and/ or US Postal Service Priority Mail (only to those consignees who provided a PO Box as an address) to all customers who received the affected product and advised them of the issue and the required actions. On 5/06/2015, Distributor Letter (Ref. DL15-095) was sent via US Postal Service Priority Mail and/or FedEx overnight courier to US Distributor consignees and instructed to notify their consignees of the issue and actions. For questions,questions, please contact Customer Technical Services at 1-800-421-3311 at any time. |  | 898 | Domestic: 898 units | Nationwide Distribution. | Open, Classified |  |  | 898 | Not_Computer
Z-1718-2015 | 71058 | No | ITH | Power traction equipment. | Physical Medicine | Physical Medicine | 510(k) | Extentrac Elite User Manual DCN: 14-001 (Powered decompression table, for traction) |  |  | May 13, 2015 | 2015 | Advanced Back Technologies, Inc. | Item is marketed for use in patients weighing up to 350 lbs, but the device was only approved for use in patients up to 300 lbs | Error in labeling | Advanced Back Technologies, Inc. sent an Urgent Medical Device Brocure &amp; User Manual Recall letter dated May 13, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed the discard the DCN's identified as DCN 13-001, DCN 09-001, DCN 14-001 and replace with the newly issued DCN 15-001, DCN 15-002. The new DCN's would be provided to customers in pdf format via email for issuance. Customers were asked to contact their customers who purchased the affected product and inform them that they should not treat patients with a maximum weight exceeding 300 lbs until the firm is able to make an amendment to their 510(k). Customers with questions should call 631-231-0076. |  | 18 | 18 units | Worldwide Distribution - US including CA, FL, NJ, TX, WA, Internationally: Korea, Russia, Kuwait, Canada, Greece, Lebanon, Turkey, Egypt, Saudi Arabia, China, Armenia, Azerbaijan, Australia. | Open, Classified |  |  | 54 | Not_Computer
Z-1227-2015 | 70128 | dispenser, liquid medication | KYX | Liquid medication dispenser. | General Hospital | General Hospital | 510(K) Exempt | Baxter, Self-Righting Syringe Tip Caps, Self-Righting DISCPAC Syringe Tip Caps. Product codes H93866100, H93869100, H93866025, H93867025, H93867100, H938671025, H938671100, H938673025, H938673100, H938674025, H938674100, H938676025, H938676100, H938677025, H938677100, H938679025, H938679100, H938690025, H93869025, H938693025, H938696025, H93868025, H93868100, H938672025, H938672100, H938675025, H938675100, H938678025, and H938678100. |  |  | December 29, 2014 | 2014 | Baxter Corporation Englewood | Self-Righting Luer Tip Caps were manufactured and shipped with a missing latex symbol on both single and case labels. | Labeling Change Control | Baxter sent an Safety Alert letter dated January 15, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Baxter is requesting that you take the following actions:1.Complete the enclosed customer reply form and return it to Baxter by faxing it to 224-270-5457 or scanning and e-mailing it to fca@baxter.com. Returning the customer reply form promptly will prevent you from receiving repeat notices.2. Please forward a copy of this Safety Alert letter to other facilities or departments within your institutions to ensure that those locations are aware of this action.Action to be taken if you are an end-user and purchased product from a distributor or reseller:1. Please forward a copy of this Safety Alert letter to other facilities or departments within your institutions to ensure that those locations are aware of this action.2. Follow your supplier's reply process. Please do not return the customer reply form to Baxter.Action to be taken if you are a distributor:If you are a dealer, wholesaler, or distributor/reseller that distributed any product to other facilities, please notify your customers of this Safety Alert in accordance with your customary procedures. Further information and supportFor general questions regarding this communication, contact The Center for One Baxter at 800-422-9837, Monday through Friday, between the hours of 8:00 AM and 5:00 PM Central Time. |  | 1404230 | 1,404,230 | Nationwide Distribution | Open, Classified |  |  | 1404230 | Not_Computer
Z-1216-2015 | 69970 | No | IKZ | Powered patient rotation bed. | Physical Medicine | Physical Medicine | 510(K) Exempt | RotoRest Delta Advanced Therapy System. |  |  | December 03, 2014 | 2014 | Arjo Hospital Equipment AB | The recalled devices labeling and instructions for use contain unapproved medical claims. | Labeling design | Devices in the rental fleet will be delivered to the customer with revised documents including the updated information. The customers who purchased the devices will be notified of the recall via letter sent by registered mail. The field correction notice will include a recall letter, response form, and updated labeling. |  | 271 | 271 units | Worldwide Distribution. US nationwide, Germany, Austria, France, Italy, India, EU, Mexico, Switzerland, Canada, Saudi Arabia, Kuwait, Middle East, China, Japan, and Qatar. | Terminated | October 30, 2015 | 331 | 1064 | Not_Computer
Z-1676-2015 | 71229 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | D2-40 (Lymphatic Marker); For In Vitro Diagnostic; |  |  | May 06, 2015 | 2015 | Biocare Medical, LLC | Firm discovered that incorrect Instructions for Use were shipped with one lot of D2-40. | Labeling mix-ups | Biocare Medical sent a Medical Device Notification letter dated May 6, 2015 to their customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete all actions listed on the enclosed Return Response Form. |  | 19 | 19 units | Worldwide Distribution - US Nationwide in the states of: MI, SC, NC, AZ, CA, MO, WI and Hungary, Spain and Austria. | Terminated | June 04, 2015 | 29 | 19 | Not_Computer
Z-1664-2015 | 71256 | joint, ankle, external brace | ITW | Limb orthosis. | Physical Medicine | Physical Medicine | 510(K) Exempt | Elastic Ankle Support. |  |  | April 13, 2015 | 2015 | United Surgical Associates | The product does not bear the caution label regarding latex content. Potential for Allergic skin reaction | Error in labeling | On 4/13/2015 United Ortho emailed the affected distributor regarding the need for correction due to latex in the product and that updated labeling and instructions would be provided. United-Ortho personnel will travel to Hawaii and accompany the distributor to label the product at the sites to which it was distributed. Records will be maintained at each location. |  | 1278 | 1278 | Domestic: HIInternational; None | Terminated | July 14, 2015 | 92 | 3294 | Not_Computer
Z-1200-2015 | 70322 | No | OEZ | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | Enforcement Discretion | NAMIC Convenience Kit, 3V WMMII KIT, UPN H965601852521, REF/Catalog No. 60185252, STERILE, Rx ONLY --- For single use only. --- Device Listing Number D136588. |  |  | January 09, 2015 | 2015 | Navilyst Medical, Inc | The labelled expiration dating of the affected product is not fully supported by shelf life verification testing, which may result in a sterility breach. Use of a Fluid Management Convenience Kit, where sterility has been breached, may result in an increased risk of infection. | Packaging | Navilyst Medical sent an Urgent Medical Device Recall Notification letter dated January 9, 2015 to US customers (end users) via Federal Express. On January 20, 2015, Recall Notification Packages were sent via Federal Express to 2 International consignees (distributors). the letter identified the affected product, problem and actions to be taken. Customers are instructed to: " Immediately remove the recall product from your inventory (whether in labs, Central Supply, Shipping and Receiving or ANY other location)." Segregate this product in a secure location for return to Navilyst Medical." Immediately forward a copy of this recall notification to all sites to which you have distributed affected product." Complete and return the Reply Verification Tracking Form. |  | 7 | 7 units | Worldwide Distribution - US Nationwide and the countries of: Canada, and Guam. | Open, Classified |  |  | 16508 | Not_Computer
Z-1245-2015 | 70194 | Yes | FIK | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | AMIA v1.0 APD System. For Automated Peritoneal Dialysis (APD) therapy for pediatric and adult renal patients. |  |  | April 03, 2014 | 2014 | Baxter Healthcare Corp. | Additional warning in the AMIA APD System labeling for patients vulnerable to unintended Increased Intraperitoneal Volume (IIPV), also known as overfill. The patients include neonates, infants, small children, and any patient with significant heart and/or lung disease. | Error in labeling | Baxter sent IMPORTANT PRODUCT INFORMATION letters dated April 3, 2014 to all peritoneal dialysis healthcare providers of record. The letters advised healthcare providers of the potential to develop Increased Intraperitoneal Volume (IIPV) and instructed the healthcare providers to complete and return the attached Customer Reply Form. Healthcare providers with clinical questions can contact Baxter's Renal Division Clinical Helpline at 888-736-2543, option 2, 8:00 AM - 5:00 PM CT. Healthcare providers with general questions can contact The Center for One Baxter by telephone at 800-422-9837, Monday - Friday, 8:00 AM - 5:00 PM CT. Additionally, an IMPORTANT PRODUCT INFORMATION letters dated April 8, 2014 were sent to all peritoneal dialysis patients of record. The letters advised patients of the potential to develop Increased Intraperitoneal Volume (IIPV) and instructed the patients to complete and return the attached Home Patient Reply Form. Patients with questions about the letter can contact The Center for One Baxter by telephone at 800-422-9837, Monday - Friday, 8:00 AM - 5:00 PM CT. |  | 55 | 55 systems | US: Nationwide in the states of (GA, MA, NC, NY, PA, WA) | Terminated | July 21, 2017 | 1205 | 55 | Not_Computer
Z-1135-2105 | 70401 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | CoCr Femoral Head - indicated for hip arthroplasty for noncemented use in skeletally mature invididuals undergoing primary surgery for rehabilitating hips damaged as a result of noninflammatory degenerative joint disease |  |  | December 17, 2014 | 2014 | Zimmer Gmbh | Potential compromise of the inner sterile tray. If a compromised tray is detected, a delay in surgery may result. If undetected breach in both the inner and outer trays were to occur, and there is contamination present, a periprosthetic infection is possible, which could lead to revision. | Package design/selection | On 12/17/2014, URGENT MEDICAL DEVICE RECALL notifications were sent to the affected distributors and hospital Risk Managers with instructions for returning the affected product. All distributors were notified via electronic mail, and all hospital Risk Managers and distributors with affected inventory were also notified via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. For questions, call the customer call center at 1-800-348-2759. |  | 6109 | 6109 | Domestic:AK, AL, AR, AZ, CA, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MS, MA, MD, ME, MI, MN, MO, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, WV, WYInternational: Canada, Japan | Terminated | September 23, 2015 | 280 | 18403 | Not_Computer
Z-1757-2015 | 71171 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Biomet Cobalt-Chrome Femoral Components, 32MM MOD HEAD COCR -6MM NECK; 32MM MOD HEAD COCR -3 MM NECK; 32MM MOD HEAD COCR STD NECK ; 32MM MOD HEAD COCR +3MM NECK; 32MM COCR MOD HD +6MM NO SKIRT |  |  | April 16, 2015 | 2015 | Biomet, Inc. | Biomet hip joint, knee joint and shoulder joint replacement components device labeling contains incorrect GTIN (Global Trade Identification Number) codes. | Labeling mix-ups | Biomet sent an Urgent Medical Device Recall Notice letter dated April 29, 2015 to affected customers via Fed-Ex. The letter identifed the affected product, problem and actions to be taken. Customers were asked to complete the attached response form and return it along with recalled product to Biomet. For questions callt (574) 372-1570, Monday through Friday, 8 a.m. to 5 p.m. |  | 1315 | 1315 | Worldwide Distribution - US Nationwide in the states of AR, CA, IL, IN, PA, FL, KS, LA, WI, VA, MA, MD, NJ, UT, NC, TX, AL, MI, KY, OH, SD, MO, GA, NV, WA and the countries of: Europe, Canada, Japan, Chile, Korea, Mexico, New Zealand, Trinidad/Tobago, Thailand, Argentina, Australia, and China. | Open, Classified |  |  | 2632 | Not_Computer
Z-1822-2015 | 71257 | No | GAT | Nonabsorbable poly(ethylene terephthalate) surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ULTRA-FAST FIX Knot Pusher Suture Cutters |  |  | April 30, 2015 | 2015 | Smith &amp; Nephew, Inc. | Sterility of device maybe compromised | Package design/selection | Smith &amp; Nephew sent an Urgent - Product Recall 1st Notification letter dated April 27, 2015, via Federal Express on April 30, 2015. Users are requested to inspect inventory and locate any unused devices from the above listed product and batch numbers, and quarantine them immediately. Complete the the Inventory Return Certification and obtain a RGA and obtain replacement product.Contact Smith &amp; Nephews Field Action Department via email at FieldActions@smith-nephew.com or fax +1-901-566-7975 to obtain a return authorization (RA) number |  | 202279 | 202,279 units | Worldwide Distribution - US (nationwide) Distribution including Puerto Rico and to the countries of : Argentina, Austria, Australia, Belgium, Brazil, Canada, Switzerland, Chile, China, Colombia, Costa Rica, Dubia UAE, Germany, Denmark, Spain, Finland, France, United Kingdom, Greece, Guam, Hong Kong, Israel, India, Italy, Japan, Korea, Mexico, Malaysia, Netherlands, Norway, Portugal, Qatar, Sweden, Singapore,Thailand, Turkey, Taiwan, Venezuela and South Africa. | Open, Classified |  |  | 202279 | Not_Computer
Z-1124-2015 | 70087 | Yes | BTR | Tracheal tube. | Anesthesiology | Anesthesiology | 510(k) | RÂżSCH Curved Reinforced Endotracheal Tube: Rusch Reinforced Endotracheal Tube is designated for oral or nasal intubation and are indicated for airway management. The product may be used to reduce the potential for kinking whenever an unusual positioning of the head or neck is required following intubation. |  |  | December 12, 2014 | 2014 | Teleflex, Inc. | Teleflex Medical has issued an advisory notification for the RUSCH Curved Reinforced Endotracheal Tube because of a discrepancy between the labeled cuff diameter and the actual cuff diameter. The actual product cuff diameter meets required specification. | Labeling mix-ups | Teleflex Medical sent an Urgent Medical Device Notification dated December 12, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Please provide this Advisory Notice to all those who need to be aware of it within your organization. Please consider, clinicians, risk managers, supply chain/distribution centres, etc. in the circulation of this notice. There is no further action required.Clinical recommendation with regard to patient use is to continue the standard clinical practice of basing ET tube selection based on the tube size, carefully checking for the adequacy of ventilation and to change the ET tube size up or down as appropriate when intubating the patient.INSTRUCTION FOR DISTRIBUTORS OF AFFECTED PRODUCTIf you are a distributor, provide this advisory notice to all of your customers who have received product in scope of this action.Teleflex Medical is committed to providing high quality, safe and effective products. We sincerely apologize for any inconvenience this action may cause your operations. If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-866-246-6990. |  | 2292361 | 2,292,361 | Nationwide Distribution | Terminated | July 30, 2015 | 230 | 2292361 | Not_Computer
Z-1797-2015 | 71161 | No | MGQ |  |  | General &amp; Plastic Surgery | 510(k) | Amerigel Hydrogel Gauze Dressing 2 x 2, A5002. |  |  | February 25, 2015 | 2015 | Amerx Health Care Corp. | Nonapproved labeling claims (Sterile) | Labeling False and Misleading | On March 9, 2015 the firm contacted its customers via telephone. Within that conversation the firm asked its customers to examine their inventory to determine if they have any of the affected lot numbers on hand and notify the firm for an immediate product substitution, credit or refund. Customers were asked to discontinue dispensing any product on hand, and notify any customers that the product may have been further distributed to. The firm also asked that the response form be completed and returned. If no contact was made via telephone, the firm sent a certified letter asking the firm(s) to do the same actions as the phone request. If you have any questions please call 1-800-448-9599 or e-mail marketcorrection@amerxhc.com. |  | 6491 | 6,491 units | Worldwide Distribution -- US, including the states of AK, AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MO, MS, MT, NC, NE, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WV, WY and Hawaii; and, the countries of Canada and Bahrain. | Open, Classified |  |  | 6491 | Not_Computer
Z-1121-2015 | 69782 | No | KRB | Thermodilution probe. | Cardiovascular | Cardiovascular | 510(k) | Pulsiocath PiCCO ProAQT Monitoring Kit, single-use cardiac output (CO) sensor, one unit per plastic sterilizable package. |  |  | November 04, 2014 | 2014 | Pulsion Medical Inc | Internal testing and investigation of the packaging revealed cracks and holes in primary packaging | Packaging | Pulsion sent an Urgent Medical Device Recall notification letter on 11/04/14 to affected customers. The letter identified the affected product, problem and actions to be taken. Customers where instructed to:" Piease examine your inventory immediately to determine if you have any of the products listed." lf so, please remove the affected products; quarantine them in a secure location." Piease complete and return the enclosed Field Recall Response Form, acknowleding your receipt and understanding of this communication " Send all affected products back to PULSION Medical Systems." Upon return of the affected product you will receive a credit note.For questions contact your PULSION Medical lnc Recall Coordinator. |  | 90 | 90 units | US Nationwide Distribution in the states of CT, NJ and NY. The affected product was distributed to one government consignee. | Terminated | October 02, 2015 | 332 | 90 | Not_Computer
Z-1117-2015 | 70275 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | HARMONIC ACE Curved Shears Instructions for Use, 5MM, IS1200; for use with the da Vinci Standard Surgical System. |  |  | January 21, 2015 | 2015 | Intuitive Surgical, Inc. | Internal labeling review found that the IFU for the Harmonic Ace incorrectly instructs the user on how to manually open the grips. | Error in labeling | A Field Safety Notice entitled, Urgent Medical Device Correction, and dated, January 21, 2015, were sent to customers. The letter identified the problem and the reason for the recall as well as the risk to health. Customers are asked to ensure that all affected personnel are informed of the notice. Additionally, customers are to discard their current Harmonic Ace IFUs. Customers are to refer to the specified manual, reference guide, and instruction card mentioned in the letter. A copy of the notice is to be retained with the da Vinci Standard, S or Si user manual. Intuitive Surgical will remove the Harmonic ACE IFU from all future shipments of the product. Representatives will be available by phone to answer questions related to the Correction. |  | 12467 | 12,467 - both sizes | Worldwide Distribution -- US, Australia, Austria, Belgium, Brazil, Canada, Chile, China, Colombia, Czech Republic, Denmark, Dominican Republic, Ecuador, Finland, France, Germany, India, Israel, Italy, Japan, Lebanon, Mexico, Monaco, Netherlands, Pakistan, Panama, Puerto Rico, Qatar, Romania, Russian Federation, Saudi Arabia, Singapore, Slovakia, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, and the United Kingdom. | Terminated | June 01, 2015 | 131 | 12467 | Not_Computer
Z-1114-2015 | 69723 | No | HSD | Shoulder joint humeral (hemi-shoulder) metallic uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Arthrosurface HHXL (OVO) 15.6x32mm Taper Post Fixation Component; Catalog Number: 8156-0032, Part Number: 156-0032. HemiCAP Contoured Articular Prosthetic for Humeral head resurfacing. Orthopedic shoulder prosthesis component. |  |  | November 24, 2014 | 2014 | Arthrosurface, Inc. | Fully threaded Taper Post Fixation components not cleared for marketing in the US with current indications. | No Marketing Application | Arthosurface sent an Urgent Medical Device Recall letter on November 24, 2014 via UPS, to all affected customers. The letter informed consignees of the recall and asked them to return two completed forms, 1. Acknowledgement &amp; Receipt Form and 2. Inventory Record via fax, email, or mail. Customers with questions were instcted to contact their local sales representatiave or email RA@arthrosurface.com. For questions regarding this recall call 508-528-4604. |  | 187 | 187 | Worldwide Distribution - USA nationwide. Internationally distributed to Europe, Canada, and Australia.Only products in US affected by this recall. | Terminated | March 27, 2015 | 123 | 786 | Not_Computer
Z-1094-2015 | 70055 | No | MTJ |  |  | General &amp; Plastic Surgery | 510(k) | Hemostatic Bone Putty, for use as a water-soluble implant material and for use in the control of bleeding from bone surfaces. |  |  | January 06, 2012 | 2012 | Synthes, Inc. | Revision A of the Hemostatic Bone Putty Brochure (J10847A) was sent to sales consultants instead of Revision B (J10847B), which corrected an error in the indication statement. | Labeling mix-ups | On/about 12/22/2011, emails containing Revision B brochure, as well as instructions to delete/destroy Revision A, were sent to 38 affected individuals (sales consultants). On/about 1/04/2012, follow up phone calls were made to each member of the sales force who had received Rev. A (110847 A) to confirm that they had replaced the Rev. A with the Rev. B (110847 B). In addition, the members of the sales force were asked whether or not the Rev. A had been disseminated to any HCP's. During the follow up conversation, one member of the sales force stated that he had provided the previous version to a surgeon. On 01/06/2012, the sales consultant spoke to the surgeon in person and the surgeon confirmed that he had disposed of the Rev. A. |  | 38 | 38 | Distributed in the state of WI | Terminated | March 30, 2015 | 1179 | 38 | Not_Computer
Z-1080-2015 | 70177 | No | CAZ | Anesthesia conduction kit. | Anesthesiology | Anesthesiology | 510(k) | ARROW Epidural Needle Component (product number AN-05501). Epidural needles permit access to the epidural space and are used as a conduit for epidural catheter placement. |  |  | January 05, 2015 | 2015 | Arrow International Inc | The product labeling does not specify the 5 year shelf life of the product. | Labeling design | Urgent Medical Device Recall Notifications, dated January 2015, were sent to direct accounts that identified the product, problem, and actions to be taken. Customers were asked to discontinue use and return affected product. |  | 63022 | 63022 | Distribution US Nationwide, Canada and Mexico. | Open, Classified |  |  | 63022 | Not_Computer
Z-1264-2015 | 70562 | stain, papanicolau | HZJ | Dye and chemical solution stains. | Hematology | Pathology | 510(K) Exempt | Carton: Richard-Allan Scientific CYTO-STAIN; REF 7501R; Quantity: 1x1 Gallon; Size: 1 Gallon/3.8L; LOT 313833; 11/2016; Bottle: Richard-Allan Scientific CYTO-STAIN G; REF 7501G; 1 Gallon/3.8L; Thermo Scientific. |  |  | February 19, 2015 | 2015 | Richard-Allan Scientific Company | The case was correctly labeled as REF 7501R, however the gallon bottle was incorrectly labeled as REF 7501G Cyto-Stain G. The bottle label has the correct lot number of 313833 and EXP date 11/2016. The gallon bottle contains the correct intended contents of 7501R: Cyto-Stain R Stain. | Error in labeling | Urgent Medical Device Recall letters sent to Richard Allan distributors and customers. The recall notification includes the reason for the voluntary recall, risk to health, product information and actions to be taken by the distributor/customer. |  | 132 | 132 | US, Canada and Europe | Terminated | January 12, 2016 | 327 | 132 | Not_Computer
Z-1814-2015 | 71329 | No | HDW | Contraceptive diaphragm and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | MILEX OMNIFLEX DIAPHRAGM SIZE 80 |  |  | May 20, 2015 | 2015 | CooperSurgical, Inc. | A diaphragm size 80 labeled box was incorrectly packaged with a size 85. | Labeling mix-ups | CooperSurgical notified consignees by letter on 5/18/15 via Fedex with confirmed delivery receipt.. Accounts requested to discontinue use and complete the attached Acknowledgement and Receipt Form for replacement. CooperSurgical provided further instruction on the disposition of the product at the time of replacementIf you have any further questions please feel free to contact 203.601.5200.. |  | 200 | 200 units | AK, CA, DC, FL, IL, ME, MT, NE, NJ, NM, NY, OK, OR, P A, VA, W A, and WIForeign: CANADA, ENGLAND, and POLAND | Terminated | January 20, 2016 | 245 | 200 | Not_Computer
Z-1077-2015 | 70332 | No | FYA | Surgical apparel. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Flyte Togas are components of the Stryker Flyte System and are intended to be worn over any Stryker Flyte Helmet. This toga provides Level 1 and 4 protection according to the Association for the Advancement of Medical Instrumentation (AAMI) Liquid Barrier Performance and Classification System (PB70:2012). |  |  | December 22, 2014 | 2014 | Stryker Instruments Div. of Stryker Corporation | A single lot (11121171) of Flyte Toga Peel-away X-Large (catalog number 0408-830-100) was labeled with an incorrect expiration date of 2412-12-01. The correct expiration date should have read 2014-12-01. | Error in labeling | Stryker sent an Urgent Medical Device Recall Notification letter dated December 22, 2014, to all affected customers. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and responding to the formal recall notification. Customers were instructed to do the following:1. Immediately review this Recall Notification2. Immediately check all stock areas and/or operating room storage and quarantine any affected product found.3. Complete the enclosed Business Reply Form (BRF) to confirm receipt of this notification and identify howmany, if any, affected items are currently in their inventory. Complete and return the BRF even if theydo not have any affected product on hand.Note: Your signature on the BRF indicates that you received and understand this Notification andhave followed the instructions in this Notification.4. If customers further distributed this product, they were instructed to forward this letter and the attached Business Reply Form (BRF) to all affected locations. Please indicate each location on the BRF. Completed Business Reply Forms should be faxed to 866-521-2762.Upon receipt of the Business Reply Form, Stryker will email a pre-paid shipper. This shipper can be used to return the recalled Flyte Toga(s) to Stryker.Upon receipt of the recalled Flyte Toga(s), credit will be applied to the customer's account.For questions regarding this recall customers were instructed to contact Stryker Instruments: Monday-Friday 8am-5pm (EST) at 269-389-2458.For questions regarding this recall call 269-389-2458. |  | 1530 | 1530 units (170 cases) | Nationwide Distribution including AL, AR, AZ, IA, IL, IN, LA, MA, NC, NE, NM, SC, TN, TX, WI, and WV. | Terminated | May 05, 2015 | 134 | 1530 | Not_Computer
Z-1076-2015 | 70051 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | The SurgiWrap Bioresorbable Adhesion Barrier Film. SurgiWrap 100x130x0.05mm |  |  | August 26, 2014 | 2014 | MAST Biosurgery USA, Inc | The expiration date of the bar code on the envelope label did not match the expiration date next to the hourglass. The printed expiration date was August of 2017, while the bar code date read as July 31, 2017. The correct expiration date was July 31, 2017. | Labeling Change Control | MAST Biosurgery labeled 49 new envelopes and sent them to the distributor, and subsequently had the distributor remove the contents of the outer envelope (instructions for use and the foil pouches containing the product) and place them into the newly supplied and properly labeled outer envelopes provided by them. The distributor requested to send back the original 49 enveloped.The second batch of new labels was inspected and attached to the new outer envelope to create a properly labeled outer envelope prior to shipping. The firm requested that once the distributor replaced the product and the IFU into the newly labeled envelopes that they provide documentation that the product had been repackaged into the new envelopes. The distributor was also requested to verify that they did not ship any product to customers with the incorrect label. The firm had the distributor return the 49 improperly with the incorrect labels and the first set of "corrected" labels to them for destruction.Further questions call (858) 550-8050. |  | 49 | 49 pieces | Worldwide Distribution to the country of China. | Terminated | February 24, 2015 | 182 | 49 | Not_Computer
Z-1818-2015 | 71132 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Viking M, L, and XL Mobile Lifts. |  |  | March 20, 2015 | 2015 | Hill-Rom, Inc. | Complaints (including one reported death) allegedly of the lift arm drifting down suddenly. If the lift arm assembly is manually lifted, the actuator can become damaged and get stuck in the highest position. If a patient is lifted into the sling while the actuator is stuck, there is a potential risk of a free fall of the patient, resulting in minor or potentially catastrophic injuries. | Labeling design | On 3/20/2015, URGENT Field Safety Notice notifications, dated 3/19/2015 were sent to US customers by e-mail and or courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Customers are forward a copy of the letter to any other appropriate facility personnel. Additionally, the attached form should be completed and returned to Hill-Rom. Upon receipt, the firm will contact customers to schedule an inspection of the potentially affected devices. For questions, call 800-445-3720 (Mon - Thurs, 8 am - 6:30 pm; Friday, 8 am - 6 pm). |  | 24798 | 24, 798 | Worldwide Distribution -- AE, AO, AR, AT, AU, AZ, BD, BE, BG, BH, BR, CA, CH, CL, CN, CO, CS, CZ, DE, DK, EG, ES, FI, FR, GR, HK, HR, HU, ID, IE, IL, IN, IQ, IR, IS, IT, JO, JP, KR, KW, LB, LI, LT, LU, LV, MA, MT, MX, MY, NL, NO, NZ, OM, PH, PL, PT, QA, RO, SA, SE, SG, SI, SK, SY, TH, TN, TR, TW, UK, US, VE, and ZA. | Terminated | August 04, 2017 | 868 | 24798 | Not_Computer
Z-1826-2015 | 71519 | No | DRE | Vessel dilator for percutaneous catheterization. | Cardiovascular | Cardiovascular | 510(k) | Merit Medical Prelude Pro Sheath Introducter, 6F (2.0mm) 11cm, REF PRO-6F-11-038, Sterile EO |  |  | June 04, 2015 | 2015 | Merit Medical Systems, Inc. | Some of these introducer sheaths were packaged with a 0.035 inch dilator, rather than the labeled 0.038 inch. This affects one lot, H767887, of catalog number PR0-6F-11-038. | Packaging process control | Merit Medical sent an Urgent Product Recall letter on June 4, 2015, to all affected consignees. Consignees were instructed to immediately quarantine any affected devices and discontinue use, ensure all personnel to whom devices were distributed are made aware of this field action, and instructions to contact their Merit representative. An email notification was sent to Merit sales representatives with accounts that had received affected product. The affected sales representatives were provided a Customer Response Form applicable to their account and a representative copy of communications sent directly to their affected accounts. Merit sales representatives were instructed to visit the account and assist with the return of the affected units back to the Merit Field Assurance Department. Customers with questions were instructed to contact their Merit Sales Representative or call (801) 208-4365. |  | 1445 | 1,445 | Worldwide Distribution - US including TN, NE, MA, SC, WI, UT and Internationally to Belgium, Switzerland, United Kingdom, and Germany. | Terminated | October 26, 2015 | 144 | 1445 | Not_Computer
Z-1078-2015 | 70248 | Yes | BTR | Tracheal tube. | Anesthesiology | Anesthesiology | 510(k) | Hudson RCI, Sheridan, Preformed Tracheal Tubes, Cuffed and Uncuffed, Rx Only, Sterile, Teleflex Medical. |  |  | January 09, 2015 | 2015 | Teleflex Medical | Labeling; The affected product packaging is incorrectly labeled as being a cuffed oral ET tube; The correct product labeling is uncuffed nasal ET tube. | Labeling mix-ups | Teleflex Medical sent an " Urgent Medical Device Recall Notification" dated January 9, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.They were instructed to immediately discontinue use and quarantine any of the affected products, return the product and to complete the enclosed Recall Acknowledgement Form even if they do not have the affected product in stock. Distributors were instructed to communicate the recall to any of their customers who may have received product included within the scope of the recall. Have the customers return any affected product to them, together with a completed Recall Acknowledgement Form.Teleflex Medical is committed to providing high quality, safe and effective products. We sincerely apologize for any inconvenience this action may cause your operations. If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-866-246-6990. |  | 300 | 300 units | Worldwide Distribution - US Distribution to the states of FL, GA, IL, KY, MA, MI, MO, MS, OH, TX, WA and WV., and to the countries of Canada and the Netherlands. | Open, Classified |  |  | 300 | Not_Computer
Z-1940-2015 | 71383 | No | HQC | Phacofragmentation system. | Ophthalmic | Ophthalmic | 510(k) | Accurus 2500 Standalone Vitreous Probe REF 8065741018; Accurus 23 GA Standalone Vitreous Probe REF 8065750821; Accurus 25+ 2500 CPM Standalone Vitreous Probe REF 8065751122; Anterior Accurus Probe with Infusion Needle REF 8065803650; Single Use Sterile Probes Rx only. |  |  | June 02, 2015 | 2015 | Alcon Research, Ltd. | Insufficient seal on the outside packaging, potentially affecting the sterility of the product. | Packaging | Recall notification letters were sent to consignees on June 2, 2015. |  | 35475 | 35,475 units | Worldwide Distribution-US (nationwide) and the countries of Argentina, Australia, Austria, Brazil, Canada, Chile, China, Colombia, Czech Republic, France, Germany, Hong Kong, Hungary, India, International APL, Italy, Japan, Kenya, Korea, Kosovo, Malaysia, Mexico, Morocco, Netherlands, New Zealand, Poland, Russia, Spain, Sri Lanka, Switzerland, Taiwan, Thailand, Turkey, and United Kingdom. | Terminated | April 20, 2017 | 688 | 35475 | Not_Computer
Z-1234-2015 | 70466 | No | KGX | Medical adhesive tape and adhesive bandage. | General Hospital | General &amp; Plastic Surgery | 510(K) Exempt | Zip 8i Surgical Skin Closure Device; |  |  | February 05, 2015 | 2015 | Zipline Medical | The manufacturers seal may not be present on the Nylon pouch resulting in lack of assurance for sterility. | Packaging process control | Zipline sent an Urgent Voluntary product Recall letters to customers on February 5, 2015 by e-mail with possible follow up on site by firm representatives. The letter identified the affected product, problem and actions to be taken. Customers are asked to quarantine inventory immediately and examine the affected lots for complete pouch seals. Any product found with an incomplete seal should be destroyed or returned to ZipLine Medical. Please complete the enclosed response form as soon as possible, even if you no longer have stock of the lot numbers subject to this correction. This information is required to reconcile the inventory. Please contact Customer Service at 1-888-326-8999 to coordinate any product returns, replacement stock, or credit notes. |  | 3109 | 3109 - total - all model numbers | Worldwide Distribution - US Nationwide and the countries of Panama, Tehran, Ireland, Sweden, Japan, Hong Kong, Finland, Pakistan, Taiwan, Malaysia. | Terminated | April 10, 2015 | 64 | 3109 | Not_Computer
Z-0998-2015 | 70080 | No | LIT | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Cordis POWERFLEX PRO PTA Dilatation Catheter |  |  | December 22, 2014 | 2014 | Cordis Corporation | One unit from a demo lot of POWERFLEX PRO PTA Dilatation Catheter was left at a hospital and was not sterilized. | Labeling design | The firm, Cordis, sent an "URGENT MEDICAL DEVICE RECALL (Removal)" letter dated December 29, 2014 to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to immediately check their inventory to confirm they do not have any units from the demo lot in possession; identify and set aside any unit from the identified lot; review, complete, sign and return the enclosed Acknowledgement Form directly to Cordis at the fax number on the form; return any affected product to Cordis; share this letter with others in your facility; contact any other facility that has been provided with units of POWERFLEX PRO product; maintain awareness of this notice until all affected product has been returned to Cordis and keep a copy of this notice witn any affected product.For any physician with medical concerns, please contact the Cordis Endovascular Medical Information Line 877-DEVICES Monday through Friday from 9am-5pm. For questions related to the recall that are not adequately addressed in the letter, please contact the Cordis Field Action Coordinator Hal Baden at (786) 313-2365. For questions related to the Acknowledgement Form and its return, please contact Cordis QA at (786) 313-8730. |  | 120 | 120 | Nationwide Distribution: US including states of: CT, NJ, NY, PA, MD, VA, WV, NC, SC, GA, FL, AL, TN, MS, KY, OH, IN, MI, IA, WI, SD, IL, MO, KS, NE, LA, OK, TX, CO, UT, AZ, CA, OR, WA, AR, IN, MN, WY, ND, NV, MA, RI, NH, CT, DE, District of Columbia and Hawaii. | Terminated | July 09, 2015 | 199 | 120 | Not_Computer
Z-2076-2015 | 71593 | Yes | DWC | Cardiopulmonary bypass temperature controller. | Cardiovascular | Cardiovascular | 510(k) | Sorin Stockert Heater-Cooler 3T, 230 V |  |  | June 15, 2015 | 2015 | Sorin Group USA, Inc. | Potential colonization of organisms, including Mycobacteria, in Sorin Heater Cooler Devices, if proper disinfection and maintenance is not performed per Instructions for Use. | Error in labeling | Sorin Group issued a Field Safety Notice dated June 15, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review their inventory and identify any affected devices. For each unit customers were instructed to determine if the device has been maintained according to the Instructions for Use. If yes, customers should strictly adhere to the new Instructions for Use. Customers were also provide with a Response form to confirm they received, read and understood the Field Notice. Customers were instructed to return the completed form to assist in monitoring the effectiveness of the communication. For technical support customers should call 1-800-221-7943, ext 6355. For questions regarding this recall call 303-467-6306. |  | 2837 | 2837 | Worldwide Distribution - US (nationwide) and Internationally to AL, AT, BE, BG, CH, CY, CZ, DE, DK, EE, ES, FI, FR, GB, GR, HR, HU, IE, IT, LT, LU, LV, MC, MK, MT, NL, NO, PL, PT, RO, RS, SE, SK, AE, AR, AU, AZ, BD, BH, BR, BY, CA, CL, CN, CO, CR, DZ, EC, EG, ET, GE, GY, HK, ID, IL, IN, IQ, IR, JO, JP, KR, KW, KZ, LB, LK, LY, MA, MN, MU, MX, MY, NG, NP, NZ, OM, PA, PE, PH, PK, PR, PS, QA, RE, RU, SA, SG, SV, SY, TH, TN, TR, TT, TW, UA, VN, ZA. | Open, Classified |  |  | 5055 | Not_Computer
Z-1516-2015 | 70838 | No | MGQ |  |  | General &amp; Plastic Surgery | 510(k) | Iamin Moist Dressing and the Iamin Hydrating Gel: intended to cover a wound or burn on the skin, provide a moist wound environment and protect against abrasion, friction, desiccation and contamination. |  |  | March 26, 2015 | 2015 | PhotoMedex, Inc. | The labeling for the Iamin family of medical devices (e.g. printed and internet text) is not consistent with the intended uses and limitations identified at the time of the clearance to market these devices. | Labeling design | PhotoMedex sent an Urgent product Recall notification, dated March 2015, to direct consignees requesting they discontinue use, complete the response form, and return affected products. |  | 200188 | 200188 | Nationwide and Internationally, including Canada. | Terminated | August 18, 2015 | 145 | 200188 | Not_Computer
Z-0937-2015 | 69979 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Stryker Orthobiologics Imbibe Aliquot Needle Bone Cement Needle- Model Numbers: 2110-0505 Beveled Needle, 11 gauge x 4 inch |  |  | December 04, 2014 | 2014 | Orthovita, Inc., dBA Stryker Orthobiologics. | There is the potential for a breach in the inner or outer packaging pouches of all lots of the Stryker Orthobiologics Imbibe Aliquot Needle Bone Cement Needle. | Package design/selection | The firm, Stryker, sent an "URGENT PRODUCT RECALL" letter dated December 4, 2014, to end users/customers. The letter identified the product, problem, and actions to be taken. Customers were instructed to complete and return the URGENT MEDICAL DEVICE RECALL NOTIFICATION ACKNOWLEDGMENT FORM within 5 days via fax to 215-253-5020, contact other hospitals in your territory that have the affected products to arrange return of the product, and return the affected product to: Stryker Orthobiologics, 67 Great Valley Parkway, Malvern, PA 19355. If you have any questions, feel free to contact the Associate Manager, Quality Assurance and Regulatory Compliance at 610-640-1775 x 5299. |  | 2155 | 2155 | Worldwide Distribution: US (nationwide) and country of: Canada. | Terminated | July 28, 2015 | 236 | 91925 | Not_Computer
Z-2092-2015 | 71366 | No | HTO | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Reamer / Irrigator / Aspirator (RIA): RIA Drive Shafts, RIA Tube Assemblies, sterile, RIA Reamer Head, sterile (12.0mm-19.0mm); intended to clear the medullary canal of bone marrow and debris and to effectively size the medullary canal for the acceptance of an intramedullary implant or prosthesis |  |  | May 28, 2015 | 2015 | Synthes, Inc. | It was reported that the RIA Drive Shaft, Tube Assembly, and Reamer Head have the potential to break when incorrectly assembled or used improperly. Precautionary statements are being added to the Reamer / Irrigator / Aspirator (RIA) Surgical Technique Guide . | Labeling False and Misleading | DePuy Synthes sent an Urgent Medical Device Safety Notification dated May 28, 2015, to direct accounts and sales consultants to advise them of the patient risk associated with misassembly or misuse of the device. Customers and sales force were instructed to remove the Surgical Technique Guide J4352 (all versions) dated 11/2011 or earlier and replace it with the updated Technique Guide (DSUS/TRM/0515/0620, Date 06/2015),and to return the completed verification response form.If you have any further questions, please call (610) 719-5450. Thank you for your attention to this matter. |  | 132564 | 132564 | Nationwide Distribution and Internationally to Canada and Switzerland. | Open, Classified |  |  | 132564 | Not_Computer
Z-2090-2015 | 71502 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | BD Flow Cytometers; used on all BD FACSCount, BD FACSCalibur, BD FACSCanto, BD FACSCanto II, BD LSR II, BD LSRFortessa,BD FACSAria, BD FACSAria II, BD FACSAria Ill, BD lnflux, BD FACSArray,BD Pathway 855, BD Pathway 435 and all SORP instruments. |  |  | June 15, 2015 | 2015 | BD Biosciences, Systems &amp; Reagents | Instruments installed between June 2007 and September 2011 do not contain labeling pursuant to the Food and Drug Administration's (FDA's) Guidance on Laser Products- Conformance with IEC 60825-1 and IEC 60601-2-22 (Laser Notice No. 50). | Labeling Change Control | BD Biosciences sent an Important Product Correction letter dated June 2015 to all affected customers. Customers were informed that a Field Service Engineer will schedule a time to perform the necessary correction. Customers with questions were instructed to contact BD Customer Support at 877-232-8995 (prompt 2). For questions regarding this recall call 408-954-6307. |  | 2188 | 2188 units | Nationwide Distribution | Open, Classified |  |  | 2188 | Not_Computer
Z-0939-2015 | 69468 | Yes | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | HomeChoice Automated PD system and HomeChoice Pro Automated PD system Baxter's |  |  | October 06, 2014 | 2014 | Baxter Healthcare Corp. | There are additional and updated warning and cautions that are not in the Patient At-Home Guide for HomeChoice Automated PD Systems and HomeChoice PRO Automated PD Systems. | Labeling design | The firm, Baxter, sent an "IMPORTANT PRODUCT INFORMATION" letter dated October 6, 2014 to all peritoneal dialysis provider direct accounts. Baxter also sent an "IMPORTANT PRODUCT INFORMATION" letter dated October 9, 2014 to all peritoneal dialysis patient direct accounts The letters described the product, problem and actions to be taken. The letters included instructions for the recipients to complete the attached Customer Reply Form and to return it to Baxter via scanned e-mail to: fca@baxter.com or fax : 224-270-5457. In addition, the letter to patients instructed them to contact their healthcare provider to discuss this issue.Customers with clinical questions should contact Baxter's Renal Division Clinic Helpline at 888-736-2573, option 2 (8:00 AM - 5:00 PM CST). Adverse reaction or quality problems experienced with the use of these product can be reported by contacting Baxter Product Surveillance at 800-437-5176 (Monday - Friday, 8:00 AM - 5:30 PM CST) or e-mail to corporate_product_complaints_round_lake@baxter.com. Customers with general questions about the letter should contact The Center for One Baxter at 800-422-9837 (Monday - Friday, 8:00 AM - 5:00 PM CST). |  | 96458 | 96,458 systems | Worldwide Distribution: US (nationwide) Including states of: AK, AL, AR, AZ, CA, CO, CT, DE, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY and District of Columbia; and countries of: Algeria, Argentina, Aruba, Australia, Austria, Bahamas, Bahrain, Barbados, Belarus, Belgium, Bosnia, Brazil, Bulgaria, Cayman Islands, Chile, Colombia, Croatia, Cuba, Czech Republic, Denmark, Ecuador, Egypt, El Salvador, Estonia, Finland, France, Georgia, Germany, Greece, Guatemala, Honduras, Hong Kong, Hungary, India, Iran, Iraq, Ireland, Israel, Italy, Jamaica, Japan, Jordan, Kazakhstan, Kuwait, Lebanon, Macedonia, Malaysia, Malta, Mexico, Morocco, Netherlands, New Zealand, Nicaragua, Norway, Oman, Panama, Poland, Portugal, Quatar, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, and Venezuela. | Terminated | March 28, 2016 | 539 | 96458 | Not_Computer
Z-1380-2015 | 70777 | No | KKX | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | da Vinci S Surgical System IS2000, Arm Drape, 20 Pack. |  |  | March 16, 2015 | 2015 | Intuitive Surgical, Inc. | Cloudy/waxy appearance with potential to transfer waxy substance to the patient, and potential for drape to tear. | Packaging process control | Firm issued field safety notice to all customers on March 16, 2015. Field safety notice includes acknowledgement form to be returned to recalling firm. A second letter was sent on May 12, 2015 to inform users that the date that the drape product unaffected by the issue will not be availalble untisl October/November 2015, not May 2015 as stated in the original communication. |  | 442475 | 442,475 total all drape models | Worldwide Distribution-Argentina, Australia, Austria, Belgium, Brazil, Bulgaria, Canada, Chile, China (including Hong Kong), Colombia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, Finland, France, France, Germany, Greece, Iceland, India, Indonesia, Ireland, Israel, Italy, Japan, Kuwait, Lebanon, Luxembourg, Malaysia, Mauritius, Mexico, Monaco, Netherlands, New Zealand, Norway, Pakistan, Panama, Philippines, Poland, Portugal, Puerto Rico, Qatar, Romania, Russia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, United States, Uruguay, Venezuela, Vietnam. | Terminated | December 03, 2015 | 262 | 442475 | Not_Computer
Z-1325-2015 | 70605 | No | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | NexGen TRABECULAR METAL FEMORAL AUGMENT BLOCK DISTAL ONLY, 20 MM Augment, Size D, Sterile, Rx Only. |  |  | February 04, 2015 | 2015 | Zimmer Trabecular Metal Technology, Inc. | One unit of the 20 mm augments was mislabeled as 5 mm and it was subsequently distributed. | Error in labeling | Zimmer Trabecular Metal Technology sent an email dated 12/10/2014 to their sole customer. The customer was requested to return the affected product. |  | 1 | 1 unit | US Distribution in the state of CA | Terminated | July 16, 2015 | 162 | 1 | Not_Computer
Z-1538-2015 | 70960 | No | FAE | Penile rigidity implant. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Genesis Malleable Penile Prosthesis, 9.5 mm, Catalog /Item Number 91-9509SC/519250. |  |  | April 09, 2015 | 2015 | Coloplast Manufacturing US, LLC | Several lots of the 9.5 mm Genesis Malleable Penile Prosthesis were mislabeled. The Retail Box and Patient Labels indicate a size 9 mm, instead of 9.5 mm. The device size is printed directly on the individual prosthesis tray during the packaging process and states the correct size, 9.5 mm. | Error in labeling | Consignees were notified via Field Representatives starting on 4-9-2015 with a Coloplast "Urgent Correction Notification" letter dated April 7, 2015. The letter described the problem and the product involved in the recall. Requested consignees to complete the Return Form if returning inventory. For questions contact Customer Service at 800-258-3476. |  | 11 | 11 (6 USA, 5 OUS) | Worldwide - US Nationwide Distribution in the states of FL, NY, TX and the country of FRANCE | Terminated | September 11, 2015 | 155 | 11 | Not_Computer
Z-2067-2015 | 71527 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Dual Delivery 4DD-1 of GYC-1000 Green Laser Photocoagulator system |  |  | March 25, 2015 | 2015 | Nidek Inc | Accessories to the GYC-1000 laser were missing Laser Aperture labels. | Error in labeling | Nidek sent Recall Notifications to doctors on March 31, 2015, by Federal Express. Letters were specific to the accessory involved.NIDEK detected that the laser aperture label has not been affixed on the delivery unit of GYC-1000, Green Laser Photocoagulator System and the device does not meet FDA regulation 21 CFR Subchapter J.In order to comply with FDA regulation, Nidek requests that the customer affix the enclosed laser aperture label to the Dual Delivery 4DD-1 of GYC-1000, according to the instructions attached.Please see the enclosed Laser Aperture label, instruction for affixing to the device and the response card.Nidek requests that you complete and return the response card as soon as possible. For additional information, please call Nidek at 1-800-223-9044 or visit http://usa.nidek.com |  | 22 | 22 units | US Distribution to the states of : CA, IL, MN, MO, OR, IL, UT RI, MA, WA, NE, NY, NJ, TX and WI. | Terminated | July 14, 2015 | 111 | 69 | Not_Computer
Z-2064-2015 | 71407 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Apex ARC Hip Stem, HA Coated, Size 5; Product Code: HC-21005-HA |  |  | June 03, 2015 | 2015 | Omnilife Science Inc. | Product may breach the inner and outer sterile pouches during shipping or handling. | Packaging | All product is consigned to US Agents until implanted or returned to OMNI. All US Agents holding the potentially affected inventory were notified that the product is being recalled via letter on June 3, 2015. The letter instructs them to immediately discontinue use of the product and contact OMNI Customer Service at 800-448-6664 to request a return materials authorization, and to return the Acknowledgement and Receipt form by FAX to 508-819-3390 or email to cflores@omnils.com. |  | 146 | 146 devices | US Nationwide Distribution in the states of: UT, CA, FL, TX, MA, GA, OK, VA, CT, OK, and WY. | Terminated | September 08, 2015 | 97 | 146 | Not_Computer
Z-2066-2015 | 71493 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | G7 Variant Elution Buffer HiS no. 1 (M) component of the Tosoh Automated Glycohemoglobin Analyzer HLC 723G7. |  |  | February 10, 2015 | 2015 | Tosoh Bioscience, Inc. | The G7 Variant Elution Buffer HiS no. 1 (M) is incorrectly labeled with an expiration date of October 2015. The correct expiration date is October 2014. | Error in labeling | Tosoh Bioscience sent an Urgent Recall Notification letters dated February 10, 2015 to their customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to discard any remaining pouches of the product, complete and return the attached response form. For questions call Tosoh Technical Support at 1-800-248-6764. |  | 45 | 9 boxes/5 pouches | Us Nationwide Distribution in the state ID including PR. | Terminated | October 01, 2015 | 233 | 45 | Not_Computer
Z-2062-2015 | 71271 | No | FQH | Jet lavage. | General Hospital | General Hospital | 510(K) Exempt | 0210-312-000-Woundcare Replacement kit with soft tip; 0210-318-000; Woundcare Replacement with fan tip; 0210-318-100- Woundcare Replacement Kit; 0210-318-200-Woundcare Replacement with retract coax tip. Accessory to Stryker InterPulse Irrigation System, used for wound debridement, soft tissue debridement, and cleansing of medical, clinical, or surgical sites. |  |  | May 08, 2015 | 2015 | Stryker Instruments Div. of Stryker Corporation | Potential sterility breach in the packaging. | Package design/selection | Stryker sent Customer Notification Letters on 05/20/2015 through Certified mail via USPS.Customers are asked to do the following: Immediately review this Recall Notification. Immediately check all stock areas and/or operating room storage and quarantine any affected product found. Complete the enclosed Business Reply Form (BRF) to confirm receipt of this notification and identify how many, if any, affected items are currently in your inventory. Please complete and return the BRF even if you do not have any affected product on hand. If you have further distributed this product, please forward this letter and the attached Business Reply Form (BRF) to all affected locations. Please indicate each location on the BRF. Fax (866-521-2762) or email (julie.forsyth@stryker.com) the completed Business Reply Form to Stryker Instruments Regulatory Department.Upon receipt of the Business Reply Form, if you have recalled product on hand, Stryker will email a pre-paid shipper. This shipper can be used to return the recalled WoundCare Replacement kits to Stryker. |  | 5328 | 5,328 each (444 boxes) | Nationwide Distribution-including the states of FL, IA, IL, IN, LA, MD, MN, MO, NC, NM, NY, PA, TN, TX, VA, and WI. | Terminated | October 19, 2015 | 164 | 5328 | Not_Computer
Z-2058-2015 | 71423 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | A.L.P.S., 3.5mm x 60 mm Cortical Locking Screw. Bone fixation screw |  |  | June 04, 2015 | 2015 | Biomet, Inc. | Sterile packaging was not sealed. The implantation of an unsterile screw may lead to infection. | Packaging | On 6/4/2015, URGENT MEDICAL DEVICE RECALL NOTICE notifications were sent to the affected distributors with instructions for removing the affected product. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Questions related to this notice should be directed to (574) 372-1570, Monday through Friday 8 a.m.to 5 p.m. |  | 50 | 50 units | Domestic: None.International:Netherlands and Japan | Terminated | March 22, 2016 | 292 | 50 | Not_Computer
Z-1532-2015 | 70893 | No | JWY | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | MicroScan Microbiology Systems: Dried Gram Negative MIC/Combo Panels |  |  | August 08, 2014 | 2014 | Beckman Coulter, Inc. | Technical Support Bulletin issued in February 2013 did not mention performance limitation for Proteus mirabilis and Imp when using the CLSIM100-S20 revised interpretive criteria. | Labeling False and Misleading | The firm, Siemens, sent an "Urgent Medical Device Correction" letter dated August 7, 2014 to affected customers by trackable method. The letter describe product, problem and actions to be taken. The customers were instructed to confirm receipt of the letter; complete the attached "Field Correction Effectiveness Check" form via fax to Siemens Healthcare Diagnostics at 302-631-8467;TSB 190 has been removed from use and archived from the Siemens Document Library, if you have a hardcopy, discard it appropriately;and follow the listed in instructions and retain and forward this letter to those who may have received this product.If you have any questions, please contact your Siemens Customer Care Center; Quality Systems &amp; Compliance Director at 916-374-3031 or email: juntalan@beckman.com or your local Siemens technical support representative. |  | 84249 | Approx ~(216240 boxes U.S. and 84249 boxes OUS) total | Worldwide Distribution: US (Nationwide) and to countries of: Argentina, Bahamas, Brazil, Canada, China, Ecuador, Egypt, Germany, Lithuania, Mexico, Paraguay, Peru Poland, Portugal, Romania, Russian Federation, Singapore, South Africa, Spain, Taiwan, Thailand, Trinidad and Tobago, United Arab Emirates, Uruguay, and Venezuela. | Terminated | April 27, 2015 | 262 | 84249 | Not_Computer
Z-2098-2015 | 71582 | No | HDX | Vaginal stent. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Milex Vaginal-Hymenal Silicone Dilators Set of 4 |  |  | June 19, 2015 | 2015 | CooperSurgical, Inc. | Incorrect expiration date on outer carton kit label | Incorrect or no expiration date | CooperSurgical sent an Urgent Medicall Device Recall dated June 19, 2015 to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to discontinue use of the affected products and complete the attached Acknowledgement and Receipt Form for a replacement. For questions call 203.601.5200. |  | 40 | 40 kits | US Nationwide Distribution in the states of: CA, FL, GA, HI, IL, KY, MD, MN, NC, NY, OH, TX, VT, and W A. | Terminated | February 01, 2016 | 227 | 40 | Not_Computer
Z-2103-2015 | 71498 | No | LCJ |  |  | General &amp; Plastic Surgery | 510(k) | Natrelle 133 Series Tissue Expanders |  |  | June 18, 2015 | 2015 | Allergan Inc | Allergan is recalling the NATRELLE 133 and CUI Series Tissue Expanders because the product shipped was beyond the 47 month shelf life. | Incorrect or no expiration date | A recall letter will be sent on 6/18/15 to all customers who received the NATRELLE 133 and CUI series tissue expanders. The letter provides the problem identified and the action to be taken. Customers are instructed to contact GENCO Pharmaceutical Services if the recall actions on the recall letter are unclear. For product returns, customers are instructed to contact GENCO at (877) 319-8966, 7am to 5pm CST. Customers with credit/reimbursements are instructed to contact Allergan at (800) 811-4148, 7am to 5pm PST. Customers with medical inquiries are instructed to contact (800) 433-8871, option 2, 8am-5pm, PST. Customers with adverse events/product complaints are instructed to call (800) 624-4261, option 3, 8am-5pm CST. |  | 167 | 167 units total (114 units in US) | US in the state of Canada | Terminated | December 16, 2015 | 181 | 167 | Not_Computer
Z-1397-2015 | 70368 | No | IMD | Hot or cold disposable pack. | Physical Medicine | Physical Medicine | 510(K) Exempt | CooperSurgical TransWarmer Infant Transport Mattress with WarmGel |  |  | May 01, 2013 | 2013 | CooperSurgical, Inc. | Trans Warmer Infant and Transport Mattress not labeled with exp.date may not meet specification | Package design/selection | CooperSurgical sent a formal recall notification letter dated April 4, 2015 to all affected customers. The letter identified the product problem and the action needed to be taken by the customer.CooperSurgical notified consignees affected by this recall via Fedex with confirmed delivery receipt. The letter identified the affected product and detailed the necessary actions to be taken by customers who carried the affected lots in their inventory. CooperSurigcal will replace any of the affected products thatcustomers had on hand. CooperSurgical provided further instruction on the disposition of the product at the time of replacement.Contact CooperSurgical at 203 601 5200 Ext. 3100 if you have questions.Previously CooperSurgical notified consignees affected by this recall on 5/1/13. |  | 4807566 | 4,807,566 | Worldwide Distribution - US Nationwide - in the countries of:ARGENTINA, AUSTRALIA,BAHRAIN, ,BERMUDA,CANADA,COLOMBIA ,ENGLAND,GERMANY,HUNGARY,INDONESIA, KUWAIT,LATVIA,LIBYA,MALAYSIAMEXICO,NEW ZEALAND,SAUDI ARABIA. TURKEY, UNITED ARAB EMIRATES | Open, Classified |  |  | 4807566 | Not_Computer
Z-1329-2015 | 70668 | [6 empty] | Ebola GP IgX Blood, Serum, Plasma, Cassette, Catalog No. I-101 |  |  | March 13, 2015 | 2015 | Lusys Laboratories, Inc. | LuSys Laboratories is recalling Ebola Virus One Step Test Kits to stop and prevent further use of these devices. The Ebola Virus One-Step Test Kits have not yet been cleared, approved or authorized by the FDA for diagnostics purposes. | No Marketing Application | LuSys Laboratories sent an Urgent Medical Device Recall letter dated March 13, 2015, to all affected customers informing them that the Ebola Virus One Step Test Kits are not approved by FDA for diagnostics purposes. The letter informs the customers of the problems identified and the actions to be taken. Customers were instructed to discontinue use and return the products to LuSys Laboratories. Customers are instructed to fax back the acknowledgement and Receipt Form to 1-858-866-1688.For questions regarding this recall call 858-546-0902. |  | 1015 | 1,015 units | Worldwide Distribution - US including CA and Internationally to Canada, Sierra Leone, and Denmark. | Terminated | May 31, 2017 | 810 | 2275 | Not_Computer
Z-2147-2015 | 71597 | No | MPO | Hot or cold disposable pack. | Physical Medicine | Physical Medicine | 510(k) | Infant Heel Warmer and Infant Gel Warmer, sold under the brand names CooperSurgical and Fisher brand as follows : |  |  | July 06, 2015 | 2015 | CooperSurgical, Inc. | Product marketed without a 510(k) | No Marketing Application | CooperSurgical sent an " Urgent WarmGel Infant Heel Warmer" letter dated June 30, 2015, FEDEx, confirmed delivery receipt on July 6, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.CooperSurgical is advising all WarmGel users to review your inventory of this product, remove it from your inventory and discontinue its use. informing them of the reason to cease using the product and either discard or return product to CooperSurgical. Accounts are provided with a response form to complete. Any questions contact CooperSurgical at 203 601 5200 Ext. 310 |  | 6344700 | 6,344,700 | Worldwide Distribution - Nationwide Distribution and to the countries of : AUSTRALIA, BERMUDA, ECUADOR , FRANCE, NEW ZEALAND, SAUDI ARABIA and UNITED KINGDOM. | Open, Classified |  |  | 6344700 | Not_Computer
Z-0907-2015 | 70044 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Ti-Bond Marketing Literature/Labeling, Catalog No. LT10000, MP-14000 |  |  | December 18, 2014 | 2014 | Spinal Elements, Inc | Spinal Elements is recalling the Ti-Bond related brochures that includes LT-10000 and MP-14000 (all revisions) the content regarding the Ti-Bond coating within the literature is unsubstantiated in context and/or misleading for the coating's use in spine applications. | Labeling False and Misleading | Spinal Elements sent an Urgent Field Action letter dated December 22, 2014, to all affected customers. The letter informed the customers to cease all distribution of the literature material. The letter informed the customers that the content regarding the Ti-Bond coating in the literature is unsubstantiated in context and/or misleading for the coating's use in spine applications. In addition to the customer notiification letter, it will include a list of claims related to the field action. The letter also informed customers of the actions to be taken. Customers were instructed to immediately examine their literature inventory, discontinue the use and /or distribution of the affected brochure, destroy all copies or send them back to:Spinal Elements, Inc.3115 Melrose DriveCarlsbad, CA 92010If product was further distributed customers should identify those customers and notify them at once. Customers were also instructed to complete and return the enclosed "Customer Response Form."For questions regarding this recall call 760-607-0121, ext 206.Customers with questions were instructed to contact Customer Service at 1-877-SPINALS, ext 213.Customers with questions were instructed to contact Customer Service at (877) SPINAL5 Ext 213. |  | 1200 | 1200 units | Nationwide Distribution | Terminated | March 10, 2015 | 82 | 1200 | Not_Computer
Z-2189-2015 | 71533 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | T-Lok Bone Marrow Biopsy Needle 8G x 4 Ext Can, Part Number DBMNJ0804TL; T-Lok Bone Marrow Biopsy Needle 11G x 4 Ext Can, Part Number DBMNJ1104TL; Bone Marrow Biopsy Needle 13G x 4, Part Number DBMNJ1304 |  |  | June 18, 2015 | 2015 | Argon Medical Devices, Inc | Potential defect in the forming of the primary package of the blister that may compromise the sterility of the product inside the tray. | Packaging process control | The recalling firm's sales force was notified of the recall on 6/17/15. The recalling firm will begin to send out Urgent Medical Device Recall letters by registered mail to affected customers on 6/18/15. Customers are instructed to inspect and quarantine and return any affected product and complete the Recal Response Form. |  | 10290 | 10290 units | US | Open, Classified |  |  | 16680 | Not_Computer
Z-1481-2015 | 70948 | No | HGS | Intrauterine pressure monitor and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | HSG Trays/HS 5F Catheter Combo; HSG Procedure Tray+ Shapeable 5F HSG catheter. For hysterosonography and hysterosalpingography, gynecological use. |  |  | April 07, 2015 | 2015 | Catheter Research, Inc | HSG Trays and Catheters were mislabeled. Trays with the flexible HS catheter were labeled as containing the shapeable, stiffer HS catheter, and vice versa. | Packaging process control | Catheter Research, Inc. (CRI) has initiated a voluntary medical device removal of HSG Tray/ 5F HS Catheter Combo, and issued letters to customer son 04/07/2015 by e-mail and certified mail. Customers are asked to inspect their inventory and return recalled devices to CRI for replacement. If the devices were further distributed CRI requested customers provide notice of the recall. Customers may contact Babacar Diouf, VP of RA and QS, at (317) 872-0074 ext. 3512. |  | 880 | 880 units | Domestic: MN, CA, OH, IL, FL, MA, WI, VAForeign: NoneDOD/VA: None | Terminated | July 17, 2015 | 101 | 880 | Not_Computer
Z-1479-2015 | 70488 | No | FTY | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | SpineFrontier MISquito Percutaneous Pedicle Screw System Surgical Technique; Part Number: DOC80023. The Rod Caliper is a manual surgical instrument that may be used as part of the PedFuse Pedicle Screw System and its MISquito components during percutaneous, minimally invasive surgical (MIS) technique of posterior pedicle screw fixation. |  |  | March 05, 2015 | 2015 | SpineFrontier, Inc. | The PedFuse Pedicle Screw Rod Caliper included with the MISquito Pedicle Screw System may be used incorrectly as the Surgical Technique does not adequately convey the intended use of the design. | Labeling design | SpineFrontier sent an Urgent Advisory Notice letter to customers via email. The letter identified the affected product, problem and actions to be taken. Customers were asked complete and return the Acknowledgement Form to confirm receipt of the recall notification. For questions call 978-232-3990. |  | 5 | 5 devices | US Nationwide Distribution in the states of TX, KS, CO, and MA | Terminated | June 15, 2015 | 102 | 5 | Not_Computer
Z-2201-2015 | 71650 | No | GXH | Cryogenic surgical device. | Neurology | Neurology | 510(k) | iovera 155 Smart Tip; |  |  | July 01, 2015 | 2015 | Myoscience Inc | The expiration date on the outer box label and the pouch label for the iovera Smart Tip product is incorrect. Instead of indicating the correct expiration or 2015-12. the labels were incorrectly labeled 2016-05. | Incorrect or no expiration date | Myoscience sent an Urgent Medical Device Correction letters dated July 1, 2015 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers are requested to quarantine any affected product at their site. Myoscience will make arrangements to either correct the incorrect expiation date label or remove the product and replace it with correctly labeled product immediately.Question or concerns may be directed to Customers Service at 510-933-1500. |  | 36 | 12 boxes - 3 boxes of 10 count and 9 boxes of 5 count. | US Nationwide Distribution in the states of TX, NY, MD, CO and CA. | Terminated | September 08, 2015 | 69 | 36 | Not_Computer
Z-2192-2015 | 71577 | No | JSO | Culture medium for antimicrobial susceptibility tests. | Microbiology | Microbiology | 510(k) | HardyCHROM VRE Agar-RSR |  |  | June 24, 2015 | 2015 | Hardy Diagnostics | Hardy Diagnostics is recalling HardyCHROM VRE Agar due to lack of 510(k) clearance. | No Marketing Application | The firm initially called customers to inform them of the recall. The firm's phone script stated that they were discontinuing the product and as part of the process were recalling anything that still might be within expiration date. The phone script stated that the firm was going to be sending something in writing.On 06/24/15 the firm sent out a notification letters to customers to inform them that they were discontinuing the product, and were requested anything remaining within expiration to be discarded. The firm requests that customers complete, sign, and fax back the enclosed form stating their compliance with the above action. For any questions the firm states to contact Technical Services Department at 800-266-222, option 2 or via email at TechService@HardyDiagnostics.com. |  | 1500 | 1,500 | Distributed to one U.S. customer in SC. | Terminated | August 28, 2015 | 65 | 1500 | Not_Computer
Z-0906-2015 | 69466 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Optetrak Asymmetric Hi-Flex Posterior Stabilized Cemented Femorals, size 3, Catalog #244-02-03, Catalog #244-03-03. |  |  | October 09, 2014 | 2014 | Exactech, Inc. | Mislabeled | Labeling mix-ups | The firm, Exactech, sent an "Urgent Field Safety Notice" letter dated September 30, 2014 to its Consignees/customers by electronic mail on October 9, 2014. The letter described the product, problem and actions to be taken. The customers were instructed to: immediately cease distribution or use of the products; extend information to your accounts that may have this product in their possession; verify whether you have any of the subject femoral components from the specified serial number ranges; complete and return the attached inventory RESPONSE FORM via fax to: 011 61 2 9638 0778 within 48 hrs of receipt of the notice, and contact your Exactech representative to confirm quantities at your location and arrangements for product return and inventory restocking.If you have any questions, contact Manager, Regulatory Affairs at 352-377-1140. |  | 96 | 96 | International Distribution only to countries of: Argentina, Austria, Australia, China, Cyprus, France, Greece, India, Japan, Lebanon, The Netherlands, Spain and Tunisia. | Open, Classified |  |  | 96 | Not_Computer
Z-2208-2015 | 71615 | No | GAM | Absorbable poly(glycolide/l-lactide) surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Silhouette Lift brochure |  |  | July 02, 2015 | 2015 | Silhouette Lift, Inc. | Silhouette Lift is recalling the Silhouette Suture brochure because it does not include a complete list of risks and side effects. Therefore, it can potentially mislead the patient regarding the safety of the product. | Labeling False and Misleading | Silhouette lift sent an "Correction of Information for Silhouette Suture" letter dated July 2, 2015, in response to FDA untitled letter. The letter identified the product the problem and the action needed to be taken by the customer.Action Requested:All brochures for Silhouette Suture should be destroyed and no longer be used for discussions with patients.Customers with questions are instructed to call (951) 734-7701. |  | 2260 | 2,260 units | US Distribution to the states of : CA, TX, IA, MA, RI, GA, WI, PA, VA, MO, FL, NE, MI, NY and CO., and Internationally to Canada. | Terminated | February 05, 2016 | 218 | 2260 | Not_Computer
Z-1455-2015 | 70739 | No | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | DBX Mix, 20 cc, Product Code: 058200. Demineralized Bone Matrix Mix, bone void filler. |  |  | December 04, 2014 | 2014 | Musculoskeletal Transplant Foundation, Inc. | The seal on the DBX Mix outer PETG/Aclar tray with foil lid packaging for lot identified above may not be fully intact. | Packaging process control | MTF notified their customers on 12/4/2014 and 12/8/2014 via phone and letter. |  | 25 | 25 units | US Nationwide distribution to CA, KY, IL, LA, PA, NC, TN, WV, and Puerto Rico. | Open, Classified |  |  | 25 | Not_Computer
Z-1459-2015 | 70722 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Equinoxe 44mm Short Humeral Head |  |  | February 13, 2015 | 2015 | Exactech, Inc. | Incorrectly packaged. Outer and inner labeling may incorrectly identify the size of the enclosed device. | Labeling mix-ups | Exactech Agents sent a Product Recall Advisory Notice dated February 12, 2015 by mailed to their customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete the Recall Inventory Respponse form and return it to Exactech, per the instructions provided on the notice. |  | 112 | 112 devices | Worldwide Distribution - US Nationwide in the states of AK, AZ, CA, CO, FL, ID, ME, MN, NC, OH, OK, PA, RI, TN, VA and WA. and countries of Brazil and Canada. | Open, Classified |  |  | 112 | Not_Computer
Z-2209-2015 | 71459 | No | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | InjectionPlasty 1.0 Bone Void Filler; an injectable and biocompatible calcium phosphate bone void filler. Sterilized using irradiation. The single-use InjectionPlasty 1.0 Kit contains the necessary components for mixing and delivery of the bone void filler. The InjectionPlasty 1.0 sterile kit contains: Calcium Phosphate Powder, Dilute Sodium Silicate Liquid, a Mixing System, Targeting K-Wires and a Delivery Needle (Cannula). |  |  | May 29, 2015 | 2015 | Skeletal Kinetics, Llc | Reports of venous uptake, DVT, pulmonary embolism in patients who were implanted with Injection Plasty Bone Void Filler. | No Marketing Application | Skeletal Kinetics sent letters to distributors by overnight mail with instructions to forward a Dear Doctor notification letter on May 14, 2015. Second letters dated May 29, 2015, advised that the firm is removing the product from the marker. Distributor letters asked that the Dear Doctor letter be forward within 24 hours of receipt of the letter. All Injection Plasty kits in their possession and those with customers should be retrieved and returned by Fed Ex. Dear Doctor letters advised that their distribution agent has been requested to return all unused kits. Questions may be directed to 408-366-5002 of duran@skeletalkinetics.com. |  | 67 | 67 kits | US Distribution including Puerto Rico. | Terminated | July 24, 2015 | 56 | 67 | Not_Computer
Z-2210-2015 | 71564 | No | JDW | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Cortical Pin 2.0 mm,Sterile, FREEZE DRIED: Store at ambient temperature. Do Not freeze. |  |  | May 27, 2015 | 2015 | Musculoskeletal Transplant Foundation, Inc. | Musculoskeletal Transplant Foundation (MTF) is recalling one lot of MTF's Allograft Cortical Pins because the lot was labeled as a 2.0 mm diameter pin and it may contain a 2.4 mm diameter pin. | Labeling mix-ups | Musculoskeletal Transplant Foundation initiated this recall by either telephoning their customers or by sending a recall letter dated 5/27/15 to the customer. |  | 22 | 22 units | Distributed in the states of AZ, CO, MI, MN, NC, OH and TX. | Terminated | December 17, 2015 | 204 | 22 | Not_Computer
Z-1365-2015 | 70846 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | Covidien Trellis 6 and Trellis 8 Peripheral Infusion Systems, Sterile EO, Rx only. Models BVT608010V01, BVT608030V01, BVT612010V01, BVT612030V01, CVT808015, CVT808025, CVT812015, CVT812025, EVT808015, EVT808025, EVT812015, EVT812025, BVT808015, BVT808030, BVT812015, BVT812030, EUT808015, EUT808030, EUT812015, EUT812030. |  |  | March 23, 2015 | 2015 | Medtronic Inc. Cardiac Rhythm Disease Management | Medtronic is conducting a voluntary recall of all former Covidien Trellis 6 and Trellis 8 Peripheral Infusion Systems due to potential for a sterility breach of the outer packaging or pouch material. A breach of the outer pouch may compromise the outside surface sterility of the inner pouch and does not directly affect the sterility of the device components within the inner pouch. | Package design/selection | Consignees were sent on 3/23/2015 a Medtronic "Urgent Medical Device Recall" letter dated March, 2015. The letter described the problem and the product involved in the recall. Consignees were advised to quarantine and return the product and to complete and return the Customer Confirmation Certificate to Medtronic. For questions contact your Medtronic Sales Representative or Customer Service a 1-800-716-6700. |  | 8269 | 8,269 ( 6,786 US, 1,483 OUS) | Worldwide Distribution: US (nationwide) including all states in continental USA, except DE and including DC and PR; and countries: AUSTRIA, AUSTRALIA, NEW ZEALAND, BELGIUM, CZECH REPUBLIC, DENMARK, EGYPT, FINLAND, FRANCE, GEORGIA, GERMANY, IRELAND, ITALY, NETHERLANDS, POLAND, PORTUGAL, ROMANIA, RUSSIA, SPAIN, SWEDEN SWITZERLAND, TURKEY, UNITED KINGDOM, CHILE, EL SALVADOR, PANAMA, and CANADA. | Open, Classified |  |  | 8269 | Not_Computer
Z-1368-2015 | 70633 | No | HNL | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | LIS27T - Lacrimal Intubation Set RX Only Sterile |  |  | February 26, 2015 | 2015 | Quest Medical, Inc. | Expiration date incorrectly printed in Manufacturing Date field on both pouch and carton labels | Labeling Change Control | The consignees were notified by letter on 3/2/2015. The recalling firm requested the consignees to return the recalled product for a replacement device. |  | 106 | 106 units | Worldwide Distribution: US (nationwide) including DC and the states of IL, KY, TX, CA, FL, MI, GA, MD, NH, OH, NY, UT, NJ, MO, TN, VA, NC, CO, MI, and PA, and the country of Ireland. | Terminated | February 26, 2016 | 365 | 106 | Not_Computer
Z-2053-2015 | 71474 | No | JDQ | Bone fixation cerclage. | Orthopedic | Orthopedic | 510(k) | Songer Spinal Cable System, Single Cable, Crimp, Bar, Stainless Steel, Sterile. |  |  | May 29, 2015 | 2015 | RTI Surgical, Inc. (dba Pioneer Surgical Technology, Inc.) | The box end label is incorrectly labeled as Titanium instead of Stainless Steel on one lot of implantable Songer Spinal Cable System, Single Cable, Crimp, Bar, Stainless Steel, Sterile, distributed by DePuy Spine Inc. | Labeling Change Control | On 5/29/15, RTI Surgical sent DePuy/Synthes Spine Inc. with an URGENT: Medical Device Voluntary Recall Notification identifying the recalled product, reason for recall, potential risks, and immediate actions required. |  | 25 | 25 | Distributed in MA. | Terminated | March 17, 2016 | 293 | 25 | Not_Computer
Z-2266-2015 | 71541 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Invivo Expression MRI Patient Monitoring System. |  |  | June 17, 2015 | 2015 | Philips Electronics North America Corporation | The device's labeling inaccurately indicates Endotracheal as a body temperature application site. The FlexTEMP reusable temperature sensor is not intended to be used for endotracheal placement. The application site should be listed as Esophageal. | Labeling False and Misleading | An URGENT Medical Device Correction was sent to all consignees by letter June 2015 regarding the Invivo Expression MRI Patient Monitoring System FlexTEMP Application Site explaining the reason for the affected labeling. |  | 697 | 697 devices. | Worldwide Distribution-US (nationwide) including the states of AL, AZ, CA, CO, CT, DE, FL, GA, IA, IL, IN, KS, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NV, NY, OH, OR, OK, PA, SC, TN, TX, UT, VA, VT, WA, WI, WV, WY, and Alaska, Hawaii and the District of Columbia, and the countries of Argentina, Australia, Austria, Canada, Czech Republic, Denmark, Egypt, France, Germany, Hong Kong, Iceland, Israel, Italy, Japan, Netherlands, Norway, Oman, Poland, Portugal, Russian Federation, Saudi Arabia, Slovenia, South Africa, Spain, Sweden, United Arab Emirates, and United Kingdom. | Open, Classified |  |  | 697 | Not_Computer
Z-1560-2015 | 70834 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | PowerLoc MAX Safety Winged Infusion Set, Product Code 0132215 |  |  | March 13, 2015 | 2015 | Bard Access Systems | Bard Access Systems is conducting a field action due to the potential that an incorrect Maximum Flow rate label associated to the power injection IFU for the PowerLoc MAX Safety Infusion Set (SWIS) was attached to the outside of the unit pouch. | Labeling False and Misleading | The firm, Bard, sent an "URGENT: Notification-Potential Incorrect Infusion IFU Label for PowerLocÂż MAX Safety Winged Infusion Set (SWIS)" letter dated March 25, 2015 to its customers. The letter described the product, problem and actions taken. The firm has corrected the cause of the issue and all future lots will be manufactured with the correct maximum flow rate of 2mL/second. Customers may return for a replacement or keep items.Please direct any return requests and questions to our coordinator at 801-522-5640. |  | 700 | 700 units | US Distribution to states of: GA and MN. | Terminated | June 24, 2015 | 103 | 700 | Not_Computer
Z-1960-2015 | 71497 | No | ETN | Surgical nerve stimulator/locator. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | Spine Wave Access Kit, Catalog Number 10-1530 |  |  | May 04, 2015 | 2015 | Spine Wave, Inc. | Mislabeled outer carton:Monopolar Probe is listed as 200 x 1.5mm when it should be listed as 280 x 1.5mm | Packaging change control | Spine Wave sent an Urgent: Medical Device Recall letter dated June 11, 2015 to affected customers. The letter identified the affect product, problem and actions to be taken. Customers were instructed to acknowledge receiptof notification by returning the enclosed Recall Notice Acknowledgement Form in the stamped envelope provided.For questions, contact Spine Wave directly at 203-944-9494and ask for the Recall Coordinator or by email at rsmith@spinewave.com. |  | 121 | 121 units | US Nationwide Distribution in the state of CA, FL, GA, LA, MA, MD, MI, MO, NC, NH, OH, OR, PA, TX, and VA. | Open, Classified |  |  | 121 | Not_Computer
Z-1557-2015 | 70753 | No | BSR | Tracheal tube stylet. | Anesthesiology | Anesthesiology | 510(K) Exempt | Introduction Pack. A device inserted into an endotracheal tube to aid in the intubation procedure. |  |  | March 13, 2015 | 2015 | Vital Signs Colorado Inc. | Vital Signs Colorado (dba CareFusion) is recalling Vital Signs LightWand Stylet catalog numbers 3910 and 3960 due to a potential patient safety risk involving a breach in the sterile packaging that may not be recognized prior to use. | Packaging | Customers were notified on 3/13/15 via letter to perform an inventory, complete and return the Recall Response Form within 15 days, and to destroy the product on site (do not return to CareFusion). |  | 399 | 399 | Worldwide Distribution. US Nationwide, Australia, Canada, Chile, Colombia, Dominican Republic, Ecuador, Egypt, Germany, Guatemala, Hong Kong, India, Israel, Italy, Japan, South Korea, Lebanon, Mexico, Netherlands, New Zealand, Panama, Peru, Poland, Russia, Slovenia, Turkey, United Kingdom, and Qatar. | Terminated | October 07, 2015 | 208 | 83219 | Not_Computer
Z-1988-2015 | 71428 | No | MRD |  |  | General &amp; Plastic Surgery | 510(k) | Natrelle Re-sterilizable Breast Implant Sizer, Size: 410 cc, Style No. MSZFX410. |  |  | June 08, 2015 | 2015 | Allergan Inc | Allergan is recalling the NATRELLE 410cc Re-Sterilizable Breast Implant Sizer because the expired product was shipped to the user level. | Error in labeling | A recall letter dated 6/8/15 was sent to all customers who received the NATRELLE 410 cc Re-Sterilizable Breast Implant Sizer. The letter provides the problem identified and the action to be taken. Customers are instructed to contact GENCO Pharmaceutical Services if the recall actions on the recall letter are unclear. For product returns, customers are instructed to contact GENCO at (877) 319-8966, 7am to 5pm CST. Customers with credit/reimbursements are instructed to contact Allergan at (800) 811-4148, 7am to 5pm PST. Customers with medical inquiries are instructed to contact (800) 433-8871, option 2, 8am-5pm, PST. Customers with adverse events/product complaints are instructed to call (800) 624-4261, option 3, 8am-5pm CST.On June 12, 2015, Actavis sent follow up letters to customers to include one additional serial number. |  | 9 | 9 units | CA, IN, MO, NJ, VA | Terminated | December 16, 2015 | 191 | 9 | Not_Computer
Z-1545-2015 | 70906 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Insorb Subcuticular Skin Stapler. INSORB staples are made from an absorbable copolymer which is a synthetic polyester derived from lactic and glycolic acids. It is chemically similar to other surgical glycolide/lactide-based copolymers. Polyglycolic/polylactic acid copolymers degrade in vivo by hydrolysis to glycolic acid and lactic acid which are then absorbed and metabolized by the body. |  |  | March 19, 2015 | 2015 | Incisive Surgical Inc | A previous Turkish distributor over-labeled a portion of the lot with an expiration date of 2016-05. The expiration date at the time of manufacturing was 2014-05. | Labeling False and Misleading | Firm is working with EU Rep on English and Turkish notification letters. |  | 1260 | 210 cartons (containing 6 INSORB staplers per carton) | International Distribution Only - Turkey. | Terminated | July 08, 2015 | 111 | 1260 | Not_Computer
Z-2047-2015 | 71420 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | PowerPICC SOLO Catheter with Sherlock 3CG TPS Stylet 4F Maximal Barrier Tray with Microintroducer |  |  | May 27, 2015 | 2015 | Bard Access Systems | Bard Access Systems is recalling PowerPICC Solo Catheter Kits due to incorrect expiration dates on the primary labeling. | Labeling Change Control | BARD sent an Urgent Notification letter dated May 27, 2015, to all affected consignees via FedEx. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed that they may choose to return the affected product to Bard Access Systems or use with the understanding that the correct expiration date is November 2015 (2015-11). Customers with questions were instructed to call 1-800-290-1689. For questions regarding this recall call 801-522-5974. |  |  |  | Nationwide Distribution | Terminated | February 10, 2016 | 259 |  | Not_Computer
Z-0965-2015 | 70010 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | MUM-1 (Multiple myeloma oncogene-1), catalog number PRM352 AA; |  |  | December 12, 2014 | 2014 | Biocare Medical Llc | Three lots of MUM-1 may not meet their labeled shelf life, as a drop in staining intensity over time has been observed. | Incorrect or no expiration date | Biocare sent a Medical Device Recall notifications letter to all customers by traceable method. The letters identified the affected product, problem and actions to be taken. Users are asked to immediately discontinue used of the affected lots and dispose of them according to applicable national, state and local laws. Users are requested to complete and return the response form so that proper credit can be issues 800-799-9499. |  | 59 | 59 | Worldwide Distrinution - US Nationwide in the states of TX, WI, MO, WV, SC, MA, WA, TN, AZ, CA, IL, NY, LA and countries of Kuwait, India, Venezuela, New Zealand, Turkey, Uruguay. | Terminated | March 30, 2015 | 108 | 59 | Not_Computer
Z-1541-2015 | 70737 | [6 empty] | Portable FAR Infrared Sauna |  |  | December 23, 2014 | 2014 | Therasage, LLC. | Product distributed without 510(k) approval | Labeling False and Misleading | On January 20, 2015 Therasage sent a notification to its consignees asking them to discard any of the labeling materials that state the following: 1) Use of the sauna allows one to "naturally detoxify." 2) Use of the sauna expels toxins from the body 3) Any reference to an exact amount of calories one uses in an hour when enjoying the benefits of this product 4) Although this product is enhanced with a negative ion environment, we cannot say that the product is enhanced with negative ion therapy 5) Any reference to the maximum infrared absorption available cannot be explicitly stated, as it may vary for each individual person. Customers with questions were instructed to call 888-416-4441 or email info@therasage.com.For questions regarding this recall call 561-416-9300. |  | 8 | 8 devices | Nationwide Distribution including SD, IL, TX, OR, NY, WI, GA, and CA. | Terminated | August 06, 2015 | 226 | 8 | Not_Computer
Z-0988-2015 | 70024 | No | HQF | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | Femto LDV Z-Generation Femtosecond Surgical Lasers are used to perform ophthalmic surgeries. |  |  | December 18, 2014 | 2014 | Ziemer Usa Inc | The devices were not shipped with the Starter Kits. | Labeling False and Misleading | Notification letters dated December 18, 2014 were sent to all direct account s(customers). The letters were included with the Starter Kit referred to in the Surgical Procedure Manual. |  | 26 | 26 lasers | US: Nationwide (AL, CA, FL, GA, IL, MN, MO, NC, ND, NH, NJ, NM, NY, PA, TX, VA, WA) | Open, Classified |  |  | 26 | Not_Computer
Z-1948-2015 | 71531 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Natural Knee II Revision Femoral Stem, 125mm, item 6215-12-185. packaged in configuration 803/791. Knee Prosthesis component. |  |  | October 02, 2014 | 2014 | Zimmer, Inc. | Failed to meet established requirements for sterile barrier integrity due to a breach of the inner cavity. If the outer tray is compromised in addition to the inner tray a periprosthetic infection may occur. The risk associated with a periprosthetic infection is a revision or potential multi-stage revision to treat the infection. | Package design/selection | On 10/2/2014, URGENT MEDICAL DEVICE RECALL notifications were sent to the affected distributors and hospital Risk Managers. All distributors were notified via electronic mail, and all hospital Risk Managers and distributors with affected inventory were also notified via courier. The recall notification included a description of the reason for the recall, affected product, and consignee responsibilities. If after reviewing this notification you have further questions or concerns please call the customer call center at 1-800-348-2759. Hours of operation are Monday through Friday, 8 a.m. through 8 p.m. ET. |  | 352 | 352 | US Nationwide Distribution, including DC and the states of AL, AZ, CA, CO, FL, GA, IL, IN, KS, LA, MD, MI, MN, MO, MT, NC, ND, NJ, NY, OH, OK, OR, PA, TN, TX, UT, VA, WA, WI, and WY. | Terminated | July 01, 2015 | 272 | 352 | Not_Computer
Z-2056-2015 | 71147 | No | BSP | Anesthesia conduction needle. | Anesthesiology | Anesthesiology | 510(k) | IntegraÂż Long Epidural Steroid Injection Needle (Tuohy Needle), 20 gauge x 5 inch. Plastic Hub, Metal Stylet, calibrations (10BX). IntegraÂż Long Epidural Steroid Injection Needle (Catalogue #PITRW2050). Peel Pouch. The IntegraÂż Long Epidural Steroid Injection Needle is packaged in a labeled and sterilized sealed tyvek / plastic pouch. The sterilized pouches are subsequently packaged in a case of 10 that is labeled, sealed and distributed to customers. |  |  | May 01, 2015 | 2015 | Integra LifeSciences Corp. d.b.a. Integra Pain Management | Integra LifeSciences Corp. d.b.a. Integra Pain Management Announces Voluntary Recall of Tuohy Needles (20 Gauge, 5 inch) due to incorrect packaging (labeled as 18 Gauge, 3. 5 inch). | Labeling design | Integra sent an Urgent Voluntary Medical Device Recall letter dated May 1, 2015, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Integra is asking customers to do the following:1. Review your inventory and determine if you have Tuohy Needle, 20 Gauge X 5 inches (Cat. No. PITRW2050) with an affected Lot Number that is listed on page 2. If so, stop using them immediately and remove them from service.2. Complete the attached form. If you do not have affected product, check the box: I do not have any product(s) with a Lot Number from the list below.3. If you do have affected product, check the box:  I do have product(s) with a Lot Number from the list below.4. Complete the other information on the form and return it by email or fax as indicated on the form. Keep a copy of the form for your records.When your form is received, Customer Service will contact you and provide an RMA number and directions to return and replace the affected products. Should you have any questions regarding these instructions, please contact Customer Service at 1-800-241-2210. For questions regarding this recall call 609-936-6822. |  | 123 | 123 Tuohy needles | Nationwide Distribution | Terminated | December 04, 2015 | 217 | 123 | Not_Computer
Z-1348-2015 | 70721 | No | FYD | Air-handling apparatus for a surgical operating room. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | LASER RESISTANT WAND - STERILE, 7/8" X 8" (2.22cm x 20.32cm), REF 24013, QTY 10 EA, STERILE EO, NO LATEX |  |  | March 10, 2015 | 2015 | Microtek Medical Inc | This recall is being conducted because a cohesive band used to package the devices contains latex and the product is not labeled as required by 21 CFR 801.437; it does not contain Caution: The Packaging of This Product Contains Natural Rubber Latex Which May Cause Allergic Reactions. | Error in labeling | The firm, Ecolab - Microtek Medical, sent an "Urgent: Recall - Mislabeled Product Immediate Action Required" letter dated March 2015 to its US customers on 03/10/2015 via letter and email. The letter describes product, problem and actions to be taken. The customers were instructed to immediately examine your inventory and promptly quarantine the lot(s) listed; complete and return the enclosed Recall Response Form via fax to Quality Assurance Manager, Ecolab-Microtek Medical at (662) 244-3011 or email to: recall@ecolab.com RE: Recall - Mislabeled Product ASAP, and if you have further distributed the products subject to this recall, please notify your customers at once. Upon receipt of the completed response forms, Ecolab-Microtek Medical will issue a return material authorization (RMA) and ask that you return any unopened product listed above to Ecolab-Microtek Medical at the address listed in the attached form.If you have any questions in regard to this recall, please contact Customer Service by phone at 1-800-824-3027 or email; customerservice@microtekmed.com. |  | 6540 | 6540 units | US (nationwide) Distribution. | Open, Classified |  |  | 145758 | Not_Computer
Z-1354-2015 | 70596 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo.plaza. A Picture Archiving and Communication System (PACS) intended to display, process, read, report, communicate, distribute, store and archive digital medical images. It supports the physician in diagnosis and treatment planning. |  |  | February 18, 2015 | 2015 | Siemens Medical Solutions USA, Inc | Possibly incomplete archived studies during pre-fetch. In a server farm setup, when pre-fetch/retrieve operation is performed for partially archived studies, the series that have not yet been archived, will remain unarchived. | Software design | Safety Advisory Notice, dated February 18, 2015, was sent to end users that identified the product, problem, and action to be taken. Users were advised not to manually remove the delete protection in the STS. |  | 4 | 4 | US Distribution: MA, WI, FL | Open, Classified |  |  | 4 | Computer
Z-0274-2016 | 72423 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Hill-Rom 100 Low Bed (GPAC) AC Powered adjustable hospital bed |  |  | October 07, 2015 | 2015 | Hill-Rom, Inc. | Hill-Rom is voluntarily recalling the Hill-Rom 100 Low Bed due to complaints of patient injury when exiting the beds set at the lowest position setting. | Device Design | Hill-Rom initiated Phase 1 of the recall on October 2, 2015, which included notices of the recall being issued to customers via certified mail, which included safety information on how to properly use the bed, and instructions to forward the information on if devices have been transferred. Phase 2 for the recall will include the distribution of caps for the edges of the bed to be placed on by the customers. Customers may contact Hill-Rom Technical Support at 800-445-3720. |  | 12251 | 12251 units | Worldwide Distribution - USA (nationwide) to the states of : AL, AR, AZ. CA, CO, CT, DC, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MS, MT, NC, ND, NE, NJ, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, VA, WA, WI and WV., and to the countries of : Canada and France. | Open, Classified |  |  | 12251 | Computer
Z-0224-2016 | 72224 | No | OLO | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | Robotic Arm Interactive Orthopedic (RIO) system component PN#205143 (locking pin) and PN# 112220 (connector). Indicated for use in surgical knee resurfacing and total hip Arthroplasty joint replacement procedures. |  |  | July 16, 2015 | 2015 | Mako Surgical Corporation | The locking pin on the RIO Base Array is damaged during assembly onto the RIO Base Array Connector. | Device Design | As of July 16, 2015, Mako communicated by mail the nature of the recall to their Product Specialists and the OR Administrators describing the specifics regarding the issue for correction. |  | 1322 | 1,322 units | Worldwide Distribution-US (nationwide) including the states of AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KY, LA, MA, MD, MI, MS, NC, NH, NJ, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, WI, WV, WY and HI, and the countries of Italy, Japan, Hong Kong, Scotland, Germany, Turkey, Greece, Australia, Singapore, Taiwan, and South Korea. | Open, Classified |  |  | 1322 | Computer
Z-0283-2016 | 72411 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX, PACS, Picture Archive and Communications System, IMPAX CV 7.8.x and IMPAX CV 12.x. Corrects certain demographic information (Patient Name, Patient Medical Record Number and/or Accession Number in previously signed reports that have since changed. |  |  | September 01, 2015 | 2015 | AGFA Healthcare Corp. | Customers have experienced IMPAX CV software, specifically, Demographics Manager (DM), not producing correct pdf reports after the reports, at the time of physician signature, were correct. | Software design | An Urgent Field Safety Notice was sent to accounts via email on 9/1/2015. The letter described the safety alert. Customers were informed of the problem and that Service would be contacting them to arrange a date to deploy the corrective action and install the Service Update at their facility. An acknowledgement was included which was to be faxed back or emailed back to Agfa. |  | 19 | 19 | Distributed in the states of CA, KY, NH, NY, NC, OH, OR, PA, SC, SD, TN, TX, WA and the country of Canada. | Open, Classified |  |  | 19 | Computer
Z-1247-2015 | 68235 | mattress, flotation therapy, non-powered | IKY | Nonpowered flotation therapy mattress. | General Hospital | General Hospital | 510(K) Exempt | Pressure IQ Evolve Mattress; Non-powered mattress replacement system. |  |  | September 01, 2014 | 2014 | Arjo Hospital Equipment AB | The firm received complaints regarding multiple mattresses were showing failures resulting in pressure ulcers. | Device Design | On 9/1/14 the recalling firm sent out a letter to their customers notifying them of an updated IFU for the recalled device. The letter also requested the customers to verify the quantity of the product they currently have on hand. |  | 1034 | 1,034 units | Nationwide Distribution. | Terminated | January 04, 2016 | 490 | 1034 | Computer
Z-0265-2016 | 72406 | No | LNS | Extracorporeal shock wave lithotripter. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | MODULARIS VARIOSTAR; Lithotripter device designed to treat urolithiasis. |  |  | September 24, 2015 | 2015 | Siemens Medical Solutions USA, Inc | Display freeze of MODULARIS hand control results in information not being updated on the display. Current treatment data is not shown to the user. | Software design | Siemens sent a Field Safety Notice (XP044/15/S) dated September 24, 2015 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. |  | 15 | 15 | Distributed in PR and the states of MO, NC, GA, MS, LA, and KY. | Open, Classified |  |  | 15 | Computer
Z-1248-2015 | 70514 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Proteus XR/a with wall stand model number 2260354. |  |  | January 30, 2015 | 2015 | GE Healthcare | Potential breakage of the steel cable which supports the bucky device in certain wall stands of Proteus XR/a X-ray imaging systems. A fall of a bucky while the system is in use could result in an injury to a patient or operator. | Device Design | Consignees were sent on 1/30/2015 a GE Healthcare "Urgent Medical Device Correction" letter GEHC Ref#17108 dated January 30, 2015. The letter was addressed to Hospital Administrators/Risk Managers, Managers of Radiology/Cardiology, and Radiologists/Cardiologists.The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. For questions contact 800-437-1171. Outside the US contact your local GE Healthcare Service Representative. |  | 192 | 192 (131 US, 61 OUS). | Worldwide Distribution-US (nationwide) including the states of AK, AZ, CA, CO, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MD, MI, MO, NC, NE, NH, NJ, NM, NV, NY, OH, OK, PA, TN, TX, UT, VA, WA, WI,and WV, and the countries of ARGENTINA, CANADA, CHILE, CZECH REPUBLIC, FRANCE, GERMANY, GREECE, HONG KONG, HUNGARY, IRELAND, ITALY, JAPAN, NEW ZEALAND, PORTUGAL, RUSSIA, SAUDI ARABIA, SLOVAKIA, SPAIN, and UNITED KINGDOM. | Terminated | September 15, 2015 | 228 | 192 | Computer
Z-1252-2015 | 70203 | No | MQM | Bilirubin (total and unbound) in the neonate test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | RAPIDPoint 405 Blood Gas Analyzer; Siemens Material Number: 10310464, 10322347, 10314817, 10317193, 10318999, 10320055, 10322347, 10328278, 10328302, 10336784. |  |  | December 30, 2014 | 2014 | Siemens Healthcare Diagnostics Inc | Neonatal bilirubin (nBili) parameter may have increased variability when the nBili concentration is greater than 12 mg/dl and the tHb concentration exceeds the upper reportable range of greater than 25 g/dl. | Software design | An Urgent Field Safety Notice was sent by e-mail to all affected Siemens Healthcare Diagnostics Regional offices on December 30, 2014 both in the United States and elsewhere for communication with affected customers. In addition, all affected domestic customers were sent a hard copy of the Urgent Field Safety Notice via Federal Express on December 30, 2014. All affected customers outside the US will be provided a copy of an Urgent FieldSafety Notice via hard copy, e-mail, and/or fax as determined by each countrieslocal regulations and procedures. This notice informs consignees of the issue. Consignees are also asked to complete a field effectiveness check form and return it Siemens via fax (312) 275-7795. |  | 3420 | 3420 devices | Worldwide Distribution - US nationwide including Puerto Rico, Canada, Mexico, Albania, Algeria, Argentina, Austria, Bahrain, Belgium, Bosnia Herzegovina, Botswana, Brazil, Bulgaria, Chile, Colombia, Croatia, Czech Republic, Denmark, Egypt, Estonia, Finland, France, Georgia, Germany, Greece, Hungary, Iceland, Iraq, Ireland, Israel, Italy, Ivory Coast, Kazakhstan, Kuwait, Latvia, Lebanon, Lesotho, Lithuania, Luxembourg, Macedonia, Netherlands, New Caledonia, Norway, Peru, Poland, Portugal, Qatar, Romania, Russian Fed., Saudi Arabia, Serbia, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Turkey, U.A.E., United Kingdom, Vatican City, Venezuela, White Russia/Belarus. | Terminated | August 02, 2017 | 946 | 10900 | Computer
Z-1240-2015 | 70220 | No | KMI | Bed-patient monitor. | General Hospital | General Hospital | 510(K) Exempt | The I-Cart is used to store, transport, and power video equipment used in medical procedures. |  |  | December 10, 2014 | 2014 | Richard Wolf Medical Instruments Corp. | The transformers installed on carts used for Richard Wolf medical video systems may overheat causing smoke and black plastic material to leak from the cart chassis. | Device Design | Field Action Notice letters dated December 10, 2014 were sent to both direct accounts (customers). The letters instructed customers to: 1) Instruct the biomed to bypass the integrated power grid on the cart and not to use the electrical system on the affected carts until a qualified Omni Corporation service technician is able to inspect and repair the carts; and 2) Omni Corporation will schedule a qualified Omni field service technician, in conjunction with Richard Wolf and the customer's biomed unit, to visit and replace the transformer on the cart. |  | 83 | 83 carts | Nationwide | Terminated | July 31, 2017 | 964 | 83 | Computer
Z-1249-2015 | 70477 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | EVOLUTION MP TIBIAL BASE, REF ETPKN8SR, SIZE 8 RIGHT, TYPE PRIMARY, FIT CEMENTED, STYLE KEELED, STERILE R, Rx ONLY, MicroPort Orthopedics, Inc. Knee prosthesis component. |  |  | February 02, 2015 | 2015 | Microport Orthopedics INC. | Some units of the EVOLUTION Tibial Base was affected by a casting tool issue which caused varying retaining wall thickness and keel offset in the finished product. | Device Design | MicroPort Orthopedics sent an Urgent Action Required letter dated January 29, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their stock including all consignment stock to determine if they have any affected product. If they have affected product they should return immediately to the MicroPort Orthopedics Inc. Distribution Center at 11481 Gulf Stream, Arlington, TN 38002. Mark all return shipping boxes with "RECALL" for identification and processing. Customers with questions were instructed to contact MPO's Customer Experience Department at 1-866-872-0211. For questions regarding this recall call 901-867-4324. |  | 17 | 17 units | Worldwide Distribution - USA including AL, AZ, ID, KS, KS, MO, TX, IA, WV and Internationally to Canada and The Netherlands. | Terminated | March 19, 2015 | 45 | 17 | Computer
Z-0273-2016 | 72477 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | ECG Out Cable,for use with a Philips ALS monitor/ defibrillator and bedside monitor for Synchronized cardioversion and Transcutaneous pacing. M1783A/M5526A 12-pin Sync cables, used with HeartStart MRx and HeartStart XL and HeartStart XL+ |  |  | November 19, 2014 | 2014 | Philips Electronics North America Corporation | When a Philips monitor/defibrillator is receiving an ECG signal from an auxiliary bedside monitor via a sync cable, the following can occur if the monitor/defibrillator experiences interference from electrical fast transients (EFTs) while connected to AC power: On the HeartStart MRx and HeartStart XL, EFT noise can be misinterpreted as an R-wave. On the HeartStart XL+, EFT noise can disable ECG mo | Device Design | Philips sent an Urgent - Field Safety Notice dated November, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Philips plans to take to remedy the issue. A replacement cable will be provided free of charge for all units affected by this issue. A Philips Healthcare representative will contact customers with affected devices to arrange for installation of the upgrade. Philips is asking customers to follow the: Action to be Taken by Customer/User section of the Urgent Medical Device Correction notification/Field Safety Notice: Prior to receiving a replacement sync cable, you can continue to use affected sync cables, provided that they are only used to connect the ECG Out jack on the monitor/defibrillator to the ECG (input) port on a bedside monitor. The issues described in this letter do not impact this application of the sync cable.Affected sync cables should not be used to connect the ECG Out jack of a bedside monitor to the ECG (input) port on the monitor/defibrillator.For further questions please call (800) 722-9377. |  | 18604 | 18,604 | Worldwide Distribution - US (nationwide) and to the countries of : Argentina, Australia, Austria, Bahrain, Belgium, Bolivia, Brazil, Bulgaria, Canada, Chile, China, Colombia, Cyprus, Czech Republic, Ethiopia, Finland, France, Gabon, Germany, Hong Kong, Hungary, Iceland, India, Indonesia, Ireland, Israel, Italy, Japan, Kenya, Republic of Korea, Latvia, Lebanon, Macau, Malaysia, Mexico, Netherlands, New Zealand, Nigeria, Norway, Oman, Philippines, Poland, Portugal, Qatar, Reunion, Russian Federation, Saudi Arabia, Singapore, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Syrian Arab Republic, Taiwan, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom, Venezuela and Yemen. | Open, Classified |  |  | 18604 | Computer
Z-0209-2016 | 72488 | component, external, limb, ankle/foot | ISH | External limb prosthetic component. | Physical Medicine | Physical Medicine | 510(K) Exempt | iOS Galileo Application Version 1.1.1 or lower that programs the Triton Smart Ankle; 1C66 Triton Smart Ankle, an external prosthetic ankle/foot. |  |  | October 23, 2015 | 2015 | Otto Bock Healthcare GmbH | A software issue can make the foot move into dorsiflexion, unnoticed by the user, who could fall. When sitting, the prosthetic foot can go into the Relief Function to move the footplate towards the ground. When operating machines with foot pedals (e.g. a car) the foot could get stuck under the foot pedal and block it or, with the foot on the pedal, it could cause unintended vehicle acceleration. | Software design | OttoBock sent an Urgent Medical Device Correction letter dated October 23, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The US Agent will notify US customers via letter on October 23, 2015, of the correction on behalf of the Manufacturer. Customers will be requested to contact all users and advise them of the hazards when operating machines or vehicles with foot pedals. They may forward a copy of the insert if needed. Customers will also be requested to contact all users with TF-fittings, where the settings were done via iOS Galileo app., inform the user about this issue, and assure that settings in the device are as expected.The customer will be requested to complete a reply form and return to the US Agent. For further questions call (800) 328-4058. |  | 321 | 273 in US: 48 Foreign | Worldwide Distribution - US (nationwide) and to the countries of : Austria, Belgium, Germany, Israel, Luxembourg and Sweden. | Terminated | March 21, 2017 | 515 | 321 | Computer
Z-0276-2016 | 72469 | No | JDQ | Bone fixation cerclage. | Orthopedic | Orthopedic | 510(k) | Synthes Application Instrument for Sternal ZIPFIX. Indications for use include primary or secondary closure/repair of the sternum following sternotomy or fracture of the sternum to stabilize the sternum and promote fusion. |  |  | October 19, 2015 | 2015 | Synthes (USA) Products LLC | Certain lots of the Application Instrument for Sternal ZIPFIX that were previously recalled on August 19, 2015 and subsequently repaired and returned to customers between August 26, 2015 and September 23, 2015. The screws on the repaired instruments have the potential to break or loosen. DePuy Synthes is requesting the return of the repaired device in exchange for a new device. | Device Design | Communication to the accounts having record of receiving this product was sent on October 19, 2015. |  | 267 | 267 | Distributed in the states of KY, UT, CA, WI, OH, CO, IL, TX, OK, WA, MI, LA, MA, PA and NE. | Open, Classified |  |  | 267 | Computer
Z-1244-2015 | 70433 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Focal Sim radiation therapy treatment planning system |  |  | December 19, 2014 | 2014 | Elekta, Inc. | Incorrect CT to ED when using the Monaco Image Statistics Tool. | Software design | An Important User Notice (IUN 382-03-FCL-001) was issued to customers on 12/19/2014. The notice included a workaround to be used until a new software release. The notice also included an acknowledgement form to be returned. |  | 907 | 907 | Worldwide distribution. US nationwide (AL, AK, CA, CO, CT, FL, IL, IN, LA, MD, MA, MI, MO, MT, NJ, NY, NC, OH, OK, OR, PA, SC, TN, TX, VA, WA, WY, DC), Puerto Rico, Argentina, Austria, Australia, Bulgaria, Bahrain, Brazil, Belarus, Canada, Chile, China, Colombia, Cuba, Cyprus, Czech Republic, Germany, Algeria, Egypt, Spain, France, Great Britain, Greece, Croatia, Indonesia, Israel, India, Iraq, Italy, Jordan, Japan, South Korea, Sri Lanka, Lithuania, Libya, Morocco, Macedonia, Mexico, Malaysia, New Zealand, Philippines, Portugal, Romania, Serbia, Russia, Singapore, Slovenia, Suriname, Thailand, Turkey, Ukraine, Venezuela, South Africa, Kosovo. | Terminated | August 07, 2017 | 962 | 907 | Computer
Z-1243-2015 | 70464 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | DynaLOC Breast Interventional Planning Software Versions 3.1 and 3.2 |  |  | January 21, 2015 | 2015 | iCAD, Inc. | A device malfunction may cause the biopsy needle to fail to reach the intended target. If location confirming radiographs are not taken and properly evaluated, biopsy site mislocation can result in non-target tissue being sampled rather than target tissue. May result in an incorrect diagnosis. | Device Design | iCAD notified Invivo of the recall via letter on January 21, 2015. Invivo sent a Customer Information letter dated January 23, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letters from both Invivo and iCAD (device manufacturer) recommend that DynaCAD Breast Interventional Software not be used with Invivo ds Breast 16 channel lateral grid plate part number 4535-303-31491. Invivo and iCAD are developing DynaCAD Breast Interventional 3.3 that corrects this error and will be available for installation beginning in March 2015.If you need any further information or support concerning this issue, please contact your local Invivo representative:1-877-Invivo1 (1-877-468-4861, option #3)Invivo apologizes for any inconveniences caused by this problem. |  | 153 | 153 devices | Worldwide Distribution - Nationwide Distribution to the states of : AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, NC, NE, NH, NJ, NY, OR, PA, SC, TN, TX, VA, VT, WA, WI, WV and W., and to the countries of : BELGIUM, GERMANY, GREAT BRITAIN, INDIA, ISRAEL, NETHERLANDS and SWITZERLAND. | Open, Classified |  |  | 153 | Computer
Z-0223-2016 | 72459 | No | MIH |  |  | Cardiovascular | PMA | FLAIR Endovascular Stent Graft, Product Code: FAS08070, FAF08070, Implant Size: 8mmx70mm |  |  | October 19, 2015 | 2015 | Bard Peripheral Vascular Inc | Bard Peripheral Vascular (BPV) is recalling the Bard Flair Endovascular Stent Graft because of the harm associated with failure to deploy or deployment related issues that could occur during use. | Device Design | Bard sent an Urgent Medical Device Recall Notification letter dated October 19, 2015, to customers who received the Bard Flair Endovascular Stent Grant. The letter informs customers that the harm associated with failure to deploy or deployment related issues that could occur during use may pose an incremental risk to the already inherit risk associated with stent graft procedures. The letter also provides customers with the actions to be taken. Customers were instructed to call Bard Peripheral Vascular's (BPV) Recall Coordinator at 1-800-321-4254 Option #2 Ext 2501 (M-F 6am to 3pm MST). Customers were instructed to fax the Recall and Effectiveness Check Form to BPV at 1-800-994-6772. |  | 292 | 292 units | Nationwide in US | Terminated | April 12, 2016 | 176 | 292 | Computer
Z-1449-2015 | 70847 | No | DXX | Steerable catheter control system. | Cardiovascular | Cardiovascular | 510(k) | Magellan Robotic System; Model No(s). 11132 and 11139. A steerable catheter control system used at clinical sites. |  |  | March 18, 2015 | 2015 | Hansen Medical Inc | Based on investigation of two complaint incidents, all Hansen Medical Magellan Robotic System (Model No. 11132 and 11139) may experience an unexpected event during a retraction operation of catheter. Retraction of catheter may not stop when equipment button is released. | Software design | Hansen Medical sent an Urgent Notice - Field Safety Notice dated March 24, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. We request that you review the information in pages 1 and 2 with Hansen Medical personnel and acknowledge your understanding by signing the Field Safety Notice Training Form on page 3.If you need any further information regarding this issue, please do not hesitate to contact me via telephone at 650-404-2784. |  | 19 | 19 total installed systems | Worldwide Distribution - US Distribution and the countries of : Saudi Arabia, Australia, France, Italy, Germany, and United Kingdom. | Terminated | September 23, 2015 | 189 | 19 | Computer
Z-1250-2015 | 70473 | No | ILY | Infrared lamp. | Physical Medicine | Physical Medicine | 510(K) Exempt | Apollo (cold) Laser Desktop Control Units, Model AP2-DT. |  |  | December 17, 2014 | 2014 | Pivotal Health Solutions | Control units were equipped with an internal mounting kit that does not meet medical safety standards, and are conductive, increasing the risk of electric shock to the user and patient. These units were manufactured prior to Pivotal Health Solutions acquisition of the Apollo product line. | Device Design | The firm, Pivotal Health Solutions, sent a letter dated December 17, 2014 on 12/17/2014 and an amended Pivotal "Urgent Medical Device Recall" letter dated February 9, 2015 to its consignees/customers. The letters described the product, problem and actions to be taken. The consignees/customers were instructed to stop using the unit and contact Pivotal Health Solution's at 1-800-743-7738 to arrange for return and repair asap; immediately examine your device inventory and quarantine any product subject to recall; if you have further distributed the product, identify your customers and notify them at once of this product recall, and complete and return the enclosed RETURN AUTHORIZATION FORM with the units to Pivotal Service Center, 1654 Mardon Drive, Dayton, OH 45432 and the DECLARATION OF CONTAMINATION STATUS form via Fax to: 605-882-8398.If you have any questions, contact Pivotal Health Solutions Service Repair Coordinator at 800-743-7738 or email Robin@PivotalHealthSolusitons.com. |  | 35 | 35 | US Distribution to states of: AZ, AR, CA, GA, IL, IN, IA, KS, MI, MN, MO, NY, OH, OR, PA, UT and WA. | Terminated | March 09, 2016 | 448 | 35 | Computer
Z-0278-2016 | 72285 | No | LFL |  |  | General &amp; Plastic Surgery | 510(k) | HARMONIC ACE+ 7 Laparoscopic Shears with Advanced Hemostasis (23cm, 36cm, 45cm length), Ethicon Endo-Surgery. Indicated for soft tissue incisions when bleeding control and minimal thermal injury are desired. |  |  | September 23, 2015 | 2015 | Ethicon Endo-Surgery Inc | Complaints of holes in Tyvek which can compromise the sterile barrier provided by the packaging. | Device Design | On October 5, 2015, the firm sent Urgent Medical Device Recall Removal Letters to their customers. |  | 441490 | 441,490 pieces | Distributed US (nationwide) including the District of Columbia &amp; Puerto Rico and in the countries of ARUBA, AUSTRALIA, BELGIUM, BRAZIL, CANADA, CHILE, CHINA, COLOMBIA, ECUADOR, HONG KONG, INDIA, JAPAN, KOREA, MALAYSIA, MEXICO, PANAMA, PERU, PHILIPPINES, SINGAPORE, SINGAPORE, THAILAND, TAIWAN &amp; VIETNAM. | Open, Classified |  |  | 441490 | Computer
Z-1242-2015 | 70590 | No | DRF | Electrode recording catheter or electrode recording probe. | Cardiovascular | Cardiovascular | 510(k) | LASSO NAV Duo Loop eco Catheters, Catalog No. D-1354-01, D-1354-02, D-1354-03, D-1354-04, D-1354-05, D-1354-06. |  |  | February 25, 2015 | 2015 | Biosense Webster, Inc. | Biosense Webster is recalling the LASSO NAV Duo Loop eco Catheters because of reports of difficulty retracting and damage to the spine cover. | Device Design | Biosense Webster sent an Urgent Voluntary Medical Device Recall letter dated February 24, 2015 , to all affected customers. The letter informed customers that Biosense Webster is recalling the LASSO NAV Duo Loop eco Catheters because of reports of difficulty retracting and damage to the spine cover. The letter informed the customers of the problems identified and the actions to be taken. Customers with questions were instructed to contact Biosense Webster sales representatives or call (866) 473-7823, Monday through Friday from 7am to 8pm EST. Customers with questions related to the Voluntary Field Removal Certification Form and its return, are instructed to call (909) 839-8468.For questions regarding this recall call 909-839-8500. |  | 464 | 464 units | Nationwide Distribution | Terminated | October 16, 2015 | 233 | 464 | Computer
Z-1321-2015 | 70410 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ADVIA Chemistry Systems Drug Calibrator I; For in vitro diagnostic use in the calibration of Phenobarbital_2 (PHNB_2), Phenytoin_2 (PHNY_2), and Theophylline_2 (THEO_2) methods on the ADVIA Chemistry systems. |  |  | January 19, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc. | Incorrect calibrator lot values were assigned for ADVIA Chemistry Systems Drug Calibrator I, lot 3JD018, for the ADVIA Chemistry Systems Phenytoin 2 (PHNY_2) and Phenobarbital 2 (PHNB_2) assays which was the result of calibrator value miscalculation. | Device Design | Siemens issued an Urgent Medical Device Correction, on/about January 19, 2015, notifying customers about the product, problem, and action to be taken. Customers were instructed to update the calibrator values for Drug Calibrator 1, Lot 3JD018. All customerswere requested to complete and return the Effectiveness Check questionnaire attached to the Urgent Medical Device Correction letter within (30) days. Field service personnel have been sent an SKB (Service Knowledge Base) describing the issue and instructing them how to deal with customer questions. Customers with questions were instructed to contact their local Siemens technical support representative. For questions regarding this recall call 914-631-8000. |  | 657 | 657 | Worldwide Distribution - US (nationwide) and Internationally to Canada and Mexico. | Terminated | September 21, 2015 | 245 | 657 | Computer
Z-1356-2015 | 70763 | No | DTE | External pacemaker pulse generator. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Dual Chamber Temporary External Pacemaker, Model 5392. |  |  | March 17, 2015 | 2015 | Medtronic Inc. Cardiac Rhythm Disease Management | Possible performance issue when used with specific AA-sized (LR6) batteries. The negative terminal of a commercially available AA (LR6) battery did not maintain a sufficient connection with the battery drawer electrical contact. This issue could prevent the EPG from powering on or cause the EPG to abruptly lose primary battery power, potentially stopping delivery of pacing therapy. | Device Design | Medtronic sent an "Urgent Medical Device Correction" letter dated March 2015 to affected customers. The letter was addressed to Risk Manager or Health Care Professional. The letter described the Issue and Customer Actions to be taken. The letter requested customers to complete and return the Customer Confirmation Certificate. For questions contact your Medtronic representative, for technical questions call Instrument Technical Service at 800-638-1991. |  | 6040 | 6040 devices | Worldwide Distribution - US Nationwide in the states of: AL, AK, AZ, AR, CA, CO, CT, DC, FL, GA, ID, IL, IN, KS, KY, LA, MA, MI, MM, MS, MO, MT, NE NV,N, NM, NY, NC, ND, OH,OK, PA, RI, SC, TN, TX, UT, VT, VA, WA, WV, WI, WY, PR, VI and the countries of:Australia, Austria, Bahamas, Belgium, Bolivia ,Botswana, Brazil ,Brunei, Darussalam, Bulgaria, Burkina Faso, Canada, Colombia, Croatia, Cyprus, Czech Republic, Denmark,, Dominican Republic, Egypt, France, Georgia, Germany, Greece, Guyana, Honduras, Hong Kong, Hungary, Iceland, India, Indonesia, Ireland, Israel, Italy, Jordan, Kazakhstan, Korea, Kuwait, Lebanon, Luxembourg, Macedonia, Malaysia, Mauritius, Morocco, Mozambique, Namibia, Nepal, Netherlands, New Zealand, Norway, Oman, Pakistan, Papua, New Guinea, Peru, Philippines, Poland, Portugal, Qatar, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, St. Pierre and Miquelon,Sweden, Switzerland, Tanzania, Thailand, Trinidad and Tobago, Tunisia, United Arab Emirates, United Kingdom, Uruguay, Uzbekistan, Venezuela, Viet Nam, and Yemen.. | Terminated | October 08, 2015 | 205 | 6040 | Computer
Z-0180-2016 | 72229 | No | HWA | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Reverse Shoulder Prosthesis Stem Impaction Fixture REF 804-03-053 |  |  | September 14, 2015 | 2015 | Encore Medical, Lp | During the assembly, the impaction forces caused the polymer, black acetal copolymer from the Impaction Fixture to wear off on the lateral surface of the Humeral Stem titanium plasma spray coating. | Device Design | DJO Global sent an Urgent Field Safety Notice dated September 15, 2015, to all affected consignees. The notice identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed become familiar with the new techniques and notify all of their applicable users of the RSP Monoblock and Turon stems of these revised methods. When each user has been notified of these new techniques, they should complete the form that was provided with the notification indicating that they have been informed of these changes and understand the new methods. Consignees were asked to complete and return forms. Consignees with questions should call 512-834-6255. |  | 634 | 634 units | Nationwide Distribution | Open, Classified |  |  | 911 | Computer
Z-0219-2016 | 70002 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | HeartStart MRx monitor/defibrillator |  |  | December 23, 2014 | 2014 | Philips Medical Systems, Inc. | MRx monitor/defibrillator could reboot at an indeterminate time, potentially causing therapy to be interrupted or delayed | Software design | Philips Healthcare issued on 12/23/14 The Urgent Medical Device Correction notification/Field Safety Notice informs customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue.A software upgrade will be provided free of charge for all units affected by one or more of these issues. A Philips Healthcare representative will contact customers with affected devices to arrange for installation of the upgrade. Philips is asking customers to follow the Action to be Taken by Customer/User section of the UrgentMedical Device Correction notification/Field Safety Notice:You can continue to use your MRx prior to receiving the software upgrade. If the MRx reboots during clinical use, continue to treat the patient according to your organizations protocol, reinitiating therapy if necessary. |  | 75693 | 75,693 units | Worldwide Distribution - US Nationwide and the countries of AFGHANISTAN, ALBANIA, ARGENTINA, ARUBA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BANGLADESH, BELGIUM, BOLIVIA, BOSNIA AND HERZEGOWINA, BRAZIL, BRUNEI DARUSSALAM, BULGARIA, CAMBODIA, CHILE, CHINA, COLOMBIA, COSTA RICA, CROATIA (local name: Hrvatska), CYPRUS, CZECH REPUBLIC, DENMARK, ECUADOR, EGYPT, EL SALVADOR, ESTONIA, ETHIOPIA, FINLAND, FRANCE, GERMANY, GHANA, GREECE, GUADELOUPE, HONG KONG, ICELAND, INDIA, INDONESIA, IRAN (ISLAMIC REPUBLIC OF), IRAQ, IRELAND, ISRAEL, ITALY, JAPAN, Jordan, KAZAKHSTAN, KENYA, REPUBLIC OF KOREA, KUWAIT, LAOS, LATVIA, LEBANON, LIBYAN ARAB JAMAHIRIYA, LITHUANIA, LUXEMBOURG, MACAU, MACEDONIA, MALAYSIA, MALTA, MEXICO, MONGOLIA, MOROCCO, MOZAMBIQUE, MYANMAR (Burma), NAMIBIA, NEPAL, NETHERLANDS, NETHERLANDS ANTILLES, NEW CALEDONIA, NEW ZEALAND, NICARAGUA, NIGERIA, NORWAY, OMAN, PAKISTAN, PALESTINE, PANAMA, PAPUA NEW GUINEA, PARAGUAY, PERU, PHILIPPINES, POLAND, PORTUGAL, PUERTO RICO, QATAR, REUNION, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SERBIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SRI LANKA, SWEDEN, SWITZERLAND, TAIWAN, UNITED REPUBLIC OF TANZANIA, THAILAND, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, TURKMENISTAN, UGANDA, UKRAINE, UNITED ARAB EMIRATES, and UNITED KINGDOM. | Open, Classified |  |  | 75693 | Computer
Z-0218-2016 | 72236 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Digital Lightbox, BrainLAB system, image processing, radiological |  |  | September 21, 2015 | 2015 | Brainlab AG | Potentially incorrectly displayed objects when actively deselecting a fused reference dataset. | Software design | Brainlab sent A FIELD SAFETY NOTICE / PRODUCT NOTIFICATION letter dated 8/27/2015 to medical users regarding Content Manager, Patient Browser and the DICOM Viewer. The Product Notification letter decribes the corrective action information and to advise the medical users of the action Brainlab is taking to address the issue. If you require further clarification, the firm advises to contact the local Brainlab Customer Support Representative. Customer Hotline: +49 89 99 15 68 44 or +1 800 597 5911 (for US customers)E-mail: support@brainlab.com (for US customers: us.support@brainlab.com)Fax: Brainlab AG: + 49 89 99 15 68 33Address: Brainlab AG (headquarters), Kapellenstrasse 12, 85622 Feldkirchen, Germany. |  | 897 | 897 | Worldwide Distribution - US Nationwide and the countries of: Algeria, Argentina, Australia, Austria, Belgium, Brazil, Cambodia, Canada, Chile, China, Colombia, Czech Republic, Denmark, Ecuador, Egypt, Finland, France, Germany, Hong Kong, Hungary, India, Iran, Ireland, Israel, Italy, Japan, Lithuania, Malaysia, Malta, Mexico, Moldova, Morocco, Myanmar, Netherlands, New Zealand, Norway, Philippines, Poland, Portugal, Qatar, Russian Federation, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom and Vietnam. | Open, Classified |  |  | 897 | Computer
Z-1441-2015 | 70749 | No | KPY | Personnel protective shield. | Radiology | Radiology | 510(K) Exempt | Kenex Radiation Shield installed with GE Healthcare Interventional, Surgery, CT and X-Ray Imaging systems. |  |  | March 09, 2015 | 2015 | GE Healthcare | GE Healthcare has recently become aware of an improper installation of a ceiling suspended pendent supporting a radiation shield and/or surgical lamp. The pendent system was not installed per specified installation requirements. There has been a reported incident of a radiation shield falling from the ceiling causing a minor injury. | Equipment maintenance | GE Healthcare sent an Urgent Medical Device Correction letter dated March 9, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. For questions in USA contact 1-800-437-1171. For other countries contact your local GE Healthcare Service Representative. |  | 797 | 797 (3 US, 794 OUS). | Worldwide Distribution - US Distribution to the states of : CA and MA., and to the countries of : SPAIN, INDIA, JAPAN, AZERBAIJAN, BULGARIA, BELARUS, SWITZERLAND, CZECH REPUBLIC, GERMANY, EGYPT, FINLAND, FRANCE, UNITED KINGDOM, GUYANA, IRELAND, ISRAEL, ITALY, NIGERIA and NORWAY. | Terminated | December 21, 2015 | 287 | 797 | Computer
Z-1270-2015 | 70443 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | GE Healthcare Engstrom Carestation, Engstrom Pro, Cart Only S/5 EC. GE model numbers 1505-8538-000, 1505-9000-000, and 1505-9003-000. Ventilators. |  |  | January 27, 2015 | 2015 | GE Healthcare | Potential failure of the caster mounting hardware, which may result in the ventilator falling/tipping (overbalancing). | Device Design | GE Healthcare sent an Urgent Medical Device Correction letter dated January 26, 2015, Ref#34063. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. For questions contact GE Healthcare Service at 1-800-437-1171, or your local Service Representative. |  | 3213 | 3,213 (437 US, 2,776 OUS) | Worldwide Distribution. US nationwide, ALBANIA, ARGENTINA, AUSTRALIA, AUSTRIA, BAHRAIN, BANGLADESH, BELGIUM, BOSNIA &amp; HERZEGOVINA, BRAZIL, CANADA, CAMEROON, CANDA, CHILE, COLOMBIA, COSTA RICA, CROATIA, CZECH REPUBLIC, DENMARK, ECUADOR, EGYPT, EL SALVADOR, FINLAND, FRANCE, GERMANY, CHANA, GREECE, GUATEMALA, HONDURAS, INDIA, INDONESIA, IRAQ, ISRAEL, ITALY, JAPAN, KAZAKHSTAN, KENYA, REPUBLIC OF KOREA, KUWAIT, LEBANON, LIBYAN ARAB JAMAHIRIYA, MACEDONIA, MALAYSIA, MALTA , MEXICO, MOROCCO, NETHERLANDS, OMAN, PANAMA, PAPUA NEW GUINEA, PERU, POLAND, PORTUGAL, QATAR, RUSSIAN FEDERATION, SAUDI ARABIA, SERBIA, SINGAPORE, SLOVAKIA, SOUTH AFRICA, SPAIN, SRI LANKA, SWITZERLAND, SYRIAN ARAB REPUBLIC, THAILAND, TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, VIETNAM, VENEZUELA, and URUGUAY. | Terminated | February 03, 2016 | 372 | 3213 | Computer
Z-0264-2016 | 72475 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | CARESCAPE VC150 Vital Signs Monitor |  |  | October 22, 2015 | 2015 | Innokas Medical Oy | Monitor may shut down unintentionally without restarting. | Software design | The firm, Innokas Medical, sent a "Urgent Field Safety Notice(s)" dated October 22, 2015 and October 23, 2015 by e-mail to their consignees. The notice informed the distributor of the device, about the deficiency, and correction of the device. The notice included the instructions regarding what to do with the product and information that the devices will be updated. The consignees were instructed to complete and return the Medical Device FSCA/Recall Return Response Acknowledgement and Receipt Form via fax to +358 8 562 3151, mail to: Innokas Medical Oy, Vihikari 10, FI-90440-Kempele, Finland or email to: feedback@innokasmedical.fi.If you have any questions, contact the QA/RA Director at +358 44 562 3131. |  | 424 | 424 units (38 in US) | Worldwide Distribution: US (nationwide) and Europe. | Open, Classified |  |  | 424 | Computer
Z-1316-2015 | 70657 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ExacTrac versions 6.x patient positioning systems are used to position patients during radiosurgery or radiotherapy procedures.. |  |  | February 16, 2015 | 2015 | Brainlab AG | The ExacTrac 6.x Patient Positioning System may incorrectly position the patient when using the ExacTrac Cone Beam CT (CBCT) with a TrueBeam-specific optional subvolume-CBCT. | Software in the Use Environment | BrainLab sent an FIELD SAFETY NOTICE / PRODUCT NOTIFICATION letters dated February 16, 2015, to all affected customers. The letters included instructions for customers to: 1) Do not use any actively re-reconstructed CBCT subvolumes from the Varian TrueBeam system with the Brainlab ExacTrac 6.x CBCT Import &amp; Alignment Software module; 2) Import exclusively original, not modified, CBCT volumes into ExacTrac to use for patient positioning; and, 3) Continue to verify any ExacTrac CBCT-based correction using the ExacTrac X-ray verification and/or an external IGRT procedure as mandatory. Brainlab provided existing potentially affected ExacTrac v.6.x CBCT Import &amp; Alignment Software customers (with a Varian TrueBeam system) with the product notification information. Brainlab will provide a software update with this issue solved to affected customers. Brainlab will actively contact the customer starting August 2015 to schedule the update installation.For further questions please call Customer Hotline + 49 89 99 15 68 44 or 1-800 597-5911 ( For US Customers only) |  | 24 | 24 systems (US); 36 systems (Foreign) | Worldwide Distribution - USA and to the countries of : Australia, Austria, Canada Denmark, France, Germany, Hong Kong, Japan, Russia, Singapore, South Korea and United Kingdom. | Open, Classified |  |  | 24 | Computer
Z-0238-2016 | 72395 | research use only/microbiology | OTT |  |  | Microbiology | Enforcement Discretion | Focus Diagnostics Anaplasma Phagocytophilum IFA IgM Test Kit, Model No. IF1450M in vitro diagnostic. |  |  | October 02, 2015 | 2015 | Focus Diagnostics Inc | Focus Diagnostics is recalling the Anaplasma phagocytophilum IFA IgM kit because of the potential increase in invalid results. | Device Design | A recall letter dated 10/2/15 was sent to distributors who distributed the Anaplasma phagocytophilum IFA IgM (IF1450M) kit lots 26431 and 27011. The letter informs the customers that Focus Diagnostics received customer complaints of IF1450M with low reactivity, which may result in higher invalid rate. The letter also provides the customers with Focus Diagnostics' recommendations and actions to be taken by the distributors. Customers with questions or require additional information are instructed to contact Focus Diagnostics' Technical Services department at (562) 240-6550 from 7am to 5pm (PST) or send an email to DxTS@focusdx.com. |  | 38 | 38 kits | Non US (worldwide): Czech Republic, Denmark, Germany, Italy, Spain and Colombia | Terminated | January 29, 2016 | 119 | 38 | Computer
Z-0984-2015 | 69880 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARTISTE MV System, Linear Accelerator (LINAC) with SysVC10A software. |  |  | November 19, 2014 | 2014 | Siemens Medical Solutions USA, Inc | There may be an existing dark current phenomenon on ARTISTE LINAC in combination with IMRT or mARC treatments using unflat beams. Software issue. | Software design | Siemens Planned Action: 1. Providing recommendations for user actions in terms of routine dosimetric quality control checks for any IMRT treatment at very high dose rates (Multiple X) as well as for mARC treatments.2. Software update for the Function Controller #3, which controls parts of the beam generation. This software update enables the correct switching between Pulse Forming Network (PFN) and reduced Pulse Forming Network (iPFN) in pause state, to reduce the dark current during the movement as much as possible. This software will be installed by a Siemens service representative.3. Siemens will contact the user to make an appointment to implement the modification system at a date and time when the system will be available. These corrections shall be made free of charge.The deployment of the permanent correction shall be completed by July 1, 2015.For further questions please call (610) 219-6300. |  | 3 | 3 | US Distribution to the states of UT, WI, and NY. | Terminated | April 28, 2015 | 160 | 3 | Computer
Z-1211-2015 | 70478 | No | RCE |  |  |  |  | GE Inspection Technologies x¦tome¦x L240 an Cabinet X-ray system |  |  | January 13, 2015 | 2015 | GE Inspection Technologies, LP | It was discovered that the upper left corner-shaped lead shield in the window in one of the doors of the v¦tome¦x L240 installed had lost adhesion and fall out of position, which could lead to gaps in shielding. | Device Design | GE Inspection Technologies, Planned Action: 1. GE will notify customers of the affected cabinet x-ray systems through a Technical Information Letter. GE notified its affected customers via email on January 13, 2015, via letter dated January 14, 2015, and via telephone on January 19, 2015. 2.GE will install an interim insert within 30 days of receipt of this letter. The insert will cover any existing or interim gaps in the original lead shielding on the door.3. GE will replace the doors of affected cabinet x-ray systems no later than 180 day of receipt of this letter.4. These corrections will be made free of charge.The Technical Information Letter used for notification of all dealers and purchasers is to be made within 15 working days of receipt of this letterFor further questions, please call (717) 447-1278 |  | 2 | 2 | US Distribution to the states of OH and PA. | Terminated | August 17, 2015 | 216 | 2 | Computer
Z-0242-2016 | 72331 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Triton Infusion Pump (model 300000) and Triton fp Infusion Pump (model 400000) (proprietary name: AFF (Mark I) Volumetric Infusion Pump). Used to pump fluids into a patient in a controlled manner. |  |  | October 01, 2015 | 2015 | WalkMed Infusion, LLC | The device might fail to detect air in line while infusing the medication Venofer. This may occur when a drop of Venofer gets stuck in the tubing adjacent to the bubble detector and the Total Volume Limit is programmed higher than the actual bag volume. | Device Design | A Field Notification (letter) was provided to users on 10/01/15 to remind them of the following:use air elimination filters, as appropriate, not to over program into the device the volume to be infused, remove air from the IV bag prior to starting the infusion, issues associated with the use of Venofer or similar drugs. No products are being returned. |  | 4695 | 4695 units | Distributed in the states of AL, CA, CO, DE, FL, GA, IL, MA, MD, MI, MN, MS, NC, NJ, NM, NV, NY, OH, PA, TN, TX, and WA. | Open, Classified |  |  | 4695 | Computer
Z-0205-2016 | 71518 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Arthroscopic Energy 50Âż Probe with Suction, Catalog number: AES-50S |  |  | June 17, 2015 | 2015 | ConMed Corporation | Probe programming error. The incorrect program can generate more heat at the surgical site, resulting in an elevated probability of tissue damage. | Software design (manufacturing process) | On June 17, 2015 Con Med Corporation distributed Urgent Device Recall Notification Letters and Recall Reply forms to their customers via courier service. The letter identified the affected product, problem and actions to be taken. Customers were instructed to examine their inventory for any of the affected lot numbers referenced on the recall notice. If any of the recalled product are in stock, customers should discontinue the use or distribution of these items. If product was further distributed, customers should notify those individuals of this recall action. Customers with questions can call Gregory Connell at 727-399-5276 or email AES50S@conmed.com. Customers should also complete and return the Recall Reply Form by mail to the address listed on the recall notice. |  | 276 | 276 devices | Worldwide - US distribution in the states of FL, OH, and the country of Spain. | Terminated | April 11, 2017 | 664 | 276 | Computer
Z-1284-2015 | 70490 | Yes | OAE |  |  | Cardiovascular | PMA | TactiSys Quartz Pack, PN-004 400. |  |  | November 25, 2014 | 2014 | St Jude Medical | A configuration update needs to be done on TactiSys to appropriately recognize all TactiCath catheters. | Software design | St. Jude Medical sent a Service Bulletin to inform of a TactiSys Configuration File update that needs to be performed. A technical service personnel will visit each consignee and performed the software change on-site. For questions about this Service Bulletin, contact Technical Support(ETechsupport@sjm.com). |  | 89 | 89 | US Nationwide Distribution in the states of: AL, AZ, CA, CO, CT, FL, GA, ID, IL IN, KS, LA, MD, MA, MI, MN,MS, NE, NV,NH, NM, NY, NC, OH, OK OR, PA, TN, TX, UT, VA,WA. | Terminated | May 06, 2015 | 162 | 89 | Computer
Z-0210-2016 | 72217 | No | FLE | Steam sterilizer. | General Hospital | General Hospital | 510(k) | AMSCO 400 and AMSCO C Small Steam Sterilizers |  |  | September 03, 2015 | 2015 | Steris Corporation | In AMSCO Small Steam Sterilizers equipped with the CS-iQ feature, the sterilizer software inadvertently prevents the capability to operate the foot pedal, which allows the Operator to open and close the sterilizer door. | Software design | On 9/21/2015 the firm sent Recall Correction Notices to their customers. |  | 32 | 32 units | US nationwide distribution. | Terminated | January 22, 2016 | 141 | 32 | Computer
Z-1305-2015 | 70687 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE Healthcare, SIGNA 1.5T TWINSPEED, SIGNA INFINITY MRISYSTEM, SIGNA 3.0T INFINITY WITH EXCITE, SIGNA 3.0T WITH EXCITE, SIGNA EXCITE 1.5T, SIGNA EXCITE 3.0T, GE 1.5T AND 3.0T SIGNA HDX MR SYSTEM, Discovery MR750 3.0T, Discovery MR450 1.5T, Discovery MR750w 3.0T, Optima MR450w 1.5T, SIGNA Contour, Brivo MR355, Optima MR360, GE 1.5T SIGNA HDE MR SYSTEM, Signa Openspeed 0.7T MR SYSTEM, Vectra, SIGNA 0.35T OVATION WITH EXCITE, SIGNA 0.2T PROFILE With EXCITE |  |  | March 09, 2015 | 2015 | GE Healthcare | GE Healthcare has become aware of a potential safety issue involving MRI systems due to software versions not being maintained properly at some sites. | Software Manufacturing/Software Deployment | GE Healthcare sent an "Urgent Medical Device Correction" letter dated March 9, 2015, to all affected customers. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. For questions contact your local service representative or call 800-437-1171 (available 24 hours a day). |  | 9369 | 9,369 (2,937 US, 6,432 OUS). | Worldwide Distribution - USA (nationwide) Distribution including PR, GU and DC., and to the countries of: ALBANIA. ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BELARUS, BELGIUM, BOLIVIA, BOSNIA AND HERZEGOVINA, BRAZIL, BULGARIA, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, COTE DLVOIRE, CROATIA, CYPRUS, CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, ECUADOR, EGYPT, EL SALVADOR, ESTONIA, FINLAND FRANCE, FRENCH POLYNESIA, GEORGIA, GERMANY. GHANA, GREECE, GUADELOUPE, GUATEMALA, HONDURAS, HONG KONG, HUNGARY, ICELAND, INDIA, INDONESIA, IRAN IRAQ, IRELAND, ISRAEL, ITALY, JAMAICA, JAPAN, JORDAN, KAZAHHSTAN,KENYA,REPUBLIC OF KOREA, KUWAIT, LATVIA, LEBANON, LIBYAN ARAB JAMAHIRIYA, LITHUANIA, LUXEMBOURG MACEDONIA, MALAYSIA, MEXICO, MOROCCO, NETHERLANDS, NEW CALEDONIA, NEW ZEALAND, NICARAGUA, NIGERIA, NORWAY, OMA N,PAKISTAN, PALESTINE, PANAMA, PARAGUAY, PERU, PHILIPPINES, POLAND, PORTUGAL, REUNION, ROMANIA, RUSSIA, SAUDI ARABIA, SERBIA, SIERRA LEONE, SINGAPORE, SLOVAKIA, SOUTH AFRICA, SPAIN SUDAN, SWEDEN, SWITZERLAND, SYRIA, TAIWAN, THAILAND, TUNISIA, TURKEY, UGANDA , UKRAINE, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY, UZBEKISTAN, VENEZUELA, VIETNAM and ZAMBIA. | Terminated | September 08, 2015 | 183 | 9369 | Computer
Z-1283-2015 | 70341 | No | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Triathlon Distal Capture Assembly |  |  | February 12, 2015 | 2015 | Stryker Howmedica Osteonics Corp. | Potential for weld disassociation of the cross pin, which could lead to loose or disassociated trigger mechanism on the Distal Capture Assembly. | Device Design | On March 20, 2015, Stryker sent Urgent Medical Device Recall Notifications/Urgent Medical Device Recall Notification Acknowledgement Forms dated March 19, 2015 to all affected customers. The letter identified the product, the related issue and the action to be taken by the customers. Customers were asked to return the attached acknowledgement form confirmed they received and reviewed the notification. Customers with questions were instructed to call 201-831-6365. Stryker sent an Urgent Medical Device Correction letter dated February 11, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to return the attached acknowledgement form confirming they have received and reviewed the letter. Customers with questions were instructed to call 201-831-6365. |  | 2030 | 2030 units | Worldwide Distribution - USA (nationwide) Puerto Rico, and Internationally to Canada, Hong Kong, Netherlands, Sweden, Switzerland, Spain, France, Italy, UK, India, Australia, Korea, Panama, Polan, Ireland, Kuala Lumpar, New Zealand and South Africa. | Open, Classified |  |  | 2030 | Computer
Z-0243-2016 | 72502 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Amphirion Deep OTW PTA Balloon Catheter: |  |  | October 23, 2015 | 2015 | Medtronic Vascular, Inc. | Medtronic has discovered that select lots of the Amphirion Deep PTA Balloon Dilation Catheter have a potential breach in the pouch seal enclosing the sterilized product. | Equipment maintenance | Beginning October 23, 2015, Medtronic will send the customer letter andconfirmation form to all identified US customers via 2-day UPS mail delivery. Internationally, communications will be sent per locally approved methods. |  | 5895 | 1652 units in US; 4243 units OUS. | Worldwide Distribution-US (nationwide) including Puerto Rico and the countries of Australia Austria Belgium Brazil China Colombia Costa Rica Czech Republic Denmark Dominican Republic Egypt France Georgia Germany Hong Kong Hungary India Iran Ireland Israel Italy Japan Jordan Kazakhstan Korea Kuwait Latvia Lebanon Lithuania Luxembourg Mexico Netherlands Norway Oman Pakistan Panama Peru Philippines Poland Portugal Russian Federation Saudi Arabia Singapore Slovakia Slovenia South Africa Spain Sri Lanka Sweden Switzerland Syria Taiwan Thailand Trinidad and Tobago Turkey United Arab Emirates United Kingdom and Uruguay.. | Terminated | July 31, 2017 | 647 | 5895 | Computer
Z-1267-2015 | 70556 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | SIEMENS SOMATOM Force with software version VA50A and /or VA50A_FP1 and/or VA50A_SP0; Computed tomography system. |  |  | February 06, 2015 | 2015 | Siemens Medical Solutions USA, Inc | Possibility of image artifacts during data acquisition when using Adaptive Cardio Sequence, Turbo Flash, and Head modes. The update improves system start-up behavior and resuming, ECG handling with visual feedback of correct placement of ECG leads and contact quality, and other safety related issues. | Software design | Siemens sent a Customer Advisory Notice dated February 6, 2015, to all customers via USPS Certified Mail. The notice identified the affected product, problem and actions to be taken. |  | 12 | 12 | US Nationwide Distribution in the states of PA, MA, MN, NY, SC, MD, NY, IL, OH, TX and KY. | Terminated | August 31, 2015 | 206 | 12 | Computer
Z-0250-2016 | 72500 | No | NBH | Arthroscope. | Orthopedic | Orthopedic | 510(K) Exempt | NovoCut Suture Manager; Ceterix catalog number CTX-001. |  |  | October 13, 2015 | 2015 | Ceterix Orthopedics, Inc. | The NovoCut Suture Manager device may cut the suture during knot tensioning and this may cause the suture knot to unravel. | Device Design | Ceterix Orthopaedics sent an Urgent Field Safety notification letter dated October 13, 2015, to all affected sites. The product was identified, reason for recall explained and letter contains list of product lot numbers affected. Customers were instructed to identify and quarantine the affected product, contact Ceterix Orthopaedics to arrange a no cost return, and complete and return the attached acknowledgement form. Customers with questions were instructed to call 1-888-305-0615 or email: docs@ceterix.com. For questions regarding this recall call 650-396-3478. |  | 1364 | 1364 units | Nationwide Distribution | Terminated | February 12, 2016 | 122 | 1364 | Computer
Z-0251-2016 | 72376 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Shimadzu C-arm Table, Catalog No. IVS-110, Medical Imaging Table. |  |  | September 30, 2015 | 2015 | Shimadzu Medical Systems | The C-arm may interfere with the table top under certain circumstances. | Software design | An Urgent Voluntary Medical Device Recall notice was sent on 9/30/2015 to all the customers who purchased the Shimadzu C-arm Table (Catalog No. IVS-110). The notification informs the customers of the description of the issue, the actions to be taken by Shimadzu to correct the issue, and the actions to be taken by the customers. Customers with questions or concerns about the device correction, are instructed to contact their local Shimadzu representative or contact ShimadzuNational Technical Support at (800) 228-1429. |  | 26 | 26 units total (4 units in US) | Worldwide Distribution. US (Texas, Wisconsin, Ohio, and Connecticut), Iran, India, Saudi Arabia, Serbia, Thailand, Germany, Croatia, Turkey, Lebanon, Korea, and China. | Open, Classified |  |  | 26 | Computer
Z-1369-2015 | 70775 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ACCU-CHEK Connect Diabetes Management App; Instruction Manual |  |  | October 30, 2014 | 2014 | Roche Diabetes Care, Inc. | Roche Diabetes Care has become aware of an issue with the Accu-Chek Connect diabetes management app that could potentially lead to inaccurate bolus advice being provided to the user. A thorough investigation of the situation revealed that this issue may occur if the user changes the screen orientation of the phone from portrait to landscape or vise versa while looking at the Bolus Advisor or Carb | Software design | Urgent Field Safety Notices were sent to the customer and healthcare professional starting on 10/30/2014. The letter provided a description of the reason for the notification; and recommended that all users should update their app immediately. |  | 72772 | The application has been downloaded total of 644 times (Italy-219, South Africa-24, Germany-401) and there are 113 bolus advice activations. | There are no US distributors as the affected device is not marketed in the US. The affected device was distributed in Germany, Italy and South Africa. The application has been downloaded total of 644 times (Italy-219, South Africa-24, Germany-401) and there are 113 bolus advice activations. | Terminated | June 08, 2015 | 221 | 72772 | Computer
Z-1314-2015 | 70578 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Variable Angle Locking Peg, D:1.8mm x L:14.0mm; Variable Angle Locking Peg, D:1.8mm x L:16.0mm, Variable Angle Locking Peg, D:1.8mm x L:18.0mm; Variable Angle Locking Peg, D:1.8mm x L:20.0mm; Variable Angle Locking Peg, D:1.8mm x L:22.0mm, Variable Angle Locking Peg, D:1.8mm x L:24.0mm; Variable Angle Locking Peg, D:1.8mm x L:26.0mm. |  |  | February 16, 2015 | 2015 | Flower Orthopedics Corporation | Locking pegs were not locking during distal radius procedures. There has been one (1) report of revision surgery scheduled to remove a loosened peg. | Device Design | On February 16, 2015, Flower Orthopedics generated and distributed a field memo/advisory notice (FOCMB-15-02) to all distributors with locking peg inventory on February 6, 2015. The field memo provided technique guidance to end users. All 12 consignees with remaining inventory (ie.not reported as used) were contacted directly via phone and all have agreed to ship the locking pegs back to Flower Orthopedics. Flower Orthopedics determined that a generic statement to all customers was not necessary for the following reasons detailed above and summarized again here: product was distributed only through direct channels, the health hazard evaluation has determined there is a very low risk to health, and all consignees have been directly contacted already and product inventory is being returned. |  | 901 | 901 | Distributed in the states of VA, TX, IL, OK, PA, CA, NY, KY, and FL. | Terminated | August 06, 2015 | 171 | 901 | Computer
Z-1263-2015 | 70264 | [6 empty] | VITROS Hand-held Barcode Scanner, Catalog Number 6844210, for use with the following VITROS Systems: VITROS 5,1 FS Chemistry System, VITROS 3600 Immunodiagnostics System, VITROS 4600 Chemistry System, VITROS 5600 Integrated System. This product is an accessory for use with VITROS Analyzer Systems. |  |  | December 29, 2014 | 2014 | Ortho-Clinical Diagnostics | In combination with the Batch Programming option (only) for programming samples, the hand-held barcode scanner unexpectedly skips sample cup positions within a sample tray. This issue only occurs when using the Batch Programming option; all other sample programming options function as intended. | Software design | On 12/29/2014, Customer letter (Ref. CL2014-315) was sent via FedEx overnight courier or ORTHO PLUS e-Communications and/or US Postal Service Priority Mail (for PO Boxes only) to US consignees (OCD Direct, Drop-Ship and US Federal Government) to inform them of the issue and provided with instructions for reconfiguring their Barcode Scanners. |  | 72 | Domestic: 62 units; Foreign: 10 units | Worldwide Distribution-US (nationwide) and the countries of Canada, Australia and England. | Terminated | March 27, 2017 | 819 | 72 | Computer
Z-0151-2016 | 72319 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | HAMILTON-G5 Ventilator with software versions between V2.00 and V2.31. |  |  | April 22, 2014 | 2014 | Hamilton Medical, Inc. | Ventilation and alarms of a HAMILTON-G5 ventilator can be suppressed unintentionally after the activation of a suctioning maneuver by the operator; this situation can occur regardless of the selected patient group (neonatal, pediatric and adult). | Device Design | Hamilton Medical issued a letter dated April 22, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were notified that their Hamilton Medical Account Manager would contact them to coordinate the deactivation of the Ventilation Suppression feature. Customers were asked to review the Medical Device Field Safety Corrective Action.Customers with questions were instructed to call 800-426-6331, Ext 215. |  | 1128 | 1128 | Nationwide Distribution | Terminated | August 24, 2017 | 1220 | 1128 | Computer
Z-0225-2016 | 72472 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | The Centricity PACS-IW with Universal Viewer is a device that displays medical images (including mammograms) and data from various imaging sources. Images and data can be viewed, communicated, processed and displayed within the system or across computer networks at distributed locations. Typical users of this system are trained professionals, including but not limited to radiologists, physicians, nurses, medical technicians, and assistants. |  |  | September 28, 2015 | 2015 | GE Healthcare | Inaccurate distance measurements with magnified projection X-ray images. | Software design | An URGENT MEDICAL DEVICE CORRECTION letter was sent to all medical users on 9/28/2015 regarding Centricity PACS-IW with Universal Viewer and Centricity Universal Viewer- Inaccurate distance measurements with magnified projection X-ray images.GE Healthcare explains that for all projection X-rays, GE Healthcare recommends that the user perform the necessary measurements on the standard contact images, and not rely on measurements made on the magnified views unless the measurements have been independently verified for each X-Ray machine installed or manually calibrated using a radio-opaque fiducial of a known size. Attention should be given to prior cases acquired before the application of the safety instructions given in this communication as they may be affected by the safety issue. GE Healthcare will correct all affected systems by providing a software update at no cost to the customer. GE Healthcare advises customers if they have any questions or concerns regarding the notification, to please contact GE Healthcare Service at1-800-437-1171 or their local Service Representative. |  | 5 | Centricity PACS-IW with Universal Viewer Versions 5.0 SP2 and higher 1052 units | Worldwide Distribution-US (nationwide) including the states of AZ, CA, CO, DE, FL, GA, IA, IL, IN, LA, MA, MI, MN, MO, NJ, NY, OR, PA, SC, SD, TN, TX, VA, WA, WI and WV, and the countries of Australia, Austria, Bangladesh, Belgium, Brazil, Canada, Chile, China, Ecuador, France, Germany, Ghana, Hong Kong, India, Israel, Italy, Japan, Kenya, Korea, Malta, Malaysia, Namibia, Netherlands, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Turkey, Taiwan, United Arab Emirates, United Kingdom and Venezuela. | Open, Classified |  |  | 11 | Computer
Z-0252-2016 | 72160 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | Smooth Muscle Action (SMA) Immunohistochemistry reagent. Catalog numbers: CME305A-0.1ml; CME305B - 0.5 ml, PME305AA - 6.0 ml prediluted and ready to use. They are the same product with the same intended use. |  |  | September 08, 2015 | 2015 | Biocare Medical, LLC | Certain catalog numbers and lots of Smooth Muscle Actin may stain some cell types that are typically expected to be negative for SMA. | Device Design | Biocare sent a Medical Device Recall Notification letter dated September 8, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately discontinue and dispose of all affected product. Customers were asked to complete an return the enclosed Return Response Form. Customer's accounts would be credited for all affected product.For questions regarding this recall call 925-603-8033. |  | 151 | 151 units | Worldwide Distribution - US including CO, FL, OH, SC, AZ, AL, FL, CT, NJ, CA, DC, KY, NC, TX, ID, MA, MD, OR, LA, PA, KS, NY, UT, GA and Internationally to Spain, Sweden, Switzerland, Hungary, Italy, Germany, Turkey, Canada, India, Bolivia, Mexico, Indonesia and Australia. | Terminated | January 11, 2016 | 125 | 151 | Computer
Z-1394-2015 | 70748 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | INFX-8000V Interventional Angiography System. CAT-880B catheterization table. |  |  | March 11, 2015 | 2015 | Toshiba American Medical Systems Inc | The table base rotational brake that holds the table top may not always be engaged when any force or side impact is applied to the table top. | Device Design | Toshiba sent a Urgent Medical Device Correction letter dated March 11, 2015, to all customers who purchased the CAT-880B catherization table for the Tosiba Infinix X-Ray Interventional System. The letter informed the customers of the problems identified and the actions to be taken. Customers were instructed to complete and return the attached form and fax it to (877) 349-3054 or email to raffairs@tams.com. Customers with questions were instructed to call (800) 421-1968 or contact their local Toshiba Representative at (800) 521-1968.For questions regarding this recall call 714-730-5000. |  | 28 | 28 units | Nationwide Distribution | Terminated | September 18, 2015 | 191 | 28 | Computer
Z-0291-2016 | 71818 | Yes | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Baxter HomeChoice and HomeChoice Pro APD systems, designed to provide Automated Peritoneal Dialysis (APD) therapy for pediatric and adult renal patients. |  |  | June 29, 2015 | 2015 | Baxter Healthcare Corp. | Loud operating sounds, which was unacceptable to the end users when the device was powered on. The HomeChoice devices continue to perform as intended, as the issue is only associated with the noise coming from the device and not the function of it. May delay treatment if changing to a different system. | Device Design | Baxter sent an Important Product Information communication to affected customers (patients and clinicians) via U.S.P.S., first-class mail on June 29, 2015. The letter states that Baxter Healthcare Corporation is updating the HomeChoice Patient At-Home Guide to include a discussion of the expected operating sounds associated with the HomeChoice and HomeChoice PRO devices. The new labeling will address the following:The HomeChoice cycler utilizes an air (pneumatic) pump; some sounds related to the operation of this pump are expected as part of normal operation of the HomeChoice cycler. There will be sounds like humming, swishing, clicking, and venting (air being released) that are normal for the operation of the cycler. In certain portions of the therapy, the sound level is expected to increase which is also part of normal operation. If there is a significant change in the sound level or a new, previously unheard sound when using the HomeChoice cycler, please contact your doctor and/or nurse or Baxter Technical Services. The firm advises customers to keep the letter with the Patient At-Home Guide.The firm also requests that customers complete the enclosed home patient reply form (enclosure) and return it to Baxter by faxing the reply form to 224-270-5457, scanning and e-mailing it to fca@baxter.com or by mail in the enclosed, preaddressed, stamped envelope. Returning the home patient reply form promptly will prevent you from receiving repeat notices. If the customer has any questions about their Peritoneal Dialysis (PD) therapy, they are advised to please contact their doctor and/or nurse.Additionally the letter states that any adverse reactions or quality problems experienced with the use of this product may be reported using one of the following options:" Calling Baxter Product Surveillance at 800-437-5176, Monday through Friday, between 8:00 am and 5:30 pm Central Time Baxter will advise customers of the additional product information regarding |  | 29355 | 29355 | Nationwide Distribution, District of Columbia, and Puerto Rico. | Open, Classified |  |  | 29355 | Computer
Z-1355-2015 | 70604 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Various Lots of Expel Drainage Catheter with TwistLoc Hub Biliary Drainage Catheter and Biliary Drainage Catheter Kits |  |  | February 25, 2015 | 2015 | Boston Scientific Corporation | Complaints of device fragmentation after the catheter was implanted in the biliary system. The most serious patient risk for this issue is additional intervention for fragment retrieval using minimally invasive methods and, in some of the reported cases, this was performed to retrieve the device fragments. | Device Design | Boston Scientific sent an Urgent Medical Device Correction - Immediate Action Required letter dated February 5, 2015. to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letter described the problem and the product involved in the recall. Advised consignees to complete and return the "Recall Removal Reply Verification Tracking Form".Your local Sales Representative can answer any questions you may have regarding this Field Correction notice.For further questions please call (763) 494- 7971. |  | 289 | 289 | Worldwide Distribution - US Distribution to the states of : CA, WI, FL and OH., and to the countries of: Belgium, Czech Republic, Finland, France, Germany, Great Britain, Italy, Spain, Switzerland, Austria, Hungary, Netherlands and Portugal. | Terminated | June 30, 2017 | 856 | 289 | Computer
Z-0255-2016 | 72301 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE Healthcare, SIGNA 3.0T. |  |  | August 31, 2015 | 2015 | GE Healthcare | When performing head or neck scans, the currently displayed SAR values could be lower than the actual SAR in the head as predicted by SAR modeling. The predicted HEAD SAR value from the modeling suggests the actual SAR in the head could exceed IEC60601-2-33 limit of 3.2 W/kg for some scans. | Software design | GE Healthcare sent an Urgent Medical Device Correction letter dated August 31, 2015. The letter was addressed to Hospital Administrators/Risk Managers, Radiology Department Managers, &amp; Radiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. For questions contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 19 | 19 | Worldwide Distribution - US: Nationwide including DC, PR and GU and the countries of ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BELARUS, BELGIUM, BERMUDA, BOSNIA HERZEGOVINA, BRAZIL, BULGARIA, CANADA, CHINA, CZECH REPUBLIC, DENMARK, EGYPT, ESTONIA, ETHIOPIA, FINLAND, FRANCE, GERMANY, GREECE, HONG Kong, HUNGARY, INDIA, INDONESIA, IRAN, IRELAND, ISRAEL, ITALY, JAMAICA, JAPAN, KAZAKHSTAN, REPUBLIC OF KOREA, KUWAIT, LEBANON, LUXEMBOURG, MALAYSIA, MEXICO, NETHERLANDS, NEW ZEALAND, NORWAY, PERU, POLAND, PORTUGAL, QATAR, ROMANIA, RUSSIA, SAUDI ARABIA, SERBIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, SYRIA, TAIWAN, THAILAND, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY, and VENEZUELA. | Open, Classified |  |  | 1509 | Computer
Z-0262-2016 | 69714 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart MRx Monitor/Defibrillator |  |  | November 19, 2014 | 2014 | Philips Medical Systems, Inc. | MRx Defib can be susceptible to one or both issues: 1. The C02 Inlet Port associated with end-tidal carbon dioxide (EtCO2) monitoring on MRx Monitor/Defibrillators can be pushed into the MRx housing, making it inaccessible. 2. The handle can separate from the MRx housing due to breakage of mounts on the rear case. | Device Design | Philips Heatlhcare issued on November 19, 2014, an Urgent Medical Device Correction notification/Field Safety Notice. The Urgent Medical Device Correction notification/Field Safety Notice informs customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue.A software upgrade will be provided free of charge for all units affected by one or more of these issues. A Philips Healthcare representative will contact customers with affected devices to arrange for installation of the upgrade. Philips is asking customers to follow the Action to be Taken by Customer/User section of the Urgent Medical Device Correction notification/Field Safety Notice:You can continue to use your MRx prior to receiving the hardware upgrade provided you follow the steps given.For questions regarding this recall call 978-687-1501Customers with questions were instructed to contact their local Philips representative. |  | 27918 | 27,918 units | Worldwide Distribution - US (nationwide) and Internationally to ARGENTINA, ARUBA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BELGIUM, BOLIVIA, BRAZIL, CANADA, CHILE, CHINA, COLOMBIA, CROATIA (local name: Hrvatska), CZECH REPUBLIC, DENMARK, EGYPT, ESTONIA, FINLAND, FRANCE, GERMANY, GREECE, GUADELOUPE, HONG KONG, ICELAND, INDIA, INDONESIA, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KENYA, KOREA, REPUBLIC OF, KUWAIT, LATVIA, LEBANON, LITHUANIA, MALAYSIA, MEXICO, NETHERLANDS, NETHERLAND ANTILLES, NEW CALEDONIA, NEW ZEALAND, NORWAY, OMAN, PHILIPPINES, POLAND, PORTUGAL, PUERTO RICO, QATAR, REUNION, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SPAIN, SRI LANKA, SWEDEN, SWITZERLAND, TAIWAN, PROVINCE OF CHINA, THAILAND, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, UGANDA, UNITED ARAB EMIRATES, UNITED KINGDOM, VENEZUELA, AND YEMEN. | Open, Classified |  |  | 27918 | Computer
Z-1399-2015 | 70591 | No | JGS | Sodium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA Chemistry XPT System, Siemens Material Number 10723034, Software Version 1.0.2, Siemens Material Number 11219493, IVD. |  |  | February 11, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has confirmed an issue with ADVIA Chemistry XPT software version 1.0.2. The system may cause samples to remain in an Inprocess state. Test resultson a sample that is held Inprocess will not transmit to the LIS. Manual intervention is necessary to complete the processing of the samples that are held Inprocess. | Software design | Siemens Healthcare Diagnostics sent Urgent Medical Device Correction Letters (dated 2/11/2015) via FedEx to the domestic consignees informing users of the software issue with ADVIA Chemistry XPT Software Version 1.0.2. The letters included work instructions to handle the bug. Customers are requested to complete and return the Effectiveness Check questionnaire attached to the Urgent Medical Device Correction letter within 30 days. For questions about in process samples, the event message 02 981 8728, or need assistance with mitigation solutions, please contact your Siemens Customer Care Center or your local Siemens Technical Support Representative. Urgent Field Safety Notice Letters were sent to the foreign consignees (distribution is determined at the country level). Field service personnel were notified of the Urgent Medical Device Correction Letter. |  | 32 | 16 units: 3 Domestic, 13 Foreign | Distributed in the states of AZ, CA, and WA, and the countries of Germany, Italy, Spain, and UK. | Open, Classified |  |  | 32 | Computer
Z-0263-2016 | 72171 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Samsung GM60A-40S |  |  | September 14, 2015 | 2015 | NeuroLogica Corporation | When the system is driving over a non-conductive floor and then over a conductive, grounded part, proper control of the drive behavior may not be possible. The system can be stopped by releasing the drive handle or by pressing the emergency switch. Proper behavior is re-enabled when the system is rebooted. | Device Design | An Urgent - Field Safety Notice identifying the affected products, problem involved, root cause, and potential hazards, was sent to customers. In the event the issue occurs, users can release the drive handle or press the emergency switch and the unit will stop. Then, reboot the system. If unintended movements occur, please contact local representative. To correct the issue, service engineers will visit to apply the corrective action. For further information or support concerning this problem, contact local representative. |  | 7 | 7 devices | Distributed to the states of NC, OH, MA, and MO. | Terminated | January 19, 2017 | 493 | 7 | Computer
Z-1260-2015 | 70580 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Synthes Inserter for Titanium Elastic Nails; The Inserter for Titanium Elastic Nails (TEN) is intended for fixation of diaphyseal fractures where the canal is narrow or flexibility of the implant is paramount. This includes: lower extremity fractures in pediatric and small-statured patients and upper extremity fractures in patients. |  |  | February 03, 2015 | 2015 | Synthes, Inc. | The Inserter has the potential for mechanical failures such as breakage of the main shaft of the Inserter for the TEN, malfunction of the chuck of the Inserter for the TEN in the form of the chuck getting stuck/jammed, breakage of the cross bar of the Inserter for the TEN. | Device Design | DePuy Synthes sent an Urgent Notice Medical Device Recall letter dated February 3, 2015, to all affected customers to alert them about the recall and explain the benefits and risks associated with removing this device from the market. It stated that the immediate removal of the product would prevent the performance of emergency surgery, thus DePuy Synthes is not requiring an immediate removal of affected product(s). DePuy Synthes is in the process of developing the replacement &amp; recovery plan for indicated devices. Due to their current but temporary lack of replacement inventory, Synthes proposed to remove and replace the device in increments to prevent any adverse impact on the availability of the device for use on patients in trauma-related surgery events.DePuy Synthes will contact customers as replacements become available.Customers were advised to complete the Verification Section by checking the appropriate box indicating affected product has been located. Also,indicate the number of devices found. If your facility chooses to return the product(s) at this time, please take the following additional steps:Call DePuy Synthes at 1-800-479-6329 to obtain a Return Authorization (RA) Number, include a copy of Verification Section with the product to: Credit/Returns, DePuy Synthes, 1101 Synthes Avenue, Monument, CO 80132.Send the completed Verification Section to the following fax or email address: Fax: (866) 782-7287 or Scan/email: Synthes5095@stericycle.com Customers with questions were instructed to call 610-719-5450 or contact their DePuy Synthes Sales Consultant.For questions regarding this recall call 810-719-5000. |  | 1550 | 1550 | Nationwide Distribution | Open, Classified |  |  | 1550 | Computer
Z-1359-2015 | 70683 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Computed Tomography X-ray Systems (Brilliance CT 64-channel w/Essence technology, Brilliance iCT, Brilliance iCT SP, Ingenuity Core, Ingenuity Core128 &amp; Ingenuity CT), Philips Medical Systems, Cleveland, OH |  |  | February 05, 2015 | 2015 | Philips Medical Systems (Cleveland) Inc | Philips discovered that a software defect exists in marketed product wherein the sign indication of the longitudinal position of some types of scan is inverted. | Software design | Philips sent an Customer Information letter dated February 5, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. They informed their customers that they should refrain from using Multi Planar Reconstruction (MPR) derived images for identifying the absolute position of the intervention for feet-first patient orientation. A Philips Field Service Engineer will contact you to schedule the software installation at your site.If you need any further information or support concerning this issue, please contact your local Philips representative.Philips apologizes for any inconveniences caused by this problem. For further questions please call (440) 483-7600. |  | 424 | 424 units | Worldwide Distribution - Nationwide Distribution to the states of : AR, AZ, CA, CT, DE, FL, GA, IA, IN, KS, LA, MA, MD, MI, MN, MO, MS, MT, NH, NJ, NM, NV, NY, OH, OR, PA, SC, TN, TX, UT, VT, WV and WY., and to the countries of : Argentina, Australia, Austria, Belgium, Bolivia, Brazil, Canada, Cambodia, China, Colombia, Denmark, Finland, France, Germany, Hong Kong, India, Indonesia, Iran, Iraq, Ireland, Italy, Japan, Latvia, Libyan Arab Jumeirah, Lithuania, Malaysia, Mexico, Myanmar, Netherlands, New Zealand, Norway, Panama, Philippines, Russia, Singapore, Slovakia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Training, Tunisia, Turkey, United Arab Emirates,United Kingdom, Uzbekistan and Viet Nam. | Open, Classified |  |  | 424 | Computer
Z-1408-2015 | 70927 | No | OBJ | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | VasoNova Vascular Positioning System (VPS ) G4 Console |  |  | March 30, 2015 | 2015 | VasoNova, Inc. | A limited number of VPS G4 Consoles are being recalled due to noncompliance with International Electrotechnical Commission Standard 60101-1-1. | Device Design | Teleflex sent an Urgent Medical Device Recall Notification letters dated on March 30, 2015 to all affected customers. The letters identified the affected product, problem and actions to be taken. The risk to users is identified and users are asked to discontinue use and quarantine affected serial numbers. Teleflex sales reps will help with return of product. The Acknowledgeable form should be completed and returned as per instructions. Questions should be directed to 1.877.236.6869. |  | 17 | 17 | US Nationwide Distribution in the states of AZ, DC, PA, IN, PA. GA, NY, KS, CA, VA, NC. | Terminated | April 11, 2016 | 378 | 17 | Computer
Z-0231-2016 | 72220 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS 5,1 FS Chemistry System, Catalog Number/Product Code 6801375, Unique Device Identifier No. 10758750001132, and VITROS 5,1 FS Chemistry System Refurbished, Catalog Number/Product Code 6801890, Unique Device Identifier No. 10758750001644. Intended for use in the in vitro quantitative measurement of a variety of analytes of clinical interest. |  |  | September 10, 2015 | 2015 | Ortho-Clinical Diagnostics | Software Anomaly during ADD Installation on VITROS 5,1 FS Chemistry Systems using Software Versions 2.2.1 through 2.8. The user modified parameters would not be saved as expected, but remained changed to the default values obtained from the ADD disk. | Software design | The firm, Ortho Clinical Diagnostics, sent a "FOLLOW UP to URGENT PRODUCT CORRECTION NOTIFICATION" letter dated 9/10/2015 and response forms to customers via FedEx overnight courier, e-mail (ORTHO PLUS e-Communications) and/or via US Postal Service overnight letter on 9/10/2015. Foreign affiliates were informed by email on 9/10/2015 of the issue and instructed to notify their consignees of the issue and required actions.The letter described the product, problem, and actions to be taken. The customers were instructed to follow the required actions and complete and return the Confirmation of Receipt-Response Required form, no laterthan 9/18/15, by fax at 1.888.557.3759 or 1.585.453.4110, or scan PDF and email to: ConfirmationAdmin@its.jnj.com. Customers with any additional questions, can contact Customer Technical Services at 1-800-421-3311 at any time. |  | 210 | Catalog # 6801375: Domestic 919; Foreign 1251; Catalog 6801890: Domestic 63, Foreign 147 | Worldwide Distribution: US (nationwide) including Puerto Rico, and countries of: Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Republic of Panama, Singapore, Spain, and Venezuela. | Open, Classified |  |  | 210 | Computer
Z-0357-2016 | 72564 | No | HJO | AC-powered slitlamp biomicroscope. | Ophthalmic | Ophthalmic | 510(k) | IOL Master 5.5: Software versions 7.5 and 7.7; |  |  | October 27, 2015 | 2015 | Carl Zeiss Meditec AG | IOL Master software versions 7.5 and 7.7 calculation printouts and exported reports can contain the wrong IOL power data. | Software design | Urgent Medical Device Correction letters will be mailed on Friday, November 13, 2015 as the IOL SW update material is now available. Letter informs customers of the issue and provide instructions for identifying the software version and for installing the software update with illustrations. All customers are requested to complete and return the confirmation form to Stericycle as soon as the upgrade is complete or it has been determined that your unit is not running the affected version of the software. |  | 49 | 49 | Nationwide Distribution. | Open, Classified |  |  | 3862 | Computer
Z-1228-2015 | 70424 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | CareLink Pro |  |  | November 10, 2014 | 2014 | Medtronic MiniMed Inc. | The firm is informing customers of a software error that might result in an inaccurate display of the Temp Basal time and duration in CareLink Clinical reports with data uploaded from the MiniMed 620G or 640G insulin pump. | Software design | Medtronic MiniMed Inc. sent an Urgent Medical Device Correction letter in December 2014 to all affected customers. The firm states in their letter that patients can continue to use their MiniMed 620G and 640G insulin pump systems, and that there is no risk to patients caused by this issue. The firm instructs that when reviewing the CareLink Pro 4.0 or CareLink Clinicial reports from the MiniMed 620G and 640G insulin pumps not to use the Temp Basal graphic for clinical consideration.For questions regarding this recall call 818-567-4700. |  | 9 | 5 CareLink Pro 4.0 CD ROMs | Worldwide Distribution to Japan only.CareLink Clinical is for clinical trials only (US and International) | Terminated | March 10, 2015 | 120 | 9 | Computer
Z-0986-2015 | 70093 | No | LWH | Surgical sponge scale. | General Hospital | General &amp; Plastic Surgery | 510(K) Exempt | The SurgiCounter Charging Cradle |  |  | June 16, 2013 | 2013 | Stryker Instruments Div. of Stryker Corporation | Faulty SurgiCounter Charging Cradle can prevent the scanner from charging or downloading information after a procedure. | Device Design | On 6/16/2013, Improvements to the SurgiCounter" charging base notifications were emailed to the affected users with instructions for requesting a replacement product. The notification included a description of the reason for the recall, affected product, and instructions for responding to the notification. For questions, call the Support team any time at (877) 520-2300 option 1, or support@scmd.com. |  | 3625 | 3625 | US Nationwide Distribution in the states of AK, AL, AR, AZ, CA, CT, FL, GA, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, NC, ND, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, WV, and WY. | Terminated | August 20, 2015 | 795 | 3625 | Computer
Z-0971-2015 | 69957 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | SOMATOM Force; the intended use of computed tomography is to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data from either the same axial plane taken at different angels or spiral planes taken at different angles. |  |  | December 03, 2014 | 2014 | Siemens Medical Solutions USA, Inc | Artifacts are found in acquired imaging on the SOMATOM Force. | Software Manufacturing/Software Deployment | A customer advisory notice, dated December 3, 2014, was sent to end users that identified the product, problem, and action to be taken. Customers were informed that new software would be installed via remote access. |  | 7 | 7 | PA, MA, MN, SC, MD, and TX. | Terminated | April 06, 2015 | 124 | 7 | Computer
Z-0977-2015 | 70065 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | lnline Persuader- PedFuse Pedicle Screw System; Part Number: 11-80164 |  |  | December 19, 2014 | 2014 | SpineFrontier, Inc. | A crack may develop in the handle to the Inline Persuader, which may progress if the user continues to use the affected instrument. | Device Design | Spinefrontier sent an Urgent Voluntary Recall Notification letter dated December 16, 2015, to all affected customers. The notice informed customers of the recall and requests that they initiate the immediate return of the parts listed. In addition, SpineFrontier is requiring confirmation via recall acknowledgement form that consignees have received this notice and that they will initiated the immediate return of the parts. Customers can contact SpineFrontier at 978-232-3990 Monday through Friday 8:30AM to 5:30PM EST with any questions. |  | 4 | 4 devices | Nationwide Distribution including TX, OH, VA, and Washington, DC. | Terminated | January 11, 2016 | 388 | 4 | Computer
Z-0967-2015 | 69830 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS 3600 Immunodiagnostic System, Catalog Number 6802783, IVD --- Ortho Clinical Diagnostics. |  |  | November 19, 2014 | 2014 | Ortho-Clinical Diagnostics | Software Anomaly: the firm has identified an anomaly with VITROS System Software Version 3.1 and below, and determined that the software may not properly identify an expired calibration. | Software design | On 11/19/2014, IMPORTANT PRODUCT CORRECTION NOTIFICATION letter (Ref. CL2014-296, dated 11/19/2014) was sent via FedEx overnight courier or ORTHO PLUS e-Communications and/or US Postal Service Priority Mail (for PO Boxes only) to the US consignees (OCD Direct, Drop-Ship and US Federal Government) to inform them of the issue and advise them of the required actions. Immediately install the most recent ADD using the All Assay Data option. If any of your calibrations are expired, the system will post PZB-005 condition codes. Perform calibrations prior to using those reagents. As expected, calibrations that will expire within one hour will also be flagged. Continue to install the most recent ADD at least once each week using the All Assay Data load option so that the system properly flags expired calibrations. --- Foreign affiliates were informed by e-mail on 11/19/2014 of the issue and instructed to notify their consignees of the issue and required actions. |  | 658 | 658 total: USA - 144, Foreign - 514 | Worldwide Distribution-US (nationwide) including Puerto Rico and the countries of Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Singapore, Spain and Venezuela. | Open, Classified |  |  | 2827 | Computer
Z-0972-2015 | 70016 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Artis Q, and Artis Q.zen ceiling systems; model numbers: 10848281; 10848354: The Artis family is dedicated angiography systems developed for single and biplane diagnostic imaging and interventional procedures including, but not limited to, pediatric and obese patients. Procedures that can be performed include cardiac angiography, neuro angiography, general angiography, rotational angiography, multipurpose angiography and whole body radiographic/fluoroscopic procedures. |  |  | December 03, 2014 | 2014 | Siemens Medical Solutions USA, Inc | Suboptimal routing of the cable may result in increased wear over time. Without additional measures, damage of a cable may result in limited functionality and in rare cases, failure of a system. An ongoing procedure could be terminated in rare cases. | Device Design | Siemens sent an important customer safety information notice, dated December 3, 2014, to all affected customers. The letter identified the product, problem, and action to be taken. The following hardware modifications will be implemented with the corrective action AX042/14/S:1. A modified cable outlet to provide additional space for the cabling.2. A modified plastic cover at the cable outlet providing optimal cable guidance in the area of the arc. In general all systems will be checked for existing damage prior to the implementation of this action.For questions regarding this recall call 610-219-6300. |  | 50 | 50 | Nationwide Distribution | Terminated | April 06, 2015 | 124 | 50 | Computer
Z-0966-2015 | 70073 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Discovery IGS 730 &amp; Discovery IGS 740. |  |  | December 10, 2014 | 2014 | GE Healthcare | Potential unintentional system motion or rotation due to a positioning user interface issue with Discovery IGS 730 and 740 products. | Software design | Consignees were sent a 12/10/2014 a GE Healthcare "Urgent Medical Device Correction" letter GEHC Ref# 12223 dated December 10, 2014. The letter was addressed to Hospital Administrators / Risk Managers, Managers of Radiology / Cardiology and Radiologists / Cardiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information ( US call 800-437-1171, other countries contact your local GE Healthcare Service Representative). |  | 24 | 24 ( 7 US, 17 OUS) | Worldwide Distribution-Distributed in the states of CO, CT, MS, NJ, NY, OH, and WA, and the countries of BELGIUM, CANADA, CHINA, FRANCE, GERMANY, JAPAN, MEXICO, PERU, UNITED ARAB EMIRATES, and UNITED KINGDOM. | Terminated | June 01, 2015 | 173 | 24 | Computer
Z-0454-2016 | 72654 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Aegis Torque Handle CATALOG NO. : 2871-10-800. A spinal surgical instrument used on the AEGIS Anterior Lumbar Plate System. |  |  | November 20, 2015 | 2015 | DePuy Spine, Inc. | Manufacturing error within the Torque Handle may prohibit some handles from being able to fully assemble with the tightener shaft. The design of the inspection gage does not account for the additional shift length | Device Design | DePuy sent an Urgent Voluntary Product Recall Notification dated November 17, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. They were asked to do the following: 1. Review inventory located within your facility and quarantine any products listed in this recall notice. Please contact your DePuy Synthes Spine* Sales Consultant to return any inventory located within your facility. 2. Review, complete, sign and return the attached business reply form to in accordance with the directions on the business reply form within 5 business days of receipt of this notification.3. Forward this notice to anyone in your facility that needs to be informed.4. If any product listed below has been forwarded to another facility, contact that facility and provide them with this letter. 5. Keep a copy of this notice.For further questions please call (508) 828-3774 |  | 16 | 16 devices | US Distribution to the states of : NC, MI, CO, MD, FL, ID, TX, CA and NY. | Terminated | July 03, 2017 | 591 | 16 | Computer
Z-0983-2015 | 70070 | No | OVD | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Timberline Anchored Lateral Fixed Retractable Sleeve Model No. 8630-0205. Timberline Anchored Lateral Fixed Retractable Sleeves are used with the Timberline MPF System to help establish access into the vertebrae for the screws utilized to anchor the plate to the vertebrae. The Timberline Retractable Drill and the Timberline Retractable Awl fit within this Sleeve. |  |  | August 22, 2013 | 2013 | Biomet Spine, LLC | The inner diameter of the Sleeve shaft is undersized, resulting in interference fit when inserting the Anchored Lateral Straight Retractable Awl or the Anchored Lateral Straight Retractable Drill into the Sleeve. | Device Design | The firm, Lanx, sent a letter to their Consignees/customers dated October 15, 2013. The letter describes the product, problem and actions taken. Lanx and their distributors replaced the affected inventory with conforming product. Please contact Chief Compliance Officer at 303-501-8534 or by email: Randy.fesmire@lanx.com if you have any questions regarding this recall. |  | 14 | 14 units | US Distribution to states of: AZ, CA, NY, TN and TX including PR. | Terminated | March 13, 2015 | 568 | 14 | Computer
Z-0452-2016 | 72688 | No | HAB | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Zimmer UNIVERSAL Power System Loaner and Modular Electric/Battery Double Trigger Handpiece. Rx only Made in Switzerland |  |  | November 10, 2015 | 2015 | Zimmer, Inc. | The firm identified a malfunction which made it possible for a handpiece to start by itself when the power source is connected. | Device Design | Zimmer sent an Urgent Medical Device Correction letter dated November 9, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.1. Review this notification.2. You will receive copies of notification sent directly to hospital risk managers and surgeons in your territory. As necessary, review and facilitate understanding of this notification by those entities.3. Further identify hospital risk managers and/or surgeons in your territory who should also receive notification of this update, and supply the information to the entities you have identified.4. If after reviewing this notification you have further questions or concerns, please call 330-364-0989 between the hours of 8 a.m. and 5 p.m. EST, Monday through Friday.Zimmer instructed their direct customers that they will upgrade the device with the new electronic during their annual preventative maintenance activities. Until the new upgrade is installed, the firm instructed their customers to follow the Instructions for Use. |  | 87 | 87 | US Distribution to the states of : AR, CA, CO, CT, FL, GA, HI, IL, KS, KY, LA, ME, MI, MN, MO, MS, NC, NE, NJ, NV, NY, OH, PA, RI, SC, TX, UT, VA, WA, WI, WV and WY. | Terminated | February 19, 2016 | 101 | 176 | Computer
Z-0957-2015 | 69960 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | iPlan RT is a radiation treatment planning system that is intended for use in stereotactic, conformal, computer planned, Linac based radiation treatment of cranial, head and neck, and extracranial lesions. |  |  | November 19, 2014 | 2014 | Brainlab AG | iPlan RT Radiation Treatment Planning Software: Potentially incorrect patient positioning when using multiple localized CT image data sets. | Software in the Use Environment | Brainlab sent a FIELD SAFETY NOTICE / PRODUCT NOTIFICATION letter dated November 19, 2014, to all affected customers. The letter included instructions to: 1) If clinically not required, do not use multiple localized CT scans within one treatment plan (to avoid potentially incorrect Reference and Alignment Set assignments in the first place). 2) If you must use multiple localized CT scans, e.g. for recurrent treatment planning of the same patient, always guarantee that the latest (see note *) CT scan is defined as both Alignment Set and Reference Set during treatment planning. Brainlab will provide a software solution to prevent the described scenario from occurring. Brainlab will actively contact affected customers tentatively starting January 2016 to schedule the update. Brainlab will additionally refine the instructions for use of the iPlan RT treatment planning software in regard to implications of Alignment Set and Reference Set selections and will provide this user guide update to existing affected customers together with the software update.Customers with questions can contact Brainlab at 800-597-5911 or via e-mail at us.support@brainlab.com. |  | 1412 | 1412 systems total | Worldwide Distribution - USA including Alabama, Arizona, Arkansas, California, Colorado, Connecticut, Florida, Georgia, Idaho, Illinois, Indiana, Iowa, Kansas, Kentucky, Louisiana, Maine, Maryland, Massachusetts, Michigan, Minnesota, Mississippi, Missouri, Montana, Nebraska, Nevada, New Hampshire, New Jersey, New York, North Carolina, North Dakota, Ohio, Oklahoma, Oregon, Pennsylvania, Rhode Island, South Carolina, South Dakota, Tennessee, Texas, Utah, Vermont, Virginia, Washington, Wisconsin, Wyoming and Internationally to Algeria, Angola, Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, China, Colombia, Czech Republic, Denmark, Dominican Republic, Ecuador, Finland, France, Germany, Greece, Honduras, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Kazakhstan, Latvia, Lebanon, Lithuania, Malaysia, Mexico, Netherlands, Norway, Oman, Panama, Poland, Portugal, Qatar, Romania, Russian Federation, Saudi Arabia, Singapore, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Syrian Arab Republic, Taiwan, Thailand, Tunisia, Turkey, Ukraine, United Kingdom, Uruguay, and Venezuela. | Open, Classified |  |  | 1412 | Computer
Z-0451-2016 | 72683 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Mobile X-ray system MobileDaRt Evolution/FDR Go Software |  |  | November 10, 2015 | 2015 | Shimadzu Medical Systems | The firm is recalling the Shimadzu Mobile X-ray system because an image may not transfer to image server properly. | Software design | Shimadzu sent an Urgent Voluntary Medical Device Recall Notice dated November 15, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letter was sent to customers to inform them that Shimadzu is recalling the Mobile X-ray system MobileDaRt Evolution because of reports that an image was not transferred to image server properly. Customers with questions are instructed to contact their local Shimadzu representative or contact Shimadzu National Technical Support at (800) 228-1429. |  | 297 | 297 units | US Distribution and Internationally to Canada | Open, Classified |  |  | 297 | Computer
Z-1016-2015 | 69993 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Airo Mobile Intraoperative CT I Airo; Model #: MobiCT-32. Computed Tomography x-ray system. |  |  | December 08, 2014 | 2014 | Mobius Imaging, LLC | There is a risk that during the transfer of an image and navigation data to the Brainlab Curve Image Guided Surgery Navigation System after a CT scan, an error may occur, causing either no navigation data or incorrect navigation data being transferred to the curve. | Software design | Mobius sent an Urgent Field Safety Notice dated December 8, 2014, to all affected customers. The letter explains the recall and provides advise on action to be taken by the user. The letter also informs consignees that Mobius Imaging is working on a software fix for the software bug/communication error. The final fix is anticipated in early 2015.Advise on action to be taken by the user:The user should be prepared for the potential necessity to perform a manual navigation registration in order to use navigation.In general, do not use navigation data transmitted by the AIRO Mobile CT Scanner of any scan the Brainlab AIR App displayed the "Received Wrong Dataset" error for. To use such scans with the Curve, a manualnavigation registration is required.To avoid the possibility of incorrect data from being transferred to the Curve, Mobius Imaging is recommending that customers only select "NO" when they come to the "Continue with same patient" screen after each scan. They can select the same patient under the "Existing Patients" tab and restart the workflow. Additionally, always adequately verify the accuracy of the navigation registration data as required by the Brainlab Curve" Navigation system software and the AERO Operator's Manual. If you require further clarification, please feel free to contact your local Brain lab Customer SupportRepresentative.Customer Hotline: +49 89 99 15 68 44 or +1800 597 5911 (for US customers) or byE-mail: support@brainlab.com (for US customers: us.support@brainlab.com)Fax Brainlab AG: + 49 89 99 15 68 33Address: Brainlab AG (headquarters), Kapellenstrasse 12, 85622 Feldkirchen, Germany.For questions regarding this recall call 978-615-5025. |  | 315 | 21 devices distributed. 15 of these devices affected by the software v. 1.1.1+ patch | USA including AK, IL, WA, MA, NY, WV, NC, TX, CA, PA and Internationally to Germany, Netherlands, Belgium, Switzerland, Saudi Arabia, Egypt. | Terminated | July 06, 2015 | 210 | 315 | Computer
Z-0994-2015 | 70020 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Ysio Max, Luminos dRF Max and Agile Max systems with software version VE10 and SmartOrtho license. |  |  | December 15, 2014 | 2014 | Siemens Medical Solutions USA, Inc | Potential for composed images to be flipped before being sent to PACS on systems with software version VE10 and SmartOrtho license. Flipped images may be reversed so the associated annotations, e.g. labels (R/L), may be displayed incorrectly, potentially leading to misdiagnosis. | Software design | The firm, Siemens, sent a "Safety Advisory Notice - XP080/14/S" letter dated December 8, 2014, to end users that identified the product, problem, and actions to be taken. Siemens provided instructions to avoid the issue and indicated a software fix would be released as a permanent resolution. The customers were instructed observe this safety notice and comply with the corresponding measure until the update has been fully completed.If you have any questions, contact the Regulatory Technical Specialist at 610-219-2119. |  | 11 | 11 | US Distribution to states of: ID, IL, MI, MN, MO, ND and PA. | Open, Classified |  |  | 11 | Computer
Z-0993-2015 | 69560 | No | OUG | Medical device data system. | General Hospital | General Hospital | 510(K) Exempt | TEGRIS System manufactured by MAQUET GMBH in Germany |  |  | August 21, 2014 | 2014 | Maquet Medical Systems USA | The wrong movement results on the MAGNUS OR table. The button commands on the TEGRIS touchscreen for Lower Leg Up and Upper Leg Down are switched in the software for the integration with the MAGNUS operating table system. | Software design (manufacturing process) | Maquet sent an Urgent Medical Device Field Correction Letter dated August 21, 2014, to all affected customers via Federal Express. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to take the following measures until the issue is corrected: The TEGRIS equipped with a software module version up to 2.1.X.X should not be used in conjunction wiht a MAGNUS operating table system. All other MAQUET operating tables are not affected.A MAQUET service representative would contact them and arrange to have their TEGRIS software module updated. Customers with questions were instructed to contact MAQUET Technical Support at 1-888-627-8383 (press option 3, followed by option 1, then option 1 again.For questions regarding this recall call 973-709-7000. |  | 6 | 6 units | US Nationwide Distribution including NE, NY and PA. | Terminated | November 17, 2015 | 453 | 6 | Computer
Z-1004-2015 | 69727 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS Software Version 3.1 utilized on the following systems: 1) 4600 Chemistry Systems (VITROS 5,1 FS System family member, 2) 5600 Integrated System |  |  | September 29, 2014 | 2014 | Ortho-Clinical Diagnostics | Calibration may not occur when using calibrator barcode labels supplied with VITROS Chemisty Products Calibrator Kit 2. | Software Manufacturing/Software Deployment | Ortho Clinical Diagnostics sent an Important Product Correction Notification letter dated September 29, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed Calibration may not occur when using calibrator barcode labels supplied with VITROS Chemisty Products Calibrator Kit 2.The resolution to this software anomaly will be in a future version of software. Until the new software is installed on their VITROSÂż System, customers were instructed not to use the calibrator barcode labels for VITROSÂż Calibrator Kit 2. Customers were advised that you manually program the calibration.Customers with questions were instructed to contact Customer Technical Services at 1-800-421-3311. |  | 12 | 12 units | Worldwide Distribution - USA including CO, NC, NJ, NY, PA, MI, WV and Internationally to Canada, Japan and Spain. | Terminated | April 03, 2017 | 917 | 12 | Computer
Z-0447-2016 | 72725 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MOSAIQ Oncology Information System. |  |  | November 24, 2015 | 2015 | Elekta, Inc. | First Databank (FDB) is a third party product incorporated into MOSAIQ. FDB routinely modifies drugs in their database to reflect market and formulary changes. In FDB Framework v2.0, those drugs are not being included in drug-drug, drug-allergy and duplicate drug therapy interaction checking. | Software design | Elekta sent an Urgent Important Field Safety Notice 371-01-MSQ-004 to all affected customers on 11/24/2015. The notice informs users to the issue, the work around and the solution. The customers are instructed to complete and return the Important Field Safety Notice Acknowledgement form to their local Elekta Office or Representative, as soon as possible but no longer than 30 days from receipt. |  | 211 | 211 | Nationwide Distribution -- AL, AK, AZ, AR, CA, CO, CT, DE, DC, FL, GA, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NJ, NM, NY, NC, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, WY, &amp; Puerto Rico. | Open, Classified |  |  | 211 | Computer
Z-0434-2016 | 72773 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | UniCel DxH 600 Coulter Cellular Analysis System, Software Version 1.1.1.0, Part No. B23858 |  |  | December 01, 2015 | 2015 | Beckman Coulter Inc. | Beckman Coulter is recalling the DxH 800 Slidemaker Stainer Software version 3.0.2.0 and DxH 600 software version 1.1.1.0 because the software for the DxH systems allows the creation of multiple orders with the same Specimen identification (ID) but different Patient identification when manually editing pendingorders at the System Manager. | Software design | Beckman Coulter sent an Urgent Medical Device Recall letter dated December 1, 2015, to all affected customers to inform them that Beckman Coulter is recalling the the DxH 800 Slidemaker Stainer Software version 3.0.2.0 and DxH 600 software version 1.1.1.0 because of the possibility of sample misidentificaton. Customers were informed that during internal testing, Beckman Coulter determined that the software for the DxH systems noted above allows the creation of multiple orders with the same Specimen identification (ID) but different Patient identification when manually editing pending orders at the System Manager. The issue does not occur for edits of released results or for test orders requested through host transmission. Customers were instructed to complete and return the enclosed response form within 10 days. Customers with questions were instructed to contact Customer Support Center at www.beckmancoulter.com/customersupport/support or via phone (800) 526-7694 in US and Canada. Customers outside of US and Canada are instructed to contact their local Beckman Coulter Representative.For questions regarding this recall call 714-961-5321. |  | 666 | 666 units total (327 units in US) | Worldwide Distribution - US (nationwide) and Internationally to Australia, Austria, Belarus, Belgium, Bosnia and Herzegovina, Brunei Darussalam, Canada, China, Colombia, Croatia, Czech Republic, Ecuador, France, Germany, Gibraltar, Greece, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Jordan, Republic of Korea, Kuwait, Lebanon, Lithuania, Macao, Malaysia, Mayotte, Mexico, Morocco, Myanmar, Netherlands, New Zealand, Norway, Oman, Panama Philippines, Poland, Portugal, Puerto Rico , Romania, Russian Federation, Saudi Arabia, Singapore, Slovakia, Slovenia, Spain, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Turkey, United Arab Republic, United Kingdom, Uruguay, and Vietnam. | Open, Classified |  |  | 666 | Computer
Z-1011-2015 | 69792 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | BrightView model number: 882478 |  |  | November 12, 2014 | 2014 | Philips Medical Systems (Cleveland) Inc | Unintended detector and gantry movement due to software issues. | Software design | Philips Healthcare sent an Urgent - Field Safety Notice dated November 4, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Philips is advising the customers/users to retain a copy of this Field SafetyBY CUSTOMER I USER Notice with the equipment Instruction for Use, until further notice.Problem #1:Philips is advising the customer/user to perform the following workaround prior tothe implementation of the field safety correction, to avoid contact between thedetector and the head holder as describe above in Problem #1:ONLY if the system is positioned in a Relative-90 configuration with one of thedetectors below the pallet level and the table in in the patient load position:Follow the sequence of operations prior to performing a study utilizing the headholder:- Deploy or store the x-ray panel if necessary;- Run the patient load PPM;- Install the head holder; and- Position the patient on the table and proceed with the studyIn addition, take the following preventative actions:" Monitor the patient during any and all system motions." Know locations of the multiple E-Stop button(s) to halt all system motion.(Emergency Stop: where the system immediately stops all detector andgantry motions by disconnecting power to the motors that control systemmotion) Customers with questions were instructed to contact the Customer Care Solutions Certer at 1-800-722-9377.For questions regarding this recall call 440-483-2119. |  | 1064 | 1064 units | Worldwide Distribution - USA (nationwide) and Internationally to Canada, Algeria, Argentina, Australia, Austria, Bahrain, Belgium, Brazil, China, Czech Republic, Denmark, Dominican Republic, Egypt, Finland, France, Gabon, Germany, Greece, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Lebanon, Lithuania, Malaysia, Mexico, Netherlands, New Zealand, Norway, Oman, Palestine, Panama, Philippines, Poland, Portugal, Russia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Syrian Arab Republic, Taiwan, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom, United Kingdom, and Viet Nam. | Open, Classified |  |  | 1064 | Computer
Z-0426-2016 | 72643 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | MYLA MASTER DVD 3.2 CLI DL380 |  |  | November 03, 2015 | 2015 | bioMerieux, Inc. | MYLAÂż software connected to a Laboratory Information System (LIS) and a VITEKÂż 2 system has the potential to link a test result to an incorrect patient record with the same specimen ID. | Software design | bioMerieux sent an Urgent Product Correction Notice letter dated November10, 2015 via FedEx to affected customers. The letter identified the affected product, problem and actions to be taken. For questions contact your local bioMÂżrieux Clinical Customer Service organization at (800) 682-2666. |  | 626 | 626 systems in total | Worldwide Distribution - US Nationwide in the states of: AK, AL, AR, AZ, CA, CT, FL, GA, HI, IA, ID, IL, IN, KS, LA, MA, MD, MI, MN, MO, MS, NC, ND, NE, NJ, NM, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV and countries of: Foreign accounts: Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, China, Colombia, Czech Republic, France, Germany, Greece, Hungary, India, Italy, Japan, Korea, Netherlands, Poland, Portugal, Russia, Serbia, Singapore, South Africa, Spain, Sweden, Switzerland, Turkey, United Kingdom, Angola, Puerto Rico, Republic of Belarus, Republic of Belarus, Benin, Bosnia, Brunei, Bulgaria, Burkina Faso, Cambodia, Cameroon, Chad, Democratic Republic of Congo, Costa Rica, Dominican Republic, Ecuador, Egypt, El Salvador, Estonia, Ethiopia, Gabon, Guam, Guatemala, Guatemala, Guinea, Honduras, Indonesia, Iraq, Israel, Ivory Coast, Kazakhstan, Kenya, Kenya, Kuwait, Laos, Latvia, Lebanon, Libya, Lithuania, Malaysia, Malaysia, Mali, Portugal, Myanmar, Managua, Niger, Nigeria, Palestine, Philippines, Philippines, Qatar, Kingdom of Saudi Arabia, Senegal, Serbia, Slovakia, Togo, Tunisia, Ukraine, UAE, Vietnam, Vietnam, Ivory Coast and Zimbabwe. | Open, Classified |  |  | 626 | Computer
Z-1013-2015 | 70200 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Syngo RT Therapist: The intended use of the SIEMENS branded ARTISTE, ONCOR and PRIMUS family of linear accelerator systems is to deliver X-Ray photon and electron radiation for the therapeutic treatment of cancer |  |  | January 02, 2015 | 2015 | Siemens Medical Solutions USA, Inc | Combination of CTVision with syngo RT Therapist / syngo RT Oncologist 4.3.SP1 automatic registration in Adaptive Targeting might result in wrong offset calculations. Applying this offset can lead to patient mistreatment. Cone Beam imaging is not affected by this problem. | Software Design Change | A customer advisory notice, dated January 2, 2015, was sent to end users which identified the product, problem, and action to be taken. It was recommended that customers use only manual registration with the 6 degrees of freedom option disabled, which provides the correct offset values. A software update will be deployed once a solution is available. |  | 3 | 3 | UT, WI, NY | Terminated | September 10, 2015 | 251 | 3 | Computer
Z-0418-2016 | 72258 | No | FDT | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fujifilm Medical Systems Duodenoscope, Model Numbers ED-200XU, ED-200XT, ED-250XT, ED-310XU, ED-420XL, ED-410XT, ED-410XU, ED-450XL, and ED-450XT. Medical endoscope for the duodenum and upper G.I. tract. Open elevator wire channel duodenoscopes. |  |  | August 31, 2015 | 2015 | Fujifilm Medical Systems U.S.A., Inc. | A precautionary measure because the Duodenoscopes have been redesigned to incorporate a closed elevator wire channel that facilitates reprocessing. | Device Design | FujiFilm mailed a Medical Device Recall Letter and Tracking/Verification Form dated August 31, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to check their inventory for affected product, if found, affected product should be removed from circulation and returned to Fugifilm. Customers were also asked to complete and return the enclosed tracking/verification form to Fujifilm Medical Systems USA, Inc., Endoscopy Division (FMSU-ESD), 10 High Point Drive, Wayne, NJ 07470 or via fax to 973-633-8818. Customers with questions were instructed to contact their Fujifilm sales representative. For questions regarding this recall call 973-686-2482 |  | 79 | 79 units (US) and 985 units (Internationally) | US including CA, CO, FL, IA, IL, KS, LA, MD, MN, MO, NC, TN and WA and Internationally to Asian Pacific, EU, Latin America, and Japan. | Open, Classified |  |  | 79 | Computer
Z-0462-2016 | 72721 | Yes | OAD |  |  | Cardiovascular | PMA | IntellaTip MiFi XP Asymmetric (N4) Curve, |  |  | November 20, 2015 | 2015 | Boston Scientific Corp | During routine monitoring of complaints, BSC noted that the IntellaTip MiFi XP N4 complaint rate for distal tip bends/kinks was above the anticipated rate in BSC's product risk documentation. | Device Design | Boston Scientific sent an Urgent Medical Device Removal letters to all affected customers dated November 20, 2015 via Fed Ex or other equivalent service. The letter identified the affected products, problem and actions to be taken. Customers were instructed to contact their local Sales Representative to answer any questions. |  | 1556 | 1556 - all models | Worldwide Distribution - US Nationwide in the state of: PA, FL, IN, MI, TN, AL, IL, KY, MI, TX, OH, MO, MA, MN, ID, AL, TN, VA, CA and internationally to Austria, Australia, Belgium, Switzerland, Germany, Great Britain, Greece, Hong Kong, Israel, Iran, Italy, Japan, Kuwait, Latvia, Malaysia, Netherlands, Saudi Arabia, Singapore, Thailand, and South Africa. | Open, Classified |  |  | 1556 | Computer
Z-0915-2015 | 69931 | Yes | NPT |  |  | Cardiovascular | PMA | Lotus TAVR 23mm, |  |  | November 19, 2014 | 2014 | Boston Scientific Corporation | Lotus valve became unlocked during release from the delivery system. This may lead to percutaneous insertion of another valve, resulting in a Valve in Valve (ViV) implantation. If this occurs, the first valve could embolize and it may be necessary to convert the patient to surgery. | Device Design | Boston Scientific notified consignees on November 19, 2014, of the problem via letter through overnight delivery directing consignees to segregate and return affected product to Boston Scientific utilizing the Field Safety Corrective Action (FSCA) Instructions provided. The FSCA details that consignees should: Immediately discontinue use of and segregate product; Complete and return the Account Reply Verification Tracking form and Package/Ship the affected product. |  | 278 | 278 units, all sizes | Distributed only in the countries of Finland, France, Germany, Great Britain, Italy, Norway, Spain, Sweden, and Switzerland. | Terminated | March 26, 2015 | 127 | 278 | Computer
Z-0963-2015 | 69747 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | HeartStart MRx Monitor/Defibrillator; Model. Numbers: M3535A, M3536A, M3536J, M3536M, M3536MC, M3536M2, M3536M4, M3536M5, M3536M6 |  |  | November 19, 2014 | 2014 | Philips Medical Systems, Inc. | Device may malfunction, which could cause therapy to be delayed, disabled, or delivered inadvertently. | Software design | On November 19, 2014, Philips notified consignees of the recall via letter. The Urgent Medical Device Correction notification/Field Safety Notice informs customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue.A software upgrade will be provided free of charge for all units affected by one or more of these issues. A Philips Healthcare representative will contact customers with affected devices to arrange for installation of the upgrade. Philips is asking customers to follow the Action to be Taken by Customer/User section of the Urgent Medical Device Correction notification/Field Safety Notice |  | 74499 | 74,499 devices | Worldwide Distribution - US Nationwide | Open, Classified |  |  | 74499 | Computer
Z-0947-2015 | 69872 | No | DHB | Radioallergosorbent (RAST) immunological test system. | Immunology | Immunology | 510(k) | Phadia 1000 Instrument (introduced as UniCAP 1000), Article number: 12-3800-01 (All instrument software versions since launch in 2003); Multiple ImmunoCAP assays for Allergen Testing, FDA Cleared for use on Phadia 1000 instrument with IU statement-see attachments. Software utilized in multiple ImmunoCAP assays for Allergen Testing, FDA Cleared for use on Phadia 1000 instrument with IU statement-see attachments. |  |  | November 20, 2014 | 2014 | Phadia US Inc | During an investigation of instrument logs it was determined that In specific circumstances involving multiple steps, a rack sequencing error may occur. This will result in a mismatch between the sample ID and the test result reported for all subsequent sample racks in that run. | Software design | ThermoFisher sent an Urgent Medical Device Product Correction letter to all affected customers. The firm will inform customers of the specific sequence of errors and operator actions that can trigger the event via letter and phone, and will request that they do not respond to the Barcode Read Error and let the sample with the error eject normally from the instrument and be rerun after the barcode read error has been resolved. The firm will issue a mandatory Phadia 1000 Instrument Software (ISW) update that will correct the issue and make efforts to install it within 45 days from date of availability. Customers were asked to return the written response form via fax to: 1-888-243-5214Attention: Customer Support, Phadia 1000 ISW Product Correction Response; EMAIL: Compliance-us.idd@thermofisher.comEmail Subject Line: Phadia 1000 ISW Product Correction Response.Customers with questions were instructed to contact Phadia US Technical Support at 800-346-4364, option 2.For questions regarding this recall call269-492-1957. |  | 146 | 146 | Nationwide Distribution including AZ, CA, CO, CT, FL, GA, HI IL, IN, KS, MD, MI, MN, NC, NJ, NV, NY, OH, OK, OR, PA, TN, TX, UT, VA, and WA. | Terminated | June 30, 2015 | 222 | 146 | Computer
Z-0564-2016 | 72658 | No | NDP | Endosseous dental implant accessories. | Dental | Dental | 510(K) Exempt | Radiographic Transparency for Certain and External Connection Tapered Implants. Images that assist with measurements for dental implants. |  |  | November 16, 2015 | 2015 | Biomet 3i, LLC | Images contained within radiographic transparency ART 1036 are undersized. | Software design (manufacturing process) | BIOMET3i sent an" Urgent Medical Device Recall letter " dated November, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Responsibilities:1. Please review this notice and check your inventory for the affected item by comparing any on-hand ART 1036 Transparencies to a standard 8.5" x 11" sheet of paper. The non-conforming ART 1036 is approximately 7.8" x 10.2".2. Immediately quarantine and remove all affected items from service.3. Complete the attached Business Reply Form and either: a. fax it to 561-514-6316 orb. email it to postmarket@biomet.com orc. return it along with the affected product using the included shipping label4. Return affected product to BIOMET 3i.For assistance or other questions that you may have relative to this notice, please contact BIOMET 3i at 1- 800-342-5454 or 1-561-776-6700. Team members are available to assist you 8:00am to 6:00pm (Eastern), Monday through Friday. |  | 686 | 686 | Worldwide Distribution - US Distribution including Puerto Rico., and to the states of : CA, CO, FL, IL, MA, MI, MO, NY, OH, SC, TX, VA, WI, WV and Hawaii., and to the countries of : Australia, Germany, France, Greece, Northern Mariana Islands, Taiwan and Uruguay. | Open, Classified |  |  | 686 | Computer
Z-0910-2015 | 70144 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | McKesson Cardiology ECG Management |  |  | December 19, 2014 | 2014 | McKesson Israel Ltd. | An ECG procedure time stamp is incorrect, due to a software error in the McKesson Cardiology ECG Management with software versions 13.1 and 13.1.1. | Software design | McKesson sent an Urgent Field Safety Notice dated December 16, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed how to identify the potentially affected procedures and update the procedure date (if relevant) until a software update is made advailable for installation by McKesson. McKesson will proactively identify prior procedures affected by this issue and will notify customers of any issues identified. In addition to the Field Safety Notice, McKesson Israel will provide correction to the affected SW versions and will contact customers for deployment of a software patch that prevents this scenario from occurring. Customers with questions should contact McKesson Support at 1-866-777-0202.For questions regarding this recall call 404-338-3556. |  | 9 | 9 | Worldwide Distribution - USA including MA, MS, NH, NC, TX, WA and Internationally to the United Kingdom. | Terminated | February 25, 2016 | 433 | 9 | Computer
Z-0909-2015 | 69785 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Cypher MIS Screw System-The system includes screws, various types and sizes of rods, and set screws. Various instruments are also available for use by the surgeon to facilitate implantation of the device. |  |  | October 08, 2014 | 2014 | Ebi, Llc | Inoperative screw head seat splay and/or fracture as a result of positioning of the rod which is a part of the Cypher MIS Screw System. | Device Design | The firm, EBI LLC/Biomet, sent an "URGENT PRODUCT INFORMATION" letter dated 10/6/2014 via Fed Ex to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to follow the steps outlined in the letter. If you have any questions, please feel free to contact Director of Regulatory Affairs at 303-501-8548. |  | 2816 | 2816 approx.units ( 2688 units US and 128 OUS) | Worldwide Distribution: US Distribution to states of: AL, CA, CT, IN, LA, NC, NJ TX &amp; WI; and the country of: The Netherlands. | Terminated | December 21, 2015 | 439 | 2816 | Computer
Z-0558-2016 | 72779 | No | LTT | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | MicroScan Pos Combo Panel Type 21, Part No. B1017-201 |  |  | December 01, 2015 | 2015 | Beckman Coulter Inc. | Beckman Coulter has confirmed falsely negative results for Voges-Proskauer (VP) (pale-pink/brown/colorless) with Quality Control (QC) American Type Culture Collection (ATCC) organism Staphylococcus aureus ATCC 29213 in a portion of affected lots of MicroScan Pos Combo and MicroScan Pos Breakpoint Combo panels. The expected result is positive. | Device Design | The firm, Beckman Coulter, sent an "URGENT MEDICAL DEVICE RECALL" letter dated December 1, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The customers were instructed to discontinue use and discard any remaining inventory of lots listed; return the attached form for product replacement; retain your inventory of the other lot numbers of the panel types, as they are not impacted by this issue; discuss content of the letter with your laboratory director/staff, retain the notification as part of your laboratory QS documentation; if you have forwarded any of the affected products to another laboratory provide them a copy of the letter, and complete and return the enclosed Response Form within 10 days via mail to: Beckman Coulter, Inc., 1584 Enterprise Blvd., West Sacramento, CA 95691, Attn: Quality Systems &amp; Compliance or Fax: 916-374-2119 or email: MicrobiologyCustomersHCUS@beckman.com.Customers with questions are instructed to contact their Customer Support Representative at http://www.beckmancoulter.com or by phone at (800) 677-7226 in US and Canada. |  | 284 | 284 units total (96 units in US) | Worldwide Distribution - USA (nationwide) Distribution and to the countries of : Austria, Belgium, Brazil, Canada, Germany, Italy, Japan, Kazakhstan, Kuwait, Mexico, Myanmar, Netherlands, Norway, Panama, Poland, Portugal, Romania, Russian Federation, Spain, Taiwan, Thailand and United Kingdom. | Open, Classified |  |  | 284 | Computer
Z-0557-2016 | 72801 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MOSAIQ Oncology Information System |  |  | December 09, 2015 | 2015 | Elekta, Inc. | Incorrect drug dosage due to "Age Limit" and patient weight data item issue. | Software design | The firm, Elekta, sent an "URGENT IMPORTANT FIELD SAFETY NOTICE 371-01-MSQ-005" to all customers on 12/9/2015. The notice informs users of the specific product and version numbers affected by both issues, also includes a workaround for both issues. A copy of the Important Field Safety Notice is to be kept with the most current labeling and all personnel working with the product should be made aware of the content of the letter. The customers are instructed to return the acknowledgement form to their local Elekta as soon as possible-Fax: 408-830-8003.If you have any queries about this Notice, please contact your local Elekta office, call 408-830-8000 or toll free: 855-693-5358 or www.elekta.com. |  | 399 | 399 | Worldwide Distribution: US (nationwide) including states of: AK, AZ, AR, CA, CO, CT, DE DC, FL, GA, ID, IL, IN, IA, KS, KY, LA, ME, ,MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, SC, SD, TN, TX, VA, WA, WI, and WY; and countries of: Antiqua &amp; Barbuda, Australia, Austria, Bahama, Bahrain, Canada, Czech Republic, Germany, Hungary, Saudi Arabia, Sweden, Turkey, &amp; United Kingdom. | Open, Classified |  |  | 399 | Computer
Z-0908-2015 | 70046 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | I.M. Hole Locator Instrument. Orthopedic manual surgical instrument. |  |  | November 12, 2014 | 2014 | DePuy Orthopaedics, Inc. | Use of excessive force when impacting may lead to intra-operative femoral fracture. | Device Design | The firm, DePuy Synthes, sent an "URGENT INFORMATION - DEVICE CORRECTION NOTICE HP LCS IM Hole Locator" notification dated November 12, 2014 to the affected distributors. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. The distributors were instructed to: review the instructions and the information in this notice; provide this notice to your sales consultants, and complete and return the DISTRIBUTOR RESPONSE CARD: Verify Sales Consultant Notifications via fax to (574) 371-4939 or email: kearle@its.jnj.com. For questions, please contact Recall Coordinator, at 574-371-4917 (M-F; 8 a.m. - 5 p.m. EST). |  | 1130 | 1130 | Worldwide Distribution: US (nationwide) including states of: AZ, CA, CO, FL, GA, HI, IA, IL, IN, KY, LA, MA, MD, ME, MI, MN, MT, NC, NV, NY, OH, PA, TN, TX, VA, WA, and WI; and countries of: Canada, Australia, Austria, Belgium, Czech Republic, France, Germany, India, Israel, Italy, Japan, Malaysia, Russia, South Africa, Switzerland, and United Kingdom. | Terminated | February 24, 2015 | 104 | 1130 | Computer
Z-0905-2015 | 70043 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Femoral Impactor ATTUNE INTUITION Impactor. |  |  | November 10, 2014 | 2014 | DePuy Orthopaedics, Inc. | ATTUNE INTUITION Impactors (cat. nos. 254401003, 254401004, and 254401006) have broken and produced small pieces. Should a fracture occur and not be observed during surgery, there is the potential for these small fractured pieces of the instrument to be left in the patient. | Device Design | DePuy Synthes sent an Urgent Information letter dated November 10, 2014, to all affected customers. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Customers were instructed to complete the Hospital/User Facility Acknowledgement Car and return to their Deputy Orthopaedics Sales Consultant or fax to 574-371-4939. Customer swere also asked to retain a copy of the complete Acknnowlegement Card in their files along with the notice. Customers with questions were instructed to call 574-371-4917 (M-F; 8 a.m.  5 p.m. EDT).For questions regarding this recall cal 574-371-4917. |  | 4661 | 4661 | Worldwide Distribution: US (nationwide) including P.R. except AK and RI; and Internationally to: Chile, Canada, Australia, Austria, Belgium, China, France, Ireland, India, Israel, Italy, Japan, Malaysia, Netherlands, New Zealand, Singapore, Slovakia, South Africa, South Korea, Spain, Switzerland, Thailand, United Arab Emirates (UAE), and United Kingdom (UK). | Terminated | December 08, 2015 | 393 | 13077 | Computer
Z-0550-2016 | 72838 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo Imaging XS is a Picture Archiving and Communication System (PACS) |  |  | November 17, 2015 | 2015 | Siemens Medical Solutions USA, Inc | For the Syngo Workflow SLR System with Software Ver: VB10C: Printouts may be printed in incorrect anatomical size when using syngo Imaging XS filming application in conjunction with a printer not released for anatomical print usage. This may happen due to an invalid combination of printer and the syngo Imaging XS printer configuration file. | Software in the Use Environment | 11/17/2015, CSAN sent to customers with a list of suitable printers for the anatomical printing. |  | 42 | 42 | AL; CA; CO; CT; FL; GA; IA; IL; IN; KS; MD; MI; MO; NC; NE; NJ; NV; NY; OH; PA; TN; TX; WI; | Open, Classified |  |  | 42 | Computer
Z-0913-2015 | 69114 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | The Symbia Intevo Excel is a non-diagnostic SPECT ICT system with CT support for only attenuation correction and anatomical localization. |  |  | August 12, 2014 | 2014 | Siemens Medical Solutions USA, Inc. | Siemens Medical Solutions USA, Inc. has received reports of unintended radial motion of the detectors. This unintended motion, should it occur, may cause the patient compression related injuries. | Device Design | Siemens sent Customer Safety Advisory Notice letters to all customers of record. The letters instructed customers to: 1) continue to use their system(s); 2) place the safety advisory in the system's instructions for use; and, 3) to anticipate that their local service engineer will contact them to perform an inspection and schedule any necessary repairs by December 2014. Customers were also instructed to share the letter with their support staff and, if the equipment was sold or transferred, to forward the safety notice to the new owner of the equipment. Customers with questions can contact their local service representative or Siemens Medical Solutions at 800-888-7436. |  | 6 | 6 systems | Worldwide Distribution - USA (naitonwide) and Internationally to Algeria, Argentina, Australia, Austria, Bangladesh, Belgium, Brazil, Bulgaria, Canada, Chile, Colombia, Costa Rica, Croatia, Czech Republic, Denmark, Ecuador, Egypt, Estonia, Finland, France, Germany, Greece, Hong Kong, Hungary, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kuwait, Latvia, Lithuania, Malaysia, Mexico, Morocco, Myanmar, Netherlands, New Zealand, Norway, Oman, P.R. China, Pakistan, Peru, Philippines, Poland, Portugal, Qatar, Republic Korea, Romania, Russian Fed., Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Venezuela, and Vietnam. | Open, Classified |  |  | 2213 | Computer
Z-0940-2015 | 69902 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | IS3000 Vision Side Cart (VSC) used in conjunction with the da Vinci Si Surgical System IS3000. |  |  | December 05, 2014 | 2014 | Intuitive Surgical, Inc. | Intuitive Surgical is initiating a voluntary correction related to da Vinci Si System (IS3000) Vision Side Carts, with threaded-in caster wheels, which can result in the caster wheel bolts gradually unthreading and the wheels loosening. | Device Design | The firm, Intuitive Surgical, sent an "Urgent Medical Device Correction" letter dated December 5, 2014 to all customers on December 5, 2014 by trackable method. The letter described the product, problem and actions to be taken. The customers were instructed to: ensure all affected personnel are fully informed of this notice; identify affected product, and retain a copy of this notice with your da Vinci Si user manual. Please note: the da Vinci Si Vision Side Cart can continue to be used safely. Ensure that all wheels are fully unlocked before attempting to move the Vision Side Cart. Avoid rolling the Vision Side Cart over uneven surfaces (i.e., over cables or a threshold).Intuitive Surgical is working on a newly designed wheel assembly. Uponavailability, Field Service will retrofit the wheel assembly of the affected VisionSide Carts with the new design.If you need further information or support concerning this Medical Device Correction, please contact your Clinical Sales Representative or Intuitive Surgical Customer Service at the numbers listed below:ÂżÂż North America and South America: 8008761310 Option 3 (6 a.m. to 5 p.m. PST)ÂżÂż Japan: 0120565635 or 0355751362 (9 a.m. to 6 p.m. JST)ÂżÂż South Korea: 0232713200 (9 a.m. to 6 p.m. KST)ÂżÂż Europe, Middle East, Asia and Africa: +800 0821 2020 or +41 21 821 2020 (8 a.m. to 6 p.m.) or ics@intusurg.com |  | 202 | 202 | Worldwide Distribution: US (nationwide) and countries of: Belgium, France, Germany, Romania, India, Israel, Italy, South Korea, Norway, Slovak Republic, Spain, Sweden, Turkey, and United Kingdom. | Terminated | March 03, 2015 | 88 | 202 | Computer
Z-0933-2015 | 69961 | No | GFD | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | CelluTome Epidermal Harvesting System Harvester (5.0 CM) STERILE EO Rx Only Manufactured for: KCI USA, Inc. Part Number CT-H50 or CT-H25 |  |  | November 19, 2014 | 2014 | KCI USA, INC. | KCI has received reports that, in a small number of cases, the CelluTome Harvester (a component of the CelluTome Epidermal Harvesting System), blades have the potential to drift during shipping and could prematurely protrude into openings of harvester top plate. | Device Design | KCI sent an Urgent - Voluntary Medical Device Field Notification letter to all affected consignees on December 8, 2014. The letter notified the customer of the issue, and provided instructions to inspect their harvester prior to use. If blade drift is identified, instructionswere provided to the customer to either return the device to KCI, or if they wish, how to reset the blade prior to use. The letter instructs the customer to respond to the recall notifcation by returning the product reconciliation form attached the letter. Follow-up notifcations will be send to non-responding consignees.Consignees with questions were instructed to contact their local KCI Representative or Customer Technical Service Center at 1-800-275-4524 and select option 3.For questions regarding this recall call 210-515-4108. |  | 5695 | 1,139/5 Packs | Worldwide Distribution - USA (Nationwide) and Internationally to Australia, Austria, Canada, France, Germany, Ireland, Italy, Isreal, Netherlands, South Africa, Spain, Sweden, Switerland, Turkey, UAE, and United Kingdom. | Terminated | September 18, 2015 | 303 | 5695 | Computer
Z-0545-2016 | 72684 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance iCT Computed Tomography X-ray system |  |  | October 29, 2015 | 2015 | Philips Medical Systems (Cleveland) Inc | Perfusion scan feature may not be available on machines running software versions 4.1.2, 4.1.3, and 4.1.4. Customers previously using the basic axial perfusion software (versions preceding 4.1.2) without buying the helical perfusion license key could not access the basic axial perfusion feature after upgrading to software version 4.1.2. | Software design | On 10/30/2015 the firm sent Customer Information letters to their Consignees. |  | 51 | 51 Units | Worldwide distribution. US nationwide, Australia, Austria, Bangladesh, Belgium, Brazil, Cambodia, Canada, China, Egypt, France, Germany, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Korea, Lebanon, Libya, Lithuania, Malaysia, Martinique, Myanmar, Mexico, Mozambique, Netherlands, New Zealand, Norway, Palestine, Panama, Philippines, Poland, Portugal, Russia, Saudi Arabia, Singapore, Slovakia, Spain, Sweden, Switzerland, Tanzania, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, Uzbekistan &amp; Vietnam. | Open, Classified |  |  | 436 | Computer
Z-0481-2016 | 72742 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Artis systems with large display from Siemens Medical Solutions USA, Inc. Interventional Fluoroscopic X-ray for angiographic procedures. |  |  | October 30, 2015 | 2015 | Siemens Medical Solutions USA, Inc | Due to improper soldering in a limited number of large display bypass module components of a specific production lot, a loss of video signal could potentially occur. This has not been observed in the field and only sporadic cases have been observed in factory screening. | Software design | Siemens sent an Important Customer Safety Notice dated October 29, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The issue will be remedied with a hardware update AX025/15/S. The affected bypass module hardware will be exchanged with unaffected production lots to eliminate the problem. For further questions call: (610) 219-6300. |  | 17 | 17 | US Distribution to the states of : IL, WI, MA, FL, AL, TN, ND, OH,TX, CA, NJ and MI. | Terminated | March 28, 2017 | 515 | 17 | Computer
Z-0931-2015 | 69753 | No | HTC | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Synthes Small Electric Drive (SED), part number: 05.001.175: The Small Electric Drive is designed for use in general traumatology, especially hand and food applications involving surgical procedures such as drilling, burring, reaming, pin and wire placement, cutting of bone and hard tissue. |  |  | November 06, 2014 | 2014 | Synthes (USA) Products LLC | The device may operate solely in reverse mode, not operate in reverse mode when intended, not function, or unintentionally start. Unintended start of the hand piece or wrong mode/direction could cause bone/soft tissue, peripheral nerve, and /or vascular damage. Injuries may need surgical or medical intervention. Surgical delay and postoperative local infection are also potential risks. | Device Design | DePuy Synthes sent an Urgent Notice: Medical Device Removal letter dated November 6, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to notify their consignees using the hospital personnel letter, dated November 6, 2014, as their template for their consignee notification letter, review their inventory and remove the affected lots from stock.Call DePuy Synthes at 1-800-479-6329 to obtain a Return Authorization (RA) Number. Complete the Verification Section (page 4 of this letter) by checking the appropriate box indicating affected product has been located. Also, please indicate the number of devices found and note the Return Authorization Number. Please include your name, title, address, telephone number and signature in the spaces provided. Return the Verification Section (page 4 of this letter) with the product to: Credit/Returns, DePuy Synthes, 1101 Synthes Avenue, Monument, CO 80132. Send a copy of the completed Verification Section to DePuy Synthes by: Fax: (610) 430-7083 or Scan/email: Fieldaction@synthes.comIf you DO NOT have the identified product, please take the following steps: Complete the attached Verification Section (page 4 of this letter) by checking the appropriate box indicating that no affected product has been located. Please include your name, title, address, telephone number and signature in the spaces provided. This return documentation acknowledges your receipt of medical device removal information. Return the documents to DePuy Synthes by: Fax: (610) 430-7083 or Scan/email: Fieldaction@synthes.com Note: If the Verification Section is answered on behalf of more than one facility and/or individual, please clearly indicate the name and address of the facility and/or individual on page 4 of the notification.Customers with questions were instructed to call 610-719-5450.For questions regarding this recall call 610-719-5000. |  | 12 | 12 | US Nationwide Distribution to Florida only | Terminated | August 17, 2015 | 284 | 12 | Computer
Z-0941-2015 | 69806 | No | GEH | Cryosurgical unit and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | CSA Medical truFreeze System; Model: CC3-01, a cryogenic surgical device |  |  | November 10, 2014 | 2014 | CSA Medical | An increase in complaint trending for a loss of audible sound associated with the timer has been observed. | Software design | The firm, CSA Medical, sent a letter to its consignees 11/14/2014, via Fedex overnight . The letter explains the event and asks consignees to us a verbal call of time during the cryogen spray period until the root cause is determined and a permanent correction is implemented. Consignees are also asked to complete an acknowledgment form and return it to CSA Medical via, fax, email, telephone, etc. |  | 46 | 46 | US Distribution to states of: AR, CA, CO, CT, DC, FL, GA, IL, IN, LA, MA, MD, MI, NC, NJ, NY, OK, PA, RI, SC, TN, TX, UT, VA, and WV. | Terminated | March 25, 2015 | 135 | 46 | Computer
Z-0930-2015 | 69987 | No | FST | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | HarmonyÂż vLED Surgical Lighting System, one unit packaged per box |  |  | November 26, 2014 | 2014 | Steris Corporation | STERIS has learned that groups of LED lights (modules) contained within the vLED light head may illuminate inconsistently. | Software Design Change | STERIS sent an Urgent Voluntary Field Correctionn Notice to all affected consignees via FedEx. The Consignee Notification mailing was initiated the week of December 15, 2014. Customers were instructed that their local STERIS representative would contact their facility to arrange for the installation of updated software. Customers with questions were instructed to call 1-800-548-4873.For questions regarding this recall call 440-392-7231. |  | 140 | 140 units | Worldwide Distribution - USA including Alabama, Alaska, Arkansas, California, District of Columbia, Florida, Illinois, Indiana, Kansas, Louisiana, Mississippi, Nevada, New Jersey, New York, North Carolina, Ohio, Oklahoma, Oregon, Pennsylvania, Puerto Rico, South Carolina, Tennessee, Texas, Utah, and Virginia and Internationally to Canada, China, Greece, Indonesia, Republic of Korea, Romania, Russian Federation, Taiwan, and Viet Nam. | Terminated | July 11, 2015 | 227 | 140 | Computer
Z-0436-2016 | 72754 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Optional PS500 Power Supply Unit for the Evita V500 Ventilator and Babylog VN500 Ventilator. |  |  | December 01, 2015 | 2015 | Draeger Medical, Inc. | The firm became aware of cases in which the battery run times of the optional PS 500 power supply unit with the Infinity Workstation Critical Care (Evita Infinity V 500) were unexpectedly short due to the design of the charging algorithm in the current software. Devices used for patient transport will be a priority. | Software design | Draeger Medical, Inc. sent an URGENT MEDICAL DEVICE RECALL letter dated November 2015 to all affected consignees. The letter identified the product, problem, and actions to be taken by the customers. The Battery Charging Software will be updated and, if necessary, the batteries replaced in all affected devices Free of Charge. |  | 2422 | 2422 | US Nationwide-including DC and PR, and the states of AK; AR; AZ; CA; CO; CT; DE; FL; GA; HI; IA; ID; IL; IN; KS; LA; MA; MD; ME; MI; MN; MO; MS; MT; NC; ND; NE; NH; NJ; NM; NV; NY; OH; OK; OR; PA; SC; SD; TN; TX; UT; VA; VT; WA; WI; and WV. | Open, Classified |  |  | 2422 | Computer
Z-0949-2015 | 70066 | Yes | NVZ |  |  | Cardiovascular | PMA | Merlin@home RF Remote Monitoring Transmitter, Model EX1150, with software versions 6.1M or 7.0. |  |  | December 18, 2014 | 2014 | St Jude Medical Cardiac Rhythm Management Division | Transmitter may initiate a software reset resulting in backup operation in some implanted St. Jude Medical Radio Frequency (RF) enabled Implantable Cardioverter Assura, Unify Assura, and Quadra Assura ICDs and Assurity and Allure Pacemakers. | Device Design | The firm, St. Jude Medical, sent an "Important Medical Device Correction" Dear Doctor Letter dated 12/18/14 to customers to inform them that the Merlin@home transmitters may initiate a software reset resulting in backup operation in some implanted St. Jude Medical Radio Frequency (RF) enabled Implantable Cardioverter Defibrillators (ICDs) and Pacemakers. The letter informs the customers of the problems identified and the actions to be taken. Customers with questions or concerns are instructed to contact their local St. Jude Medical representative or St. Jude Medical's Technical Services Department at 1-800-722-3774. |  | 257456 | 257,456 units | Worldwide Distribution: US (nationwide) and countries of: EU, Mideast, Canada, Australia/New Zealand, and Asia-Pacific (including Japan). | Terminated | March 30, 2016 | 468 | 257456 | Computer
Z-0953-2015 | 70014 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | The ATTUNE INTUITION Impaction Handle is a re-useable instrument utilized in knee replacement procedures. The instrument interfaces with several related ATTUNE instruments: keel punch, tibial tower, evaluation bullet, fixed bearing tibial impactor, rotating platform tibial impactor, and femoral impactor, for the purpose of impaction, extraction, and insertion applications at steps within the surgical procedure. Due to the frequency of its use, two impaction handles are supplied in each kit. |  |  | November 10, 2014 | 2014 | DePuy Orthopaedics, Inc. | Complaints have indicated that the ATTUNE INTUITION Impaction Handle lever has fractured during impaction. If the lever fractures, there is the potential for fractured pieces of the device, or the inner spring, if released, to be left in the patient if not observed during surgery. | Device Design | On 11/10/2014, URGENT INFORMATION ATTUNE INTUITION IMPACTION HANDLE AND IMPACTORS DEVICE CORRECTION NOTICE notifications were sent to the affected distributors and medical professionals. All distributors were notified via electronic mail, and all medical professionals with affected inventory were also notified by distributor or Sales Consultant via in person delivery or mail or electronic mail. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. For questions, please contact Kim Earle, Recall Coordinator, at 574-371-4917 (M-F; 8 a.m. - 5 p.m. EDT). |  | 7944 | 7,944 | Worldwide Distribution-US (nationwide) except AK and RI, including PR and the countries of Chile, Canada, Australia, Austria, Belgium, France, Hong Kong, Ireland, India, Israel, Italy, Japan, Malaysia, Netherlands, New Zealand, Saudi Arabia, Singapore, Slovakia, South Africa, South Korea, Spain, Switzerland, Thailand, United Arab Emirates (UAE), and United Kingdom (UK | Terminated | December 08, 2015 | 393 | 7944 | Computer
Z-1019-2015 | 70141 | No | IOR | Mechanical wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Zippie Zone Pediatric Rigid Wheelchair Model EIRZ1. |  |  | December 19, 2014 | 2014 | Sunrise Medical (US) LLC | There is the potential for the height adjustable handle to dislodge due to failing locking mechanisms on specific identified chairs. | Device Design | Sunrise Medical sent an Urgent Medical Device Field Correction letters dated December 19, 2014 to the two affected distributors. The letter identified the affected product, problem and actions to be taken. For question contact Sunrise Medical Regulatory Affairs at either (559)294-2840. |  | 3 | 3 | Nationwide Distribution in the states of CO and ME. | Terminated | March 10, 2015 | 81 | 9 | Computer
Z-1027-2015 | 70318 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo.plaza: Syngo.plaza is a Picture Archiving and Communication System (PACS) intended to display, process, read, report, communicate, distribute, store and archive digital medical images. It supports the physician in diagnosis and treatment planning. |  |  | January 13, 2015 | 2015 | Siemens Medical Solutions USA, Inc | In case of a system crash, images may not be written to the hard disk and this may result in inconsistencies in the database. In case of a system crash (e.g. blue screen, power outage) images may not be written from cache to the hard disk and might get lost. | Software Manufacturing/Software Deployment | Siemens sent a Safety Advisory Notice dated January 13, 2015, to all affected customers. The Safety Notice identified the product, the problem, and the action to be taken. Siemens recommended checking patient data in the event of a system crash and advised customers that a software update is planned for mid-2015 that will resolve the issue.For questions regarding this recall call 610-219-6300. |  | 50 | 50 | Nationwide Distribution | Open, Classified |  |  | 50 | Computer
Z-0284-2016 | 71680 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Prism Medical C-300 Consumer Ceiling Lift; PN 330050, 330050R, 330053, 330053R, and 330057; weight specification 300 lbs. /136 kg. |  |  | September 08, 2015 | 2015 | ErgoSafe Products, LLC (DBA) Prism Medical | Drive shaft and over speed arm may simultaneously fail potentially resulting in a patient falling during use | Device Design | The firm, Prism Medical, sent a "Field Safety Notice (FSN 0008)" letter dated 9/8/2015 to all consignees by email via their sales portal. The letter describes the product, problem and actions to be taken. The consignees were instructed to call Prism Medical directly at 1-314-219-8655 and indicate that you are participating in the 300 series lift recall; provide your existing 300 series lift model number and serial number; follow the instructions provided; complete and return the Acknowledgement Form via scanned &amp; signed to bschoenman@prismmedicalltd.com or fax to: 314-692-7858, Attention: 300 series recall.For further information please contact: Prism Medical Quality department at: 1-314-219-8655 or North American Director for Quality &amp; Regulatory Affairs at 1-314-219-8614. |  | 561 | 561 lifts | Worldwide Distribution: US (nationwide) and countries of: Australia, Canada, South Korea, Taiwan, Mexico, Japan, and Great Britain. | Terminated | May 04, 2017 | 604 | 2282 | Computer
Z-1131-2015 | 70340 | No | DSQ |  |  | Cardiovascular | PMA | HeartWare Ventricular Assist System (HeartWare Controller) |  |  | January 16, 2015 | 2015 | HeartWare Inc | The affected clinical trial Controllers exhibit a higher susceptibility to ESD than newer commercial Controllers. | Device Design | HeartWare sent an URGENT MEDICAL DEVICE RECALL dated January 16, 2015 asking clinicians to quarantine the affected controllers and complete and return the acknowledgement form. HeartWare also asks clinicians to identify patients with affected controllers and review the applicable risks with the patient and as soon as reasonably possible if medically advisable, exchange the affected controller under medical supervision with a new controller. HeartWare also asks that for each patient using an affected controller as their backup to contact the patient and arrange to have the backup controller replaced with a new controller. Lastly, HeartWare asks the firm to complete the patient tracking form and return all quarantined controllers to HeartWare along with the completion form. Should you have any questions or concerns, please contact your local HeartWare representative. Our 24-Hour Clinical Support personnel are also available at 1-888-494-6365. |  | 4845 | 4,845 | Worldwide Distribution - US Nationwide in the states of: AL, AZ, CA, CO, FL, GA, IL, IN, KY, LA, MA, MD, MI, MN, MO, NC, NY, OH, OR, PA, TX, UT, VA, WA and District of Columbia and countries of: Canada, Poland, Australia, Austria, Italy, DE, New Zealand, South Africa, France, Switzerland, Germany, Brazil, UK, Belgium, Netherlands, Israel, Malaysia, Turkey, Greece, Japan, Singapore, Poland, Lithuania, Sweden, Norway, Finland, and India. | Open, Classified |  |  | 4845 | Computer
Z-1115-2015 | 70304 | No | NDP | Endosseous dental implant accessories. | Dental | Dental | 510(K) Exempt | GemLock Long Hex Driver, Catalog RHL2.5; Also distributed within the Tapered SwissPlus &amp; SwissPlus Implant Systems Kits, Surgical, Complete, Catalog OPCST; Tapered Screw-Vent Implant System Surgical Kit, Complete, Catalog TSVKIT. |  |  | January 22, 2015 | 2015 | Zimmer Dental Inc | The Gemlock Long Hex Driver may not fit into the Fixture Mount Transfer (FMT) or the internal hex of the implant. | Device Design | An Urgent Device Recall Notice was sent on 1/22/15 to their domestic and foreign customers who purchased the Gemlock Long Hex Driver. The letter informs the customers of the problems identified, risks, and the actions to be taken. Customers are instructed to review the notice and check their inventory for the affected units. All affected product is to be immediately quarantined and removed from service. Customers are to contact Simmer Dental Customer Service at 1-800-854-7019 to obtain a Return Authorization Number and set up return of the products. Customers should complete the attached Business Reply Form and fax it to (574) 372-4265 or e-mail to corporatequality.postmarket@zimmer.com. Customers with questions are instructed to contact Zimmer Dental Customer Service at (800) 854-7019, Mon-Fri, 7 am-5pm (Pacific).Zimmer Dental is expanding the recall to include 1 additional lot for catalog no.TSVKIT. Zimmer Dental plans to inform an amended customer notification letter to their US and foreign consignees. The letter will go out to customers on 4/20/15 to inform them of the expanded lot. |  | 627 | 627 units | Worldwide Distribution -- US, Argentina, Australia, Bulgaria, Canada, Chile, Colombia, Costa Rica, Dominican Republic, Ecuador, Egypt, France, Germany, Hong Kong, Hungary, Iran, Israel, Italy, Japan, Lebanon, Lithuania, Malaysia, Mexico, Morocco, Netherlands, Pakistan, Panama, Poland, Romania, Russia, Saudi Arabia, Serbia, Taiwan, Thailand, Tunisia, Turkey, and UAE.Expanded Distribution: Spain | Terminated | August 03, 2015 | 193 | 627 | Computer
Z-1116-2015 | 70492 | Yes | DWC | Cardiopulmonary bypass temperature controller. | Cardiovascular | Cardiovascular | 510(k) | Terumo HX2 Temperature Management System provides temperature control of two independent water circuits that directly controls the temperature of patient blood and cardioplegia solution during cardiovascular surgery. |  |  | February 06, 2015 | 2015 | Terumo Cardiovascular Systems Corporation | Malfunctioning mixing valve and loss of cooling and/or heating capability. System goes to standby mode and shows error messages EOd (invalid mix valve measured range of motion) or EOE (valve unable to move or complete a motion). Mixing valves need to be replaced at regular intervals. | Device Design | On 2/6/2015, URGENT MEDICAL DEVICE RECALL notifications were mailed to the affected hospitals. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. For questions, call customer service at 1-800-521-2818. |  | 103 | 103 units | US (nationwide) including the states of AR, CA, CO, CT, FL, IA, IL, KS, KY, MD, MI, MO, MS, NY, OH, PA, TX, UT, WA, and WI and the country of Japan. | Open, Classified |  |  | 103 | Computer
Z-0342-2016 | 72618 | Yes | NKE |  |  | Cardiovascular | PMA | Medtronic InSync III, model 8042, 8042B and 8042U. Implantable pulse generator pacemaker. |  |  | November 09, 2015 | 2015 | Medtronic Inc. | Medtronic has confirmed an issue with InSync III CRT-pacemaker devices related to long-term battery performance. The root cause is unexpected high battery impedance. | Device Design | Medtronic sent an Urgent Medical Device Correction letter dated November 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Medtronic records indicate you are following one or more patients with this device. Please see the enclosed Physician/Patient Detail Report.We regret any difficulties this may cause you and your patients. We will continue to monitor device performance and provide regular updates in our product performance report available at wwwp.medtronic.com/productperformance/. Medtronic Patient Services is available to assist patients at 800-551-5544.Medtronic will offer a supplemental device warranty. Contact your sales representative for terms and conditions. If you have any questions, please contact your local Medtronic Representative or Medtronic Technical Services at 800-505-4636. |  | 96787 | 96,787 | Worldwide Distribution - USA (nationwide) including Puerto Rico and to Algeria, Argentina, Armenia, Australia, Austria, Bangladesh, Belarus, Belgium, Bermuda, Bolivia, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, El Salvador, Estonia, Finland, France, French Polynesia, Georgia, Germany, Greece, Guatemala, Honduras, Hong Kong, Hungary , India, Indonesia, Iran, Ireland, Israel, Italy, Japan Jordan, Kazakhstan, Kenya, Kuwait, Korea, Latvia, Luxembourg, Lebanon, Lithuania, Macedonia, Malaysia, Malta, Martinique, Mauritius , Mexico, Morocco, Netherlands, New Zealand, Norway,Oman, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Qatar, Reunion, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Syria, Trinidad, Tobago, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Venezuela and Vietnam. | Open, Classified |  |  | 96787 | Computer
Z-0335-2016 | 72554 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | NxStage System One S Cycler - Model no. NX1000-3. |  |  | October 29, 2015 | 2015 | NxStage Medical, Inc. | Ultrafiltration Volume (UFV) may not decrease during treatment-software error. The UF pump may continue to run and remove fluid even after the target UF volume has been removed | Software design | NxStage issued a Field Safety Notice (FSN), dated October 29, 2015, to clinic and home patients with a NX1000-3 or NX1000-4 cycler. For patients performing short daily hemodialysis, the FSN included instructions for steps to be taken so that the cycler and treatments are not affected by the software error. Patients performing nocturnal hemodialysis were instructed to contact their Center for a replacement NX1000-1 Cycler. A software update will be released to correct the error. Customers should complete and return the reply form. Contact NxStage Customer Service at 1-866-NXSTAGE (1-866-697-8243) with questions or comments. |  | 2134 | 2,134 units | Worldwide Distribution. | Open, Classified |  |  | 2274 | Computer
Z-1125-2015 | 70363 | No | DXH | Telephone electrocardiograph transmitter and receiver. | Cardiovascular | Cardiovascular | 510(k) | Alive ECG App 2.1.2 ( a medical device application for the Apple iOS operating system, intended to be used with the AliveCor Heart Monitor. The AliveCor Heart Monitor is intended to record, store and transfer single-channel electrocardiogram (ECG) rhythms. The AliveCor Heart Monitor also displays ECG rhythms and detects the presence of atrial fibrillation and normal sinus rhythm (when prescribed or used under the care of a physician). |  |  | January 09, 2015 | 2015 | Alivecor SFO | Alive ECG App version 2.1.2 (intended to be used with the AliveCor Heart Monitor) crashed upon use of the application. | Software change control | AliveCor posted information on their website, Facebook Page and Twitter to alert users of the issue. |  | 5600 | 5600 active users with Alive ECG app for iOS . | Downloaded by Apple users - locations not shared by Apple. | Terminated | February 18, 2015 | 40 | 5600 | Computer
Z-1126-2015 | 70294 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | BodyGuard 323 pump, models |  |  | January 16, 2015 | 2015 | CME America, LLC | CME America is recalling the BodyGuard and BodyGuard 323 Infusion pumps due to the potential for an over delivery. | Software design | CME America sent an "Urgent Medical Device Recall" letter dated January 16, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Required Action for Users: To date there have been no reports of adverse events or serious injury. However, there is the potential for serious injury or death. CME America recommends that you DO NOT use the BodyGuard 323 pump for PCA infusion. If your devices are currently in use, be sure that the programmed parameters DO NOT involve a patient controlled bolus dose less than 0.1 miln volume (translated into meg or mg).Follow-up Actions by CME America: CME America has identified the root cause of this issue and is recommending installation of a new version of software, USA36B4, for affected BodyGuard 323 infusion pumps. This software is identical to the previously released version, USA36A4.2. However the designation of B4 in the software name will identify updated software in each device. CME America will contact all affected customers within 60 days to initiate the scheduling process for the remediation.Your assistance is appreciated and necessary to prevent use and facilitate remediation of the recalled product. CME America can address any questions regarding the recalled product by calling the CME America recall line at 720-839-2391. |  | 3186 | 3,186 | Worldwide Distribution - Nationwide Distribution and VA/military/govt consignees and the country of Canada. | Terminated | March 31, 2017 | 805 | 3186 | Computer
Z-1128-2015 | 70374 | No | OLO | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | Restoris Partial Knee Application (PKA) RIO (TGS 2.0). |  |  | January 22, 2015 | 2015 | Mako Surgical Corporation | When using the MAKOplasty partial knee Arthroplasty application, the burr continues spinning outside of the stereotactic boundary and after the control switches (foot pedal and trigger) cease to be activated | Device Design | On January 22, 2015 the firm sent an IMPORTANT MEDICAL DEVICE CORRECTION asking consignees to complete the attached acknowledgement form and return them by email, fax or mail. The notice also noted the RIO system can continue to be used. The likelihood of occurrence is rare (.035% of cases) and the event is readily detectable and can be resolved. Should you have any questions regarding this Important Medical Device Recall Notice, to contact us at (954) 628-1721. |  | 228 | 228 units | Nationwide Distribution-including the states of AK, AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, LA, MA, MD, MI, MO, MS, NC, NH, NJ, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, WI, WV and Hawaii. | Open, Classified |  |  | 228 | Computer
Z-1137-2015 | 69378 | No | ITI | Powered wheelchair. | Physical Medicine | Physical Medicine | 510(k) | TDX Electric Wheelchair. Invacare FDX, TDX series and Storm Series (3G) power wheelchairs in 20 to 24 inch width configurations with power recline seating system and a conventional style back. |  |  | January 15, 2015 | 2015 | Invacare Corporation | If the slack in the wires is not routed and secured correctly, flexing of the back pan under the users weight may cause the headrest knob or clamp bracket to pinch, damage or cut the wire. Smoking, sparking, burning or fire may occur if a wire is repetitively pinched at the same place. | Device Design | The firm initiated the recall on 01/16/15 by sending a notification letter to customers through the mail. Communication was sent to providers to contact consignees via mail return receipt to inform them of the recall. Recall notification included a Provider letter, Consumer letter, Provider Response card, Unit Disposition form and tracking sheet and the Wiring guide. The notification instructed the provider to visually inspect all potentially impacted chairs. Not all potentially impacted chairs are expected to have improperly routed and secured wires. |  | 1942 | 1942 | Worldwide Distribution-US (nationwide) including DC and the states of AK, AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WV, and the countries of Canada and Australia. | Terminated | September 23, 2015 | 251 | 2527 | Computer
Z-1145-2015 | 70453 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Spacelabs Healthcare qube Compact Monitor, Model 91390. |  |  | January 28, 2015 | 2015 | Spacelabs Healthcare Inc | Spacelabs Healthcare qube Bedside Monitors, Model 91390, are recalled because the firm has received multiple reports of qube monitors failing to boot up or returning to factory default configuration settings following power on or reset. | Device Design | Spacelabs began to send the Urgent Medical Device Correction XPREZZON Bedside Monitor, Model 91393 qube Compact Monitor, Model 91390 letter, dated 28, January 2015, to US consignees.Spacelabs began sending the Urgent Medical Device Correction XPREZZON Bedside Monitor, Model 91393 qube Compact Monitor, Model 91390 letter (translated as necessary) to international customers via email on 2/4/15.Actions to be Implemented by Users:-Do not rely solely on the monitor for critically ill patients.-If customer facility operates any XPREZZON or qube monitors on a standalone basis, rather than via central monitoring, please contact Spacelabs technical support as these monitors will be the first repaired.Spacelabs Healthcare will contact consignees to schedule a convenient time for Spacelabs to replace the PCBAs in all of your facility's affected XPREZZON and qube monitors at no cost.If you have any questions about this corrective action, please contact Spacelabs at the toll free number of 1-800-522-7025 and select 2 for Technical Support, or via direct dial at 1-425-363-5203. |  | 2955 | total 2955 units (2117 units in US and 838 units outside US) | Worldwide Distribution. | Terminated | December 30, 2015 | 336 | 2955 | Computer
Z-0326-2016 | 72543 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | NxStage System One S Cycler (High Permeability Hemodialysis System) |  |  | October 29, 2015 | 2015 | NxStage Medical, Inc. | Ultrafiltration (UF) Volume software error inaccurate fluid removal | Software Design Change | NxStage issued issued a Field Safety Notice (FSN) on 10/29/15 to consignees who have NX1000-5 or a NX1000-5-A cycler. The FSN included Instructions for steps to be taken so that the cycler and treatments are not affected by the software error.NxStage will be sending a representative to the facility to attach detailed instructions on each cycler as a user reference. Return the attached reply form to acknowledge receipt of this letter.Questions or comments, please contact NxStage Customer Service at 1-866-NXSTAGE (1-866-697-8243). |  | 323 | 323 units | Worldwide Distribution. US Nationwide, Netherlands, and Sweden. | Open, Classified |  |  | 339 | Computer
Z-1022-2015 | 70068 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | VC150 Vital Signs Monitor (VC150 monitor equipped with Masimo SpO2 technique only); Innokas Medical CARESCAPE VC150 Rx Only Innokas Medical Oy. Catalog numbers 2067980-002, 2067980-006, 2067980-010, and 2067980-014. |  |  | December 03, 2014 | 2014 | INNOKAS MEDICAL OY | If the SpO2 or RRa value is violating the respective alarm limit at the time of switching to monitoring mode, the monitor does not issue visual and audible alarms. | Software design | A notification letter was sent to the distributors on 12/3/2014. The letter requests the devices be returned to the manufacturer for repair. |  | 20 | 20 units | Distributed in the US and Europe. | Terminated | April 24, 2015 | 142 | 20 | Computer
Z-0325-2016 | 69127 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Ingenia, Intera, Achieva and Achieve dStream MR systems on Software versions R5.1.7 and R5.1.8 with the Mobi View software option. Diagnostic imaging system. |  |  | August 21, 2014 | 2014 | Philips Medical Systems, Inc. | When a fused series of a sagittal, coronal or radial multi station scan is generated in MobiView, the resultant image order is reversed. This occurs with software release R5.1.7/R5.1.8. | Software Design Change | Philips Healthcare sent an Customer Information Letter dated July 2014, on August 21, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Philips will issue a Field Change Order to install updated software in a Service Pack that makes the slice numbering consistent for fused and unfused images.If you need any further information or support concerning this issue, please contact your local Philips representative: 1-800-722-9377 |  | 307 | 307 units | Worldwide Distribution - USA (nationwide) and to the countries of: Argentina, Australia, Austria, Bangladesh, Belgium, Brazil, Canada, Chile, Denmark Estonia, Ethiopia. France, Germany, Hong Kong, India, Indonesia, Iraq, Ireland, Italy,Japan, Jordan, Korea, Republic of, Libya, Malaysia, Mexico, Nepal, Netherlands, New Zealand, Norway, Oman, Palestinian territory, Peru, Poland, Russian Federation, Saudi Arabia, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, United Arab Emirates and United Kingdom. | Open, Classified |  |  | 307 | Computer
Z-1212-2015 | 70545 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | COMPASS SW Version 3.1, Catalog Number CS10-100, medical linear accelerator, radiological |  |  | February 02, 2015 | 2015 | Iba Dosimetry Gmbh | Error in the software. During internal tests of the current development version of the Compass SW it was found that dose reconstruction for DMLC plans when the jaws move or mlc leaves reverse during beam on will not be scaled correctly for all control points. | Software design (manufacturing process) | IBA sent an Field Safety Notice dated January 28, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Actions to be taken:Please do not use the dose reconstruction function COMPASS 3.1 and 3.1a for DMLC plans from Monaco or Eclipse -11 when the jaws move or MLC leaves reverse during beam on. A correct Version ( COMPASS 3.1b) will be delivered as soon as it is available. The field safety notice including a confirmation of receipt was sent to the affected hospitals; by email and UPS letter on February 2, 2015.Should you have questions or require additional information, please contact the Customer Service Team at South and Middle America, USA and Canada - Phone. +1 901.386.2246Asia Pacific, Australia and New Zealand - Phone. +86 10 8080 9107All other countries - Phone. 49 9128 607 - 38 |  | 124 | 124 units | Worldwide Distribution - US Distribution to the states of AZ and OK., and the countries of Germany, Austria, France, UK, Poland, Finland, The Netherlands, Spain, Turkey, Russia, and South Africa, Algeria, North and South America, Mexico, Brazil, Venezuela, Colombia, Panama, Asia, China, Japan, Philippines, Korea, India, Singapore, Thailand, Taiwan, Hong Kong, Malaysia and Guam. | Terminated | August 18, 2015 | 197 | 124 | Computer
Z-0319-2016 | 72576 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo Imaging VB36D_HF02. Radiological image processing system. |  |  | October 28, 2015 | 2015 | Siemens Medical Solutions USA, Inc | To provide supplementary information regarding the release of the syngo Imaging Software Version VB36D_HF02. The software provides improvements for all syngo Imaging installations running on the software Version VB36D. | Software change control | Siemens sent an Customer Information Letter dated October 28, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.If the device has been sold and therefore no longer in your possession, please forward this notice to the new owner. We would also request that you inform us of the identity of the device's new owner where possible. Please observe this safety notice and comply with the corresponding measure until the update has been fully completed. For further questions, please call: (610) 219-6300. |  | 3 | 3 | US Distribution to the states of : NC, NE and OH. | Open, Classified |  |  | 3 | Computer
Z-0320-2016 | 72344 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart MRx Monitor/Defibrillator; M3535A, M3536A, M3536J, M3536M, M3536MC, M3536M2, M3536M4, M3536M5, M3536M6. |  |  | October 14, 2015 | 2015 | Philips Electronics North America Corporation | The following MRx software issue has been identified: MRx model M3535A with software version F.03.06 and earlier, and model M3536A with version T.00.05 and earlier may stop the automated Ready-For-Use (RFU) test in an abnormal state when the device is turned off. | Software design | On 10/14/2015, Philips sent an Urgent Field Safety Notice to each known affected customer. The Urgent Medical Device Correction notification/ Field Safety Notice informs customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue. Philips will provide a software upgrade for M3535A (software version F.03.07 or higher) and M3536A (software version T.00.06 or higher) to correct this issue. A software upgrade will be provided free of charge for all affected products. A Philips Healthcare representative will contact customers with affected devices to arrange for installation of the upgrades.Philips is asking customers to follow the Action to be Taken by Customer/User section of the Urgent Medical Device Correction notification/Field Safety Notice: All MRx users should be informed that if the pattern appears, they can continue to use the device. Once the MRx is turned on, the pattern will disappear and the device will function normally in all monitoring and therapy modes. |  | 81161 | 81,161 | Worldwide Distribution - US (nationwide) including Puerto Rico, and the countries of AFGHANISTAN, ALBANIA, ARGENTINA, ARUBA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BANGLADESH, BELGIUM, BOLIVIA, BOSNIA AND HERZEGOVINA, BRAZIL, BRUNEI DARUSSALAM, BULGARIA, CAMBODIA, CHILE, CHINA, COLOMBIA, COSTA RICA, CROATIA (local name: Hrvatska), CYPRUS, CZECH REPUBLIC, DENMARK, ECUADOR, EGYPT, EL SALVADOR, ESTONIA, ETHIOPIA, FINLAND, FRANCE, GERMANY, GHANA, GREECE, GUADELOUPE, HONG KONG, Hungary, ICELAND, India, INDONESIA, ISLAMIC REPUBLIC OF IRAN, Iraq, IRELAND, ISRAEL, ITALY, JAPAN, Jordan, KAZAKHSTAN, KENYA, REPUBLIC OF KOREA, KUWAIT, LAOS, LATVIA, LEBANON, LIBYAN ARAB JAMAHIRIYA, LITHUANIA, LUXEMBOURG, MACAU, MACEDONIA, MALAYSIA, MALTA, Mauritius, Mexico, MONGOLIA, MOROCCO, MOZAMBIQUE, Myanmar (Burma), NAMIBIA, NEPAL, NETHERLANDS, NETHERLANDS ANTILLES, NEW CALEDONIA, NEW ZEALAND, NICARAGUA, NIGERIA, NORWAY, OMAN, PAKISTAN, State of Palestine, PANAMA, PAPUA NEW GUINEA, PARAGUAY, PERU, PHILIPPINES, POLAND, PORTUGAL, QATAR, REUNION, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SERBIA, SINGAPORE, SLOVENIA, SOUTH, AFRICA, SPAIN, SRI, LANKA, SWEDEN, SWITZERLAND, TAIWAN, UNITED REPUBLIC OF TANZANIA, THAILAND, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, TURKMENISTAN, UGANDA, UKRAINE, UNITED, ARAB, EMIRATES, UNITED KINGDOM, UZBEKISTAN, VENEZUELA, VIETNAM, and YEMEN | Open, Classified |  |  | 81161 | Computer
Z-1214-2015 | 64784 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | The Centricity PACS Workstation is intended for use as a primary diagnostic and analysis tool for diagnostic images by trained healthcare professionals, including radiologists, physicians, technologists, clinicians and nurses. It is also intended for use as a clinical review workstation throughout the healthcare facility. The workstation interface provides the user with a means to display, manipulate, archive, print and export images when connected with the Centricity PACS infrastructure. |  |  | February 18, 2013 | 2013 | GE Healthcare | Using Merge Exam in single Study Mode may result in Missing Study Record (Cannot Display Exam, Send Exam, etc). Exam merge in Single Study Mode may fail on &amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp | Software design | Consignees were notified of the issue via telephone and site visit by GE Healthcare IT Field Engineers. |  | 128 | 128 systems | Distributed US (nationwide) including the states of AK, AZ, CA, CO, FL, GA, HI, IA, ID, IL, IN, KS, LA, MA, MD, ME, MN, MO, MS, NC, NE, NJ, NY, OH, OK, OR, PA, SC, SD, TN, TX, VA, WA, WI, and WV, and the country of Bermuda. | Open, Classified |  |  | 128 | Computer
Z-1224-2015 | 70224 | No | CBN | Gas-scavenging apparatus. | Anesthesiology | Anesthesiology | 510(k) | Servo-I ventilator system CO2 analyzer. Part number 65 23 588. |  |  | January 12, 2015 | 2015 | Maquet Cardiovascular Us Sales, Llc | A calibration problem was detected in a specific number of the Servo-I ventilator system CO2 analyzer models that may have been caused by the module's internal circuit board function. | Device Design | Maquet sent an Urgent Product Recall Medical Device Action Field Removal Letter dated 1/12/2015 via FedEx. |  | 486 | 486 units | Worldwide distribution - USA nationwide, Austria, Australia, Belgium, Chile, Canada, Colombia, Ecuador, Egypt, Finland, France, Germany, Hungary, Iceland, India, Indonesia, Iraq, Italy, Japan, Libya, Netherlands, Norway, Panama, Peru, Portugal, Romania, Russia, Saudi Arabia, South Africa, Spain, Sweden, Thailand, Tunisia, Turkey, Ukraine, and United Kingdom. | Open, Classified |  |  | 486 | Computer
Z-0297-2016 | 72427 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Bulk Loader Module for cobas p 512/612 |  |  | September 22, 2015 | 2015 | Roche Diagnostics Operations, Inc. | Possible for sample tubes in the loader to be opened and contaminate the system as well as adjacent samples during processing. | Device Design | Roche Diagnostics Operations, Inc. sent an Urgent Medical Device Correction letter to affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were asked to retain the Urgent Medical Device Correction letter until a Roche Field Service Representative has installed the BLM modification kit. Consignees were also asked to complete the attached form and fax to 1-877-906-8982, ext 2015. Consignees with questions were instructed to call 1-800-428-2336.For questions regarding this recall call 317-521-3911. |  | 1 | 1 | Distributed to TX only. | Terminated | February 03, 2016 | 134 | 1 | Computer
Z-1209-2015 | 70506 | Yes | OAD |  |  | Cardiovascular | PMA | IntellaTip MiFi Open-Irrigated radiofrequency Ablation Catheters: |  |  | February 03, 2015 | 2015 | Boston Scientific Corp | Reports of formation of char adherent to the proximal part of the distal tip electrodes | Device Design | Customer notification letters were sent by overnight mail on February 10, 2015. |  | 14 | 14 units | Distributed in Germany, Great Britain, and the Netherlands. | Terminated | March 21, 2017 | 777 | 14 | Computer
Z-0294-2016 | 72484 | Yes | KPO | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius 132 Gallon GranuFlo Dissolution Unit - P/N G047-80101. |  |  | October 23, 2015 | 2015 | Fresenius Medical Care Renal Therapies Group, LLC | Dissolution units and replacement motors units do not provide full thermal protection for both the Main and Starter coils resulting in the generation of smoke and/or potentially lead to fire. | Device Design | Fresenius issued an Urgent - Medical Device Correction notification, dated October 21, 2015, on 10/27/15. The letter identified the affected device and the reason for the recall. It also contained instructions (on how to perform the required wiring modification) along with a Reply Form. Customers are to return the form after completing the modification. Inquiries regarding the correction should be directed to Fresenius Technical Services at 800-227-2572, referencing the Field Action number FA-2015-01-W. |  | 918 | 918 units | Nationwide Distribution. | Open, Classified |  |  | 2258 | Computer
Z-1457-2015 | 70835 | No | CHI | Lipase test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ADVIA Chemistry Lipase Reagent , Catalog Number B01-4840-01 REF: 01984894. |  |  | March 23, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc. | Siemens internal investigation confirmed that current contamination avoidance settings do not prevent carryover resulting in falsely elevated ADVIA Chemistry Lipase outlier results when running with Triglycerides, Triglycerides_2, Triglycerides_c, Cholesterol concentrated, and DLDL reagents. | Device Design | Siemens issued an Urgent Medical Device Correction on March 23, 2015, notifying direct consignees about the product, problem, and action to be taken. Customers were instructed to follow the directions provided in the customer letter in the event that a Lipase outlier is observed. All customers were requested to complete and return the Effectiveness Check questionnaire attached to the Urgent Medical Device Correction letter within (30) days. Field service personnel have been sent an SKB (Service Knowledge Base) describing the issue and instructing them how to deal with customer questions. For further questions please call (312) 275-7795. |  | 7814 | 7814 | Worldwide Distribution - US (nationwide) Distribution and to the countries of : Thailand, Pakistan, Vatican, Guadeloupe, Reunion, French Polynesia , Serbia, Martinique, India, China, Singapore, Brazil, South Africa, Mexico, Canada, Israel, Australia, Egypt, Russian Federation, Indonesia, Rep. Korea (S) ,Unit.Arab Emir., Taiwan ,Argentina ,Chile, Malaysia ,Colombia, Peru, Kazakhstan, Paraguay, Vietnam, Algeria, Austria, Belgium, Switzerland, Cyprus, Czech Republic, Germany, Estonia, Spain, Finland, France, United Kingdom, Greece, Hungary, Italy, Lithuania, Luxembourg, Latvia, Netherlands, Norway, Poland, Portugal, Romania, Slovenia, Slovakia and Turkey. | Terminated | December 18, 2015 | 270 | 7814 | Computer
Z-1089-2015 | 70376 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Telluride Percutaneous Rod Inserter Long, Model #7706-1904. The Inserter is a reusable device provided within an aluminum sterilization case along with other orthopedic manual surgical instruments. The case consists of three trays of instruments, and the middle tray contains a quantity of two Inserters. |  |  | December 19, 2011 | 2011 | Biomet Spine, LLC | The Telluride Percutaneous Rod Inserter Long does not contain a feature that locks the Rod in place during use. There is a potential for the Rod to disengage from the Inserter during insertion, and could lead to harm to the patient. | Device Design | Consignees were informally contacted 12/16/2011 of units in the field requesting retrieval. All units were in Lanx's (Biomet's) possession at the time of the recall initiation on 12/19/2011. |  | 10 | 10 | Distributed to TX, OH, NE and NY. | Terminated | March 23, 2015 | 1190 | 10 | Computer
Z-0361-2016 | 72524 | Yes | DTR | Cardiopulmonary bypass heat exchanger. | Cardiovascular | Cardiovascular | 510(k) | Model 5001102 MPS Delivery Set w/ arrest agent and additive cassettes, heat exchanger &amp;10 ft. delivery tubing; Model 5001102-AS MPS Delivery Set with 6 ft. delivery tubing; Model 7001102 MPS Low Volume Delivery Set for use with cardiopulmonary bypass equipment. |  |  | October 28, 2015 | 2015 | Quest Medical, Inc. | The products have been found to intermittently exhibit a seal failure during use. | Equipment maintenance | The recalling firm issued a press release on 10/28/2015. The press release was posted on FDA website on 10/29/2015. The recalling firm emailed affected consignees on 10/28/15 and sent out recall letters via USPS beginning 10/30/15. |  | 18761 | 18,761 units | Distributed US (nationwide) and the countries of Canada, Japan, Saudi Arabia, Kuwait, South Africa, Switzerland, France, and Italy. | Open, Classified |  |  | 18761 | Computer
Z-0364-2016 | 69090 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Patient Support Material Number: 9890 010 87431 used with: BuckyDiagnost ; Model: 704060/706037 |  |  | August 19, 2014 | 2014 | Philips Medical Systems, Inc. | The hook does not securely hold the footplate in vertical position. Therefore the footplate can fall or swing down and may hurt the operator or patient. | Device Design | The Field Safety Notice FCO70600086 will be mailed on August 19, 2014. Philips Medical Systems will send out certified letters and will have the receipts returned for tracking purposes as well as have a Field Service Engineer visit each site to provide the fix when the Field Change Order is released. |  | 1517 | 1517 systems, not all units have the patient support accessory | Distributed in the states of MN, KY, MA, VA, TX, NC, HI, FL, AZ, CT, and WY and the countries of Argentina, Australia, Austria, Canada, Denmark, Germany, Netherlands, Russia, Saudi Arabia, Singapore, Spain, and Switzerland. | Open, Classified |  |  | 1517 | Computer
Z-1034-2015 | 69299 | No | OYC |  |  | Clinical Chemistry | PMA | Animas Vibe Insulin Infusion Pump. |  |  | September 06, 2011 | 2011 | Animas Corporation | Calibration factors in the pump overwritten during a programming step. The force sensor could send a lower signal value to the pump processor, with loss of prime warnings, occlusion alarms and the pump unable to detect a cartridge during the prime sequence. Field action initiated 8/29/2011. | Software design (manufacturing process) | The field action was initiated on 8/29/2011 and comprised communication to patients and distributors using email, letters and verbal communication. In addition, notifications were sent to those Health Care Professionals who had patients that were affected by this field action. Health Authorities were notified in those countries where the pump had been distributed (France, Germany, Sweden and United Kingdom) or was in the hands of patients. Replacement product was provided to distributors and end users who were identified as having affected product. |  | 1235 | 1235 | No US distribution, Distributors are located in France, Germany, Sweden and United Kingdom. | Terminated | August 10, 2015 | 1434 | 1235 | Computer
Z-0999-2015 | 70058 | No | OWD |  |  | Molecular Genetics | PMA | cobas 4800 KRAS AMP/DET 24T CE-IVD Mutation Test; CE-IVD 5852170190. |  |  | December 18, 2014 | 2014 | Roche Molecular Systems, Inc. | cobas KRAS Mutation Kit T10786 is generating invalid results for the Mutant Control and/or KRAS Calibrator. | Device Design | RMS issued an Urgent Field Safety Notice on 12/18/2014.A Quality Notification was provided to all users of the of cobasÂż KRAS Mutation Kit CE-IVD, lots T13091, T15435, T15475, T14777, and W00389, informing them of the issue and instructing them to discontinue use and discard any remaining inventory immediately. |  | 660 | 660 pieces | Foreign Distributed. | Open, Classified |  |  | 660 | Computer
Z-1038-2015 | 70021 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | SoftLab with SA HIS versions 4.0.7.0-4.0.7.1 |  |  | May 20, 2014 | 2014 | SCC Soft Computer | The interface fails to send abnormal flags for Reference Lab test results. | Device Design | The firm sent out 2 CORRECTION COMMUNICATIONS informing consignees of functionality in release of 4.0.7 HIS Add-on interfaces which SCC has designated to be a potential patient Risk-to-Health issue. Further the corrective action indicated HIS releases 4.0.7.2 and higher will be corrected. Hot Fixes have been developed for HIS Add-on patches 4.0.7.0 and 4.0.7.1. Lastly the communication requests that the user "please acknowledge receipt of this task and grant permission to load the mandatory hot fix(es). Thank you." |  | 48 | 48 | Distributed US (nationwide) in the states of OH, GA, NJ, WI, FL, NH, IL, NY, IN, CA, AL, VA, MO, MI, KY, MA, PA, CO, ME, TN, WA, MN, SD, LA, NC, and NV and the country of Canada. | Open, Classified |  |  | 48 | Computer
Z-0401-2016 | 72617 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Revolution CT Scanners. |  |  | October 23, 2015 | 2015 | GE Medical Systems, LLC | GE Healthcare has recently become aware of a potential issue with routine head scans on your Revolution CT scanner. A potential hazardous situation can occur during a routine head scan with possible artifacts that may emulate pathology between the brain tissue and bone in the head images. No injuries have been reported to date related to this issue. | Software design | The firm, GE Healthcare, sent an"Urgent Medical Device Correction" letter GEHC Ref#25465 dated October 23, 2015 to its Consignees. The letter described the Safety Issues, Safety Instructions, Product Correction, Affected Product Details and Contact Information. The consignees were instructed to follow the instructions in existing CT User Manual and the safety instructions listed in the letter. GE will be delivering an addendum to the User Manual that reiterates these mitigations. If you have any questions or concerns regarding this notification, please contact GEHealthcare Service at 1-800-437-1171 or your local Service Representative. |  | 125 | 125 (US = 37; OUS = 88) | Worldwide Distribution: US (nationwide) including states of: AL, CA, CT, FL, GA, HI, IL, NY, NC, OH, OR, PA, TX, UT, VA, WA, WI; and countries of: AUSTRALIA, BELGIUM BRAZIL, CANADA, CHINA,DENMARK, FRANCE, GERMANY, HONG KONG, INDONESIA, ITALY, JAPAN, REPUBLIC OF KOREA, KUWAIT, NORWAY, PERU, POLAND, SAUDI ARABIA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, UNITED ARAB EMIRATES, and UNITED KINGDOM. | Open, Classified |  |  | 125 | Computer
Z-1039-2015 | 70062 | Yes | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Baxter Integrated APD Set with Cassette 3-Prong, Sterile, nonpyrogenic fluid path. |  |  | December 08, 2014 | 2014 | Baxter Healthcare Corp. | Baxter is issuing a recall for one lot of Integrated Automated Peritoneal Dialysis set with Cassette 3 prong due to complaints received for leakage. | Device Design | Baxter sent an Urgent Product Recall notification letter dated January 8, 2015 . The letter identified the affected product, problem and action to be taken. The recalling firm has asked the consignees to 1) locate and remove all affected product from their facilities, 2) contact Baxter to arrange for return and credit, and 3) complete the attached customer reply form and return it to Baxter. For questions call 800-422-9637. |  | 8640 | 8,640 units | US Nationwide Distribution in the states of FL, WI, AR, TN, TX, NC, MO, NM, NJ, CT, PA, NV, NY, CA, MI, WI, KS, and OH. | Terminated | October 28, 2015 | 324 | 8640 | Computer
Z-1047-2015 | 70288 | Yes | KOC | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius Crit Line in a Clip (CLiC) with SW version 2.51 Model Number: CL10041001. A continuous real-time monitor for non-invasive hematocrit, oxygen saturation and percent change in blood volume calculation during hemodialysis treatment. |  |  | December 19, 2014 | 2014 | Fresenius Medical Care Holdings, Inc. | Potential for misinterpretation of the graphic display of the Blood Volume ( BV) slope | Device Design | FMCRTG, LLC representatives contacted Clinics on 12/19/14 by visit/telephone phone) and removed all CLiC units. |  | 35 | 35 units | CT, NY | Completed |  |  | 35 | Computer
Z-0409-2016 | 69055 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance 64 Computed Tomography X-Ray System, Model number 728231, running version 2.6.0. Intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. This device may include signal analysis and display equipment, patient, and equipment supports, components and accessories. |  |  | March 03, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | Multiple software defects which could impact the performance or reliability of the system. These impacts include, but are not limited to: error messages which prevent operation, incorrect scan region, couch movement failure, inaccurate CTDI reporting, image artifacts, lost images. There is a risk of misdiagnosis. | Software design | On March 3, 2011, the firm sent Field Service Notices to their customers. |  | 1765 | 1,765 units | Worldwide Distribution. US nationwide, including Puerto Rico, Argentina, Australia, Austria, Azerbaijan, Bahrain, Bangladesh, Belgium, Brazil, Canada, Chile, China, Colombia, Croatia, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, Equatorial Guinea, Finland, France, Germany, Greece, Hungary, Iceland, India, Indonesia, Iran, Iraq, Ireland, Israel, Italy, Jamaica, Japan, Jordan, Kazakhstan, Kenya, Kuwait, Latvia, Lebanon, Libyan Arab Jamahiriya, Lithuania, Malaysia, Malta, Mexico, Morocco, Netherlands, New Zealand, Nicaragua, Norway, Pakistan, Palestine, Panama, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Serbia &amp; Montenegro, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uzbekistan, Venezuela, Vietnam &amp; Yemen. | Terminated | December 15, 2015 | 1748 | 1765 | Computer
Z-0399-2016 | 70931 | Yes | MSX | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | The eCareManager system. Software intended for use in data collection, storage and clinical information management with independent bedside devices, and ancillary systems that are connected either directly or through networks. For use in a hospital environment. eCareManager allows clinicians to generate a PDF display of an order (new, change, discontinue), which is printed at the bedside and/or hospital pharmacy. This printed PDF should be reviewed by the clinician and verified by the hospital pharmacist and entered into the hospital pharmacy system. |  |  | March 27, 2015 | 2015 | Visicu, Inc. | A software defect may cause incorrect medication order change. If the user decides to edit the order prior to electronic signature AND selects the return to Previous Screen function, the previously selected medication is replaced with whatever medication is in the number one spot on the complete medication list. At the stage where the order is signed, the medication displayed may be incorrect. | Software design | Philips Visicu sent an "Field Safety Notification" to all affected customers still using an affected version of eCareManager. (4 Customers) 03/27/2015 The letter identified the problem and circumstances in which the defect can occur; recommend actions the customer should take to prevent risks for patients; and list the actions planned by Philips to correct the problem."Philips Hospital to Home has taken corrective action to address the software defect in all supported versions of eCareManager software. Philips Hospital to Home is working with affected customers to complete the necessary upgrade.Please contact your local Philips Representative for further questions. |  | 43 | 43 | US distribution only. | Open, Classified |  |  | 43 | Computer
Z-0400-2016 | 69121 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance CT (computed tomography) scanners, Big Bore configuration, Model number 728243. software version 2.0 The products are intended to provide cross sectional images of the human body and visualization of the internal organs of the body. |  |  | April 25, 2006 | 2006 | Philips Medical Systems (Cleveland) Inc | Brilliance Big Bore- An anomaly was identified with software version 2.0 tumor localization application when printing work sheets. this causes information regarding relative locations of marked regions of interest on the patient work sheets to be incorrect and may cause mistreatment. | Software design | Philips Medical Systems sent a product safety notice to customers and corrective software for the issue was released to the field on July 25, 2006. Service Reps will contact each customer to install the new version of software. For questions regarding this recall call 440-483-7600. |  | 54 | 54 units | Worldwide distribution, including US states of CA and NV. | Terminated | December 15, 2015 | 3521 | 54 | Computer
Z-1049-2015 | 70365 | No | NFJ | Picture archiving and communications system. | Radiology | Ophthalmic | 510(k) | FORUM Archive and Viewer, version 3.1, v 3.1.1, (DVD Format) and v 3.2, v 3.2.1.(DVD and USB Flash Drive Format). |  |  | January 22, 2015 | 2015 | Carl Zeiss Meditec AG | Software defect in the FORUM Viewer versions 3.1 and 3.2 which may lead to misinterpretation of the optical coherence tomography (OCT) data. | Software in the Use Environment | Important Field Corrective Action for Forum Viewer letters were sent to all affected US customers on January 22, 2015 by certified mail. |  | 985 | 985 | Nationwide Distribution. | Terminated | March 04, 2016 | 407 | 985 | Computer
Z-1052-2015 | 70240 | No | OIY | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Persona (TASP) Tibial Articular Surface Provisional Shim, all sizes and Lots Catalog Number: 42-5279-002-00 thru 42-5279-009-04. |  |  | December 11, 2014 | 2014 | Zimmer, Inc. | Potential for delay in surgery and/or ball bearings to be left in the wound. Repeated cycles of sonic cleaning can break down the ball bearing retention feature (swage) on the shim. To prevent this, TASP shims must be manually cleaned and inspected per product insert, 87-6203-991-22. | Device Design | Zimmer sent an URGENT MEDICAL DEVICE CORRECTION notifications letter dated December 11, 2014 to affected distributors via electronic mail, and all hospital Risk Managers and distributors with affected inventory via courier. The letter identified the affected product, problem, instructions for responding to the formal recall notification.and actions to be taken. For questions call the customer call center at 1-800-348-2759. |  | 67963 | 67963 | Worldwide Distribution - USA Nationwide including DC, and PR and the countries of Australia, Austria, Belgium, Canada, Czech Republic, Dominican Republic, France, Germany, India, Iran, Israel, Italy, Japan, Luxembourg, Netherlands, Saudi Arabia, Singapore, South Africa, South Korea, Spain, Sweden, Switzerland, United Arab Emirates, and United Kingdom. | Terminated | October 23, 2015 | 316 | 67963 | Computer
Z-1055-2015 | 70079 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | BD 10 mL Syringe BD Luer-Lok Tip with Blunt Plastic Cannula, Not for Skin Injection ; Catalog #303348, |  |  | December 19, 2014 | 2014 | Becton Dickinson &amp; Company | The stopper of the 10mL BD Luer-Lok syringe is not properly seated in the syringe. | Device Design | BD sent an Urgent Product Advisory Notice letter dated 12/19/2014. The letter identified the affected product, problem and actions to be taken. Customers are advised of the issue with the product and should notify their customers. The attached Product Advisory Response Form should be completed and faxed to 1-866-551-2707. Questions can be directed to 1-866-367-4703. |  | 564000 | 564,000 units | US Nationwide Distribution | Terminated | December 07, 2015 | 353 | 564000 | Computer
Z-1058-2015 | 70239 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Puritan Bennett 980 Ventilator System, Model No. PB980 Ventilator (980xxxxxxxx), |  |  | January 12, 2015 | 2015 | Covidien LP (formerly Nellcor Puritan Bennett Inc.) | Covidien is issuing a voluntary field action for all Puritan Bennett 980 ventilators due to occasional GUI transient resets that last approximately 30 seconds. | Software design | Covidien sent an Urgent Field Safety Notice on January 12, 2015, to all affected customers. Customers were requested to forward the recall letter to anyone they may have transferred or assigned for use one of the potentially affected ventilators. Covidien is updating software that will modify ventilator alarms and status display messages during a GUI reset. The Puritan Bennett 980 Ventilator Operator Manual will be updated to delete,replace ventilator during a GUI transient reset, and to inform users to monitor the secondary status display until the primary display refreshes. Covidien will communicate with customers when the software and operator manual updates are available.Customers with questions were instructed to contact Technical Support Department at 800-255-6774, option 4, and then option 1.For questions regarding this recall call 303-876-8870. |  | 418 | 418 units | Worldwide Distribution - USA (nationwide) and Internationally to Canada. | Terminated | January 06, 2016 | 359 | 418 | Computer
Z-1062-2015 | 69863 | No | DSP | Intra-aortic balloon and control system. | Cardiovascular | Cardiovascular | 510(k) | CARDIOSAVE Intra-Aortic Balloon Pump and CS300 Intra-Aortic Balloon Pump; and, non-sterile, not for human use, SENSATION and SENSATION PLUS IABC demonstration units. Part Number 0998-00-0800-xx; 0998-00-3023-xx; 0998-UC-0800-xx; 0998-UC-3023-xx |  |  | December 16, 2014 | 2014 | Maquet Datascope Corp - Cardiac Assist Division | Ferrule Cap was not removed on some fiber-optic Sensation and Sensation Plus IABC demonstration units. If fiber-optic connector is inserted into the Maquet CS300 or CARDIOSAVE IABP, cap will prevent connection of unit and may become dislodged and remain in IABPs fiber-optic receptacle. | Device Design | Maquet Inc. sent a Product Recall Medical Device Field Correction letter/Field Correction response form, dated 12/16/2014 via Fed Ex. The letter identified the affected product and the reason for the recall. Users were asked to inspect their inventory for the affected units. If customers have any affected product, they are to follow the instructions provided. Additionally, users are to complete and return the attached Field Correction Response Form to acknowledge receipt of the letter and to indicate completion of the steps outlined. Completed forms should be returned by e-mailing a scanned copy to fieldactions@maquet.com or faxing a copy to 1-973-396-3607. Technical questions should be directed to the Technical Support Department at 1-800-777-4222. |  | 1202 | 838 units-US; 364 units - OUS | Worldwide Distribution | Terminated | June 09, 2017 | 906 | 1202 | Computer
Z-1067-2015 | 70157 | No | GWE | Evoked response photic stimulator. | Neurology | Neurology | 510(k) | EPWorks software used in the Protektor 32 |  |  | January 02, 2015 | 2015 | Natus Medical Incorporated | Software error occurs when using remote monitoring; if the remote user tries to stop the free run waveform group, the system will display a message informing the user that they do not have sufficient privilege. | Software design | Natus Medical Inc. DBA Excel-Tech Ltd (Xltek) issued medical device notification letters dated December 22, 2014 to their customers vial email and via courier service on January 2, 2015. Within the notice, customers are instructed are on to removal the previous installation and also on how to install the new program upgrade. Natus Medical Inc. has requested that the Previous Version Removed from Use Confirmation form be completed by each customer and returned. Customers who require assistance can contact Natus Technical service by calling 800-387-7516 or via email Oakville_Technical_Service@xltek.com or ots@natus.com. |  | 248 | 248 units | US Nationwide Distribution | Terminated | March 22, 2017 | 810 | 1176 | Computer
Z-1046-2015 | 70103 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ABL90 FLEX analyzer; Model number 393-090. A portable, automated analyzer that measures pH, blood gases, electrolytes, glucose, lactate, bilirubin and oximetry in whole blood. |  |  | December 08, 2014 | 2014 | Radiometer America Inc | The ABL90 analyzer does not always use the most recent calibration data to calculate patient results. This can in some cases lead to a biased patient result. | Software design | Radiometer sent a Field Safety Notice dated December 17, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. What you should do short term: " Please ensure that the operator always closes the inlet when prompted by the analyzer, as per the labeling. " In case the inlet, by mistake, is not closed as per the labeling, please ensure that the following is carried out before measuring a patient sample: " Close the inlet and wait for the analyzer to go to READY (approximately five minutes) " Open the inlet and then close again (to initiate a Rinse) " Wait for the analyzer to go to READY (approximately 1Âż minutes) " Please return the confirmation fax form no. 1 to your RADIOMETER representative when the above actions have been implemented Final Solution provided by RADIOMETER: A mechanism will be provided in an upgraded version of the analyzer software to eliminate the possibility of this error. The new software version will be installed by your local engineer when available.Customers were instructed to ensure that this letter is distributed to the final end-user. Customers with questions were instructed to contact Radiometer America Technical Support at 1-800-736-0600, opt 4. |  | 5002 | 5002 units | Worldwide Distribution -USA including CT, MN, FL, CA, PA, NY, KY, MA, WA, MI, TX, NC, IL, OH, SC, OK, NJ, MD, VA, UT, AL, NM, GA, LA, OR, IN, WI, AZ, ND, CO, MO, WV, ID, TN, VT, NE, IA, NH, MS, ME, AR and Internationally to Canada, Australia, Austria , China, Czech Republic, Denmark, Estonia, Lithuania, Romania, Slovenia, United Arab Emirates, Bolivia, Brazil, Indonesia, Israel, Kuwait, Malaysia, Mexico, Peru, Qatar, Saudi Arabia, Serbia and Montenegro, Syrian Arab Republic, Thailand, Vietnam, Finland, France, Germany, Hong Kong, Hungary, Ireland, Italy, Japan, Netherlands, New Zealand, Norway, Poland, Republic of Korea, Russia, Singapore, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, and United Kingdom. | Terminated | September 23, 2015 | 289 | 5002 | Computer
Z-1070-2015 | 70426 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Aquarius iNtuition Client Viewer. Findings Workflow module, RECIST 1.1: |  |  | January 21, 2015 | 2015 | TeraRecon, Inc. | Software anomaly related to RECIST1.1 target lesion evaluation criteria in Findings Workflow Module within the Aquarius iNtuition Client Viewer. | Software design | Medical Device Correction letters were sent to all affected users on January 21, 2015 by both e-mail and physical mail - certified mail for us customers and FEDEX for international customers. |  | 91 | 91 | Worldwide Distribution-US (nationwide) and the countries of Canada, Switzerland, Germany, France and Italy. | Terminated | September 30, 2015 | 252 | 91 | Computer
Z-1065-2015 | 69837 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | SoftPath Laboratory Information System. Versions 4.3.0.8, 4.3.0.9, 4.3.0.10, 4.3.0.11, 4.3.0.12, 4.3.0.14, 4.3.0.15, and 4.4.0.0 |  |  | November 26, 2014 | 2014 | SCC Soft Computer | Modifications to diagnostic text may be: 1) Saved to the database but not appear on the report sent to the physician; or 2) Documented on the report, but not saved to the database. | Software design | On 11/26/2014 SCC Soft Computer sent a Correction Communication to their consignees: Corrective Action: Versions 4.3.0.8 through 4.4.0.2 will have a mandatory hotfix created. Due to the amount of versions affected, the range of availability dates will be 11/24/2014 through 12/30/2014. Utility: A utility will be made available to be run on your Live system that will identify cases in which there is a disparity between the diagnosis text in the document and the database. The utility output will be provided for client review and validation. |  | 108 | 108 units | Distribution US Nationwide and CANADA. | Open, Classified |  |  | 108 | Computer
Z-0374-2016 | 70000 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Philips IntelliVue Information Center iX (release A.00, A.01, and A.02) are impacted by this issue: |  |  | December 19, 2014 | 2014 | Philips Medical Systems, Inc. | Sp02 and/or Non Invasive Blood Pressure (NBP) alarms may become disabled without visual notification | Software Design Change | The firm, Philips Healthcare, issued a "URGENT-Field Safety Notice" on 12/19/14 to inform customers of the issue, identify details of the units affected, instructs on actions to be taken by the customer and identify what action Philips plans to take to remedy the issue. Philips will provide a software correction free of charge.Philips is asking customers to follow the Action to be Taken by Customer/User section of the Field Safety Notice-To prevent this issue from occurring, customers/users should:1) Disable the Change Bed Label field in Patient Management. This is accomplished by changing the feature from Optional to Hidden2) All patient transfers must be performed using the Transfer key in the Manage Patient application.Remember that the most reliable method of patient monitoring combines close personal surveillance with correct operation of monitoring equipment.If you need any further information or support concerning this issue, please contact your local Philips Healthcare Customer Service representative or the Customer Care Solutions Center at (800) 722-9377 with questions or concerns about this Device Correction. |  | 10013 | 10,013 units | Worldwide Distribution: US (nationwide) and countries: AUSTRALIA, AUSTRIA, BELGIUM, BULGARIA, CANADA, DENMARK, ESTONIA, FINLAND, FRANCE, GERMANY, HUNGARY, ICELAND, INDIA, IRAQ, IRELAND, ISRAEL, ITALY, JAPAN, LUXEMBOURG, MOROCCO, NETHERLANDS, NEW ZEALAND, NIGERIA, NORWAY, POLAND, PORTUGAL, QATAR, ROMANIA, RUSSIAN FEDERATION, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, UNITED ARAB EMIRATES, and UNITED KINGDOM. | Open, Classified |  |  | 10013 | Computer
Z-0169-2016 | 72228 | No | HIB | Obstetric-gynecologic specialized manual instrument. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(K) Exempt | Welch Allyn KleenSpec Corded Illumination System, Vaginal Speculum Lighting System, Model Number 78x Series Illuminator, REF 901070, Catalog Numbers 78810, 78812, 78814 and 78816, Rx ONLY |  |  | October 02, 2015 | 2015 | Welch Allyn Inc | The firm became aware that the KleenSpecÂż 788 Corded Illuminator has the potential to be incorrectly connected to a commonly available AC power cord instead of the illuminator low voltage transformer power cord which is supplied with the illuminator. This may result in a shattered lamp bulb and a potential electric shock hazard. | Device Design | On 10/02/2015, Welch Allyn commenced mailing of the Urgent Medical Device Correction letters (dated 10/01/2015) via UPS Certified Next Day Air Saver Mail. The letters inform customers of the issue. To help customers avoid any such occurrences, Welch Allyn is providing supplementary User Precautions which re-emphasize safe use of this product and providing additional specificity to the warnings already on the labeling and packaging. |  | 194464 | Domestic: 179,401 units, Foreign: 15,063 units | Nationwide, Puerto Rico and foreign distribution to the following countries: Australia, Bahamas, Bahrain, Belgium, Bermuda, Cambodia, Canada, Czech Republic, China, Cyprus, Denmark, Egypt, France, Germany, Great Britain, Greece, Guam, Guatemala, Iceland, India, Israel, Japan, Kenya, Kuwait, Malaysia, Mexico, Netherlands, New Zealand, Norway, Oman, Panama, Peru, Philippines, Poland, Russia, Saudi Arabia, Serbia, Singapore, South Africa, Spain, Sri Lanka, Thailand, Turkey, United Arab Emirates. | Open, Classified |  |  | 194464 | Computer
Z-2267-2015 | 71583 | No | DQA | Oximeter. | Cardiovascular | Anesthesiology | 510(k) | OxiMax N-65 Handheld Pulse Oximeter. N65, N65-1, N65P, and N65P-1. For continuous or spot check monitoring of functional arterial oxygen saturation (SpO2) and pulse rate. |  |  | June 29, 2015 | 2015 | Covidien LP (formerly Nellcor Puritan Bennett Inc.) | Potential missing segments on the display that can result in misinterpretation of data. | Device Design | Metronic sent an Urgent Medical Device Correction letter dated June 29, 2015 to affected customers. The letter identified the adfected product, problem and actions to be taked. Customers where instructed to return the attached verification form. For questions contact your Medtronic representative. |  | 211426 | 211,426 units | Worldwide Distribution - US Nationwide including Canada. | Open, Classified |  |  | 317257 | Computer
Z-0178-2016 | 72282 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Ingenuity Flex Computed Tomography X-ray systems. Intended to produce cross-sectional images of the body. |  |  | September 14, 2015 | 2015 | Philips Medical Systems (Cleveland) Inc | The firm discovered Ring/Dot artifact due to X-ray measurement error. | Software design | On 9-23-2015 the firm sent URGENT Field Safety Notices to their customers. The letter identified the products, problem, and actions to be taken by the customers. |  | 67 | 67 units | Worldwide Distribution-US (nationwide) including DC and the states of AL, AR, AS, AZ, CA, CO, CT, FL, GA, HI, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, ND, NH, NJ, NV, NY, OH, OK, OR, SC, TN, TX, UT, VA, VT, WI &amp; WV, and the countries of Albania, Algeria, Australia, Austria, Bangladesh, Belgium, Brazil, Canada, China, Czech Republic, Denmark, Ecuador, Egypt, Ethiopia, Finland, France, Germany, Hong Kong, India, Indonesia, Iran, Iraq, Italy, Japan, Jordan, Kazakhstan, Latvia, Lebanon, Libya, Lithuania, Malaysia, Mexico, Mongolia, Myanmar, Netherlands, Nigeria, Norway, Oman, Panama, Philippines, Poland, Portugal, Romania, Russia, Rwanda, Samoa, Saudi Arabia, Singapore, South Korea, Spain, State of Palestine, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uzbekistan &amp; Viet Nam. | Open, Classified |  |  | 643 | Computer
Z-2067-2010 | 55868 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | GE Centricity Laboratory software; GE Healthcare Integrated IT Solutions, Barrington, IL 60010. The intended use: the Centricity Laboratory is used to facilitate the general clinical, anatomic pathology and cytology laboratory workflow such as order entry, results entry, instrument interfacing, results reporting and patient record retrieval. |  |  | May 26, 2010 | 2010 | Ge Healthcare It | There is a potential safety issue associated with the use of GE Centricity Laboratory software where results from interfaced external reference laboratories are not displayed or stored correctly. This happens when the test is set up in Centricity Laboratory as a decimal (numeric) type result and the test result sent from the reference lab includes a comma in the result field. | Software design | The firm, GE Healthcare, sent "Urgent Medical Device Correction" letters dated May 26, 2010 to GE Centricity Laboratory software customers on the same date. The letters advised the users of the patient safety issue involving results from interfaced external reference laboratories that are not displayed or stored correctly if an unexpected alpha character is encountered in a result message that is expected to have a numeric value. This happens when the test is set up in Centricity Laboratory as a decimal (numeric) type result and the result sent from the reference lab includes a comma in the result field. The letters provided the users with specific safety instructions to follow until the software is updated. The customers were requested to complete the enclosed customer reply form and fax the completed form to 847-620-2780. Any questions should be directed to the their GE Service Representative or the GE Helpdesk at 888-778-3375. |  | 37 | 37 units | Worldwide distribution: California, Georgia, Michigan, Minnesota, New York, Ohio, Oregon, Tennessee and Texas, and internationally to Australia, Canada, China, England, India, Malaysia, New Zealand, Qatar, Scotland and Singapore. | Terminated | July 15, 2011 | 415 | 37 | Computer
Z-2482-2015 | 71611 | No | JJX | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Contour Next Control Solution for Contour Next Blood Glucose Test System. |  |  | June 22, 2015 | 2015 | Bayer Healthcare, LLC | If the control solution bottle is not thoroughly shaken prior to use, the control test result may be higher than the control range and may not be recognized and marked as a control test by the blood glucose meter. Also, the control solution, used as a quality control test in blood glucose meter systems, may read outside the specified range too high or too low. | Device Design | On 6/22/2015, Bayer Healthcare began sending URGENT - MEDICAL DEVICE FIELD CORRECTIVE ACTION notifications to affected Health Care Providers, Pharmacies, and Users. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Customers are instructed to always check the expiration, not to use and discard of expired control solutions. Also, only CONTOUR NEXT control solution should be used with the CONTOUR NEXT test strips. The control solution bottle should be shaken about 15 times before each use. The letter instructed customers to refer to their meter's user guide for further testing instructions. If test results outside the control range are received, another test should be performed following the instructions provided in the letter. If the problem persists, customers should contact Bayer Diabetes Support. Bayer will replace the control solution or provide new control ranges for the test strips, depending on the situation. The firm has implemented improvements to provide more consistent control solution product. |  | 5000000 | 5 million bottles of control solution | Worldwide Distribution -- USA, Australia, Austria, Belgium, Canada, Croatia, Denmark, Finland, France, Germany, Greece, Ireland, Italy, Luxembourg, Netherlands, Norway, Portugal, Slovenia, Spain, Sweden, Switzerland, UK, Albania, Azerbaijan, Bulgaria, Czech Republic, Estonia, Hungary, Korea, Latvia, Lithuania, Philippines, Romania, Serbia, Slovakia, and Thailand. | Open, Classified |  |  | 5000000 | Computer
Z-2479-2015 | 71840 | No | KHP | Lactic acid test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | RAPIDPoint 500 Blood Gas Analyzer, |  |  | July 28, 2015 | 2015 | Siemens Healthcare Diagnostics Inc | When both ports (Serial and Ethernet) are configured to transmit data, the data stream from one port could potentially affect the data stream from the other port causing the data management system to include either: duplicate data, missing data, or data from a different patient. | Software design | Siemens Healthcare issued Urgent Field Safety Notice sent by e-mail to all affected Siemens Healthcare Diagnostics Regional offices on July 28, 2015, both in the United States and outside of the United States, for communication with affected customers. This notice informs customers of the issue with the RAPIDPoint 500 dual port LIS transmission feature.Users are informed to: If the system is currently set up with dual LIS transmission enabled, one of the ports (either Serial or Ethernet) needs to be disabled. Instruction provided with the notification. Customer to complete Field Effectiveness Check Form to confirm receipt of the notice.The resolution of this issue will be implemented in the next software release for the RAPIDPoint 500 system.If you have any questions, contact Siemens Customer Care Center or your local Siemens technical support representative. Please fax this completed Customer Care form to (312) 275-7795. |  | 5241 | 5241 units | Worldwide Distribution - US (nationwide) Distribution and to the countries of : Albania, Algeria, Armenia, Australia, Austria, Bahrain, Bangladesh, Belgium, Bosnia Herzeg, Brazil, Bulgaria,Canada, Chile, Colombia, Croatia, Czech Republic, Denmark, Egypt, Estonia, Finland, France, Georgia, Germany,Greece, Hong Kong, Hungary, India, Indonesia, Ireland, Italy, Japan, Kenya, Kuwait, Latvia, Lebanon, Lithuania, Luxembourg, Malaysia, Mexico, Netherlands. New Zealand, Norway, P.R. China, Peru, Poland, Portugal, Qatar, Republic Korea, Romania, Russian Fed, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Tunisia, Turkey, U.A.E., United Kingdom and Vietnam. | Open, Classified |  |  | 5241 | Computer
Z-2477-2015 | 71691 | No | HKO | Ophthalmic refractometer. | Ophthalmic | Ophthalmic | 510(K) Exempt | OPD-Scan III Refractive Power/Corneal Analyzer Opthalmic |  |  | June 15, 2015 | 2015 | Nidek Inc | Software bug was found where there was no difference in Total and Corneal high-order aberrations, but differences were found in Internal high-order aberrations. | Software design | Dear Doctor letters mailing started on July 15, 2015 - first 100. They will be mailed out as the correction is received from Nidek, Japan.Dear Doctor letter will include software upgrade kit including CD or USB, instruction for upgrade and pre-paid response card to be returned to Nidek Inc. |  | 3836 | 3836 | Nationwide Distribution. | Terminated | April 25, 2016 | 315 | 3836 | Computer
Z-2470-2015 | 71964 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Centricity Universal Viewer |  |  | August 11, 2015 | 2015 | GE Healthcare | Images from the Centricity PACS-IW with Universal Viewer andCentricity Universal Viewer may be missing when a system parameter MapRoute is set to a value greater than 1. | Software design (manufacturing process) | The firm sent out an Urgent Medical Device Correction letter dated 8/11/2015 to all affected customers. Letter states that No User activity required. GE Healthcare has completed a remote inspection of your system and found that this system parameter is currently set to the default setting of 1. The firm will correct all affected systems by a providing software update.The letter also states that if the customer has any questions they should call GE Healthcare Service at 1-800-437-1171 or their local service representative. |  | 26 | 26 systems | Worldwide Distribution - U.S. nationwide in the states of: AK, CA, CO, DE, FL, HI, ID, NY, PA, SD, TN and TX; France, Germany, Greece, India, Israel, Malaysia, Mexico, Northern Ireland (UK), Panama, Poland, Russia, South Africa, Saudi Arabia, Switzerland, Turkey, UAE and Uruguay. | Terminated | January 05, 2017 | 513 | 108 | Computer
Z-2475-2015 | 71834 | No | JDQ | Bone fixation cerclage. | Orthopedic | Orthopedic | 510(k) | Synthes Application Instrument for Sternal ZIPFIX; indications for use include primary or secondary closure/repair of the sternum following sternotomy or fracture of the sternum to stabilize the sternum and promote fusion |  |  | July 22, 2015 | 2015 | Synthes (USA) Products LLC | The end cap may loosen and detach making the instrument non-functional. No injuries reported. | Device Design | Urgent recall notification letters, dated July 22, 2015 were sent to end users and consultants to inform them of the issue, the associated risks, and provided instructions to return affected devices along with the response form. |  | 451 | 451 | Nationwide and internationally to Canada. | Open, Classified |  |  | 451 | Computer
Z-0541-2011 | 52039 | No | KLC | Dental chair and accessories. | Dental | Dental | 510(k) | A-dec Decade Plus 1221 Dental Chair |  |  | April 01, 2009 | 2009 | A-Dec Inc | The design of the Decade Plus 1221 chair backrest creates a possible pinch-point between the chair backrest and the underside of the armrests. | Device Design | A-dec sent a letter dated March 23, 2009 to consignees/customers (doctors and network dealers). The consignees were informed that the design of the Decade Plus 1221 chair backrest creates a possible pinch-point between the chair backrest and the underside of the armrests. The potential for this occurs when returning the chair to the upright position. The consignees were advised to reprogram the Position 0/Entry-Exit preset to stop the back before it reaches a full upright position. This increases the gap between the armrest and the back slightly. To reprogram a chair, the firm included excerpts from a chair Instructions for Use (IfU) guide that detail the steps to reprogram a chair. Consignees can contact the firm's Customer Service at 503-538-9471, 1-800-547-1883, or via fax at 503-538-0276.The firm will replace the chair back assembly for all affected dental chairs. |  | 2723 | 2723 units | Worldwide distribution: US and Canada. | Terminated | December 16, 2010 | 624 | 2723 | Computer
Z-1846-2015 | 71359 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Artis zee/ zeego systems; |  |  | May 22, 2015 | 2015 | Siemens Medical Solutions USA, Inc | There exists a possible position sensor fault in the swivel base axis not being detected by the system software. When moving, the system could potentially exceed its usual speed, resulting in a collision. No injuries reported. | Software design | Siemens sent an Important Customer Safety Notice letter dated May 22, 2015, to affected customers. The letter identified the affected product, problem and actions to be taken. The letter also informed customers that Siemens would contact them to schedule the installation of a software update to resolve the problem. |  | 475 | 475 | US Nationwide Distribution | Open, Classified |  |  | 475 | Computer
Z-2374-2015 | 71551 | No | PDR |  |  | Anesthesiology | PMA | SEDASYS Computer-Assisted Personalized Sedation System, SEDPRU01. |  |  | May 13, 2015 | 2015 | Ethicon Endo-Surgery Inc | Ethicon has found that disinfecting methods not specified in the Instructions For Use (IFU) have led to the degradation of the plastic material of the SEDASYS System Control Unit. | Device Design | Sedasys, a Division of Ethicon Endo-Surgery, Inc. sent an Urgent Customer Notification letter dated May 27, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to do the following:1. Please distribute this information to all staff within your facility who use the SEDASYS System2. Complete the Business Reply Form (BRF) (Attachment A) confirming receipt of this notice within three (3) business days. The BRF may be given to your sales or service representative or sent to Sedasys by email at SedasysTechSupport@its.jnj.com.3. Keep this notice visibly posted in your facility for awareness.4. If you have any questions or concerns regarding the condition of your SEDASYS System including the Control Units, please contact your Sales or Service Representative.If you need clinical or product support, please contact your local sales representative or call our Customer Support Center at 1-800-SEDASYS (1-800-733-2797). If you need additional communications packages, contact 1-800-SEDASYS and reference Event 052715SED01.For questions regarding this recall call 513-337-7439. |  | 14 | 14 units | Nationwide Distribution including OH, RI, TX &amp; WA. | Terminated | May 02, 2016 | 355 | 14 | Computer
Z-2467-2015 | 71829 | No | IOR | Mechanical wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Quickie IRIS Wheelchair. model EIZ4-2 in combination with option code 183M02 - MONO Backrest system with Dynamic Backrest option. |  |  | July 27, 2015 | 2015 | Sunrise Medical (US) LLC | Quickie IRIS and Zippie IRIS tilt-in-space wheelchairs with MONO Backrest System with Dynamic Backrest may break over time resulting in a fall or injury to occupant. | Device Design | Sunrise Medical sent an Urgent Medical Device Field Correction letter dated July 27, 2015, to all affected dealers. The letter identified the product, the problem, and the action to be taken by the dealer. Dealers were instructed to immediately contact their customers to make arrangements to have the correction made with a replacement kit which will be supplied by Sunrise Medical to each dealer. Each dealer is requested to send back to Sunrise Medical the Acknowledgment and Response Form(s) by fax, email or regular mail once the work is completed. Customers with questions were instructed to contact Sunrise Medical Regulatory Affairs at (888) 208-4901. |  | 64 | 64 total wheelchairs, both models. | Worldwide Distribution - US (nationwide) and Internationally to Australia, Germany and Canada. | Terminated | December 10, 2015 | 136 | 64 | Computer
Z-2060-2010 | 55224 | Yes | DWC | Cardiopulmonary bypass temperature controller. | Cardiovascular | Cardiovascular | 510(k) | Terumo HX2 Temperature Management System 120V, 50/60 Hz catalog number 809810, Terumo Cardiovascular Systems Corp., Ann Arbor MI. |  |  | June 11, 2010 | 2010 | Terumo Cardiovascular Systems Corporation | The HX2 Temperature Management System mixing valve may fail resulting in an unusable water channel and then the machine would go into standby mode. | Device Design | The firm, TERUMO Cardiovascular Systems, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated June 11, 2010, to the consignees. The letter describes the product, reason for correction, what to do in the event of failure, potential hazard, correction, affected population and action to be taken by the customers. The customers are instructed to review the correction notice, assure that all users are aware of this notice, and complete, sign and fax the attached Customer Response Form to Terumo CVS Customer Service at 800-292-6551. The Terumo CVS Customer Service Dept. will contact you upon receipt of the Customer Response Form to schedule the replacement. The firm plans to inspect and correct all units in the field within 5 months.We encourage you to contact Terumo CVS Customer Service at 1-800-521-2818, Monday-Friday, 8 AM-6 PM EST, with any questions or concerns. |  | 100 | 100 | Nationwide Distribution: AR, CA, CO, CT, FL, HI, IA, IL, KS, LA, MD, MI, MO, MS, NY, OH, PA, TX, UT, WA and WI. | Terminated | April 13, 2011 | 306 | 100 | Computer
Z-2419-2015 | 71777 | No | MEC | Medical washer-disinfector. | General Hospital | General Hospital | 510(K) Exempt | Getinge 86 Series Washer/Disinfector with G1 control system software, Model 8666 and Model 8668. For use in healthcare facilities and laboratories where reusable items such as instruments, utensils, anesthesia sets, and glassware are handled for decontamination purposes. |  |  | June 26, 2015 | 2015 | Getinge Disinfection Ab | Getinge Disinfection AB has initiated a Field Correction due to the potential for unauthenticated remote access to the product over the network. | Software Manufacturing/Software Deployment | Getinge Disinfection AB has initiated a Field Correction due to unauthenticated remote access to the product over the network. |  | 20 | 20 devices | Distributed in CA, CT, GA, IA, ID, MD, MI, MN, NY, PA, TX, and WA. | Open, Classified |  |  | 20 | Computer
Z-2442-2015 | 71816 | No | MDS |  |  | Clinical Chemistry | PMA | Guardian Real-Time Monitor, Model No. CSS7100. Indicated for continuous or periodic monitoring of glucose levels in the fluid under the skin, in adults (18 years and older) with diabetes mellitus, for the purpose of improving diabetes management. |  |  | July 22, 2015 | 2015 | Medtronic MiniMed Inc. | Medtronic MiniMed is recalling the Guardian Real-Time Continuous Glucose Monitoring System because it has a language translation error that impacts the Finnish language. Specifically, the Monitor has a Finnish translation error in Predictive Alerts setting: the on-screen Finnish translation for LOW/HIGH is reversed as HIGH/LOW. | Software design | A Urgent Medical Device recall letter was sent to customers on 7/22/15 to inform them that the Guardian Real-Time Continuous Glucose Monitoring System System has a language translation error that impacts the Finnish language. The letter informs the customers of the problems identified and the actions to be taken. Customers with questions or concerns are instructed to contact Nordic Support Line 0-800-164-064. |  | 254 | 254 units | Distributed in Finland. | Terminated | September 18, 2015 | 58 | 254 | Computer
Z-2457-2015 | 71414 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | RayStation 4.7, version 4.7.1, used with Structure Definition or Automatic Breast Planning. Radiation Therapy Treatment Planning System, for treatment planning and analysis of radiation therapy. |  |  | May 27, 2015 | 2015 | RAYSEARCH LABORATORIES AB | Issue with the algorithm for ROI contraction in RayStation 4.7 when non-uniform contraction distances are used. The ROI contraction tool uses six distances as input: right/left, inferior/superior and posterior/anterior. These contraction distances are, in error, pairwise reversed, i.e. right is interpreted as left, inferior is interpreted as superior, and posterior is interpreted as anterior. | Software design | On May 27, 2015 RaySearch Laboratories distributed Medical Device Correction notices to their customers via courier service. Within the Medical Device Correction notice, RaySearch Laboratories included a workaround to the problem for their customers. Customers should not use structure templates containing derived ROI definitions with non-uniform contractions. They should avoid using non-uniform contractions. In the case non-uniform contractions must be used, enter switched distances manually and review the result carefully before using it for treatment planning. Lastly, do not use humerus shielding in Automatic Breast Planning. Customers are also expected to complete and return the lower portion of the correction notice via fax to 888-501-7195. **CAPA** Correct the software problem in patch version 4.7.2 of RayStation 4.7, scheduled for market release June 2015. |  | 126 | 126 programs | Distributed in CO, IL, MI, NC, OH,TN, and WA. | Open, Classified |  |  | 126 | Computer
Z-0492-2011 | 56405 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris PC unit Model 8000 only when used with Alaris Pump module or Alaris Syringe module |  |  | August 02, 2010 | 2010 | CareFusion Corporation | Than recall was initiated because CareFusion has identified identified a potential risk associated with the Alaris PC unit formerly known as Medley PC unit (PC Unit) model 8000 only when used with Alaris Pump module (Pump module) or Alaris Syringe module' (Syringe module), This recallnotification details the potential risk and recommended steps for users to take if they encounter a specific PC | Device Design | CareFusion will send by registered return receipt mail the customer notification letter to all customers usingthe affected Alaris PC Unit beginning August 24, 2010. The customer notification letter will be addressed tothe Directors of Nursing, Risk Management, and Biomedical Engineering. Customers will be required toconfirm receipt of the notification by returning the Recall Response Card to CareFusion by postage-paid, selfaddressedmail, fax or email. The recall notice informed the customers of the reason for the notice, the unitsaffected, the potential risk, a step-by-step description of the issue and the required actions for users.The customers were also told that CareFusion is working to release a software update to address this display error and will contact your facility as soon as the software update is available. CareFusion does not require that you return your devices. Clinicians should weigh the risk/benefits to the patients beforecontinuing the use of the device.Customers were also told that any adverse reactions experienced with the use of this product, and/or quality problems shouldalso be reported to the FDA's MedWatch Program. The recall notice provided contact information for CareFusion's Recall Center, Customer Advocacy and Technical Support units. |  | 94146 | 94,146 units | Nationwide in the United States, Canada, Mexico, and the European Union | Terminated | July 25, 2013 | 1088 | 94146 | Computer
Z-2459-2015 | 71708 | No | MED | Liquid chemical sterilants/high level disinfectants. | General Hospital | General Hospital | 510(k) | CIDEX Activated Dialdehyde Solution, 4 bottles/case, P/N: 2266 |  |  | July 27, 2015 | 2015 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) is recalling the CIDEX Activated Dialdehyde Solution because it may lack the adequate data to support the entire 24-month shelf life. | Device Design | A Urgent Product Notification Letter dated July 27, 2015, will be sent to customers and distributor who purchased the CIDEX Activated Dialdehyde Solution because it may lack the adequate data to support the entire 24 month shelf life. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to complete the requested information on the enclosed postage paid Business Reply Card and return it to Stericycle. Customers with questions or need additional packing labels are instructed to contact Stericycle at (877) 550-0701, M-F, 8am-5pm ET. Customers are instructed to report complaints or suspected problems with CIDEX Activated Dialdehyde Solution by contacting (888) 783-7723. |  | 23426 | 23,426 cases | US (nationwide) Distribution | Open, Classified |  |  | 23426 | Computer
Z-2455-2015 | 71470 | No | FZX | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Offset Femoral Alignment Guide 4mm x 6Âż; Product Code: KS 67024 |  |  | June 09, 2015 | 2015 | Omnilife Science Inc. | The markings on both versions of the OFFSET FEMORAL ALIGNMENT GUIDES are incorrect. Specifically, the markings on the alignment guides are reversed from the intended position that should provide proper alignment for the offset revision stems when used with the femur. | Device Design | Omni sent an Urgent Medical Device -Field Safety Notice letters to their customers. The letter identifed the affected product, problem and actions to be taken. Customers were instructed to Return the attached Acknowledgement and Receipt Form by FAX to 508-819-3390. |  | 15 | 15 | Worldwide Distribution - US Nationwide in the states of OK, CO, OH, MA and the countries of Spain and Australia. | Terminated | September 28, 2015 | 111 | 15 | Computer
Z-1728-2011 | 57814 | No | JJQ | Colorimeter, photometer, or spectrophotometer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Hc2 System Software Suite 4.0 version 3.0, available as component in Qiagen user guide, catalog #5050-00162, and as a component within a Qiagen PC, available as catalog # 5050-11140A. |  |  | January 12, 2011 | 2011 | QIAGEN Gaithersburg, Inc. | The software component of a diagnostic medical device may be defective and could cause incorrect patient values. | Software design | The firm, Qiagen, sent a "Notice of Product Field Action" letter dated January 12, 2011 to all customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately discontinue the process of running multiple assay protocols on a single microplate to eliminate the potential for inaccurate patient result reporting, and to complete and return the acknowledgment section via fax to QIAGEN 661-702-3854 or scan the document and email to techservice-na@qiagen.com. The firm informed users that they will provide further detailed instructions for a workaround and software fix. A QIAGEN representative will contact the customers who indicated that they run multiple assays on a plate and are currently using the HC2 System Software Version 3.0 (Suite 4.0), to assist in the identification of those patients that may have been affected by the software defect.If you have any questions or concerns, please contact Qiagen Technical Services at 800-344-3631, Option 2. |  | 163 | 163 | The recalled product was sold to medical laboratories nationwide and in Canada. | Terminated | November 14, 2011 | 306 | 163 | Computer
Z-0014-2012 | 59949 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Dynamic Multileaf Collimator (DMLC) |  |  | May 25, 2011 | 2011 | Elekta, Inc. | It is possible to load different patients in the record and verify (R&V;) system and in the multilead control software (MCS). | Device Design | Elekta, Inc. sent an "IMPORTANT NOTICE" letter dated May 25, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. The letter recommends several precautions customers should use until Elekta finds a solution. Customers will be contacted when a solution is found. |  | 27 | 27 | Nationwide Distribution-USA including the state of CT, FL, GA, IA, MN, MO, NJ, NY, ND, OH, OK, PA, RI, TX, VA , and WA. | Terminated | January 14, 2014 | 965 | 27 | Computer
Z-0336-2011 | 56939 | No | DRC | Trocar. | Cardiovascular | Cardiovascular | 510(K) Exempt | ENDOPATHÂż XCEL" with OPTIVIEW" Technology trocar |  |  | September 27, 2010 | 2010 | Ethicon Endo-Surgery Inc | This recall has been initiated because Ethicon Endo-Surgery is aware of an increase in reports of leaking in the ENDOPATHÂż XCEL" with OPTIVIEW" Technology trocar, which potentially can lead to loss of pneumoperitoneum during a surgical procedure. Pneumoperitoneum leaks may lead to conversion from laparoscopic to open surgery. | Device Design | Ethicon Endo-Surgery sent an URGENT: DEVICE RECALL EVENT 2244 letter, dated October 8, 2010. The letter identified the product, the problem, and the action to be taken by the customer. The customers were instructed to stop using the product, examine all inventory to determine if they had the affected product on hand, and remove the affected products from inventory. Ethicon Endo-Surgery requested that each customer complete the Business Reply Form and fax it back to Stericycle at (877) 272-6141 with in 3 business days and also return any affected product to Stericycle.For question regarding this recall call (877) 598-5706 and reference 2244. |  | 671 | 671 lots totalling 2,479,134 instruments | Worldwide distribution: USA, ARGENTINA, AUSTRALIA,BELGIUM,BRAZIL,CANADA,CHILE,CHINA,COLOMBIA,COSTA RICA, DOMINICAN REPUBLIC,ECUADOR, EL SALVADOR, GREECE, GUATEMALA, HONG KONG, ISRAEL, ITALY, JAPAN,JORDAN, KOREA, REPUBLIC OF, LEBANON, NETHERLANDS ANTILLES, NICARAGUA, PANAMASAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICASPAIN, TRINIDAD AND TOBAGO, UNITED ARAB EMIRATESURUGUAY, and VENEZUELA. | Terminated | January 25, 2012 | 485 | 671 | Computer
Z-2449-2015 | 71848 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | CARESTREAM DRX-EVOLUTION X-Ray System |  |  | August 07, 2015 | 2015 | Carestream Health Inc | Software defect that impacts image alignment when using the automatic stitching option for long length images. The defect could cause stitching inconsistencies which may go undetected on a radiograph. | Software design | Carestream sent an Urgent Medical Device Correction letters dated July 29, 2015 and Customer Notification Acknowledgement Forms to customers via Certified Mail. The letter identified the affected product, problem and actions to be taken. Carestream will update the software on your device to correct this issue. For questions or concerns, contact the Carestream Customer Care Center in the U.S. at 1-800-328-2910, available 7 days per week on a 24 hour basis. Outside of the U.S., please call your local Service support number. |  | 236 | Domestic: 76 units, Foreign: 160 units | Worldwide Distribution - US Nationwide and in the countries: Australia, Canada, Chile, China, India, Korea, Malaysia, New Zealand, Singapore, Switzerland, Taiwan and Thailand. | Open, Classified |  |  | 236 | Computer
Z-3163-2011 | 59427 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | TomoMobile, Hi-Art System, H-0000-0003, TomoTherapy 1240 Deming Way, Madison, WI 53717 |  |  | July 07, 2011 | 2011 | Accuray Incorporated | TomoTherapy Inc. is sending this Field Safety Notice to make you aware of an anomaly which may affect the performance of the TomoMobile Hi-Art System.TomoTherapy has discovered that when attempting to open the TomoMobile shielding doors, while the door hinges are in the locked position, the hinges may fail allowing the door to disengage from the shielding. | Device Design | The firm, TomoTherapy, sent an "Urgent Medical Device Correction" letter dated July 7, 2011 to its consignees/customers. The letter described the Issue, Product Affected, Recommended Action (advised consignees to not attempt to open the shielding doors when the hinges are in the locked position. The Hinges should always be in the unlocked position when opening the doors) and Resolution. A TomoTherapy field service engineer or an authorized service representative will contact the customers when a resolution is available.If you have any questions, please contact the TomoTherapy Technical Solutions Center by email at support@tomotherapy.com or phone at: North America: 1-866-368-4807; Belgium: 0800 38783; France: 0805 631 565; Germany: 0800 000 7541; Italy: 800 986 399; Netherlands: 0800 020 1364; Spain: 800 300049; Switzerland: 0800 001927; United Kingdom: 0808 238 6035; China/CNC (Northern): 10 800 712 1701; China (Southern): 10 800 120 1701; Hong Kong: 800 967912; Japan: 0044 22 132374; Singapore: 800 1204 683; South Korea: 0079 81 4800 7204, all other locations: +1 608 824 2900 or +32 2 400 44 44. |  | 4 | 4 | Worldwide distribution: USA including states of: AZ, CA and OK; and country of: United Arab Emirates | Terminated | June 29, 2012 | 358 | 4 | Computer
Z-2440-2015 | 71898 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ExacTrac 6.0.x |  |  | August 05, 2015 | 2015 | Brainlab AG | ExacTrac 6.0 Patient Positioning System:Display of potentially incorrect Digitally Reconstructed Radiograph (DRR) for x-ray correction andverification. | Software in the Use Environment | Brainlab intends to distribute the Field Safety Notice / Product Notification letter CAPA-20150713-001446, dated July 20, 2015, to the customers. The letter provides instructions for the user on how to address this issue, and to avoid a patient and treatment target position potentially different than intended due to this issue.The Field Safety Notice / Product Notification letter states that as generally required, always verify that the correct DRRs for the current patient are displayed in ExacTrac and an accurate fusion to the X-Ray images is possible for patient positioning. In case incorrect DRRs are displayed, exit and re-start the ExacTrac software. To minimize the probability of this issue occurring for the affected ExacTrac versions: -Reduce the amount of contours exported to ExacTrac as far as possible.-Avoid exporting large contours, such as couch top models, to ExacTrac.-After every (re-)start of the ExacTrac application on any workstation, always first load a phantom plan not containing contours. The Field Safety Notice / Product Notification letter states that as a general reminder the user should always follow the instructions and warnings as described in the user guide. In the context of this notification please specifically consider the safety notes relevant for X-ray correction and verification:-The image overlay functions must be used to verify fusion accuracy in both image views, especially if the images contain a series of similar structures such as vertebrae.-Do not perform patient treatment unless accurate image fusion is possible.-Large correction shifts (e.g. greater than 10 mm) may indicate incorrect automatic fusion. In such cases, verify the correction shift a second time.Brainlab Corrective Action:1. Existing potentially affected ExacTrac 6.0.x customers receive the product notification information.2. Brainlab will provide a software revision (ExacTrac v. 6.0.6) with this issue solved to affe |  | 361 | 361 systems | Worldwide Distribution - USA including the states of AL, AZ, AK, CA CO, CT, FL, GA, ID, IL IN, KY, LA, MD, MA, MI, MN, NE, NY, NC, ND, OH, OK, OR, PA, RI, SC, TN, TX, VA, WA, and WI; and, the countries of Australia, Austria, Belgium, Brazil, Canada, China, Denmark, Finland, France, Germany, Hong Kong, India, Israel, Italy, Japan, Netherlands, Panama, Qatar, Russia, Slovenia, South Africa, South Korea, Spain, Switzerland, Taiwan, Thailand, Turkey, and United Kingdom. | Terminated | January 05, 2017 | 519 | 361 | Computer
Z-1653-2008 | 46686 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ABL700 Series analyzer - Human Blood Analyzer (Laboratory Device). |  |  | January 18, 2008 | 2008 | Radiometer America Inc | Reported calibration errors are not displayed on the analyzers parameter bar screen or on the patient results when the calibration number is between 32768 - 65535 and 98303 - 131070. This may cause the operator to perform measurements and obtain questionable results. | Software design | On February 28, 2008, a Field Action Notes Letter was issued advising them about a Mandatory Software Update. Customers were notified of the problem by fax and letter (Fed Ex). The Recall Letter instructed the user to disable the parameter bar feature and only rely on the traffic light system indicator which properly indicates system errors. The letter includes instructions to temporarily fix the problem and notice of a update of the software module that updates the parameter bar. If you have any questions contact Vincent Sigmund at (440)-871-8900 ext. 209. |  | 2481 | 2481 Units | Worldwide Distribution - USA and Canada - Products were sold in all fifty (50) US States. including twenty-nine 29 government accounts and 138 Canadian accounts. | Terminated | February 24, 2012 | 1498 | 2481 | Computer
Z-0327-03 | 4680 | [6 empty] | Alumina V40 Head |  |  | June 26, 2002 | 2002 | HOWMEDICA OSTEONICS CORP. | V40 Alumina Heads may not assemble properly onto V40 hip stems | Device Design | On June 26, 2002, Howmedica sent out recall letters to their customers instructing them to remove and quarantine affected product from their inventory. |  | 190 | 190 (+676 28mm heads) | Howmedica distributed this product throughout the nation to various hospitals. There were 4 international custmers in United Kingdom, Australia, France, and Italy. | Terminated | June 29, 2010 | 2925 | 190 | Computer
Z-1831-2015 | 71385 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | SERFAS 90 degree Energy Probe, Part Number 279-350-101; |  |  | June 03, 2015 | 2015 | Stryker Endoscopy | Stryker Endoscopy is recalling all non expired SERFAS 90 degree Energy Probes due to reports of fragments of the probe breaking off into the patient. | Device Design | Stryker sent an Urgent Medical Device Recall letter dated June 5, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to do the following:1. Inform individuals within their organization who need to be aware of this device removal.2. Review inventory of lots of part number 279-350-101 and determine if you have the affected product (all non-expired devices) in stock. Response is required.3. If no product is found, complete acknowledgement form located on the Stryker Endoscopy recall website endorecall.stryker.com by logging in using the account number and zip code on this letter.4. If you do have product, segregate the product and call Stryker customer service at 1-800-624-4422 Option 3 to arrange for product return and issuance of credit.For questions regarding this recall call 408-754-2664. |  | 22063 | 22,063 devices | Worldwide Distribution - US (nationwide) and Internationally to US Argentina, Australia, Bolivia, Brazil, Chile, China, Colombia, Hong Kong, India, Italy, Japan, Republic of Korea, Malaysia, Netherlands, Poland, Romania, Singapore, South Africa, Spain, Sweden, Switzerland, United Arab Emirates, and United Kingdom. | Terminated | February 10, 2016 | 252 | 22063 | Computer
Z-2491-2015 | 71882 | No | OUO | Emission computed tomography system. | Radiology | Radiology | 510(k) | Ingenuity TF PET/MR, Diagnostic Imaging System , Philips Medical Systems |  |  | July 23, 2015 | 2015 | Philips Medical Systems (Cleveland) Inc | Philips Healthcare has become aware of a problem in which the Ingenuity TF PET/MR PET Reconstruction Server (PRS) database may lock up after an MR acquisition is completed and before a PET acquisition is begun, even though no actual database lockup issue has been reported in the field-installed Ingenuity TF PET/MR systems. | Software design | On July 23, 2015 the firm sent Field Safety Notices to their customers. |  | 11 | 11 | Product was shipped to the following states: NY, OH &amp; TX. Product was also shipped to the following countries: Finland, Germany, Hong Kong, Japan, Netherlands, South Korea, Spain &amp; Switzerland. | Open, Classified |  |  | 11 | Computer
Z-2742-2015 | 71937 | No | DRY | Cardiopulmonary bypass on-line blood gas monitor. | Cardiovascular | Cardiovascular | 510(k) | CDI 500 Blood Parameter Monitoring System. |  |  | August 07, 2015 | 2015 | Terumo Cardiovascular Systems Corporation | Inaccuracies in SvO2, temperature, pH, pCO2, pO2, Hematocrit, and Potassium readings following a software upgrade to version 1.69. | Software design | A voluntary Urgent Medical Device Correction notice that clearly explains initial in-vivo calibration requirements, device operating ranges, and temperature measurements was sent on 08/17/2015, via express mail to consignees of CDI System 500 v1.69. Following the initial notice, Terumo will be updating the Operators Manual and will send the new manual to each consignee when available. One manual per unit at each facility will be provided. Customers with questions and return response forms may contact: Terumo Recall Email: tcvs.recall@terumomedical.comTerumo Recall Fax: 734-741-6149Terumo CVS Customer Service: 1.800.521.2818, Monday - Friday, 8 a.m.  6 p.m. ET. |  | 4638 | 4638 | Worldwide Distribution-US (nationwide) including DC and PR and the states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, WV, and the countries of Australia, Belgium, Canada, Chile, Colombia, Hong Kong, Indonesia, Japan, Malaysia, Mexico, Philippines, Singapore, South Korea, Taiwan, Thailand, United Arab Emirates. | Open, Classified |  |  | 4638 | Computer
Z-1758-2015 | 71269 | No | DXC | Vascular clamp. | Cardiovascular | Cardiovascular | 510(k) | Edwards Lifesciences IntraClude Intra-Aortic Occlusion Device, model number ICF100, 10.5 Fr (3.5 mm) x 39.4" (100 cm) ThruPort System. Packaged sterile and non-pyrogenic in a sealed, peel-type pouch |  |  | May 05, 2015 | 2015 | Edwards Lifesciences, LLC | Potential device damage on the strain relief near the hub area, which may result in blood leakage. | Device Design | Urgent Medical Device Safety Notice letters were sent via Fed-Ex by May 05, 2015. As part of the customer instructions in the safety notice, customers were requested to review the notice to understand the potential hazard, complete and return the acknowledgement form, and forward copies of the Safety Notice to any organization where the potentially affected devices have been transferred. The device will only be distributed to customers who have returned the acknowledgement form. Questions should be directed to Edwards Customer Service at 800-424-3278. |  | 3360 | 3,360 units. | Nationwide Distribution -- US, including the states of Alabama, Arizona, California, Colorado, Connecticut, Florida, Georgia, Idaho, Illinois, Kentucky, Louisiana, Massachusetts, Maine, Michigan, Missouri, Mississippi, North Carolina, Nebraska, New Jersey, New York, Ohio, Oklahoma, Oregon, Pennsylvania, South Carolina, South Dakota, Tennessee, Texas, Utah, Virginia, Washington, Wisconsin, and West Virginia. | Open, Classified |  |  | 3360 | Computer
Z-2741-2015 | 71725 | No | JFY | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | DimensionÂż CREA (Creatinine) and Dimension VistaÂż CREA (Creatinine); The CREA method used on the DimensionÂż clinical chemistry system is an in vitro diagnostic test intended for the quantitative determination of creatinine in human serum, plasma and urine. |  |  | July 15, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc. | Siemens has confirmed that the Dimension CREA and Dimension Vista CREA may produce falsely depressed results up to - 12.7% and -22% respectively, for a creatinine concentration of 1.5 mg/dL when hemolysis exceeds 200 mg/dL and 300 mg/dL respectively. | Device Design | Siemens sent an Urgent Medical Device Correction letter, dated July 2015, was provided to all DimensionÂż CREA and Dimension VistaÂż CREA customers who received the products to notify them of the issue, the potential risk to health, and actions to be taken by customers.If you have any questions please contact your Siemens Customer Care Center Technical Solutions at (800) 441-9250, or your local Siemens Technical Support Representative. |  | 32165 | Total number Dimension Vista - 32,165, Dimension - 238,634 | Worldwide Distribution - US (nationwide) and to the countries of : Afghanistan, Algeria, Argentina, Armenia, Australia, Austria, Bahamas, Bahrain, Bangladesh, Belgium, Bolivia, Bosnia Herzegovinia, Bulgaria, Brazil, Canada, Canary Islands, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Ecuador, Egypt, El Salvador, France, Georgia, Germany, Ghana, Great Britain, Greece, Guatemala, Hungary, Iraq, Ireland, Honduras, Hong Kong, Iceland, Indonesia, India, Israel, Italy, Japan, Jordan, Kazakhstan, Kuwait, Latvia, Lebanon, Lithuania, Libya, Macedonia, Malaysia, Mali, Mexico, Myanmar, Morocco, Netherlands, Nepal, Nicaragua, New Zealand, Norway, Oman, Pakistan, Paraguay, Peru, Philippines, Poland, Portugal, Qatar, Russian Federation, Saudi Arabia, Senegal, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Switzerland, Taiwan, Thailand, Tunisia, Turkmenistan, United Arab Emirates, Uruguay, Venezuela, Vietnam and Yemen. | Open, Classified |  |  | 32165 | Computer
Z-1726-2015 | 71242 | No | DSQ |  |  | Cardiovascular | PMA | HeartWare Ventricular Assist System (HVAD). |  |  | April 29, 2015 | 2015 | HeartWare Inc | Complaints with the HVAD Retraction of Pins within the driveline connector. | Device Design | Urgent Medical Device Correction letters, dated APR2015, were sent to users. The letters identified the affected device, as well as the reason for the recall and risks involved. Users were asked to review and familiarize themselves with the notice and "Patient Communication." The notice should be forwarded to anyone within customers' organization that need to be aware. Customers are to identify their patients who are currently supported by the HVAD System and distribute the "Patient Communication" to them. A HeartWare representative can assist with that process is needed. Customers should continue to reinforce the messages in the notice with their patients. The "Acknowledgement and Completion Form" should be completed, signed, and returned within 30 days of receipt of the notice. Questions or concerns should be directed to a local HeartWare representative. Clinical Support is also available at 1-888-494-6365. |  | 3747 | 3,747 currently in use | Worldwide Distribution -- US, including the states of AL, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, NC, NH, NJ, NY, OH, OR, PA, SC, TN, TX, UT, and VA; and, the countries of Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, Croatia, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, Hong Kong, Hungary, India, Israel, Italy, Japan, Korea, Lebanon, Lithuania, Luxembourg, Malaysia, Netherlands, New Zealand, Norway, Poland, Romania, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Turkey, and UK. | Open, Classified |  |  | 3747 | Computer
Z-1796-2015 | 71405 | No | DPS | Electrocardiograph. | Cardiovascular | Cardiovascular | 510(k) | ELI 380 Electrocardiograph. |  |  | May 29, 2015 | 2015 | Mortara Instrument, Inc | Mortara Instrument, Inc. has recently becomeaware of a potential safety hazard involving our ELI380 electrocardiograph. When used in a particularworkflow, acquired ECG waveforms for one patientmay become associated with the patientdemographics for a different patient when therecord is transmitted to a records managementsystem. | Software design | Mortara Instrument, Inc. sent an "Urgent Medical Device Correction" letter dated May 29, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were advised that users in their facilities be made aware of the safety notification and that required actions listed on the letter be taken immediately. A Field service Engineer will be dispatched to their facility to update the affected unit(s) on-site. For questions they can contact Technical support at 1-888-667-8272 or Techsupport@mortara.com or contact your local service Representative. |  | 176 | 176 | Worldwide Distribution - US Distribution to the states of : AZ, CA, FL, GA, IL, MI, NE, SC, UT, WA and WI., and to the countries of : CHINA, PORTUGAL, CANADA, OMAN, MALAYSIA, NEW ZEALAND, INDIA, POLAND, LEBANON, SWITZERLAND, FINLAND, ROMANIA,UNITED ARAB EMIRATES, UNITED KINGDOM, AUSTRALIA, THE NETHERLANDS, ITALY, FRANCE, GERMANY, THAILAND,SAUDI, ARABIA, NORWAY and IRELAND. | Terminated | August 05, 2015 | 68 | 176 | Computer
Z-1795-2015 | 71264 | No | GWN | Nystagmograph. | Neurology | Neurology | 510(k) | I-PortalÂż NOTC and VNG; a device used to measure, record, or visually display the involuntary movements (nystagmus) of the eyeball, indicated for use as a diagnostic tool to assist trained clinicians in their analysis of vestibular disorders, which requires the separation of central vs. peripheral nervous system deficits. |  |  | April 30, 2015 | 2015 | Neuro Kinetics, Inc. | complaints of system malfunction and unintended, sudden movement at start up. No injuries reported. | Software design | Neuro Kinetics sent a Field Correction Advisory letter dated April 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers with the affected product were instructed to contact Customer Service if a malfunction is encountered and not to use the device until it has been assessed and corrected by Nuerokinetics. Customers with questions should call 412-963-6649. |  | 140 | 140 | Worldwide distribution - US (nationwide) and Internationally to Canada and Mexico. | Terminated | April 17, 2017 | 718 | 140 | Computer
Z-1731-2015 | 71239 | No | DSQ |  |  | Cardiovascular | PMA | HeartWare Ventricular Assist System (HVAD). |  |  | April 29, 2015 | 2015 | HeartWare Inc | Complaints with the HVAD Internal Controller Alarm Battery failures. | Device Design | Urgent Medical Device Correction letters were sent to customers. The letter identified the affected product, as well as the reason for the recall and risks involved. HeartWare requests that customers review and familiarize themselves with the letter and "Patient Communication," and ensure that all necessary individuals are aware. Patients currently being supported by the device should be identified and given the "Patient Communication." A HeartWare representative can provide assistance with that process if needed. The "Acknowledgement and Completion Form" should be completed, signed, and returned within 30 days of the receipt of the letter. Customers can contact their local HeartWare representative with questions or concerns. Clinical Support personnel are also available at 1-888-494-6365. |  | 3747 | 3,747 curently implanted | Worldwide Distribution -- US, including the states of AL, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, NC, NH, NJ, NY, OH, OR, PA, SC, TN, TX, UT, and VA; and, the countries of Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, Croatia, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, Hong Kong, Hungary, India, Israel, Italy, Japan, Korea, Lebanon, Lithuania, Luxembourg, Malaysia, Netherlands, New Zealand, Norway, Poland, Romania, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Turkey, and UK. | Open, Classified |  |  | 3747 | Computer
Z-1801-2015 | 71247 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | T.A.L.O.N (Tactically Advanced Lifesaving Intraosseous Needle) 15G x 38.5 mm 10 NEEDLE SET STERILE EO Single Use Rx only Vidacare Corporation |  |  | May 12, 2015 | 2015 | Vidacare Corporation | Potential breach in sterility | Device Design | Teleflex sent an Urgent Medical Recall Notification dated May 12, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letter directed the customers to quarantine the affected product and return it to the manufacturer. If you have any further questions please call 1-866-246-6990. |  | 2104 | 2,104 boxes | US Distribution to the states of : NC, MD, MS, TX and MO., and Internationally to Canada and Denmark. | Terminated | April 11, 2016 | 335 | 2104 | Computer
Z-2722-2015 | 71907 | No | DPW | Cardiovascular blood flowmeter. | Cardiovascular | Cardiovascular | 510(k) | Transonic Flow Probe, 6 mm, Catalog Number HQC6FMV -- |  |  | July 27, 2015 | 2015 | Transonic Systems Inc | Software defect; programming errors in the blood flow probes EPROM. Transonic received several complaints from users indicating that when the flow probe is connected to a flow meter the meter reads Used Probe or Irrelevant Use Number. | Software Manufacturing/Software Deployment | Transonic sent an Urgent Medical Device Recall letters and Response Forms dated July 27, 2015 to customers via Certified Mail, email and/or fax. The letter identified the affected product, problem and actions to be taken. Customers are instructed to contact Transonic Systems' Customer Service at 1-800-353-3569 or 1-607-257-5300 to arrange for replacement (repair/exchange) immediately. |  | 1 | Domestic: 1 unit | Worldwide Distribution - US Nationwide and the countries of: Canada, Japan, Russia and The Netherlands. | Terminated | July 18, 2016 | 357 | 143 | Computer
Z-1807-2015 | 71432 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Philips Ultrasound, Model Q-Station, with software version 3 or higher, Catalog number: 795088; Part Number: 989605382391 |  |  | May 15, 2015 | 2015 | Philips Ultrasound, Inc. | Philips Healthcare has discovered a problem in the Philips Ultrasound Q-Station version 3.0 or higher that couldresult in measurements from a study (Structured Report) to be appended to subsequent studies for other patients. | Software design | PhilipsHealthcare sent an URGENT - Medical Device Correction Philips Ultrasound Q-Station, Q-STATION STRUCTURED REPORTING AND MEASUREMENTS letter, FCO 79500335, dated 2015 May 15, to all affected consignees.Philips is sending a CD with Q-Station software version 3.3.2 to consignees for a free of charge. Consignees must uninstall previous versions of Q-Station and install version 3.3.2 on your system carefully following the installation instructions that accompany the CD. It is important that consignees register their Q Station 3.3.2 installation as described in those instructions, so that Philips can confirm that they have completed this correction.Consignees were instructed to complete the enclosed Customer Reply Form to confirm that they have (1) read and understood this important Medical Device Correction and (2) have either installed Q-Station version 3.3.2 or declare that this FCO 79500335 does not apply to their institution. Consignees were asked to return the enclosed Customer Reply form to Philips Healthcare within 5 days of receipt of the replacement software. Consignees with questions can call local Philips representative or Philips Customer Service at 1-800-722-9377. |  | 125 | 125 units (30 in US and 95 international) | Worldwide Distribution - US (nationwide) and Puerto Rico; internationally to Australia, Canada, Colombia, France, Germany, Great Britain, India, Italy, Mexico, Netherlands, Norway, Pakistan, Poland, Russia, Singapore, Spain, Switzerland, Thailand, and Zimbabwe. | Open, Classified |  |  | 125 | Computer
Z-2734-2015 | 72026 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | da Vinci Xi Surgical System P4 software;with Xi Stapler; model number IS4000 : |  |  | August 17, 2015 | 2015 | Intuitive Surgical, Inc. | Complaints regarding the Exposed Knife Blade recoverable fault for the Xi Stapler with Xi System P4 version software. | Software design | Intuitive sent an Urgent Medical Device Correction letters dated August 21, 2015 to their customers. The letter identified the affected product, problem and actions to be taken. For questions contact Clinical Sales or ISI Customer service at 800-876-1310 Option 3. |  | 198 | 198 affected devices | Worldwide Distribution - US Nationwide and in the countries of Italy and Norway. | Terminated | February 08, 2016 | 175 | 198 | Computer
Z-1811-2015 | 71081 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris PC unit model 8015, Infusion pump. software version 9.17 |  |  | May 13, 2015 | 2015 | CareFusion 303, Inc. | An issue with the cancel functionality used during atypical infusion programming to cancel user inputted values. An infusion may start that is greater than or less than the hospital established limits for the specific medication. | Software design | CareFusion sent an Urgent Medical Device Recall Notification letter dated May 13, 2015 to customers. The letter identified the affected product, problem and the recommended actions to be taken. Customers are informed that CareFusion has released a software upgrade to address this issue. CareFusion will contact all affected customers within 60 days to initiate the scheduling process for the software upgrade installation. Customers with recall related questions are instructed to contact CareFusion Support Center at (888) 562-6018, 7am-4pm PT, supportcenter@carefusion.com. Customers with adverse event report questions are instructed to contact customer advocacy at (888) 812-3266, 24 hrs/day, 7 days a week, customerfeedback@carefusion.com. Customers with technical questions regarding the Alaris System are instructed to contact Technical Support at (888) 812-3229, 6am-5pm PT, Mon-Fri, dl-us-inf-techsupport@carefusion.com. |  | 63 | 63 units | US Nationwide Distribution | Open, Classified |  |  | 63 | Computer
Z-1816-2015 | 70164 | No | OOI | Instrumentation for clinical multiplex test systems. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | BD Viper LT System, catalog number 442839 when used in conjunction with the BD ProbeTec Chlamydia trachomatis (CT) Qx Assay (cat. #442959) and the BD ProbeTec Neisseria gonorrhoeae (GC) Qx Assay (cat. #442842). In vitro diagnostic. |  |  | November 07, 2014 | 2014 | Becton Dickinson &amp; Co. | Potential for false positive Chlamydia trachomatis (CT) results and Neisseria gonorrhoeae (GC) results when processing the CT Qx and GC Qx Assays on the BD Viper LT Instrument. | Device Design | BD sent an Urgent Field Corrective Action letter dated November 2014, to all affected customers. The letter identified the problem the product and the action needed to be taken by the customer. Customers were advised: a. Include the date recall was initiated, if it is already underway or the date your firm plans to start the recalling process. b. How do you plan to notify all the consignees affected by this recall? (press release, letter, telefax, telephone, e-mail, visit, etc.) The customer was initially contacted via a phone call and subsequently provided with a written communication. c. How do you plan to monitor the number of consignees non-responding to the recall communication? The single US customer was contacted by phone and provided the written communication.d. How do you plan to do effectiveness checks of this recall? The affected instrument will be corrected.e. Date your firm ceased further distribution of the product(s). Product was not placed on hold. This issue does not affect new customers. f. How do you plan to store the recall product? BD has not requested that instruments or reagents be returned.g. How do you plan to dispose of the recall products? BD is not requesting that customers return the instruments or reagents.If you require further assistance or clarification, please contact the BD Technical Services Department at 1-800 638-8663. |  | 7 | 7 Viper LT Systems | Worldwide distribution to US (state of LA), Hong Kong, Austria, China, France, and the UK. | Terminated | May 09, 2016 | 549 | 7 | Computer
Z-0417-2015 | 69545 | No | JZH | Properdin factor B immunological test system. | Immunology | Immunology | 510(K) Exempt | MicroVue Bb Plus EIA is packaged as a kit which contains all reagent components to perform the test. In vitro diagnostic. |  |  | October 16, 2014 | 2014 | Diagnostic Hybrids Inc | Samples are quantitating incorrectly, with approximately 50% higher concentration values than previously reported and historically documented. | Device Design | Quidel sent an Urgent Medical Device Recall letter dated October 16, 2014, to all consignees. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately review their inventory and determine if they have of the affected product. If so, they should remove it from their inventory and destroy it immediately by discarding it into their normal biomedical waste stream. Complete the enclosed Certificate of Destruction. Upon receipt of the Certificate of Destruction, Quidel, will then schedule their shipment of replacement product. If consignees further distributed the affected product, they should identify the customers and provide them with the Urgent Medical Device Recall. Consignees were instructed to monitor and reconcile the quarantine of product with the customers as this information may need to be provided to the authorities. For questions regarding this recall call 740-589-3300.For technical support consignees should contact technicalsupport@quidel.com or call 800-874-1517. |  | 171 | 171 | Worldwide Distribution - USA, including AL, CO, CT, FL, KY, MN, NC, NV, OH, PA, and WA; and, the countries of Canada, China, India, and Switzerland. | Terminated | September 03, 2015 | 322 | 171 | Computer
Z-2719-2015 | 71978 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo.plaza; picture archiving and communications system |  |  | July 06, 2015 | 2015 | Siemens Medical Solutions USA, Inc | Potential issue leading to data loss and patient data mix-up | Software design | Siemens sent a Customer Information letter, dated July 6, 2015, to customers The letter identified the affected product, problem and action to be taken. The letter informed cusotmers about the Hotfix software update that was being released. |  | 37 | 37 | US Nationwide Distribution - in the states of CA, OR, FL, MO, NY, WA, MD, MA, NY, WI, KS, MN and PA. | Terminated | January 06, 2017 | 550 | 37 | Computer
Z-2544-2015 | 71619 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS 5,1 FS Chemistry System, Catalog Number/Product Code 6801375, Unique Device Identifier No. 10758750001132, and VITROS 5,1 FS Chemistry System Refurbished, Catalog Number/Product Code 6801890, Unique Device Identifier No. 10758750001644, IVD --- Ortho Clinical Diagnostics Inc. |  |  | June 30, 2015 | 2015 | Ortho-Clinical Diagnostics | Software anomaly may occur during the installation of an Assay Data Diskette (ADD) using Software Versions 2.2.1 through 2.8. If the user is utilizing User- Configured (Modified) parameters, ALL default settings are restored in the VITROS 5,1 FS Chemistry System from the ADD instead of retaining the user modified parameters. | Software design | The firm, Ortho Clinical Diagnostics, sent an "URGENT PRODUCT CORRECTION NOTIFICATION" letter dated 6/30/15 and response forms to customers via FedEx overnight courier, e-mail (ORTHO PLUS e-Communications) and/or via US Postal Service overnight letter on 6/30/2015. Foreign affiliates were informed by email on 6/30/2015 of the issue and instructed to notify their consignees of the issue and required actions. The letter described the product, problem, and actions to be taken. The customers were instructed to follow the required actions and complete and return the Confirmation of Receipt-Response Required form, no later than 7/8/15, by fax at 1.888.557.3759 or 1.585.453.4110, or scan PDF and email to: ConfirmationAdmin@its.jnj.com. Customers with any additional questions, can contact Customer Technical Services at 1-800-421-3311 at any time. |  | 210 | Catalog 6801375: Domestic 919; Foreign 1251; Catalog 6801890: Domestic 63, Foreign 147 | Worldwide Distribution: US (nationwide) including Puerto Rico, and countries of: Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Republic of Panama, Singapore, Spain, and Venezuela. | Open, Classified |  |  | 210 | Computer
Z-1817-2015 | 71398 | No | DRQ | Transducer signal amplifier and conditioner. | Cardiovascular | Cardiovascular | 510(k) | CLEARSIGN II Amplifier for LABSYSTEM PRO EP Recording System, 120 channels, using firmware version 2.08. The channel numbers in the amplifier are as follows: |  |  | June 02, 2015 | 2015 | Boston Scientific Corporation | Boston Scientific has received complaints indicating that, when using the blood pressure (BP) channels on the CLEARSIGN II Amplifier, the surface Electrocardiogram (ECG) channels become over-written to a variable degree, with the result that it appears shifted from baseline on the system's output screen. This may, in turn, manifest as an uninterpretable ECG signal in the affected channel. | Software design | Boston Scientific sent an Important Field Safety Notification letter to all affected customers beginning June 3, 2015, by 24 hour overnight delivery. The letter identified the product the problem and the action needed to be taken by the customer. INSTRUCTIONS:1. Please read carefully the Field Safety Notice letter and immediately post this information in a visible location near the product to ensure it is easily accessible to all users of the device.2. Please complete the attached Verification Form even if you do not have any product to return.3. When completed, please return the Verification Form to your local Boston Scientific Office to the attention of ÂżCustomer_Service_Fax_NumberÂż on or before DATE 2015.4. If you have products to return, and once a loaner unit is available, a Boston Scientific representative will coordinate with you to package the product being returned using a special shipping container.Your Competent Authority is being notified of this Field Safety Notice.We regret any inconvenience that this action may cause and we appreciate your understanding as we take action to ensure patient safety and customer satisfaction. If you have any questions or would like assistance with this Field Safety Notice, please contact your local Sales Representative. For further questions please call (978)-805-3200 |  | 43 | 43 units | No US distribution. Worldwide Distribution to the countries of : Austria, Belgium, France, Germany, Great Britain, Greece, Italy, Netherlands, Poland, Romania, South Africa, Spain, Switzerland and Turkey. | Open, Classified |  |  | 43 | Computer
Z-1819-2015 | 71262 | No | FZX | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ORTHOFIX Centurion POCT System Drill Guide PN 69-1010 and PN 69-1012. Drill Guides are used to interface with various drill tips during the drilling procedure of bone preparation. |  |  | May 11, 2015 | 2015 | Orthofix, Inc | There is a possibility that the thumb button remains in the depressed (unlocked) position allowing the threaded sleeve (drill stop) to unintentionally rotate and change the depth setting. | Device Design | Orthofix sent an Urgent Medical Device Product Advisory Notification letter dated May 11, 2015, to the affected distributors and hospitals. The consignees were directed to inspect thier inventory and perform checks before surgical use of the drill guides. Customers were instructed to provide a copy of the Product Advisory Notification letter if product was further distributed. Customers with questions were instructed to contact their local Orthofix representative and to for returns they should contact Orthofix customer service representative at 888-298-5700. |  | 36 | 36 units | Nationwide Distribution including CA, VA, FL, ID, CO, OR, PA, OK, SC, and TX. | Terminated | February 17, 2016 | 282 | 36 | Computer
Z-1825-2015 | 71449 | No | IOL | Powered exercise equipment. | Physical Medicine | Physical Medicine | 510(K) Exempt | GE Healthcare, Treadmill T2100. |  |  | March 12, 2015 | 2015 | GE Healthcare | GE Healthcare has recently become aware of a potential safety issue due to uncontrolled motion of the walking belt on your T2100 Treadmill. | Device Design | General Healthcare sent an Urgent Medical Device Correction letter (GEHC Ref# 30074) dated March 12, 2015, to all affected consignees. The letter was addressed to Director of Biomedical Engineering, Chief of Nursing and Healthcare Administrator / Risk Manager. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction, and Contact Information. For questions contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 11423 | 11,423 (5,540 units US, 5,883 units OUS | Nationwide Distribution including Guam, DC, and Puerto Rico.OUS: VIET NAM, VENEZUELA, UZBEKISTAN, URUGUAY, UNITED KINGDOM, UNITED ARAB EMIRATES, TURKEY, TRINIDAD &amp; TOBAGO, THAILAND, TAIWAN EL SALVADOR, SRILANKA, SPAIN, SOUTH AFRICA, SINGAPORE, SAUDI ARABIA, RUSSIA, PORTUGAL PHILIPPINES, PERU, PARAGUAY, PANAMA, PAKISTAN, NEW ZEALAND, NICARAGUA, NEPAL, MYANMAR, MOROCCO, MEXICO, MARTINIQUE, MALAYSIA, LEBANON, KUWAIT, REPUBLIC OF KOREA, KAZAKHSTAN, JAPAN, JAMAICA, ITALY, ISRAEL, IRELAND, IRAQ, INDONESIA, INDIA, HON GONG, GERMANY, FRANCE, FINLAND, ETHIOPIA, EGYPT, ECUADOR, DOMINICAN REPUBLIC, COSTA RICA, COLOMBIA, CHINA, CHILE , CAYMAN ISLAND , CANADA, CAMBODIA, BRUNEI DARUSSALAM, BRAZIL, BOLIVIA, BHUTAN, BANGLADESH, BAHRAIN, AUSTRIA, AUSTRALIA, ARGENTINA, ALGERIA. | Open, Classified |  |  | 11423 | Computer
Z-2538-2015 | 71653 | No | EHD | Extraoral source x-ray system. | Dental | Radiology | 510(k) | Panoramic X-ray Model PC-1000 |  |  | June 04, 2015 | 2015 | Panoramic Rental Corp. | Jackscrew nut part failure potentially leading to device malfunction and patient/operator injury. | Equipment maintenance | The firm, Panoramic Corporation, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated 7/17/2015 via certified mail to customers on 7/20/2015, with response forms for customers to complete and return. The firm previously issued "IMPORTANT SERVICE BULLETIN" notices in 2003 and 2008 via FED-EX. The letter describes the product, problem and actions to be taken. The customers are instructed to select to use several options including replacing the failing part with one containing the safety feature, or trading in their machine for a newer one; discontinue using the devices immediately regardless the option you select, and complete and return the attached response card via email to recall@pancorp.com or fax to 260-489-5683. If the machine has been sold or donated to another dental outfit, the customer should notify Panoramic so that they can update their records. When the customer is reached, the firm will confirm whether or not the machine is still in service.Customers with questions may contact 800-654-2027 Monday through Thursday 8:00 AM to 6:00 PM and Friday 8:00 AM to 5:00 PM EST to discuss these options and determine which is best for their practice. |  | 7885 | 7885 | Worldwide distribution: US (nationwide) all 50 states, including Puerto Rico; and countries of Antigua, Armenia, Aruba, Austria, Canada, China, Colombia, Ecuador, Egypt, England, France, Germany, Greece, Haiti, Honduras, Indonesia, Japan, Kenya, Lithuania, Mexico, Panama, Saudi Arabia, Singapore, Taiwan, Turkey, UAE, Uganda, Venezuela, Vietnam, and Virgin Islands. | Terminated | May 30, 2017 | 726 | 7885 | Computer
Z-2486-2015 | 71803 | No | LBI | Neonatal phototherapy unit. | General Hospital | General Hospital | 510(k) | Bright Embrace, a Single Bili Light for newborn phototherapy |  |  | August 31, 2013 | 2013 | Physician Engineered Products, Inc. | Upper limit for ambient temperature for use of the device was printed incorrectly in the instructions | Device Design | Physician Engineered Products, Inc. sent an Bright Embrace Recall notification dated September 10, 2013, to notify customers of the recall of the Bright Embrace .PEP is issuing a Return Authorization to you. Upon receipt PEP will reimburse to you:1) The cost of the device(s) 2) Shipping costs-both from and to PEPFor further questions please call (207) 935-1256. |  | 97 | 97 units | Worldwide Distribution - US Distribution to the states of : CA, CO, FL, KY, MA, MD, NM, PA, TN, TX and VA., and to the countries of : Egypt, Honduras, Saudi Arabia and Singapore. | Terminated | September 02, 2015 | 732 | 97 | Computer
Z-1917-2015 | 71240 | No | DSQ |  |  | Cardiovascular | PMA | HeartWare Ventricular Assist System (HVAD) |  |  | April 29, 2015 | 2015 | HeartWare Inc | Complaints with the HVAD Abnormal Power Source Switching. | Software design | HeartWare sent an Urgent Medical Device Correction letter to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the enclosed Patient Communication and familiarize the contents, forward the notice to those individuals withing their organization who need to be aware of its contents, identify their patients currently supported by the HVAD System, distribute the Patient Communication to their patients via FedEX or some other reliable means of communication. Customers with questions were instructed to contact their HeartWare representative should they need any assistance with the process. Customers were asked to complete, sign, and return the Acknowledgement and Completion Form to HeartWare. Customers with quesitons were instructed to call 1-888-494-6365. |  | 3747 | 3,747 currently in use. | US including AL, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, NC, NH, NJ, NY, OH, OR, PA, SC, TN, TX, UT, VA and Internationally to Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, Croatia, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, Hong Kong, Hungary, India, Israel, Italy, Japan, Korea, Lebanon, Lithuania, Luxemburg, Malaysia, Netherlands, New Zealand, Norway, Poland, Romania, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Turkey, and UK. | Open, Classified |  |  | 3747 | Computer
Z-2439-2015 | 71395 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | PTW UNIDOS webline: 1) REF T10023, 2) REF T10022 and 3) REF T10023 Firmware version 2.05. Dosimeter for radiation therapy, diagnostic radiology and health physics. |  |  | May 27, 2015 | 2015 | PTW-Freiburg | In some cases, especially with high dose-per-pulse-rates typical for particle (specifically proton) therapy, measuring errors can result. Individual measurement intervals can be lost that are not visible to the user when there is a high dose-per-pulse rate resulting in a too-low dose being displayed on the monitor. This could result in a radiation over-exposure. | Software design | PTW issues Urgent Safety Notice due to an error with firmware version 2.05 of the UNIDOS webline in some cases, especially with the high dose-per-pulse-rates typical for particle therapy, measuring errors can result. |  | 242 | 242 firmwares | Distributed in CA, MI, MA, NC, NJ, NY, PA, TN, TX, PA &amp; WI. | Terminated | March 31, 2017 | 674 | 242 | Computer
Z-0182-2016 | 72372 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON SC2000 ultrasound systems |  |  | October 07, 2015 | 2015 | Siemens Medical Solutions USA, Inc. | Siemens will provide all customers a new version of system software (VA35E or VB10C depending on the hardware configuration). Hardware fixes will be implemented to address situations where the metal component used to engage the wheel lock becomes disconnected, and where the system locks up due to misalignment between the bezel and MPI board, in all units that have not already been corrected. | Device Design | On August 19, 2015, Siemens Ultrasound began sending a customer notification to users, advising them to correct the uppercase/lowercase differences in the patient name at the time of registration until the issue is resolved with a future software release. Siemens believed that the other issues in this correction will begin approximately November 30, 2015 and December 31, 2015. Field service reps will visit sites to install software updates or hardware corrections. |  | 2099 | 2,099 units | Worldwide Distribution -- US and Albania, Angola, Argentina, Armenia, Australia, Austria, Azerbaijan, Bahrain, Bangladesh, Belarus, Belgium, Bolivia, Bosnia Herzeg., Brazil, Brunei, Bulgaria, Canada, Colombia, Croatia, Czech Republic, Denmark, Dominican Rep., Ecuador, Egypt, El Salvador, Finland, France, Georgia, Germany, Hungary, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Lebanon, Lithuania, Malaysia, Maldives, Malta, Mexico, Netherlands, New Zealand, Nigeria, Norway, P.R. China, Pakistan, Philippines, Poland, Portugal, Republic Korea, Romania, Russian Fed., Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Turkey, Turkmenistan, U.A.E., Ukraine, United Kingdom, Uzbekistan, Venezuela, and Vietnam. | Open, Classified |  |  | 2099 | Computer
Z-2340-2015 | 71711 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Siemens Artis zee systems; |  |  | June 29, 2015 | 2015 | Siemens Medical Solutions USA, Inc | The possibility exists for a large amount of liquid to seep through the top cover of the table and penetrate the components inside the table. If this occurs, table movements may no longer be possible. No injuries reported. | Device Design | A customer safety notice, dated June 29, 2015, was sent to direct accounts to alert them about the issue and that Siemens' service organization will be contacting them to perform the corrective action. |  |  |  | US Nationwide | Open, Classified |  |  |  | Computer
Z-2059-2015 | 71528 | No | MNH | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Altiva / Exactech Spine, Large Set Screw, Catalog Number 28202, Lot Number AS05107. |  |  | March 06, 2009 | 2009 | Exactech, Inc. | Incorrect thread form on the screw. | Device Design | Consignees were notified with a Product Recall Notice on 03/06/2009 that all screws must be checked to ensure that any discrepant products to be removed from the field. |  | 749 | 749 devices | Distributed in the states of FL, GA, IL, NY, TX, and WA. | Terminated | July 14, 2015 | 2321 | 749 | Computer
Z-2364-2015 | 71674 | No | GAA | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Perfectum Tuberculin Glass Syringe, Non sterile |  |  | July 09, 2015 | 2015 | Cadence Inc. | Graduation marks of the syringe are not centered with the numeric position marks on the syringe barrel. | Device Design | Cadence issued the recall letter, specific to the customer dated 7/9/15 via registered mail on July 17, 2015. Accounts are requested to Recalled products must not be used.2. Locate and quarantine all affected products.3. Contact Cadence Science, Inc. using the Response Form provided.4. If this information is received by a dealer, wholesaler, or distributor/reseller that distributed any of these products to other facilities, then notify their customers of this action so that those customers can locate and remove all affected products. Any questions contact: Phone: 401-942-1031 |  | 766 | 766 units | NationwideForeign Canada, Denmark. Italy | Open, Classified |  |  | 2431 | Computer
Z-2358-2015 | 71613 | No | ITI | Powered wheelchair. | Physical Medicine | Physical Medicine | 510(k) | InvacareÂż ProntoÂż Air Personal Transporter (power wheelchair), Invacare Corporation, One Invacare Way, Elyria, Ohio, 44035-4190, (800) 333 -6900. |  |  | February 02, 2015 | 2015 | Invacare Corporation | The batteries used in the InvacareÂż ProntoÂż Air Personal Transporter models power wheelchairs have the possibility to overheat during charging which could potentially result in leaking battery acid and/or swelling battery, loss of use, chemical burns and/or property damage. | Device Design | Invacare sent an " Urgent Medical Device Recall " letter on April 6, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.You should immediately contact your provider at the number below so a replacement batteries can be installed.If you have any questions concerning these instructions, please call Invacares recall support line at (877) 413-6008, Monday  Friday, 8 a.m.  5 p.m. EST, for assistance. Visit www.invacare.com/recall for more information about this recall. |  | 1057 | 1,057 finished devices | Worldwide Distribution - US (nationwide) Distribution including Puerto Rico, and to the states of : AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA and WI., and to the countries of : AUSTRALIA, CANDA, FRANCE, NEW ZEALAND, SOUTH AFRICA, TRINIDAD AND TOBAGO | Terminated | September 14, 2015 | 224 | 1057 | Computer
Z-2063-2015 | 71573 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MOSAIQ Oncology Information System |  |  | July 01, 2015 | 2015 | Elekta, Inc. | A problem exists in MOSAIQ resulting in the incorrect field size being sent to the treatment machine for stereotactic plans using cones. | Software design | Elekta sent an Urgent Important Field Safety Notice 371-01-MSQ-003 on July 1, 2015, to all affected customers. The notice informed users of the specific product and version numbers affected and that the problem only exists if all the following criteria is met - (1) Running Mosaiq 2.41; (2) with at least one linac equipped with an Agility 160 MLC head (Versa HD, Infinity, Axesse (EOS) or Synergy); and (3) who are treating with Stereotactic Cones. The notice also advises users of the workaround to avoid the problem. The customers were also instructed to return the Important Field Safety Notice Acknowledgment form. For questions regarding this recall call 770-670-2422. |  | 74 | 74 Total (45 US) | Worldwide Distribution - US including AZ, CA, GA, ID, IL, MO, NJ, NY, NC, OR, SC, TN, WV, and Internationally to Austria, Brazil, Bulgaria, Canada, Colombia, Germany, Israel, Italy, Japan, Mexico, Russia, Taiwan, and Venezuela. | Open, Classified |  |  | 74 | Computer
Z-2094-2015 | 69760 | Yes | NOU | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Breas Vivo 50 Home Care Ventilator |  |  | November 07, 2014 | 2014 | Breas Medical AB | Unintended treatment termination could result from a keypad malfunction in some situations. The device erroneously interprets this as a Stop Treatment Instruction. An alarm will not sound, or be registered. Accessories and monitoring equipment connected to the Vivo 50 will stop functioning as the device enters stand-by mode. | Device Design | Breas Medical sent Urgent Release: Field Safety Notice letters dated November 7, 2014 and Field Safety Notice dated November 21, 2014to their customers. The letters identified the affected product, problem and the actions to be taken. |  | 846 | 846 US | Worldwide - United States Nationwide Distribution and the countries: Argentina, Belgium, Chile, Cyprus, Denmark, Ecuador, Finland, France, Germany, Greece, Hong Kong, India, Ireland, Israel, Italy, Japan, Kuwait, Lebanon, Malaysia, Netherlands, Norway, Panama, Poland, Portugal, Russia, Saudi Arabia, Serbia, Singapore, Slovenia, Spain, Sweden, Switzerland, Thailand, Turkey, UAE and United Kingdom. | Terminated | April 07, 2016 | 517 | 846 | Computer
Z-2347-2015 | 71814 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Merge Hemo, 9.10, 9.20.0, 9.20.1,9.20.2, 9.30, 9.40.0, 9.40.1, 9.40.2 with Massimo PHASEIN End Tidal CO2 (EtCO2) module. Merge Hemo (formerly named HeartSuite Hemodynamics) monitors, measures, and records physiologic data from a human patient undergoing a cardiac catheterization procedure. |  |  | January 21, 2015 | 2015 | Merge Healthcare, Inc. | A system freeze-up of the Merge Hemo system, that included the PHASEIN End Tidal CO2 unit, resulted in the user needing to move the patient to another cath lab. | Software design | Merge Healthcare sent an Important Product Information letter dated January 21, 2015, to all affected consignees. The letter described the Issue, Affected Product, Instructions, Product Updates, and Contact Information. For questions contact Merge Customer Service at (877) 741-5369 or support@merge.comA second UPDATE letter was sent April 16, 2015. The letter included the Product Updates.For questions regarding this recall call 866-261-5761. |  | 451 | 451 | Nationwide Distribution including IL, SC, WI, AL, CT, CA, LA, NM, TX, OK, GA, IA, OH, FL, CO, TN, NY, NC, MA, AR, MO, HI, ND, AK, UT, WV, WA, AZ, MI, PA, ME. | Terminated | December 09, 2015 | 322 | 451 | Computer
Z-2085-2015 | 71555 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Patient Handling System (Motion Control Software), |  |  | April 01, 2015 | 2015 | Viewray Incorporated | ViewRay received a report that the couch moved unexpectedly into the bore after performing a RTCS reboot. | Software Design Change | ViewRay sent a Customer Advisory Notification Letters dated April 1, 2015 to thier customers. The letter identified the affected product, problem and actions to be taken. For questions or concerns regarding this issue, contact ViewRaylM Customer Support at support@viewray.com or call the ViewRaylM Customer Support center at 855-286-8875. |  | 3 | 3 | US Nationwide Distribution in the states of CA, MO &amp; WI | Terminated | February 04, 2016 | 309 | 3 | Computer
Z-2087-2015 | 71628 | No | PGY | Picture archiving and communications system. | Radiology | Radiology | 510(k) | RadiForce GX540, 5MP Monochrome LCD Monitor, Part Code No. 0FTD1826A5, 0FTD1826B5, 0FTD1826C5, UPC 690592034752. |  |  | July 07, 2015 | 2015 | Eizo Corporation | EIZO is recalling the RadiForce LCD Monitors could display abnormally while connecting to particular graphics boards with DisplayPort. | Device Design | EIZO sent An Urgent Field Safety Notice dated July 7, 2015, to all customers who purchased the RadiForce RX440, RX650, RX850 and GX540. The notice informs customers that the RadiForce monitors could display abornormally in 6-bit (64 gradations) instead of normal 8-bit (256 gradations) or 10-bit (1024 gradations) while connecting to particular graphics boards with DisplayPort. The notice provides the customers with the actions to be taken and contact information (1-800-800-5202, M-F, 8am-4pm, PST, us.feedback@eizo.com, 5710 Warland Drive, Cypress, CA 90630). |  | 19 | 19 units | Nationwide Distribution including OH, FL, IL, TX, OH, NY, AZ, and NJ. | Terminated | November 30, 2016 | 512 | 24 | Computer
Z-2097-2015 | 71446 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Magnus Hybrid OR Table Column (1180.01B2) used in conjunction with Siemens Artis Zeego Angiography System. Maquet GmbH &amp; Co. |  |  | June 10, 2015 | 2015 | Maquet Cardiovascular Us Sales, Llc | There is a potential issue that could result to collision between the C-arm of the angiography system and the MAGNUS OR table system. After switching on the angiography-system, in rare cases, MAGNUS column provides wrong position data for the longitudinal movement leading to a collision. | Device Design | Maquet sent a field correction letter dated June 8, 2015, via Fed Ex. Customers are advised of the issue with the device and provided instructions. Questions or additional information can be obtained from the local Maquet representative or Maquet Technical Support at 1-888-627-8383 (option 3 followed by option 1) Monday through Friday between the hours of 8:00 am and 6:00 pm EST. |  | 12 | 12 units | Worldwide Distribution - US Nationwide Distribution and to the countries of : France, Australia, Belgium, Brazil, Switzerland, Chile, China, Germany, Spain, United Kingdom, Hong Kong, Ireland, Israel, India, Italy, Japan, Kuwait, Kazakhstan, Mexico, Netherlands, Norway, Russia, Sweden, Turkey, Taiwan and South Africa. | Terminated | April 18, 2016 | 313 | 12 | Computer
Z-1838-2015 | 71056 | No | LHN | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | PROBEAT WITH DISCRETE SPOT SCANNING SYSTEM |  |  | April 21, 2015 | 2015 | Hitachi America, Ltd., Power Systems Division | At a PROBEAT III site installed in Japan, the operator of the system found that the radiation was delivered inconsistent with the treatment plan and stopped the irradiation. The irradiation control system may be potentially affected when rebooting a part of the control system. | Device Design | The recalling firm notified the one consignee about the incident that occurred in Japan on 4/7/15, then gave them an update on 4/13/15 and 4/21/15. The written notification had details for the consignee on how to prevent the device from malfunctioning. The recalling firm also had held conversations with the consignee to discuss the issue and revised work instructions via telephone conference and on site Hitachi personnel. |  | 3 | 3 | Distribution in US (Texas) and Japan | Terminated | March 24, 2016 | 338 | 3 | Computer
Z-2153-2015 | 71411 | No | HQL | Intraocular lens. | Ophthalmic | Ophthalmic | PMA | Bausch &amp; Lomb SofPort Advanced Optics Aspheric Lens |  |  | May 26, 2015 | 2015 | Bausch &amp; Lomb Surgical, Inc. | Haptics breaking during lens loading and insertion. | Device Design | Bausch + Lomb sent an Urgent Medical Device Recall letter dated May 29, 2015 to their custometers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to: segregate the identified IOLs, return the IOLs through your B&L; representative or by using the prepaid shipping label included in this package, order replacement product through customer service (800) 338-2020 or your B&L; representative and complete &amp; return the acknowledgement form included in this package. Customers were advised that if replacement product is needed for a scheduled surgery before your full return can be processed, please contact your B&L; representative to expedite shipment of the IOL. For questions contact Bausch + Lomb at (800) 338-2020. |  | 72773 | 72,773 units. 52,787 units (US) - 19,986 units (Foreign) | Worldwide Distribution - US Nationwide in the states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, ID, IL, IN, KS, KY, LA, MA, ME, MI, MD, MN, MO, MS, NC, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, WV, Hawaii and Puerto Rico and countries of: Canada, Brazil, Argentina, Austria, Belgium, Switzerland, Germany, Spain, France, Great Britain, Georgia, Greece, Kyrgyzstan, Norway, Russia, Ukraine, United Arab Emirates, China, Australia, India Malaysia, Korea, and Taiwan. | Open, Classified |  |  | 72773 | Computer
Z-2448-2015 | 71941 | No | HRM | Surgical microscope and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | LX3 Floor Stand for Endure Luxite Envision Optical System; Endure LuxOR Q-Vue. |  |  | August 11, 2015 | 2015 | Alcon Research, Ltd. | Alcon is initiating a voluntary recall of all LX3 Microscope Floor Stands after receiving reports concerning loss of functionality due to issues with the cabling between the power, host and illumination modules, potentially causing the system to display a system message, and/or resulting in illumination failure. | Device Design | Alcon sent an Urgent Medical Device Correction letter dated August 11, 2015 to affected customers. The letter identified the affected product, problem and actions to be taken. For questions contact Alcon at 800-862-5266. |  | 688 | 688 units | Worldwide Distribution - US Nationwide and the countries Argentina, Australia, Austria, Belarus, Belgium, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Czech Republic, Denmark, Dominican Republic, France, Germany, Greece, Guatemala, Hungary, India, Italy, Japan, Jordan, Korea, Lebanon, Malaysia, Mexico, Mongolia, New Zealand/Fiji, Oman, Panama, Philippines, Poland, Portugal, Puerto Rico, Qatar, Romania, Saudi Arabia, Singapore, Slovakia, South Africa, Spain, Switzerland, Taiwan, Thailand, Turkey, Ukraine and United Kingdom. | Open, Classified |  |  | 688 | Computer
Z-2328-2015 | 71584 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart XL+ Defibrillator/Monitor |  |  | June 05, 2015 | 2015 | Philips Electronics North America Corporation | Multiple software and hardware issues with device that can affect its function. | Device Design | On June 29, 2015, Philips sent an Urgent Medical Device Correction notification/Field Safety Notice (FSN) to inform customers of the issue. The FSN identifies details of the units affected, gives instructions on actions to be taken by the customer, and identifies what action Philips plans to take to remedy the issue. An upgrade consisting of both software and hardware will be provided free of charge for all units affected by these issues. A Philips Healthcare representative will contact customers with affected devices to arrange for installation of the upgrades. Philips is asking customers to follow the "Action to be Taken by Customer/User" section of the Urgent Medical Device Correction notification/Field Safety Notice.For questions contact your local Philips representative. |  | 13168 | 13,168 devices | Worldwide Distribution - US Nationwide including PUERTO RICO, Canada, ALGERIA, ANGOLA, ARGENTINA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BANGLADESH, BELGIUM, BOLIVIA, BULGARIA, CHILE, CHINA, COLOMBIA, COSTA RICA, CROATIA (local name: Hrvatska), CYPRUS, CZECH REPUBLIC, DENMARK, DOMINICA, DOMINICAN REPUBLIC, EGYPT, EL SALVADOR, ESTONIA, ETHIOPIA, FINLAND, FRANCE, GABON, GERMANY, GREECE, GUATEMALA, HONG KONG, HUNGARY, ICELAND, INDIA, INDONESIA, IRAN (ISLAMIC REPUBLIC OF), IRAQ, IRELAND, ISRAEL, ITALY, JAPAN, Jordan, KENYA, KOREA, REPUBLIC OF, KUWAIT, LATVIA, LEBANON, LIBERIA, LIBYAN ARAB JAMAHIRIYA, LITHUANIA, MACAU, MALAYSIA, MALTA, MAURITIUS, MEXICO, MOROCCO, MYANMAR (Burma), NAMIBIA, NEPAL, NETHERLANDS, NEW ZEALAND, NIGERIA, NORWAY, OMAN, PAKISTAN, PANAMA, PERU, PHILIPPINES, POLAND, PORTUGAL, QATAR, ROMANIA, RUSSIAN FEDERATION, RWANDA, SAUDI ARABIA, SENEGAL, SERBIA, SEYCHELLES, SINGAPORE, SLOVAKIA (Slovak Republic), SLOVENIA, SOUTH AFRICA, SPAIN, SRI LANKA, SWEDEN, SWITZERLAND, SYRIAN ARAB REPUBLIC, TAIWAN, United Republic of TANZANIA, THAILAND, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, UZBEKISTAN, VIETNAM and ZIMBABWE. | Open, Classified |  |  | 13168 | Computer
Z-2137-2015 | 71421 | No | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | ATTUNE FB PS ARTICULATION SURFACE SIZES 3-8 |  |  | June 12, 2015 | 2015 | DePuy Orthopaedics, Inc. | Potential for the Balseal, a small wire spring coil located on the post feature of the ATTUNE Knee Tibial Articulation Surface Instruments used in trialing to become damaged and disassociate (come off). If the Balseal is separated from the post, it has the potential to enter the surgical site and be left in the patient if the surgeon is unaware of the disassociation. | Device Design | URGENT INFORMATION: MEDICAL DEVICE CORRECTION was issued for all lots of the affected devices. The device correction is to inform US Distributors and Medical Professionals about the issue, how to prevent the issue, and any actions to take. US Distributors will provide Hospitals/User Facilities with the Medical Professional Notice and Acknowledgment Cards. US Distributors will also work with the sales consultants to ensure all consultants complete sales training regarding this issue. |  | 34676 | 34676 units | Worldwide Distribution - US including AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY and Internationally to CANADA , AUSTRALIA, AUSTRIA, BELGIUM, ENGLAND/ UK, FRANCE, GERMANY, HONG KONG, INDIA, INDONESIA, IRELAND, ISRAEL, ITALY, JAPAN, JAPAN, MALAYSIA, NETHERLANDS, NEW ZEALAND, SINGAPORE, SLOVAKIA, SOUTH AFRICA, SOUTH KOREA, SPAIN &amp; IBERIA, SWEDEN, SWITZERLAND, THAILAND, UAE, CHILE, COLOMBIA, MEXICO, PERU, and PUERTO RICO. | Terminated | November 12, 2015 | 153 | 280591 | Computer
Z-2133-2015 | 71654 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | Volcano s5/s5i/CORE Intravascular Ultrasound Systems software versions 3.3 and 3.4 (with and without iFR Scout Technology) and v.3.4 software kits; |  |  | June 22, 2015 | 2015 | Volcano Corporation | A software issue was discovered where an inaccurate FFR/iFR value could be calculated under certain circumstances. | Software design | Urgent Medical Device Correction letters were sent on June 22, 2015 to all affected sites via FED Ex. |  | 4007 | A total of 4007 devices, all models | Worldwide Distribution-US (nationwide) and regions in Europe and the countries of Japan, Canada, South Africa, Taiwan, Thailand, Singapore, Philippines, Australia, New Zealand, El Salvador, Malaysia, Mexico, Lebanon, Kuwait, South Korea, Kenya, Jordan, Iran, India, Israel, Hong Kong, Egypt, Algeria, Dominican Republic, Costa Rica, Chile, Brazil, Bolivia, and Argentina | Terminated | April 14, 2017 | 662 | 4007 | Computer
Z-2197-2015 | 71699 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | GEMINI TF Ready CT/PET System. |  |  | July 01, 2015 | 2015 | Philips Medical Systems (Cleveland) Inc | Philips has identified four (4) software defects in the Tumor LOC software application that may potentially result in irradiation of healthy tissue or non-irradiation of diseased tissue. | Software design | On 7/6/2015 the firm sent Field Safety Notices to their customers. |  | 1 | 1 unit | Worldwide Distribution - US Nationwide, Australia, Belgium, Brazil, Canada, China, Denmark, France, Germany, Israel, Italy, Japan, Lebanon, Netherlands, Portugal, Saudi Arabia, South Africa, Switzerland, Thailand, Venezuela &amp; Vietnam. | Open, Classified |  |  | 86 | Computer
Z-2321-2015 | 71688 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON S2000 and ACUSON S3000 ultrasound systems with the Virtual Touch IQ option; |  |  | June 25, 2015 | 2015 | Siemens Medical Solutions USA, Inc. | Potential measurement error on ACUSON S Family ultrasound system. When repositioning the Virtual Touch IQ region of interest from the original (default) location, the lateral position of the measurement tool may not align with the lateral position of the shear velocity data. | Software Design Change | On June 25, 2015 Customer Safety Advisory Notification letters were sent to all affected customers by Certified mail. The letter identified the affected product, problem and actions to be taken. The Siemens Service organization coordinates the implementation of the field correction with the distributors for all countries. For questions contact your Siemens Service for information. |  | 629 | 629 units | Worldwide Distribution - US Nationwide and the countries of Algeria, Australia, Austria, Belgium, Brazil, Canada, Chile, Colombia, Croatia, Czech Republic, Denmark, Dominican Rep., France, Germany, Greece, Hong Kong, India, Indonesia, Israel, Italy, Japan, Kenya, Malaysia, Mexico, Nepal, Netherlands, New Zealand, Norway, P.R. China, Philippines, Poland, Portugal, Republic Korea, Romania, Russian Fed., Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, U.A.E., Ukraine, United Kingdom. | Open, Classified |  |  | 629 | Computer
Z-2200-2015 | 71689 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Pinnacle3 Software Version 9.0, 9.2, 9.4 and 9.6, Model Numbers 453560446041, 459800091001, 459800220161, 459800232931, 459800235871, 459800338451. |  |  | July 31, 2014 | 2014 | Philips Electronics North America Corporation | Philips, Pinnacle Radiation Treatment Planning System version 8.0h, 8.0k, 8.0m, 8.0n, 9 0, 9 2 9 4, 9 6 is being recalled because the dose may be inconsistent with the density of a density-overridden ROl. | Software design | Consignees were sent on 7/31/2014 a Philips "Urgent Field Safety Notice" dated 2014-July 20. The letter described the problem, product involved in the recall and the action that should be taken by the customers. For questions contact Customer Care Solutions Center at 1-800-722-9377.UPDATE: A second Philips letter dated 2 July 2015 was sent in August 6, 2015. This letter updated the recall to incorporate the additional products being recalled. |  | 1642 | 1642 | Worldwide Distribution-US: All states in the US including DC and PR. OUS: Australia, Austria, Bangladesh, Belgium, Brazil, China, Denmark, France, Germany, Hong Kong, India, Indonesia, Ireland, Italy, Japan, Malaysia, Mexico, Netherlands,New Zealand, Oman, Philippines, Singapore, South Africa, South Korea, Spain, Switzerland, Taiwan, Thailand, United Kingdom, Viet Nam.UPDATED: Canada, Bulgaria, Egypt, Hungary, Jordan, Luxembourg, Poland, Reunion, Romania, Russia, Sauth Arabia, Turkey. | Open, Classified |  |  | 1642 | Computer
Z-2314-2015 | 71687 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON S 1000, ACUSON S 2000, or ACUSON S 3000 ultrasound systems with software version C3, C3, C3, or C1. |  |  | June 25, 2015 | 2015 | Siemens Medical Solutions USA, Inc. | Potential measurement error on ACUSON S Family ultrasound system when using the Doppler manual trace measurement tool in full screen format that has a low probability of misdiagnosis. | Device Design | Customer Safety Advisory notifications were sent by certified mail to US customers on June 25, 2015. Siemens Service organization coordinate the implementation of the field correction with the distributors for all countries.Letters advised users of the potential miscalculation on ACUSON S1000, ACUSON S2000, or ACUSON S3000 ultrasound systems with software version VC30A, VC30B, VC30C, or VC31A. The manual trace tool may incorrectly calculate the Time Average Velocity and subsequently incorrectly calculate the Pulsatility index. This could lead to potential misdiagnosis.To mitigate the problem users should not perform measurement utilizing the manual Doppler trace tool in a full screen display format.This issue will be resolved in a future software release. If you have any questions, please contact your Siemens Service for more information.Until the necessary repairs have been completed on your system, please share this information with all personnel within your organization who need to be aware of this issue.Siemens is continuing to ship in forward production and will be including the customer safety advisory notice with the shipment. Siemens will update the software for those customers as well once it is released. For further questions please call (650) 694-5398. |  | 1789 | 1789 units | Worldwide Distribution - Nationwide Distribution including Puerto Rico and to the countries of : Algeria, Angola, Australia, Austria, Azerbaijan, Bahrain, Bangladesh, Belarus, Belgium, Bolivia, Bosnia Herzeg., Brazil, Bulgaria, Canada, Chile, Colombia, Croatia, Cyprus, Czech Republic, Dem. Rep. Congo, Denmark, Dominican Rep., Ecuador, Egypt, El Salvador, Finland, France, Germany, Greece, Hong Kong, Hungary, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Lebanon, Malaysia, Mexico, Morocco, Nepal, Netherlands, New Zealand, Norway, P.R. China, Peru, Philippines, Poland, Portugal, Republic Korea, Romania, Russian Fed., Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Trinidad,Tobago, Turkey, Turkmenistan, U.A.E., Ukraine, United Kingdom, Uzbekistan, Venezuela and Vietnam. | Open, Classified |  |  | 1789 | Computer
Z-2280-2015 | 71212 | No | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | ANSPACH - Fluted Drum, Side Cut only 10MM x 14MM , REF L-10DRC-G1 |  |  | April 13, 2015 | 2015 | The Anspach Effort, Inc. | G1 Dissection Tools are not compatible with all attachments listed on the labeling; G1 cutting tools are labeled with a Specified Shaft Exposure which may not be consistent when used with each attachment on the product label. Shaft exposure may be greater. | Device Design | Anspach sent an Urgent Notice: Medical Device Recall (Correction) letter of Anspach G1 Dissection Tools dated May 6, 2015, to all affected customers. Customers were not required to return their inventory of GI Dissection tools. Customers were instructed to complete, sign, and return the attached reply form confirming receipt of the letter by fax or email. Customers were also instructed to forward the letter as appropriate immediately. Customers with questions should contact the Complaint Handling Unit Manager at 561-494-3673 or contact their Anspach Sales Representative.For questions regarding this recall call 561-627-1080. |  | 65 | 65 tools. | Worldwide Distribution - US including AZ, CA, CO, FL, GA, IL, KY, LA, MN, MO, NC, NJ, NY, OH, OR, PA, TX, and Internationally to Canada and China. | Open, Classified |  |  | 5320 | Computer
Z-2263-2015 | 71570 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Paragon Laboratory Management |  |  | June 15, 2014 | 2014 | McKesson Technologies, Inc. | For Paragon Laboratory Management 12.1 and 12.1.1 releases, if used with Microsoft SQL Server 2012, the Final Cumulative Report may incorrectly display the data. | Software design | McKesson issued a Clinical Alert to notify customers of the issue on June 15, 2015, and requested that customers disable Cumulative Reporting. McKesson is identifying customers currently in a live environment with Microsoft SQL Server 2012 and the affected configuration. McKesson will provide those customers with corrective software, and provide assistance in generating correct reports and updating those to the patient folder. For further questions, please call (404) 338-3556. |  | 128 | 128 | Nationwide Distribution including Puerto Rico and to the states of : CA, CO., CT, FL, GA, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NH, NJ, NM, NY, NC, OH, OK, OA, PA, SC, SD, TN, TX, VT, VA, WV and WY. | Open, Classified |  |  | 128 | Computer
Z-2254-2015 | 71124 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Computed Tomography X-ray Systems Brilliance CT 64-channel with Essence technology. |  |  | February 05, 2015 | 2015 | Philips Medical Systems (Cleveland) Inc | 5 issues. Fast Get Ready incorrectly enabled during service procedures. Contrast annotation missing for some images when manual contrast was administered. Dot artifact present, intermittently, after startup. Heart Rate measured value incorrectly remains constant for some heart rate change conditions. DoseRight incorrectly enabled when the Exam Card specified disabling DoseRight. Tube heat predict | Software change control | On 2/5/2015 the firm sent Customer Information Letters to their customers. |  | 4 | 4 Units | Worldwide Distribution-US (nationwide) including the states of AR, AZ, CA, CT, DE, FL, GA, IA, IN, KS, LA, MA, MD, MI, MN, MO, MS, MT, NH, NJ, NM, NV, NY, OH, OR, PA, SC, TN, TX, UT, VT, WV &amp; WY, and to the following countries: Argentina, Australia, Austria, Belgium, Bolivia, Brazil, Cambodia, Canada, China, Colombia, Denmark, Finland, France, Germany, Hong Kong, India, Indonesia, Iran, Iraq, Ireland, Italy, Japan, Latvia, Libyan Arab Jamahiriya, Lithuania, Malaysia, Mexico, Myanmar, Netherlands, New Zealand, Norway, Panama, Philippines, Russia, Singapore, Slovakia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, Uzbekistan &amp; Viet Nam. | Open, Classified |  |  | 424 | Computer
Z-2057-2015 | 71234 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Hamilton-G5 Ventilators with software versions V2.40/2.41 |  |  | March 24, 2015 | 2015 | Hamilton Medical, Inc. | Customer reports that the ventilator display can freeze. Ventilation continues but the information is no longer displayed on the screen and the user can no longer operate the device. | Device Design | *** Please note *** This is not a new recall. An error occurred, and this recall was originally classified as a class I on July 10, 2015. No further action is required.Hamilton Medical AG notified their customers with a letter dated February 19, 2015 and had a fix (SW 2.42) that was available as of February 20th 2015. Hamilton Medical proceeded to send out notification letters the 24th of March 2015 once they compiled the complete and validated list of customers that had the affected software. Customers with questions were instructed to call 800-426-6331, ext.215. |  | 248 | 248 | Nationwide Distribution | Terminated | December 23, 2015 | 274 | 248 | Computer
Z-2054-2015 | 71315 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens Healthcare Diagnostics Dimension Vista Systems, used with the Dimension Vista Intelligent Lab System software version 3.6.1. In vitro diagnostic analyzer. |  |  | May 21, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc. | Two software issues may occur in Vista software version 3.6.1. Issue #1: Samples may stop processing without notification. Issue #2: Timing issue in Vista 1500 may cause wrong reagent or no reagent delivery. This may lead to unflagged, unexpected low results and results flagged with assay errors. | Software design (manufacturing process) | Siemens sent an Urgent Medical Device Correction letter dated May 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Actions to be Taken by the Customer:For the corresponding issues listed above, please perform the following:Issue 1: If you experience this situation, sample processing can be resumed using the step(s) below:Step 1: From the Operation Icon select PAUSE, then immediately return to the Operation Icon and select Cancel PAUSE. Sample aliquotting should immediately resume.Step 2: If sample processing does not resume after Step 1, wait until the instrument displays System Ready, then restart the system software per the Vista Operator Guide,Section 2, Controlled Shutdown-Restart procedure.Issue 2:To prevent the Reagent Servers from temporarily losing synchronization, Siemens recommends setting the Auto Dispose option to No for methods assigned to Reagent Server 2 (see instructions below). These empty or expired reagent cartridges will now remain on the instrument and require the operator to unload them.Please consult your Vista Operators Guide or method Instructions for Use when results with flags and/or errors occur.We apologize for the inconvenience this situation may cause. If you have any questions, please contact your Siemens Customer Care Center-Technical Solutions at (800) 441-9250 or your local Siemens technical support representative. |  | 2315 | 2315 | Worldwide Distribution - Nationwide Distribution including Puerto Rico., and the countries of : Austria, Australia, Belgium, Canada, Denmark, France, Finland, Germany, Japan, Italy, New Zealand, Netherlands, Czech Republic, Portugal, Republic Korea, Norway, Spain, Switzerland, Slovakia, Saudi Arabia, Slovenia, and United Kingdom, | Open, Classified |  |  | 2315 | Computer
Z-2376-2015 | 71658 | No | RCE |  |  |  |  | PSW 1203.U/1, PSW 1301.U, PSW 1307.U, PSW 1403.U, and PSW 1501.U Pacemaker Programmer software. Software approved for use with BIOTRONIK device programmers (trade names: Renamic and ICS 3000). |  |  | June 30, 2015 | 2015 | BIOTRONIK, Inc. | Ventricular packing: LV software programming versions for BIOTRONIK CRT-P and CRT-D devices are recalled because they are not approved for use in the US. | Software Manufacturing/Software Deployment | BIOTRONIK sent the Important Product Information letter and the Acknowledgement of Important Product Information Form, dated June 20, 2015, to affected physicians with the following recommendations: -Review programming settings for your patients with CRT-P or CRT-D devices.-If you have a patient with a device programmed to Ventricular pacing: LV, arrange to have the patients device reprogrammed to an appropriate approved mode (i.e. "Ventricular pacing: RV" or "Ventricular pacing: BiV").-Physicians are requested to sign the Acknowledgement of Important Product Information Form and immediately fax to 800-913-6993 or email to UScompliance@biotronik.com.BIOTRONIK'S representative will update your programmer with approved 1502.U programmer software. Following this update, the Ventricular pacing: LV feature will no longer be available. However, a device previously programmed with"Ventricular pacing: LV will remain programmed in this mode unless and until the pacing mode setting is changed. If you have any questions concerning this Important Product Information, please contact the BIOTRONIK Technical Services Department:Telephone: 1-(800) 284-6689 (24-hour service) |  | 768 | 768 units (software) | Nationwide Distribution. | Terminated | December 24, 2015 | 177 | 768 | Computer
Z-1950-2015 | 71197 | Yes | JOH | Tracheostomy tube and tube cuff. | Anesthesiology | Anesthesiology | 510(k) | Shiley Neonatal, Pediatric and Long Pediatric Tracheostomy Tube, Cuffless. Model Numbers: 2.5NEF,3.0NEF,3.5NEF,4.0NEF,4.5NEF,2.5PEF,3.0PEF,3.5PEF,4.0PEF,4.5PEF,5.0PEF,5.5PEF,5.0PELF,5.5PELF,6.0PELF,6.5PELF,2.5NCF,3.0NCF,3.5NCF,4.0NCF,4.5NCF,2.5PCF,3.0PCF,3.5PCF,4.0PCF,4.5PCF,5.0PCF,5.5PCF,5.0PLCF,5.5PLCF,6.0PLCF,6.5PLCF,3.0NEF-P,3.5NEF-P,4.0NEF-P,4.5NEF-P,3.0PEF-P,3.5PEF-P,4.0PEF-P,4.5PEF-P,5.0PEF-P,5.5PEF-P, |  |  | May 08, 2015 | 2015 | Covidien LP (formerly Nellcor Puritan Bennett Inc.) | Reports where patients who recently switched from the current Shiley Neonatal and Pediatric products to the affected products experienced discomfort and in some cases, breathing difficulties, corresponding to a negative effect on oxygen levels. | Device Design | Consignees were notified 5/8/15 via letter. Consignees were asked to return products. |  | 69461 | 69,461 units total | Worldwide Distribution. US nationwide, Canada, Australia, Belgium, Chile, Germany, Israel, Italy, Japan, Saudi Arabia, Singapore, South Africa, Turkey, and Uruguay. | Terminated | April 12, 2017 | 705 | 69461 | Computer
Z-1849-2015 | 71451 | No | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | The VariSource iX series afterloader systems are computer controlled remote electro/mechanical systems used for medical purposes, for placing a cable incorporating an irradiated iridium seed internally or close by a malignant tumor or tumor bed in a practice known as brachytherapy. |  |  | May 18, 2015 | 2015 | Varian Medical Systems Inc | After an application freeze and restart of the VariSource iX series by power cycling, the Partial Fraction generated by thesystem will not be correct. The application freeze issue affects only the systems equipped with magnetic Hard Disk Drives (HDD). | Software Design Change | All the consignees affected by this recall will be notified by:The Urgent Medical Device Correction Letter/ Field Safety Notification will bedistributed to all affected users, with a description of the problem.The Letter will also be distributed to the Varian Sales, Marketing, and Serviceorganizations, informing them of the issue.Additionally, the letter will be posted to the customer support site:http:/ /www.MyVarian.com. |  | 55 | 55 | Worldwide Distribution. | Open, Classified |  |  | 55 | Computer
Z-1918-2015 | 71346 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | StoneLight 30 Laser System (NS3000). |  |  | May 13, 2015 | 2015 | New Star Lasers, Inc. | During internal testing a situation was discovered where when changing the settings, a new setting could not be accepted and potentially result in a discrepancy in the energy delivered (higher or lower) from the value displayed on the Graphical User Interface display. | Software design (manufacturing process) | Medical Device Correction letters were sent to all affected consignees by Federal Express on May 13, 2015. The letter identified the affected product as well as the issue and risk involved. The software will be upgraded to address the issue; and, the firm will contact customers to coordinate installation of the upgraded software. Customers may continue to use their system safely by taking the precautions provided in the letter. Customers are to complete the attached response form. Questions should be directed to Nina Davis at 916-677-1909. |  | 52 | 52 units | Worldwide Distribution -- US, including the states of CA, IL, AZ, FL, OH, MA, MI, and AL; and the country of the Netherlands. | Terminated | October 28, 2016 | 534 | 52 | Computer
Z-1934-2015 | 71484 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | SOMATOM Emotion 6; produce cross-sectional images of the body by computer reconstruction of x-ray transmission data from either the same axial plane taken at different angels or spiral planes taken at different angles. |  |  | June 04, 2015 | 2015 | Siemens Medical Solutions USA, Inc | A potential safety issue involving two (2) broken screws and one (1) bowed, which is used to fix the Balance Weight of the DMS in the gantry, was found during a service call at the customer site. | Device Design | A safety advisory notice, dated June 4, 2015, was sent to direct accounts explaining the problem and potential risk to patients and operators. Siemens engineers will begin replacing the bolts in May 2015. |  | 40 | 40 total | Nationwide Distribution. | Terminated | March 24, 2017 | 659 | 40 | Computer
Z-1942-2015 | 71204 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | EasyLink Data Management System; a clinical data management system that assists medical laboratory professionals with preanalytic and post-analytic functions in conjunction with multiple instruments, the laboratory information system (LIS) and Siemens StreamLAB Automation Solutions. |  |  | April 29, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc. | System Limitations and Software Issues related to the following features: Custom reference ranges, User-initiated sample and result actions, Auto-verification and quality control rules, and Sample reports. | Software design | Siemens sent an Urgent Field Safety Notice (UMDC # ISW-15-01.A.US) on May 5, 2015 to all affected customers. The letter instructed customers to follow specific work around instructions, review this letter with their Medical Director, complete and return the Field Correction Effectiveness Check attached to this letter within 30 days, retain the letter with laboratory records, and forward it to those who may have received this product.For questions regarding this recall call 800-441-9250. |  | 1355 | 1355 | Worldwide Distribution - US (nationwide) and Internationally to Austria, Belgium, Canada, Switzerland, Germany, Denmark, Spain, France, Hong Kong, Italy, Japan, Korea, Malaysia, Netherlands, Norway, New Zealand, Philippines, Portugal, Saudi-Arabia, Slovenia, and Taiwan. | Terminated | May 11, 2017 | 743 | 1355 | Computer
Z-1939-2015 | 71455 | No | FPO | Wheeled stretcher. | General Hospital | General Hospital | 510(K) Exempt | Sechrist Low Profile Wheeled Stretcher, Model No. 21762-110-33H, 21762-110-36H, 21762-110-41H, 21762-135-33HJ, 21762-010-36H, 21762-010-41H, |  |  | June 01, 2015 | 2015 | Sechrist Industries Inc | Sechrist Industries is recalling the Sechrist Low Profile Wheeled Stretcher because it may experience an unexpected incline angle during the height adjustment of the Low Profile Wheeled Stretcher when lifting the patient upward or when lowering the Wheeled Stretcher to a height at which the patient can be loaded/unloaded. | Device Design | Seachrist sent an Urgent: Medical Device Correction letter dated June 1. 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Our records indicate the Sechrist Low Profile Wheeled Stretcher identified in the attached serial number list (Attachment A) are affected by the field correction and are located at your facility. Within the following weeks, a Sechrist trained Field Service Technician will schedule a visit to your facility to install a Hydraulic Wheeled Stretcher Alignment kit which will greatly minimize the degree of the incline angle that can be experienced after it is installed.Customers are informed that within the following weeks, a Sechrist trained Field Service Technician will schedule a visit to their facility to install a Hydraulic Wheeled Stretcher Alignment kit which will greatly minimize the degree of the incline angle that can be experienced after it is installed. For further questions, please call: (714) 579-8318, or (714) 579-8311.We apologize for any disruption that this correction may cause. Sechrist is dedicated to ensuring the safety and quality of our products. If you have any questions about what to do, please contact us. |  | 81 | 81 US Distribution | US Distribution and Internationally to Japan and Jordan. | Terminated | October 20, 2015 | 141 | 81 | Computer
Z-1949-2015 | 71448 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | INSTRUMENT ARM DRAPE, IS4000, 20 PACK; used on the da Vinci Xi Surgical System. |  |  | June 16, 2015 | 2015 | Intuitive Surgical, Inc. | Field Safety Notification to advise of an issue that may occur involving the combination of the da Vinci Xi Drapes and Xi instruments. A pinched drape may cause a message on the Vision Cart touchscreen - Missing reload error. | Device Design | Intuitive Surgical sent an Urgent Medical Device Notification letter on June 16 , 2015 to all affected consignees. Consignees were instructed to take the following actions:1. Ensure that all affected personnel are fully informed of this notice. Forward thisnotice to your Risk Manager, OR Director, Purchasing Manager, BiomedicalEngineering staff and members of your medical staff who perform da Vinci Surgeryprocedures.2. Complete and return the attached Acknowledgment Form to Intuitive Surgicalusing the instructions provided.3. Please retain a copy of this notice with your system.4. Refer to the da Vinci Xi System User Manual for more information on draping thesystem.Consignees with questions or needing spport should contact their Clinical Sales Representative or contact Intuitive Surgical Sustomer Service at the numbers listed below:Sales Representative or contact Intuitive Surgical Customer Service at the numbers listed below:Âż North and South America: (800) 876-1310, Option 3 (6 AM to 5 PM PST) or mail:customersupport-servicesupport@intusurg.comÂż Europe, Middle East, Asia and Africa: +800 0821 2020 or +41 21 821 2020 (8 AM to6 PM CET) or ics@intusurg.comÂż South Korea: 02-3271-3200 (9 AM to 6 PM KSTJ)Âż Japan: 0120-56-5635 or 003-5575-1362 (9 AM to 6 PM JST). |  | 310 | 310 | Worldwide Distribution - US (nationwide) and Internationally to Australia, Austria, Belgium, Czech Republic, France, Germany, Greece, India, Israel, Italy, Japan, Netherlands, Norway, Puerto Rico, Qatar, Saudi Arabia, Singapore, South Korea, Spain, Sweden, Switzerland, Turkey, and United Kingdom. | Terminated | October 05, 2015 | 111 | 310 | Computer
Z-1961-2015 | 71073 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | The Rapidlab 1260 system is a blood gas analyzer used for laboratory testing of blood gases, electrolytes, metabolites and CO-oximetry in arterial, venous and capillary whole |  |  | May 06, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc | D50 and D51 Diagnostic error codes are not functional. | Software design | Siemens sent an Urgent Field Safety Notice April 2015, to all affected customers. Customers were notified of the recall beginning on May 6, 2015 via Federal Express. Foreign customers were notified starting on May 5, 2015 via method determined by each countries local regulations and procedures. The notification informs customers of the issue and asks customers to do the following: Ensure that the Glucose and Lactate sensor contacts are aligned and completely seated with the contacts on the analyzer, Analyze a minimum of 2 levels of quality control material after installing the sensors, as indicated in the RAPIDLab 1200 Operator's Guide, Review this letter with your Medical Director. Complete and return the Field Correction Effectiveness Check Form attached to this letter within 7 days. Retain this letter with your laboratory records, and forward this letter to those who may have received this product. For further questions please call (800) 242-3233. |  | 122 | 122 devices | Worldwide Distribution - Nationwide Distribution including Puerto Rico and to the countries of : Argentina, Australia, Austria, Belgium, Bosnia, Herzog,Brazil, Brunei, Canada, Chile, Colombia, Croatia, Czech Republic, Egypt, Finland, France, Germany,Georgia, Greece, Hong Kong, Iceland, India, Indonesia, Ireland, Italy, Japan, Latvia, Lithuania, Malaysia, Mexico, Netherlands, New Zealand, Norway, P.R. China, Peru, Poland, Portugal, Qatar, Republic Korea, Romania, Russian Fed, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, U.A.E., United Kingdom, Venezuela and Vietnam. | Terminated | August 02, 2017 | 819 | 2835 | Computer
Z-1965-2015 | 71434 | No | BZO | Ventilator tubing. | Anesthesiology | Anesthesiology | 510(K) Exempt | Hamilton Medical Infant Flow Sensor, single use, (3.1m), Part number 260179. The infant flow sensor is an accessory for the Hamilton Medical Ventilators. |  |  | June 01, 2015 | 2015 | Hamilton Medical, Inc. | An issue has been discovered with the Hamilton Medical Infant Flow Sensor, single use. When a ventilation mode with adaptive volume control (APVcmv/CMV+) is used, the ventilator may adjust the ventilation on the inaccurate flow measurement and decrease the amount of ventilation delivered. | Device Design | Hamilton Medical Inc sent a Medical Device Safety Notices letter dated May 8, 2015 by traceable method to affected customers. The letter identified the affected product, problem and actions to be taken. For questions contact Technical Support Manager at 800-426-6331, Ext 215. |  | 154 | 154 total, all part numbers | US Nationwide Distribution. | Terminated | August 12, 2015 | 72 | 154 | Computer
Z-1972-2015 | 71507 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | SYNGO IMAGING (VERSION V30 and V31); Syngo Imaging is a Picture Archiving and Communication System (PACS) intended to display, process, read, report, communicate, distribute, store and archive digital medical images, including digital mammography images. |  |  | June 11, 2015 | 2015 | Siemens Medical Solutions USA, Inc | During certain clinical workflows safety issues may occur involving monitor assignment order and incorrect display and measurements for images with non square pixel size. | Software design | Siemens sent a Customer Advisory Notice, dated June 11, 2015, to their customers. The letter identified the affected product, problem and actions to be taken. Customers were also asked to check the software version and consider whether the issue impacts clinical routine and if a software update is necessary. |  | 5 | 5 | US in the states of NC and NE | Terminated | March 24, 2017 | 652 | 5 | Computer
Z-1959-2015 | 71469 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON SC2000 Ultrasound Systems between software versions VA16A and VA30A and with the Stress Echo Option. |  |  | May 29, 2015 | 2015 | Siemens Medical Solutions USA, Inc. | When performing a 2D Stress Echo study, some following keystrokes result in a potential loss of data, where some selected clips may not be saved as part of the study. | Software design | Siemens sent an Customer Safety Advisory Notification on May 29,2015, to all affected customers. The letter identified the product the problems and the action needed to be taken by the customer.Customer Safety Advisory notifications inform users of the potential issue with the ACUSON SC2000 systems with software versions between VA16! and VA30A and the Stress Echo feature.Letters describe the steps that result in the potential loss of data and suggest that the following steps be taken to remedy the situation:1. Select Stress Echo Selection2. If applicable relabel the clips for all Stages and Views3. Do not return to Stress Echo Selection to select additional relabeled clips.The issue will be resolved in an upcoming software release for customers who did not receive the original software update. Until the necessary repairs have been completed, all personnel in the Organization who need to be aware of the issue should be notified. For further questions please call (650) 694-5398. |  | 573 | 573 devices | Worldwide Distribution - US Distribution and to the countries of : Albania, Angola, Armenia, Australia, Austria, Bangladesh, Belarus, Bolivia, Bosnia Herzog, Brazil, Brunei, Bulgaria, Canada, Colombia, Czech Republic, Ecuador, Egypt, Finland, France, Germany, Hungary, Iceland, India, Ireland, Israel, Italy, Japan, Kazakhstan, Kenya, Kuwait, Lithuania, Malaysia, Maldives, Mauritius, Mexico, Netherlands, Nigeria, Oman, P.R. China, Philippines, Poland, Portugal, Qatar, Rep. of Yemen, Republic Korea, Romania, Russian Fed, San Marino, Saudi Arabia, Serbia, Slovakia, South Africa, Spain, Suriname, Sweden, Switzerland, Taiwan, Thailand, Turkey, U.A.E., United Kingdom and Venezuela | Open, Classified |  |  | 573 | Computer
Z-1958-2015 | 71270 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Cios Alpha; The Cios Alpha is a mobile X-ray system designed to provide X-ray imaging of the anatomical structures of patient during clinical applications. Clinical applications may include but are not limited to: interventional fluoroscopic, gastro-intestinal, endoscopic, urologic, pain management, orthopedic, neurologic, vascular, cardiac, critical care and emergency room procedures. The patient population may include pediatric patients |  |  | May 05, 2015 | 2015 | Siemens Medical Solutions USA, Inc | under certain circumstances the Cios Alpha system may freeze during a procedure. | Software design | Siemens sent an Safety Advisory Notice dated May 5 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. They were alerting them about the issue and risks involved. Customers were provided work-around instructions until a software fix is released. For further questions please call (610) 219-6300. |  | 37 | 37 | Nationwide Distribution | Terminated | January 13, 2017 | 619 | 37 | Computer
Z-1967-2015 | 71352 | No | KHO | Fluorometer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Infinite F500, in vitro diagnostic. |  |  | May 29, 2015 | 2015 | Tecan US, Inc. | A firmware bug causes incorrect values to be displayed for specific measured luminescence values. | Device Design | Tecan sent an Urgent Field Corrective Action letter dated May 29, 2015 to their customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete the attached Customer Response Form. For questions contact your local Tecan Helpdesk. |  | 868 | 868 in total | US Nationwide Distribution in the states of AL, AK, AZ, AR, CA, CO, CT, DE, FL, GA, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MO, NE NH, NJ, NM, NY, NC, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, including DC | Terminated | August 18, 2017 | 812 | 868 | Computer
Z-2404-2015 | 71762 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Artis zee/ zeego systems; dedicated angiography systems developed for single and biplane diagnostic imaging and interventional procedures including, but not limited to, pediatric and obese patients. |  |  | June 30, 2015 | 2015 | Siemens Medical Solutions USA, Inc | in case a system error occurs and the system enters the "Bypass Fluoro" mode while the X-ray locking function is active, the only way to exit the X-ray locking function, would be to either resolve the root cause of the system being in "Bypass Fluoro" or to restart the system. | Software design | A customer advisory notice, dated June 30, 2015, was sent to consignees informing them of the corrective action that would be performed to prevent the possible loss of the image generation function. Siemens plans to contact end users to install a software patch. For questions regarding this recall call 610-219-6300. |  | 586 | 586 | Nationwide Distribution | Open, Classified |  |  | 586 | Computer
Z-1954-2015 | 71505 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Treatment Planning and Delivery System Software version 3.6. ViewRay. |  |  | May 07, 2015 | 2015 | Viewray Incorporated | The software was failing to determine new patient locations if imaging is not enabled during treatment. | Software design | On May 7, 2015 the firm sent Customer Advisory Notification Letters to their customers. |  | 3 | 3 | Distributed in the states of CA, MO &amp; WI. | Terminated | February 04, 2016 | 273 | 3 | Computer
Z-1986-2015 | 71443 | No | ISD | Measuring exercise equipment. | Physical Medicine | Physical Medicine | 510(K) Exempt | Primus (THE BTE WS30), model PrimusRS (PRRS) and Primus+ (PR30) a measuring exercise equipment device designed to provide active and passive torque resistance to patient motion during rehabilitation therapy. The Primus that includes the optional Chop / Lift Bar. |  |  | April 09, 2015 | 2015 | BTE Technologies, Inc. | On the Primus that includes the optional Chop / Lift Bar, there is a remote possibility that the Chop / Lift Bar could get accidently detached from the snap hook in specific use scenarios when the bar is rotated and twisted around the snap hook. The snap hook could open, resulting in possible injury to the user due to a fall. | Device Design | BTE sent an Medical Device Correction letter dated April 10, 2015, to all affected customers.The letter identified the product the problem and the action needed to be taken by the customer.In order to correct this issue, please immediately install the provided retrofit kit according to BTE instructions 751W2519 included in the kit. From this point on, the ring on the Chop / Lift Bar shall not be used to connect the Chop / Lift Bar to the snap hook on the 191 Tool or any other device with a snap hook. Only the Eye part of the Jaw and Eye swivel will be used to connect to the snap hook.Please pass this Medical Device Recall notice to all those who need to be aware of it within the customer organization. Once the correction is completed, notify BTE using the provided Field Correction Confirmation card.We apologize for any inconvenience. Please do not hesitate to call BTE Customer Support at 410-850-0333 if you have any questions or concerns. |  | 510 | 510 Devices | Worldwide Distribution - Nationwide Distribution and to the countries of : Austria, Belgium, Canada, France, Germany, Greece, Hong Kong, Italy, India, Japan, South Korea, Mexico, P.R. China, Poland, Republic of South Africa, Romania, Saudi Arabia, Singapore, Spain, Switzerland, Thailand, and United Kingdom. | Open, Classified |  |  | 510 | Computer
Z-1987-2015 | 71208 | No | ILQ | Powered communication system. | Physical Medicine | Physical Medicine | 510(K) Exempt | Philips GoSafe Communicator 7000L |  |  | April 22, 2015 | 2015 | Lifeline Systems, Incorporated | Home communicator Model 7000L may not sense the phone line under certain unusual landline phone conditions. | Software design | Philips Lifeline notified all consignees of the recall via phone starting on April 22, 2015. The call informs consignees of the issue and actions to be taken. For questions, consignees can contact customer service at 1-800-451-0525 ext. 1608. Consignees were also sent a letter to all consignees instructing them that Lifeline would send out a qualified Installiation Representative for an equipment exchange. |  | 3189 | 3189 devices total, 38 are subject to recall | Nationwide Distribution including AK AZ CA CT FL IA IL KY MA MD MT NC NE NH NY OH OR SC VA WA and WI. | Completed |  |  | 3189 | Computer
Z-2329-2015 | 71675 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Puritan Bennett 980 Ventilator System, Universal (with neonatal functionality enabled) and Neonatal Models. |  |  | July 16, 2015 | 2015 | Covidien LP (formerly Nellcor Puritan Bennett Inc.) | Reports in which tidal volumes reaching patients were lower than set tidal volumes in neonatal Volume Control Plus (VC+) Mode with active humidification. This situation may potentially lead to respiratory compromise if not recognized. | Software design | Consignees will be notified via letter on July 17, 2015. No products will be returned, the reconfiguration of affected ventilators will occur at customer sites. |  | 657 | 657 units | Worldwide Distribution-US (nationwide) and the countries of Canada, United Arab Emirates, Argentina, Australia, Bahrain, Brazil, Chile, Colombia, Costa Rica, Egypt, Ireland, Israel, Jordan, Japan, Kuwait, Lebanon, Morocco, Mexico, Panama, Peru, Qatar, Russia, Saudi Arabia, Singapore, Slovenia, South Africa, Turkey, and Uruguay. | Open, Classified |  |  | 657 | Computer
Z-1990-2015 | 71351 | No | FSY | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | The PowerLED / X-TEN / VOLISTA Surgical Light Systems are surgical light systems intended to be used to provide visible illumination of the surgical area or the patient. Ceiling surgical light systems typically include one or more arms that are able to rotate 360 degrees around a central axis. |  |  | May 29, 2015 | 2015 | Maquet Cardiovascular Us Sales, Llc | Maquet Suspension Arms used with PowerLED/X-TEN/Volista Surgical Light Systems may grind at the junction of the suspension arms which can cause paint chips or particles to fall into the surgical field, when the suspension arms are mounted above the surgical table. | Device Design | Maquet Medical Systems sent an Urgent Medical Device Field Correction letter dated May 29, 2015, to all affected customers via Fedex. The letter identified the product the problem and the action needed to be taken by the customer.The customer was informed that the correction will include inspection of PowerLED / X-TEN / VOLISTA Surgical Light Systems mounted above the surgical table and replacement of the suspension arms if found to be defective. As a remedial measure, pending replacement of the defective suspensions, Maquet personnel will install a protective plastic strip at each junction where the arms are grinding. The correction will be performed by an authorized Maquet Service Technician. The customer will be contacted by a representative of the Maquet Service team to schedule an on-site inspection and possible service of the PowerLED / X-TEN / VOLISTA Surgical Light Systems. If the customer has any questions or require additional information, they can contact their local Maquet representative, or Maquet Customer Service at 1-888-627-8383 (option 3), Monday through Friday, between the hours of 8:00 a.m. and 6:00 p.m. EST. |  | 646 | 646 units | Worldwide Distribution - Nationwide Distribution and to the countries of : Argentina, Austria, Australia, Belgium, Bulgaria, Bahrain, Bolivia, Brazil, Belarus, Canada, Switzerland, Chile, China, Colombia, Croatia, Hungary, Indonesia, Ireland, Iraq, Jordan, Japan, Czech Republic, Germany, Denmark, Ecuador, Egypt, Spain, Finland, France, Greenland, Greece, Hong Kong, United Arab Emirate, United Kingdom, Kenya, Kuwait, Sri Lanka, Luxembourg, Libya, Mongolia, Martinique, Mexico, Malaysia, Namibia, Nigeria, Netherlands, Norway, Oman, Panama, Peru, Philippines, Pakistan, Poland, Saint Pierre and Miquelon, Portugal, Reunion, Romania, Russia, Saudi Arabia, Sweden, Singapore, Slovenia, Slovakia, Senegal, Thailand, Tunisia, Turkey, Trinidad and Tobago, Taiwan, Uruguay, Venezuela, Serbia, Yemen, Mayotte and South Africa. | Open, Classified |  |  | 646 | Computer
Z-1989-2015 | 71457 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Nakoma-SL Anterior Cervical Plating System's Variable Single Barrel Drill Guide |  |  | June 09, 2015 | 2015 | Alliance Partners LLC | Inadequate interference between the screw and the plate when screw implanted into the plate at angulation greater than 7 degrees can result in the screw advancing through the plate during implantation. | Device Design | Alliance Spine sent an Urgent Medical Device Field Action Notification letter dated June 15, 2015 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete the response form and fax to 210-314-2524. |  | 21 | 21 units | US Nationwide Distribution in the states LA, TX, and PR. | Terminated | October 09, 2015 | 122 | 21 | Computer
Z-2390-2015 | 71714 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | RadPro 40kW Digital X-Ray Panel and Mobile/Portable Diagnostic X-Ray. |  |  | July 13, 2015 | 2015 | Sedecal S.A. | Electrostatic energy may be stored in Systems that can cause a short-cicuit which, in turn, can generate an Electomagnetic Pulse which might disturb the electronics of the Digital Motion Control Board producing uncontrolled movements and such uncontrolled movements could cause the system to bump into the user or ppatient. | Device Design | FIELD CORRECTION NOTIFICATION letters dated July 13, 2015 were sent to all dealers. The letters instructed the dealers: to schedule the replacement of the current arrestor by a new design of resistive arrestor with a resistance of 1 M ohm between frame and floor (ground). Although the metallic arrestor is only implemented in the equipments manufactured from April 2013 till March 2015, the upgrade kit has to be implemented in all the installed base of Mobiles with the same configuration in order to ensure that even the arrestors installed as spare part are replaced. Sedecal will provide the necessary upgrade kits to all the customers or distributors. *** U.S. Distributor, Virtual Imaging (a Canon U.S.A. Company) will use letter entitled, FIELD CORRECTION NOTIFICATION, RADPRO 40kW - -DIGITAL MOBILE X-RAY SYSTEM and dated 7/13/2015 to communicate to end users. The letter states "Once Virtual Imaging receives the resistivearrestor, we will contact you to schedule the replacement on the Affected System that you service."Please contact Virtual Imaging, Inc.'s Technical Support for any questions at 561.893.8500 or technicalsupport@vifla.com. **** |  | 361 | 361 systems (U.S.); 5 systems (Foreign) | Distributed US (nationwide) and in Canada. | Open, Classified |  |  | 361 | Computer
Z-2052-2015 | 71456 | No | EBF | Tooth shade resin material. | Dental | Dental | 510(k) | Block Versions Used by Dental Offices often referred to as chairside milling: |  |  | June 15, 2015 | 2015 | 3M Company - Health Care Business | 3M is initiating a Field Correction to remove the crown indication from all Lava Ultimate products. These products will continue to be indicated for onlays, inlays and veneers. | Device Design | 3M ESPE Dental sent an URGENT: MEDICAL DEVICE CORRECTION letter dated June12, 2015 to affected customers. The letter described the affected product, problem and action to be taken. The letter instructed customers to complete and return the Acknowledgement Return/ Refund Form. For questions call 3M ESPE Dental Customer Care center at 1-800-634-2249. |  | 1800000 | 1.8 million (US 1,017,834) | Worldwide Distribution - US: Nationwide and the countries of: Australia, India, Japan, Korea, New Zealand, Singapore, Thailand, Vietnam, Canada, Bulgaria, Czech Republic, Estonia, Hungary, Israel, Latvia, Lithuania, Poland, Romania, Slovakia, Switzerland, Turkey, China Hong Kong, Taiwan, Brazil, Colombia, Costa Rica, Guatemala, Mexico, Egypt, Morocco, South Africa, Austria, Belgium, Denmark, Finland, France, Germany, Greece, Italy, Netherlands Norway, Portugal, Spain, Sweden, UK and Ireland. | Terminated | April 26, 2016 | 316 | 1800000 | Computer
Z-1751-2015 | 71241 | No | DSQ |  |  | Cardiovascular | PMA | HeartWare Ventricular Assist System (HVAD) For use as a bridge-to-cardiac transplantation in patient who are at risk of death from refractory end-stage left ventricular heart failure. |  |  | April 29, 2015 | 2015 | HeartWare Inc | Complaints with the HVAD Discolored and Cracked Driveline Outer Sheath. | Device Design |  |  | 3747 | 3,747 currently in use | AL, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, NC, NH, NJ, NY, OH, OR, PA, SC, TN, TX, UT, and VA, Foreign: Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, Croatia, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, Hong Kong, Hungary, India, Israel, Italy, Japan, Korea, Lebanon, Lithuania, Luxemburg,Malaysia, Netherlands, New Zealand, Norway, Poland, Romania, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Turkey, and UK, | Open, Classified |  |  | 3747 | Computer
Z-2753-2015 | 72036 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Needle Driver; used with the da Vinci Si Surgical System (IS3000); |  |  | August 08, 2015 | 2015 | Intuitive Surgical, Inc. | In certain circumstances involving select da Vinci Si Single-Site instruments, the jaws of grip actuated instruments may become fixed in a closed position. | Device Design | Intuitive Surgical sent an Urgent Device Correction letter to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to do the following:1. Ensure that all affected personnel are fully informed of this notice. Forward this notice to your Risk Manager, OR Director, Purchasing Manager, Biomedical Engineering staff and members of your medical staff who perform da Vinci Surgery procedures.2. Replace any instrument that shows signs of damage and contact Intuitive SurgicalCustomer Service.3. Complete and return the attached Acknowledgment Form to Intuitive Surgical usingthe instructions provided.4. Please retain a copy of this notice as well as the attached instructions with your Single-Site User Manual.If you need further information or support concerning this issue, please contact your Clinical Sales Representative or contact Intuitive Surgical Customer Service at the numbers listed below:Âż North and South America: (800) 876-1310, Option 3 (6 AM to 5 PM PST) or mail:customersupport-servicesupport@intusurg.comÂż Europe, Middle East, Asia and Africa: +800 0821 2020 or +41 21 821 2020 (8 AM to 6PM CET) or ics@intusurg.comÂż South Korea: 02-3271-3200 (9 AM to 6 PM KSTJ)Âż Japan: 0120-56-5635 or 003-5575-1362 (9 AM to 6 PM JST)For questions regarding this recall call 408-523-2100. |  | 89 | 89 total units | Worldwide Distribution - US (nationwide) and Internationally to Australia, Belgium, Brazil, Canada, Chile, China, Cyprus, Denmark, France, Germany, Greece, Israel, Italy, Monaco, Netherlands, Panama, Romania, Saudi Arabia, Singapore, South Korea, Spain, Sweden, Switzerland, Taiwan, Turkey and United Kingdom. | Terminated | December 03, 2015 | 117 | 89 | Computer
Z-2747-2015 | 71985 | No | NXT |  |  | Orthopedic | PMA | Smith &amp; Nephew BIRMINGHAM HIP (TM) RESURFACING DYSPLASIA CUP W/IMPACTOR, REF Numbers: 74120246 (46 MM), 74122248 (48 MM), 74120250 (50 MM), 74122252 (52 MM), 74120254 (54 MM) |  |  | June 03, 2015 | 2015 | Smith &amp; Nephew, Inc. | The firm observed that the BIRMINGHAM HIP(TM) Resurfacing (BHR) System had revision rates which were higher than established benchmarks. | Device Design | Smith &amp; Nephew sent an notification letter dated June 3, 2015 to customers via E-mail and Fed- Ex. The letter identified the affected product, problem and actions to be taken. Customers were asked to quarantine the product and return it. Distributors were asked to notified their customers of the recall and ask them to return the product. |  | 624 | 624 units | Worldwide Distribution - US Nationwide and countries of: Australia, Austria, Bahrain (UAE), Belgium, Canada, China, Denmark, Finland, Germany, Hong Kong, Italy, Japan, Malaysia, Mexico, Netherlands, New Zealand, Norway, Poland, Portugal, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Thailand and United Arab Emirates (UAE). | Open, Classified |  |  | 16778 | Computer
Z-0112-2016 | 72221 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco Radiation Treatment Planning System. |  |  | September 15, 2015 | 2015 | Elekta, Inc. | Dose and MU are incorrect when CT images are viewed from the head, and, when using multiple prescriptions with forced densities. | Device Design | The firm issued an URGENT IMPORTANT FIELD SAFETY NOTICE on 09/15/2015. The letter identified the affected device, as well as the reason for the recall, and clinical impact. Customers were informed that both issues would be resolved in Monaco Patch Release 5.10.01. The attached acknowledgement form should be completed, signed, and returned to a local Elekta Office or representative within 30 days. Questions regarding the notice should be directed to a local Elekta office. |  | 154 | 154 units | Worldwide Distribution -- US, including the states of CA, IN, MI, MO, TX and WI; and, the countries of Australia, Canada, Germany, Greece, India, Netherlands, New Zealand, Turkey, and the UK. | Terminated | April 04, 2017 | 567 | 154 | Computer
Z-0126-2016 | 72243 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Synthes Screw Inserter; for use with Intramedullary Rod, orthopedic device. Used to secure the head element to the head element inserter to prevent disengagement during head element insertion. |  |  | September 21, 2015 | 2015 | Synthes (USA) Products LLC | Certain lots of the Screw Inserter may be etched with incorrect graphics. The orientation of the etched image of the oblique angle on the Screw Inserter may not correspond to the tip of the angle of the instrument. | Device Design | DePuy Synthes sent an urgent notice of medical device recall letter dated September 21, 2015, to all affected customers to notify them about the issue, the potential risks, and to provide instructions for return of the completed response form and affected devices. Customers with questions were instructed to call 610-719-5450. |  | 793 | 793 | US Nationwide Distribution | Open, Classified |  |  | 793 | Computer
Z-1577-2015 | 71118 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Ultraview SL Command Modules, Model 91496, with the Masimo SpO2 PCBA, PN: 010-1636-02. |  |  | April 21, 2015 | 2015 | Spacelabs Healthcare Inc | Modules resetting. This failure mode happens when the Spacelabs Command Module is configured with the Masimo SpO2 option (-M), the Spacelabs Respiration option (-R) and the Masimo SpO2 PCBA, PN: 010-1136-02. With these two options and the Masimo SpO2 PCBA present, the module may experience random resets. | Software design | Spacelabs expanded the recall to include additional 29 units and notified one additional new consignee on June 4, 2015.*******************************************************************************************The firm, Spacelabs Healthcare, sent an "URGENT-MEDICAL DEVICE CORRECTION" letter dated 27 April 2015, via priority mail to US consignees and email to foreign consignees. The letter described the product, problem and actions to be taken. The consignees were instructed to circulate this notification to all persons concerned by the information it contains; add a copy to the appropriated system user manuals until such time as corrections have been implemented by Spacelabs, and weigh the benefits versus the risks when deciding whether or not to continue to use your modules until they can be upgraded.Only modules shipped or repaired after March 8, 2014 are affected by this recall.Affected command modules will be upgraded with a new version of software by Spacelabs at no cost.Customer with questions about this corrective action, please contact Spacelabs Healthcare at 1-800-522-7025 and select 2 for Technical Support. |  | 1973 | 1973 units total (1345 modules in the US and 604modules international and 3 PCBAs in US and 21 PCBAs international) | Worldwide Distribution:US (nationwide) including states of: AL, AR, AZ, CA, FL, GA, IA, ID, IL, LA, ME, MI, MN, MO, NE, NJ, NY, OH, OR, PA, SC, TN, TX, VA, WA, WY, and Puerto Rico and countries of: Argentina, Australia, Bangladesh, Canada, Chile, France, Great Britain, India, Kuwait, Panama, Poland, Qatar, Saudi Arabia, and Turkey.. | Terminated | November 16, 2015 | 209 | 1973 | Computer
Z-0117-2016 | 71991 | No | GWQ | Electroencephalograph. | Neurology | Neurology | 510(k) | NicoletOne Software v5.94, Catalog/Part Number: 482-649600. |  |  | September 11, 2015 | 2015 | Natus Neurology Inc | When using the NicoletOne v5.94 software, after exiting the impedance check function and returning to the EEG screen, the impedance check signal remains active in waveform, obscuring the EEG signals. | Software change control | Natus Neurology sent an Ugent Medical Device Corrrection Notification letter dated September 9, 2015 to all affected consignees. The letter described the problem and the product involved in the recall. Consignees were informed consignees that a Field Service Technician will contact them to install the new software. Once the software is installed they will ask the Recipient or Field service Technician to complete and return the Verification Form.For questions regarding this recall call 608-829-8600. |  | 101 | 101 (88 US, 13 OUS) | Worldwide distribution. US nationwide, INDIA, ROMANIA, DENMARK, JAPAN, AUSTRALIA, RUSSIA, BANGLADESH, SRI LANKA, BRUNEI, and UNITED ARAB EMIRATES. | Terminated | March 08, 2016 | 179 | 101 | Computer
Z-0115-2016 | 72166 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MRIdianÂż ViewRayÂż Radiation Therapy System, ViewRayÂż |  |  | August 27, 2015 | 2015 | Viewray Incorporated | ViewRay discovered that in the event that an encoder breaks or fails on the Patient Handling System (PHS, or couch), when attempting to restart the system, the couch could move unexpectedly. | Device Design | ViewRay sent an Customer Advisory Notice dated August 27, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.It is recommended that if you have a patient on the couch and you receive the fault ''The patient table has experienced an unintended movement. This may be due to the patient or operator leaning on the table.Clear the fault and retry. If the problem persists call Service" on the TPDS,first attempt to clear the fault.If it does not clear and requires you to restart the couch software (same switch as restarting the RTCS) you should manually remove the patient from the couch before restarting the system.A software patch to address this issue is planned for early 2016. In the meantime please practice our recommendation of removing the patient from the couch prior to restarting the RTCS. If you have any questions or concerns regarding this issue, please do not hesitate to contact ViewRay Customer Support at support@viewray.com or call the ViewRayÂżÂż Customer Support center at 855-286-8875. |  | 4 | 4 units | US Distribution to the states of : CA, MO and WI., and Internationally to Korea. | Terminated | June 09, 2017 | 652 | 4 | Computer
Z-1583-2015 | 71029 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | AXIOM ARISTOS FX; Multipurpose Radiography System (MPRS) is a dedicated x-ray system with a flat panel detector which allows the acquisition of x-ray exposures without the use of conventional film/screen systems. The MPRS allows radiographic exposures of the whole body including skull, spinal column, chest, abdomen, extremities, and excluding mammography. Radiographic exposures may be taken with the patient in the sitting, standing, or supine positions. The intended use and indications for use of this modified device as described in its labeling have not changed from its predicate device. |  |  | April 14, 2015 | 2015 | Siemens Medical Solutions USA, Inc | If the cable cart and the corrugated hose are unable to move freely, the cable holder could drop down and hit patients or system operators. | Device Design | The firm, Siemens Healthcare, sent a "Safety Advisory Notice" dated April 14, 2015, to its customers. The notice described the product, problem and actions to be taken. The customers were instructed to follow the instructions in the notice; promptly notify and instruct all the staff at your organization making them aware of the problem; and forward this safety information to other organizations that could be affected. If the device has been sold and therefore no longer in your possession, forward the safety notice to the new owner. Additionally, customers were instructed observe this safety notice and comply with the corresponding measures until the update has been fully completed. If you have any questions, contact customer support at 610-219-6300. |  | 43 | 43 | Nationwide Distribution. | Open, Classified |  |  | 43 | Computer
Z-1578-2015 | 71084 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | EPIQ 5 Ultrasound System with Pediatric Cardiology option, Model: EPIQ 5G, EPIC 5C, EPIQ 5W. Catalog Number: 795204 / 795205. Part Number: 989605408541 With |  |  | April 10, 2015 | 2015 | Philips Ultrasound, Inc. | When Epiq 5 Ultrasound System, WITH the Pediatric Cardiology option, is set up to Metric and weight and/or height is entered, a unit conversion error may result in the incorrect calculation of the patients Body Surface Area (BSA). | Software design | Philips sent the Urgent-Medical Device Correction-Philips Model EPIQ5 and Ultrasound system letter, dated 2015 APRIL 10, and the Urgent-Medical Device Correction-Philips Model EPIQ7 and Ultrasound system letter, dated 2015 APRIL 10, letters to their consignees.These systems WITH Pediatric Cardiology option are affected.Philips issued two Field Change Orders (FCO) to their Field Service Engineer to install upgraded software: -FCO 79500336 for EPIQ5-FCO 79500337 for EPIQ7Customers with questions can call Philips Customer Service at 1-800-722-9377. |  | 751 | 751 units total (246 units in the US and 505 units outside the US) | Worldwide Distribution - US (nationwide and DC) and Internationally to Norway, India, New Zealand, Bahrain, Sweden, Italy, Australia, France, Netherlands, Switzerland, Germany, Austria, and Turkey. | Terminated | February 29, 2016 | 325 | 751 | Computer
Z-1580-2015 | 71050 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ViewRay System, Radiation Therapy System |  |  | January 15, 2015 | 2015 | Viewray Incorporated | The software was not correctly using the RT (Radiation Therapy) to MR (Magnetic Resonance image) coordinate correction for non HFS (Head First Supine) patient orientations, resulting in slice mismatch error. | Software Design Change | On January 15, 2014, ViewRay personnel discussed the investigation of the issue (slice mismatch error) with the customer. The firm advised the customer that until the issue was corrected, to not treat patients in non HFS orientations, or to use the MR acquired from the ViewRay system as the primary plan series for non HFS scans. ViewRay installed software version 3.4.5.28 on January 20 and 21, 2014, in order to prevent this slice mismatch error from occurring in any patient position. |  | 1 | 1 | US distribution to MO. | Terminated | September 23, 2015 | 251 | 1 | Computer
Z-0114-2016 | 72170 | No | FSM | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Disposable Paper Filter for Sterilization Containers |  |  | September 17, 2015 | 2015 | Stryker Craniomaxillofacial Division | Stryker has become aware that the inhomogeneity of the filter paper might potentially compromise the ability of the filter to maintain a sterile barrier during post-sterilization shelf-life. | Device Design | An Urgent Medical Device Recall letter has been sent to the customers. The affected product was identified, as well as the reason for the recall and risk to health. The letter states that customers should inform users of the Medical Device Removal and pass the notice to all appropriate individuals within the organization; scrap all affected products available at the location or return it to Stryker Craniomaxillofacial, Attention: Recall Coordinator, at the address provided. If there are questions regarding the impacted product, don't hesitate to Contact Stryker Customer service and refer to PFA #2015-032. Complete and sign the enclosed Business Reply Form and fax or copy to: (877) 648-7114 or e-mail a copy to CMF-custserv@stryker.com. |  | 915 | 915 packs | Worldwide distribution -- US, Canada, France, Germany, Italy, Japan, Norway, Sweden, Switzerland, and the Netherlands. | Terminated | February 26, 2016 | 162 | 915 | Computer
Z-0111-2016 | 72113 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Specialist 2 Intramedullary (SP2 IM) Rod 400mm Instrument (PN 96-6120) |  |  | September 08, 2015 | 2015 | DePuy Orthopaedics, Inc. | DePuy Orthopaedics, Inc. is initiating a voluntary recall of select lots of the Specialist 2 Intramedullary Rod 400mm Instrument (PN 96-6120) due to complaints of the rod fracturing and leaving a portion of the rod in patients. | Device Design | DePuy Synthes sent an Urgent Information Recall Notice to all affected customers. Instruments are not lot traceable in the system, so the company sent notices by e-mail on September 8, 2015, and US Distributors will hand deliver notices to all Medical Facilities that may have purchased the affected instrument. The US Distributor or US Sales Consultant will deliver the Medical Professional Notice and Reconciliation Form to the affected Medical Facilities. If assistance is needed, Sales Consultants will assist customers with inspection of all instruments, return of affected lots, receiving credit and returning the completed Medical Facility Reconciliation Forms to DePuy Orthopaedics, Inc. The affected instruments will be returned to the quarantine warehouse located at 700 Orthopaedic Drive, Warsaw, IN 46582. For questions regarding this recall call 574-371-4917. For questions regarding this recall call 574-267-8143. |  | 7488 | 7488 | Worldwide Distribution - US including AL; AZ; CA; CO; CT; DC; FL; GA; HI; IA; IL; IN; KY; LA; MA; MD; ME; MI; MN; MO; MS; MT; NC; NE; NH; NJ; NV; NY; OH; OR; PA; RI; SC; TN; TX; UT; VA; VT; WA; WI; and WV and Internationally to ARGENTINA; AUSTRALIA; AUSTRIA; BELGIUM; BRAZIL; CANADA; CHILE; COLOMBIA; CZECH REPUBLIC; EGYPT; FRANCE; GERMANY; GREECE; INDIA; ISRAEL; ITALY; KOREA; Mexico; NEW ZEALAND; NORWAY; P R CHINA; PANAMA; PERU; POLAND; RUSSIA; SINGAPORE; SLOVENIA; SOUTH AFRICA; SWITZERLAND; UAE; UK; URUGUAY; VENEZUELA; and VIETNAM. | Open, Classified |  |  | 7488 | Computer
Z-0109-2016 | 72176 | No | LTT | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | MICroSTREP plus 2 Panel, Part No. B1027-202 |  |  | September 10, 2015 | 2015 | Beckman Coulter Inc. | Beckman Coulter is recalling MICroSTREP plus 1 and MICroSTREP plus 2 panels because a manufacturing error may lead to antimicrobic degradation and the potential for elevated or false resistant results. | Equipment maintenance | Beckman Coulter Inc. sent an Urgent Medical Device Recall letter dated Setember 10, 2015 to all customers who purchased the MICroSTREP plus 1 and MICroSTREP plus 2. The letter identified the affected products, problem and the actions to be taken. Customers are instructed to complete and return the enclosed Response Form within 10 days. Customers with questions regarding the notice, are instructed to contact their Customer Support Representative:website: http://www.beckmancoulter.com or by phone: 1-800-677-7226 in the United States and Canada. |  | 438 | 438 units total (433 units in US) | Worldwide Distribution - US Nationwide and the countries of Bahamas, Belgium, Canada, Ecuador, Germany, Italy, Republic of Korea, Mexico, Peru, Poland, Romania, Saudi Arabia, South Africa, Thailand, United Sates (including Puerto Rico), and Uruguay. | Terminated | January 22, 2016 | 134 | 438 | Computer
Z-1582-2015 | 70951 | No | IYH | X-ray radiation therapy system. | Radiology | Radiology | 510(k) | ExacTrac 6.x. is software used to place patients at an accurately defined point within the treatment beam of a medical accelerator for stereotactic radiosurgery or radiotherapy procedures. |  |  | March 06, 2015 | 2015 | Brainlab AG | ExacTrac 6.x Patient Positioning System: Potentially incorrect patient positioning when using the ExacTrac Cone Beam CT (CBCT) with a CBCT acquired at a couch angle other than 0.0 degrees. | Software in the Use Environment | FIELD SAFETY / PRODUCT NOTIFICATION letters dated March 6, 2015 were sent to all customers of record. The letters included instructions for customers to: 1) With immediate effect, always make sure the couch is at exactly 0.0 degrees (displayed on the Varian monitor) when acquiring a CBCT set to import into the ExacTrac software; 2) For Varian systems providing an option to prevent a CBCT acquisition for couch angles other than 0.0 degrees, Brainlab recommends selecting that option for use in combination with ExacTrac; and, 3) Continue to verify any ExacTrac CBCT based correction using the ExacTrac X-ray verification and/or an external IGRT procedure as mandatory. The letters further advised that Brainlab will be taking the following actions: 1. Brainlab provides existing potentially affected ExacTrac v.6.x CBCT Import &amp; Alignment Software customers (with a Varian C-Series linac) with this product notification information; and, 2. Brainlab will provide a software update with this issue solved to affected customers. Brainlab will actively contact the customers starting August 2015 to schedule the update installation. Customers with questions about this issue were advised to contact Brainlab at 800-597-5911 or via e-mail by writing at us.support@brainlab.com. |  | 25 | 25 systems (US only); 53 systems (Foreign) | *** US: Nationwide; *** FOREIGN: Austria, China, Denmark, France, Germany, Hong Kong, Japan, Netherlands, Qatar, Russian Federation, South Korea, Switzerland | Open, Classified |  |  | 25 | Computer
Z-2759-2015 | 71877 | No | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Biomet Spine Lineum OCC Starter Flexible and Lineum Flexible Inner Shaft, Model Numbers 14-525050 and 14-525057, reusable devices distributed with bubble wrap inside a plastic pouch. |  |  | July 22, 2015 | 2015 | Biomet Spine, LLC | Biomet has initiated a recall on Biomet Spine Lineum OCC Starter Flexible and its inner shaft instrumentsafter identifying that these instruments were distributed prior to completion of all design control activities. | Device Design | The firm, Zimmer Biomet, sent an "URGENT MEDICAL DEVICE RECALL NOTICE" letter dated 7/23/2015 to their Consignees/Customers via certified letter. The letter described the product, problem and actions to be taken. The customers were instructed to immediately locate and discontinue use of device; remove from circulation; carefully follow instructions on enclosed "Response Form" and complete and return via fax to 303.501.8444; if you have further distributed these devices to medical facilities, you MUST notify them of this action; and return device and form to Zimmer Biomet Spine, 310 Interlocken Pkwy, Ste 120, Broomfield, CO 80021. Questions related to this notice should be directed to the Quality Engineer at 303-501-8457 Monday through Friday, 8 a.m. to 5 p.m. MST. |  | 49 | 49 | Worldwide Distribution: US (nationwide) in states of: NY, MO, WI, IN, CA, NC, and OH, and to countries of: Japan and Netherlands. | Terminated | December 10, 2015 | 141 | 49 | Computer
Z-0005-2016 | 71763 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Celesteion PCA-9000A/2 PET/CT System |  |  | May 08, 2015 | 2015 | Toshiba American Medical Systems Inc | It was found that if specific operations are performed in multi-phase helical scanning, the acquired raw data may not be saved due to a software problem. | Device Design | Planned Action by Toshiba Medical America Systems to address the defect:a. A notification will be sent to the customer with affected system in compliance with 21 CFR1 003.21 upon approval by the Secretary.b. To correct this issue revised software will be installed on the system to prevent this occurrence from happening. The consignee will be contacted by their Toshiba service representative for an appointment when the revised software becomes available.c. The customer will be advised to do the following until the corrective measures have been taken, i.e., revised software installed on their systems. If you have any questions please contact your local Toshiba Representative at (800) 521-1968. |  | 1 | 1 | US Distribution to the state of Nevada. | Terminated | December 03, 2015 | 209 | 1 | Computer
Z-1587-2015 | 70933 | No | HWE | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | Orthopedic | 510(K) Exempt | Universal Charger |  |  | March 25, 2015 | 2015 | Stryker Instruments Div. of Stryker Corporation | The Stryker Universal Battery Charger is not transmitting usage data to the Stryker Cloud as designed. The firm is initiating a software correction to address the issue. | Software design | Stryker issued e-mail notices to distributors and reps on March 26, 2015 advising that all battery charges would receive a software update to provide a field correction for the upgrade to units to be able to track device usage data to Stryker as designed. There is no customer action required, the chargers will transmit a message to Stryker once the upgrade is complete. Please contact Kelly Jo Davis at 269-389-2921 with any questions. |  | 223 | 223 chargers | US Nationwide in the states of AL, AR, AZ, CA, CO, FL, GA, IA, IL, IN, KS, KY, MD, MI, MO, MS, MT, NC, NH, NV, NY, OH, OK, PA, SC, TN, TX, UT, VA, VT, WA, WI and the foreign country of: Canada | Terminated | July 08, 2015 | 105 | 223 | Computer
Z-0106-2016 | 71979 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Syngo RT Oncologist, an optional accessory to the linear accelerator system and permits localization, contouring, segmentation, image review, and review and approval of treatment plan parameters. In addition, it includes tools and administrative functions to aid in the diagnosis, staging, and prescription of radiation therapy. |  |  | August 11, 2015 | 2015 | Siemens Medical Solutions USA, Inc | This update is intended to provide a software update for the Syngo RT Oncologist, which is currently running SW version 4.2 or 4.3. The safety-related issues, which were described in the Customer Safety Notice distributed as UFSN-RTT/RTO 4.3 Adaptive Targeting [Auto-registration]. The update contains several important safety and performance fixes. | Software design | The firm, Siemens, sent Recall letters, dated August 2015, that describe the product, the problem, and correction, will be hand-delivered to affected sites at the time the fix is implemented by a Siemens technician.If you have any questions, contact Regulatory Technical Specialist at 610-448-6471 or email: marlynne.galloway@siemens.com |  | 14 | 14 | US Distribution to states of: MA, MO, OH, PA, UT, and WI. | Open, Classified |  |  | 14 | Computer
Z-0107-2016 | 72137 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | SOMATOM Force; computed tomography x-ray system. |  |  | August 18, 2015 | 2015 | Siemens Medical Solutions USA, Inc | Software and firmware bugs | Software design | Customers will be notified via the system interface and e-mail or U.S. postal mailing via a Customer Advisory Notice. If necessary the advisory letter may be hand delivered. The letter informed customers that the safety related issues identified were resolved by SW-Version VA50A_SP2. The update will be remotely pushed to affected systems. The letter also advised customers to notify anyone in their organization that should be aware of the information. In the event that the device had been sold, the notice should be forwarded to the new owner, and the firm should be informed of the new owner. |  | 21 | 21 | Nationwide Distribution | Terminated | May 08, 2017 | 629 | 21 | Computer
Z-1606-2015 | 71211 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, CARESCAPE Monitor B850, B650 and B450 (Bx50). |  |  | May 04, 2015 | 2015 | GE Healthcare | The heart rate could be calculated from pacer pulses without indication that pacemaker detection is OFF in some combinations of Bx50 monitors, a PDM, and CIC/CSCS. Then the monitor does not show the Pacer Off message and there may not be an alarm for asystole. Undetected asystole could result in irreversible changes in the patientÂżs condition and delayed or missed life sustaining patient treatment | Software design | GE Healthcare sent an Urgent Medical Device Correction letter dated May 1, 2015, to all affected consignees. The letter was addressed to Healthcare Administrator / Risk Manager, Chief of Nursing, and Director of Biomedical Engineering. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. For questions contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 110 | 110 (83 units US, 27 units OUS) | Worldwide Distribution - US including CO, MI, WI and Internationally to NETHERLANDS, FINLAND, and BELGIUM. | Terminated | September 15, 2015 | 134 | 110 | Computer
Z-1607-2015 | 71135 | No | LMC | Medical image hardcopy device. | Radiology | Radiology | 510(K) Exempt | Nuclear gamma cameras (ADAC VERTEX PLUS/SOLUS/CARDIO, ADAC VERTEX V60 &amp; ADAC CARDIO C60) |  |  | March 11, 2015 | 2015 | Philips Medical Systems (Cleveland) Inc | The firm was notified by a customer that the collimator exchange carriage and the collimator storage cabinet were misaligned. | Device Design | Philips sent an Urgent - Medical Device Correction letter dated March 12, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The firm requested their customers to do the following: " If a misalignment on the top and/or bottom of the guide rail between the cabinet and carriage can be visually detected do not continue with the collimator exchange procedure," If the misalignment is visually detected, be aware that the customer/user cannot remedy this issue and will require service interaction prior to the continued use of the system." This letter should be placed in your User Documentation until otherwise notified.If you need any further information or support concerning this issue, please contact your local Philips Healthcare representative: For North America and Canada contact the Customer Care Solutions Center (1-800-722-9377, Option 5: Enter site ID or follow the prompts). |  | 975 | 975 units | Worldwide Distribution - US (nationwide) to the states of: AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV and WY., and to the countries of : Australia, Austria, Belgium, Brazil, Canada, China, Costa Rica, Egypt, France, Germany, Hong Kong, Ireland, Italy, Japan, Korea, Lebanon, Netherlands, Norway, Pakistan, Portugal, Russia, Singapore, South Korea, Spain,Thailand and United Kingdom. | Open, Classified |  |  | 975 | Computer
Z-0020-2016 | 72091 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | SOMATOM Definition Flash; intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data from either the same axial plane taken at different angles or spiral planes taken at different angles. |  |  | August 14, 2015 | 2015 | Siemens Medical Solutions USA, Inc | software bug issues for SW-Version VA48A_SP0. The following safety issues were resolved: 1) Correction to improve visual warning and error indication son the gantry display. 2) Correction to improve acquisition data in order to optimize image quality. 3) Correction to improve robustness and general system behavior in some exception handling procedures. 4) Correction to improve auto post processin | Software design | A customer advisory notice, dated 8/14/15, was sent to direct accounts advising them of the issue to observe the safety notice and comply with the corresponding measure until the update has been fully completed. |  | 113 | 113 total | Nationwide Distribution. | Open, Classified |  |  | 113 | Computer
Z-0023-2016 | 72212 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Centricity PACS IW by GE Healthcare Dynamic Imaging Solutions is a device that receives medical images (including mammograms) and data from various imaging sources. Images and data can be stored, communicated, processed and displayed within the system or across computer networks at distributed locations. |  |  | September 16, 2015 | 2015 | GE Healthcare | Images may be missing when a system parameter MapRoute is set to a value greater than 1. | Software Design Change | The firm, GE Healthcare, an "Urgent Medical Device Correction" by mail or hand delivered to all sites. The letter instructs the customer to ensure that all potential users in your facility are made aware of this safety notification and the recommended actions. Until GE Healthcare conducts remote inspection of the systems "MapRoute" parameter value. If the value is found to be greater than the recommended setting of 1, GE will reset this system parameter to a value of 1. The letter informs the user facilities not to make any changes to this parameter value.Additionally, until a product modification is available to correct this issue, the following actions are recommended. The firm is recommending user facilities to: 1. Utilize the image count within the QC process to alert the user of a discrepancy in the number of transmitted images from the modality to the number of images available in the PACS IW viewer.a. If a discrepancy is identified, attempt to retransmit the exam to PACS.b. If retransmission is unsuccessful, contact your GE Healthcare Service representative for assistance and instructions as to the impacted exam. Urgent cases impacted by this issue should be interpreted at the modality.2. If DICOM storage commit is configured and in use, no commit notification is sent to the modality for images that were impacted by this issue.GE Healthcare will correct all affected systems by providing a software update at no cost to the user facilities. If the user facilities have questions, they are encouraged to communicate questions or concerns regarding this notification to GE Healthcare Service at 1-800-437-1171 or their local Service Representative. |  | 921 | 921 systems are impacted. | Worldwide Distribution: US (nationwide) including states of: AL AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, ON, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, WY, and Puerto Rico; and countries of: Argentina , Austria, Bahrain, Belgium, Bosnia and Herzegovina, Brazil, Canada, Chile , China, Colombia, Costa Rica, Ecuador, Egypt, France, Germany, Great Britain, Hong Kong, Hungary, India, Israel, Italy, Japan,Korea, Kuwait, Lebanon, Libya, Lithuania, Malaysia, Mexico, Netherlands, New Caledonia, Nigeria, PanamÂż, Peru, Philippines, Poland, Qatar, Russia, Saudi Arabia, Slovenia, South Africa, Spain, Switzerland, Turkey, Venezuela, and Vietnam. | Terminated | January 05, 2017 | 477 | 921 | Computer
Z-0024-2016 | 72238 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Synthes Periarticular Insertion Handle for 4.5mm LCP Proximal Tibia Plates and Periarticular Aiming Arm for 4.5mm LCP Proximal Tibia Plates |  |  | September 14, 2015 | 2015 | Synthes (USA) Products LLC | Affected lots of Periarticular Aiming Arms and Insertion Handles (Part Numbers 03.120.001, 03.120.002, 03.120.004 and 03.120.005) for the 4.5 mm LCP Proximal Tibia Plates are not able to be secured to the subsequent versions of these devices as intended. | Device Design | An urgent notice of medical device recall, dated September 14, 2015, was sent to end users and sales consultants to alert them to the issue and potential risks. Return of the completed response form and the affected devide was also requested. |  | 172 | 172 | Nationally | Open, Classified |  |  | 172 | Computer
Z-1608-2015 | 71233 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Spacelabs qube Compact Monitor, Model 91390. The monitor is functioning as either bedside or central monitors; passively displays data generated by Spacelabs Healthcare parameter modules, Flexports interfaces, and other SDLC based products in the form of waveform and numeric displays, trends and alarms. |  |  | May 08, 2015 | 2015 | Spacelabs Healthcare Inc | The input power can fail even though the power source at the wall has not been interrupted. In addition, there is the potential for the qube Compact Monitor to fail to turn on. This is a retrospective recall related to a recall initiated in 2012 (Z-0403-2013). | Software design | The firm sent the Urgent Medical Device Correction- Spacelabs Healthcare Inc.-qube Compact Monitor, Model 91390, Power Input PCBA Failure letter, dated 8 May 2015, via priority service mail to US consignees.This letter described the affected system, problem, actions to be taken by users and corrective actions conducted by Spacelabs.For additional information or technical assistance, please contact:Technical SupportSpacelabs Healthcare, Inc.35301 SE Center St.Snoqualmie, WA 980651-800-522-7025 and select 2 for Technical Support |  | 5 | 5 units distributed in the US | distributed in North Carolina, Texas, and Wyoming | Terminated | November 05, 2015 | 181 | 5 | Computer
Z-0125-2016 | 72133 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Inserter for Titanium Elastic Nails (TEN), orthopedic device |  |  | September 03, 2015 | 2015 | Synthes (USA) Products LLC | potential for mechanical failures such as breakage. | Device Design | Synthes sent an Urgent Notice: Medical Device Recall letter dated September 3, 2015, to all affected customers. The letter identified the product the problem and the action need to be taken by the customer. ACTIONS REQUIRED:We have record that your facility has received the product(s) subject to this recall. It is important to note that the removal of the product would prevent the performance of emergency surgery, thus DePuy Synthes is not requiring an immediate removal of affected product(s). DePuy Synthes has developed a replacement &amp; recovery plan for indicated devices and replacement parts are becoming available. If you have questions on replacement timing and availability, please call DePuy Synthes at 1-800-479-6329.If you have questions regarding the return and recovery of the part, we ask that you contact your local DePuy Synthes Sales Consultant. |  | 1550 | 1550 | US (nationwide) distribution and one Internationally to Australia. | Open, Classified |  |  | 1550 | Computer
Z-1574-2015 | 70638 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MHI-TM2000 Linear Accelerator System (Software Version 3.5.0 and 3.5.1) |  |  | April 17, 2015 | 2015 | MITSUBISHI HEAVY INDUSTRIES, LTD., HIROSHIMA MACHINERY WORK | The operator console allows users to deliver therapeutic radiation to patients even though a specific communication error relevant to imaging conditions has occurred. In addition, the treatment record (the delivered radiation record) cannot be saved. | Software design (manufacturing process) | MITSUBISHI HEAVY INDUSTRIES, LTD. has issued a Safety Alert for the MHI-TM2000 Linear Accelerator System dated April 10, 2015, due to a software issue.Customers were advised not to deliver therapeutic X-ray if popup warning message show is displayed during treatment. Instructions were provided if therapeutic radiation is delivered during popup warning. Customers with questions should contact the Customer Service Representative by calling 1-800-597-5911, Monday through Friday or email: us.support@brainlab.com. **CAPA** Correction software for this issue has a tentative availability date of June 2015. Estimated software update completion date would be end of October 2015. |  | 11 | 11 units total | Worldwide Distribution - US including FL, NY, OH, &amp; TX, **Internationally - Belgium, Italy, Korea, France &amp; Germany. | Terminated | April 11, 2017 | 725 | 11 | Computer
Z-1559-2015 | 70899 | No | MDZ | Medical washer-disinfector. | General Hospital | General Hospital | 510(k) | Caviwave Pro Ultrasonic Cleaning System, |  |  | February 19, 2015 | 2015 | Steris Corporation | Should the Caviwave Pro Ultrasonic Cleaning System experience a temporary interruption in electrical power during a processing cycle, the circulation pump may stop functioning. In addition, it is possible that the ultrasonic generator may stop working. The machine does not provide an alert notifying the user of this device malfunction. | Software design | STERIS sent an Urgent Field Correction Notice letter dated March 9, 2015 customers. The letter identified the description of the product, description of the problem and actions to be taken. For questions contact STERIS Customer Service at 1-800-548-4873 or local sales representative. |  | 491 | 491 units | Worldwide Distribution - US Nationwide in the states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NH,NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI &amp; WV and the countries of: Australia, Brazil, Canada, France, Mexico, Saudi Arabia, Spain, Taiwan &amp; United Arab Emirates. | Terminated | October 21, 2015 | 244 | 491 | Computer
Z-1553-2015 | 70919 | No | HWE | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | Orthopedic | 510(K) Exempt | Synthes Small Battery Drive (SBD) Oscillating Saw Attachment Part Number: 532.021 for orthopedic surgery. |  |  | March 24, 2015 | 2015 | The Anspach Effort, Inc. | The potential exists for the Oscillating Saw Attachment to disengage from the Small Battery Drive handpiece. If the attachment disengages from the handpiece intra-operatively, there is a potential for injury to the patient and/or user. | Device Design | DePuySynthes sent a "URGENT NOTICE: MEDICAL DEVICE RECALL- CORRECTION" letter dated March 24, 2015 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. The firm requests that customers review the inventory and perform the in-process inspection outlined in the letter attachment, a review and repair will be implemented if necessary during the next visit schedule maintenance interval at the customer site.Customers having questions regarding this notice should contact the firm Complaint Handling Unit Manager at 561-494-3673 or to contact their Depuy Synthes Sales Consultant.For questions regarding this recall call 561-627-1080. |  | 12085 | 12,085 pieces | Worldwide Distribution - US (nationwide including DC, Hawaii, and Puerto Rico) and Internationally to Argentina, Austria, Australia, Belgium, Bahrain, Bolivia, Brazil, Canada, China, Chile, Costa Rica, Czech Republic, Germany, Denmark, Ecuador, Egypt, Spain, Finland, France, Great Britain, Greece, Guatemala, Hong Kong, Croatia, Hungary, Indonesia, Ireland, Israel, India, Iraq, Iran, Italy, Jordan, Japan, Kenya, Korea, Kuwait, Lebanon, Sri Lanka, Libyan Arab Jamahiriya, Morocco, Macedonia, Mexico, Malaysia, Netherlands, Norway, Nepal, New Zealand, Oman, Ontario, Peru, Philippines, Palestine, Portugal, Romania, Serbia, Russian Federation, Saudi Arabia, Sweden, Singapore, Slovenia, Slovakia, South Africa, El Salvador, Thailand, Turkey, Taiwan, Ukraine, Uruguay, Venezuela, Vietnam and United Arab Emirates. | Open, Classified |  |  | 12085 | Computer
Z-1463-2015 | 69996 | No | HIG | Hysteroscopic insufflator. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Fluid Management System P4000; for irrigation and fluid warming in laparoscopic procedures, and distention, fluid warming, and volume/deficit measurements in endoscopic procedures within gynecology, urology, and orthopedic procedures. |  |  | September 01, 2014 | 2014 | Thermedx LLC | To correct software bugs that could affect the ability to accurately measure fluid deficit. | Software design | Thermedx sent an Urgent Product Correction letter dated March 13. 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The firm informed their customers that the software update will be performed at customer locations by Thermedx trained and authorized personnel. A Thermedx and/or your Sales Representative will contact customers to schedule a time to perform the field service update.Should you have any questions regarding this notification, please call a Thermedx technical service representative. Support is available toll free at 1-888-542-9276 from 8AM to 5PM Monday to Friday Eastern Standard Time or by e-mailing techservice@thermedx.com. Our website may be found at www.thermedx.com. We sincerely apologize for the inconvenience and thank you in advance for your cooperation. |  | 127 | 127 units | US Distribution to the states of : OH, MS, MA, WA, MI, NC, NY, WV, LA, TX and IL. | Terminated | September 14, 2015 | 378 | 127 | Computer
Z-1480-2015 | 70710 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Captiva's lumbar intervertebral fusion system Reamers (color code rings on the Reamers). |  |  | February 25, 2015 | 2015 | Captiva Spine, Inc | It is possible for the titanium color-coded ring to detach from the Reamers. | Device Design | On 02/26/2015 the firm sent a Voluntary Correction Notification to its consignees via FedEx. Within that notification the firm asks to please check with your product representative and Central Supply Departments to determine if you have any of the products listed on the notificaiton. If any product is located please call 877-772-5571 from 8:30am (Eastern Time) to receive instructions as to how to remove the color-coded ring OR please follow the enclosed steps for the correction in their entirety. They contain information about returning the removed rings. Enclosed is Form 30.4 Product Tracking and Verification Form. You must complete and return this form even if you do not have any rings to remove and return. Please contact your Captiva Spine customer service representative if you have any questions regarding this correction, any of our products, or would like assistance with the correction. |  | 233 | 233 devices | GA, IN, FL, UT, AZ, CA, and NY | Terminated | June 01, 2015 | 96 | 233 | Computer
Z-0170-2016 | 72374 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARIA Radiation Oncology, versions 10, 11 , 13.0 and 13.5 with Clinical Assessment License. |  |  | September 28, 2015 | 2015 | Varian Medical Systems, Inc. | An anomaly was identified with the ARIA for Radiation Oncology software with a Clinical Assessment license. Admin Instructions may not print on drug order prescriptions or be transmitted with e-Rx. | Software design (manufacturing process) | Varian sent an Urgent Medical Device Corrections letter dated September 29, 2015 via trackable method to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were informed that Varian is developing a technical correction for this issue. Customers will be contacted by a Varian Field Service. A copy of the notification is to be kept with the most current product labeling and all appropriate personnel in the radiation oncology department should be aware of this issue. For questions contact your local Varian Medical Systems Customer Support District orRegional Manager |  | 1971 | 1971 units | Worldwide Distribution - US Nationwide and the countries of: Canada, Switzerland, South Korea, Germany, Spain, Ireland, Spain, France, Belgium, United Arab Emirates, United Kingdom, Norway, Italy, Trinidad and Tobago, Netherlands, Philippines, Portugal, Denmark, Sweden, Belgium, Saudi Arabia, Thailand, Australia, India, Taiwan, South Africa, Israel, Australia, Kuwait, Finland, Tunisia, Algeria, Argentina, Bangladesh, Brazil, Chile, China, Colombia, Costa Rica, Czech Republic, Guam, Guatemala, Hong Kong, Indonesia, Japan, Kazakhstan, Kuwait, Lebanon, Luxembourg, Madagascar, Moldova, Morocco, New Zealand, Panama, Poland, Qatar, Romania, Russia, Singapore, South Africa, Tunisia and Turkey. | Terminated | May 12, 2017 | 592 | 1971 | Computer
Z-1515-2015 | 70679 | No | OOE | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | OptiMedica Catalys Precision Laser System; Model No(s). Catalys-U (US), and Catalys-I (Outside US). |  |  | February 20, 2015 | 2015 | Abbott Medical Optics, Inc. | Device may experience loss of suction during treatment resulting in scoring the cornea during lens fragmentation. | Device Design | Abbott Medical Optics (AMO) sent an Advisory Notice letter dated March 5, 2015 via FedEx overnight delivery to each Catalys Precision Laser System owner. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete and return the attached Acknowledgement Form. For questions contact AMO Technical Support at 1-855-662-0911. |  | 202 | 202 | Worldwide Distribution - US Nationwide and the countries of, Canada, Germany, Switzerland, Australia, Spain, Japan, Great Britain, France, Netherlands, Austria, Singapore, South Korea, India, Thailand, Colombia, Ecuador, Italy, Brazil, Egypt, Tunsia, Belgium, Mexico, Greece, Guatemala, Chile, Turkey, Peru, Hong Kong, Saudi Arabia, Bangladesh, and Dominican Rep. | Open, Classified |  |  | 202 | Computer
Z-0183-2016 | 72104 | No | JGS | Sodium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA Chemistry XPT Chemistry Systems, Siemens Material Number (SMN) 10723034. |  |  | August 28, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc. | ISE module may produce discrepant, unflagged results for Sodium, Potassium or Chloride when the respective electrode is left unplugged after replacement, washing or maintenance activities. The calibration and 2-level QC will fail if the electrode is left unplugged. The IFU states to calibrate the ISEs after replacement and to run two levels of controls before running patient samples. | Device Design | Urgent Medical Device Correction letters (dated 8/27/2015) and Field Correction Effectiveness Check Forms were sent to customers via FedEx starting on 8/28/2015, informing them of the potential for discrepant ISE results. Customers must follow the instructions provided in the letter. The Actions to Be Taken section of the letter notes that customers must ensure that the electrodes are fully plugged in after replacement, washing or maintenance activities, followed by performing an ISE Calibration and running 2 levels of Quality Control material. Urgent Field Safety Notices were sent to customers outside the USA. The Field Correction Effectiveness Check form should have been completed and returned. A copy of the letter should be kept with laboratory records; and, the letter should be forwarded to anyone who may have received the affected product. Customers should contact their Siemens Customer Care Center or their local Siemens technical support representative with questions. |  | 74 | Domestic: 15 units; Foreign: 59 units | Worldwide Distribution -- US, Australia, Bahrain, Denmark, Egypt, Egypt, Finland, France, Germany, Italy, Netherlands, New Zealand, Norway, Republic Korea, Singapore, Slovakia, Spain, and the United Kingdom. | Open, Classified |  |  | 74 | Computer
Z-0168-2016 | 71207 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ONCOR Avant Garde, ONCOR Impression Plus, ONCOR Expression, ONCOR Impression, PRIMUS HI, MEVATRON M2/PRIMUS Mid-Energy, ARTISTE; intended use is to deliver x-ray radiation for therapeutic treatment of cancer. |  |  | May 05, 2015 | 2015 | Siemens Medical Solutions USA, Inc | potential safety risk of collision of the LINAC gantry with the patient or the treatment table in case of automatically sequenced treatment delivery techniques using SIMTEC" Auto Field Sequence Option. | Software design | Siemens sent Safety Advisory Notices, dated May 5th and 22nd, 2015 to all affected customers to inform them about an update of the Control Console software to correct defects affecting the Control Console with their digital Linear Accelerator. Customers were instructed to ensure everne at their organization is aware of the information in the Advisory Notice. If product was further distributed, please forward Safety Notice to all those affected.For questions regarding this recall call 610-219-6300. |  | 158 | 158 | Nationwide Distribution | Terminated | May 05, 2017 | 731 | 158 | Computer
Z-0166-2016 | 72013 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Perkin Elmer Specimen Gate, Screening Center. Data management of neonatal screening test results and demographics by qualified laboratory personnel in newborn screening programs. |  |  | August 24, 2015 | 2015 | Perkin Elmer Life Sciences Inc | It would be possible for an abnormal analytical test result value from an affected newborn to be reported with a default result interpretation of normal. The analytical test result value would be correct but the interpretation of the test result value would be incorrect. | Software design | Perkin Elmer issued the Recall Letter on 8/24/15 delivered to the customer by email .Software Services will be asked to make sure the customers receive the Recall Letter and implement the recommended corrective measures: Ensure all laboratory routine sample testing procedures are followed to avoid procedural anomalies. In the event a procedure anomaly occurs, review the final analytical result and the corresponding interpretation to confirm both are accurate before releasing the report. Once a PerkinElmer reporting error query is available, continue to run the query until the final corrective action is deployed to your site. The final corrective action will involve an installation update to the software.For further information, please contact your local PerkinElmer representative or specimen.gate.support@perkinelmer.com |  | 6 | 6 units | US distribution to FL, GA, and NV; and Canada | Open, Classified |  |  | 6 | Computer
Z-1513-2015 | 70851 | No | GEH | Cryosurgical unit and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | CSA Medical truFreeze Console- Cryosurgical Unit Cryogenic Surgical Device |  |  | March 25, 2015 | 2015 | CSA Medical | TruFreeze Console caused a higher rate of liquid nitrogen (cryogen) to be delivered and may cause: stricture,scarring, bradycardia, or pneumothorax | Software design | CSA Medical issued letter dated 3/25/15 advising users of the problem. Users provided with:The mitigations available to the active venting procedures coupled with the extremely unlikely probability of injury, rare risk, thus allowing the physician to continue with active venting procedures.In regards a passive venting procedure may not have sufficient mitigation to allow the physician to identify the potential hazard with sufficient time to preclude a potential for injury. Therefore, passive users will be instructed not to use the system until a software improvement is put into place. Additionally, no catheters or consoles will be shipped to passive venting users of truFreeze.A response form to be signed and returned confirming receipt of the notification.Questions contact: Stephen Mascioli, MD 781-538-4755 smascioli@csamedical.com. |  | 82 | 82 units | Nationwide | Terminated | August 21, 2015 | 149 | 82 | Computer
Z-1517-2015 | 64225 | No | MEA | Infusion pump. | General Hospital | General Hospital | 510(k) | The PCA 3 infusion pump allows clinicians, or patients, to (self) Administer analgesia safely and effectively within clinician programmed limits. |  |  | February 27, 2013 | 2013 | Hospira Inc. | The infusion pump has a clear plastic door that, due to its design, allows for excessive flexing. This in turn permits unauthorized access to the vial of medication. | Device Design | Hospira will contract with Stericycle Inc., in Indianapolis, IN to send letters and a reply form to Hospira customers. Receipt of the letters will be confirmed by the UPS tracking. Customers will be requested to return a reply form to Stericycle indicating the number of pumps on hand and that they have received the letter and notified the users.For questions regarding this recall call 800-441-4100. |  | 8214 | 8,214 units | Worldwide Distribution - US including AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY, District of Columbia, Puerto Rico, American Virgin Islands; and Internationally to Canada. | Open, Classified |  |  | 47096 | Computer
Z-1519-2015 | 70863 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Light Sheer Desire Diode Laser System with XC Handpiece Accessory options. |  |  | March 25, 2015 | 2015 | Lumenis Limited | Device software treatment preset parameters for the XC treatment handpiece do not match the Operator Manual, and exceed recommended settings. Operator Manual parameters are lower than indicated for specific hair color and Fitzpatrick skin type resulting in insufficient treatment effect. May result in patient burns and hypopigmentation. | Software design | Lumenis Limited sent an Urgent Safety Advisory Notice dated March 25, 2015, to all affected customers. Customers were instructed that Lumenis is eliminating the potential risk of patient harm through a software upgrade and Operator Manual replacement. The software upgrade must be completed by a Lumenis-certified engineer as soon as possible. In the Interim, customers can continue to use their LightSheer DESIRE system by following the recommendations in the operator manual to always perform a test patch prior to a full treatment. Customers with questions were instructed to contact the Aesthetic Business Unit Product Manager.For questions regarding this recall call 801-656-2690. |  | 117 | 117 units | Worldwide distribution. US nationwide (to OR, NM, NY, IL, and CO), Switzerland, Canada, Italy, Great Britain, United Arab Emirates, France, Belgium, India, Israel, Japan, Portugal, Czech Republic, Kuwait, Germany, Spain, Georgia, Armenia, Slovenia, New Zealand, Saudi Arabia, and Russian Federation. | Terminated | June 15, 2016 | 448 | 117 | Computer
Z-0118-2016 | 72138 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | CIOS ALPHA; image intensified fluoroscopic x-ray system |  |  | August 17, 2015 | 2015 | Siemens Medical Solutions USA, Inc | patient procedure interruption due to a potential system failure | Device Design | Customer Safety Advisory Notice XP038/15/S, dated August 17, 2015, was sent to direct accounts to alert them to the issue and provide work-around instructions to avoid potential risks. |  | 44 | 44 | Nationwide Distribution. | Open, Classified |  |  | 44 | Computer
Z-1521-2015 | 70685 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Software Version 2.8 &amp; Below on VITROS 5,1 FS Chemistry Systems, Catalog Number 6801375, Global Trade Item Number 10758750001132, and VITROS 5,1 FS Chemistry System Refurbished, Catalog Number 6801890, Global Trade Item Number 10758750001644; IVD. |  |  | March 10, 2015 | 2015 | Ortho-Clinical Diagnostics | Software Anomaly; It is possible for the device to process samples with a cartridge other than the intended cartridge, potentially leading to erroneous patient results. If this anomaly occurred previously, an indication would be a series of results that were similar across multiple samples (i.e., results could be believable or outside of the Reportable Range for the intended assay). | Software in the Use Environment | On 3/10/2015, URGENT PRODUCT CORRECTION NOTIFICATION letters (Ref. CL2015-047, dated 3/10/2015) were sent to users notifying them of the issue and advising them of the required actions. Customers are to complete and return the Confirmation of Receipt no later than March 23, 2015. The firm is currently working on a resolution and will issue a follow up notification upon the software's availability. For questions, please contact Customer Technical Services at 1-800-421-3311. |  | 274.1 | VITROS 5.1 System: Domestic - 909; Foreign - 1250; 5,1 Refurbished: Domestic - 66 units, Foreign - 152 | Worldwide Distribution -- USA, Puerto Rico, Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Republic of Panama, Singapore, Spain, and Venezuela. | Open, Classified |  |  | 274.1 | Computer
Z-1522-2015 | 70874 | No | JWY | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | MicroScan -Neg Breakpoint Combo Panel Type 34; |  |  | May 14, 2014 | 2014 | Beckman Coulter, Inc. | Siemens Healthcare Diagnostics has received complaints regarding panel degradation with a portion of the MicroScan Neg Breakpoint Combo 34 (NBP34) panel lot 2014-05-28. | Equipment maintenance | Siemens sent an Urgent Medical Device Recall letter dated May 14, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The risk to health is explained and users are asked to discontinue use and discard any of the recalled lot still in inventory. No other lot numbers of the NBPC 34 panels are affected and may still be used. The Field Correction Effectiveness Check form should have been completed and returned to Siemens.PLEASE FAX THIS COMPLETED FORM TO 302-631-8467. Contact Siemens Customer Care Center or local Siemens Tech support. For further questions please call (916) 374-3031 |  | 681 | 681 boxes | Worldwide Distribution - Nationwide Distribution including Puerto Rico and to the states of: KS, LA, MD, NY, WV,NM, NH, MO, CT, MA, OK, NC. IA, OH. WI, TX, PA, MS and MN and to the countries of : Canada, Kuwait, South Africa and Turkey. | Terminated | April 27, 2015 | 348 | 681 | Computer
Z-1525-2015 | 71021 | No | DRF | Electrode recording catheter or electrode recording probe. | Cardiovascular | Cardiovascular | 510(k) | WEBSTER HIS Catheter, 4 Pole Fixed Curve with Auto ID Technology, for mapping of cardiac structure; i.e., stimulation and recording only. |  |  | April 17, 2015 | 2015 | Biosense Webster, Inc. | The Webster HIS catheter, 4Pole Fixed Curve with Auto ID, is not being recognized by the CARTO 3 navigation System. | Software Manufacturing/Software Deployment | A recall letter dated 4/17/15 will be sent to customers by FedEx to inform them that Biosense Webster is recalling Webster HIS catheter, 4Pole Fixed Curve withAuto ID (Catalog No. D1085413). The letter informs the customers of the problems identified and the actions to be taken. Customers with questions related to the recall letter are instructed to contact their Biosense Webster salesrepresentative or call (866) 473-7823, Monday through Friday from 7:00AM to 8:00PM EST. Customers with questions related to the Voluntary Recall Certification Form and the process for returning the form, are instructed contact Elie Ghattas, Field Action Coordinator, at (909) 839-7281. |  | 503 | 503 units | US nationwide distribution. | Terminated | December 10, 2015 | 237 | 503 | Computer
Z-0154-2016 | 72273 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | TELETOM EQUIPMENT MANAGEMENT SYSTEM 4 and 6 series |  |  | September 23, 2015 | 2015 | Berchtold Corp. | During the use of the TELETOM Series 4 &amp; 6 there is potential for the controls and brakes not to work properly due to fluid intrusion within the control switches and/or the control board assembly. | Device Design | Stryker sent an Urgent Medical Device Correction letter to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. A Stryker Representative would visit each account for product inspection. During the inspection the Stryker Representative will remove the control switches and replace them with new control switches. Customers were instructed if any malfunction prior to inspection were noticed customers should call local Technical Support immediately at 800-243-5135. |  | 2129 | 2129 units | Worldwide Distribution - US (nationwide), Canada, and Germany | Open, Classified |  |  | 2129 | Computer
Z-1542-2015 | 71074 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Ultraview SL Command Modules, Model 91496-M, upgraded with the Masimo SpO2 Option. |  |  | April 17, 2015 | 2015 | Spacelabs Healthcare Inc | Ultraview SL Command Modules which were upgraded with the Masimo SpO2 Option, Model 91496-M was affected. Customer reported that during the time that the module is resetting (~10 seconds), monitoring of all parameters will be suspended. Upon completion of this sequence, alarms will be reset to their default values. | Software design | Spacelabs Healthcare sent an Urgent - Medical Device Correction letter, dated 17 April 2015, to the US consignees via priority mail. On 4/24/2015, the firm will send the Urgent - Medical Device Correction letter (translated as necessary) via email to international subsidiaries and distributor of record. The Identified the affected product, problem and actions to be taken. Ultraview SL Command Modules which were upgraded with the Masimo SpO2 Option, Model 91496-M are affected. Ultraview SL Command Modules originally shipped with Masimo SpO2 are NOT affected. Spacelabs Healthcare will contact you to schedule a convenient time for Spacelabs to correct all of your facility's affected Command Modules at no cost. For additional information or technical assistance, please contact:Technical SupportSpacelabs Healthcare, Inc.35301 SE Center St.Snoqualmie, WA 980651-800-522-7025 and select 2 for Technical Support |  | 1999 | 1999 units total (1932 in the US and 67 international) | Worldwide Distribution- US Nationwide in the states of AL, DE, GA, ID, IL, KY, LA, MA, ME, MI, MN, MO, NC, NJ, NY, OH, OR, PA, SC, TN, VA, WA and the country of France. | Terminated | November 13, 2015 | 210 | 1999 | Computer
Z-1550-2015 | 70574 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | 2008 Series Hemodialysis Machines: 2008T with the following product code and description: 190395 2008K at Home System - 1st Gen. -Âż Canada; |  |  | February 20, 2015 | 2015 | Fresenius Medical Care Holdings, Inc. | The Acetate value for GranuFlo on the Select Concentrate screen should be 8mEq/l but it is displayed as 4mEq/l. The correct value is shown on the dialysate screen. | Software design | The firm, Fresenius Medical Care, sent an "URGENT - MEDICAL DEVICE CORRECTION" letter dated February 19, 2015 to its Consignees via certified mail on February 26, 2015. The letter describes the product, problem and actions to be taken. The consignees were instructed to sign and return the attached "Reply Form" via fax to: 781-699-9769 Attn: Quality Department or Scan and E-mail to: NOTIFYRA@fmc-na.com to confirm that they received and understand the customer notification. The firm will be providing a software update and stated that in the near future, your local clinic will be contacting you about installing the software update.If you have any questions pertaining to this Customer Notification, please contact your local clinic, 800-227-2572 or the Corrections and Removals Manager at 800-662-1237. |  | 133751 | 133,751 in total (US: 130,337 and OUS: Canada 1,201 and Mexico 2,213) | Worldwide Distribution: US (Nationwide) and countries of: Canada and Mexico. | Open, Classified |  |  | 133751 | Computer
Z-0135-2016 | 68228 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | All Philips Ingenia, Intera, Achieva and Multiva MR systems using R5.1i and R5.1 .2 version of software. Indicated for use as a diagnostic device. It can produce cross-sectional images, spectroscopic images and/or spectra in any orientation of the internal structure of the head, body, or extremities. |  |  | May 05, 2014 | 2014 | Philips Medical Systems, Inc. | In spine clinical workflows, cross reference lines may be used to determine the position of slices. In cases, where MobiView fused Images are used to show the cross reference lines, the cross reference lines may be positioned incorrectly. | Software design | Consignees were notified of the recall by letter sent May 2014. They were instructed to do the following: Do not perform planning or review planning of transversal images on fused sagittal images which were generated through Mobi View postprocessing Instead, review the planning of transversal scans only on the unfused stations. A Field Change Order with reference FC0781 00431 is being released that PHILIPS requires Philips field service engineers to disable the MobiView post processing software. When possible this will be done remotely. Once it becomes available, a software update will be installed which will address the reference line positioning issue with MobiView Should you need to communicate with Philips with regard to this program, please reference FC0781 00431. If you need any further information or support concerning this issue, please contact your local Philips representative or the Technical Support Line: 1-800-722-9377. |  | 317 | 317 | Worldwide Distribution-US: AL AR AZ CA CO FL GA IL IN KS KY MA MI NY OH OR PA SC TX VA VT WA OUS: Argentina Australia Austria Bangladesh Belgium Brazil Canada Chile China Colombia Denmark Ecuador Estonia Ethiopia France Germany India Indonesia Iraq Ireland Italy Japan Jordan Korea, South Malaysia Mexico Netherlands New Zealand Norway Oman Palestine Poland Russian Federation South Africa Spain Sweden Switzerland Taiwan Turkey United Arab Emirates United Kingdom. | Open, Classified |  |  | 317 | Computer
Z-0131-2016 | 72208 | No | PHX | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | AltiVate Humeral Socket Shell Trial, Part Number 804-06-052 |  |  | September 15, 2015 | 2015 | Encore Medical, Lp | Difficulty assembling the shell trial to the broach and the retaining ring disassembled from the screw. | Device Design | The firm, DJO Global, sent an "URGENT FIELD SAFETY NOTICE" letter dated 9/16/15 to affected consignees/customers. The letter described the product, problem, and actions to be taken. The consignees/customers were instructed quarantine the device, pass the notice on to all who need to be aware within your organization, contact customer service to place a replacement order and receive an RMA number; and to complete a response form and return the recalled device for a replacement.If you have any questions, please call (512) 834-6302 or by email at desiree.hubby@djoglobal.com. |  | 20 | 20 units | US Distribution to states of: PA, FL, ID, and CA. | Terminated | January 10, 2017 | 483 | 20 | Computer
Z-0130-2016 | 72223 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System. |  |  | August 31, 2015 | 2015 | Elekta, Inc. | Incorrect Treatment Delivery Using Third Party Fixed Wedges. | Software design | Consignees were notified via URGENT FIELD SAFETY ALERT, on 08/31/2015. |  | 1440 | 1440 units | Worldwide Distribution-Domestic: AK, AL, AZ, CA, CO, CT, DC, FL, IN, IL, KY, MA, MD, MI, MN, MO, NJ, NC, ND, NY, OH, OK, PA, PR, TN, TX, VA, WA and WI; International: Albania, Argentina, Australia, Austria, Bahamas, Bahrain, Bangladesh, Belarus, Belgium, Bosnia and Herzegovina, Botswana, Brazil, Canada, Chile, China, Colombia, Croatia, Cyprus, Czech Republic, Ecuador, Egypt, Estonia, France, Gabon, Germany, Greece, India, Indonesia, Iraq, Ireland, Israel, Italy, Japan, Jordan, Libya, Macedonia, Malaysia, Mali, Malta, Mexico, Morocco, Myanmar, Netherlands, New Zealand, Pakistan, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, South Korea, Spain, Suriname, Switzerland, Taiwan, Thailand, Turkey, The United Kingdom, Ukraine, Venezuela and Vietnam. | Open, Classified |  |  | 1440 | Computer
Z-1561-2015 | 70938 | No | FLE | Steam sterilizer. | General Hospital | General Hospital | 510(k) | AMSCO C and AMSCO 400 Steam Sterilizers, STERIS Corporation. |  |  | February 19, 2015 | 2015 | Steris Corporation | STERIS has identified that the control board software in select AMSCO 400 and AMSCO C units will interrupt and cancel a processing cycle should the selected sterilizer shutdown time coincide with a processing cycle. This may result in a procedure delay. | Software design (manufacturing process) | On 3/9/2015, the firm sent Urgent Field Correction Notices, to their customers. The letter identified the affected product and gave a description of the problem. The letter states that the control software is being revised to ensure the utility shutdown feature will not interrupt an ongoing processing cycle. STERIS Field Representatives will contact customers to schedule installation of the new control board on affected units. Customers may direct questions regarding the issue to STERIS Customer Service at 1-800-548-4873; Marie LaFrance, Product Manager, at 440-392-7648, or their local sales representative. |  | 701 | 701 units | Worldwide Distribution -- US, including the states of AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, I, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, &amp; WY; and, the countries of Afghanistan, Canada, Japan, Kuwait, Mexico, Republic of Korea, &amp; United Arab Emirates. | Terminated | January 20, 2016 | 335 | 701 | Computer
Z-1604-2015 | 71099 | No | DSQ |  |  | Cardiovascular | PMA | HeartWare Ventricular Assist System (HVAD) Pump Driveline Splice Kit |  |  | April 29, 2015 | 2015 | HeartWare Inc | Failures of the splice repair kit if exposed to excessive force. | Device Design | HeartWare, Inc. sent an Urgent Medical Device Correction letter to all affected consignees on May 11, 2015. The letter identified the affected product, problem, and actions to be taken by the consignees. The firm requested that consignees: 1) review the notice and the Patient Communication contents, 2) forward the notice to individuals within their organization who need to be aware of the notice, 3) identify patients that are currently on support and have undergone a driveline splice repair, 4) distribute the Patient Communication to the affected patients in person, 5) continue to reinforce the messages described in the notice with patients who have experienced a splice repair during their regularly scheduled appointments, 6) complete, sign, and return the "Acknowledge and Completion Form to HeartWare within 30 days of receipt of the letter. Should any questions or concerns arise, please contact there local HeartWare representative. Their 24-hour Clinical Support personnel are also available at 1 (888) 494-6365. |  | 35 | 35 | Worldwide Distribution - US Nationwide in the states of FL, KY, MO, NY, TN,TX, and the countries of: Canada, Germany, Italy, Netherlands, Turkey and United Kingdom. | Open, Classified |  |  | 35 | Computer
Z-1612-2015 | 70890 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Fujifilm Synapse Cardiovascular I ProSolv CardioVascular |  |  | February 04, 2015 | 2015 | Fujifilm Medical Systems U.S.A., Inc. | Under certain circumstances the patientDemographics in a report exported into the EMR may not match the demographics shown in the corresponding Synapse CV clinical report. | Software in the Use Environment | Fujifilm Medical on 2/4/14 , contacted users and advised to discontinue using HL7 reports pending FMSU determination of root cause. Firm issued follow-up letters on 2/24/14 and 7/24/14 advising users of a mandatory upgrade to version 5.1.1 I base, in-place upgrade, to correct the problems. |  | 16 | 16 units | Nationwide Distribution including PR and the states of IL, IN, FL, GA, MD, MI, MO, MS, NC, NY, OR, and PA, PR and the country of Canada. | Open, Classified |  |  | 16 | Computer
Z-2826-2015 | 71983 | No | FPO | Wheeled stretcher. | General Hospital | General Hospital | 510(K) Exempt | Ferno iNX Wheeled Stretcher, Integrated Patient Transport &amp; Loading System |  |  | August 03, 2015 | 2015 | Ferno-Washington Inc | Hole elongation in the telescoping load frame channel. A hole elongation could affect the ability of the retention pin to stay seated which is necessary to preclude movement of the telescoping load frame. | Device Design | Ferno began calling their customers on 8/3/2015. The firm informed their customers of the following: To provide optimal Customer Experience and minimal disruption to your service, the iNX field action involves the following steps:" Exchange of your current iNX with a new iNX;" Participation in a brief training session; and" Inspection of your Ferno Model 175 antler and rail fastening system as requested by the Dept. of Transportation.As you await delivery of your new iNX, it is important you ensure proper function of the fastening system and iNX telescoping load frame. |  | 72 | 72 units | Worldwide Distribution - US Nationwide in the states of: CO, IA, IL, IN, KY, MD, MN, MO, NC, NJ, NY, OH, OK, OR, TX, VA, WA &amp; WY and the countries of CANADA, AUSTRALIA, UNITED KINGDOM &amp; NEW ZEALAND. | Terminated | January 20, 2016 | 170 | 72 | Computer
Z-1691-2015 | 70772 | Yes | OAE |  |  | Cardiovascular | PMA | SMARTABLATE RF Generator System, Catalog No. M490006, M490007, M4900106, and M4900107. |  |  | March 19, 2015 | 2015 | Biosense Webster, Inc. | Biosense Webster is recalling the SMARTABLATE Generator System because Biosense Webster has received two complaints reporting that during an RMT procedure the SMARTABLATE RF Generator inadvertently delivered RF energy. | Device Design | Biosense Webster sent an Urgent Field Safety Notification Letter dated March 18, 2015 via FedEx to all their customers that received the affected SMARTABLATE Generator System. The letter informs the customers of the problems identified and the actions to be taken. Biosense Webster is taking precautionary safety measures of reinforcing the Warnings in the Instructions for Use (IFU) for the SMARTABLATE System Foot Pedal. Customers with questions related to the are instructed to contact their local Biosense Webster sales representative or call (866) 473-7823, Monday through Friday from 7:00AM to 8:00PM EST. Customers with questions related to the voluntary field removal certification form and its return, are instructed to contact Elie Ghattas Field Action Coordinator at (909) 839-7281. |  | 115 | 115 units | Worldwide Distribution - US Nationwide in the state of: CA, IL, KY, NJ, PA and the countries of: Czech Republic, Denmark, Finland, France, Germany, and Sweden. | Terminated | June 17, 2015 | 90 | 115 | Computer
Z-1682-2015 | 71228 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Ingenuity TF Positron Emission Tomography (PET) and Magnetic Resonance Imaging (MRI), 882380, Philips Medical Systems. |  |  | March 11, 2015 | 2015 | Philips Medical Systems (Cleveland) Inc | It was noticed that the spectroscopy voxel map did not align with the anatomical images of the phantom. This resulted in incorrect orientation of the Chemical Shift Imaging (CSI) data. | Software in the Use Environment | On 3/11/2015 the firm sent Urgent - Medical Device Correction letters to their customers. The letters identified the affected device, explained the reason for the recall, how to identify affected product, advice on actions to be taken by customer/user, and any actions planned by the firm. Customers were advised to use software provided by Philips to view spectroscopy images; and, not to export and view images using third party software. The firm stated that it will be issuing a software update. If further information or support concerning the issue is needed, customers are asked to contact their local Philips representative or local Philips Healthcare office. |  | 10 | 10 units | Worldwide Distribution -- US, including the states of NY, OH, &amp; TX; and, the countries of Finland, Germany, Japan, Netherlands, South Korea, Spain, &amp; Switzerland. | Open, Classified |  |  | 10 | Computer
Z-1686-2015 | 71296 | No | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | MEDRADÂż Stellant CT Injector System with Certegra |  |  | September 20, 2013 | 2013 | Medrad Inc dba Bayer R&I; | The root cause investigation showed that the under-volume hazard can occur when the unit has been configured to prime with contrast, but instead primes with saline due to the configuration data not being sent to the injector head. This may create a non-optimal priming condition that could lead to under-infusion. | Software design | Bayer Healthcare sent a Medical Device Field Corrective Action letter to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to identify any affected device. Bayer Service will schedule time to update their software free of charge. Customers with questions were instructed to contact the customer support team at 1-800-633-7237. For questions regarding this recall call 412-767-2400. |  | 133 | 133 | Nationwide Distribution | Terminated | June 16, 2015 | 634 | 133 | Computer
Z-2785-2015 | 71477 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Radiation Therapy Treatment Planning System software |  |  | May 29, 2015 | 2015 | RAYSEARCH LABORATORIES AB | An error caused by a floating point precision problem tends to cause a problem with the display of isodose lines for dose and dose difference and color tablesfor dose, dose difference and PET images | Software design | The firm, RaySearch Laboratories, distributed a "Field Safety Notice, Medical Device Correction #7106" notice dated May 29, 2015,and reply forms to their customers via email. The notice describes the product, problem and actions to be taken. The Customers were provided a workaround as a temporary fix. Customers were also instructed to educate all users, inspect your product and identify all installed units, confirm you have read and understood the notice by contacting Quality and Regulatory Affairs Manager; and complete and return the reply form via email to freddie.cardel@raysearchlabs.com or by fax to 888-501-7195. The solution will be resolved in the next version of RayStation, scheduled for release in June 2015. Customers were advised to share the notice to all those who need to be aware within the organizations. Customers can contact Quality and Regulatory Affairs Manager at +46 8 510 530 12 or via email at david.hedfors@raysearchlabs.com. |  | 20 | 20 devices | Worldwide Distribution: US (nationwide) to state of WA; and countries of: France, Sweden, Austria, Spain, Germany, Italy, UK, New Zealand. | Terminated | March 13, 2017 | 654 | 20 | Computer
Z-2778-2015 | 71695 | No | FDT | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fujifilm Medical Endoscope for the duodenum and upper G.I. tract. ED-530XT, ED-250XT5, ED-250XL5, ED-450XT5, and ED-450XL5 Endoscope Operation Manuals. |  |  | June 30, 2015 | 2015 | Fujifilm Medical Systems U.S.A., Inc. | The suction button on the ED-530XT endoscope may become lodged (stuck) in the active position following the use of x-ray contrast agents during ERCP procedures. | Device Design | FujiFilm sent a Voluntary Field Correction/Verification Form dated June 30, 2015 to their affected customers. |  | 256 | 256 units | Distributed in the US and Canada. | Terminated | May 25, 2017 | 695 | 256 | Computer
Z-2783-2015 | 72094 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON SC2000 Ultrasound System with software versions VB10B and lower; |  |  | August 19, 2015 | 2015 | Siemens Medical Solutions USA, Inc. | The ACUSON SC2000 ultrasound system considers uppercase/lowercase differences in the same patient name as unique patient instances when registered on the same ultrasound system. If these differences are not corrected at the time of registration, the system does not capture images or clips. | Software design | Siemens sent a Customer Safety Advisory Notification letter on August 19, 2015 by certified mail. The letter identified the affected product, problem and actions to be taken. The letter states that this issue will be resolved in a future software release. If you have any questions,please contact your customer care center for information regarding timelines and status. Until the necessary repairs have been completed on your system, please share this information with all personnel within your organization who need to be aware of this issue. The Siemens Service organization coordinates the implementation of the field correction with the distributors for all countries. |  | 2039 | 2039 devices | Worldwide distribution. US Nationwide and the countries of Albania, Angola, Argentina, Armenia, Australia, Austria, Azerbaijan, Bahrain, Bangladesh, Belarus, Belgium, Bolivia, Bosnia Herzeg., Brazil, Brunei, Bulgaria, Canada, Colombia, Croatia, Czech Republic, Denmark, Dominican Rep., Ecuador, Egypt, El Salvador, Finland, France, Georgia, Germany, Hungary,, Iceland, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Kuwait, Lebanon, Lithuania, Malaysia, Maldives, Malta, Mauritius, Mexico, Netherlands, New Zealand, Nigeria, Norway, Oman, P.R. China, Pakistan, Philippines, Poland, Portugal, Qatar, Rep. of Yemen, Republic Korea, Romania, Russian Fed., San Marino, Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, Spain, Suriname, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Turkey, Turkmenistan, U.A.E., Ukraine, United Kingdom, Uzbekistan, Venezuela, and Vietnam. | Open, Classified |  |  | 2039 | Computer
Z-1704-2015 | 71206 | No | LMD | Medical image communications device. | Radiology | Radiology | 510(K) Exempt | syngo Workflow SLR; The information system syngo Workflow SLR is a digital radiology information system (RIS) with integrated modules for patient administration, examination, reporting, statistics and system administration. The system electronically displays stores, retrieves, transfers, exchanges, and prints. |  |  | May 06, 2015 | 2015 | Siemens Medical Solutions USA, Inc | A potential exists for order transactions from interfaced HIS (Hospital Information System) systems to be lost during a restart of interface processes when using the Order Batching feature, leading to diagnosis delay. | Software design | Siemens sent an Safety Advisory Notice, dated May 6, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were given work around instructions to follow and informed that a software fix would be implemented to permanently resolve the issue.For further questions please call (610) 219-6300. |  | 131 | 131 | Nationwide Distribution | Open, Classified |  |  | 131 | Computer
Z-1701-2015 | 71202 | No | OLO | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | Tibial Alignment Guide; Orthopedic Stereotaxic Instrument; |  |  | April 19, 2015 | 2015 | Zimmer CAS | Zimmer CAS has determined that the potential exists for the spikes on the iASSIST Tibial Alignment Guide to bend or break during insertion or extraction from the tibia during use. | Device Design | Zimmer distributed notices via certified mail on May 14, 2015. Zimmer is removing affected Tibial Alignment Guide iAssist Knee System product from distribution in a two-phased process. Based upon complaint investigation, Zimmer CAS has determined that the potential exists for the spikes on the iASSIST Tibial Alignment Guide to bend or break during insertion or extraction from the tibia during use. Customerswere asked to review the notifictaion, follow the iASSIST Surgical Knee Technique provided, further distribute the notice for any distributed product, inspect devices in use before and after procedures, report any adverse events, and complete the certificate of acknowledgment and return it to CorporateQuality.PostMarket@zimmer.com.Customers with questions concerning this notice, please contact Customer Service at the following address/phone number.Customer Service / Zimmer CAS75 Queen Street, Suite 3300Montreal, Quebec, Canada H3C 2N6Email : cas-support@zimmercas.com, Telephone : 1-514-395-8883, toll free for North America 1-866-336-7846, Fax : 1-866-978-3801.For questions regarding this recall 574-372-4487. |  | 78 | 78 devices | Nationwide Distribution including AL, AZ, CA, CA, CO, FL, IN, MI, NJ, NY, OH, OR, PA, TX, VA, and WI. | Open, Classified |  |  | 78 | Computer
Z-1705-2015 | 71314 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo.via and Syngo.x ; a software solution intended to be used for viewing, manipulation, communication, and storage of medical images. |  |  | May 15, 2015 | 2015 | Siemens Medical Solutions USA, Inc | measurements drawn on the 2nd and subsequent images of the series are not visible on printouts when the series is sent to print. | Software design | Siemens sent a Customer Advisory Notice, dated May 15, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to check the software version of their syngo.via system and consider whether the issue described in the advisory notice impact their clincal coutine and whether to arrange the installation of SY027/15/P (VA30A_HF07). Customers were advised to include the advisory notice in their operator's manual. Customers were also advised to notify anyone at their organization who should be aware of this information.For questions regarding this recall call 610-219-6300. |  | 167 | 167 | Nationwide Distribution | Terminated | March 24, 2017 | 679 | 167 | Computer
Z-2776-2015 | 71971 | No | DXO | Catheter tip pressure transducer. | Cardiovascular | Cardiovascular | 510(k) | Edwards Pressure Monitoring Kit with TruWave disposable pressure transducer. A sterile, single-use kit that monitors pressures when attached to pressure monitoring catheters. For use on patients requiring intravascular, intracranial, or intrauterine pressure monitoring. |  |  | July 31, 2015 | 2015 | Edward Lifesciences, Llc | Flow restriction issues in a trifurcated IV set included in the pressure transducer kits. A restriction of the flow of IV fluid would result in either an inability to prime the system before use or difficulty in flushing the line after a blood draw. | Device Design | Edwards sent an Urgent Product Recall Letter on August 10, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The firm's letter requests that customers return any of the units that they currently have in inventory. The letter requests that customers verify the inventory that has been pre-populated on the acknowledgment form, and complete the remainder of the form. The firm states to fax back the acknowledgment form within three days of receipt of notification to Edwards Customer Service at 1-800-422-9329. The firm requests to call Edwards Customer Service at 1-800-424-3278 to arrange for return of any unused product from the implicated lots as well as to obtain information about product replacment. The notification states to call for an RGA prior to returning any unused product and return unused product to the firm at Edwards Lifesciences Return Goods. Any questions the firm states to contact Edwards Customer Service at 1-800-424-3278. |  | 13077 | 13,077 (U.S.) and 12,306 (International) | Worldwide Distribution - US nationwide, Australia, Austria, Canada, China, Colombia, Croatia, Denmark, France, Germany, Great Britain, Italy, Netherlands, New Zealand, Sweden, and Slovenia. | Open, Classified |  |  | 13077 | Computer
Z-2770-2015 | 71851 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Lumenis VersaCut+ Tissue Morcellator GA-0007500 |  |  | July 30, 2015 | 2015 | Lumenis | There is a probability the hand piece may operate inconsistently and unpredictably during the morcellation procedure resulting in damage to non-target tissue. | Device Design | Lumenis sent an Urgent Medical Device Recall letter on August 7, 2015, to all consignees. Consignees were instructed to destroy the operator manual, and return the affected device to Lumenis along with the completed Recall Verification form. Consignees with questions were instructed to call 801-656-2663. |  | 162 | 54 devices with 108 hand pieces | Worldwide Distribution - US including CA, PA, MA, FL, and NY, and Internationally to Italy, Germany, Australia, China, Great Britian, United Kingdom, Taiwan, France, Peru, Mexico, Turkey, Spain, and Russia. | Open, Classified |  |  | 162 | Computer
Z-2771-2015 | 71657 | No | FWB | Surgical camera and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Sony LMD-1951MD, Sony LMD1951MD/BS, Sony LMD-1951MC (Foreign) 19 inch medical grade liquid crystal display (LCD) Medical Monitors with light-emitting diode (LED) backlight technology. |  |  | January 13, 2015 | 2015 | Sony Medical Systems | Tthe medical monitor has either not turned on or has lost image during clinical use. | Device Design | Sony Corporation sent an URGENT: MEDICAL DEVICE FIELD ACTION field letter dated January 13, 2015. Customers are asked to affix the warning sticker to the rear of the monitor and place the adhesive seal over the AC inlet. If the firm does not possess an AC adaptor (AC-110MD) for each of the affected monitors, stop using the monitor and take the monitor out of service until the permanent corrective action described below is implemented.SEL will be replacing the power board of all affected units with a redesigned power board that does not exhibit the appropriate failure mode. Questions can be directed to 866-789-SONY (7669) between 8:00 am and 8:00 pm Monday through Friday. |  | 4952 | 4952 units (1016 US; 3936 OUS) | Worldwide Distribution - US Distribution to the states of : AZ, CA, NC, PA, IA, OH and FL., and to the countries of : China, Japan, Taiwan, South Korea and Indonesia. | Open, Classified |  |  | 4952 | Computer
Z-1719-2015 | 71295 | No | IWB | Radionuclide radiation therapy system. | Radiology | Radiology | 510(k) | Laksell GammaPlan, a computer based dose planning system specifically designed for use with Leksell Gamma Knife, radiation therapy treatment. |  |  | June 01, 2015 | 2015 | Elekta, Inc. | Memory can become corrupted when creating a fused study via drag and drop in Leksell GammaPlan 10.2. | Software in the Use Environment | The firm, Elekta, sent an "URGENT IMPORTANT FIELD SAFETY NOTICE 100-01-102-015" to all customers on 6/1/2015. The notice described the product, problem and actions to be taken. The customers were instructed to not use drag-and-drop to fuse images as a solution until a permanent fix can be developed and implemented on all affected devices. A solution is in development that will resolve the issue and is expected to be released September 2015. Elekta Service will then work with customers to implement on affected devices - estimated 6 month completion. Customers were also instructed to complete and return the IMPORTANT FIELD SAFETY NOTICE ACKNOWLEDGEMENT form to your local Elekta Office or Representative as soon as possible and within 30 days at the latest. If you have any queries about this Notice, please contact your local Elekta representative or the Global Post market Surveillance Manager at 770-300-9725 or email: Elekta_Global_PMS@elekta.com. |  | 10 | 10 | Worldwide Distribution: US (nationwide) including states: IL, OH, PA, TX, and WI; and countries of: Germany, Japan and United Kingdom. | Terminated | January 12, 2016 | 225 | 10 | Computer
Z-2775-2015 | 71994 | No | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | Angiographic Injector. Medrad Mark 7 Arterion Injection System, which utilizes Software Version SW 005.006_SH; to be used specifically for the purposes of injecting contrast medium and common flushing solutions into humans for angiographic studies. |  |  | August 04, 2015 | 2015 | Bayer Healthcare | Software Version SW 005.006_SH, resulted in the removal of purge enforcement from traditional New-Case, Power Up and Syringe Change use cases while the injector head is in the upright position. | Software design | Direct accounts were notified by UPS Notification letter dated August 4, 2015 and advised when using the Arterion Injection System to ensure that purges are completed to avoid the injection of air. |  | 71 | 71 | Worldwide Distribution. US Nationwide, Germany, France, UK, Netherlands, Sweden, and Canada. | Open, Classified |  |  | 71 | Computer
Z-1698-2015 | 71164 | No | DSQ |  |  | Cardiovascular | PMA | HeartWare Ventricular Assist System (HVAD). |  |  | April 29, 2015 | 2015 | HeartWare Inc | HeartWare has received complaints relating to damage or bent connection pins within the power supply ports of the VAD controller. | Device Design | HeartWare sent an Urgent Medical Device Correction letter to customers on May 25, 2015. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to:1. Review the enclosed notice and Patient Communication .2. Forward the notice to those individuals within their organization who need to be aware of its contents.3. Identify the patients currently supported by the HVAD System.4. Distribute the "Patient Communication" to their patients directly via FedEx or some other reliable means of communication.5. Continue to reinforce the messages set forth in the notice with their patients during their regularly scheduled appointments.6. Complete, sign and return the Acknowledgement and Completion Form to HeartWare within 30 days of receipt of this letter.For questions contact your local HeartWare representative or the 24-Hour Clinical Support personnel at 1-888-494-6365. |  | 3747 | 3,747 currently implanted worldwide | Worldwide Distribution - US Nationwide in the states of AL, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, NC, NH, NJ, NY, OH, OR, PA, SC, TN, TX, UT, VA and countries of Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, Croatia, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, Hong Kong, Hungary, India, Israel, Italy, Japan, Korea, Lebanon, Lithuania, Luxembourg, Malaysia, Netherlands, New Zealand, Norway, Poland, Romania, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Turkey, and UK. | Open, Classified |  |  | 3747 | Computer
Z-1728-2015 | 70918 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS 5,1 FS Chemistry System, Software Version 2.8 &amp; Below |  |  | April 06, 2015 | 2015 | Ortho-Clinical Diagnostics | Software anomaly regarding urine samples that require acidified pretreatment. The software has allowed urine samples for assays that require acidified pretreatment to be metered from the same sample container as assays that required no pretreatment when the sample was programmed. | Software design | Ortho Clinical Diagnostics distributed correction notification letters dated April 6, 2015, and response forms to their customers via courier service. All urine patient results that require pretreatment was advised to refer to the notice of correction for instructions. Correction notices should also be posted beside the device user documentation for easy reference. Customers were advised to complete and return the Confirmation of Receipt before April 20, 2015. Customers with questions may contact the Customer Technical Service at 1- 800-421-3311 anytime. |  | 2387 | 2,387 units Total (987 domestically &amp; 1400 internationally) | Worldwide Distribution - US Nationwide, Brazil, Chile, Mexico, Australia, Hong Kong, India, Taiwan, Belgium, Denmark, France, Ireland, Italy, Portugal, Romania, Spain, Sweden and the UK. | Open, Classified |  |  | 2387 | Computer
Z-2769-2015 | 71956 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Voyant Open Fusion Device |  |  | July 31, 2015 | 2015 | Applied Medical Resources Corp | Potential to improperly characterize tissue within the jaws, which may result in nonoptimal power delivery. | Device Design | The firm sent out customer notification letters to International customers the beginning of July 31st, 2015 and domestic customers the week of August 3rd, 2015. The firm's letter requested customers complete the following actions:- check inventory for recalled product- complete the attached Recall Notification Confirmation Form to acknowledge the recall and indicate if facility is returning or has already used any of the products- Return the confirmation form(s) to Applied Medical by emailing torecall60669166@appliedmedical.com or faxing to (949) 713-8936.- Return affected product and a copy of the confirmation form(s) to Applied MedicalFor product return questions contact Karen Mitchell, Sales Operations Supervisor, RGA Dept. at (949) 713-8622 or by e-mail at kmitchell@appliedmedical.com.For regulatory questions contact Lauren Contursi, at (949) 713-8767 or bye-mail at lcontursi@appliedmedical.com. |  | 101 | 101 units (U.S) and 38 units (International) | Worldwide Distribution - U.S Nationwide in the states of:FL, GA, NV. and countries of: France, Netherlands, Spain, and United Kingdom. | Terminated | November 20, 2015 | 112 | 101 | Computer
Z-1723-2015 | 71344 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Vilex, Inc., FUZE 10.0 mm x 250 mm, Product # FZ100-250T-25. |  |  | April 22, 2015 | 2015 | Vilex In Tennessee Inc | Incidence of breakage is higher than expected. | Device Design | The initial telephone call (4/9/2015) informed the distributor that the product was on hold (not to be used or transferred). The second telephone call (4/29/2015) and letter informed the customer that the product was being removed from market and that the customer was to return the product to the corporate office.The letter was disseminated on 05/05/2015 to distributors via certified mail. The letter contains the instructions for them to follow, which as summarized are: inventory all products on hand, sign off on inventory sheet, send products back to the corporate office with the identified return number on the shipping container, inform the corporate office that the process has been completed and provide the corporate office with the list of the products returned.A second letter was disseminated to sales representatives (distributors), surgeons, and hospitals on 05/28/2015. |  | 4 | 4 units | Distributed in the states of MS, MD, PA, and TX. | Open, Classified |  |  | 24 | Computer
Z-1727-2015 | 71308 | No | LKF | Obstetric-gynecologic specialized manual instrument. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Ally Uterine Positioning System (Ally UPS) used in the mounting and positioning and holding of uterine manipulators |  |  | May 18, 2015 | 2015 | CooperSurgical, Inc. | Design of device may expose user to injury to fingers or body parts | Device Design | CooperSurgical initiated the recall on May 18, 2015 via Fedex with confirmed delivery receipt. The letter identified the affected product and asked users to discontinue use and complete the Acknowledgement and Receipt Form to schedule a replacement.Questions contact 203.601.5200. |  | 5 | 5 units | Distributed in the states of CA, IL, NC, and NY. | Terminated | August 07, 2015 | 81 | 5 | Computer
Z-1742-2015 | 70917 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS 4600 Chemistry System, Software Version 3.2 &amp; Below. |  |  | April 06, 2015 | 2015 | Ortho-Clinical Diagnostics | Software anomaly allows testing of multiple assays using a single urine specimen, regardless of the pretreatment requirements for each selected assay, potentially leading to biased patient results. | Software design | On April 6, 2015, Ortho Clinical Diagnostics distributed correction notification letters and response forms to their customers via courier service. All urine patient results that require pretreatment was advised to refer to the notice of correction for instructions. Correction notices should also be posted beside the device user documentation for easy reference. Customers were advised to complete and return the Confirmation of Receipt before April 20, 2015. Customers with questions may contact the Customer Technical Service at 1-800-421-3311 anytime..***CAPA***Since the Assay Data Disk (ADD) provides data tables that determine which assays can be tested from the same sample container, an updated ADD is being created to address the anomaly on the VITROS 4600 and 5600. The anomaly will be addressed with an updated ADD is uploaded and the revised ADD mitigates the issue totally. Anticipated date of availability TBD. |  | 255 | 255 units total (131 domestically &amp; 255 internationally) | Worldwide Distribution-US (nationwide), Brazil, Chile, Mexico, Australia, Hong Kong, India, Taiwan, Belgium, Denmark, France, Ireland, Italy, Portugal, Romania, Spain, Sweden and the UK. | Open, Classified |  |  | 255 | Computer
Z-1732-2015 | 71246 | No | DXN | Noninvasive blood pressure measurement system. | Cardiovascular | Cardiovascular | 510(k) | ClearSight, Model No. EVHRS Heart Reference Sensor. The EV1000 Clinical Platform Nl and ClearSight finger cuffs noninvasively measures blood pressure and associated hemodynamic parameters. |  |  | May 18, 2015 | 2015 | Edwards Lifesciences, LLC | Edward Lifesciences is recalling the ClearSight System because the finger component of the Heart Reference Sensor may be damaged. | Device Design | Edwards Lifesciences sent an URGENT Product Notification letter dated May 12, 2015, to all affected customers. The letter informed customers that Edwards Lifesciences has received reports that a component of the ClearSight System can be damaged if not handled properly. The letter informs customers of the problems identified and the actions to be taken. Customers with questions are instructed to call Edwards Technical Support. For questions regarding this recall call 949-250-2500. |  | 1309 | 1,309 units total (696 units in US) | Worldwide Distribution - US (nationwide) and Internationally to Australia, Estonia, Italy, Romania, Sweden, Austria, Finland, Kuwait, Saudi Arabia, Switzerland, Bahrain, France, Netherlands, Slovenia, Turkey, Belgium, Great Britain, Oman, Slovakia, South Africa, Bulgaria, Greece, Poland, South Africa, United Arab Emirates, Czech Republic, Hungary, Portugal, Spain, Denmark, Ireland, and Qatar. | Terminated | November 06, 2015 | 172 | 1309 | Computer
Z-1734-2015 | 71235 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | AccuChek Spirit insulin infusion pump. |  |  | April 22, 2015 | 2015 | Roche Diabetes Care, Inc. | Roche has become aware of an increase of mechanical errors experienced by customers using the Accu-Chek Spirit Combo insulin infusion pump and Accu-Chek Spirit insulin infusion pump when the insulin pumps display E6 and El0 error messages.If the user does not act upon the E6 and E10 error messages appropriately, insulin delivery will be stopped and if unnoticed may lead to severe hyperglycemia u | Device Design | Roche sent an URGENT MEDICAL DEVICE CORRECTION notification dated April 22, 2015, to all affected customers. The letter identified the problem the product and the action needed to be taken by the customer. The recall notification included a description of the reason for the recall, customer responsibilities, training chart, and instructions for responding to the formal recall notification. Please contact the ACCU-CHEK Customer Care Service Center, 24 hours a day, seven days a week at 1-800-688-4578, if you have questions about the information contained in this notification. |  | 35795 | 35,795 | Nationwide Distribution | Open, Classified |  |  | 35795 | Computer
Z-1687-2015 | 67989 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | 100/120V AC, Advanced Perfusion System Platform (APS) The TerumoÂż Advanced Perfusion System 1 is indicated for use for up to 6 hours on the extracorporeal circulation of blood for arterial perfusion, regional perfusion, and cardiopulmonary bypass procedures, when used by a qualified medical professional who is experienced in the operation of this or similar equipment. |  |  | May 01, 2015 | 2015 | Terumo Cardiovascular Systems Corporation | Terumo CVS is implementing field correction activities to address the identified causes of the Terumo System 1 experiencing loss of system power. These activities will be implemented in a phased approach beginning late 2015:- Implementing design changes and upgrades to improve the reliability of power switching and battery backup.- Adding a switch protector to the On/Off power switch to preven | Device Design | On 5/1/2015 Terumo issued an URGENT MEDICAL DEVICE CORRECTION letter to users explaining that Terumo has identified causes of the Terumo System 1 experiencing loss of system power. These activities will be implemented in a phased approach beginning late 2015: - Implementing design changes and upgrades to improve the reliability of power switching and battery backup.- Adding a switch protector to the On/Off power switch to prevent inadvertent shut off.- Replacing the On/Off power switch with a new design to prevent switch failure. A Terumo Field Service Representative will contact users to schedule the field correction activities once the activities are ready to be implemented.CUSTOMER INSTRUCTIONS: Review this Medical Device Correction and assure that all users have received notice of this issue.Âż- Complete and return the enclosed Customer Response Form.Âż- Terumo CVS will contact users to schedule the field correction activities.We encourage you to contact us with any questions or concerns:Terumo CVS Customer Service1-800-521-2818Fax 1-734-741-6149Customer Service Hours:Monday  Friday, 8 a.m.  6 p.m. ET |  | 1713 | 1713 units total | Worldwide Distribution-US (nationwide) including DC and the states of AL , AR , AZ , CA , CO , CT , DE , FL , GA , HI , IA , ID , IL , IN , KY , LA , MA , MD , ME , MI , MN , MO , MS , MT , NC , ND , NE , NJ , NM , NV , NY , OH , OK , OR , PA , SC , TN , TX , UT , VA , VT , WA , WI , and WV, and the countries of Mexico , AUSTRALIA , UNITED ARAB EMIRATES (UAE) , Hong Kong , Indonesia , Singapore , Taiwan , Thailand , Philippines , South Korea , CHILE , Vietnam , Malaysia , BELGIUM , Japan , and CANADA. | Open, Classified |  |  | 1713 | Computer
Z-1683-2015 | 70672 | No | OOE | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | Catalys Precision Laser System-Catalys - U - US product; Catalys-1 - International product; software version 3.00.05 |  |  | February 20, 2015 | 2015 | Optimedica Corporation | Software anomaly on OptiMedica CATALYS System version 3.00.05 that may result in incorrect parameters provided for cataract incision surgeon templates and both eyes have the same parameters. | Software design | Abbott sent an Advisory Notice dated March 5, 2015, vis Fed Ex overnight to all affected customers. Out of State accounts will be notified through their distributor after letters are translated into local language.Recall notification letters advise customers of issue with CATALYS System using software version 3.00.05. The letter actually covers two issues, the cataract templates being the same for both eyes, and a loss of suction during treatment that may result in scoring the cornea.In both cases the issue is explained and users are advised to use mitigation listed to prevent injury. AMO is developing an update to software that will correct these issues and users will be advised when it is available for installation.All users of the device should be made aware of the issues and each surgeon is requested to sign the acknowledgement form to verify their receipt and understanding. The acknowledgement for should be returned to Abbott Medical Optics per instructions. Questions in the US may be directed to AMO Technical support at 1-855-662-0911. |  | 171 | 171 distributed Worldwide | Worldwide Distribution - US Nationwide, Germany, Switzerland, Australia, Spain, Japan, Great Britain, France, Netherlands, Austria, Singapore, South Korea, India, Brazil, Italy, Belgium, Egypt, Colombia, Hong Kong, and Canada. | Terminated | December 08, 2015 | 291 | 171 | Computer
Z-0003-2016 | 72167 | No | OBJ | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | SOUNDSTAR eco 8F and 10F Diagnostic Ultrasound Catheters. For intra-cardiac and intra-luminal visualization of cardiac and great vessel anatomy and physiology as well as visualization of other devices in the heart. |  |  | September 10, 2015 | 2015 | Biosense Webster, Inc. | Image disappeared from the cardiac ultrasound system when the CARTO 3 EP Navigation System needed restarting while the patient was experiencing pericardial effusion. Affects the CARTOSOUND Module of the CARTO 3 EP Navigation System when used with the SOUNDSTAR eco 8F and 10F Diagnostic Ultrasound Catheters. New precautions added. | Device Design | A Customer Notification Letter and Acknowledgement Form was distributed to customers on 09/10/15. One letter and form was sent via express mail, addressed to the Risk Management Office and one additional letter and form was sent by the firm's account representative to the Electrophysiology (EP/Cardiology Lab directly at the customer site.The firm's notification letter stated that they want to emphasize the following statement "The intra-cardiac ultrasound image will disappear if the CARTO 3 EP Navigation System power is disrupted and this may present a safety issue if the EP is using the ultrasound to monitor the patien during EP procedure. The ultrasound image will not reappear until the CARTO 3 System is restored."The firm will be updating the product labeling to further reinforce the precautionary statement.The firm requests that customers take the following actions:1. Read the Field Safety Notification carefully.2. Pass on this notification to anyone in the facility that needs to be informed of this issue, including appropriate clinical personnel involved in the use of SOUNDSTARÂż eco Catheters.3. Review, complete, sign and return the attached Acknowledgement Form in accordance with the instructions on the form. 4. Maintain a copy of this letter with the product.5. Maintain awareness of the Field Safety NotificationFor any questions the firm states to contact the BWI sales representative. |  | 1035 | 1035 (U.S.) | Worldwide Distribution. US nationwide, Italy, Canada, Japan, China, Singapore, Colombia, England, Russia, Spain, France, Germany, Finland, Ireland, Australia, Sweden, Hungary, India, Brazil, Slovenia, Netherlands, Mexico, and South Africa. | Open, Classified |  |  | 1035 | Computer
Z-1679-2015 | 71289 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | RaySearch RayStation 2.5, 3.0, 3.5 and 4.0; Radiation Therapy Treatment |  |  | May 08, 2015 | 2015 | RAYSEARCH LABORATORIES AB | RaySearch Laboratories AB became aware of an issue with the dose calculation for some imported VMAT plans where controls points are defined with wide gantry angle spacing. | Software Design Change | RaySearch sent a Device Correction letter dated May 8, 2015 via e-mailed to their customers. The letter identified the affected product, problem, actions to be taken by users and contact information. |  | 539 | Domestic: 539 units | US Nationwide Distribution | Terminated | April 11, 2017 | 704 | 539 | Computer
Z-1611-2015 | 70891 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | FujiMedical Synapse Cardiovascular I ProSolv CardioVascular. |  |  | February 04, 2015 | 2015 | Fujifilm Medical Systems U.S.A., Inc. | Measurement values in the EMR may not accurately reflect the actual measurement value on the Synapse Cardiovascular clinical report. | Software in the Use Environment | FujiMedical Systems North America issued intial recall letter dated February 4, 2014, advising users of the HL7 report exported to the EMR wherein measurement values presented in the Hl7 reports in your EMR may not accurately reflect the actual measurement value on the Synapse Cardiovascular clinical report.Follow-up letters were issued February 12, 2014, and July 24, 2014. A mandatory upgrade to software version 5.1.1/ base, in-place upgrade required. A customer response form attached to acknowledge receipt and arrange for the on site upgrade. For questions regarding this recall call 203-324-2000. |  | 30 | 30 units | Worldwide Distribution - US including CA, FL, IL, IN, MD, MI, MN, MO, NC, NY, OH, OR, PA, PR, TX, VA, WIForeign: Canada | Open, Classified |  |  | 30 | Computer
Z-1613-2015 | 70995 | aid, transfer | IKX | Daily activity assist device. | Physical Medicine | Physical Medicine | 510(K) Exempt | ReTurn belt, wipeable, sizes Small through Extra Large, Plastic pouch, 1 item per pouch. Patient transfer aid. |  |  | March 26, 2015 | 2015 | Handicare USA, Inc. | Complaints were received that the heat-welded seam on the wipeable belt's strap tore when force was applied. | Device Design | Handicare sent an Voluntary Product Recall Notification letter dated April 2, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were requested to quarantine any remaining stock and to contact customers if product was further distributed and to return the affected product for a replacement and/or full refund.Also, enclosed is a response form. Please complete and return the form even if you do not have any product to return. Your Sales representative can assist you in completing this form. This information is essential in order to maintain recall effectiveness information (required by FDA).Thank you for your support. We are committed to re-releasing the wipeable ReTurn belt as soon as we believe that it meets the quality standards that you expect from Handicare. Please contact our Customer Experience team at 1-866-276-5438 if you have any questions regarding the recall or would like assistance with the recall. We regret any inconvenience that this action may cause, but we appreciate your understanding as we take action to ensure patient and customer satisfaction. |  | 64 | 64 | US Distribution to the states of : FL, CA, NJ, NC and MD. | Terminated | August 17, 2015 | 144 | 64 | Computer
Z-2822-2015 | 72106 | No | MUU | Suction lipoplasty system. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ASP-1021LipoTower with Aspiration Pump and Tumescent Pump. |  |  | June 12, 2015 | 2015 | MicroAire Surgical Instruments, LLC | There is potential risk of electrical shock if lead wires inside the device become detached and the ground connection fails. | Device Design | The firm has issued FDA recall letters to all consignees. The letter have been sent out by UPS to expedite the delivery. |  | 1189 | 168, ASP-1021 LipoTower with Aspiration Pump and Tumescent Pump | Nationwide Distribution. | Open, Classified |  |  | 1192 | Computer
Z-1620-2015 | 71027 | No | GWM | Intracranial pressure monitoring device. | Neurology | Neurology | 510(k) | The Integra Licox Pt02 Monitors are compact, portable device for use with Pressure/Temperature, Pressure and Oxygen Partial Pressure Transducer-Tipped Catheters. |  |  | April 14, 2015 | 2015 | Integra LifeSciences Corp. | Integra LifeSciences Corporation is recalling Integra's Camino ICP Monitor (CAM 02) &amp; Licox PtO2 Monitor ( LCX02 &amp; LCX02R) due to complaints that the audible speaker malfunctions. | Device Design | Integra initiated this recall by sending an Urgent Voluntary Medical Device Recall Letter/Recall Acknowledgement Form dated 4/14/15 via traceable courier service to their affected customers. |  | 59 | 59 monitors | Worldwide Distribution. | Open, Classified |  |  | 1814 | Computer
Z-1621-2015 | 71049 | [6 empty] | MYLA server HP Proliant: DL380-G8 and ML350-G6 version 3.xx computer application software. |  |  | April 14, 2015 | 2015 | bioMerieux, Inc. | The MYLAÂż server could slow down due to the volume ( weight ) of the data to manage and it could also stop processing, impacting the server start-up phase. The problem occurs in high volume settings. The MYLA server slows down and will stop processing completely if any task requires more than 5 minutes to complete. The server then fails to reboot and requires service, provided remotely or on-site | Software design | Consignees were notified electronically on April 14, 2015 and letter via FedEx on 04/20/2015. |  | 730 | 730 units | Worldwide Distribution-Algeria, Argentina, Australia, Austria, Brazil Canada, Chile China, Colombia, Czech Republic, France Germany , Greece, Hungary, India, Italy Japan, Korea, Mexico, Netherlands, Poland, Portugal, Russia, Serbia, Singapore, South Africa, Spain, Sweden, Switzerland, Thailand, Turkey, Angola, Armenia, Kocaeli, Bahrain, Bangladesh, Republic of Belarus, Benin, Bolivia, Bosnia, Brunei, Bulgaria, Burkina Faso, Cambodia, Cameroon, Chad, Democratic Republic of Congo, Costa Rica, Germany Dominican Republic, Dutch Antilles, Ecuador, Egypt, El Salvador, Estonia, Ethiopia, Gabon, Georgia, Guam, Guatemala, Guinea, Honduras, Indonesia, Iran, Iraq, Israel, Ivory Coast, Kazakhstan, Kenya, Kuwait, Laos, Latvia, Lebanon, Libya, Lithuania, Macedonia, Malaysia, Mali, Malta, Morocco, Portugal, Myanmar, Managua, Niger, Nigeria, Oman, Palestine, Panama, Paraguay, Peru, Philippines, Qatar, Romania, Kingdom of Saudi Arabia, Senegal, Slovakia, Sri Lanka, Tanzania, Togo, Tunisia, Ukraine, UAE, Uraguay, Venezuela, Vietnam, Yemen and Zimbabwe. | Open, Classified |  |  | 730 | Computer
Z-1628-2015 | 70946 | No | NEP | Intravascular administration set. | General Hospital | General Hospital | 510(K) Exempt | ABACUS SE (Single-Workstation Edition) and ABACUS ME (Multi-Workstation Edition) |  |  | April 08, 2015 | 2015 | Baxter Corporation Englewood | Baxter Corporation is conducting a field action for the ABACUS SE and ME models due to the possibility that the compounder will load the incorrect formula upon scanning the bag label when two or more different formula files have been created with the same order number. | Software design | Baxter sent a Safety Alert letter dated April 8, 2015 to affected customers via USPS first class mail. The letter identified the affected product, problem and actions to be taken. Customers were asked to 1) contact Baxter Technical Services if you need to restore the ABACUS databse, 2) foword a copy of this Safety Alert letter to other facilities or departments within your institution to ensure that those locations are awae of this action, and 3) complete the customer reply form and return it to Baxter by faxing to 224-270-5457 or scanning and e-mailing it to fca@baxter.com. For questions contact the Center for One Baxter at 800-422-9837. |  | 882 | 882 | Worldwide Distribution - US Nationwide | Open, Classified |  |  | 882 | Computer
Z-1555-2015 | 70916 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips Healthcare DuraDiagnost X- Ray |  |  | February 09, 2015 | 2015 | Philips Medical Systems, Inc. | The system is designed to emit a beep upon termination of an exposure. However, if the system has been powered on for over 12 hours, the system will no longer emit this signal. This is a failure to comply with 21CFR 1020.31(j). | Software design (manufacturing process) | Philips Healthcare Planned action: 1. Philips healthcare will contact customers and initiate a software update to correct the software deficiency causing the problem.2. Philips engineers will install the software.3. Philips engineers will perform testing to ensure the software update was effective.4. Corrections are to be made free of charge5. Corrective actions should be completed by August 15, 2015For further questions call (978) 659-7892 |  | 3 | 3 DuraDiagnost | Worldwide Distribution - Nationwide Distribution and the countries of Australia, Austria, Argentina, Canada, China, Chile, Colombia Czech Republic, Great Britain, Germany, Egypt, Finland, France, India, Ireland, Japan, Malaysia, Macedonia, Mexico, Netherlands, New Zealand, Nepal, Peru, Poland, Russia, Saudi Arabia Singapore, Switzerland, Sweden, Spain, United Arab Emirates and United Kingdom. | Open, Classified |  |  | 57 | Computer
Z-2784-2015 | 71995 | Yes | NVZ |  |  | Cardiovascular | PMA | Merlin@home RF Remote Monitoring Transmitter Model EX1150. |  |  | August 18, 2015 | 2015 | St Jude Medical Cardiac Rhythm Management Division | Some devices exhibit backup VVI operation due to a backup reset. | Software change control | The notification material is in the form of a Physician Letter was delivered to physicians by FedEx overnight service. The mailing went out on August 17, 2015 for receipt on August 18, 2015.The firm's notification letter identifies clinical implications.Clinical ImplicationsIn the event that an Ellipse, Fortify Assura, Unify Assura, or Quadra Assura ICD enters the back-up safety mode, the nominal operational settings will be VVI pacing mode, 67 ppm, 5.0v/0.6ms with bipolar pacing output and defibrillation settings of aVF detection rate of 146 bpm and 36J high voltage therapy. In the event an Assurity or Allure pacemaker enters backup mode, it will have output settings of VVI pacing mode, 67 ppm, 5.0v/0.6ms with unipolar pacing.This anomaly can only occur when the patient is being actively monitored by a Merlin@home bedside transmitter. If a device enters backup mode, the ICD will deliver a patient vibratory alert and the pacemaker will deliver a patient audible alert. If backup operation is encountered, St. Jude Medical Technical Services (1-800-722-3774) can assist with non-invasively restoring the device to normal operation.The firm's notification also idenfities recommendations and mitigations.Recommendations and MitigationsOnce available, a Merlin@home transmitter software update will be performed automatically over its telephone, broadband or cellular connection without requiring any action from you or your patients. No changes to the patients remote or in-clinic follow up schedules are required. In the event a patients device reverts to back-up mode, the firm recommends bringing the patient back in the clinic to clear the condition and return the device to full functionality.If you have any questions or concerns contact the local St. Jude Medical representative or St. Jude Medical's Technical Services Department. |  | 259214 | 259,214 transmitters (U.S. distribution) 23,604 (foreign distribution) | Worldwide Distribution-US (nationwide) and foreign distribution to the following: Japan, France, Austria, Canada, Italy, Spain, Netherlands, South Africa, Great Britain, Australia, Sweden, Belgium, Finland, Germany, Portugal, Singapore, Norway, Poland, Ireland, Czech Republic. Denmark, Hong Kong, Saudi Arabia, Hungary, Switzerland. | Open, Classified |  |  | 259214 | Computer
Z-2818-2015 | 72051 | No | ODG | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | EVIS EXERA II Gastrovideoscope OLYMPUS GF type UCT180; designed for endoscopic real-time ultrasound imaging, ultrasound guided needle aspiration and other endoscopic procedures within the upper gastrointestinal tract and surrounding organs. |  |  | August 21, 2015 | 2015 | Olympus Corporation of the Americas | Insufficient information for specific cleaning accessories used for the cleaning and reprocessing of the GF-UCT180 ultrasound endoscope | Device Design | Olympus sent an Urgent Device Safety Information letter dated August 2015, to all affected customers. Corrected instructions were provided with the letter and users were asked to begin implementing them immediately. Customers were asked to indicate on the enclosed questionnaire and fax the completed from to 484-896-7128. Customers with questions were instructed to call 484-896-5688. |  | 1014 | 1014 | Nationwide Distribution | Terminated | May 08, 2017 | 626 | 1014 | Computer
Z-1632-2015 | 71105 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | EPIQ 7 Ultrasound System versions 1.3.2 or lower, WITHOUT the Pediatric Cardiology option, Model: EPIQ 7G, EPIC 7C, EPIQ 7W. Catalog Number: 795200 / 795201 |  |  | April 14, 2015 | 2015 | Philips Ultrasound, Inc. | When EPIQ 7 Ultrasound System, WITHOUT the Pediatric Cardiology option, is set up to "Metric" and weight and/or height is entered, a unit conversion error may result in the incorrect display of their values and incorrect calculation of the patients Body Surface Area (BSA). | Software design | Philips sent the Urgent-Medical Device Correction-Philips Model EPIQ5 Ultrasound system and the Urgent-Medical Device Correction-Philips Model EPIQ7 Ultrasound system letters, both dated 2015 APR 14, to their US consignees. Philips sent the Urgent-Field Safety Notice-Philips Model EPIQ5 Ultrasound system letter and the Urgent-Field Safety Notice-Philips Model EPIQ7 Ultrasound system letter (both letters dated 2015 APRIL 14) to their foreign consignees. The EPIQ5 or EPIQ7 Ultrasound systems WITHOUT the Pediatric Cardiology option are affected. Philips issued two Field Change Orders (FCO) to their Field Service Engineer to install upgraded software: -FCO 79500340 for EPIQ5 -FCO 79500341 for EPIQ7 Customers with questions can call Philips Customer Service at 1-800-722-9377.Field Service Engineers (FSE) to install software version 1.3.3 per FCO 79500340 instructions for EPIQ5 5 and per FCO 79500341 instructions for EPIQ7. |  | 1692 | 1692 units total (902 units in the US and 790 units outside the US) | Worldwide Distribution | Open, Classified |  |  | 2800 | Computer
Z-2812-2015 | 71915 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARTISTE, ONCOR Avantgarde, ONCOR Impression Plus, ONCOR Expression, PRIMUS, PRIMUS Plus, Syngo RT Therapist; therapeutic treatment of cancer |  |  | August 05, 2015 | 2015 | Siemens Medical Solutions USA, Inc | A software fix has been released to prevent automatic movement resulting in a collision safety risk for patients. | Software design (manufacturing process) | Siemens sent Notification letters to affected customers. The letter identified the affected product, problem and actions to be taken. The letter advise customers that the software fix was rolled out to address multiple safety issues. |  | 26 | 26 | Worldwide Distribution and US Nationwide | Open, Classified |  |  | 26 | Computer
Z-1636-2015 | 71127 | No | OYD |  |  | General &amp; Plastic Surgery | PMA | MelaFind Device, Model MF 100 US, optical diagnostic device for melanoma detection. |  |  | April 02, 2015 | 2015 | MELA Sciences, Inc. | The probability and histogram data within the Melafind's device displayed user interface lacks approval (not included in the PMA supplement). | Software design | The probability and histogram data within the Melafind's device displayed user interface is not included in the PMA supplement |  | 65 | 65 units | AL, CA, CT, CO, FL, HI, IL, MA, MD, MI, NH, NJ, NY, OH, PA, RI, SC, TN, TX, VA and WA | Open, Classified |  |  | 65 | Computer
Z-1642-2015 | 71193 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | DeRoyal(R) Medium Specimen Retrieval Bag, REF 5500-000-003, Rx Only |  |  | April 17, 2015 | 2015 | DeRoyal Industries Inc | The firm received reports of specimen retrieval bags tearing, ripping, and coming apart during use. | Device Design | DeRoyal sent an Urgent: Voluntary Recall letter dated April 17, 2015 to their customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to 1. Check your inventory for the product and specific lot number from the attached Recall Response Form.2. If you have affected inventory, please destroy it and dispose of it according to your facilitys procedures.3. Record all information on the included Recall Response and Notice of Destruction forms.4. If you purchase through a distributor, please indicate their name/location on the Notice of Destruction Form so that we may issue credit appropriately. 5. Return both forms to recalls@deroyal.com or 865-362-3716 (fax) even if you no longer have affected inventory.6. Once we have received your Recall Response and Notice of Destruction forms, you will receive credit for destroyed product as indicated on the form.7. If you are a distributor of these products, please forward this recall notice on to your end-users, or provide us with your customer listing and we will gladly contact them for you.For questions regarding this recall, please contact your DeRoyal Sales Representative or DeRoyal Customer Service at 1-800-251-9864. |  | 5359 | 5359 units | US Nationwide Distribution | Terminated | December 29, 2015 | 256 | 5359 | Computer
Z-1643-2015 | 71117 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Pinnacle3 Software Version 10.0, Model 459800200841. |  |  | April 23, 2015 | 2015 | Philips Medical Systems, Inc. | A problem (the dose computed in planning mode is calculated incorrectly ) has been detected in the Philips Pinnacle3 software version 10.0 that, if it were to re-occur, could pose a risk for patients or users. Specifically, the dose engine is being passed the wrong snout position. The snout position is used to determine the penumbra of the beam which includes the calculation of the source si | Software design | Philips Healthcare sent a "Field Safety Notification PROST-T-300-01-06" letter dated 20 April 2015 to their customers. The letter described the Problem Description, Hazard Involved, How to Identify Affected Products, Action to be Taken by Customer / User, Actions Planned by Philips and Further Information and Support (Contact Philips Customer Care Solutions Center at 1-800-722-9377). |  | 4 | 4 | US Nationwide Distribution in the states of MO, FL, OH | Open, Classified |  |  | 4 | Computer
Z-1641-2015 | 71215 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ENDOPATH ECHELON" Vascular White Reload for Advanced Placement Tip, |  |  | February 26, 2015 | 2015 | Ethicon Endo-Surgery Inc | During an internal inspection of the production process the firm discovered an issue which may cause the cartridge to deploy an incomplete staple line. | Software in the Use Environment | Ethicon sent an Urgent Medical Device Recall letter dated February 27, 2015 to their customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to examine your inventory immediately to determine if you have recalled product on hand, remove and quarantine the recalled product, communicate the issue to relevant operating room or materials management personnel or anyone else in your facility whoneeds to be informed, if any product included in this recall has been forwarded to another facility, contact that facility to arrange return, and complete the Business Reply Form (BRF) (Attachment B) confirming receipt of thisnotice within three (3) business days. The BRF may be sent to Stericycle by fax at 1-800-807-5967or by email at ees8760@stericycle.com. Please return the BRF even if you do not have affected product.For clinical or product support, please contact your local sales representative orcall our Customer Support Center at 1-877-ETHICON (1-877-384-4266) (7:30 a.m.  6:30 p.m. EST). If you need additional shipping labels or a communications package, contact Stericycle at 1-877-643-8419 (8:00 a.m.  5:00 p.m. EST) and reference Event8760. |  | 6744 | 6,744 devices | Worldwide Distribution - US Nationwide in the states of: AR, AZ, CA, CO, DE, FL, GA, IA, IL, KS, KY, LA, MA, ME, MI, MN, MO, NC, NE, NJ, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA &amp; WI and the country of Hong Kong. | Terminated | November 17, 2015 | 264 | 6744 | Computer
Z-1645-2015 | 71101 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | SOMATOM Emotion 16, the intended use of computed tomography is to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data from either the same axial plane taken at different angels or spiral planes taken at different angles. |  |  | April 20, 2015 | 2015 | Siemens Medical Solutions USA, Inc | Software bugs in VC20B SP0a or SP1 software versions may cause issues that could make it necessary to rescan patients. Syngo Main UI may crash if the Patient Browser is scrolled with arrow keys of the keyboard. Tomo images may freeze on the screen when zoomed in or out under the mode of CAREVision. Sporadic displaying error in WorkStream4D application. DB may lock under heavy and multitasks. | Software design | Siemens sent an Customer Safety Advisory Notice dated April 20, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. A remote software update is scheduled for release that will resolve the issue. For further questions, please call (610) 219-6300. |  | 75 | 75 | Nationwide Distribution including Puerto Rico. | Open, Classified |  |  | 75 | Computer
Z-1669-2015 | 71115 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | SIEMENS Luminos Agile Max; a universal imaging system for radiographic and fluoroscopic studies. Using either film cassettes or a digital mobile flat detector, it can perform a range of applications including general R/F, angiography and pediatric examinations. Luminos Agile is applicable to emergency treatment on an outpatient basis, as well as for bedside examinations. |  |  | April 13, 2015 | 2015 | Siemens Medical Solutions USA, Inc | Study and all acquired images deleted when using systems with software version VE10E. Acquisitions have to be repeated. Connection with portable detectors (MAX wi-D, MAX mini) is sporadic in image recovery. Pressing undo button on the RAD subtask card or previous series button on the Image subtask card during image readout. Sporadically, during an automatic or a manual RIS update. | Software design | A safety advisory notice, dated April 13, 2015, was sent to end users with affected devices to alert them about the issue and associated risks,and to provide instructions for avoiding potential risk. A new software version will be released to permanently resolve the issue. |  | 133 | 133 total | Nationwide Distribution. | Terminated | March 20, 2017 | 707 | 133 | Computer
Z-2806-2015 | 71706 | No | FET | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fujifilm Processor EPX 2500 High Definition Endoscopy; (Sales Brochure). |  |  | June 30, 2015 | 2015 | Fujifilm Medical Systems U.S.A., Inc. | The EPX-2500 Operation Manual and sales brochure incorrectly identify that the ED-530XT endoscope is meant to be used with the EPX-2500 Processor. The Operation Manual and sales brochure state that when the ED-530XT endoscope is used with the 2500 processor, a "super image" is displayed when it is actually a "standard image" that is displayed. | Device Design | Fujifilm Medical Systems USA, Inc. initiated this recall by sending a Voluntary Field Correction Letter and Tracking/Verification Form dated June 30, 2013 to their affected customers. |  |  | multiple | Nationwide US and Canada | Terminated | May 25, 2017 | 695 |  | Computer
Z-1678-2015 | 71110 | No | BSY | Tracheobronchial suction catheter. | Anesthesiology | Anesthesiology | 510(K) Exempt | KimVent Turbo-Cleaning Closed Suction System for Adults, 10 F T-Piece (Product Code 227105); KimVent Turbo-Cleaning Closed Suction System for Adults, 12 F, DSE (Product Code 2271603); KimVent Turbo-Cleaning Closed Suction System for Adults, 12 F, T-Piece (Product Code 227135). |  |  | April 14, 2015 | 2015 | Halyard Health, Inc | The internal blue flapper valve of certain KimVent Turbo-Cleaning Closed Suction System for Adults may have a potential defect. If present, the defect potentially allows air to leak out of the closed ventilator circuit. | Device Design | Halyard Health Regulatory Affairs sent an Urgent Voluntary Product Recall letter dated April 16, 2015, to all affected consignees. Consignees were instructed to quarantine and discontinue use of the recalled product. Specific return instructions were provided in the Notice. The consignees were also requested to provide a response to to the recall notification. Wholesalers were instructed to follow their internal recall procedures to identify and address affected products. Consignees with questions were instructed to call Halyard Health Regulatory Affairs at 470-448-5700. |  | 220 | 11 cases each containing 20 individual endotracheal tubes | Nationwide Distribution including CA, MA, MI, NC, TX, VA , and WA. | Terminated | July 07, 2015 | 84 | 220 | Computer
Z-1675-2015 | 70959 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Panorama Patient Monitoring Network, Multi- Parameter Patient Monitor (with Arrhythmia Detection and Alarm. Part number 0998-00-0708-01. |  |  | March 12, 2015 | 2015 | Mindray DS USA, Inc. dba Mindray North America | Software anomaly in the Panorama System software version 8.9 that manifests when the Panorama is in use with Mindray Passport V Monitor. Arrhythmia alarms which were previously displayed on both the Passport V and Panorama will not be displayed on the Panorama following a restart of either system (due to communication loss, restart, etc.). | Software design | Mindray sent a Corrective Action letter dated March 12, 2015 to their affected customer via FedEx. The letter identified the affected product, problem and actions to be taken. For questions call 201-995-8407. |  | 4 | 4 units | Worldwide Distribution - US Nationwide in the state of TX. | Open, Classified |  |  | 4 | Computer
Z-1671-2015 | 71185 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | EliteCore Full Core Biopsy Device. EliteCore 18G, 20cm with and without HiLiter tip, 15 cm with and without HiLiter tip, and 10 cm with and without HiLiter tip. |  |  | April 24, 2015 | 2015 | Stryker Instruments Div. of Stryker Corporation | Potential for the device cannula to overthrow past the intended length. | Device Design | Stryker sent an Urgent Medical Device Recall Notification letter dated May 5, 2015, to all affected customers via Fed Ex overnight. Customers were instructed to qaurantine any product found, return the repsonse form even if there is no product on hand, forward the letter if further distribution of recalled product occurred, return recalled product to Stryker. Customers were instructed to 269-389-2458 with any questions. |  | 1420 | 1,420 Units | US Nationwide Distribution including AL, AR, AZ, CA, CO, FL, GA, IL, IN, KS, KY, LA, MI, MN, MO, MS, NJ, NM, NV, NY, OK, OR, PA, TN, TX, UT, VA, and WI. | Terminated | October 15, 2015 | 174 | 1420 | Computer
Z-0121-2016 | 72168 | No | IMJ | Shortwave diathermy. | Physical Medicine | Physical Medicine | 510(k) | Monode, Shortwave Accessory |  |  | August 21, 2015 | 2015 | DJO, LLC | Overheating and burning of the wire insulation inside the applicator. | Device Design | The firm sent out two customer notification letters; one to distributors and one to direct users. Letters were sent September 1, 2015 via USPS certified return receipt.Distributors were asked to obtain a list of their customers with contact information that have purchase or received the monode accessory or Intelect Shortwave device that included the monode as an accessory item. Once the list has been assembled the firm requests that it be forwarded to DJO Global Chattanooga Product Support by phone 877-733-4888, or by e-mail at monode.support@djoglobal.com. The firm's notification states that customers will be sent an upgraded monode accessory along with a pre-paid shipper for the return of the current used unit.Any questions the letter states to contact Product Support per the contact information above or William Fisher, Global Regulatory Compliance at 760-734-3126.Letters sent to direct users received the same instruction of returning the product for a replacement. Any questions the letter states to call 877-733-4888 or email at monode.support@djoglobal.com. |  | 2887 | 2887 | Worldwide Distribution - US Nationwide in the states of WY, OH, WA, NC, FL, NE, GA, KS CA, MO, IN, MA, IL, OK, AZ, TX, AR, UT, NJ, WI, PA, SC, WV, NY, MI, MN, VA, AK, TN, OR, NC, NH, KY, CO, NV, ID and the countries of El Salvador, China, Ecuador, South Korea, Chile, Israel, Peru, Nicaragua, Bolivia, Indonesia | Terminated | October 24, 2016 | 430 | 2887 | Computer