## Sheet1
Recall Number | Recall Event ID | Device Type | Product Code | Regulation Description | Medical Specialty | Review Panel | Submission Type | Trade Name/Product | Main Name | Recall Class | Date Posted | Year Posted | Recalling Manufacturer | Reason for Recall | FDA Determined Cause | Action | Consumer Instructions | Clean Quantity | Quantity in Commerce | Distribution | Recall Status | Termination Date | Time to Terminate | Merged Quantities | Fault_Class | 0
Z-2327-2016 | 74669 | No | FZP | Implantable clip. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Horizon 24 Large Titanium Clips with Adhesive Backing, Ref 004204, Rx Only, Sterile. |  |  | August 08, 2013 | 2013 | Teleflex Medical | The manufacturing and expiration dates are not printed on the blister pack label. | Error in labeling | Teleflex sent an Urgent Medical Device Recall Notification dated August 8, 2013, to all affected customers. The letter requested that they quarantine the product and to return it to Teleflex. An acknowledgment form was included with the letter which also was to be returned to Teleflex. A separate letter was sent to distributors requesting they conduct a sub-recall.If you have any further questions, please call (610) 378-0131. |  | 8220 | 8,220 ea. | USA Distribution to the states of :AL, AZ, CA, CT, FL, IL, KY, MD, MI, MN, NM, NY, NC, OR, PA, SC, TX, VA, WA, WI, and Internationally to Canada | Completed |  |  | 8220 | Not_Computer | Not_Computer
Z-0078-2017 | 75350 | No | NAW |  |  | Radiology | PMA | TheraSphere(R) YTTRIUM-90 CLASS MICROSPHERE INJECTION |  |  | September 21, 2015 | 2015 | Biocompatibles U.K., Ltd. | One mislabeled dose of TheraSphereÂż Y-90 Glass Microspheres, which was shipped to Hong Kong (total radioactivity incorrect). The product ws labeled as 10 GBq and the actual does was 5 GBq. | Packaging | Biocompatibles UK Ltd sent an Urgent Field Safety Notice dated September 21, 2015, to all affected customers. The firm did not request return or destruction of the product. Customers with questions were instructed to call 613-410-1870.For questions regarding this recall call 613-801-1848. |  | 64 | 64 doses | Worldwide Distribution - US to CA, FL, IL, WI, IN, OR, PA, MA, NY, KY, TX, MI, MD, NV, WA, OH, MN, and Internationally to Germany, Turkey, Hong Kong, Italy, Portugal | Completed |  |  | 64 | Not_Computer | Not_Computer
Z-0401-2017 | 75256 | No | JIY | Iron (non-heme) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension Vista IRON Flex reagent cartridges |  |  | August 25, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has become aware of deferoxamine interference with Dimension and Dimension Vista IRON assays at a different concentration than listed in the Instructions for Use (IFU). | Error in labeling | Siemens mailed an Urgent Medical Device Correction letter to all customers who received the Dimension and Dimension Vista IRON Flex reagent with an update to the Limitation of Procedure section for the Dimension and Dimension Visa IRON assays Instructions For Use (IFU) to indicate that: patients treated with metal-binding drugs (e.g. deferoxamine) may have depressed iron values, as chelated iron may not properly react in the iron assay. Customers were told that the information provided in the letter, related to deferoxamine, supersedes the information related to deferoxamine int he current IRON IFU's for Dimension and Dimension Vista products until the IFU's are updated. |  | 30247 | 30,247 units | Distributed Nationwide | Open, Classified |  |  | 68477 | Not_Computer | Computer
Z-1909-2016 | 73979 | No | PFX |  | Immunology | Immunology | 510(k) | CytoScanÂż Dx Assay. The CytoScan Dx Assay consists of five reagent modules, wash buffers (WS), a microarray kit, and analysis software. |  |  | April 14, 2016 | 2016 | Affymetrix Inc | During a recent Ohio site inspection by the FDA it was discovered that warning statements needed to be included with the product shipment. | Labeling design | A letter dated April 14, 2016 was sent to customers. The letter stated the reason for the recall. It also informed customers that there is no impact on the performance of the assay. Customers were asked to either visit the IVD web portal (www.affymetrix.com/ivd) and review the warning statements; or, request a hard copy free of charge. |  | 507 | 507 kits | Distributed in the US to the states of AR, CA, DC, MI, MA, MN, NC, NM, NY, PA, SC, TN, TX, WA and WV. | Open, Classified |  |  | 507 | Not_Computer | Not_Computer
Z-2282-2016 | 74452 | No | JIT | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Instructions for Use (IFU) for Progesterone 3 Calibrator Set, repackaged as one set/package. |  |  | June 07, 2016 | 2016 | Tosoh Bioscience, Inc. | Instructions For Use had incorrect lot number and calibrator values | Error in labeling | Tosoh Bioscience sent an Urgent Correction Notification letter dated June 7, 2016 to all affected customers. The letter identified the affected product, problem and actions to be taken. For questions call Tosoh Technical Support at 1-800-248-6764. |  | 123 | 41 single sets of Progesterone 3 calibrators | US Nationwide Distribution in the states of AL; CA; CO; FL; KY; MA; NY; OH; SC; TX; UT; VA and WI. | Open, Classified |  |  | 123 | Not_Computer | Not_Computer
Z-2767-2016 | 75006 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Sterile Drape, SK6000 PAQ drape pouch (for use with the SPY Fluorescent System (SPY 2000/2001)) |  |  | August 18, 2016 | 2016 | NOVADAQ TECHNOLOGIES INC. | Novadaq is informing SK6000 PAQ users that a pinhole through the drape pouch film material was identified in some SK6000 PAQs during shipping validation testing. | Packaging | The firm, Novadaq, sent an "URGENT DEVICE CORRECTION: SK6000 PAQ Drape Pouch" letter dated 8/18/2016 to affected consignees. The Firm instructed consignee to cease distribution of affected product; accept the shipping-validated replacement PAQ; complete and sign the enclosed "Acknowledgment of Receipt" form, scan it and email to quality@novadaq.com. If you have any questions, contact Director of Regulatory Affairs at 1-905-629-3822 ext 205 or email: jpendlebury@novadaq.com.. |  | 20 | 20 units | US Distribution in states of: FL, MI, and VA. | Open, Classified |  |  | 20 | Not_Computer | Not_Computer
Z-2086-2016 | 74173 | No | HSN | Ankle joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Cadence Size 5 Left Tibial Tray Implant, designed to treat ankle arthritis through replacement of the ankle joint with a prosthesis, thereby reducing pain, restoring alignment, and allowing for movement at the replaced joint. Catalog Number: 10207205; Lot Number #QJ0248 |  |  | April 29, 2016 | 2016 | Integra LifeSciences Corp. | As a result of an internal review of all labeling for the Cadence System, it was determined that the outer package label for the Size 5 Left Tibial Tray Implant reflects the material composition incorrectly. The labeling incorrectly reflects the material as Cobalt Chromium instead of Titanium. | Labeling mix-ups | On 4/29/16, US consignees were notified by electronic mail delivered by traceable method. |  | 12 | 12 units | NC and PA | Open, Classified |  |  | 12 | Not_Computer | Not_Computer
Z-1178-2016 | 73267 | No | GDM | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Pilling Modular Verres Body/Obturator, Product Code 140880 |  |  | January 08, 2015 | 2015 | Teleflex Medical | Packaging error in which customers who ordered product code 140880 (Verres cannula for use with a separately available needle tip) were incorrectly provided with product code 140878 (Verres cannula with integral needle tip). | Packaging process control | Teleflex sent an Medical Device Advisory Notification letter dated January 8, 2016, to all affected customers notifying them of the packaging error.Our records indicate that you have received products that are subject to this notification.Teleflex Medical is committed to providing high quality, safe and effective products. We sincerely apologize for any inconvenience this action may cause your operations. If you have any other questions, feel free to contact your local sales representative or Customer Service at recalls@teleflex.com or 1-866-246-6990. |  | 8 | 8 | US Distribution to the state of : UT | Terminated | April 14, 2017 | 827 | 8 | Not_Computer | Not_Computer
Z-2759-2016 | 75020 | No | GDO | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | WECKÂż Horizon" Manual-Load Ligating Clip Applier, For Medium-Large Size Horizon Ligation Clips Curved Jaw, 11 (28cm), Rx Only, REF 337111, |  |  | August 24, 2016 | 2016 | Teleflex Medical | Misbranded: Incorrect etching on the device. | Error in labeling | Teleflex sent an Urgent Medical Device Notification letter dated August 24, 2016 to customers via FedEx. The letter identified the affected product, problem and actions to be taken. For questionscontact your local sales representative or Customer Service at 1-866-246-6990. |  | 49 | 49 units | Worldwide Distribution US Nationwide in the states of FL, IL, KY, MI, NY and OK and countries of Mexico and Canada. | Open, Classified |  |  | 49 | Not_Computer | Not_Computer
Z-2745-2016 | 74771 | No | MAI | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Smith &amp; Nephew Whipknot Soft Tissue Cinch (USP#5) suture. PN: 7211015 Used during ligament reconstruction. |  |  | July 19, 2016 | 2016 | Smith &amp; Nephew, Inc. | Sterility compromised due to breach in sterile packaging | Package design/selection | Smith &amp; Nephew issued recall letters via Federal Express on July 19, 2016 to the customers. Customers asked to inspect inventory and locate any unused devices, and quarantine immediately. Contact Smith &amp; Nephews Field Actions Department via email at FieldActions@smith-nephew.com or fax +1-901-566-7975 to return the completed Inventory Return Certification Form and obtain a return authorization (RA) number if you have product to return. |  | 34065 | 34,065 units | Worldwide distribution. US nationwide, and the countries of: AU, AT, BE, CA, CH, CR, DB, DE, ES, FI, FR, GB, GR, IC, IL, IN, IT, MY, NL, NO, PL, PR, PT,SE, SG, TH, and ZA. | Open, Classified |  |  | 34065 | Not_Computer | Not_Computer
Z-1916-2016 | 74107 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Liko Universal SlingBars |  |  | March 08, 2016 | 2016 | Hill-Rom, Inc. | The center bolt of the sling bar, which connects the bar to the patient lift, has failed during use. If the sling bar is not used as intended the bolt may be weakened. If the bolt is weakened there is a potential risk for breakage with the result of a free fall of the patient. This hazard could cause potentially minor to catastrophic injuries to a patient. | Labeling design | On 3/8/2016, MEDICAL DEVICE CORRECTION notifications were sent to the affected US customers via email and courier. On 3/24/2016, MEDICAL DEVICE CORRECTION notifications were sent to the affected Canadian customers via email and courier. On 3/29/2016, MEDICAL DEVICE CORRECTION notifications were sent to the affected Global customers via email and Docapost. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. The notification provided new instructions for users with a lift manufactured prior to 11/19/2013, instructed the user on how to inspect for damage, and offers replacement at a 50% reduction in cost. If you have any questions concerning this Field Safety Notice, please contact Hill-Rom Technical Support, your distributor, or your Hill-Rom representative. |  | 163561 | 163,561 | Worldwide and US Nationwide Distribution AE, AO, AR , AT , AU , AZ , BD , BE , BG , BH , BM , BN , BR , CA , CH , CL , CN , CO , CS , CZ , DE , DK , EG , ES , FI , FR , GB , GR , HK , HR , HU , ID , IE , IL , IN , IR , IS , IT , JO , JP , KR , KW , LB , LI , LK , LT , LU , LV , LY , MA , MT , MX , MY , NL , NO , NZ , OM , PH , PK , PL , PT , QA , RO , RS , RU, SA , SE , SG , SI , SK , SY , TH , TN , TR , TW , UA , US , UY and ZA. | Terminated | May 15, 2017 | 433 | 163561 | Not_Computer | Not_Computer
Z-2764-2016 | 74893 | No | GWL | Physiological signal amplifier. | Neurology | Neurology | 510(k) | Patient Cable. Part Numbers: 4460, 6735. Sterile within Chipboard Boxes within corrugated cartons. |  |  | August 10, 2016 | 2016 | Blackrock Microsystems, LLC | Blackrock Microsystems, LLC announces a voluntary field action for the Blackrock NeuroPort Biopotential Signal Processing System because the Neural Signal Amplifier and the Patient Cable are incorrectly labeled as Type CF Applied Parts. | Labeling False and Misleading | Blackrock Microsystems sent an Urgent Medical Device letter dated August 12, 2016 via email. The letter identified the affected product, problem and actions to be taken. Customers were instructed not use the amplifier in application requiring a type CF applied part; perform a physical count of the inventory of amplifiers and patient cables with the part numbers listed. For questions regarding this Correction call 801-582- 5533, ext. 240. |  | 120 | 120 total cables | Worldwide Distribution - US Nationwide and Canada, | Open, Classified |  |  | 120 | Not_Computer | Computer
Z-0674-2017 | 75589 | No | MUU | Suction lipoplasty system. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Adiprep Adipose Transfer System. Product Codes: 51431; 51432; 51433; 514001-347. |  |  | November 03, 2016 | 2016 | Terumo BCT, Inc. | Terumo BCT announces a voluntary field action for the Harvest : Adiprep Adipose Transfer System because the affected product is incorrectly labeled and has the potential for operator confusion. | Error in labeling | Consignees will be notified on 11/07/16 via letter to contact the North America Terumo BCT support center at +1.877.842.7837 or +1.303.542.5050 if they have any affected product to issue a Return Goods Authorization for product return. |  | 660 | 660 | Nationwide, Canadian, and other foreign consignees. No VA/govt/military consignees. | Open, Classified |  |  | 660 | Not_Computer | Not_Computer
Z-0732-2016 | 73007 | No | NWV | Gynecologic laparoscope and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(K) Exempt | PS-500 Scope Warmer; 10 units per box. |  |  | December 01, 2015 | 2015 | JosNoe Medical, Inc. | Damage to package seals potentially compromising sterility. | Packaging process control | On 12/1/2015, notifications were sent to the affected customers via e-mail. The recall notification included a description of the reason for the recall, the affected product, and instructions for identifying and returning the affected product. |  | 1008 | 1008 boxes | Distributed US (nationwide) and in Canada. | Terminated | March 22, 2016 | 112 | 1008 | Not_Computer | Not_Computer
Z-2751-2016 | 74798 | No | KKY | Polytetrafluoroethylene with carbon fibers composite implant material. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Hard Tissue Replacement (HTR) Implant, Rx only, for use in single procedure only. Custom made. |  |  | June 13, 2016 | 2016 | Biomet Microfixation, LLC | Potential sterility deficiency. | Packaging | Zimmer Biomet Customers were sent an Urgent Product Recall letter dated June 13, 2016, via via FedEx mail. The letter notified customers of a recall of Hard Tissue Replacement Polymer Implants due to voids in sterile packaging. The firm instructed customers to locate and discontinue use of the implants, complete the Fax Response form and fax it to 904-741-9425, send back any quarantined product using the FedEx labels provided along with the recall notification, and to notify facilities where the implant was further distributed by forwarding them a copy of the provided notice. The firm instructed distributors to forward the notice to officials of their customers who are in charge of receiving recall notices. Additionally, the firm informed distributors that they are responsible for locating and returning distributed product. Questions are to be directed to 1-800-874-7711 or 904-741-4400 x9574, Monday thru Friday from 8am to 4pm ET. Her after hours phone is 812-614-9499.For questions regarding this recall call 904-741-9225. |  | 54 | 54 units | Worldwide Distribution - US including AZ, CA, FL, MO, NH, PA, TX, &amp; VA, and Internationally to Austria, Canada, England, Germany, Greece, Italy, Netherlands, &amp; South Africa. | Open, Classified |  |  | 54 | Not_Computer | Not_Computer
Z-0737-2016 | 72945 | No | DTK | Cardiovascular intravascular filter. | Cardiovascular | Cardiovascular | 510(k) | Denali Filter-Jugular/Subclavian DL950J |  |  | March 13, 2015 | 2015 | Bard Peripheral Vascular Inc | IFU missing contraindications: Patients with uncontrolled sepsis and Patients with known hypersensitivity to nickel-titanium alloys. | Labeling Change Control | The firm sent out Medical Device Recall customer notification letters on 03/13/15.The notification letter stated that there is no required action on the customer's part. Customers were told that if the device did not meet their needs then the firm will provide a replacement unit. Effective immediately the firm is implementing a change in the ordering code for the device to accurately reflect the approved information with the IFU and corresponding retrieval kit.Any questions the firm states to call Bard Canada:1-800-268-2862 Western Canada1-800-387-9473 Ontario1-800-387-7851 Eastern Canada |  | 1183 | 1,183 units | Distributed to CA only. | Terminated | February 01, 2016 | 325 | 1183 | Not_Computer | Not_Computer
Z-1397-2016 | 72992 | No | KKX | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Catalog Number: 9001771 MAYO STAND COVER STERILE |  |  | January 07, 2016 | 2016 | Customed, Inc | Recall initiated due to inadequate validation/qualification of sterilization processes, and package integrity for products released for distribution as sterile. | Packaging | Recall Notification Letter sent to direct accounts via certified mail with return receipt on January 7, 2016 and January 8, 2016. |  | 540 | 540 | Distributed only in Puerto Rico. | Terminated | May 03, 2017 | 482 | 62224 | Not_Computer | Not_Computer
Z-0403-2017 | 75492 | No | FAD | Ureteral stent. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Bard Fluoro-4 Silicone Ureteral Stent, BardÂż Fluoro-4", Silicone Ureteral Coil Stent; BardÂż Silicone Uretera l Coil Stent with Figure Four, End, Manufacturer: C.R. Bard, Inc., Covington, GA 30014. |  |  | October 13, 2016 | 2016 | C.R. Bard, Inc. | Labeling: Statement was added to the Warnings/Precautions section requesting end users to consider the presence of a knot if significant resistance is encountered during attempts at removal. | Labeling Change Control | Bard Medical sent an Urgent: Medical Device Communication Letter dated October 13, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Action required:Examine your inventory and identify any product subject to this notification. If you have any remaining inventory, include a copy of this communication with the unit(s) for future use.If you may have further distributed or transferred this product, please identify the respective organizations and notify them at once of this communication. Your notification to these organizations may be enhanced by including a copy of this product communication letter.This notification should be carried out to the user level.This product notification is being made with the knowledge of the Food and Drug administration.We appreciate your cooperation and assistance in dealing with this matter and sincerely apologize for any inconvenience that may result from this action.For further questions, please call (770) 784-6220. |  | 75089 | 75,089 units | U.S. Nationwide. | Open, Classified |  |  | 75089 | Not_Computer | Not_Computer
Z-2004-2016 | 74263 | No | MBP | Resorbable calcium salt bone void filler device. | Orthopedic | Physical Medicine | 510(k) | Musculoskeletal Transplant Foundation, AFT Diverted Tube (3/4/Filled), indicated for treatment of surgically-created osseous defects or osseous defects created from traumatic injury. It can be used in the: extremities, pelvis, and spine. Product Code: 227005. |  |  | April 21, 2016 | 2016 | Musculoskeletal Transplant Foundation, Inc. | AFT Diverted Tubes are being recalled due to a customer complaint involving one unit which had an unsealed Tyvek pouch. The non-sterile foil pouch may introduce contamination into the sterile field. | Packaging process control | Musculoskeletal Transplant Foundation (MTF) sent an Voluntary Field Action Notification Letter dated April 21, 2016, and via telephone to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were notified of the situation with the product and provided instructions on how to handle the product. Questions can be directed to Customer Service Customer at 1-800-433-6576. |  | 114 | 114 units | nationwide | Open, Classified |  |  | 114 | Not_Computer | Not_Computer
Z-2292-2016 | 74693 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Avenir MÂżller Stem 4 Standard |  |  | June 30, 2016 | 2016 | Zimmer Gmbh | Product mix-up. The Avenir MÂżller Stem 6 lateral uncemented might be placed in the packaging of the Avenir MÂżller Stem 4 standard uncemented and vice versa.Potential for delay in surgery, and/or potential for use of incorrect implant which may in turn lead dismetry in patient and possibly the need for revision. | Labeling mix-ups | Zimmer Biomet sent an Urgent Medical Device Recall letter dated June 30, 2016 to all affected customers via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Customers were instructed to review the notification immediately and ensure affected personnel are aware of the contents and assist Zimmer Biomet sales representative with the quarantine of any affected device. Your Zimmer Biomet sales representative will remove any affected device. Customers were also instructed to complete the Certification of Acknowledgement form and return to corporatequality.postmarket@zimmerbiomet.com. Customers with further questions or concerns were instructed to contact the customer call center at 1-800-348-2759 between 8:00 am and 8:00 pm EST.For questions regarding this recall call 800-348-2759. |  | 2 | 2 | Nationwide Distribution to OH and NM | Terminated | March 27, 2017 | 270 | 2 | Not_Computer | Not_Computer
Z-0893-2017 | 75737 | No | KGO | Non-powdered surgeon's glove. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Gammex Non-Latex Sensitive Neoprene Surgical Gloves |  |  | October 24, 2016 | 2016 | Ansell Healthcare Products LLC | Ansell Healthcare Products is initiating a voluntary recall involving 15 cases of Gammex Non Latex Sensitive Surgical Gloves because the packaged contents incorrectly included a natural rubber Latex surgical glove. | Packaging change control | Ansell Healthcare Products, LLC. sent an Urgent: Voluntary Surgical Glove Recall letter dated October 28, 2016, to all affected customers. The letter identified the problem the product and the action the customer needs to take.Please check your product stock at the warehouse and shelf level in order to advise us directly if any gloves having the recalled lot numbers are found. This information, along with your tracking number, must be communicated to our Customer Service Department by calling 1-800-952-9916, or by returning the attached Response Form to Ansell by fax at1-800-722-8155, or via email customerserviceus@ansell.com . Returned product will either be replaced or credited without delay. Your Ansell representative will be available to follow up with you regarding any concerns you may have. |  | 15 | 15 Cartons | US Distribution to the states of : PA, FL, NJ, IL and MA. | Open, Classified |  |  | 15 | Not_Computer | Not_Computer
Z-2294-2016 | 74222 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Endotine Transbleph 3.5 (CFD-080-0167) |  |  | May 06, 2016 | 2016 | MicroAire Surgical Instruments, LLC | MicroAire has initiated a limited recall (29 Units)on the Endotine Transbleph 3.5 (CFD-080-0167), Lot # 285373, due to minor mislabeling ( the front label is correct, but the side label lists an incorrect part number, i.e. CFD-080-0177). | Labeling Change Control | The firm is sending recall letters to the customers who have received the Endotine Transbleph 3.5 (CFD-080-0167), lot #285373. Letters sent by UPS to expedite the delivery. |  | 29 | 29 | The product was distributed in the United States and Sweden | Completed |  |  | 29 | Not_Computer | Not_Computer
Z-0080-2017 | 75240 | No | EAJ | Personnel protective shield. | Radiology | Radiology | 510(K) Exempt | Zero-Gravity Radiation Protection System |  |  | September 09, 2016 | 2016 | Contour Fabricators Inc | Two customer complaintssuggest that excessive rotation of the cable in one direction (clockwise or counter-clockwise)without allowing the cable to unwind to its normal condition may cause the cable to breaksuddenly and the body shield to fall. While no injuries or deaths occurred in either failure, theZero-GravityTM body shield weighs 55 pounds, and in the event of a cable failure, the bodyshield will fall immediately. A fall of this nature places the user, associated staff and the patientat risk for a potential injury. | Labeling design | On 9/9/2016, URGENT: MEDICAL DEVICE CORRECTION letters dated 9/8/2016 were sent to the affected consignees via certified mail. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. TIDI Products is taking the following action to correct this issue:1. Revising instructions for use to include the step to allow the cable to unwind after useand prior to storing the body shield in a locked or secured position.2. Modifying user and staff training protocols to include the requirement to allow the cableto unwind after each use.3. Adding additional warning labels to the system.4. Replacing all cables in the field to ensure that any potentially stressed cable is removed.A representative from TIDI Products will be contacting you to schedule a visit from a technicianto implement this corrective action.Customer ActionsCables should be carefully inspected for signs of wear or rotational tension, which will create abend or buckle in the cable. Practitioners and medical institutions should review their routineuse of the Zero-Gravity Radiation Protection System to determine if the body shield undergoesrepeated rotation without allowing the cable to relax back to its normal condition. If the cablehas been consistently rotated in one direction without being allowed to relax back to its normalcondition or if there is any doubt about excessive rotation in one direction, TIDI recommendsthat practitioners and medical institutions discontinue use until such time as corrective actioncan be implemented by TIDI Products.If facilities continue to use their Zero-Gravity" Radiation Protection System until correction andinstall of the replacement cable can be performed, they should make sure that they are allowingthe unit to go back to its relaxed state.Please complete and return the enclosed response for |  | 447 | 447 | Nationwide. Canada.AUSTRALIAAUSTRIABELGIUMBRUNEICHINAFRANCEGERMANYINDAINDONESIAISRAELITALYKSAMALAYSIANETHERLANDSPOLANDQATARQUEBECSAUDI ARABIA SINGAPORESOUTH AFRICATAIWAN (R.O.C.) | Open, Classified |  |  | 447 | Not_Computer | Not_Computer
Z-1237-2016 | 73398 | No | LJE |  |  | Gastroenterology/Urology | 510(k) | UltraxxTM Nephrostomy Balloon Set |  |  | January 27, 2016 | 2016 | Cook Inc. | Cook Medical is voluntarily recalling specific lots of the Ultraxx" Nephrostomy Balloon Set and the AscendÂż AQÂż Ureteral Dilation Balloon Catheter Set because of a labeling mix-up. Cook Medical has received reports of packages labeled as Ultraxx" Nephrostomy Balloon Set actually containing theAscendÂż AQÂż Ureteral Dilation Balloon Catheter Set and vice versa. | Labeling mix-ups | Cook Medical sent an Urgent Medical Device Recall Notification letter dated January 27, 2016, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Customers were instructed to quarantine any unused affected product. Immediately collect and return all unused affected products to Cook Medical If you have any questions or concerns, please do not hesitate to contact Cook Medical Customer Relations at (800) 457-4500 or 1(812) 339-2235.For questions regarding this recall call 800-457-4500. |  | 20 | 20 | Worldwide Distribution - US including GA, ID, LA, OR, SC, TX and Internationally to CHINA. | Terminated | May 17, 2017 | 476 | 20 | Not_Computer | Not_Computer
Z-1570-2016 | 73748 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | BD Trucount Absolute Counting Tubes: |  |  | March 31, 2016 | 2016 | Becton, Dickinson and Company, BD Biosciences | One lot of lot of BD Trucount tubes may contain tubes from another lot. There is a 3.3% difference in bead counts between the two lots found within some of the pouches and presents a low risk to patient safety or impact to data if this product is used. | Packaging | BD Biosciences sent an Urgent Product Recall letter dated March 2016 to all affected customers on March 31, 2016 by Certified Mail, e-mail and direct delivery by Sales Associates.Letters instructed that any remaining inventory of the affected lot be discarded and that date files obtained using reagents from the affected lot be reviewed for the possibility of incorrect data.Customers with questions were instructed to contact BD Customer Support at 877-232-8995 (prompt3, 3) in the united State. For customers outside the US, customers should contact their local BD Biosciences representative or distributor.For questions regarding this recall call 408-954-6080. |  | 1743 | 1743; 467 US and 1276 International | Worldwide Distribution - US including MD, CA GA. WA. PA. CO. TN. FL. NC. DE. VA, CT, LA. NC, NY, OH, MN, NH and Internationally to China, Japan, and India. | Open, Classified |  |  | 1743 | Not_Computer | Not_Computer
Z-1049-2016 | 73248 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Discovery LS, 2286327 |  |  | February 10, 2016 | 2016 | GE Medical Systems, LLC | GE Healthcare has updated the Service Manuals for a number of CT and PET CT Scanners to update an incorrect torque value. | Labeling design | GE Healthcare sent an Urgent Medical Device Correction letter " GEHC Ref# 25467-ABT or 25467" dated February 12, 2016, to all affected consignees . The letter was addressed to Hospital Administrators / Risk Managers, Managers of Radiology/Cardiology, &amp; Radiologist/Cardiologist. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. Consignees were instructed to ensure that the organization that service their equipment uses the latest version of the Service Manual to maintain the safe and proper performance of their medical device. For questions call in the United States: 1-800-437-1171 or your local GE Healthcare Service Representative. For questions regarding this recall call 262-513-4122. |  | 817 | 817 | Worldwide Distribution - US including all states in including PR except ND, VT &amp; WY.Foreign Distribution: ALBANIA, ARGENTINA, AUSTRALIA, AUSTRIA, BELGIUM, BOSNIA &amp; HERZEGOVINA, BRAZIL, BULGARIA, BURKINA FASO, CAMERON, CANADA, CHILE, CHINA, COLOMBIA, DOMINICAN REPUBLIC, ESTONIA, FINLAND, FRANCE, GERMANY, GREECE, GUADELOUPE,HONG KONG, HUNGARY, ICELAND, INDIA, INDONESIA, IRA N, IRAQ, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KENYA, REPUBLIC OF KOREA, KUWAIT, LATVIA, LEBANON, LIBYAN ARAB JAMAHIRIYA, LITHUANIA, LUXEMBOURG, MALAYSIA, MEXICO, MOROCCO, NETHERLANDS, NEW ZEALAND, NORWAY, OMAN, PERU, PHILIPPINES, POLAND, PORTUGAL, REUNION, ROMANIA, RUSSIA, SAUDI ARABIA, SINGAPORE, SLOVAKIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, TUNISIA, TURKEY, UKRAINE, UNITED ARAB EMIRATES, UNITED KINGDOM, VENEZUELA. | Open, Classified |  |  | 1793 | Not_Computer | Not_Computer
Z-0388-2017 | 75388 | No | KNS | Endoscopic electrosurgical unit and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | SB Knife -Electrosurgical Knife, Sterile |  |  | September 09, 2016 | 2016 | Akita Sumitomo Bakelite Co., Ltd. | Mislabeled product: Sterilization pouch and carton box labeled as Standard type instead of Jr type-Long | Error in labeling | Sumitomo Bakelite on September 9, 2016 initiated removal of all SB Knife from US accounts.The information was emailed to the accounts and Sumitomo Bakelite requested all customers not to use the mislabeled SB Knife and return product for replacement. |  | 15 | 15 units | US in the states of MA | Terminated | November 10, 2016 | 62 | 23 | Not_Computer | Not_Computer
Z-1900-2016 | 73200 | No | MRA |  |  | Orthopedic | PMA | Novation Ceramic AHS Cluster-Hole Shell, Plasma Coated, 54mm, O.D. (Acetabular Cup), Catalog Number 140-01-54; For use in primary total hip arthroplasty. |  |  | April 12, 2016 | 2016 | Exactech, Inc. | Devices were inadvertently sent to packaging and finished goods prior to the PMA site-change supplement approval. | Packaging process control | Customers were notified 4/12/2016 with a Product Recall Advisory Notice of affected products.In order to comply with applicable regulations and Exactech policies: " Immediately cease distribution or use of these products. " Extend this information to your accounts that may have this product in their possession. " Identify and quarantine any of the subject devices in your inventory " Complete and Return the attached Recall Inventory Response Form to Exactech, per the instructions provided on the notice itself. According to policy, please contact your Exactech inventory representative within 5 business days to confirm quantities at your location. If you have any inventory restocking questions related to this issue, please contact Kaya Davis at kaya@exac.com or 1-800-392-2832. Please return the products to the following address:Exactech Distribution Center 2411 NW 66th Court Gainesville, Florida 32653Attention: Kaya Davis kaya@exac.com 352-337-3915 (Fax) |  | 3 | 3 devices | Distributed in FL, NJ, NY, OH, OK, RI, TN, VA and HI. | Open, Classified |  |  | 20 | Not_Computer | Not_Computer
Z-0887-2016 | 73077 | Yes | KPO | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius Naturalyte Liquid Acid Concentrate |  |  | January 18, 2016 | 2016 | Fresenius Medical Care Renal Therapies Group, LLC | On the label calcium (Ca) listed in the right upper corner of the label (red background) incorrectly indicates the product contains 2.5 mEq/L Ca. The concentration indicated in the list of constituents represents the actual calcium concentration of the final dialysate, 3.0 mEq/L. | Error in labeling | The firm, Fresenius Medical Care, issued an "URGENT-MEDICAL DEVICE RECALL" letter dated 1/14/2015 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately examine your stock for the product; if found, discontinue use immediately; place all units in a secure, segregated area; contact your FMCNA Customer Service Team at 1-800-323-5188 for instructions how to return the affected lot (lot# 15STAC072), and promptly complete and return the attached reply form by Fax to: 781-699-9796 Attn: Quality Department or scan and email to: NOTIFYRA@fmc-na.com. Fresenius will pick up and replace your product.If you have any additional medical concerns or questions, please contact Medical Information and Communication at 855-816-2309 or Website: www.fresenius-medinfo.com.Further information or support concerning this issue, contact the Customer Service Care Team at 800-323-5188 and reference the and Support Field Action Number. |  | 848 | 848 (212 cases) | US Distribution to states of: KS, LA, MS, OK, and TX. | Terminated | August 14, 2017 | 574 | 848 | Not_Computer | Not_Computer
Z-2126-2016 | 73198 | No | MOQ | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Battery Power Line II Battery Oscillator, DePuy Synthes Power Tools, The Anspach Effort Inc. |  |  | November 25, 2015 | 2015 | The Anspach Effort, Inc. | Has the incorrect operating directions on the mode switch ring, i.e. its is labeled "FWD, LOCK, REV" instead of "ON, LOCK, ON". | Labeling design | DePuy Synthes sent an Urgent Notice Recall Notification dated November 24, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.They were instructed to review their inventory and immediately remove any affected devices, complete the verification section of the letter and return it to the Anspach Effort Inc., and return any affected product to their DePuy Synthes Sales Consultant. Any questions regarding this recall / removal contact the Complain Handling Unit Manager at (561) 494-3673 or their DePuy Synthes Sales Consultant.For further questions regarding this, please call (561) 627-1080. |  | 20 | 20 hand pieces | US Distribution to the state of : VA and Internationally to:Thailand and UK | Open, Classified |  |  | 20 | Not_Computer | Not_Computer
Z-0880-2017 | 75792 | No | NDE | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | GAM Kirschner Wire, Sterile R, T2 K-Wire, Sterile R and T2 K-Wire Recon, Stryker Trauma GmbH, Prof.-Kuntscher-Strasse 1-5, Schonkirchen, Germany 24232; T2 K-Wire, Sterile R |  |  | November 09, 2016 | 2016 | Stryker Howmedica Osteonics Corp. | A review of packaging revealed the seal integrity of the pouch may be compromised. More specifically, there is a potential that the sterile pouch is not sealed at one end due to a manufacturing error. | Packaging process control | Stryker sent an Urgent Medical Device Removal letter dated November 9, 2016, to all affected customers via UPS (with return receipt). Stericycle will be handling the returns. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to inform users of the Medical Device Recall and forward the notice to all those individuals who need to be aware within their organization. Return all affected products available at their location to:Stryker C/O Stericycle2670 Executive Dr., Suite AIndianapolis, IN 46241Attn: RA 2016-169 - Event XXXCustomers were also instructed to complete and sign the enclosed Business Reply Form and return to Stericycle.Customers with questions should call +1-201-851-6683.For questions regarding this recall call 201-831-5000. |  | 2084 | 2,084 units | Worldwide Distribution - US (nationwide) Internationally to AU, CA, CL, CH, ES, FR, GB, SE, JP, FR, ZA, and NL. | Open, Classified |  |  | 2084 | Not_Computer | Not_Computer
Z-1578-2016 | 73680 | No | JJH | Clinical sample concentrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Hand-Held Scanner USB IT3800 |  |  | March 03, 2016 | 2016 | Roche Molecular Systems, Inc. | The hand-held barcode scanner model IT3800 used with the COBAS AmpliPrep instrument mis-identified a sample barcode ID. | Labeling design | Roche Molecular Systems sent an "Urgent Medical Device Correction (UMDC) notification letter dated March 3, 2016, to their affected customers. Customers were instructed to provide the UMDC to other affected parties who may need to be aware of this issue. Customers were asked to complete the attached fax form and fax it to 1-866-737-5396. Customers with questions should contact Roche Support Network Customer Support Center at 1-800-526-1247. For questions regarding this recall call 908-253-7200. |  | 6939 | 6,939 pieces | Nationwide Distribution | Terminated | May 25, 2017 | 448 | 6939 | Not_Computer | Not_Computer
Z-1830-2016 | 73703 | No | OWD |  |  | Molecular Genetics | PMA | cobasÂż EGFR Mutation Test, v2 and cobasÂż cfDNA Sample |  |  | March 15, 2016 | 2016 | Roche Molecular Systems, Inc. | An error was found within the Hungarian translations of the cobasÂż EGFR Mutation Test v2 Instructions for Use (M/N07340761001-01HU, Doc Rev. 1.0, Dated 08/2015) and the cobasÂż cfDNA Sample Preparation Kit Instructions for Use(M/N 07573758001-01HU, Doc. Rev. 1.0, Dated 05/2015). | Packaging | Roche Molecular Diagnostics Inc. issued an Urgent Field Safety Notice dated March 15, 2016 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Actions to be taken by the customer:" If you have utilized 2 ÂżL of DNA PBB during testing of plasma specimens, the samples must be retested using the correct volume (2mL DNA PBB).Communication of this Field Safety Corrective ActionThis notice must be passed on to all those who need to be aware within your organization or to any organization/individual where the potentially affected devices have been distributed/supplied.Please maintain awareness of this notice and resulting action for an appropriate period to ensure the effectiveness of the corrective action.The undersigned confirms that this notice has been notified to the appropriate Regulatory Agency.For questions regarding this recall call 908-253-7200. |  | 8 | 8 kits | Hungary | Open, Classified |  |  | 8 | Not_Computer | Not_Computer
Z-2719-2016 | 75033 | No | FRO |  |  | General &amp; Plastic Surgery | 510(k) | ReliaMedÂż Silver Alginate/CMC. Each dressing is packed in a pouch (primary packaging), each pouch is packed in a carton (secondary packaging) holding multiple (depending on stock keeping unit) pouches and each carton contains one (1) Instructions For Use (IFU) leaflet. |  |  | January 15, 2015 | 2015 | Advanced Medical Solutions Ltd. | On 08 January 2015 AMS Sole Customer for the products notified AMS of the issue after discovering the misplaced IFUs during its review of packaging in its warehouse | Package design/selection | Advanced Medical Solutions' (AMS) Sole Customer has already received notification via telephone conference call, followed by emails and correspondence. |  | 47569 | 47,569 cartons | OH | Terminated | November 23, 2016 | 678 | 76773 | Not_Computer | Not_Computer
Z-0069-2017 | 75297 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | V. Mueller Neuro/Spine, SHADOW-LINE, SPINAL DISTRACTION SCREW, STERILE, various sizes (12mm, 14mm, 16mm and 18mm) 5 PER BOX. |  |  | September 23, 2016 | 2016 | Cardinal Health 200, LLC | Potential risk associated with corrosion demonstrated on the distraction screw surface. | Package design/selection | On September 23, 2016, Becton Dickinson sent a Field Safety Notice to all affected customers describing the details of the recall and necessary action to take. ACTIONS TO BE TAKEN BY BD:A credit will be applied to your account for the affected product upon receipt of the affected unit(s) or receipt of certification of destruction documented on the customer response form.Complete the Recall Response form certifying that all affected product in your inventory has been destroyed and mail/email a copy of the response form.: The firm requests customers to promptly return the enclosed response form to expedite the correction process and acknowledge receipt of this notification. Or email: GMB-GLBÂżVMFieldActions@carefusion.com For further questions, please call (847) 473-1500. |  | 72076 | 72,076 units | Worldwide Distribution - USA (nationwide) and the countries of Australia, Canada, Germany, New Zealand, Saudi Arabia, South Africa, Switzerland and Thailand. | Open, Classified |  |  | 72076 | Not_Computer | Not_Computer
Z-0704-2017 | 75796 | No | EBC | Pit and fissure sealant and conditioner. | Dental | Dental | 510(k) | UltraSeal XT Plus Clear Refill. The package contains 4x 1.2 ml UltraSeal XT Plus. |  |  | November 17, 2016 | 2016 | Ultradent Products, Inc. | There was a labeling error with product, UltraSeal XT Plus Clear Refill, Part Number (PN) 565, Lot Number (Lot) BD79T. The package was mislabeled as PQ1, PN 641; while the barcode, IFU and products were labeled for UltraSeal XT Plus Clear Refill, PN 565, Lot BD79T. | Labeling mix-ups | Consignees notified via phone on 17Nov2016 and by letter on 22Nov2016 to identify and return products. Consignees with any questions regarding this field action can be sent to 801-553-4858 or 801-553-4220. |  | 31 | 31 | Worldwide Distribution - US Nationwide | Terminated | June 05, 2017 | 200 | 31 | Not_Computer | Not_Computer
Z-2120-2016 | 74273 | No | GZO | Galvanic skin response measurement device. | Neurology | Neurology | 510(K) Exempt | The ZYTO Tower is the input device to program the software with various virtual items. |  |  | November 23, 2015 | 2015 | ZYTO Technologies, Inc. | ZYTO Technologies Inc. Announces a Voluntary Recall of the ZYTO Tower and ZYTO Select and Elite Software Due to Claims Exceeding the 510(k) Clearance. | Labeling False and Misleading | Consignees were notified on 11/23/2015 via letter and email of an update to the ZYTO Select and Elite Software that would disable all tower functionality.. |  | 1252 | 1252 total | Nationwide. Canadian and other foreign consignees. No VA/govt/military. | Open, Classified |  |  | 1252 | Not_Computer | Computer
Z-2733-2016 | 75028 | No | HBG | Manual cranial drills, burrs, trephines, and their accessories | Neurology | Neurology | 510(k) | Integra Cranial Access Kits, Catalogue Nos. INSHITH; INSHITHRZN; INSHITHND |  |  | August 24, 2016 | 2016 | Integra LifeSciences Corp. d.b.a. Integra Pain Management | Integra LifeSciences has been notified of a medical device recall by Adaptive Surgical, LLC involving their Riverlon" Nylon 3-0 Sutures that are packaged in Integras Cranial Access Kits. It was found that product seals may become open during extreme distribution conditions and render the device unsterile. | Packaging | Consignees notified on August 12, 2016 via letter to identify and return product. |  | 710 | 710 | Nationwide and VA/govt/military. No foreign consignees. | Terminated | March 30, 2017 | 218 | 710 | Not_Computer | Not_Computer
Z-2096-2016 | 74301 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | AngioDynamics Soft Vu Omni Flush Angiographic Catheter |  |  | June 01, 2016 | 2016 | Stryker Sustainability Solutions | Product defect; tip of two (2) Angiographic Catheters model # 10732203 broke off, prior to being used on the patient. | Incorrect or no expiration date | URGENT MEDICAL DEVICE RECALL RESTERILIZED ANGIODYNAMICS SOFT VU OMNI FLUSH ANGIOGRAPHIC CATHETER customer notification letters were sent 06/01/16.Customers were instructed to discontinue use of the recalled product immediately. Customers were told to complete the Recall Effectiveness Check Form and list the affected lots shipped to the facility. The form should be completed, signed, and returned to the local Stryker Sustainability Sales Representative, emailed to ssspfa@stryker.com, or mail to:Stryker Sustainability Solutions1810 West Drake DriveTempe, AZ 85283Attn: Jodie Rueckert Customers will receive a credit for all affected devices returned. |  | 167 | 167 | Worldwide Distribution -- USA, to the states of TX, UT, PA, CA, VT, MA, WA, VA, NY, MD, and CT; and, the country of CANADA. | Terminated | October 19, 2016 | 140 | 167 | Not_Computer | Not_Computer
Z-0702-2017 | 75630 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | MANOS EX (CAT#TY-12-101) |  |  | November 04, 2016 | 2016 | Thayer Intellectual Property, Inc. | The firm failed to notify current customers of IFU update/warning. During a recent FDA inspection, it was discovered that Thayer Intellectual Property, Inc. had revised their IFU to include a warning not to bend, deform or otherwise modify any part of the device. The warning was added as a result of a MAUDE report filed in December 2014 where a piece of the MANOS EX broke during use and was left in the patient's hand. The patient suffered a tendon injury and an additional surgery was required to remove the broken piece of the device and repair the tendon. FDA inspection found that while the firm had updated their IFU, they had not notified customers who had been shipped the device with the older version that did not include the warning. Risk: It is important for MANOS EX users to be aware that bending, deforming, or modifying the device could result in device breakage which could impact patient safety and device performance. | Packaging change control | Thayer sent an Urgent Medical Device Correction letters dated November 4, 2016 by Federal Express. The letter identified the affected product, problem and actions to be taken. Customers were asked to check if any of the affected lots were still available, and if so, to change out the IFU with version 5. For questions call 415-923-9375. |  | 339 | 339 devices | US Nationwide Distribution in the states of CA, TX. MD, MI,VA, DC, UT, CT, ID, SD, IN, and GA. | Open, Classified |  |  | 339 | Not_Computer | Not_Computer
Z-0698-2017 | 75632 | No | PAX |  |  | Neurology | HDE - Humanitarian Device Exemption | Elana Surgical KitHUD |  |  | November 09, 2016 | 2016 | Elana, Inc. | Elana has issued this Field Action to voluntarily remove specific serial numbers of our Elana Surgical KitHUD due to potentially compromised integrity of the outer sterile packaging of the Elana Catheter 2.0. | Packaging | Elana sent an Urgent Medical Device Recall letter dated November 7, 2016, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to not use the affected device and to return the product for replacement. Customers were also instructed to complete the attached Acknowledgement and Receipt Form to facilitate the product return and exchange. Customers with questions were instructed to call 301-215-7010. |  | 4 | 4 | Worldwide Distribution - US including NY and Internationally to Switzerland and the Netherlands. | Open, Classified |  |  | 4 | Not_Computer | Not_Computer
Z-0393-2017 | 75474 | No | DXF | Septostomy catheter. | Cardiovascular | Cardiovascular | 510(k) | NRG Transseptal Needle and NRG RF Transseptal Kit. |  |  | October 11, 2016 | 2016 | Baylis Medical Corp * | The firm is voluntarily recalling lots of the NRG Transseptal Needle and NRG RF Transseptal Kit due to the possibility for an integrity defect to the Tyvek pouch of the NRGTransseptal Needle. | Packaging process control | A recall notification, marked URGENT Voluntary Medical Device Recall was sent by Baylis Medical Company via overnight courier to all customers affected by the recall with an Acknowledgement Form. Customers were asked to examine their inventories, quarantine any affected lots, and complete the attached Acknowledgement Form. The form should be returned to Baylis Medical Company via the instructions provided. Upon receipt, the firm will contact customers with information on how to return affected product. Questions regarding the letter can be directed to the Customer Service Department at 1-800-276-4416. |  | 28093 | 28,093 units | Worldwide Distribution -- US, Canada, Chile, France, Germany, Hong Kong, Ireland, Italy, Mexico, Netherlands, Spain, UAE, and UK. | Terminated | July 10, 2017 | 272 | 28093 | Not_Computer | Not_Computer
Z-2743-2016 | 74613 | No | NOV | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Achillon Minimally Invasive Device for Achilles Tendon Rupture Treatment |  |  | June 17, 2016 | 2016 | Integra LifeSciences Corp. | Data from internal studies suggests there may be a potential that the secondary package seal integrity does not provide the level of sterility intended for Minimally Invasive Achilles Tendon Suture System Achillon. | Packaging | Integra sent an Urgent Voluntary Medical Device Recall letter and recall acknowledgment and return form dated June 17, 2016 to their customers via traceable courier. Customers were asked to review their inventory and complete the attached form. Customers with questions should contact Customer Service at 1-800-654-2873.For questions regarding this recall call 717-840-3431. |  | 6729 | 6,729 units | Nationwide distribution | Open, Classified |  |  | 6729 | Not_Computer | Not_Computer
Z-2744-2016 | 74710 | No | KKX | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Microtek Franz Angiography Drape with Radiation Shield, Item # RBA202, and Microtek Angiography Drape with Radiation Shield and Pouches, Item # RBA203. Sterile EO. Microtek Medical, Inc., Columbus, MS 39702. |  |  | July 14, 2016 | 2016 | Ecolab Inc | Ecolab is recalling 2 models of Angiography Drapes with Radiation Shields because the sterile packaging may contain small holes in the film of the pouch. | Packaging process control | Consignees were sent on 7/14/2016 an Ecolab "Urgent Medical Device Recall" letter dated July 14, 2016. The letter described the problem and the product involved in the recall. Advised consignees to immediately quarantine the product and to notify their customers. For questions contact Customer Service at 1-800-824-3027. |  | 6750 | 6750 eaches (675 cases) | US: OH, NY, CA, MN. | Open, Classified |  |  | 6750 | Not_Computer | Not_Computer
Z-2747-2016 | 74895 | No | LCJ |  |  | General &amp; Plastic Surgery | 510(k) | MENTOR 2" Tall Height TE With Sutures 350cc PN: 354-9312 |  |  | August 09, 2016 | 2016 | Mentor Texas, LP. | Tissue expander product mix found: CPX4 350cc was placed in a carton for the CPX4 450cc | Labeling mix-ups | The firm, Mentor, sent an "MEDICAL DEVICE REMOVAL" letter dated 8/11/2016 to affected consignees. The letter will be followed up with a respective sales representative visit. The letter described the product, problem and actions to be taken. The consignees were instructed to immediately identify all unused units and return them per the normal return process; forward the letter to anyone your facility that needs to be informed; if you have sent these units to another facility, forward the letter to that facility; and maintain awareness of the letter until all units in your inventory have been returned.If you have any questions, please contact your local Mentor sales representative or Mentor Customer Service at 1-800-235-5731. |  | 65 | 65 units | Worldwide Distribution: US (nationwide) and countries of: South Korea, Argentina, Mexico, Canada, Australia and New Zealand. | Open, Classified |  |  | 135 | Not_Computer | Not_Computer
Z-1163-2016 | 72900 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Product 65 consists of all product under product code: KTT and same usage: |  |  | January 11, 2016 | 2016 | Zimmer Manufacturing B.V. | As a result of the insufficient sealer calibration data for product packaged in firm's Building II between August 2010 and April 2013. The affected products are sterile. | Packaging | Zimmer Biomet sent an Urgent Medical Device Recall Letter, dated January 11, 2016, to Distributors, Sales Representatives, Operation Managers, and Risk Managers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to 1.Review the notification and ensure that relevant personnel are aware of the contents.2. Assist your Zimmer Biomet sales representative with the quarantine of any affected product.3. Your Zimmer Biomet sales representative will remove the recalled product from your facility.4. Complete and return the attached Certificate of Acknowledgment form to corporatequality.postmarket@zimmerbiomet.com.5. If after reviewing this notification you have further questions or concerns please call Customer Service at 1-800-348-2759, or contact your Zimmer Biomet Sales Representative. |  | 5538 | 5538 | US Nationwide Distribution | Terminated | April 13, 2017 | 458 | 665035 | Not_Computer | Not_Computer
Z-2778-2016 | 74940 | No | LBS | Lipoprotein test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | LDX Optical Check Cassettes (OCC) |  |  | August 02, 2016 | 2016 | Alere San Diego, Inc. | Incorrect expiration date on the magnetic strip that does not match the labeled expiration date. | Packaging | Alere issued a recall notification to customers via telephone and email. Customers was informed of the affected product, problem and actions to be taken. Customers were instructed if they have the affected lot contact the Alere technical service department at 877-308-8289 to report this lot and receive a free replacement. For questions call 877-308-8289 and reference their specific case number. |  | 12951 | 12,951 | Worldwide Distribution - U.S Nationwide in the states of: TN, MA, GA, OK, VA, FL, TN, MO, AL, TX, IL, NY, VT, RI, KY, WI, AL, NC, LA, SD, TX, VA, MI, MN, IN, IA, CA, AR, PA, HI, WV, KS, CO, CT, UT and the countries of Japan, England, Canada | Open, Classified |  |  | 12951 | Not_Computer | Not_Computer
Z-2779-2016 | 73989 | No | NBC | B-type natriuretic peptide test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | VIDAS NT-pro-BNP |  |  | April 15, 2016 | 2016 | Biomerieux France Chemin De L'Or | Due to a packaging error; The kit does not contain the correct number of C1 or C2 control vials. | Packaging | bioMerieux. sent an Urgent Product Removal Notice letter dated April 26, 2016 to customers via FedEx. The letter identified the affected product, problem and actions to be taken. Customers were instructed to check for affected products in their inventory and discard the incomplete products. Customers were asked to complete the attached Acknowledgement Form and return it via fax to bioMerieux. For questions contact your local bioMerieux Clinical Customer Service organization at (800) 682-2666. |  | 229 | 229 kits | US Nationwide Distribution in the states of: AL, AZ, CA, FL, GA, IA, KS, KY, MA, MI, MN, NC, NE, NJ, NY, OH, SC, TX, VA and WV. | Terminated | July 21, 2017 | 462 | 229 | Not_Computer | Not_Computer
Z-0757-2016 | 73059 | No | FDT | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Olympus TJF-Q180V flexible gastrointestinal Duodenoscope |  |  | January 08, 2016 | 2016 | Olympus Corporation of the Americas | Olympus America Inc is conducting a voluntary removal/corrective action of all TJF-Q180V duodenoscopes in order to replace the forceps elevator mechanism Olympus is replacing the forceps elevator mechanism with a new forceps elevator design consistent with the design specification in the recently cleared TJF-Q180V 510k. | No Marketing Application | Olympus America Inc. sent an Urgent Notification letter dated January 15, 2016 to customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to inspect inventory for the affected product, Olympus will contact facilities to make arrangements for return of the TJF-Q180V duodenoscope(s) for the elevator mechanism replacement. For questions call (484) 896-5688. |  | 4436 | 4436 | US Nationwide Distribution except PR | Terminated | January 09, 2017 | 367 | 4436 | Not_Computer | Not_Computer
Z-0750-2016 | 73027 | No | MDB | Microbial growth monitor. | Microbiology | Microbiology | 510(k) | BacT/ALERTÂż FA, bioMÂżrieux, Inc., |  |  | December 24, 2015 | 2015 | bioMerieux, Inc. | Bottles may have been exposed to non-normal shipping conditions which may lead to media degradation and impact growth performance. | Packaging process control | Biomerieux sent a Product Recall Notice letter to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to discontinue use of the referenced part and lot number and discard remaining inventory according to their facility's procedure for culture media. Customers were also instructed to contact their local bioMÂżrieux representative for product replacement. |  | 2741 | 2741 units in total | Distributed in the countries of Angola, Austria, Bosnia, Bulgaria, Cameroon, Czech Republic, Democratic Republic of Congo, France, Germany, Greece, Hungary, Italy, Kazakhstan, Kingdom of Saudi Arabia, Kosovo, Lebanon, Netherlands, Niger, Pakistan, Poland, Republic of Belarus, Romania, Russia, Serbia, Sri Lanka, Spain, Sweden, Switzerland, United Arab Emirates, Turkey and Zimbabwe. | Open, Classified |  |  | 2741 | Not_Computer | Not_Computer
Z-2046-2016 | 74208 | tray, impression, preformed | EHY | Preformed impression tray. | Dental | Dental | 510(K) Exempt | 6534-08N InterActiveÂż Implant Open-Tray Transfer Narrow |  |  | May 19, 2016 | 2016 | Implant Direct Sybron Manufacturing, LLC | The incorrect sized product, InterActive Implant Open-Tray Transfer Wide, 6534-08W, was packaged instead of the InterActive Implant Open-tray Transfer Narrow, 6534-08N, causing a possible issue with the impression. This discrepancy may lead to the possibility of an oversized crown fabrication. | Packaging | The firm sent out customer notificaiton letters on 05/19/16 via FedEx.Customers were asked to reivew the table to determine if they have any of the affected product in inventory. Customers were being informed that the proper product was being sent with the notificaiton. Customers were asked to complete and return the the Acknowledgement and Recall Return Form within 48 hours. If customers are an authorized Implant Direct Sybron Manufacturing distributor, the firm requests that they identify those customers that may have been shipped the affected product lot and contact these customers to inform them of this issue within fortyeight (48) hours of receipt of this notification in order to provide the customers with the correct tool. |  | 48 | 48 | Worldwide Distribution - U.S. Nationwide in the states of OR, NE, VA, FL, MO, WA, TX and the countries of CZ, GB, NL, MF, DK, IT, IE, DE | Open, Classified |  |  | 48 | Not_Computer | Not_Computer
Z-1083-2016 | 73293 | No | EKM | Gutta percha. | Dental | Dental | 510(K) Exempt | TF Adaptive Gutta Percha Point Size SM3 50pk, Part No. 815-1541 |  |  | February 08, 2016 | 2016 | Ormco/Sybronendo | Ormco Corporation is recalling the TF Adaptive Gutta Percha Point Size SM3-50pk because the inner generic label is identified as ML3. The products outer package is identified as SM3 which matches the product inside. | Error in labeling | An Urgent Medical Device Recall letter dated February 8, 2016 was sent to all customers. The letter identified the affected product, problem and actions to be taken. Customers are instructed to complete and fax back the enclosed acknowledgement form to 909-962-5605 to confirm receipt of the recall notification . |  | 5568 | 5,568 units total (2,070 units in US) | Worldwide Distribution - US Nationwide and the countries of Canada, Australia, Mexico, Europe, Hong Kong, MY, IN, JP, ID, SG, TH | Open, Classified |  |  | 5568 | Not_Computer | Not_Computer
Z-1075-2016 | 72979 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS 5,1 FS Chemistry System, Catalog Number 6801375, Unique Device Identifier Number 10758750001132; and VITROS 5,1 FS Chemistry System (Refurbished), Catalog Number 6801890, Unique Device Identifier Number 10758750001644; IVD. |  |  | December 21, 2015 | 2015 | Ortho-Clinical Diagnostics | Increased U90-382 or 6LU condition codes generated by VITROS 250, 350, 5,1 FS, 4600 and 5600 Chemistry Systems when using Calibrator Kit 9, Lot 954. A trend of complaints regarding customers actions following U90-382 or 6LU condition codes was noted (See RES 72289 - VITROS Calibrator Kit 9, lot 954 recall). | Error in labeling | Ortho Clinical Diagnostics sent an URGENT PRODUCT CORRECTION NOTIFICATION Letter (Ref. CL2015-201, dated December 21, 2015) via FedEx overnight courier and/or ORTHO PLUS e-Communications and/ or US Postal Service Priority Mail (only to those consignees who provided a PO Box as an address). The notification letter was sent to all customers who received the affected product. Foreign affiliates were informed by e-mail on 12/20/2015 of the issue and instructed to notify their consignees of the issue and required actions. The letter identified the affected product, problem and actions to be taken. For questions contact Customer Technical Services at 1-800-421-3311. |  | 1405 | US = 994; Foreign = 1405 | Worldwide Distribution - US Nationwide including Puerto Rico, and the countries of Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Singapore, Spain, and Venezuela | Open, Classified |  |  | 8387 | Not_Computer | Not_Computer
Z-2871-2016 | 75158 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | CADstream software. |  |  | November 01, 2012 | 2012 | Merge Healthcare, Inc. | An incorrect biopsy or missed target could result if the incorrect grid is selected within the application. | Error in labeling | Merge sent Urgent Medical Device Recall dated January 28, 2016 to affected customers. The letter identified the affected product, problem and actions to be taken. For questions Merge Customer Service at (877) 741-5369 or support@merge.com. |  | 1349 | 1,349 sites potentially have the affected versions | US Nationwide Distribution | Open, Classified |  |  | 1349 | Not_Computer | Computer
Z-2875-2016 | 74954 | No | EOQ | Bronchoscope (flexible or rigid) and accessories. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | Pentax Video Bronchoscope |  |  | May 31, 2016 | 2016 | Pentax of America Inc | Pentax America Inc. is recalling various bronchoscopes which contain an incorrect Operation and Reprocessing IFUS. | Packaging process control | Pentax Medical sent a U.S. Urgent Field Correction notice to all affected customers with a recall letter/return response form on May 31, 2016 via Fed Ex.Customers with questions were instructed to call 800-431-5880, ext. 2064. |  | 1719 | 1719 units | Nationwide Distribution | Open, Classified |  |  | 1719 | Not_Computer | Not_Computer
Z-1082-2016 | 73140 | No | LXT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Crosslock DVR ePAK -Sterile, Ref 81180050, are complete surgical sets including plates, screws, and instruments for orthopedic fixation |  |  | January 29, 2016 | 2016 | Biomet, Inc. | Components of the DVR ePAKs can oxidize during shipment and storage prior to use. A delay in surgery, less than 30 minutes, may occur if the surgical staff proceeds to re-sterilize the product. If oxidative material would flake off a component contained within the ePAK and enter the surgical field/wound an inflammatory reaction may occur. The inflammatory reaction may result in a foreign body r | Package design/selection | The firm, Zimmer Biomet, sent an "URGENT MEDICAL DEVICE RECALL NOTICE" dated 1/29/2016 to the affected customers. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. The customers were instructed to immediately locate and remove the identified device(s); carefully follow instructions on enclosed Response Form; complete and return Response Form within three (3) business days via email to CPWARFieldAction@zimmerbiomet.com prior to return of product, use priority carrier for your shipment; if you have further distributed the product, you MUST notify hospital personnel of this action. Questions related to this notice should be directed to (574) 372-1570, Monday through Friday, 8 a.m. to 5 p.m. |  | 21268 | 21,268 | Worldwide distribution: US (Nationwide) and countries of: Panama; Korea; Japan; Netherlands; Ecuador; Chile; Canada; and Australia. | Open, Classified |  |  | 21268 | Not_Computer | Not_Computer
Z-2048-2016 | 74276 | Yes | NIP |  |  | Cardiovascular | PMA | EV3 EverFiex" Self-expanding Peripheral Stent System, Sterilized using Ethylene Oxide, Rx Only, PRB35-08-040-080. |  |  | May 23, 2016 | 2016 | Ev3, Inc. | Medtronic is issuing a medical device recall to return product potentially associated with an outer boxmislabel issue for two Protege EverFiex" Self-expanding Peripheral Stent Systems, model PRB35-08-040-080, from lot A164580. The size on the outer box is labeled as 8x40x80mm while potentially containing a6x40x120mm device as indicated by both the inner pouch and device labels. This potential issue is limited totwo EverFiex stent system devices from lot A164580 that were repackaged by Medtronic and distributed.This does not affect any other Medtronic product or implantable devices. | Labeling mix-ups | Medtronic sent an "Urgent Medical Device Recall" letter dated May 2016, to all affected customers. The letter described the problem and the product involved in the recall. Consignees were instructed to identify, quarantine and return unused product to Medtronic. Consignees were requested to complete the Customer Confirmation Certificate and email it to RS.CFQFCAÂżMedtronic.com or fax it to Medtronic at 651-367-0612 to the attention of Customer Focused Quality. Customers with questions were instructed to call 651-367-0612.For question regarding this recall call 763-398-7000. |  | 2 | 2 | Nationwide Distribution to LA and WI. | Open, Classified |  |  | 2 | Not_Computer | Not_Computer
Z-1627-2016 | 73908 | No | DPW | Cardiovascular blood flowmeter. | Cardiovascular | Cardiovascular | 510(k) | GEM(TM), FLOWCOUPLER(R), 2.0 mm, STERILE EO, Rx only, 1 FLOWCOUPLER, REF GEM2752-FC |  |  | April 21, 2016 | 2016 | Baxter Healthcare Corp. | Instructions for use booklet may puncture the outer Tyvek lid. | Package design/selection | Baxter sent an Urgent Product Recall letter dated April 21, 2016 to customer. The letter identified the affected product, problem and actions to be taken. The letters requested that direct consignees contact Synovis MCA to arrange for return and credit. Companies that further distributed the product were asked to notify other facilities or departments. For questions contact Synovis MCA at 800-510-3318 or 205-941-0111. |  | 315 | 315 units | Nationwide and Germany, Sweden, Netherlands, Israel, Denmark, United Kingdom, Canada | Open, Classified |  |  | 2058 | Not_Computer | Not_Computer
Z-1041-2016 | 73254 | No | KKX | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Volcano R100 Sterile Equipment Cover; Sterile; |  |  | February 05, 2016 | 2016 | Volcano Corporation | Volcano Corporation has identified that certain lot numbers of Volcano R-100 Sterile Equipment Covers may have the incorrect expiry date on the outer shipping box; individually wrapped inner product bears the correct expiry date. | Labeling False and Misleading | Urgent Medical Device letters were sent to affected consignees the week of February 8, 2016 by Fed Ex or UPS so delivery can be tracked. For OUS consignees, the same letter will be sent and they will execute delivery to customers and track responses. |  | 1355 | 1355 total | Distributed US (nationwide) and the countries of Australia, South Africa, Canada, New Zealand, Belgium, Japan, and Paraguay. | Terminated | March 21, 2017 | 410 | 1355 | Not_Computer | Not_Computer
Z-0002-2017 | 75127 | No | MAV | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | Merit Inflation Syringe Kit. Catalog Number KOS-02857 |  |  | August 25, 2016 | 2016 | Merit Medical Systems, Inc. | Merit Medical Systems, Inc. announces a voluntary field action for Merit Inflation Syring Kit due to unit labels missing required device labeling information. | Labeling design | Merit Medical Inc. notified all affected customers on August 25, 2016, via verbal and written notification to return the affected units. The returned units will be re-labeled.For further questions, please call (801) 316-3690. |  | 1297 | 1,297 total units | One consignee located in CA. No Canadian, other foreign or VA/govt/military consignees. | Open, Classified |  |  | 1297 | Not_Computer | Not_Computer
Z-0584-2016 | 72719 | No | KWL | Hip joint femoral (hemi-hip) metallic cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | BIRMINGHAM HIP RESURFACING MODULAR HEAD, REF numbers: |  |  | November 16, 2015 | 2015 | Smith &amp; Nephew, Inc. | The 16 devices distributed in the United States were incorrectly labeled as the monoblock Birmingham Hip Modular Head, with labeling not cleared for marketing in the US. The actual devices were monoblock Modular Femoral Head for hemiarthroplasty use. | Labeling design | The firm sent their field safety notices on 11/16/2015 by letter delivered by Fed Ex (overnight) and email. Non- responding consignees will be contacted within 4 weeks of the initial notice |  | 16 | 16 units in US | Worldwide distribution. US nationwide, France, UAE, Austria, Australia, Belgium, CANADA, Switzerland, Germany, Denmark, Spain, Finland, UK, Hong Kong, ISRAEL, India, Italy, Japan, South Korea, Netherlands, Norway, Poland, PORTUGAL, Sweden, SINGAPORE, Thailand, Turkey, and SOUTH AFRICA. | Open, Classified |  |  | 16 | Not_Computer | Not_Computer
Z-1389-2016 | 73568 | No | MNL | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Cygnus Medical First StepBedside ADDWATER Kit,non-sterile. |  |  | March 17, 2016 | 2016 | Cygnus Medical | Foreign material in container | Packaging | Cygnus Medical LLC sent an Recall Notice letter certified mail to direct customers and receivers on March 17, 2016. The letter identified the product the problem and the action needed to be taken by the customer. The letter was preceded with an email or fax to expedite notification. Customers to respond via fax or email regarding the quantities in their possession, including a response where the quantity is zero. Returns to made to firm. Contact Quality Manager Cygnus Medical LLC, phone: (203) 488-4554 x 105 fax: (203) 764-1858.Email: qualitycontrol@madpoly.com.For further questions, please call (203) 488-4554. |  | 4680 | 39 cases of 120 units | US Distribution to the states of : CA , CO, FL, IL. MN, MT, NY, OH, PA, TX, VA, VT, WA and WI. | Open, Classified |  |  | 4680 | Not_Computer | Not_Computer
Z-1043-2016 | 72853 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Columbus Revision Knee System, EnduRo Knee System |  |  | November 30, 2015 | 2015 | Aesculap, Inc. | Aesculap Inc. US has initiated a recall on Tibial and Femur extension sterile pressfit stem implant packaging that are used for Columbus Revision/Enduro Knee Implants because it is labeled as Cementless. US product is indicated in the US for use with Bone Cement only. | Packaging change control | Aesculap sent an Urgent Medical Device Correction letter to Distributors and Sales Rep. The letter identified the affected product, problem and actions to be taken. Customers were notified to add the over label to product packaging. Telephone conferences with Distributors and Sales Rep reiterated need for representative to inform Surgical Handler of the label when providing product for review. Customers were instructed to complete the attached Product Correction Acknowledgement Form and return to Aesculap Quality Assurance department by faxing the form to 610-791-6882 or e-mail to val.strawn@aesculap.com, two (2) weeks of receipt, even if the total inventory in your possession is zero (0). |  | 1232 | 1232 | US Nationwide Distribution | Open, Classified |  |  | 1232 | Not_Computer | Not_Computer
Z-2888-2016 | 75160 | No | GWM | Intracranial pressure monitoring device. | Neurology | Neurology | 510(k) | Spiegelberg Brain Pressure Monitoring System |  |  | December 29, 2009 | 2009 | Spiegelberg Gmbh &amp; Co. KG | In the vast majority of products specified, a shrinkage of the air chamber during sterilization has occurred. Because of these false high readings were displayed in the lower measuring range from 0 to 20 mmHg.The observed measurement error was particularly at the lowest end of the range (to 0 mmHg) up to 8 mmHg. | Packaging | Spiegelberg Gmbh &amp; Co. KG sent an Urgent Medical Device Recall letter dated December 29, 2009, to all affected customers. Customers were instructed to quarantine and return stored products.Furthermore, they were asked to inform all their customers which have received the affected products and recall the products accordingly. Furthermore, they were asked to tell us about the actions taken by them. |  | 22 | 22 units | Nationwide Distribution to PA only | Completed |  |  | 22 | Not_Computer | Not_Computer
Z-1064-2016 | 73045 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Super Arrow-Flex Percutaneous Sheath Introducer Sets |  |  | January 12, 2016 | 2016 | Arrow International Inc | Arrow has issued a recall for these products due to potential packaging damage. The sterility of the products cannot be guaranteed. If non-sterile product is used there is a potential for infection to occur. | Packaging process control | Arrow International sent an Urgent Medical Device Recall Notification letter dated January 13, 2016, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately discontinue use and quarantine any affected product. Customers were asked to complete the enclosed Recall Acknowledgement Form and fax it to 1-855-419-8507 Attn: Customer Service or email it to recall@telefex.com. A customer service representative will contact them with a Return Goods Authorization (RGA) Number and will provide instructions for the return of product to the firm. If customers have no affected stock, they should also complete and the enclosed Recall Acknowledgement Form. Customers with questions should call Customer Service at 1-866-246-6990. For questions regarding this recall call 610-378-0131. |  | 56521 | 56,521 units | Nationwide Distribution including ALARAZCACOCTDEFLGAIAIDILINKYLAMAMDMIMNMOMSMTNCNDNENHNJNMNVNYOHOKORPARISCTNTXUTVAVTWAWIWV | Open, Classified |  |  | 56521 | Not_Computer | Not_Computer
Z-1040-2016 | 73172 | No | GDF | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | 50001 (Inrad) AccuPlace Drace Stereotaxic Needle Guide; Assists the user with inserting a needle at a predetermined angle for computer tomography. |  |  | February 04, 2016 | 2016 | Stryker Instruments Div. of Stryker Corporation | During an evaluation of the product packaging, it was determined that there is a potential for the product to cause damage to the packaging, creating a possible breach in sterility. | Package design/selection | Stryker sent URGENT MEDICAL DEVICE RECALL NOTIFICATIONS to Customers via Overnight service. Stryker Sales Representatives were notified via email.1. Immediately review this Recall Notification.2. Check all stock areas and/or operating room storage to determine how manyAccuPlace" Needle Guides are at your facility. Discontinue use of theAccuPlace" Needle Guides.3. Complete the enclosed Business Reply Form (BRF) to confirm receipt of thisnotification and identify how many, if any, affected items are currently in yourinventory. Please complete and return the BRF even if you dont have anyaffected product on hand.Note: Your signature on the BRF indicates that you received andunderstand this Notification and have followed the instructions in theNotification.4. If you have further distributed this product, please forward this letter and theattached Business Reply Form (BRF) to all affected locations. Pleaseindicate each location on the BRF.5. Fax the completed BRF to Stryker Instruments Regulatory Department at866-521-2762, or scan and email a copy to kara.spath@stryker.com.6. If the BRF for your facility indicates that recalled product is currently onhand, a FedEx label will be emailed to you. This shipping label should beused to return recalled product. Upon receipt of the recalled product, a credit will be issued to your account.In addition, Stryker Instruments sent recall notifications directly to facilities that received recalled product(s) originally shipped by Inrad. Facilities that received products from both Stryker and Inrad will receive one recall notification. The Business Reply Form will be pre-filled to indicate specifically what recalled products were sent to each facility, including the quantity and Product Number(s). |  | 5800 | 5,800 units total | Nationwide Distribution-including the states of AZ, CA, CO, CT, FL, GA, IA, KS, MA, MI, MN, MO, NY, OH, OK, PA, TX, VA, WA, and WI. | Terminated | May 09, 2017 | 460 | 5800 | Not_Computer | Not_Computer
Z-2918-2016 | 75050 | No | HTN | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Arthrex Suture Washer |  |  | August 12, 2016 | 2016 | Arthrex, Inc. | The device was assembled incorrectly. Suture and washer were not assembled; components were placed loose in the package. | Packaging process control | Consignees were notified 8/12/2016 of the potentially affected devices; Product/Batch Numbers: AR-7000-818T Batch #10028617. For Healthcare Professionals: immediately discontinue use of all affected products with stated numbers in their possession. Call (855) 552-5056 or email returns@arthrex.com with the product number, lot, and quantities so they may ship replacement products and expedite return of recalled product. To please complete the "Recall Acknowledgement of Receipt: postal card and return to Arthrex, Inc.Distributing Agencies: Immediately identify and hold all stated products in your inventory. Call 888-519-7255 ext. 72067 or e-mail AIM@arthrex.com Attn: Christi Zimmer. Have product number, lot number and quantities you may have so replacement product may be shipped and expedite the return of the recalled product. Please complete the "Recall Acknowledgement of Receipt" postal card and return to Arthrex. Inc. Any questions about the return of products to call 866-267-9138 and ask for Ellen Barry. Also by email to complaints@arthrez.com. |  | 45 | 45 devices | OR, NY, MN, KY, CA, ID, CO, MA, and FL.Thailand. | Open, Classified |  |  | 45 | Not_Computer | Not_Computer
Z-0610-2017 | 75499 | No | OLO | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | OMNIBotics Tracker Kit -Model: RM-10000 |  |  | October 14, 2016 | 2016 | OMNI LIFE SCIENCE | Incorrect lot number on outer kit | Error in labeling | The firm, OMNIlife Science, Inc., issued an "URGENT MEDICAL DEVICE ADVISORY NOTICE" dated 10/14/16 by email to its consignees/customers. U.S. and foreign agents were also sent the notification packet via overnight mail. Acknowledgement of receipt is requested. The notice described the product, problem and actions to be taken. The customers were instructed to immediately quarantine lot 19-09-021 for OMNIBotics Tracker Kit RM-1000; overlabel the outside packaging with the label provided following the provided training paperwork and rework instructions; note advisory notice only includes listed product, it does not affect other lots; and complete and return the Acknowledgement and Receipt Form and accompanying FN-102 by Fax to 508-519-3390.Accounts are to hold item RM-10000 with the lot # 19-09-021 until a corrected label is provided for the outer box. The patient labels inside the box are correct.For further information , Regulatory Affairs Specialist at 774-226-1847 or 800-448-6664 Monday through Friday, 8:00 AM to 4:30 PM, Easter Time. |  | 15 | 15 units | Worldwide Distribution: US distribution to OK, MD, TX and country of: Australia. | Open, Classified |  |  | 15 | Not_Computer | Not_Computer
Z-1392-2016 | 73552 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Oblique Posterior Atraumatic Lumbar (OPAL) System, Polymeric spinal fusion cage, non-sterile, OPAL Spacer 10 mm x 24 mm, 11 mm Height - Revolve |  |  | February 18, 2016 | 2016 | Synthes (USA) Products LLC | The label on the outer packaging does not match the product inside and intended use of the device. | Labeling mix-ups | DePuy Synthes sent an Urgent Notice letter dated February 18, 2016 to customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to review inventory, immediately remove the affected lots from their stock and complete the attached Verification Section. For questions call 610-719-5450 or contact your DePuy Synthes Sales Consultant |  | 6 | 6 units | Distributed in the US: OH and ND | Open, Classified |  |  | 6 | Not_Computer | Not_Computer
Z-2914-2016 | 75048 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | CT3Plus (a.k.a. CT3PZ); |  |  | September 19, 2016 | 2016 | New Star Lasers, Inc. | FDA inspection found that customers who had received devices with promotional material with outside use indications were not notified when labeling was revised to remove these indications. | Labeling False and Misleading | Letters sent on September 19, 2016 via US mail advised customers of the unapproved claims that may be in their literature and advised of the approved uses for each device. |  | 97 | 44 systems CT3Plus; 47 systems CT3PZ | US distribution only. | Open, Classified |  |  | 119 | Not_Computer | Not_Computer
Z-1054-2016 | 73329 | No | HDW | Contraceptive diaphragm and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Milex Arcing Diaphragm Size 95; Model Number: MXWS95 |  |  | February 16, 2016 | 2016 | CooperSurgical, Inc. | The products have been identified to contain an incorrect size Diaphragm. | Package design/selection | Cooper Surgical sent an "Urgent - Medical Device Recall" letter dated February 16, 2016, to all affected customers. The letter identified the problem, the product and the action needed to be taken by the customer.Customers were instructed to discontinue use of the affected products and complete the attached Acknowledgement and Receipt form for replacement.For further questions, please call (203) 601-9818. |  | 2 | 2 units | US Distribution to the states of : OR and NY | Terminated | March 29, 2016 | 42 | 2 | Not_Computer | Not_Computer
Z-2057-2016 | 73584 | container, specimen, non-sterile | NNI | Specimen transport and storage container. | Pathology | General Hospital | 510(K) Exempt | dynarex Specimen Containers, Screw-On, Leak-Resistant Lid, Tamper-Evident Seal, Sterile, Single Use Only |  |  | March 03, 2016 | 2016 | Dynarex Corporation | Exterior carton is mis-labeled as Sterile. Non-sterile, bulk specimen containers were packed in incorrect exterior carton. | Labeling False and Misleading | On March 2, 2016, dynarex Corporate Headquarters distributed Voluntary Device Recall notices and Recall Verification Forms to their customers via email. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to examine their stock immediately to determine if they have any of the lots listed on hand. If so, discontinue distributing and quarantine the product immediately. Promptly contact the Dynarex Credit &amp; Return representative Avi Celnick (845-365-8200 Ext 6644) to arrange for product return. Customers account will be credited for all returned products. All customers are advised to complete the Voluntary Recall Verification Form and returm immediately via fax or e-mail to Vijay Sachdev (Fax 845-365-8238; v.sachdev@dynarex.com). |  | 9975 | 399 cases/25 sleeves/20 cups &amp; lids per sleeve | US Nationwide in the states of AL, AR, FL, GA, MA, NC, NY, NJ, NH, PA, OH, TX &amp; VA; International: Canada | Open, Classified |  |  | 9975 | Not_Computer | Not_Computer
Z-1033-2016 | 73115 | No | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Covidien Kangaroo Polyurethane Feeding Tube Radiopaque Line, Safe Enteral Connections 6.5 Fr/Ch (2.2 mm) x 36" (91 cm), Sterile; Catalog Number: 461438. |  |  | January 21, 2016 | 2016 | Covidien LLC | 8FR 16 PVC tubes were packaged in a 6.5 FR 36 Polyurethane packs and 8FR 20 PVC tubes were packaged in an 8FR 20 Polyurethane packs | Labeling Change Control | Medtronic/Covidien Customers notified via an Urgent Medical Device Recall letter, dated January 22, 2016. The letter identifies the affected product, instructs customers to discontinue use and quarantine any of the product, and return all affected inventory. Customers are also to complete the "Recalled Product Return Form." Questions or concerns can be directed to a Medtronic representative or Customer Service at 800-882-5878. |  | 1010 | 1,010 units | Worldwide Distribution -- US, to the state of CA; and, to Australia and Europe. | Terminated | June 01, 2017 | 497 | 2340 | Not_Computer | Not_Computer
Z-1503-2016 | 73685 | No | IKD | Electrode cable. | Physical Medicine | Physical Medicine | 510(K) Exempt | Ad-Tech Electrode Connection System/Cables. Lightweight TECH-ATTACH Cable Models: |  |  | March 29, 2016 | 2016 | Ad-Tech Medical Instrument Corporation | Ad-Tech is recalling unopened/unused sterile cables intended to be used only with all compatible Ad-Tech electrodes for the purpose of connecting these electrodes to third party monitors/stimulators. This recall has been initiated due to package failures (i.e. compromised sterile barriers) that were identified during the execution of a shipping simulation investigation. | Package design/selection | Consignees were sent a "Medical Device Recall" letter dated March 29, 2016. The letter described the problem and the product being recalled. Advised consignees to immediately examine their inventory, quarantine and return the product to AD-Tech. If the product has been further distributed identify customers and notify them of the recall. Requested consignees to complete and return the "Medical Device Recall Return Response" form to their Ad-Tech Clinical Specialist at customersupport@adtechmedical.com; FAX: 262-634-5668 or Telephone 262-634-1555. For questions contact Ad-Tech Clinical Specialist.. |  | 197 | 197 | Distributed in the states of GA and MI, and the countries of FRANCE, UNITED KINGDOM, COLOMBIA, and BRAZIL. | Terminated | August 24, 2016 | 148 | 197 | Not_Computer | Not_Computer
Z-2834-2016 | 74825 | No | HQC | Phacofragmentation system. | Ophthalmic | Ophthalmic | 510(k) | LAUREATE World Phaco System |  |  | July 28, 2016 | 2016 | Alcon Research, Ltd. | Steps 16-25 are missing from the LAUREATE Multipak FMS Directions for Use (DFU) | Labeling Change Control | All remaining inventory of the identified affected products within Alcon DCs will be supplied corrected directions for use material prior to any further shipments. |  | 8646 | 8,646 UNITS | Burundi, China, Iran, South Korea, Taiwan, Thailand, USA. *** The recalled product is not sold in the USA or its territories. However, donated units are in use by a reduced number of USA clients for Medical Missions outside the country. | Terminated | June 07, 2017 | 314 | 8646 | Not_Computer | Computer
Z-0063-2017 | 75177 | No | DXN | Noninvasive blood pressure measurement system. | Cardiovascular | Cardiovascular | 510(k) | Blood Pressure Monitor UA-767PBT-Ci for small cuff. Designed to be used by end users who are eighteen (18) years and older at home to monitor their blood pressure (systolic and diastolic) and pulse rate. |  |  | September 12, 2016 | 2016 | A &amp; D Enginering Inc | A&D; may have shipped blood pressure monitors that were not labeled for use with small cuff. | Labeling Change Control | Notification letters were sent on September 12, 2016 to all affected customers. Letters identify the product and reason for the recall. Letters include a serial number list that is customer specific and ask that inventory be reviewed and that other organizations that may be impacted by this issue be notified. |  | 378 | 378 units | to distributors in NY, MA, VA. TX, | Open, Classified |  |  | 378 | Not_Computer | Not_Computer
Z-2218-2016 | 74449 | No | NBH | Arthroscope. | Orthopedic | Orthopedic | 510(K) Exempt | Smith and Nephew TWINFIX QUICK-T Knot Pusher Suture Cutter; Part Number: 7209417, for arthroscopic use. |  |  | June 15, 2016 | 2016 | Smith &amp; Nephew, Inc. | The device could be subject to breaches of its sterile packaging. | Package design/selection | Urgent Recall Notification Letters were sent via Federal Express on June 15, 2016, to the customers listed in the firm's distribution report. Customers were asked to inventory their devices and contact the firm for return authorizations. |  | 563 | 563 | Worldwide distribution. US nationwide, AT, AU, DE, FR, GB, JP, PT, and RU. | Open, Classified |  |  | 563 | Not_Computer | Not_Computer
Z-1582-2016 | 73676 | Yes | CAC | Autotransfusion apparatus. | Anesthesiology | Anesthesiology | 510(k) | Pleur-evac Chest Drainage System, Dry Suction/Dry Seal Control, REF A-6000-08LF, Rx Only, Teleflex Medical. |  |  | March 30, 2016 | 2016 | Teleflex Medical | The label on the Tyvek bag is missing. This is used by customers to identify material code, lot number, and expiration date once the product is removed from the shipper box. | Error in labeling | Teleflex sent an Urgent Medical Device Recall Notification dated March 30, 2016 to customers and distributors. The letter identified the affected product, problem and actions to be taken. The letter requested that they discontinue use and quarantine any products. The letter instructed customers to complete the enclosed Recall Acknowledgment Form and fax it to 1-855-419-8507, Attn: Customer Service or email to recalls@teleflex.com.. For questions contact your local sales representative or Customer Service at 1-866-246-6990. |  | 12124 | 12,124 ea. | Worldwide Distribution - US Nationwide in the states of: AZ, AR, CA, CO, CT, FL, GA, IL, IN, IA, KS, LA, MD, MA, MI, MS, MO. MT, NH, NJ, NM, NY, NC, OH, OK, OR, PA, RI, SC, TN, TX, UT, VT, VA, WV, Puerto Rico, Belgium, Canada, Chile, Dominican Republic &amp; Germany | Open, Classified |  |  | 12124 | Not_Computer | Not_Computer
Z-2791-2016 | 74727 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Supracondylar Nail T2 SCN 09x300 mm |  |  | June 30, 2016 | 2016 | Stryker Howmedica Osteonics Corp. | Stryker discovered that a Tibial Nail and femoral Nail were incorrectly labeled and distributed to the field. | Labeling mix-ups | The firm, Stryker, sent an "URGENT MEDICAL DEVICE RECALL" letter dated June 30, 2016 by e-mail to its customer requesting that they quarantine affected devices and sent the letter via UPS (with return receipt) on July 1, 2016. The letter described the product, problem and actions to be taken. The customer was instructed to inform users of the Medical Device Removal and forward the notice to all who need to be aware; inspect your inventory and return affected products to Stryker Orthopaedics/PFA Product Returns, Attn: Distribution Inventory Team; 325 Corporate Drive; Dock M-East , Mahwah, NJ 07431 REF: PFA #2016-094; complete and return the Business Reply Form via email to SO_M_PRODUCT_FIELD_ACTION_RESPONSE@stryker.com within 5 days; and keep a copy of the completed and executed Business Reply Form for your records. If you have any questions, feel free to contact the Mgr., Divisional Regulatory Compliance, Marketing Quality at (201) 831-6693. |  | 1 | 1 unit in total | US distribution to state of: CA. | Terminated | June 21, 2017 | 356 | 1 | Not_Computer | Not_Computer
Z-0663-2017 | 75617 | No | MJN | Vascular clamp. | Cardiovascular | Cardiovascular | 510(k) | Bridge Occlusion Balloon. Model: 590-001 |  |  | October 31, 2016 | 2016 | Spectranetics Corporation | Spectranetics Corporation announces a voluntary field action for the Turbo Elite Atherectomy Catheter, ELCA Coronary Atherectomy Catheter, and Bridge Occlusion Balloon due to potentially compromised integrity of the outer sterile packaging. | Packaging | Spectranectics sent an Urgent Medical Device Recall letter dated October 31, 2016, to all affected consignees. Consignees were instructed to discontinue use of the potentially impacted lots and return the product for replacement. Consignees were also instructed to complete the attached Acknowledgement and Receipt Form to facilitate the product return of any remaining product in inventory. Customers with questions were instructed to call 1-800-231-0978, Option 2. For questions regarding this recall call 719-651-8517. |  | 873 | 873 units | Worldwide Distribution - US (nationwide and PR) Internationally to Austria, Belgium, Cyprus, Denmark, France, Germany, Hungary, Italy, Netherlands, Norway, Poland, Spain, Sweden, Switzerland, United Kingdom, Australia, Bahrain, Brazil, Canada, Chile, Colombia, Dominican Republic, Israel, Japan, Kuwait, Taiwan, Thailand. | Terminated | April 18, 2017 | 169 | 37276 | Not_Computer | Not_Computer
Z-2793-2016 | 75125 | No | LLG | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(k) | Syngo Plaza picture archiving and communication system (PAC) (PACS) intended to display, process, read, report, communicate, distribute, store and archive digital medical images. |  |  | August 02, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Siemens is releasing a Customer Advisory Notice to explain meaning of a suffix o.p. that shows on some DICOM images on area, distance or perimeter values. This suffix stands for on projection and indicates that a shown measurement value may be different from the real value. | Error in labeling | Siemens mailed to customers on August 04, 2016 a Customer Advisory Notice that provided interpretations about the meaning of measurement values like area, distance and perimeter shown as "o.p." by the syngo.plaza. |  | 101 | 101 Systems | Nationwide Distribution | Terminated | March 20, 2017 | 230 | 101 | Not_Computer | Not_Computer
Z-0658-2017 | 75553 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | ConforMIS iTotal Cruciate Retaining (CR) Knee Replacement System, Sterile |  |  | October 18, 2016 | 2016 | ConforMIS, Inc. | Mislabeled: incorrect serial number on the F4 A/P cut guide iJig. | Labeling Change Control | ConforMIS notified the sales representative on 10/18/16 after it became aware of the device mix-up. The sales representative was requested to return any remaining components from iTotal CR kit S/N: 0381042. |  | 2 | 2 units | Distributed to the states of CO and NV. | Completed |  |  | 2 | Not_Computer | Not_Computer
Z-0786-2016 | 73022 | No | DSY | Vascular graft prosthesis. | Cardiovascular | Cardiovascular | 510(k) | Bard Distaflo Vascular Bypass Graft; Product Code: DF8006SC; 6 mm x 80 cm. |  |  | January 05, 2016 | 2016 | Bard Peripheral Vascular Inc | Size mislabeling; the graft is a 6 mm diameter graft when the actual packaged product is an 8 mm graft. | Error in labeling | Customer notification letters were sent on 01/05/2016. Customers were asked to not use the affected product, or further distribute it. Customers were asked to check all inventory locations for the affected product code, lot number combination listed. If customers further distributed any of the product they are requested to contact that location, advise them of the recall, and forward the listed instructions. The letter states to remove all identified product from the shelves. If no affected product is going to be returned the firm states to complete and return the attached form stating so. If customers have product it is requested that they also complete the attached form, and state the quantities and lot numbers in stock. The letter requests that customers call the firm's Recall Coordinator, Raye Seisinger, at 1-800-321-4254, option #2 ext. 2501 (M-F 6am to 3pm MST) or email raye.seisinger@crbard.com. Once all information has been verified the customer will be issued a Return Authorization or Consignment Recall Number to facilitate the expedient return of the product. The firm will provide replace product for returned product.Fax the attached form to the firm at 1-800-994-6772. If the customer cannot fax the firm they request that the Recall Coordinator be contact and provided the information verbally.The letter states a mailing label is enclosed to return the affected product. Customers are requested to amek the package as "recalled product" and include the XC or XH number. All products should be returned to the following shipping address:Bard Peripheral Vascular, Inc.1415 W. 3rd St.Tempe, AZ 85281 |  | 4 | 4 units | Distributed in the states of AL, TX, and CA. | Terminated | March 25, 2016 | 80 | 4 | Not_Computer | Not_Computer
Z-0655-2017 | 75558 | No | HHM | Obstetric-gynecologic specialized manual instrument. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(K) Exempt | SURESOUND22US Multipack (US Distribution) containing 2 packs of six devices. |  |  | October 28, 2016 | 2016 | Hologic, Inc | Breach in sterile barrier compromising product sterility | Package design/selection | Hologic notified customers in the United States through Certified letter on 10/28/16. The letter contains instructions with the option of returning or destroying the product, provides contact information, and includes a postage paid return card for the consignee to complete and return to Hologic.Questions or require clarification of the response process please contact 1- 800-442-9892 and selection option 6 and then option 2 for surgical product technical support.For questions regarding this recall call 508-263-8920. |  | 600 | 600 units (US) ; 24 units (OUS) | Worldwide Distribution - US (Nationwide) and Internationally to Finland and Austria. | Terminated | August 18, 2017 | 294 | 600 | Not_Computer | Not_Computer
Z-0590-2016 | 72750 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | NexGEN System REF 5980-37-01 Tibial Component Size Stemmed 3 Precoat, and REF 5986-37-01 Nonaugmentable Tibial Component Option CR/PS/LPS Size Stemmed 3. Tibial component of implanted knee prosthesis. |  |  | December 03, 2015 | 2015 | Zimmer Manufacturing B.V. | A product complaint was received stating that a carton labeled as a Precoat Stemmed Tibia Size 3 was opened for use and the contents were a non-augmented Stemmed Tibia Plate size 3. | Labeling mix-ups | Zimmer Biomet sent an "Urgent Medical Device Recall-Lot Specific" letter to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to do the following:Review the notification and ensure affected personnel are aware of the contents.2. Locate all affected product identified above and quarantine them immediately.3. Carry out a physical count of all affected product in your territory and complete the Inventory Return Certification Form (Attachment 1). Email a completed copy of Attachment 1 to corporatequality.postmarket@zimmer.com.4. Return the recalled product along with the completed Inventory Return Certification Form (Attachment 1). Clearly mark the outside carton of each product return shipment made as Recall. 5. Please notify Zimmer Biomet of any hospitals to which you have further distributed the affected product. Supply the information for any hospitals that you have identified, as well as the affected surgeons, using the provided spreadsheet template. The template will be emailed to you for completion. Please return to corporate quality.postmarket@zimmerbiomet.com.6. If after reviewing this notification you have further questions or concerns please call the customer call center at 1-800-348-2759 between 8:00 am and 8:00 pm EST. |  | 768 | 384 distribution events; 1 unit was returned leaving 383 units . | Worldwide Distribution to Germany, Belgium, Denmark, France, UK, Greece, Ireland, Netherlands, Norway, Russian Fed., Saudi Arabia, Sweden, and Thailand. | Open, Classified |  |  | 768 | Not_Computer | Not_Computer
Z-0767-2016 | 72210 | No | LHN | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MEVION S250 Radiation therapy system |  |  | November 18, 2015 | 2015 | Mevion Medical Systems, Inc. | The treatment table positioning belt can slip gear positions, resulting in deviation of patient setup and treatment positions from the time of slippage until the discovery of slippage and the halting of treatment. | Labeling design | Customer notification letter was sent via certified mail on November 18, 2015. The notification explains the issue and actions to be taken by the customer. |  | 4 | 4 | US nationwide distribution including FL, NJ, OK, and MO. | Terminated | April 03, 2017 | 502 | 4 | Not_Computer | Not_Computer
Z-1924-2016 | 74176 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Tearaway Introducer, Model # VS203, VS303, 510 K # 130687, packaged individually in a pouch, 5 pouches per carton, lot # MBZ140, MBZL450, MBZV930 MBZZ490 |  |  | May 05, 2016 | 2016 | Medical Components, Inc dba MedComp | This recall has been initiated due to the product labeled with the incorrect expiration date. | Labeling mix-ups | MedComp sent a Product Recall letter dated May 5, 2016 to customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to evaluate their inventory and quarantine for the return of all un-used affected product. They were asked to contact Customer Service for a Return Goods Authorization (RGA) number at 215-256-9191. |  | 3122 | VS203 Lot# MBZX140 (100 units); VS303 Lot # MBZL450 (60 units); VS303 Lot # MBZV930 (65 units); VS303 Lot # MBZZ490 (33 units). | US Distributed to: FL, TX, NC | Terminated | June 16, 2017 | 407 | 3122 | Not_Computer | Not_Computer
Z-1647-2016 | 73413 | No | LTW | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | Vitek 2 Gram Positive Susceptibility Test Cards (AST-P646 REF 420144), 20 cards/kit. |  |  | February 24, 2016 | 2016 | Biomerieux Inc | Product insert fails to identify performance limitation related to EUCAST breakpoints for Teicolplanin. | Labeling Change Control | Biomerieux sent an Urgent Product Correction Notice letter via email on February 24, 2016 to affected customers . The letter identified the affected product, problem and actions to be taken. Customers were asked to complete the attached Acknowledgement Form (Attachment A) and return it to Biomerieux, Inc. For questions contact your local bioMÂżrieux representative. |  | 1035 | 1,035 kits | Foreign distribution in the countries of Czech Republic, Hong Kong (China), Hungary, Japan, Poland, South Africa, Switzerland, and Turkey. | Open, Classified |  |  | 37476 | Not_Computer | Computer
Z-1587-2016 | 73486 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | The Livongo Health In TouchÂż Blood Glucose Monitoring System (BGMS), is an Over-The-Counter (OTC) system designed for the self- monitoring of blood glucose by persons with diabetes in home settings to aid in their diabetes management. |  |  | February 18, 2016 | 2016 | Livongo Health | The instructional insert for the In Touch Control Solution for use with the In Touch meter and In Touch strips as a quality control check did not contain the limitation of sensor accuracy: (1) with the use of acetaminophen (Tylenol)(2) with the use of Ascorbic Acid (vitamin C) supplementation(3) in the presence of uric acid | Error in labeling | The recall was conducted in three distinct events. Members were first sent an e-mail of the Recall Notification - Livongo Test Strip Insert (Event 1) containing information regarding the test strip insert discrepancy. Following the e-mail distribution, a hard copy of the same information was distributed (Event 2) to members via United States Postal\ Service (USPS) mail. Lastly, a second e-mail (Event 3) containing the same information was sent to members. The distribution list for the second (hardcopy distribution) and third (e-mail distribution) events was a subset of the initial member distribution. |  | 63000 | 63,000 units | US consignees: Arizona, California, Delaware, Florida, Georgia, Illinois, Kentucky, Maryland, Michigan, Minnesota, Mississippi, Missouri, Montana, New Jersey, New York, Ohio, Pennsylvania, South Carolina, Tennessee, Texas, Virginia and Washington. | Open, Classified |  |  | 63000 | Not_Computer | Not_Computer
Z-0584-2017 | 75485 | No | GXD | Radiofrequency lesion generator. | Neurology | Neurology | 510(k) | RF MultiGenÂż, 0406-900-000; Refurb RF MultiGenÂż, 0406-900-000U |  |  | October 11, 2016 | 2016 | Stryker Instruments Div. of Stryker Corporation | Stryker Instruments is voluntarily recalling the Care Instructions/Instructions for Use packaged with the Radiofrequency (RF) MultiGenÂż Generator, because the instructions for use incorrectly contained instructions supporting the sterilization of the MultiGenÂż cables. | Labeling Change Control | Stryker Instruments initiated a voluntary recall of the RF MultiGen, and Refurb RF MultiGen because the instructions for use contained incorrect instructions on sterilizing the device. Stryker distributed customer notices via certified mail on 10/11/2016. Customer's were instructed to do the following: Immediately review this recall notification.2. Enclosed was a revised version of the IVS Care Instructions. Please distribute the revised version to the appropriate person or department.3. Remove and discard any previously distributed IVS Care Instructions or Instructions for Use.4. Please complete, sign, and return the enclosed Business Reply Form (BRF) acknowledging your receipt and understanding of this notification.5. Fax the completed BRF to Stryker Instruments Regulatory Department, 866-521-2762, or scan and email a copy to kellyjo.davis@stryker.com. Customers with questions can contact: 269 323 7700 . |  | 2675 | 2,675 | Domestic:AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, WV, WYForeign:France, Germany, Canada, UK, Lebanon, South Africa, Chile, Hong Kong, Poland, Argentina, Spain, Greece, Italy, India, Brazil, Switzerland, Netherlands, SwedenVA/DOD: | Terminated | July 25, 2017 | 287 | 2675 | Not_Computer | Not_Computer
Z-2835-2016 | 74849 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | FPS 35mm x 3.5 mm non-locking screw |  |  | July 21, 2016 | 2016 | Ortho Solutions Inc | Mislabeling of a single batch of FPS 3.5 x 35 mm Locking Screw as being Non-Locking Screws. | Labeling mix-ups | Ortho Solutions sent a "Field Safety Notice" letter dated July 21, 2016 to their affected customers.The letter identify the Solutions Business Development Manager will contact the customer to arrange for the isolation and removal of the unused product from their inventory. Replacement product will be provided. The customer was advised to passed on the "Field Safety Notice" to all who need to be aware within their organization and to any organization to which their have further distributed or transferred the affected screw. Upon removal of the affected device, customers are asked to please sign and return the "Field Safety Corrective Action Declaration Form", by which to confirm they have received the "Field Safety Notice" and that they declare that the required recall has been completed at their hospital. For further questions, call (201) 258-7305. |  | 15 | 15 | US Distribution | Open, Classified |  |  | 15 | Not_Computer | Not_Computer
Z-2843-2016 | 74928 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | 3.5mm Bowed Locking Compression Femur Plates, 14 hole |  |  | July 27, 2016 | 2016 | OrthoPediatrics Corp | OrthoPediatrics Corp.is voluntarily recalling multiple lots of the 3.5mm Bowed Locking Compression Femur 14-hole Plate and 18-hole Plate due to the devices being mislabeled. | Labeling mix-ups | Orthopediatrics initiated the recall of 3.5mm Bowed Locking Compression Femur Plates by distributing e-mails on 0July 27, 2016, followed by telephone calls to all customers on July 28, 2016. All distributors and direct representatives were sent an e-mail to provide them with a written notification of the voluntary recall and instruct them to quarantine all identified devices from the affected lot numbers immediately. Customers were instructed to quarantine recalled product and response via a hard copy Reply Form that was included in a hard copy of the Recall Notice. Each sales representative was also asked to contact Orthopediatrics Logistics Department to receive a Return Authorization Number for returning recalled product.in order to track their affected plates. If customers have any questions about this recall, they were instructed to reach out to the firm via phone (office: 574-267-0865) or email (acargill@orthopediatrics.com). |  | 41 | 41 | Nationwide Distribution to AZ, CA, CO, DE, FL, GA, KY, MI, MN, MO, KY, RI, NY, OH, OR, RI, TN, TX, VA, WAForeign: NoneVA/DOD: None | Terminated | March 24, 2017 | 240 | 41 | Not_Computer | Not_Computer
Z-2845-2016 | 74896 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | 3I OSSEOTITE CERTAIN DENTAL IMPLANTS, 4x11.5 mm, Sterile, Rx only. |  |  | August 08, 2016 | 2016 | Biomet 3i, LLC | Product was incorrectly packaged. | Packaging process control | An Urgent Medical Device Removal Notice is being sent to Customers August 12, 2016, notifying them of the packaging of OSSEOTITE Certain Implant (4x11.5mm) in T3 Implant boxes. This recall is specific to Lot No. 2016031461. The material shipped within the box was described as being correct ,but only the box it was shipped in was incorrect. Customers were notified that there are no immediate or long term health consequences. In the notification letter, customers are instructed to complete the Certificate of Acknowledgement, and to return it via email (postmarket@zimmerbiomet.com) if no product is on hand or to return the certificate with product by shipping materials to: Biomet 3i - Post-Market Returns. Replacement product will be sent to customers upon return. Customers were also advised to keep a copy of their Certificate of Acknowledgement. Further questions should be directed to 1-800 443-8166 between the hours of 8:00 AM and 6:00 PM, Monday thru Friday. For further questions, please call 1-800 342-5454. |  | 92 | 92 units | US Distributions to the states of : CA and FL., and Worldwide to : Colombia, France, Germany, Italy, and Spain. | Open, Classified |  |  | 92 | Not_Computer | Not_Computer
Z-0758-2016 | 73098 | No | MCW | Intraluminal artery stripper. | Cardiovascular | Cardiovascular | 510(k) | 2.3mm Turbo-Power OTW, Manual, Model No 423-050 |  |  | January 15, 2016 | 2016 | Spectranetics Corp. | The product is labeled with an expiration date that is past its shelf life. | Error in labeling | Spectranectics sent a Medical Device Field Action letter dated January 15, 2016 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to locate and set aside the affected product and contact Customer Service at 1-800-231-0978 to arrange a product exchange. Customers were asked to document the return on the enclosed Field Action Acknowledgement and Receipt Form. Customers with additional questions were instructed to contact their local Spectranetics sales representative.For questions regarding this recall call 719-447-2318. |  | 8 | 8 units | Nationwide distribution to Florida, North Carolina and Pennsylvania. | Open, Classified |  |  | 8 | Not_Computer | Not_Computer
Z-2850-2016 | 75094 | No | MQL | Reproductive media and supplements. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Sage In-Vitro Fertilization- CSC (Choline Substituted Cryopreservation) Freezing Medium . Product Number: ART -8017 |  |  | August 25, 2016 | 2016 | CooperSurgical, Inc. | Products contain international package insert and notthe correct USA package insert | Packaging change control | CooperSurgical sent an Advisory Notice dated August 25, 2016, to affected customers via UPS with confirmed delivery receipt. The letter identified the affected product and detailed the necessary actions to be taken by customers who carried the affected lots in their inventory. The advisory notification asked customers to replace the previously inserted incorrect Instructions for use (IFU) with the correct IFU sent with the advisory notification.For questions regarding this recall call 203-601-5200. |  | 1 | 1 unit | Nationwide Distribution to CA, GA, and NY | Open, Classified |  |  | 5 | Not_Computer | Not_Computer
Z-0582-2017 | 75347 | No | GBZ | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Atrium Medical 28 Fr Trocar catheter, Sterile |  |  | October 26, 2016 | 2016 | Atrium Medical Corporation | Trocar Catheters are packaged with a label containing an icon that shows more holes (eyelets) than the product actually contains. | Labeling design | Atrium Medical issued an "Urgent Medical Device Recall (Removal)" communication to customers on 10/26/16 via FedEx 2 day delivery. The recall identifies the problem, product, and risk factors. In addition, it provides instructions on how to handle product that may be affected. Accounts are to read the letter completely and share with all users and staff. They have been requested to complete the enclosed Recall Reply Form, acknowledging receipt of the notice; and, remove and quarantine any affected product for return. Additionally, obtain product from an alternate supplier because replacement product is not available as all Trocar catheters have been discontinued. Contact Atrium Medical Customer Service at 1-800-370-7899, Monday through Friday between 9:00 am to 5:00 pm, for a Return Goods Authorization to return the product and receive credit. |  | 7043 | 7043 cases (10/cs) | Worldwide Distribution -- USA, Australia, Austria, Belgium, Canada, Costa Rica, Dominican Republic, Ecuador, Finland, France, Germany, Honduras, Hong Kong, Ireland, Jamaica, Kuwait, Mexico, Netherlands, New Zealand, Panama, Peru, Portugal, Puerto Rico, Qatar, Reunion, Romania, Saudi Arabia, South Africa, Spain, Sweden, Switzerland, and Venezuela. | Open, Classified |  |  | 31247 | Not_Computer | Not_Computer
Z-2869-2016 | 74989 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Sertera Biopsy Device, Sterile; Used primarily in ultrasound-guided breast biopsies. |  |  | August 22, 2016 | 2016 | Hologic, Inc | Incorrect label on the Sertera Biopsy Kit | Labeling Change Control | Hologic notified customers in the United States through Certified letter on 8/22/16. The letter contains instructions of returning the product, provides contact information, and includes a postage-paid return card for the consignee to complete and return to Hologic. Contact technical support at 508-263-2900 or l-800-442:9892 Press Option 6 ). |  | 570 | 570 units | Nationwide. | Open, Classified |  |  | 570 | Not_Computer | Not_Computer
Z-1870-2016 | 73743 | No | LRS | Intravascular catheter. | General Hospital | General &amp; Plastic Surgery | Enforcement Discretion | General Surgery Pack, part number 006880-3. This is a custom convenience surgical instrument disposable kit. |  |  | March 21, 2016 | 2016 | Windstone Medical Packaging, Inc. | These packs contain item #4037 Kerlix X-ray Detectable Laparotomy Sponges which were recalled by Medtronic due to the potential for sterile barrier to be comprised at the packaging seal. | Packaging | Aligned Medical Solutions (AMS) sent the Urgent Medical Device Recall letter, dated March 21, 2016, and the Recall Reply Form to their consignees via email.AMS informed consignees that their AMS packs (surgical kits) contained Kerlix X-Ray Detectable Laparotomy Sponge which was recalled by Medtronic.Consignees were advised to follow the over-labeling instructions and dispose the Kerlix X-Ray Detectable Laparotomy Sponge at the time the surgical kit is opened for use. Consigneeswere instructed to follow the "Instruction for placing Recall Notice Label to Packs" and label any kits remaining inventory.Consignees should return the Reply Form via fax at 407-865-7456 or call 321-527-7714 for any questions about this recall. |  | 170 | 170 units | Nationwide distribution to Louisiana and Montana. | Open, Classified |  |  | 1192 | Not_Computer | Not_Computer
Z-1672-2016 | 74096 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Visualase Cooled Laser Applicator System (VCLAS). The VCLAS system consists of the Visualase Laser Diffusing Fiber (LDF) and the Visualase Cooled Catheter System (CCS). |  |  | May 09, 2016 | 2016 | Medtronic Navigation, Inc. | Potential for the Visualase Cooled Laser Applicator System (VCLAS) catheters to melt during procedures. | Labeling design | Medtronic will be sending a Medical Device Safety Notification letter FedEx on May 18, 2016. Each customer will be asked to complete and return a response form to Medtronic. This is a safety alert notification and does not require return of devices.We are committed to patient safety and appreciate your prompt attention to this matter. If you have any questions regarding this action, please contact Medtronic Technical Services at 1-800-595-9709. |  | 120 | 120 | Distributed within the US, Netherlands, Canada and Australia and to govt/VA/military consignees. | Open, Classified |  |  | 120 | Not_Computer | Not_Computer
Z-1523-2016 | 73752 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | Mach 4 Mouse Probe vial within the Mach 4 Universal HRP-Polymer Kit; |  |  | March 31, 2016 | 2016 | Biocare Medical, LLC | Biocare Medical has identified a single lot (082715) of its Mach 4 HRP Polymer kit in which some few Mach 4 Mouse Probe vials may have been mislabeled as HRP-Polymer. | Labeling mix-ups | Biocare Medical sent a Medical Device Recall Notification letter dated March 31, 2016 to affected US customers by USPS. Foreign consignees were also notified per their various government regulations. The letter identified the affected product, problem and actions to be taken. Customers were requested to complete and return by April 15, 2016 the enclosed Return Response Form. |  | 93 | 93 | US Nationwide Distribution in the states of: KS, CA, AZ, OR, FL, KS, CO, MA, KY, NY, IL, OR, MO, NC, TX, PA, OK. SC, TN, PA, TN, GA, and OR.Hungary, Canada, Boliva | Open, Classified |  |  | 93 | Not_Computer | Not_Computer
Z-0764-2017 | 75735 | No | GDT | Removable skin staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | WECK Visistat 35W, Disposable Skin Stapler, Rx Only, Sterile, Teleflex Medical, NC 27709. |  |  | November 18, 2016 | 2016 | Teleflex Medical | Sterility Deficiency: The packages are not sealed properly and thus the sterility is not assured. | Packaging process control | The firm, Telexflex Medical, sent an "URGENT MEDICAL DEVICE RECALL NOTIFICATION" letter dated 11/18/2016 to its Consignees/Customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately discontinue use and quarantine any affected products; complete and return the Recall Acknowledgement Form by fax to 1-855-419-8507, Attn: Customer Service or email to: recalls@teleflex.com (even if you do not have any affected product) and return the product with the completed Recall Acknowledgement Form. A customer service representative will contact you with a RGA # and provide instructions for the return of the product.If you have any other questions, feel free to contact your local salesrepresentative or Customer Service at 1-866-246-6990. |  | 96402 | 96,402 units (9,120 ea US, 87,282 ea OUS) | Worldwide Distribution- U.S. (Nationwide) and countries of: Australia, Belgium, Canada, China, Germany, Hong Kong, India, Japan, Korea, Malaysia, New Zealand, Singapore and Thailand. | Open, Classified |  |  | 96402 | Not_Computer | Not_Computer
Z-0600-2016 | 72753 | No | OMP | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ActiV.A.C. 300ml Canister (with Gel) for ActiV.A.C. Therapy System for the application of V.A.C. Therapy in the acute extended and home care settings. The software-controlled therapy unit applies negative pressure wound therapy. The dressing, to which the therapy unit is connected, enables distribution of the negative pressure across the surface of the wound, while the tubing transfers accumulated fluids to the canister. |  |  | November 25, 2015 | 2015 | KCI USA, INC. | There is potential for small holes, the size of a pin, to develop in the pouches of the ActiV.A.C. Sterile Canisters. | Package design/selection | Customer notification letters were sent out the week of November, 23, 2015. The international distributors will be responsible for notification to their sub-consignees. |  | 12728 | 12,728 units | Japan and Australia | Open, Classified |  |  | 12728 | Not_Computer | Not_Computer
Z-0812-2016 | 72977 | syringe, irrigating (non dental) | KYZ | Irrigating syringe. | General Hospital | General Hospital | 510(K) Exempt | REF 0038460 Toomey Syringe 70cc Syringe with Catheter Tip and Luer Tip Adapters, Single Use, Rx Only, Latex Free, 50 units. |  |  | January 04, 2016 | 2016 | C.R. Bard, Inc. | Report by a customer of an open package seal. | Packaging process control | BMD sent an Urgent Medical Product Recall letter to each of the customers via FedEx with proof of delivery notice. The leter identified the affected product, problem and actions to be taken. All consignees were instructed to carry out the recall to the user level. The letter included a Recall and Effectiveness Check Form which was to be returned to Bard. |  | 15400 | 15,400 units | US Nationwide Distribution - in the states of: AL, AZ, CA, CO, FL, GA, IL, IN, IA, KS, KY, LA, MD, MI, MS, MO, NE, NV, NJ, NM, NY, NC, OH, OK, PA, SC, TN, TX, VA, WA, WV, and Puerto Rico | Terminated | June 06, 2017 | 519 | 15400 | Not_Computer | Not_Computer
Z-2148-2016 | 74030 | [6 empty] | Breathe Right Clear, 30 ct. carton., UPC 7-57145-00246-7 (US), UPC 9-300673-871749 (AU/NZ) |  |  | April 25, 2016 | 2016 | GSK Consumer Healthcare | Some cartons used in the secondary packaging of Breathe Right Clear, Large, 30 ct. Nasal Strips were incorrectly labeled for cartons intended for the AU/NZ market. | Error in labeling | GSK Consumer Healthcare sent an Urgent - Product Recall Letter dated April 4, 2016, to all affected customers and an updated recall letter dated April 25, 2016, via FedEx.The recall letters identified the product and problem. The April 25, 2016, was sent to further clarify, the lot number on the Breathe Right carton. The customers were asked to immediately examine their inventory, quarantine and cease distribution of the specific lot of the affected product. Customers were asked to carry out a physical count and record the data on the response form. The response form should be returned to GSK within five (5) business days, even if you do not have the recalled lot. The response form should be faxed to (336) 499-8795. If customer needs shipping assistance or questions about the recall process, they can call 1-844-722-2104, and refer to Event Number 535STRIPS16. If any customer has questions, they can direct them to GSK Consumer Healthcare Consumer Relations team at 800-858-6673. |  | 1440 | 1,440 cases ( 34, 560 retail - units ) | US Nationwide | Terminated | July 19, 2017 | 450 | 1440 | Not_Computer | Not_Computer
Z-1320-2016 | 73497 | No | OHG | Obstetric-gynecologic specialized manual instrument. | Obstetrics/Gynecology | Obstetrics/Gynecology | Enforcement Discretion | Gomco Style Circumcision Clamp Trays is an instrument used to compress the foreskin of the penis during circumcision of a male infant. |  |  | February 29, 2016 | 2016 | Cardinal Health | Two lots of Gomco Style Circumcision Clamp Trays, SSI-0034 and 10-1100 contained a 1.3cm circumcision clamp. They both were labelled as containing a 1.1cm clamp. | Labeling Change Control | Cardinal Health sent an URGENT: PRODUCT RECALL letter dated February 29, 2016, to affected consignees explaining the reason for the product recall and actions that should be taken by the end user. These included the following actions:1. STOP using affected product.2. RETURN the enclosed acknowledgment form via fax to 847-689-9101 or 614-652- 9648, whether or not you have affected product, as Cardinal Health is required to confirm receipt of this notification from you, and to prevent further notices.3. NOTIFY any customers to whom you may have distributed product affected by this recall.4. CONTACT the appropriate Customer Service group to arrange for return of any affected product you may have:" Hospital800.964.5227" Distributor800.635.6021 The firm asks that consignees who have question or requests of special assistance relating to the recall should contact Cardinal Health Customer Advocacy at 800.292.9332.For questions regarding this recall call 847-473-1500. |  | 440 | SSI-0034: 380 units and 10-1100: 60 units | Nationwide Distribution including AR, KS, WV, , NC, TX, AL, MS and VA. | Open, Classified |  |  | 440 | Not_Computer | Not_Computer
Z-1608-2016 | 73318 | No | GAD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | TLCÂż Self Retaining Retractor System Contents 1-RETRACTOR FRAME 6-ELASTIC STAYS STERILE EO Rx ONLY |  |  | February 17, 2016 | 2016 | Applied Medical Technology Inc | The firm received a complaint of an open pouch which breaches the sterilization validation. | Packaging | AMT sent an Urgent Medical Device Recall letter dated February 17, 2016 to their customers. The letter identified the affected product, problem and actions to be taken. For questions contact Customer Service at 440-717-4000. |  | 5334 | 5,334 units | Product was shipped to the following states: AL, AR, AZ, CA, CT, FL, GA, IA, IN, KS, KY, MA, MD, MI, MN, MO, NC, NE, NH, NJ, OH, PA, SC, TN,TX, UT, VA &amp; WI.Product was also shipped to the following countries: Austria, Denmark, Germany, Italy, Japan, Norway, Spain, Sweden, Switzerland &amp; United Kingdom. | Open, Classified |  |  | 12451 | Not_Computer | Not_Computer
Z-1938-2016 | 74034 | Yes | BTR | Tracheal tube. | Anesthesiology | Anesthesiology | 510(k) | Sterile Sheridan/HVT Cuffed, Murphy Eye, Tracheal Tube, Ref: V5-10314, ID: 7.0 mm, Rx Only, Manufactured by: Teleflex Incorporated. |  |  | May 05, 2016 | 2016 | Teleflex Medical | The labeling of some units indicates that it contains a size 7mm tracheal tube, but may contain a size 7.5mm tracheal tube. | Error in labeling | The firm, Teleflex Medical, sent an "Urgent Medical Device Recall Notification" letter dated May 5, 2016 to its accounts/customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately discontinue distribution and quarantine product; communicate this recall to any of your customers who have received product; have your customers return to you any affected product with Recall Acknowledgement Form; contact Teleflex Medical Customer Service for more information at 1-866-246-6990; and complete and return Recall Acknowledgment Form via fax to 1-855-419-850, Attn: Customer Service or email to recalls@teleflex.com, even if you and your customers have no affected stock/product. Teleflex will provide you with instructions for the return of product. If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-866-246-6990. |  | 6390 | 6,390 ea. | US Distribution to states of: AL, AZ, CO, FL, GA, ID, IL, IN, IA, KY, LA, MD, MA, MI, MN, MO, NY, NC, OK, PA, RI, TX and UT. | Open, Classified |  |  | 6390 | Not_Computer | Not_Computer
Z-1957-2016 | 74156 | No | MUU | Suction lipoplasty system. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | The Red Head 2 Reusable Fat Harvesting Canister, Product # RH-2, Non-Sterile. |  |  | April 21, 2016 | 2016 | Miami Fat Supply, Inc | During an FDA inspection it was found that the products are marketed without a cleared 510k, nor is approved for multiple use. | No Marketing Application | Customers were sent an Urgent: Medical Device Recall/Correction letter, dated 04/19/2016, on 04/21/2016 that served as a notification of the Red Head 2 voluntary recall. Customers were informed that there were previous complaints of compromised product integrity and performance following sterilization. Lids are subject to fracture and fail after steam sterilization. Customers were advised to destroy used product and to ship unused product back to the manufacturer. Customers were to return an acknowledgement form to confirm their notification of the voluntary recall. |  | 443 | 443 units | Worldwide Distribution -- US, including the states of AL, AZ, CA, CO, FL, GA, IL, MD, MI, MO, NM, NY, OH, TX, UT, and VA; and, the country of Dominican Republic. | Open, Classified |  |  | 443 | Not_Computer | Not_Computer
Z-2158-2016 | 74283 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | 3.5 mm and 4.0 mm Cannulated Screw, Product code HWC, Device Listing Number D029094, Screw, Fixation Bone |  |  | February 16, 2016 | 2016 | Synthes (USA) Products LLC | Labeling does not match the cleared indications for use in the United States and Canada. | Labeling False and Misleading | DePuy Synthes initiated a Field Safety Notice to remove the affected Technique Guides and marketing literature from the field. Sales Consultants and Synthes Trauma Marketing will receive training on affected technique guides and marketing literature which promoted off label use. All technique guides and literature have been put on hold and removed from DePuy Synthes website. Effectiveness of this field action will be tracked using their training system. They will confirm training when completed as a means to measure effectiveness. |  | 19497844 | 19,497,844 units in total | US Nationwide Distribution in the states of :AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, N/A, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY. | Terminated | June 01, 2017 | 471 | 19497844 | Not_Computer | Not_Computer
Z-2470-2016 | 73184 | No | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | Implant EP Pick-Up Coping |  |  | November 25, 2015 | 2015 | Biomet 3i, LLC | Pouches may not have been sealed during packing. | Packaging | Biomet sent an Urgent Device Recall Notice to all affected consignees the week of November 30, 2015, communicating the issue and customer responsibilities. International consignees were notified via traceable courier or email week of November 30, 2015. Customers were instructed to return affected product by calling BIOMET at 1-800-342-5454 to obtain a RMA number to return products. Customers were also asked to complete the Business Reply Form and fax to 561-514-6316.For questions regarding this recall call 561-776-6700. |  | 1 | 1, 648,273 devices (all products subject to this recall) | Worldwide Distribution Nationwide andHawaii, Guam, Puerto Rico and Virgin Islands (U.S.) .FOREIGN CONSIGNEES: Albania, Andorra, Argentina, Australia, Austria, Barbados, Belgium, Bolivia, Brazil, Canada, Cayman Islands, Chile, China, Colombia, Costa Rica, Denmark, Dominican Republic, France, Germany, Great Britain, Greece, Hong Kong, Hungary, India, Ireland, Israel, Italy, Jamaica, Japan, Korea (South), Kuwait, Lebanon, Liechtenstein, Lithuania, Martinique, Mexico, Mongolia, Netherlands, Northern Mariana Islands, Paraguay, Peru, Poland, Portugal, Qatar, Reunion, Russian Federation, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Syria, Taiwan, Thailand, Trinidad and Tabago, Turkey, United Arab Emirates, Uruguay and Viet Nam. | Open, Classified |  |  | 1648274 | Not_Computer | Not_Computer
Z-0879-2017 | 75758 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | OrthoPediatrics PediNail Intramedullary Nailing System |  |  | November 18, 2016 | 2016 | OrthoPediatrics Corp | Labeled as a 48mm Cortical Screw but measuring at 50mm | Labeling mix-ups | OrthoPediatrics sent an URGENT MEDICAL DEVICE RECALL notification on November 18, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. . It is important that you carefully follow these recall instructions:" Please immediately take possession of any of these screws your customer may have or that you have, and quarantine them. Contact logistics@orthopediatrics.com to receive a Return Authorization Number. It is very important that we contain these devices rapidly and maintain them in quarantine. Replacement devices can be ordered from Customer Service." The FedEx package you receive will contain a pre-addressed FedEx return label for the return of the recalled screw(s)." Included in this FedEx package is a Mandatory Reply Form for you to indicate the total number of affected screw(s) in your sets.We will be providing any additional follow up instructions as needed. If you have any questions call (574) 267-0865.For further questions regarding this recall , please call (574) 268-6379 |  | 22 | 22 units | Worldwide Distribution - US Distribution to the states of : TX, MO, FL, GA, NC, AZ, NY., and to the countries of France and Ireland. | Terminated | July 27, 2017 | 251 | 22 | Not_Computer | Not_Computer
Z-1319-2016 | 73334 | apparatus, melting point, paraffin | IDT | Tissue processing equipment. | Pathology | Pathology | 510(K) Exempt | Richard-Allan Scientific Paraffin Type 3 |  |  | February 16, 2016 | 2016 | Richard-Allan Scientific Company | Multiple complaints were received from customers indicating that they ordered product 8335(Type 3 Paraffin) but when they opened the box, they found product 8336 (Type 6 Paraffin) inside. | Labeling mix-ups | ThermoFisher sent an URGENT MEDICAL DEVICE RECALL NOTIFICATION letter dated February 15, 2016, to all affected customers. The letter contained the following instructions: If you have the above lot in your possession, please determine whether the bags are labeled as Type 3 (PN 8335) or Type 6 (PN 8336).Âż If the product is labeled as Type 3 (PN 8335) please complete the attached Response form with the quantity in inventory and return the form to Richard-Allan Scientific. The product may be used and no further action is required.Âż If the product is labeled as Type 6 (PN 8336) please complete the attached Response form with the quantity in inventory. The product may be used as is or returned to Richard-Allan Scientific. Please indicate on the Response form if the product will be used or a return is desired. Return the completed form to Richard-Allan Scientific.Âż If affected product has been used previously, please complete the attached Response form with the quantity that was used. No further action is required.By returning the attached Acknowledgment Form will facilitate reporting of this matter to the FDA. Customers with questions were instructed to call 269-544-5628. |  | 29 | 29 cases (10 bags per case) | Worldwide distribution to AZ , CA , FL , GA , IL , MI , MN , NM , NY , OH , TX , UT and Internationally to Germany and the United Kingdom.6 distributed to OUS accounts: Germany, and the United Kingdom | Open, Classified |  |  | 29 | Not_Computer | Not_Computer
Z-2311-2016 | 74600 | No | DIO | Cocaine and cocaine metabolite test system. | Toxicology | Toxicology | 510(k) | AmediCheck Instant Test Cup, Part No. C-C0501A-3A |  |  | July 05, 2016 | 2016 | Branan Medical Corporation | Branan Medical is recalling the AmediCheck Instant Test Cup, a 5 panel drug cup, because it may erroneously contain 12 panel cups due to a packaging error. | Labeling mix-ups | BMC sent an Urgent Medical Device Recall Notice letter on July 6, 2016 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers are instructed to complete and return the form even if they do not have affected product. The letter informs the customers that replacement product will be provided free of charge for any unused portion of the affected lot based on completion and return of the attached Customer Verification Form. Customers are instructed to return the completed Customer Verification within 10 business days of receiptby FAX at 949-598-7166, or by email at branan.complaints@brananmedical.com. Customers with anyquestions about the information contained in the notification are instructed to contact Branan Medical Corporation at 949-598-7166 or by e-mail atbranan.complaints@brananmedical.com. |  | 96 | 96 units | US Nationwide distribution in the states of: TN, NC, FL, and PA | Open, Classified |  |  | 96 | Not_Computer | Not_Computer
Z-0765-2017 | 75675 | No | NLM | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Reprocessed Xcel Bladeless Trocar (H12LP) |  |  | November 11, 2016 | 2016 | Stryker Sustainability Solutions | An increase in reports indicating that the reprocessed H12LP Trocar adjustable plug may break in the packaging or during use. | No Marketing Application | Stryker sent an Urgent Medical Device Recall letter to all affected customers on November 11, 2016.Customers were instructed to discontinue use of the affected lots and were asked to follow the link below to complete the Recall Effectiveness Check Form and indicate if any devices from the lots referenced in Attachment 1 remain in their inventory. This form must be completed even if no affected product is found.www.novasyte.com/sss142Use the following reference number as you complete the form: [Ship to Account #].If the form indicates that any affected devices remain in your inventory, a prepaid shipping label will be issued for the return of the product. Credit for all affected devices returned will be issued to your facility.Address any questions regarding the recall to Novasyte by calling (760) 279-3180 or emailing sss142@novasyte.com. |  | 40476 | 40,476 devices | Worldwide Distribution - US (nationwide) and Internationally to Israel and CanadaForeign distribution to the following; IL and CA | Terminated | May 05, 2017 | 175 | 40476 | Not_Computer | Not_Computer
Z-1316-2016 | 73281 | No | IMD | Hot or cold disposable pack. | Physical Medicine | Physical Medicine | 510(K) Exempt | evaluemed Infant Heel Warmer 301-1223; Distributed by Tri-anim |  |  | February 10, 2016 | 2016 | Philips Electronics North America Corporation | eValueMed Infant Heel Warmers and eValueMed Infant Transport Mattresses bear no expiration date, nor any warning that product used beyond its expected shelf life may warm to temperature above its specification. | Labeling design | Philips sent an Urgent Medical Device Recall leter dated January 2016. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to do the following: Actions for Tri-Anim Customers to Take1.) Discontinue use of eValueMed Infant Heel Warmers and eValueMed Infant Transport Mattresses.2.) Check for any inventory of eValueMed Infant Heel Warmers or eValueMed Infant Transport Mattresses.3.) Destroy all products on hand (e.g., by activating or puncturing each unit). Make a note of the total number of cases of products that were destroyed as this will be needed for step 4 below.4.) Complete and sign the Customer Reply Form provided in this letter or Tri-anims version thereof, to notify Tri-anim of how many eValueMed Infant Heel Warmers and eValueMed Infant Transport Mattresses were destroyed. Submit your signed reply form only after you have completed all steps. Please do not return any product to Philips.Customers with questions should call 336-449-5230 or 262-798-9802.For questions regarding this recall call 978-659-3000. |  | 539.156 | 539.156 | Nationwide distribution to Ohio only. | Terminated | August 01, 2017 | 538 | 539.156 | Not_Computer | Not_Computer
Z-2547-2016 | 74663 | No | FCG | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Bard TruGuide Disposable Coaxial Biopsy Needle  Mixed Product |  |  | July 06, 2016 | 2016 | Bard Peripheral Vascular Inc | May contain two trocar tip stylets with coaxials or two blunt tip stylets instead of one of each. | Packaging | Bard Peripheral Vascular (BVP) sent an Urgent Medical Device Recall Notification letter dated July 6, 2016 to all affected customers via FedEx with proof of delivery notification.. The letter identified the product, the problem, and the action to be taken by the customer. The notification provided customers the following instructions:1. Do not use or further distribute any affected product.2. Check all inventory locations within your institution for affected product code I lot number combination listed in the recall notice. If you have further distributed any of the product code I lot numbers, please immediately contact that location, advise them of the recall, forward these instructions and have them return the affected product to BPV.3. Please remove any identified product from shelves.4. If used the affected product, complete and return the attached Recall and Effectiveness Check Form indicating no product will be returned.Once the product affected by recall has been removed from inventory;5. Fill out the Recall and Effectiveness Check Form. Be sure to state the quantities and lot numbers of each recalled product in stock even if no longer have possession of the recalled product.6. Call the firm's Recall Coordinator Raye Seisinger at 1-800-321-4254 Option #2 Ext 2501 (M-F 6am to 3pm MST) or email at raye.seisinger@crbard.com. Once all information has been verified, the Recall Coordinator will issue a Return Authorization (XC) Number or Consignment Recall Number (XH) to facilitate the expedient return of the product. BPV will provide a credit for your returned product.7. Fax the Recall and Effectiveness Check Form to BPV at 1-800-994-6772. If cannot FAX the form, please call the BPV recall coordinator at 1-800-321-4254 Option #2 Ext 2501) and report the required information verbally.8. A mailing label is enclosed to return the affected product. Please mark the outside package as "RECALLED PRODUCT" and i |  | 1250 | 1250 units | Nationwide Distribution to the following; CA, LA, VA, IA, AR, SC, PA, MA, IL, MO.Foreign distribution to the following; Canada, Belgium, Australia, Taiwan, China, Argentina. | Terminated | November 17, 2016 | 134 | 1250 | Not_Computer | Not_Computer
Z-0753-2017 | 75826 | No | KKX | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Ecolab, Microtek Rolled Probe Cover PC0907NB, Rx Only. |  |  | November 30, 2016 | 2016 | Ecolab Inc | The reason for the recall is that the inner carton box of the Rolled Probe Cover PC0907NB has the incorrect expiration date (2015-10-01) on the label while the individual pouch labels and outer shipper have the correct expiration date (2018-10-01). Themanufacturing date was inadvertently used as the expiration date on the inner carton box labels. | Incorrect or no expiration date | Ecolab sent an "Urgent Medical Device Recall" letter dated November 30, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were advised to quarantine the product for return. If the product was further distributed notify them of the recall. Requested customers to complete and return the Recall Response Form. For questions contact Customer Service by phone at 1-800-824-3027 or e-mail customerservice@microtekmed.com.For further questions please call (651) 250-4385. |  | 40 | 40 cases (144 eaches per box) | US Distribution to the states of : CA, FL, MI, NJ, TX, UT and VA.. | Open, Classified |  |  | 40 | Not_Computer | Not_Computer
Z-0763-2017 | 75770 | No | FRO |  |  | General &amp; Plastic Surgery | 510(k) | Ceramax Skin Barrier Cream: |  |  | November 21, 2016 | 2016 | Oculus Innovative Sciences, Inc. | Firm became aware that professional samples of Ceramax are not meeting stability standards set for the product. | Packaging process control | Oculus issued Medical Device Recall letters dated November 21, 2016, to Sales reps notifying them to stop all further distribution and immediately remove the affected lots of Recall product from sample inventory. Customers were also instructed to sign the Response form verifying receipt of the letter. Customers with questions were instructed to call Oculus Innovative Science-Regulatory Affairs at 707-559-7181. For questions regarding this recall call 559-7191. |  | 68970 | 68,970 units | Nationwide Distribution | Terminated | June 13, 2017 | 204 | 68970 | Not_Computer | Not_Computer
Z-0687-2016 | 73038 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | The e.cam Dual Signature Gamma Camera system. Used to detect or image the distribution of radionuclides in the body or organ, using the following techniques: planar imaging, whole body imaging, tomographic imaging for isotopes with energies up to 588keV. |  |  | December 31, 2015 | 2015 | Siemens Medical Solutions USA, Inc. | To provide operator instruction manuals for the e.cam Dual Signature; Gamma Camera that were not sent in its entirety when the e.cam systems were refurbished and shipped to the customer. | Packaging process control | An updated Customer Advisory Notice Letter will be provided to each impacted consignee. The receipt of the letter will be documented and additional attempts will be made until such receipt is acknowledged. The Customer Advisory Notice asks that the letter be placed with systems instructions for use along with all provided user documentation. It also advises the consignee to forward the notice to the new customer if the equipment was sold.Customers with questions regarding the Customer Advisory Notice can contact their local Service representative or use the following contact number 1-800-888-7436. |  | 10 | 10 | Distributed in the states of Florida, New York, Indiana, Maryland, South Dakota, Missouri, Tennessee, and Georgia, and the countries of Australia and India. | Open, Classified |  |  | 10 | Not_Computer | Not_Computer
Z-2550-2016 | 74736 | No | BZD | Noncontinuous ventilator (IPPB). | Anesthesiology | Anesthesiology | 510(k) | Quattro Air FFM MED  AMER, Product Code: 62702, full face mask for use with active-exhaust-valve ventilators |  |  | July 12, 2016 | 2016 | Resmed Corporation | Masks in this lot, labelled and sold as 62702 vented Quattro Air FFM size Medium actually contain a non-vented (NV) Quattro Air FFM Large mask. | Packaging process control | An urgent product recall notification letter was sent to all customers on 7/12/16 to inform them that ResMed is recalling NV Quattro Air &amp; Quattro Air Mask due to a mislabeling Issue. The letter informs the customers that during the investigation of a product complaint, ResMed identified that a small number of masks labelled and sold as 62702 Quattro Air FFM MED  AMER, Lot 1142623 actually contain an NV Quattro Air FFM Large mask. The letter informs the customers to notify anyone they may have sold the products to. Customers with any questions are instructed to call the ResMed Customer Service team: 858-836-5968. |  | 42 | 42 units | US nationwide distribution | Open, Classified |  |  | 42 | Not_Computer | Not_Computer
Z-1313-2016 | 73325 | No | JSI | Selective culture medium. | Microbiology | Microbiology | 510(K) Exempt | NEL GBS Medium 20 ml in 100x15mm plate (Standard Petri Dish) |  |  | December 21, 2015 | 2015 | Northeast Laboratory Services, Inc. | Product marketed without a 510 (k) | No Marketing Application | Northeast Laboratory services issued recall letter dated 2/29/16and advised users to: Immediately examine your inventory for all lots of these products you may have in stock. Please destroy all remaining product in stock. Users requested to fill out the questionnaire immediately and return the questionnaire via email or Fax.Previously Northeast Laboratory Services notified users by letter dated 12/21/15 that the firm has discontinued the products effective January 2016. Alternative products were referenced in the letter for users to transition to replacement product. |  | 9300 | 9300 | Nationwide Distribution. | Terminated | April 07, 2016 | 108 | 42040 | Not_Computer | Not_Computer
Z-0756-2017 | 75481 | No | FMH | Specimen transport and storage container. | Pathology | General Hospital | 510(K) Exempt | Covidien Precision Specimen Container- 4 oz (118 mL), Plastic Screw Top Cap, OR Packaged, Sterile |  |  | October 14, 2016 | 2016 | Covidien LLC | Sterility of the outer surface of the container compromised due to packaging defect. | Packaging process control | Covidien initiated recall by letter on October 14, 2016. Customers were notified via Federal Express and the letter instructs customers to discontinue use of the product and return all inventory of the affected lots to Medtronic (Covidien), 195 McDermott Road, North Haven, CT 06473, Attention: Field Returns Department.Accounts requested to completed Recalled Product Return From to obtain an RGA.Questions: Contact your Medtronic representative or Customer Service at (800) 882-5878. |  | 5187100 | 5,187,100 Units | NationwideForeign: Japan | Open, Classified |  |  | 5187100 | Not_Computer | Not_Computer
Z-0650-2016 | 72521 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Biomet Integral Centralizer Hip System, Lateralized Femoral Stem (Co-Cr-Mo), |  |  | November 17, 2015 | 2015 | Biomet, Inc. | PMMA is listed as a material on the label but the product does not contain PMMA. | Labeling Change Control | On 11/17/2015, URGENT MEDICAL DEVICE RECALL NOTICE notifications were sent to the affected distributors via courier with instructions for identifying and returning the affected product. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Questions related to this notice should be directed to (574) 372-1570, Monday through Friday, 8 a.m. to 5 p.m. |  | 21 | 21 | Worldwide distribution. US nationwide including IA, UT, MN, PA, MO, OK, IL, OH, and WA; CANADA, and NETHERLANDS | Open, Classified |  |  | 21 | Not_Computer | Not_Computer
Z-1528-2016 | 73205 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Trinity Acetabular Shell |  |  | January 29, 2016 | 2016 | Corin USA Limited | Potential for abrasion of the sterile barrier pouches of the new packaging system. | Packaging process control | Corin sent an Urgent Medical Device Recall letter dated February 3, 2016. The letter identified the affected product, problem and actions to be taken. On January 29, 2016, Eight (8) of 11 distributors were spoken to during a Corin National Sales Meeting informing them of the recall. The 3 remaining distributors were notified by phone calls same date. Also, on February 2, 2016 Corin sent a Recall Letter to their Foreign consignee informating them of the recall of the Trinity Acetabular Shell and the actions to be take by the costumers including the return of the Customer Response Form and Contact information as Corin LTD Compliance Director at +44 (0) 1285 649190 Monday thru Friday 8:30am to 5:00pm UK Time |  | 629 | 629 devices in total | Worldwide Distribution - US Nationwide in the states of CA, FL, ID, TX, OK, OH, MA, and NC and the countries of Australia Belgium, Egypt, Germany, Ghana, Italy, Japan, Poland, South Africa and UK. | Open, Classified |  |  | 629 | Not_Computer | Not_Computer
Z-1321-2016 | 73525 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | TOPAZ MicroDebrider 0.8 mm REF AC4040-01; TOPAZ MicroDebrider IFS 0.8 mm REF ACH4040-01; TOPAZ MicroDebrider 0.8 mm REF Q6000-01; TOPAZ EPF MicroDebrider 0.8 mm 45 REF Q6002-01 |  |  | March 03, 2016 | 2016 | ArthroCare Corporation | During functional testing of devices after real time aging, it was determined that product failed to meet shelf-life at the expiration date indicated on the product labels. | Package design/selection | Consignees were notified via letter on 3/3/2016. Consignees were instructed to quarantine the recalled products and complete an Inventory Return Certification Form. |  | 26903 | 26,903 devices | Worldwide Distribution-US (nationwide) and the following countries-Austria, Belgium, Bulgaria, Canada, Denmark, Estonia, Finland, France, Germany, Gibraltar, Great Britain, Greece, Ireland, Israel, Italy, Latvia, Netherlands, Norway, Poland, Portugal, Spain, Sweden, Switzerland, China, Hong Kong, India, Lebanon, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Taiwan, Thailand, Turkey, United Arabic Emirates, Vietnam, Australia, Japan, New Zealand, Brazil, and Chile. | Terminated | May 17, 2017 | 440 | 26903 | Not_Computer | Not_Computer
Z-0771-2017 | 75781 | No | KOD | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | RUSCH, Pocket PAC IC Intermittent Cath Kit, 8 FR, 12 FR, 14 FR, 10 FR and 16 FR, Rx Only, Sterile, |  |  | November 22, 2016 | 2016 | Teleflex Medical | Labeling Inconsistency: The products have beenlabeled with the incorrect expiration date which exceeds the actual, validated 35-month shelf life. | Error in labeling | Teleflex sent an Urgent Medical Device Recall Notification letter dated November 22, 2016, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the customer. Customers with affected stock, immediately discontinue use and quarantine any affected products. To return product, customers were advised to complete the enclosed Recall Acknowledgment Form and fax it to 1-855-419-8507, Attn: Customer Service or email to recall@teleflex.com. Customers with no affected stock should also complete the enclosed Recall Acknowledgment Form to confirm receipt of the letter. Customers with questions should contact Customer Service at 1-866-246-6990. |  | 102843 | 102,843 units | Nationwide Distribution to CA, CO, FL, GA, IL, KS, KY, MA, MO, MS, NC, NJ, NY, OH, OK, OR, PA and TX. | Open, Classified |  |  | 102843 | Not_Computer | Not_Computer
Z-0317-2017 | 75371 | No | HAB | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ZimmerÂż Gender Solutions" Patellofemoral Joint Prosthesis |  |  | September 28, 2016 | 2016 | Zimmer Biomet, Inc. | Complaints have been reported that the handpiece is inoperable. The investigation found that the complaints were related to the lack of preventative maintenance contrary to what is recommended in the instructions for use included with the handpiece. An inoperable device may lead to a delay in surgery | Labeling design | On 9/28/2016, URGENT MEDICAL DEVICE FIELD REMOVAL notifications were sent to the affected distributors and Risk Managers / Surgeons via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Risk Managers and SurgeonsRisk Manager Responsibilities: 1. Review this notification and ensure affected personnel are aware of the contents. 2. Assist your Zimmer Biomet sales representative with quarantining all affected product. 3. Your Zimmer Biomet sales representative will remove the affected product from your facility for the initial preventative maintenance and arrange for a loaner device. It is estimated that your handpieces will undergo preventative maintenance sometime between November 2016 and March 2017. 4. Complete Attachment 3  Certificate of Acknowledgement. a. Return a digital copy via email to CPWARFieldAction@zimmerbiomet.com or via fax to 574-372-1683 within three (3) days. b. Retain a copy of the Acknowledgement Form with your field removal records in the event of a compliance audit of your facilities documentation. 5. If after reviewing this notice you have further questions or concerns please call the customer call center at 1-877-946-2761 between 8:00 am and 5:00pm EST, Monday through Friday. Calls received outside of the call center operating hours will receive a prompt to record a voicemail or be transferred to an on-call representative in the case of an emergency. Alternatively, your questions may be sent by email to CPWARFieldAction@zimmerbiomet.com. Surgeon Responsibilities: 1. Review this notification for awareness of the contents. 2. There are no specific patient monitoring instructions related to this field removal that are recommended beyond your existing follow up schedule. 3. Complete Attachment 3  Certificate of Acknowledgement. a. Return a digital copy via email to |  | 703 | 703 | Nationwide.CANADAHONG KONGKOREASINGAPORECHINAVENEZUELAAUSTRALIAJAPANNETHERLANDSGERMANYINDIATAIWAN | Open, Classified |  |  | 703 | Not_Computer | Not_Computer
Z-2333-2016 | 74610 | No | KNX | Urine collector and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | SureStep Tray, with Add-a-Foley, COMPLETE CARE, Drainage Bag, and STATLOCK Stabilization |  |  | June 29, 2016 | 2016 | C.R. Bard, Inc. | Labeling problem | Labeling mix-ups | Bard Medical sent an Urgent Medical Device Product Recall letter dated June 29, 2016. The letter identified the product, the problem, and the action to be taken by the customer. The letter requested that customers examine their inventory and quarantine the product. The letter also requested a sub-recall. The letter included a reply form which is to be returned to Bard. Customers were also instructed if product was further distributed they identify customers and notify them at once. Customers with questions should contact the BMD Recall Coordinator at 1-800-793-8110. |  | 530 | 530 units | Nationwide Distribution to CA, FL, IL, IA, KY, NC, MN, NC, OR, RI, SD, TX, UT, and WA. | Open, Classified |  |  | 530 | Not_Computer | Not_Computer
Z-1937-2016 | 73828 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | Henry Schein Single Use Syringe Luer-Lock, 10 mL, Intended for use by health care professionals for general purpose fluid aspiration/injection |  |  | March 25, 2016 | 2016 | Becton Dickinson &amp; Company | BD 10 mL Syringe Luer-Lok Tip with BD PrecisionGlide Needle 21G x 1 (0.8 mm x 25 mm) and Henry Schein 10 mL Single Use Syringe Luer-Lock are being recalled due to possible adulteration and non-conforming manufacturing practices. | Error in labeling | Becton Dickinson sent via email and UPS Ground an "Urgent Voluntary Product Recall" Letter dated March 25, 2016 along with a Recall Response Form and packaging instructions to their sole customer. Customers were instructed to immediately review their inventory , quarantine any affected product, and immediately discontinue shipment of the affected product. Customers should complete the Recall Response Card from and fax it to BD at 1-201-847-6990. Return all affected product with the completed Recall Response Card. If product was further distributed please identify your customers and notify them immediately of the product recall. Customers were also instructed to complete the Recall Response Form even if they do not have any affected lots in their inventory. Customers with questions were instructed to call 1-201-847-4267. For questions regarding this recall call 201-847-5612. |  | 811200 | 811,200 units | Nationwide Distribution to NY only | Open, Classified |  |  | 987200 | Not_Computer | Not_Computer
Z-2326-2016 | 74565 | No | FTR | Silicone gel-filled breast prosthesis. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | PMA | Mentor MemoryGel Breast Implant; Indicated for females for breast augmentation and breast reconstruction. |  |  | July 01, 2016 | 2016 | Mentor Texas, LP. | The box of 300cc MemoryGel Breast Implant was labeled with null manufacturing and expiration dates. | Labeling mix-ups | The affected consignees were notified via customer letter on 7/1/16, and a follow-up with respective sales representative via visit. |  | 35 | 35 units | Distributed in the states of NY, WI, CA, TX, CO, NC, NJ, OH, TN, RI, and IL. | Open, Classified |  |  | 35 | Not_Computer | Not_Computer
Z-0816-2016 | 73164 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | LEGION(TM) HEMI STEPPED, TIBIAL SCREW-ON WEDGE, REF 71421147, QTY 1, 5 MM, SIZE 5-6 LT-LAT / RT-MDL; Tibial knee prosthesis, augment. |  |  | January 18, 2016 | 2016 | Smith &amp; Nephew, Inc. | The package for catalog item, a LEGION Hemi Stepped Tibial Wedge, 5 MM Size 5-6 LT-LAT/ RT-MDL, should include both a screw-on hemi stepped tibial wedge and two tibial wedge screws to affix the wedge to the tibial baseplate. A complaint was received indicating that the screws packaged with the wedge were too long. | Packaging change control | The firm notified their direct consignee of the issue on 01/18/2016 by email. The notice requested return of the product. |  | 8 | 8 units | Distributed in the states of TN, WI, OH, and KS, and in the countries of Belgium, Germany, and United Kingdom. | Open, Classified |  |  | 8 | Not_Computer | Not_Computer
Z-0611-2016 | 72570 | No | DAP | Fibrinogen/fibrin degradation products assay. | Hematology | Hematology | 510(k) | Siemens Acute Care DDMR TestPak |  |  | November 11, 2015 | 2015 | Siemens Healthcare Diagnostics Inc | Stratus CS Cartridge Barcode labels may not scan on the Stratus CS systems | Labeling design | Siemens Healthcare Diagnostics issued an Urgent Field Safety Notice by e-mail to all affected Siemens Healthcare Diagnostics Regional offices on November 11, 2015 both in the United States and internationally for communication with affected customers. This notice informs customers of the issue and provides mitigation instructions. Customers with questions were advised to contact their Siemens Customer Care Center or their local Siemens technical support representative. For questions regarding call 781-269-3000. |  | 2647 | 2647 | Worldwide Distribution - US (nationwide), Algeria, Armenia, Australia, Austria, Bangladesh, Bosnia Herzegovina, Canada, Croatia, Czech Republic, Denmark, Finland, France, Germany, Greece, India, Ireland, Italy, Japan, Libya, Malaysia, Poland, Portugal, Romania, Saudi Arabia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, U.A.E., and United Kingdom. | Open, Classified |  |  | 12288 | Not_Computer | Not_Computer
Z-0827-2017 | 75855 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Zeus-P Lumbar Interbody Device |  |  | November 22, 2016 | 2016 | Amendia, Inc | Labeling error where three boxes in inventory had conflicting patient and primary labels. Evaluation indicated that the primary label and patient label did not match. The primary label matched the label on the inner pouch, but the patient label was incorrect. | Error in labeling | Amendia Inc. sent an Urgent - Medical Device Recall letter dated November 22, 2016, to all affected customers. The letter requested that they examine inventory and quarantine the product. The letter also requested that they contact Amendia to confirm receipt of the letter and to obtain a Return Material Authorization for any product on hand.For further questions please call (404) 610-7215. |  | 25 | 25 units | US Distribution to the states of : GA. MI, OK, SC, TX, VA | Terminated | June 05, 2017 | 195 | 25 | Not_Computer | Not_Computer
Z-1699-2016 | 73678 | [6 empty] | Orapharma (Onpharma) Onset Mixing Pen Rx only, Product/Part Number 1054-02. |  |  | January 29, 2016 | 2016 | Valeant Pharmacueticals International | It has been determined that the Onset Mixing Pen has an optimal use period of 18 months from start of use. However, the current labeling and IFU does not stipulate this time period. | Labeling False and Misleading | Valeant sent a letter entitled "Field Correction to Orapharma Onset Mixing Pen," dated 1/29/2016, to customers via UPS. The letter identified the affected device and the issue involved. The firm recommends that users stop using the device if it had been in use for more than 18 months. Customers were to refer to the list of the affected lots to determine if their device needed to be replaced. If the device was one of the ones identified, users were to contact Valeant Pharmaceuticals to arrange for a replacement at no charge. Questions should be directed to customer service 800-321-4576. |  | 7716 | 7,716 | Nationwide Distribution | Open, Classified |  |  | 7716 | Not_Computer | Not_Computer
Z-1357-2016 | 72922 | No | MJC | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | BARDEX I.C. Anti-Infective 2-Way 3cc Foley Catheter, Rx only. Intended for urinary bladder drainage in female patients requiring catheterization for management of incontinence, voiding dysfunction and surgical procedure. |  |  | December 18, 2015 | 2015 | C.R. Bard, Inc. | Labeling error: Catalog contains Instructions for Use (IFU) error; incorrect inflation volume. | Error in labeling | C. R. Bard, Inc. sent a "NOTIFICATION OF MISLABELING" letter, dated 12/18/2015 to all affected consignees. The letter identifies the product, problem, and actions to be taken by the customers. |  | 1428000 | 1,428,000 units (total) | Distributed US (nationwide) and the countries of Canada and Australia. | Open, Classified |  |  | 1428000 | Not_Computer | Not_Computer
Z-2237-2016 | 74486 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Replace Select Tapered TiU NP 3.5x11.5mm, Article No. 36105 |  |  | June 22, 2016 | 2016 | Nobel Biocare Usa Llc | Incorrect top label on an implant package. | Labeling False and Misleading | Nobel Biocare sent a customer notification letter dated June 23, 2016 to all affected customers to inform them that Nobel Biocare has detected that the lot#13030677 of Replace Select Tapered TiU NP 3.5 x 11.5 mm (product#36105) carries an incorrect top label. The letter informs the customers that the incorrect top label states "Replace NP 3.5 x 11.5 CC" (ie. Conical Connection). The product, Replace Select Tapered TiU NP 3.5 x 11.5 mm, has a Tri-lobe connection. Customers were instructed to complete the attached Quality Message Acknowledgement form. Customer with questions were instructed to contact customer service at 714-282-4808. |  | 53 | 53 units total (33 units in US) | Worldwide Distribution - US including OR, OH, NY and Internationally to Taiwan and BrazilNon US: Taiwan, SC | Completed |  |  | 53 | Not_Computer | Not_Computer
Z-1848-2016 | 73427 | No | LRS | Intravascular catheter. | General Hospital | General &amp; Plastic Surgery | Enforcement Discretion | Kerlix" X-Ray Detectable Laparotomy Sponges, Tray Pack, 16" x 17" (40.6 cm x 43.2 cm) |  |  | March 03, 2016 | 2016 | Covidien LLC | Compromised sterility due to breach of barrier | Packaging process control | Covidien initiated recall by letter on March 4, 2016. Customers were notified via Federal Express and the letter instructs customers to discontinue use of the product and return all inventory to Medtronic (Covidien), 195 McDermott Road, North Haven, CT 06473, Attention: Field Returns Department. Any questions or concerns, contact your Medtronic representative or Customer Service at (800) 882-5878 |  | 114720 | 114720 | Distributed US (nationwide) and the countries of Canada, Austria, Belgium, Columbia, Croatia, France, Germany, Guadeloupe, Ireland, Italy, La Reunion, New Zealand, Norway, Poland, Portugal, Saudi Arabia, South Africa, Spain,Sweden and United Kingdom. | Terminated | June 15, 2017 | 469 | 8748710 | Not_Computer | Not_Computer
Z-1347-2016 | 73555 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | SIGNA Pioneer MR Systems |  |  | March 01, 2016 | 2016 | GE Medical Systems, LLC | Ferrous material in the SIGNA Pioneer Table Control Box (TCB) could cause unexpected attraction to the MR system's high strengthmagnetic field during servicing activities. This could create a potential risk of serious injury to service personnel who attempt to inappropriately remove the TCB fromthe magnetic field. There have been no events reported as a result of this issue. Clinicalscanni | Labeling Change Control | Consignees were sent a GE Healthcare "Urgent Medical Device Correction" letter GEHC Ref# 68002 dated March 1, 2016. The letter was addressed to Hospital Administrators / Risk Managers, Radiology Department Managers, and Radiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. For questions or concerns regarding this notification,contact GE Healthcare Service at 18004371171 or your local Service Representative. |  | 16 | 16 (2 US; 14 OUS) | Worldwide Distribution - US Nationwide in the states of SC, NC and the countries of Australia, Canada, France. Germany, Japan and Turkey. | Open, Classified |  |  | 16 | Not_Computer | Not_Computer
Z-0821-2017 | 75892 | No | KRO | Knee joint femorotibial metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Tibial Component Metal Cased Small Long; Tibial Component Metal Cased Small Short; Tibial Component Metal Cased Std Long; Tibial Component Metal Cased Std Short |  |  | December 07, 2016 | 2016 | Stanmore Implants Worldwide Ltd. | The recalling firm has received 5 customer inquiries reporting that at least one pouch of the double pouch packaging configuration has been punctured by the tip of the device's stem. | Packaging | Stanmore Implants sent an Urgent Field Safety Notice dated November 25, 2016, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to immediately check their internal inventory and quarantine all subject devices, circulate the Field Safety Notice internally to all interested/affected parties, maintain awareness of the notice internally until all required actions have been completed within their facility, inform Stanmore Implants if any of the subject devices have been distributed to other organizations, complete the attached customer response form and list of affected batches form, return the completed form and any affected devices to their S/W Representative. Customers with questions were instructed to contact their local sales representative. |  | 51 | 51 units | Nationwide Distribution to OH, NE, FL, LA, NY, TX, MS, MD, CT, AL, PA | Terminated | July 17, 2017 | 222 | 105 | Not_Computer | Not_Computer
Z-2211-2016 | 74297 | No | OIJ | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(K) Exempt | Mammotome Revolve Dual Vacuum Assist Biopsy System Model number: MST1009 &amp; MST0815. Each disposable probe is packaged within a thermoformed tray and covered with a pre-printed (product code and graphics) Tyvek lid. Each tray is labeled on-demand after sealing with a lot number and expiration date. Trays are kitted into sets of five probes and packaged within a sales carton. Two sales cartons are then packed into a corrugated case and shipped to the global distribution center. Device is labeled as sterile. |  |  | May 19, 2016 | 2016 | Devicor Medical Products Inc | A 10 Gauge needle, 9cm long was shipped in an 8 Gauge needle, 15 cm package. | Packaging | An Urgent: Medical Device Recall of Mammotome Revolve Dual Vacuum Assist Biopsy Probe due to Potential Product Code Mislabel communication was sent to the customers via 1st class USPS on June 6, 2016. The labeling and the Customer response form was also attached. Customers were asked to examine their inventory and quarantine the affected product. If the product was further distributed, those customers are to be identified and notified. The letter stated that the assigned sales rep would be in contact to support the identification, inspection and replacement of the recalled products. Customers are to complete and return the enclosed response form. Questions are to be directed to Mammatome Customer Service 1-877-926-2666. |  | 360 | F11537107D1 = 360; F11536423D1 = 80; F11537109D1 = 360 | Worldwide Distribution -- USA, including the states of AL, AR, CA, CT, DC, FL, GA, IL, KS, KY, LA, MA, MD, ME, MI, MN, MS, NC, NJ, NY, OH, PA, SC, TN, WA and WI; and, the countries of Canada and Japan, Korea, and The Netherlands. | Terminated | August 03, 2017 | 441 | 360 | Not_Computer | Not_Computer
Z-2203-2016 | 74154 | No | OCC | Respiratory viral panel multiplex nucleic acid assay. | Microbiology | Microbiology | 510(k) | cobasÂż Liat Influenza A/B Quality Control Kit; For use with cobasÂż Liat Influenza A/B Assay. |  |  | May 17, 2016 | 2016 | Roche Molecular Systems, Inc. | Kit labeled with the incorrect expiration date. | Incorrect or no expiration date | An Urgent Medical Device Recall (UMDR) notice (Customer Letter) was sent to the 35 individual affected customer sites by UPS on May 17 2016.Customers were asked to take the following actions: " Discontinue use of cobasÂż Liat Influenza A/B Quality Control kit lot numbers 51222E, 60121H, 60201D, and 60229F, and discard them per your local regulations. " Complete the enclosed UMDC faxback form (6941-00-0516) and fax it to 1-844-449-8506. " File this UMDC for future reference. |  | 194 | 194 distributed kits | Distributed in the states of: TX, NY, OR, HI, MD, PA, NC, MA, IA, WA, OR, AK, ME, MN, VA, MT, PA, and LA. | Open, Classified |  |  | 194 | Not_Computer | Not_Computer
Z-2375-2016 | 74782 | No | DZI | Bone cutting instrument and accessories. | Dental | Dental | 510(k) | Instructions for Use (IFU) and Manual NobPrl CC Guided Surgery Kit |  |  | July 20, 2016 | 2016 | Nobel Biocare Usa Llc | Incorrect drill protocol in the Instructions for Use (IFU) | Labeling design | The firm, Nobel Biocare, sent "IMPORTANT-QUALITY MESSAGE" Field Correction Letters and response forms dated July 22, 2016. The letter describes the product, problem and actions to be taken and informs the customers that the titles of the IFU and the Manual are "NobelGuide for NobelParallel Conical Connection" (ID number 2002) and "Guided Surgery with Nobel Parallel Conical Connection" (ID number 79423). The Customers are instructed to review if you have a printed version of the IFU or the Manual related to the article and discard them; and to complete the attached "Quality Message Acknowledgement" form and fax to 714-998-9348 or email to: joann.radovich@nobelbiocare.com. Customers are informed that an English version of the corrected IFU is attached to the letter and translated versions will be available on the Nobel Biocare website for electronic IFU's http://www.nobelbiocare.com/IFU within 30 days. If you have further questions, contact our customer service department at 714-322-5001. |  | 231 | 231 kits | Worldwide Distribution: US (nationwide) and Countries of: Austria, Belgium, Canada, Germany, Denmark, Spain, France, Great Britain, Italy, Netherlands, Norway, Poland, Portugal, Russia, and Turkey. | Open, Classified |  |  | 231 | Not_Computer | Not_Computer
Z-2191-2016 | 74416 | No | HWE | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | Orthopedic | 510(K) Exempt | Part # SD313.011, Matrix Neuro Screwdriver Blade MQC/MR Conditional/Short, Lot # UR75220, mfg 7/18/2007 |  |  | May 31, 2016 | 2016 | Synthes (USA) Products LLC | DePuy Synthes is voluntarily initiating this recall due to affected devices that were distributed, etched and or labeled with MR Safety information not meeting the current ASTM standard. | Labeling False and Misleading | DePuy Synthes mailed an URGENT MEDICAL DEVICE FIELD SAFETY NOTIFICATION letter to their customers (Hospital Personnel, including Imaging Department Personnel and DePuy Synthes Colleague) to inform them of the recall and actions to take. The letter identified the affected, problem and actions to be taken. The letter instructed customers not to use the affected Instruments which are labeled or etched MR Safe and MR Conditional during MRI scans. Customers were asked to review the information contained in the Field Safety Notification letter and complete the Verification Section located on page 3 of this notification. For questions call 610-719-5450 or contact your DePuy Synthes Sales Consultant. |  | 18 | 18 units | US Nationwide Distribution in the states of: AR,CA,CO,CT,DC,ID,IL,KY,LA,MA,MD,MN,MO,MS,MT,NY,PA and TX. | Open, Classified |  |  | 749 | Not_Computer | Not_Computer
Z-1843-2016 | 73664 | No | LDH |  |  | General Hospital | 510(k) | AllergiEnd ST-9 Multiple Skin Test Applicator (Item 1000) and AllergiEnd ST-9 Multiple Well Test Tray (Item (1001). Contents: 8 applicators to perform up to 72 skin tests. |  |  | March 22, 2016 | 2016 | Medscience Inc | During an FDA inspection it was found that the products are marketed without a cleared 510k. | No Marketing Application | MedScience sent an Urgent Recall letter dated March 28, 2016, to their customers to inform them of the product recall. The letter instructed customers to discontinue and refrain from further use of recalled products and to contact MedScience to arrange the return of products still in possession by the consignee. MedScience set up a phone line and email address for customers to ask questions and obtain further direction: (800) 393-8817 and recall@medscienceinc.com. |  | 11400 | 11,400 (Item 1000) &amp; 380 (Item 1001) | Nationwide Distribution to AZ, CA, CO, CT, FL, GA, LA, MD, MI, NJ, NY, PA, TX, VA, and WA. | Open, Classified |  |  | 11400 | Not_Computer | Not_Computer
Z-2202-2016 | 74489 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Femoral Arterial Line |  |  | June 02, 2016 | 2016 | Arrow International Inc | Labeling error. Some shipping cartons labeled as ASK-04018-VUH contained ASK-04020-VUH kits; and some shipping cartons labeled as ASK-04020-VUH contained ASK-04018-VUH kits. | Labeling mix-ups | Arrow International mailed an Urgent Medical Device Recall Notification letter and acknowledgement form dated June 20, 2016 to affected customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to 1. Immediately discontinue distribution and quarantine any products with the product codes and lot numbers listed above.2. Inspect affected product within your control to identify if the shipping corrugate label correctly identifies the products inside the shipping carton. Products with an incorrect label on the shipping carton can be returned to Arrow. 3. Once you have completed inspection of all of the products from your own inventory, check the box on the enclosed Acknowledgment Form that indicates that you have completed the inspection and fax it to 1-855-419-8507, Attn: Customer Service, or email it to recalls@teleflex.com.For questions contact your local sales representative or Customer Service at1-866-246-6990. |  | 50 | 50 units in total | US distribution in the states of TN and KY. | Terminated | August 21, 2017 | 445 | 50 | Not_Computer | Not_Computer
Z-2378-2016 | 74788 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | BD Tritest CD3 FITC/CD4 PE/CD45 PerCP (IVD); |  |  | July 18, 2016 | 2016 | Becton, Dickinson and Company, BD Biosciences | Internal review found that some vials of CD3/CD4/CD45 state that the product contains gelatin. The product does not contain gelatin and should state in buffer. | Labeling Change Control | BD Life Sciences sent an Urgent Product notification letter dated July 18, 2016 to all affected customers via UPS and/or e-mail. The letter informed the user of the issue and states that no action is required since product performance is not impacted. The response form is to be returned to BD as instructed. Customers requiring further assistance were instructed to contact BD Customer Support at 855-236-2772.For questions regarding this recall call 408-954-6080. |  | 14089 | 14,089 | Nationwide distribution to AR, CA, CO, IL, IN, LA, MA, MD, Ml, MN, NC, NJ, NM, NY, OR, PA, TN, TX, WA | Open, Classified |  |  | 14089 | Not_Computer | Not_Computer
Z-0318-2017 | 75332 | No | FDS | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Pentax Video Gastroscope is intended to be used with a Pentax video processor (including light source), documentation equipment, monitor, EndoTherapy Device such as a Biopsy Forceps, and other ancillary equipment for endoscopy and endoscopic surgery within the upper digestive tract including the esophagus, stomach and duodenum. |  |  | August 15, 2016 | 2016 | Pentax of America Inc | PENTAX Medical is initiating this field action to provide customers that have purchased the affected endoscopes with the most recent Operation and Reprocessing IFUs. | Labeling Change Control | Pentax sent as "Urgent Field Correction" and Field Correction Response Form dated August 15, 2016 to their affected customers. |  | 3436 | 3,436 (US) and 14,488 (OUS) | US Nationwide and Japan | Open, Classified |  |  | 3436 | Not_Computer | Not_Computer
Z-0331-2017 | 75372 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | HANDS-OFF Infusion Port Thermodilution Catheter consists of the Arrow IPTD thermodilution catheter enclosed in a contamination shield (Arrow Cath-Gard) with integral flushing/balloon test chamber, enabling the practitioner to prepare, test, and insert the catheter without exposing it to external contamination. |  |  | September 02, 2016 | 2016 | Arrow International Inc | Labeling inconsistency | Labeling mix-ups | Arrow sent an Urgent Medical Device Recall Notification letter dated September 20, 2016, to all affected customers via FedEx 2-day air. The letter identified the problem and provided instructions to immediately discontinue use and quarantine any products with the associated lot numbers indicated in the letter. If product was found, customers were asked to complete the Recall Acknowledgement Form and a Customer Service Rep will issue a Return Goods Authorization (RGA) Number for the product's return. Disposition of recalled product will be scrapped. For further questions, please call (610) 378-0131 |  | 1361 | 330 units in US and 1,031 units OUS | Nationwide distribution | Open, Classified |  |  | 1361 | Not_Computer | Not_Computer
Z-1289-2016 | 73380 | orthosis, limb brace | IQI | Limb orthosis. | Physical Medicine | Physical Medicine | 510(K) Exempt | Rebound Cartilage |  |  | February 24, 2016 | 2016 | Ossur H / F | Ossur Americas is recalling , Knee Shield / Patella Cup, Knee Undersleeve Gray, NecLoc Extrication Collar, Rebound Cartilage, Rebound Cartilage Custom, Rebound Air Walker, Rebound Diabetic Walker, Back Support Suspenders Clip on Black, Back Support Suspenders Sew On Black, Extreme Custom, Unloader Custom XT, Unloader Custom Lite, Unloader Custom, and Papoose Infant Spine Immobilizers because the p | Labeling False and Misleading | The firm sent out customer notification letters on 02/26/16.The notification letter instructed customers to include the latex caution notice in any existing inventory of affected products. The latex caution notice can be retrieved at the conclusion of the response survey or by going directly tosafetyalerts.ossur.com. Should customers require Âżssur to provide pre-printed notices for existing inventory, contact Âżssur customer service. For the Âżssur Miami JÂż Cervical Collars with lot numbers before MX150727 customers are instructed to remove and discard the assist strap, which contains natural rubber latex, located in the patient pack.Recipients of the notice with should take the following actions:1. Pass the notice to those who need to be aware within the organization or to any organization where the affected devices have been transferred.2. Maintain awareness on this notification for an appropriate period.3. If have further distributed the products, identify customers and notify them at once of the product alert. It is recommended to include a copy of the alert notification letter.4. If any of your customers are currently wearing a product, check for any skin irritations or other reactions consistent with natural rubber latex allergies.5. To verify have received the notification, visit safetyalerts.ossur.com to complete the Response Survey.If any questions, contact Âżssur Customer Service at (800) 233-6263.The notification letter for the Papoose Infant Spine Immobilizer contained the following additional actions/instructions:Examine inventory, quarantine products subject to the recall and contact customer service for a return authorization to return the product to Âżssur.If any of customers are currently wearing a product, it is recommended replacing the unit with one shipped after the transition lot number MX150727. Check for any skin irritations or other reactions consistent with natural rubber latex allergies. |  | 1030 | 1030 | U.S. distribution to the following; AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, RI, TN, TX, UT, VA, VT, WA, WI, WV | Open, Classified |  |  | 27899 | Not_Computer | Not_Computer
Z-0646-2016 | 72938 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Zimmer Periarticular Locking Plate System 3.5mm Locking Screw 58 mm Length; Zimmer Periarticular Locking Plate System 2.4mm Locking Screw 14 mm Length. |  |  | December 09, 2015 | 2015 | Zimmer Biomet, Inc. | Product on the package label did not match what was inside the package. The label was for part number 00-2359-058-35 (Zimmer Periarticular Locking Plate System 3.5mm Locking Screw 58mm Length) lot number 62918019 while the product inside the package was part number 00-2359-014-24 (Zimmer Periarticular Locking Plate System 2.4mm Locking Screw 14mm Length) lot number 62896038. | Labeling Change Control | On 12/9/2015, URGENT MEDICAL DEVICE RECALL notifications were sent to the affected distributors and hospital risk managers via courier with instructions for identifying and returning the affected product. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. If after reviewing this notification you have further questions or concerns please call the customer call center at 1-877-946-2761. Hours of operation are Monday through Friday, 8 a.m. through 5 p.m. EST. |  | 400 | 400 | Worldwide Distribution-US including the states of CA, NY, PA, AL, FL, GA, ID, WI, MO, SC, NC, and ME, and the countries of Brazil, Dominican Republic, Taiwan, China, Uruguay, Australia, Japan, Malaysia, India, Belgium, France, Italy, Netherlands, Poland, United Arab Emirates, and Canada. | Open, Classified |  |  | 400 | Not_Computer | Not_Computer
Z-0299-2017 | 75224 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Keystone Dental PrimaConnexÂż Straight Implant TC, RD Internal Connection 4.0 x 11.5mm. |  |  | September 15, 2016 | 2016 | Keystone Dental Inc | Small diameter (SD) cover screw was packaged with the regular diameter implant and do not match the implant diameter (RD) | Packaging process control | Keystone Dental initially called all customers on September 15, 2016. Forms were completed to document these calls. In addition a letter has been sent via Federal Express/or hand delivered on September 19, 2016, to all distributors and customers. Customers with questions were instructed to contact Customer Relations.For questions regarding this recall call 781-328-3390. |  | 55 | 55 units | Worldwide distribution - US (Nationwide Distribution to CA, CO, IL, MA, MN, ND, NY, NV, PA, and VA.) and Italy | Terminated | January 03, 2017 | 110 | 55 | Not_Computer | Not_Computer
Z-1784-2016 | 73734 | No | HBA | Neurosurgical paddie. | Neurology | Neurology | 510(k) | American Surgical Delicot 4 mm x 50 mm |  |  | March 31, 2016 | 2016 | American Surgical Company | Product packaging defective compromising sterility | Packaging process control | American Surgical issued an Urgent Recall and Field Safety Notice letterdated March 31, 2016 to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were provided instruction to address affected devices in their inventory. They should address this by either selecting Action A or B: Action A: Users should perform a check of the device before use and, as stated in the IFU, patties should not be used if in an open or damaged pouch. For the affected lots, special attention should be given to lower portion of the pouch where the Tyvek and clear film are sealed.Action B: Place the affected lots in quarantine and contact American Surgical Company for an RMA and replacement. All affected devices are eligible for replacement by ASC.Complete and return confirmation of actions described in Field Safety Notice.Contact:Tel: 1-800-343-0060 Option 2, Fax: 781-595-5460Email: CustomerService@AmericanSurgical.com |  | 99 | 99 units | Worldwide Distribution - US Nationwide and the countries of Australia, Belgium, Canada, France, Germany, Greece, Hong Kong, Hong Kong, Israel, Italy, Italy, Japan, Netherlands, New Zealand, Portugal, Saudi Arabia, South Africa, Spain, Sweden, Switzerland, and United Kingdom. | Open, Classified |  |  | 7701 | Not_Computer | Not_Computer
Z-1602-2016 | 73869 | No | HBB | Pneumatic cranial drill motor. | Neurology | Neurology | 510(k) | Medtronic Midas Rex Dissecting Tool Legend 10 cm 4mm BA DIAM. |  |  | April 08, 2016 | 2016 | Medtronic Sofamor Danek Usa, Inc - Dallas Distribution | Incorrect color coding and attachment reference were included on the labeling of 10BA40D dissecting tools with lot number H5237248. | Labeling mix-ups | A notification letter was provided to each consignee, translated as necessary, on April 8, 2016. |  | 81 | 81 tools | Distributed in France, Ireland, Netherlands, Portugal, Switzerland, and United Kingdom. | Open, Classified |  |  | 81 | Not_Computer | Not_Computer
Z-2609-2016 | 66606 | Yes | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | 1) The HomeChoice automated peritoneal dialysis system. Dataplate located on the device is labeled in part: Baxter. |  |  | October 07, 2013 | 2013 | Baxter Healthcare Corp. | In systems with version 10.4 software Initial Drain logic, the device will attempt to drain the patient to empty at the beginning of therapy (initial drain). The operator can stop, but cannot bypass, the active Initial Drain. This is to mitigate against Unintended Increased Intraperitoneal Volume (IIVP). This can cause serious problems in patients with unrelated a co-morbid condition of ascites. | Labeling design | An IMPORTANT PRODUCT INFORMATION letter dated October 7, 2013, was sent to all affected customers. The letter included instructions for clinicians to: 1) be aware that HomeChoice will attempt to drain the patient to empty at the beginning of therapy (Initial Drain); 2) review HomeChoice instructions for use for additional details on HomeChoice 10.4 Initial Drain Logic; 3) acknowledge your receipt of the notification by completing the attached Customer Reply Form and returning by fax to either 224-270-5457 or scanning and e-mailing it to fca@baxter.com; 3) if the recipient of the letter is a dealer, wholesaler, or distributor/reseller that distributed any product to other facilities, to forward the communication to all customers who may have received these products. Clinical questions should be addressed to Baxter's Renal Division Clinical Helpline at 888-736-2543, option 2 (Monday - Friday, 8:00 AM - 5:00PM CST). General questions about this recall should be addressed to The Center for One Baxter at 800-422-9837 (Monday - Friday, 8:00 AM - 5:00 PM CST). |  | 65590 | *** 1) Product Codes 5C4471 and 5C4471R: Approximately 48,600 units; *** 2) Product Codes 5C8310 and 5C8310R: Approxiamtely 16,990 units | Worldwide distribution. US nationwide including Puerto Rico; Australia, Canada, New Zealand, and Japan | Completed |  |  | 65590 | Not_Computer | Computer
Z-2137-2016 | 74078 | No | IRT | Powered heating pad. | Physical Medicine | Physical Medicine | 510(K) Exempt | HeaterBaum Foot &amp; Ankle Heating Pad, Heating Soft Compression Therapeutic Boot with Precise Temperature Control Thermostat, 110 V with AC Adapter. |  |  | January 19, 2016 | 2016 | Ergoactives, Llc | Products are marketed without a cleared 510k. | No Marketing Application | Customers were notified by phone on 01/19/2016. Two follow-up notifications were sent to customers via e-mail. Customers were asked to immediately examine their inventory for the affected product; and, to notify and identify any customers they may have further distributed to. Ergoactives LLC requests that all Heaterbaum products be sent back to their headquarters: 20725 NE 16th Avenue, Suite A-26, Miami, FL 33179, along with the order number and note of where the product was purchased. The enclosed response form should be completed and returned as well. If customers have any questions, they can call Ergoactives LLC at (305)359-5031. Customers will receive a full refund for the return of their Heaterbaum products. |  | 278 | 278 units | Worldwide Distribution -- US, including the states of AL, AR, CA, FL, GA, KS, MA, MD, ME, MI, MO, MS, NC, NE, NJ, NM, NV, NY, OH, OK, PA, TN, TX, UT, VA, and WA; and, the countries of Ecuador &amp; Venezuela. | Open, Classified |  |  | 278 | Not_Computer | Not_Computer
Z-0859-2016 | 73139 | No | LGW |  |  | Neurology | PMA | Specify 5-6-5 Lead, Specify 2x8 Lead, Model Numbers: 39565, 39286. Spinal cord stimulation leads. Part of a neurostimulation system for treating chronic, intractable pain of the trunk and/or limbs-including unilateral or bilateral pain. |  |  | January 11, 2016 | 2016 | Medtronic Neuromodulation | Some Specify 5-6-5 and 2x8 surgical lead product labeling does not adequately distinguish between permanent surgical implant and trialing/screening use. | Labeling design | Consignees were sent a Medtronic "Medical Device Correction" letter dated January 2016. The letter was addressed to Healthcare Professional. The letter described the problem and the product involved in the recall. Advised consignees to ensure their staff of the action that is being recommended. Requested consignees to complete and return the "Physician Confirmation Form". For additional information they can contact Medtronic Neuromodulation Technical services at 1-800-707-0933 weekends 7 am - 6 Pm Central time. |  | 51138 | 51,138 (44,006 US; 9,132 OUS) | Worldwide Distribution, including US Nationwide. | Open, Classified |  |  | 51138 | Not_Computer | Not_Computer
Z-0662-2016 | 72878 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Excella II Rod Cat. No.: SR 55100 5.5 x 100 mm LOT 1502 and Excella II Rod Cat. No.: SR 55110 5.5 x 110 mm LOT 1502 |  |  | December 07, 2015 | 2015 | Innovasis, Inc | Innovasis recalled two lots of Excella II titanium rods which were packaged using each other's labels. | Labeling mix-ups | Innovasis sent a Voluntary Product Recall letter, dated December 7, 2015 to customers regarding two lots of Excella II titanium rods which were packaged using each other's labels. Customers were asked to check their inventory and either remove the rod from the mislabeled package and request additional rods, or return the unopened packages immediately. The firm will replace all returned inventory upon receipt. |  | 1519 | 17 units Labeled SR55100 Lot 1502 with SR55110 rods in package | Distributed in US -- CA; KS; OH; &amp; TX. | Terminated | April 07, 2016 | 122 | 1519 | Not_Computer | Not_Computer
Z-0111-2017 | 75237 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Vanguard Distal Femoral Augment 75 X 5 RL/LM |  |  | August 09, 2016 | 2016 | Zimmer Biomet, Inc. | Contents of the package do not match the product labeling.Product is labeled: Vanguard Distal Femoral Augment 75 X 5 RL/LMCarton contains: Vanguard Distal Femoral Augment 75 X 5 LL/RMRisks:Adverse health outcomes that could occur as a result of this event are a delay less than 30 minutes and a different size augment being used to complete the procedure.As the augment will not assemble with femoral component due to its shape, it would be easily identified during surgery. | Labeling mix-ups | On 8/9/2016, URGENT MEDICAL DEVICE RECALL REMOVAL notifications were sent to the affected consignees via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. If after reviewing this notice you have further questions or concerns please call 1-800-348-9500 ext. 1251 between 8:00 am and 5:00pm EST, Monday through Friday. Calls received outside of operating hours will receive a call back the next business day. Alternatively, your questions may be sent by email to CPWARFieldAction@zimmerbiomet.com. |  | 5 | 5 | Worldwide Distribution - US Nationwide in the states of NJ, MI, FL and the countries of Sweden and Canada. | Terminated | May 17, 2017 | 281 | 5 | Not_Computer | Not_Computer
Z-1252-2016 | 73386 | No | KYF | Aqueous shunt. | Ophthalmic | Ophthalmic | 510(k) | Molteno3 Glaucoma Implants G-Series: |  |  | February 22, 2016 | 2016 | Innovative Ophthalmic Products, Inc | Innovative Ophthalmic Products (IOP) is recalling the Molteno 3 G-Series Glaucoma Implant the package insert/Instructions for Use (IFU) for the device has been updated to add a caution statement to minimize off-label use. | Labeling False and Misleading | Innovative Ophthalmic Products, Inc ( IOP ) sent an Customer Notification letter dated February 22, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken. The letter informs the customers that IOP is recalling because the package insert/Instructions for Use (IFU) for the device has been updated to add a caution statement. Specifically, the new IFU caution statement is "Caution should be taken with supra-Tenon implantation of this device as erosion of the conjunctiva is known to result from the raised edge of the oval pressure ridge."For further questions please call (202) 857-6464 |  | 278 | 278 units | Nationwide Distribution | Open, Classified |  |  | 278 | Not_Computer | Not_Computer
Z-1262-2016 | 73504 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | The E.CAM gamma camera system |  |  | February 26, 2016 | 2016 | Siemens Medical Solutions USA, Inc. | We received a customer complaint that Siemens reported via the MedWatch process (1423253-2015-00002). The result of a Hazard Analysis and CAPA investigation, we have determined that for e.cam or Symbia E systems that although the existing user manual is correct, additional supplemental instructions, with the inclusion of visual aids, may further clarify and emphasize the proper collimator change | Labeling design | Siemens sent a Customer Advisory Notice CAN 001 -2016 to all affected customers describing the recall and additional actions to take.Siemens updated user instructions, drafting customer advisory letter, and obtaining internal and external approvals, and translation of this content. Customer Advisory Notice Letter provided to each impacted consignee will document receipt of the letter and additional attempts will be made until such receipt is acknowledged. Instructions provided in letter explain that Hospital staff should inform end users of the enhancements that have been made to the user manual and continue to use your system. It also instruct end user that the advisory notice and addendum instructions should be placed with the system's instructions for use. The system use information should be available to all operators of the system.Siemens advises that if the end user has sold their equipment and it is no longer in their possession, They kindly ask that the end user forward the notice to the new owner of this equipment. They should inform Siemens about the new owner of the equipment.Adverse events or quality problems experienced with the use of this product should be reported to Siemens through the contact information provided below and may be reported to the FDA's MedWatch Adverse Event Reporting program either online, by regular mail or by fax. If the end user has any questions regarding the important notice, they should contact their local Service representative at the contact numbers provided below.-America: 1-800-888-7436-Europe, Middle East, and Africa: +49 9131 940 4000-Asia and Australia: +86 (21) 3811 2121 |  | 3090 | 3090 units | Worldwide Distribution - US Nationwide in the states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY and Puerto Rico.***** Foreign Consignees: Albania, Algeria, Angola, Argentina, Australia, Austria, Bahrain, Bangladesh, Belarus, Belgium, Bolivia, Botswana, Brazil, Bulgaria, Canada, Chile, Colombia, Costa Rica, Croatia, Czech Republic, Denmark, Dominican Rep., Ecuador, Egypt, El Salvador, Finland, France, Georgia, Germany, Ghana, Greece, Hong Kong, India, Indonesia, Iran, Ireland, Israel, Italy, Jamaica, Japan, Jordan, Kazakhstan, Kosovo, Kuwait, Libya, Malaysia, Mauritius, Mexico, Morocco, Myanmar, Netherlands, New Zealand, Nigeria, Norway, Oman, P.R. China, Pakistan, Palestinian Ter, Peru, Philippines, Poland, Portugal, Qatar, Rep. of Yemen, Republic Korea, Romania, Russian Fed., Saudi Arabia, Serbia, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Syria, Taiwan, Thailand, Trinidad,Tobago, Tunisia, Turkey, U.A.E., Ukraine, United Kingdom, Uruguay and Vietnam. | Open, Classified |  |  | 3816 | Not_Computer | Computer
Z-0112-2017 | 75018 | No | HWE | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | Orthopedic | 510(K) Exempt | Sagittal Blade 18.0X0.97X90MM, Model Number: 6118-097-090 |  |  | August 17, 2016 | 2016 | Stryker Instruments Div. of Stryker Corporation | Stryker Instruments is voluntarily recalling the Sagittal Blade 18.0X.97X90MM due to the potential for a sterile barrier breach in the product's packaging. | Packaging | Stryker Instruments is voluntarily recalling the Sagittal Blade 18.0X.97X90MM, product number 6118-097-090, lot number 16029017, due to a breach in the packaging causing a breach in sterility. There is a potential for a bone or soft tissue infection which may require medical or surgical intervention. Customer notices were issued via certified mail on 08/17/2016. Actions to be taken by the Customer/User:1.Immediately review this Recall Notification. 2.Check all stock areas and/or operating room storage to determine how many Sagittal Blades from the affected lot (listed above) are at your facility. Quarantine and discontinue use of the recalled Sagittal Blades.3.Complete the enclosed Business Reply Form (BRF) to confirm receipt of this notification and identify how many, if any, affected items are currently in your inventory. Please complete and return the BRF even if customers dont have any affected product on hand. Note: Customer signature on the BRF indicates that you received and understand this Notification and have followed the instructions in the Notification.4.If customers have further distributed this product, please forward the recall letters and the attached BRF to all affected locations. Please indicate each location on the BRF. 5.Fax the completed BRF to Stryker Instruments Regulatory Department at 866-521-2762, or scan and email a copy to kellyjo.davis@stryker.com.6.If the BRF for customer facilities indicates that recalled product is currently on hand, a FedEx label will be emailed to you. This shipping label should be used to return recalled product. 7.Upon receipt of the recalled product, replacement product will be shipped to customer accounts.For questions regarding this recall, please contact Stryker Instruments: Kelly Jo Davis, 269-389-2921, kellyjo.davis@stryker.com |  | 225 | 225 | US Nationwide in the states of RI, WA, CA, CO, WI,FL, ID, ND, MD,TXForeign: Australia, New Zealand, Netherlands, France, Italy | Terminated | March 29, 2017 | 224 | 225 | Not_Computer | Not_Computer
Z-2001-2016 | 73169 | No | LBI | Neonatal phototherapy unit. | General Hospital | General Hospital | 510(k) | Natus neoBLUE Phototherapy system, Catalogue Part Number: 010066. |  |  | April 11, 2016 | 2016 | Natus Medical Incorporated | The firm is adding the warning statement to the Instruction for Use for the neoBLUE LED Phototherapy System and provide a laminated Quick Guide to affix to each device. | Labeling False and Misleading | The URGENT: FIELD SAFETY NOTICE PRODUCT LABELING CHANGEneoBLUE LED Phototherapy System, dated April 2016, was went to US customers between April 11, 2016 and April 27, 2016. Natus distribution partners outside the US will receive the Letter to Distributors and they are requested to contact their downstream customers to distribute the information. The estimated date of notifying international consignees is May 31, 2016.Enclosed with this letter, customers would receive:1) Quick Guides (2 each) which provide operating steps, information about measuring intensity, and warnings (PN 051693)2) User manual addendum3) Customer reply formNatus requests that this information be distributed to all associated parties within your organization that use or service this product. Please note: this communication is in addition to any prior labeling changes that you may have received from Natus.Customers should submit the Reply Form via fax to 847-949-6031 or email to mailto:Natus_Quality_Programs@natus.com.If there are any questions regarding this letter and addendum, please call Natus Technical Service at 888-496-2887. |  | 13541 | 13541 units total (7840 units in the US and 5701 units outside the US) | Worldwide Distributed - US Nationwide in the US including American Samoa, Guam Puerto Rico and in the following countries:Argentina, Australia, Austria, Bangladesh, Belgium, Bermuda, Bolivia, Brunei Darussalam, Canada, China, Colombia, Costa Rica, Croatia, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, Finland, France, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Indonesia, Iraq, Ireland, Israel, Italy, Japan, Jordan, , Kenya, , Kuwait, Latvia, Lithuania, Malaysia, Moldova Republic of, Morocco, Namibia, Netherlands, Norway, Oman, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovakia (Slovak Republic), Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Uruguay, Viet Nam, Virgin Islands (British), and Yemen. | Open, Classified |  |  | 13541 | Not_Computer | Not_Computer
Z-0740-2017 | 75802 | No | OFT | Anesthesia conduction catheter. | Anesthesiology | Anesthesiology | Enforcement Discretion | BD Spinal Anesthesia Tray |  |  | November 07, 2016 | 2016 | Becton Dickinson &amp; Company | BD is conducting a voluntary product recall of the BD Spinal Anesthesia Tray, Cat (Ref) 405671, lotB01K089D since a very low number of trays are incorrectly labeled as BD Single Shot Epidural Tray,Cat (Ref) 406069. | Error in labeling | BD sent an Urgent Product Recall letter dated November 7, 2016, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to immediately review their inventory for the affected product. Return the affected products with the Business Response Card form following the instruction on the enclosed packing instruction. The completed form should be faxed to 1-866-240-5212 or email the completed form to bd6549@stericycle.com. Customers with questions were instructed to call 1-866-792-5451. For questions regarding this recall call 866-792-5451. |  | 4500 | 4500 | Nationwide Distribution | Open, Classified |  |  | 4500 | Not_Computer | Not_Computer
Z-1996-2016 | 71372 | No | GEL |  |  | General &amp; Plastic Surgery | 510(k) | EVARREST Fibrin Sealant Patch, Manufactured by: Omrix Biopharmaceuticals, Ltd, Israel. |  |  | April 11, 2016 | 2016 | Ethicon, Inc. | Ethicon is recalling EVARREST Fibrin Sealant Patch because the characters "0-bad" were printed in the expiration date field on the foil pouch labels of the Fibrin Patch units instead of the valid expiration date (2016-12-28). | Error in labeling | Ethicon sent an "Urgent: Product Recall (Removal)," dated April 14, 2016 via UPS next day mail to their affected customers. The letter identified the affected product and the reason for the recall. The letter also states that product from the affected lot is still within the expiration date, and health care practitioners, who have treated patients with the product should continue to follow patients in the usual manner. Users are to examine their inventory and quarantine the affected product. The affected product should be returned per the instructions outlined in the letter. Ethicon will provide replacement product upon receipt. Customers are to complete and return the Business Reply Form. If customers need assistance returning the product, they can contact Stericycle at 1-866-761-9472. If there are questions regarding the replacement product, customers can contact Ethicon Customer Support Center at 1-877-ETHICON. |  | 228 | 228 units | Nationwide Distribution | Terminated | August 02, 2017 | 478 | 228 | Not_Computer | Not_Computer
Z-0686-2016 | 72836 | No | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Wilson Penile Implantation System Retractor Frame, Retractor Stays &amp; Penile Elevation Strap, SINGLE USE ONLY Rx ONLY. |  |  | November 25, 2015 | 2015 | Abeon Medical Corporation | The firm was notified by some of their customers that some of the sealed pouches contain tears compromising the sterility of the product. | Packaging | On 12/17/2015 the firm sent Recall Notification Letters to their customers. The letter identified the product, problem, and actions to be taken by the customer. |  | 575 | 575 units | Worldwide Distribution-US (nationwide) including the states of AL, AR, AZ, CA, CT,FL, GA, IA, IN, KS, KY, MA, MD, MI, MN, NC, NH, NJ, OH, SC, TN, TX, UT, VA &amp; WI., and the countries of Austria, Germany, Italy, Japan, Spain, Sweden &amp; Switzerland. | Open, Classified |  |  | 4180 | Not_Computer | Not_Computer
Z-0113-2017 | 75119 | No | LYC | Bone grafting material. | Dental | Dental | 510(k) | Cytoplast Titanium-Reinforced Posterior Singles Membrane |  |  | September 07, 2016 | 2016 | Osteogenics Biomedical, Inc. | The device has the potential to be packaged in the wrong product box. | Labeling mix-ups | Osteogenics Biomedical sent an Urgent Customer Notification letter dated September 7, 2016, to all affected customers. This recall has been extended to the user level (clinician). All units sold domestically were to the end users (clinicians). End users were instructed to inspect inventory and return any affected product.For further questions, please call 1-888- 796-1923. |  | 32 | 32 boxes | Worldwide Distribution - US Distribution and to the countries of : Colombia, Venezuela/Ecuador, Japan, India, Canada, and UK. | Open, Classified |  |  | 32 | Not_Computer | Not_Computer
Z-2673-2016 | 74666 | No | ELC | Ultrasonic scaler. | Dental | Dental | 510(k) | Diamond Coated Tips and Burs are accessories to the air powered Bone drill (VARIOSURG with and without light) |  |  | July 08, 2016 | 2016 | Nakanishi Inc. | Through an error, the labeling failed to include the symbol or other text indicating that the diamond coated tips were single use only. A contributing factor to this error is that in other markets such as Japan, the diamond coated tips are multiple use and are not limited to single use only. The tips are stainless steel with a coating of fine diamond powder. The tips are sold non-sterile and are to be sterilized prior to first use. The tips were cleared as single use only.This problem was discovered during an ongoing, proactive, quality system improvement plan that was initiated by Nakanishi in October 2014. Because the correct instructions for sterilization are not available, it is possible that the devices might not be sterile prior to use. The absence of the single use statement introduces a risk of cross-contamination if the tips are reprocessed and reused. | Error in labeling | Naknishi sent out an URGENT : MEDICAL DEVICE CORRECTION letter on July 18, 2016 describing the potential safety issue regarding their Diamond Coated Products. The firm is requesting customers take the following actions: 1. Examine their inventory immediately to determine if they have any product, both used and unused, bearing any of the Nakanishi Product Codes listed in List of Products enclosed with this letter.2. Immediately dispose of any USED product according to applicable regulations.3. Segregate any UNUSED product for return to Nakanishi.4. In order for the firm to complete their records, they are requesting that customers complete the enclosed Correction Activity Sheet and FAX or Mail to: NSK America, Corp.1800 Global Parkway, Hoffman Estates, IL 60192 US Facsimile: 1-800-838-9328Customers with questions were advised to Telephone: 1-888-675-1675.NSKrecall@NSKDental.comFor questions regarding this recall call 847-843-7664. |  | 2574 | 2574 units | Nationwide Distribution to Illinois, Florida and Georgia. | Open, Classified |  |  | 37577 | Not_Computer | Not_Computer
Z-0772-2016 | 73013 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | IMPLANT,TM, MTX |  |  | January 06, 2016 | 2016 | Zimmer Dental Inc | Information specific to the 3.7 mmD and 4.1 mmD Trabecular Metal Zimmer Dental Implants was omitted from the latest revision of the applicable IFU in the Indications, Contraindications, and Precautions sections. | Labeling Change Control | Zimmer sent an Urgent Device Correction letter dated January 6, 2016 to all affected customers.. The letter identified the affected product, problem and actions to be taken. Customers were asked to do the following; review the notice and check inventory for the affected units, ff they wish to keep the product familiarize themselves with the Technical Bulletin information, if they wish to return the affected product call Zimmer Dental Customer Service at 1-800-854-7019 to obtain a Return Authorization Number and to setup a FedEx pick-up to return the products, and complete Attachment 1, the Business Reply Form.It was requested that customers maintain a copy of the noticeand a signed copy of Attachment 1 for your records. For assistance or any other questions contact Zimmer Dental Customer Service at 1-800-854-7019. Customer Service is available 7:00am to 5:00pm (Pacific), Monday through Friday. |  | 1059 | 1059 | Worldwide Distribution. US Nationwide in the states of NY, CA, KS, NJ, CO, MI, MD, WA, CT, KY, OR, OH, IA, SC, AZ, AK, IN, TN, WI, FL, AL, UT, WY, MA, PA, MN, IL, MI, GA, ND, RI, including Puerto Rico; and the countries of Thailand, France, Germany, Colombia, Hungary, Netherlands, Saudi Arabia, Poland, Spain, and Italy | Open, Classified |  |  | 11631 | Not_Computer | Not_Computer
Z-0864-2016 | 73195 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | EkoSonic Endovascular System, Catalog Number 500-56112. |  |  | January 21, 2016 | 2016 | EKOS Corporation | The EkoSonic kit, including the individually labeled MicroSonic Device (MSD) and Intelligent Drug Delivery Catheter (IDDC) are labeled with a 135 cm working length, and programmed with a 135 cm family code, however the MSD actual working length is 106 cm. | Error in labeling | All affected consignees were contacted by telephone January 20- 22, 2016.The URGENT  VOLUNTARY PRODUCT RECALL letter, dated January 21, 2016, was sent to consignees. Consignees were advised that replacement devices were shipped to their facility on January 21, 2016 and should arrive there on January 22, 2016. Consignees with questions can call: Jocelyn KerstenVice President, Regulatory and Clinical AffairsEKOS Corporation, a BTG International Group Company425-415-3132Jocelyn.kersten@ekoscorp.com |  | 10 | 10 units | Distributed in the states of AZ, CA, MA, MI, OH, and VA. | Terminated | March 15, 2016 | 54 | 10 | Not_Computer | Not_Computer
Z-1676-2016 | 73981 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Periarticular plating system, cancellous bone screw, partially threaded, hex head, 4.0 mm diameter, 16 mm length |  |  | May 04, 2016 | 2016 | Zimmer Manufacturing B.V. | A small hole may be present in the inner tray of specific lots of sterile Hip and Trauma Implant Screws. | Packaging change control | All distributors will be notified via electronic mail. Distributors with affected lots will also be notified via courier. Hospitals risk managers and surgeons will be notified via recall notification letter, dated May 4, 2016. |  | 99 | 99 | Worldwide distribution. US nationwide, Korea, Japan, Argentina, Brazil, Singapore, Taiwan, China, El Salvador, Australia, India, Hong Kong, Chile, Germany, Italy, France, Spain, UK, Sweden, Canada. | Terminated | July 05, 2017 | 427 | 11400 | Not_Computer | Not_Computer
Z-2297-2016 | 74617 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | SPHERICAL HEAD REFLECTION(R) 6.5 MM CANCELLOUS SCREW, REF 71332525, QTY: (1), STERILE R, smith&nephew; |  |  | July 06, 2016 | 2016 | Smith &amp; Nephew, Inc. | The labels correctly state that the product should be a 25mm screw, but the box contained a longer, 35mm screw. | Packaging process control | Smith &amp; Nephew sent an Urgent Product Recall Notification letter dated July 6, 2016, to all affected consignees via email and FedEx Overnight delivery on July 6, 2016. The letter identified the product, the problem, and the action to be taken by the consignee. Customers were asked to inspect their inventory and located any unused affected devices and quarantine them immediately, complete the Inventory Return Certification Form and fax to 1-901-566-7975 indicating the quantities that need to be returned. Consignees should contact Smith &amp; Nephew's Field Actions Department via email at FieldActions@smith-nephew.com to obtain a return authorization (RA) number.For questions regarding this recall call 901-399-5520. |  | 185 | 185 units | Worldwide Distribution - US (nationwide) and Internationally to AU, CA, GB, IN, ZA, and TR | Terminated | July 01, 2017 | 360 | 185 | Not_Computer | Not_Computer
Z-1234-2016 | 73337 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | BioTex Reusable Adapter Kit, Part Number 401-021-1010, packaged non-sterile in a dedicated case for steam sterilization |  |  | February 22, 2016 | 2016 | Medtronic Navigation, Inc. | Medtronic Navigation is recalling the Biotex Adapter Kit because it was commercially distributed by BioTex without a cleared premarket [510(k)] from FDA. | No Marketing Application | The firm, Medtronic, sent an "URGENT Medical Device Recall" letter dated 2/2016 to its consignees. The letter described the product, problem and actions to be taken. The identified consignees will be visited by the assigned Medtronic field representative to explain the need for the field action. The consignees are instructed to locate and remove/secure the kit from use. The Medtronic field representative will assist the consignee with identifying and removing the kit, and complete a Customer Visit record for each consignee they are assigned/visit to document the results of their visit. The consignee is instructed to sign the customer visit form presented by the representative. A copy of this record and a letter explaining the FCA will be left with the consignee.Please contact Medtronic Technical Services at 1-800-595-9709 if you have any questions regarding the subject action or the content of the letter. |  | 4 | 4 | US Distribution to states of: FL, NJ, AZ, and NC. | Open, Classified |  |  | 4 | Not_Computer | Not_Computer
Z-2600-2016 | 74855 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Large QWIX Compression screw, dia. 7.5 mm L 90 mm, Lot No. FH86, Rx only, |  |  | July 21, 2016 | 2016 | Integra LifeSciences Corp. | Labeling on the Qwix Screw product states that it is a "compression screw" instead of a "positioning screw". | Labeling mix-ups | Integra sent an Urgent Voluntary Medical Device Recall letter and acknowledgement return response form dated July 21, 2016 to their affected customers. The recall letter identifies the product and lots involved, problem and actions to be taken. All customers are asked to return the completed acknowledgement and return response form by email to FCA1@integralife.com or fax to 1-609-275-9445. Should any customer has any questions regarding the instructions, they can contact Customer Service at 1-800-654-2873. |  | 120 | 120 units from multiple lots | US Nationwide Distribution in the states of AL, GA, ID, MA, MD, MN, NC, NJ, NY, OH, OK, PA, PR, TX, VA and WA | Open, Classified |  |  | 120 | Not_Computer | Not_Computer
Z-1524-2016 | 73549 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Cervical Spine Locking Plate (CSLP) System, 4mm Cancellous Self Tapping Screw, Length 16mm Lot # 9666776 and 4mm Cortex Standard Screw, Length 16mm Lot # 9671879 |  |  | January 06, 2016 | 2016 | Synthes (USA) Products LLC | DePuy Synthes is initiating a Voluntary Medical Device Recall of lot # 9666776 and 9671879 of Cancellous and Cortex 4.0 mm Screws, which are a part of the Cervical Spine Locking Plate (CSLP) System due to discovery of a single lot of 450.136 screws (4mm Cancellous Self Tapping Screw, Length 16mm, lot 9666776) being swapped with a single lot of 450.016 (4mm Cortex Standard Screw, length 16 mm, lot | Packaging process control | DePuy Synthesis sent an Urgent Notice Medical Device Recall letter dated January 6, 2016 to customers. The letter identified the affected product, problem and actions to be taken. The letter requested customers to complete and return the attached Verification Section (page 3 of the letter). For questions call 610-719-5450 or contact your DePuy Synthes Sales Consultant. |  | 290003280 | 30 units of Lot # 9666776 and 11 units of Lot # 9671879 | Worldwide Distribution - US Nationwide in the states of NC, CA, NY, TX, NJ, CO and the countries of: Ontario and Canada. | Open, Classified |  |  | 290003280 | Not_Computer | Not_Computer
Z-0597-2017 | 75516 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Ingenuity Core Computed Tomography X-ray system, Model 728321 |  |  | October 13, 2016 | 2016 | Philips Medical Systems (Cleveland) Inc | The product label does not include the correct current rating. | Labeling Change Control | The firm will send a customer letter via certified mail and will have the receipts returned for tracking purposes. Also, a Field Service Engineer will implement the Field Change Order on affected systems to resolve the issue. Implementation of the Field Change Order will be monitored. |  | 2 | 2 | Worldwide distribution. The products were shipped to the following states: AL, AZ, CO, GA, IN, KY, LA, MA, MN, NY, OH, PA, SC, TX, and UT.There are five (5) government accounts for this recall: DEPARTMENT OF VETERANS AF, VETERANS WAY AND LAMONT ST, Mountain Home, TN 37684; US ARMY EVANS COMMUNITY H, 1650 Cochrane Cir, Colorado Springs, CO 80913; US VETERANS ADMINISTRATION, 3710 SW US VETERANS HOSPITAL, PORTLAND, OR 97239; VA MEDICAL CENTER, 4100 W 3RD ST, DAYTON, OH 45428; VAMC - PHILADELPHIA, UNIVERSITY &amp; WOODLAND AVE, PHILADELPHIA, PA 19104.There are five (5) Canadian accounts for this recall: FIVE HILLS HEALTH REGION, 55 DIEFENBAKER DR, MOOSE JAW, SASKATCHEWAN S6J 0C2; INSTITUT UNIV CARDIOLOGIE, 2725 CH STE FOY, SAINTE FOY, QUEBEC G1V 4G5; LAKERIDGE HEALTH, 1 HOSPITAL CRT, OSHAWA, ONTARIO L1G 2B9; LAKERIDGE HEALTH BOWMANVI, 47 LIBERTY ST S, BOWMANVILLE, ONTARIO L1C 2N4; MCGILL UNIVERSITY HEALTH, 1051 BOUL DÂżCARIE, MONTRÂżAL, QUEBEC H4A 0A5.There were zero (0) Mexican accounts for this recall.The recalled products were shipped to the following countries: Argentina, Australia, Belgium, China, Egypt, France, Germany, Hong Kong, India, Ireland, Israel, Italy, Japan, Republic of Korea, Lebanon, Netherlands, New Zealand, Norway, Russian Federation, Spain, Sweden, Switzerland, Taiwan, Turkey, United Arab Emirates, and United Kingdom. | Open, Classified |  |  | 117 | Not_Computer | Not_Computer
Z-1552-2016 | 73706 | No | LIT | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | ev3, PowerCross" 0.018" OTW PTA Dilatation Catheter, model: AB18W030200150, Sterile EO, Rx Only. |  |  | March 29, 2016 | 2016 | ev3, Inc. | Medtronic is conducting a voluntary Urgent Medical Device Recall of two lots of the PowerCross 0.018" Âż Over the Wire (OTW) Percutaneous Transluminal Angioplasty (PTA) Dilatation Catheter because the product is mislabeled. | Labeling False and Misleading | Consignees were sent on 3/29/2016 a Medtronic "Urgent Medical Device Recall" letter dated March 2016. The letter was addressed to Risk Manger or Health Care Professional.The letter described the problem and the product involved in the recall. |  | 27 | 27 (25 US, 2 OUS) | US: CA, GA, LA, MI, MS, MO,NV,NH. NY, OH, PA, SD, TX.OUS: SINGAPORE | Terminated | May 19, 2017 | 416 | 27 | Not_Computer | Not_Computer
Z-0795-2016 | 73053 | No | CAF | Nebulizer. | Anesthesiology | Anesthesiology | 510(k) | PharmaCaribe PulmoSal 7% Sodium chloride, Inhalation Solution, 7% Sterile, Preservative -free, Non-pyrogenic, pH 7.4 |  |  | December 29, 2015 | 2015 | Pharmacaribe llc | Vials were labeled as USP 7% Hypertonic saline instead of PulmoSal 7% (pH+) | Error in labeling | PharmaCaribe, together with its Contract Manufacturing firm Asept Pak, Inc., have initiated or completed the following actions to issued an Urgent Notice: Field Notification to Medical Device letter dated January t7, 2016. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to: Determine what is on hand and what has been distributed.Using the Return Notice provided, report this information as per the instructions on the notice. Please record units as either cases or as cartons (if the cases have been opened and are partial). This document is to be returned via mail or email as per the guidance provided. This will inform the Manufacturer of your inventory status and of your intentions regarding product replacement. Cases of this product may be returned for new cases of the same product (different lot) that have the correct embossing on the vials. Instructions for return are provided in the Field Notification Package Cover Letter. Alternatively, customers may choose to be issued a full refund for all returned product. For cases which have already been distributed to the retail/clinician level and/or are in use by patients, it is acceptable to continue to use the product, with the understanding (and proper communication to all parties) that the tabs on the vials are incorrectly labeled.Contact information:Lisa Early, QA Manager Asept Pak, Inc.Phone: 518-651-2026 Fax: 518-651-2046 E-mail: learly@aseptpak.com |  | 7020 | 7,020 vials | US Nationwide Distribution in the states of FL, LA, TX and VA | Terminated | March 28, 2017 | 455 | 7020 | Not_Computer | Not_Computer
Z-2252-2016 | 74537 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | BD Trucount Tubes: |  |  | June 24, 2016 | 2016 | Becton, Dickinson and Company, BD Biosciences | One lot of BD Trucount Absolute Counting Tubes (IVD) may not be sealed properly and exposure to ambient environment can rapidly degrade the performance of the product. | Packaging process control | Urgent Product Correction notifications letters were sent on June 24, 2016 to all affected customers. Letters identify the issue and product and provide illustrations of the Kit Box and pouch labels.Correction letters identified the issue and the product involved with photos of product labels. BD letters states that they want to reinforce the need to visually inspect the desiccant contained in each pouch . If the desiccant has turned from blue to lavender, discard the remaining tubes.The acknowledgement form should be completed and returned regardless if any inventory remains. If you require further assistance, please contact BD Customer Support at 855.236.2772 (prompt 3) in the Unites States. For customers outside the US, contact your local BD Biosciences representative or distributor |  | 844 | 328 US, 516 International | USA Distribution to the states of : NC, WA, WI, NJ, PR, MD, NY, MA, CA, NY , MD, KY, IN, VA, SD, PA, OH, AR, AZ and Foreign to : Canada, Belgium and Chile | Open, Classified |  |  | 844 | Not_Computer | Not_Computer
Z-2606-2016 | 74800 | No | LOD | Polymethylmethacrylate (PMMA) bone cement. | Orthopedic | Orthopedic | 510(k) | Cobalt HV Bone Cement |  |  | July 26, 2016 | 2016 | Encore Medical, Lp | The outer packaging was mislabeled on the box indicating "Cobalt HV with Gentamicin". The bone cement does not contain antibiotics. | Labeling mix-ups | The recalling firm sent an Urgent Field Safety Notice letter dated July 26 2016 to Sales Agents. The letter identified the affected product, problem and actions to be taken. The sales agents were instructed to pass on the notice to any organization where the potential affected product has been transferred. Agents were instructed to identify hospitals in their territory, complete a visually inspection of the affected product in the hospital's inventory, complete an Acknowledgement and Receipt form. If affected products is found contact Customer Service at 1-800-456-8696 for replacement. |  | 866 | 866 units | US Nationwide Distribution | Terminated | January 13, 2017 | 171 | 866 | Not_Computer | Not_Computer
Z-2588-2016 | 74752 | No | LWJ | Hip joint femoral (hemi-hip) metallic cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | PROFEMUR(R) MODULAR NECK, HIP FEMORAL NECK MicroPort total hip systems intended for use in total hip arthroplasty for reduction or relief of pain and/or improved hip function in skeletally mature patients. |  |  | June 30, 2016 | 2016 | MicroPort Orthopedics Inc. | Profemur(R) neck, Item Number PHA01222 (Short AR VV2 Ti Neck) was labeled as Item Item Number PHA01212 (Short AR VV1 Ti Neck). | Error in labeling | The firm initiated their recall by email on 06/30/2016. |  | 14 | 14 units | Belgium, Italy, Spain, and Poland | Open, Classified |  |  | 14 | Not_Computer | Not_Computer
Z-0654-2016 | 72944 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Zimmer Universal Locking System 3.5mm Locking Reconstruction Plate |  |  | December 09, 2015 | 2015 | Zimmer Biomet, Inc. | The product on the package label did not match what was inside the package. The label was for part number 00-4936-011-13 lot number 62968956 while the product inside the package was part number 00-4936-008-07 lot number 62968974. | Labeling Change Control | On 12/9/2015, URGENT MEDICAL DEVICE RECALL notifications were sent to the affected distributors and hospital risk managers via courier with instructions for identifying and returning the affected product. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Specifically, customers were asked to review the notification and ensure all affected personnel are aware of the information; assist the Zimmer Biomet sales representative with the quarantine and removal of any of the identified affected product; and complete and return the Certificate of Acknowledgement form. If after reviewing this notification customers have further questions or concerns please call the customer call center at 1-877-946-2761. Hours of operation are Monday through Friday, 8 a.m. through 5 p.m. EST. |  | 52 | 52 | Worldwide Distribution -- USA, including the states of MN, NY, MS, AL, OH, KY, ME, and NM; and, the countries of Taiwan, Korea, Singapore, El Salvador, Uruguay, Czech Republic, France, Italy, Poland, United Kingdom, South Africa, and Slovakia. | Open, Classified |  |  | 52 | Not_Computer | Not_Computer
Z-0298-2017 | 74809 | No | FRO |  |  | General &amp; Plastic Surgery | 510(k) | Neptune Plus (Hemostatic Pad), Part Numbers: 8870-03 |  |  | July 07, 2016 | 2016 | TZ Medical, Inc. | Neptune Plus was mis-labeled and included the statement WITH ANTIMICROBIAL BARRIER . This statement should not appear on label. | Labeling Change Control | TZ Medical notified their consignees by telephone on July 7, 2016.Customers with the boxes with the identified statement, "With Antimicrobial barrier, the firm plans to overlabel the statement on-site. If customer does not have the mis-labeled box upon contact, no action is planned. The firm also determined to overlabel devices on-hand.For questions regarding this recall call 503-639-0282. |  | 1410 | 1410 (141 boxes of 10) | Nationwide Distribution to AL, CA, FL, GA, IL, IN, KY, MO, MS, NC, NV, NY, PA, SC, TN, TX, and WV. | Open, Classified |  |  | 1410 | Not_Computer | Not_Computer
Z-0414-2017 | 75491 | No | DRS | Extravascular blood pressure transducer. | Cardiovascular | Cardiovascular | 510(k) | 4 Port "ON" Manifold (600 psi) w/4 ft. TranspacÂż IV, Item No. 46300-14NS |  |  | September 27, 2016 | 2016 | ICU Medical, Inc. | ICU Medical Inc. has identified a potential for leak with certain SafeSet Reservoir and blood sampling kits containing the Transpac IV transducers. | Device Design | ICU Medical sent an Urgent Medical Device Recall Notification letter dated September 27, 2016, to all affected customers to inform them of the expanded recall of the monitoring kits with Transpac IV transducers. Customers were informed of updated communications from August 10, 2016, when ICU Medical sent a communication indicating a potential leak with certain SafeSet Reservoir and blood sampling kits containing the Transpac IV transducers. Based on continued evaluation, ICU Medical is expanding the notification to now include certain other monitoring kits that contain the Transpac IV transducers. A small number of these devices may have the potential to leak or potentially introduce small amounts of air into the line. Customers with questions were instructed to contact ICU Customer Service Monday through Friday between the hours of 8:30 AM and 4:00 PM Pacific time: 866-829-9025 and select option 2, or e-mail the following address: productreturns@icumed.com. Customers are instructed to complete the attached Medical Device Recall Response Form and fax to (801) 264-1755 or e-mail to recall@icumed.com. Customers with questions were instructed to call 866-829-9025. |  | 20 | 20 units | Nationwide Distribution | Terminated | July 11, 2017 | 287 | 210737 | Computer | Not_Computer
Z-0411-2017 | 75494 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MEVION S250 |  |  | October 20, 2016 | 2016 | Mevion Medical Systems, Inc. | Software defect:2D projection of contours in Verity. The defect causes structures to appear off from their true position in the DRR and the radiograph with an error that increases with the distance of the structure from isocenter | Software design | Mevion issued an Important Safety Notice letter on October 14, 2016 to their customers. The letter identified the affected product, problem, containment, mitigation and actions taken. Customers were informed that Mevion will be updating software when it is released. For further Information, contact: srosenthal@mevion.com. |  | 6 | 6 units | US Nationwide Distribution in the states of: FL, MO, NJ, OH, OK | Terminated | May 15, 2017 | 207 | 6 | Computer | Computer
Z-0806-2016 | 73086 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MRIdian ViewRay Radiation Therapy System, VIEWRAY Inc. |  |  | December 24, 2015 | 2015 | Viewray Incorporated | The system loaded a completion fraction in the incorrect order after a treatment interruption. | Software design | ViewRay sent an Customer Advisory Notice dated December 24, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Workaround: If you want to move a completion plan to a later date in the Delivery Calendar then do not schedule it until right before you are ready to treat that completion plan.While we will address this issue in the next release of software we recommend that in the event of a completion plan you always review the plan name loaded on the TPDS and verify the remaining time and beams are as expected with the completion plan.If you have any questions or concerns regarding this issue, please do not hesitate to contact ViewRay Customer Support at support@viewray.com or call the ViewRay" Customer Support center at 855-286-8875. |  | 5 | 5 units | US Distribution to the states of : CA, FL, MO &amp; WI., and Internationally to Republic of Korea. | Open, Classified |  |  | 5 | Computer | Computer
Z-0392-2017 | 75475 | No | LTI |  |  | Gastroenterology/Urology | PMA | ReShape Integrated Dual Balloon System, Gen 1, Model No. 01-0013-001, Catalog No. RSM110 |  |  | October 14, 2016 | 2016 | Reshape Medical Inc | ReShape Medical is recalling the Integrated Dual Balloon System Gen 1 due to the potential leak of distillation fluid during balloon inflation. | Device Design | ReShape sent a customer notification letter dated October 14, 2016 was sent to customers to inform them that ReShape Medical is recalling the Integrated Dual Balloon System due to customer reports of instillation fluid leaking during the device inflation procedure. The letter informs the customers that any instillation fluid leakage during the inflation process will cause an incomplete balloon fill and require immediate removal of the device. Customers are instructed to contact ReShape Medical Customer Service at (844) 9377374 for any questions related to the product returns. |  | 114 | 114 units | US Nationwide Distribution in the states of: MO, TX, AZ, CA, NY, TN, GA, LA, FL, MA, OH, IN, IL | Open, Classified |  |  | 114 | Computer | Not_Computer
Z-0586-2017 | 75241 | No | NDC | Predictive pulmonary-function value calculator. | Anesthesiology | General &amp; Plastic Surgery | 510(k) | Version 1.2.0 of Accu-Chek Connect Diabetes Management App (iOS) released on July 11, 2016 |  |  | September 14, 2016 | 2016 | Roche Diabetes Care, Inc. | Roche Diabetes Care, Inc. initiated a voluntary recall for the version 1.2.0 of Accu-Chek Connect Diabetes Management App (iOS), released on July 11, 2016, due to the software containing a programming error (bug) in the Bolus Advisor feature, which could lead to an incorrect insulin bolus recommendation. | Software design (manufacturing process) | Roche Diabetes Care, Inc., initiated a voluntary recall on September 14, 2016, of the version 1.2.0 of the Accu-Chek Connect app on iOS products (iPhone, iPad). The Prescription Activation Service (PAS) which allows the users to activate the bolus advice feature was placed on product hold immediately. This measure prevented any new bolus advice activations. An updated version (V 1.2.2) of the Accu-Chek Connect Diabetes Management App for iOS that contains the fix for the issue was released on August 29, 2016. Patients were instructed to:Âż If you use the Accu-Chek Connect App on an IOS product (iPhone, iPad), please verify that you have installed the newest version of the app.Âż Contact Accu-Chek Customer Care at 1-800-628-3346 if you have additional questions.HCPs were instructed to provide a copy of the UMDC to those patients to whom they have prescribed the Bolus Advisor on the Accu-Chek Connect App. No product returns are expected. Customers with questions may contact Accu-Chek Customer Care at 1-800-628-3346. For further questions please call (317) 521-4312, |  | 7909 | 7,909 downloads of iOS version 1.2.0 of the Accu-Chek Connect Diabetes Management App in the US and 8,775 downloads Internationally | US Distribution to the states of :Alabama, Arizona, Arkansas, California, Colorado, Connecticut, DC, Delaware, Florida, Georgia, Idaho, Illinois, Indiana, Iowa, Kansas, Kentucky, Louisiana, Maine, Maryland, Massachusetts, Michigan, Minnesota, Mississippi, Missouri, Nebraska, Nevada, New Hampshire, New Jersey, New Mexico, New York, North Carolina, North Dakota, Ohio, Oklahoma, Oregon, Pennsylvania, Rhode Island, South Carolina, South Dakota, Tennessee, Texas, Virginia, Washington, West Virginia, WisconsinForeign: None provided | Terminated | June 06, 2017 | 265 | 7909 | Computer | Computer
Z-0387-2017 | 75357 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Roche COBAS INTEGRA c111 |  |  | September 27, 2016 | 2016 | Roche Diagnostics Operations, Inc. | cobas c 111 analyzers (catalog numbers 04777433001 and04528778001) with software versions up to and including 4.20 may encounter the following alarm: 7002: 108000572, A software error occurred. This alarm is generated due to a measurement timingerror. Under very rare conditions, the instrument may process two tests in the same cuvette if the runrestarts. If a used cuvette is used again result of the test(s) will be erroneous.These erroneous results may not be flagged.Falsely low or high patient results may lead to incorrect diagnostic measures and medical therapeuticdecisions. The medical risk depends on the parameter. | Software design | Roche sent an URGENT MEDICAL DEVICE CORRECTION letter dated September 27, 2016 to the affected customers via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Customers were asked to complete the attached fax form and fax it to 1-866-808-1159. For questions contact the Roche Support Network Customer Support Center at 1-800-428-2336. |  | 139 | 139 | US Nationwide Distribution | Terminated | July 21, 2017 | 297 | 139 | Computer | Computer
Z-0478-2016 | 72715 | No | KDQ | Vacuum-powered body fluid suction apparatus. | General Hospital | General Hospital | 510(K) Exempt | The Fuhrman Pleural/Pneumopericardial Drainage Set |  |  | November 17, 2015 | 2015 | Stryker Sustainability Solutions | Stryker Sustainability Solutions is recalling the Fuhrman Pleural &amp; Pneumopericardial Drainage Set due to receiving complaints that it broke off in the pleural cavity while inserting the device into the patient. | Device Design | The firm, Stryker Sustainability Solutions, sent a "URGENT MEDICAL DEVICE RECALL" letter dated 11/17/2015 to its customers. The letter identified the product, problem, and actions to be taken by the customers. Customers were instructed to discontinue the use of the product, and complete the Recall Effectiveness Check Form even if no product has been found in inventory. Customers should return the completed Recall Effectiveness Check Form to their local Stryker Sustainability Sales Representative, email to ssspfa@stryker.com, or mail to: Stryker Sustainability Solutions, 1810 West Drake Drive,Tempe, AZ 85283.If the firm indicates that affected devices remain in inventory, a prepaid shipping label will be issued for the return of the product. Customers will receive credit for all affected devices returned.Adverse reactions or quality problems experienced with the use of this productmay be reported to:- Stryker Sustainability Solutions Complaint Hotline: +1(888) 888-3433x5555- http://www.stryker.com/productexperience/- The FDAs MedWatch Adverse Event Reporting program either online, byregular mail, or by fax.If you have any questions, contact Senior Regulatory Affairs Specialist at (888) 888-3433x5506 or email: chelsea.cullen@stryker.com. |  | 34 | 34 devices | US Distribution in the states of: AK, CA, WI, IN, FL, OR, NY, and NC. | Completed |  |  | 34 | Computer | Not_Computer
Z-1624-2016 | 73816 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Hamilton G5 with software version between v2.41, v2.42, v2.50. |  |  | March 09, 2016 | 2016 | Hamilton Medical, Inc. | After performing the suctioning maneuver, including disconnecting the patient, suctioning , and reconnecting the patient, the preset pattern of ventilation many not continue as expected. | Software design | Hamilton Medical Inc. sent an" Medical Device Field Safety Corrective Action" letter dated March 7, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Hamilton Medical then had to compile a list of to whom they had distributed the affected ventilators and sent letters on April 6, 2016. Letters included the Medical Device Field Safety Corrective Action from Hamilton AG. Letters included immediate actions to be take by use and that reps will be visiting each site to upgrade ventilators to v 2.60 to alleviate the issue.Distributors will update software and will provide an updated Operator's manual as soon as it is available.Questions may be directed to 800-426-6331, Ext, 215. |  | 1115 | 1115 | US distribution only. | Terminated | June 29, 2017 | 477 | 1115 | Computer | Computer
Z-1613-2016 | 73977 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | NovaPACS  Diagnostic Viewer versions 8.3.7, 8.4.2, 8.4.3, and 8.4.4. Novarad Corporation |  |  | November 16, 2015 | 2015 | Novarad Corporation | The SUV values that are being calculated in the PET/CT fusion tool are incorrect. | Software design | Consignees were notified on 11/16/15 via letter to contact Novarad Technical Support to arrange for an upgrade to a corrected version of NovaPACS software. |  | 2386 | 2,386 | Nationwide. 3 Canadian and 33 foreign consignees. No VA/gov/military. | Open, Classified |  |  | 2386 | Computer | Not_Computer
Z-0377-2017 | 75178 | No | FZP | Implantable clip. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | AtriClip LAA Exclusion System with preloaded Gillinov-Cosgrove Clip |  |  | September 22, 2016 | 2016 | AtriCure, Inc. | There is a potential for a component failure of the device which could necessitate medical intervention. | Device Design | AtriCure sent an Urgent Advisory Notice dated September 21, 2016, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately examine their inventory and determine if they have any affected product. If so customers should ensure that all users of the affected devices are aware of the Advisory Notice. If customers further distributed the product, they should identify their customers and notify them at once of the Advisory Notice. Any units from the affected lots should be returned to AtriCure, Inc. Customers should contact AtriCure customer service at 1-866-349-2342, and select option 6 to return any affected product(s). Customers were asked to complete and return the enclosed Device Notification Acknowledgement Form as soon as possible. Customers with questions should call 513-755-4563. |  | 298 | 298 | Nationwide Distribution to the follwoing states: AL, AR, AZ, CA, CT, FL, GA, IL, KS, MD, MI, MN, MO, MS, NC, NM, NY, OH, OK, PA, SC, TN, TX, VA, VT &amp; WI.Product was also shipped to the following countries: Poland and Switzerland | Open, Classified |  |  | 298 | Computer | Not_Computer
Z-0939-2017 | 75876 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | DR Systems Unity PACS software, now known as Merge Unity PACS software. |  |  | October 15, 2015 | 2015 | Merge Healthcare, Inc. | The software fails to associate to the correct MG image if there are two images for the same view. | Software design | Merge sent an Urgent : Medical Device Recall letter dated July 11, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Be aware of the following workaround: Customer I User:Ensure both" tome" images are reviewed when more than one is captured for the same view. Merge has a released fix available for this issue.YOUR RESPONSE TO THIS NOTIFICATION IS REQUIREDPlease reply using the enclosed form and the return addressed envelope. Your response is required no later than August 2,2016. For further questions regarding this recall please call (877) 741-5369 |  | 9 | 9 sites potentially have the effected software | Distribution was made to medical facilities located in MT, CA, PA, and TX. | Open, Classified |  |  | 9 | Computer | Computer
Z-0771-2016 | 73193 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | CyberKnife Robotic; Catalog/part number 032000 and 033000 Cosmetic cover package. |  |  | January 22, 2016 | 2016 | Accuray Incorporated | The gun box mounting bracket may fail to support the weight of the gun box when in vertical (inverted) position. If this failure occurs the gun box may become loose and could come into contact with a patient. | Device Design | Accuray sent an Urgent Safety Advisory Notification letters dated January 22, 2016 to all affected customers by trackable mail method. The letters identified the affected product, problem and actions to be taken. Customers were instructed to complete and return the Acknowledgement form. For questions contact Customer Support at 1-877-668-8667 or visit the Accuray Technical Solution Center at http://www.accuray.com/services-support |  | 226 | 226 | Worldwide Distribution. US Nationwide, Venezuela, United Kingdom, Ukraine, Turkey, Thailand, Taiwan, Switzerland, Spain, Saudi Arabia, Russia, Poland, Pakistan, Myanmar, Mexico, Korea, Japan, Italy, Ireland, India, Greece, Germany, France, Finland, Czech Republic, Colombia, China, Canada, and Belgium. | Terminated | October 20, 2016 | 272 | 226 | Computer | Not_Computer
Z-0332-2017 | 75477 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Xhibit Central Station, Model 96102. |  |  | October 18, 2016 | 2016 | Spacelabs Healthcare Inc | The firm received reports of telemetry SpO2 numerics dropping off the Xhibit Central display. Desaturation, high, and low limit alarms work normally. | Software design | The firm sent the Urgent - Medical Device Correction letter, dated 18 October 2016, via priority service, return receipt requested, to all U.S. customers on 10/18/16. The firm will e-mail a customer letter (translated as necessary) to all international subsidiaries and distributors of record on October 25, 2016.Xhibit Central Stations used without Xhibit Telemetry Receivers, are not affected.Customers were asked to circulate the letter to all person affected and place a copy with the appropriate monitor user manual until corrections have been implemented. Spacelabs Healthcare will contact affected consignees to schedule a convenient time for Spacelabs to update the affected Xhibit Central Station(s) with a new software version 1.1.6 at no cost.Customers with questions can call Technical Support at 1-800-522-7025 and select 2. |  | 348 | 348 units (292 units in the US and 56 units outside the US) | Worldwide Distribution -- USA, to AL, CA, CT, FL, GA, IA, ID, IL, IN, KY, MI, MN, MS, NC, NE, NJ, NY, OH, OK, PA, RI, SC, SD, TX, VA, WA, WY, and Puerto Rico; and, to the countries of AUSTRALIA, CANADA, CHILE, CZECH REPUBLIC, FRANCE, MEXICO, NETHERLANDS, PANAMA, SAUDI ARABIA, SWITZERLAND, TAIWAN, and UNITED KINGDOM. | Terminated | April 03, 2017 | 167 | 348 | Computer | Not_Computer
Z-0581-2016 | 72763 | No | GWQ | Electroencephalograph. | Neurology | Neurology | 510(k) | Natus Quantum System with NeuroWorks Software. Catalog /Part Numbers: 013926. |  |  | November 11, 2015 | 2015 | Natus Neurology DBA Excel Tech., Ltd. (XLTEK) | During an internal testing, Natus Medical Incorporated, identified that in Neuro Works 8.1 with Quantum hardware, the incoming signal is displayed with reversed polarity. | Software Manufacturing/Software Deployment | The firm, Natus Neurology, sent an "Urgent Medical Device Correction Notification" letter on 11/11/2015. The letter described the the problem, the product involved in the recall and actions to be taken. Natus will provide further information once an updated software is released. The consignees were instructed to complete and return the confirmation form by mail to Senior Manager Quality and Regulatory Affairs, Excel-Tech, Ltd (XLTEK) Division, Natus Medical Incorporated, 2568 Bristol Circle, Oakville, ON L6H 5S1 or by email to :ots@natus.com.If you have any questions, contact Senior Manager Quality and Regulatory Affairs at 905-287-5055 or email: sanjay.mehta@natus.com. |  | 56 | 56 (43 US, 13 OUS) | Worldwide Distribution: US Distribution including states of: AL, AZ, AR, CO, FL, GA, MD, MA, MN, MS, NJ, NY, OH, SC, TN, TX and UT; and countries of: INDIA, KOREA, SPAIN, SWITZERLAND, THAILAND, UNITED KINGDOM, and CANADA.. | Open, Classified |  |  | 56 | Computer | Computer
Z-0348-2017 | 75271 | Yes | DWC | Cardiopulmonary bypass temperature controller. | Cardiovascular | Cardiovascular | 510(k) | Heater Cooler Unit |  |  | September 27, 2016 | 2016 | Maquet Cardiovascular Us Sales, Llc | There is the potential for contaminated water to enter parts of the Heater-Cooler Unit circuits. There is a potential risk of bacterial contamination including Mycobacteria species. | Device Design | Maquet issued an Urgent Medical Device Recall letter/Return Response form dated September 27, 2016 to affected customers. The letter identified the affected product, problem and actions to be taken. Customer were instructed to complete the return response form via fax to 1-973-629-1518. For questions contact Maquet Customer Service at 1-888-627-8383 option 2 followed by option 2 Monday through Friday between 8 amd and 6 pm. EST. |  | 305 | 305 units | Worldwide Distribution - US Nationwide | Open, Classified |  |  | 305 | Computer | Not_Computer
Z-1933-2016 | 73921 | No | CHH | Cholesterol (total) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Cholesterol, Catalog No. OSR6x16 |  |  | April 13, 2016 | 2016 | Beckman Coulter Inc. | Beckman Coulter has identified that Dipyrone (Metamizole) may cause erroneous, false low results in the following assays: Cholesterol OSR6X16, Triglyceride OSR6X118, Uric Acid OSR6X98, Lactate OSR6X93, and Lipase OSR6X30. | Device Design | Beckman Coulter sent an Urgent Medical Device Recall letter dated April 13, 2016 to affected customers. The letter identified the affected product, problem and actions to be taken. Customers are instructed to complete and return the enclosed Response Form within 10 days. Customers with questions regarding the customer notification letter are instructed to contact Customer Support Center at http://www.beckmancoulter.com/customersupport/support; (800) 223-0130 in the US and Canada. Outside of the US and Canada, customers are instructed to contact their local Beckman Coulter Representative. |  | 43877 | 43,877 units total (10,496 units in US) | Worldwide Distribution - US Nationwide including Puerto Rico and the countries of: Albania, Algeria, Australia, Austria, Bahrain, Bangladesh, Belarus, Belgium, Bolivia, Bosnia-Herzegovina, Botswana, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, El Salvador, Eritrea, Ethiopia, Finland, France, Germany, Ghana, Gibraltar, Greece, Guadeloupe, Guam, Honduras, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Jordan, Kazakhstan, Kenya, Korea, Kuwait, Lebanon, Lesotho, Liberia, Libyan Arab Jamahiriya, Lithuania, Malawi, Malaysia, Mexico, Morocco, Mozambique, Myanmar, Namibia, Netherlands, New Zealand, Nigeria, Oman, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Qatar, Romania, Russian Federation, Saudi Arabia, Serbia, Sierra Leone, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Venezuela, Vietnam, Zambia, and Zimbabwe. | Open, Classified |  |  | 43877 | Computer | Not_Computer
Z-0309-2017 | 75358 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | Marathon Flow Directed Micro Catheter, UltraFlow HPC Flow Directed Micro Catheter |  |  | October 05, 2016 | 2016 | Micro Therapeutics Inc, Dba Ev3 Neurovascular | Delamination and detachment of the PTFE (polytetrafluoroethylene) coating material may lead to PTFE coating in the blood stream. | Device Design | Medtronic sent an Urgent Medical Device Recall letter dated October 5, 2016, to all affected customers to inform them that Medtronic has identified the potential for an issue with a specific set of lot numbers of Medtronic Neurovascular products as noted where the PTFE (polytetrafluoroethylene) coating could delaminate and separate from the delivery wire. Customers are instructed to complete the attached customer confirmation certificate and fax it to Medtronic at (949) 434-5020 to the attention of Neurovascular Quality.On October 14, 2016, Medtronic announced that it has notified customers of a voluntary recall of certain lots of its Pipeline embolization device, Alligator retrieval device and X-Celerator hydrophilic guidewire. The recall also includes the stylet containing UltraFlow flow directed micro catheters and Marathon flow directed micro catheters. This voluntary recall is being conducted due to the potential separation and detachment of the polytetrafluoroethylene (PTFE) coating on parts of these devices. Should the PTFE separate from the delivery wire or stylets, PTFE particulate could enter the blood stream of the patient. PTFE in the blood stream, based on the size and quantity, could lead to a thromboembolic event.Customers were asked to complete the attached Customer Confirmation Certificate and fax it to Medtronic at 1-949-434-5020 to the attention of Neurovascular Quality. Customers with questions were instructed to contact their Medtronic representative. |  | 84278 | 84,278 units total (17,811 in US) | Worldwide Distribution - US (nationwide) and Internationally to Argentina, Australia, Austria, Bahrain, Bangladesh, Belgium, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, Finland, France, Germany, Greece, Hong Kong, Hungary, Iran, Ireland, Israel, Italy, Japan, Korea, Latvia, Lithuania, Malaysia, Mexico, Nepal, Netherlands, Nicaragua, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Puerto Rico, Russian Federation, Saudi Arabia, Singapore, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Uruguay, Venezuela, and Vietnam. | Open, Classified |  |  | 84278 | Computer | Not_Computer
Z-0768-2016 | 72974 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Optima XR220amx, Mobile Digital Ready Radiographic System |  |  | December 23, 2015 | 2015 | GE Medical Systems, LLC | Potential safety latch mechanism failure. The safety latch, within the column that supports the horizontal arm, is a backup in the unlikely event the primary safety cable breaks. A fall of the arm assembly could result in bodily harm to a person. | Device Design | Consignees were sent on 12/23/2015 a GE Healthcare "Urgent Medical Device Correction" letter GEHC Ref# 10906 dated December 18, 2015. The letter was addressed to Director of Biomedical Engineering, Director of Radiology and Chief of Cardiology. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. Customers may continue to use the system. The firm will correct the affected devices and no charge to the customer; and a GE Healthcare representative will contact customers to arrange for the correction. For questions or concerns regarding this notification, please contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 2389 | 2389 | Worldwide Distribution -- US, including Washington D.C. Puerto Rico, Virgin Islands, and Guam; and the countries of Afghanistan, Albania, Anguilla, Antigua and Barbuda, Armenia, Australia, Austria, Bahamas, Bahrain, Belgium, Bermuda, Bolivia, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Czech Republic, Denmark, Dominican Rep., Ecuador, Egypt, Ethiopia, Finland, France, Germany, Greece, Guadeloupe, Guatemala, Guyana, Honduras, Hong Kong, Hungary, Iceland, India, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Korea (Republic of), Kuwait, Lebanon, Malaysia, Martinique, Mexico, Netherlands, New Zealand, Nigeria, Norway, Pakistan, Panama, Papua New Guinea, Paraguay, Philippines, Poland, Portugal, Romania, Russian Federation, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Taiwan, Thailand, Trinidad , Turkey, United Arab Emirates, United Kingdom, Uruguay, Venezuela, and Vietnam. | Open, Classified |  |  | 3410 | Computer | Not_Computer
Z-0413-2017 | 75456 | No | KRG | Pacemaker programmers. | Cardiovascular | Cardiovascular | PMA | The ORCHESTRA PLUS Programmer |  |  | September 29, 2016 | 2016 | Sorin Group USA, Inc. | Sorin Group USA, Inc. announces a voluntary field action for the ORCHESTRA Programmer and ORCHESTRA PLUSProgrammer because the residual longevity estimate may be less accurate. | Software Design Change | The firm, LivaNova, sent an "Urgent Field Safety Notice" dated 9/29/2016 to its customers. The notice described the product, problem, actions to be taken. The notice states that a Sorin representative will contact customers to upgrade the software in the facility's programmer.If customers have any questions, they are urged to call their LivaNova representative or contact LivaNova Technical Services at 877.663.7674. |  | 586 | 586 programmers | Programmers were distributed nationwide, to VA/govt/military, Canadian, and other foreign consignees. | Open, Classified |  |  | 1433 | Computer | Computer
Z-0794-2016 | 72978 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | RadSuite, a picture archive and communications system. Model Numbers Versions 5.30.0, 5.30.1, 5.30.2, 5.30.2 SB0344, 5.30.3, 5.30.4, 5.30.4 HF0418E, 5.30.5, 5.30.5 HF0404E, 5.30.5 HF0413E, 5.30.5 HF390E, 5.30.5 SB0368, 5.30.6, 5.30.6 HF0405E, 5.30.6 HF0406E, 5.30.7, 5.30.7 HF0429E, 5.30.7 HF0441E, 5.30.7 HF0446E, 5.30.7 HF0470E, 5.30.7 HF0473E, 5.30.8, 5.30.8 HF0477E, 5.30.8 HF0490E, 5.30.8 HF0495 SB0508E, 5.30.8 HF0495E, 5.30.8 HF0495E SB504E, 5.30.8 HF0550, 5.30.8 SB0485E, 5.30.8 SB0486E, 5.30.X SB0381E, 5.35.1, 5.35.1 HF0487E, |  |  | December 17, 2015 | 2015 | Merge Healthcare, Inc. | Potential incorrect Standardized Uptake Values (SUV) measurements in RadSuite. | Software design | Merge Healthcare sent an " Urgent - Medical Device Recall letter dated January 21, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer .If you may have further distributed this product, please identify your customers and notify them at once of this product recall. Your notification to your customers may be enhanced by including a copy of this recall notification letter. Merge Healthcare is committed to improve efficiencies and enhancing the quality of healthcare worldwide. If you have any additional questions, call Merge Customer Service at (877) 741-5369 or support@merge.com. |  | 164 | 164 | Nationwide Distribution to the states of : MN, WI, IL, PA, AL, MO, MT, AZ, CA, OH, FL, SC, MI, TN, CO, NV, IN, NY, MD, AR, KS, NY, IN, IA, TX, OK, NC, WA and DC. | Open, Classified |  |  | 164 | Computer | Not_Computer
Z-0361-2017 | 64005 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Symbiq One Channel Infuser; an Rx medical device infusion pump used to administer I.V. fluids. List number 16026. Note that Hospira retired this pump June 30, 2015. |  |  | March 16, 2011 | 2011 | Hospira Inc. | May go to a white screen due to a problem with the SOM2 module. A white screen is a diagnostic screen displayed to the user as part of the safety shutdown procedure of the device. The infusion and the processor stop, sound a 3 minute audible alarm from the backup buzzer and deactivate the couch screen and the automated loading/unloading of cassettes. Delay or interruption of therapy. | Device Design | Hospira sent an Urgent Device Recall Letters, dated February 5, 2013, to all affected customers. Customers were directed to notify their consignees if they further distributed the product and confirm they have done so by return of the reply form. Healthcare professionals were to contact Stericycle. Hospira discontinued support of these devices as of June 30, 2015. Customers with questions were instructed to call 1-800-201-9064.For questions regarding this recall call 224-212-4892. |  | 10337 | 10,337 | Worldwide Distribution - US (nationwide) and Canada | Open, Classified |  |  | 21234 | Computer | Computer
Z-0399-2017 | 75419 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Merge LIS software. |  |  | March 05, 2015 | 2015 | Merge Healthcare, Inc. | There is a potential for duplicate container numbers to be created for patients. | Software design | The recalling firm issued a letter dated 1/27/2016 via e-mail on 1/30/2016. If no email address was available, the letters were issued via certified mail. Letters to nonresponders were issued 5/26/2016. |  | 413 | 413 sites potentially have the affected software | Distribution was made nationwide and to the Virgin Islands. Foreign distribution was made to the Bahamas. | Open, Classified |  |  | 413 | Computer | Not_Computer
Z-0582-2016 | 72794 | Yes | MSX | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Remote Patient Monitoring System, MDP 2000 Series, Part Numbers: MDP2040-0100, PGM340R3, or PGM340R4. |  |  | November 08, 2015 | 2015 | Vidco, Inc. | Testing at customer site showed unit Remote Patient Monitoring System MDP2040-0100 in a continuous trap condition, not allowing system to reset and reboot. Two customers recently complained of the system freezing and it could only be restarted if the user re-applied power. | Device Design | The firm began notifying consignees on Nov. 8, 2015 and followed up with the Advisory Notice letter, dated November 10, 2015.VIDCO identified the root cause and has corrected software that is currently going through validation. This software will be made available for updating the MDP2040-0100 as soon as it is released. The units may be updated on-site with the software program that will become available on the File Transfer Protocol (FTP) site. If you are unable to perform the procedure yourself, VIDCO will assist you in this effort.The current software version can still be used for monitoring as long as the user confirms the displays are still dynamically writing waves to the screen (i.e. waves will be changing in real-time).Consignees with questions can call (503) 641-1804 or 800-638-4326, or send email to sales@vidcoinc.com |  | 59 | 59 units | US Nationwide distribution in the states of AZ, CA, MD, NM, NJ, and OH. | Open, Classified |  |  | 59 | Computer | Computer
Z-0410-2017 | 75498 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MEVION S250 |  |  | October 21, 2016 | 2016 | Mevion Medical Systems, Inc. | Recovery of Incorrect Isocenter -operational change | Software Design Change | Mevion issued an User Notice letter on October 24, 2016 to their customers. The letter identified the affected product, problem and actions taken. Customers were informed that Mevion will be updating software when it is released. For further clarification or assistance, please contact Rob Cessac by email at rcessac@mevion.com or by phone at (978) 540-1501. |  | 6 | 6 units | US Nationwide Distribution in the states of FL, MO, NJ, OH and OK, | Terminated | May 10, 2017 | 201 | 6 | Computer | Not_Computer
Z-0349-2017 | 73286 | No | IRT | Powered heating pad. | Physical Medicine | Physical Medicine | 510(K) Exempt | Model 055: Thermophore Classic, Professional Quality, Deep-Heat Therapy, Model 055 Large 14' x 27", |  |  | September 15, 2016 | 2016 | Battle Creek Equipment Co. | Products manufactured with the SPT cord have a known failure mode of sparking,causing smoke, fire and burns. | Device Design | All consumers for which Battle Creek has contact information and that purchased a recalled pad will be contacted via first class mail. Consumers will be asked to contact us by phone after verifying that the model number and lot code is part of the recall. Battle Creek plans to add information about the recall on its web site including a pdf of the consumer letter. All customer consignees (non-consumer) will be informed of the recall. Email will be the most effective method , and mail and telephone calls will be used as needed. |  | 99060 | 99,060 | Distributed within the United States, Canada, Mexico, Puerto Rico and Columbia | Open, Classified |  |  | 99060 | Computer | Not_Computer
Z-0805-2016 | 73092 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Eclipse Treatment Planning System versions 11, 13.0, 13.5 and 13.6; |  |  | January 11, 2016 | 2016 | Varian Medical Systems, Inc. | When using PBC 11.0.31 to calculate the dose for a conventional arc field with more than 100 segments for Eclipse versions 11.0, 13.0, 13.5 or 13.6, the displayed dose does not correspond to the calculated Monitor Units (MU). Potential for unintended radiation exposure. | Software design | Varian sent an Urgent Medical Device Correction letters dated January 11, 2016 to all affected customers. Letters were also sent to Varian Sales and Marketing and Service organizations to notify them of the recall. The letter identified the affected product, problem and actions to be taken. For questions customers were instructed to contact their local Varian Medical Systems Customer Support District or Regional Manager. |  | 9499 | 9499 | Worldwide Distribution - US Nationwide and the countries of Albania, Guadeloupe, Norway,Algeria, Guam, Pakistan, Argentina, Guatemala, Panama,Australia, Hong Kong, Paraguay, Austria, Hungary, Peru,Azerbaijan, Iceland, Philippines, Bangladesh, India, Poland,Belgium, Indonesia ,Portugal, Belorussia, Iran, Puerto Rico,Brazil, Iraq, Qatar, Brunei Darussalam, Ireland, Reunion,Bulgaria, Israel, Romania, Burma, Italy, Russia,'Canada, Jamaica, Saudi Arabia, Chile, Japan, Singapore, China, Kazakhstan, Slovakia, Colombia, Kuwait, South Africa, Costa Rica, Latvia, South Korea,Czech Republic, Lebanon, Spain, Denmark, Libya, Sweden, Dominican Republic, Lithuania, Switzerland, Ecuador, Luxembourg, Taiwan, Egypt, Macau, Thailand, El Salvador, Macedonia, Trinidad and Tobago, Estonia, Madagascar, Tunisia, Finland, Malaysia, Turkey, France, Malta, Ukraine, French Polynesia, Mexico, United Arab Emirates, Georgia, Monaco, United Kingdom, Germany, Morocco, Venezuela, Ghana, Netherlands, Vietnam,Greece, New Zealand, and Zimbabwe. | Open, Classified |  |  | 9499 | Computer | Computer
Z-0814-2016 | 73093 | No | BRY | Anesthetic cabinet, table, or tray. | Anesthesiology | Anesthesiology | 510(K) Exempt | Jadak Barcode Scanner, Part No. 350587-32, 350587-41, 351377-32. |  |  | January 15, 2016 | 2016 | CareFusion 303, Inc. | Customers reported issues when scanning medications with the scanners. When a user scans a medication using the affected barcode scanner, information on a different medication could be communicated to the Pyxis product. | Device Design | BD sent an Urgent Medical Device Recall Notification letter dated January 15, 2015 to all affected customers The letter identified the affected product, problem and actions to be taken. Customers are instructed to complete the enclosed Customer Response Card by selecting the desired deployment option and return the card to CareFusion. Customers with questions are instructed to contact the CareFusion Support Center at (888) 562-6018. Customers with adverse event reports are instructed to contact CareFusion Customer Advocacy at (800) 836-7842. Customers with technical questions are instructed to contact CareFusion Technical Support at (800) 727-6102. |  | 9345 | 9,345 units | Worldwide:Distribution - US Nationwide and the countries of Saudi Arabia, Bahrain, Australia, United Arab Emirates, Qatar, Mexico, Guam and the Bahamas. | Terminated | May 08, 2017 | 479 | 9345 | Computer | Computer
Z-0398-2017 | 75394 | No | MRG | Vitamin D test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Siemens ADVIA CentaurÂż Vitamin D Total Assay |  |  | October 05, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc | Vitamin D Total- Change in Correlation Between Serum and Plasma Specimen Tubes- negative bias with plasma vs. serum samples | Device Design | Siemens issued an Urgent Medical Device Correction (UMDC) was sent to all affected Siemens Healthcare Diagnostics customers in the United States on October 5, 2016 to be delivered to customers on October 6th, 201 6. An Urgent Field Safety Notice ~UFSN) was sent out for distribution to all customers outside the United States on October 5, 2016. These notices inform the customer ofthe change in correlation between serum and plasma specimen tube results generated on the ADVIA Centaur/ADVIA Centaur XP/ADVIA Centaur XPT Vitamin D Total assay. A revised letter was issued on November 7, 2016 by Fedx for delivery on November 8, 2016 and advises users to: In the interim, Siemens requires that customers discontinue use of plasma tubes for specimen collection for the ADVIA Centaur Vitamin D Total assay until further notice.Customers can continue to use the ADVIA Centaur Vitamin D Total assay with serum specimen collection tubes.Accounts to Complete and return the Field Correction Effectiveness Check Form attached within 30 days. Review the letter with your Medical Director.and retain this letter with your laboratory records, and forward this letter to those who may have received this product.Question contact your Siemens Customer Care Center or your local Siemens technical support representative. |  | 37437 | 37437 kits | NationwideForeign:Afghanistan, Algeria, Argentina, Australia, Austria, Bahrain, Bangladesh, Belgium, Brazil,Canada, Chile, China, Colombia, Croatia, Cyprus, Czech Republic, Denmark, Ecuador, Egypt,Finland, France, French Polynesia, Georgia, Germany, Great Britain, Greece, Guadeloupe,Guatemala, Hong Kong, Hungary, India, Indonesia, Iraq, Italy, Japan, Jordan, Latvia, Lebanon,Lithuania, Luxembourg, Macedonia, Malaysia, Malta, Martinique, Mexico, Nepal, Netherlands,New Zealand, Norway, Oman, Pakistan, Paraguay, Philippines, Poland, Portugal, Qatar,Romania, Russian Federation, Saudi Arabia, Singapore, Serbia, Slovakia, South Africa, SouthKorea, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Turkey, United Arab Emirates,Uruguay, Vietnam | Open, Classified |  |  | 107218 | Computer | Not_Computer
Z-0360-2017 | 74242 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | SureT G29 6mm, 18", Paradigm; SureT G29 6mm 23"; SureT G29 6mm, 23", Paradigm; SureT G29 6mm 32"' SureT G29 6mm, 32", Paradigm; SureT G29 8mm, 23"; SureT G29 8mm, 23", Paradigm; SureT G29 8mm, 32"; SureT G29 8mm, 32", Paradigm; SureT G29 10mm, 32", Paradigm |  |  | November 07, 2014 | 2014 | Unomedical As | Tubing detachment | Device Design | ConvaTec sent an Urgent Medical Device Correction letter to all affected consignees on November 7, 2014. The letter notified the consignees of what action they need to take if they experience a detached tubing. You can continue to use your infusion sets by following the advice below.1. When changing your infusion set, closely follow the instructions for use included in the product box. Check the tubing at the connect/disconnect location identified in the drawings above to make sure it is not loose.2. As always, it is essential to monitor your blood sugar levels frequently using your blood glucose meter. Proactively check your tubing connections occasionally throughout the day to ensure tubing is secure. It is especially important to check your blood sugar and tubing connections at bedtime to confirm insulin delivery is occurring.3. If you experience a high blood sugar, check your tubing connections and infusion site closely to ensure your tubing is secure.If you discover the tubing is detached:1. Do not attempt to reattach the tubing. Replace the infusion set immediately.2. Treat any high blood sugar based on guidelines provided by your healthcare professional.3. Call your distributor's Helpline to report the issue. They will give you instructions on how to return the affected infusion set.The letter included a reply card to be returned. |  | 4748000 | 4,748,000 units | Worldwide Distribution - US including CA, IN, PA, and Internationally to Canada, Korea, and Japan. | Completed |  |  | 17165789 | Computer | Not_Computer
Z-1929-2016 | 74265 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ExacTrac Vero is a Patient Positioning System for Radiation therapy. |  |  | May 04, 2016 | 2016 | Brainlab AG | Potentially incorrect positioning when using Implanted Marker Detection with Brainlab ExacTrac Vero 3.5 | Device Design | A FIELD SAFETY NOTICE / PRODUCT NOTIFICATION was sent to customers on May 4, 2016 and distributed via email. An active reply by the consignees will be followed, if necessary via further different communication media.As of now for new installations- the Product Notification letter is handed over to the customer during installation (before clinical use) by the according Brainlab Service representative,- a signature is required by the customer that the contained User Corrective Action will be implemented before clinical use by the facility (to be adhered to by the users).Brainlab Instructions and actions required to be performed by the customer are described below.1. For Positioning and Verification, carefully review plausibility of implant marker fusions resulting in rotation angles of more than 10Âż. If in doubt consider the following options before proceeding:- Deselect one or several markers defined in the CT data set, if you suspect these markers to have migrated- Apply bony fusion instead of marker based fusion- Apply ConeBeam CT correction and verification instead of marker based fusion- Acquire a new CT scan and repeat treatment planning and definition of implanted marker positions based on the new CT data. For further details regarding the suggested proceedings, please consult the according Clinical User Guide. 2. For Dynamic Tracking continue to follow the instructions of the Clinical User Guide regarding the review of the correlation model based on implanted marker detections. Specifically: - Examine the correctness of localization of markers and targets in the X-ray sequence.- Verify the consistency of the Target Detection Curve and/or Correlation Model Curve. Irregularities of the curves are an indication for incorrect implanted marker detection. In this case please consider the options described in 1. before proceeding.If the customers require further clarification, they should contact their local Brainlab Customer |  | 27 | 27 Systems | Distributed in the states of Florida, New York, Texas and Ohio, and in the countries of Belgium, France, Germany, Italy, Japan and South Korea. | Open, Classified |  |  | 27 | Computer | Computer
Z-0409-2017 | 75496 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MEVION S250 |  |  | October 31, 2016 | 2016 | Mevion Medical Systems, Inc. | Completing Setup Without Program Move Complete Can Cause An Isocenter Error | Software design | Mevion issued an Important Safety Notice letter on October 31, 2016 to customers. The identified the affected product, problem, mitigation and advisory information. The letter informed customers that Mevion will release and install a software patch that removes the reliance on tracking isocenter changes for the correct reporting of the Couch position to the user displays and to the OIS. For further information mailto:srosenthal@mevion.com |  | 6 | 6 units | US Nationwide in the states of FL, MO, NJ and OK | Terminated | June 29, 2017 | 241 | 6 | Computer | Computer
Z-1648-2016 | 73669 | No | LZH | Infusion pump. | General Hospital | General Hospital | 510(k) | Covidien Kangaroo Connect Enteral Feeding Pump |  |  | April 01, 2016 | 2016 | Medtronic | Kangaroo Connect Feeding Pump Occlusion alarms fail to alarm | Device Design | Medtronic(formerly) Covidien notified accounts via Urgent Medical Device Recall letter, dated 4/1/16, that specific serial numbers with software versions 1.12 and 1.13 have a software anomaly. The firm shipped replacement units with software V 1.14. Affected devices must be returned to the firm once the replacement units have been received. A Recalled Product Return Form required to be completed and returned. Questions or concerns, contact Medtronic representative or Customer Service at (800) 882-5878. |  | 223 | 223 units | Worldwide Distribution -- USA, Australia, Canada, France, and Singapore. | Open, Classified |  |  | 368 | Computer | Not_Computer
Z-0793-2016 | 73084 | No | NBH | Arthroscope. | Orthopedic | Orthopedic | 510(K) Exempt | Smith &amp; Nephew BEAVER Blade 4.0mm |  |  | January 11, 2016 | 2016 | Smith &amp; Nephew, Inc. | Blade exhibiting fracture during hip arthroscopy procedures | Device Design | Smith &amp; Nephew issued letters via Federal Express on January 11, 2016. requested to inspect inventory and locate any unused devices and quarantine them immediately.Complete Inventory Return Certification Form on the following page, indicating the quantities that need to be returned.Contact Smith &amp; Nephews Field Action Department via email at FieldActions@smith-nephew.com or fax +1-901-566-7975 to obtain a return authorization (RA) number. |  | 69043 | 69043 units | Worldwide distribution. US Nationwide and countries of AT, AU, BE, CA, CH, CL, CN, CZ, DB, DE, DK, EE, ES, FI, FR, GB, HK, IE, IL, IN, IT, JP, LT, LV, NL, NO, PL, PR, PT, RO, SE, SG, SI, TH, TR, and ZA. | Open, Classified |  |  | 69043 | Computer | Not_Computer
Z-0669-2017 | 75370 | Yes | NIP |  |  | Cardiovascular | PMA | MISAGO RX Self Expanding Peripheral Stem |  |  | August 04, 2016 | 2016 | Terumo Medical Corp | During recent routine testing, the MISAGO Peripheral Stent was found not to conform to product specification related to stent diameter at the proximal/distal ends and/or the shape of the stent. | Device Design | Terumo Medical Corporation issued an Urgent Product Recall Bulletin dated August 4, 2016, via Fed Ex.Customers must immediately discontinue use and quarantine any available stock of the Misago peripheral stent to prevent further use. The customers should complete the enclosed Inventory Confirmation form as soon as possible and email or fax the completed formto misago.rc@terumomedical.com or (734) 330-2034, Attn: MISAGO Recall.Customers are encourage you to contact Terumo Medical Corporation with any questions or concerns-Terumo Medical Corporation Quality Department: 1.800.283.7866, Press 4 for list of departments, Press 7 for Quality Assurance; Monday  Friday, 8 a.m.  5 p.m. ET- Email Quality Assurance: misago.rc@terumomedical.com- Customer Care: 1.800.888.3786- Contact your local Terumo representative |  | 7965 | 7,965 units | Nationwide | Open, Classified |  |  | 7965 | Computer | Not_Computer
Z-0764-2016 | 72571 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Philips Allura Xper FD20 Ceiling version only; 722028; imaging applications. |  |  | November 03, 2015 | 2015 | Philips Electronics North America Corporation | Normally, movement of the C-arc is initially stopped by the motor (holding torque) then by the electromagnetic brake, leading to a full stop of the C-arc. Because of improper adjustment and tolerances, when the C-arc is initially stopped in an unbalanced position, the air gap of the brake can become too large, resulting in continued C-arc movement (C-arc is looking for balanced position). | Device Design | The letters for the customers in the United States were sent on 11/3/2015, via certified return receipt and were addressed to the Risk Manager.Philips will perform a test on the 15 affected Allura Xper R8.2 FD 20C systems to see if the motion controller amplifier of the brake motor works correctly. If the test will not pass, the motor will be replaced. The Recall Package is sent to all Key Markets worldwide. The Key Market of an affected country will contact the facility where these systems were installed. |  | 15 | 15 subject to correction and removal | Worldwide Distribution. US states of TN and IN and the countries of Australia, Belgium, Egypt, France, Germany, Japan, Republic of Korea, Netherlands, and Saudi Arabia. | Open, Classified |  |  | 17 | Computer | Not_Computer
Z-1614-2016 | 73849 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | NeuroBlate System SideFire Select 2.2 mm Directional Laser Probe, Sterile EO. Model #s SFS000-01,SFS002-01, SFS122-01, SFS222-01, SFS322-01, SFS422-01, and SFS522-01. |  |  | March 25, 2016 | 2016 | Monteris Medical Corp | Monteris received a report that the sapphire lens on a 2.2SF probe fractured during a LITT procedure, with resulting introduction of C02 into the patient's brain. | Device Design | Monteris Medical Corporation field personnel visited each consignee starting March 25, 2016 and April 1, 2016. The devices were physically removed from each site. Customers from whom the devices were removed will either receive a credit or receive replacements of other models of laser delivery probes. |  | 121 | 121 | Nationwide Distribution in the states of KS, MO, NC, OH, and VA. | Terminated | March 10, 2017 | 350 | 121 | Computer | Not_Computer
Z-0750-2017 | 75786 | No | POA |  |  | Microbiology | EUA - Emergency Use Authorization | LightMixÂż Zika rRT-PCR Test |  |  | November 16, 2016 | 2016 | Roche Molecular Systems, Inc. | The LightCyclerÂż 480 algorithm used for the LightMixÂż Zika rRT-PCR Test, EUA (catalog number07987897001) occasionally (~1% of the time) calls a positive result for a negative sample if there isan inflection point in the fluorescence signal. | Software design | Roche Molecular Diagnostics Inc. sent an Urgent Medical Device Correction Letter dated November 16, 2016, to all affected customers via UPS. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use of the affected product until Roche provides updated Instructions for Use to assist in result interpretation. If product was further distributed, ensure the Urgent Medical Device Correction letter is provided to that site. Also, customers were instructed to completed the attached fax form and fax it to 1-888-912-2188. For questions customers should call 1-800-526-1247. |  | 24 | 24 pieces | Nationwide Distribution | Open, Classified |  |  | 24 | Computer | Computer
Z-0745-2017 | 75252 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | BD 23 G x .75 in. BD VacutainerÂż Push Button Blood |  |  | August 08, 2016 | 2016 | Becton Dickinson &amp; Company | BD hasreceived multiple reports of units associated with the 23 G x .75 in. BD VacutainerÂż Push Button BloodCollection Set with 12 in. Tubing whose lots are exhibiting separation offront and rear barrels upon activation of the safety feature, which retracts theneedle. | Device Design | The firm, BD, sent a "URGENT PRODUCT RECALL" letter dated 8/8/2016 to its customer with a return response form via UPS 2nd Day mail on 8/8/2016.The letter described the product, problem and actions to be taken. The customers were instructed to immediately review your inventory for the specific lot numbers, quarantine the product and discontinue the shipment of the affected product; complete and return the Recall Response Form by fax to BD at 1-866-913-0661 or email to bd6835@stericycle.com (even if you do not have any product); return all affected products with completed response form per packing instructions; if you have distributed this product, identify your customers and notify them immediately of this product recall; if you would like BD to conduct the notification to your customers, email your customer list within 3 business days to Matthew.Kelleher@bd.com.if you have any questions or require assistance with the return of the recalled product and/or availability of credit product, contact 1-888-628-0732 between 8Am and 5 PM ET Monday through Friday. |  | 1788150 | 1,788,150 units | Worldwide distribution- US (nationwide) and countries of: Canada, Belgium, Brazil. | Open, Classified |  |  | 1788150 | Computer | Not_Computer
Z-0659-2016 | 72896 | No | LYJ |  |  | Neurology | PMA | VNS Therapy AspireSR Generator Model 106. |  |  | December 18, 2015 | 2015 | Cyberonics, Inc | Recall being initiated in response to three reports of "Burst Watchdog Timeout" events occurring with the Model 106 AspireSR Generator, resulting in a device reset condition where stimulation output is disabled. | Device Design | The firm notified consignees of the issue via letter on 12/18/15. The letter identified the affected device, the issue involved, and actions to be taken. Physicians are to contact Clinical Technical Support at 866-882-8804 to report if a patient's generator has been disabled due to the issue identified. Users are to complete and return the effectiveness card as soon as possible. If further information is needed, customers can contact Clinical Technical Support at 866-882-8804 or via e-mail at cservices@livanova.com. |  | 4935 | 4,935 units | Worldwide Distribution -- United States, Austria, Belgium, Croatia, Czech Republic, Finland, France, Germany, Iceland, Italy, Netherlands, Norway, Poland, Slovakia, Spain, Sweden, Switzerland, United Kingdom, Cyprus, Israel, Kuwait, Lebanon, Oman, Qatar, Saudi Arabia, and United Arab Emirates. | Open, Classified |  |  | 4935 | Computer | Computer
Z-0734-2017 | 75863 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | CyberKnife Robotic Radiosurgery System using software version 10.6; |  |  | November 22, 2016 | 2016 | Accuray Incorporated | Accuray has become aware of a potential safety issue involving unexpected treatment robot motion when removing an accessory. This issue is limited to certain CyberKnifeÂż Systems with software version 1 0.6. | Software design | Accuray sent an Urgent Safety Advisory Notification letter dated November 22, 2016, to all affected customers. Letters alerted users to the issue and the sequence of events necessary for it to occur. Accuray reps will visit all sites to perform a software upgrade to correct the issue.For more information or any questions about this notification, please contact Accuray Customer Support at 1-877-668-8667 or visit the Accuray Technical Solution Center at http://www.accuray.com/services-support.For further questions, please call (408) 716-4600. |  | 40 | 40 affected devices | Worldwide Distribution - US Distribution to the states of : PA, WA, MT, NY, SC, CA, MN, AK, NJ, FL, KY, WI., and to the countries of : Bulgaria, Canada, France, Germany, India, Japan, Luxembourg, Netherlands, Portugal, Switzerland and Turkey | Terminated | May 22, 2017 | 181 | 40 | Computer | Computer
Z-0742-2017 | 75812 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | SOMATOM Definition AS System, x-ray, tomography, computed |  |  | November 15, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Software update that provides software and firmware bug-fixes to improve system performance | Software Design Change | Siemens mailed a Customer Information Letter to affected customers on November 21, 2016 informing them of software updates that would provide bug-fixes to improve system performance and stability for customers with SOMATOM Definition AS, Definition Edge and Definition Flash systems using Software versions VA48A_SP1 and VA48A_SP2. The update will remotely be pushed to their systems or customers will be contacted by one of Siemens service organization to arrange a date for installation of the update on site. Customers with questions were instructed to call 1-800-888-7436. |  | 104 | 104 systems | Nationwide Distribution | Open, Classified |  |  | 104 | Computer | Computer
Z-1883-2016 | 73948 | No | HXX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Torque Limiting Screwdriver: |  |  | April 26, 2016 | 2016 | Greatbatch Medical | Greatbatch Medical has initiated a global Field Recallfor Torque Limiting Devices. Greatbatch identified through laboratory testing that the required sterility assurance level (SAL) of 10-6 is notachieved on torque limiting devices when subjected to steam sterilization cycles identified in the InstructionFor Use (IFU) provided with the device. | Device Design | Consignees were sent on 4/26/2016 a Greatbatch "Urgent Medical Device Recall" letter dated April 26, 2016. The letter described the product involved in the recall, the reason, and the risk to health. The letter advised consignees to identity and quarantine the product and to complete and return the Filed Recall Response Form. For questions contact your local Greatbatch Medical representative, at 1-763-951-8235, or e-mail FieldActionCenter@Greatbatch.com. |  | 1110 | 1110 | US: IN, MI, TX.OUS: GERMANY, FRANCE, ITALY, SWITZERLAND. | Open, Classified |  |  | 14128 | Computer | Not_Computer
Z-1881-2016 | 73978 | No | HXX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Firebird NXG Spinal Fixation System Set Screw Driver used to insert a set screw into a spinal fixation construct. Rx Only Provided non Sterile. |  |  | April 27, 2016 | 2016 | Orthofix, Inc | There is a possibility that the subject set screw driver could disengage from the torque limiting handle when the counter torque wrench is allowed to contact the counter torque handle. | Device Design | Affected distributors will be provided a written notification of field correction via USPS Certified Mail or FedEx on 4/27/16. Distributors will be additionally instructed to provide field correction notifcations to hospitals, surgeons, or other distributors that they may have further ditributed the product to. |  | 36 | 36 units | KS, NV, TX, CA, ID, FL, MA, MI, NJ, AZ, NH, SC, UT, WA | Terminated | November 23, 2016 | 210 | 36 | Computer | Not_Computer
Z-0899-2017 | 75836 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | AQURE System |  |  | November 23, 2016 | 2016 | Radiometer America Inc | Design error when displaying additional information in the patient view window; error may result in misreading a parameter and its value. | Software design | Radiomet4er sent a letter dated November 30, 2016, to all affected customers via email with a request for customers to complete a return form. Customers/Users were advised to not use the "more" button to expand patient results when viewing them in the AQURE patient view. Customers/Users Radiometer Representative will contact them to schedule a visit of a remote session. During the visit or remote session the representative will run a script on the AQURE system, which serve to remove the misalighnment possibility.It is noted in the letter to customers that expanding the results by using the More button will not affect the stored results.If customers are not the end-user of the affected product ensure that the notification letter is distributed to the final end-user.If any questions regarding the notification letter contact Radiometer America Technical Support at 1-800-736-0600 opt.1.For questions regarding this recall call 440-871-8900. |  | 7 | 7 total (2 in US) | Worldwide Distribution - US including MI and NY and Internationally to England, Switzerland, and United Arab Emirates. | Open, Classified |  |  | 7 | Computer | Computer
Z-0598-2016 | 71924 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Brainlab Cranial Navigation System: An Image Guided Surgery System / Stereotactic. Radiology Departments. |  |  | April 22, 2013 | 2013 | Brainlab AG | Software Error: The effect of setup on overall navigation accuracy could potentially intensify small inaccuracies arising from individual steps of a complex navigation procedure that may cause an inaccurate display of instruments by the navigation system compared to the actual patient anatomy. | Software in the Use Environment | The firm, Brainlab AG, sent out a "Field Safety Notice / Product Notification" CAPA-20130417-000315 Brainlab Cranial Navigation System: Effect of setup on overall navigation accuracy for all versions of software on 4/22/2013 and an update on 5/29/2015.- From April 22, 2013 has been distributed via email to the consignees.- UPDATE from May 29, 2015 was distributed via email to the medical users.An active reply by the consignees will be followed, if necessary via further different communication media.FIELD SAFETY NOTICE / PRODUCT NOTIFICATION letter describes the product, problem and actions to be taken. The customers were instructed to adhere to the measures described in document Measures to Improve Cranial Navigation Accuracy in the Product Notification letter when using the Brainlab Cranial Navigation System.Brainlab is requesting that 1. Existing potentially affected customers receive this Product Notification Letter.2. These customers receive the attached supplement to the Instructions for Use regarding the Brainlab Cranial Navigation System in hardcopy version as an amendment to their current User Manuals. Tentative planned timeline for availability: June 2013Additionally to advise the appropriate personnel working in your department of the content of Product Notification letter.If the medical users have the need for further clarification, they are asked to contact their local Brainlab Customer Support Representative. Customer Hotline: +49 89 99 15 68 44 or +1 800 597 5911 (for US customers) or byE-mail: support@brainlab.com. Fax Brainlab AG: + 49 89 99 15 68 33 Address: Brainlab AG (headquarters), Kapellenstrasse 12, 85622 Feldkirchen, Germany. |  | 1430 | 1,430 units | Worldwide Distribution: US (Nationwide) including Puerto Rico; and countries of: Algeria, Argentina, Australia, Austria, Bahrain, Belarus, Belgium, Bolivia, Brazil, Canada, Chile, China, Colombia, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, Estonia, Finland, France, Germany, Greece, Honduras, Hong Kong, Hungary, India, Iran, Iraq, Ireland, Israel, Italy, Jamaica, Japan, Jordan, Latvia, Lithuania, Malaysia, Mexico, Morocco, Netherlands, Norway, Oman, Pakistan, Panama, Philippines, Poland, Portugal, Qatar, Russian Federation, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Turkey, Ukraine, United Arab Emirates, United Kingdom, Venezuela, and Vietnam. | Open, Classified |  |  | 1430 | Computer | Computer
Z-0728-2017 | 74488 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Carry bar which is an accessory to a lift used for patients in hospitals and other facilities. Recall impacts the following model numbers: 360741, 360750, 360751, 360755, 360755 Rev B, 360756, and 360757. |  |  | August 30, 2016 | 2016 | ErgoSafe Products, LLC (DBA) Prism Medical | The black plastic puck on the carry bar is breaking resulting in potential for patient harm. | Device Design | The recalling firm sent field safety notices to their customers by mail on August 30, 2016. The notice instructed customers to inspect their carry bar to see if it used a black plastic puck. If it did, they were instructed to contact the recalling firm to receive a metal insert to install. If the puck is cracked or damaged, they can request a new, replacement carry bar at a discounted price. For further questions, please call (866) 891-6502. |  | 17964 | not greater than 17,964 units | Distributed US nationwide and to Canada, Japan, Australia, Thailand, Great Britain, and South Korea | Open, Classified |  |  | 17964 | Computer | Not_Computer
Z-0655-2016 | 72897 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Ergon 2 Articulating Equipment Boom Monitor Brackets |  |  | December 02, 2015 | 2015 | Skytron, Div. The KMW Group, Inc | The hardware that holds the VESA plate and the display interface / monitor to the bracket may fail causing the display interface to fall from the bracket. Over time and with use the threaded rod hardware material can weaken and detach from the bracket. A display interface falling from its mounting location may result in injury to the patient or provider. | Device Design | Skytron sent an Urgent Medical Device Recall letter dated December 02. 2015, to all affected customers via certified mail. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Customers with questions were instructed to contact at 616-656-2900 with any questions. For after-hours support, please call 1-800-SKYTRON. |  | 286 | 286 | Worldwide Distribution - US including WI, TX, FL, NC, MD, OH, AZ, CA, UT, MT, OK, HI, CO, WY, MO, NJ, IL, NY, MA, IA, DE, PA, MI, IN, and WV and Internationally to Canada, and Japan. | Open, Classified |  |  | 286 | Computer | Computer
Z-0730-2017 | 75714 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge Cardio software. |  |  | October 01, 2015 | 2015 | Merge Healthcare, Inc. | Reporting feature times out after inactivity for more than an hour sending the user back to the study list, which causes all reporting data being entered to be lost. | Software design | MERGE sent an Urgent Medical Device Recall letter dated March 28, 2016, to all affected customers. Non-responders were issued a second letter dated July 11, 2016, The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to reply using the enclosed form and the return addressed envelope and also to ensure that all users of the product were provided with the notification. If product was further distributed, customer were asked to identify their customers and notify them at once of the product recall. Customers with questions were advised to email recall@merge.com.For questions regarding this recall call 877-741-5369. |  | 116 | 116 sites potentially have the effected versions | Nationwide Distribution | Open, Classified |  |  | 116 | Computer | Computer
Z-1707-2016 | 73968 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Discovery IGS 740. Indicated for use in generating fluoroscopic and rotational images of human anatomy. |  |  | April 15, 2016 | 2016 | GE Medical Systems, LLC | Potential non-recoverable loss of image acquisition. The affected Discovery systems may experience multiple X-Ray abort errors before or during a real-time fluoroscopic Interventional procedure. | Device Design | Consignees were sent on 4/15/2016 a GE Healthcare "Urgent Medical Device Correction" letter GEHC Ref#12246 dated April 14, 2016. The letter was addressed to Hospital Administrators/Risk Managers, Managers of Radiology/Cardiology and Radiologists/Cardiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. For questions or concerns regarding this notification, please contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 25 | 25 | Worldwide Distribution-US (nationwide) including the states of AL, CA, CO, CT, IL, IN, MD, MN, MS, MO, NE,NV,NH, NM, NY, OH, OR, PA, SC, TN, TX, UT, VA, WA, and WI, and the countries of ALGERIA, BELGIUM CANADA, CHINA, FRANCE, GERMANY, HUNGARY, IRELAND, ITALY, JAPAN, MEXICO, NETHERLANDS, PERU, SAUDI ARABIA, SINGAPORE, UNITED ARAB EMIRATES, and UNITED KINGDOM. | Open, Classified |  |  | 79 | Computer | Computer
Z-0729-2017 | 75825 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge Cardio software. |  |  | March 28, 2016 | 2016 | Merge Healthcare, Inc. | If images are sent without an order in the system, they will be matched with the latest order on the current patient/modality matching potentially resulting in the matching of the report to the incorrect accession number (but still associated to the correct patient). | Software design | MERE sent an Urgent Medical Device Recall letter dated July 11, 2016, to all affected customers via email or via certified mail if they did not have an email address for the customer. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to respond to the notification using the enclosed form and the return addressed envelope. Customers were also instructed to ensure that all users of the product are provided with the notification and if product was further distributed customers need to identify their customers and notify them at once of the product recall. Customers with questions should send an email to recall@merge.com. For questions regarding this recall call 877-741-5369. |  | 198 | 198 sites potentially have the effected versions | Nationwide Distribution | Open, Classified |  |  | 198 | Computer | Computer
Z-0602-2016 | 72840 | No | PFJ | Hygroscopic Laminaria cervical dilator. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Gynecare Thermachoice III Uterine Balloon Therapy System Ethicon, Inc. |  |  | December 03, 2015 | 2015 | Ethicon, Inc. | Stability data does not substantiate the labeled two-year shelf life of affected product. | Device Design | Ethicon, Inc. sent "Urgent: Medical Device Recall (REMOVAL)" notifications and Business Reply Forms dated 12/3/2015 on same date via UPS Next Day Mail to its customers. The notification informed the customers of the issue with the product; how to identify affected product and action required. The customers were instructed to examine their inventory immediately to determine if they have affected product on hand and quarantine the affected product(s); remove the affected product and communicate the issue to relevant operating room or materials management personnel, or anyone else in their facility who needs to be informed; if any affected product has been forwarded to another facility, contact that facility to arrange return; complete and return the enclosed Business Reply Form (BRF) confirming receipt of this notice within three (3) business days to Stericycle by fax at 1-866-792-5453 or by email at Ethicon7427@stericycle.com, return BRF even if you do not have affected product. Stericycle is handling returns. If the customer requires any assistance with returning product or have any questions, they were instructed to contact the Customer Support Center at 1-877-ETHICON (1-877-384-4266).Ethicon, Inc. issued an update to their 12/3/2015 letter in a notification dated 12/23/2015. The 12/23/2015 notification clarified the lots of the product subject to their recall previously communicated in their 12/3/2015 letter. There are three (3) lots of product subject to the recall that were inadvertently omitted from the original lot listing. The notification stated that "All THERMACHOICEÂż Catheter lots with an expiration date prior to December 10, 2017 are subject to this recall and are required to be returned. The lot number listing of all product subject to the recall is contained in Attachment 1. The expiration date and lot number can be determined by using the Product Identification Tool attached hereto at Attachment 2." The attachments included the updated pr |  | 96124 | 96,124 units | Worldwide Distribution: US (nationwide) including Puerto Rico and countries of: India, Belgium, Argentina, Mexico, Brazil, Canada, Colombia, Ecuador, Singapore, Aruba and Venezuela. | Terminated | May 30, 2017 | 544 | 96124 | Computer | Not_Computer
Z-0731-2017 | 75853 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MOSAIQ Oncology Information System |  |  | December 02, 2016 | 2016 | Elekta, Inc. | Edits to Particle field definition parameters may not be saved when the field definition window is saved. | Software design | The firm, Elekta, sent an "URGENT IMPORTANT FIELD SAFETY NOTIFICATION" letter dated 11/2016 was sent to affected customers on 12/2/2016. The letter describes the product, problem and actions to be taken. The customers were instructed to save and re-open the field definition form when manual edits have been performed to confirm that the edits were save; post this notice in a place accessible to all users; advise the appropriate personnel working with the product the content of the letter; and complete and return the Acknowledgement Form via fax to: 408 830 8003 immediately upon receipt but no later than 30 days.Elekta will resolve the problem with the MOSAIQ 2.60 SP7F, MOSAIQ 2.60 SP14 and 2.64 SP3. Customers are encouraged to contact Elekta support to get the corrected version.If you have any questions, contact RA Manager Americas at 1-770-300-9725 or email to: Elekta_Global_PMS@elekta.com. |  | 47 | 47 | Worldwide Distribution-US (nationwide) to states of: FL, IL, MO, NJ, OH, OK, TN, TX, VA, and countries of: Italy, Japan, South Korea, Sweden and United Kingdom. | Open, Classified |  |  | 47 | Computer | Computer
Z-0725-2017 | 75263 | No | JGJ | Magnesium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Magnesium on RX Imola analyser |  |  | April 14, 2016 | 2016 | Randox Laboratories, Limited | According to the firm, Carry over was observed when the amylase or pancreatic amylase test is run directly before or after the Magnesium assay. An update is required to the RX Imola analyser running order and an acid wash recommended when setting up the magnesium assay. A correction was made to the IFU for the Magnesium assay and all RX Imola customers in the USA were contacted with the updated instruction. | Device Design | Randox sent an Urgent Field Safety Notice to all affected consignees. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to review and update Rx Imola running order where required. Rework any remaining stock with updated IFU and Important Notice. Complete and return the vigilance response section of the form to technicalservikces@randox.com. |  | 15 | 15 | Nationwide Distribution The product was only distributed to US Consignees. | Open, Classified |  |  | 15 | Computer | Not_Computer
Z-0721-2017 | 75731 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | RayStation 4.0, 4.5, 4.7 and 5.0; Radiation Therapy Treatment Planning System |  |  | November 09, 2016 | 2016 | RAYSEARCH LABORATORIES AB | Software anomaly; an issue was found with the proton Pencil Beam Scanning (PBS) dose calculation in RayStation 4.0, 4.5, 4.7 and 5.0. For treatment plans with a combination of range shifter, large air gap and beams that enter the patient surface at an oblique angle, the dose calculation accuracy may be less than expected. | Software design | Ray Search sent an Urgent Field Safety Notice Medical Device Correction letter (dated 11/8/2016) via e-mailed to customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to Inspect their product and identify all installed units with the above software version number, then confirm you have read and understood this notice (contact information included). |  | 66 | US: 66 units | US Nationwide Distribution in the states of California, Florida, Illinois, Louisiana, Michigan, New Jersey, Tennessee, Texas, Washington | Open, Classified |  |  | 66 | Computer | Computer
Z-0723-2017 | 75652 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | ADVIA 560 Hematology Systems, Siemens Material Number (SMN) 11170842, IVD. |  |  | October 21, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc. | Siemens is investigating an issue which may cause an incorrect result to be reported. Siemens received two reports of multiple discordant records for the same Sample ID in the ADVIA 560 Hematology System database that occurred during the installation of the systems. The database should only contain one record of a Sample ID number for any given time and date. If there are multiple records for the same Sample ID, it is possible that multiple results may be manually or automatically sent to the Laboratory Information System (LIS), printed or displayed on the results report screen. | Software design | An Urgent Medical Device Correction Letter HI17-02.A.US (dated October 21, 2016) and Response Form were sent to US Customers via Federal Express. An Urgent Field Safety Notice was sent to customers outside the US on October 21, 2016, (distribution is determined at the country level). The letters instruct customers to contact their local Siemens technical support representative if they obtain multiple results for one Sample ID. It also informs the customer what to do if the ADVIA 560 Hematology System generates any of the six error messages listed in the letter. For further questions, please call (312) 275-7795. |  | 164 | US: 23 systems; Foreign: 141 systems | Worldwide Distribution - US Distribution to the states of including Puerto Rico and : CA, CO, FL, IL, MI, NJ, NY, OH, TX, WA, WI, WI and WY., and to the countries of : Angola, Australia, Belgium, Brazil, Canada, Chile, Colombia, Croatia, Czech Republic, Egypt, France, Germany, Greece, India, Italy, Japan, Kenya, Malaysia, Mexico, Myanmar, Netherlands, Norway, Pakistan, Portugal, Saudi Arabia, Slovenia, Spain, Switzerland, Turkey, Uganda, United Kingdom, Vietnam. | Open, Classified |  |  | 164 | Computer | Computer
Z-1885-2016 | 73885 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension Vista 500 Intelligent Lab System running on software versions V.3.6.1 SP1 or V.3.6.2, Device listing # D011374, All serial numbers and lots are affected. |  |  | March 24, 2016 | 2016 | Siemens Healthcare Diagnostics Inc. | Discrepant patient results on Dimension Vista Intelligent Lab Systems. Siemens Healthcare confirmed a software defect which, in a very specific set of circumstances, results in the Dimension Vista System omitting an aliquot probe rinse between sample aspirations when processing tubes in Sample Racks that are front loaded on the Dimension Vista System. | Software design | Siemens sent an Urgent Medical Device Correction letter dated March 24, 2016, to all affected customers. Customers were advised to follow the instructions contained in the "Urgent Medical Device Correction" notification letter to help minimize or eliminate the impact that this specific software defect could cause until a software correction could be made (i.e instructions to minimize the impact of the software defect due to "Bad Sample Detect " False Transition" errors when processing front loaded samples) and (how to eliminate the impact of the software defect when processing front loaded samples). Customers were also asked to review the recall notification letter with their Medical Director and complete and return the Field Correction Effectiveness Check Form within 30 days. Customers with questions were advised to contact Siemens Customer Care Center - Technical Solutions at 800-441-9250 or their local Siemens technical support representative.For questions regarding this recall call 302-631-8248. |  | 1301 | 1301 units | Nationwide Distribution to AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, , ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY and Puerto Rico. | Terminated | March 13, 2017 | 354 | 5202 | Computer | Computer
Z-0653-2016 | 72776 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | 30 deg, 8mm Endoscope, IS4000: |  |  | November 23, 2015 | 2015 | Intuitive Surgical, Inc. | This correction to the da Vinci Xi I&A; User Manual communicates proper use of the da Vinci Xi 8 mm Endoscopes in combination with the da Vinci 8 mm and 12 mm cannulae, cannula reducers, and monopolar energy instruments. It informs customers of scenarios that may result in unintended patient harm due to capacitive coupling from monopolar energy sources. | Device Design | Intuitive Surgical sent an Urgent Medical Device Correction letter to affected customers on December 4, 2015,. Letters explained the issue and provided action to follow to prevent the issue. Customers were instructed to do the following:Distribute a copy of this letter to all da Vinci Xi users at your facility.3. Place a copy of this letter with your user manual.4. Complete the attached Acknowledgement Form and return it to Intuitive Surgicalas instructed.5. Retain a copy of this letter and the Acknowledgement Form for your files.If you need further information or support concerning this Medical Device Notification ,please contact your Clinical Sales Representative or contact Intuitive Surgical CustomerService at the numbers listed below:Âż North and South America: (800) 876-1310, Option 3 (6 AM to 5 PM PST) ormail: customersupport-servicesupport@intusurg.comÂż Europe, Middle East, Asia and Africa: +800 0821 2020 or +41 21 821 2020 (8AM to 6 PM CET) or ics@intusurg.comÂż South Korea: 02-3271-3200 (9 AM to 6 PM KSTJ)Âż Japan: 0120-56-5635 or 003-5575-1362 (9 AM to 6 PM JST) |  | 2624 | 2624 total all models | Worldwide Distribution - US (nationwide) and Internationally to Australia, Austria, Belgium, Cyprus, Czech Republic, Denmark, France, Germany, Hong Kong, India, Israel, Italy, Japan, Monaco, Netherlands, Norway, Puerto Rico, Qatar, Romania, Saudi Arabia, Singapore, South Korea, Spain,Sweden, Switzerland, Taiwan, Turkey, United Kingdom, and Venezuela.. | Terminated | April 05, 2016 | 134 | 2624 | Computer | Computer
Z-0755-2017 | 75584 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | GE Healthcare, Avance, Avance CS2, Amingo. |  |  | November 03, 2016 | 2016 | GE Healthcare, LLC | GE Healthcare has become aware of a potential safety issue where certain Avance CS2, Avance and Amingo anesthesia devices can transition to a System Malfunction state ifthe lower storage drawer containing the optional large tray insert accessory closed with an abnormally high amount of force. | Device Design | GE Healthcare sent an "Urgent Medical Device Correction" letter GEHC Ref# 34079 dated November 3, 2016, to all affected customers. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction &amp; Contact Information. For questions contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 23098 | 23,098 (US = 8,711; OUS = 14,387) | USA ( nationwide) Distribution including Puerto Rico and DC. and to the countries ofAustraliaAUSTRIAAzerbaijanBahrainBangladeshBarbadosBelarusBelgiumBOLIVIABOSNIA-HERZEGOVINABrazilBRUNEIBULGARIACanadaCHILEChinaCOLOMBIACOSTA RICACroatiaCyprusCzech RepublicDenmarkECUADOREgyptEstoniaEthiopiaFinlandFRANCEGeorgiaGERMANYGhanaGibraltarGREECEHONG KONGHUNGARIAIcelandIndiaIndonesiaIRAQIrelandIsraelITALYJAPANJORDANKAZAKHSTANKenyaKOREA, REPUBLIC OF KuwaitLatviaLebanonLibyan Arab JamahiriyaLITHUANIAMACEDONIAMALAYSIAMaltaMexicoMOLDOVAMOROCCONAMIBIANetherlandsNew ZealandNicaraguaNigeriaNORWAYOmanPakistanPanamaPARAGUAYPERUPHILIPPINESPolandPORTUGALQatarROMANIARussian FederationSaudi ArabiaSerbiaSERBIA-MONTENESingaporeSlovakiaSLOVENIASOUTH AFRICASPAINSWEDENSWITZERLANDSYRIATAIWANThailandTunisiaTURKEYUkraineUnited Arab EmiratesUNITED KINGDOMUruguayUZBEKISTANVenezuelaVietNamYemen | Open, Classified |  |  | 23098 | Computer | Not_Computer
Z-0720-2017 | 75633 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Radiation Therapy Treatment Planning System |  |  | October 11, 2016 | 2016 | RAYSEARCH LABORATORIES AB | An error may occur with the display of dose computed on images other than the planning CT (auxiliary CT) when using multiple patient cases in RayStation 5. If a CT image set with the same Frame of Reference as the displayed auxiliary CT set exists in another case, the dose display may be incorrect. When the error occurs, the display of dose in patient views, including the maximum dose position, the dynamic isodose lines and the dose grid may be incorrect. The dose value normally displayed in the upper left corner when pointing in the 2D view may be incorrect or missing. | Software design | On October 11, 2016 RaySearch Laboratories distributed a Field Safety Notice, Medical Device Correction #13529 to customers via email. Customers were instructed to take caution when a patient has multiple cases. The following may be incorrect if there are image sets that shre the same Frame of Reference residing in different cases: *Display of evaluation dose computed on another CT image set and *Dose tracking fraction dose and deformed dose. Until a corrected version has been installed, all affected users must maintain awareness of this field safety notice. In these cases, the display of the dose values in the 2D view, including the dose value shown when pointing in the view, maximum dose position, dynamic isodose lines and dose grid display may be incorrect. Customers are encouraged to educate staff and all users about this. For regulatory information, please contact David Hedfors, at +46 8 510 530 12 or david.hedfors@raysearchlabs.com. |  | 540 | 540 units | US Nationwide Distribution in the states of AZ, AR, CA, CT, DE, FL, GA, HI, IL, LA, MA, MD, MN, MI, MS, MO, NJ, NY, NC, OH, PA, RI, TX &amp; WA | Open, Classified |  |  | 540 | Computer | Computer
Z-0639-2016 | 72540 | Yes | BTL | Powered emergency ventilator. | Anesthesiology | Anesthesiology | 510(k) | Mercury Medical Neo-Tee T-Piece Resuscitator and T-Piece Circuit. |  |  | October 22, 2015 | 2015 | Mercury Enterprises, Inc. dba Mercury Medical | Incorrect use of T-Piece resuscitators will result in complete occlusion of the patient circuit and consequent failure to ventilate the patient. | Device Design | Mercury Medical sent notification letters dated October 22, 2015, to all affected Domestic Consignees by courier and for the foreign consignees the notifications were sent via e-mail.The letter identified the product, the problem, and the action to be taken by the consignee. Customer Immediate Actions - Immediately notify all personnel subject to using these devices to remove protective cap before connecting a mask. Further this filed safety advisory notice should be read in conjuction with the manufacturer supplied instructions for use.In addition, if you further distributed this product, please identify your customers and notify them at once of this product correction. Notificatin to your customers should include a copy of this letter. For questions contact Mercury Medical at 787-573-0088, 727-573-4966 fax or via e-mail at regulatoryaffairs@mercurymed.com |  | 753245 | 753,245 devices | Worldwide Distribution - US (nationwide) and Internationally to Australia, Austria, Bahrain, Canada, Chile, China, Colombia, Denmark, Ecuador, Egypt, Finland, France, Germany, Guam, Hong Kong, Hungary, India, Iran, Ireland, Israel, Italy, Japan, Kuwait, Latvia, Lebanon, Lithuania, Malaysia, Netherlands, Norway, Poland, Portugal, Puerto Rico, Qatar, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, United Arab Emirate, and United Kingdom.. | Open, Classified |  |  | 753245 | Computer | Computer
Z-0878-2017 | 75909 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge OrthoCase software. The firm name on the label is Merge Healthcare, Hartland, WI. |  |  | October 08, 2015 | 2015 | Merge Healthcare, Inc. | Measurements in the software are changing after saving a plan with a measurement result, which may result in potential patient injury or delay in diagnosis or treatment. | Software design | Merge issued a letter dated March 28, 2016, via email on April 4, 2016, or via certified mail if they did not have an email address for the customer. The letter notifies the customer of the issue, informs them of a workaround and that a fix is not yet available but they will be notified once one is available. The notification instructs the customer to pass the information onto all other users, as well as any downstream customers if it was further distributed. The notification also informs the customer their response is required using the enclosed form and the return addressed envelope. The response form asks, in part, if they are interested in accepting the fix. If they decline the fix, they are asked to provide the reason why. The response was required no later than April 18, 2016.Nonresponders were issued a second letter dated July 11, 2016, asking for a response by August 2, 2016.For further questions, please call (877) 741-5369. |  | 30 | 30 sites potentially have the effected versions | Distribution was made to medical facilities in CA, CO, CT, MA, MD, MN, MT, NE, NV, NY NJ, PA, SC, WA, and WI. Foreign distribution was made to Canada, as well as other countries. | Open, Classified |  |  | 30 | Computer | Computer
Z-0604-2016 | 72869 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Oncentra External Beam - VMAT - Radiation therapy planning system |  |  | December 18, 2015 | 2015 | Elekta, Inc. | When using the option "Tumor Overlap Fraction" in VMAT planning it has been observed that in rare cases the system does include an organ at risk as target volume. This could result in open MLC, and open jaws in areas away from the target volume. | Software design | Elekta sent an Important Field Safety Notice 806-01-ETP-001 to all affected customers on December 18, 2015. The notice informs users of the issue and that should the issue occur it could result in overdosage of organs at risk outside the intended treatment volume. Users are strongly advised to performproper Quality Assurance for all treatment plans before delivery of the first fraction to the patient. To prevent the issue from occurring, the workaround is to not use the Tumor Overlap Fraction for VMAT planning.and not to use the Tumor Overlap Fraction for VMAT planning. The issue will be solved in Oncentra External Beam version 4.5.2. The notice included an acknowledgement form which is to be returned to Elekta. For questions contact your local Elekta representative. |  | 154 | 154 | Worldwide Distribution - US Nationwide in the states of CA, NY, PA, and the countries of Austria, Belarus, Canada, China, Denmark, Finland, France, Germany, Greece, India, Italy, Japan, Netherlands, Norway, Poland, South Korea, Spain, Sweden and United Kingdom. | Open, Classified |  |  | 154 | Computer | Computer
Z-0865-2017 | 75891 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | CyberKnife Robotic Radiosurgery Systems, models: G3, G4, VSI. |  |  | December 06, 2016 | 2016 | Accuray Incorporated | Accuray has become aware of a potential safety issue involving the possibility of electric shock duringmaintenance activities from the lon Pump HV Power Supply at the back of the modulator cabinet. This issue is limited to certain G3, G4, and VSI models of the CyberKnifeÂż Robotic Radiosurgery System. | Device Design | Accuray Incorporated sent an Urgent Safety Advisory Notification letter to all affected customers on December 1, 2016 by traceable method.. Letters identified the product, the problem, and the action to be taken by the customer. provided users with affected product and described the issue. The final resolution will result in a modulator cabinet inspection and installation of the Plexiglas cover - as necessary. Customers will be contacted by Accuray to schedule the inspection. Prior to this modulator cabinet inspection, inform all applicable personnel of the potential danger and need to closely follow electrical safety procedures. Customers with questions were instructed to contact Accuray Customer Support at 1-877-668-8667 or visit the Accuray Tehnical Solution Center at http://www.accuray.com/services-support or call:North America +1-877-666-8667, Europe +800-4141-9595, Hong Kong +852-2247-8688, Japan +81-3-6265-1530. Customers were also asked to complete the attached acknowledgement and return to Accuracy.For questions regarding this recall call 608-824-3362. |  | 115 | 115 affected devices. | Worldwide Distribution - US including CA. MN, OK. FL. MO, TN, KS, TX, AZ, MS, IA, MA. NC, NY. IN, NJ, VA. LA, WA, WI, PA. CT, G,. MT, MD, LA., WV, OH, and DC.Foreign distribution to Taiwan, Japan, Korea, Turkey. Malaysia, Vietnam, Greece, China, Spain, France, Italy, Pakistan, and Georgia, | Terminated | June 09, 2017 | 185 | 115 | Computer | Computer
Z-0624-2016 | 72701 | Yes | FJK | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | NIPRO SET / SLIMLINE BLOOD TUBING SET for Hemodialysis with Transducer Protectors and Priming Set |  |  | March 30, 2015 | 2015 | Nipro Medical Corporation | Blood foaming was found during evaluation of the product. | Device Design | NIPRO contacted their consignees by email 9/25/2013 of the affected product. Customers were instructed to return product back to Nipro Medical Warehouse or discard at point of use. Please contact Director of QA/RA at Nipro Medical Corporation for any additional information or return correspondence to JessicaO@nipromed.com / 305.599.7174x249. |  | 1484 | 1,484 pieces. | US Distribution to states of: TN and NJ. | Terminated | February 04, 2016 | 311 | 1484 | Computer | Not_Computer
Z-0597-2016 | 72751 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Software version VD10E for Syngo X-Workplace; Picture archiving and communication system. |  |  | November 18, 2015 | 2015 | Siemens Medical Solutions USA, Inc | Potential post-processing software issue when using tabcard "4D" on X-Workplace with software version VD10E. A too small measurement in the MPRs of a volume when images are acquired with a CT scanner with a tilted gantry. In a 2x2 layout, the 4D tabcard will show an incorrect, too small length measurement. This could result in selecting a device of the wrong size, which then needs to be exchanged. | Software design | Siemens sent an Important Customer Safety Notice dated November 17, 2015. The Notice identified the product, the problem, and the action to be taken by the customer. The potential issue will be remedied with a software update via Update Instructions AX067/15/S. Siemens service organization will contact customers to arrange a date for the installation of the software update. Customers with questions were advised to call 800-888-7436. |  | 73 | 73 units | Nationwide Distribution including:AL, AZ, CA, CO, FL, IA, ID, IL, IN, MA, MD, MI, MN, MO, MT, NC, NY, OH, OK, PA, SC, TN, TX, VA, and WI. | Open, Classified |  |  | 73 | Computer | Computer
Z-0787-2017 | 75816 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | AMS-651, 120" Microbore extension set with slide clamp; Product Code: AMS-651 |  |  | November 18, 2016 | 2016 | Vygon MFG, Inc., dba/ Churchill Medical Systems, Inc. | Incorrect priming volume is printed on the device package. | Device Design | Vygon sent an Urgent Medical Device Recall letter dated November 18, 2016 to customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to take the following actions: 1. Check your stock for the products included within the scope of the recall. Immediately cease use and distribution, and quarantine all affected product immediately; 2. Count and document your affected inventory on the "Recall Acknowledgement and Inventory Return Form; 3. Fax or email the form to Vygon including customer contact information; 4. Vygon customer service will contact the customer and issue a Return Merchandise Authorization number; 5. Return the product as well as the original signed Recall Acknowledgement and Inventory Return Form using the instructions from Customer Service involving product pick-up; 6. All inventory will be credited or replaced at no cost to you. For questions contact local sales representative or Vygon's Customer Service Department at 1-800-473-5414.or by email at customerservice@vygonus.com. |  | 31500 | 31,500 | US distribution in the state of CA | Open, Classified |  |  | 31500 | Computer | Not_Computer
Z-1865-2016 | 73826 | No | CHP | Estradiol test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | The ARCHITECT Estradiol Kit |  |  | March 24, 2016 | 2016 | Abbott Laboratories | Interaction of Fulvestrant with the ARCHITECT Estradiol assay, leading to falsely elevated estradiol results in patients treated with Fulvestrant. | Device Design | The firm sent a Product Correction letter dated 3/24/2016 to all affected consignees describing the nature of the recall as well as instructions to follow. The necessary actions included the following:Consignees are asked to review the letter with their Medical Director. Patients undergoing Fulvestrant therapy should not be tested with the ARCHITECT Estradiol assay. Complete the Customer Reply Form. If the consignee has forwarded the product listed to other laboratories, a copy of Product Correction letter should be provided to them.The customer is asked to retain the letter for their laboratory records.Customers with questions should contact either Customer Service at 1-877-4ABBOTT for the US; or, the local area customer service for outside of the US. |  | 39346 | 39,346 units | Worldwide Distribution -- US, to the following states: AK, AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NH, NJ, NV, NY, OH, OK, PA, SC, SD, TN, TX, UT, VA, WA, WI, WV and WY; and, the countries of ALBANIA, ALGERIA, ANGOLA, ARGENTINA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BANGLADESH, BELARUS, BELGIUM, BOLIVIA, BOSNIA AND HERZEGOVI, BRAZIL, BRUNEI, BULGARIA, CAMEROON, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, CROATIA, CYPRUS, CZECH REPUBLIC, DENMARK, ECUADOR, EGYPT, EL SALVADOR, ESTONIA, ETHIOPIA, FINLAND, FRANCE, GAZA &amp; JERICHO, GEORGIA, GERMANY, GREECE, GUATEMALA, HONDURAS, HUNGARY, INDIA, INDONESIA, IRAN, IRELAND, ISRAEL, ITALY, IVORY COAST, JAMAICA, JAPAN, JORDAN, KAZAKHSTAN, KENYA, KUWAIT, KYRGYZSTAN, LATVIA, LEBANON, LIBYA, LITHUANIA, LUXEMBOURG, MACEDONIA, MALAYSIA, MALDIVES, MALI, MAURITIUS, MEXICO, MONTENEGRO, MOROCCO, NETHERLANDS, NEW ZEALAND, NIGERIA, NORWAY, OMAN, PAKISTAN, PARAGUAY, PERU, PHILIPPINES, POLAND, PORTUGAL, QATAR, REPUBLIC OF YEMEN, ROMANIA, RUSSIA, SAUDI ARABIA, SENEGAL, SERBIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SOUTH KOREA, SPAIN, SRI LANKA, SWEDEN, SWITZERLAND, TAIWAN, TANZANIA, THAILAND, TOGO, TRINIDAD &amp; TOBAGO, TUNISIA, TURKEY, UAE, UKRAINE, UNITED KINGDOM, URUGUAY, VENEZUELA, VIETNAM and ZIMBABWE. | Open, Classified |  |  | 39346 | Computer | Not_Computer
Z-1709-2016 | 73945 | No | HTW | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | 3.2 mm Three-fluted drill bit, sterile, Catalog ID # 03.010.060S |  |  | April 18, 2016 | 2016 | Synthes (USA) Products LLC | The sterile packaging may be compromised. A protective cap is to be secured on the tip of the drill bit. If the protective cap is not secured, the drill bit may puncture the sterile barrier. | Device Design | DePuy Synthes sent an Urgent Notice Medical Device Letter dated April 19, 2016 to customers. The identified the affected product, problem and actions to be taken. The letter instructed customers to examine their inventory to determine if they currently have the affected product. Customers were asked to complete and return the Verification Section of the recall notification letter. For questions call 610-719-5450. |  | 7363 | 7363 units | US Nationwide distribution in the states of: AE, AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX,UT, VA, VT, WA, WI | Open, Classified |  |  | 8637 | Computer | Not_Computer
Z-0786-2017 | 75794 | No | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | AbbVie PEG Kit 20 FR |  |  | November 22, 2016 | 2016 | AbbVie Inc. | AbbVie is issuing an urgent field correction notification due to an increased number of reports of disconnections of the AbbVie 20FR PEG tube from the Y-connectors. | Device Design | The firm sent an URGENT: MEDICAL DEVICE FIELD CORRECTION NOTIFICATION AbbVie PEG Kit 20 FR to affected customers describing the reason for the notification. The letter asks affected customers to forward this letter to all Movement Disorder Specialists and physicians (e.g. GIs, IRs, and surgeons) that place PEG- J Tubes within your facility.The letter describes actions to be taken for patients still using the AbbVie PEG 20FR Tube. The letter requests user facilities to continue to monitor the performance of the Tube. The firm states that there is no need to replace the AbbVie PEG 20FR Tube if it is functioning as expected and maintaining a viable connection with the Y-Connector. The letter provides the following instructions when establishing a connection between Abbvie PEG 20FR Tube with the 20FR Y-connector, always refer to AbbVie J Intestinal Tube 9 FR to PEG 15 and 20 FR Instructions for Use for 62918-001. The letter notes the following:1. Ensure that the set-up process clearly follows the sequence of events as described in the instructions for use and the following clarifications. 2. Ensure that the AbbVie PEG 20FR tube is connected to the corresponding 20 Y- connector (see photograph below which shows the difference between the 15FR and 20FR connection diameter). 3. DO NOT USE ALCOHOL CONTAINING PRODUCTS with the AbbVie PEG 20 FR tube as this may result in damage to the tube.4. Ensure that the Tube Clamp and Tube Clip on Fixation Plate are both closed. 5. Cut off the proximal cone and Fixation Loop of the AbbVie PEG Tube (15FR or 20FR). Note: Ensure that the cut is perpendicular (90 degree angle) to the tube and not on a diagonal as this ensures proper seating of the tube when engaged into the pin of the Y-connector fitting. 6. Slide the colored Fixation Screw (including the outer ring) over the Abbvie PEG Tube:a. For 15 FR, use the Blue/White screwb. For 20 FR, use the Violet/White screw7. Grasp near the |  | 541 | 541 units | U.S. consignees located throughout the country and Puerto Rico. Foreign Consignees: Worldwide. | Open, Classified |  |  | 541 | Computer | Not_Computer
Z-1872-2016 | 73961 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance 64 Computed Tomography X-ray Systems |  |  | April 08, 2016 | 2016 | Philips Medical Systems (Cleveland) Inc | The firm became aware of a problem where the system may not map Varian drive after CT user logout/login. | Software design | The firm sent a Customer Information letters dated March 8, 2016 to their customers. The letter identified the affected product, problem and actions to be taken. For information or support concerning this issue, please contact your local Philips representative or local Philips Healthcare office. For North America and Canada, contact the Customer Care Solutions Center (1-800-722-9377, Option 5: Enter Site ID or follow the prompts). |  | 158 | 158 units | Worldwide Distribution - US Nationwide in the states of: AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WV &amp; WY. Product was also shipped to the following countries: Argentina, Australia, Austria, Bangladesh, Belgium, Bolivia, Brazil, Cambodia, China, Colombia, Cuba, Estonia, Finland, France, Germany, India, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Korea, Latvia, Lebanon, Libya, Lithuania, Malaysia, Myanmar, Netherlands, New Zealand, Norway, Palestine, Panama, Philippines, Poland, Portugal, RÂżunion, Russian Federation, Senegal, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Tanzania,Thailand, Turkey, United Arab Emirates, United Kingdom, Uzbekistan &amp; Vietnam. | Open, Classified |  |  | 1097 | Computer | Computer
Z-0647-2016 | 72804 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | RT Elements are applications for radiation treatment planning for use in stereotactic, conformal, computer planned, Linac based radiation treatment of cranial, head and neck, and extracranial lesions. The simulated plan is intended for treatment evaluation for example in tumor board meetings or operating rooms |  |  | November 05, 2015 | 2015 | Brainlab AG | Large objects with fine resolution are potentially displayed cropped when imported into Brainlab Brain Metastases 1.0.0. | Software design | Brainlab sent a FIELD SAFETY NOTICE / PRODUCT NOTIFICATION letter dated November 5, 2015, to all direct accounts informing the of the recall and actions to be taken.User Corrective Action:Users of Brain Metastases Version 1.0.0 or Adaptive Hybrid Surgery Analysis 1.0.0 shall adhere to the following:After import always carefully review and verify all objects for correctness and validity.For reviewing an object, select the image set which was used to outline or to modify the objects in the Data menu. Also compare the contours of an object in Brain Metastases or Adaptive Hybrid Surgery Analysis software to the contours in SmartBrush. Please note that the error will not be visible in SmartBrush. Always verify the correct shape of the objects in Brain Metastases or Adaptive Hybrid Surgery Analysis software. If possible do not use image data sets with large slice distance and/or with a low amount of slices. To retrospectively identify if cropped objects have been used for treatment planning, the object contours of a specific treatment plan can be reviewed in Brainlab Dose Review, where the error will also be visible.Brainlab will provide a software update with this issue solved to affected customers. Brainlab will actively contact affected customers tentatively starting February 2016 to schedule the update. The request that users advise the appropriate personnel working in their department of the content of the Notification letter. If the user facilities require further clarification, thay can contact their local Brainlab Customer Support Representative. Customer Hotline: +49 89 99 15 68 44 or +1 800 597 5911 (for US customers)E-mail: support@brainlab.com (for US customers: us.support@brainlab.com)Fax: Brainlab AG: + 49 89 99 15 68 33 |  | 20 | 20 systems | Worldwide Distribution - US (nationwide) and Internationally to Australia, Belgium, Brazil, Canada, China, France, Germany, Hong Kong, Ireland, Israel, Italy, Japan, Mexico, Portugal, South Korea, Spain, Switzerland, Turkey and the United Kingdom. | Open, Classified |  |  | 68 | Computer | Computer
Z-0769-2017 | 75228 | No | LBI | Neonatal phototherapy unit. | General Hospital | General Hospital | 510(k) | Updated Natus neoBLUE blanket LED Phototherapy Systems. It provides phototherapy underneath the infant and can be used in a bassinet, open bed, radiant warmer, incubator, or while holding the infant. |  |  | September 16, 2016 | 2016 | Natus Medical Incorporated | The firm sent the Technical Bulletin to consignees who received the reliability updated neoBLUE blanket LED Phototherapy Systems between March 2016 to August 2016. The reliability update to prevent the early failure of the fiberoptic pads experienced by some customers on older systems. These failures on previous systems involved discoloration and melting of the fiberoptic bundle at the connector that is inserted in the neoBLUE blanket light box. | Device Design | Natus sent the Urgent Recall Notification/Field Safety Notice and Customer Reply Form (DOC-021042A), dated September 16, 2016 and the Technical Bulletin (DOC-021016) to US consignees via email.The letters and technical bulletin for foreign consignees was sent via email to the appropriate Natus Distribution Partner (DP) on November 7, 2016.US Customers can email the customer reply form to Natus_Quality_Progams@natus.com or fax to (847-949-6031). US Customers with questions can call Natus Technical Service at 888-496-2887. International customers should contact their local distributor with any questions. |  | 713 | 713 units total (403 units in US and 310 units outside US) | Worldwide Distribution - USA (nationwide) Distribution to the states of : AL, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MS, NC, ND, NJ, NM, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WI, WV and Puerto Rico. , and to the countries of : Argentina, Austria, Canada, Chile, China, Denmark, Dominican Republic, Finland, France, Germany, Hong Kong, Hungary, Indonesia, Italy, Latvia, Malaysia, Namibia, Paraguay, Poland, Qatar, Romania, Russia, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Taiwan, Trinidad and Tobago, Turkey, United Arab Emirates, United Kingdom, Uruguay, Viet Nam, and Virgin Islands (British). | Open, Classified |  |  | 713 | Computer | Not_Computer
Z-0891-2017 | 75856 | No | OVD | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Optimus Fixed Awl |  |  | November 22, 2016 | 2016 | Amendia, Inc | Complaints that the tip of the awl broke after impaction during surgery. | Device Design | Amendia Inc. sent an Urgent : Medical Device Recall letter dated November 22, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The letter requested that they examine inventory and quarantine the product. The letter also requested that they contact Amendia to confirm receipt of the letter and to obtain a Return Material Authorization for any product on hand.For further questions, please call (404) 610-7215. |  | 9 | 9 devices | US Distribution to the states of : FL, GA. IL, MO, TX | Open, Classified |  |  | 9 | Computer | Not_Computer
Z-0768-2017 | 75576 | No | OUG | Medical device data system. | General Hospital | General Hospital | 510(K) Exempt | MicroScan LabPro Data Management System. |  |  | October 25, 2016 | 2016 | Beckman Coulter Inc. | Beckman Coulter has received and confirmed reports of an intermittent and unexpected behavior when loading new panels into a WalkAway instrument usingLabPro Data Management System version 4.42. The issue could cause workflow interruption with a potential of delaying reporting results due to the inability to begin processing new panels. | Device Design | An Urgent Medical Device Recall letter dated 10/25/16 was sent to all their customers to inform them that Beckman Coulter has received and confirmed reports of an intermittent and unexpected behavior when loading new panels into a WalkAway instrument using LabPro Data Management System version 4.42. Customers are informed that the issue could cause WalkAway workflow interruption with a potential of delaying reporting results due to the inability to begin processing new panels. The letter informs the customers of the actions to be taken and the resolution for the product on recall. Customers are instructed to complete and return the enclosed Response Form within 10 days. Customers with any questions regarding the notice are instructed to contact their Customer Support representative: http://www.beckmancoulter.comor call 1-800-677-7226, option 1 in the United States. |  | 131 | 131 units total (130 units in US) | Worldwide Distribution -- United States and Mexico. | Terminated | May 01, 2017 | 188 | 131 | Computer | Computer
Z-0770-2017 | 74587 | No | LBI | Neonatal phototherapy unit. | General Hospital | General Hospital | 510(k) | Natus neoBLUE blanket LED Phototherapy system, neonatal phototherapy. Used to provide treatment for neonatal Hyperbilirubinemia. |  |  | October 28, 2016 | 2016 | Natus Medical Incorporated | neoBLUE blanket LED Phototherapy System distributed since release of the product in 2011 is recalled due to customer complaints regarding discoloration/ degradation of fiberoptic bundle at connection of pad and box. | Device Design | Natus began sending the URGENT RECALL NOTIFICATION / FIELD SAFETY NOTICE neoBLUE blanket system performance letter dated October 27, 2016 and the Technical Bulletin to US customers on October 28, 2016.Natus plans to notify their international customers on December 2, 2016. Customers were informed in March 2015 of a potential issue of discoloration/degradation/melting of the fiber optic bundle at the pad connector which is inserted in the neoBLUE blanket light box. Included with that letter was a Technical Bulletin which described the issue and addressed many of your potential questions.The purpose of this current letter (dated October 27, 2016) is to communicate Natus' plan for a follow-on field correction for neoBLUE blanket systems shipped between 2011 through March 21, 2016 (Serial Numbers up to and including xxxx004282), and to remind customers of the instructions contained in the neoBLUE blanket Technical Bulletin.Natus is in the process of confirming an updated neoBLUE blanket configuration which will not be susceptible to the degradation described above. However, the date of project completion is uncertain at this time. Natus will contact customers when the new configuration is available, at which time Natus intend to initiate a field corrective action to replace or upgrade your device(s).Please contact Natus Quality_Programs@natus.com if you have any further questions. Please call (800) 426-0353. |  | 2619 | 2619 units (1925 units in the US and 694 units OUTSIDE the US) | Worldwide Distribution - USA (nationwide) Distribution and to the countries of : including American Samoa and Guam, and in the following countries:Albania, Argentina, Australia, Austria, Bangladesh, Canada, Chile, Costa Rica, Croatia, Czech Republic, Denmark, Ecuador, Finland, France, Germany, Hong Kong, Indonesia, Iraq, Israel, Italy, Korea, Kuwait, Latvia, Malaysia, Mongolia, Namibia, Netherlands, Nigeria, Norway, Panama, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Singapore, Slovakia (Slovak Republic), Slovenia, South Africa, Spain, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Uruguay, Viet Nam, and Virgin Islands (British). | Open, Classified |  |  | 2619 | Computer | Not_Computer
Z-1690-2016 | 73691 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | 1) 24 G x 0.75 in. BD Angiocath" Autoguard" shielded IV |  |  | March 18, 2016 | 2016 | Becton Dickinson &amp; Company | The device may have a defect in the catheter. In some instances this defect could result in catheterseparation or breakage. | Device Design | BD isssued an urgent voluntary product recall lettter/recall response card on 3/18/2016. |  | 1305150 | 1,305,150 units | Nationwide | Terminated | July 19, 2017 | 488 | 1305150 | Computer | Not_Computer
Z-0596-2016 | 72879 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Philips Lumify Diagnostic Ultrasound, Catalogue Number: 795216 |  |  | December 09, 2015 | 2015 | Philips Ultrasound, Inc. | Color Flow direction is displayed incorrectly in Lumify 1.0. The system displayed Color Flow direction does not correctly represent the annotated Color Bar or Velocity Markers. (e.g.) When the Color Bar conveys that Red is to be displayed for color flow toward the Transducer, the system displays Blue for color flow toward the Transducer. This could result in misdiagnosis in some studies. | Software design | Philips began mailing the URGENT - Medical Device CorrectionPhilips Model Lumify Ultrasound System, FSN 79500372, dated 2015 DEC 11, to consignees.Consignees are instructed to do the following:1. Check the version of software on your system using the instructions provided above on How to Identify Affected Products. Philips has corrected this issue in version 1.0.1 and higher.2. Connect the tablet to the internet.3. The system will alert you to download the upgrade from Google Play.4. Follow all on screen to upgrade the software version.If consignees cannot promptly connect to the internet to download the upgraded software, Philips recommends that you do not use Color Flow. In accordance with the subscription terms and conditions, you will be required to reconnect the system to the internet within no more than 30 days. If consignees do not connect their device to the internet after 30 days, they will be unable to use the system.Consignees with questions should contact your local Philips representative or Philips Customer Service at 1-800-722-9377. |  | 21 | 21 units in the US only | distributed in CA, CT, ND, NV, OR, RI, TN, and WA. | Open, Classified |  |  | 21 | Computer | Computer
Z-0766-2017 | 75645 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Empowr PS Knee System Box Cut Guide |  |  | September 19, 2016 | 2016 | Encore Medical, Lp | It has been discovered that during Empowr PS Total Knee replacement, while chiseling for bone removal in the box area, femoral chipping has occurred in surgery. | Device Design | DJO Global sent an Urgent Field Safety Notice dated September 19, 2016, to all affected customers. The field safety notice provided a newly revised Surgical Technique to educate users on proper use of the device. Customers were instructed to become familiar with the new surgical techniques and notify all of their applicable users of the revised methods. Customers were also instructed to sign the Acknowledgement and Receipt Form. Customers with questions were advised to call 760-734-3551.For questions regarding this recall call 512-832-6302. |  | 162 | 162 units | Nationwide Distribution | Open, Classified |  |  | 162 | Computer | Not_Computer
Z-0642-2016 | 72597 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS 4600 Chemistry System, Catalog Number 6802445, Unique Device Identifier Number 10758750012343, Software Version 3.2 and below; IVD. |  |  | November 03, 2015 | 2015 | Ortho-Clinical Diagnostics | Software anomaly may allow VITROS Systems to sample and process assays while the fluids and reagent temperatures are not in the required temperature range for optimal processing. When this intermittent issue occurs, the VITROS Systems continue to operate and process results without notification to the User, possibly producing erroneous results. | Software design | Ortho-Clinical Diagnostics sent an Urgent Product Correction Letters (Ref. CL2015-211, dated November 3, 2015) via FedEx overnight courier and/ or ORTHO PLUS e-Communications and/ or US Postal Service Priority Mail (only to those consignees who provided a PO Box as an address) to all customers who received the affected devices (VITROS 3600 and 5600 Systems). The notification identified the affected products, problem and actions to be taken. Foreign affiliates were informed via e-mail on 11/03/2015 of the issue and instructed to notify their consignees of the issue and required actions.Ortho-Clinical Diagnostics sent an Urgent Product Correction Letters (Ref. CL2015-221, dated November 11, 2015) via FedEx overnight courier and/ or ORTHO PLUS e-Communications and/ or US Postal Service Priority Mail (only to those consignees who provided a PO Box as an address) to all customers who received the affected devices (VITROS 5,1 FS and 4600 users). The notification identified the affected products, problem and actions to be taken. Foreign affiliates were informed via e-mail on 11/11/2015 of the issue and instructed to notify their consignees of the issue and required actions. |  | 427 | Domestic: 157, Foreign: 270 | Worldwide Distribution - US Nationwide and the countries of Puerto Rico, and foreign distribution to: Argentina, Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Israel, Italy, Japan, Kenya, Mexico, Panama, Philippines, Poland, Romania, Russia, Singapore, Spain, Taiwan, Turkey, and Venezuela. | Open, Classified |  |  | 4399 | Computer | Computer
Z-1688-2016 | 73883 | No | OOE | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | LenSx Laser System. Indicated for use in patients undergoing cataract surgery and refractive surgery. |  |  | April 14, 2016 | 2016 | Alcon Research, Ltd. | Alcon is conducting a voluntary recall of all LenSx Lasers with Corneal Flap Capability after receiving reports concerning incomplete corneal flap creation, which manifests as an un-dissected area(s) within the side cut. | Software design (manufacturing process) | Recall notification letters were sent to consignees on April 14, 2016. The letter details precautions to be taken by the surgeon and requests the consignee sign and return an acknowledgement form. |  | 81 | 81 units | United States, Albania, Australia, Brazil, Chile, Costa Rica, Czech Republic, France, Germany, Greece, Hungary, India, Italy, Romania, Slovakia, Spain, Taiwan, Thailand, Turkey, UAE, Ukraine | Open, Classified |  |  | 81 | Computer | Not_Computer
Z-0672-2016 | 72202 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Accusorb MRI |  |  | October 06, 2014 | 2014 | MWT Materials, Inc. | When not used correctly, Accusorb MRI products may overheat and lead to patient injury. | Device Design | MWT issued a letter dated 10/4/2014 via email to all distributors for further distribution to consignees. Tghe letter identified the affected product, problem and actions to be taken. Questions can be direct to info@MWT-Materials.com. |  | 1 | 1 units | US Nationwide Distribution in the states of NY and MA. | Open, Classified |  |  | 187 | Computer | Not_Computer
Z-0689-2016 | 72965 | Yes | NIO |  |  | Cardiovascular | PMA | ZilverÂż 518RX Vascular Stent with Rapid Exchange Delivery System |  |  | December 22, 2015 | 2015 | Cook Medical Incorporated | Manufacturing of the joint between the inner catheter and stylet of the delivery system does not consistently meet specifications, and the joint may separate during deployment. Failure of this joint during a procedure may result in a partial stent deployment. The clinical effect of a partial deployment may be a need for surgical intervention, or may result in vessel trauma or an embolic event | Device Design | Cook Medical sent an "URGENT: MEDICAL DEVICE RECALL" letter dated December 22, 2015, to all affected customers. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Should you have any questions or concerns, please do not hesitate to contact Cook Medical Customer Relations at 1-800-457-4500 or 1-812-339-2235.On 2/5/2016 a second recall notification was sent to affected consignees, expanding the recall to All Lots of the affected product. |  | 753 | 753 | Nationwide Worldwide | Terminated | July 25, 2017 | 581 | 753 | Computer | Not_Computer
Z-0611-2017 | 75497 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | iConnect Access software. The firm name on the label is Merge Healthcare, Hartland, WI. |  |  | May 25, 2016 | 2016 | Merge Healthcare, Inc. | Software displayed incorrect prior reports in the viewport area, only when more than one prior study (2 or more) was viewed. | Software design | The recalling firm issued a notification letter dated 1/27/2016 via email or via certified mail if the firm did not have an email address for the customer. Nonresponders were issued a second letter dated 5/25/2016 on 5/26/2016. |  | 60 | 60 sites potentially have the affected versions | Distribution was made to medical facilities nationwide. Foreign distribution was also made. | Open, Classified |  |  | 60 | Computer | Computer
Z-0755-2016 | 72391 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ADVIA Chemistry XPT Systems, Software Version 1.0.3, SMN 11127538 |  |  | October 07, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc. | Multiple Software issues. Auto Start-UP Fail, Calibration Interval Resets when a Reagent Blank is run, Control Definition screen assumes range defined is 2 SD, Laboratory Information System (LIS) communication/Laboratory Automation (LAS) issue, Printer Driver Resets, ISE Calibration Ranges are too conservative for Urine Sodium, Archiving and deletion may fail and Workstation services may restart. | Software design | On October 7, 2015 Siemens Healthcare Diagnostics distributed Urgent Field Safety Notices to their custoemrs via Fed Ex in the U.S. and via email to regional countries for implementation outside of the U.S. Customers were provided a work-around if they have experienced any of the following problems: 1) Intermittent auto-startup function fails, 2) calibration interval resets as though a full calibration was performed, 3) Users may be unaware that the QC definition screen with +/- suggests a 2 Standard Deviation range value, 4) a rescanned and re-queried sample may become stuck in a communication loop and all tests may read as "inprocess" if the the initial reaction process is incomplete, 5) The printer stops printing if multiple printouts are requested in quick succession, 6) the ISE calibration for Urine Sodium (Na) may fail due to the ranges being set conservatively during development, 7) The automatic or manually started, archiving and deletion maintenance activity may fail and 8) A workstation services restart screen may appear under multiple scenarios causing the interface to be temporarily inaccessible. If customers have experienced any of the previously mentioned workflow problems, please perform the following actions: 1. Auto Start Failing: Manually perform the Startup Wash, run controls and any calibrations that are due. 2. Calibration interval resets when an RBL is run: Calibration should be run according to the recommendations in the Instructions for Use (IFU) for each assay. Keep a record of required calibrations and ensure that all assays are calibrated before use. 3. Control definition screen assumes that the range is defined as 2 SD values: When creating QC definitions, enter the values associated with a +/- 2 Standard Deviation (SD) range in the +/- ranges field.4. Laboratory Information System (LIS) communication / Laboratory Automation (LAS) issue: Be aware of errors in the event log and take action as needed. Manual intervention and rerunning |  | 87 | 87 units | Worldwide Distribution: US (nationwide) in states of: AZ, CA, MA, NY, TX &amp; WA, and to countries of: Bahrain, Egypt, Slovakia, France, Spain, Italy, Norway, Singapore, New Zealand, Germany, Australia, Republic Korea, Netherlands, Denmark, Finland &amp; United Kingdom. | Open, Classified |  |  | 87 | Computer | Computer
Z-0736-2016 | 72972 | No | GGP | Factor deficiency test. | Hematology | Hematology | 510(k) | HemosIL Plasminogen, Part No. 0020009000, IVD. Instrumentation Laboratory Company. Kit contents: 2 x 2.5 mL Streptokinase reagent and 2 x 2 mL Chromogenic substrate |  |  | December 18, 2015 | 2015 | Instrumentation Laboratory Co. | This Lot is not meeting labeled on-board instrument stability claims of 5 days. | Device Design | Instrumentation Laboratory Co. sent an " Urgent Medical Device Removal" letter dated December 2015, was sent on December 18, 2015. via USPS to US and Canadian customers. The notification requires US and Canadian customers to respond using a Mandatory Response Tracking Form. A Field Safety Notice was e-mailed on December 18, 2015, to the affected IL affiliates, notifying them of the product lot removal due to a stability issue and requesting them to translate the Field Safety Notice, to notify their Competent Authorities and to provide the notification to their affected customer base and secondary distributors. For further questions, please call 1(800) 955-9525. |  | 305 | Domestic: 108 kits, Foreign: 197 kits | Worldwide Distribution - US Distribution and to the countries of : Canada, Argentina, Austria, Belgium, Brazil, China, Czech Republic, France, Germany, Hong Kong, Italy, Lithuania, Panama, Poland, Russia, Saudi Arabia, United Kingdom, and Venezuela. | Open, Classified |  |  | 305 | Computer | Not_Computer
Z-1910-2016 | 74070 | No | NBH | Arthroscope. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker FlowPort II Adapter: Model number: 00CAT00778 |  |  | May 04, 2016 | 2016 | Stryker Corporation | Complaints were received for the Stryker FlowPort II Adapter, and investigation found that the scope sat in the adapter 180 degrees in the wrong direction | Device Design | A recall letter was not sent as all affected devices have been returned. A ship hold was placed on the device on October 16, 2015. |  | 21 | 21 units | GA, UT, MD, NJ, MT | Terminated | June 03, 2016 | 30 | 21 | Computer | Not_Computer
Z-0745-2016 | 73023 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | DS/US Proton Feature with Pinnacle3 Software Version 10.0.0 and 14.0.0 Radiation Therapy Planning System Model #459800200621 |  |  | January 08, 2016 | 2016 | Philips Medical Systems | In DS/US proton planning, there is a correction factor used internally to the dose engine that is calculating improperly in some cases. This can result in an incorrect dose calculation. | Device Design | Philips sent an "Urgent Medical Device Correction" letter dated 16-DEC-2015 to affected customers. The letter described the Affected Products, Problem Description, Hazard Involved, How to Identify Affected Products, Action to be Taken by Customer / User, Actions Planned by Philips and Further Information and Support (contact Philips Customer Care Solutions Center at 1-800-722-9377 ). |  | 3 | 3 | US: Nationwide Distribution in the states of FL, OH, and MO. | Open, Classified |  |  | 3 | Computer | Not_Computer
Z-0748-2016 | 72959 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | AQURE System; Model Number: 933-599. The AQURE System manages blood gas and immunoassay analyzers. |  |  | December 29, 2015 | 2015 | Radiometer America Inc | The AQURE System has a design error regarding sample type in which sample type may be specified for some results and left blank for other results. The fact that the sample type field is left blank in some cases could lead to misinterpretation of results. | Device Design | Customers were sent a notification letter on 12/29/15.The notification states that the AQURE System can be used, but to eliminate the potential confusion, blank fields for sample type are to be interpreted as not specified. Customers are told that a Radiometer representative will be in contact to schedule a visit or a remote session. During the visit or remote session the representative will run a database script on the AQURE system which will serve for two purposes:1. Update existing patient results where the sample type field is blank to include Notspecified. 2. Install a so-called database-trigger, which will include the text Not specified in ablank sample type field for future patient results.An upgraded version of the AQURE software will include modified device driversto ensure that Not specified is inserted if the device transmits an empty fieldfor sample type. The new software version will be installed by local engineerwhen available. |  | 375 | 375 | Worldwide Distribution-US (nationwide) including the states of PA, GA, MI, OH, VA, FL, AL, TX, OR, OK, TN, NY, and MA, and the countries of Austria, Czech Republic, Denmark, Estonia, United Emirates, Kuwait, Qatar, Saudi Arabia, Oman, Jordan, Finland, France, Germany, Hungry, Ireland, United Kingdom, Italy, Netherlands, Norway, Poland, Spain, Sweden, Switzerland, Australia, New Zealand, China, Hong Kong, Taiwan, Japan, Russia, Singapore, South Africa, Republic of Korea, and Malaysia. | Terminated | March 29, 2016 | 91 | 375 | Computer | Computer
Z-0645-2017 | 75478 | No | CHP | Estradiol test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | LIAISON Estradiol II Gen, 310680, 310680-CN |  |  | July 13, 2016 | 2016 | Diasorin Inc. | DiaSorin is issuing an urgent medical device correction for the LIAISONÂż Estradiol II Gen assay. Due to the risk of a recently identified cross reactivity, the LIAISONÂż Estradiol II Gena assay should not be used to test patients being treated with fulvestrant (Faslodex). | Device Design | DiaSorin Inc. sent an Customer Notification letter dated July 13, 2016, to all affected customers. The letter described the problem and the product involved in the recall. The letter listed the "Recommended Actions" and requested customers to complete and return the response portion. For questions contact Product Support at 1-800-328-1482. |  | 10739 | 10,739 | USA (nationwide) Distribution and Worldwide Distribution including the states of : AR, AZ, CA, FL, ID, NJ, NV, NY and PA.., and to the countries of : Saudi Arabia, Colombia, Mauritius, Iraq, Mongolia, Greece, Albania, Austria, Bosnia, Bosnia Herzegovina, Belarus, Bulgaria, Macedonia, Poland, Serbia, Dominican Rep., Tunisia, Paraguay, Hungary, Jordan, Egypt, Belgium, Caribbean, Germany, Brazil, Mexico, Italy, India, South Korea, Cyprus, Bangladesh, Morocco, Chile, Bahrain, Palestine, Utd.Arab.Emir., Kosovo, Uruguay, Venezuela, Hungary, Australia, Vietnam, Sri Lanka, Turkey, Vietnam, Romania, Mauritius, Ecuador, Lebanon, Nepal, Argentina, Hong Kong, China and Canada. | Open, Classified |  |  | 10739 | Computer | Not_Computer
Z-0692-2016 | 72914 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Emergency Transport Ventilator |  |  | December 22, 2015 | 2015 | Draeger Medical, Inc. | Drager became aware of situations where the error message "Poti unplugged" was generated causing acoustic and visual alarms, the breathing system releases pressure and the ventilation function stops operating. | Device Design | The firm, Draeger Medical, Inc., sent an "Urgent Medical Device Recall" letter dated 12/21/2015 and a copy of the Instructions for Use (IFU) Supplement to consignees in a method that confirms delivery. The letter described the product, problem and actions to be taken. The consignees were instructed to include a copy of the supplement with their Oxylog IFU; inform all affected users in their facility; and contact Draeger Customer Support at 1-800-543-5047 (press 4 at the prompt) to obtain additional copies of the IFU Supplement Free of Charge.If there are any questions regarding this letter, please contact Michael Kelhart at 1-800-543-5047 (press 1 at the prompt, then 2, then 32349). For questions regarding the operation and/or servicing of your DrÂżger Oxylog ventilator please contact DrÂżgerService Technical Support at 1-800-543-5047 (press 4 at the prompt).On 12/09/2016, Draeger mailed an Urgent Medical Device Recall notification letter to affected customers to inform them of a revised new software for the Oxlog 3000/3000plus (version 1.06 for Oxylog 3000 plus and version 1.23 for Oxylog 3000) that reduces the impact of the error condition. Customers will be contacted by a Drager Service Representative to schedule a time to perform the software update free of charge. The new software changes the device behavior in the event of the error (a controller failure). With the new software, in the event of a controller failure, ventilation will not stop. The device will continue to ventilate with the last setting. An alarm will annunciate and be displayed, alerting the user of the controller failure. |  | 402 | 402 units (314 - Oxylog 3000; 88 - Oxylog 3000 Plus) | Nationwide Distribution -- including states of AR, AZ, CA, CT, FL, GA, IA, ID, IL, IN, KS, KY, LA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WI, WV, and WY. | Terminated | April 05, 2017 | 470 | 402 | Computer | Computer
Z-1650-2016 | 73832 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | NeuViz 64 Multi-Slice CT Scanner System (consist if two variants: NeuViz 64e, NeuViz 64i) |  |  | May 02, 2015 | 2015 | Neusoft Medical Systems Co., Ltd. | It was found by R&D; in April, 2014 that there was a defect in NeuViz 64 system software version 1.0.5+P09. The icons of side decubitus didn't meet the patient position description, when the scanner position was set to "Right" in the system setting - scanner options. | Device Design | An upgraded software version (1.0.5.2619+P11) was released on May 7, 2015. The FCO was released to the service engineers of Neusoft Medical Systems USA, which is subsidiary of the reporting firm, and the engineers performed FCO for each affected device together with the distributors' engineers. The corrections (FCOs to upgrade software) on all the affected systems in the USA have been completed. At present (4/27/16) there is no longer this issue on the systems, so no further actions are needed for customers/distributors. However, the recalling firm shall inform the direct accounts of this fix before May 5, 2016. The recalling firm will submit the communication evidence ASAP once they have evidence of the communication. For further questions, please call (281) 453-1205. |  | 7 | 7 units | US Distribution - Including Puerto Rico and the states of IL, CT, SC, NE | Open, Classified |  |  | 7 | Computer | Computer
Z-1654-2016 | 73667 | No | JFY | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA CHEMISTRY ENZYMATIC CREATININE_2 |  |  | March 17, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has become aware of N-Acetylcysteine (NAC) and Metamizole (Dipyrone) interference with Trinder and Trinder-like reaction assays. Siemens has confirmed that falsely depressed results may occur on samples drawn from patients receiving N-Acetylcysteine (NAC) or Metamizole. NAC is the accepted antidote for acetaminophen toxicity and is justified in patients at signific | Device Design | Siemens mailed an Urgent Medical Device Correction letter to customers making them aware of the problem. Special instructions were also given for the customer to review the letter with their Medical Director and ensure Venipuncture occurs before drug administration of NAC or Metamizole as indicated in the letter. Baseline assay values before administration of NAC or Metamizole therapy would not be affected. Customers were asked to complete and return the Field Correction Effectiveness Check Form attached with the letter within 30 days. Furthermore, if customers received any complaints of illness or adverse events associated with the specified lots, they were to immediately contact their local Siemens Customer Care Center or local Siemens technical support representative. |  | 137422 | 137422 units total | Nationwide: AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY, Puerto Rico, St. Thomas, VI, Guam, Pacific Mariana Islands | Open, Classified |  |  | 137422 | Computer | Not_Computer
Z-0746-2016 | 73071 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Astral 100, 150 |  |  | January 13, 2016 | 2016 | Resmed Corporation | Resmed Corporation is recalling Astral 100, 150 External Battery because of a start-up issue. | Software design | ResMed sent an Urgent Safety Alert dated January 13, 2016, to all affected customers. The Alert identified the product, the problem, and the action to be taken by the customer. The notification letter provided the following instructions to distributorsField devices can be upgraded using ResMed Service Software or at an authorized ResMed Service Center. Contact Astral Support for further details.The SR2 software update will be available free of charge and should be implemented for all devices within the two-year service cycle or before at the customers preference.To acknowledge receipt of this Astral Safety Alert, submit this form.For any questions call Astral Support:855-245-4640 |  | 14736 | 14,736 | Worldwide Distribution - US including VA, FL, MO, CA, LA, TX, NJ, NH, NC, WV, MI, ID, OK, MI, MN, UT, CT, IL, AZ, UT, KY, MT, PA, RI, TN, DE, NE, LA, NY, SC, WY, IN, VA, GA, CO and Internationally to CA, CO, and AR. | Terminated | April 22, 2016 | 100 | 14736 | Computer | Computer
Z-1914-2016 | 73887 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS 5600 Integrated System, Catalog Number 6802413, Unique Device Identifier No. 10758750002740; and VITROS 5600 Integrated System (refurbished), Catalog Number 6802915, Unique Device Identifier No. 10758750007110; IVD. |  |  | April 06, 2016 | 2016 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics confirmed a software timing anomaly in which the VITROS 3600 and 5600 Systems may unexpectedly aspirate from or dispense into an unintended sample that may lead to erroneous or contaminated results. | Software design | On 4/6/2016,URGENT PRODUCT CORRECTION NOTIFICATION letters (Ref. CL2016-076, dated 4/06/2016) were sent via FedEx overnight and/or US Postal Service Priority Mail (only to those consignees who provided a PO Box as an address) to all customers who received the affected product and advised them of the issue and the required actions. Follow enclosed instructions to help decrease the probability of the software anomaly until the software update is installed. Discuss any concerns with previously report results with their Laboratory Medical Director to determine the appropriate course of action. Install Software Version 3.2.3 at their earliest convenience. Foreign affiliates were informed of the issue by e-mail on 4/6/2016, and instructed to notify their consignees of the issue and required actions. |  | 7881 | VITROS 5600: US: 1161 units, Foreign: 1063 units; VITROS 5600 Refurbished: US: 19 units, Foreign: 38 units | Nationwide, Puerto Rico and Foreign distribution to the following countries: Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Singapore, Spain, and Venezuela. | Open, Classified |  |  | 8686 | Computer | Computer
Z-0660-2017 | 75664 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco RTP System; Used to make treatment plans for patients with prescriptions for external beam radiation therapy. |  |  | November 03, 2016 | 2016 | Elekta, Inc. | Incorrect Dose when using the reset function. | Software design | Important Field Safety Notice 382-01-MON-006 was sent to all affected customers on 11/3/2016. The notice informs users of the specific product and version numbers affected by the issue, and any work around that can be used to avoid the issue. An acknowledgement form was included which is to be returned to Elekta. |  | 671 | 671 | AK, AZ, CA, CO, FL, GA, IL, IN, MA, MI, MN, MO, NV, NH, NY, NC, OH, OK, OR, PA, SD, TX, UT, VA, WA, WI, Algeria, Australia, Austria, Brazil, Canada, China, Cuba, Cyprus, Czech Republic, Ecuador, Egypt, France, Germany, Greece, Hong Kong, Hungary, India, Israel, Italy, Japan, Laos, Malaysia, Mexico, Morocco, Myanmar, Nepal, Netherlands, New Zealand, Norway, Peru, Poland, Portugal, Russia, Serbia, Singapore, South Africa, Spain, Sweden, Switzerland, Thailand, Turkey, United Kingdom, Venezuela, | Open, Classified |  |  | 671 | Computer | Computer
Z-0614-2017 | 75267 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge RadSuite,f/k/a: Emageon Advanced Visualization , RadSuite versions 8.30.7.8, 8.30.7.9, 8.30.8, and 8.31. |  |  | July 14, 2016 | 2016 | Merge Healthcare, Inc. | It was reported by a customer that RadSuite images are not appearing as readable images. This product malfunction has the potential to result in delay in care and a possible re-exposure of the patient to radiation (x-ray). | Software design | Merge sent an Urgent Medical Device Recall letter dated October 19, 2016,(Recall #2016-075) to all affected customers. Initially, before it was considered a recall, the firm provided customers with a software upgrade to fix the issue. The letter described the Product, Issue, Potential Harm, Containment by the Customer/User, Action by Merge and Actions by Customers. Requested consignees to complete and return the response form. For questions email recall@merge.com. Customers with questions were instructed to send an email to recall@merge.com.For questions regarding this recall call 262-367-0700. |  | 29 | 29 | Nationwide Distribution including AL, CA, FL,IN, MI, MO, NY, PA, TN, and WI. | Open, Classified |  |  | 29 | Computer | Computer
Z-0622-2017 | 75661 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Aquilion Prime TSX-303A |  |  | October 25, 2016 | 2016 | Toshiba American Medical Systems Inc | Toshiba American Medical Systems (TAMS) is recalling the Aquilion CT system because ithas been reported that the cooling oil in the x-ray tube heat exchanger may leak during system operation. | Device Design | Toshiba American Medical Systems, Inc. sent an Urgent Medical Device Correction letter dated October 25, 2016, to all affected customers who received the Aquilion CT system to inform them that Toshiba America Medical System has received reports that the cooling oil in the x-ray tube heat exchanger may leak during system operation. The letter informs the customers of the description found and corrective action. Customers were instructed to complete the attached customer reply form and fax it to (877) 349-3054 or send it to raffairs@tams.com. Customers with questions are instructed to call (800) 421-1968 or contact their local Toshiba Representative at (800) 521-1968.For questions regarding this recall call 714-730-5000. |  | 291 | 291 units | Nationwide Distribution | Terminated | August 19, 2017 | 298 | 291 | Computer | Not_Computer
Z-0603-2017 | 75606 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | AXIOM Artis, Angiographic x-ray system |  |  | November 01, 2016 | 2016 | Siemens Medical Solutions USA, Inc | The possibility exists that the semi-automatic focus switch may not work as intended when pressing the footswitch multiple times. | Software in the Use Environment | Siemens mailed a Customer Safety Advisory Notice to affected customers to inform them of a corrective action they plan to perform to prevent a possible hazard to persons and equipment. The notice identified the associated risks, underlying issues, impact on system operation and actions to be taken. |  | 4657 | 4,657 systems Worldwide | US Nationwide Distribution. | Terminated | June 21, 2017 | 232 | 4657 | Computer | Computer
Z-1705-2016 | 74087 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | iGUIDE System, for patient positioning, with assistance of a 30 Tracking System in a radiotherapy environment. |  |  | May 12, 2016 | 2016 | Elekta, Inc. | If by mistake the initial pre-treatment imaging was performed BEFORE the HexaPOD was moved to the DRIVE (*START) position, iGUIDE offers the possibility to branch into a specific catch-up workflow, which was implemented to use the initial scan data. | Software design | Elekta sent an Important Field Safety Notice that was distributed to all affected customers on May 12, 2016. The notice informs end users about the problem and the potential clinical impact. It also gives a recommendation how the risk can be mitigated. The notice also advises the customer that a patch will be released that will introduce a patch that avoids an inaccurate position and thus eliminates an unsafe situation.Corrective Action #2: Permanent Solution  Software UpgradeA software patch will be released that will correct the behavior of the iGUIDE software. The target release date for this patch is January 2017.Service teams will have 6 months from the date of release to correct all affected devices in the field. For further questions, call (770) 300-9725. |  | 20 | 20 | IL, LA, Austria, Australia, Botswana, Germany, Denmark, France, India, Japan | Open, Classified |  |  | 20 | Computer | Computer
Z-0599-2017 | 75629 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge PACS software. |  |  | March 28, 2016 | 2016 | Merge Healthcare, Inc. | The software did not show unviewed images when the last view was skipped. The physician may potentially misdiagnose when not all images are available for viewing. | Software design | The recalling firm issued a letter dated March 28, 2016 via e-mail or via certified mail to customers. The letter identified the affected product, problem and actions to be taken. The notification instructs the customer to pass the information onto all other users, as well as any downstream customers if it was further distributed. The notification also informs the customer their response is required using the enclosed form and the return addressed envelope. |  | 534 | 534 sites potentially have the affected versions | Worldwide Distribution - US Nationwide including Canada. | Open, Classified |  |  | 534 | Computer | Computer
Z-1637-2016 | 73873 | No | HXC | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Axial Torque Limiting Driver (1st generation) |  |  | February 02, 2015 | 2015 | Bradshaw Medical Inc. | A component failure could cause a situation where the applied torque was significantly higher than the calibrated torque setting. | Device Design | Consignees were sent a Bradshaw Medical "Product Withdrawal Notice" letter, date marked starting February 2, 2015. The letter specified the Part Number, Lot Number (s) and Shipped date for each consignee. The letter described the product and advised consignees to discontinue use of the product. Requested consignees to complete and return the Customer /Distribution verification Form to BMI Regulatory Affairs FAX (262) 925-1380 and to further send the "Withdrawal Notice and Customer Verification Form" to all customers to whom they have distributed the product. |  | 3760 | 3,760 units | US Nationwide Distribution in the states of: CA, FL, IN, MA, MI, OH, PA, TN UT, VA and the country of UNITED KINGDOM | Open, Classified |  |  | 3760 | Computer | Not_Computer
Z-0608-2017 | 75677 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | SOMATOM Definition Flash, Computed tomography x-ray system |  |  | November 02, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Siemens is releasing a software update that provides bug-fixes to improve system performance for customers with SOMATOM Definition AS, SOMATOM Definition Edge, SOMATOM Definition Flash Systems with software version VA48A_SP2 and Care Contrast license correction of potential safety issues. | Software design | Siemens mailed a Customer Safety Notice on November 9, 2016, to customers who were affected by the recall to inform them of a possible malfunction with the systems. The notification explained the underlying issue, the effect if has on the operation of the system, potential risks involved; and the actions they plan to implement to correct the issues. Customers with questions were instructed to call 1-800-888-7436.For questions regarding this recall call 610-448-6471. |  | 7 | 7 systems | Nationwide distribution to MI, NY, CA, KY, ND, and NE. | Open, Classified |  |  | 7 | Computer | Computer
Z-1638-2016 | 73578 | Yes | MSX | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Ascom Mobile Monitoring Gateway (versions 4.1.1 and 4.2.0) and Ascom Unite Connect for Nurse Call (version 2.2.0). |  |  | December 10, 2015 | 2015 | Ascom US, INC. | Due to a malfunction of software, the secondary module will not automatically take over messaging if the primary module fails after a software upgrade to one of the other affected versions. Messages from GE Carescape/Nurse Call will not forward to handset until primary module is restored. | Software design | On January 4, 2016 , the firm is initiating a field corrective action to update the software in the installed base and resynchronize the primary and secondary servers. All customers were sent a letter January 4, 2016, informing them of the voluntary recall of Ascom Mobile Monitoring Gateway and Unite Connect for Nurse Call. The reason stated for why the recall was initiated was that the secondary module would not take over the messaging function when the primary module failed, following a software upgrade. Ascom instructed customers to check if their software was one of the affected versions they should contact their local Ascom representative to review the MMG configuration and determine if the system is subject to malfunction. The same instructions were directed towards those who were not able to determine their software version. If it is determined that the program is subject to malfunction, the primary and secondary modules should be re-synchronized. A test of module redundancy should be conducted on re-synchronized modules. The second module should take over for the primary module within thirty seconds. Model redundancy should be tested frequently and following system upgrades. For more information call (919) 535-6143 , 8am till 4:30pm. |  | 319 | 319 | CA, CT, FL, IA, ID, IL, IN, KY, LA, MA, MD, ME, MI, MO, NJ, NM, OH, OK, PA, TN, TX, VA, WA, WI and Hawaii.Australia, Austria, Belgium, Canada, France, Germany, Netherlands, Poland, Qatar, Sweden, Switzerland, and United Arab Emirates. | Open, Classified |  |  | 319 | Computer | Computer
Z-1634-2016 | 73846 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | CinchLock SS Knotless Anchor with Inserter; |  |  | April 12, 2016 | 2016 | Stryker Corporation | Stryker Endoscopy is recalling the CinchLock SS Knotless Anchor with Inserter and the CinchLock Flex Anchor with Inserter after 13 reports of s of an internal wire breakingupon deployment of the anchor. | Device Design | Urgent Medical Device Recalls letters were sent on April 12, 2016 to US customers by certified mail. Stryker International distribution sites will be notified by e-mail. Letters describe the reason for recall with the affected part and lot numbers. Any recalled product should be segregated and returned to Stryker. |  | 3100 | 3100 total device, both model numbers. | Distributed US (nationwide) and in Australia, Switzerland, Netherlands, Spain and Sweden. | Open, Classified |  |  | 3100 | Computer | Not_Computer
Z-1669-2016 | 73765 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Lumenis M22 System Model Number: GA-0005400 (M22 IPL + YAG Module) with Acne Filter (KT-1014971). |  |  | November 17, 2015 | 2015 | Lumenis Ltd | Lumenis Ltd Announces a Field Action of the M22 IPL Acne Filters for the Lumenis M22 IPL Hand Piece due to the Risk of Superficial Burns When Using the Device. | Device Design | Customers were notified on November 17, 2015 by Customer Notification Letter and customers were asked to return the device. |  | 33 | 33 filters | Distributed in the states of GA, CA, CO, CT, KY, LA, MA, MN, NJ, NY, RI, and SC, and the countries of Italy, Germany, France, and China. | Terminated | November 30, 2016 | 379 | 33 | Computer | Not_Computer
Z-0668-2017 | 75718 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco TRP System |  |  | November 17, 2016 | 2016 | Elekta, Inc. | Contours are not handled correctly if the CT dataset contains slices with DICOM Z coordinates of xx.x50mm. | Software design | Elekta sent an Important Field Safety Notice 382-01-MON-007 was sent to all affected customers on November 17, 2016, to all affected customers. The notice informs users of the specific product version numbers affected by the issue, and any work arounds that can be used to avoid the issue. The notice included a reply form which is to be returned. |  | 2148 | 2148 | Worldwide Distribution - US including AK, AZ, CA, CO, FL, GA, IL, IN, KY, LA, MD, MA, MI, MN, MO, MT, NV, NH, NJ, NY, NC, ND, OH, OK, OR, PA, SD, TX, UT, VA, WA, WI, Puerto Rico and Internationally to Algeria, Antiqua &amp; Barbuda, Argentina, Australia, Austria, Bahrain, Belarus, Botswana, Brazil, Bulgaria, Canada, Chile, China, Columbia, Croatia, Cuba, Cyprus, Czech Republic, Ecuador, Egypt, Estonia, Finland, France, Germany, Greece, Hong Kong, Hungary, India, Indonesia, Iran, Iraq, Israel, Italy, Japan, Laos, Lithuania, Malaysia, Malta, Mexico, Morocco, Myanmar, Namibia, Nepal, Netherlands, New Zealand, Norway, Panama, Peru, Philippines, Poland, Portugal, Romania, Russia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Sweden, Switzerland, Thailand, Tunisia, Turkey, Unite Kingdom, Ukraine, Venezuela, Vietnam, and Zimbabwe. | Open, Classified |  |  | 2148 | Computer | Computer
Z-0651-2016 | 70765 | No | MUH | Extraoral source x-ray system. | Dental | Radiology | 510(k) | Vantage Panoramic X-Ray System used to perform dental and orthodontic radiological imaging |  |  | September 04, 2014 | 2014 | Midmark Corp dba Progeny Inc | Failure of the primary collimator may result in unintentional movement of the device. The unintentional movement will not result in the generation of an error message. | Device Design | MidMark sent an Urgent Medical Device Field Correction letter dated September 4, 2014, to all affected customers. The letter included instructions for distributors to: 1) encourage end users to contact Midmark to arrange the field correction after the final solution becomes available; 2) if a partial image is reported to have occurred with a machine, the machine should be removed from use and the end user should initiate a service request per the normal procedure; 3) quarantine any affected machines in the distributor's inventory and contact Midmark Technical Support within 10 business days to provide the Serial numbers to on-site servicing of the affected machines can be scheduled. Distributors with questions about this market action can contact Midmark Technical Support by calling 888-294-3800 x 3 or by e-mail at imagingtechnsupport@midmark.com. |  | 832 | 832 systems | Worldwide Distribution - US including AL, AR, AZ, CO, CT, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MI, MN, MO, MS, NE, NJ, NM, NV, NY, OH, OK, PA, SC, TN, TX, UT, VA, WA, WI; and Internationally to Australia, Belgium, Canada, Czech Republic, France, Hungary, Lithuania, New Zealand, Poland, Russia, Saudi Arabia, South Africa, Switzerland, Turkey, United Arab Emirates, United Kingdom, and Vietnam. | Open, Classified |  |  | 832 | Computer | Computer
Z-0712-2016 | 72599 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Allura Xper FD20 Biplane OR Table; Model Numbers: 722025 722020 |  |  | November 03, 2015 | 2015 | Philips Electronics North America Corporation | The use of Allura Xper, Allura Clarity, Allura CV, Allura Centron and UNIQ systems regarding live images and still images can lead to still images being interpreted as live Images. | Device Design | The firm, Philips, issued "URGENT-Field Safety Notices" dated 10/26/2015 and 11/19/2015 to its customers. The notices described the product, problem and actions to be taken. The customers were instructed to ensure the audible signal (X-ray buzzer) is always enabled, to activate only ONE fluoroscopy pedal at time, and to configure the footswitch to the default Philips configuration. Customers were also instructed to review the information with all members of your staff who need to be aware; retain a copy with the equipment IFU and complete and return the Confirmation Respond form to Mail Stop 4302, 3000 Minuteman Road, Andover, MA 01810 or email: Debra.Gibbs@Philips.com. Philips is tracking the delivery of these notices via certified return receipt.If you need any further information or support concerning this issue, please contact your local Philips representative: Technical Support Line: 1-800-722-9377. |  | 13297 | 13297 in total | Worldwide Distribution: US (Nationwide) including Puerto Rico and countries of: Albania, Algeria, US Virgin Islands, Argentina, Australia, Austria, Bahrain, Bangladesh, Belarus, Belgium, Bosnia and Herzegovina, Brazil, Brunei Darussalam, Bulgaria, Canada, Cayman Islands, Chile, China, Colombia, Democratic Congo, Costa Rica, Croatia, Cuba, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, Estonia, Fiji, Finland, France, French Polynesia, Germany, Greece, Guadeloupe, Guatemala, Hong Kong, Hungary, Iceland, India, Indonesia, Iran, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Korea, Republic of Kuwait, Latvia, Lebanon, Libya, Lithuania, Macedonia, Malaysia, Malta, Martinique, Mauritius, Mexico, Moldova, Morocco, Myanmar, Nepal, Netherlands, New Zealand, Norway, Oman, Pakistan, Palestine, Panama, Peru, Philippines, Poland, Portugal, Qatar, RÂżunion, Romania, Russian Federation, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sudan, Sweden, Switzerland, Syria, Syrian Arab Republic, Taiwan, Tanzania, Thailand, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Uzbekistan, Venezuela, and Vietnam. | Open, Classified |  |  | 13297 | Computer | Computer
Z-1816-2016 | 74038 | No | GEA | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Various trauma and sports medicine instruments and implants. |  |  | April 01, 2016 | 2016 | Biomet, Inc. | Insufficient seal strength on sterile barrier packaging.The seal could be jeopardized causing the sterile barrier to be compromised and the product could become nonsterile.If the product becomes non-sterile and is used or implanted, it may lead to an infection. | Equipment maintenance | Zimmer Biomet sent an Urgent Medical Device Recall Notice dated April 1, 2016, to all affected customers. Customers were instructed to locate and return the affected product, and implanting surgeons were instructed to monitor patients. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Questions related to this notice should be directed to (574) 372-1570, Monday through Friday, 8 a.m. to 5 p.m. |  |  |  | Worldwide Distribution - US (Nationwide) and globally to Argentina, Australia, Brazil, Canada, Chile, Colombia, Costa Rica, Ecaudor, The Netherlands, Japan, Korea, Malaysia, Mexico, Peru. | Terminated | May 23, 2017 | 417 |  | Computer | Not_Computer
Z-0707-2017 | 75270 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge, eFilm Workstation and eFilm Lite |  |  | October 20, 2016 | 2016 | Merge Healthcare, Inc. | A product issues happens for RF projection images. If user measures on RF image, .cal (Calibration) is not displayed and measurement is not correct. The .cal label is not displayed on the projection images after calibration is performed which should be present if the measurement is presented based on Pixel Spacing. | Software design | Merge Healthcare sent an Urgent Medical Device Recall letter dated October 12, 2016, to all affected consignees. The letter described the Product, Issue, Potential Harm, Containment by the Customer / User, Actions by Merge, and Actions by Customers. Advised consignees to discontinue using the product, and identify &amp; notify all customers. Requested consignees for a response to the notification. Customers with questions were instructed to send an email to recall@merge.com. |  | 2163 | 2163 | Worldwide Distribution - US: Nationwide including DC, GU, PR and Internationally to Canada, Argentina, Armenia, Australia, Austria, Belgium, Brazil, Chile, China, Colombia, Congo, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Ecuador, Estonia, Finland, France, Germany, Greece, Guatemala, hong Kong, Hungary, Iceland, India, Indonesia, Ireland, Israel, Italy, Japan, Kazakhstan, Latvia, Lebanon, Lithuania, Luxembourg, Malaysia, Mauritania, Mexico, Netherlands, New Caledonia, New Zealand, Norway, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Serbia and Montenegro, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Vietnam, and Yemen. | Open, Classified |  |  | 2163 | Computer | Computer
Z-0695-2016 | 72927 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | D-10 Rack Loader; |  |  | December 15, 2015 | 2015 | Bio-Rad Laboratories, Inc. | On a rare occasion, there is a potential to assign a patient result to an incorrect sample ID when running in D10 Rack Loader configuration. | Software design | Urgent Medical Device Correction letters were sent on December 2015 to all affected consignees and subsidiaries. Each subsidiary will be given a Subsidiary Medical Device Correction Notice, Customer Medical Device Correction Notice, and Return Response Form and the number of customers distributed by e-mail. Each Subsidiary much contact customers and translate the Customer Medical Device Correction Notice and Return ResponseForm as needed. A revised letter to subsidiaries and customers was sent on January 20, 2016. |  | 365 | 165 in US, 200 Internationally | Worldwide Distribution-US (nationwide) and the countries of Canada, Australia, Belgium, Brazil, Brunei Darussalam, China, Czech Republic, France, Germany, Greece, India, Israel, Italy, Malaysia, Maldives, New Zealand, Norway, Austria, Poland, Portugal, Russia, Spain, Sweden, Switzerland, Thailand, and UK. | Terminated | May 02, 2016 | 139 | 365 | Computer | Computer
Z-0700-2017 | 75458 | No | MHY |  |  | Neurology | PMA | Model 37751 Recharger |  |  | October 03, 2016 | 2016 | Medtronic Neuromodulation | Medtronic has identified an increased number of complaints from customers involving reports of Rechargers (Medtronic Model 37751 Recharger, which is included in Models 37754 and 97754 Charging Systems,Spinal Cord Stimulation)that are in an unresponsive error state, where the Recharger is non-functional with a blank display screen and is beeping every 5 seconds. Medtronic has determined all Rechargers manufactured starting in November 2014 (indicated by serial numbers beginning with NKA4 or NKU4 ) are more susceptible to this error state. | Software change control | Consignees were sent a series of documents which included an introductory Medtronic and "Urgent Medical Device Correction" letters dated September 2016. The urgent letters included one for Deep Brain Stimulation (DBS) and one for Spinal Cord Stimulation(SCS) . Each letter was addressed to Healthcare Professional. Only physicians associated with both SCS and DBS therapy received the cover letter. The letters provided information on "Background, Issues Mitigation, Recommendation, Additional Information and Patient Notification (DBS).Requested consignees to complete and return the Physician /HCP Reply Form.For questions contact Medtronic Technical Services at 1-800-707-0933 weekdays 7 am to 6pm, CT. |  | 42887 | 42,887 | Worldwide Distribution - US Nationwide including DCOUS: Canada and others. | Open, Classified |  |  | 42887 | Computer | Computer
Z-0693-2016 | 72926 | Yes | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius Liberty Cycler-Peritoneal Dialysis Cycler |  |  | December 15, 2015 | 2015 | Fresenius Medical Care Renal Therapies Group, LLC | The door latch may not fully close and the door may open unexpectedly. The door may make physical contact with anything near the door of the machine, including the patient. May cause injury, delay in treatment, or Increased Intraperitoneal Volume (IIPV) (overfill). | Device Design | Fresenius Medical Care issued on December 15, 2015, a n Important Product Notification to all affected customers informing them of the problem and how to recognize the proper latch closing. Questions regarding the contents of the Customer Notification or questions regarding the issue, contact Technical Services at 1-800-227-2572 and reference the Field Action Number FA-2015-10-W.A reply form is included as part of the customer communication withinstructions to fill out and return via fax or scanned/emailed copy. |  | 14094 | 14,094 units | Nationwide Distribution | Open, Classified |  |  | 14094 | Computer | Not_Computer
Z-0701-2017 | 75644 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | 3M Universal Electrosurgical Pad, REF 9130 (100 pads / case) and |  |  | October 26, 2016 | 2016 | 3M Company - Health Care Business | 3M is recalling Universal Electrosurgical Pads because a report of an unintended material (process liner) in the product. This may prevent or impede the safe return of electrosurgical current following Electrosurgical Unit (ESU) activation. This defect has the potential to increase the risk of a patient burn directly under the pad or at an alternate site of the body. | Equipment maintenance | 3M sent an Urgent Medical Device Recall letter dated October 26, 2016 to affected customers. The letter identified the affected product, problem, and actions to be taken. Customers were instructed to complete and return the attached Customer Product Return and Recall Form. Customers with questions can contact 3M Health Care Customer Helpline at 1-800-228-3957, 8:30 am - 7:00 pm EST. |  | 32000 | 32,000 pads | Worldwide Distribution - US Nationwide in the states of AR, AZ, CA, FL, GA, ID, IL, KY, NC, NY, TN, TX, VA and the countries of Canada, Chile, Costa Rica, Germany, Japan, Mexico and South Africa. | Terminated | August 01, 2017 | 279 | 32000 | Computer | Not_Computer
Z-1879-2016 | 74164 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco RTP System |  |  | May 18, 2016 | 2016 | Elekta, Inc. | When DICOM exporting a 3D Monaco plan and the "Composite Field Sequencing" (CFS) checkbox is selected, the Y jaws will snap to the port shape when they should remain where they were defined. | Software design | Field Safety Notice 382-01-MON-004 was sent to customers on May 18, 2016. The notice informs users of the specific product and version numbers affected by both issues, and any work around that can be used to avoid the issue. The customers were instructed to complete and return the Field Safety Notice Acknowledgement form to their local Elekta as soon as possible. |  | 1401 | 1401 units | Worldwide Distribution - US Nationwide in the states of AK, AZ, CA, CO, FL, GA, IL, IN, KY, MI, MN, MO, MT, NV, NH, NJ, NY, NC, OK, OR, PA, SD, TX, VA, WA, WI, and the countries of: Alger, Antiqua &amp; Barbuda, Argentina, Australia, Austria, Belarus, Botswana, Brazil, Bulgaria, Canada, China, Columbia, Croatia, Cyprus, Czech Republic, Ecuador Egypt, Estonia, France, Germany, Greece, Hong Kong, India, Indonesia, Iran, Israel, Italy, Japan, Lithuania, Malaysia, Malta, Mexico, Morocco, Nabibia, Netherlands, New Zealand, Norway, Panama, Peru, Poland, Portugal, Russia, Serbia, Singapore, Slovakia, South Africa, South Korea, Spain, Sri Lanka,Sweden, Switzerland, Thailand, Turkey, United Kingdom, Ukraine, Venezuela, Zimbabwe | Open, Classified |  |  | 1401 | Computer | Computer
Z-0696-2017 | 75756 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Ingenuity Core 128 728323 |  |  | November 12, 2016 | 2016 | Philips Medical Systems (Cleveland) Inc | Software error due to the filament on timer | Software design | Philips sent a Customer Information Letter dated November 14, 2016, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers with questions were instructed to contact their local Philips representative or local Philips Healthcare office. For North America and Canada, contact the Customer Care Solutions Center (1-800-722-9377). |  | 509 | 509 | Worldwide Distribution - US including AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, VT, WA, WI, WV, WY and Internationally to Algeria, Argentina, Australia, Austria, Bangladeshm Belgium, Bolivia, Brazil, Cambodia, Canada, Chile, China, Colombia, Croatia, Cuba, Czech Republic, Denmark, Djbouti, Dominican Republic, Egypt, Estonia, Finland, France, Germany, Gibraltar, Greece, Hong Kong, Hungary, India, Indonesia, Iran, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Korea, Republic of Korea, Latvia, Lebanon, Lithuania, Malaysia, Martinique, Mexico, Monaco, Mongolia, Mozambique, Myanmar, Netherlands, New Zealand, Norway, Oman, Pakistan, Palestinian territory, Panama, Peru, Philippines, Poland, Portugal, Qatar, Reunion, Russian Federation, Rwanda, Saudi Aarabia, Senegal, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, Uzbekistan, and Vietnam. | Open, Classified |  |  | 1419 | Computer | Computer
Z-0817-2016 | 73114 | No | GWF | Evoked response electrical stimulator. | Neurology | Neurology | 510(k) | MEE-1000A Neuromaster, Model No. Accessories: MS-110B (Electric Stimulation), MS-120BK (Current Stimulator Box) JB-116B, JB-132B (Amplifiers) Electromyography. (EMG) |  |  | January 19, 2012 | 2012 | Nihon Kohden America Inc | The MEE-1000A Neuromaster had defective printed circuit boards (PCBs). Gas arrestors on these boards which were intended to provide protection from electrostatic discharge failed to prohibit unintentional current flow. | Device Design | A customer notification letter dated 9/25/12 was sent to customers informing them of a potential malfunction of NIHON KOHDEN amplifier units, models JB-116B andJB-132B and Extension Unit MS-110B, used with the MEE-1000A Neuromaster Neural Function Measuring System. The letter informs the customers of the potential malfunction and that the probability that the malfunction occurs is extremely rare. The letter informs the customers of the preventive measure and the preventive corrective action. Customers are instructed to inform all potential users and operators of the Neural Function Measuring System MEE-1000A Neuromaster about the Corrective Action notice. Customers are instructed that NIHON KOHDEN technical support representative will contact them to arrange an update of their amplifier products. |  | 477 | 477 units total (137 units in US) | Worldwide distribution. US nationwide, Australia, Austria, France, Germany, Greece, Japan, Romania, South Africa, South Korea, Spain, and United Kingdom. | Terminated | February 22, 2016 | 1495 | 477 | Computer | Not_Computer
Z-0902-2017 | 75817 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Specimen Gate Laboratory 5002-0180, Specimen Gate eReports 5002-0370, Software Version 1.0 |  |  | December 01, 2016 | 2016 | PerkinElmer Health Sciences, Inc. | Issue may cause eReports to consume an excessive amount of disk space causing the software to stop operating. | Software design | PerkinElmer Health Sciences issued a Medical Device Correction letter dated December 1, 2016, to all affected customers email facilitated through our PerkinElmer Software Services group located in Akron Ohio. Software Services will be asked to make sure the customers receive the Recall Letter and implement the recommended corrective measures.Actions to be takenA field correction of the issue is recommended. PerkinElmer will correct the issue with an update installed through remote connection. You will be contacted by your PerkinElmer representative to make the necessary arrangements for implementing the update.In the interim period, please do not attempt to login multiple times in quick succession. Allow the software time to respond to first login command.Please contact your local PerkinElmer representative or Specimen Gate Support for further information.Other InformationPlease inform those affected in your organization accordingly.To comply with regulatory requirements we request that you complete the enclosed response form and return it by fax to number 1-330-825-8520 (United States) /+358 2 2678 357 (outside United States) or as scanned by e-mail to TurkuQMresponse@perkinelmer.com as soon as possible, but not later than December 22, 2016.For questions regarding this recall call 781-663-6900. |  | 21 | 21 | Worldwide Distribution - US including CO, FL, GA, IL, MA, MI, NV, NJ, OH, SC, AZ, KY and Internationally to Italy, Ireland, UK, Spain, and Peru | Terminated | May 01, 2017 | 151 | 21 | Computer | Computer
Z-0666-2017 | 75336 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Hamilton-MR1 Ventilator: Catalog# 161010 |  |  | September 26, 2016 | 2016 | Hamilton Medical, Inc. | Oxygen tubing and the oxygen connector of the Hamilton-MR1 could potentially become loose during the preparation for ventilation. | Device Design | Hamilton Medical sent a letter dated September 6, 2016, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Hamilton Medical will update each affected Hamilton-MR1 ventilators by installing a new oxygen mixer mounting plate with the correct screws at no cost to each facility. A Hamilton Medical Field Service Technician will contact customers soon to schedule a convenient time to install the hardware upgrade. Customers may also call 1-800-426-6331, option #2 to be issued an RGA for return to the service center for the upgrade kit to be installed. Customers with questions should call 817-909-0308. |  | 44 | 44 units | Nationwide Distribution | Open, Classified |  |  | 44 | Computer | Not_Computer
Z-0589-2016 | 72652 | Yes | DWC | Cardiopulmonary bypass temperature controller. | Cardiovascular | Cardiovascular | 510(k) | Maquet HCU 30 Coolant Type R134A Filling weight 250 g 120 V 50-60 Hz 1920 W |  |  | November 13, 2015 | 2015 | Maquet Cardiovascular Us Sales, Llc | Failure of the HEATER-COOLER UNIT (HCU 30) Type 2 power supply board. A drop of more than 10% voltage can cause overheating and subsequent failing of the R56 thermistor of the power supply board. Board failure typically occurs when the compressor is switched on (prior to patient use); however, it can occur during patient use. This can shut down the compressor or the entire HCU 30 Type 2 unit. | Device Design | MAQUET sent "Product Recall Urgent Medical Device Field Correction" notices and "Medical Device Field Correction Response Forms" dated 11/13/2015 via UPS to affected customers. The notification identified the issue and how to identify the affected unit. The notification instructed customers to always keep a replacement heater-cooler unit on standby in the event of a unit failure; complete the attached Field Correction Response form (fax to 1.973.629.1518 or email to HCU30@maquet.com; and that they would be contacted by a representative of the Maquet Service Team to schedule an onsite service to determine if the power supply board in their HCU 30 Type 2 unit is affected adn replace the power supply board as necessary. If the customer had any questions they are instructed to contact their Maquet Service Representative or Maquet Technical Support at 1.888.627.8383 (Press option 3, followed by Option 1 and then option 1 again), Monday through Friday between the hours of 8:00 am and 5:00 pm. |  | 100 | 100 Units (distributed in United States) | United States Nationwide Distribution | Open, Classified |  |  | 100 | Computer | Computer
Z-0683-2017 | 75063 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Lineage Liners: |  |  | June 21, 2016 | 2016 | MicroPort Orthopedics Inc. | Revision rate trends from within the MicroPort Orthopedics Inc. database evaluated for the modular head Metal on Metal THA systems show an increasing overall trend from 2009 to present and it was found that there was a specific hazard/harm for suspected tissue reaction to metal debris over one percent for modular head CONSERVE(R) systems and DYNASTY(R)metal acetabular liner systems. | Device Design | MicroPort sent an Urgent Field Safety Notice dated June 22, 2016, to all affected customers notifying them of the change in the instructions for use on July 7th, 2016 by email. The firm also notified US hospitals where the implantation surgeries occurred within the last twelve (12) months by letter on July 13th, 2016 delivered by FedEx Priority. Customers were asked to complete and return the enclosed form even if they do not have any affected product to return by fax to 1-901-451-6032 or PostMarket@ortho.microport.com.For questions regarding this recall call 901-867-4771. |  | 5009 | 5,009 units | Worldwide Distribution - US (nationwide) Argentina, Australia, Austria, Belgium, Brazil, Bulgaria, Canada, China, Columbia, Cyprus, Czech Republic, Denmark, Finland, France, Georgia, Germany, Greece, Iran, Italy, Jamaica, Japan, Lithuania, Malaysia, Philippines, Poland, Portugal, Russia, Slovenia, South Africa, South Korea, Spain, Switzerland, Taiwan, Netherlands, Turkey, Arab Emirates, United Kingdom, and Vietnam. | Terminated | July 01, 2017 | 375 | 513126 | Computer | Not_Computer
Z-0676-2017 | 75637 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | ARROW OnControl Bone Marrow Biopsy System Comprehensive Tray, Sterile, The Arrow OnControl Bone Marrow Aspiration System is intended for bone marrow aspiration of the iliac crest of adult and pediatric patients. The ArrowÂż OnControl Bone Marrow Biopsy System is intended for bone marrow core biopsy of the anterior or posterior iliac crest of adult patients |  |  | November 08, 2016 | 2016 | Teleflex Medical | Sterility: Due to a potential incomplete seal on the outer sterile package. | Device Design | Teleflex sent an Urgent Medical Device Recall Notification letter dated November 8, 2016, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately discontinue use and quarantine any affected product, complete the enclosed Recall Acknowledgement Form and fax to 1-855-419-8507, Attn: Customer Service or email to recalls@teleflex.com. Customers with questions were instructed to call their local sales representative or Customer Service at 1-866-246-6990.For questions regarding this recall call 919-433-4940. |  |  |  | Worldwide Distribution - US (nationwide) Internationally to Brazil, Canada, Ecuador, Guatemala, Mexico and Peru. | Open, Classified |  |  |  | Computer | Not_Computer
Z-0866-2017 | 75716 | No | HIF | Laparoscopic insufflator. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Insufflator, laparoscopic |  |  | November 21, 2016 | 2016 | Northgate Technologies, Inc. | Device could cause a patient overpressure situation without any visual or auditory indication or warning | Software design | Northgate Technologies sent an Urgent Medical Device Recall Notification letter dated November 17, 2016, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customers. Customers were instructed to examine their records and ascertain the location of affected devices, cease distribution or use and quarantine those products appropriately. Ship all affected products in their possession to Northgate Technologies, Inc., 1591 Scottsdale Court, Elgin, IL 60123. To return product, complete the enclosed Recall Acknowledgement Form and fax it to 847-608-9405. If product was further distributed, those customers should be notified immediately. Customers with questions should call 224-856-2253. For questions regarding this recall call 800-348-0424. |  | 49 | 49 devices | Worldwide Distribution - US including FL and MA. Internationally to Canada | Open, Classified |  |  | 49 | Computer | Computer
Z-0673-2017 | 75569 | No | CDS | Urea nitrogen test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | MSS CASSETTE GLU/LAC/UREA (BUN) as a part of the |  |  | October 31, 2016 | 2016 | Roche Diagnostics Operations, Inc. | Incorrect Urea results may be generated sporadically on cobas b 221&lt;6&gt;Roche OMNI S6 system and cannot be detected by QC. QC measurement results can be affected as well. This issue is only detectable if comparison result of a second measurement is available. The cause for this issue is an error in the software measurement procedure.In children, decreased urea results can point at inborn disorders or urea cycle. Therefore, false normal urea values could lead to delayed/incorrect diagnosis of such a disorder.Besides the patients most at risk, the risk for the overall patient population with falsely low or high values is considered remote. No severe adverse health effects are expected; but, unnecessary diagnostic measures may be initiated.Falsely normal results may result in further necessary examination not being carried out, with the possibility of a delay in the diagnosis and treatment. A medical risk cannot entirely be excluded. | Software design | Roche sent an Urgent Medical Correction letter dated October 31, 2016, to all affected customers. The recall notifications included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Actions Required" Use the workaround outlined in this Urgent Medical Device Correction (UMDC) for Ureameasurement until the updated software version is available." If your facility has distributed the affected product to another site, please ensure this UMDC isprovided to that site." Complete the attached fax form and fax it to 1-877-766-7471." File this UMDC for future reference.QuestionsPlease contact the Roche Support Network Customer Support Center, 24 hours a day, seven days a week at 1-800-526-2272 if you have questions about the information contained in this UMDC. |  | 97 | 97 | Nationwide Distribution including ALCA CO CT FL IL IN KS MD MI MO NC NE NJ NM NY OH OK PA PR SC TX VA and WI | Open, Classified |  |  | 97 | Computer | Not_Computer
Z-0665-2017 | 75564 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Merge Hemo software. |  |  | April 30, 2014 | 2014 | Merge Healthcare, Inc. | The application may crash during the cath lab procedure. | Software design | Merge sent an Important Alert letter to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers with questions were instructed to contact Merge Support at 866-261-5761 or support@merge.com. For questions regarding this recall call 877-741-5369. |  | 192 | 192 sites potentially have the affected software | Nationwide Distribution | Open, Classified |  |  | 192 | Computer | Not_Computer
Z-0664-2017 | 75334 | No | OUG | Medical device data system. | General Hospital | General Hospital | 510(K) Exempt | Merge FlexConnect software, a component of Merge LIS. |  |  | January 30, 2016 | 2016 | Merge Healthcare, Inc. | Communication protocols interfacing with the affected software version with some select instruments were not properly handled, preventing data captured by these instructions to be processed by the rest of the LIS software. | Software design | The recalling firm issued a letter dated 1/27/2016 via e-mail on 1/30/2016. Nonresponders were issued a second letter dated 5/25/2016 via email on 5/26/2016. |  | 324 | 324 sites potentially have the affected version | Nationwide Distribution | Open, Classified |  |  | 324 | Computer | Computer
Z-0876-2017 | 75383 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | NeuViz 128 Multi-slice CT Scanner System with |  |  | August 31, 2016 | 2016 | Neusoft Medical Systems Co., Ltd. | The Firm discovered during contrast agent tracking scan, when the expected concentration of contrast agent is reached, later planned scan may be interrupted. The system must be powered down and restarted to clear the interruption error before scanning can resume. | Software design | Neusoft Medical Systems Co. Ltd. planned action to bring into compliance.1.Affected Customers will be notified by letter to provide them a description of the affected device, instructions for using the device prior to correction, hazards associated with the defect, and a statement of corrective actions to be taken at no cost to the owner.2.Field Service Engineers will visit owner sites to install a software update which resolves the defect.3.Corrections will be made at no cost to the owner4.Corrections will be completed by May 11, 2017. CDRH approves the CAP subject to the following conditions:.Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR 1003.21 and 1003.22. This office and the Food and Drug Administration (FDA) district office coordinator noted below are to be included in the notification.Further questions please call (281) 453-1205. |  | 23 | 23 units | US Distribution | Open, Classified |  |  | 23 | Computer | Computer
Z-0667-2017 | 75639 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Z6Ms transesophageal transducers used on ACUSON SC2000 ultrasound systems, Z6Ms model 10436113. |  |  | November 07, 2016 | 2016 | Siemens Medical Solutions USA, Inc. | Firm received reports about "acquisition errors" and the deterioration of material covering the articulating section of the Z6Ms transesophageal transducer. | Device Design | Siemens sent a Customer Safety Advisory Notification letter to all affected customers on November 7, 2016 by certified mail advising them of the issue and identified steps that can be taken to avoid potential risk of this issue. This letter has been released first in English while translations are being completed. Once the translations are completecustomers in the countries who require translations will receive the letter.Prior to this mailing, ECO #641310 was released on July 5, 2016 and implemented into forward production on July 28, 2016 (SN 63028009) to correct the susceptibility to Gastro-flex circuit corrosion and electro migration. The ECO changed standard copper traces and traditionally applied soldermask on the circuit to fully gold plated traces with a dry film solder mask.ECO #645821 was released on September 8, 2016 date and implemented into forward production on September 8, 2016 (SN 63530013) to correct susceptibility to articulation sleeve damage. The ECO changed the articulation sleeve material from C-Flex to Viton.For questions regarding this recall call 650-694-5398. |  | 33 | 33 | Worldwide Distribution - US (nationwide) Internationally to AUSTRIA, GERMANY, SOUTH AFRICA, NETHERLANDS, MAURITIUS, MEXICO, and ITALY,JAPAN, SWEDEN | Terminated | May 26, 2017 | 200 | 33 | Computer | Not_Computer
Z-0701-2016 | 72864 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Panorama Patient Monitoring Network. |  |  | December 15, 2015 | 2015 | Mindray DS USA, Inc. dba Mindray North America | An issue with the Panorama Central Station may cause the system to spontaneously restart. This may occur after about 49 days of continuous operation. On restart, monitoring will be lost for about 2.5 minutes. Settings and configurations are maintained. | Software design | The firm, Mindray DS USA, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated 12/15/2015 to its customers via certified mail. The letter described the product, problem, and actions to be taken. The customers were instructed to route the recall letter to all potential users of your facility. Mindray will resolve the issue by upgrading the system software on all affected Panorama systems in your facility. Please contact your Mindray Service Representative at 1-800-288-2121 (Monday-Friday 8:30am-5:30pm) to arrange upgrade.If you have any questions, please contact Director, Quality Operations and Regulatory Affair at 201-995-8407 or email to: d.arpino@mindray.com. |  | 63 | 63 products | Nationwide Distribution. | Open, Classified |  |  | 63 | Computer | Computer
Z-2244-2016 | 74513 | No | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | Angiographic Injector and Syringe, MEDRAD MRXperion MR Injection System, MEDRAD MRXperion Sterile Disposable MRI Kit |  |  | June 09, 2016 | 2016 | Bayer Healthcare | Bayer Healthcare is initiating this recall due to complaints that were received from customer sites describing a 4205 error message when the injector is used with a 3T scanner. | Device Design | Bayer sent an Urgent Medical Device Field Corrective Action letter dated June 8, 2016 to affected customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to take the following steps should a 4205 error message is encountered. For questions contact Bayer Customer Support Team at 1-800-633-7237. |  | 42 | 42 units | Worldwide and the US state of NEW YORK | Open, Classified |  |  | 42 | Computer | Not_Computer
Z-0324-2017 | 75221 | No | KCY | Pneumatic tourniquet. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Reprocessed Pressure Tourniquet Cuff (PTC) |  |  | September 19, 2016 | 2016 | Stryker Sustainability Solutions | Issue with the process for detection of leaking Pressure Tourniquet Cuffs (PTC). Leaking PTC devices may result in excessive blood loss requiring surgical or medical intervention or systemic leakage of local anesthetic during Bier Block requiring medical intervention. | Device Design | Stryker Sustainability Solutions sent an Customer Notification letter September 19, 2016. Customers were instructed to discontinue use of the affected products. The letter stated to follow the link below to complete the Recall Effectiveness Check Form and indicate if any devices containing expiration dates from May 2016-May 2018 remain in inventory. This form must be completed even if no affected product is found.Use the following reference number as you complete the form: [Ship to Account #]If the form indicates that any affected devices remain in inventory, a prepaid shipping label will be issued for the return of the product. Credit for all affected devices returned will be issued to the facility..Adverse reactions or quality problems experienced with the use of this product may be reported to:o Stryker Sustainability Solutions Complaint Hotline: +1(888) 888-3433 X5555o http://www.stryker.com/productexperience/o The FDAs MedWatch Adverse Event Reporting program either online, byregular mail, or by fax. |  | 1695760 | 1,695,760 | U.S. distribution nationwide.No foreign distribution | Open, Classified |  |  | 1695760 | Computer | Not_Computer
Z-1299-2016 | 73548 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | system, x-ray, angiographic, model # 10094135, 10094137, 10094139, 10094141, 10280959 |  |  | February 26, 2016 | 2016 | Siemens Medical Solutions USA, Inc | In Artis zeego systems, angulations in the vicinity of the C-arm collision area can cause the cable inlet to become clamped and can result in mechanical damage. In Artis systems with A100Plus generators, an attempt to resume operation following detection of a fault (such as a short circuit in the X-ray tube) can result in the failure of a module in the high-voltage generator. | Device Design | The firm, Siemens, sent an "Important Customer Safety Notice" dated 2/16/2016 to its customers. The notice informed them of a corrective action that will be performed to prevent a possible hazard to patients, operators, other persons or equipment. The notice described the product, problem and actions to be taken. Siemens plans to contact each customer to arrange a date to add additional protection mechanisms by means of a system software update.If you have any questions, contact Engineer, Regulatory Technical Specialist at 610-448-6461 or email: Meredith.adams@siemens.com. |  | 552 | 552 units | US Nationwide Distribution including states of: AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, GU, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, WY, and PUERTO RICO. | Terminated | June 21, 2017 | 481 | 552 | Computer | Not_Computer
Z-2627-2016 | 74919 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | RadSuite software. The firm name on the label is Merge Healthcare, Hartland, WI. |  |  | January 31, 2016 | 2016 | Merge Healthcare, Inc. | Potential for RadSuite AV viewer to skip image slices when the user presses page up or page down when the viewer displays multiple viewports and the user has mouse focused on any viewport other than the first viewport. | Software design | Merge Healthcare sent an Urgent Medical Device Recall letter dated January 28, 2016, and via email on January 31, 2013, to all affected customers. the letter identified the product the problem and the action needed to be taken by the customer.The notification instructs the customer to pass the information onto all other users, as well as any downstream customers if it was further distributed. In addition, the customer is requested to call Customer Service or send them an email to schedule the upgrade. If the customer chooses not to upgrade, they are requested to complete the attached form or go online and complete it. If you have any additional questions call Merge Customer Service at (877) 741-5369.For further questions, please call (262) 367-0700. |  | 25 | 25 sites potentially have the affected versions | Distribution was made to AL, AZ, CA, IL, MD, MA, MI, MO, NV, NJ, NC, OH, PA, TN, and VA. Government distribution was made to OK. There was no foreign/military distribution. | Open, Classified |  |  | 25 | Computer | Computer
Z-2143-2016 | 74354 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | enGen Laboratory Automation System, Product Code engen, Software Versions enGen Select 3.2 and Above and enGen Custom 2.0 and Above; IVD |  |  | May 16, 2016 | 2016 | Ortho-Clinical Diagnostics | If the user selects the Comment or Interpretation option (in the enGEN .gsb file), and the numerical results fall between Negative and Reactive, the IM incorrectly displays the interpretive text as Borderline and sends the incorrect text to the LIS. The IFU states the interpretation term for these numerical result values is Retest?.Âż The difference in terminology exists for US Markets only. | Software design | Ortho Clinical Diagnostics sent an "IMPORTANT PRODUCT CORRECTION NOTIFICATION" Customer Letter dated May 16, 2016, to all affected customers in the US with enGen Systems via FedEx or US Priority Mail (for Post Office Box addresses) or by e-mail to subscribers of ORTHO PLUS e-Communications to advise them of the issue and the required actions . Required Actions: 1. Use the numerical test results to assess actions to be taken for these assays, per the IFU.2. Post this notification by your enGen System or with your user documentation.3. Complete and return the Confirmation of Receipt form by 25 May 2016. with the Instructions for Use (IFU) directions).If you have any questions regarding this notification, please call the Ortho Technical Solutions Center at any time at 1-800-421-3311. |  | 30 | US: 30 units | Nationwide | Open, Classified |  |  | 30 | Computer | Computer
Z-1279-2016 | 69339 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Foot Switches used with Philips MultiDiagnost Eleva with Flat Detector; Product Codes: 708037, 708038 |  |  | September 24, 2014 | 2014 | Philips Medical Systems, Inc. | Loss of key image functionality due to a bent pedal of the Footswitch. | Device Design | Field Safety Noticed was mailed on September 24, 2014 via Certified Return Receipt. This FCC consists of a Mechanical Hardware update of the Footswitch in order to prevent bent pedals and therefore prevents an intermittent or continuous inability of making live fluoro images or exposures.Customers with questions were instructed to contact their local Philips representative. |  | 168 | 168 | Worldwide Distribution - US (nationwide) and Internationally to:CZECH REPUBLICALBANIAALGERIAARGENTINAAUSTRALIAAUSTRIAAZARBAIJANBAHAMASBANGLADESHBELGIUMBRAZILBULGARIACAYMAN ISLANDSCHILECHINACOLUMBIACOSTA RICACROATIACUBACZECHOSLOVAKIADENMARKDOMINICAN REPUBECUADOREGYPTESTONIAFINLANDFRANCEGABONGERMANYGREECEHONG KONGHUNGARYICELANDINDIAINDONESIAIRANIRAQIRELANDISRAELITALYJAPANJORDANKAZAKHSTANKENYAKOREA SOUTHKUWAITLATVIALEBANONLIBYALITHUANIAMACEDONIAMALAYSIAMAURITIUSMEXICOMOLDAVIAMONGOLIAMOROCCONEPALNETHERLANDSNETHERLANDS ANTNEW ZEALANDNORWAYOMANPAKISTANPALESTINEPANAMAPapua New GuineaPARAGUAYPERUPHILIPPINESPOLANDPORTUGALPUORTO RICOQATARREUNIONROMANIARUSSIAN FEDERATSAUDI ARABIASERBIASINGAPORESLOVENIASOUTH AFRICASPAINSUDANSWEDENSWITZERLANDSYRIATAIWANTANZANIATHAILANDTUNISIATURKEYUKRAINEUNITED ARAB EMIUnited Arab EmiratesUNITED KINGDOMURUGUAYUSAUZBEKISTANVENEZUELAVIETNAMYEMENGUADELOUPESYRIAN ARAB REPUBLIC | Open, Classified |  |  | 4348 | Computer | Not_Computer
Z-1276-2016 | 73411 | No | CHP | Estradiol test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Estradiol III; Elecsys and cobas e analyzers; REF 06656021 190; M 6.5 mL, R1 9 mL, R2 9 mL, |  |  | March 07, 2016 | 2016 | Roche Diagnostics Operations, Inc. | Due to the risk of a recently identified cross reactivity, the Elecsys Estradiol assay should not be used when monitoring patient being treated with fulvestrant. Customers should be informed that fulvestrant will increase the apparent concentration of estradiol in women being treated with this drug. If treatment with fulvestrant has been altered or discontinued as a result of falsely elevated e | Device Design | Roche sent an Urgent Medical Device Correction letter dated March 7, 2016 via UPS Ground (receipt required) to the consignee list. The letter identified the affected product, problem and actions to be taken. Consignees were instructed to complete the attached fax form and fax it to 1-866-891-1984. For questions contact the Roche Support Network Customer Support Center, 24 hours a day, seven days a week at 1-800-428-2336. |  | 23630 | 23,630 | US Nationwide Distribution including Puerto Rico. | Open, Classified |  |  | 125942 | Computer | Not_Computer
Z-1308-2016 | 73436 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | VERO MHI-TM2000 Linear Accelerator System --- Common Name: Medical Linear Accelerator |  |  | March 07, 2016 | 2016 | MITSUBISHI HEAVY INDUSTRIES, LTD., HIROSHIMA MACHINERY WORK | In rare cases the VERO/MHI-TM2000 Operator Console could set an incorrect Gantry angle for the first beam of an IMRT treatment if certain specific conditions are met. If Gantry angle for the first beam of IMRT is incorrectly set and if nobody is aware of it, an erroneous treatment would be given to the patient. | Software design | A Field Safety Notice dated March 2, 2016, was sent to all affected customers via e-mail on March 7, 2016. The Field Safety Notice informed users of the problem and provided a safe workaround until corrected software is available. The correction will be free of charge and a Customer Service Representative will contact all affected customers to schedule its installation. Customers with questions were advised to contact their Customer Service Representative. |  | 29 | 25 Worldwide, US: 4 units | Nationwide Distribution including New York, Texas, Florida, and Ohio | Terminated | April 20, 2017 | 409 | 29 | Computer | Computer
Z-2605-2016 | 74737 | No | FNM | Alternating pressure air flotation mattress. | General Hospital | General Hospital | 510(K) Exempt | Panacea Air Pro Mattress Control Unit, Model: Platinum 5000, US Model Number 5000. |  |  | July 11, 2016 | 2016 | American National Mfg Inc | The Panacea Air Pro Mattress Control Unit is being recalled due to a firmware problem, which may result in the device not having the firmness range required. | Software Manufacturing/Software Deployment | An Urgent Medical Device Recall letter, dated 7/11/16, was sent to customers to inform them of a recall involving the Panacea Air Pro Mattress Control Unit also labeled as Platinum 5000, US Model Number 5000, Serial numbers 9500051 through 9500300. The letter informs the customers that the recall has been identified due to firmware problem. The letter informs the customers of the problem identified and the actions to be taken. Customers were asked to immediately examine their inventory and quarantine the affected device. A shipping label would be provided for return and update of the affected devices. All customers that the device was further distributed to would need to be notified of the recall. |  | 31 | 31 units | Distributed to the US -- WI, OH, IL, TX, AZ, MD, IA, PA, CO, WA, ID, MN, LA, WY, and MS. | Open, Classified |  |  | 31 | Computer | Not_Computer
Z-1303-2016 | 73477 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Siemens branded SOMATOM Emotion and SOMATOM Sensation families of CT systems intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data from either the same axial plane taken at different angles or spiral planes taken at different angles. |  |  | February 15, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Siemens is initiating this recall due to a potential malfunction and hence potential hazard to patients and personnel when using the SOMATOM Emotion or SOMATOM Sensation systems in conjunction with the use of the cardio licenses "Cardio Pro" and "Cardio Retro". | Device Design | A Safety Advisory Notice was sent to customers informing them of a potential malfunction and hence potential hazard to patients and personnel when using the SOMATOM Emotion or SOMATOM Sensation systems. Siemens informed them that they will deliver an Update Instruction for their Service Engineers to modify the installed Firmware. This will correct the rare synchronization errors of Scanner Firmware and Scanner Software to start the application of the x-ray in compliance with the optimum triggering point for best image quality. |  | 124 | 124 units | US (nationwide) including Guam and the states of AL, AR, AZ, CA, CO, FL, GA, HI, IA, ID, IL, KY, MA, MD, ME, MI, MN, MO, NC, NJ, NM, NY, OH, PA, SC, TN, TX, VA, and WA. | Open, Classified |  |  | 124 | Computer | Not_Computer
Z-2595-2016 | 73607 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Adult ThermoGardÂż Dual Dispersive Electrodes (for patients &gt;15 kg.), 10' (3.05m) Cable, Catalog Number 51-7310. |  |  | May 18, 2016 | 2016 | Conmed Corporation | A version of the dual dispersive electrodes may not be compatible with some electrosurgical generators, placing patients at risk for undetected pad lift and a potential burn. | Device Design | On May 17, 2016 CONMED Corporation distributed Urgent Device Recall notification letters and Business Reply Forms to their customers via courier service. CONMED Corporation has decided to recall these ThermoGardÂż Dual Dispersive Electrodes for all customers who do not use an electrosurgical generator that employs the Contact Quality Monitoring systems listed. If your electrosurgical generator is not listed, your generator should not be used with the lot codes of the ThermoGard Dual Dispersive Electrodes identified. Customers are advised to inspect inventory for any of the recalled devices. If you use the ThermoGardÂż Dual Dispersive Electrodes in conjunction with an electrosurgical generator listed, no further action is required. However, if your electrosurgical generator and CQM system is not listed above, do not use the affected lot codes of ThermoGardÂż Dual Dispersive Electrodes. Please note that ThermoGardÂż Dual Dispersive Electrodes produced after July 5, 2015, are not subject to this recall. Customers should contact all of those departments or organizations within their facility and any other facilities that you may have supplied or given these affected products to. It is imperative that all end users of these devices receive this notice and respond immediately. If you have questions, please contact Patricia Cotter, CONMED Recall Coordinator 315-624-3237 (8AM-5PM EST, M-F) or fax to 315-624-3225 or email thermog@CONMED.com. |  | 307684 | 307,684 units | Worldwide Distribution -- USA, including AK, AL, AR, AZ, CA, CO, CT, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, WY, &amp; Puerto Rico; and, the countries of Australia, Canada, Belgium, Brunei Darussalam, China, France, Guatemala, Hong Kong, Israel, Italy, Japan, South Korea, Malaysia, Mexico, New Zealand, Norway, Portugal, Saudi Arabia, Spain, Sweden, and United Arab Emirates. | Open, Classified |  |  | 447697 | Computer | Not_Computer
Z-2604-2016 | 74884 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | CARESCAPE VC150 Vital Signs Monitor; Intended to monitor a single patient's vital signs at the site of care. |  |  | July 27, 2016 | 2016 | INNOKAS MEDICAL OY | A software error on released software versions 1.6.12, 1.6.12F and 1.6.16 may give wrong time data to measurements. | Software design | Innokas Medical has notified the one US distributor on 7/27/16 via letter of the deficiency and correction of the device. |  | 1715 | 1715 units | AR, AZ, FL, IN, LA, MA, MI, NC, NJ, NM, NY, OH, PA, TN, WI | Terminated | April 19, 2017 | 266 | 1715 | Computer | Computer
Z-1278-2016 | 73605 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Cios Alpha, mobile X-ray system |  |  | February 19, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Software issues on Cios Alpha mobile C-Arm system | Software in the Use Environment | Siemens sent an Customer Safety Advisory Notice on February 26, 2016, to inform them of some issues that may arise on the Cios Alpha unit. The letter identified the product the problem and the action needed to be taken by the customer. Users were instructed to change the date format to DDMMYYYY to prevent issues, not to release radiation as long as the DVD icon is blinking and to disconnect and re-connect the trolley if the system is needed immediately. In the meantime they are preparing a modification of Cios Alpha that will resolve the above-mentioned issues. As soon as the modification is available, a service engineer will get in touch with each customer and implement the modification to the Cios Alpha system free of charge. If the device was sold, customers were instructed to forward the recall notification to the new owner.For further questions please call (610) 448-6478. |  | 89 | 89 systems | Nationwide Distribution to the states of : AL, AR, CA, CO, CT, DE, FL, GA, IA, IL, KY, MD, MI, MN, MO, NC, NE, NJ, NV, NY, OH, OR, PA, RI, SC, TN, TX, UT, WA and WI | Terminated | March 20, 2017 | 395 | 89 | Computer | Computer
Z-2587-2016 | 74595 | No | HJO | AC-powered slitlamp biomicroscope. | Ophthalmic | Ophthalmic | 510(k) | OPTICAL BIOMETER AL-Scan: Software Version: 1.09 and earlier. |  |  | July 01, 2016 | 2016 | Nidek Inc | Nidek Inc. received information from our manufacturer NIDEK CO. Japan that the shown axiallengths were much shorter at AL-Scan measurement. | Software in the Use Environment | Dear Doctor/Recall letters were sent to affected customers on July 7, 2016. The letters describe the issue and provide instruction should the issue occur. The letters inform users that Nidek will reach out in the next few weeks to schedule installation of new software that will solve the problem.Questions or concerns may be directed to:Customer Service 1.800.722.0219; Monday- Friday 8 am to 5pm PSTE-mail: AL-Scan_support@noritsuservice.com. |  | 17 | 17 | US Nationwide Distribution in the states of: NY, PA. GA. AZ, VA, DE, ME,. MO, TN, and NM. | Open, Classified |  |  | 17 | Computer | Not_Computer
Z-1309-2016 | 73550 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MRIdian ViewRay Radiation Therapy System, ViewRay Treatment Planning and Delivery System (also known as the MRIdianÂż System) is indicated for stereotactic radiosurgery and precision radiotherapy for lesions, tumors, and conditions anywhere in the body when radiation treatment is indicated |  |  | March 07, 2016 | 2016 | Viewray Incorporated | When editing the isocenter or the couch position of the plan while in the treatment workflow (in the Points screen) and re-optimizing, the software will not prompt the user to shift the couch to the new isocenter. Asa result there is the potential to deliver dose to the initial isocenter rather than the new location. | Software design | ViewRay sent an Urgent Device Correction letter dated March 7, 2016. The letter identified the product, the problem, and the action to be taken by the customer. Customers with questions were instructed to contact ViewRay Customer Support at support@viewray.com or call the ViewRay Customer Support center at 855-286-8875. |  | 5 | 5 units | Worldwide Distribution - US including CA, FL, MO &amp; WI., and International distribution to the Republic of Korea and The Netherlands. | Open, Classified |  |  | 5 | Computer | Computer
Z-1945-2016 | 70054 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | The XN-Series modules (XN-10, XN-20) are quantitative multi-parameter automated hematology analyzers intended for in vitro diagnostic use in screening patient populations found in clinical laboratories. |  |  | December 18, 2014 | 2014 | Sysmex Corporation | There is a potential for elevated hemoglobin (Hgb) and impedance platelet count (PLT-I) values. | Device Design | Sysmex Corporation sent a Product Notification Field Correction letter were to all affected customers of record. The letters included instructions for customers to: 1) Review results with an unexpected increase in Hgb, MCH, MCHC, and/or PLT-I. The increase may result in a low or normal value being artificially elevated to either a normal result or a result above the reference ranges, or a high value being artificially increased. There is negligble impact to the RBC when the issue occurs; 2) The relationship with an elevated Hgb would cause the indices (MCH and MCHC) values to appear to be abnormally elevated and may result in the results being flagged for verification or indicated to be outside of reference ranges; 3) Review any "Turbidity/Hgb Interf?" flag which will automatically trigger at an MCHC of 37.5 g/dL; 4) If Hgb is affected, the PLT-I may be also. Review any PLT-I result when an unexpected increase in MCH/MCHC or "Turbidity/Hgb Interf?" flag occurs; 5) Gross interference is likely to trigger a flag, however there is a potential any interference could increase the Hgb or the PLT-I count from a lower level to an "normal" level. Follow laboratory protocol for confirming unexpected results. Customers were further advised that they will be contacted by a Sysmex Field Representative within 180 days to schedule an appointment to replace the rinse cup on the analyzer(s) and that version 00-16 software will be installed to change the rinsing sequence for the aspiration probe. Customers with questions are requested to e-mail the Sysmex Technical Assistance Center. Urgent questions can be answered by calling the Technical Assistance Center at 888-879-7639 (24 hours a day / 7 days a week). |  | 1137 | 1,137 units (US) 54 units (Canada) 9 inventory | Worldwide Distribution - US to AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WY + District of Columbia + Guam; Internationally to Australia, Bahrain, Bangladesh, Cambodia, Canada, China, Germany, Hong Kong, India, Indonesia, Malaysia, Maldives, Mongolia, Mynamar, Nepal, New Zealand, Pakistan, Philippines, Singapore, South Korea, Sri Lanka, Taiwan, Thailand, and Vietnam.THE CORRECTION IS LIMITED TO XN-10/XN-20 UNITS DISTRIBUTED BY SYSMEX AMERICA, INC., INSTALLED AT END-USER SITES IN US, PUERTO RICO, AND CANADA. | Terminated | July 25, 2016 | 585 | 1137 | Computer | Not_Computer
Z-2553-2016 | 74403 | No | KRO | Knee joint femorotibial metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | METS SMILES Total Knee Replacement |  |  | May 31, 2016 | 2016 | Stanmore Implants Worldwide Ltd. | The manufacturer has identified that the stated Femoral Plateau Plates may not fit into the intended Femoral Knee mating component and thus sit flush upon the Femoral Knee component prior to implantation. | Device Design | The recalling firm sent recall notification letters to their consignees on 05/31/2016. Follow up communciation was sent out on 7/26/2016 after a technical and medical assessment was completed. |  | 3 | 3 units | Directly distributed to TX, further distributed to AL and MS | Open, Classified |  |  | 3 | Computer | Not_Computer
Z-1282-2016 | 73606 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Artis One, Interventional, Fluoroscopic, x-ray system |  |  | March 11, 2016 | 2016 | Siemens Medical Solutions USA, Inc | After importing segmentation results of the left atrium created on the Artis One system, the possibility exists for the results to be mirrored to an electro-anatomical 3D mapping system, CARTO from Biosense Webster Inc. | Software in the Use Environment | Siemens mailed an Important Customer Safety Notice to customers to inform them of the issue and the corrective actions they plan to take to address the problem. The letter stated that Siemens service organization will contact customers to arrange a date for the installation of the software update. Customers were instructed to call Siemens service organization for an appointment at 1-800-888-7436. |  | 1 | 1 angiography system | Nationwide Distribution to Michigan only | Open, Classified |  |  | 1 | Computer | Computer
Z-1558-2016 | 73838 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | SOMATOM Force, Computed Tomography x-ray system |  |  | April 01, 2016 | 2016 | Siemens Medical Solutions USA, Inc | The Neonate Head protocol with the pediatric kernel Hp38 could result in artefacts and possibly lead to a misdiagnosis (i.e. either non-existing blood or liquid is mimicked in the images or actual existing blood or liquid is not depicted as expected). There is also a risk of a potential misdiagnosis in using this protocol for surgery planning. Investigations revealed an incorrect parameterization of the reconstruction algorithm applied for the pediatric head kernel as cause of the problem. | Software design | Siemens distributed a Safety Advisory Notice dated April 1, 2016, to affected customers The letter identified the affected product, problem, how to avoid potential risks, how the issue will be resolved. and actions to be taken. Customers were asked to observe the safety notice and comply with the corresponding measures until further notice. |  | 37 | 37 CTsystems | US Nationwide Distribution : NC, NH, IL, MA, WA, PA, MD, MN, TX, NY, SC, OH, CA, KY, MT, IA | Terminated | April 17, 2017 | 381 | 37 | Computer | Not_Computer
Z-2579-2016 | 64785 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Centricity PACS Workstation, versions 3.1.1.x through 3.2.1.x Picture Archiving and Communication System |  |  | November 15, 2012 | 2012 | GE Healthcare It | While merging exams in a test instance, merges carry forward in the production system, when the middle tier is configured to the wrong IP address of the Centricity Exam Manager. When a current and historical exam are opened at the same time, the system asynchronously tries to access the operation specifying how each image should be grouped, causing random image-series grouping errors. | Software design | GE sent an Urgent Medical Device Correction letter dated November 15, 2012. The software has been modified to correct these defects. The modified software was released and installed in user sites under GE Healthcare FMI #85196. |  | 498 | 498 units | Worldwide distribution including US nationwide, Puerto Rico, Australia, Austria, Belgium, Bermuda, CANADA, Denmark, Egypt, ENGLAND, France, Germany, India, Ireland, Hong Kong, Korea, Kuwait, Italy, Malta, Netherlands, Portugal, Qatar, Saudi Arabia, SCOTLAND, South Africa, South Korea, Spain, Sweden, Switzerland, Thailand, Turkey, and UK. | Open, Classified |  |  | 498 | Computer | Computer
Z-1312-2016 | 73600 | No | FZX | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Synthes Interlocking Bolt, 151 mm (General and Plastics Surgery Devices) Guide, Surgical Instrument |  |  | February 16, 2016 | 2016 | Synthes (USA) Products LLC | DePuy Synthes is initiating a voluntary medical device recall of certain lots of the Interlocking Bolt which are a part of the Synthes 4.5 mm LCP Proximal Tibia Plate System. The Synthes 4.5 mm LCP Proximal Tibia Plate are indicated for the treatment of nonunions, malunions and fractures of the proximal tibia including: simple comminuted, lateral wedge, depression, medial wedge, bicondylar, combi | Device Design | DePuy Synthes sent an Urgent Notice to inform customers that they were initiating a voluntary recall of lot numbers: 1776197, 5706588, 1782340 and 1868678 of the Interlocking Bolt which are part of the Synthes 4.5mm LCP Proximal Tibia Plate System. Customers were asked to review their inventory for the affected lots, call DePuy Synthes at 1-800-479-6329 to obtain a Return Authorization (RA) Number, complete the Verifcation Section (page 3 of the letter) identifying that the affected product has been located and return it to Credit/Returns, DePuy Synthes, 1101 Synthes Avenue, Monument, CO 80132. A copy of the Verification Section of the letter was to be faxed to 866-737-4701 or emailed to: Synthes2856@stericycle.com. |  | 42 | 42 plate systems | Nationwide Distribution including AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NJ, NV, NY, OH, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, and WV. | Open, Classified |  |  | 42 | Computer | Not_Computer
Z-2551-2016 | 74846 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge Unity Z3D software. |  |  | July 23, 2015 | 2015 | Merge Healthcare, Inc. | The software is unable to accurately determine the calcium score of scans with a slice thickness not equal to 3 mm. | Software design | The firm, Merge, sent a "canned" email on 7/23/2015 to its customers which also contained a personal message. The recalling firm issued a second letter entitled "URGENT: MEDICAL DEVICE RECALL" dated 1/27/2016 via e-mail on 1/30/2016. The letters described the product, problem and actions to be taken. The letter informed customers that an upgrade is available. The customers were instructed to call Merge Customer Service at (877) 741-5369 or send an email to support@merge.com to schedule the upgrade-Reference this recall #2016-14. If you decide to decline this maintenance release, complete the attached form and return to Merge Healthcare via recall@merge.com. Alternatively, complete the online form: http://www.merge.com/Landing-Pages/Customer-Upgrade-Decline-Form.aspx no later than February 22, 2016. Ensure that all users of the product are provided with this notification. If you have further distributed this, identify your customers and notify them at once of this product recall and provide a copy of the recall notification letter.If you have any additional questions, call Merge Customer Service at (877) 741-5369 or support@merge.com. |  | 9 | 9 sites potentially have the affected version | US Distribution to states of: CA, PA, TX, and MT. | Open, Classified |  |  | 9 | Computer | Not_Computer
Z-2546-2016 | 74796 | Yes | NPT |  |  | Cardiovascular | PMA | Medtronic EnVeo R Delivery Catheter System (DCS), Models: ENVEOR-L, ENVEOR-L-C, and ENVEOR-US. |  |  | August 03, 2016 | 2016 | Medtronic Cardiovascular Surgery-the Heart Valve Division | Potential for vascular trauma with the use of the Medtronic EnVeo R Delivery Catheter System (DCS). | Device Design | Firm initiated a recall on 8/3/16 by sending an Urgent Medical Device Correction letter to inform them of the potential for vascular trauma with the use of the Medtronic EnVeo R Delivery Catheter System (DCS). The letter informs the customers of the recommendations to reduce the rate of vascular trauma events and the procedural uses. The letter informs the customers that the Evolut R System Instructions for Use (IFU) will also be updated appropriately consistent with the Dear Doctor communication. Patients with questions can contact Medtronic Cardiovascular Patient Services at 877-526-7890 (Monday-Friday, 7:30am-5pm Central Time). Customers with any questions are instructed to contact their Medtronic Field Representative or Lifeline Technical Services at 1-877-526-7890.On 8/19/16, firm sent a copy of a customer notification letter that will be sent to their customers to inform them that of the communication letter that was released in August 2016 by Medtronic regarding potential for aortic vascular trauma with the use of the Medtronic EnVeo R Delivery Catheter System (DCS). The letter informs the customers that FDA classify the recall as a Class II recall and the recall does not require customers to return the product. The recall letter informs the customers that Medtronic has taken further action by updating all customer training materials and the Instructions for Use (IFU). |  | 7319 | 7,319 units | Worldwide Distribution -- US, Australia, Austria, Belgium, Brazil, Colombia, Costa Rica, Czech Republic, Denmark, Dominican Republic, Egypt, Estonia, Finland, France, Germany, Greece, Hong Kong, Hungary, Iceland, India, Iran, Ireland, Israel, Italy, Kazakhstan, Korea, Kuwait, Latvia, Lebanon, Lithuania, Luxembourg, Macedonia, Malaysia, Morocco, Netherlands, New Zealand, Norway, Oman, Panama, Peru, Philippines, Poland, Portugal, Puerto Rico, Qatar, Romania, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Thailand, Trinidad and Tobago, Tunisia, Turkey, United Arab Emirates, United Kingdom, and Venezuela. | Terminated | March 31, 2017 | 240 | 7319 | Computer | Not_Computer
Z-2542-2016 | 74876 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Xario 100 Diagnostic Ultrasound System, TUS-X100; Xario 200 Diagnostic Ultrasound System, TUS-X200. |  |  | August 04, 2016 | 2016 | Toshiba American Medical Systems Inc | Toshiba American Medical Systems (TAMS) is recalling the Xario Diagnostics Ultrasound System because it may become hot because of a software error. | Software design | An Urgent Medical Device Correction letter was sent to all customers to bring to attention a potential problem of transducer heating with your Xario Ultrasound System. The letter informs the customers that the surface of the acoustic lens of a transducer may become hot because of a software error. The error occurs when a transducer is disconnected, reconnected, or another transducer is selected. The letter informs the customers of error message and the corrective action. Customers are instructed to complete the attached form and fax it to (877) 349-3054 or TAMS-RegulatoryAffairs@toshiba.com. Customer with any questions are instructed to call (714) 730-5000 or their Toshiba service representative at (800) 521-1968. |  | 93 | 93 units | US Distribution | Open, Classified |  |  | 93 | Computer | Computer
Z-1315-2016 | 73381 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | TeleSentry Battery, 3.7 Li-Ion Polymer Battery, CAUTION Do not expose to high temperature. Do not disassemble. Dispose of properly. Use specified charger only. 60pC/140pF The ScottCare Corporation. |  |  | February 25, 2016 | 2016 | Scottcare Corporation | The firm received a complaint of the battery exploding while recharging. | Device Design | On 2/26/2016 the firm sent "Urgent: TeleSentry Battery Recall" notification letters to their customers. The letter identified the affected device and the reason for the recall. Customers were asked to immediately examine their inventory and quarantine the affected device. If the product had been further distributed, then a copy of the letter should be provided to those customers. Customers are to follow the steps provided for returning the affected device, including contacting ScottCare Service at 1-800-243-9412. In addition, the attached response form should be completed and returned. Questions should be directed to ScottCare Customer Service at 1-800-243-9412. |  | 9067 | 9,067 batteries | Worldwide Distribution -- USA, including the states of CA, CT, DE, FL, IL, KS, KY, MA, MD, MI, MN, MO, NC, NH, NJ, NM, NV, NY, OH, PA, SC, TN, TX, &amp; WV; and, the country of Canada. | Open, Classified |  |  | 9067 | Computer | Not_Computer
Z-2144-2016 | 74522 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge OfficePACS software. The firm name on the label is Merge Healthcare, Hartland, WI. |  |  | January 30, 2016 | 2016 | Merge Healthcare, Inc. | Potential data loss occurs as a result of the software archiving not working properly. | Software design | Merge Healthcare sent an Urgent Medical Device Recall letter dated January 26, 2016, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customers. Customers were instructed to provide a copy of the notification to all users. For additional questions customers were advised to call Merge Customer Service at (87) 741-5369. For questions regarding this recall call 262-367-0700. |  | 216 | 216 sites have the affected software | The software was distributed nationwide to medical facilities. There was no government/military/foreign distribution. | Open, Classified |  |  | 216 | Computer | Computer
Z-1318-2016 | 73489 | No | KRR | Knee joint patellofemoral polymer/metal semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | All instrumentation associated with the Prelude PF Resurfacing Knee System. |  |  | March 09, 2016 | 2016 | Biomet, Inc. | Inadequate design control. | Device Design | Zimmer Biomet sent an "Urgent Medical Device Recall Notice" dated March 9, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The following actions are REQUIRED:1. Immediately locate and remove the identified device(s) listed below from circulation.2. Carefully follow the instructions on the enclosed Response Form .3.Email a copy of the response form to CPW ARFieldAction@zimmerbiomet.com prior to return of product.4. Use priority carrier for your shipment.Please confirm receipt of this notice by sending back the response form within three (3) business days. Thank you in advance for your assistance and prompt attention. On behalf of Zimmer Biomet, I apologize for any inconvenience this may cause. Questions related to this notice should be directed to (574) 372-1570, Monday through Friday, 8 a.m. to 5 p.m. |  | 576 | 576 | US Distribution to the states of : CA, CO, GA, IN, KS, MI, MO, NC, NJ, OH, TN and TX. | Terminated | May 23, 2017 | 440 | 576 | Computer | Not_Computer
Z-2624-2016 | 74840 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | SOMATOM Definition AS with software version VA48A-SP2; Model # 10430603, computed tomography x-ray system |  |  | July 13, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Due to an internal communication error between the firmware and the software of the components, the planned CT scan executes properly, but the injector is not started. Therefore, the contrast agent is not injected and the desired examination result is not achieved. This error only affects the automatic mode or coupled mode and does not affect the manual control of the injector. | Software in the Use Environment | Siemens mailed to customers an Important Customer Safety Notice to inform about a possible malfunction when using the affected systems and what actions to take to eliminate the problem. Also, Siemens informed their customers that they were currently developing a new firmware to eliminate the internal communication error, which should be released in August 2016. Customers were asked to contact their service organization at 800-888-7436, if they had any questions. |  | 1 | 1 unit | Distributed to: MI, NY, CA, KY, ND, NE | Open, Classified |  |  | 7 | Computer | Computer
Z-2628-2016 | 74808 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | OEC UroView 2800, fluoroscopic x-ray system |  |  | July 22, 2016 | 2016 | GE OEC Medical Systems, Inc | A firmware issue in the supplied LCD workstation monitor(s). Under some circumstances, the monitor(s) may appear dark with no video displayed. | Software design (manufacturing process) | On July 22, 2016 GE Healthcare Surgery began distributing an Urgent Medical Device Correction letter to customers via mail to inform them of the issue and provide instructions for immediate mitigation of the issue. The letter was sent to 3 titles at each facility affected: 1.) Facility Administrator, 2.) Director/Manager of Radiology and Radiologists, 3.) Radiology Department. Consignees instructured that local GE Healthcare Field Engineer will contact them to coordinate replacement of the affected monitor(s), and will ensure that system meets specifications. |  | 47 | 47 total | Distribution US nationwide and Korea. | Open, Classified |  |  | 47 | Computer | Computer
Z-1560-2016 | 73857 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Aquilion Lightning CT System iStation cover, SID: 70000246, Item Code: TSX-0354A/2D.03 |  |  | April 12, 2016 | 2016 | Toshiba American Medical Systems Inc | Toshiba American Medical Systems, Inc. is recalling the TSX-035A because of a potential problem with the Aquilion Lightning CT System iStation cover falling off of the gantry. | Device Design | An Urgent Medical Device Correction letter dated 4/12/16 was sent to a customer to bring to the customer's attention with a potential problem with the Aquilion Lightning CT System iStation cover. The letter informs the customer that the iStation cover may come off due to deficient dimensional geomtry. The letter informs the customer of the problem identified and corrective action to be taken. The customer is instructed to complete and return the attached form and fax it to (877) 349-3054. If the customer has any questions regarding the recall letter, then the customer is instructed to contact Paul Biggins at (800) 421-1968 or contact their local Toshiba Representative at (800) 521-1968. |  | 1 | 1 unit | US in the state Nevada | Open, Classified |  |  | 1 | Computer | Computer
Z-2639-2016 | 74013 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Ziehm Vision RFD 3D Mobile C-arm for Mobile Fluoroscopic Imaging |  |  | April 15, 2016 | 2016 | Ziehm Imaging Inc | Malfunction of the radiation-warning timer and the 3D-specific interval warning. No 5-minute timer dialogs or beeper will occur until the unit is rebooted. | Device Design | Ziehm Imaging, Inc planned action to bring these defects into compliance: 1. You will contact customers and initiate a software update to correct the defects.2. You will install the software and perform testing to ensure the software update was effective.3. The customer notification letter which includes a statement that you will without charge, remedy the defect or bring the product into compliance.For further questions, please call (866) 949-4346. |  | 31 | 31 devices | US Distribution including Puerto Rico and to the states of :FL, GA, VA,MA. ME, MI, NY, NJ, PA, RI, IL, TX, AZ, CA, NV, UT, and WA. | Open, Classified |  |  | 62 | Computer | Computer
Z-2715-2016 | 74853 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge RadSuite software. Radiological image processing system. |  |  | April 09, 2015 | 2015 | Merge Healthcare, Inc. | When RadSuite is used with IPID (Issuer of Patient ID) as a part of the "Patient Identifier," it is possible in some circumstances that the demographics of one patient will be applied to a study or studies for another patient. | Software design | The recalling firm issued letters dated 4/3/2015 via email on 4/9/2015 notifying the customer of the issues. |  | 10 | 10 sites have the affected software that is configured with the specific conditions listed in Code Information | Distribution was made to medical facilities in AL, MI, MO, PA, TN, and TX. | Open, Classified |  |  | 10 | Computer | Computer
Z-2732-2016 | 74912 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | UniCel DxH 800 and DxH 600 Coulter Cellular Analysis System. |  |  | August 08, 2016 | 2016 | Beckman Coulter Inc. | There is a possibility of specimen misidentification for the stand-alone DxH 800 and DxH 600 where the Primary Identifier is configured to Tube Position ID. | Software design | An Urgent Medical Device Recall letter dated 8/8/16 was sent to all customers who purchased the Unicel DxH 800 and DxH 600 Coulter Cellular Analysis System to inform them that there is a possibility of specimen misidentification for the stand-alone DxH 800 and DxH 600 where the Primary Identifier is configured to Tube Position ID. Customers are instructed to complete and return the enclosed response form within 10 days. Customers with any questions regarding the notice are instructed to contact Customer Support Center: http://www.beckmancoulter.com/customersupport/support; 800-526-7694 in the US and Canada; outside the US and Canada, contact local Beckman Coulter representative. |  | 5282 | 5,282 units total (2,877 units in US) | Affected geographies: Andorra, Australia, Belarus, Belgium, Bosnia and Herzegovina, Brunei Darussalam, Canada, China, Columbia, Croatia, Czech Republic, Egypt, Ecuador, France, Germany, Gibraltar, Greece, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Korea, Kuwait, Lebanon, Macao, Malaysia, Mayotte, Mexico, Morocco, Myanmar, New Zealand, Norway, Oman, Pakistan, Panama, Paraguay, Philippines, Poland, Portugal, Puerto Rico, Qatar, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, Spain, Sweden, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Turkey, United Arab Republic, United Kingdom, Uruguay and Vietnam. | Terminated | April 29, 2017 | 264 | 5282 | Computer | Not_Computer
Z-2731-2016 | 75081 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE Healthcare SIGNA Creator, SIGNA Explorer. Diagnostic Imaging Device. |  |  | August 31, 2016 | 2016 | GE Healthcare, LLC | GE Healthcare is conducting a recall dealing with an incorrect RF Power Monitor Function involving SIGNA CREATOR and SIGNA EXPLORER which could result in higher than expected thermal dose to the patient and higher than expected localized heating. There have been no failures of this type reported, and no injuries reported as a result of this issue. | Software design | Consignees were sent on 8/31/2016 a GE Healthcare "Urgent Medical Device Correction" letter GEHC Ref# 67921 dated August 31, 2016. The letter was addressed to Hospital Administrators / Risk Managers, Radiology Department Managers, and Radiologists.The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. For questions contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 383 | 383 (US = 24;OUS = 359) | US: AZ, CA, CT, FL, IL, MA, MS, MO, NC, PR, TX, WI.OUS: Algeria, Australia, Bangladesh, Belgium, Brazil, Canada, Colombia, Dominican republic, Egypt, France, Germany, Ghana, Guatemala, HONG KONG, Hungary, India, Indonesia, Iran, Italy, Jamaica, Japan, Korea, Kuwait, LEBANON, Lithuania, Maldives, Martinique, Mexico, MOROCCO, Netherlands, New Zealand, OMAN, Pakistan, Philippines, Poland, Portugal, Romania, SAUDI ARABIA, Singapore, South Africa, Spain,Sweden, Tanzania, Thailand, Turkey, United Arab Emirates, United Kingdom, Vietnam. | Open, Classified |  |  | 383 | Computer | Computer
Z-2709-2016 | 75027 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge Cardio software. The firm name on the label is Merge Healthcare, Hartland, WI. Image processing system. |  |  | February 09, 2016 | 2016 | Merge Healthcare, Inc. | A report can be confirmed with the incorrect patient demographics resulting in it being saved under the incorrect patient record in the Electronic Medical Record (EMR). | Software design | MERGE sent an Urgent Medical Device Recall letter dated February 9, 2016, to all affected customers via email or via certified mail if they did not have an e-mail address for the customer.The notification instructs the customer to pass the information onto all other users, as well as any downstream customers if it was further distributed. In addition, the customer is requested to call Customer Service or send them an email to schedule the upgrade. If the customer chooses not to upgrade, they are requested to complete the attached form or go online and complete it.For further questions, please call (800) 724-5970 ext. 8027. |  | 116 | 116 customers potentially have the affected versions | Nationwide Distribution | Open, Classified |  |  | 116 | Computer | Computer
Z-1233-2016 | 73292 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | cobas p 512 pre-analytical system |  |  | February 23, 2016 | 2016 | Roche Diagnostics Operations, Inc. | Due to a false triggering or detection of the lifting gripper READY signal, sample tubes are not correctly placed back in the Rack Tube Transport (RTT) after the decapping process. Therefore, open sample tubes can be dropped in the cobas p 512, spilling the sample material. | Software design | Roche Diagnostics sent an "Urgent Medical Device Correction" letter via UPS Ground (receipt required) on February 23, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.To prevent this malfunction, a software patch has been developed and validated. The installation of this software patch is mandatory and will be performed by a Roche Field Engineering Specialist (FES). Your Roche FES will contact you to schedule the installation of the software patch. Until the software patch is implemented, Roche recommends that instrument operators monitor their systems for the occurrence of Error 140 Gripper has lost tube. If this error occurs, all components in contact with a spill must have extra cleaning. Refer to the cobas p 512 pre-analytical instrument Operators Manual Version 1.5.Please contact the Roche Support Network Customer Support Center, 24 hours a day, seven days a week at 1-800-428-2336 for further questions. |  | 35 | 35 Units | US Distribution including Puerto Rico and to the states of :TX, OH, TN, AZ, WA, MI and GA | Open, Classified |  |  | 35 | Computer | Not_Computer
Z-2711-2016 | 74900 | No | NQE | AC-powered slitlamp biomicroscope. | Ophthalmic | Ophthalmic | 510(k) | SPECULAR MICROSCOPE CEM-530; Software version 1.08 and 1.09. |  |  | August 08, 2016 | 2016 | Nidek Inc | Software version 1.08 and 1.09 for the Specular Microscope CEM 530 included a change of analysis results feature that was not reviewed and approved by the FDA. | Software change control | The firm, Nidek, sent a Dear Doctor letters entitled "RECALL NOTIFICATION" dated August 15, 2016 to affected sites by Federal Express. The letter informs users of the issue and stated that Nidek field service representatives will visit to install updated software.If you have any questions, please call Customer Service at 1-800-722-0219; Monday-Friday 8am to 5pm PST or email: cem530update@noritsuservice.com. |  | 29 | 29 affected devices | US Distribution to states of: TX, KS, FL, NV, CA, PA, KY, OR, TN, AL, VA, MO, OK, CT, NY. LA, NM, and IN. | Terminated | January 09, 2017 | 154 | 29 | Computer | Not_Computer
Z-1232-2016 | 73481 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo X Workplace is a medical workstation for real-time viewing, image manipulation, 3D-visualization, communication, and storage of medical images and data on exchange media. It is used for diagnostic image viewing and post processing during interventional procedures. |  |  | February 24, 2016 | 2016 | Siemens Medical Solutions USA, Inc | After importing, the segmentation results appear mirrored at the CARTO system and can't be used for the ablation procedure. | Software design | Siemens sent an Important Customer Safety Notice dated February 24, 2016, to all affected customers notifying them that the issue will be remedied with a software update via AX004/16/S. Following the installation of this update, the segmentation results can be used at the CARTO system as described in the operator manual. Customers with questions were instructed to call 800-888-7436. For questions regarding this recall call 610-448-6461. |  | 6 | 6 | Nationwide Distribution to IL, NY, MT, and MN. | Open, Classified |  |  | 6 | Computer | Not_Computer
Z-1244-2016 | 72991 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MHI-TM2000 Linear Accelerator System (Software Version 3.0.0 and after) |  |  | December 15, 2015 | 2015 | MITSUBISHI HEAVY INDUSTRIES, LTD., HIROSHIMA MACHINERY WORK | Due to Operator Console software anomaly, a change in treatment completion status of the last patient of the day may be altered from "Completed (or Discontinued)" to "Untreated" under specific conditions. | Software design | Due to Operator Console software anomaly, a change in treatment completion status of the last patient of the day may be altered from "Completed (or Discontinued)" to "Untreated" if all the following conditions are met: (1) At least one patient has been treated on the day; (2) A major error occurs when VeroTM/MHI-TM2000 is operated in a mode, other than Clinical Mode, which means Daily Check Mode or Maintenance Mode. (3) The major error is NOT associated with a function that uses therapeutic/ MV radiation parameters, such as an error in Isocenter Check function, etc. Note 1: If the major error is associated with a function that uses therapeutic/MV radiation parameters, the problem that treatment completion status is changed will never occur. The Operator Console changes the treatment completion status in the course of processing the major error. Note 2: Although treatment completion status is changed to "Untreated", its delivered dose result has been correctly saved into R&V; System. It has not been lost. Advice on Actions by Customer/Users: 1. Please make sure to perform the operations as described below in case a major error occurred in a mode other than Clinical Mode. (1) Cancel the error.(2) Click [Clinical] button on the left of Operator Console screen. (3) Verify, on R&V; System screen, that treatment completion status of the patients who were treated on the day are all correct. The treatment completion status is indicated in different colors as follows depending on the status: Untreated: BLUE, Completed: GRAY, Discontinued: RED. Customers requiring support concerning this issue, please contact the Customer Service Representative: 1-800-597-5911 (Mon - Fri) or via email us.support@brainlab.com. If any other problem is detected, please stop operation and contact the Customer Service representative. |  | 25 | 25 devices (4 domestically in U.S.A. and 21 internationally) | Distributed in the states of Florida, New York, Ohio &amp; Texas, and the countries of France, Germany, Japan, Italy, Korea, &amp; Belgium. | Terminated | April 20, 2017 | 492 | 25 | Computer | Computer
Z-2125-2016 | 74344 | Yes | NVZ |  |  | Cardiovascular | PMA | Medtronic, MyCareLink" Patient Monitor, Model 24950, Rx Only. |  |  | May 26, 2016 | 2016 | Medtronic Inc., Cardiac Rhythm and Heart Failure | Recently, a new software version was automatically sent to a subset of Model 24950 MyCareLink Monitors. After release, Medtronic identified an issue with the software that prevents implanted device data from being available to clinicians on the CareLink" Network. While the transmission appears successful to the patient, the transmitted data, including CareAlerts, are not visible to the clinic. | Software design | Medtronic sent an Urgent Field Safety Notice dated May 2016 to all affected customers. The letter described the issue and the product involved in the recall. The letter also included the "Customer Actions" for consignees to follow. Requested consignees to complete and return the Customer Confirmation Certificate to |  | 262 | 262 | Worldwide Distribution to Austria, Belgium, China , Czech Republic , Denmark , Finland , France, Germany, Hong Kong, India , Ireland , Israel, Italy , Malaysia, Netherlands, Norway, Poland, Portugal , Saudi Arabia, Slovakia, South Africa, Spain, Sweden, Switzerland, UAE, and United Kingdom. | Open, Classified |  |  | 262 | Computer | Computer
Z-1568-2016 | 73736 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | EndoWrist Stapler 45 and Stapler 30 instruments used on the da Vinci Xi systems (IS4000) with p5 software. |  |  | March 25, 2016 | 2016 | Intuitive Surgical, Inc. | Potential for unexpected motion of the Xi Stapler jaws on the da Vinci Xi System with p5 software relating to a combination of the p5 software and the surgeon quickly transitioning from the clamp to the fire pedal during use. | Software design | Intuitive Surgical sent an Urgent Product Safety notification to customers via e-mail on March 25, 2016. Customers were advised to STOP USE on the EndoWrist Stapler 45 and 30 instruments on da Vinci XI systems with p5 software. Product letters followed and were mailed on March 31, 2016 by Fed Ex.Urgent Medical Device letters were sent via Fed Ex on March 31, 2016 (dated April 1, 2016). These letters reiterated what was explained in the previous e-mail and provide more information of the unexpected motion issue. In addition to the stop use request, letters requested that all personnel be fully informed of the issue.Further information or support concerning this issue, please contact your Clinical Sales Representative or contact ISI Customer Service at the number:North America: 800-876-1310 Option 3 (6 am to 5pm PST) |  | 43 | 43 | US Nationwide Distribution in the states of: WI, FL, OH, TX, CA, AZ, IL, MD, MN, NY, OK, HI, MO, IA, KS, PA, and AL. | Open, Classified |  |  | 43 | Computer | Not_Computer
Z-2135-2016 | 74454 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | MEXL-3010/A5- Vantage Titan 3T Magnetic Resonance Imaging System |  |  | April 14, 2016 | 2016 | Toshiba American Medical Systems Inc | A potential problem associated to the gradientcoil (ASGC) in the Titan 3T MRI System. | Device Design | An Urgent Medical Device Notification letter was sent to customers to alert them to a potential problem associated to the gradient coil (ASGC) in the Titan 3T MRI System. The letter informs the customers of the problem identify and the corrective action. Customers are informed that their Toshiba service representative will contact them for an appointment to install the modifiedsoftware on their system when it becomes available. Customers are instructed to complete the attached customer response form and fax it to the toll-free number shown at the top of the form; sent via email to TAMS-RegulatoryAffairs@toshiba.com. Customers with any questions are instructed to contact Paul Biggins at (800) 421-1968 or to contact their local Toshibarepresentative at (800) 521-1968. |  | 33 | 33 units | US Nationwide Distribution | Terminated | March 30, 2017 | 350 | 33 | Computer | Not_Computer
Z-2704-2016 | 74830 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | ADVIA Chemistry XPT, SMN 10723034, IVD. --- This issue affects only the ADVIA Chemistry Hemoglobin A1c_3 Automated Pretreatment (A1c_3) results when HbA1c values are reported in International Federation of Clinical Chemistry (IFCC) equivalent units (HbA1cR). |  |  | July 20, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc. | Siemens identified an issue with the ratio equation provided on the ADVIA Chemistry XPT System Software Test Definition (TDef) Version 1.0 disks (SMN: 11127343 and 11222123). This issue affects only the ADVIA Chemistry Hemoglobin A1c_3 Automated Pretreatment (A1c_3) results when HbA1c values are reported in International Federation of Clinical Chemistry (IFCC) equivalent units (HbA1cR). The ratio equation used to calculate ADVIA Chemistry Hemoglobin A1c_3 Automated Pretreatment (A1c_3) results in IFCC units contains an error. The error results in falsely depressed HbA1cR results. | Software design | Siemens sent an Urgent Field Safety Notice dated July 20, 2016, to all affected customers via Fed Ex and the Urgent Field Safety Notice (CHSW16-02 July 2016) was e-mailed to regional countries for implementation Outside US (OUS). Customers have been advised to stop reporting results and call Siemens immediately to schedule service. The file with the typographical error is located on ADVIA Chemistry XPT Test Definition (TDef) disk. -- Field service personnel were sent a support bulletin describing the issue and instructing them how to correct the equation. Service engineers have been instructed to remove all version 1.0 TDef disks from customer sites regardless of whether they run the ADVIA Chemistry Hemoglobin A1c_3 Automated Pretreatment assay and report in IFCC units.Customers with questions were advised to contact their Siemens Customer Care Center or their local Siemens technical support representative. |  | 229 | Domestic: 32 units; Foreign: 197 units | Worldwide Distribution - US to MA, NY, MI, IL, TX, CO, AZ, CA, WA and Foreign distribution to the following countries: Australia, Austria, Bahrain, Brazil, China, Denmark, Egypt, Finland, France, Germany, Greece, Hungary, Italy, Malaysia, Netherlands, New Zealand, Norway, Portugal, Portugal, Republic Korea, Singapore, Slovakia, Spain, Sweden, Turkey, United Kingdom, and Vietnam. | Open, Classified |  |  | 229 | Computer | Not_Computer
Z-2690-2016 | 74525 | No | MUE | Full-field digital mammography system. | Radiology | Radiology | 510(k) | FUJIFILM Digital Mammography System Aspire HD (FDR MS-1000) with Biopsy Positioner (FDR-1000BPY) |  |  | June 09, 2016 | 2016 | Fujifilm Medical Systems U.S.A., Inc. | If an exposure is aborted during a stereo biopsy workflow, an error may occur which would prevent the system from continuing to operate properly. | Software design | The firm, Fujifilm, sent a "FIELD SAFETY NOTICE" letter dated June 9, 2016 via Federal Express to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to read and follow the instructions in the letter; call your local service representative, in case an error occurs, and complete and return the Customer Feedback Form via fax (203)-251-7862. Both affected systems will be corrected via on-site visits by FMSU Customer Service Engineers.If you have any questions about this matter, please contact the Product Manager, Women's Health directly at (303)-888-4549 or call your local FUJIFILM office. |  | 2 | 2 Units | US Distribution to: NC and HI. | Open, Classified |  |  | 2 | Computer | Computer
Z-2683-2016 | 74793 | No | BYD | Heat and moisture condenser (artificial nose). | Anesthesiology | Anesthesiology | 510(K) Exempt | MAQUET Servo Humidifier 163, Inmed Mfg. Sdn. Bhd. Lot PT 2577, Jalan Perusahaan, 4, 34600 Kamunting, Perak, Malaysia. |  |  | December 03, 2015 | 2015 | Teleflex Medical | Cracking may occur at the 15mm and 22mm ISO femaleconnector. | Device Design | INMED sent an Urgent - Field Safety Notice dated December 3, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.They were instructed to cease use and distribution of stock of the affected product batch and quarantine immediately. Customers are to return all affected product to MAQUET. An Acknowledgement form was included to be completed and returned to their MAQUET representative. If the product was distributed, they were instructed to provide the field safety notice to all of their customers who have received product in scope of the Field Action. Their customer is then required to complete the Acknowledgement form and return it to them.Should you require any further information or support concerning this issue, please contact your local MAQUET representative at fieldactions@maquet.com.INMED is committed to providing high quality, safe and effective products. W e sincerely apologize for any inconvenience this action may cause your operations. If you have any other questions, feel free to contact your local sales representative or Customer Service. For further questions, please call (610) 378-0131. |  | 2700 | 2,700 units | US Distribution to the states of : FL, KS, and WV. | Completed |  |  | 2700 | Computer | Not_Computer
Z-2685-2016 | 74924 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | RadSuite software. The firm name on the label is Merge Healthcare, Hartland, WI. |  |  | January 30, 2016 | 2016 | Merge Healthcare, Inc. | Interventional Radiology (IR) images are stored as JPEG2k Lossless in Merge Enterprise Archive (EA) and are not displaying correctly in RadSuite. | Software design | The firm, Merge, sent an "URGENT: MEDICAL DEVICE RECALL" letters dated 1/29/2016 and 5/25/2016 via e-mail or via certified mail if they did not have an e-mail address for the customers. The letters described the product, problem and actions to be taken. The customers were instructed to do the following: you must discontinue using the affected version of RadSuite or iConnect Enterprise Archive with GE AW Workstation; use the workaround listed in the letter; complete and return the enclosed form no later than June 17, 2016; ensure that all users of the product are provided with the notification; if you have further distributed this product, please identify your customers and notify them at once of this product recall Merge has a released fix available for this issue.If you have any additional questions, please send an email to recall@merge.com or call 800-724-5970 ext 8027. |  | 96 | 96 sites potentially have the affected versions for both products | US Distribution to: CO and OK. | Open, Classified |  |  | 96 | Computer | Computer
Z-1243-2016 | 73052 | No | CHP | Estradiol test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | IMMULITE 2000 EstradiolTest Code E2 |  |  | January 13, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc | The drug fulvestrant (FaslodexÂż) may cause falsely elevated estradiol results. | Device Design | Siemens issued an Urgent Medical Device Correction (UMDC) to affected Siemens Healthcare Diagnostics customers in the United States on January 13, 2016 to be delivered to customers on January 14, 2016 via Federal Express. An Urgent Field Safety Notice (UFSN) was sent out for distribution to all customers outside the United States on January 13, 2016. The letters identified affected product, problem and the actions to be taken. Customers may continue to use the estradiol assays to report patient results however, the assays should not be used for patients being treated with fulvestrant. Customers where instructed to complete and return the Correction Effectiveness Check attached. For questions contact your Siemens Customer Care Center or your local Siemens technical support representative. |  | 9334 | 9334 units | Worldwide Distribution  US Nationwide , AR, AU, AU, BD, BO, BR, CA, CH, CL, CN, CO, DO, EC, EG, HK, ID, IE, IL, IN, JP, KR, LK, LY, MX, MY, NI, NP, NZ, PE, PH, PY, SA, SG, SV, TH, TW, UY, VE, VN, and ZA. | Open, Classified |  |  | 54134 | Computer | Not_Computer
Z-1266-2016 | 73157 | No | KCT | Sterilization wrap. | General Hospital | General Hospital | 510(k) | The S.C.O.R.E.S. Unit (self contained operating room equipment |  |  | January 13, 2016 | 2016 | PMBS, LLC | Water retention issues. Units with filters that remain wet after the sterilization may compromised sterility. | Device Design | On January 13, 2016 PMBS, LLC sent letters to their customers informing them of the correction. Per product's instructions for use, the SCORES cabinet is to be run for a minimun of a 4 minute sterilization cycle, followed by a 30 munites dry time followed by a 30 minutes door crack time. not all facilities generate completely dry SCORES filters following these instructions. It is imperative that the filters be checked in the Operation Room prior to useing the surgical instruments sterilised within the cabinet. If the filters are wet, do not use the contents. In the letter is a IFU addendum including instructions for running the equiptment with extended cycles. For more information please contact Ms. Michele Mauzerall, CEO at 908-300-7093. |  | 58 | 58 units | FL, GA, MN, NY, OH, OR SC, and TN | Open, Classified |  |  | 58 | Computer | Not_Computer
Z-2134-2016 | 74515 | No | FET | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Image1 SPIES X-LINK Camera Control Unit (CCU), Product Model No. TC301, when used with Image1 SPIES D1 camera head, Product Model No. TH116 &amp; TH115, and DC1 Video Mediastinoscope |  |  | June 20, 2016 | 2016 | Karl Storz Endoscopy | Potential electrical interference issue when the Image1 SPIES X-LINK and Image1 SPIES D1 camera head are used in conjunction with the DCI Video Mediastinocope and an electro-surgery generator (ESG). | Device Design | KARL STORZ sent an Customer Notification letter dated June 22, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. KARL STORZ has become aware become aware of a potential electrical interference issue when the lmage1 SPIES X-LINK and lmage1 SPIES D1 camera head are used in conjunction with the DCI Video Mediastinoscope and an electro-surgery generator (ESG). Customers are instructed to complete the attached acknowledgement Form and to return it to KARL STORZ. Customers with questions are instructed to contact Technical Support Dept. at (800) 421-0837, press 2, and then press 1. |  | 619 | 619 units total (90 units in US) | Worldwide Distribution - US Distribution to the states of : IL, NM, TX, MD, TN, PA, FL, KY, MO, NY, MN, DC, MA, WI, CT, MI, ID, CO, GA, NJ and AR., and to the countries of : Argentina, Australia, Belgium, Bulgaria, China, Germany, Finland, Great Britain, India, Italy and Japan. | Terminated | January 05, 2017 | 199 | 619 | Computer | Computer
Z-1265-2016 | 73400 | No | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Vanguard XP Knee System Porous Femoral Component |  |  | February 05, 2016 | 2016 | Biomet, Inc. | The anterior/posterior (AP) dimension may be oversized by 0.016 . If the femoral component does not fit as tightly as intended, bony ingrowth may not occur causing loosening of the femoral component. A revision surgery may be necessary. | Device Design | On 2/5/2016, URGENT MEDICAL DEVICE RECALL NOTICE notifications were sent to the affected distributors, hospital risk managers, and implanting surgeons via courier. Distributors were instructed to locate and return the affected product, hospitals to locate and discontinue use, and implanting surgeons to monitor patients. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Questions related to this notice should be directed to (574) 372-1570, Monday through Friday, 8 a.m. to 5 p.m. |  | 890 | 890 | UTNetherlands | Terminated | March 28, 2017 | 417 | 890 | Computer | Not_Computer
Z-1249-2016 | 73494 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Reveal LINQ insertable Cardiac Monitor, Model LNQ11. |  |  | February 23, 2016 | 2016 | Medtronic Inc. | Medtronic has identified an issue with the sensitivity of an algorithm used in the Reveal LINQ ICM that may prematurely trigger the Recommended Replacement Time (RRT) alert in some devices. Battery capacity is not affected and the device will continue to support data collection and manual data transmissions. As part of the normal behavior of the device, 30 days after RRT status is reached, Reveal | Software design | Consignees in US were hand delivered a Medtronic "Medical Device Correction" letter dated February 2016. The letter was addressed to Physician or Healthcare Professional. The letter described the Issue &amp; Patient Management Guidance. For questions contact your local Medtronic Representative or Medtronic Diagnostic Technical Services at 800-929-4043.Outside the U.S., the letter will be mailed or hand delivered to impacted consignees. |  | 189009 | 189,009 | Worldwide Distribution-US (nationwide) including DC and the countries of Andorra, Australia, Austria, Belgium, Botswana, Brunei Darussalam, Bulgaria, Canada, China, Colombia, Croatia, Czech Republic,Denmark, Faroe Islands , Finland, France, Germany, Greece, Hong Kong,Hungary, Iceland, Indonesia, Iran, Ireland, Israel, Italy, Kazakhstan, Kenya,Kuwait, Latvia, Lithuania, Luxembourg, Malaysia, Malta, Martinique, Namibia,Netherlands, New Caledonia, New Zealand, Norway, Oman, Philippines, Poland, Portugal, Qatar, Russian Federation, Saudi Arabia, Serbia, Singapore,Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Tanzania,Thailand, Turkey, United Arab Emirates and United Kingdom. | Open, Classified |  |  | 189009 | Computer | Computer
Z-2671-2016 | 64796 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Centricity Laboratory Core Lab System 4.1 |  |  | November 20, 2012 | 2012 | GE Healthcare It | The Centricity Laboratory removes free-text notes on a master panel when the user deletes the slave panel. The deletion of a slave panel will (auto slash) the corresponding test (item) results on the master panel. Additionally, the displayed and HL7 results show neither the appropriate abnormal flag value nor the reference range values, when using large_num.g &gt;7 to report numeric results. | Software design | GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated Nov 20, 2012 to all affected customers. The letter described the product, problems and actions to be taken. The customers were instructed to follow the safety instructions and to ensure that all potential users in their facility are made aware of this safety notification and the recommended actions. The software has been modified to correct these defects. The modified software will be released and installed in user sites through GE Healthcare FMI #86071.Customers were instructed to contact their GE Service Representative with any questions regarding this issue. For other questions, contact GE Helpdesk:North America; (905) 479-1249 or 1-888-778-3375; email: tgg.nahelpdesk@med.ge.comEurope; +44 1603 877595; email: tgg.ukhelpdesk@med.ge.comAsia Pacific; +61 3 5327 4000; email: tgg.ozhelpdesk@med.ge.com.For questions regarding this recall call 847-277-5000. |  | 17 | 17 | Worldwide Distribution - US Distribution to GA, OH, NY, TX, TN ,MN and Internationally to India, Australia, Canada, England, Qatar, and Scotland. | Open, Classified |  |  | 17 | Computer | Computer
Z-2659-2016 | 74937 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | NovaPACS versions 8.4.7, 8.5.3, and 8.5.6. Distributed by Novarad Corporation. Picture archiving and communication system (PACS) |  |  | August 12, 2016 | 2016 | Novarad Corporation | Novarad Corporation announces a voluntary field action for the NovaPACS versions 8.4.7, 8.5.3, and 8.5.6 due to a "pop-up screen" that was prematurely released. | Software design | Novarad sent an Urgent Medical Device Correction letter dated August 12, 2016, to all affected consignees. As a workaround, customers are advised to use the HTMLS Viewer for review purposes only and not for diagnostic purposes. The issue has been fixed in all future releases of NovaPACS software. You can eliminate the possibility of incorrect information appearing by contacting Novarad Technical Supoort to arrange for an upgrade to a corrected version of NovaPACS software. This correction is available to all affected customers at no cost and will be included din all future upgrades.Customers will need to fax, mail, or email th eincluded Acknowledgement and Receipt Form back to Novarad. Customers with questions or concerens regarding this notification or would like to request an upgrade, should contact Novarad Technical Support at 801-221-5895.For questions regarding this recall call 801-642-1001. |  | 90 | 90 | Worldwide Distribution - US (Nationwide) Canada and other foreign consignees. | Open, Classified |  |  | 90 | Computer | Not_Computer
Z-2658-2016 | 74820 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Ingenuity Core128, Model number 728323 |  |  | July 20, 2016 | 2016 | Philips Medical Systems (Cleveland) Inc | After upgrading to 3.6.7 software version via FCO72800643, during reconstruction of gated helical scans, the planned anatomy can be cut off on the CT images. This issue can occur on gated CT helical reconstruction that is not planned at iso-center (0,0). If the region of interest is not visualized in the images, a CT rescan may be performed. | Software Design Change | Customers were sent a letter dated July 28, 2016 informing them of the problem, under what circumstances it can occur and the actions that the customer can take to minimize the effect of the problem. The letter also includes the actions planned by Philips to correct the problem. Philips plans on releasing a field change order (FCO) to correct the affected systems.If you need any further information or support concerning this issue, contact your local Philips representative or local Philips Healthcare office. ForNorth America and Canada, contact the Customer Care Solutions Center (1 -800-722-9377). |  | 3 | 3 | US Nationwide Distribution in the states of: AL,AR,AZ,CA,CO,CT,DC,DE,FL,GA,HI,IA,ID,IL,IN,KS,KY,LA,MA,MD,ME,MI,MN,MO,MS,NC,ND,NE,NJ,NM,NV,NY,OH,OK,OR,PA,PR,SC,TN,TX,UT,VA,VT,WA,WI,WV,WY | Open, Classified |  |  | 615 | Computer | Computer
Z-2532-2016 | 74734 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | iConnect Access used with Ortho PACS software provides medical specialists with access to diagnostic quality images, reports, and various types of patient data over conventional TCP/IP (e.g., internet) networks. |  |  | January 30, 2016 | 2016 | Merge Healthcare, Inc. | Studies that are viewed in iConnect Access that originate on Merge PACS / Ortho PACS may not display the correct current patient demographics if there has been a PDE (Patient / Study Demographic Edit) performed after the images were ingested in PACS. | Software design | The recalling firm issued a letter dated 1/28/2016 via e-mail on 1/31/2016 or via certified mail if they did not have an e-mail address for the customer. |  | 146 | 146 sites may have the affected version for both products | Distribution was made to medical facilities nationwide and to one foreign medical facility in New Zealand. There was no military or government distribution. | Open, Classified |  |  | 146 | Computer | Not_Computer
Z-2526-2016 | 74520 | No | DOP | Mass spectrometer for clinical use. | Clinical Chemistry | Toxicology | 510(K) Exempt | AB Sciex API 3200MD" LC/MS/MS System. In-Vitro Diagnostic |  |  | June 16, 2016 | 2016 | Ab Sciex | MultiQuant MD software where under certain conditions a usercan be presented with incorrect quantitative results when using the Sum Multiple Ions feature. | Software design | AB SCIEX issued recall notification by letter dated June 16, 2016, to all affected customers. The letter identified the problem, provided temporary measures until a permanent software is available. A response form was to be completed and returned. questions regarding this noticecontact SCI EX at +1 289 982 2712. |  | 124 | 124 units | Worldwide Distribution - US (Nationwide) Internationally to Canada, Austria, China, Netherlands, Norway, India, Korea, Germany, France, Switzerland, Ireland, Italy, Singapore, Russia, UK, Malaysia, | Open, Classified |  |  | 275 | Computer | Computer
Z-1325-2016 | 73487 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius 2008 Series Hemodialysis Delivery Systems: 2008T, 2008K2, 2008K and 2008K@home: |  |  | March 10, 2016 | 2016 | Fresenius Medical Care Renal Therapies Group, LLC | Added Warning Statement: Fresenius 2008 Series Hemodialysis Machines False Blood Leak Alarm When Dialyzing Patients Treated with Hydroxocobalamin (or any form of Vitamin B-12) | Device Design | Fresenius Medical Care issued an Urgent Medical Device Alert dated March 10, 2016, to all affected consignees. Consignees were instructed to place notification and Warning Addendum with their 2008Âż Series Hemodialysis Machine Operators Manual. A reply form was included as part of the customer communication with instructions to fill out and return via fax or scanned/emailed copy.Further information or support concerning this issue, customers were instructed to contact Fresenius Technical Services at 800-227-2572. |  | 8384 | 8384 units (8369 US) (15 OUS) | Nationwide Distribution | Open, Classified |  |  | 8384 | Computer | Not_Computer
Z-1336-2016 | 73412 | No | MDS |  |  | Clinical Chemistry | PMA | Dexcom G4 PLATINUM (Professional) Receiver |  |  | February 23, 2016 | 2016 | Dexcom Inc | Customers may not receive an intended audible alarm or alert if relying on hearing the alarm or alert. | Device Design | Dexcom sent out a customer notification letter by certified mail to all direct customers on 02/23/16 notifying customers that they may not receive an intended audible alarm or alert if relying on hearing the alarm or alert. As a result customers may not detect a severe hypoglycemic (low glucose) or hyperglycemic (high glucose) event.The letter instructs the customer to peridically test the alarms and alerts on their receiver to make sure that the alarms and alerts are functioning properly (especially if the receiver gets wet of is dropped). Receiver part numbers and instructions for testing the alarms and alerts can be found on the back of the notification as well as in the User's Guide.Should customers find that alarms and alerts are not functioning properly or if any questions regarding the notification contact the deciated hotline tool free at 844-607-8398 or dial 858-291-1700.The healthcare professional notification that was sent out provided similar instructions as the direct customer notificaiton in which it is recommended that healthcare professionals periodically check the speaker function by following the testing instructions attached in the notification.The distributor notification letter instructed them to take the following actions:Hold shipments of Dexcom receivers and sticker all existing Dexcom receiver boxes on the outside front of the box with the enclosed notification stickers (see the attached visual guide for sticker placement). Once completed can begin shipping the products.In order to ensure that the notification is timely received by patients select one of the following two (2) options by no later than February 25, 2016:Send, via certified mail, the notification directly to those patients that have purchased a Dexcom receiver from you. If elect this option, Dexcom will provide the required quantity of notifications enclosed in certified mail envelopes. In order to facilitate this effort email the number of notifi |  | 5722 | 5,722 | Worldwide Distribution-US (nationwide), Canada, Italy, Netherlands, Belgium, Norway, New Zealand, Poland, Sweden, Denmark, Slovenia, Turkey, South Africa, India, Israel, Hungary, Hong Kong, UK, Ireland, France, Finland, Spain, Portugal, Colombia, Chile, Germany, Austria, Switzerland, Kuwait, United Arab Emirates, Australia, Czech Republic, Bahrain, Lebanon, and Qatar. | Terminated | August 11, 2017 | 535 | 270111 | Computer | Computer
Z-1553-2016 | 73349 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ABX Pentra 400 (version 5.0.8 or lower) |  |  | February 19, 2016 | 2016 | Horiba Instruments Inc | Horiba Instruments, Inc. is recalling ABX Pentra 400( version 5.0.8 or lower) and Pentra C400 (version 1.1.2 or lower) because clinical chemistry analyzer malfunctions when the following certain alarms appear on the system. | Software design | HORIBA Medical sent an "Urgent Safety Notice" dated February 19, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. A new software version to solve the issue is currently under development; however the firm requests customers use the provided procedure until the new software version is installed on the system, or be extremely vigilant with results of one of the alarms mentioned appears on the system. The letter instructs customers to share the information with laboratory staff and retain the notification as part of laboratory documentation. The letter states it is mandatory for customers to complete and return the enclosed response form within 10 days, in order to verify understanding of it.Any questions or concerns contact the local Horiba Medical representative.For further questions, please call (949) 250-4811. |  | 201 | 201 | USA (nationwide Distribution) to the states of : CA, NY, AL,OK, MS, MD, NV, MN, TX, KY, WV, MI, NY, MT, IL, FL,WA, ID, MO, WY, NJ, VA, KY, GA, PA, TN, MS, KS, AZ, NE, SC, OR, OH, IA, RI, MI and NC. | Open, Classified |  |  | 201 | Computer | Computer
Z-1377-2016 | 73649 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | da Vinci S Surgical System (IS4000) user manual English and OUS. |  |  | March 16, 2016 | 2016 | Intuitive Surgical, Inc. | Intuitive Surgical has found that the use of instruments other than the Cardiac Probe Grasper (e.g. LargeNeedle Driver Instrument) to manipulate the ablation probes during cardiac surgery can result in an increase of microscopic metallic particulates generated from the contact with the probes. | Device Design | Intuitive sent a Field Safety Notice/Urgent Medical Device Correction letters dated March 22, 2016 by trackable mail method to their customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to 1. In addition to existing instructions for use, observe the following warning when manipulating the cardiac ablation probe, Use only the Cardiac Probe Grasper to manipulate a cardiac ablation probe.2. Distribute a copy of this letter to all da Vinci S, Si, and Xi users at your facility.3. Place a copy of this letter with your user manual.4. Complete the attached Acknowledgement Form and return it to Intuitive Surgical as instructed.5. Retain a copy of this letter and the Acknowledgement Form for your files.Intuitive Surgical will provide a user manual addendum to incorporate the warning related to this risk.For questions contact your Clinical Sales Representative or contact Intuitive Surgical Customer Service. |  | 248 | 248 - total all models | Worldwide Distribution - US Nationwide in the countries Australia, Belgium, Brazil, Canada, Chile, Finland, France, Germany, Greece, Iceland, India, Israel, Italy, Netherlands, Qatar, Russia, Singapore, South Korea, Sweden, Switzerland, Taiwan, Turkey, United Arab Emirates, and United Kingdom. | Open, Classified |  |  | 248 | Computer | Not_Computer
Z-2307-2016 | 74731 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | MultiDiagnost Eleva; x-ray system. Product Codes: 708032, 708036 |  |  | July 19, 2016 | 2016 | Philips Electronics North America Corporation | The two bolts that attach the clamping plate to the lateral shaft may break, which could cause the c-arm to move laterally forward, slide out of its bearings, and sag | Device Design | On July 19, 2016, Philips sent a Field Safety Notice to all affected customers. Philips Field Service Engineers will add an independent safety catch free of charge. |  | 993 | 993 total | Worldwide distribution. US nationwide, Canada, Algeria, ARGENTINA, Australia, Austria, Belgium, China, Colombia, Czech Republic, Denmark, Egypt, Estonia, Faroe Islands, Finland, France, French Guiana, Germany, Greece, Hong Kong, Indonesia, Ireland, Israel, Italy, Jersey, Jordan, Republic of Korea, Latvia, Lebanon, Luxembourg, Malaysia, Mexico, Netherlands, New Zealand, Norway, Oman, Philippines, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Syrian Arab Republic, Taiwan, Thailand, TURKEY, Ukraine, United Arab Emirates, and United Kingdom. | Open, Classified |  |  | 993 | Computer | Not_Computer
Z-1375-2016 | 73630 | No | DRC | Trocar. | Cardiovascular | Cardiovascular | 510(K) Exempt | Boston Scientific TSX Transseptal Needle, Model Number M004TSX60o; |  |  | March 15, 2016 | 2016 | Merit Medical Systems, Inc. | Merit Medical Systems, Inc. is voluntarily conducting a recall of one lot of the TSX" Transseptal Needle due to a discrepancy between the labeling and the actual configuration of the device in the packaging. The curve dimensional values on the labels of this lot incorrectly indicate a 50Âż curve, rather than the correct 86Âż curve | Software design | The consignee was provided verbal and written notification of the field action on March 15, 2016. There was a total of one(1) consignee, an OEM customer in the USA, who received 25 units. Written notification included product identification information, instruction to immediately quarantine any devices and discontinue distribution, and instructions to initiate a product recall in accordance with their company's established policies for customers to whom they shipped the affected Merit product. An email notification was sent to the Merit sales representative of the affected account. The sales representative was provided a Customer Response Form applicable to their account. The sales representative was instructed to assist the account with the completion of the Customer Response Form and the return of the affected units back to the Merit Field Assurance Department. |  | 25 | 25 | Distributed to one customer in the US. | Open, Classified |  |  | 25 | Computer | Not_Computer
Z-2301-2016 | 74648 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge PACS software. The firm name on the label is Merge Healthcare, Hartland, WI. |  |  | January 30, 2016 | 2016 | Merge Healthcare, Inc. | Studies coming over via telmed were missing patient's DOB, procedure, and referring physician. | Software design | The recalling firm issued Urgent: Medical Device Recall letters dated 1/27/2016 and 5/26/2016. The letter informed the customer of the affected product, the nature of the issue, the potential harm involved, and actions to be taken. The firm released a fix for the issue; and, customers were asked to complete and return the enclosed form by June 17, 2016. Questions can be directed to recall@merge.com. |  | 109 | 109 sites have the affected versions | Worldwide Distribution -- USA, New Zealand, and United Kingdom. | Open, Classified |  |  | 109 | Computer | Not_Computer
Z-1529-2016 | 73602 | No | LHN | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MEVION S250 for Proton Radiation Therapy |  |  | March 15, 2016 | 2016 | Mevion Medical Systems, Inc. | Possibility of Gantry Motion with Aperture or Compensator only partially inserted. An Aperture may fall out during Gantry motion. It is possible that it can hit a patient and cause traumatic injury or death. | Device Design | Customer notification letters were sent to customers on March 15, 2016, by certified mail, return receipt requested, as well as by email. The letter explains the problem, description of potential harm, and the cause containment, and mitigation. The cause of the malfunctioning sensors was discovered to be a non-conforming limit switch assembly. A Mevion Field Service Engineer will immediately inspect all such parts in the field for similar non-conformance, and replace those parts if necessary. In addition, a new periodic preventive maintenance check will be performed to individually test each of the redundant sensors at all sites. In the interim period, on a daily basis, until the inspection, repair, and the preventive maintenance check is completed at any site, Mevion Field Service will check that the Applicator Door interlock sensors are working properly. Mevion will also update the inspection procedure for incoming parts to ensure that any non-conformance is identified early in the manufacturing process. |  | 5 | 5 | US nationwide distribution to MO, OK, FL, and NJ. | Open, Classified |  |  | 5 | Computer | Not_Computer
Z-2303-2016 | 74651 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge PACS software. The firm name on the label is Merge Healthcare, Hartland, Wisconsin. |  |  | January 30, 2016 | 2016 | Merge Healthcare, Inc. | A migrated study that has annotations will display the annotations, but when a prior study is loaded that should have annotations, the second study will not display annotations. | Software design | The recalling firm, Merge, issued an "URGENT: MEDICAL DEVICE RECALL" letter dated 1/27/2016 via e-mail on 1/30/2016 or via certified mail if they did not have an e-mail address for the customers. The letter described the product, problem and actions to be taken. The customers were instructed to not discontinue use of Merge PACS; to open prior comparison studies in the secondary viewer to check if there were annotations and read the report or visit note pertaining to that study to see if there are noted findings that were not otherwise apparent; to receive an upgrade call Merge Customer Service at (877) 741-5369 or send an email to support@merge.com to schedule the upgrade (Reference this call #2016-010); complete the online form: http://www.merge.com/Landing-Pages/Customer-Upgrade-Decline-Form.aspx no later than 2/22/2016; If you decide to decline the maintenance release, complete the attached Customer Decline Response form and return to Merge Healthcare via recall@merge.com, and ensure all users of the product are provided with this notification. If you have further distributed this product, identify your customers and notify them at once of this product recall.If you have any additional questions, call Merge Customer Service at (877) 741-5369 or support@merge.com. |  | 208 | 208 sites have the affected versions | Worldwide Distribution: US (Nationwide) and countries of: Australia, Belgium, Canada, Jordan, New Zealand, and the United Kingdom. | Open, Classified |  |  | 208 | Computer | Computer
Z-1380-2016 | 73665 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | EPIQ DIAGNOSTIC ULTRASOUND SYSTEM, Model EPIQ 5C, EPIC 5G, EPIQ 5W, EPIQ 7C, EPIC 7GC, and EPIQ 7W. |  |  | March 28, 2016 | 2016 | Philips Ultrasound, Inc. | The fasteners securing the control panel assembly to the base of the Philips EPIQ Ultrasound System may loosen over time, which could subsequently lead to the detachment of the entire assembly from the ultrasound system. | Device Design | The firm, Philips, sent an "Urgent-Medical Device Correction" Philips EPIQ Ultrasound System (MDC 79500381/2), letter dated 2016 MAR 23 to consignees on 3/28/16. The letter describes the product, problem and actions to be taken. The customers were instructed to review the information with all members of your staff who need to be aware of the contents of this communication.If, at any time, the control panel assembly on your ultrasound system wobbles or feels loose, stop using your system immediately and contact your local Philips representative or Philips Customer Service at 1-800-722-9377. Otherwise, you may continue to use your system.Philips will contact all EPIQ customers to arrange for service to replace the fasteners connecting the control panel assembly to the system with fasteners less likely to loosen with repeated handling. This service will be performed free of charge. If you need any further information or support concerning this issue, please contact your local Philips representative or Philips Customer Service at 1-800-722-9377. |  | 11085 | 11,085 units total (4909 units in the US and 6176 units outside the US) | Worldwide Distribution: US (nationwide) including Washington, D.C., and countries of: Algeria, Argentina, Armenia, Australia, Austria, Bahrain, Bangladesh, Belgium, Bermuda, Bolivia, Brazil, Brunei Darussalam, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cuba, Czech Republic, Denmark, Ecuador, Egypt, Estonia, Finland, France, French Guiana, French Polynesia, Georgia, Germany, Greece, Guadeloupe, Guatemala, Hong Kong, Hungary, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jersey, Jordan, Kazakhstan, Kenya, Korea, Republic of, Kuwait, Latvia, Lebanon, Libya, Lithuania, Luxembourg, Macedonia, Malaysia, Malta, Martinique, Mayotte, Mexico, Monaco, Mongolia, Morocco, Myanmar, Netherlands, New Caledonia, New Zealand, Nicaragua, Norway, Oman, Pakistan, Palestine, State of, Panama, Peru, Philippines, Poland, Portugal, Puerto Rico, Qatar, RÂżÂżunion, Romania, Russia , Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uzbekistan, and Viet Nam. | Open, Classified |  |  | 11085 | Computer | Not_Computer
Z-2300-2016 | 64758 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Imagecast PACS with Centricity RIS-IC versions prior to 10.6 Update Package 18 |  |  | February 18, 2013 | 2013 | Ge Healthcare It | A software defect was discovered that causes images to be out of context with clinical information. | Software design | GE sent an "Urgent Medical Device Correction" letter to all affected customers on April 2, 2013.The letter identified the product the problem and the action customers need to take.The software has been modified to correct these defects. The modified software will be released and installed in user sites through GE Healthcare FMI #85204.For further questions, please call ( 847) 277-4468. |  | 13 | 13 | US Nationwide Distribution. | Open, Classified |  |  | 13 | Computer | Computer
Z-1379-2016 | 73595 | No | LWH | Surgical sponge scale. | General Hospital | General &amp; Plastic Surgery | 510(K) Exempt | SurgiCounter scanner, part of the SurgiCount Safety-Sponge System. |  |  | March 21, 2016 | 2016 | Stryker Instruments Div. of Stryker Corporation | The affected scanner does not correctly interact with the SC360 software. When docked pre/postoperatively, the affected SurgiCounter scanner may display an error message (code=2). Also, the SurgiCounter scanner software version number is not appropriately displayed in the SC360 application during installation. | Software design | Stryker sent an URGENT MEDICAL DEVICE RECALL NOTIFICATION letters dated March 21, 2016 to affected customers via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. For questions regarding this recall, please contact Stryker Instruments:Kelly Jo Davis269-389-2921kellyjo.davis@stryker.com |  | 286 | 286 | US Nationwide Distribution in the states of AR, AZ, CA, CT, FL, GA, IN, LA, MD, MN, MO, MS, NJ, NM, NY, OH, SC, SD, TN, TX and UT. | Terminated | September 06, 2016 | 169 | 286 | Computer | Computer
Z-2296-2016 | 74583 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Integra External Fixation System Rocker Bottom Model # 12225400. |  |  | July 01, 2016 | 2016 | Integra LifeSciences Corp. | The silicone insert embedded in the rocker bottom has the potential to become dislodged or loose from the rocker bottom mold. | Device Design | All affected consignees were notified via an "Urgent: Voluntary Medical Device Recall" letter sent on 7/1/16. The letter identified the affected product as well as the reason for the recall. Customers were asked to check their traceability records for shipments of the affected product. If the affected parts are found, customers are to stop use and distribution. A copy of the letter should be provided to any customers that the product may have been distributed to. The attached "Acknowledgement and Return Form" should be completed and returned regardless if customers have affected product or not. Questions regarding the instructions should be directed to Customer Service at 1-800-654-2873. |  | 27 | 27 units | Distributed in the US to the states of FL, VA, OK, CO, DC, and IN. | Open, Classified |  |  | 27 | Computer | Not_Computer
Z-2293-2016 | 74668 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MOSAIQ Oncology Information System |  |  | July 15, 2016 | 2016 | Elekta, Inc. | It is possible that a change to an Order Set will not be saved in the current open Care Plan. This can result in a wrong chemo dose in the order. | Software Design Change | Elekta sent Important Field Safety Notice 371-03-MSQ-001 to all customers on 7/15/2016. The notice informs users of the specific product and version numbers affected by the issue, and any work around that can be used to avoid the issue. The letter also included a response form which is to be returned to Elekta. |  | 383 | 383 | Worldwide distribution. US nationwide, Antigua and Barbuda, Australia, Austria, Bahamas, Bahrain, Canada, Czech Republic, Germany, Hungary, South Africa, Spain, Sweden, Turkey, and United Kingdom | Open, Classified |  |  | 383 | Computer | Not_Computer
Z-2291-2016 | 74529 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Molift Air 205, Molift Air 205 Propulsion, Molift Air IRC, Molift Air 205 IRC Propulsion, Molift Air 300, Molift Air 300 Propulsion, Molift Air 300 IRC, Molift Air 300 IRC Propulsion |  |  | June 24, 2016 | 2016 | Moller Vital | Potential for the bolt on the suspension coupling on the Molift Air to come loose from the housing. | Device Design | Moller Vital sent a Field Safety Corrective Action on June 22, 2016. Field correction instructions were provided to affected consignees. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to read the instruction, and return the recipient confirmation, report all serial numbers distributed in their region/country to Etac Supply Gjovik, order the upgrade kit from Etac Supply Gjovik, Customers were also instructed to follow up with the customers and users and confirm that all hoists in their market are upgraded. Customers with questions were instructed to contact Etac Supply Gjovik.at molift.qa@etac.com. |  | 516 | 516 units (374 US) | Worldwide Distribution - US (nationwide) and Internationally to SZ, TR, BE, CH, DE, DK, ES, FI, FR, GB, IS LV, NL, NO, SE, SI, AU, CN, IL, JP, NZ, IN, QA, PE, AE | Terminated | June 23, 2017 | 364 | 516 | Computer | Not_Computer
Z-2290-2016 | 74596 | No | IOR | Mechanical wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Quickie Q7 wheelchair. Model EIR4 |  |  | July 01, 2016 | 2016 | Sunrise Medical (US) LLC | There is an observed failure of the backrest bracket which can fatigue over the life time of the product. When this occurs there is the possibility of the bracket fracturingcausing the backrest to detach from the chair frame. | Device Design | The firm began calling dealers to alert them to the recall and obtain primary contacts information. Dealers were then sent their list of Quickie Q7 wheelchairs that records show being sold by them and dealers are asked to return a list of chairs they are able to locate for correction. The first dealer letter was sent via e-mail on July 7, 2016. |  | 13852 | 12,142 devices in US, 1,710 devices internationally. | US and Mexico, Japan, Chile, Kuwait, Columbia, Hong Kong, China, Uruguay, South Africa, Costa Rica, South Korea, Singapore, New Zealand, El Salvador, Bahrain, Uruguay, Lebanon, United Arab Emirates, Brazil, Argentina, Netherlands Antilles, British Virgin Islands, | Open, Classified |  |  | 13852 | Computer | Not_Computer
Z-2217-2016 | 74652 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Carestream Touch Prime, Catalog Number 1738830, and Carestream Touch Prime XE, Catalog Number 1738822 --- Common Name: Touch Ultrasound |  |  | July 08, 2016 | 2016 | Carestream Health Inc | A software issue related to the generic volume measurement functionality could result in the volume of the subject anatomy being overestimated. For example, measurement results of the anatomy may appear enlarged or distended when it actually is not. | Software design | The firm, Carestream Health Inc., sent an "URGENT: MEDICAL DEVICE CORRECTION" Letter dated 7/7/2016 and response form the end users/customers on 7/8/2016 via Certified Mail, Return Receipt Requested. The letter described the product, problem and actions to be taken. The Customers were instructed not use the generic volume measurement tool until Carestream has updated the system software. In the interim, calculate volumes by using the specific volume tool from the desired calculation package. If you have distributed the device outside your facility, please alert your customer(s) of this field correction and contact Carestream Customer Care Center as listed above.Carestream has created a software update for Touch Ultrasound to resolve this issue and will update your software. Carestreams representative will contact you and schedule a convenient time to upgrade the software.For questions or concerns, please contact the Carestream Customer Care Center in the US at 1-800-328-2910 (available on a 24/7 basis). Outside of the US, please call your local Service support number. |  | 7 | 7 Units (US: 5 units, OUS: 2 units) | Worldwide Distribution: US Distribution to states of: GA, IA, and TX; and country of: Italy. | Terminated | May 02, 2017 | 298 | 7 | Computer | Not_Computer
Z-2279-2016 | 73985 | No | ODG | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Ultrasound Gastroscope |  |  | April 06, 2016 | 2016 | Pentax Medical Company | Pentax Medical Video Processor EPK-i5010 has been shipped with a reusable Pentax Medical water bottle assembly (OS-H4) which has not been cleared for use with this devices in the United States. | Device Design | Pentax Medical sent an "Urgent Field Correction Letter and Response Form" dated April 6, 2016 via Certified Mail to their affected customers. The recall letter identified the product, problem and action to be taken by customer. The customer is asked to complete the response form even if the customers does not have any OS-H2/OS-H4 water bottles to discard. If customer needs assistance in completing the response form, they can contact their local sales representative. For questions regarding the recall, customers can contact Mr. Paul Silva, Pentax Americas Recall Coordinator at: (800) 431-5880 ext. 2064, from (8:30 am - 5:00 pm, Monday - Friday, EST), via fax: (201) 799-4063 or via email to: recall.coordinator@pentaxmedical.com. |  | 819 | 819 units | US Nationwide Distribution | Terminated | July 17, 2017 | 467 | 23755 | Computer | Not_Computer
Z-1394-2016 | 73430 | No | NFG | Medical image communications device. | Radiology | Ophthalmic | 510(K) Exempt | VERION Reference Unit (Vision Planner) |  |  | March 01, 2016 | 2016 | Alcon Research, Ltd. | Alcon is conducting a voluntary medical device correction of all VERION Reference Units (Vision Planner) that are shared with the Alcon LenSx Laser System after receiving reports concerning the inclusion of unplanned arcuates on printed, saved (.pdf) or exported surgical plans. | Software change control | Alcon sent an Urgent Medical Device Correction letter dated February 29, 2016 to affected customers. The letter identified the affected product, problem and actions to be taken. The letter stated that an Alcon Technical Services Engineers will contact customers, schedule Service Calls, and perform onsite corrections. For questions call 1-800-862-5266. |  | 97 | 97 units | Worldwide Distribution - US Nationwide throughout the US and the countries of: Argentina, Australia, Austria, Bangladesh, Belgium, Brazil, Bulgaria, Czech Republic, Chile, China, Colombia, Costa Rica, Denmark, Dominican Republic, Ecuador, Egypt, El Salvador, France, Georgia, Greece, Guatlemala, Hungary, India, Indoesia, Isreal, Italy, Japan, Korea, Malaysia, Mexico, Netherlands, Nicaragua, Norway, Panama, Philippines, Portugal, Romania, Singapore, Slovakia, South Africa, Spain, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Turkey, UK, Ukrain, United Arab Emirates | Terminated | April 13, 2017 | 408 | 97 | Computer | Computer
Z-2219-2016 | 74436 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | The pump (with cushioning foam inserts) is packaged in a single pump box. Four pump boxes are placed in an over-shipper for distribution. |  |  | June 14, 2016 | 2016 | WalkMed Infusion, LLC | WalkMed Infusion, LLC Announces a Voluntary Field Action of the Triton Infusion Pump (model 300000) and Triton FP Infusion Pump (model 400000) Due to Unapproved Changes to the Software and Specifications of the Triton Infusion Pump (model 300000) and Unapproved Changes to the Intended Use of the Triton FP Infusion Pump (model 400000). | Device Design | WalkMed sent an Urgent : Medical Device Recall letter dated June 14, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Action to be taken by Users: WalkMed recommends the following actions be taken with respect to the recalled product:Our records indicate that you have received Triton pumps. Immediately review your inventory and determine whether you have any of the Triton or Triton fp infusion pumps (model numbers 300000 and 400000).Please complete and promptly return the enclosed recall response form, even if you do not have any affected stock subject to this recall notice.WalkMed recommends that users stop use of and replace the Triton and Triton fp Infusion pumps with alternative devices as soon as possible.To ensure that patients who use WalkMed Triton infusion pumps can continue to receive infusion therapies without interruption, WalkMed will support customers in their transition to alternative devices. To ensure continuity of patient care, WalkMed will continue to provide IV administration sets to existing customers, provide preventive maintenance and service for distributed pumps and make available any parts necessary for the repair or service of units for up to 3 months from the date of this recall notice.Contact WalkMed at (303) 420-9569 to arrange for shipment of the product being returned.If an owner of a Triton or Triton fp infusion pump is interested in information about options for a replacement product, please contact WalkMed. For further questions please call (303) 420-9569. |  | 2482 | 2482 units | Nationwide to AL, AZ, CA, FL, IL, KS, MA, MD, MS,NJ, OH, PA, TN, TX, UT, and WA. No foreign/VA/govt/military. | Terminated | April 28, 2017 | 318 | 5448 | Computer | Not_Computer
Z-2263-2016 | 74043 | No | NEP | Intravascular administration set. | General Hospital | General Hospital | 510(K) Exempt | Dual Chamber Container fill with gravity/attached transfer sets, 250 mL / 1250 mL, 300 mL / 1700 mL, 500 mL / 2500 mL, 500 mL / 3500 mL |  |  | June 23, 2016 | 2016 | The Metrix Company | An issue was identified after a redesign of the auto compounding device. The compounder cannot meet the designed accuracy specification when dispensing into bags. After reviewing the initial design of the device, it was determined the dispensing volume validation was not properly done and none of the compounders on the market can meet the stated specifications. | Device Design | On 6/23/2016, the recalling firm sent written recall notices to their distributors and end users notifying them of the recall. Customers and consignees were instructed to not use the devices, locate and quarantine all affected devices, and to complete the response form. |  | 18474 | 18,474 case | Product distributed to medical facilities in the following US states: AL, AZ, CA, CO, CT, FL, IA, IL, KS, LA, MA, MD, MN, NC, OH, PA, PR, TN, TX, UT, and WA. Product was also distributed internationally to the following countries: Australia, Canada, Dominican Republic, Malaysia, Turkey, and United Kingdom | Open, Classified |  |  | 180049 | Computer | Not_Computer
Z-1504-2016 | 73776 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Extended Brilliance Workspace NM Nuclear medicine image display and processing application suite. AutoSPECT Pro Software Application. Philips Medical Systems, Cleveland, OH. |  |  | March 04, 2016 | 2016 | Philips Medical Systems (Cleveland) Inc | The AutoSPECT Pro application was only designed to reconstruct cardiac SPECT data obtained with detectors positioned at 90Âż or 180Âż relative to one another. However, certain gamma cameras allow for other relative detector angles. Data acquired at these other angles will not be correctly reconstructed by AutoSPECT Pro and the results will likely be erroneous. | Software design | On March 4, 2016, the firm sent an URGENT - Field Safety Notice Medical Device Correction letter to their customers. The letter identified the affected device, the reason for the recall, and under what circumstances the problem occurs. Customers are to contact Philips as described if they are unsure whether their gamma cameras supports the modes. The firm will implement the AutoSPECT Pro software upgrade and contact customers to schedule the appointment. If further assistance or support is needed, customers are to contact their local Philips representative. |  | 126 | 126 Units | Worldwide Distribution -- US, including the states of CO, DE, FL, GA, IL, IN, KS, MO, NC, NV, NY, OH, OR, PA, TN &amp; TX; and, the countries of Australia, Austria, Belgium, Canada, China, Denmark, Egypt, France, Germany, Greece, Ireland, Israel, Italy, Lebanon, Malaysia, Saudi Arabia, Slovenia, South Africa, Switzerland, Turkey, &amp; United Kingdom. | Open, Classified |  |  | 196 | Computer | Computer
Z-1519-2016 | 73593 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Siemens SOMATOM Definition AS; Intended to produce cross-sectional images of the body. Computed tomography (CT) x-ray system. |  |  | March 02, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Software bugs in these SOMATOM systems could possibly cause scan aborts, rescans of patients with additional dose, delayed diagnosis and as worst case scenarios, could possibly cause the need of additional contrast medium, accidental x-radiation to the user or impede customer workflow. | Software design | Siemens initiated a Customer Advisory Notice on 03/07/2016 to inform customers about actions for bug-fixing the following systems: SOMATOM Definition AS, SOMATOM Definition Flash and SOMATOM Definition Edge, which could possibly cause scan aborts, res cans of patients with additional dose, delayed diagnosis and as worst case scenarios, could possibly cause the need of additional contrast medium. Siemens mailed certified as well as via email distribution a Customer Advisory Notice to their customers to inform them about corrective actions for bug-fixing the following systems running with software version VA44A_SP3a, VA44A_SP3b, VA44A_SP3c, correction of potential safety issues. The issues described will be corrected with software VA44_SP4, handled as update CT051/15/P. Following the installation of this update, the mentioned malfunctions will be corrected. |  | 1096 | 1096 | Nationwide Distribution-including DC, and Puerto Rico and the states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY. | Open, Classified |  |  | 1096 | Computer | Computer
Z-1507-2016 | 73505 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | VariAx Compression Plating System |  |  | October 10, 2013 | 2013 | Stryker Howmedica Osteonics Corp. | Stryker received two reports indicating that during screw insertion under a high lateral angle a bone screw went through a oblong hole. | Device Design | Stryker Orthopaedics sent notification letters/product accountability forms on October 10, 2013, via email. Customers are provided the product issue, hazards and risk mitigation along with actions needed. Customers should receive the affected products and complete the BRF and return to Stryker. Questions can be directed to Stryker. For further questions, please call (201) 831-5000. |  | 118 | 118 units | Nationwide Distribution | Terminated | April 21, 2016 | 924 | 118 | Computer | Not_Computer
Z-2245-2016 | 74620 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo.via, picture archiving and communications system software controlled. |  |  | June 23, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Incorrect values for the volume calculation. Software update VB30B via Update Instructions SY018/16/P to resolve software errors. | Software design | Siemens released and mailed to customers a notice with software update VB30B instructions via update instruction SY018/16/P for customers to follow until the update can be made. The letter also described the software issues involved; and asked customers to include the letter in the operator's manual until the update is performed. Customers are to ensure that all necessary personnel are aware of the situation. |  | 50 | 50 units | Nationwide Distribution -- AR, CA, CO, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, ND, NE, NJ, NM, NV, NY, OH, OK, PA, SC, TX, and UT. | Open, Classified |  |  | 50 | Computer | Computer
Z-2316-2016 | 74715 | No | OLO | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | Synergy Spine. The software application is sent in CD format with an IFU, wrapped in plastic with a label for shipping purposes. |  |  | July 18, 2016 | 2016 | Medtronic Navigation, Inc. | Medtronic Navigation, Inc. announces a voluntary field action for the Medtronic Navigation StealthStation Software applications affected by Neurologica BodyTom/CereTom floor-based scanners. | Software design | Medtronic sent an Urgent Medical Device Safety Alert to all affected customers on August 20, 2016. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete the attached confirmation form and return to RS.NavFCA@medtronic.com or fax it to Medtronic Technical Service at 651-367-7075. For questions contact Neurologica Technical Support at 1-888-564-8561 (US/Canada) or +1-978-564-8561 (International) or email support@neurologica.com |  | 27 | 27 in total | Worldwide Distribution - US Nationwide in the state of: NY, AZ, AR, NC, CA, OH, MI, IN, MA, TX, OH, OR, CO, MO, UT, IL. and the countries of: India, Australia, and Russian Federation. | Open, Classified |  |  | 27 | Computer | Computer
Z-2325-2016 | 74691 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | BK Medical ApS Ultrasound System Scanner bk2300 - |  |  | June 28, 2016 | 2016 | B-K Medical A/S | In the system (bk3500) software generic volume measurement (HxWxLxl) function the user may beunaware of the calculation factor (1) and assume the generic volume calculation factor is similar to that ofthe other (prostate, adenoma, testis, kidney, and bladder) volume calculation factors provided in thesoftware. | Software design | BK Medical sent an Important Corrective Action Notification letter dated June 27, 2016, to all affected customers. The letter explains the issue and asks customers to discontinue the use of the generic volume measurement and contact BK medical ApS Service for assistance. Customers with questions were instructed to call 800-645-4442. |  | 6 | 6 | Worldwide Distribution - US to TX, FL, and MA. Internationally to Australia | Open, Classified |  |  | 6 | Computer | Computer
Z-2149-2016 | 74585 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Optical Surface Monitoring System{OSMS, Varian Private Label} This is used on Varian's TrueBeam EDGE; |  |  | June 23, 2016 | 2016 | Varian Medical Systems, Inc. | Varian Medical Systems has received reports that unintended changes were made to planned couch parameters, specifically the couch rotation parameter, during patient set-up. This can occur when a user sets a different couch rotation from the plan couch rotation while performing patient alignment with Optical Surface Monitoring System [OSMS]. | Software design | The firm, Visionrt, sent an "URGENT MEDICAL DEVICE CORRECTION NOTIFICATION-URGENT FIELD SAFETY NOTIFICATION" letter dated June 24, 2016 to all affected customers. Letters describe the product, problem and actions to be taken. The customers were instructed as follows: DO NOT USE the "Move Couch" function if the couch rotation value set in the AlignRT device DOES NOT match the planned couch rotation for the treatment field; retain a copy of this document along with your most current product labeling; pass the letter to all those who need to be aware of it within their organization; always confirm that the patient is correctly aligned prior to commencing treatment; is not required to return their devices; promptly inform Vision RT if they believe that patient harm occurred, and complete Appendix 1- Customer Acknowledgment Response and return it via email to corrections@visionrt.com.Should you have any queries on this letter, please do not hesitate to contact Vision RT by telephone on +44 20 83464300 (866-778-2379 from the US) or as per http://www.visionrt.com/contact/details. |  | 189 | 166 OSMS in US, 23 OSMO - International, | Worldwide Distribution: US (nationwide) and countries of: Austria, Azerbaijan, Brunei Darussalam, Bulgaria, Denmark, Georgia, Germany, India, Italy, Malaysia, Morocco, Norway, Portugal, Switzerland, and United Kingdom. | Open, Classified |  |  | 211 | Computer | Computer
Z-1362-2016 | 73171 | No | LIT | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | USE (Ebony) PTA .014" rapid exchange (RX) Peripheral Dilatation Catheter. Distributed by US Endovascular, St. Louis, MO. |  |  | February 11, 2016 | 2016 | Us Endovascular | Product may kink during use rendering the product unusable. | Device Design | NATEC Medical ltd.sent an Urgent Medical Device Advisory Notice dated March 18, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Warnings:This device should be used only by physicians who are experienced and have a thorough understanding of the clinical and technical aspects of angioplasty.What action do you need to take:1. Always use a guiding catheter when contralateral access is needed.2. Always insert cautiously the PTA catheter inside the guiding catheter, and if any resistance is felt, the cause should be determined via fluoroscopy.3. For contralateral procedure, the USE014OTW has to be preferred to Ebony 014 Rx.NATEC Medical Ltd and our distributor US Endovascular are committed to keeping you and your healthcare professional informed of issues and solutions concerning our products and services. Should you have any questions regarding this or any other matter regarding your USE 014Rx PTA catheter, please do not hesitate to contact our distributor: US Endovascular at (314) 341-4930. |  | 843 | 843 catheters | US Distribution to the states of : AR, CA, FL, LA, MO, MI, MS, NY, TN, and TX. | Open, Classified |  |  | 843 | Computer | Not_Computer
Z-2405-2016 | 74724 | No | HQF | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) |  |  |  | July 15, 2016 | 2016 | Topcon Medical Laser Systems, Inc | Topcon discovered an anomaly with titration, only when EpM is turned on during the titration process.This event may occur if the user turns on EpM, turns off the Landmark (LM) function (with any EpM supported pattern) and then enters titration mode. | Device Design | The firm, PASCALS Synthesis, sent an "Urgent Field Safety notification" letter dated July 15, 2016 to all affected sites by traceable method. The letter described the product, problem and actions to be taken. The customers were instructed to Acknowledge the receipt of the field notice via an official letter; locate all systems with Endpoint Management software feature enabled; for all affected units- inform the physician, hospital or institution customers of the addendum and additional warnings issued as a result of this notice; obtain signed confirmation from the end-user regarding receipt of the addendum and additional warnings; send the signed confirmation from ALL end-users to TMLS, no later than October 15, 2016; and complete to the following record "Products lines affected" for all affected units and either fax to (408) 235-8259 or scan and email to tmlsinfo@topcon.com.Contact Customer Care if you have questions at: 1.408.235.8200 or email: tmlsinfo@topcon.com. |  | 205 | 205 units (21 in US, 184 OUS) | Worldwide Distribution: US (nationwide) and countries of: Canada, Mexico, Japan, Netherlands, Algeria, Australia, Brazil, Chile, China, France, Guatemala, Hong Kong, Italy, New Zealand, Pakistan, Philippines, Singapore, Taiwan, and United Kingdom. | Open, Classified |  |  | 205 | Computer | Computer
Z-2396-2016 | 74743 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Orthopedic manual surgical instrument |  |  | July 01, 2016 | 2016 | Stryker Howmedica Osteonics Corp. | Malfunctioning of the Exeter Rasp/trial introducer/extractor Handle. | Device Design | Stryker notified Branches/Agencies of the Urgent Medical Device Recall by e-mail on July 1, 2016 and an "Urgent Medical Device Recall Letter and Business Reply Form" was sent to branches, agents and hospital risk management via UPS (with return receipt) on July 5, 2016. The letter identified then affected product, problem and actions to be taken. For questions call (201) 831-6693. |  | 241 | 241 units | Worldwide Distribution - US Nationwide | Terminated | June 08, 2017 | 342 | 241 | Computer | Not_Computer
Z-2525-2016 | 74534 | No | JHX | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | StratusÂż CS STAT Fluorometric Analyzer-microprocessor-controlled instrument |  |  | June 22, 2016 | 2016 | Siemens Healthcare Diagnostics Inc | Software defect, where either an Above Assay Range or an inaccurate value could potentially be reported, in the Stratus CS Acute Care Diagnostics System regarding TestPak calibration and quality control (QC) status when a TestPak is due to expire within 48 hours. | Software design | Siemens Healthcare issued an Urgent Field Safety Notice by e-mail to all affected Siemens Healthcare Diagnostics Regional offices on June 22, 2016 both in the United States and Outside the US for communication with affected customers. This notice informs customers of the issue and provides mitigation instructions. Customers are to contact Siemens Customer Care Center or their local Siemens technical support representative at 800-405-6473. Siemens Healthcare is currently developing a software update to address this issue and will be providing new information as it becomes available |  | 977 | 977 units | Worldwide Distribution -- USA, U.A.E., Algeria, Armenia, Australia, Austria, Bangladesh, Belgium, Bosnia Herzeg., Botswana, Canada, Colombia, Croatia, Czech Republic, Denmark, Finland, France, Germany, Greece, India, Ireland, Italy, Japan, Libya, Lithuania, Malaysia, Mexico, Namibia, P.R. China, Poland, Portugal, Romania, Saudi Arabia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, U.A.E., and United Kingdom. | Open, Classified |  |  | 977 | Computer | Not_Computer
Z-2159-2016 | 74510 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge OfficePACS software. The firm name on the label is Merge Healthcare, Heartland, WI. |  |  | January 31, 2016 | 2016 | Merge Healthcare, Inc. | Potential data loss occurs as a result of product archiving not working properly. | Software design | The recalling firm issued a letter dated 1/27/2016 via e-mail on 1/31/2016 explaining the reason for recall and instructing the customer to provide a copy of the notification to all users. |  | 49 | 49 sites have the affected software | Distribution was made to medical facilities located in AL, AZ, CA, CO, CT, FL GA, ID, IL, KY, MA, MO, NE, NH, NJ, NY, NC, OK, PA, SC, TN, VA, and WA. There was no government/military/foreign distribution. | Open, Classified |  |  | 49 | Computer | Computer
Z-2386-2016 | 74757 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Integra External Fixation System Universal Wire Fixation Bolt - 17.5 mm Catalogue #12224218 ; Integra External Fixation System Universal Wire Fixation Bolt - 23 mm Catalogue #12224219 |  |  | July 20, 2016 | 2016 | Integra LifeSciences Corp. | Through the investigation of reported complaints Integra verified that there have been instances where the Universal Wire Fixation Bolts the Slotted Post broke during use at the surgical site. | Device Design | Integra sent an Urgent - Voluntary Medical Device Recall letter dated July 20, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Integra is asking you to take the following actions:1. Please check your IntegraÂż External Fixation System inventory to determine if you have the 17.5mm or 23mm UniversalWire Fixation Bolts and/or the large or small Slotted Posts.2. If you do have 17.5mm or 23mm Universal Wire Fixation Bolts and/or large or small Slotted Posts, stop using them immediately.3. Complete the attached "Acknowledgement and Return Form" and check the box: I do have affected products on the list and record the lot number.4. Or, complete the attached "Acknowledgement and Return Form" and check the box: I do not have affected products.5. Complete the other information as indicated on this form. Keep a copy of the form for your records.Return the completed "Acknowledgement and Return Form" by email or fax indicating your receipt and review of this notification.When your form is received, if you have noted you have affected products, an Integra Representative will contact you and provide you with directions to return the product, as well as input an order to replace the quantity you indicated on the form.Receipt of this form ensures that Integra has achieved a level of effectiveness in communicating this information. We recommend you also maintain a copy of this notification and signed copy of the acknowledgement form for your records. Regulatory agencies such as the FDA perform audits of field actions of this nature to verify that our customers have been notified and understand the nature of the field action.Should you have any questions regarding these instructions, please contact Sales Operations at 888-601-0203, option 2. We apologize for any inconvenience this may cause and thank you for your cooperation in this effort.Fo |  | 495 | 495 units | US Distribution to the states of : FL, VA, OK, CO, DC and IN. | Open, Classified |  |  | 606 | Computer | Not_Computer
Z-2379-2016 | 74773 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge PACS software. Product was distributed under the AMICAS label, Brighton, MA, but labeling was later changed to Merge Healthcare, Hartland, WI. |  |  | January 30, 2016 | 2016 | Merge Healthcare, Inc. | The patient name in the Halo title bar and the thumbnails do not match the name on displayed images. | Software design | The recalling firm issued a letter dated 1/27/2016 via e-mail on 1/30/2016 explaining the reason for recall and instructing the customer to provide a copy of the notification to all users. A second letter was issued to nonresponders dated 5/25/2016 |  | 93 | 93 sites have the affected version | The software was distributed to medical facilities nationwide. Government distribution was made to MN. Foreign distribution was made to Australia and to the United Kingdom. There was no military distribution. | Open, Classified |  |  | 93 | Computer | Not_Computer
Z-2380-2016 | 74131 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance 64 CT Model 728231; To produce cross-sectional images of the body. |  |  | April 29, 2016 | 2016 | Philips Medical Systems (Cleveland) Inc | Software issues found in software versions v4.1.3/4.1.4/4.1.5 in the Philips Brilliance 64 products that could affect the performance of the equipment. | Software design | The customers will be sent a letter with information about, what the problems are and under what circumstances they may occur, the actions customers can take to minimize the effect of the problem and the actions planned by Philips to correct the problem. To correct these issues, a Philips field service representative will install a software update on the affected systems. Reference field change order (FCO) 72800652. |  | 150 | 150 | Product was shipped to the following states: AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WV &amp; WY.Product was also shipped to the following countries: Argentina, Australia, Austria, Bangladesh, Belgium, Bolivia, Brazil, Cambodia, China, Colombia, Cuba, Estonia, Finland, France, Germany, India, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Korea, Latvia, Lebanon, Libya, Lithuania, Malaysia, Myanmar, Netherlands, New Zealand, Norway, Palestine, Panama, Philippines, Poland, Portugal, RÂżunion, Russian Federation, Senegal, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Turkey, United Arab Emirates, United Kingdom, Uzbekistan &amp; Vietnam. | Open, Classified |  |  | 1403 | Computer | Computer
Z-2371-2016 | 73750 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Ingenuity Core Computed Tomography X-ray systems is intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. |  |  | March 24, 2016 | 2016 | Philips Medical Systems (Cleveland) Inc | Software defect causing intermittently slow response of Host. | Software change control | On 3/24/2016 the firm mailed Customer Information letters to their customers. |  | 29 | 29 units | Product was sent to the following states: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WV &amp; WY.Product was also sent to the following countries: Algeria, Argentina, Australia, Austria, Bangladesh, Belgium, Bolivia, Brazil, Cambodia, Canada, China, Colombia, Cuba, Denmark, Djibouti, Egypt, Estonia, Finland, France, Germany, Ghana, Gibraltar, Hong Kong, Hungary, Iceland, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Korea, Latvia, Lebanon, Libya, Lithuania, Macedonia, Malaysia, Martinique, Mexico, Monaco, Mongolia, Mozambique, Myanmar, Netherlands, New Zealand, Norway, Oman, Palestine, Panama, Philippines, Poland, Portugal, RÂżunion, Russian Federation, Senegal, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, Uzbekistan &amp; Viet Nam. | Open, Classified |  |  | 635 | Computer | Computer
Z-2376-2016 | 74615 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Volker 2080, 2082, S 280 and S 282 series beds with lateral release button. |  |  | June 14, 2016 | 2016 | Hill-Rom, Inc. | The current construction of the side rail latch may not be able to support the full patient weight and may give way allowing the side rail to abruptly lower and the patient to fall, potentially resulting in harm. To date there has been one report of a related injury. | Device Design | Hill-Rom sent an URGENT: MEDICAL DEVICE CORRECTION letter dated June 14, 2016 to the affected US and Canadian customers via email and courier, international customers. The letter included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. The letter requests the number of beds in the consignees possession so that repair kits can be provided. If you have any questions regarding this safety notice, contact Hill-Rom Technical Support at 1-800-445-3720, your distributor, your Hill-Rom representative or contact us on www.hill-rom.com. |  | 24262 | 24262 | Worldwide Distribution - US Nationwide and the countries of: Canada, AUSTRALIA, CHINA, HONG KONG, THAILAND, POLAND, RUSSIAN FEDERATION, SLOVENIA, UNITED ARAB EMIRATES, AUSTRIA, BELGIUM, DENMARK, FINLAND, FRANCE, GERMANY, ICELAND, IRELAND, ISRAEL, ITALY, LUXEMBOURG, NETHERLANDS, NORWAY,SPAIN, SWITZERLAND, and UNITED KINGDOM. | Terminated | May 15, 2017 | 335 | 24262 | Computer | Not_Computer
Z-2151-2016 | 74556 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | NovaPACS  versions 7.4, 7.5, 7.6, and 8.0, Diagnostic Viewer. Distributed by Novarad Corporation. Picture archiving and communication system (PACS) |  |  | May 09, 2013 | 2013 | Novarad Corporation | Potential for images to be flipped while streaming, which could incorrectly display image orientation markers. | Software design | Consignees were notified on 05/09/13 via letter and email to upgrade software and to turn off streaming function. |  | 2695 | 2695 | Worldwide distribution including US nationwide. No Canadian distribution. | Open, Classified |  |  | 2695 | Computer | Computer
Z-2377-2016 | 74634 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker Orthopaedics Modular Handle Triathlon Instruments, Non-Sterile, REF 6541-4-808, Lot 123, NDC M546654148080S |  |  | June 23, 2016 | 2016 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedic is voluntarily recalling the Triathlon Modular Handle because analysis revealed that the press fit specifications between the dowel pin and the mating hole in the shaft were not being achieved potentially resulting in disassociation from the instrument. | Device Design | Stryker sent an Urgent Product Recall letter dated June 24, 2016 to Branches/Agencies by e-mail and Business Reply Forms was sent to Branches/Agents/Hospital Risk Management via UPS (with return receipt) on June 24, 2016. The letter identified the affected product, problem and actions to be taken. For questions call (201) 831-6693. |  | 12469 | 12,469 units | Worldwide Distribution - US Nationwide | Open, Classified |  |  | 12469 | Computer | Not_Computer
Z-2363-2016 | 74776 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge PACS software. The firm name on the label is Merge Healthcare, Inc., Hartland, WI. |  |  | February 05, 2016 | 2016 | Merge Healthcare, Inc. | When measuring a lesion on an unmagnified mammography image then performing the same measurement on an image magnified by the Hologic imager (different image), the measurements are not the same. | Software design | The recalling firm's customer support notified customers via phone call beginning 2/5/2016 to inform them of the issue. Customers were informed that a software fix that corrects the issue was available. Customers were asked not to use the measurements on mammo images until upgraded. If customers want an upgrade, then someone would contact them to schedule the upgrade. |  | 49 | 49 sites potentially have the affected versions | Worldwide Distribution -- USA and Australia. | Open, Classified |  |  | 49 | Computer | Not_Computer
Z-1349-2016 | 73475 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | A Picture Archiving and Communication System (PACS) intended to display, process, read, report, communicate, distribute, store and archive digital medical images, including digital mammography images. |  |  | January 18, 2016 | 2016 | Siemens Medical Solutions USA, Inc | To inform users about the possible incorrect values for Distance Measurements when using certain modalities in combination with syngo Imaging. | Software design | Siemens mailed a Customer Advisory Notice to their customers instructing them to check their modalities for the usage of the mentioned DICOM attributes with different values. They recommended upgrading the workplace(s) to syngo Imaging Studio version VB36D if they have such modalities and it cannot be ensured by the modality provider that only one attribute is filled or both attributes are filled with the same value. Customers were asked to contact their local service organization to clarify the compatibility of their syngo Imaging server version with workplace version VB36D. |  | 49 | 49 imaging systems | US (nationwide) including the states of AL, CA, CO, CT, FL, IA, IN, KS, LA, MA, MD, MI, MN, MO, MS, NC, NE, NJ, NY, OK, OR, PA, TN, TX, and WI. | Terminated | August 03, 2017 | 563 | 49 | Computer | Computer
Z-1270-2016 | 73468 | No | HQF | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | Ellex Medical Fixed eye Safety filters Model Integre S LP5532. |  |  | February 01, 2016 | 2016 | Laserex Systems Inc. | The Firm discovered that there is a design error which is traceable back to 2012 when a design change on the safety filter optic adaptor was implemented. | Device Design | Ellex Medical Pty Ltd. Planned Action to bring into Compliance:1.Field Corrective Action replacing all affected safety filters of all units manufactured between July 2012 and February 2016, 2.All repairs will be made at no cost to the purchaser; and3.The corrections will be completed by March 30, 2017.CDRH approves the CAP subject to the following conditions:Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR 1003.21 and1003.22. For further questions please call (800) 824-7444. |  | 74 | 72 units plus 2 spare units | Worldwide Distribution - US Distribution | Open, Classified |  |  | 74 | Computer | Not_Computer
Z-2350-2016 | 74705 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | CURVED-TIP STAPLER 30,IS4000, model number 470530; |  |  | July 14, 2016 | 2016 | Intuitive Surgical, Inc. | An increase in the number of da Vinci EndoWrist Stapler complaints where the customer experienced aStapler Firing Failed message as well as review of da Vinci Xi Surgical System logs found an increased trend | Software design | Field Safety Notifications - Urgent Medical Device Corrections and Removal letters were sent on July 15, 2016 (dated July 14, 2016 advised users of the Stapler Firing Failed message and actions to take should the message be received. |  | 2065 | 2065 total | US, including Puerto Rico and Australia, Austria, Belgium, Chile, Cyprus, Czech Republic, Denmark, Finland, France, Germany, Hong Kong, Ireland, Italy, Japan, Monaco, Netherlands, Norway, Portugal, , Singapore, South Korea, Spain, Sweden, Switzerland, Taiwan, Turkey, United Kingdom, Venezuela. | Terminated | January 03, 2017 | 173 | 2065 | Computer | Not_Computer
Z-2352-2016 | 74764 | No | CHP | Estradiol test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS Immunodiagnostic Products Estradiol Reagent Pack, Unique Device Identifier No. 10758750005017, REF/Catalog Number/Product Code 855 2630; IVD -- Ortho-Clinical Diagnostics, Pencoed, Bridgend CF35 5PZ, UK --- The VITROS Estradiol test is performed using the VITROS Estradiol Reagent Pack and Calibrators on VITROS ECi/ECiQ Immunodiagnostic Systems, VITROS 3600 Immunodiagnostic Systems and VITROS 5600 Integrated Systems. |  |  | July 14, 2016 | 2016 | ORTHO-CLINICAL DIAGNOSTICS | Ortho Clinical Diagnostics has issued a recall of their VITROS immunodiagnostics Estradiol Reagent assay. The firm was made aware of a potential cross-reactivity between the drug Fulvestrant and theVITROS Immunodiagnostic Estradiol Reagent assay. The firm conducted an internal investigation and was able to confirm biased positive (ranging from 1303-4046%) on Estradiol results, obtained from postmenopausal females taking Fulvestrant. | Device Design | On 7/14/2016, URGENT PRODUCT CORRECTION NOTIFICATION Customer Letter (Ref. CL2016-141, dated 7/14/2016) was sent via FedEx overnight courier and/ or ORTHO PLUS e-Communications and/ or US Postal Service Priority Mail (for PO Boxes only) to all customers who received product since 6/1/2015 to inform them that Fulvestrant is an interferent of the Estradiol Reagent Pack and causes positively biased results. Discuss any concerns you may have regarding previously reported estradiol results with your Laboratory Medical Director to determine the appropriate course of action. The results from this or any other diagnostic test should be used and interpreted only within the context of the overall clinical picture. For questions, contact Customer Technical Services at 1-800-421-3311. Foreign affiliates were informed of the issue via e-mail on 7/14/2016 and instructed to notify their consignees of the issue. |  | 7597 | Domestic: 1620; Foreign: 7597 | Nationwide, Puerto Rico, and foreign distribution to the following countries: Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Poland, Singapore, Spain, Venezuela. | Open, Classified |  |  | 7597 | Computer | Not_Computer
Z-2346-2016 | 74402 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | MX 16-slice whole body computed tomography X-ray system. Imaging diagnostic tool. |  |  | April 20, 2016 | 2016 | Philips Medical Systems (Cleveland) Inc | The following issues have been found in MX 16-slice systems with software version 1.1.4.21426:1. During the filming operation on MX16-slice console software, the clipboard used for copying and pasting images is not cleared between patients. If the operator fails to copy the current patients images before pasting, a previous patients images may be present in the clipboard and be copied onto the film of the current patient. 2. During the Bolus Tracking scan, if the Auto Voice in Tracker scan is enabled, the Tracker scan will be aborted unexpectedly and the diagnostic scan after the tracker scan needs to be manually started.3. If the scan protocol with SAS (Spiral Auto Start) function is selected to plan the scan, the SAS option may not be displayed on the Contrast tab of the scan protocol parameter area or the SAS option is displayed but not enabled as pre-configured in the protocol. This issue only occurs on the first helical scan after system startup that applies SAS function | Software design | On April 20, 2016 the firm sent Field Safety Notices to their customers. Philips will release a Field Change Order to upgrade the software free of charge. A Philips Field Engineer will contact their customers to schedule this correction. |  | 899 | 899 units | Products were sold in the following states: AL, AR, CA, FL, GA, IL, IN, MS, NC, ND, NJ, NV, NY, OH, PA, SC, TN &amp; VA.Products were also shipped to the following countries: Albania, ALGERIA, Argentina, Australia, Austria, Bangladesh, Belgium, Brazil, Bulgaria, Burundi, Cape Verde, Chile, China, Colombia, Congo, Democratic, Costa Rica, Cyprus, Dominican Republic, Ecuador, Egypt, France, Germany, Greece, Guatemala, Hungary, India, INDONESIA, Ireland, Italy, Japan, Jordan, Latvia, Lebanon, Libyan Arab Jamahiri, Lithuania, Malawi, Mali, Mexico, Morocco, Myanmar, Namibia, Nepal, Netherlands, Nigeria, Oman, Palestine, Paraguay, Philippines, Poland, Romania, Russia, Senegal, Slovakia, South Africa, South Korea, Spain, Sri Lanka, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Uganda, United Kingdom, Venezuela, Viet Nam &amp; Zambia. | Open, Classified |  |  | 979 | Computer | Computer
Z-1556-2016 | 73131 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Spirit One A-C Powered Hospital Bed |  |  | July 02, 2015 | 2015 | Stryker Medical Division of Stryker Corporation | Inaccurate scale systems | Device Design | Stryker Corp. will initiate the recall by distributing letters on 07/21/2015 via FEDEX. Customers are asked to do the following: 1. Locate the units listed on the attached business reply form. 2. Remove these units from service. (Note: The original mattress purchased with the bed is not within scope of this action. The mattress will be compatible with your replacement option.) 3. Your Stryker Sales Representative will contact your facility to coordinate removal of the recalled beds and present replacement options. 4. Return the enclosedbusiness reply form to your Stryker Sales Representative to confirm receipt of this notification or fax (269)488 -8691 or email productfieldaction@stryker.com 5. If you have loaned or sold any of the beds listed in this letter, please forward a copy of this notice to the new users and advise us of their new location in the space provided on the business reply form. 6. If you have disposed of any of the beds and they are no longer in use, please advise us of their obsolescence by providing us with their serial number in the space provided on the business reply form. *For additional information, please refer to the products operations and/or maintenance manuals. Adverse reactions or quality problems experienced with the use of this product may be reported tothe FDA's MedWatch Adverse Event Reporting program either online, by regular mail or by fax. If there are any questions or concerns, please contact Stryker Customer Service (1-800-327-0770) during normal business hours are Monday-Friday 8 a.m-6 p.m. (EST). |  | 143 | 143 | Distributed in the states of AL, AZ, FL, GA, HI, IA, IL, KS, LA, MI, MN, MO, NE, NJ, OH, PA, SC, TN, and TX, and the countries of Canada and UK. | Terminated | March 24, 2017 | 631 | 143 | Computer | Not_Computer
Z-2215-2016 | 74469 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Oxylog 3000, Disposable Pediatric Patient Circuit, Catalog Number: 5704964 |  |  | May 31, 2016 | 2016 | Draeger Medical, Inc. | Isolated episodes of leakage at the check valve were found | Device Design | Drager sent an Urgent Medical Device Recall letter dated May 31, 2016 and Customer Reply and Order Card to all customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to inspect stock and dispose of affected products. For questions regarding this letter call 1-800-543-5047 (press 1 at the prompt, then 2, then 32349). For questions regarding the operation and/or servicing of your DrÂżger device/accessory contact DrÂżgerService Technical Support at 1-800-543-5047 (press 4 at the prompt). |  | 42922 | 42,922 units worldwide (US distribution 1,530) | US Nationwide Distribution including states of: AK, CA, FL, IA, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NH, NY,OH,OK,OR,PA,TN,TX,UT,VA,WA, and WI. | Open, Classified |  |  | 42922 | Computer | Not_Computer
Z-1351-2016 | 73338 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance 64:Computed Tomography X-ray Systems, Intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes |  |  | February 05, 2016 | 2016 | Philips Medical Systems (Cleveland) Inc | The firm was notified of a software error in which the system may not automatically send all image/data series to remote devices. | Software design | Philips Medical Systems sent a Customer Information dated February 5, 2016, to all affected customers via certified mail. The letter identified the product, the problem, and the action to be taken by the customer. To correct the issue Philips will release field change orders (FCO) 72800652 and 72800644 to install a software update on the affected systems.Customers with questions were instructed to contact their local Philips representative or local Philips Healthcare office. For North America and Canada, contact the Customer CareSolutions Center (1-800-722-9377, Option 5: Enter Site ID or follow the prompts).For questions regarding this recall call 440-483-2997. |  | 285 | 285 units | Worldwide Distribution - US to the District of Columbia: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA &amp; WV.Product was also shipped to the following countries: Afghanistan, Australia, Austria, Bangladesh, Belgium, Brazil, Canada, China, Cuba, Denmark, Egypt, Ethiopia, Finland, France, Germany, Ghana, Gibraltar, Greece, Hungary, Iceland, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kyrgyzstan, Lebanon, Lithuania, Macedonia, Malawi, Mexico, Myanmar, Netherlands, New Zealand, Norway, Oman, Palestinian, Philippines, Poland, Republic of Korea, RÂżunion, Romania, Russian Federation, Singapore, Spain, Sweden, Switzerland, Taiwan, Thailand, Ukraine, United Arab Emirates, United Kingdom, Uzbekistan &amp; Viet Nam. | Open, Classified |  |  | 1002 | Computer | Computer
Z-2341-2016 | 74577 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Merge HEMO software. |  |  | July 02, 2012 | 2012 | Merge Healthcare, Inc. | The International Normalized Ratio (INR) value displayed in the study report does not match the value that is imported on the pre-procedure labs screen from HL7 | Software design | The firm issued an "Important Product Information" bulletin to customers dated 7/2/2012 via e-mail letting them know there was an upgrade available. On 1/30/2016, the recalling firm initiated phone calls to their customers letting them know there was an upgrade available and if they declined it, they would like a written response. |  | 37 | 37 | The medical facilities who currently have the affected versions are located nationwide. There was government distribution but no military or foreign distribution. | Terminated | June 26, 2017 | 1820 | 37 | Computer | Computer
Z-1555-2016 | 71667 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Spirit One A-C Powered Hospital Bed |  |  | July 02, 2015 | 2015 | Stryker Medical Division of Stryker Corporation | Customer complaints associated with faulty brake system | Device Design | Stryker Corp.sent an Urgent Medical Device Recall letter on July 21, 2015, via FedEx to all affected customers. The letter will identified the product the problem and the action needed to be taken by the customer. Customers are instructed the following actions: 1.Locate the units listed on the attached business reply form.2.Remove these units from service. (Note: The original mattress purchased with the bed is not within scope of this action. The mattress will be compatible with your replacement option.)3.Your Stryker Sales Representative will contact your facility to coordinate removal of the recalled beds and present replacement options.4.Return the enclosed business reply form to your Stryker Sales Representative to confirm receipt of this notification or fax (269)488-8691 or email productfieldaction@stryker.com 5.If you have loaned or sold any of the beds listed in this letter, please forward a copy of this notice to the new users and advise us of their new location in the space provided on the business reply form.6.If you have disposed of any of the beds and they are no longer in use, please advise us of their obsolescence by providing us with their serial number in the space provided on the business reply form.*For additional information, please refer to the products operations and/or maintenance manuals. Adverse reactions or quality problems experienced with the use of this product may be reported to the FDA's MedWatch Adverse Event Reporting program either online, by regular mail or by fax.If there are any questions or concerns, please contact Stryker Customer Service (1-800-327-0770) during normal business hours are Monday-Friday 8 a.m-6 p.m. (EST). |  | 143 | 143 total | US Distribution to the states of : AL, AZ, FL, GA, HI, IA, IL, KS, LA, MI, MN, MO, NE, NJ, OH, PA, SC, TN and TX., and Internationally to the countries of : Canada and UK. | Terminated | March 24, 2017 | 631 | 143 | Computer | Not_Computer
Z-1361-2016 | 73410 | No | GXD | Radiofrequency lesion generator. | Neurology | Neurology | 510(k) | Neurotherm NT2000iX RF Generator |  |  | October 09, 2013 | 2013 | Neurotherm, Inc. | NT2000iX software shipped with the international setting turned on. The international settings include access to Corodotomy, Bi-Polar, and No Temperature modes were not cleared in the United States for use. | Software Manufacturing/Software Deployment | Neurotherm notified consignees via telephone and guided through the process to update the software manually. Thecalls were initiated by Neurotherm representatives on October 09,2013. |  | 25 | 25 | CA, CO, IA, ME, NJ, NY, SC, TX, WI | Terminated | April 14, 2016 | 918 | 25 | Computer | Computer
Z-2118-2016 | 74427 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Ingenuity Core 128-Computed Tomography X-ray system |  |  | April 01, 2016 | 2016 | Philips Medical Systems (Cleveland) Inc | Philips Healthcare received reports from the field that certain Ingenuity Core128 systems running software version 3.5.4 exhibited intermittent swirl-like ring artifacts that may appear on reconstructed images. A patient rescan may be required if the images cannot be used for interpretation due to the swirl-like artifact. | Software design (manufacturing process) | Philips Medical Systems (Cleveland) Inc. sent an Urgent Field Safety Notice letter to affected customers , via certified letter with receipts returned for tracking purposes. The letter identified the affected product, problem and actions to be taken. Philips Healthcare is implementing software update version 3.5.5 to correct the problem. For questions contact your local Philips representative or local Philips Healthcare office. |  | 19 | 19 | US Nationwide Distribution in the states of CA, FL, ID, KS, MI, NY,OH,TX | Open, Classified |  |  | 262 | Computer | Computer
Z-2734-2016 | 74886 | [6 empty] | 2 over 2 Elevating Monitor Suspension - Omega Medical Imaging, LLC Elevating Monitor Suspension Systems. |  |  | June 21, 2016 | 2016 | Omega Medical Imaging, Inc. | Elevating Monitor Suspension separated from the lifting column bracket | Device Design | Consignees were notified by a Field Safety Advisory Notice - FSAN 16-0001 of this immediate Recall on 6/21/2016. |  | 26 | 26 units | CO, MN, GA, NC, IN, OH, KY, IL, NY, AL, MO, LA, PA, and SD Denmark | Open, Classified |  |  | 26 | Computer | Not_Computer
Z-2736-2016 | 74894 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | BrightView XCT Model 882482, gamma camera for Single Photon Emission Computed Tomography (SPECT) |  |  | May 25, 2016 | 2016 | Philips Medical Systems (Cleveland) Inc | The firm became aware of an issue where during a non-Auto Body Contouring (ABC), non-circular, rel-180 scan, the detector may come in contact with the patient. The issue was caused by a software defect in the workflow setup script. | Software design | The firm is informing the consignee of this issue through Customer Information Letter. |  | 424 | 424 units | Worldwide distribution. US nationwide, Algeria, Argentina, Armenia, Australia, Austria, Belgium, Brazil, Czech Republic, Denmark, Finland France, Gabon, Germany, Greece, India, Indonesia, Ireland, Japan, Korea, Republic of Korea, Republic of Lebanon, Lithuania, Malaysia, Netherlands, New Zealand, Norway, Oman, Panama, Philippines, Poland, Portugal, Russian Federation, Saudi Arabia, Slovakia, Spain, Sweden, Switzerland, Taiwan, Turkey, United Arab Emirates, United Kingdom, Viet Nam. | Open, Classified |  |  | 544 | Computer | Computer
Z-0093-2017 | 75163 | No | ORZ | Eyelid thermal pulsation system. | Ophthalmic | Ophthalmic | 510(k) | LipiFlow thermal Pulsation System, consisting of Console and Activator, Model LFD-1100 |  |  | September 13, 2016 | 2016 | Tearscience, Inc | Complaints related to an error condition occurring during patient treatment where the FipiFlow System automatically stopped treatment after detecting a problem with the Activator. | Device Design | The firm, TearScience, sent a "Medical Device Recall" letter dated September 9, 2016 to its consignees/customers on 9/13/2016. The letter described the product, problem and actions to be taken. The customers were instructed to complete and return the enclosed Acknowledgement and Receipt form by fax to (919) 882-9603, email to: Customerservice@tearscience.com or mail to TearScience, ATTN: CUSTOMER SERVICE RCL001, 5151 McCrimmon Parkway, Ste 250, Morrisville, NC 27560; to discontinue use of the product; to remove all a product with lot #201615220020 from your inventory; and call TearScience Customer Service at 919-459-4891 to return the product for above mentioned lot# and request new product for replacement.For any questions regarding this letter, please call (919) 459-4891 Monday through Friday, 8:30AM to 5:30PM, Easter Time. |  | 12000 | 1,200 Activators (120 boxes of 10 units) | US Distribution including states of: AK, AR, CA, CO, FL, GA, ID, IL, IN, LA, MD, MA, MN, MO, NV, NM, NY, NC, OH, OK, OR, PA, SC, TN, TX, VA , WA, WI and Puerto Rico. | Open, Classified |  |  | 12000 | Computer | Computer
Z-1992-2016 | 73692 | No | JAI | Powered radiation therapy patient support assembly. | Radiology | Radiology | 510(k) | Protura Robotic Couch System Software, versions 1.6.1 (Part Number MT6XSM1.6.1), 1.7.0 (Part Number MT6XSM1.7.0), and 1.7.1 (Part Number MT6XSM1.7.1), CIVCO, Orange City, IA 51041 |  |  | March 31, 2016 | 2016 | Med Tec Inc | When using the recalled product in split shift mode, the non-sub-millimeter positioning has been found to be inconsistent with the setting parameters. | Software design | Civco Medical Solutions sent a "Urgent Field Correction Notice" dated March 31, 2016, to all affected customers. The letter identified the product the problem and the action need to be taken by the customer.The software's other configuration, "All 6 Degrees Protura" does not have the same translational issues. The firm send notification letters on February 23, 2016. In this letter, the Product Manager verified each site's software is configured to this operational mode which operates without the identified problem. In a later and separate notification on March 31, 2016, the firm informed customers that a new software version will be installed to address the problem. No timeframe was given in the letter.For further questions, please call ( 712) 737-8688. |  | 9 | 9 | Worldwide Distribution - US Distribution including the countries of Switzerland, Germany.. | Open, Classified |  |  | 9 | Computer | Computer
Z-0004-2017 | 75181 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Skintact Electrodes for Defibrillation, DF29N. |  |  | September 01, 2016 | 2016 | Leonhard Lang Medizintechnik GmbH | There is a risk that defibrillation electrodes model DF29N will be connected with the defibrillator Welch Allyn AED 10 only with delay or not at all. This may cause a situation in which a patient, who is in a life threatening condition and requires adefibrillation shock, cannot be treated in good time. | Device Design | Consignees was sent via e-mail a Leonhard Lang "Important Safety Notice" dated September 1, 2016. The letter described the product being recalled, Description of the defect, Actions and time frame of the recall, &amp; Compensation for the recalled electrodes. Advised consignees to inform all users within their organizations of the recall; and, secure and destroy the unused electrodes in their inventory. Confirmation of destruction can be done by completing and returning the "Confirmation of Destruction / Consumption" form by October 14, 2016. If the product was further distributed, they were to forward a copy of the Safety Notice. For questions contact sales staff. |  | 11110 | 11,110 ( US 8,040; OUS 3,070) | Worldwide Distribution - US, to the state of Florida; and, the countries of South Africa, Germany, Slovenia, France, United Arab Emirates, Great Britain, Italy, Israel, Serbia, Thailand, Thailand, South Africa, Lebanon, and Poland. | Open, Classified |  |  | 11110 | Computer | Computer
Z-0079-2017 | 74750 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Radiation Therapy Treatment Planning System, Model 5.0 |  |  | May 30, 2016 | 2016 | RAYSEARCH LABORATORIES AB | For a treatment plan consisting of multiple beam sets, the table for ROI plan dose statistics in the report may show the statistics for a beam set dose. The error can only be triggered when using a report template where statistics for a beam set dose is included as the last dose statistics prior to the plan dose statistics. | Software design | On May 19th, 2016, RaySearch Laboratories distributed Field Safety Notice, Medical Device Correction notices and reply forms to their customers via email. Actions to be taken by the customer include: 1) Inspect all report templates that are intended to be used for plans with multiple beam sets. 2) Make sure that POI Dose statistics (Plan) is included immediately before ROI Dose statistics [Plan dose] in the Report content area. 3) Please educate planning staff and all users about this workaround. 4) Inspect your product and identify all installed units with the above software version number. Finally, complete and return the reply form via email to: freddie.cardel@raysearchlabs.com or by fax to: 888-501-7195 to confirm that you have read and understand the information. Customers with questions can contact David Hedfors via email: david.hedfors@raysearchlabs.com |  | 491 | 491 units (187 domestically &amp; 304 internationally) | AZ, CA, FL, MT, NC NY, TX, OH &amp; WA | Open, Classified |  |  | 491 | Computer | Computer
Z-1599-2016 | 73668 | No | IXW | Automatic radiographic film processor. | Radiology | Radiology | 510(K) Exempt | Smart CR Digital Computed Radiography System; Model Number: CR-IR346RU |  |  | April 01, 2016 | 2016 | Fujifilm Medical Systems U.S.A., Inc. | Reports of failure of the device's power supply. | Device Design | Fujifilm sent an Medical Device Correction letter dated March 22, 2016, via FedEx to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Instructions for turning the unit off when not in use as an interim action. FMSU will conduct a field replacement of the power supply in all installed units.For further questions, please call (888) 385-4633. |  | 2062 | 2,062 devices | Nationwide Distribution. | Open, Classified |  |  | 2062 | Computer | Computer
Z-0888-2016 | 73307 | No | OTE |  |  | Radiology | PMA | Mammomat Inspiration with Tomosynthesis functionality |  |  | February 12, 2016 | 2016 | Siemens Medical Solutions USA, Inc | A potential system fatal error may occur during tomosynthesis reconstruction with large breasts that cover nearly the entire detector surface and with a breast thickness larger than 90 mm. The data to be processed for the tomosynthesis application can exceed the capacity of the graphical processing unit (GPU). If the fatal error message occurs, reconstruction of the data will be aborted. The syste | Software design | Siemens sent a Field Safety Notice dated February 11, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.How will the issue finally be resolved?Siemens is releasing a field modification to resolve this issue as soon as possible. The upgrade software (VB30P) will be available to the affected customers free of charge. A Siemens service engineer will contact you to schedule implementation of the modification for your system. According to regulations set forth in 21 CFR 7.49(d), we advise that you follow the instructions in this notice and where necessary request that you promptly notify and instruct all the staff at your organization who have to be aware of this problem.If the device has been sold and therefore no longer in your possession, please forward this notice to the new owner. We would also request that you inform us of the identity of the device's new owner where possible. Please observe this safety notice and comply with the corresponding measure until the update has been fully completed.We thank you for your cooperation. For further questions, please call (610) 448-6478. |  | 76 | 76 units | Nationwide Distribution including Puerto Rico and to the states of : AR, CA, CO, FL, GA, IA, IL, IN, KS, MA, MD, MI, MO, NC, NE, NH, NJ, NY, OH, PA, SC, TN, TX, VA and WI. | Open, Classified |  |  | 76 | Computer | Computer
Z-0077-2017 | 75296 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Head Frame - Model number HHF -ALL Head Frame Posts (PN1008016 ) and Post Kits (PN1008016). |  |  | September 13, 2016 | 2016 | Varian Medical Systems, Inc. | Varian Medical Systems has received a report that a user was able to easily rotate the head frame posts when attached to the mounting cam on the metal head ring. There was no report of serious injury due to this issue | Device Design | Varian Medical Systems sent an Urgent Medical Device Removal letter dated September 14, 2016, to all affected sites. Each site was also contacted by telephone to advice users of the issue and request of return of the head frame posts on September 14 and 15th. Calls enabled Varian to prioritize replacement parts for those with patients scheduled in the immediate future.Recommended User Actions:1.CEASE USE of the Varian Head Frame Assembly and QUARANTINE the Varian Head Frame Assembly and components.2.Collect, Clean, Disinfect all Head Frame Posts [PN1008016] if used previously.3.Contact Varian 1.888.827.4265 for Return Material Authorization number, and return all Head Frame Posts in accordance with instructions provided with the enclosed Recall Return Material Form.We sincerely apologize for any inconvenience and thank you in advance for your cooperation. If you require further clarification, please feel free to contact your local Varian Medical Systems Customer Support District or Regional Manager. This notice has been provided to the appropriate Regulatory Authorities. For further questions, please call (650) 424-6833. |  | 75 | 75 devices | US Distribution and Internationally to 1 site in Canada, India and Morocco | Open, Classified |  |  | 75 | Computer | Not_Computer
Z-0076-2017 | 75303 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco RTP System |  |  | September 29, 2016 | 2016 | Elekta, Inc. | When creating 3D plans using either MU or Dose weighting modes, if the user changes the Physician's Intent Rx Dose and/or the number of fractions, and then modifies the wedge angle, the MU value is scaled incorrectly. | Device Design | Elekta sent an Important Field Safety Notice 382-01-MON-005 was sent to all affected customers on 9/29/2016. The notice informs users of the specific product and version numbers affected by the issue,and any work arounds that can be used to avoid the issue. The notice included an acknowledgment form which is to be returned.For questions regarding this recall call 314-993-0003 or 800-878-4287. |  | 365 | 365 | Worldwide Distribution - US to CA, FL, KY, MD, MO, MT, NJ, NC, PA, WA, WI. Internationally to Algeria. Antiqua &amp; Barbuda, Argentina, Australia, Austria, Belarus, Botswana, Brazil, Bulgaria, China, Columbia, Croatia, Czech Republic, Ecuador Estonia, France, Germany, India, Iran, Israel, Italy, Japan, Malaysia, Mexico, Morocco, Namibia, Netherlands, New Zealand, Norway, Panama, Peru, Philippines, Portugal, Romania, Russia, Slovakia, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Thailand, Turkey, United Kingdom, Ukraine, Venezuela, and Zimbabwe. | Open, Classified |  |  | 365 | Computer | Computer
Z-0827-2016 | 73241 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Evita Infinity V500 Ventilator; Babylog Infinity VN500 Ventilator. |  |  | February 04, 2016 | 2016 | Draeger Medical, Inc. | The battery capacity of the optional PS500 power supply unit for the Infinity (ACS) Workstation Critical and Neonatal Care was rapidly and unexpectedly reduced, despite using the most recent power supply firmware version 1.50 released in November 2015 (recall Z-0436-2016). The batteries were prematurely discharged, even though an adequate battery charge status was displayed on the device. | Software design | Draeger Medical, Inc. will send an Urgent Medical Device Recall Letter to consignees in a method that confirms delivery. For consignees that have received new devices with firmware version 1.50 (127 devices) or through device updates (644 devices; reference Recall Z-0436-2016), those devices will be downgraded to firmware version 1.49 and the batteries replaced. Batteries will continue to be replaced for all consignees every six months until a final solution is available. |  | 2553 | 2553 | US Nationwide Distribution including Puerto Rico. | Open, Classified |  |  | 2553 | Computer | Computer
Z-0073-2017 | 75192 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Fujifilm Synapse PACS, Software versions: 4.0.xxx, 4.1.xxx, version 4.2.xxx, version 4.3.xxx |  |  | September 15, 2016 | 2016 | Fujifilm Medical Systems U.S.A., Inc. | Image data for a patients image may not be correct | Software design | The firm, FMSU, sent a "Medical Device Correction" letter dated September 15, 2016 to its customers. The letter describes the product, problem and actions to be taken. The firm will install a patch that corrects the issue (Synapse IPSS v1.3.1.0 Patch).Customers with Version 4.0 may upgrade to version 4.4 (will not require the patch) or upgrade to 4.1, 4.2, or 4.3 and install the patch. FMSU will install the patch on the server remotely if your site uses Active Monitoring. Your notification that the patch was installed will be your service ticket. For sites not using Active Monitoring, FMSU will schedule a time to install the patch on your server.Questions or concerns, contact the FUJIFILM Customer Support Center at 1-888-FUJIMED (1-800-385-4633) Option #1. International callers can access the FUJIFILM Customer Support Center by calling 1-203-602-3580. |  | 251 | 251 units | US Nationwide Distribution. | Open, Classified |  |  | 251 | Computer | Computer
Z-0070-2017 | 75279 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Medtronic Navigation O-arm O2 Surgical Imaging System Product Catalog Number: BI-700-02000 |  |  | September 22, 2016 | 2016 | Medtronic Navigation, Inc.-Littleton | O-arm O2 Surgical Imaging System Spatial calibration may be erroneous in Stealth Station navigated images | Software design | Medtronic Navigation contacted customers via telephone beginning September 23, 2016. Customers were informed of the affected product, problem and actions to be taken. Medtronic requests that users immediately discontinue the use of the O-Arm O2 Surgical Imaging system for automatic registration in Stealth Station navigated surgical procedures until further notice.Medtronic Service personnel have been scheduled to begin the necessary inspections and re-calibration activities on customer systems beginning on September 23, 2016. |  | 14 | 14 units | Worldwide Distribution - US Nationwide in the states of CO, KA, NE, MI, MO, OH, WA and country of Switzerland | Terminated | May 12, 2017 | 232 | 14 | Computer | Computer
Z-2003-2016 | 74206 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Dome Centric component (Reference 01.04227.005) from the Anatomical Shoulder (AS) System 2.0. |  |  | May 17, 2016 | 2016 | Zimmer Gmbh | In some cases it has been difficult or not possible to disassemble the adjusted Dome Centric from the AS Humeral Rasp after initial positioning of the AS Humeral Trial Head which resulted in the whole construct of the Dome Centric with the still assembled AS Humeral Rasp taken out of the humeral canal. This could cause a delay in surgery time or the surgeon could decide to close the wound without finishing the surgery, increasing the infection risk or second use of anesthesia necessary. | Device Design | Zimmer Biomet sent an URGENT MEDICAL DEVICE RECALL  LOT SPECIFIC notification dated May 17, 2016, to all affected consignees via courier with instructions for identifying and returning the affected product. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Consignees were instructed to review the notification and ensure affected personnel were aware of the contents. Assist their Zimmer Bioment sales representative with the quarantine of any affected product from the table above. Complete the Certificate of Acknowledgement Form and return to corporatequality.postmarket@zimmerbiomet.com.If after reviewing this notification consignees have further questions or concerns please call the customer call center at 1- 800-348-2759 between 8:00 am and 8:00 pm EST. |  | 177 | 177 | Nationwide Distribution to AZ, CA, CO, GA, IL, KS, KY, MA, MD, MO, OH, PA, TN, and TX. | Open, Classified |  |  | 177 | Computer | Not_Computer
Z-0966-2016 | 73024 | No | PIF | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | EndoVive 3s Low Profile Balloon Kits |  |  | December 23, 2015 | 2015 | Xeridiem Mediem Medical Devices Inc | A feeding (reflux) valve leakage issue. The feeding valve appears to close inconsistently, or close in a delayed manner, when an extension set is disconnected from the Low Profile Balloon device, thus allowing feeding solution or gastric contents to leak out of the device in some instances. | Device Design | Xeridiem sent a customer notification letter dated December 23, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer.Customer was provided the following instructions:1. Segregate all devices from the above table that in Boston Scientific Corporation's inventory and restrict them from being distributed to end users.2. Determine the list of end users who have received devices from the above list, notify them of the quality issue, and have these devices returned to Boston Scientific Corporation. Provide Xeridiem Medical Devices with this list of end users.For any questions the customers is instructed to contact the firm.For questions regarding this recall call 520-882-7794. |  | 38 | 38 | Nationwide distribution to MA only. | Open, Classified |  |  | 7171 | Computer | Not_Computer
Z-1022-2016 | 73010 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | BrightView XCT &amp; BrightView XCT Upgrade, gamma cameras; For Single Proton Emission Computed Tomography. |  |  | December 17, 2015 | 2015 | Philips Medical Systems (Cleveland) Inc | The firm discovered four issues with the BrightView Family systems.Issue 1: During a cardiac non-Auto Body Contouring (ABC) scan, using Cardiac High Resolution (CHR) collimators, the collimator may come in contact with the bottom edge of the patient pallet. If this occurs, collision sensors would pause the scan.Issue 2:During a scan, the gantry motion may halt prematurely. This cause | Software design | On 12/22/2015, the firm sent Customer Information Letters to their consignees.The letter identified the product, problem and actions to be taken by the customers. |  | 427 | 427 units | Worldwide Distribution-US (nationwide) including the states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, VA, VT, WA, WI, WV &amp; WY, and in DC and PR and in the countries of Algeria, Argentina, Armenia, Australia, Austria, Bahrain, Belgium, Brazil, Canada, Cayman Islands, China, Czech Republic, Denmark, Dominican Republic, Egypt, Finland, France, Gabon, Germany, Greece, India, Indonesia, Iran, Ireland, Italy, Japan, Korea, Lebanon, Lithuania, Malaysia, Mexico, Netherlands, New Zealand, Norway, Oman, Palestinian, Panama, Philippines, Poland, Portugal, Russian Federation, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Syrian Arab Republic, Taiwan, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom &amp; Viet Nam. | Open, Classified |  |  | 1221 | Computer | Computer
Z-2039-2016 | 74245 | No | JOY | Automated cell-locating device. | Hematology | Hematology | 510(k) | CytoVision Image Analysis and Capture System, an automated cell-locating device. The Cytovision is a rapid metaphase finder, image acquisition and computer aided chromosome analysis system which assists the operator in viewing chromosomes and looking for cellular anomalies. CytoVision enables a qualified Cytogeneticist to rapidly and accurately analyze the chromosome banding pattern. |  |  | May 02, 2016 | 2016 | Leica Biosystems Richmond Inc. | Systems have an improperly activated Windows 7 OS, even though a valid Windows 7 license was associated with each manufactured system. The system appears to be unlicensed. These systems cannot be re-activated using standard Windows licensing procedures and are requiring checks of each potentially affected system to confirm Windows 7 OS is properly activated. | Software Manufacturing/Software Deployment | Leica Biosystems Richmond sent Medical device Field Correction notification letters to the customers of the potentially affected product. The initial written communication was sent to the customers via first class USPS on May 2, 2016. According to the firm a decision to conduct a Field Corrective Actions of the Ariol Image Analysis and Scanning Systems was done out of an abundance of caution. The Medical Device Field Correction letter describes the issues identified with the CytoVision Image Analysis and Capture System and the Windows 7 operating system. The letter provided the customer with actions they are required to take which include the following: 1) whether they have a properly activated Windows 7 operating system. 2) If they do not have a properly activated operating system they should to contact Leica Technical Support for additional information and to request a service appointment for correction to their system. For USA customers, they ask that they call 18005374669. International customers are asked to go tohttp://www.leicabiosystems.com/contactus/contactusonline/ and then choose your country. 3) Customers are asked to confirm receipt of the notice by signing and dating the attached Acknowledgement Form and emailing a copy to RichRA@leicamicrosystems.com or faxing it back to us at 18472363800. Service personnel will be sent to the affected customer locations to replace the hard drives of the affected systems. |  | 283 | 283 systems | Worldwide distribution. US nationwide; Portugal, Spain, Thailand, Italy, China, Romania, Egypt, India, France, Hungary, Mexico, Germany, Norway, Turkey, Pakistan, Malaysia, UK, Luxembourg, Israel, Canada, Switzerland, Taiwan, South Africa, Saudi Arabia, Korea, Singapore and Japan. | Open, Classified |  |  | 283 | Computer | Computer
Z-1600-2016 | 73884 | No | LGW |  |  | Neurology | PMA | Nuvectra, Algovita Spinal Cord Stimulation System Clinical Programmer, Model 4500. For use with Algovita Spinal Cord Stimulation Systems. The version of the software in the Clinician Programmer is v1.1.5. |  |  | April 06, 2016 | 2016 | Nuvectra | If used, the optional Swap feature copies programming parameters from the external pulse generator (EPG) as well as unintended calibration data to the implantable pulse generator (lPG). This may result in the user being unable to recharge the lPG.This safety notice only applies to CPs Model 4500. | Software design | Nuvectra sent an "Urgent Field Safety Notice" letter dated April 6, 2016 to customers. The letter identified the affected product, problem, recommendations and Mitigations, and actions to be taken. |  | 17 | 17 OUS | Germany | Open, Classified |  |  | 17 | Computer | Computer
Z-1594-2016 | 73742 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | CARESTREAM Image Suite V4; Image Suite V4: DICOM STORE SCP: REF/Catalog # 1056191, DICOM STORE SCP/FOR IMAGE SUITE V4: REF/Catalog # 6566988; Image Suite V4 Bundles: IMAGESUITE STANDALONE PACS: REF/Catalog # 1741289, IMAGESUITE STANDALONE PACS FOR INDIA: REF/Catalog # 1741297; Carestream PRO (Image Suite V4): PRO Medical Wireless GOS System-Desktop: REF/Catalog # 1741891, PRO Wireless System Laptop: REF/Catalog # 1741925, PRO Wireless System - w/o Computer: REF/Catalog # 1741933, PRO Tethered System Desktop: REF/Catalog # 1741941, PRO Tethered System Laptop: REF/Catalog # 1741958, PRO Tether System - w/o Computer: REF/Catalog # 1741966, PRO Fixed System - w/o Computer: REF/Catalog # 1741974, PRO Medical Wireless CsI System-Desktop: REF/Catalog # 1741982, PRO Wireless System Laptop: REF/Catalog # 1742006, PRO Wireless System - w/o Computer: REF/Catalog # 1742014, PRO Tethered System Desktop: REF/Catalog # 1742022, PRO Tethered System Laptop: REF/Catalog # 1742055, PRO Tether System - w/o Computer: Catalog # 1742063; PRO Fixed System - w/o Computer: REF/Catalog # 1742089 -- Made in U.S.A. by: Carestream Health, Inc., 150 Verona Street, Rochester, NY 14608 |  |  | April 15, 2016 | 2016 | Carestream Health Inc | Carestream Health received a complaint related to CARESTREAM Image Suite 4 from a foreign hospital stating that the annotation on the overlay is displayed as "L (Left)", when it should be "R(Right)". | Software design | Carestream Health sent an URGENT MEDICAL DEVICE CORRECTION letter dated April 6, 2016, and Consignee Notification Acknowledgement form to the US consignees on April 15, 2016, via Certified Mail, Return Receipt Requested. Foreign consignees are being notified based on the individual country regulations. Customers who have this Image Suite configuration but do not use it for CT or MR Imaging do not need to take any action. Customers who use the product for CT or MR Images are asked to make users aware of the problem and take additional steps to check correct orientation of any markers, until the software is updated. Customers with questions or concerns were instructed to contact the Carestream Customer Care Center in the US at 1-800-328-2910. Customers outside the US were instructed to contact their local Service Support number.For questions regarding this recall call 1-800-328-2917. |  | 306 | US: 37 units, Foreign: 269 units | NY and OH for further distribution nationwide; Foreign distribution to the following countries: Afghanistan, Angola, Argentina, Australia, Austria, Bahrain, Bangladesh, Belgium, Bosnia &amp; Herzgovena, Brazil, Bulgaria, Cambodia, Canada, Chad, Chile, China, Colombia, Croatia, Czech Rep., Djibouti, France, Germany, Greece, Hungary, India, Indonesia, Iran, Iraq, Ireland, Italy, Jamaica, Japan, Jordan, Kuwait, Kyrgyzstan, Laos, Latvia, Lebanon, Liechtenstein, Lithuania, Macao, Malaysia, Mexico, Montenegro, Myanmar (Burma), Nepal, Netherlands, New Zealand, Niger, Norway, Oman, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Scotland, Serbia, Singapore, Slovakia, South Africa, South Korea, Spain, Sri Lanka, Taiwan, thailand, United Arab Emirates, United Kingdom, Uruguay, Uzbekistan, Vietnam. | Terminated | March 16, 2017 | 335 | 2855 | Computer | Computer
Z-2044-2016 | 74241 | Yes | KOC | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Crit Line in a Clip (CLiC). CliC P/N: CL 10041001 |  |  | May 24, 2016 | 2016 | Fresenius Medical Care Renal Therapies Group, LLC | Use of the Blood Pressure Graph located on the Crit-Line screen of the 2008T Hemodialysis Machine may result in a frozen screen, possibly causing the blood pump to stop. | Software design | Fresenius Medical Care Renal Therapies Group, LLC issued Recall Letter issued 5/24/16 via certified mail. Clinics actively using the CLiC option have been verbally communicated to stop using the CLiC option at . The other clinics that did not activate the CLiC option have been verbally communicated to not activate the CLiC option at this time. A Software update to all 2008T Hemodialysis machines will be required.A reply form is included as part of the customer communication with instructions to fill out and return via fax or scanned/emailed copy.For technical Inquiries, please contact Technical Services at 1-800-227-2572. Additional medical concerns or questions, please contact the Medical Information and Communication line at 855-616-2309 or website www.fresenius-medinfo.com. |  | 875 | 875 units | Distributed US (nationwide) | Open, Classified |  |  | 875 | Computer | Computer
Z-1592-2016 | 73854 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON SC2000 Ultrasound imaging system with software version VB10C and using transesophageal (TEE) transducer Z6Ms, V5Ms or V7M. |  |  | April 08, 2016 | 2016 | Siemens Medical Solutions USA, Inc. | While imaging with a transesophageal (TEE) transducer (Z6Ms, V5Ms, or V7M), user may lose the ability to control the color region of interest, the pulsed wave or continuous wave Doppler gate or cursor, the M-mode cursor, the 2D field of view, or the RES region of interest with the trackball. | Software design | Under Siemens update program US013/16/S a Customer Advisory Notification will be mailed to all affected customers by Certified mail with return receipt beginning March 30, 2016. Letters provide customers information on how to avoid the issue until the software patch is installed. |  | 87 | 87 | Worldwide distribution. US nationwide, Austria, Canada, China, Germany, India, Italy, Spain, Sweden, and Switzerland. | Open, Classified |  |  | 87 | Computer | Computer
Z-1595-2016 | 73601 | No | LKD | Cutaneous carbon dioxide (PcCO2) monitor. | Anesthesiology | Anesthesiology | 510(k) | TCM CombiM |  |  | November 18, 2015 | 2015 | Radiometer America Inc | Radiometer America, Inc. is recalling TCM CombiM on rare occasions they may have an error that can cause the alarm to sound higher than the pre-set maximum volume. | Software design | RadioMeter sent an Product Notification letter out on November 18, 2015, January 10, 2016 and March 29, 2016.The notification sent on November 18, 2015 and January 10, 2016, instructed customers a temporary solution of switching off the analyzer and then on again in order to re-establish the pre-set alarm volume. If alarm sound is higher than the pre-set volume customers are instructed to contact their local Radiometer service representative. If any questions contact Radiometer America Technical Support at 1-800-736--600 opt. 4.The notification sent on March 29, 2016, provided the same instructions and temporary solution. The notification also informed customers that the firm will be in contact with them to discuss the upgrade process. |  | 1488 | 1488 | Worldwide Distribution - USA (nationwide) Distribution to the states of : MO, NE, AL, PA, IL, MN, TX, ID, CO, UT, NV, CA, NY, RI, OR, NJ, AZ, GA, IN, VA, LA, NM, MD and AR., and to the countries of : Canada, Austria, Czech Republic, Denmark, Finland, France, Germany, Hungary, Ireland, United Kingdom, Italy, Netherlands, Norway, Spain, Sweden, United Arab Emirates, Kuwait, Qatar, Saudi Arabia, Oman, Jordan, Croatia, Malta, Iceland, Brazil, Thailand, Philippines, Peru, Honduras, Mexico, Kenya, Israel, Romania, Russia, Poland, Turkey, Australia, New Zealand, China, Hong Kong, South Africa, India, Korea, Singapore and Taiwan. | Open, Classified |  |  | 1488 | Computer | Not_Computer
Z-2045-2016 | 74191 | No | GEY | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Myriad handpiece |  |  | February 24, 2014 | 2014 | Nico Corp. | This recall has been initiated due to a defective component within the handpiece which willcause the handpiece to fail. The failure will result in the inability to cut tissue with thehandpiece, rendering it non-functional. Should this failure occur, a replacement Myriadhandpiece will be required. If a replacement handpiece is not available, alternate instrumentation for tissue removal will be necessary. | Device Design | NICO sent an Urgent: NICO Myriad Handpiece Recall Notification on February 24, 2014, to all affected customers requesting return of the affected product. The recall notification included a description of the reason for the recall, affected product, customer responsibilities, and instructions for responding to the formal recall notification. If you have any questions regarding this matter, please call (317) 660-7118, ext. 104. |  | 63 | 63 units | US Distribution to the states of : OR, IL, IA, NC, TX, CA, GA, OH, WI, AR, MI, MD, IN. | Completed |  |  | 63 | Computer | Not_Computer
Z-2240-2016 | 74576 | No | JDO | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | TSP Hip Fracture Plating System |  |  | June 10, 2016 | 2016 | Biomet, Inc. | The screw may fracture at the junction of the shaft and the head during implantation. The patient may retain a foreign body if the fractured screw cannot be removed. Soft tissue irritation will not occur as the retained portion of the screw is below the surface of the plate. A delay in surgery less than 30 minutes may occur. | Device Design | Zimmer Biomet sent an Urgent Medical Device Recall letter dated June 10, 2016, to all affected customers via courier. Distributors and hospitals were instructed to locate and return the affected product, and implanting surgeons were instructed to monitor patients. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Questions related to this notice should be directed to (574) 372-6725, Monday through Friday, 8 a.m. to 5 p.m. |  | 5458 | 5458 units | Worldwide Distribution - US (Nationwide) Internationally to Australia, The Netherlands, and Japan | Terminated | July 20, 2017 | 405 | 5458 | Computer | Not_Computer
Z-2043-2016 | 74372 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Fujifilm Synapse PACS software version 4.4.000, |  |  | May 10, 2016 | 2016 | Fujifilm Medical Systems U.S.A., Inc. | Synapse cannot display image files, DICOM SR files, and/or Annotation files. The "Image Not Loaded" message displays instead | Software design | The firm, FujiFilm, sent an "MEDICAL DEVICE CORRECTION" letter dated May 10, 2016 via Federal Express to its customers. The letter describes the product, problem and actions to be taken.The customers were instructed to: Extend the Ad hoc deletion retention period or discontinue use of Ad Hoc Deletion, and complete and submit the attached Customer Response Form to Fujifilm via fax to (203)-251-7863. Affected systems will be corrected via on-site visits by FMSU Customer Service Engineers upgrading to Version 4.4.100.Questions or concerns, contact the FUJI FILM Customer Support Center at 1-888-FUJIMED Option #1 . |  | 27 | 27 units | US Distribution to states of: CA, FL, KS, NE, NY, and PA. | Open, Classified |  |  | 27 | Computer | Computer
Z-0045-2017 | 75068 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | ROI-A Anterior Delivery Device. Orthopedic manual surgical instrument. Model number: SI-ROIA-0023 |  |  | August 29, 2016 | 2016 | LDR Spine USA, Inc. | The ROI-A Anterior Delivery Device T-Handle, part number SI-ROIA-0023 has experienced binding of the rotation of the t-handle. There have been no reports of patient injury | Device Design | The firm, LDR, contacted affected consignees via phone and then sent a follow up email dated 9/8/16. The phone script and email described the product, problem and actions to be taken. The consignees were instructed to promptly return the instrument using the RMA#1180. The firm will perform a rework that prevents the binding to occur.If you have any question, contact the Quality Engineering Manager by email: ron.musselman@zimmerbiomet.com, phone: 512-344-3436 or Field Inventory Manager - US Supply Chain or phone: 512-344-3436. |  | 24 | 24 units | US Distribution to states of: GA, OR, FL, KY, NC, AR and PA. | Terminated | January 13, 2017 | 137 | 24 | Computer | Not_Computer
Z-1993-2016 | 73473 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | LIFEPAK15 monitor/defibrillator with End-Tidal C02 (EtC02) feature installed. |  |  | January 27, 2016 | 2016 | Physio-Control, Inc. | The firm became aware that when using EtC02 in the kPa or % setting and in a situation where the reading is above 9.9 kPa, the display of the LIFEPAK 15 respiratory rate may partially obscure a portion of the leading digit of the EtC02 value. This affects all LIFEPAK 15 with an EtC02 feature installed and configured to the kPa or % setting. | Software design | Physio-Control sent the letter with the title Information Regarding LIFEPAK 15 MONITOR/DEFIBRILLATOR WITH End-tidal CO2 OPTION, dated January 2016, to customers in the United Kingdom (UK) as REQUESTED by the Medicines and Healthcare product Regulatory Agency (MHRA) on January 27, 2016.Physio initiated a voluntary field action on February 23, 2016 and started notifying US consignees on March 22, 2016 and started notifying international consignees on April 18, 2016. The Information Regarding LIFEPAK 15 MONITOR/ DEFIBRILLATOR WITH End-tidal CO2 OPTION letter, dated March 2016, was sent.Customer Letters began mailing in EMEA region, beginning on 06/20/2016 via designated Courier.Customer Letters began mailing in Canada on 09/09/2016 via FedEx.Customer Letters began mailing in Latin America on 09/16/2016 via email.Customer Letters began mailing in APAC region on 09/22/2016 via email.Second Customer letters with Confirmation Sheets began mailing in the US October 2016 via FedEx and USPS Certified Mail.Actions should be taken by users:1. Please forward this information to all of your sites, trainers and users that may have a LIFEPAK 15 monitor/defibrillator.2. If you use your LIFEPAK 15 monitor/defibrillator to display EtC02 values you should also display the EtC02 waveform which includes the range along the right side of the display.3. If you use either kPa or % settings, activate EtC02 alarms to identify EtC02 values outside the normal range.4. If your protocols allow this option, consider changing the EtC02 to measure in the mmHg setting.The firm reported that the revised software for LifePak15 version 4 (v4) and the customer install instructions would be available on LIFENET by April 25, 2016.The revised software for LifePak 15 versions 1 and 2 (v1 and v2) and the customer install instructions are estimated to be available on LIFENET in mid May 2016.Customers who have any questions about this subject, please contac |  | 44798 | 44798 units total (23189 units in US; 21609 units international) | Worldwide Distribution - US (nationwide) including Guam and Puerto Rico and Internationally to Âżland Islands, Algeria, Argentina, Aruba, Australia, Austria, Azerbaijan, Bahrain, Barbados, Belgium, Bosnia and Herzegovina, Botswana, Brazil, Brunei Darussalam, Cambodia, Canada, Cayman Islands, Chile, China, Costa Rica, Croatia, CuraÂżao, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, Eritrea, Faroe Islands, Finland, France, French Guiana, French Polynesia, Georgia, Germany, Greece, Guadeloupe, Guinea, Guyana, Hong Kong, Hungary, Iceland, India, Indonesia, Iran, Islamic Republic of, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Korea, Republic of, Kosovo, Kuwait, Latvia, Lebanon, Libya, Lithuania, Luxembourg, Macedonia, the former Yugoslav Republic of, Malaysia, Maldives, Malta, Martinique, Mauritius, Mayotte, Mexico, Morocco, Mozambique, Namibia, Netherlands, New Caledonia, New Zealand, Norway, Oman, Pakistan, Panama, Papua New Guinea, Poland, Portugal, Qatar, RÂżunion, Romania, Russian Federation, San Marino, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Suriname, Sweden, Switzerland, Taiwan, Province of China, Thailand, Trinidad and Tobago, Turkey, Ukraine, United Arab Emirates, United Kingdom, and Yemen. | Terminated | March 31, 2017 | 429 | 44798 | Computer | Computer
Z-0118-2017 | 75161 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | CADstream software |  |  | October 19, 2011 | 2011 | Merge Healthcare, Inc. | Customers may experience an issue with the software study preferences when changes are made to the study protocol, resulting in incorrect patient follow-up. | Software design | Merge sent an Urgent Medical Device Recall dated January 28, 2016 to affected customers via email and certified mail if no email address was available. The letter identified the affected product, problem and actions to be taken. For questions Customer Service at 877-741-5369 or support@merge.com. |  | 844 | 844 sites potentially have the affected versions | US Nationwide Distribution | Open, Classified |  |  | 844 | Computer | Computer
Z-0016-2017 | 75017 | No | EHD | Extraoral source x-ray system. | Dental | Radiology | 510(k) | Panoramic X-ray Model PC-1000, Laser-1000 packed in a crate |  |  | August 19, 2016 | 2016 | Panoramic Rental Corp. | Panoramic Corporation is performing a voluntary Medical Device Correction on certain PC-1000 X-ray systems due to the potential for a safety nut to be ineffective, which could result in blunt force trauma that may impact the head, neck or shoulders. | Device Design | Panoramic Corporation sent an Urgent Medical Device Correction letter dated August 19, 2016 to customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete the attached form and return to Panoramic. For questions contact Panoramic at correction@pancorp.com or 888-381-2577 Monday through Thursday 8:00 AM to 6:00 PM and Friday 8:00 AM to 5:00 PM EST in order to verify the effectiveness of the safety component. |  | 13340 | 13,340 | US Nationwide Distribution including Puerto Rico | Open, Classified |  |  | 13340 | Computer | Not_Computer
Z-1605-2016 | 73780 | Yes | NVZ |  |  | Cardiovascular | PMA | Medtronic CareLink" Monitor (2490C) and Medtronic CareLink Express" Monitor (2020B) |  |  | March 31, 2016 | 2016 | Medtronic Inc., Cardiac Rhythm and Heart Failure | A recent firmware update developed by Medtronic for the 2490C CareLink Monitors and 2020B CareLink Express Monitors included incorrect data on the country analog modem dial-up configuration table used to dial into the CareLink Network. This incorrect data resulted in the removal of the 0 prefix necessary for select countries as well as an incorrect alignment of phone numbers for other select cou | Software design (manufacturing process) | Medtronic sent an "Urgent Field Safety Notice" dated April 2016to their customers. The letter was addressed to Healthcare Professional. The letter identified the affected product, problem and customer actions to be taken. For questions contact your local Medtronic Representative. |  | 5038 | 5038 | AUSTRIA, AUSTRALIA, BELGIUM, SWITZERLAND, GERMANY, DENMARK, FRANCE, GREAT BRITAIN, HUNGARY, JAPAN, NETHERLANDS, NEW ZEALAND, POLAND, PORTUGAL, SWEDEN. | Open, Classified |  |  | 5038 | Computer | Computer
Z-0321-2017 | 75062 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | TD60 display Telemetry Transmitter, SHENZHEN Mindray Bio-Medical Electronics Co., LTD, Mindray Building, Keji 12th Road South, Hi-Tech Industrial Partk, Nanshan, Shenzhen, 518057 P.R. China |  |  | September 08, 2016 | 2016 | Mindray DS USA, Inc. dba Mindray North America | Mindray has become aware that the TD60 Display Telemetry Transmitter used with the BeneVision Central Station may crack on the bezel surrounding the display and also the transmitter may not power on or will power on but display an error message. | Device Design | Mindray sent a "Medical Device Action" letter dated September 8, 2016 via UPS to their affected customers. |  | 1278 | 1,278 units | US and Canada | Open, Classified |  |  | 1278 | Computer | Computer
Z-1616-2016 | 73657 | No | MQP | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Atlas Spine Torque Limiting Handle 6N-m, P/N 70089 Model # ITN-001-000 |  |  | July 26, 2013 | 2013 | Atlas Spine, Inc. | The device was not appropriately set to mitigate the risk of implant fracture upon rotation inside the disc space. | Device Design | Atlas Spine, Inc. informed all consignees of the recall and sub-recall on July 26, 2013, in person, electronically, and/or via email. Firm takes the responsibility of tracking of set return and inventory control of the product.For questions regarding this recall call 561-354-1108, ext 102. |  | 16 | 16 | Nationwide Distribution to FL, MO, and WV. | Completed |  |  | 16 | Computer | Not_Computer
Z-0826-2016 | 73274 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | software for Syngo Dynamics |  |  | January 29, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Siemens is releasing a software update that addresses an issue of mixing data from multiple patients. In rare situations, echo trend graphs may mix data from multiple patients. | Software design | The firm, Siemens Healthcare, sent a "Customer Advisory Notice" dated 1/19/2019 to its customers.The notice described the product, problem and actions to be taken. The customers were instructed to create a new department if immediate resolution is necessary, use workplace for reporting and hand over the customer letter before or during installation of field modification(PSL) distributed with Update Instruction SY013/16P. A software update to resolve reported issues is being released via an Update Instruction SY013/16/P. If you have any questions, contact Regulatory Technical Specialist at 610-448-6478 or email: anastasia.sokolova@siemens.com. |  | 151 | 151 | US Nationwide Distribution including states of: AL AR AZ CA CO CT DE FL GA IL IN KS LA MA MD MI MN MO NC ND NE NJ NM NY OH OR PA SC TN TX UT VA WA and WI. | Open, Classified |  |  | 151 | Computer | Computer
Z-0315-2017 | 75473 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | da VinciÂż Xi" Surgical System, model number IS4000, A70_P5x with P5 Software; |  |  | October 05, 2016 | 2016 | Intuitive Surgical, Inc. | Intuitive Surgical has identified a software anomaly in the da Vinci Xi P5 software that can result inunexpected master movement and potential instrument tip movement under certain circumstances | Software design | Urgent Medical Device Correction letters were sent on October 12, 2016 by traceable method to all affected customers. Letters notify customers of the issue and provided instructions to prevent it from happening until ISI software can install new software upgrade in 1st Q 2017. |  | 677 | 677 devices | Australia, Austria, Belgium, Colombia, Cyprus, Czech Republic, Denmark, Finland, France, Germany, Hong Kong, India, Ireland, Israel, Italy, Japan, Monaco, Netherlands, Norway, Portugal, Puerto Rico, Qatar, Romania, SouthKorea, Spain, Sweden, Switzerland, Taiwan, Turkey, United Kingdom, United States | Terminated | May 01, 2017 | 208 | 677 | Computer | Computer
Z-0824-2016 | 73099 | No | JPA | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(k) | BCS Automated Blood Coagulation Analyzer System, Multipurpose system for in vitro coagulation studies, Catalog Numbers: 10454728 (BCS RECONDITIONED), 10454729 (BCS RECONDITIONED), 10454742 (BEHRING COAGULATION SYSTEM), 10459303 (BCS RECONDITIONED), 10460659 (BCS INSTRUMENT), 10461881 (BCS INSTRUMENT), IVD --- Device Listing D018966 |  |  | January 14, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc. | False short clotting times for PT on BCS and BCS XP with Dade Innovin can occur. This has also an effect for false low PT INR values. These false short clotting times were observed only with turbid, hemolytic and/or icteric samples from intensive care patients. The corresponding reaction curves of these samples showed a pre-peak, which can lead to falsely short clotting time. The result was not flagged by the existing check algorithms. | Software design | Urgent Medical Device Correction Letter (dated 1/12/2016) was sent to US consignees via FedEx on 1/14/2016 and to foreign consignees (distribution is determined at country level). The letter re-emphasizes the importance of checking the interference levels given in the respective Application Sheets of the BCS and BCS XP systems. For questions, please contact your Siemens Customer Care Center or your local Siemens technical support representative. |  | 1680 | 1680 units total | Worldwide Distribution-US (nationwide) and the countries of Argentina, Australia, Austria, Bahrain, Belarus, Belgium, Bosnia Herzegovina, Brazil, Canada, Chile, China, Croatia, Cyprus, Czech Republic, Denmark, Finland, France, Germany, Guadeloupe, Hungary, Italy, Kuwait, Latvia, Lebanon, Libya, Macedonia, Mexico, Montenegro, Netherlands, Peru, Poland, Portugal, Qatar, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Turkey, United Arab Emirates, and Uruguay. | Terminated | March 16, 2017 | 427 | 1680 | Computer | Not_Computer
Z-1620-2016 | 73882 | No | FFT | Stomach pH electrode. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Given Imaging BravoÂżÂż pH capsule delivery device, 1-pack. |  |  | April 15, 2016 | 2016 | Covidien LLC | Potential of an allergic reaction in a patient with nickel sensitivity. | Device Design | The firm, Medtronic, sent an "URGENT: Medical Device Recall" letter dated 4/15/2016 to its Customers via Federal Express. The letter describes the product, problem and actions to be taken. The customer was instructed to promptly forward the information from the letter to recipients, if you have distributed any of the affected devices; and complete and return the Acknowledgement and Receipt Form by email to: Quality Compliance MITGFCA@Covidien.com or Fax to: (203) 492-7719. Medtronic is adding the following statement to the BravoÂż pH monitoring system User Manual and product package insert so that physicians are aware of the nickel content and the potential of an allergic reaction in a patient with nickel sensitivity. "The Bravo capsule contains a trocar needle that is made of stainless steel. Use caution in patients with known sensitivities or allergies to the metals that are contained including chromium, nickel, copper, cobalt and iron. The BravoÂż pH test lasts from 48-96 hours."If you have any questions, please contact Technical Service at 1-800-448-3644, option 3. |  | 1305 | 1,305 units | Worldwide Distribution: US (nationwide) and countries of: Argentina, Australia, Belgium, Brazil, Canada, Cyprus, El Salvador, Greece, Hong Kong, Israel, Mexico, Panama, Singapore, South Korea, and Thailand. | Terminated | July 13, 2017 | 454 | 14200 | Computer | Not_Computer
Z-0094-2017 | 75186 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Roche/Hitachi MODULAR Analytics Combination Systems |  |  | July 15, 2016 | 2016 | Roche Diagnostics Operations, Inc. | Roche has received four reports of injury caused by the sharp edge of the center guide rail of ananalyzers sample rack tray. The cleaning procedure for the sample rack tray has been revised to clarify the cleaning process. | Device Design | On 7/15/2016, URGENT MEDICAL DEVICE CORRECTION notifications were sent to the affected customers via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, an updated cleaning procedure, and instructions for responding to the formal recall notification. Please contact the Roche Support Network Customer Support Center at 1-800-428-2336 if you have questions about the information contained in this UMDC.On 10/5/2016, URGENT MEDICAL DEVICE CORRECTION UPDATE notifications were sent to the affected customers via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. The notification updated the 7/15/16 notification to include additional affected systems and to provide updated cleaning instructions for all systems via the enclosed Operators Manual addendum. |  | 9014 | 9014 in total | US Nationwide Distribution | Open, Classified |  |  | 9014 | Computer | Not_Computer
Z-0305-2017 | 63994 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Symbiq One Channel Infuser; an Rx medical device infusion pump used to administer I.V. fluids |  |  | February 04, 2013 | 2013 | Hospira Inc. | Symbiq Infusers have the potential to experience a white screen during titration of a Keep Vein Open (KVO) delivery. This can allow the clinician to select the "NEXT" button prior to completion of the programming. If this error occurs, a delay in therapy could occur and potentially result in significant injury. | Software design | Hospira sent an Urgent Device Field Correction notice dated February 4, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to follow the steps provided in the letter while Hospira is developing a software change. Customers were asked to complete the attached reply form and return it via fax to the number on the form. Customers with questions were instructed to contact Stericycle at 1-888-201-9332.For questions regarding this recall call 224-212-4892. |  | 35596 | 35,596 devices in total | Worldwide Distribution - USA (nationwide) and Canada | Open, Classified |  |  | 35596 | Computer | Computer
Z-0858-2016 | 72720 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Philips Healthcare IntelliVue Info Center iX, B.0, 866389 |  |  | November 25, 2015 | 2015 | Philips Electronics North America Corporation | Reconstructed ECG leads viewed or printed at the Information Center iX may misrepresent the ECG waveform in specific leads. | Software Design Change | Philips Healthcare issued a Field Safety Notice on 11/25/15 to inform customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue.Philips will provide a software correction free of charge. On 1/28/16 the Philips issued an updated letter to users.Philips has initiated a correction to address this issue. A software correction will be provided to customers with impacted devices at no charge.Philips is asking customers to follow the Action to be Taken by Customer/User section of the Field Safety Notice: Arrhythmia alarms are not affected. ST alarms, ST values and measurements are also not affected and could be used to assist with the decision to obtain a diagnostic ECG.The ECG waves at the bedside are not affected. Printing ECG waves from the bedside is an alternative to printing from the PIIC iX.If you need any further information or support concerning this issue, contact your local Philips Healthcare Customer Service representative or the Customer Care Solutions Center at (800) 722-9377 with questions or concerns about this Device Correction. |  | 1880 | 1880 units | Worldwide Distribution-US (nationwide) and the countries of Algeria Australia Austria Belgium Bulgaria Canada Denmark Estonia Finland France Germany Hong Kong Hungary Iceland India Indonesia Iraq Ireland Israel Italy Japan Kenya Korea Kuwait Luxembourg Malaysia Mongolia Morocco Netherlands New Zealand Nigeria Norway Poland Qatar Romania Russian Federation Saudi Arabia Singapore Slovenia South Africa Spain Sweden Switzerland Taiwan Turkey United Arab Emirates and United Kingdom. | Open, Classified |  |  | 13120 | Computer | Computer
Z-0301-2017 | 75384 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | FemTouch delivery system includes the FemTouch Kit (AC-1007415) and FemTouch" Tip (SPSA-20002590) |  |  | September 25, 2016 | 2016 | Lumenis Limited | The cleaning instructions provided in FemTouch Operator's Manual might be ineffective for a fully assembled device and may lead to a potential for cross-contamination. | Software Manufacturing/Software Deployment | On 9/25/16 consignees were notified via an Urgent - Field Safety Corrective Action letter with a FemTouch disassembly tool kit as well as FemTouch disassembly instructions. The instructions describe the simple disassembly process to be performed as part of the Pre-cleaning phase and the process for reassembling the device prior to re-use. It is recommended that customers clean and sterilize all of their handpieces according to the pre-cleaning disassembly process before each use. Consignees were asked to complete and return a verification form. Questions about the FSCA, should be directed to the Lumenis field correction administrator, Rafi Barel at Rafi.Barel@lumenis.com. |  | 69 | 69 units | Worldwide Distribution -- USA, Italy, Russian Fed., Australia, United Arab Emirates, Spain, United Kingdom, India, Kuwait, and Belarus. | Terminated | April 11, 2017 | 198 | 69 | Computer | Not_Computer
Z-0300-2017 | 75167 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | enGen Laboratory Automation System (Product Code ENGEN) configured with Thermo Scientific Centrifuge Module and TCAutomation Software Version 3.6.1 and Below; IVD. --- Thermo Scientific Centrifuge Module has Product Code |  |  | September 07, 2016 | 2016 | Ortho-Clinical Diagnostics | Ortho confirmed that it is possible for the mis-association of sample identification and assay results due to an error with Thermo Fishers centrifuge module. It is possible for mis-associated assay results to be reported from the laboratory prior to the identification of the Cross Check error on the enGen Laboratory Automation System. Assay results associated with the wrong patient and reported out of the laboratory could lead to inappropriate intervention with the potential for serious injury to the patient. | Software design | The firm, Ortho Clinical Diagnostics, sent an "URGENT PRODUCT CORRECTION NOTIFICATION" Letter (Ref. CL2016-183, dated 9/7/2016) was sent via FedEx overnight courier and/ or ORTHO PLUS e-Communications and/ or US Postal Service Priority Mail (for PO Boxes only) to all enGen customers with centrifuge modules to inform them of the issue and advise them to check samples associated with cross check errors as instructed in the operator manual, and perform the preventative maintenance and cleaning procedures for that centrifuge module as provided in the maintenance manual and user guide. The customers were also instructed to Complete and return the Confirmation of Receipt form by 16 September 2016 via Fax:1.888.557.3759 or 1.585.453.4110 or email:RA-OCDUS-CONFIRMAD@ORTHOCLINICALDIAGNOSTICS.COM. For questions regarding this notification, please call the Ortho Care Solutions Center any time at 1-800-421-3311. Foreign affiliates were informed by e-mail on 9/7/2016 of the issue and instructed to notify their consignees of the issue and appropriate actions to take. |  | 60 | 60 (US: 24; Foreign: 36) | Worldwide Distribution-(Nationwide) and foreign distribution to the following countries: Belgium, Bermuda, Canada, Chile, Denmark, France, Italy, Mexico, Norway, Portugal, Sweden, Taiwan and United Kingdom. | Open, Classified |  |  | 60 | Computer | Not_Computer
Z-0117-2017 | 75393 | Yes | NIK |  |  | Cardiovascular | PMA | Quadra Assura, Model No. CD3265-40, CD3265-40Q, CD3365-40C, CD3365-40Q. |  |  | October 10, 2016 | 2016 | St Jude Medical Inc. | St Jude Medical is recalling implantable cardiac defibrillator (ICD) and cardiac resynchronizationtherapy defibrillator (CRT-D) devices because they are subject to premature battery depletion due to lithium cluster with the battery causing a short circuit. | Device Design | St. Jude Medical, Inc., sent a customer letter to all affected customers to inform them that St. Jude Medical recently announced a global medical device advisory for a subset of their company's Fortify, Fortify Assura, Quadra Assura, Unify, Unify Assura and Unify Quadra implantable cardiac defibrillator (ICD) and cardiac resynchronization therapy defibrillator (CRT-D) devices. The letter informs customers that the cause for Premature Battery Depletion is attributed to the formation of lithium clusters within the battery causing a short circuit. There is a possibility that affected devices may lose battery power within days. Thus far, there have been two (2) deaths and ten (10) serious events (fainting) that may have been associated with premature battery depletion. Customers with questions were instructed to contact www.sjm.com/batteryadvisory, (866) 915-5065. |  | 193954 | 193,954 units | Worldwide Distribution - US (nationwide) and OUS | Open, Classified |  |  | 603524 | Computer | Computer
Z-0297-2017 | 75376 | No | MJN | Vascular clamp. | Cardiovascular | Cardiovascular | 510(k) | The PAXWIRE Occlusion Balloon System consists of an occlusion catheter, inflation handle and syringes. It is designed for use in the iliofemoral artery to provide temporary occlusion as well as a path for dilator and/or sheath introduction, removal or repositioning during catheterization procedures. |  |  | October 06, 2016 | 2016 | AccessClosure, Inc., A Cardinal Health Company | AccessClosure, Inc. (ACI) is voluntarily recalling the PAXWIRE Occlusion Balloon System device because it was made aware from customers that the device may be removed in the absence of vascular sheath across the arteriotomy. This is a deviation from the firm's standard catheterization technique. | Device Design | Cardinal Health sent an Urgent Medical Device Field Safety Notification letter to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed complete the attached Field Action Return Response Form as soon as possible indicating the number of devices that have been used as their location and the number that will be returned. Customers should discuss the reimbursement for their purchased and return product with their sales representative, who will coordinate the process. Customers with questions were instructed to call 408-610-6570.For questions regarding this recall call 408-610-6500. |  | 6 | 6 in USA | US distribution to FL and NJ | Open, Classified |  |  | 6 | Computer | Not_Computer
Z-0294-2017 | 75425 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Fusion Workstation.; Indicated for the transmission and review of radiological images. |  |  | October 28, 2011 | 2011 | Merge Healthcare, Inc. | After a period of time running Fusion Workstation, the Hounsfield measurement tool will report incorrect values. | Software design | The recalling firm initiated phone calls to affected sites beginning 1/30/2016. |  | 13 | 13 sites potentially hae the affected version | Distribution was made to AZ, CA, ID, IL, IN, MI, MN, and NJ. | Open, Classified |  |  | 13 | Computer | Computer
Z-0291-2017 | 75209 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Philips IntelliVue MX40 Patient Monitor: IntelliVue MX40 WMTS Smart-hopping (1.4GHz) |  |  | September 19, 2016 | 2016 | Philips Electronics North America Corporation | Philips IntelliVue MX40 Patient Wearable Monitor Configuration Setting Disables Generation and Delivery of ECG Alarms to PIIC iX | Software design | The firm, Philips Healthcare, sent a "Field Safety Notice (FSN)" dated 9/14/2016 to its customers. The notice described the product, problem and actions to be taken. The customers were instructed to make sure that the MX40 Patient Wearable Monitor local configuration setting 'Alarms On' is configured to its factory default until your software is upgraded; and follow the instructions listed in the letter. A Philips Healthcare representative will contact customers with affected devices to arrange for the installation of the software upgrade.If you need any further information or support concerning this issue, please contact Philips Customer Support at 1-800-722-9377. |  | 9804 | 9804 units | Worldwide Distribution-US (Nationwide) and countries of: Aruba, Australia, Austria, Belgium, Bulgaria, Canada, China, Czech Republic, Denmark, Finland, France, Germany, Ghana, Greece, Hong Kong, Iceland, India, Iraq, Ireland, Italy,Japan, Korea, Republic of, Malaysia, Mexico, Netherlands, New Zealand, Norway, Philippines, Poland, Portugal, Singapore, Slovenia, South Africa, SpainSweden, Switzerland, Thailand, Turkey, United Arab Emirates and United Kingdom. | Open, Classified |  |  | 13840 | Computer | Computer
Z-0852-2016 | 73179 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | AccuLIF PL Tubing Set and AccuLIF TL Tubing Set |  |  | December 16, 2015 | 2015 | Stryker Spine | There were two potential interference conditions identified with the way the tubing set attaches to the inserter. | Device Design | Stryker Spine sent an Urgent Product Removal letter dated December 17, 2015 to affected customers via Fed Ex priority overnight. The letter identified the affected product, problem and actions to be taken. Customers are asked to follow instructions included in the letter. The customer response form should be completed and sent back to email via Spine-RegulatoryActions@Stryker.com or Fax 855-632-9049. For questions call (201)749-8389. |  | 413 | 413 units | US Nationwide Distribution | Open, Classified |  |  | 413 | Computer | Not_Computer
Z-0296-2017 | 74701 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Endoscopic Articulating Linear Cutter, 38ATW35. VASCULAR/THIN 35MM |  |  | July 15, 2016 | 2016 | Ethicon Endo-Surgery Inc | Quality control identified a component issue where the pinion gear in the device could fail under extreme use cases. If this condition occurs, staples will be formed past the cut line and the device can be opened and removed from the patient. The firing stroke may be interrupted and the knife may not fully return to the home position potentially exposing the healthcare professional to a sharps injury. | Device Design | Ethicon sent an Urgent Medical Device Recall letter dated July 15, 2016, to all affected consignees via UPS next day mail. The letter identified the product, the problem, and the action to be taken by the customer. Consignees were instructed to examine their inventory immediately and quarantine any affected product and return to ENDOPATH. If product subject to the recall was further distributed consignees were instructed to notify the other facility and arrange for return of the affected product. Customer were instructed to complete the Business Reply Form confirming receipt of this letter and fax to Stericycle at 1-888-228-5719 or email to ethicon5182@stericycle.com.Customers with questions were instructed to contact the Ethicon Customer Support Center at 1-877-384-4266.For questions regarding this recall call 513-337-3582. |  | 213 | 213 | Worldwide Distribution - US including AL, AR, AZ, CA, CO,CT, DC, FL,GA, IA,IL,KS,KY,LA,MA,MD, ME, MI,MN,MO, MS, MT,NC, ND, NH,NJ, NM, NV, NY, OH, RI, TN, TX, VA, VT, WA, WI and Internationally to Brazil and Canada. | Open, Classified |  |  | 318 | Computer | Not_Computer
Z-0853-2016 | 73209 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Philips IntelliVue Measurement Module X1 |  |  | January 04, 2016 | 2016 | Philips Electronics North America Corporation | The ST elevation alarm on the Patient Monitor or standalone X2 Measurement Module will not sound when indicated for all chest leads derived using Hexad 12-Lead ECG Monitoring in the Host Monitor | Software Design Change | Philips Healthcare issued the Field Safety Notice on January 4, 2016, and informed customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue.Philips is asking customers to follow the Action to be Taken by Customer/User section of the Field Safety Notice: Until your software is upgraded, please make sure that the ST Analysis is switched ON when using STE measurement in the Hexad 12-lead ECG Monitoring. This can be done by entering the ST Analysis Menu and select ST Analysis to On . For more detailed information, please refer to the Instructions for Use (IFU) of your host monitor or to the X2 Measurement Module IFU.Philips will provide a software correction free of charge. For further questions please call (978) 659-3000. |  | 13300 | 13,300 | Worldwide Distribution - Nationwide Distribution and to the countries of : Argentina, Australia, Austria, Bahrain, Bangladesh, Belgium, Bolivia, Brazil, Canada, Chile, China, Colombia, Czech Republic, Denmark, Dominican Republic, Estonia, Finland, France, Germany, Greece, Hong Kong, Hungary, Iceland, India, Indonesia, Iran ( Islamic Republic of ) Ireland, Israel, Italy, Japan, Jordan, Korea, Republic of, Kuwait, Latvia, Lebanon, Libyan Arab Jamahiriya, Lithuania, Luxembourg, Malaysia, Morocco, Myanmar(Burma) Netherlands Antilles, Norway, Oman, Pakistan, Philippines, Poland, Qatar, Romania, Russian Federation, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Province of China, Thailand, Turkey, United Arab Emirates, United Kingdom and Vietnam. | Open, Classified |  |  | 13300 | Computer | Computer
Z-0137-2017 | 75239 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | CADstream software. The firm name on the label is Merge Healthcare, Hartland, WI. |  |  | January 30, 2016 | 2016 | Merge Healthcare, Inc. | The software was not presenting the PURE filter user interface message on DICOM transfer studies. If a site starts using PURE filter, the user interface message will not display. This may result in incorrect patient follow-up. | Software design | Merge Healthcare Inc. sent an Urgent Medical Device Recall letter dated January 28, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The notification instructs the customer to pass the information onto all other users, as well as any downstream customers if it was further distributed. In addition, the customer is requested to call Customer Service or send them an email to schedule the upgrade. If the customer chooses not to upgrade, they are requested to complete the attached form or go online and complete it.Merge Healthcare is committed to improve efficiencies and enhancing the quality of healthcare worldwide.If you have any additional questions, call Merge Customer Service at (877) 741-5369 or support@merge.com. |  | 844 | 844 sites potentially have the affected versions | Worldwide Distribution - USA (nationwide) Distribution. Foreign/military/government distribution was also made. | Open, Classified |  |  | 844 | Computer | Computer
Z-0862-2016 | 73041 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo Plaza Picture archiving and communication system. |  |  | December 22, 2015 | 2015 | Siemens Medical Solutions USA, Inc | Siemens is releasing an updated software version to address several software issues including RGB images will show "?" since calculation of HU is not possible; save as option enabled; changes in access for loading studies; breast region is now properly fitted to segment boundary when clicking fit breast to screen. | Software design | Siemens sent a Voluntary Correction and Removal Action letter dated December 30, 2015. The letter identified the product, the problem, and the action to be taken by the customer. Customers were notified that a software update VB10A_HF01/06 is being released simulataneously to eliminate the issue. Customers with questions were instructed to call 1-888-826-9702. |  | 68 | 68 | Nationwide Distribution including ALCACTFLINKSLAMAMDMIMNMOMSNENYOHORPATXWAWI | Open, Classified |  |  | 68 | Computer | Computer
Z-1607-2016 | 74005 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Toshiba American Medical Systems Celesteion, Item Code: PCA-9000A/2, 9000A/3B, 9000A/2B |  |  | April 24, 2016 | 2016 | Toshiba American Medical Systems Inc | Potential degradation of detector performance. | Device Design | Toshiba sent an Urgent Medical Device Correction letter dated April 24, 2016 to all affected customers. The letter informed customers that Toshiba America Medical Systems is bringing attention to a potential problem with the Celestion PET/CT System. The letter informed the customers of a potential degradation of detector performance and informed the customers of the problems and the corrective actions to be taken. Customers were instructed to complete and return the attached form and fax it to (877) 349-3054 or be sent via email to regulatoryaffairs@toshiba.com. Customers with questions were instructed to call (800) 421-1968 or contact their local Toshiba Representative at (800) 521-1968. |  | 3 | 3 units | US Distribution to WI and NV. | Open, Classified |  |  | 3 | Computer | Computer
Z-1966-2016 | 73598 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | Vaxcel PICC with PASV Catheter Kit under the following labels: 1) 3F, UPN H965454350, Rx ONLY, 2) 4F, UPN M001454550, Rx ONLY &amp; 3) 5F, UPN M001454700, Rx ONLY. |  |  | March 14, 2016 | 2016 | Navilyst Medical, Inc., an AngioDyamics Company | PICC catheters that contain valve(s) manufactured prior to July 15, 2014 lack a manufacturing specification change that reduces the incidence of hemolysis during blood sampling through the PICC catheter. | Device Design | Navilyst Medical distributed Urgent Voluntary Medical Device Recall notification letters on March 14, 2016, and Reply Tracking forms to their customers via courier service. Customers were asked to immediately remove any affected (recalled) product from their inventory (whether in labs, central supply, shipping and receiving or any other location). Segregate this product in a secure location for return to Navilyst Medical. Immediately forward a copy of therecall notice to all sites to which you have distributed affected product. If affected product is located in your institution, please call Navilyst Medical Customer Service at 1-800-772-6446 Monday to Friday, 8 am - 7 pm, EST to obtain a replacement or credit for your returned product. Promptly complete, sign and return the enclosed Reply Verification Tracking Form (even if you do not have any product to return) by fax to 1-800-782-1357.For questions regarding this recall call 518-792-4112. |  | 2716 | 2,716 units | Worldwide Distribution US (Nationwide) Internationally to AU, GR, SE, SA, IN, CA, GB, BR, BE, NL, ES, KR, HK, AE, IT and ES. | Open, Classified |  |  | 17283 | Computer | Not_Computer
Z-1601-2016 | 73907 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo Dynamics; Kinetdx Picture Archiving and Communications System |  |  | April 14, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Siemens has identified changes in default carry forward workflow using VA10B_HF03 version of Syngo Dynamics. If the template design has measurements inside a carry forward-configured report section, this will result in carrying forward the entire section of the report including data from previous studies. This may result in the physician making analysis based on the old measurements. | Software design | Siemens mailed out a "Customer Advisory Notice" #SY009/16/P to customers informing them of the issue and what to do if it was to occur. Siemens informed customers that they are working on the "carry forward" feature enhancement and is planning to release a software update VA10B_HF06 and they should follow the instructions in the advisory notice until the update is performed. |  | 84 | 84 units | Nationwide: AZ, CA, IL, IN, MA, MD, MI, MN, NJ, NY, OH, OR, PA, TN, TX, UT, WA, WI | Terminated | March 20, 2017 | 340 | 84 | Computer | Computer
Z-0010-2017 | 75157 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Ingenuity Core Model Number 728321 |  |  | August 24, 2016 | 2016 | Philips Medical Systems | The infant performance test did not meet the test specification for the Infant Head Uniformity test with High Resolution. As a result, when the user performs High Resolution head scans, the reconstructed images may exhibit degraded image quality manifested as non-uniformity of the image, CT number shift greater than 5 Hounsfield units, and reduction in gray/white matter differentiation. The issue is most sever in the High Resolution mode, and the factory reference pediatric scan protocols are set to High Resolution default. Adults head scans have a factory reference protocol that defaults to standard resolution which doesn't experience the issue. | Software design | Philips sent an Urgent Field Safety Notice (FSN) CLE 16-045 letter dated August 26, 2016, via a certified letter to customers. The letter identified the affected product, problem, actions to be taken by customer/user and actions planned by Philips to correct the problem. The letter informed customers a Philips field service representative will install a software update on the affected systems. For questions contact your local Philips representative or local Philips Healthcare office. For North America and Canada, contact the Customer Care Solutions Center (1-800-722-9377, follow the prompts).. |  | 358 | 358 | Worldwide Distribution - US Nationwide in the states of:: AK, AL, AR, AZ, CA, CO, CT, DE, GA, IA, IL, IN, FL, KS, KY, LA, MA, MD, ME, MI, MO, MS, MT, NC, ND, NE, NJ, NH, NM, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, VT, WA, WI, WV, &amp; WY. and the countries of Algeria, Argentina, Australia, Bangladesh, Belgium, Bolivia, Brazil, Cambodia, China, COlumbia, Cuba, Denmark, Dijbouti, Dominican REpublic, Egypt, Estonia, Finland, France, Germany, Gilbraltar, Greece, Hong Kong, Hungary, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Korea, Latvia, Lebanon, Libya, Lithuania, Malaysia, Martinique, Monaco, Mongolia, Mozambique, Myanmar, Netherlands, New Zealand, Norway, Palestine, Panama, Peru, Philippines, Poland, Portugal, Qatar, Reunion, Russian Federation, Senegal, Singapore, Slovakia, South Africa, Spain, Sweden, Seitzerland, Taiwan, Tanzania, Thailand, Tunisia, Turkey, Uniter Arab Emirates, United Kingdom, Uzbekistan, and Vietnam. | Open, Classified |  |  | 1302 | Computer | Computer
Z-2056-2016 | 63990 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Symbiq Two Channel Infuser; an Rx medical device infusion pump used to administer I.V. fluids |  |  | February 04, 2013 | 2013 | Hospira Inc. | Frequent Proximal Occlusion Alarms (POAs) occurring on Symbiq devices with list numbers 16026 and 16027. When there are false alarms, the pump will continue to alarm even though there is no occlusion present. False POAs could result in a delay/interruption in therapy, and require the clinician to reset the alarm multiple times or to replace the pump. | Software design | Hospira sent letters on May 2, 2013, to customers informing them of the discontinuation of sales, service, and support for the Symbiq pump were sent via traceable mailing to all affected customers. For questions regarding this recall call 224-212-4892. |  | 15285 | 15,285 | Worldwide Distribution - USA (nationwide) | Completed |  |  | 35596 | Computer | Computer
Z-2739-2016 | 64761 | No | LMB | Medical image storage device. | Radiology | Radiology | 510(K) Exempt | Centricity Enterprise Archive is a software product for receiving, archiving and sending of medical data. Qualified system administrators install, monitor and maintain the system. DICOM devices (e.g. modalities, workstations) communicate with the archive using the DICOM protocol (published by ACR-NEMA). The clinical use is limited to the interaction other systems have with the Centricity Enterprise Archive. Centricity Enterprise Archive is not intended for use in patient vicinity. The intended users for Centricity Enterprise Users service personnel and System administrators, and non clinical users. As of V4.0 XDS enabled systems can communicate with the archive using the XDS and XDS-I profiles (published by IHE). |  |  | November 16, 2012 | 2012 | Ge Healthcare It | There is a potential for data loss associated with empty path names resulting from HL-7 updates when the cache disk is almost full. In an unlikely event that the cache disk is almost full (&lt;100MB) and HL-7 update events are received, the update action fails and an empty path name is written to the EA database. As a result the instance is not deleted from the storage device but no longer referenced. When trying to retrieve the instance an error is returned. | Software design | URGENT MEDICAL DEVICE CORRECTION letters dated November 16, 2012 and April 19, 2013, were sent to all affected customers.The letters requested that customers verify that all Enterprise Archive partitions have enough free space available and that the cache disk has at least 10% free space available. The letters further requested that customers whose disks are not being purged to the configured watermarks and the cache disk is almost full to contact their local GE service person. A permanent product correction (Service Pack 12 and release notes) will be provided by GE Healthcare IT. Customers with questions can contact their GE service representative or the GE Healthcare IT Remote Online Center Customer Care Center at 1-800-437-1171. |  | 1379 | 1,379 units | Nationwide Distribution | Terminated | January 11, 2017 | 1517 | 1379 | Computer | Computer
Z-2832-2016 | 74991 | No | DTN | Cardiopulmonary bypass blood reservoir. | Cardiovascular | Cardiovascular | 510(k) | The Bigger-Better-Bladder, TM 2 With 3/8"ID tubing, 72" long tubing, ITEM BBB38-72 |  |  | July 01, 2016 | 2016 | Circulatory Technology Inc | The seal between the balloon and the housing may leak. | Device Design | The firm, Circulatory Technology Inc., sent an amended "Urgent: Expanded Device Recall" letter dated July 1, 2016, and response form to their customers via email. The letter described the product, problem and actions to be taken. The customers were instructed to immediately inspect their inventory, quarantine and prepare to ship the product back to Circulatory Technology Inc. Customers were also instructed to complete and return the response form via email to ytamari@cirtec.com. Customers with questions can call 516-624-2424, Monday through Friday 9:00 AM - 7:00 PM, EST. |  | 20 | 20 units | US Distribution to states of: LA, MI, NY, OH, OR &amp; TX. | Open, Classified |  |  | 20 | Computer | Not_Computer
Z-2830-2016 | 74712 | No | EZL | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Option-vf Urinary Catheter. Female use only. Latex Free. Sterile, Rx only. |  |  | July 11, 2016 | 2016 | C.R. Bard, Inc. | During an FDA inspection it was found out that the Practical Foley Catheters to be potentially nonsterile. | Device Design | C.R. Bard sent an Urgent - Medical Device Product Recall letter dated July 8, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The firm notifies customers of the recall, provides a clinical risk statement, and gives instructions regarding product disposition. Facilities are instructed to examine their inventory and quarantine any recalled product. The firm requested that customers complete the Recall &amp; Effectiveness Check Form if product is or is not in inventory. If product was further distributed, customers should be forwarded the recall notification letter and Recall&amp; Effectiveness Check Form.If you or the patient using these catheters has had an adverse event related to the recalled catheters, please contact Bard Medical Division Field Assurance at 1-800-526-4455 (option 5, then option 4) or via email at BMD.FieldAssurance@crbard. com. |  | 274 | 274 units | US Distribution to the states of : AL, CA, CO, FL, IL, IN, MA, MD, NJ, NY, OR, PA, TN, WA, and WI. | Terminated | March 22, 2017 | 254 | 274 | Computer | Not_Computer
Z-2802-2016 | 74910 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | RAPIDLab 1260 Blood Gas Analyzer |  |  | August 11, 2016 | 2016 | Siemens Healthcare Diagnostics Inc | There is a potential for the first and/or last name of one patient to be printed with Patient ID and result data from a different patient, even though those fields have been turned Off for the system | Software design | Siemens Healthcare issued Urgent Field Safety Notice to (US) customers via Federal Express. . All affected customers outside the US provided a copy of an Urgent Field Safety Notice via hard copy, e-mail, and/or fax as determined by each countries local regulations and procedures.This notice informs customers of the issue, Actions to be Taken by the Customer and to return the response form.Questions contact your Siemens Customer Care Center or your local Siemens technical support representative |  | 114 | 114 units | NationwideForeign:AlbaniaAlgeriaAngolaArgentinaArmeniaAustraliaAustriaBahrainBangladeshBelarusBelgiumBosnia Herzeg.BotswanaBrazilBruneiBulgariaCanadaChileColombiaCroatiaCzech RepublicDenmarkEcuadorEgyptEstoniaFinlandFranceFren.PolynesiaGeorgiaGermanyGreeceHong KongHungaryIcelandIndiaIndonesiaIraqIrelandIsraelItalyIvory CoastJapanKazakhstanKenyaKuwaitLatviaLebanonLesothoLithuaniaLuxembourgMacedoniaMalaysiaMexicoNetherlandsNew CaledoniaNew ZealandNorwayP.R. ChinaPeruPhilippinesPolandPortugalQatarRep. of YemenRepublic KoreaRomaniaRussian Fed.Saudi ArabiaSerbiaSingaporeSlovakiaSloveniaSouth AfricaSpainSwedenSwitzerlandTaiwanThailandTunisiaTurkeyTurkmenistanU.A.E.United KingdomUruguayVatikancityVietnam | Open, Classified |  |  | 13245 | Computer | Computer
Z-2796-2016 | 74815 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | RadSuite software. The firm name on the label is Merge Healthcare, Hartland, WI. |  |  | May 06, 2013 | 2013 | Merge Healthcare, Inc. | Multiple studies were missing at a customer location that never made it to the archive before watermarking | Software design | The recalling firm first issued a letter via email dated May 6, 2013. A letter dated January 28, 2016, was also issued via email. Both letters informed the customer a software upgrade was available and of the issue and makes recommendations about using the software. They are also informed an upgrade is available to correct the issue. The notification instructs the customer to pass the information onto all other users, as well as any downstream customers if it was further distributed. In addition, the customer is requested to call Customer Service or send them an email to schedule the upgrade. If the customer chooses not to upgrade, they are requested to complete the attached form or go online and complete it.For further questions please call (800) 724-5970 Ext. 8027. |  | 45 | 45 sites potentially have the affected versions | Distribution was nationwide to medical facilities. Government distribution was made to LA and OK. | Open, Classified |  |  | 45 | Computer | Computer
Z-2062-2016 | 66372 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | *** 1) Product labeled in part: 19683-28; LIFESHIELD; MICRODRIP SOLUSET WITH 1.2 MICRON FILTER 150ML BURETTE PLUMSET, CONVERTIBLE PIN, CLAVE SECONDARY PORT, OPTION-LOK, NON-DEHP; NOMINAL LENGTH: 118 IN; PRIMING VOLUME: 14 ML; |  |  | September 09, 2013 | 2013 | Hospira Inc. | Intravenous fluid may leak at the CLAVE secondary port of Hospira PlumSet IV administration sets. | Device Design | Hospira sent an URGENT MEDICAL DEVICE CORECTION letters dated September 9, 2013 to customers. The letter identified the affected product, problem and actions to be taken. the letter included additional instructions on the proper handling/use of the CLAVE devices. The letters request that all incidents of separation or breakage of CLAVE from the secondary port of the cassette be reported to Hospira by calling 1-800-441-410 (Monday - Friday, 8:00 AM - 5:00 PM CST) or by e-mail at ProductComplaintPP@hospira.com. Customers are further requested to complete and return a response form included wih the URGENT MEDICAL DEVICE CORRECTION letter and to notify any customers of this advisory. Hospira will also be supplementing the instructions for CLAVE use as an interim measure until a new method for bonding the CLAVE to the secondary port of the Plum cassettes can be devised and implemented. |  | 3920932 | 3,920,932 units in total | US (Nationwide) Distribution | Completed |  |  | 10212296 | Computer | Not_Computer
Z-2790-2016 | 74903 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Hospira MedNet Medication Management Suite software, version 6.1, List Numbers 16037-64-02 and 16037-64-03, in combination with the SapphirePlus 13.1x infusion pump |  |  | July 28, 2016 | 2016 | Hospira Inc., A Pfizer Company | Hospira MedNet 6.1 software, in combination with the SapphirePlus 13.1x infusion pump, may result in incorrect bolus amount calculations for drugs whose concentrations are listed in Million Units. | Software design | Hospira sent an Urgent Medical Device Correction letter dated July 28, 2016, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The letter provided instructions for step-by-step process to mitigate the issue. Customers with questions were instructed to contact Hospira Global Complaint Management at 1-800-441-4100 or Hospira Technical Service Center at 1-800-241-4002, option 4. |  | 20 | 20 | Worldwide Distribution - US including ND, TX, MI, MN, WI, IA, ME, HI, PA, MA, NJ, NY, CT, and Internationally to Canada. | Open, Classified |  |  | 20 | Computer | Not_Computer
Z-1584-2016 | 73581 | No | MQJ | Assisted reproduction micromanipulators and microinjectors. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(K) Exempt | Integra 3; Model Numbers: 6-54-100, 6-54-110, 6-54-150, 6-54-160 |  |  | March 11, 2016 | 2016 | RESEARCH INSTRUMENTS LTD - SITE B | The units have been fitted with a potentially faulty power supply cable. | Device Design | Research Instruments sent a recall notification dated 03/11/2016, via UPS with confirmed delivery receipt, to all affected customers. The letter identified the affected product, problem and the action needed to be taken by the customer. The letter asked customers to contact Research Instruments First Line Support if they had any questions. |  | 380 | 380 units | Worldwide Distribution - US Nationwide in the states of: MA, NY and countries of : Algeria, Australia, Austria, Bahrain, Belgium, British Virgin Islands, Bulgaria, Canada, Denmark, Egypt, Estonia, Finland, France, Georgia, Germany, Greece, Hong Kong, India, Iran, Iraq, Ireland, Israel, Italy, Japan, Jordan, Korea, Kuwait, Libya, Lithuania, Macedonia, Malaysia, Malta, Mexico, Nigeria, Norway, Pakistan, Peru, Poland, Romania, Serbia, Singapore, Slovakia, South Africa, Spain, Sri Lanka, Sweden, Taiwan, Thailand, Tunisia, Turkey, UAE, Ukraine, United Kingdom, and Vietnam, | Open, Classified |  |  | 380 | Computer | Computer
Z-1585-2016 | 73745 | No | IXR | Radiologic table. | Radiology | Radiology | 510(K) Exempt | X-Ray Diagnostic Table, Model No. YSF-300 |  |  | March 31, 2016 | 2016 | Shimadzu Medical Systems | Report of unintentional movement of table to reverse tilting. | Device Design | Shimadzu sent an Urgent Voluntary Medical Device Recall Notice to all affected customers on April 8, 2016. Shimadzu Corporation has identified a potential issue in the Shimadzu X-ray Diagnostic Table. The letter informs the customers of the problems identified and the actions to be taken. Customers with questions or concerns are instructed to contacted Shimadzu National Technical Support at (800) 228-1429.Shimadzu will send out an updated customer notification letter to inform customers that the liquid penetration may cause an electrical short circuit, causing the table to unintentionally tilt in reverse. If this unintentional reverse tilting occurs, there is a potential risk of injury to the patient and/or the user.For questions regarding this recall call 310-217-8855, ext 139. |  | 185 | 185 units | Nationwide Distribution | Terminated | January 02, 2017 | 277 | 185 | Computer | Not_Computer
Z-1098-2016 | 73421 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Intra operative MRI Systems: Neuro II-SE , Neuro III-SV, iMRI 1.5T A, and iMRI 3T S. |  |  | February 10, 2016 | 2016 | Deerfield Imaging | Emergency helium venting lines (i.e. quench lines) may not have been properly installed. | Device Design | Consignees were sent a IMRIS "Urgent Field Safety Notice" letter dated February 10, 2016. The letter described the problem and the product involved in the recall. The firm also stated that it is important that customers review the portion of the quench line vent described and ensure that it is installed to the specifications supplied as part of the Planning Guide. Customers were reminded that it is important to have an emergency plan in place in the event Helium gas escapes into the magnet room or any other area. For questions customers can contact their local customer support hotline or +1-866-475-0525. |  | 49 | 35 USA, 14 OUS | Worldwide Distribution -- US, including the states of CA, CT, FL, GA, IL, IN, MD, MA, MN, MO, NH, NC, OH, OR, PA, TX, UT, VA, and WI; and ,countries of CHINA, CANADA, AUSTRALIA, FRANCE, JAPAN, and GERMANY. | Terminated | April 05, 2017 | 420 | 49 | Computer | Not_Computer
Z-2775-2016 | 74837 | No | DOP | Mass spectrometer for clinical use. | Clinical Chemistry | Toxicology | 510(K) Exempt | Endura MR Mass Spectrometer u using software versions Endura MD Software 1.0 and 1.0 SP1. |  |  | July 19, 2016 | 2016 | Thermo Finnigan LLC | Thermo Fisher has determined that the Endura MD mass spectrometer instrument control software versions 1.0 and 1.0 SP! have a software defect which affect data accuracy. | Software design | ThermoFisher Scientific sent an Urgent Medical Device Field Correction letter dated July 19, 2016, to all affected customers via FedEx. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to discontinue use of the affected product until a company authorized Field Service Engineer installs corrected software to the instrument system. Customers were also instructed to communicate to all who operate the instrument and to complete the provided Acknowledgment Form, sign, and return it via email FieldActionTeam@thermofisher.com. Customers with questions should contact the Field Action Team at FieldAction@thermofisher.com. |  | 6 | 6 (5 in Us distribuion and 1 in Japan) | Worldwide distribution - US including FL, NY and Internationally to Japan. | Open, Classified |  |  | 6 | Computer | Computer
Z-2088-2016 | 74390 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo.plaza, Picture archiving and communication system (PACS), Model Numbers - 10863171, 10863172, 10863173, |  |  | May 09, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Software error in previous software versions in which two references for the same image may exist in the database. | Software design | Siemens sent an Customer Information letter dated May 10, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letter with software instruction to customers informing them about an upcoming software VB10B under(SY003/16/P) update change and issue that the new software version will address. For further questions, please call (610) 448-6478. |  | 74 | 74 systems | US Distribution to the states of : AL, CA, CT, FL, IN, KS, LA, MA, MD, MI, MN, MO, MS, NE, NY, OH, OR, PA, TX, WI | Terminated | March 20, 2017 | 315 | 74 | Computer | Computer
Z-1172-2016 | 73253 | No | LFR | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Accu-Chek Inform II Base Unit as a part of the Accu-Chek Inform II Blood Glucose Monitoring System |  |  | February 11, 2016 | 2016 | Roche Diagnostics Operations, Inc. | Accu-Chek Inform II Base Unit might produce physical transmission errors in the form of data loss in the communication between the meter and the Data Management Systems (DMS). The issue can lead to the data loss or in the worst case to an erroneous assignment of the patient data (patient mismatch). The issue will only occur at sites using POTCT1-A communication via USB. | Software Design Change | Roche sent an URGENT MEDICAL DEVICE RECALL CORRECTION letter dated February 11, 2016, via UPS Ground (receipt required) to the consignee. The Consignee instructions were as follows:Âż If currently using USB connectivity, follow the instructions outlined in this Urgent Medical Device Correction (UMDC) to determine if your base unit is affected by the issue.Âż Until the new software is available, do not use USB connectivity if the base unit is affected by this issue; switch the connectivity method to Ethernet (TCP/IP) or Wi-Fi.Roche is not requesting return of affected product. Any returned product will be managed according to current procedures.Consignees with questions were instructed to contact the ACCU-CHEK Customer Care Service Center at 1-800-440-3638.For questions regarding this recall call 317-521-3911. |  | 97529 | US 5,604 devices, OUS 91,925 | Nationwide Distribution including AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, SC, TN, TX, UT, VA, WA, WI, and WV. | Terminated | May 30, 2017 | 474 | 97529 | Computer | Computer
Z-1171-2016 | 73129 | No | LDP | Acetaminophen test system. | Toxicology | Toxicology | 510(k) | Sekure Chemistry Sekisui Diagnostics Acetaminophen L3K Assay, In-vitro Diagnostic |  |  | January 22, 2016 | 2016 | Sekisui Diagnostics P.E.I. Inc. | Crystal formation in the Color Reagent (R2) when left on board automated analyzers causing samples to under recover | Device Design | Sekisui sent an Important Field Correction - Customer Notification letter dated January 22, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.1. Discard all reagent in house with a current expiration date of 2016-09.2. Update the expiration on all remaining kits on hand with the new expiration dates provided 3. When adding additional reagent to the analyzer, a new reagent wedge should be used.4. If crystals are observed within reagent stored on-board, reagent should be discarded and replaced with fresh reagent.5. Complete, sign and return the attached CONFIRMATION OF NOTIFICATION form by FAX to the number provided. This will indicate receipt of this Field Correction, act as confirm action of compliance with the actions,and allow for issuance of replacement material for expired product on hand (lots 47677, 47694, 47695).Questions contact Technical Services at 1-800-565-0265 or viaemail at PEIDiagnosticTechnical@sekisui-dx.com |  | 14407 | 14,407 kits. | Worldwide Distribution - US Distribution to the states of : AZ, CA, FL, IA, ID, IL, MA, , MN, MT and ND., and to the countries of Canada, France and Hong Kong. | Open, Classified |  |  | 14407 | Computer | Not_Computer
Z-2769-2016 | 74649 | No | BZC | Pulmonary-function data calculator. | Anesthesiology | Anesthesiology | 510(k) | Innocor Cardiopulmonary Exercise and Lung Function Testing System |  |  | May 13, 2016 | 2016 | Innovision A/S | The purpose of the recall is to replace data transmission cable (USB-cable) connecting the lnnocor device with an external computer.The replacement USB cable will provide galvanic isolation of the lnnocor from the external computer via opto-couplers, as opposed to the grounding protection currently in use.The replacement is intended to prevent risk of electric shock. | Device Design | Recall notifications described the issue and actions to be taken by the firm. The notice also stated that replacement costs would be carried by Innovision ApS. Acknowledgment of the notice can be made by contacting service manager Knud Pedersen at kp@innovision.dk or (+45) 65 95 91 00) to schedule the service activity and to answer any questions. |  | 8 | 8 | Worldwide Distribution -- US, to Massachusetts and Washington, DC; and, the countries of Canada, Germany, and United Kingdom. | Open, Classified |  |  | 8 | Computer | Computer
Z-1572-2016 | 73886 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON X700 Ultrasound System with software version 1.0.04. |  |  | September 05, 2015 | 2015 | Siemens Medical Solutions USA, Inc. | Due to a communication error between the software and V5Ms transducers rotation function, ACUSON X700 ultrasound systems at software versions 1.1.04, display an IMG_15 error message causes the system to lock up, which requires a reboot to recover system operation. | Software design | September 5, 2015 we initiated a correction via update program US020/15/S where we visited all four units and updated their software to version 1.0.05 |  | 4 | 4 | Distributed in Germany, U.A.E., Hungary and Brazil. | Terminated | April 27, 2016 | 235 | 4 | Computer | Computer
Z-1176-2016 | 73422 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | The Spirit TM Select bed is a Med-Surg bed intended to support and transport |  |  | February 18, 2015 | 2015 | CHG Hospital Beds Inc | It was identified that in some situations the bed exit alarm may not function as intended. The software code for the bed exit system has the potential to auto-reset erroneously. In some situations, the software code does not allow enough time for the weight value to fall to zero once the patient egresses from the bed, in which case, the bed has the potential to reset the bed exit alarm even without a patient on the bed.The bed exit on the Spirit Select bed may incorrectly reset if the Auto Reset option is selected. A patient fall may result in soft tissue or bone injury which may require medical or surgical intervention to prevent permanent disability. If the bed exit alarm incorrectly resets then medical assistance after a fall or during entanglement may be delayed. Patients with dementia or disorientation are at greatest risk. There have been two reported injuries associated with these issues to date. A Medical Risk Assessment determined that there is a remote probability of occurrence of harm requiring medical intervention. | Software design | On 2/18/2016, URGENT MEDICAL DEVICE RECALL notifications were sent to the affected customers via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. If you have any questions or concerns, please contact Stryker Customer Service (1-800-327-0770). Our normal business hours are Monday-Friday 8 a.m-6 p.m. (EST). |  | 1892 | 1892 | Distributed US (nationwide) and in Canada. | Terminated | May 23, 2017 | 825 | 1892 | Computer | Computer
Z-2753-2016 | 74722 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Axiom REG implants Sterile |  |  | June 20, 2016 | 2016 | Anthogyr | Following a FDA inspection showing that the firm failed to validate the design and process of its implants. | Device Design | The firm, Anthogyr, contacted physicians by phone and sent an " URGENT: MEDICAL DEVICE RECALL" notification by mail, dated 06/20/2016, informing users of a recall due to inadequate validation of design and process of implants by the FDA. The notice described the product, problem and actions to be taken. The physicians were instructed to isolate the parts and return any Anthogyr implants from lots subject to recall to Business Area Manager, 247 SouthWest 28th road, Miami, FL 33129;and complete and return the RETURN RESPONSE form by fax to: +33 (0)4 50 93 78 60 or e.mail: j.duthil@anthogyr.com or mail: Anthogyr- Director of Quality &amp; Regulatory Affairs- 2237 avenue- Andre Lasquin-France. The firm stated that they will not market their products in the United States until they obtain validation through 21 CFR Part 820. If you have any questions, call 404-718-9689 or email: jduthil@anthogyr.com or mfoley@anthogyr.com. |  | 111 | 111 units | US Distribution to sates of: NC, NY, and OR. | Terminated | March 28, 2017 | 281 | 121 | Computer | Not_Computer
Z-1813-2016 | 73984 | No | OBJ | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Volcano s5, s5i, CORE and CORE Mobile systems with software version 3.4 and v3.4 software kits. |  |  | April 25, 2016 | 2016 | Volcano Corporation | Volcano has become aware of an incompatibility issue between Impacted Systems and hospital network scans. In specific circumstances, an impacted System will encounter unexpected data from the hospital network and be forced to reboot | Software in the Use Environment | The firm, Philips, sent an "URGENT MEDICAL DEVICE FIELD CORRECTIVE ACTION" letter dated 4/25/2016 to its customers beginning April 29, 2016 which include a technical service bulletin to explain how to assure the incompatibility issue does not affect system performance. The customers were instructed to do the following:-If you have an Impacted System but it is not connected to a network, this issue does not affect your system. Additionally, if network scans are not performed, this issue does not affect your system.-If you have an Impacted System that is connected to a network and staff performs network scans while the Impacted System is in operation, this issue does affect you and the circumstances exist for a manual reboot to occur during an active procedure. To avoid this, the following corrections must promptly be taken: 1. Disconnect your Impacted Systems from the facilities network. If you need to reconnect it while not being used in a procedure, make sure it is disconnected again before starting any procedure; OR 2. Do not permit your staff to perform network scans while the Impacted Systems are in operation.-Please complete, sign, and return the attached form indicating that you received this Field Corrective Action notification via Fax to Volcano Customer Service at (916) 638-8812 or email to CS@volcanocorp.com.If you have any questions, please call (800) 228-4728, option 5. |  | 5875 | 5875 (2685 in US; 3190 ROW) | Worldwide Distribution: US (nationwide) and countries of: Japan, South Africa, Germany, Sweden, Belgium, France, Italy, Spain, India, Russian Federation, Hungary, Turkey, Poland, Portugal, Costa Rica, Finland, Great Britain, Croatia, India, Chile, Thailand, Mauritius, Iraq, Denmark, Norway, Ireland, Netherlands, Bulgaria, Czech Republic, Luxembourg, UAE, Austria, Poland, Pakistan, Sri Lanka, Kazakhstan, Slovakia, Saudi Arabia, Unite Arab Emirates, Egypt., Jordan, Kuwait, Lebanon, Bulgaria, Romania, Latvia, Lithuania, Bosnia and Herzegovina, Greece, Malta, Cyprus, Romania, Tunisia, Kenya, Algeria, Israel, Jordan, Hong Kong, Republic of Korea, Malaysia, New Zealand, Venezuela, Brazil, Philippines, Australia, Taiwan, Peru, Singapore, Panama, Mexico, Argentina, Costa Rica, Dominican Republic, China, and Paraguay, | Open, Classified |  |  | 5875 | Computer | Computer
Z-2758-2016 | 74395 | No | IYX | Scintillation (gamma) camera. | Radiology | Radiology | 510(k) | e.cam and Symbia E gamma camera systems with patient beds |  |  | August 05, 2016 | 2016 | Siemens Medical Solutions USA, Inc. | "We have received reports of patients hair being caught in the e.cam and Symbia E patient bed up/down drive. The purpose of this letter is to remind users of the patient positioning instructions for your e.cam or Symbia E system and to inform you of an update that will be performed by your local Siemens service representative." | Device Design | The firm, Siemens, sent a "Customer Advisory Notice" letter on 8/5/2016 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to review the "safety" sections of the user manual with your staff; continue using your system under the guidelines set forth in the letter; place advisory in the system's instructions for use; and if you have sold this equipment and it is no longer in your possession, forward the notice to the new owner and inform us about the new owner of the equipment.Siemens has designed an enhancement kit to further reduce the gap in the patient bed up/down drive. Your local Siemens service organization will be contacting you to schedule this update.If you have any questions regarding this important notice, contact your local Service representative at the following numbers: America: 1-800-888-7436; Europe, Middle East, and Africa: +49 9131 940 4000 or Asia and Australia: +86 (21) 3811 2121. |  | 5451 | 5451 | Worldwide Distribution: US (Nationwide) including states of: District of Columbia and Puerto Rico and countries of: Algeria, Argentina, Austria, Australia, Bangladesh, Bahrain, Belarus, Belgium, Brazil, Cameroon, Canada, Chile, China, Colombia, Congo-The Democratic Republic of the, Costa Rica, Croatia, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, El Salvador, Finland, France, Germany, Great Britain, India, Iran, Ireland, Israel, Italy, Japan, Kazakhstan, Korea, Kuwait, Lebanon, Mauritius, Malaysia, Mexico, Morocco, Myanmar, Nepal, Netherlands, Nigeria, Norway, Oman, Pakistan, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Spain, Serbia (Republic of Serbia), South Africa, Sweden, Switzerland, Syrian Arab Republic, Taiwan, Thailand, Trinidad and Tobago, Tunisia, Turkey, Uganda, United Arab Emirates, Uzbekistan, Viet nam, and Yemen. | Open, Classified |  |  | 5451 | Computer | Not_Computer
Z-1182-2016 | 73452 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | McKesson Horizon Medical Imaging (HMI) products versions 4.6.1 to including 11.9 and McKesson Radiology (MR) products 12.0 and 12.1.1. |  |  | February 18, 2016 | 2016 | Mckesson Medical Immaging | McKesson has identified a design deficiency where under rare circumstances, imported images/studies may re-use a non-unique image directory. This issue may cause incorrect images to be displayed for a patient. | Software design | McKesson sent an Urgent Field Safety Notice dated February 18, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Immediate actions to take to reduce the risk to patientAll customers receiving this safety notice must contact McKesson Support as soon as possible at 1-800-663-2533 US/Canada and International Toll Free Radiology - 00 800 626 20009:1) To acknowledge that they have read and understood this Field Safety NoticeAnd2) Arrange installation of software update/upgrade.This notice should be distributed to all personnel within your organization who need to be aware of this issue and the required actions, and alert other affiliated parties that are affected by this Field Safety Notice.For further questions, please call (404) 338-3556 |  | 631 | 631 devices | Worldwide Distribution - USA (nationwide) including Puerto Rico and to the states of : AL, AK, AZ, AR, CA, CO, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OR, PA, RI, SC, TN, TX, VT, VA, WA, WV, WI, WY., and the countries of : Guam, Great Britain, France, Canada, Ireland and Australia | Open, Classified |  |  | 631 | Computer | Computer
Z-1183-2016 | 73479 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Siemens Syngo Dynamics-a Picture Archiving and Communication System (PACS) Model Numbers: 10091805, 10091807, 10091637, 10091673 |  |  | January 25, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Siemens' conducting a recall due to a potential issue when using the measurement package of the VA10 version of syngo Dynamics. | Software design | The firm, Siemens Healthcare, sent a "Customer Advisory Notice" dated 1/19/2016 to customers. The notice described the product, problem and actions to be taken. The customers were instructed to follow the instructions provided on what steps the user can take to avoid the issue and to forward the notice to the new owner, if you sold this device/equipment and its no longer in your possession. Siemens is preparing a modification of syngo Dynamics that will resolve this potential malfunction.If you have any question, contact Regulatory Technician Specialist at 610-448-6478 or email: Anastasia.sokolova@siemens.com. |  | 228 | 228 units | US Nationwide Distribution including states of: MN, NC, NE, FL, TX, CA, AR, MI, MA, IL, DE, VA, SC, OH, CO, WI, MO, NY, CT, UT, PA, WA, NJ, GA, IN, MD, TN, LA, OR, AZ, ND, KS, AL, NM. | Open, Classified |  |  | 228 | Computer | Not_Computer
Z-1186-2016 | 73352 | Yes | DWC | Cardiopulmonary bypass temperature controller. | Cardiovascular | Cardiovascular | 510(k) | Stockert Heater-Cooler 3T; Item Number 16-02-8, H3T/240V/60Hz ; 16-02-82, H3T208V/60Hz ; 16-02-85, H3T/120V/60Hz. Device Listing E572259 |  |  | August 15, 2012 | 2012 | Sorin Group USA, Inc. | On August 15, 2012, Sorin informed consignees via letter of changes made to create Version 12 of the Heater-Cooler 3T IFU. The revised IFU included enhancements to achieve water quality that meets the standard for German drinking water. | Device Design | Sorin Germany, via Sorin Colorado, distributed an important information letter to its US customers on August 15, 2012 informing them about an update of the cleaning and infection procedure for the U.S. market that were included in Revision 12 of the IFU. |  | 1125 | 1125 units | Worldwide Distribution - US Nationwide, VA/govt/military | Terminated | May 26, 2017 | 1745 | 1125 | Computer | Not_Computer
Z-2738-2016 | 73833 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | PINNACLEÂż Introducer Sheath |  |  | March 16, 2016 | 2016 | Terumo Medical Corporation | Terumo Medical is initiating this recall due to complaints of valve leakage that have been related to a manufacturing issue. If affected introducer sheaths are used, there is possible risk of incremental blood loss through the valve. | Device Design | Terumo sent an Urgent Product Recall Bulletin dated March 7, 2016, to all affected customers. All affected customers will be notified regarding the required recall actions (stop using/quarantine the affected products). TMC will strive for 100% verification to confirm that all customers have acknowledged the recall. Follow-up communication is being sent specifying the mechanisms to be used for labeling and returning the involved product from the customer level.We regret any inconvenience this action may cause your operations and we appreciate your understanding. Please contact us if you need assistance procuring additional inventory. We are committed to continuing to offer products that meet the highest quality standards that you expect from Terumo Medical Corporation. If you have questions or need further information, please call Customer Care at 1-800-888-3786 or contact your local sales representative.For further questions regarding this recall please call (410) 392-7226. |  | 1370995 | 1,370,995 eaches have been distributed to the field | Nationwide Distribution | Open, Classified |  |  | 1370995 | Computer | Not_Computer
Z-2108-2016 | 74393 | No | CDP | Total triiodothyronine test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access Free T3 Assay, Ref No. A13422 |  |  | June 08, 2016 | 2016 | Beckman Coulter Inc. | Demonstrate an upward shift in patient results. | Device Design | Beckman Coulter sent an " Urgent Medical Device Recall" letter dated June 8, 2016, to all affected customers. Beckman Coulter has determined through customer feedback and internal testing that the Access Free T3 Reagent lots listed above demonstrate an upward shift in patient results. The preliminary results indicate this upward shift may be related to a June 2015 formulation design change that was introduced to improve the Access Free T3 open reagent pack stability. The upward shift in patient test results, therefore, is expected to be maintained for all future lots. The letter informs the customers of the impact and the actions to be taken. Customers are instructed to complete and return the enclosed response form within 10 days. Customers with any questions are instructed to contact Customer Technical Support Center at http:www.beckmancoulter.com, call (800) 854-3633 in US and Canada. Customers outside US and Canada, are instructed to contact local Beckman Coulter representative. |  | 205744 | 205,744 kits total (13,358 kits in US) | Worldwide Distribution - US Distribution and to the countries of : Algeria, Australia, Austria, Bangladesh, Belgium, Bolivia, Bosnia and Herzegovina, Botswana, Brazil, Brunei Darussalam, Bulgaria, Canada, China, Colombia, Costa Rica, Croatia, Czech Republic, Denmark, Djibouti, Ecuador, Egypt, El Salvador, Estonia, France, Germany, Ghana, Gibraltar, Greece, Guam, Honduras, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Kuwait, Lebanon, Lithuania, Macao, Malaysia, Mayotte, Mexico, Morocco, Namibia, Netherlands, New Zealand, Nigeria, Norway, Oman, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Puerto Rico, Qatar, Romania, Russian Federation, Rwanda, Saudi Arabia, Serbia, Sierra Leone, Singapore, Slovakia, South Africa, Spain, Swaziland, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Uruguay, Venezuela, Viet Nam, Zambia, Zimbabwe | Terminated | April 18, 2017 | 314 | 205744 | Computer | Not_Computer
Z-1091-2016 | 72486 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | BladderScan BVI 9600, Model Numbers: 0270-0451, 0270-0452, 0270-0754, and U270-0451. |  |  | November 30, 2015 | 2015 | Verathon, Inc. | The firm is providing customers with an updated Operations and Maintenance Manual for each of the BladderScan BVI 9600 and AortaScan AMI 9700 devices. The updated Operations and Maintenance Manuals clarify that these devices should not be used for the screening, detection, or diagnosis of abdominal aortic aneurysms (AAAs). | Device Design | Verathon sent an URGENT: MEDICAL DEVICE CORRECTION BLADDERSCAN BVI 9600 WITH AORTASCAN MODE &amp; AORTASCAN AMI 9700 letter dated November 30, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Verathon will manage other international consignees outside Europe internally and follow a similar approach as the firm requested Stericycle to perform. Verathon will instruct the companys foreign distributors to communicate the Customer Correction Letter and provide the revised OMMs to their international customers.This current Medical Device Correction is intended to reinforce the message in the May 2014 communication regarding the potential variance in measurement and to provide you with updated Operation and Maintenance Manuals that further reinforce the devices intended use and indications for use and that revise warning information, including warnings specified in the May 2014 communication.Customers are instructed to destroy all existing Operations and Maintenance Manuals and replace them with updated versions and complete the information on the Medical Device Correction Reply Form and fax or email the completed form within 10 business days to: Fax: (877) 907-9956 or Email: verathon3108@stericycle.com.Customers with any questions can contact your Verathon representative or Verathon Customer Care at (866) 875-0382. Verathon has live agent support on Monday through Friday from 8am to 5 pm Pacific Time.Customers should report any suspected malfunction or adverse event related to any AortaScan AMI 9700 devices and any BladderScan BVI 9600 devices to Verathon Customer Care at (800) 331-2313 |  | 24230400 | 2524 units total (1441 units of BVI 9600 and 820 units of AMI 9700 in US; 209 units of BVI 9600 and 54 units of AMI 9700 OUTSIDE the US | Worldwide Distribution - US (nationwide) and to the countries of : Guam, United Arab Emirates, Australia, Belgium, Canada , China, Czech Republic, France, Germany, Luxembourg, Netherlands, Norway, Saudi Arabia, Sweden , Taiwan, and United Kingdom | Open, Classified |  |  | 24230400 | Computer | Not_Computer
Z-1092-2016 | 73393 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Spacelabs Healthcare Xhibit Central Station, Model 96102 is used to provide clinicians with central monitoring of patient data for those patients connected to networked Spacelabs Healthcare patient monitors and telemetry transmitters. |  |  | February 26, 2016 | 2016 | Spacelabs Healthcare Inc | The firm has received one report of values for patient height and weight being switched when input at the Xhibit Central Station, Model 96102, causing a bedside monitor Body Surface Area (BSA) calculation to be in error. | Software in the Use Environment | Spacelabs sent the letter Urgent Medical Device Correction -Spacelabs Healthcare Inc. - Xhibit Central Station, Model 96102, dated 26 February 2016 to US consignees.The firm will send a notification letter to international consignees on March 4, 2016.Consignees who have any questions about this corrective action, please contact Spacelabs at 1-800-522-7025 and select 2 for Technical Support. |  | 1075 | total 1075 units (614 units in the US and 461 outside the US) | distributed in AL, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KY, LA, MA, MI, MN, MO, MS, MT, NC, NJ, NY, OH, OK, OR, PA, SC, SD, TN, TX, VA, WA, WY and Guam and in the following countries AFGHANISTAN, ARGENTINA, AUSTRALIA, BAHRAIN, BOLIVIA, BRAZIL, BRUNEI DARUSSALAM, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, DOMINICAN REPUBLIC, EL SALVADOR, FRANCE, GERMANY, INDIA, IRAN (ISLAMIC REPUBLIC OF), ITALY, KUWAIT, MALAYSIA, MEXICO, MOROCCO, NETHERLANDS, PALESTINE, PANAMA, PARAGUAY, PERU, PHILIPPINES, POLAND, QATAR, ROMANIA, SAUDI ARABIA, SINGAPORE, SWITZERLAND, TAIWAN, THAILAND, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, and UNITED KINGDOM. | Open, Classified |  |  | 1075 | Computer | Not_Computer
Z-2050-2016 | 73608 | No | DTN | Cardiopulmonary bypass blood reservoir. | Cardiovascular | Cardiovascular | 510(k) | Bigger-Better-Bladder With 3/8" ID tubing, ITEM BBB38 |  |  | March 03, 2016 | 2016 | Circulatory Technology Inc | The seal between the balloon and the housing may leak. | Device Design | Circulatory Technology Inc. sent an Urgent Product Recall notices dated March 3, 2016, to their customers via email. Circulatory Technology Inc. has identified this action as a voluntary recall, however they did not withdraw the product from the market. Instead, Circulatory Technology Inc. has advised customers to check the integrity of the seal between both the inlet and outlet tube and the housing as per the Instructions For Use. In addition, for this lot, Circulatory Technology Inc. recommends that you bend the inlet as well as the outlet tube away from the housing before conducting the leak test. There should be no visible separation between the tubing and the housing. Customers with questions can call 516-624-2424. |  | 72 | 72 units | US Distribution to the states of: CA, IN, MI, NY, PA, TX &amp; VA | Open, Classified |  |  | 72 | Computer | Not_Computer
Z-2085-2016 | 74275 | No | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | Saw Sagittal Catalog Number: ES 300; Saw Sagittal Refurbished Catalog Number: ES300-R; Saw Saggital, Finger Catalog Number: ES310; Saw Saggital, Finger Refurbished Catalog Number: ES310-R; Used to dissect bone in surgical applications. |  |  | May 13, 2016 | 2016 | Medtronic Sofamor Danek Usa, Inc - Dallas Distribution | Medtronic Powered Surgical Solutions testing of the Sagittal Saw Rocker Arm Seal identified a potential for leakage of fluid into the motor during use. | Device Design | Recall notification letters were sent to consignees on 05/13/2016. |  | 368 | 368 devices | Worldwide Distribution-Including the states of AL, AR, AZ, CA, CT, FL, GA, IA, ID, IL, IN, MA, MD, NC, NH, NY, SC, TN, TX, and UT. | Terminated | June 13, 2017 | 396 | 368 | Computer | Not_Computer
Z-2900-2016 | 75111 | No | OYE | Automated differential cell counter. | Hematology | Hematology | 510(k) | AQUIOS CL Flow Cytometer, Catalog No. B30166; Intended for use with in vitro diagnostic flow cytometry applications using up to four florescent detection channels. |  |  | September 06, 2016 | 2016 | Beckman Coulter Inc. | Beckman Coulter is recalling the AQUIOS CL System because the use of the reagent carousel can lead to misalignment of the carousel and damage to the prep probe. | Software Design Change | An Urgent Medical Device Recall letter dated 9/6/16 was sent to customers to inform them that the use of Positions 9 and 10 of the reagent carousel can lead to misalignment of the carousel and damage to the prep probe. The letter informs the customers of the impact of the recall, action to be taken, and the resolution by Beckman Coulter. Customers are instructed to complete and return the enclosed Response Form within 10 days. Customers with any questions regarding this notice, are instructed to contact Customer Support Center: http://www.beckmancoulter.com/customersupport/support; call 1-800-369-0333 in the US and Canada; outside the US and Canada, contact their local Beckman Coulter Representative. |  | 148 | 148 units total (12 units in US) | US: IL, PA, NY, GA, WA, FL, SC.Affected geographies: Angola, Australia, Belgium, Botswana, Canada, Chile, Columbia, Czech Republic, Denmark, Finland, France, Germany, Ghana, Israel, Italy, Malaysia, Nambia, Netherlands, Panama, Poland, Romania, Slovakia, South Africa, Spain, Sweden, Switzerland, Uganda, United Kingdom and Zambia | Terminated | August 07, 2017 | 335 | 148 | Computer | Not_Computer
Z-2898-2016 | 74955 | No | PHV |  | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | CareLink iPro Version 1.10, Catalog No. MMT-7340 |  |  | August 05, 2016 | 2016 | Medtronic Inc. | Medtronic MiniMed is recalling the CareLink iPro Therapy Management Software due to a time stamp error. | Software design | Medtronic sent an Urgent Field Corrective Action letter dated August 5, 2016, to affected customers who received the CareLink iPro Therapy Management Software. The letter informed the customers that the recall affects the users of CareLink iPro software who have downloaded reports that include data that was uploaded from certain models of Roche BG meters. Customers were informed that the recall is due to time stamp error and data from certain Roche meters uploaded to CareLink iPro may resulted in incorrect reports. Customers were informed that if they manually enter Logbook data from the Roche ACCU-CHEK meters or upload data from a blood glucose meter other than the Roche meters are not affected by the recall situation. Customers with any questions regarding the recall letter are instructed to contact the Medtronic 24-Hour HelpLine at 1.800.646.4633.For questions regarding this recall call 818-576-4700. |  | 33 | 33 units | Nationwide Distribution to TN, MN, and WA | Open, Classified |  |  | 33 | Computer | Computer
Z-2892-2016 | 75171 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo.plaza VB10A, Picture Archiving and Communication System |  |  | September 06, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Software upgrade to eliminate several issues | Software design | Siemens notified customers via mail on 09/07/2016 informing them of the latest syngo.plaza software change that was available to address several issues. |  | 47 | 47 systems | Distributed Nationwide | Terminated | May 08, 2017 | 244 | 47 | Computer | Computer
Z-1066-2016 | 72325 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Philips X-Ray Systems, Allura Xper with R8.2.16 |  |  | October 01, 2015 | 2015 | Philips Electronics North America Corporation | Upon initiating Fluoroscopy the user may encounter a user message Fluoro failed. | Software design | Philips sent a Field Safety Notice to all affected customers on October 1, 2015. The notice informs customers of the issue and the actions planned by Philips. Philips Field Service Engineers are to install Software R8.2.16.1 which addresses the issue. For further information or support concerning the issue, consignees should contact their local Philips Representative or call the Technical Support Line at 800-722-9377. |  | 196 | 196 devices | US Nationwide in the states of : AK, AL, AZ, CA, FL, GA, HI, IL, IN, KS, KY, LA, MA, MD, MI, MO, MS, NC, NJ, NY, OH, OR, PA, TN, TX, UT, WI, WV and the country of Canada. | Open, Classified |  |  | 196 | Computer | Computer
Z-2783-2016 | 74835 | No | HWE | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | Orthopedic | 510(K) Exempt | Light Adaptor for Small Battery Drive and Small Battery Drive II. |  |  | June 24, 2016 | 2016 | The Anspach Effort, Inc. | There was a potential for the Adaptor and Light Adaptor for Small Battery Drive and Small Battery Drive II to generate excessive internal pressure that may cause the products to burst. | Device Design | An Urgent Notice for this Recall was provided to the consignees on 6/24/2016. They were to review their inventory and return the products with instructions. IF YOU HAVE ANY PRODUCT: To immediately quarantine product so it's not used; ensure all in their facility read the Urgent Notice for the Recall; Call the DePuy Synthes Customer Support to obtain a Return Materials Authorization (RMA) Number; to complete the Verification Section indicating the product subject to recall was located; to return the Verification Section with the product to Product Returns, The Anspach Effort Inc., 4500 Riverside Drive, Palm Beach Gardens, FL 33410; Send a copy of the completed Verification Section to DePuy Synthes, Customer Quality Dept by: FAX (561) 627-2682 or Scan/email: DPYUS-PowerToolsFieldActions@its.jnj.com;Keep the notice visibly posted for awareness until all products have been returned; maintain a copy of the notice subject to this action and keep a copy for your records. The return documentation acknowledges your receipt of the medical device removal information.IF YOU DO NOT HAVE the identified product: complete the Verification Section of the recall, by checking the box indicating no affected product located; Include your name title, address, telephone #, signature and date; Send a copy of the completed Verification Section to DePuy Synthes Customer Quality Dept by: Fax: (561) 627-2682, or Scan/e-mail: DPYUS-PowerToolsFieldActions@its.jnj.com; This return document acknowledges your receipt of medical device removal information;ADVISORY: For Alternative products: This product does not have any inventory for replacement. Since the recall of this device may have an impact on your surgical schedule, DePuy Synthes Power Tools would like to propose alternative products to assist you with a replacement for the Recalled Adaptor (05.100.024) or Light Adaptor (05.001.108) . The products are the Small Battery Drive II Hand-piece (532.110) can also be operated |  | 1833 | 1,833 devices (05.001.024 and 05.001.108) | Worldwide Distribution - US Nationwide and including the countries of Canada, Austria, Australia, Belgium, Brazil, Chile, China, Colombia, Costa Rica, Czech Republic, Denmark, Finland, France, Germany, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Kuwait, Libya, Malaysia, Mexico, Nepal, Netherlands, New Zealand, Norway, Poland, Portugal, Romania, Russia, Saudi Arabia, Singapore, Slovenia, Slovakia, Spain, South Africa, South Korea, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, and Vietnam, and areas of Middle East, Africa, Asia &amp; Pacific. | Open, Classified |  |  | 1833 | Computer | Not_Computer
Z-1065-2016 | 73177 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Trilogy 100, Trilogy 200, Trilogy O2, Trilogy 202, |  |  | January 18, 2016 | 2016 | Philips Respironics | Software Issue | Software design | Philips Respironics has determined that affected devices may continue to be used in accordance with the device Instructions for Use and the Field Safety Notification that is being issued to customers. Philips Respironics will be providing a software update for affected devices free of charge. Customers will be notified by mail when the software update is available for download.For further questions please call ( 724 ) 387-7500. |  | 104508 | 104,508 units | Nationwide Distribution including Puerto Rico and Internationally to the US Virgin Islands. | Terminated | May 26, 2017 | 494 | 104508 | Computer | Computer
Z-2890-2016 | 75100 | No | FOX | Infusion stand. | General Hospital | General Hospital | 510(K) Exempt | The EKOS cart has the catalog number 700-60101. |  |  | August 22, 2016 | 2016 | EKOS Corporation | Casters on the EKOS carts may be missing washers which can cause premature wear and dislodgement of the casters/wheels leading to failure. | Device Design | The firm, EKOS, sent an "URGENT FIELD SAFETY ACTION" letter dated August 22, 2016 to its customers. The letters were sent to customers informing them of the field safety action, how to identify the affected carts at their facility and instructing customers locate the cart, quarantine it and do not use the cart for any procedures. A representative from EKOS Corporation will be contacting customers in the near future to arrange a date and time when the cart at a facility can be inspected and any affected caster(s) replaced. As the cart is inspected, it will be tagged to easily identify it and confirm it has been inspected and repaired as necessary. Customers will be provided with a copy of the inspection/repair form for the cart they inspect. Please retain that form for records.If customers have any questions, please contact EKOS Customer Service at 1-888-400-3567 or your local EKOS representative. |  | 455 | 455 units | US Nationwide Distribution in states of: AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MT, NC, NE, NH, NJ, NM, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, WI, and WY. | Terminated | December 30, 2016 | 130 | 455 | Computer | Not_Computer
Z-2886-2016 | 75078 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Spacelabs Healthcare Xhibit Telemetry Receiver (XTR), Model 96280. |  |  | August 25, 2016 | 2016 | Spacelabs Healthcare Inc | The firm received multiple reports of Xhibit Telemetry System going offline or locking up. In addition, short-duration asystole alarms may not display an audio and visual alarm as reported from a single Canadian hospital. | Software design | Spacelabs sent a letter dated August 25, 2016 to affected US consignees informing them of the going offline or locking up issue when Xhibit Central Station is used as part of an Xhibit Telemetry System. On September 1, 2016, a customer letter (translated as necessary) will be emailed to all international subsidiaries and distributors of record.Spacelabs sent the second Urgent - Medical Device Correction letter, dated September 15, 2016, to affected US consignees informing them that they have received reports of certain short-duration asystole alarms generated by the Xhibit Telemetry Receiver (XTR) where there may be no audible or visual alarm. On September 22, 2016, a second letter (translated as necessary) will be emailed to all international subsidiaries and distributors of record.Spacelabs Healthcare will contact consignees to schedule a convenient time for Spacelabs to update the recalled Xhibit Telemetry System at no cost. Customers with Xhibit Telemetry Receivers will get all Xhibit Central Stations updated regardless if connected to an Xhibit Telemetry Receiver.Consignees are directed to technical support for additional information or technical assistance:Technical SupportSpacelabs Healthcare, Inc35301 SE Center St.Snoqualmie, WA 980651-800-522-7025 and select 2 for Technical Support |  | 316 | 316 XTR units total (248 units in the US and 68 units outside US) | Worldwide distribution. US nationwide including Puerto Rico, Australia, Canada, China, France, Germany, Ireland, Italy, Netherlands, Philippines, Poland, Romania, Switzerland, and United Kingdom. | Terminated | March 13, 2017 | 200 | 316 | Computer | Computer
Z-2052-2016 | 74378 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MOSAIQ Oncology Information System |  |  | June 17, 2016 | 2016 | Elekta, Inc. | It is possible that that the selected Plan-of-the Day could be accidently changed by scrolling the mouse wheel up and down and inadvertently selecting another plan. | Software design | Elekta sent an Important Field Safety Notice to all customers on June 17, 2016. The notice informs users of the specific product and version numbers affected by the issue, and any work around that can be used to avoid the issue. A software patch will be issued for Mosaiq version 2.62 in order to the correct the defect. For further questions please call (404) 993-5886. |  | 69 | 69 | US Distribution including Puerto Rico and to the states of : AK, AZ, CA, FL, GA, IL, IN, LA, MD, MI, MO, NJ, NY, OH, OR, PA, TN, TX, UT, VA, WA, and to the countries of : Canada, Denmark, Finland, Germany, Israel, Japan, Netherlands, Norway, Sweden, United Kingdom | Open, Classified |  |  | 69 | Computer | Computer
Z-2879-2016 | 74508 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris Syringe Module Model 8110 with software version 9.15 |  |  | August 08, 2016 | 2016 | CareFusion 303, Inc. | A software anomaly with the Alaris Syringe module software version 9.15 may cause an infusion to unexpectedly stop when the Syringe module is transitioning from one rate to another. | Software design | The firm, BD, sent an "Urgent: Medical Device Recall Notification" letter dated 8/8/16 to customers to inform them that BD has identified a software anomaly associated with the Alaris Syringe module software version 9.15. The letter describes the product, problem and actions to be taken. The customers were instructed to follow the instructions and promptly complete and return the Customer Response Card to expedite the corrective action process.Customers with recall related questions are instructed to contact the BD Support Center at (888) 562-6018. Customers with adverse event reports are instructed to contact customer advocacy at (888) 812-3266. Customers with technical questions regarding the Alaris System are instructed to contact Technical Support at (888) 812-3229. |  | 12000 | 12,000 units | Worldwide Distribution: US(nationwide) and countries of: Canada, Australia and United Arab Emirates. | Open, Classified |  |  | 12000 | Computer | Computer
Z-2072-2016 | 73894 | No | CAL | Flexible laryngoscope. | Anesthesiology | Anesthesiology | 510(K) Exempt | Intubation Scope |  |  | April 06, 2016 | 2016 | Pentax Medical Company | Pentax has become aware that operational/cleaning accessories and therapeutic devices (e.g., clips, stents, balloons, etc.) can become lodged in an endoscope's instrument channel. | Device Design | Pentax sent an Urgent Field Correction letter dated April 6, 2016 to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were reminded to carefully check that all accessories are intact, that no parts have fallen off and become lodged within the endoscope's instrument/suction channel and to ensure that any therapeutic devices passed through the instrument channel are accounted for after use. If the instrument/suction channel does become blocked or clogged due to the accumulation of debris, an accessory that cannot be removed, or other cause, do NOT attempt to correct the blockage or continue to use the endoscope. Customers were instructed to contact their local PENTAX Medical Service facility to have the endoscope repaired. Customers with questions were instructed to call 800-431-5880, ext 2064. |  | 1220 | 1,220 units | Worldwide Distribution - US (Nation Wide) including Puerto Rico and Internationally to Denmark, Guam | Open, Classified |  |  | 31293 | Computer | Not_Computer
Z-2878-2016 | 75113 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Syngo RT Therapist, Accelerator, Linear, Medical |  |  | July 26, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Software patch installation to address several safety issues. | Software design | Siemens sent a Field Safety Notice to all affected customers informing them of a new software patch that they plan to install to address software issues. Customers were advised to notify anyone at their organization who should be aware of the information contained in the Notice. Customers were also advised to include the Field Safety Notice in their System Owner Manual, chapter "Safety Advisory Letters" where it should remain. For questions regarding this recall call 610-219-4834. |  | 103 | 103 accelerators | Nationwide distribution | Open, Classified |  |  | 103 | Computer | Computer
Z-2872-2016 | 75194 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | McKesson Radiology, versions 12.1.1 and 12.2 |  |  | July 12, 2016 | 2016 | Mckesson Medical Imaging Group | The STAT priority icon was not displayed in the Study List when the study was opened. | Software design | McKesson sent an Urgent Field Safety Notice on July 12, 2016 to customers with potentially affected software versions. The letter requested that they contact McKesson to acknowledge that they read and understood the Field Safety Notice and to arrange installation of software update/upgrade. Customers with questions were instructed call the McKesson Support department and ask to speak to their Support Manager.For questions regarding this recall call 404-338-3556. |  | 164 | 164 devices | Worldwide Distribution - US including AL, AK, AZ, AR, CA, CO, FL, GA, ID, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MS, MO, NE, NV, NH, NJ, NY, NC, OH, OK, PA, RI, SC, SD, TN, TX, VA, WA WV, WI, WY, and Internationally to Canada, France, Great Britain and Ireland | Open, Classified |  |  | 164 | Computer | Computer
Z-2865-2016 | 75034 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge Cardio software. |  |  | January 30, 2016 | 2016 | Merge Healthcare, Inc. | Patient information in the header is only found on the first page of the report and not on the subsequent pages. | Software design | Merge sent an Urgent Medical Device Recall letter dated January 28, 2016 to affected customers via e-mail or via certified mail if they did not have an e-mail address. The letter identified the affected product, problem and actions to be taken. For questions call Merge Customer Service at (877)741-5369 or support@merge.com. |  | 42 | 42 sites potentially have the affected version | US Nationwide Distribution | Open, Classified |  |  | 42 | Computer | Computer
Z-2852-2016 | 75143 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo.x picture archiving and communication system |  |  | August 05, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Software error. Incorrect values for the volume calculation from a freehand VOI at the customer site. In volume calculations of prostate as well as in volume calculations of liver were too high. | Software design | Siemens sent an Customer Letter dated August 5, 2016, to all affected customers. The letter identified the product the problem and the action need to be taken by the customer. The letter informed customers of a software change VB10B_HF02 that they have implemented to address several issues.If the device has been sold and therefore no longer in your possession, please forward this notice to the new owner. We would also request that you inform us of the identity of the device's new owner where possible. Please observe this safety notice and comply with the corresponding measure until the update has been fully completed. For further questions please call (610) 219-4834. |  | 338 | 338 systems | Distributed Nationwide | Open, Classified |  |  | 338 | Computer | Computer
Z-1067-2016 | 73100 | No | KQO | Automated urinalysis system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | iChem VELOCITY Automated Urine Chemistry, Catalog No. 800-3049, 800-3050, 800-3079, 800-3080, 800-3530, 800-7167, 800-7190, 800-7713, 800-7714, 800-7719, 800-7720 |  |  | January 13, 2016 | 2016 | Beckman Coulter Inc. | Iris International noted loose or missing analyte pads found in several locations (strip vial, strip provider module, strip conveyor system or waste container) that are undetectable by the iChemVELOCITY System. | Equipment maintenance | Iris sent an Urgent Medical Device Recall letter dated January 13, 2016 and an updated notification letter dated February 19, 2016 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers are instructed to complete and return the enclosed Response Form within 10 days. Customers with questions or concerns are instructed to contact their local support representative:Âż Via our website, at http://www.beckmancoulter.comÂż By phone, call 800 854-3633 in the United States and Canada. |  | 2661 | 2,661 units total (1,424 units in US) | Affected geographies: Argentina, Australia, Austria, Bangladesh, Bermuda, Bosnia and Herzegovina, Brazil, Brunei Darussalam, Canada, Chile, China, Columbia, Croatia, Czech Republic, Ecuador, France, Germany, Greece, Hong Kong, Hungary, India, Indonesia, Israel, Italy, Japan, Jordan, Republic of Korea, Kuwait, Lebanon, Lithuania, Macao, Malaysia, Mexico, Montenegro, Netherlands, New Zealand, Nigeria, Oman, Panama Paraguay, Peru, Philippines, Poland, Portugal, Puerto Rico , Qatar, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, Spain, Switzerland, Taiwan, Thailand, Turkey, United Arab Republic, United States (including Puerto Rico), and Venezuela | Open, Classified |  |  | 2661 | Computer | Not_Computer
Z-2867-2016 | 74296 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Philips Expression Information Portal (IP5) Model 865471. The IP5 is intended to be used as a secondary display control unit for the Expression MRI Patient Monitoring System and the Expression MR200 MRI Patient Monitoring System. |  |  | September 14, 2014 | 2014 | Invivo Corporation | Frozen Display Numerics and Disabled Menu Keys after extended run time. This customer notification was sent August 22, 2014. | Software in the Use Environment | A recall notification letter was sent to hospitals who purchased the recalled product via UPS March 2014. The letter instructed users to power cycle the IP5 unit every day. This is done by pressing the power button on the side of the unit. Customers were instructed to communicate the recall to all those subject and to attach an addendum to the IP5 Instructions for Use. The firm advised customers to call their Philips representative, option 1, then option 2, if they had any questions or concerns. |  | 639 | 639 units | Worldwide distribution. US nationwide, Australia, Austria, Denmark, France, Germany, Iceland, Ireland, Italy, Netherlands, Norway, Poland, Saudi Arabia, Singapore, Slovenia, Spain, Sweden, Switzerland, United Arab Emirates, and England | Completed |  |  | 639 | Computer | Computer
Z-2064-2016 | 74142 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris System PC Unit Model 8015 with software versions 9.17 and 9.19; |  |  | May 12, 2016 | 2016 | CareFusion 303, Inc. | A patient weight can be populated incorrectly under certain conditions when using the RESTORE feature to restore infusions running on the Alaris LVP module model 8100 and the Alaris Syringe module model 8110. | Software design | An Urgent Medical Device Recall Letter was sent to customers on 5/31/16 to inform them that BD had identified an issue with the Alaris System PC unit model 8015 with software versions 9.17 and 9.19. Firm has received reports that a patient weight can be populated incorrectly under certain conditions when using the RESTORE feature to restore infusions running on the Alaris LVP module model 8100 and the Alaris Syringe module model 8110. The letter informs the customers of the potential risk and the required action for users. Customers are informed that BD will contact them within 60 days to initiate the scheduling process for the software upgrade installation. |  | 69693 | 69,693 units | Distributed US (nationwide) and in Canada. | Open, Classified |  |  | 69693 | Computer | Computer
Z-1084-2016 | 73428 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Mobile wDR motorized portable diagnostic X-ray systems. |  |  | January 18, 2016 | 2016 | Sedecal S.A. | This is related to previous 806 Notifications about uncommanded movements of a motorized mobile x-ray system. Reason for Recall: Along with the previously noted encoder and gauges problems and corresponding fixes, we have observed that static electricity can also elicit the unwanted behavior. | Device Design | Philips Medical Systems sent an URGENT - Field Safety Notice letter dated August 6, 2014 was sent to all consignees of record informing them of a potential safety issue on the SEDECAL mobile units manufactured from May 2013 on. The firm requests the direct acounts with affected units should take the following steps:1) The unit is equipped with a safety switch in the drive handle. If the drive handle isreleased the unit will stop; 2) Do not try to correct unintended movement using the controls on the drive handle; 3) Instead, stop the system immediately by releasing the drive handle entirely; 4) Alternatively, the emergency button can be pressed; 5) Take special care when the system crosses metal strips or plates on the floor or enters an elevator, as these are situations where electrostatic discharge is likely to occur; and, 6) Should the customer feel uncertain regarding this action, he/she can contact Philips Healthcare.Philips is preparing a Field Action Kit to replace the existing arrestor strap by a new arrestor with higher resistivity. This will reduce the electromagnetic pulse to a level, which is safe to not disturb the electronics. A Philips Service Engineer will contact the direct accounts to distribute Field Action Kit.The firm states that should direct accounts need to communicate with Philips in regards to this recall that they reference FC0-71200138.If they need any further information or support concerning this issue, please contact their local Philips representative. |  | 349 | 349 systems | Nationwide Distribution including AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NM, NY, OH, OR, PA, SC, TN, TX, UT, VA, VT, WA, WI. | Open, Classified |  |  | 349 | Computer | Computer
Z-2849-2016 | 75080 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | TriMed Semi-Tubular Plate, 10-Hole, Model No. SMTP-10 |  |  | August 29, 2016 | 2016 | Trimed Inc | TriMed is recalling the Semi-Tubular Plates with 10 holes (SMTP-10) because a non-locking screw passed through a hole in a SMTP-10 plate being implanted. | Device Design | The firm, TriMed, sent a "Medical Device Recall Notification" Letter dated 8/29/16 to customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately examine your inventory and quarantine the product subject to this recall; notify your customers at once of this product recall; discontinue use; return product, and complete and return the enclosed MEDICAL DEVICE RECALL RETURN RESPONSE Acknowledgement and Receipt Form as soon as possible by Fax to: 1-661-254-8485 or email to: davidmedoff@trimedortho.com. Customers with questions are instructed to contact Director of Operations at (800) 633-7221, davidmedoff@trimedortho.com, M-F, 9am-5pm, Pacific Standard. |  | 30 | 30 units | Worldwide Distribution-US (nationwide) including states of: Alabama, Florida, Indiana, Connecticut, California, Wisconsin, Missouri, Florida, Hawaii; and countries of: Sweden and South Africa. | Open, Classified |  |  | 30 | Computer | Not_Computer
Z-2842-2016 | 75041 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | HeartStart MRx Monitor/Defibrillator |  |  | August 26, 2016 | 2016 | Philips Electronics North America Corporation | The handle can separate from the MRx housing due to breakage of mounts on the rear case. | Device Design | Philips Healthcare issued customer an Urgent Medical Device Correction notification/Field Safety Notice on 11/19/14. The Urgent Medical Device Correction notification/Field Safety Notice FSN86100158A informs customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer, and identifies what action Philips plans to take to remedy the issue.You can continue to use your MRx prior to receiving the hardware upgrade. If one or both sides of the handle are found to have separated from the MRx housing, continue to treat your patient per existing protocols. Do not use the handle to carry the MRx. The MRx can be carried in another manner. For example, the MRx shoulder strap or bed rail hook can be used if available on the unitA hardware upgrade will be provided free of charge for all units affected by one or more of these issues. A Philips Healthcare representative will contact customers with affected devices to arrange for installation of the upgrade. Contact your local Philips representative or call 1-800-722-9377. |  | 394 | 394 units | Worldwide distribution. US Nationwide, Canada, AUSTRALIA, FRANCE, GERMANY, INDIA, JAPAN, NETHERLANDS, SOUTH AFRICA, SPAIN, SWEDEN, THAILAND, and UNITED KINGDOM. | Open, Classified |  |  | 394 | Computer | Not_Computer
Z-2839-2016 | 75121 | No | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | G1 Dissection Tool - 3mm FLUTED BALL; For cutting and shaping bone including spine and cranium. |  |  | August 26, 2016 | 2016 | The Anspach Effort, Inc. | The firm discovered on March 4, 2016 that several product codes were inadvertently left off of Hold 1503. | Device Design | On April 26, 2016 DePuy Synthes sent a letter informing their customers of the recall. ACTION REQUIREDOur records show that your facility has received one or more of product(s) subject to the recall (correction). The AnspachEffort, Inc. asks that you:1) There is no requirement for retaining your inventory of Gl Dissection Tools.2) To ensure Gl Cutting Tool &amp; Attachment compatibility, please reference Table 1 to review which Attachments arecompatible. The Cutting TooVAttachment combination noted with a check mark in Table I will function as intended.The Cutting Tool/Attachment combinations noted with an "X" in Table 1 will not secure effectively and should not beused.3) Review, complete, sign and return the attached reply form confirming your receipt of this letter by fax or email detailsprovided on the form. Returning the form promptly will prevent you from receiving repeat notices.1. Share this letter with others in your facility that needs to be made aware of this recall (correction).2. If you distribute any of the product(s) subject to the recall (correction) to other services or facilities, please forwardthis letter as appropriate immediately.3. Keep a copy of this notice with the product(s) subject to the recall (correction).If you have any questions regarding this correction, please contact the Complaint Handling Unit Manager at 561-494-3673or contact your Anspach Sales Representative. |  | 63 | 63 | AZ, DC, CA, MD, MI, and ORSwitzerland | Open, Classified |  |  | 397 | Computer | Not_Computer
Z-1586-2016 | 73433 | No | DOP | Mass spectrometer for clinical use. | Clinical Chemistry | Toxicology | 510(K) Exempt | AnalystÂż MD Version 1.6.1 and 1.6.2 Software used with the following instruments: API 3200MD" LC/MS/MS System, Instrument Part Number: 5024501; 3200MD QTRAPÂż LC/MS/MS System, Instrument Part Number: 5024500; Triple Quad" 4500MD LC/MS/MS System, Instrument Part Number: 5031257; QTRAPÂż 4500MD LC/MS/MS System, Instrument Part Number: 5031231 |  |  | February 24, 2016 | 2016 | Ab Sciex | Wrong quantitative results may be displayed in a report from the device, which may potentially lead to an incorrect patient diagnosis. | Software design | Sciex sent an Urgent Medical Device Correction letter dated February 29, 2016, to all affected US customers via Fedex on March 2, 2016. International customers were notified beginning on March 3, 2016. The letter informed customers of the issue and requests that they take temporary steps to eliminate the potential for erroneous events. In addition, a software update is in development and is expected to be available in approximately 4 weeks. Customers with questions were instructed to contact SCIEX at 1-289-982-2531. |  | 279 | 279 | Worldwide Distribution - US (Nationwide) and Internationally to the following countries:AUSTRIACANADACHINAFRANCEGERMANYINDIAIRELANDITALYKOREAMALAYSIANETHERLANDSNORWAYRUSSIASINGAPORESPAINSWITZERLANDUNITED KINGDOM | Open, Classified |  |  | 279 | Computer | Computer
Z-0046-2017 | 74871 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Toshiba America Medical Systems, Inc Aquilion Lightning TSX-035A CT Scanner. |  |  | April 25, 2016 | 2016 | Toshiba American Medical Systems Inc | It has been found that if the scan range extends beyond the maximum Field of View (FOV), a reconstruction operation error may occur. As a result, scanning may be interrupted and rebooting of the system may be necessary. Please note that this issue has not occurred in the U.S. | Software design | Toshiba America Medical System, Inc. planned action to bring the defect into Compliance:1. A notification will be sent to customers with affected systems in compliance with 21 CFR 1003.21 upon approval by the Secretary. The notification letter includes a statement by the correction will be made at no cost to the customer.2. To correct this issue, modified software will be installed on the systems to prevent this occurrence from happening. The consignee will be contacted by their Toshiba service representative for an appointment when the revised software becomes available.3. To avoid this problem in scan planning, the consignee will be advised to do the following until the modified software is installed:a. Confirm that the scan range does not extend beyond the maximumFOV.b. If the ROI for the scan range (indicated by the yellow solid lines) is set as the same size as the maximum FOV (indicated by green dashed lines), the ROI position must be changed. Instructions for changing ROI will be provided as attachments to the notification letter.4. The modification to the software will be completed within 90 days from the date that the letters are submitted to Toshibas customers.5. Progress letters will be provided to the Secretary, including the number of electronic products repaired. For further questions, please call (714) 730-5000. |  | 165 | 165 | US Distribution to the states of: PA, IA, MA, SD, MI, LA, KY, NC, NY, AR, IN, WI, OH, TY, FL, WV, WA, CT, MS, AL, ND, CO, HI, NO, MN, ID, OH, GA, MT, LA, NM, VT, TN, CA, NH, NJ, NE, SC and KS. | Terminated | July 13, 2017 | 444 | 165 | Computer | Computer