## Sheet1
Recall Number | Recall Event ID | Device Type | Product Code | Regulation Description | Medical Specialty | Review Panel | Submission Type | Trade Name/Product | Main Name | Recall Class | Date Posted | Year Posted | Recalling Manufacturer | Reason for Recall | FDA Determined Cause | Action | Consumer Instructions | Clean Quantity | Quantity in Commerce | Distribution | Recall Status | Termination Date | Time to Terminate | Merged Quantities | Fault_Class
Z-1075-2017 | 76089 | No | EQL | Laryngostroboscope. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(K) Exempt | Rigid Laryngostroboscopes |  |  | December 12, 2016 | 2016 | Pentax of America Inc | Pentax is initiating a Field Correction to revise the Instructions for Use of the Rigid Laryngostroboscopes Model 9106 and 9108. | Labeling design | The firm, Pentax Medical, sent an "U.S. URGENT FIELD CORRECTION" letter and a Return Response form dated December 12, 2016 to their affected customers via USPS Certified Mail. The letter described the product, problem and actions to be taken. The customers were instructed to ensure that all users read and carefully follow the revised IFU in order to properly operate and reprocess the affected product; contact your PENTAX Medical sales representative to arrange for training regarding the new procedures, and complete and return the FIELD CORRECTION RESPONSE FORM via fax to: PENTAX Recall Coordinator at 201-799-4063 (alternate 201-391-4189) or email a pdf copy to customeradvisories@pentaxmedical.com even if you no longer possess the product.If you have any questions regarding this action, please feel free to contact us at: 800-431-5880 (8:30 AM-5:00 PM, Monday-Friday, EST) or email: customeradvisories@pentaxmedical.com. |  | 2056 | 2,056 (1,545 (US) and 511 (OUS)) | Worldwide Distribution-US (nationwide) and countries of: Albania, Australia, Belgium, Brazil, Bulgaria, Canada, Chile, Croatia, Denmark, Ecuador, Egypt, Finland, France, Germany, Hong Kong, India, Ireland, Israel, Italy, Japan, Kuwait, Lebanon, Lithuania, Malaysia, Mexico, Netherlands, New Zealand, Oman, Philippines, Portugal, Saudi Arabia, Singapore, South Africa, South Korea, Spain, Sweden, Taiwan, Thailand, Trinidad, Turkey, United Arab Emirates, and United Kingdom. | Open, Classified |  |  | 2056 | Not_Computer
Z-1812-2017 | 76787 | No | IRT | Powered heating pad. | Physical Medicine | Physical Medicine | 510(K) Exempt | Volt Resistance Heated Slipper (All sizes) |  |  | April 15, 2016 | 2016 | 2k Innovations Inc. | During an FDA inspection it was found that the firm claims unapproved therapy and medical claims on product information and website. | Labeling False and Misleading | 2k Innovations sent e-mail communication from 04/15/2016 thru 09/09/2016 informing their consignees that the firm is no longer linking their products to any medical and therapeutic claims and that new labels and boxes will be corrected and the claims will be taken out of the website. For questions customers should call 888-518-6871, ext 700. |  | 22757 | 22,757 items in total | Nationwide Distribution | Open, Classified |  |  | 22757 | Not_Computer
Z-1835-2017 | 76604 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | INFINITY Internal Hex Healing Abutment, 4.5mm Platform, 4.7mm CW-3mm CH Straight; Product Code: 20145473 |  |  | February 22, 2017 | 2017 | Ace Surgical Supply Co., Inc. | Incorrect device was packaged. | Labeling mix-ups | ACE sent a Customer Notice letter dated February 22, 2017 via certified mail to customers. The letter identified the affected product, problem, risk to health and actions to be taken. Customers were instructed to complete and return the enclosed Business Reply Form. For question call 1-800-441-3100 ext. 208. |  | 87 | 87 | Nationwide Distribution in the states of CA, CT, FL, NY, OH, PA, and UT. | Terminated | August 07, 2017 | 166 | 87 | Not_Computer
Z-1834-2017 | 76509 | No | GFJ | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Soft Face "No Bounce" Mallet |  |  | February 20, 2017 | 2017 | Signal Medical Corporation | Potential for poly tip of the mallet to come off during use. Potential for foreign material in the surgical field or delay in surgery. Reminder for customers that the poly tip must be removed for cleaning and sterilization to assure proper cleaning and sterilization of the entire mallet. | Labeling design | On 2/20/2016, URGENT: MEDICAL DEVICE RECALL notifications dated February 10, 2017 were sent to the affected consignees via e-mail. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Customers are to retain the notice and replace the old Instructions For Use with the updated version. Also, they are to acknowledge receipt of the instructions by signing the notice and emailing to SignalMedical2017@gmail.com or faxing to 810-364-7072. The firm has attached the revised labeling to the letter. |  | 1640 | 1,640 | Nationwide Distribution. | Open, Classified |  |  | 1640 | Not_Computer
Z-2725-2017 | 77481 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Medica ISE Module Calibrant A; 500 mL - 140 mmol/L Na+, 125 mmol/L Cl-, buffers and preservatives; For use with ISE Module; |  |  | February 29, 2016 | 2016 | AMS Diagnostics, LLC | Lot was not functioning properly and not exhibiting the correct physical properties. | Labeling mix-ups | AMS Diagnostics contacted consignees by telephone on 02/29/2017. Representatives advised consignees to discard any recalled product on hand, to be replaced with working product. Customers should not return recalled product to the firm. |  | 9 | 9 units | US Nationwide Distribution in the states of AZ, CA, FL, MI, NC, NY, and SC. | Open, Classified |  |  | 9 | Not_Computer
Z-1079-2017 | 75599 | chest drainage kit | PAD | Vacuum-powered body fluid suction apparatus. | General Hospital | General Hospital | Enforcement Discretion | Atrium Medical Ocean Drains as follows: |  |  | November 29, 2016 | 2016 | Atrium Medical Corporation | Outer Packaging is not sterile | Package design/selection | Atrium Medical issued recall letter on 11/29/16 via two-day express mail delivery advising users: of the problem, updated instructions for use; the new Warning Label applied to the packaging; and requesting to post the warning label near all Atrium chest drain product inventory.A Field Notification Reply Form acknowledging receipt of the Field Notification was requested to be completed and returned.Questions contact Atrium/Maquet/Getinge chest drainage representative or Atrium Customer Service at 1-800-528-7486 (603-880-1433), Monday through Friday between 8:00 am and 5:00 pm (EST). |  | 964463 | 964,463 cases (US) 521,975 cases (OUS) in total | Worldwide Distribution - US Nationwide and the countries of Algeria, Australia, Austria, Belgium, Brazil, Canada, Chile, China, China, Columbia, Ecuador, Egypt, Finland, France, Germany, Great Britain, Greece, Guatemala, Hong Kong, Ireland, Italy, Kuwait, Lebanon, Libya, Mexico, Netherlands, New Zealand, Pakistan, Panama, Peru, Poland, Portugal, Puerto Rico, Saudi Arabia, South Korea, Spain, Sweden, Thailand, Tunisia, and Venezuela, | Open, Classified |  |  | 964463 | Not_Computer
Z-1829-2017 | 76872 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | Arrow EZ-IO Needle Set 45 mm 15 ga, Ref 9079. |  |  | March 29, 2017 | 2017 | Teleflex Medical | Some of the individual unit packaging (pouches) may be missing the lot number an expiration date. | Incorrect or no expiration date | Consignees were notified of the recall by letter on March 29, 2017. The letter requested that they check their inventory and immediately discontinue use and quarantine any product and to inspect on hand for missing lot numbers and expiration date. The letter included a reply form to be returned to Teleflex. A separate letter was sent to distributors requesting a sub-recall. |  | 4760 | 4760 eaches | AL, AK, AZ AR, CA, CO, CT, FL, GA, HI, ID, IL IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, WY, Canada and Costa Rica | Open, Classified |  |  | 4760 | Not_Computer
Z-1854-2017 | 76729 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Peri-Loc(R) VLP 3.5 MM X 6 MM SELF-TAPPING CORTEX SCREW, QTY: (1), STERILE R, REF 71801306 |  |  | March 09, 2017 | 2017 | Smith &amp; Nephew, Inc. | The expiration dates on the labeling are incorrect, the actual expiration dates are before the date on the labeling. | Labeling Change Control | Smith &amp; Nephew sent an Urgent Medical First Recall Notice dated March 9, 2017, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. The letter requested the quarantine and return of the devices, and for the distributors to notify their consignees and ensure the same actions. International consignees were notified by email. Consignees with questions were instructed to contact FieldAction@smith-nephew.com.For questions regarding this recall call 978-749-1440. |  | 1 | 1 unit | Worldwide Distribution - US (Nationwide) Internationally to Colombia, Denmark, Finland, France, Great Britain, Italy, and Singapore | Open, Classified |  |  | 317 | Not_Computer
Z-2715-2017 | 77565 | Yes | LWQ |  |  | Cardiovascular | PMA | On-X Prosthetic Heart Valve |  |  | June 12, 2015 | 2015 | CryoLife, Inc. | Incorrect product label. The type of Sewing Ring was mislabeled as being a Standard Sewing Ring instead of an Anatomic Sewing Ring. | Labeling mix-ups | On-X LTI Management Team had assessed this event to not be reportable. Correction were done by the On-X LTI in 2015. |  | 52 | 52 | Worldwide Distribution - US Nationwide in the states of AL, CA, FL, IA, IN, LA, MA, MN, MO, MS, NC, OR, PA, and TN.and the countries of Finland, United Kingdom and Israel | Open, Classified |  |  | 52 | Not_Computer
Z-1738-2017 | 76840 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | DX-D600 - DXD Imaging Package |  |  | May 02, 2016 | 2016 | AGFA Healthcare Corp. | One of the dent markers from the DX-D600 dropped down from the overhead rail. | Labeling design | On May 2, 2016, an "URGENT FIELD SAFETY NOTICE" letter was either emailed or sent via FedEx to the US and Canadian consignees. The notice included an acknowledgment to be sent back. The letter described background information and actions to be taken. |  | 81 | 81 units | AL, CA, DC, FL, GA, HI, IL, IN, KS, LA, MA, MN, MO, NV, NJ, NM, NY, OH, OK, OR, SC, TN, TX, VA, and Canada | Terminated | April 10, 2017 | 343 | 81 | Not_Computer
Z-1023-2017 | 76034 | No | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Medacta International SA, M.U.S.T. Rod, REF 03.50.433, Size 5.5mm, Length 480mm, Type Straight, Mat Co-Cr-Mo |  |  | December 07, 2016 | 2016 | Medacta Usa Inc | Damaged packaging, lack of sterility assurance | Packaging | The firm, Medacta International, immediately notified all customers by phone on 12/07/16 and further notified via email on 12/9/16. The firm also issued an "URGENT- Field Safety Notice" dated 7/12/2016 to all its customers. The notice described the product, problem and actions to be taken. The customers were instructed to return all units to Medacta International SA and forward the notice to all people and organizations potentially affected.If you have any questions contact Regulatory Affairs Manager at +41 (0) 91 696 60 60. |  | 82 | 82 units | Worldwide Distribution- US (nationwide) distribution to states of: TX and CO; and countries of: Australia, Belgium, Spain, German, Italy and Switzerland. | Open, Classified |  |  | 82 | Not_Computer
Z-2660-2017 | 77603 | No | IEO | Automated tissue processor. | Pathology | Pathology | 510(K) Exempt | Nussloch GmbH ASP6025 Tissue Processor |  |  | June 19, 2017 | 2017 | Leica Microsystems, Inc. | Incorrect labeling for specified voltage for the Alarm Connectors on the rear side of the unit. | Labeling mix-ups | Customers were notified via email on 06/19/2017. Instructions to customers included to clean surface around the Alarm Connectors and place the newly provided label at the Alarm Connectors (over the engraved specification). They were also instructed to update their current Instructions For Use and update the online help function in the software, as well as return the response form. |  | 812 | 812 units | Domestic: AR, AZ, CA, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MN, MO, MT, NC, NE, NJ, NM, NV, NY, OH, PA, SC, TN, TX, UT, VA, WA, WI, WV, &amp; DC. Foreign: Australia, Austria, Bangladesh, Belgium, Brazil, Canada, Chile, China, Colombia, Costa Rica, Czech Republic, England, France, Georgia, Germany, India, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Korea, Kuwait, Lithuania, Malaysia, Maldives, Mexico, Morocco, Myanmar, Netherlands, Norway, Peru, Philippines, Poland, Portugal, Qatar, Reunion, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovenia, South Africa, Spain, Switzerland, Taiwan, Thailand, Turkey, &amp; United Arab Emirates. | Open, Classified |  |  | 812 | Not_Computer
Z-1737-2017 | 76698 | No | PJC | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | Bond Elut OMS (Dried Matrix Spotting)' Product; |  |  | June 17, 2015 | 2015 | Agilent Technologies, Inc. | The product was marketed without 510(k) clearance. | No Marketing Application | Agilent Technologies Letters via e-mail were sent to customers on June 17, 2015.Letters advised that the Bond Elut DMS Product was being withdrawn from the market. |  | 364 | 364 in total | US and Australia, Thailand, Great Britain, Germany, Japan, Netherlands, France, Taiwan, Canada, South Africa, China, Slovenia, Spain, Italy, Norway, India, Russia, Finland, Czech Republic, Singapore, Denmark, Korea, Ireland, United Arab Emirates, .Belgium, Switzerland and Sweden. | Terminated | April 05, 2017 | 658 | 364 | Not_Computer
Z-2664-2017 | 77004 | No | HTX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | EXACTECH Equinoxe Glenoid Reamer - Catalog #: 321-05-38, Catalog #: 321-05-42, and Catalog #: 321-05-46 |  |  | May 25, 2017 | 2017 | Exactech, Inc. | Instructions for use (IFU) is updated to clarify the proper technique and instrumentation to ream the Glenoid, i.e., to include a caution statement that the surgeon should ream on-axis and that reaming off-axis may result in instrument breakage. | Error in labeling | On May 22, 2017 Exactech sent a Notification to all their Direct Accounts an Distributors telling them that a new updated technical techniques LIT # 718-01-30 Rev J is now available in their website (www.exac.com) or can be requested by contacting the Customer Service Representative at 800-392-2832, fax 352-378-2617 or via e-mail at Kaya@exac.com. This updated the technical techniques for implantation for Equinoxe Platform Shoulder System to clarify the proper technique for use of the Pilot Tip Reamers, i.e., to include caution statement that the surgeon should avoid applying a bending force to the pilot tip reamer and should avoid using the reamer to retract the humeral head, as this may result in instrument breakage. The notification also request to forward and informed this notice to all the accounts that have received this product and to complete and return the attached Field Advisory Notice Respond Form within 5 days of receipt. For more information please contact: Kaya Davis, Inventory Representative at kaya@exac.com or by phone at (352) 377-1140 |  | 813 | 813 | AL, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KS, MA, MD, ME, MN, NC, NJ, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, VA, WA, WI, Hawaii and Puerto RicoAustralia, Austria, Brazil, Canada, China, France, Germany, Greece, India, Italy, Korea, Luxembourg, Singapore, South Africa, Spain, Sweden, Switzerland, The Netherlands and United Kingdom | Open, Classified |  |  | 4390 | Not_Computer
Z-1249-2017 | 75829 | No | HIS | Condom. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Durex Pleasure Pack . |  |  | November 22, 2016 | 2016 | Reckitt Benckiser LLC | Reckitt Benckiser LLC is recalling Durex Condoms Pleasure Packs due to missing information on the package labeling associated with the active ingredient Benzocaine. | Error in labeling | Reckitt Benckiser LLC sent an Urgent Medical Device Recall letter/business response form dated November 22, 2016 to the affected customers. Stericylcle is handling returns. |  | 108342 | 108342 units | Nationwide and Foreign | Open, Classified |  |  | 108342 | Not_Computer
Z-1891-2017 | 77100 | No | MAV | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | DOLPHIN INFLATION DEVICE CALIBER INFLATION DEVICE; Catalog Number: CL3030- 018SNA- 018SND 0185NA 0185ND; INTERNAL CODE NUMBER: 0185TR- 0185NA- 018SND 0185NA 0185ND |  |  | April 05, 2017 | 2017 | PEROUSE MEDICAL | Complaints regarding broken blister. | Packaging change control | Consignees were notified of the recall on April 5, 2017. They were informed of the issue and asked to take the following steps: 1. Inspect your stock to determine if you still have in your ownership devices among which the references and batch codes are listed in front page; 2. Control the potential damage of the blister according to the pictures above and the instructions mentioned on the labeling and the Instruction For Use: any damaged and/or cracked blister must be immediately discarded; 3. Use the compliant products without risk, after control. Indeed, the primary packaging will not deteriorate in time. If, to date, there is no crack in the blister, then the integrity of the product is not compromised; 4- Do not use the non-compliant products; 5. Complete and sign the attached reply form, then return it within 5 working days following the reception of this letter, and keep a copy; 6. Contact customer service for the modalities of return, replacement and associated support at +33 (0)4.72.39.74.13 or by fax at +33 (0}3.44.08.17.67 or by email at iblayon@vvgon.com (CC mhpourriere@vygon.com); 7. Communicate the safety information to any person concerned in your company and to all end customers to which the product were transferred. |  | 18431 | 18,431 | US: AZ (shipped through Belgium to USA), MN, TN | Open, Classified |  |  | 18431 | Not_Computer
Z-1894-2017 | 77075 | No | MMZ | Urological table and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Sonesta S2; Article numbers: 525-S2-110V and 525ÂżÂżS2-220V |  |  | March 08, 2017 | 2017 | Sonesta Medical Ab | The user / service manual for the Sonesta S2 Chair contain wrong information concerning the maximum patient weight. | Labeling mix-ups | Sonesta notified customers (distributors) on March 8, 2017, of the faulty labeling. Corrected labeling was sent to customers, they were requested to apply the new labels and confirm that the correction was implemented. For further information regarding this recall, please call (512) 327-9997. |  | 28 | 28 devices | Worldwide Distribution - US Distribution to the state of : MA., and to the countries of : Canada, Netherlands, Dubai, Taiwan, Lebanon, Italy | Open, Classified |  |  | 28 | Not_Computer
Z-1818-2017 | 76782 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Keystone Dental Inc. PrimaConnex Straight Implant TC, WD Internal Connection 5.0 x 11.5mm. |  |  | March 09, 2017 | 2017 | Keystone Dental Inc | Straight PrimaConnex Implants contained Tapered PrimaConnex Implants | Labeling mix-ups | Keystone Dental initiated a market removal communication by telephone to customers on 3/9/2017. Customer Relations Fax Back Forms were completed to document these calls. In addition, a letter was sent via Federal Express/or hand delivered to all distributors and customers on 3/13/2017. Product returns requested and if implanted to implement healing and prosthetic protocol. Toll Free: 866-902-9272 |  | 59 | 59 units | CA, IL, IN, FL, MA, MN, NJ, NV, PA, | Terminated | July 05, 2017 | 118 | 59 | Not_Computer
Z-1884-2017 | 76842 | No | MNH | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | 5.5mm Diameter x 500mm CoCr Straight Rod, Model number X077-0420. The product is packaged in plastic bags with one piece per bag. |  |  | April 12, 2017 | 2017 | XTANT MEDICAL INC | The material type on the label may incorrectly state Ti6Al-4V ELI . The rods are composed of Cobalt Chromium. | Labeling mix-ups | XTANT sent an Urgent Medical Device Removal letter dated 4/12/2017 to customers with a Return Response Acknowledgement and Receipt form. The letter identified the affected product, problem and actions to be taken. For questions call 406-570-7028. |  | 97 | 97 units | Worldwide distribution - US Nationwide in the states of: AZ, CA, CT, GA, LA, NV, OH, TX and in the country of Mexico. | Terminated | August 16, 2017 | 126 | 97 | Not_Computer
Z-1871-2017 | 76775 | No | EYZ | Urine collector and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | SURESTEP Tray, Bardex I.C. Complete Care Drainage Bag, 2000 ml Drainage Bag, 16 Fr., Sterile. |  |  | March 09, 2017 | 2017 | C.R. Bard, Inc. | Misbranded: Product labeled, Not made with natural rubber latex, may in fact contain natural rubber latex. | Error in labeling | Consignees were notified by letter on/about 03/09/2017. |  | 5030 | 5030 units | CA, CO, FL, GA, IA, IL, IN, KY, LA, MI, MN, MS, NC, NE, NJ, NM, OH, OR, PA, RI, TN, TX, VA and WA. | Open, Classified |  |  | 5030 | Not_Computer
Z-2669-2017 | 77190 | [6 empty] | Electrophysiology catheter cables are designed |  |  | May 08, 2017 | 2017 | Sterilmed Inc | Sterilmed is recalling all lots of specific product codes of unexpired Sterilmed Reprocessed EP Ablation Cables because Sterilmed determined that the cables are Class III devices and thus require, but do not currently have, premarket review and approval. | No Marketing Application | Sterilmed sent an Urgent Notice Medical Device Recall letter dated April 25, 2017, to all affected customers. The letter identified affected product, stated reason for recall. Affected devices were asked to be quarantined, until returned to Stericycle, by August 31, 2017 to qualify for credit. A response form was asked to be returned. Questions can be directed to Sterilmed's Customer Care Team at 1-888-541-0078. For further questions, please call ( 877 ) 787-0369. |  | 963 | 963 devices | US Distribution to the states of : AL, AR, CO, FL, IL, MI, MO, MS, MT, NC, OH, PA, TN, TX, VA WI and WV. | Open, Classified |  |  | 963 | Not_Computer
Z-1112-2017 | 76111 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | SwishTapered Implant |  |  | December 12, 2016 | 2016 | Implant Direct Sybron Manufacturing, LLC | The extender product code (9037-03) that is packaged with the SwishTapered (Part Number 933308; Lot Number 73241) Implant was packaged in the incorrect position. This discrepancy may lead to the possibility of extender not seating properly on the implant, which may lead to micro-leakage and a possible infection to the patient. | Packaging process control | Customer notification letters were sent on 12/12/16.Customers were requested to review inventory for the affected product. Complete and return the Acknowledgement and Recall Return Form within 48 hours. If customers are an authorized distributor the firm requests that they identify those customers that may have been shipped the affected product lot and contact those customers to inform them of the issue within 48 hours of receipt of the notification in order to provide the customers with replacement product.If customers have any of the affected product the firm will send a replacement part. Any questions contact Implant Direct Sybron Manufacturing LLC Customer Care at 1-888-649-6425. |  | 17 | 17 | U.S. distribution to the following; NM, TX, CA, OR, OH, ME.Foreign distribution to the following; Netherlands. | Terminated | May 22, 2017 | 161 | 17 | Not_Computer
Z-1105-2017 | 76190 | No | LHC | Powered heating unit. | Physical Medicine | General Hospital | 510(K) Exempt | NORMFlow H-110 Irrigation Fluid Warmer, Continental Europe, Model H-1100, Reorder H-1100-INT-230. |  |  | October 25, 2016 | 2016 | Smiths Medical ASD, Inc. | Smiths Medical became aware that that one (1) NORMOFLOÂż H-1100 Irrigation Fluid Warmers was incorrectly labelled with Spanish language labels instead of English labels. Specifically five (5) labels were incorrect; the Operators Manual, Outer Base Box label, Pole Assembly Box label, and two (2) O-ring Lube labels. | Labeling mix-ups | Consignee were hand delivered a Smiths Medical Urgent Medical device Filed Safety Notice letter dated October 25, 2016. The letter described the Reason for Field Corrective Action, Risk to Health, and Instructions to Customers. Advised consignees to Provide the device to the Smiths Medical Representative who delivered this notice to you. The Smiths Medical Representative will replace the labeling on the device and packaging, and provide you with the correct English language Operators Manual. Complete the Recall Confirmation Form and return it to Smith Medical via e-mail to FCA.Response@Smiths-medical.com. For questions regarding this notification, please ask the Smiths Medical Representativewho delivered the notification or contact Smiths Medical Customer Service Department at +46 859477250. |  | 1 | 1 | Distribution in the foreign country of Sweden | Terminated | July 25, 2017 | 273 | 1 | Not_Computer
Z-1060-2017 | 76074 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | OEC 9900 Elite. MDL Numbers: D148942, D155043 |  |  | December 27, 2016 | 2016 | GE OEC Medical Systems, Inc | GE Healthcare Surgery announces a voluntary field action for the OEC 9800 and OEC 9900 Elite due to additional supplementation to the user documentation to include areas for enhanced planned maintenance and component replacement frequencies. | Packaging change control | The customer communication is expected to begin distribution in January 2017. Consignees will be notified of the availability of the Operator Manual supplement via a letter. No product is being returned. |  | 26256 | 26,256 total | Worldwide including govt/VA/military and foreign consignees. | Open, Classified |  |  | 26256 | Not_Computer
Z-2662-2017 | 77001 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Logic Fit Tibial Tamp Head |  |  | May 25, 2017 | 2017 | Exactech, Inc. | The firm determined that the labeling (surgical technique) should be updated to clarify the proper technique and instrumentation to remove the tamp assembly from the bone, i.e., to include a caution statement about the potential for instrument breakage if the tamp handle/guide is misused by impacting in retro-grade. | Error in labeling | Consignees were sent a Field Advisory Notice dated 5/22/2017 by mail informing them updated surgical techniques are available on Exactech's website (www.exac.com). The notice requested consignees:Extend the information to accounts in possession of the productComplete &amp; return the attached Field Advisory Notice Response Form to Exactech |  | 573 | 573 | AL, AZ, CA, CA, CO, FL, GA, IA, ID, IL, IN, KS, LA, MA, MD, ME, MN, NC, NJ, NV, NY, OH, OK, OR, RI, SC, TN, TX, VA, WA, WI, Hawaii and Puerto RicoAustralia, France, Germany, India, Italy, Japan, Korea, Luxembourg, Singapore,Spain, Switzerland, The Netherlands and United Kingdom | Open, Classified |  |  | 573 | Not_Computer
Z-1869-2017 | 76446 | No | FMG | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Elcam Minimal Residual Volume Luer-activated Swabable-stockcock (MRVLS) 4w gxl:w1/r rbls |  |  | February 02, 2017 | 2017 | Elcam Medical, Inc. | The stopcock was found to have punctures. | Package design/selection | Elcam Medical sent their customers a recall letter/return response form dated February 2, 2017. Elcam is asking customers to examine their inventory for any affected products and to complete the return response form provided. For questions call 201-457-1120 ext. 102 or 201-457-1120 ext. 112. |  | 810 | 810 | U.S Nationwide Distribution | Open, Classified |  |  | 810 | Not_Computer
Z-1147-2017 | 75338 | No | CAF | Nebulizer. | Anesthesiology | Anesthesiology | 510(k) | Thayer MiniSpacer Dual Spray MDI Adapter Counter Incrementing Actuator with 15mm O.D./I.D. Connections, Reference no. 1543A. |  |  | September 29, 2016 | 2016 | Thayer Medical Corporation | The insert and case label description of the device incorrectly reads Counter Incrementing Actuator with 22mm O.D./I.D. Connections. | Labeling design | Thayer Medical sent an Urgent Medical Device Recall letter dated September 30, 2016, to all affected customers. The letter informed customers that Thayer Medical has identified a labeling error on REF 1543A, Thayer MiniSpacer Dual Spray MDI Adapter with Counter Incrementing Actuator and 15mm O.D./I.D. Connections. Customers were informed to complete and return the attached Customer Reply Form by email to info@thayermedical.com or fax to 1-520-790-5854.For questions regarding this recall call 520-790-5393, ext 2210. |  | 2550 | 2,550 units | Nationwide Distribution | Terminated | February 07, 2017 | 131 | 2550 | Not_Computer
Z-1767-2017 | 76749 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | CADDÂż Administration Set, 21-7339-24, Rx Only, Sterile EO. |  |  | March 17, 2017 | 2017 | Smiths Medical ASD Inc. | Smiths Medical has become aware that the drawing on the Instructions for Use (IFU) on certain yellow striped CADDÂż Administration Sets (part number 21-7339-24), incorrectly represent the actual filter size of the product. Although the text description on the IFU correctly identifies the product to include a 0.2 micron filter, the illustration on the IFU identifies the filter size as 2.0 microns. | Packaging process control | Smiths Medical sent an "Urgent Medical Device Field Safety Notice" dated March 13, 2017. The letter was addressed to "Risk Manager". It provided a list of the Affected devices, Risk to Health, Instructions to Distributors &amp; Customers. Advised consignees to notify their customers. There is no need to return the devices to Smiths Medical. Requested consignees to complete the Urgent Medical Device Field Safety Notice (FSN) Response Form and return to smithsmedical3947@stericycle.com. For questions contact Stericycle at 844-265-7407. For further questions regarding this recall , please call (844) 265-7407. |  | 38415 | 38,415 (38145 US; 2268 OUS) | US:AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MS, MT, NC, NE, NJ, NM, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, VT, WA, WV, WY.OUS: Canada | Open, Classified |  |  | 38415 | Not_Computer
Z-3067-2017 | 77380 | No | ODG | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Ultrasound Video Gastroscope, Model EG-3670URK, |  |  | May 23, 2017 | 2017 | Pentax of America Inc | Error in reprocessing instructions in Revisions 13 and 14 of IFU Z845, used with EG-3670URK and EG-3870UTK ultrasound gastroscopes. In the disinfection step, the IFU has a warning indicating the cleaning detergent solution should remain in contact with the all internal channels and external endoscope surfaces. The disinfecting solution, not the cleaning detergent, should be used for this step. | Error in labeling | Recall Notifications were sent to 104 consignees via USPS Certified Mail on 5/23/2017. Included with the Field Safety letter was the customer response form, and the addendum to IFU Z845. These were also posted on Pentax of America web page until the recall is closed. |  | 118 | US: 118 units | Nationwide | Open, Classified |  |  | 234 | Not_Computer
Z-2780-2017 | 77490 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Oxford/Oxbridge AFN End Cap |  |  | April 19, 2017 | 2017 | Ortho Solutions Inc | There is a material mislabeling of the AFN End Cap. | Error in labeling | Orthosol notified their customers on 4/19/2017 via USPS. |  | 277 | 277 units | US distribution to MD, NY &amp; PA | Open, Classified |  |  | 277 | Not_Computer
Z-1487-2017 | 76687 | No | JDR | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | MotoCLIP(TM) Super Elastic Fusion System, Size 18mm x 14mm x 14mm, REF 1442-1814, Rx Only, STERILE R. |  |  | February 16, 2017 | 2017 | Crossroads Extremity Systems Llc | Seven 18mm x 18mm x 18 mm implants had been placed in kits, packed and labeled as 18mm x 14mm x 14mm. | Labeling mix-ups | The firm notified their distributors by email on 02/16-17/2017. The distributors were asked to return the unused lot #101866. The firm also contacted the implanting surgeons either by visit or telephone about the potential for an incorrect size of the implant and asked if there were any adverse events. |  | 20 | 20 units | Al, FL, IN, Il, NE, TX, NM, CO, AZ | Terminated | April 27, 2017 | 70 | 20 | Not_Computer
Z-1407-2017 | 76560 | No | HTY | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | MicroAire K-Wires |  |  | February 21, 2017 | 2017 | MicroAire Surgical Instruments, LLC | Mislabeling on three lots of K-Wires. The incorrect product description " dual trocar" instead of "single trocar" was on the product label on the following parts: 1600-9355NS (lot 61005), 1600-9625NS (lot 60297), and 1604-162NS (lot 62856). | Labeling mix-ups | The firm will mail recall letters to the customers who received the mislabeled k-wires on February 21, 2017. Letters will be sent by UPS to expedite the delivery. The firm will contact the non-responding accounts at least two more times. |  | 362 | Total 362 | United States. | Terminated | July 14, 2017 | 143 | 362 | Not_Computer
Z-3066-2017 | 77518 | [6 empty] | Tissu-Trans SFILL 360, 3-TT-SFILL 360. |  |  | June 01, 2017 | 2017 | Shippert Medical Technologies | Seal integrity of the Tissu-Trans product sterile packages cannot be assured. | Packaging | Customers were notified of the recall between 06/01/2017 and 06/14/2017 via FedEx letter. Instructions included remove and quarantine all affected devices, complete and return the response form, arrange for the return of affected devices, and notify any customers if further distributed. |  | 5188 | 5,188 units total | Worldwide Distribution - US Nationwide in the states of: AL, AR, AZ, CA, CO, CT, FL, GA, HI, IL, IN, KS, KY, MD, MI, MN, MO, NC, NJ, NV, NY, OH, OR, PA, TN, TX, UT, VA, WA, WV. and the countries of: Italy, Austria, Turkey, United Arab Emirates, Germany, Spain, Saudi Arabia, Switzerland, Japan Kuwait, Denmark, Sweden, United Kingdom, Belgium, France, Australia, Netherlands, Greece, South Korea. | Open, Classified |  |  | 5188 | Not_Computer
Z-2566-2017 | 76629 | No | KRO | Knee joint femorotibial metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Segmental malefemale taper, 200mm, Sterile, |  |  | February 28, 2017 | 2017 | Zimmer Biomet, Inc. | Zimmer Biomet Inc. initiated a field action on 02/28/2017 for sterile-packaged implants packaged in two different package configurations due to packaging design verification test failures. | Packaging | Zimmer Biomet initiated a field action for sterile-packaged implants packaged in two different package configurations due to packaging design verification test failures. Zimmer Biomet sent letters via Fed-Ex and e-mail dated 02/28/2017 to customers informing them of the recall, and instructing them to do the following: 1. Review this notification and ensure affected team members are aware of the contents.2. Complete the Certification of Acknowledgement portion of Attachment 1a. Return a digital copy to corporatequality.postmarket@zimmerbiomet.com within three(3) days.3. Assist the Zimmer Biomet sales representative quarantine all affected product.4. If after reviewing this notice customers have further questions or concerns, please call 574-371-3071 between 8:00 am and 5:00pm EST, Monday through Friday. Calls received outside of the call center operating hours will receive a prompt to record a voicemail or be transferred to an oncall representative in the case of an emergency. Alternatively, questions may be sent byemail to corporatequality.postmarket@zimmerbiomet.com.Adverse reactions or quality problems experienced with the use of this product may be reported to the FDA:Âż MedWatch Reporting: Adverse reactions or quality problems experienced with the use of this product may be reported to the FDAs MedWatch Adverse Event Reporting program either online, by mail, or by fax.Âż Online: www.fda.gov/medwatch/report.htmÂż Mail: Use postage paid, pre-addressed form FDA 3500, available at: www.fda.gov/MedWatch/getforms.htmÂż Fax: 1-800-FDA-0178Under 21 CFR 803, manufacturers are required to report any serious injuries where a product has contributed or may have contributed to the event. Please keep Zimmer Biomet informed of any adverse events associated with this product or any other Zimmer Biomet product by emailing zimmer.per@zimmerbiomet.com.For questions regarding this recall call 574-371-3071. |  | 9867 | 9867 | Worldwide Distribution - US (nationwide) Foreign: Argentina, Australia, Belgium, Brazil, Bridgend Wales, Canada, Chile, Colombia, Dominican Republic, England, Eschbach, France, India, Italy, Japan, Jordan, Korea, Malaysia, Mexico, Netherlands, Nicaragua, Panama, Peru, Russia, Saudi Arabia, Singapore, Spain, Switzerland, Taiwan, Thailand | Open, Classified |  |  | 9867 | Not_Computer
Z-1402-2017 | 76318 | No | LRJ | General purpose disinfectants. | General Hospital | General Hospital | 510(K) Exempt | CaviWipes1 Extra Large, Part No. 13-5150 |  |  | January 20, 2017 | 2017 | Metrex Research, LLC. | Metrex is recalling the CaviWipes Extra Large because they may have been contaminated during the packaging process. | Packaging process control | Metrex sent an Urgent Medical Device Recall letter dated January 20, 2017, to all affected customers to inform them that Metrex is recalling one lot of CaviWipes Extra Large because they may have been contaminated during the packaging process. Customers were informed that while handling sealed containers, a production worker sustained a minor finger cut which may have come in contact with the exterior surface of the container. Customers were instructed to contact Metrex Customer care at (800) 537-7123 to receive a RMA number. The RMA will allow for a quick return and replacement or credit. Customers were instructed to identify any customers that may have been shipped the affected product and contact their customers using the customer notification letter that Metrex issued. Metrex requests that those customers contact Metrex Customer Care at (800) 537-7123 and complete the attached acknowledgement form.For questions regarding this recall call 734-947-6700. |  | 960 | 960 cases | Nationwide Distribution to PA, FL, IL, MN, TX, AL, NY | Open, Classified |  |  | 960 | Not_Computer
Z-2799-2017 | 77671 | No | NRY | Percutaneous catheter. | Cardiovascular | Neurology | 510(k) | Stroke Fast Pack(TM) Trevo XP ProVue Retriever 3 x 20, Trevo Pro 14 Microcatheter, AXS Catalyst 060x132CM - US, UPN M0033PK32021001 |  |  | June 26, 2017 | 2017 | Stryker Neurovascular | Stryker Neurovascular has become aware that some 3-Pack Stroke Fast Packs weremanufactured using a carton sleeve where the pre-printed contents on the back of the sleeve did not match the physical contents of the pack. | Packaging process control | Stryker notified their consignees by letter on 06/21/2017. The letter stated the following:This potentially impacts all 3-pack Stroke Fastpacks manufactured in the US. The product quality of the individual products in the Stroke Fastpack is not impacted. All units were manufactured to specification.We request that you read this notice carefully and complete the following actions:1. Immediately check your internal inventory for impacted Catalog numbers.2. Remove and discard the Stroke Fastpack carton sleeve.3. Circulate this Field Safety Notice internally to all interested/affected parties.4. Maintain awareness of this notice internally until all required actions have been completed within yourfacility.5. Inform Stryker if any of the subject devices have been distributed to other organizations.a) Please provide contact details so that Stryker can inform the recipients appropriately.6. Please inform Stryker of any adverse events concerning the use of the subject devices.7. Complete the attached customer response form. It may be that you no longer have any physical inventory on site. Completing this form will allow us to update our records and will also negate the need for us to send any further unnecessary communications on this matter. Therefore please complete the form even ifyou no longer have any of the subject devices in your physical inventory.8. Return the completed form to your nominated Stryker Representative or to NVFieldActions@stryker.com |  | 43 | 43 packs | nationwide | Open, Classified |  |  | 293 | Not_Computer
Z-0961-2017 | 75838 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | GoDirect Screw-Retaining Top |  |  | November 28, 2016 | 2016 | Implant Direct Sybron Manufacturing, LLC | Consignees received product 6060-62, Replant Overdenture Abutment with the label for 4500-61 GoDirect Screw-Retaining Top. | Labeling mix-ups | On 11/28/16 customers were called and instructed to return the product. |  | 14 | 14 | U.S. distribution to the following; WA.Foreign distribution to the following; Netherlands. | Open, Classified |  |  | 14 | Not_Computer
Z-1484-2017 | 76607 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | AMS-636, 60" Microbore extension set. |  |  | February 20, 2017 | 2017 | Vygon MFG, Inc., dba/ Churchill Medical Systems, Inc. | Incorrect priming volume printed on the device package. | Error in labeling | Recall letters were sent to the customer via email and fax on 2/20/2017. They were asked to take the following actions: The recall will be conducted at no cost to you. To effect return of the recalled product and replacementwith new inventory, we request your cooperation in the following steps:1. Check your stock for the products included within the scope of the recall. Immediately cease use and distribution, and quarantine all affected product immediately; 2. Count and document your affected inventory on the "Recall Acknowledgement and Inventory Return Form; 3. Fax or email the form to Vygon including customer contact information; 4. Vygon customer service will contact the customer and issue a Return Merchandise Authorization number; 5. Return the product as well as the original signed Recall Acknowledgement and InventoryReturn Form using the instructions from Customer Service involving product pick-up; 6. All inventory will be credited or replaced at no cost to you. If you have questions, please contact your local sales representative or Vygon's Customer Service Department at 1-800-4 73-5414, or by e-mail at customerservice@vygonus.com. |  | 30700 | 30,700 | US Nationwide Distribution in the states of NJ, TX, VA, FL, VA, MA, IL, MI, NY, OH, WA, PA, WV | Open, Classified |  |  | 30700 | Not_Computer
Z-2866-2017 | 77653 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | OptiLITE(TM) Products for Laser Surgery, Catalog Numbers: HTM1020F, HTM1040F, HTM1060F, HTM1010F; Manufactured by CONVERGENT LASTER TECHNOLOGIES |  |  | June 23, 2017 | 2017 | Xintec Corp | It has been determined that the reprocessing instructions may not provide sufficiently detailed information for the cleaning, disinfection and sterilization of the products listed below. The preliminary investigation indicates that validation data related to the reprocessing of these devices may not meet newly enacted standards. | Labeling False and Misleading | The firm, Convergent Laser Technologies, sent an "URGENT - MEDICAL DEVICE RECALL" letter dated 06/23/2017 to its Consignees. The letter described the product, problem and actions to be taken.The consignees were instructed as follows: "Please immediately review your inventory and quarantine any affected product. Please return any affected products within 30 days according to the instructions provided in the attached Acknowledgement and Receipt Form. Upon receipt of the product and response form, we will arrange for a credit to your account. Also, whether or not you are returning any affected product, please send a copy of the Acknowledgement and Receipt Form to our Regulatory Department via fax at 510-832-1600 or via e-mail to regulatory@convergentlaser.com."If you have any further questions or concerns, please contact our Customer Service Department at 510-832-2130 or 800-848-8200. |  | 4343 | 4343 units | Worldwide distribution- US (Nationwide) and countries of: Canada and Sri Lanka. | Open, Classified |  |  | 4343 | Not_Computer
Z-0967-2017 | 75952 | No | KNX | Urine collector and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | SURESTEP Tray, BARDEX I.C. COMPLETE CARE, Urine Meter, Temperature STATLOCK Stabilization Device |  |  | July 11, 2016 | 2016 | C.R. Bard, Inc. | The product contained a latex catheter instead of a silicone catheter. | Labeling mix-ups | Bard Medical sent an Urgent : Medical Device Product Recall Letter dated July 11, 2016, to all affected customers. The letter requested that they check inventory and quarantine product, and if the product had been further distributed, they need to conduct a sub-recall. The letter included a reply form to be returned. For further questions, please call (770) 784-6220. |  | 2690 | 2690 units | US Distribution to the states of : CA, CT, FL, GA, IL, IN, KS, KY, MD, MA, MI, NE, NV, NJ, NM, NY, OR, SC, TN, TX, UT, VA, WA | Open, Classified |  |  | 2690 | Not_Computer
Z-1543-2017 | 76593 | No | MIH |  |  | Cardiovascular | PMA | Endurant / Endurant II 23mm and 25mm Bifurcated Stent Graft Systems |  |  | February 27, 2017 | 2017 | Medtronic Vascular, Inc. | Medtronic is initiating a voluntary Urgent Medical Device Recall for a subset of Endurant/ Endurant IIBifurcated Stent Graft Systems of specific models and serial numbers as this specific subset of stent grafts has greater susceptibility to fabric permeability variations that may be associated with endoleaks observed during the initial implant procedure. | Packaging change control | Medtronic sent an Urgent Medical Device Recall letter dated March 2017, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Our records indicate that your facility has received a potentially affected Endurant /Endurant II 23mm and / or 25mm Bifurcated Stent Graft System from the identified subset. As a result, Medtronic is asking that you take the following actions:Identify and quarantine unused Endurant /Endurant II 23mm and 25mm Bifurcated Stent Graft System from the identified subset of models and serial numbers that are in your inventory .Return unused products from the attached model and serial number list that are in your inventory to Medtronic by contacting Customer Service at 1-800-854-3570 (#4) and referencing this communication to initiate a return and credit of unused product. Your Medtronic sales representative can assist you in the return of this product as necessary.Complete the attached Customer Confirmation Certificate and email it to RS.CFQFCA@Medtronic.com or fax it to Medtronic at 651-367-0612 to the attention of Customer Focused Quality.For further questions please call Customer Service at 1 (800) 633-8766. |  | 3331 | 3331 devices | Worldwide Distribution - US Distribution including Puerto Rico and to the countries of :ArgentinaArmeniaAustraliaAustriaBelgiumBrazilBulgariaCanadaChinaColombiaCosta RicaCroatiaCyprusCzech RepublicDenmarkDominican RepublicEcuadorEgyptEl SalvadorFinlandFranceFrench PolynesiaGermanyGreeceHong KongHungaryIcelandIndiaIndonesiaIranIrelandIsraelItalyJapanJordanKazakhstanKoreaKuwaitLatviaLebanonLithuaniaLuxembourgMalaysiaMexicoMontenegroNetherlandsNew ZealandNorwayOmanPanamaParaguayPeruPhilippinesPolandPortugalReunionRomaniaRussian FederationSaudi ArabiaSerbiaSingaporeSlovakiaSloveniaSouth AfricaSpainSri LankaSwedenSwitzerlandTaiwanThailandTurkeyUnited Arab EmiratesUnited KingdomUruguayVenezuelaViet Nam | Open, Classified |  |  | 3331 | Not_Computer
Z-1221-2017 | 76138 | No | GBZ | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Atrium PVC 8 Fr Straight Thoracic Catheter |  |  | January 20, 2017 | 2017 | Atrium Medical Corporation | Labeling icon depicts a catheter having six (6) eyelets; however the correct number of eyelets on the 8 Fr and 10 Fr thoracic catheters is two (2) | Labeling design | Marquet/Atrium notified consignees by letter dated 01/20/17 informing of the Labeling Icon Clarification. The thoracic catheter labeling is being updated for future production. Users should be made aware of this Field Notification and may continue to use the affected thoracic catheter products with the current labeling.Questions, please contact your local Maquet/Atrium thoracic catheter representative or Maquet Customer Service at 1-800-528-7486 or 603-880-1433, Monday through Friday between 8:00 am and 5:00 pm (EST). |  | 660 | 660 cases (10/case) | Nationwide | Open, Classified |  |  | 2054 | Not_Computer
Z-1465-2017 | 76433 | Yes | JOH | Tracheostomy tube and tube cuff. | Anesthesiology | Anesthesiology | 510(k) | BivonaÂż TTS (Tight to Shaft) Pediatric Tracheostomy Tube, Model No. 67SP035, Reorder No. 67SP035. |  |  | February 09, 2017 | 2017 | Smiths Medical ASD Inc. | Smiths Medical became aware that one lot of the BivonaÂż 3.5mm TTS Cuffed Pediatric tracheostomy tubes was incorrectly labeled for neonatal use instead of correctly labeled for pediatric use. | Labeling False and Misleading | Consignees were sent a Smiths Medical "Urgent Medical Device Field Safety Notice" dated February 08, 2017. The letter was addressed to "Clinician". The letter listed the Affected devices, Reason for Recall, Risk to Health, and Instructions to Customers. Advised customers to complete and return the "Urgent Medical Device Field Safety Notice (FSN) Response Form" to smithsmedical4390@stericycle.com. All affected devices must be returned to Stericycle for processing. If the product was further distributed, consignees were asked to immediately notify them of the Recall and retrieve all affected devices. For questions contact Stericycle via e-mail at the following address: smithsmedical4390@stericycle.com. |  | 142 | 142 | US: CA, ILM NC, GA, NE, MA, CO, CT, MO, TN, LA, MD, MN, FL, SC,M OK, MI, PA, TX, WA, WI. | Open, Classified |  |  | 142 | Not_Computer
Z-1408-2017 | 76327 | No | MMP | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General Hospital | 510(k) | Premier Guard Sterile Drape For Use with Blair-Port Wand (02-0029) and Blair-Port Wand X (01-0046) |  |  | January 23, 2017 | 2017 | Covidien LLC | Sterility compromised due to breach in sterile barrier | Packaging process control | Medtronic issued recall letter on January 23, 2017. Customers were notified via Federal Express and the letter instructed customers to discontinue use of the product and return all inventory to Medtronic (Covidien)  195 McDermott Road, North Haven, CT 06473 Attn: Field Returns Department. |  | 1111600 | 1,111,600 units | Nationwide | Open, Classified |  |  | 1154350 | Not_Computer
Z-1126-2017 | 76130 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Intact Disposable Wands (part of the Intact System) |  |  | December 28, 2016 | 2016 | Medtronic Advanced Energy, LLC | Product sterility compromised due to breach of sterile barrier | Package design/selection | Medtronic Advanced Energy notified customers by letter on December 27, 2016 via Next Day FEDEX. The communication advises users to immediately stop using the affected Intact Disposable Wands and to quarantine them until they can be returned to Medtronic. Medtronic Sales representatives will arrange for the return of all recalled product. |  | 10 | 10 units | Worldwide Distribution - US Nationwide in the countries of China, Greece, Netherlands, Egypt, Turkey, and Switzerland. | Terminated | May 11, 2017 | 134 | 3303 | Not_Computer
Z-1485-2017 | 76571 | bag, urine collection, leg, for external use, non-sterile | NNW | Urine collector and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | PureWick External Catheter for Women-Wick |  |  | February 24, 2016 | 2016 | Purewick Corporation | PureWick is recalling external catheter for women Wicks because the labeling did not contain the phrase "This product contains dry natural rubber" and/or erroneously included the phrase "Latex-free." The white gauze component of the PureWick contains dry natural rubber. | Labeling design | A Urgent Medical Device recall letter dated March 2, 2017 will be sent to customers to inform them that PureWick Corporation is recalling specific lots of PureWick external catheter for women Wicks, Item #625. Customers are informed that PureWick is recalling PureWick Wicks because labeling on product shipped prior to February 22, 2017 either did not contain the phrase This product contains dry natural rubber and/or erroneously included the phrase Latex-free. The white gauze component of the PureWick Wick contains dry natural rubber. PureWick Corporation has received zero complaints or reports of adverse events related to the error in labeling. Customers are informed of the risk to health and the actions to be taken. Customers are instructed to complete the enclosed PureWick Field Correction form and fax to PureWick Corporation at (619) 660-5459 or email to response@purewick.com. Customers with questions or would like assistance completing the PureWick Field Correction Form are instructed to call PureWick toll-free at (844) 584-0734 Monday  Friday from 8:00 a.m. to 4:30 p.m., Pacific Time. |  | 176863 | 176,863 units | US: Nationwide Distribution | Terminated | May 09, 2017 | 440 | 176863 | Not_Computer
Z-1446-2017 | 76476 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | LeMaitre Vascular Reddick Scoop Tip Cholangiogram Catheter (eIFU) |  |  | February 07, 2017 | 2017 | LeMaitre Vascular, Inc. | The Contains Latex symbol was not printed onto the box label or the sterile (inner) product pouch | Labeling design | LeMaitre Vascular, Inc. issued a Dear Doctor letter, on 2/7/17, to distributors and hospitals via FedEx overnight. The letter contains a form that is requested to be returned to LeMaitre Vascular, Inc. as a record of notification and reconciliation. Any unused product requested to be returned and replaced.Questions concerning this notice contact 781-221-2266 ext. 183. |  | 4673 | 4673 | Worldwide distribution. US Nationwide, AUSTRIA, BELGIUM, FRANCE, GERMANY, ITALY, LUXEMBOURG, Morocco, Saudi Arabia, Singapore, SPAIN, SWITZERLAND, UAE, and UK | Open, Classified |  |  | 21150 | Not_Computer
Z-1411-2017 | 76158 | No | OLO | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | MAKO RIO THA Application User Guides, for total hip arthroplasty, orthopedic surgery. |  |  | December 22, 2016 | 2016 | Mako Surgical Corporation | Five (5) AFMEA Risk Control Measures are missing from the MAKO RIO THA Application User Guides | Labeling design | Letter from Stryker dated 1/28/2017 was sent via USP Next Day Air with return receipt to all their customers. With the following actions; Inform users of this medical device correction and forwar this notice to all those individuals who need to be aware within your organization and return the attached Business Replay Form within 5 business days, to Fax 844-782-5561 or via e-mail to stykerortho4253@stericycle.com |  | 777 | 777 | US nationwide distribution. | Open, Classified |  |  | 777 | Not_Computer
Z-3046-2017 | 77917 | sudan iv | KKK | Dye and chemical solution stains. | Hematology | Pathology | 510(K) Exempt | Convertors Robotic Drape with Attached Leggings; Product is packaged as Single Sterile Drape/8 per Case |  |  | July 28, 2017 | 2017 | Cardinal Health 200, LLC | Product 29474 LAVH drape was packaged as 29574 Robotic Drape with Leggings. | Packaging process control | Cardinal Health sent an Urgent: Product Recall letter dated July 28, 2017.Distributors are requested to notify their customers of the action. An acknowledgement form is included as part of the notification. Responses are tracked. All returned product will be quarantined. All inventory recovered from the field will be destroyed. For further questions, please call (800) 292-9332. |  | 200 | 200 units | US Distribution to the states of : NC, UT, OR, NY,TX, GA, NJ and FL. | Open, Classified |  |  | 200 | Not_Computer
Z-3030-2017 | 77696 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | VersaSafe Plastic Cannula - 11 Gauge, Model 9391-0200 |  |  | June 30, 2017 | 2017 | CareFusion 303, Inc. | Punctures in blister packaging that was detected during the packaging process. | Packaging process control | BD sent an Urgent Medical Device Recall letter dated June 2017 to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to discontinue use and return the affected lot to CareFusion for replacement. For recall related questions contact BD Support Center 1-888-562-6018. |  | 1546800 | 1,546,800 units | Worldwide Distribution in the states of : CO, IL, IN, MD, MI, NY, OH, &amp; TX. and foreign countries of: Australia, New Zealand, &amp; Japan. | Open, Classified |  |  | 1546800 | Not_Computer
Z-1416-2017 | 76153 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Roth Net Polyp Retrieval |  |  | December 13, 2016 | 2016 | US Endoscopy Group Inc | The firm was notified that the expiration date on the outer carton is different than the expiration date on the inner package. The expiration date is shorter than the actual expiration. | Labeling mix-ups | US Endoscopy Group sent an Urgent: Roth Net Polyp Retriever Recall letter dated January 6, 2017, to all affected customers.The firm requested their customers to do the following:1. Quarantine and destroy the affected Roth Net devices in your inventory.2. Contact US Endoscopy Customer Service at 1-800-769-8226 to obtain no cost replacement product.3. Complete and return the attached Response Card.For further information please call (440) 358-6055. |  | 2131 | 2,131 units | US Distribution to the states of : AL, CA, CO, IL, KS, MA, MT, NC , NE , NY, OH, PA, TN, TX and WA.Product was also shipped to the following countries: Australia, Belgium, Canada, Czech Republic, Ireland, Kuwait, Norway, Poland &amp; Sweden. | Terminated | June 13, 2017 | 182 | 2131 | Not_Computer
Z-1684-2017 | 76501 | No | GIM | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | Cyto-Chex BCT - blood specimen collection device - Immunophenotyping Preservative. IVD. Sterile. 2.0ml only. Packaged 6-tube pack REF 213559 and 100-tube box REF 218980. Streck. |  |  | February 23, 2017 | 2017 | Streck | Complaints of the 2.0ml tube stopper coming loose during phlebotomy when small diameter tube adapters were not utilized. Such adapters aid with centering small diameter tubes to properly pierce the middle of the stoppers. | Error in labeling | On February 23, 2017, Streck electronically notified customers of the updated device correction. The recalling firm used an electronic mail alert to notify all of their customers the IFU for the recalled product had been revised. The specified the revision included phlebotomy guidelines specific to the 2.0 ml blood collections tubes. The revised IFU was provided as a link in the email. The recalling firm is reportedly able to monitor the access of the revised IFU to ensure each customer has received the information. For further questions, please call (402) 333-1982. |  | 128692 | 128,692 tubes | Worldwide Distribution - USA (nationwide) and to the countries of : Belgium, Canada, Germany, Israel, Japan, the Republic of Korea, the Netherlands, and the United Kingdom. There was no military or government distribution. | Open, Classified |  |  | 128692 | Not_Computer
Z-1441-2017 | 76473 | No | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | RTI Biologics BioSet IC RT Paste 2 cc |  |  | January 23, 2017 | 2017 | RTI Surgical, Inc. | Incorrect instructions for use (IFU) included. | Labeling Change Control | The firm, RTI Surgical, Inc., issued and "IMPORTANT NOTIFICATION Field Action" letter dated January 20th, 2017 to its consignee via e-mail on 01/23/2017. Verification and confirmation of the e-mail was received on the same date. The letter described the product, problem and actions to be taken. The consignee were instructed to immediately quarantine any grafts in your possession; return any unused RTI and include return authorization along with instructions on how to return the product.If you have any questions, Sr. Manager, Regulatory Compliance at 386-418-8888 ext. 4396 or email: amcconnell@rtix.com. |  | 24 | 24 units | International Distribution: Colombia only | Open, Classified |  |  | 24 | Not_Computer
Z-0955-2017 | 75782 | No | KLE | Resin tooth bonding agent. | Dental | Dental | 510(k) | VOCO Futurabond M+ adhesive, RF 1515, Manufactured by VOCO GmbH, Futurabond M+ is a dental adhesive for use with methacrylate-based restorative, core build-up or luting materials. It can be applied with phosphoric acid etching (selective enamel etching or total-etch) of the dental hard tissue as well as without (self-etch). A1] Adhesive for direct restorations of all classes of cavities using light-curing composite /compomer / Ormocer materials with a methacrylate base. [A2] Adhesive for core build-ups made of light-curing composite materials. [A3] Adhesive for indirect restorations with light-curing luting composites. [B] Intraoral repair of composite, compomer or Ormocer restorations, ceramic veneers and allceramic restorations without an additional primer. [C] Treatment of hypersensitive tooth necks. [D] Protective varnish for glass ionomer cement restorations. [E1] Sealing of cavities prior to amalgam restorations. [E2] Sealing of cavities and core preparations prior to the temporary luting of indirect restorations. In combination with Futurabond M+ DCA (Dual Cure Activator): [F1] Adhesive for direct self-curing or dual-curing composite restorations of all classes of cavities and for core build-ups. |  |  | November 03, 2016 | 2016 | Voco GmbH | Labeling mix-up: Futurabond M+ DCA (Dual Cure Activator) was marked labeled as Futurabond M+ Universal Adhesive. | Labeling mix-ups | Voco Gmbh notified consignees via email on/about November 3, 2016. The consignees were informed about the incorrect labeling and requested not to use this affected lot as well as to return it. They were also requested to pass this information to their customers |  | 50 | 50 units | Nationwide Distribution to FL, IA, NH, PA, SC and TX | Open, Classified |  |  | 50 | Not_Computer
Z-1685-2017 | 76619 | No | LHL | Haemophilus spp. serological reagents. | Microbiology | Microbiology | 510(K) Exempt | Legionella IFA (IF0950) intended for qualitatively detecting and semi-quantitating human serum antibodies to Legionella pneumophila. |  |  | February 15, 2017 | 2017 | Focus Diagnostics Inc | After a submission for CLIA database update from Focus Diagnostics to DiaSorin Molecular,the company found that their Legionella IFA device (IF0950) is not 510(k) exempt. | No Marketing Application | Focus Diagnostics/DiaSorin Molecular is issuing a medical device correction for the Legionella IFA (Immuno Fluoresence Assay). On February 1, 2017, the firm was notified that the Legionella IFA was not properly classified according to 21 CFR Part 807. The product is being withdrawn from the market. Use of the product should be immediately discontinued, and any remaining inventory should be immediately destroyed. The firm sent notice to its consignees, which should be reviewed with the facility's laboratory or mediclal director. Patient results obtained with all lots of the affected product should be evaluated by the clinician and if recommended, retested by another method. |  | 196 | 196 units | Product was distributed to three testing laboratories in California, New Jersey and Virginia | Terminated | April 13, 2017 | 57 | 1596 | Not_Computer
Z-1739-2017 | 76606 | No | HMP | Eye pad. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Covidien Curity Eye Pad |  |  | March 03, 2017 | 2017 | Covidien LLC | Product sterility is compromised due to breach of sterile barrier | Packaging process control | Medtronic/Covidien) initiated recall by letter on March 3, 2017 via Federal Express. The letter instructs customers to discontinue use of the product and return all inventory to Medtronic (Covidien), 195 McDermott Road, North Haven, CT 06473, Attention: Field Returns Department. Contact Medtronic representative or Customer Service at (800) 882-5878For questions regarding this recall call 203-492-5000. |  | 7168000 | 7168000 | Worldwide Distribution. US (Nationwide), Internationally to Canada, Australia, Brunei, Japan, Mexico, Philippines, Singapore, Andorra, Belgium, Botswana, Czech Republic, Denmark, Egypt, France, Greece, Hungary, Ireland, Israel, Italy, Kenya, Libya, Luxembourg, Mauritius, Namibia, Netherlands, Norway, Poland, Portugal, Russia, Russian Federation, Spain, Saudi Arabia, Seychelles, Slovakia, South Africa, Switzerland, UAE, and United Kingdom. | Open, Classified |  |  | 70254852 | Not_Computer
Z-1243-2017 | 76414 | No | HAB | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Smith &amp; Nephew Dyonics Saw Blade, Narrow, Product Number 3705 |  |  | February 02, 2017 | 2017 | Smith &amp; Nephew, Inc. | The single use devices are provided sterile and do not have an expiration date on the label. | Labeling Change Control | Smith &amp; Nephew, Inc. initiated recall by letter on 2/2/17 advising of the problem, requesting inspection of inventory and return of product. Contact Smith &amp; Nephews Field Actions Department via email at FieldActions@smith-nephew.com or fax +1-901-566-7975 to return. |  | 3462 | 3462 | NationwideForeign:Brazil, Canada, Chile , China, Denmark, Hong Kong, India, Mexico, Peru, Singapore, Venezuela | Open, Classified |  |  | 23262 | Not_Computer
Z-1238-2017 | 76362 | No | EGJ | Iontophoresis device. | Physical Medicine | Physical Medicine | 510(k) | Daavlin Aquex (DAAV1000) |  |  | January 30, 2017 | 2017 | Hidrex GmbH | The device was reclassified from a class III device to a class II device and there are new documentation requirements. | Error in labeling | Distributor sent recall notification letter to affected dealers and end users on 2/10/17. Notification letter will include a copy of the addendum to the user manual. The device is not instructed to be returned. |  | 250 | 250 units | Nationwide throughout the US | Terminated | August 04, 2017 | 186 | 494 | Not_Computer
Z-1041-2017 | 75966 | No | OMP | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Smith &amp; Nephew, RENASYS GO Negative Pressure Wound Therapy, Rx only. |  |  | November 11, 2016 | 2016 | Smith &amp; Nephew, Inc. | Lack of 510k clearance for design modification. | No Marketing Application | On letter sent to all their consignees, Smith &amp; Nephew asked to the consignees to removed all existing Renasys Go NPWT Pumps and returned following the instructions on the letter and also removed the affected power supply units identified on the recall letter. Each consignee needs to complete the Acknowledge |  | 15960 | 15,960 units in the US | Nationwide | Open, Classified |  |  | 15960 | Not_Computer
Z-1145-2017 | 75557 | No | LGZ | Infusion pump. | General Hospital | General Hospital | 510(k) | HOTLINE Blood and Fluid Warmer, Mole l HL-90. Designed for use with the Hotline Fluid Warming Set (L-80, L-70 and L-70N) to warm blood and intravenous (I.V.) fluids and deliver them to the patient's intravenous access site at normothermic temperatures under gravity flow conditions. |  |  | October 31, 2016 | 2016 | Smiths Medical ASD, Inc. | HOTLINE Blood and Fluid Warmers included Greek and Danish language labeling instead of the required Finnish labeling. | Labeling mix-ups | Smiths Medical sent an Urgent-Medical Device Field Safety Notice dated October 28, 2016, to all affected customers. The letter described the Reason for Field Corrective Action, Risk to Health, and Instructions to Customers. Advised consignees to notify customers and to review and complete the Recall Confirmation Form and return it to Smith Medical via e-mail to FCA.Response@Smiths-medical.com. For further questions, please call (651) 633-2556. |  | 33 | 33 | Internationally to Finland | Open, Classified |  |  | 33 | Not_Computer
Z-1150-2017 | 76146 | No | GCY | Nonpowered, single patient, portable suction apparatus. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | 9805 |  |  | January 03, 2017 | 2017 | Stryker Leibinger GmbH &amp; Co. KG | An incorrect sterility status on the label, non-sterile products labeled as sterile | Labeling mix-ups | Stryker sent an " URGENT MEDICAL DEVICE RECALL " notifications dated January 3, 2017, to the affected customers. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, instructions for returning affected product, and instructions for responding to the formal recall notification.Please report any adverse events or product quality problems to Stryker CMF Customer Service: 1 (800) 962-6558. |  | 619 | 619 | Worldwide Distribution - USA (nationwide) Distribution and to the countries of :Australia, Canada, Chile, Spain, South Africa and Taiwan. | Terminated | July 27, 2017 | 205 | 619 | Not_Computer
Z-1425-2017 | 76228 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Biograph 16 TruePoint TrueV - 4R, Material Number 10249556 |  |  | February 20, 2017 | 2017 | Siemens Medical Solutions USA Inc. | The mobile-specific instructions was not shipped to a limited population of Biograph Mobile systems. | Labeling Change Control | Siemens sent a Customer Advisory Notice letter on February 20, 2017 to affected customer. The letter identified the affected product, problem and actions to be taken. For questions contact your local Siemens. |  | 4 | 4 units | Worldwide distribution - US Nationwide in the states of CA, CT, FL, GA, IL, KY, MA, MI, ND, NJ, OH, OK, OR, PA, TN, WA, WI, WY and the countries ofSwitzerland, Great Britain, Italy | Open, Classified |  |  | 74 | Not_Computer
Z-0983-2017 | 75915 | No | OHW | Powered corneal burr. | Ophthalmic | Ophthalmic | 510(K) Exempt | Algerbrush-II |  |  | October 12, 2016 | 2016 | Bausch &amp; Lomb, Inc. | According to firm ( Bausch &amp; Lomb, Inc. ) on June 17, 2016 a packaging operator notified his supervisor he had not placed the operating instruction into the last repackaging of SKU 0816 C (Pterygium Burr Medium Diamond). Further investigation found that repackaged Algerbrush-II and its accessories from April 30, 2012 to present, contained out of date operating instruction. Therefore Bausch &amp; Lomb is proposing a Field Correction to provide customer purchasing product from April 30, 2012 to June 20, 2016 with the current operating instruction. | Error in labeling | Bausch + Lomb sent an Urgent Medical Device Voluntary Field Correction letter dated October 12, 2016 to affected customers. The letter identified the affected product, problem and actions to be taken. For questions contact Bausch + Lomb customer service at 1-800-338-2020 |  | 2487 | 2487 units Domestic (USA), 1,211 units Foreign (outside USA) | Worldwide Distribution - US Nationwide | Terminated | July 14, 2017 | 275 | 2487 | Not_Computer
Z-1539-2017 | 76693 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Panta Nail, Rx only, Sterile, |  |  | March 03, 2017 | 2017 | Integra LifeSciences Corp. | Voluntary recall/Removal of Panta and Panta XI Nails because there may be the potential that the internal and external package seal does not provide the level of sterility intended. | Packaging change control | Integra sent "Urgent Medical Device Recall Letters and Return Response Forms" dated March 3, 2017 to their distributors and customers.All customers were asked to examine their inventory to determine if they had the recalled lots and if they did they are instructed to stop using them immediately. Distributors are asked if the recalled product was further distributed, those customers should be notified and instructed to follow the recalling procedures. The return response form should be completed, even if they do not have any recalled lots, annotating the amount of product and list a record of the lot numbers. The return response form should be sent back to Integra via email to: FCA2@integralife.com or via fax to: 1-609-750-4220. Upon receipt by Integra, if your company has identified recalled lots, Customer Service will contact the customer and provide an RMA number and directions to return the product, as well as input an order to replace the quantity indicated on the return response form. Should any customer have any questions regarding the recall instructions, they can contact Customer Service at 1-800-654-2873. To return units under field inventory, customer can contact 1-888-601-0203, option 2.For further questions, please call (717) 309-3966. |  | 4449 | 4,449 units | US Nationwide and Internationally | Open, Classified |  |  | 4449 | Not_Computer
Z-1535-2017 | 76365 | No | ORZ | Eyelid thermal pulsation system. | Ophthalmic | Ophthalmic | 510(k) | LipiFlow Thermal Pulsation System/LipiFlow, System Activator (Disposable), Rx Only, Sterile |  |  | January 30, 2017 | 2017 | Tearscience, Inc | Labeling: The labeled storage temperature is not consistent with the temperature conditions tested for the entire four year shelf life duration. | Labeling Change Control | Tearscience sent an MEDICAL DEVICE LABELING CORRECTION/RECALL letter dated January 30, 2017..Actions to be Taken by Customer: Within 10 business days of receipt of this letter, please take the following actions: 1) Read the enclosed Summary of Important Labeling Changes for the LipiFlowÂż System. 2) Complete the Acknowledgement and Receipt form at www.TSVerify.pmcprograms.com or return the enclosed completed form by email or mail. 3) If you have stored Activators in an environment that is not temperature controlled (without heat and air conditioning), discontinue use of these Activators and call TearScience Labeling Change Hotline at (844) 872-5658.For any questions regarding this letter, please call (844) 872-5658 Monday through Friday, 8:00AM to 6:00PM, Eastern Time. We want to thank you for being a valuable TearScience partner, and as always would like to thank you for your business, your trust, and your commitment to treating MGD.For further questions please call (919) 459-4831. |  | 94095 | 94,095 units. | US (nationwide) | Open, Classified |  |  | 94095 | Not_Computer
Z-3081-2017 | 77480 | No | HRM | Surgical microscope and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | LX3 Floor Stand for LuxOR Ophthalmic Microscope; Endure LuxOR Q-VUE Ophthalmic Microscope. |  |  | May 25, 2017 | 2017 | Alcon Research, Ltd. | Multiple shipments or forceful impact to a caster on a LuxOr/LX3 microscope may stress the microscopebase, which could contribute to a caster breaking or detaching from the base. | Packaging process control | Alcon sent an Urgent Medical Device Correction letter dated May 25, 2017, to all affected customers. Demonstration units in inventory will be scheduled for retrofit of base and/or upgraded caster configuration through onsite service calls. Customers with questions or concerns were instructed to call 1-800-832-7827.For questions regarding this recall call 817-293-0450. |  | 1074 | 1,074 units | Worldwide Distribution - US (nationwide and Puerto Rico): Units were also distributed globally to the following countries: Algeria, Argentina, Australia, Austria, Belarus, Belgium, Bosnia, Herzegovina, Bulgaria, Canada, Chile, Columbia, Costa Rica, Croatia, Curacao, Czech Republic, Denmark, Dominican Republic, Egypt, El Salvador, Ecuador, Finland, France, Georgia, Germany, Greece, Guatemala, Honduras, Hungary, India, Italy, Jamaica, Japan, Jordan, Kazakhstan, South Korea, Kuwait, Lebanon, Lithuania, Luxembourg, Malaysia, Mexico, Mongolia, Morocco, New Zealand, Nicaragua, Nigeria, Norway, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, UAE, Ukraine, United Kingdom, and Uzbekistan | Open, Classified |  |  | 1074 | Not_Computer
Z-2756-2017 | 77668 | No | PJC | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | PerkinElmer 226 Sample Collection Device |  |  | May 31, 2017 | 2017 | PerkinElmer Health Sciences, Inc. | PerkinElmer 226 Sample Collection Devices Lot Number 317114 had serial numbers printed on the pack labels that did not match the serial number range of the devices contained within the pack. | Packaging | PerkinElmer sent an Urgent Medical Device Correction letter dated May 30, 2017, to the sole consignee. PerkinElmer requested the consignee implement the following actions:1. Discontinue use of the pack label to track the device serial numbers distributed to your sample submitters.2. Verify submitter of devices received without provider information prior to demographic entry.3. Verify the serial number of the pack label matches the device serial numbers within the pack of your undistributed inventory prior to distributing to your sample submitter or,4. Return undistributed inventory to PerkinElmer for inspection , and rework if required.For further questions please call (864) 299-8787. |  | 99968 | 99,968 units | US Distribution to the state of : AL | Terminated | July 27, 2017 | 57 | 99968 | Not_Computer
Z-3077-2017 | 77802 | No | DXE | Embolectomy catheter. | Cardiovascular | Cardiovascular | 510(k) | Fogarty Adherent Clot Catheter Model Numbers 1408010 and 140806 |  |  | July 06, 2017 | 2017 | Edwards Lifesciences, LLC | The diameter of the Fogarty Catheter balloon, measures 6mm instead of 10mm. | Packaging | Recall notification letters were sent to affected customers on 7/6/2017. The recall notice included a response form and instructed consignee to return any unused devices.. |  | 200 | 200 units | Worldwide Distribution - US Nationwide GB, NL, DE, IT, FR, BE, ES, NO, CZ, BG, SE, AT, CH, PL, RU, FI | Open, Classified |  |  | 200 | Not_Computer
Z-1250-2017 | 76363 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | GrebSet Micro-Introducer Kit, model 7941, Sterilized with Ethylene Oxide. |  |  | January 12, 2017 | 2017 | Vascular Solutions, Inc. | Four Lots of GrebSet Micro-introducer Kits, model 7941, contain guidewires that have a shelf life shorter than the kit expiration date. | Error in labeling | Consignees were sent on 1/12/2017 a Vascular Solutions "Urgent Medical Device Recall" letter dated January 11th, 2017. The letter identified the problem and product involved in the recall. The letter provided "Your Immediate Action Is Required". This included to remove the product from inventory and to complete and return the VSI's Account Inventory Form to VSIs Customer Service Dept. |  | 633 | 633 | US Nationwide Distribution in the states of NH, ME, VT, NY, LA, MS, AL CA, IA, WA, OR, MD, DE, PA, NJ, ID, TX, OK, MO, WI, GA, FL, AR, OH, AZ. | Open, Classified |  |  | 633 | Not_Computer
Z-1233-2017 | 76368 | No | GBX | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Jackson-Pratt Channel Drain Silicone Round, Fully Fluted w/10FR Trocar. Sterile. Rx |  |  | January 31, 2017 | 2017 | Degania Silicone, Ltd. | Product sterility compromised due to breach of sterile barrier | Packaging process control | Degania Silicone Ltd. issued an Urgent Recall letter dated 1/31/17 and requested return of the product to Degania Silicone 14 Thurber Boulevard, Suite A, Smithfield, RI 02917, USA.. For questions regarding this recall call +1-401-349-5373. |  | 3040 | 3040 units | To PR only | Terminated | April 04, 2017 | 63 | 3040 | Not_Computer
Z-3050-2017 | 77823 | No | NCF | Ophthalmic refractometer. | Ophthalmic | Ophthalmic | 510(K) Exempt | ORA System with VerifEye, Catalog Number 8065998300 For use during intraocular lens surgery |  |  | June 30, 2017 | 2017 | Alcon Research, Ltd. | Some ORA Carts have the potential to return an incorrect IOL power measurement during cataract surgery. This issue appears to have been caused by a software coding error that results in the lens coefficients for an IOL model being downloaded from the Alcon server in an incorrect order. | Software design | The Market Action was initiated with Initial telephone contact to all affected customers starting June 30, 2017. These customers were informed of the issue, and advised not to use their ORA Carts for calculations with the lens model affected on their specific system. A confirmatory letter (Attachment 4) was also sent via overnight mail to these customers July 12, 2017. As of July 12, 2017, Alcon has identified 8 ORA Carts that are affected by the software coding error. Alcon has reset the IOL databases on all 8 of the identified affected ORA Carts. |  | 367 | 367 units | Worldwide distribution. US Nationwide and Argentina, Australia, Belgium, Brazil, Canada, China, Colombia, France, India, Japan, Mexico, Netherlands, Panama, Portugal, Romania, Spain, Thailand, UAE, and United Kingdom | Open, Classified |  |  | 796 | Computer
Z-2685-2017 | 77267 | No | JJN | Electrophoresis apparatus for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Multiphor II Electrophoresis System |  |  | May 15, 2017 | 2017 | Perkinelmer Life Sciences, Inc. | Potential for electric shock in the case of a failure to install the grommets, supplied with the system. | Device Design | Perkin Elmer issued URGENT: Field Correction letter on 5/15/17 and accompanied with the GE HealthCare Important Product Information letter The notifications will be delivered to the end customer by email facilitated through PerkinElmer Technical Support group located in Akron Ohio.Technical Support will ensure the customers receive the both letters and understand the recommended actions. Accounts requested to complete the enclosed Response Form.Questions regarding the actions to be taken, as outlined in the GE HealthCare letter, should be directed via email to orderFI@ge.com or via phone (in the US) 011 44 800 515313, (in the UK) 0800 515 313, and (rest of world) +44-800 515 313. |  | 241 | 241 instruments | US Nationwide Distribution | Open, Classified |  |  | 241 | Computer
Z-2703-2017 | 77244 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | Carestation 620 A1, Carestation 650 A1, and Carestation 650c A1 |  |  | May 02, 2017 | 2017 | GE Healthcare, LLC | Carestation 620, 650 and 650c has Potential for Elevated FiCO2 and Unexpected System Malfunction | Device Design | Consignees were sent on 5/2/2017 a GE Healthcare "Urgent Medical Device Correction" letter GEHC Ref# 34082 dated May 2, 2017. The letter was addressed to Chief of Anesthesia, Director of Biomedical / Clinical Engineering&amp; Health Care Administrator / Risk Manager. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction &amp; Contact Information. GE Healthcare will correct affected products. GE Healthcare representative will contact the firms to arrange for the correction. For questions contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 3448 | 3,448 (334 US; 3,114 OUS) | US Distribution including PR and to the states of : AL, AR, AZ, CA, CO, CT, FL GA, ID,IL, IN, KY, LA, , MA, MD, MI, MO, MS, MT, MN, NC, NH, NJ, NV, NY, OH, OK, PA, TN, TX, VA, WA, WI, and UT. OUS: AlgeriaArgentinaAustraliaAustriaBelgiumBoliviaBosnia and HerzegovinaBrazilBulgariaCambodiaCanadaChileColombiaCosta RicaCote d'IvoireCroatiaCzech RepublicEcuadorEgyptEl SalvadorEstoniaFinlandFranceGermanyGreeceHondurasHong KongHungaryIndiaIndonesiaIraqIrelandIsraelItalyJapanJordanKenyaKorea (Republic Of)KuwaitLatviaLebanonLithuaniaMalaysiaMaldivesMaltaMauritiusMexicoMoldovaMoroccoMyanmarNetherlandsNew ZealandNigeriaNorwayOmanPakistanParaguayPeruPhilippinesPolandPortugalRomaniaRussiaSaudi ArabiaSenegalSerbiaSingaporeSlovakiaSloveniaSouth AfricaSpainSwedenSwitzerlandTaiwanTanzaniaThailandTurkeyUnited Arab EmiratesUnited KingdomUruguayVietnamZimbabwe | Open, Classified |  |  | 3448 | Computer
Z-3045-2017 | 77608 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Neusoft NeuViz 128 Multi-slice |  |  | May 19, 2017 | 2017 | Neusoft Medical Systems Co., Ltd. | Software defect | Software design | Neusoft Medical Systems will bring defect into compliance: The defect will be remedied by updating software. Field Change Order I related to software updated will be released to the affected systems in ' Aug, 2017. The affected systems will be upgraded with updated software in the field free of char e in 6 months after the FCO release. |  | 154 | 154 | US Nationwide distribution | Open, Classified |  |  | 388 | Computer
Z-3038-2017 | 77803 | Yes | LWS |  |  | Cardiovascular | PMA | EMBLEM MRI S-ICD Model 219, Subcutaneous-Implantable Cardioverter Defibrillator, Rx. |  |  | June 29, 2017 | 2017 | Boston Scientific Corporation | The device can deliver an atypical amount of energy due to memory corruption inside the device. | Software design | The firm issued notifications dated June, 2017, beginning 6/29/2017 to physicians (implanting and patient follow-up physicians) via hand delivery by their sales representatives. The firm estimated approximately 30% of the U.S. physicians would be receiving the notification via hand delivery. Hand delivery by affiliates in foreign countries to customers began approximately 6/30/2017. Overnight mail letters were issued to the U.S. and foreign physicians starting on/about 7/7/2017 who did not receive hand- delivered notifications. |  | 9200 | Approximately 9,200 devices | Worldwide - US Nationwide distribution, including Puerto Rico, U.S. Virgin Island, and Guam, There was also worldwide foreign distribution, including Canada. | Open, Classified |  |  | 38400 | Computer
Z-3036-2017 | 77384 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Samsung GU60A system- |  |  | May 24, 2017 | 2017 | NeuroLogica Corporation | When performing the Stitching function on GU60A system, unintended area is exposed to radiation | Device Design | Neurologica will bring defect into compliance : 1.All facilities were notified via a customer notification letter.2.Samsung will release improved software to avoid unintended exposure. When setting the stitching area, if there is an over exposed area, the height of the swivel arm will be lowered over the exposed area and the first region (first image) will be collimated by the bottom blade.3.A software patch was developed: SW3.01 was released on 05/08/2017 and SW1.00 was released on 07/28/17.4.The software patch will be installed on the affected devices.5.Samsung has provided the statement and assurance that all corrections will be performed free of charge.6.All units will be provided with updated software by November 11, 2017.For further questions, please call (978) 564-8500. |  | 87 | 87 units | Nationwide Distribution | Open, Classified |  |  | 87 | Computer
Z-3048-2017 | 77655 | [6 empty] | VASER PRO Amplifier, VASER 2.2 VASERlipo System stem, Rx Only, 115/230V - 60/50Hz 300VA |  |  | June 01, 2017 | 2017 | Solta Medical Inc | Potential inaccurate calibration by service depot repair, which could contribute to patients being burnt by the device. This only impacts units serviced between the dates of 09/08/2016 and 05/23/2017. | Equipment maintenance | Consignees were telephoned on June 1, 2017, as initial recall notification. At that time, all consignees agreed to discontinue use of the affected product until it was returned or repaired. An Urgent Medical Device Recall letter dated June 14, 2017, requested that consignees that have not been contacted regarding this recall and have a device repaired between the dates of 09/08/2016 and 05/23/2017 contact Solta Medical immediately at 877-782-2286. Any questions regarding the recall can be directed to Solta Medical Product Support Team at 877-782-2286 or 510-259-5299, option 2. |  | 7 | 7 units | US Distribution to the states of : AL, CA, CO, FL, NY, PA, &amp; TX. | Open, Classified |  |  | 11 | Computer
Z-1037-2017 | 76086 | No | MIH |  |  | Cardiovascular | PMA | AFX Endovascular AAA System, Endoleak Type IIIA |  |  | December 27, 2016 | 2016 | Endologix | Endologix updated the Instructions for Use (IFU) and implemented modifications to the AFX Endovascular AAA System (AFX System). | Device Design | Endologix sent a dear physician letter dated June 2, 2015, to inform them of the enclosed 2014 Clinical Update on the Endologix Endovascular Systems for AAA Repair. The update reports the cumulative clinical data and information on the Powerlink System and the AFX Endovascular AAA System (AFX System). The letter informs the Physicians of the year's update on the clinical programs and commercial experience (i.e. post-market surveillance).A dear physician letter dated January 2016 was seen to inform them of the enclosed 2015 Clinical Update covering the Endologix Endovascular Systems for AAA Repair. The update reports the cumulative clinical data and commercial experience (i.e. post-market surveillance) of the Powerlink System and the AFX Endovascular AAA system and contains important information and recommendations regarding the continued safe and effective use of the devices.A dear physician letter dated December 27, 2016, will be sent to provide important information related to the AFX Endovascular AAA System (AFX System). Endologix is informing the customers of a series of updates to the Instructions for Use (IFU) and modifications to the product. Customers with any questions are instructed to contact their local Endologix representative. |  |  |  | Worldwide Distribution - US (nationwide) and Internationally to Austria, Belgium, Czech Republic, France, Germany, Greece, Ireland, Israel, Italy, Latvia, Luxembourg, Monaco, Netherlands, Poland, Portugal, Spain, Sweden, Switzerland, Turkey, and United Kingdom. | Terminated | April 28, 2017 | 122 |  | Computer
Z-0940-2017 | 76014 | No | OGZ | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | TVC Imaging System model TVC-MC9 , for the near-infrared examination of coronary arteries in patients undergoing invasive coronary angiography |  |  | December 15, 2016 | 2016 | Infraredx, Inc. | Near-infrared spectroscopic image (Chemogram) generated by the TVC Imaging System TVC-MC9 in some cases demonstrated an unusual appearance in that the yellow in the image (indicating the presence of lipid core plaques) was horizontally-oriented | Device Design | Infraredx, Inc. sent an Urgent Product Recall Medical Device Field Correction letter dated December 15, 2016, describing the problem and replacing the defective lasers. A Field Service Representative will contact customers to schedule a replacement of the system's laser rack module. If customers receive this letter and are not contacted within 48 hours, please call Infraredx Technical Support@ 1-800-596-3104.For questions regarding this recall call 781-221-0053. |  | 7 | 7 units | Worldwide Distribution - US including MI and Internationally to South Korea | Open, Classified |  |  | 7 | Computer
Z-2671-2017 | 77077 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris PC Unit, Model 8015 |  |  | June 12, 2017 | 2017 | CareFusion 303, Inc. | BD initiated the recall of Alaris PC unit model 8015 after the firm identified five scenarios which can result in the occurrence of Systems Error Code 255-16-275 and can potentially result in interrupted infusions. | Software design | BD sent an Medical Device Safety Notification Letter dated June 12, 2017 , to customers to inform them that BD is issuing this Safety Notification letter to inform you of reports of customers experiencing a System Error (error code: 255-16-275) with Alaris System PC unit model 8015 that can result in interruption of infusions. Customers are informed of the issue, potential risk, and required actions to be taken by users and by BD. Customers with questions with the safety notification are instructed to call (888) 562-6018 or email supportcenter@carefusion.com. Customers with adverse event reports are instructed to call (888) 812-3266 or email customerfeedback@carefusion.com. Customers with technical questions are instructed to call (888) 812-3229 or email dl-us-inf-techsupport@carefusion.com. For further questions, please call (858) 617-5925. |  | 568283 | 568,283 units | Worldwide Distribution - USA (nationwide) and to the countries of : Europe, Australia, New Zealand,South Africa, Greater Asia, Middle East, and Canada. | Open, Classified |  |  | 568283 | Computer
Z-2701-2017 | 76427 | No | MEC | Medical washer-disinfector. | General Hospital | General Hospital | 510(K) Exempt | The Olympic Pasteurmatic 3000 and Olympic Pasteurmatic 3500 are intended for cleaning and high-level disinfection of respiratory therapy and anesthesia equipment by killing vegetative bacteria and viruses, in a hot water bath at ~160Âż - ~170Âż F for 30 minutes. |  |  | May 22, 2017 | 2017 | Natus Medical Incorporated | Olympic Pasteurmatic 3000/3500 Systems are recalled because they do not fully meet todays regulatory requirements. Given the age of the products design, component and technology obsolescence, and the time involved in fully complying with the latest design and development standards, the firm decided to recall of the Olympic Pasteurmatic 3000/3500 System devices by means of disabling the devices. | Device Design | Natus Medical sent an URGENT: RECALL NOTIFICATION MEDICAL DEVICE RECALL letter dated May 2017.Please refer to the accompanying Instructions to Disable the Pasteurmatic (DOC-022838), then complete and return the attached reply form. Should you require assistance in decommissioning your affected system please indicate option 2 on the reply form and you will be contacted to schedule an in- service. Natus may reimburse facilities for Olympic Pasteurmatic" 3000/3500 System purchases based on the age of the affected system as determined by the system serial number.There is no evidence to suggest there is a safety risk associated with the Olympic Pasteurmatic" 3000/3500 System device, nor have there been any safety incidences reported to Natus.Natus apologizes for any inconvenience. Please understand that this is part of our commitment to providing only the highest quality products and information to our customers. Please contact us at Natus_Quality_Programs@natus.com or (866) 940-7143 option #2 if you have any further questions. |  | 275 | total 275 units (97 units in US and 178 units international) | Worldwide Distribution - USA (nationwide) Distribution to the states of : AL, AR, CA, CO, FL, GA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, ND, NH, NJ, NM, NY, OH, OK, OR, PA, TN, TX, UT, VA, WA, WI, and WV and in the following countries: Canada, Colombia, Costa Rica, Ecuador, Hong Kong, India, Indonesia, Japan, Korea, Malaysia, Mexico, Panama, Peru, Qatar, Saudi Arabia, Taiwan, Thailand, and Viet Nam. | Open, Classified |  |  | 275 | Computer
Z-3056-2017 | 77739 | No | PGW | Stereotaxic instrument. | Neurology | Ear Nose &amp; Throat | 510(k) | Medtronic Instrument Kit 9733908 Fusion ENT, reusable, RX. This kit includes Kit 9733452 (containing 9733449, Straight Suction, 9733450, 70 deg Curved Suction, and 9733451, 90 deg Curved Suction). The firm name on the label is Medtronic Navigation, Inc., Louisville, CO. |  |  | June 29, 2017 | 2017 | Medtronic Navigation, Inc. | The instruments are not able to pass instrument verification which is performed prior to the use of each instrument. | Device Design | Medtronic sent an Urgent Medical Device Recall letter dated June 25, 2017, via overnight mail on June 29, 2017, to all affected customers. Letters to foreign customers are being handled by geography-specific Regulatory Specialists. Customers were instructed examine their inventory for the affected product and immediately quarantine them for return for Medtronic. Customers with affected product were instructed to contact Medtronic Technical Services at 1-800-595-9709 to receive a return material authorization (RMA) and arrange for their return and no charge replacements. Customers were asked to complete and follow instructions on the attached Consignee Response Form. Sign and date the bottom of the form and then return the form to Medtronic at RS NavFCA@medtronic.com or fax it to Medtronic Technical Services at 651-367-7075. Customers with questions were instructed to call 1-800-595-9709. |  | 15 | 15 kits | Worldwide Distribution - US (nationwide) and Internationally to Canada. | Open, Classified |  |  | 603 | Computer
Z-1034-2017 | 76039 | No | EID | Dental hand instrument. | Dental | Dental | 510(K) Exempt | NanoTips 30g 20pk, 100pk PN 1240, 1426 |  |  | December 14, 2016 | 2016 | Ultradent Products, Inc. | Ultradent Products, Inc. announces a voluntary field action for the NanoTips 30g 20pk, 100pk because the cannula attached to the tip hub may become loose or detach from the blue base that connects to the syringe. | Device Design | Ultradent Products, Inc. sent an Urgent Medical Device Correction letter dated December 14, 2016, to all affected consignees. The letter instructed consignees to discard remaining inventory and return the response form. For any questions, please call at 801-553-4858. |  | 14300 | 143 packs of 100 tips; 104 packs of 20 tips | Worldwide Distribution - US (Nationwide) - Internationally to Belgium, Switzerland, Germany, Italy, South Korea, Netherlands, and Japan. | Terminated | June 05, 2017 | 173 | 14300 | Computer
Z-0941-2017 | 75882 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | G3 Opticage Expandable Interbody Fusion Device, Model No. 9180-00, 9180-01, 9180-06, 9180-07, 9180-17, 9101-00, 9101-06 |  |  | November 22, 2016 | 2016 | Interventional Spine Inc | During implantation, the proximal end of the device could separate if the angle between the device and insertion instrument is too steep. This separation will prevent the device from expanding. | Device Design | Interventional Spince, Inc. sent an An Urgent Medical Device Recall letter dated November 22, 2016, to all affected customers to inform them that during implantation, the proximal end of the device could separate if the angle between the device and insertion instrument is too steep. This separation will prevent the device from expanding. The letter informs the customers of the risk to health and actions to be taken by the customer/user. Customers were instructed to discontinue use and return device per the enclosed instructions. Complete the enclosed acknowledgment letter and fax, email or mail it to Interventional Spine. Customers with questions were instructed to call Monday through Friday 8am to 5pm (Pacific Time), 949-472-0006 or email cs@i-spineinc.com. |  | 239 | 239 units | Nationwide Distribution to NY, NC, SC, PA, TX, OH, GA, NJ, DE, AZ, DC, CA, and VA. | Completed |  |  | 239 | Computer
Z-0944-2017 | 75873 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Modular REDAPT(TM) Hip Systems |  |  | November 15, 2016 | 2016 | Smith &amp; Nephew, Inc. | All lots of modular neck hip prostheses are being recalled due to a higher than anticipated complaint and adverse event trend. | Device Design | Smith &amp; Nephew sent an Urgent Product Recall Notification letter dated November 15, 2016, to all affected customers. Customers were instructed to inspect their inventory and locate any unused devices and quarantine immediately. Complete the Inventory Return Certification Form and obtain a return authorization (RA) number if there was product to return.For questions regarding this recall call 978-749-1440. |  | 6266 | 6266 units | Worldwide Distribution - US (nationwide) and Internationally to Austria, Australia, Belgium, Canada, Switzerland, China, Cyprus, Czech Republic, Dubai, Germany, Denmark, Spain, France, United Kingdom, Greece, Hong Kong, India, Italy, Mexico, Netherlands, Poland, Puerto Rico, Portugal, Russia, Turkey, and South Africa. | Open, Classified |  |  | 31936 | Computer
Z-2700-2017 | 77279 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris Pump Module model 8100 |  |  | June 12, 2017 | 2017 | CareFusion 303, Inc. | There is a potential risk that could cause unintended flow in the older, centered sear door latch design in the Alaris Pump module model 8100. | Device Design | An Urgent Medical Device Recall letter will be sent on 6/12/17 to customers to inform them that BD is issuing this letter to inform you of a potential risk associated with the Alaris Pump module. Firm identified a specific scenario that could cause unintended flow in the older, centered sear door latch design in the Alaris Pump module model 8100. This scenario is reproducible under the following situations: a) the user does not close the roller clamp on the IV administration set before the pump door is opened as recommended, and b) opening the pump door using an atypical technique with the door latch with the centered sear design. The letter instructs customers If their device has a centered-sear that was manufactured between June 2002 to June 2004, clearly mark and sequester (e.g. Biomed department) the Alaris Pump module that exhibited the issue. Notify CareFusion Customer Advocacy at 888-812-3266 or customerfeedback@carefusion.com to coordinate an onsite remediation or send the devices to the BD Service Depot. |  | 35940 | 35,940 units | US and Canada | Open, Classified |  |  | 35940 | Computer
Z-0978-2017 | 75049 | No | MSS | Intraocular lens guide. | Ophthalmic | Ophthalmic | 510(k) | Alcon Monarch III (1) "C" Cartridge |  |  | September 05, 2016 | 2016 | Alcon Research, Ltd. | The intraocular lens could become lodged within the cartridge | Device Design | Alcon sent an Urgent Field Safety Action letter dated August 2016 to all affected consignees who received the affected product. Customers were instructed to check their inventory and segregate any of the affected cartridges. Customers were also asked to complete and return the Product Recall Response Form. , requesting the return of any units in their inventory. If customers had transferred product to other organizations, they were asked to provide them with a copy of the notice. Customers with questions were instructed to contact Alcon.For questions regarding this recall call 817-551-3058. |  | 1130 | 1,130 units | Distributed to Mexico only | Open, Classified |  |  | 1130 | Computer
Z-3035-2017 | 77695 | No | DOF | Lead test system. | Toxicology | Toxicology | 510(k) | LeadCare II Blood Lead Test Kit, containing LeadCare II Lead Controls |  |  | January 20, 2016 | 2016 | Magellan Diagnostics, Inc. | LeadCare II Test Kits controls out of range | Device Design | Magellan notified customers initially by Email, fax and/or verbal communication with Magellan Product Support representatives beginning on January 20, 2016 to advise of the problem and provide revised Assignment Control Value Matrix.. |  | 12931 | 12,931 kits | Worldwide - US Nationwide distribution and the countries of:COLOMBIA, CANADA, THAILAND, INDIA, UAE, AUSTRALIA, GERMANY, CANADA, SOUTH AFRICA, SPAIN, NEW ZEALAND, KOREA | Completed |  |  | 12931 | Computer
Z-0962-2017 | 75943 | Yes | NIK |  |  | Cardiovascular | PMA | Amplia MRI" CRT -D SureScan", |  |  | December 09, 2016 | 2016 | Medtronic Inc., Cardiac Rhythm and Heart Failure | During internal testing conducted as part of next generation product development, it was discovered that the AdaptivCRT (aCRT) feature of Amplia MRI and Claria MRI CRT -D devices may appear to be enabled when it is not active. | Software Design Change | Consignees were hand delivered beginning on 12/9/2016 a Medtronic "Urgent Medical Device Correction" letter dated December 2016. The letter was addressed to "Dear Physician". The letter described the issue and the product involved in the recall. The letter informed consignees that a software update is being developed to address the issue and it will also address other issue not covered in this recall. The letter also included Appendix A which listed Patient Management Recommendations (Until the software update has been approved and the device models listed in Table A receive the update, follow the programming recommendations found in Appendix A. These recommendations also apply to any new device implants).Additionally, two other "Urgent Medical Device Correction" letters were hand delivered to the consignees, "Principal Investigator Contact Record" and a supplement to the advisory notification.For questions contact your local Medtronic Representative or Medtronic Technical Services at 800-723-4636. |  | 21111 | 21,111 ( 13,198 US; 7,913 OUS). | US: All states in USA including DC &amp; PR. OUS: Australia, Finland, Japan, Saudi Arabia, Austria, France, Korea,Singapore,Bangladesh, Germany,Luxembourg, South Africa, Belgium,Greece, Malaysia, Spain, Brunei Darussalam, Guadeloupe, Netherlands, Sweden, Cayman Islands, Hong Kong,New Zealand, Switzerland, Chile, India, Norway, Thailand, Cyprus, Iran, Poland, Turkey, Denmark, Ireland, Portugal, United Arab Emirates,Dominican Republic, Italy, Reunion,United Kingdom. | Open, Classified |  |  | 21111 | Computer
Z-0991-2017 | 75969 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | ConFORMIS ITOTAL PS-IPOLYIMPLANT KIT-RIGHT |  |  | October 31, 2016 | 2016 | ConforMIS, Inc. | iView report displayed femoral cut measurements in the incorrect unit of measure (inches vs. mm). | Software Design Change | ConForMIS intiated the recall by telephone on October 31, 2016, to the Sales Rep to request return of the kit.For questions regarding this recall call 781-345-9001. |  | 1 | 1 unit | Distribution to TX only | Terminated | January 11, 2017 | 72 | 1 | Computer
Z-2772-2017 | 77717 | Yes | DWC | Cardiopulmonary bypass temperature controller. | Cardiovascular | Cardiovascular | 510(k) | STOCKERT Heater-Cooler System 3T, 240V/60 Hz |  |  | June 30, 2017 | 2017 | Sorin Group USA, Inc. | Plastic heat exchanger fibers separating the blood compartment from the water compartment in certain oxygenators may allow diffusion of hydrogen peroxide in a quantity that exceeds allowable limits. | Device Design | "Consignees were sent a recall notification, dated 06/30/2017, via mail. Instructions included a recommendation from LivaNova to use its oxygenators, or other oxygenators that are equipped with compatible heat exchangers, to ensure that the permeability rate is within the prescribed allowable limits (ALs). Depending on whether or not the oxygenator was from LivaNova, there were specific instructions to follow for each type and source. Consignees were also instructed to complete and return the Confirmation Form.If product was further distributed, the firm asks that further customers be notified of the recall and that customers contact the LivaNova Quality Assurance Team at usfsn@livanova.com. Questions regarding the recall can be directed to 800-986-4702 or 3T.US@livanova.com. Any adverse reactions are to be reported to FDA's MedWatch Adverse Event Reporting program." |  | 7546 | 7,546 units | "Domestic: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, PR, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, &amp; WY. Foreign: AL, AE, AR, AT, AU, AZ, BD, BE, BG, BH, BR, BY, CA, CH, CL, CN, CO, CR, CY, CZ, DE, DK, DZ, EC, EE, EG, ES, ET, FI, FR, GB, GE, GR, GY, HK, HR, HU, IE, ID, IL, IN, IQ, IR, IT, JO, JP, KR, KW, KZ, LB, LK, LT, LU, LV, LY, MA, MC, MK, MN, MT, MU, MX, MY, NG, NL, NO, NP, NZ, OM, PA, PE, PH, PK, PL, PR, PS, PT, QA, RE, RO, RS, RU, SA, SE, SG, SK, SV, SY, TH, TN, TR, TT, TW, UA, VN, &amp; ZA." | Open, Classified |  |  | 7546 | Computer
Z-2696-2017 | 77509 | No | MHY |  |  | Neurology | PMA | Infinity 7, Implantable Pulse Generator, REF 6663, Rx only, STERILE EO |  |  | June 02, 2017 | 2017 | St. Jude Medical, Inc. | The firm discovered a risk of loss of Neuromodulation therapy due to exposure of the IPG to monopolar electrosurgical devices (ESU) during medical procedures. | Device Design | The notification material is in the form of a Physician Letter and Patient letter which will be delivered to physicians and patients in the US. The physician mailing will be initiated June 2, 2017 for delivery on June 7-9, 2017. The physicians were provided with patient management recommendations, Step for the patient controller updates, and Clinician's Manual Revised Warning. The Patient mailings will begin the week of June 12, 2017 to allow for clinician notification and are expected to be delivered within 3-5 business days. The patient letter informs of them of the potential for the loss of therapy and the steps to receive the software upgrade. |  | 40 | 40 units | US Nationwide Distribution | Open, Classified |  |  | 19219 | Computer
Z-0990-2017 | 76018 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Eclipse Treatment Planning System version 13.MR2 [13.06.31 with Smart Segmentation Knowledge Based Contouring version 2.3 [2.3.12] |  |  | December 02, 2016 | 2016 | Varian Medical Systems, Inc. | Modifications in version 13.6MR2 for Contouring, SmartAdapt, and SmartSegmentation workspaces resulted in contours not being saved consistently in Eclipse. Treatment Planning System. The issue only occurs if certain conditions are fulfilled. | Device Design | Varian Medical Systems sent an Urgent Medical Device Correction Letter dated December 5, 2016 , that were sent to all 5 affected sites on December 5, 2016 by a trackable method.Letters advised users of the issue and the specific scenario that must take place for the issue to occur. Varian is investigating the issue and will provide a correction once it is available.We sincerely apologize for any inconvenience and thank you in advance for your cooperation. If you require further clarification, please feel free to contact your local Varian Medical Systems Customer Support District or Regional Manager. This notice has been provided to the appropriate Regulatory Authorities. For further questions please call (650) 424-6833.Varian Oncology Help Desk Contact Information:Phone:+1.888.827.4265Email:support @varian.com |  | 42 | 42 devices are affected | US Distribution to the states of : NJ, TN. OR and FL. | Terminated | May 12, 2017 | 161 | 42 | Computer
Z-2744-2017 | 77137 | No | GCW | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | 9400 |  |  | April 13, 2017 | 2017 | Pentax of America Inc | Pentax Medical did not always provide transformers with 9175 isolation transformers are used with 7245C, 7245C/E, 7245D, 9200Cs, 9200Ds, 9310HDs, and 9400s computer systems. | Device Design | Pentax notified their customers on 4/13/2017 via USPS. |  | 241 | 241 units in total | Worldwide Distribution - US Nationwide | Open, Classified |  |  | 241 | Computer
Z-0995-2017 | 76127 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Biograph mCT S(40)-4R, Material Number 10248671 |  |  | December 13, 2016 | 2016 | Siemens Medical Solutions USA Inc. | Siemens Medical Solutions, Molecular Imaging has become aware, through our supplier that some laser cover windows may fall off. | Device Design | Siemens Healthineers sent a Customer Safety Advisory Notice letter on December 13, 2016. The letter states that the light marker windows used for the positioning laser and integrated into the front cover of the Biograph mCT or Biograph mCT Flow system may loosen and potentially drop out and warning the consignee not to use the system of the front cover is loose or missing. Customers with questions were instructed to contact their local Siemens representative at the following numbers:America 1-800-888-7436Europe, Middle East and Africa +49 9131 940 4000Asia and Australia +86 (21) 3811 2121For questions regarding this recall call 865-218-2000. |  | 88 | 88 units | Worldwide Distribution - US (nationwide) and Internationally to U.A.E., AT, AU, AZ, BR, CA, ,CH, CL, CN, CO, CZ, DE, DK, EG,FI, FR, GB, HK, HR, IL, IN, IR, IT, JP, KR, MX, NL, NO, NZ, PK, PL, QA, RO RU, SE, SG, SI, SK, TH, TR, AND TW. | Terminated | May 08, 2017 | 146 | 409 | Computer
Z-2754-2017 | 77627 | No | NFJ | Picture archiving and communications system. | Radiology | Ophthalmic | 510(k) | Merge Eye Care PACS Viewer 5.2 |  |  | June 13, 2017 | 2017 | Merge Healthcare, Inc. | When the user has not set up any user preference on the sorting order to render the study images (OS/OD/etc.) and reports, the ECP may get an incorrect study index. If the user attempts to delete an image, ECP may then read that image as OS when it is, in fact, OD. | Software design | The firm initiated the recall by certified letter the week of 06/26/2017. The notice provided the work-around and requested a response. |  | 18 | 18 units | FL, IL, CO, MI, CA, NY, ND, OH, CT | Open, Classified |  |  | 18 | Computer
Z-2614-2017 | 77581 | No | RCR |  |  |  |  | Standex International Bakers Pride Microwave Ovens Model E300 |  |  | January 01, 2017 | 2017 | Standex Intl Corp | It was discovered that all five microwave ovens fail to have concealed or inaccessible safety interlocks. | Device Design | The Standex Cooking Solutions Group plans to replace all 149 Bakers Pride Model E300 Speed Ovens remaining in the field as follows :1.Product design changes include (a) an updated interlock switch mounting that reduces the effects of shipping and handling on switch positioning; (b) a revised mechanism for the secondary interlock switch to prevent operation by any part of the human body or any object with a straight insertable length of 3.9 inches. A 200,000-cycle leakage measurement is conducted and will be finished by July 31, 2017. 2.The updated replacement oven will be shipped with the improved, stronger bracing materials and added cover over the handle with markings and changes in the door mechanism.3.New compliant ovens will be delivered by an installer to distributors and owners.4.All shipping and installation costs for delivery of the new units and return of the old units will be paid by Standex International Corporation. 5.The updated and corrected pages of the Service Manual and Owners Manual will be included with the delivery of the replacement oven, and the corrected warning label and Certification label will be permanently affixed to the replacement ovens.6.Standex International Corporation expects to deliver the replacement units after testing, which will be completed by July 31, 2017. Distributors and customers will then be given a period of 4 weeks to return the old unit. 7.All returned ovens will be updated to revised construction and subjected to the full Unit Test Procedure required for new products. Ovens that cannot be updated will be destroyed. 8.Progress reports will be provided to FDA on the 15th of every month until all non-compliant ovens have been replaced with new compliant ovens, returned to the factory and modified to the revised and compliant construction, as verified by conducting the Unit Test Procedure.CDRH approves the CAP subject to the following conditions:1.Notification of all dealers and |  | 149 | 149 | US Distribution | Open, Classified |  |  | 149 | Computer
Z-2863-2017 | 77474 | No | MCW | Intraluminal artery stripper. | Cardiovascular | Cardiovascular | 510(k) | Spectranetics Turbo-Elite Laser Atherectomy Catheter Peripheral Rapid Exchange Catheters |  |  | June 06, 2017 | 2017 | Spectranetics Corporation | Potential for flouroscopially visible marker band at the tip of the Rapid Exchange Turbo Elite Atherectomy Catheters to become loose or disengaged rom the catheter when the device is in use in an off-label manner. | Device Design | Spectranetics sent a Field Safety Notice letter via certified mail on June 7, 2017, to all customers who have ordered the affected model numbers since 23May2016 reiterating the risks associated with off-label use of the device (i.e., not activating the laser within a stent and not until all contrast media is flushed from the treatment area). The product continues to meet specifications and does not need to be returned.Customers with questions were encouraged to call their local Spectranetics Sales Representative or Spectranetics Customer Service Department. Customer Service can be reached by phone at 800-231-0978 (Option 2), by fax at 877-447-2022, or by email at customer.service@spnc.com. Customer Service is available Monday thru Friday, From 7AM-5PM MST. |  |  |  | Worldwide Distribution - US including AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, IL, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, NJ, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, WV, Hawaii and Puerto RicoForeign distribution to Bahrain, Brazil, Canada, Italy, Japan, and Netherlands | Open, Classified |  |  |  | Computer
Z-2710-2017 | 77599 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | APLIO 300 TUS-A300 |  |  | May 01, 2017 | 2017 | Toshiba American Medical Systems Inc | When Continuous Trace method is used as the tracing method on Spectral Doppler in velocity trace measurement, an incorrect value may be displayed in the ultrasound systems and in the ultrasound workstations listed below. | Software Manufacturing/Software Deployment | Toshiba Medical sent an Urgent Medical Device Correction letter on May 1, 2017. The consignee is instructed not to use the Continuous Trace method for cardiac velocity trace measurements to quantify the MPG values.Modified software is being provided. A Toshiba service representative will contact the consignee to make an appointment to install the modified software to fix the problem.Customers were asked to complete the form and fax it to the toll-free number shown at the top of the form. The form can also be sent by email to RAffairs@tams.com.Customers with questions were instructed to call 800-521-1968. |  | 90 | 90 units | Nationwide Distribution | Open, Classified |  |  | 433 | Computer
Z-2707-2017 | 77086 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | The Schiller PB 1000, also known as Merge Hemo Patient Data Module (PDM), a major component of the Merge Hemo software. |  |  | March 10, 2016 | 2016 | Merge Healthcare, Inc. | If the PDM is utilized to output a pressure for computing fractional flow reserve (FFR), the analog output may not reflect the true representation of the selected pressure waveform being provided by the PDM to the Link Assembly which provides the analog output. | Software design | The recalling firm issued recall letters dated 3/10/2016, 5/3/2016, and 8/5/2016 via email or via certified mail if the email address is not available. |  | 555 | 555 PDM's | Distribution US nationwide. | Open, Classified |  |  | 555 | Computer
Z-2713-2017 | 77308 | No | FDT | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Video Duodenoscope |  |  | January 17, 2017 | 2017 | Pentax of America Inc | Pentax initiated a field correction/safety alert for two (2) models of the Video Duodenoscope to determine how soiling may have occurred on the surface of the suction cylinder and under the distal cap during testing. | Device Design | The firm, Pentax Medical, sent a "FIELD CORRECTION' letter dated January 17, 2017 to their affected customers via USPS Certified Mail. The letter described the product, problem and actions to be taken. The customers were instructed to immediately remove any affected product from use; follow product labeling; ensure all reprocessing personnel are knowledgeable and thoroughly trained on the instructions of Use for manual reprocessing of the devices; clean elevator recesses and follow all reprocessing instructions; and complete and return the FIELD CORRECTION RESPONSE FORM via a member of the PENTAX Service Department as part of an on-site-visit or Fax to: QA/RA Department at 201-799-4063 (alternate 201-391-4189) or a pdf copy to customeradvisiories@pentaxmedical.com. The firm will replace any affected product with its current generation model.If you have questions or request further information or assistance contact PENTAX Medical Customer Support at 1-800-431-5880 (8:30AM - 5:00PM, Monday-Friday, EST) or email: customeradvisories@pentaxmedical.com. |  | 2015 | 2,015 (US - 519 (ED-3490TK) and 5 (ED-3270K); OUS - 1,491) | Worldwide Distribution-US (Nationwide) including states of: AL, AZ, CA, CO, CT, DC, DE, FL, GA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MO, MS, MT, NC, NH, NJ, NV, NY, OH, OK, PA, PR, TN, TX, UT, VA, WA, WI, WV, WY;. and country of: CANADA. | Open, Classified |  |  | 2015 | Computer
Z-2734-2017 | 77505 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Draeger Medical Systems Infinity Acute Care System (M540) |  |  | June 06, 2017 | 2017 | Draegar Medical Systems, Inc. | Cockpits with revision index 06 or higher that contain 4GB RAM modules may not annunciate audio or visual alarms on the Cockpit and Central Station. | Software design | Draeger issued on June 6, 2017 Urgent Medical Recall letter with Addendum US serial numbers via Certified Mail addressed to Risk Manager and Director Clinical Engineering The letter instructs the consignees, if the serial number of the cockpit is on the serial number list, to remove the Cockpit from use until it can be determined whether the 4GB RAM Module is among those affected. The IACS can be used without the Cockpit and the M540 will function as the primary monitoring device once the Cockpit is disconnected. A DrÂżgerService Representative will schedule a time to inspect the Cockpit and, if necessary, replace the 4GB RAM Module. The inspection and 4GB RAM replacement (if necessary) will be performed free of charge. For questions regarding the Operation and/or servicing the DrÂżger InfinityÂż Acute Care System Cockpit contact DrÂżgerService Technical Support at 800-437- 2437 (press 4 at the prompt and then press 4 again).For OUS the responsible local Draeger Sales and Service will mail the letter with Addendum serial number list. |  | 719 | 37 US and 682 OUS | AZ, NH, SD, TN, TXForeign:AustraliaAustriaBelgiumCanadaChinaCosta RicaFranceGermanyGreat BritianHong KongIndiaIrelandItalyJapanMalaysiaNamibiaNetherlandsPolandPortugalRussian Fed.South AfricaSouth KoreaSpainSwedenSwitzerlandThailandTurkeyUnited Kingdom | Open, Classified |  |  | 719 | Computer
Z-2732-2017 | 77532 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Nova Biomedical Stat Profile Prime Analyzer as follows: |  |  | June 19, 2017 | 2017 | Nova Biomedical Corporation | Software- scheduled A-QC analysis could initiate priorto a patient result being released by the operator and result in the pending patient result being overwrittenwith the A-QC test results. | Software design | Nova Biomedical issued Urgent Field Correction Notice on June 19, 2017, to all affected customers. The letter advises users of the problem and Nova Biomedical or an authorized Nova Biomedical representative will schedule a no charge update of your Stat Profile Prime Analyzer software to version VI1.0.883.0. A representative will contact customers to schedule this update. Questions: contact Nova Biomedical Technical Support at 1-800-545-6682, in Canada contact us at 1-800-263-5999, or outside of N. America contact your authorized Nova Biomedical subsidiary or distributor.For questions regarding this recall call 781-647-3700. |  | 179 | 179 units (US) and 904 (OUS) Total 1083 | Worldwide Distribution - US (Nationwide)Foreign:AlbaniaAlgeriaArgentinaAustraliaAustriaBangladeshBeninBoliviaBrazilCanadaChileChinaColombiaCosta RicaCzech RepublicECUADOREgyptEl SalvadorFinlandFranceGermanyGhanaGreeceGuatemalaGuyanaHondurasHong KongHungaryICELANDIndiaIranIsraelItalyIvory CoastJapanJordanKenyaKoreaLatviaMacedoniaMalaysiaMexicoNepalNetherlandsNigeriaPakistanPakistanPalestinePeruPhilippinesPolandPuerto RicoRepublic of PanamaRepublica DominicanaRussiaSaudi ArabiaSlovak RepublicSloveniaSouth AfricaSri LankaSwitzerlandTaiwanThailandThe NetherlandsTrinidad,W.I.TunisiaTurkeyUKUruguayVietnam | Open, Classified |  |  | 179 | Computer
Z-2735-2017 | 77632 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Agfa Healthcare NX 3.0.8950 Imaging Processing Software |  |  | December 21, 2016 | 2016 | AGFA Healthcare Corp. | A customer reported that when using an NX workstation with software version NX 3.0.8950 software and selecting the affected patient/exam from closed exams, initially the wrong image was linked to the exam and appeared. After a short time the wrong image was replaced by the correct image, however the wrong image was used for transmitting to PACS. | Device Design | On December 22, 2016, an URGENT FIELD SAFETY NOTICE letter was sent either by mail or email. The letter included an acknowledgement form to be returned. The letter informed the customer of the problem and actions to be taken. |  | 221 | 221 units | AL, AZ, AR, CA, DC, FL, GA, IL, IN, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, NE, NJ, NY, NC, OH, OK, PA, SC, SD, TN, TX, UT, VA, WA, WV and Canada | Open, Classified |  |  | 221 | Computer
Z-2769-2017 | 77712 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Diagnostics Ultrasound System |  |  | April 24, 2017 | 2017 | Konica Minolta, Inc. | The Doppler velocity measurement on the Sonimage HS1 ultrasound system is not accurate. | Software design | Konica Minolta notified their customers on April 24, 2017. |  | 36 | 36 units | CA, CO, KY, FL, NJ, NY, IA TX, &amp; WA | Open, Classified |  |  | 36 | Computer
Z-2690-2017 | 77362 | No | PER | Pipetting and diluting system for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Coulter PrepPlus 2, Part Number 378600, |  |  | May 22, 2017 | 2017 | Beckman Coulter Inc. | When samples are prepared using the PrepPlus or PrepPlus 2 sample preparation systems and then run on a Flow cytometer, there is an imprecision in the absolute counts on only the first tube in a run. | Software design | The recalling firm initiated their recall by letter disseminated via email on 05/22/2017. The letter requested the following actions be taken:1. Insert an empty daughter tube in place of a patient sample as the first tube in the carousel that is programmed for the addition of Flow-Count Fluorospheres.2. Refer to the COULTER PrepPlus Operators Guide / PrepPlus 2 Instructions for Use (IFU) for instructions on adding Flow-Count Fluorospheres to daughter tubes.3. Review all results following the flow cytometer data review instructionsreferenced in the associated IFU.Beckman Coulter also requested the consignee share this information with their laboratory staff and retain this notification as part of their laboratory Quality System documentation. If the consignee forwarded any of the affected product(s) listed above to another laboratory, They were asked to provide them a copy of the letter. |  | 473 | 473 units | Worldwide Distribution | Open, Classified |  |  | 516 | Computer
Z-2779-2017 | 77508 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | Myelotec Video Guided Catheters, REF 2010 3.0 mm OD x 30 Cm Working Length, Two 1.3mm Working Channels |  |  | March 02, 2017 | 2017 | Myelotec, Inc. | Obstructed/blocked port from defective VGC Access Port Body component. | Device Design | Myelotec sent an Urgent Product Recall Notification letter dated April 10, 2017, to all affected customers. Customers are instructed to notify customers of any product further distributed, identify any recalled product on hand, and to complete the attached recall response, to be sent back to the distributor. For further questions, please call (770) 664-4656 Ext. 140. |  | 2980 | 2980 units | Worldwide Distribution - US Distribution to the states of : CA, FL, KY, MN, and NJ ., and to the countries of : Brazil, Iran, Kuwait, Malaysia, Singapore, South Korea, Taiwan, and Turkey | Open, Classified |  |  | 2980 | Computer
Z-2873-2017 | 77500 | No | MNI | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | PALLAS M/MAXIMIS 5.5 mm Instrument Tap. Part of MAXIMIS Pedicle Screw Spinal Fixation System. Valorem Surgical, Chicago, IL. Product labeled as Pallas M Spinal Fixation System through 2015. |  |  | June 21, 2017 | 2017 | Valorem Surgical LLC | Devices are not consistent with Quality System Requirements. | Device Design | Valorem Surgical sent an Urgent Medical Device Recall letter dated June 21, 2017, informing consignees of the product removal and instructing consignees to return any product to Valorem Surgical LLC to minimize potential risk to patients. Consignees are to fill out the attached Medical Device Recall Return Response form and send it back to the firm by email at global.safety@ValoremSurgical.com or by fax at 312-275-7117, or by mail. For questions pertaining to this recall, please submit an email inquiry to global.safety@ValoremSurgical.com |  | 2225 | 2,225 units total | Nationwide Distribution - US including CA and VA. | Open, Classified |  |  | 2225 | Computer
Z-1028-2017 | 76037 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius 2008K2 Series : Hemodialysis System |  |  | December 21, 2016 | 2016 | Fresenius Medical Care Renal Therapies Group, LLC | When the UF Rate, Goal or Time is adjusted using the up and down arrow keys, and the change is cancelled by using the esc key, the cancelled UF Rate is actually being executed rather than rate displayed on the machine | Device Design | Fresenius issued Customer Notification via Certified Mail on 12/21/2016. Users provided with problem and recommendation, when adjusting the UF Rate, UF Goal, or Tx Time, is to use the numeric data entry keys instead of the up or down arrow keys. Change the values by using data entry keys, then save by pressing CONFIRM. Fresenius will make arrangements to install software on the hemodialysis machine. Confirmation of the noticed by the consignee will be required via a Reply Form. |  | 60151 | 60,151 units (US) 1773 (OUS) | NationwideForeign: Canada, Mexico, Jamaica, Germany, Guam | Open, Classified |  |  | 205393 | Computer
Z-2746-2017 | 77354 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Install CD, Spine tools, Plus and S7, Version 25. Model Number 9731958. |  |  | May 10, 2017 | 2017 | Medtronic Navigation, Inc. | Software issue related to the StealthStation S7 system and the Synergy Spine application Version 2.1 configured with Spine Tool Install CD version 25. Issue may result in user being unable to navigate the Navigated Elevate Inserter with the StealthAiR Spine Frame during spine surgical procedures. | Software in the Use Environment | Medtronic sent an Urgent - Product Advisory Notice dated May 10, 2017, to all affected customers. Instructions indicated that a Medtronic representative will be in contact to update the system with version 26 and will remove any copies of the Spine Tools Install CD, version 25. Please quarantine any copies of version 25 until they can be removed.If you have further questions regarding this communication, please contact your local Medtronic Navigation Technical Services Representative, your local Medtronic Spine Sales Representative, or call (800 595-9709. |  | 108 | 108 | Worldwide Distribution - USA (nationwide) Distribution and to the countries of : Brazil, Germany, Switzerland and Great Britain. | Open, Classified |  |  | 108 | Computer
Z-1017-2017 | 75602 | No | HKI | Ophthalmic camera. | Ophthalmic | Ophthalmic | 510(k) | Merge Eye Station |  |  | December 09, 2016 | 2016 | Merge Healthcare, Inc. | This recall has been initiated due to an issue related to the potential accidental deletion of record(s) by an Eye Station user. | Software design | Consignees were sent on 12/9/2016 a Merge "Urgent Medical Device Recall" letter. The letter was addressed to Office Manager and Recall # 216-078. The letter described the Product, Issue, Potential Harm, Containment by the Customer / User, Action by Merge and Actions by Customer. Requested consignees to provide response to the "Response to this Notification is Required". If the product was further distributed advised consignees to notify them at once. For questions send an e-mail to recall@merge.com. E-mails to be sent to consignees on 12/19/2016. |  | 1597 | 1597 (1451 US; 146 OUS) | US: Alabama, Alaska, Arizona,Arkansas, California,Colorado,Connecticut,Delaware,District of Columbia, Georgia,Hawaii, Idaho, Illinois,Indiana, Iowa, Kansas, Kentucky, Louisiana, Maine, Maryland,Massachusetts, Michigan, Minnesota, Mississippi, Missouri, Montana, Nebraska, Nevada, New Hampshire, New Jersey, New Mexico, New York, North Carolina, North Dakota, Ohio, Oklahoma, Oregon, Pennsylvania, Puerto Rico, Rhode Island, South Carolina, South Dakota, Texas, Utah, Vermont, Virginia, Washington, West Virginia, Wisconsin, Florida, Tennessee, Wyoming.OUS:Czech RepublicAlgeriaBelgiumBrazilCanadaDenmarkFranceGermanyGreeceIndiaIsraelItalyKorea, Republic ofLuxembourgNetherlandsPolandRussian FederationSingaporeUnited Kingdom | Open, Classified |  |  | 1597 | Computer
Z-2699-2017 | 77604 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Arrow Glide Thru Peel-Away Sheath/Dilator Introducer |  |  | June 12, 2017 | 2017 | Arrow International Inc | Arrow is recalling additional lots that were identified as part of an active recall. Arrow is recalling the affected product due to the possibility that the catheter peel-away component hub tabs may prematurely detach when the practitioner begins to peel apart the sheath body from the catheter. | Device Design | Teleflex/Arrow International mailed an Urgent Medical Device Recall Notification Letter to affected customers on 06/12/2017 to inform them of the issue. Arrow requested that customers examine their inventory immediately for the affected lots and discontinue use and quarantine any products with the associated product codes identified in the notice and complete the Recall Acknowledgement Form and fax back to the number included in the notice. |  | 13542 | 9,037 units in the U.S. and 4,505 Internationally | Distributed to SC, AL, NJ, IN, MA, GA, CA, PA, AZ, VA , WA and Bangkok | Open, Classified |  |  | 13542 | Computer
Z-1020-2017 | 75889 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Reunion TSA Peg Alignment Sound |  |  | December 02, 2016 | 2016 | Stryker Howmedica Osteonics Corp. | Reunion TSA Peg Alignment Sound broke during surgery. | Device Design | Stryker Orthopaedics sent an Urgent Medical Device Recall Notification letter dated December 2, 2016, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to inform users of the Urgent Medical Device Recall and forward the notice to all those individuals who need to be aware within their organization. Return all affected products available to: Stryker Orthopaedics/PFA Product ReturnsAttn: Distribution Inventory Team325 Corporate DriveDock M-EastMahwah, NJ 074311360184/RA2016-112Complete and sign the enclosed Business Reply Form and fax a copy to 855-261-3635 or email to SO_M_PRODUCT_FIELD_ACTION_RESPONSE@stryker.com.Customers with questions were instructed to call 201-831-6693. |  | 131 | 131 units | Nationwide Distribution | Open, Classified |  |  | 131 | Computer
Z-2099-2017 | 77099 | No | MIH |  |  | Cardiovascular | PMA | Zenith Alpha Thoracic Endovascular Graft |  |  | March 22, 2017 | 2017 | Cook Medical Incorporated | There have been five reports of graft thrombosis/occlusion duringglobal commercial use, each following treatment for BTAI.One case resulted in patient death, and three cases resulted inreintervention. | Device Design | The firm, Cook Medical, sent an "URGENT: MEDICAL DEVICE CORRECTION" letter dated March 22, 2017 to the affected consignees via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. On 6/22/2017, updated "URGENT: MEDICAL DEVICE CORRECTION AND REMOVAL" notifications were sent to the affected consignees via courier. The update included removal of the indication for use in BTAI, as follows: The Zenith Alpha Thoracic Endovascular Graft is indicated for the endovascular treatment of patients with aneurysms or ulcers of the descending thoracic aorta having vascular morphology suitable for endovascular repair including:" Iliac/femoral anatomy that is suitable for access with the required introduction systems " Nonaneurysmal aortic segments (fixation sites) proximal and distal to the aneurysm or ulcer: o with a length of at least 20 mm, and o with a diameter measured outer-wall-to-outer-wall of no greater than 42 mm andno less than 20 mm.Additionally, the following warning has been added to describe the thrombus risk that has been observed when the device is used to treat BTAI:" Risk of in-graft thrombus has been observed when the Zenith Alpha ThoracicEndovascular Graft has been used to treat BTAI.Because of the IFU correction to remove BTAI from the indication, it is necessary to remove specific sizes of this device (grafts with a proximal or distal diameter 18  22 mm) that would likely be used only for BTAI.This notice must be shared with appropriate personnel, including down to the user level, within your organization or to any organization where the potentially affected devices have been transferred.Should you have any questions or concerns, please contact Cook MedicalCustomer Relations at 800.457.4500 or 812.339.2235. |  | 4916 | 4916 | Nationwide Distribution. | Open, Classified |  |  | 4916 | Computer
Z-0972-2017 | 76016 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Artis Q/Q.zen, Angiographic x-ray system |  |  | December 08, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Siemens initiated a corrective action to address two possible, mutually independent causes of a system defect related to the following:- In Artis Systems with A100 Plus or A100G generators, an attempt to resume operation following detection of a fault can result in the failure of a module in the high-voltage generator.- For biplane systems delivered with software version VD11B since April of this year, software problems in conjunction with graphics card may, in rare cases, result in the loss of image display in the examination room. | Software in the Use Environment | Siemens mailed a Customer Safety Advisory Notice (CSAN) dated December 8, 2016, to all affected customers with Customer Safety Advisory Notice AX002/16/S. Additionally, Siemens has initiated a software update via Update Instruction AX001/16/S. Their Service Organization will contact each customer to schedule an appointment to perform the corrective action; however, customers can make contact for an earlier appointment. Customers with questions should call 1-888-826-9702.For questions regarding this recall call 610-448-6461. |  | 1500 | 1,500 distributed Worldwide | Nationwide Distribution | Open, Classified |  |  | 1500 | Computer
Z-2778-2017 | 77607 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | INVUITY(R) PhotonBlade(TM), REF PB1 |  |  | June 12, 2017 | 2017 | Invuity, Inc. | A small number of customer observations were recently received which related to possible unintended energy discharge with the potential to cause tissue damage to patients. | Device Design | Official notices were sent to customers via FedEx overnight on 6/14/17. Email notices were sent to the sales personnel 06/12/2017 regarding the return of their trunk stock inventory and asking them to also contact customers who had been sent letters. |  | 1328 | 1328 units | nationwide | Open, Classified |  |  | 1328 | Computer
Z-2869-2017 | 77781 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Non-Locking screws in the following lengths: |  |  | July 13, 2017 | 2017 | In2Bones, SAS | The firm received several customer complaints reporting some perioperative breakages of NE0411 Non-Lockingscrews when inserting the screw into the radius diaphysis | Device Design | In2Bones sent a Field Safety Corrective Action letter dated July 13, 2017, to all affected customers. The consignee was directed to remove the products from distribution and return it to the recalling firm. Customers with questions were instructed to contact the Quality and Regulatory Affairs team at: +33 4 72 29 26 26/ +337 61 88 44 30 or by email: qualite@in2bones.com. |  | 287 | 287 screws | Worldwide Diribution - US including Tennessee; Foreign - Italy, Portugal, Spain, Switzerland, UK, and France | Open, Classified |  |  | 287 | Computer
Z-0989-2017 | 76102 | No | NEP | Intravascular administration set. | General Hospital | General Hospital | 510(K) Exempt | i.v. Station and i.v. Station ONCO Automated Drug Compounding System, Pharmacy Compounding System |  |  | November 23, 2016 | 2016 | HEALTH ROBOTICS S.R.L. | The firm has initiated a product correction due to a bag carousel motor error causing misalignment and resulting in a mislabeling of final IV bag preparations. | Software design | Omincell mailed an Urgent Medical Device Field Correction Notification letter to affected customers on 11/21/2016. The letter provided information of the event and instructions on how to continue operating the device to prevent risks. The issue will be corrected with a software update. |  | 121 | 121 systems | Shipped throughout United States. | Terminated | March 15, 2017 | 112 | 121 | Computer
Z-1006-2017 | 76024 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | IQon Spectral CT with software version 4.7.0 |  |  | December 16, 2016 | 2016 | Philips Medical Systems (Cleveland) Inc | Multiple issues have caused the device to result in CT rescans or incorrect scan location or misrepresentation of image results. | Software design | Philips sent a Customer Information Letter dated December 13, 2016 to customers via certified mail. The letter identified the affected product, problem and actions to be taken. Additionally, a Field Service Engineer will contact each site to schedule a time to implement the Field Change Order. For further information or support concerning this issue contact your local Philips representative or local Philips Healthcare office. For North America and Canada, contact the Customer Care Solutions Center (1-800-722-9377). |  | 35 | 35 | Worldwide Distribution - US Nationwide in the states of AR, IN, LA, MA, MN, OR, TN &amp; TX. and the countries: Australia, Belgium, Denmark, France, Germany, Israel, Japan, Republic of Korea, Poland, Portugal, Slovenia, Switzerland, Thailand and United Kingdom. | Open, Classified |  |  | 35 | Computer
Z-0987-2017 | 76085 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Oncentra External Beam Oncentra Brachy |  |  | December 23, 2016 | 2016 | Elekta, Inc. | Cross profile for Varian 60 degree wedge shows "horns." | Software design | Elekta sent an Urgent Important Safety Notification letter dated December 23, 2016 to affected customers. The notice informs users of the specific product and version numbers affected by the issue and any work arounds that can be used to avoid the issue. The letter included an acknowledgement form to be returned to Elekta. |  | 433 | 433 | Worldwide Distribution - US Nationwide in the states of AL, CA, DC, NE, NY, ND, OR, PA, RI, SC, TX, and the countries of Angola, Argentina, Armenia, Australia, Austria, Barbados, Belarus, Brazil, Bulgaria, Canada, China Columbia, Curacao, Czech Republic, Denmark, Finland, France, Germany, Greece, Honduras, Hungary, Iceland, India, Indonesia, Ireland, Italy, Japan, Kazakhstan, Kenya, Malaysia, Netherlands, Norway, Oman, Peru, Poland, Portugal, Romania, Russia, Slovakia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom, Yemen, Zambia. | Open, Classified |  |  | 433 | Computer
Z-2684-2017 | 77621 | No | LHN | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Proteus 235, Proton Therapy System |  |  | December 13, 2016 | 2016 | Ion Beam Applications S.A. | Software issue | Software design | An Urgent Field Safety Notice was mailed to affected customers to inform them of the issue. The notice indicated possible software failure, the associated risks and actions users should implement to correct the issue. IBA plans to follow-up with a software update which will be deployed on impacted sites to correct this issue. |  | 11 | 6 worldwide and 5 in the U.S. | Distributed to FL, VA, IL, NJ, WA and South Korea | Open, Classified |  |  | 11 | Computer
Z-2688-2017 | 77324 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Sensis Vibe Systems with Software Version VD10B, Model Numbers 10765502, 10910620, 11007641, 6648161 --- Programmable diagnostic computer, |  |  | May 16, 2017 | 2017 | Siemens Medical Solutions USA, Inc | Software error. In Sensis Vibe systems with software version VD10B, a software error can result in: problems generating a report and/ or- information from different examinations of the same patient being combined into one report. --- The error causes information from two examinations to be combined into one report. This does not result in a hazardous situation; however, if the erroneous report has been used for treatment or diagnosis, this can potentially result in an incorrect treatment or diagnosis. | Software design | Siemens sent an Customer Safety Advisory Notice dated May 16, 2017, was mailed to all users on May 16, 2017. The letter informs customers of the software issue and potential impact. Siemens Medical Solutions will contact users to make arrangements to update the system with updated software to correct the issue. For further questions, please call (610) 448-6461. |  | 13 | US: 13 units | US Distribution to the states of : IN, MI, IA | Open, Classified |  |  | 13 | Computer
Z-1246-2017 | 75914 | No | NFJ | Picture archiving and communications system. | Radiology | Ophthalmic | 510(k) | Merge Eye Station Import Utility (ESIU) when used with Merge Eye Station and Merge Eye Care PACS. The firm name on the labeling is Merge Healthcare, Hartland, WI. |  |  | March 28, 2016 | 2016 | Merge Healthcare, Inc. | Eye Station images were not importing properly and were imported under "unknown" due to an issue when validating patients using only an Medical Record Number (MRN) | Device Design | The recalling firm issued a letter dated 3/28/2016 via e-mail. If no email address was available, the letters were issued via certified mail. A follow-up letter dated 7/11/2016 was also issued. |  | 1627 | 1,627 sites potentially have the affected versions | Distribution was nationwide to medical facilities. There was also government and foreign distribution, including Canada. | Open, Classified |  |  | 1627 | Computer
Z-1816-2017 | 76506 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | TotalCare Bed Systems, Model Numbers P1900, PR1900, and U1900 with 3-Level Patient Position Monitoring and Scale Upgrade Kits, Model Numbers P1947A and P1948A). |  |  | February 23, 2017 | 2017 | Hill-Rom, Inc. | The bed exit alarm may not function properly if the alarm has been armed for a period longer than seventy-two (72) hours without the alarm being triggered by patient movement or removal of patient weight within that time period. | Software change control | On 2/23/2017, URGENT: MEDICAL DEVICE CORRECTION notifications were sent to the affected customers via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Customers without a service contract received the following instructions:Action to be taken by the user: A software update for the TotalCareÂż beds GCI has been released to resolve this issue. Enclosed with this letter are a USB stick with the updated software, instructions on how to update the GCI software, and a response form to be returned within 30 days. 1) Identify any affected TotalCareÂż beds with 3-Level Patient Position Monitoring in your facility by using the criteria above. 2) Please use the enclosed instructions and update the software on each potentially affected TotalCareÂż bed immediately. 3) Record the serial numbers of the beds being updated on the response form and return to Hill-Rom within 30 days of receipt of this letter. Additional USB sticks with the updated software may be obtained by contacting Technical Support at 1-800-445-3720 (select Option 2, then Option 2 again). Until each of the affected TotalCareÂż beds are updated with the new software, we recommend you do one or more of the below to avoid any potential occurrences of bed exit not functioning. 1. Power cycle (unplug and plug in the bed mains power cord) once every 48 hours. 2. Remove the Patient and re-zero the scale once every 48 hours. 3. Turn off and do not use the Bed Exit Alarm functionality. Continued use of the bed exit alarm without the mitigations listed above may cause risk to patient. If you have an upgrade kit (P1957A) that has not been installed, please contact Hill-Rom for return and replacement. Customers with a service contract received the following instructions:Action to be taken by the user: A Hill-Rom representative per |  | 3037 | 3037 | US Nationwide. AUSTRALIABANGLADESHBRAZILCANADACHINAGERMANYINDIAITALYKUWAITPAKISTANSAUDI ARABIATHAILANDUNITED ARAB EMIRATESCYPRUSGREECENETHERLANDS | Open, Classified |  |  | 3037 | Computer
Z-1899-2017 | 76870 | No | NDC | Predictive pulmonary-function value calculator. | Anesthesiology | General &amp; Plastic Surgery | 510(k) | Accu-Chek Connect Diabetes Management App versions 2.0.0, 2.0.1 and 2.1.0 for iOS and Android |  |  | March 14, 2017 | 2017 | Roche Diabetes Care, Inc. | A program error (bug) in the Bolus Advisor feature, which could result in incorrect bolus advice and a potential insulin over-delivery. | Software design | Roche Diabetes Care, Inc. discovered a software issue in the Accu-Chek Connect app versions 2.0.0, 2.0.1, and 2.1.0 that can result in incorrect bolus recommendations. This software bug may result in errors in the active insulin calculation, creating the potential for a false high insulin bolus recommendation. An insulin bolus delivered based on potentially incorrect bolus advice could lead to serious health consequences such as hypoglycemia. Customer notices were distributed via mail on 03/30/2017, and the firm's website was also updated with recall information on that day. Customers were instructed to do the following: discontinue use of the bolus advisor feature in app versions 2.0.0, 2.0.1, and 2.1.0 (iOS &amp; Android) until the corrected version is available.Âż Provide patients with a copy of the enclosed Urgent Medical Device Correction 17-001.Âż Complete the enclosed faxback form and fax it to Roche Diabetes Care, Inc., at 1-888-912-8456.Âż If you have questions about the information in this notification, please contact Accu-Chek Customer Care by visiting accu-chek.com or by calling 1-800-688-4578.The corrected version of the app is expected to be available on or about April 13, 2017. New patients will be able to begin using the bolus advisor feature once the corrected version is available. Please advise your patients to visit accu-chek.com for updates regarding the availability of the corrected version of the app. |  | 64324 | 64,324 | US Nationwide in the states of: AL, AR, AZ, CA, CO, CT, DE, DC, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV | Open, Classified |  |  | 64324 | Computer
Z-1045-2017 | 76187 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Ambra PACS |  |  | December 16, 2016 | 2016 | DICOM GRID INC | A software error caused the window/level to become the same in one series regardless if the image had different levels; image results have a washed-out grey appearance. | Software change control | On 12/16/16 a fix is the software issue was conducted via remote cloud system. On 01/11/17 customers were sent Release Notes that addressed the software correction. The issue was resolved and required no further action. |  | 209 | 209 | U.S. distribution to the following; TX, FL, NE, CA, WI, AZ, NH, MO, MA, CO, PA, NV, PR, IL, OH, NY, MD, NJ, MN, MI, CT, VA, GA, UT, KS, WA, MI, HI.Foreign distribution to the following; Canada, Spain, Israel, Singapore, Kosovo, Austria, Brazil, Malaysia, Australia. | Terminated | April 07, 2017 | 112 | 209 | Computer
Z-1204-2017 | 75942 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | CentraLinkÂż Data Management System Software Versions: v13x,v14x,v15x |  |  | November 29, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc. | There is a remote possibility CentraLink may download an order to the ADVIA Automation System without specifying the sample type. This can occur when an order is received from the LIS without a sample type, requiring that the sample type be set in CentraLink based on the sample type of the test in the order. | Device Design | Siemens Healthcare Diagnostics, sent an Urgent Medical Device Correction Notice dated November 29, 2016, to all affected customers within the USA, and and issued notices titled "Urgent Field Safety Notices" to their international customers. Effectiveness Check Forms were also sent to the affected customers. Siemens recommends the LIS specify the sample type in the work order. The CentraLink system is designed to operate in an environment in which the sample type is specified in the work order that is downloaded from the LIS. A Siemens Technical Support Representative will contact affected customers to review your CentraLink configuration. If your LIS does not include sample type information in the sample orders downloaded to CentraLink and Pre-Tests assigned by sample type are configured on the ADVIA Automation System your representative will update your configuration. Complete and return the Field Correction Effectiveness Check Form attached to the letter within 30 days. Please retain this letter with your laboratory records and forward this letter to those who may have received this product. If you have any questions, please contact your Siemens Customer Care Center or your local Siemens Technical Support Representative.For further questions, please call (914) 631-8000. |  | 3893 | 3,893 units | Worldwide Distribution - USA (nationwide) to the states of : AL, AR, AZ, CA, CO, CT FL, GA, IL, IN, KS, LA, MA, MD, ME, MI, MN, MO, NC, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, WA and WV; and to the country of Canada and others. | Open, Classified |  |  | 3893 | Computer
Z-1457-2017 | 76447 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Merge Hemo software. |  |  | April 04, 2016 | 2016 | Merge Healthcare, Inc. | In some instances, the system will lock tabs within a study, even when a second user does not have the study open on a different workstation. | Software design | The recalling firm issued a letter dated March 28, 2016 via e-mail. If no email address was available, the letters were issued via certified mail. Non-responders were issued letters on July 11, 2016. The letter identified the affected product, problem and actions to be taken. For question send an email to recall@merge.com. |  | 361 | 361 sites potentially have the affected versions | US Nationwide Distribution. | Open, Classified |  |  | 361 | Computer
Z-1460-2017 | 76429 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Arkon Anesthesia Delivery System, Model 99999. |  |  | February 08, 2017 | 2017 | Del Mar Reynolds Medical, Ltd. | Spacelabs has received reports of the Arkon Anesthesia Delivery System, Model 99999, shutting down without warning. The power supply has failed in a way that does not trigger the normal on battery alarms. | Software design | On 08 February 2017, a customer letter will be sent via priority service, return receipt requested, to all U.S. customers. Records of notification will be maintained to ensure complete notification.On 15 February 2017, a customer letter (translated as necessary) will be emailed to all international subsidiaries and distributors of record. Records of notification will be maintained to ensure complete notification.Upgrade is expected to be available in April 2017; all affected Arkon Anesthesia Delivery Systems will receive a software upgrade onsite. |  | 450 | 325 devices sold to 35 US consignees and 76 devices sold to 14 international consignees | Worldwide Distribution - US nationwide and internationally in the following countries: BOLIVIA, CANADA, CHINA, ECUADOR, FRANCE, IRELAND, MALAYSIA, MEXICO, MOROCCO, NETHERLANDS, PANAMA, PERU, POLAND, UNITED KINGDOM. | Open, Classified |  |  | 450 | Computer
Z-1463-2017 | 76339 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | The ABL800 Model#:All |  |  | January 26, 2017 | 2017 | Radiometer America Inc | Reports that the sample type "cord blood" has been changed into the two sample types "cord blood arterial" and "cord blood venous. | Device Design | Customer notification letters were sent 01/27/16. The letter provided the following instructions:The firm recommends that customers ensure that operators and clinicians using the "cord blood sample results from a HIS/LIS or a middleware system to determine treatment are made aware that arterial or venous is not transmitted.An upgraded version of the ABL90 and ABL800 series software will ensure that the full sample type cord blood arterial and cord blood venous will be transmitted to the HIS/LIS or middleware system. The new software version will be installed by local engineer when available.If not the end user customers are instruction to provide a copy of the recall notification letter to the end user.Any questions regarding this letter contact Radiometer America Technical Support at 1-800-736-0600 opt.1. |  | 17522 | 17,522 | U.S. distribution to the following; nationwide.Foreign distribution to the following; Netherlands, China, Czech Republic, Denmark, Germany, Spain, France, Hungary, India, Japan, Poland, South Africa, Singapore, Switzerland, Turkey, Norway, Austria, Italy, Sweden, Finland. | Open, Classified |  |  | 27210 | Computer
Z-1200-2017 | 76281 | No | LFZ | Cytomegalovirus serological reagents. | Microbiology | Microbiology | 510(k) | VIDAS 3 software v. 1.1.4 |  |  | January 11, 2017 | 2017 | BioMerieux SA | During development of the VIDAS 3 software version 1.2, some anomalies have been identified and observed to be already present in the current software version VIDAS 3 version 1.1.4. available in the field. | Software in the Use Environment | Consignees were notified via letter on January 11th, 2017. |  | 1161 | 1161 units | U.S. | Open, Classified |  |  | 1161 | Computer
Z-1201-2017 | 76344 | No | MWS | Cardiovascular surgical instruments. | Cardiovascular | Cardiovascular | 510(K) Exempt | Hercules 360 Universal Stabilizing Arm, top mount, 360-degree rotation, reusable stainless steel |  |  | August 20, 2014 | 2014 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (Terumo CVS) is recalling certain lots of Hercules TM 360 Universal Stabilizing Arms because the two pins that hold the sleeve attached to the distal end of the device may come loose, causing the pins and sleeve to separate from the device. While the arm can still function without the pins and sleeve in place, there is potential for the loose components to fall into the surgical site. | Device Design | Terumo Cardiovascular Systems has contact all affected parties (user facilities and consignees) via a Customer Notification Letter that advises of the recall activity. |  | 260 | The total number of distributed devices stands at 260 | Terumo Cardiovascular has identified 16 domestic consignees and 30 international consignees . | Completed |  |  | 260 | Computer
Z-1198-2017 | 76257 | apparatus, traction, non-powered | HST | Nonpowered orthopedic traction apparatus and accessories. | Orthopedic | Neurology | 510(K) Exempt | ArthrexÂż Shoulder Suspension Tower; Catalog Number: AR-1650-01 |  |  | January 03, 2017 | 2017 | Allen Medical Systems, Inc. | During use, there is a potential scenario which could allow the fastening screws attaching covers to come loose and potentially land in the sterile field causing a delay in surgery. | Device Design | The distributor was notified of the recall via email on January 3, 2017. The letter identified the affected product, problem and actions to be taken. The letter instructs the consignee to remove the product from service and follow instructions to return the device to Allen Medical for repair. |  | 112 | 112 devices | US state of FL | Open, Classified |  |  | 112 | Computer
Z-1710-2017 | 76688 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access 2 assay Immunoassay System, Remanufactured, Part No. 386220 |  |  | March 07, 2017 | 2017 | Beckman Coulter Inc. | Beckman Coulter is recalling the Access 2 Family of Immunoassay Systems because a circuit board that controls motor functions can affect device positioning on the pipettor and incubator belt while the instrument is running. The position errors cause the instrument to stop running and cancel tests that are in process. | Software Design Change | An Urgent Medical Device Recall letter was sent on 3/7/17 to inform customers that Beckman Coulter has determined through internal investigation that a circuit board that controls motor functions in the instrument systems listed above can affect device positioning on the pipettor and incubator belt while the instrument is running. Customers are informed of the impact, action to be taken, and the resolution. Customers are instructed to complete and return the enclosed Response Form within 10 days. Customers with any questions are instructed to contact Customer Technical SupportCenter:Âż From our website: http://www.beckmancoulter.comÂż By phone: call 1-800-854-3633 in the United States and Canada.Âż Outside the United States and Canada, contact your local Beckman Coulterrepresentative. |  | 851 | 851 units total for all products (291 units in US) | Affected geographies: Algeria, Austria, Australia, Bangladesh, Bulgaria, Brazil, Canada, China, Costa Rica, Cyprus, Colombia, Czech Republic, El Salvador, France, Germany, Ghana, Greece, Guatemala, India, Indonesia, Italy, Japan, Jordan, Kazakhstan, Kenya, Lebanon, Lithuania, Malawi, Mexico, Morocco, Myanmar, Nepal, Oman, Panama, Philippines, Pakistan, Paraguay, Poland, Puerto Rico, Romania, Russian Federation, Saudi Arabia, Serbia, South Africa, South Korea, Spain, Switzerland, Tunisia, Turkey, United Kingdom, United States, Vietnam, Zimbabwe | Open, Classified |  |  | 851 | Computer
Z-1197-2017 | 76194 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | STERRAD NX Sterilization System, Product Code: 10033, 10033-002 |  |  | January 09, 2017 | 2017 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) has identified that in a rare sequence of events following an external power outage, the system door interlock can be become disengaged during system reboot for a 20-second duration. | Software design | Advanced Sterilization will send an Field Safety Notice dated January 9, 2017,to inform customers that Advanced Sterilization Products (ASP) has identified that in a rare sequence of events following an external power outage, the system door interlock can be become disengaged during system reboot for a 20-second duration. ASP is sending the letter to remind customers to follow the instructions in the STERRAD NX System Users Guide. Specifically, after cycle initiation, the user should only open the chamber door of the STERRAD NX System whenprompted by the Graphical User Interface (GUI) per the Users Guide. Customers are instructed to complete and return the attached Business Reply Card (double-sided) for their acknowledgement within 3 business days of receipt of this notification letter via mail by using the pre-paid postage, by emailat ASP5816@stericycle.com, or by fax to Stericycle: 8662297775. For questions regarding the business reply card, customers are instructed to contact Stericycle at (877) 497- 2425 and reference event # 5816.For further questions regarding this recall please call (949) 453-6400, |  | 7694 | 7,694 units total (5,935 units in US) | Worldwide Distribution - USA (nationwide) including Puerto Rico and to the countries of : Australia, China, Hong Kong, India, Indonesia, Korea (South), Malaysia, New Zealand, Philippines, Singapore, Sri Lanka, Taiwan, Thailand, Vietnam, Canada, Belgium, Czech Republic, Egypt, France, Germany, Great Britain, Greece, Hungary, Ireland, Israel, Italy, Netherlands, Poland, Portugal, Qatar, Russian Federation, Saudi Arabia, Slovenia, South Africa, Spain, Sweden, Switzerland, Turkey, United Arab Emirates, Japan, Argentina, Bolivia, Brazil, Chile, Colombia, Ecuador, Guatemala, Mexico, Panama, Peru, Uruguay, and Venezuela. | Open, Classified |  |  | 7694 | Computer
Z-1194-2017 | 76292 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | BONE MARROW ASPIRATION NEEDLE 15G x 2.688 MAX and 15G x 4 MAX |  |  | January 18, 2017 | 2017 | Argon Medical Devices, Inc | The recalling firm received a complaint from one of their customers of a potential defect in the forming ofthe primary package of the blister that may compromise the sterility of the product inside the tray. | Equipment maintenance | Argon Medical Devices, Inc. is sending an Urgent Product Recall Notice letter dated January 17, 2017, by registered mail to all affected customers. Customers were asked to determine if their inventory shows the package defect and discontinue use of the device and promptly return the affected product. For questions regarding this recall call 903-403-0749. |  | 497 | 497 boxes | Worldwide Distribution - US (nationwide) and Internationally to Switzerland, Canada, Australia, New Zealand, Ecuador, Colombia, and Japan. | Terminated | May 17, 2017 | 119 | 1284 | Computer
Z-1480-2017 | 76483 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | 882480: BrightView designed for single or dual detector nuclear imaging accommodating a range of ECT studies. In addition it can be used to perform planar static, dynamic, gated, total body, circular-orbit and noncircular orbit SPECT, gated SPECT (circular and noncircular)studies, computer-programmed protocol strings, and reference scans (dual detectors). |  |  | January 31, 2017 | 2017 | Philips Medical Systems (Cleveland) Inc | Four issues: 1. Motion controller problem stops scan and no data image produced. 2. Door interlock switch problem disables CT scan. 3. Detector contacts head holder when performing Patient Unload. 4. JETStream freezes during gated planar scan | Software design | Consignees received a customer information letter via certified mail, 1/31/2017. International consignees were notified by Philips Healthcare Global Market representatives. Field Service Engineers are contacting consignees to conduct appropriate corrections. Non-responders will be contacted multiple times. |  | 1218 | 1218 total | Worldwide distribution. US nationwide, ALGERIA, ARGENTINA, ARMENIA, AUSTRALIA, AUSTRIA, BAHRAIN, BELGIUM, BRAZIL, CANADA, CAYMEN ISLANDS, CHILE, CHINA, CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, EGYPT, FINLAND, FRANCE, GABON, GERMANY, GREAT BRITAIN, GREECE, INDIA, INDONESIA, IRAN, IRELAND, ISRAEL, ITALY, JAPAN, KOREA, LEBANON, LITHUANIA, MALAYSIA, MEXICO, NETHERLANDS, NEW ZEALAND, NORWAY, OMAN, PALESTINE, PANAMA, PERU, PHILIPPINES, POLAND, PORTUGAL, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, SYRIAN ARAB REPUBLIC, TAIWAN, THAILAND, TURKEY, UKRAINE, UNITED ARAB EMIRATES, UNITED KINGDOM, &amp; VIETNAM. | Open, Classified |  |  | 1218 | Computer
Z-1103-2017 | 75971 | No | PAO | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Comprehensive Reverse Shoulder System Humeral Tray |  |  | December 20, 2016 | 2016 | Zimmer Biomet, Inc. | Higher than anticipated rate of fracturing due to design. Risk of revision surgery due to fracturing. | Device Design | On 12/20/2016, URGENT MEDICAL DEVICE RECALL-LOT SPECIFIC notifications were sent to the affected Risk Managers and Distributors via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Risk Manager Responsibilities: 1. Review this notification and ensure affected personnel are aware of the contents. 2. Assist your Zimmer Biomet sales representative quarantine all affected product. 3. Your Zimmer Biomet sales representative will remove the affected product from your facility. 4. Complete Attachment 1  Certificate of Acknowledgement. a. Return a digital copy within three (3) days to corporatequality.postmarket@zimmerbiomet.com. b. Retain a copy of the Acknowledgement Form with your recall records in the event of a compliance audit of your facilities documentation. 5. If after reviewing this notice you have further questions or concerns please call 411 Technical Services at 574-371-3071 between 8:00 am and 5:00pm EST, Monday through Friday. Calls received outside of the call center operating hours will receive a prompt to record a voice mail or be transferred to an on-call representative in the case of an emergency. Alternatively, your questions may be sent by email to corporatequality.postmarket@zimmerbiomet.com. Distributor Responsibilities 1. Review this notification and ensure affected team members are aware of the contents. 2. Immediately locate and quarantine affected product in your inventory. 3. Complete the Certification of Acknowledgement portion of Attachment 1  Inventory Return Certification Form a. Return a digital copy to corporatequality.postmarket@zimmerbiomet.com within three (3) days. 4. Immediately return all affected product from your distributorship and affected hospitals within your territory along with a completed Attachment 1  Inventory Return Certification Form |  | 268 | 268 | Nationwide. CANADA AUSTRALIA CHILE COLOMBIA KOREANetherlands | Open, Classified |  |  | 268 | Computer
Z-1499-2017 | 76047 | No | ILQ | Powered communication system. | Physical Medicine | Physical Medicine | 510(K) Exempt | Arial Wireless Water-Resistant Call Pendant |  |  | June 22, 2016 | 2016 | Stanley Security Solutions Inc | Devices were incorrectly programmed during manufacturing therefore depressing the pendant button may result in an alarm not sounding as intended. | Software design (manufacturing process) | Stanley Healthcare sent an Important Recall Notice dated June 22, 2016, to all affected customers. The notice instructed customers to discontinue use of the device and to contact the recalling firm to arrange for return of the device. The recalling firm also asked the customer to complete a response form indicating they read, understood and complied with the recall notice. Customers with questions were instructed to contact Arial Customer Service at 888-716-0077. For questions regarding this recall call 781-216-9035. |  | 500 | 500 individual pendants | Nationwide Distribution to AR, AZ, CA, FL, GA, IA, ID, IN, KS, MD, MN, MO, NC, NE, NJ, NY, OH, OR, PA, SD, UT, VA, and WI. | Open, Classified |  |  | 500 | Computer
Z-1498-2017 | 76312 | No | OUG | Medical device data system. | General Hospital | General Hospital | 510(K) Exempt | Merge Eye Station Import Utility (ESIU). The firm name on the labeling is Merge Healthcare. |  |  | April 04, 2016 | 2016 | Merge Healthcare, Inc. | System locks up which may result in potential patient injury or delay in diagnosis or treatment. | Software design | Merge Healthcare, Inc. sent an Urgent: Medical Device Recall letter dated March 28, 2016, via email on April 4, 2016. If no email address was available, the letters were issued via certified mail. Nonresponders were issued letters dated July 11, 2016, via email.The letter notifies the customer of the issue and informs them of the workaround and that a fix was not yet available, although they were working on one. The notification instructs the customer to pass the information onto all other users, as well as any downstream customers if it was further distributed. The notification also informs the customer their response is required using the enclosed form and the return addressed envelope. The response form asks, in part, if they understand the workaround. The response was required no later than April 18, 2016.Nonresponders were issued a second letter dated July 11, 2016, asking for a response by August 2, 2016. The letter and response form contained the same information as the first letter. If the second letter results in additional unresponsive customers, a call campaign will be managed by Customer Support and documented via SalesForce. Three call attempts will made before closing the case.The recalling firm reported emails will be tracked for delivery confirmation and the letters will be tracked for delivery by USPS certified services.For further questions, please call (877) 741-5369. |  | 78 | 78 sites potentially have the affected accessory | Distribution was made USA (nationwide) to medical facilities. Government distribution was made. Foreign distribution was made to Canada. There was no military distribution. | Open, Classified |  |  | 78 | Computer
Z-1500-2017 | 76405 | No | NFJ | Picture archiving and communications system. | Radiology | Ophthalmic | 510(k) | EyeSuite i.8.2.1.0 Software for ophthalmic use including selection of Intra Ocular Lenses (IOLs). |  |  | January 05, 2017 | 2017 | Haag-Streit USA Inc | There is a possibility for data to be stored under the wrong patient on the DICOM Server following a non-standard workflow. | Software design | On January 5, 2017 the firm sent Field Safety Notice to their customers. |  | 163 | 163 units | US Nationwide distribution. | Open, Classified |  |  | 163 | Computer
Z-1492-2017 | 76683 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance BigBore Oncology CT - 728243 |  |  | February 23, 2017 | 2017 | Philips Medical Systems (Cleveland) Inc | When the user performs a retrospective respiratory gated 4D CT scan with the Bellows pulmonary gating device, the CT images might be reconstructed at a single phase, while the annotations on the images incorrectly indicate that the reconstruction is from the requested specific phases (0% to 90%). | Software design | Philips sent an Urgent Field Safety Notice dated February 24, 2017, to all affected consignee's via certified mail with returned receipts for tracking purposes. Philips Key Market representatives will distribute the notices internationally. A Field Service Engineer will contact each site to schedule a time to implement the Field Change Order. Implementation of the software update will be monitored. Consignees requiring further information or support concerning this issue were instructed to contact their local Philips representative. For North America and Canada, contact the Customer Care Solutions Center at 1-800-722-9377. |  | 11 | 11 | Worldwide Distribution - US to AR, CA, CO, CT, FL, GA, IL, IN, KY, LA, MA, MD, MI, MO, NC, NE, NJ, NM, NY, OH, TX &amp; WI; Internationally to the following countries: Australia, Austria, Canada, China, Czech Republic, Germany, India, Israel, Jordan, Korea, Mexico, New Cldonia, Romania, Russian Federation, Turkey, &amp; United Kingdom. | Open, Classified |  |  | 78 | Computer
Z-1190-2017 | 76236 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Offset Reamer Handle, T5766 |  |  | October 19, 2016 | 2016 | Greatbatch Medical | On June 14, 2016, Greatbatch discovered a significant increase in the number of field complaints for the T5766 Offset Reamer Handle in which the U-joints of the drive chain component seizes during orthopedic procedures rendering the device non-functional. The 2016 failure rate for this failure exceeds the occurrence rate as defined in the DFMEA. | Device Design | Consignee was sent via e-mail on 10/16/2016 a Greatbatch "Urgent Medical Device Recall" letter dated 10/19/2016 The letter described the purpose of the letter, Reason for Notification, Risk to Health, Recommendation, Instruction to Zimmer. Advised consignee to examine their inventory, quarantine the product and provide a list of the affected countries where the product was distributed. Also requested consignee to complete the Customer Notification Response form and return to Greatbatch at FieldActionCenter@Greatbatch.com. For questions contact your local Greatbatch Medical representative, Monday through Friday via telephone at 1-763-951-8235 8:00 AM (CT) to 5:00 PM (CT), or e-mail us at FieldActionCenter@Greatbatch.com . |  | 1441 | 1,441 | Worldwide Distribution - US Nationwide in the states of IN and the countries of Canada, Austria, Germany, Japan. | Open, Classified |  |  | 1441 | Computer
Z-1188-2017 | 76378 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Discovery MR450 |  |  | January 20, 2017 | 2017 | GE Healthcare, LLC | Potential safety issue with the patient bore heating on the Discovery MR450 narrow bore 1.5T MRI. In the event of a RF component failure, typically happening after extended system use, a small area on the bore surface may become warmer than normal. If this happens in an area where there is patient contact with the bore, and proper padding is not used, it may cause a serious burn. | Device Design | The firm, GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION" letter GEHC Ref# 60913 Dated January 20, 2017. The letter was addressed to Hospital Administrators / Risk Managers, Radiology Department Managers &amp; Radiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction &amp; Contact Information. For questions contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 60 | 60 | Worldwide Distribution- US (nationwide) including states of: CA, DC, IL, MN, MO, NY, SC, TX, WI, WV; and countries of: Canada, Czech Republic, Denmark, Finland, France, Germany, Italy, Netherlands, Norway, Poland, Saudi Arabia, Spain, Sweden, Switzerland, Taiwan, United Kingdom, and Uruguay. | Open, Classified |  |  | 60 | Computer
Z-1801-2017 | 76727 | No | CHP | Estradiol test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VIDAS Estradiol II, Ref 30431-01, 60 tests |  |  | March 17, 2017 | 2017 | BioMerieux SA | Fulvestrant cross reacts with the VIDAS Estradiol II immunoassay and may cause falsely elevated estradiol results in patients treated with Fulvestrant. | Device Design | bioMerieux sent an Important Product Safety Notice A Product Safety Notice to all affected customers on March accounts on March 17, 2017, informing them not to use VIDAS Estradiol II test to test estradiol level in patients under Fulvestrant therapy. The letter did not request return of any of the product, and requested that they return an acknowledgement form acknowledging receipt of the notice and that they have followed the instructions and implemented the actions in the field notice. Customers with questions were instructed to contact their bioMerieux Customers Service representative. For questions regarding this recall call 314-731-8805. |  | 834 | 834 kits | Nationwide Distribution including AZ, CA, CO, FL, GA, IN, IA, LA, MI, MS, NY, NC, OH, OK, OR, PA, TX, WA | Open, Classified |  |  | 834 | Computer
Z-1494-2017 | 76376 | No | LZS |  |  | Ophthalmic | PMA | Final Fit Software Version 1.11 and 1.12; |  |  | February 01, 2017 | 2017 | Nidek Inc | During treatment planning, the procedure was programmed with an unintended (wrong) correction. | Software change control | Nidek sent a Recall notification letter dated February 1, 2017 to all affected consignees by certified mail method beginning February 2, 2017. Letters advised of the issue received from a complaint and the steps the user took that created the issue. To prevent the issue, the letters provide steps that should be taken. Nidek's Customer Support team will contact customers by phone or email. Customers with questions were instructed to contact Customer Service at 1-800-223-9044 or email ff_support@nidek.com. |  | 29 | 29 | Nationwide Distribution to MN, NV, OH, NC, CA, CO, CA, GA, VA, MI, AZ, PA, TN, WA, TX, NY. | Open, Classified |  |  | 29 | Computer
Z-1490-2017 | 76712 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ONCOR" Avant-garde |  |  | February 14, 2017 | 2017 | Siemens Medical Solutions USA, Inc | Software update | Software design | Siemens mailed an Customer Information Letter on 02/15/2017 to affected customers to inform them of a software update they plan to implement to reduce risk of collision when delivering automatically sequenced treatments with automatic movements of the gantry and/or the treatment table, using the SIMTEC-Auto Field Sequence Option. |  | 80 | 80 systems | Distributed throughout the United States | Open, Classified |  |  | 80 | Computer
Z-1517-2017 | 76719 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge Cardio software. |  |  | September 22, 2015 | 2015 | Merge Healthcare, Inc. | Users can merge a device import file with an image study that already has a confirmed report, which may result in including information not present at the time of physician interpretation. | Software design | On 9/22/2015, the recalling firm issued letters dated 9/8/2015 and flagged "PRODUCT INFORMATION" via email to their customers. The recalling firm also issued letters dated 3/28/2016 via email the week of 4/4/2016. If no email address was available, the letters were issued via certified mail. A letter dated 7/11/2016 was issued to nonresponders. |  | 89 | 89 sites potentially have the affected software | US Nationwide Distribution. | Open, Classified |  |  | 89 | Computer
Z-1179-2017 | 75936 | No | JGP | Total protein test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS Chemistry Products UPRO Slides, UPN 10758750000593, |  |  | December 08, 2016 | 2016 | Ortho-Clinical Diagnostics | There is a potential for variable negatively and positively biased urine protein results using VITROS UPRO Slides. | Device Design | On December 8, 2016, Ortho Clinical Diagnostics distributed Product Correction Notification letters to their customers via FedEx, ORTHO PLUS e-Communications and/ or US Postal Service Priority Mail. Customers are advised to check their stock for affected product and to immediately discontinue and discard all remaining inventory. Ortho will credit your account for any discarded VITROS UPRO Slides, VITROS Chemistry Products Calibrator Kit 10 and/or VITROS Chemistry Products UPRO Performance Verifier I &amp; II. Please complete the Confirmation of Receipt and return by December 22, 2016. Customers with questions, please contact Ortho CareTM Technical Solutions Center at 1-800-421-3311. |  | 3381390 | 37,571 packs/90 slides totalling 3,381,390 slides | Worldwide distribution - Nationwide U.S.A; International - Bermuda, Australia, Canada, Chile, China, Colombia, India, Japan, Mexico, Republic of Panama, Singapore, Venezuela, England, France, Germany, Italy, Spain and Poland. | Open, Classified |  |  | 3381390 | Computer
Z-1180-2017 | 76258 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge RadSuite software. The firm name on the label is Merge Healthcare, Hartland, WI. |  |  | November 17, 2015 | 2015 | Merge Healthcare, Inc. | The values provided from the Pixel Value tool do not appear to be correct, which may result in potential patient injury or delay in diagnosis or treatment. | Software design | Merge sent an Urgent Medical Device Recall letter dated March 28, 2016 to all affected customers. The letter notifies the customer of the issue and informs them an upgrade is available. The notification instructs the customer to pass the information onto all other users, as well as any downstream customers if it was further distributed. The notification also informs the customer their response is required using the enclosed form and the return addressed envelope. The response form asks, in part, if they are interested in accepting the fix. If they decline the fix, they are asked to provide the reason why. The response was required no later than April 18, 2016.A second letter dated July 11, 2016, was issued asking for a response by August 2, 2016, even though all upgrades were made because the effectiveness checks were not yet completed. The letter and response form contained the same information as the first letter.The recalling firm reported emails were tracked for delivery confirmation and the letters were tracked for delivery by USPS certified services. For further questions, please call 877-744-5369. |  | 8 | 8 sites have the potentially affected version | US Distribution was made to medical facilities in AL, IN, MI, PA and WI. | Completed |  |  | 8 | Computer
Z-1178-2017 | 76323 | Yes | LWS |  |  | Cardiovascular | PMA | Boston Scientific, EMBLEM S-ICD Programmer Model 3200 Programmer. |  |  | January 12, 2017 | 2017 | Boston Scientific Corporation | There is a potential for radio frequency (RF) interference to alter wireless communication from a programmer, which in rare instances may cause an S-ICD to perform an unintended command. This behavior can only occur during an active, in-clinic interrogation/programming session with the Model 3200 S-ICD programmer. There is no risk of this behavior occurring when the LATITUDE Patient Management System communicates with an S-ICD in an ambulatory setting. | Software design | Boston Scientific sent an "Important Medical Device Information" letter dated on January 12, 2017. The letter was addressed to "Dear Doctor". The letter included information on the Description and Clinical Implications, Rate of Occurrence, Recommendations, Affected Programmer &amp; Additional Information. For additional questions regarding this communication or to report clinical events, contact your Boston Scientific representative or Technical Services at 1 800-227-3422. |  | 4500 | 4,500 | Worldwide Distribution - USA ( nationwide) including PR and to the states of : AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, WY., and to the countries of :Canada, Austria, Belgium, Czech Republic, Denmark, Finland, France, French Polynesia, Guadeloupe, Monaco, New Caledonia, Reunion, Germany, Greece, Hungary, Ireland, Israel, Italy, Latvia, Lithuania, Luxembourg, Netherlands, Norway, Poland, Portugal, Slovakia, Spain, Sweden, Switzerland, United Kingdom, Australia, China, Egypt, Hong Kong, India, Iran, Japan, Jordan, Kuwait, Lebanon, Malaysia, New Zealand, Qatar, Saudi Arabia, Singapore, South Africa, Thailand, Tunisia, Turkey, United Arab Emirates, Argentina, Bermuda and Brazil. | Open, Classified |  |  | 4500 | Computer
Z-1681-2017 | 76309 | No | NFJ | Picture archiving and communications system. | Radiology | Ophthalmic | 510(k) | Merge Eye Care Systems, consisting of Merge Eye Station, Merge Eye Care PACS, and Merge Eye Station Import Utility (ESIU). |  |  | April 04, 2016 | 2016 | Merge Healthcare, Inc. | The system will completely stop allowing the downloading or acquisition of new images due to incompatible software application versions. | Software Design Change | The recalling firm issued a letter dated 3/28/2016 via e-mail the week of 4/4/2016. If no email address was available, the letters were issued via certified mail. |  | 1627 | 1,627 sites potentially have the affected software | Distribution was made to medical facilities nationwide. Government distribution was made. Foreign distribution was made to Canada, as well as other countries. There was no military distribution. | Open, Classified |  |  | 1627 | Computer
Z-1495-2017 | 76442 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Orthosize Templating Version 1.2.6 |  |  | February 14, 2017 | 2017 | Zimmer Biomet, Inc. | Digital templates were created with the incorrect files. | Software Manufacturing/Software Deployment | On 2/14/2017 URGENT MEDICAL DEVICE CORRECTION notifications were sent to the affected users and 3rd party users via email. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for 3rd party users to respond to the formal recall notification. Orthosize Templating UsersIn approximately 6 weeks you will receive a notice that an update is available for the Orthosize Templating App. Upon selecting the app you will be required to update the app in order to use it.1. Review this notification and ensure affected personnel are aware its contents.2. Immediately update the app when the update becomes available.3. If after reviewing this notice you have further questions or concerns please call 411 TechnicalServices at 574-371-3071 between 8:00 am and 5:00pm EST, Monday through Friday.Alternatively, yourquestions may be sent by email to corporatequality.postmarket@zimmerbiomet.comDigital Template Users (3rd party)1. Review this notification and ensure affected personnel are aware of the contents.2. Immediately update your systems to include the revised digital templates included with notice.3. Complete Attachment 1  Certificate of Acknowledgement.a. Return a digital copy to corporatequality.postmarket@zimmerbiomet.com.b. Retain a copy of the Acknowledgement Form with your records.4. If after reviewing this notice you have further questions or concerns please call 411 TechnicalServices at 574-371-3071 between 8:00 am and 5:00pm EST, Monday through Friday. Callsreceived outside of the call center operating hours will receive a prompt to record a voicemail orbe transferred to an on-call representative in the case of an emergency. Alternatively, yourquestions may be sent by email to corporatequality.postmarket@zimmerbiomet.com |  | 3232 | 3232 | Worldwide distribution. US Nationwide including Puerto Rico, Canada, Argentina, Australia, Austria, Belgium, Chile, China, Denmark, Finland, France, Germany, Greece, India, Italy, Japan, Korea, Lithuania, Malaysia, Mexico, Netherlands, Nicaragua, Norway, Pakistan, Poland, Portugal, Romania, Singapore, South Africa, Spain, Sweden, Switzerland, Turkey, United Kingdom. | Open, Classified |  |  | 3232 | Computer
Z-1761-2017 | 76648 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Specimen Gate Laboratory; Product Number: 5002-0180 |  |  | March 07, 2017 | 2017 | PerkinElmer Life and Analytical Sciences, Wallac, OY | Potential for errors in patient results generated by the Specimen Gate Laboratory Software. | Software design | PerkinElmer Life and Analytical Sciences sent an Urgent Medical Device Correction letter dated March 7, 2017, to all affected consignees. The letter contains the following instructions: Until PerkinElmer's corrective measures are completed, we request our customers to close the Specimen Gate Laboratory Result Viewer between viewing and especially between manually calculating assay results. Closing and reopening the Result Viewer between processing the results from different assays effectively prevents the described issues from occurring. We also request our customers to follow unreported specimens with available tools to avoid any delay in reporting screening results. The available tools will vary depending on their process or the LIMS system that is used for activities associated with management of patient demographics, patient reports and the overall management of data generated from newborn screening testing. The final corrective action will involve an installation update to the software. PerkinElmer will deploy an installation update that will prevent the described issue when assays are calculated during the same Result Viewer session without closing it between the processing of different assays. Once the installation update is available, Software Services will contact you and arrange for the deployment. The installation update will be provided to you free of charge.Customers were instructed to complete the enclosed response form and return it by fax to 1-330-825-8520 (US) =358 2 2678 357 (OUS).For questions regarding this recall call 317-418-1735. |  | 29 | 29 | Worldwide Distribution - US including AL, AZ, AR, CA, CO, CT, FL, GA, IL, IN, KS, KY, LA, MD, MI, MO, NV, OH, OK, OR, TN, TX, VA; and Internationally to Canada and Brazil | Open, Classified |  |  | 29 | Computer
Z-1403-2017 | 76590 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge Cardio software. The firm name on the label is Merge Healthcare, Hartland, WI. |  |  | April 04, 2016 | 2016 | Merge Healthcare, Inc. | When taking measurements from images on the Cardio workstation or from the US cart, numbers are not crossing to the report in the correct unit of measure. | Software design | Merge sent an Urgent : Medical Device Recall letter dated March 28, 2016. If no email address was available, the letters were issued via certified mail. Nonresponders were issued a second letter dated July 11, 2016.For further questions, please call (877) 741-5369. |  | 110 | 110 sites potentially have the affected versions | Distribution was made to medical facilities nationwide. There was no foreign/military/government distribution. | Open, Classified |  |  | 110 | Computer
Z-1228-2017 | 76237 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | Vascular catheter introduction kit |  |  | January 11, 2017 | 2017 | Arrow International Inc | There have been complaints for peel away sheaths flaring | Device Design | Arrow International mailed an Urgent Medical Device Advisory Notification letter to affected customers on 01/13/2017 to inform them of the issue. Customers with questions were instructed to contact their local sales representative or Customer Service at 1-866-246-6990. |  | 215703 | 215,703 units | Nationwide Distribution | Open, Classified |  |  | 215703 | Computer
Z-1405-2017 | 76351 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | A-series lift - A625 Manual w/Load Cell &amp; Wi-Fi, Omni Part No. 324124. Prism Medical. |  |  | January 30, 2017 | 2017 | ErgoSafe Products, LLC (DBA) Prism Medical | During internal cycle test at max load (625 lbs.) 2 units had straps break prematurely. | Device Design | The recalling firm called each customer on Jan 30, 2017 to review the recall and schedule a time for the firm to come to their facility to personally swap the recalled product for a new device. The FSNs were sent on Jan 31, 2017. |  | 21 | 21 | Product was distributed to MT and NV; one VA consignee. | Open, Classified |  |  | 194 | Computer
Z-1245-2017 | 76173 | No | PQQ | Instrumentation for clinical multiplex test systems. | Clinical Chemistry | Microbiology | 510(K) Exempt | Keyspan High-High Speed USB to Serial Adapter |  |  | December 22, 2016 | 2016 | Tosoh Smd Inc | Power outages causes reporting software to shutdown. | Software design | Tosoh Bioscience sent a Customer Recall Notification letter dated December 21, 2016 to customers. The letter identified the affected product, problem and action to be taken. For questions call Tosoh Technical Support at 800-248-6764. |  | 34 | 34 | US Nationwide Distribution in the states of CT, MA, ME, MN, MO, NC, NE, NH, PA, SD, TX, VA, &amp; WV. | Open, Classified |  |  | 34 | Computer
Z-1929-2017 | 76488 | No | HET | Gynecologic laparoscope and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Olympus ENDOEYE HD II Video Telescopes |  |  | February 10, 2017 | 2017 | Olympus Corporation of the Americas | A damaged temperature sensor at the endoscope's tip can cause the endoscope's distal end to become abnormally hot. | Device Design | A letter to consignees was sent February 10, 2017. Olympus will be repairing the device at user sites by disabling the temperature sensor. |  | 863 | 863 units | Nationwide | Open, Classified |  |  | 863 | Computer
Z-1384-2017 | 76374 | No | KQO | Automated urinalysis system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | iChemVELOCITY Urine Chemistry System, Catalog No. 800-3530, 800-7167, 800-7719, 800-7720, 800-3564, 800-7100, 800-7103, 800-7106, 800-7162, 800-7163, 800-7166, 800-3061. |  |  | January 30, 2017 | 2017 | Beckman Coulter Inc. | The probe misalignment or bending has the potential to lead to delay in reporting results or incorrect patient results. | Device Design | Beckman Coulter/Iris sent an Urgent Medical Device Recall letter dated January 30, 2017 to all affected customers to inform them that Iris International has determined that sample probe misalignment or bending on the iChemVELOCITY may lead to a remote possibility of false negative results due to inadequate strip dosing. Customers were informed that Inadequate dosing may lead to incorrect patient results which may not be detected by quality control, or to a delay in reporting results. Incorrect patient results might be observed as an unexpected discrepancy between instrument results and the patient's clinical picture. The greatest impact could occur when proteinuria is not detected. Customers were instructed to complete and return the enclosed Response Form within 10 days. Customers with questions were instructed to contact their local support representative:Âż Via our website, at http://www.beckmancoulter.comÂż By phone, call 800 854-3633 in the United States and Canada |  | 3175 | 3175 units total (1697 units in US) | Worldwide Distribution - US (nationwide) and internationally to Argentina, Australia, Austria, Bangladesh, Belarus, Bolivia, Bosnia and Herzegovina, Brazil, Brunei, Canada, Chile, China, Columbia, Croatia, Czech Republic, Ecuador, France, Germany, Greece, Hong Kong, Hungary, India, Indonesia, Israel, Italy, Japan, Jordan, Kuwait, Lebanon, Macau, Malaysia, Mexico, Moldova, Montenegro, Netherlands, New Zealand, Nigeria, Oman, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Puerto Rico, Qatar, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, South Korea, Spain, Switzerland, Taiwan, Thailand, Turkey, United Arab Republic and Venezuela. | Terminated | August 19, 2017 | 201 | 3175 | Computer
Z-1385-2017 | 76471 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | DR Systems Unity PACS software, now known as Merge Unity PACS. |  |  | April 08, 2016 | 2016 | Merge Healthcare, Inc. | The software is not adding the correct tomo image. | Software design | Merge sent an Urgent Medical Device Recall dated July 11, 2016 to affected customers. The letter identified the affected product, problem and actions to be taken. For questions send an email to recall@merge.com |  | 28 | 28 sites potentially have the affected software | US Nationwide Distribution in the states of AL CA, CO, FL, IL, MD, MN, MT, NV, NY, OR, PA SC, TN, and TX. | Terminated | March 06, 2017 | 332 | 28 | Computer
Z-1386-2017 | 76052 | No | DSP | Intra-aortic balloon and control system. | Cardiovascular | Cardiovascular | 510(k) | CARDIOSAVE Hybrid IABP and CARDIOSAVE Rescue IABP affected part number is 0998-:XX-0800-XX |  |  | December 23, 2016 | 2016 | Maquet Datascope Corp - Cardiac Assist Division | Maquet has received 3 confirmed complaints related to the Li-lon Battery used with the Maquet CARDIOSAVE Hybrid IABP and Maquet CARDIOSAVE Rescue IABP of the battery venting after being accidentally dropped. | Device Design | Maquet contacted their customers via letter on 12/23/16 and additional customers identified were notified on 1/10/17 and issued a recall letter/return response form. The letter identified the affected product problem and actions to be taken. Customers were instructed to complete and return the response form via fax to 1-973-629-1518 or email to cardiosavebattery@getinge.com |  | 8000 | 8000 | Worldwide Distribution - US Nationwide | Open, Classified |  |  | 8000 | Computer
Z-1692-2017 | 76540 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | UOC Femoral Driver, Product number PE 93045101, orthopedic manual surgical instrument/accessory for use during surgery to implant femoral components and femoral trial, and to remove femoral trial. |  |  | February 10, 2017 | 2017 | United Orthopedic Corporation | The locking lever pin component of the femoral driver may become deformed and break, resulting in malfunction during surgery and causing the device to become inoperable. | Device Design | United Orthopedic Corporation sent an Urgent Field Safety Notice dated January 19, 2017, to all affected consignees via e-mail on February 10. 2017. The Urgent Field Safety Notice instructed consignees to recover any affected units and return them to the manufacturer. A redesigned device will be sent for replacements. Consignees with questions were instructed to call 949-328-3366. |  | 33 | 33 units | Worldwide Distribution - US including State of CA and Internationally to Pakistan, Spain, and Taiwan. | Terminated | June 28, 2017 | 138 | 33 | Computer
Z-1396-2017 | 76500 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | McKesson Radiology 12.2 - Picture Archive Communication System (PACS) |  |  | December 02, 2016 | 2016 | Mckesson Medical Imaging | Issue for customers that use an EMR login or legacy web URL login or legacy web URL login for McKesson Radiology PACS that may result in missing images in a newly imported study, and/or study imports that remain in an "in-progress" status. | Software design | McKesson sent an Urgent Field Safety Notice on December 2, 2016 to customers with potentially affected software versions. The notice indicated that they were to contact McKesson to install a software update to correct the issue, and that they read and understood the Field Safety Notice. |  | 17 | 17 | AK, CA, FL, HI, KY, MD, MI, MS, NH, OH, PA, TN, TX, VT | Open, Classified |  |  | 17 | Computer
Z-1395-2017 | 76503 | pylon, post surgical | ISM | Prosthetic and orthotic accessory. | Physical Medicine | Physical Medicine | 510(K) Exempt | SONNET Mini Battery Pack CableProduct Usage: |  |  | October 25, 2016 | 2016 | MED-EL Elektromedizinische Gereate, Gmbh | Unit Defect: Firm inspection revealed units did not have the compartment lock firmly welded to the housing. | Device Design | MED-EL is notifying all consignees to whom affected products have been directly shipped by Headquarters (based on first delivery shipment records) on 10/25/2016. |  | 538 | 538 units | Worldwide Distribution - US Nationwide in the states of CA, NC and TX; and the countries of Belgium, Belarus, Canada, France, Germany, Hong Kong, Israel, Netherlands, Philippines, Poland, Portugal, Turkey, United Kingdom, and Uruguay. | Terminated | April 04, 2017 | 161 | 538 | Computer
Z-1225-2017 | 75727 | No | JCM | Abnormal hemoglobin assay. | Hematology | Hematology | 510(K) Exempt | Tosoh Bioscience HLC-723G8 Analyzer for HbA1c testing. |  |  | November 07, 2016 | 2016 | Tosoh Bioscience, Inc. | The device defect associated with this recall may result in HbA1c values with a negative bias in patients with the hemoglobin variants C, D, E and S. The negative bias occurs at the higher levels of HbA1c, e.g. 9%, but not at the diagnostic medical decision point of 6.5%. | Software design | Recall notification letters were distributed to customers on 11/7/16. The letters instructed customers to review the chormatography to determine whether or not a variant peak was detected. If so, restest the sample using an alternative method. |  | 682 | 682 analyzers | Worldwide distribution. The products were shipped to the following states: AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, AND WY.The recalled products were shipped to the following countries: Brazil, Chile, Columbia, Guatemala, Costa Rica, Dominican Republic, Eduador, Nicaragua, Panama, Peru, Puerto Rico, Uruguay, | Open, Classified |  |  | 682 | Computer
Z-1257-2017 | 76252 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | The LIFEPAK 1000 defibrillator is intended for use by personnel who are authorized by a physician/medical director and are trained in CPR and the use of the LIFEPAK 1000 defibrillator. |  |  | January 13, 2017 | 2017 | Physio-Control, Inc. | The firm has received complaints that the LIFEPAK 1000 Defibrillator is unexpectedly powering off during device usage. | Device Design | Physio Control sent an Urgent Medical Device Correction letter dated January 2017 to all affected customers. The firm issued the press release on January 13, 2017. The link to press release is http://www.fda.gov/Safety/Recalls/ucm537114.htm.The firm also started mailing letter, confirmation sheet, check list to consignees on January 13, 2017. Information about this notice is available at: www.physio-control.com/lifepak1000-274. Customers were instructed to immediately remove and reinstall the battery from their LIFEPAK 1000 defibrillator. The removal and reinstallation of the battery will clean the contacts of oxidation and will restore power to the device.It is critically important that customers implement a weekly schedule of battery removal and reinstallation for all LIFEPAK 1000 devices. Removing and reinstalling the battery on a weekly basis will help ensure their device is ready for use.Customers were also advised to always carry a spare, fully charged battery. The weekly schedule of battery removal must be performed until their device correction has been completed. Customers were asked to forward this information to all of their sites, trainers and users who have LIFEPAK 1000 devices.The Customer Letter was updated on the web page to clarify the statement "It is also important to always carry a spare, fully charged battery" was provided as a reminder to customers and is consistent with what is stated in the LIFEPAK 1000 Defibrillator Operating Instructions.Customers were asked to complete and return the Confirmation Sheet to Physio-Control.Customers with questions regarding this notification, please contact Physio-Control by calling 1-866-231-1220, 6:00 a.m. to 4:00 p.m. (Pacific) Monday  Friday, or by email to rsrecalls@physio-control.com or fax to 1-866-448-9567. |  | 133330 | total 133,330 units (50,046 units in the US) | Worldwide Distribution - US (nationwide) and in the following countries: Bahrain, Belize, Cambodia, Chile, China, CÂżte d'Ivoire, Ecuador, Germany, Honduras, Hong Kong, India, Indonesia, Jamaica, Myanmar, Netherlands, Niger, Philippines, Saudi Arabia, Senegal, Singapore, Sri Lanka, Thailand, Tunisia, Ukraine, and Zambia. | Open, Classified |  |  | 133330 | Computer
Z-1533-2017 | 76639 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo.plaza, Picture Archiving and Communications System (PACS) |  |  | February 20, 2017 | 2017 | Siemens Medical Solutions USA, Inc | Software updates | Software design | Siemens mailed a notification letter dated February 21, 2017 to affected customers to inform them of the software updates that were now available to address several issues. Customers were advised to follow the instructions in the notice and where necessary promptly notify and instruct all their staff at their organization who have to be aware of the problem. Customers were also instructed to forward the information to new owners where possible. For questions regarding this recall call 610-219-4834. |  | 13 | 13 systems | Nationwide Distribution to MO,TX, FL, MA, WI, PA, IN, and CA | Terminated | August 07, 2017 | 168 | 13 | Computer
Z-1404-2017 | 76279 | No | LPL | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic | 510(k) | Color contact lenses labeled under the following brands: Diamond, Fierce, Ambition, Circle Collection, Lunatic, Fright, Desire and See Clear. |  |  | January 06, 2017 | 2017 | The See Clear Company | Sterility and Misbranding: Lenses may not be sterile and may be labeled with erroneous lot numbers and expiry dates. | Device Design | The firm, The See Clear Company, sent an "URGENT: DEVICE RECALL" letter dated 12/30/2016 to its Consignees on 01/06/2017. The letter described the product, problem and actions to be taken. The consignees were instructed to immediately examine your stock to determine if you have any of the listed product; immediately discontinue distributing the product and promptly return them via parcel post to Norcross facility, ATTENTION: RETURN GOODS; if you have further distributed any of the listed product, immediately contact your accounts, advise them of the recall, and have them return their outstanding recalled stocks to you; and immediately complete and return RECALL RETURN RESPONSE FORM via mail.If you have any questions, contact us at (678) 313-1399. |  | 119829 | 119,829 units | Worldwide Distribution: US (nationwide) including states of: AZ, CA, DE, FL, GA, HI, LA, NJ, NY, NV, ON, PA and WA; and country of: Canada. | Open, Classified |  |  | 119829 | Computer
Z-1410-2017 | 76408 | No | JDR | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | The UNI-CP System; Model Number: 330230SND. |  |  | January 25, 2017 | 2017 | Integra LifeSciences Corp. | The firm received one complaint from one non-US sales representative (France) that the label on the UNI-CP plate was incorrect. | Device Design | Integra LifeSciences Corp. issued a recall letter/return response form on 1/25/2017. |  | 39 | 39 units | Nationwide | Open, Classified |  |  | 39 | Computer
Z-1412-2017 | 76456 | No | HWA | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Turon Impaction Fixture |  |  | February 09, 2017 | 2017 | Encore Medical, Lp | During the Turon assembly, the impaction forces caused the polymer, black acetal copolymer from the Impaction Fixture to wear off on the lateral surface of the Humeral Stem titanium plasma spray coating. | Device Design | There are two field safety notices, one for Consignees who have surgeons that use the impaction fixtures (Verson 1) and one for Consignees who had previously indicated their surgeons do NOT use the impaction fixtures (Verson 2). The two recall notification letters were sent out on 2/9/17. |  | 297 | 297 units | US, South Korea, Australia, Canada, United Kingdom/Ireland, Germany | Open, Classified |  |  | 923 | Computer
Z-1680-2017 | 76036 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | JResultNet drivers using ASTM coding language. |  |  | October 23, 2014 | 2014 | Data Innovations, LLC | Software issue resulting in the results of one patient potentially being associated with another. | Software design | Data Innovations, LLC contacted consignees by phone, email, and post. Recall notifications were sent out on November 3, 2014. Consignees were provided instructions on how to reconfigure the drive to work around the issue. If consignees were unable to perform the work around, Customer Support is able to service the units. In order to request this service, consignees were to call 239-931-6004, 802-658-1955, or 800-332-0499. Attached to the recall notification was a response form that is to be faxed to 802-658-2782. Consignees with and without affected drives are to complete this response form. For questions regarding this recall call 239-931-6004. |  | 657 | 657 units | Nationwide Distribution to AK, AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, WY, &amp; DC. | Completed |  |  | 657 | Computer
Z-1415-2017 | 76401 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge Cardio software using EchoIMS. The firm name on the label is Merge Healthcare, Hartland, WI. |  |  | April 04, 2016 | 2016 | Merge Healthcare, Inc. | A situation can occur allowing two physicians to access the same study report in EchoIMS when launched from the Cardio Study List without receiving the read-only notification prompt. | Software design | Merge sent an Urgent: Medical Device Recall letter dated March 28, 2016, via e-mail on April 4, 2016. A letter was issued to nonresponders on July 11, 2016.The letter notifies the customer of the issue, informs them of the workaround, and that no further action was being taken to correct the issue, therefore, the customer needs to be aware of the possible behavior. The notification instructs the customer to pass the information onto all other users, as well as any downstream customers if it was further distributed. The notification also informs the customer their response is required using the enclosed form and the return addressed envelope. The response form asks, in part, if they understand the workaround. The response was required no later than April 18, 2016. Nonresponders were issued a second letter dated July 11, 2016, asking for a response by August 2, 2016. The letter and response form contained the same information as the first letter. If the second letter results in additional unresponsive customers, a call campaign will be managed by Customer Support and documented via SalesForce. Three call attempts will made before closing the case. For further information please call (877) 741-5369. |  | 17 | 17 sites potentially have the affected versions | US Distribution was made to medical facilities in CA, FL, IL, MD, MO, OH, OK, TX, VT, and WI. Military distribution was also made. | Terminated | March 20, 2017 | 350 | 17 | Computer
Z-1431-2017 | 76594 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge PACS software. The firm name on the label is Merge Healthcare. |  |  | April 04, 2016 | 2016 | Merge Healthcare, Inc. | Studies that have QC, PDE, or HL7 ADT updates applied may fail to propagate via telmed protocol to the destination server and may fail to import on the destination server. | Software design | Merge sent an Urgent: Medical Device Recall letter dated March 28, 2016, to all affected customers via email the week of April 4, 2016. If no email address was available, the letters were issued via certified mail. A letter dated February 21, 2017 , was issued was also issued via email or certified mail.The letter notifies the customer of the issue, informs them not receiving updated demographics or image content may result in delay in diagnosis or treatment of the patient or misdiagnosis, and they should be aware of the workaround. The customer is also notified a fix is not yet available, however, the recalling firm is working on correcting the issue and will inform them once a fix is made available. The notification instructs the customer to pass the information onto all other users, as well as any downstream customers if it was further distributed. The notification also informs the customer their response is required using the enclosed form and the return addressed envelope. The response form asks, in part, if they understand the workaround and are they interested in accepting the fix. If they decline the fix, they are asked to provide the reason why. The response was required no later than April 18, 2016.A second letter was issued dated February 21, 2017,via email or certified mail if no email address was available. The letter notified customers a fix was now available and also contained the same information and response form as the first letter. The letter informed the customer a response was required 15 calendar days after receipt of the letter.The recalling firm reported emails will be tracked for delivery confirmation and the letters will be tracked for delivery by USPS certified services.For further questions, please call (877) 741-5369. |  | 248 | 248 sites potentially have the affected versions | USA (nationwide ) Distribution was made to medical facilities. Internationally to Canada. There was no government or military distribution. | Open, Classified |  |  | 248 | Computer
Z-1212-2017 | 76454 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARTISTE MV System, medical charged-particle radiation therapy system |  |  | January 25, 2017 | 2017 | Siemens Medical Solutions USA, Inc | Siemens became aware of Incorrect values for the rotational collimator position displayed due to mechanical problems of the sensor slide. In addition, there are improper weld seam at the overhead suspension. | Device Design | Siemens mailed a Customer Safety Advisory Notice (CSAN) dated February 7, 2017 to affected customers. The letter identified the affected product, problem and actions to be taken. The letter stated that Siemens will contact customers to arrange a date for the installation of the software update. |  | 194 | 194 systems | United States distribution | Open, Classified |  |  | 194 | Computer
Z-1426-2017 | 76541 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | DR Systems Unity PACS software, now known as Merge Unity PACS |  |  | April 08, 2016 | 2016 | Merge Healthcare, Inc. | Reports were missing demographic header information when they were faxed to referring physicians. | Software design | Merge sent an Urgent : Medical Device Recall letter dated March 28, 2016, to all affected customers via email on April 8, 2016. Nonresponders were issued a second letter dated July 11, 2016.The letter notifies the customer of the issue, makes them aware of a workaround, and informs them a fix is available. The notification instructs the customer to pass the information onto all other users, as well as any downstream customers if it was further distributed. The notification also informs the customer their response is required using the enclosed form and the return addressed envelope. The response form asks, in part, if they are interested in accepting the fix. If they decline the fix, they are asked to provide the reason why. The response was required no later than April 18, 2016.Nonresponders were issued a second letter dated July 11, 2016, asking for a response by August 8, 2016. The letter and response form contained the same information as the first letter. If the second letter results in additional unresponsive customers, a call campaign will be managed by Customer Support and documented via SalesForce. Three call attempts will made before closing the case.The recalling firm reported emails will be tracked for delivery confirmation.For further questions please call ( 877) 741-5369. |  | 9 | 9 sites potentially have the affected version | US Distribution was made to medical facilities in CA, MT, PA, and TX. There was no foreign/military/government distribution. | Terminated | March 17, 2017 | 343 | 9 | Computer
Z-1439-2017 | 76189 | No | MEA | Infusion pump. | General Hospital | General Hospital | 510(k) | CADD Solis VIP Ambulatory Infusion Pump, Model 21-21210, Reorder 21-2120-0102-15, |  |  | October 31, 2016 | 2016 | Smiths Medical ASD, Inc. | 20 Pumps sold to the Finnish market contain a message in which one word in the message is mistranslated. When the user follows a specific set of key presses the pump will display the incorrect message. The message indicates that a Patient Controlled Analgesia (PCA) dose is unavailable because the pump is running. It should indicate that the PCA dose is not available because the pump is stopped. The function of the pump is unchanged and no patient injury can occur since no drug is being delivered. | Software design | Smiths Medical sent an Urgent Medical Device Field Safety Notice dated October 26, 2016, via FedEx on October 31, 2016. The letter described the Reason for Field Corrective Action, Risk to Health, and Instructions to Customers. Advised consignees to contact the Smiths Medical Service &amp; Repair Department for the software to be loaded and to complete the Recall Confirmation Form and return it to Smith Medical via e-mail to FCA.Response@Smiths-medical.com. For further questions please call (651) 633-2556. |  | 20 | 20 | Internationally to Finland | Open, Classified |  |  | 20 | Computer
Z-1470-2017 | 76654 | No | LMB | Medical image storage device. | Radiology | Radiology | 510(K) Exempt | iConnect Enterprise Archive when used with RadSuite. The firm name on the label is Merge Healthcare. |  |  | April 04, 2016 | 2016 | Merge Healthcare, Inc. | The software produced a number of "do not route" exceptions, which may result in potential patient injury or delay in diagnosis or treatment. | Software design | Merge Healthcare sent an Urgent: Medical Device Recall letter dated March 28, 2016, via e-mail the week of April 4, 2016. If no email address was available, the letters were issued via certified mail. A second letter dated July 11, 2016, was issued to nonresponders.The letter notifies the customer of the issue and informs them a fix has been released for the issue. The notification instructs the customer to pass the information onto all other users, as well as any downstream customers if it was further distributed. The notification also informs the customer their response is required using the enclosed form and the return addressed envelope. The response form asks, in part, if they are interested in accepting the fix. If they decline the fix, they are asked to provide the reason why. The response was required no later than April 18, 2016.Nonresponders were issued a second letter dated July 11, 2016, asking for a response by August 2, 2016.The letter and response form contained the same information as the first letter. If the second letter results in additional unresponsive customers, a call campaign will be managed by Customer Support and documented via SalesForce. Three call attempts will made before closing the case.The recalling firm reported emails will be tracked for delivery confirmation and the letters will be tracked for delivery by USPS certified services.For further questions, please call (877) 741-5369. |  | 187 | 187 sites potentially have the affected versions | USA (nationwide ) Distribution to medical facilities. Government distribution was also made. There was no foreign or military distribution. | Open, Classified |  |  | 187 | Computer
Z-1699-2017 | 76633 | No | FSY | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | The Stryker VisumÂż LED Surgical Lighting system is intended to illuminate the operative site during surgical procedures with high intensity light. Affected part numbers are 0682001432 (EDS LIGHT SUSPENSION, CENTRAL AXIS, 9P/FP), 0682001286 (EDS LIGHT SUSPENSION, CENTRAL AXIS, 5P/FP), 0682001298 (EDS LIGHT SUSPENSION, CENTRAL AXIS, 5P), 0682001299 (EDS LIGHT SUSPENSION, CENTRAL AXIS, 5P/5P), 0682001300 (EDS LIGHT SUSPENSION, CENTRAL AXIS, FP/5P/5P), 0682001433 (EDS LIGHT SUSPENSION, CENTRAL AXIS, 9P), 0682001434 (EDS LIGHT SUSPENSION, CENTRAL AXIS, 9P/5P), 0682001435 (EDS LIGHT SUSPENSION, CENTRAL AXIS, FP/9P/5P) |  |  | February 28, 2017 | 2017 | Stryker Communications | The potential hazard that can arise for the combination of the missing screws and the overloading by the end user is insufficient mount force that causes the equipment to fall. | Device Design | Stryker sent an Urgent Medical Device Correction letter dated March 3, 2017 to affected customers. The letter identified the affected product, problem and actions to be taken. Distributors will notify customers via tracked Customer Notification Letter. A method of written communication will include a confirmation of delivery. For questions contact Stryker Technical Support at 800-243-5135 or comm.techservices@stryker.com |  | 292 | 292 devices (in US) | Worldwide distribution - United States Nationwide - Australia, Canada, Great Britain, Japan, Netherlands and Puerto Rico | Open, Classified |  |  | 292 | Computer
Z-1176-2017 | 76260 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge PACS software. The firm name on the label is Merge Healthcare, Hartland, WI. |  |  | April 04, 2016 | 2016 | Merge Healthcare, Inc. | Potential exists for an incorrect patient image being displayed which could result in the delay in diagnosis or treatment. | Software design | The recalling firm issued a letter dated 3/28/2016 via email on 4/4/2016 or via certified mail if they did not have an email address for the customer. A second letter was issued 7/11/2016 to nonresponders. |  | 29 | 29 sites have the potentially affected software | Distribution was made to medical facilities in AZ, CA, CO, FL, HI, IL, MA, MD, MI, MO, NY, OH, PA, SC, TN, UT, and WI. There was no foreign/government/military distribution. | Open, Classified |  |  | 29 | Computer
Z-1542-2017 | 76373 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Hitachi Echelon MRI System |  |  | April 16, 2015 | 2015 | Hitachi Medical Systems America Inc | Image data transferred from the MRI system to a workstation showed errors on the slice position reference image. | Software in the Use Environment | Formal notification letters were sent to consignees on 2/13/17. Hitachi Service will update the software at each site during visits. |  | 68 | 68 units | US nationwide distribution. The products were distributed to the following states: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, WA, WI, and WY.There are four (4) government accounts for this recall: VA Medical Center/Lyster Army Health Clinic, Building 301 - Andrews Avenue, Fort Rucker, AL 36362; Parma VA Medical Center, 8787 Brookpark Rd, Parma, OH 44129; Viera VA Outpatient Clinic, 2900 Veterans Way, Viera, FL 32940; Perry Point VA Medical Center, Avenue D, Suite 1A-120, Perry Point, MD 21902.There are zero (0) Canada accounts for this recall.There are two (2) Mexico accounts for this recall: Hospital General Tipo "B" de Alta Especialidad - SMH, "Fraccionamiento 1 De La Parcela 101, Del Ejido De Atapaneo", Morelia, MX 58300; Hospital General de Occidente, Zoquipan 1050 Avenue, Zapopan Jalisco, MX 45170The product was distributed to the following countries: Brazil. | Terminated | June 29, 2017 | 805 | 303 | Computer
Z-1769-2017 | 76518 | No | DWE | Cardiopulmonary bypass pump tubing. | Cardiovascular | Cardiovascular | 510(k) | Custom Tubing Pack with Bioline Coating |  |  | February 16, 2017 | 2017 | Datascope Corporation | A complaint was received on January 1, 2017 stating the pigtail on line 2 of the oxygenator was assembled to the port where the yellow cap and the clear cap remained on the center port. | Device Design | Maquet Inc. was notified by their customer on January 1, 2017. This was a reverse notification. The one impacted customer said they will not return the product since they can easily correct the pigtail connection before use.For further questions regarding this recall, please call (973) 709-7442. |  | 21 | 21 kits | US Distribution to the state of : CA. | Open, Classified |  |  | 21 | Computer
Z-1875-2017 | 76987 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | 18L6 HD transducer on the ACUSON S Family ultrasound systems with software versions VD10A or VD10C; |  |  | March 30, 2017 | 2017 | Siemens Medical Solutions USA, Inc. | When scanning with the 18L6 HD transducer on the ACUSON HELX" Evolution with Touch Control, the ultrasound system may display a triple image or an image with a dark band. For the triple image issue, the system repeats one-third of the aperture, but does not display the full field of view. | Software design | Siemens sent an Customer Safety Advisory Notification via certified mail on March 30, 2017. OUS consignees are notified through the implementation of the field correction with the distributors for all countries. Letters advised users of the issues with software version VD10A or VD10C and provided steps to avoid the potential risk. Once available, all systems will be updated with corrected software. For further questions, please call (650) 694-5398. |  | 2045 | 2,045 systems | Worldwide Distribution - USA (nationwide) Distribution and to the countries of : United Arab Emirates, Argentina, Armenia, Austria, Australia, Azerbaijan, Bahrain,Bangladesh, Belgium, Bolivia, Brazil, Canada, Chile, Columbia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, El Salvador, Ethiopia, Egypt, France, French Polynesia, Germany, Greece, Hong Kong, Hungary, India, Indonesia, Iraq, Ireland, Israel, Italy, Japan, Korea, Kuwait, Latvia, Lebanon, Luxembourg, Malaysia, Mexico, Myanmar, Netherlands, New Zealand, Norway, Pakistan, Paraguay, Peru, Philippines, Poland, Portugal, Romania, Russia, Serbia, Singapore, Slovakia,Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, Ukraine, Venezuela, Vietnam and United Kingdom, | Open, Classified |  |  | 2045 | Computer
Z-1099-2017 | 76147 | No | NDC | Predictive pulmonary-function value calculator. | Anesthesiology | General &amp; Plastic Surgery | 510(k) | Accu-Chek Connect Diabetes Management App |  |  | December 20, 2016 | 2016 | Roche Diabetes Care, Inc. | iOS and Android: Under certain conditions the affected app versions may disregard historical bolus data potentially leading to an incorrect bolus insulin recommendation being provided to the user.iOS only: Pairing and using multiple meters with the Accu-Chek Connect app can under rare circumstances cause the bolus advisor to fail to offer a correction bolus recommendation within the eligible time window following a blood glucose measurement (10  15 minutes).Depending on the individual metabolic situation potentially incorrect bolus advice could lead to serious health consequences such as hypoglycemia. Both software issues may also cause the amount of active insulin displayed during the bolus calculation process to be incorrect and should not be used to manually calculate a bolus. | Software design | Customers were notified via Google Play Store or the Apple App Store that an update to the Accu-Chek Connect App was available on December 20, 2016.. The description of the new version available in the app stores indicated that the new version resolved issues with the Bolus Advisor. It also includes direction to the Accu-Chek website. On December 21, 2016, information regarding the issues corrected with the new version was placed on the website. The information on the website was replaced on January 10, 2017, with the patient version of the Urgent Medical Device Correction (UMDC) which was sent on January 10, 2017, to US HCPs who have received prescription pads that would allow them to prescribe the bolus advice activation.If you have additional concerns, please contact the Accu-Chek Customer Service Center at 1- 800-858-8072. All other versions of the app as well as other Accu-Chek products using the bolus advice feature are not affected by this issue. |  | 27243 | 27243 | Worldwide Distribution - Nationwide Distribution The Accu-Chek Connect App for iOS is distributed directly to end users via the Apple App Store. The Accu-Chek Connect App for Android is distributed directly to end users via the Google Play App Store. | Open, Classified |  |  | 27243 | Computer
Z-1828-2017 | 76844 | No | NFJ | Picture archiving and communications system. | Radiology | Ophthalmic | 510(k) | Merge Eye Station and Merge Eye Care PACS. |  |  | April 04, 2016 | 2016 | Merge Healthcare, Inc. | During an antivirus program scan of the Eye Station or Eye Care PACS, the antivirus program may detect a newly captured image to be an unwanted file and delete the image permanently. | Software in the Use Environment | The recalling firm issued a letter dated 3/28/2016 via email the week of 4/4/2016 or via certified mail if they did not have an email address for the customer. A second letter dated 7/11/2016 was issued to nonresponders. |  | 1627 | 1,627 sites potentially have the affected versions | Distribution was nationwide to medical facilities. Foreign distribution was made to Canada, as well as to other countries. There was also government distribution. | Open, Classified |  |  | 1627 | Computer
Z-1090-2017 | 76078 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | cobas p 612 pre-analytical system; |  |  | December 27, 2016 | 2016 | Roche Diagnostics Corporation |  | Software design | On 12/27/2016, URGENT MEDICAL DEVICE CORRECTION notifications were sent to the affected Consignees via courier fax or email. The recall notifications included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. The cobas p 612 pre-analytical system (63x) software is being revised to correct the aspiration process. We anticipate this new version (2.0.4) to be available during Quarter 1 2017.For now, discontinue using disposable non-filter tips for Aliquotting on your cobas p 612 pre-analytical system (63x). You will be contacted by a Roche Field Engineering Specialist (FES) about modifying your system to allow for the use of disposable filter tips. Filter tips will be provided at the time of service.Actions Required" Immediately discontinue using disposable non-filter tips for Aliquotting on your cobas p 612 pre-analytical system (63x)." Complete the enclosed faxback form and fax it to number 1-877-906-8982 or email it to Indianapolis.cd_tsc@roche.com." Provide copies of this UMDC to other clinicians who may need to be aware of this software issue." If your facility has distributed the affected product to another site, please ensure this UMDC is provided to that site." File this UMDC for future reference.Please contact the Roche Support Network Customer Support Center at 1-800-428-2336, 24 hours a day/7 days a week, if you have questions about the information contained in this UMDC. |  | 5 | 5 | TX and WA | Terminated | August 04, 2017 | 220 | 5 | Computer
Z-1091-2017 | 76214 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Merge Hemo software. Merge Hemo monitors, measures, and records physiologic data from a human patient undergoing a cardiac catheterization procedure |  |  | March 28, 2016 | 2016 | Merge Healthcare, Inc. | There is a potential connection issue when powering up the Merge Hemo Record Station and the Hemo Monitor does not communicate with the Client PC. | Software design | MERGE sent an Urgent Medical Device Recall letter dated July 11, 2016, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to reply using the enclosed form and ensure all users of the product are provided with the notification. Customers with questions were instructed to send an email to recall@merge.com.. |  | 242 | 242 sites potentially have the affected versions | Nationwide Distribution | Open, Classified |  |  | 242 | Computer
Z-1830-2017 | 76804 | No | LHN | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Proteus 235, Proton Therapy System |  |  | March 21, 2017 | 2017 | Ion Beam Applications S.A. | In specific circumstances, the user is able to restart an aborted irradiation without having to select the correct partial irradiation data. | Software design | A Field Safety Notice was mailed to the customer affected by the issue on 03/21/2017. The Notice informed the user when resuming from a partial irradiation, to ensure that the correct delivered dose has been selected in the partial report dialog box before proceeding with irradiation. |  | 3 | 2 worldwide, 1 in U.S. | Distribution in US (to Louisiana), and France. | Open, Classified |  |  | 3 | Computer
Z-1836-2017 | 76384 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Valleylab Laparoscopic Handset; Material # E2750; Laproscopic Handset Non-Sterile Bulk; Material # E2750NSB |  |  | January 26, 2017 | 2017 | COVIDIEN MEDTRONIC | Customer reports of the device handset continuing to operate after release of the activation button. | Device Design | A recall was initiated on January 26, 2017. Customers were notified via Federal Express and the letter instructs customers to discontinue use of the product and return all inventory to Medtronic (Covidien), 195 McDermott Road, North Haven, CT 06473, Attention: Field Returns Department. Disposition of product within Medtronic control and customer returned devices: Will be reworked. |  | 24703 | 24,703 Units | Worldwide - US Nationwide Distribution in the states of:AR AZ CA CO FL GA IA IL KS KY LA MA MI MN MO MS NC NE NM NY OH OR PA SC SD TN TX VA WA WV | Open, Classified |  |  | 24703 | Computer
Z-1076-2017 | 75878 | Yes | MRN | Nitric oxide administration apparatus. | Anesthesiology | Anesthesiology | 510(k) | Mallinckrodt Pharmaceuticals INOmax DSIR Plus delivery system. Model 10007. Delivers INOMAX (nitric oxide for inhalation). |  |  | November 21, 2016 | 2016 | INO Therapeutics (dba Ikaria) | Mallinckrodt discovered that some devices may have been manufactured with an out-of-tolerance manometer which can result in in an inaccurate proportional valve calibration table in an affected device. | Equipment maintenance | Recall letter was sent to customers in a e-mail beginning November 21, 2016 with a had copy sent via FedEx overnight with signature receipt beginning December 5, 2016. Devices will be picked up and replaced by Mallinckrodt. Questions can be directed to 1-877-KNOW-INO (1-877-566-9466). |  | 45 | 45 | US Distribution to the states of : CA, CO, DC, FL, GA, HI, IA, IL, IN, MI, MN, NC, NJ, NM, NY, OH, PA, and TX | Open, Classified |  |  | 45 | Computer
Z-1841-2017 | 76824 | No | GFG | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Smith &amp; Nephew ACUFEX TRUNAV Retrograde Drill 7MM |  |  | March 15, 2017 | 2017 | Smith &amp; Nephew, Inc. | Potential for the cutting blade to detach from the device and/or the distal drill head to fracture during retrograde reaming | Device Design | Smith &n; Nephew sent an Urgent Medical Device Recall Notice dated March 15, 2017, to all affected customers via Federal Express. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to return product and to follow instructions on the attached Response Form. Customers with questions were instructed to email FieldActions@smith-nephew.com.For questions regarding this recall call 508-261-3600. |  |  |  | Worldwide Distribution - US (Nationwide) and Internationally to AU, CL, DB, DE, ES, FR, GB, NL, SE | Open, Classified |  |  |  | Computer
Z-1077-2017 | 75709 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Cobas b 123 POC system |  |  | September 03, 2015 | 2015 | Roche Diagnostics Operations, Inc. | under specific settings, an issue may occur during simultaneous Sensor Cartridge and Fluid Pack change on the cobas b 123 &lt;2&gt; POC system and cobas b 123 &lt;4&gt; POC system.The issue occurs when the software function [AutoQC as follow-up] is configured to run all three levels of AutoQC only after a Fluid Pack change, but not after a Sensor Cartridge change. When both are changed simultaneously, starting with the Sensor Cartridge and followed by the Fluid Pack, the analyzer carries out only the follow-up actions associated with the Sensor Cartridge change after completing the change workflow. As a result, no follow-up AutoQC is performed and the three expected AutoQC measurements for the Fluid Pack change are not carried out.Without running quality control, there is a remote possibility that system issues would not be detected and wrong results would not be excluded on all parameters: pH, PO2, PCO2, Na+, K+, Ca++, Cl-, Glu, Lac, Hct, SO2, O2Hb, COHb, MetHb, HHb, and Bili. | Software design | Roche issued an Analyzer Bulletin on September 3, 2015, to all affected Consignees. The recall notifications included a description of the reason for the recall, affected product, and consignee responsibilities; the notification did not include instructions for responding to the notification. Actions Required" Follow the workaround outlined in this Analyzer Bulletin to avoid the described issue." File this Analyzer Bulletin for future reference.QuestionsPlease contact the Roche Support Network Customer Support Center at 1-800-526-2272 if you have questions about the information contained in this Analyzer Bulletin. |  | 30 | 30 | Nationwide Distribution including CA, LA, NE, PR, SC, TX, VA, and WV | Terminated | June 05, 2017 | 641 | 30 | Computer
Z-1870-2017 | 76927 | No | FCS | Fiberoptic light ureteral catheter. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Bush DL Ureteral Illuminating Catheter Set |  |  | March 07, 2017 | 2017 | Cook Inc. | Product may overheat, melt, and burn patient. | Device Design | On 3/7/2017, URGENT: MEDICAL DEVICE RECALL notifications were sent to the affected consignees via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Action To Be Taken:1. Examine inventory immediately to determine if you have affected products and quarantine affected products.2. Return the affected products to Cook Medical with a copy of the Acknowledgement and Receipt Form toreceive a product credit.3. Even if you do not have affected products on hand, you must still complete the Acknowledgement andReceipt Form and send it via fax (812.339.7316) or email to (fieldactionsna@cookmedical.com).4. Complete and return the Acknowledgement and Receipt Form attached to this letter within 30 days.5. Report adverse events to Cook Medical Customer Relations by phone at 800.457.4500 or 812.339.2235,Monday through Friday between 7:30 am and 5:00 pm (Eastern time), or by email viacustomerrelationsna@cookmedical.com.This notice must be shared with appropriate personnel, including down to the user level, within your organization or to any organization where the potentially affected devices have been transferred.Should you have any questions or concerns, please contact Cook MedicalCustomer Relations at 800.457.4500 or 812.339.2235. |  | 19756 | 19756 | The devices have been distributed within the United States and the following countries: Bahamas, Belgium, Brazil, Canada, Chile, China, Cyprus, Czech Republic, Denmark, Dominican Republic, Egypt, Finland, France, Germany, Greece, Guatemala, Honduras, Hong Kong, Hungary, Ireland, Israel, Italy, Jersey, Kuwait, Mexico, Morocco, New Zealand, Norway, Panama, Paraguay, Portugal, Romania, Saudi Arabia, Spain, Sweden, Switzerland, Trinidad and Tobago, United Arab Emirates and the United Kingdom. | Open, Classified |  |  | 19756 | Computer
Z-1061-2017 | 76080 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Velara Generator with an old Firmware version used with Philips Allura Xper Systems (Allura 15 &amp; 12 (monoplane), Allura FD10C, Allura FD10F, Allura FD 10 F/C, Allura FD20, Allura XPER FD10/10, Allura XPER FD20 Biplane) |  |  | December 09, 2016 | 2016 | Philips Electronics North America Corporation | System may lock up. | Device Design | Consignees were notified of the recall by letter, sent via certified mail on 12/9/16. Philips will upgrade the firmware of the Velara Generator, allowing the Velara Generator to recover in case of intermittent shortcuts. This service will be provided free of charge for all affected systems and will start effective January 2017. A Philips service representative will contact customers with affected devices to arrange for the service. |  | 59 | 59 devices (this number represents all products listed in recall) | US: AL, CA, DC, IL,GA, KS, MD, MO, NC, NY, OH, PA, RI, SC, TX, WAOUS: Brazil, Bulgaria, Costa Rica, France, Germany, Israel, Italy, Japan, Korea, Republic of, Mexico, Russian Federation, Singapore, Spain, Sweden, SwitzerlandSwitzerland, Turkey, United Kingdom | Open, Classified |  |  | 59 | Computer
Z-1832-2017 | 76807 | No | DSQ |  |  | Cardiovascular | PMA | HeartWare Ventricular Assist System (HVAD) Pump Driveline Splice Kit |  |  | March 10, 2017 | 2017 | Heartware | In April 2015, HeartWare, now a part of Medtronic, notified users of a safety issue related to the HVAD Pump Driveline Splice Kit. Medtronic now intends to send a follow-up communication to identified clinicians informing them that a new, improved splice kit is available. | Device Design | HeartWare, now a part of Medtronic sent an Urgent Medical Device Recall letter on April 2015 and a follow-up letter on March 2017, to customers via overnight mail and/or email.. The letters identified the affected product, problem and actions to be taken. Customers were instructed to review the recall notice, forward the notice to individuals within their organization who need to be aware of the notice, complete, sign, and return the Acknowledgement Form . For questions contact your local HeartWare representative. |  | 9 | 9 units | Worldwide Distribution - US Nationwide in the states of FL, MO, TN, KY, NY and the country of Germany. | Open, Classified |  |  | 9 | Computer
Z-1819-2017 | 76911 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | da Vinci Xi EndoWrist Suction Irrigator, 8 mm instrument; |  |  | March 31, 2017 | 2017 | Intuitive Surgical, Inc. | Intuitive Surgical has become aware that in specific scenarios with system software P6 and the da Vinci XiSuction Irrigator (PN 480299-03), users can experience unexpected motion of a system arm. | Software design | Firm sent e-mails on 3-17-17 to affected sites instructing them to STOP USE of the XI Suction Irrigator and RMA. and return all unused Xi Suction/irrigation instruments for credit. A formal letters was sent to affected sites on March 31, 2017. Letters advised of the reason for recall with details on product name and product number. Letters requested that any unused devices be returned to Intuitive via the standard RMA process. All personnel who use the Da Vinci system should be made aware of the issue. Response form should be completed online or by returning copy enclosed with the letter. |  | 15 | 15 x 6 packs (90 units) | US Only - one location each in AL, CO, KS, NV, NY, and 2 in TX | Terminated | July 28, 2017 | 119 | 15 | Computer
Z-1064-2017 | 73459 | No | KXJ | Radiologic table. | Radiology | Radiology | 510(K) Exempt | Apollo: Models: 9784000131, 9784000231, 9784020131, 9784020231, 9784100231, 9784120131, 9784120231, 9784130131 |  |  | February 19, 2016 | 2016 | Villa Radiology Systems LLC | The analysis showed that the incident was caused by a chain of four faults:1.) The injured person overcame a physical barrier (the tabletop support arm. height ca. 60 em) to access a prohibited area not intended as a working position.2) The directions for use (device labeling and instructions for use) did not contain an explicit exclusion of the prohibited area.3.) The movement was activated inadvertently and continuously. and it was not released (the movement requires continuous activation).4.) Delayed activation of the emergency stop. | Device Design | Consignees were notified of the recall via letter starting on 2/19/16. The letter explains the issue and actions to be taken by the consignee and manufacturer. |  | 66 | 66 distributed in US | ALAZCACOCTDCFLIAILINLAMAMDMIMNMOMSNJNYOHPAPuerto RicoTNTXUTWI | Open, Classified |  |  | 66 | Computer
Z-1055-2017 | 76043 | No | JZQ | Radial immunodiffusion plate. | Immunology | Immunology | 510(K) Exempt | Digital RID Plate Reader and Software |  |  | October 09, 2012 | 2012 | The Binding Site Group, Ltd. | If a control ring is marked after reading, the software will not flag results that are out of the specified QC range. | Software Design Change | On 10/2012 notification letter was sent to customers requesting they take the following actions:1. If have not already installed software version Vl.0.3 then please continue to work with version V1.0.2 until the above issues are resolved.2. If have already installed software version V1.0.3, when processing results always mark the controlprior to reading the ring.3. If are uncertain about the validity of any assays you have performed using software version Vl.0.3 please review the controlresults obtained and compare them to the specified control range for the RID assay and;"If the control results are acceptable then your assay is valid and results will be unaffected."If the control results are not acceptable your assay should be considered invalid and your patient results should be reviewed and repeated as appropriate.On 02/2013 a second notification was sent in which customers were informed the new software was released.Customers were asked to take the following actions;1. Upon receipt install the new software and uninstall the previou version2. Install following the onscreen instructions. The installation process has not changed from previous versions. The instruction guide received with the RID Plate reader is still the current version.3. After receipt of the new software disc complete and return the Eback form to The Binding Site.For further assistance contact Technical Services at The Binding Site via the email address: technicalservices@bindingsite.co.uk. |  |  |  | U.S. distribution to the following; MA, GA, NY, IA, CA, IL, UT, NJ, NC, MD, MN.Foreign distribution to the following; UK, Germany, Spain, France. | Terminated | April 04, 2017 | 1638 |  | Computer
Z-1039-2017 | 76113 | No | MIH |  |  | Cardiovascular | PMA | AFX Endovascular AAA System, Endoleak Type: IIIB |  |  | December 27, 2016 | 2016 | Endologix | Endologix updated the Instructions for Use (IFU) and implemented modifications to the AFX EndovascularAAA System (AFX System). | Device Design | Endologix sent a Dear Phycisicans Letter dated June 2, 2015, to all affected customers to inform them of the enclosed 2014 Clinical Update on the Endologix Endovascular Systems for AAA Repair. The letter informs the customers of the update reports the cumulative clinical data and information on the Powerlink System and the AFX Endovascular AAA System (AFX System).On January 2016, a Dear Physician letter was sent to inform the customers of the 2015 Clinical Update covering the Endologix Endovascular Systemsfor AAA Repair. Customers were informed the update reports the cumulative clinical data and commercial experience (i.e., post-market surveillance) of the Powerlink System and the AFX Endovascular AAA system and contains important information and recommendations regarding the continued safeand effective use of these devices.On 12/27/16, a Dear Physician letter was sent to inform customers of the important information related to the AFX Endovascular AAA System (AFX System). The letter informs the customers of a series of updates to the Instructions for Use (IFU) and modifications to the product were implemented,including introduction of a graft material processing improvement known as Duraply. The letter informs the customers that at the time of the submission of these modifications to FDA, they were not identified as measures intended to address Type III endoleaks. The on-going Endologix investigation has determined that these changes may help prevent the occurrence of Type III endoleaks reported with the AFX device. Customers with any questions areinstructed to contact their Endologix representative. |  | 4143 | 4,143 devices | Worldwide Distribution - US (nationwide) and Internationally to Austria, Belgium, Czech Republic, France, Germany, Greece, Ireland, Israel, Italy, Latvia, Luxembourg, Monaco, Netherlands, Poland, Portugal, Spain, Sweden, Switzerland, Turkey, and United Kingdom. | Terminated | April 28, 2017 | 122 | 4143 | Computer
Z-1046-2017 | 75957 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge Cardio software. |  |  | March 28, 2016 | 2016 | Merge Healthcare, Inc. | If a reader selects Multi-study review prior to the original study completing its loading process and then immediately returns to image review of the original study, it is possible that not all current images will be present for review, which could result in incorrect treatment/diagnosis. | Software design | The recalling firm issued a letter dated 3/28/2016 via email or via certified mail if they did not have an email address for the customer. A second letter dated 7/11/2016 was issued to nonresponders. |  | 188 | 188 sites potentially have the effected versions | US Nationwide Distribution | Open, Classified |  |  | 188 | Computer
Z-1052-2017 | 76213 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Carestream Touch Prime, Catalog # 1738830, and Touch Prime XE, Catalog # 1738822, Ultrasound System |  |  | November 21, 2016 | 2016 | Carestream Health Inc | Software error; Carestream Health Inc, received a complaint stating that when a user accidentally obtains a measurement value of 0 and corrects the value in the report, the resulting measurement unit is not displayed, i.e., centimeters or millimeters. As such, the user expects that the measurement is taken calculated in centimeters, consistent with other values in the report. In actuality, the measurement is taken in millimeters. When this updated measurement is used in an average calculation, the result appears incorrect as two measurements are interpreted as centimeters while the user corrected value is interpreted as millimeters. If the user selects a Calc Result display as Min or Max, the values are also interpreted as millimeters when centimeters were expected. | Software design | Carestream sent an URGENT: MEDICAL DEVICE CORRECTION Letter dated November 17, 2016, with a Consignee Notification Acknowledgement (HRA # MA-2016-019) to all affected US consignees via UPS on November 21, 2016. The letter identified the product, the problem, and the action to be taken by the consigneeAction to be Taken:When reviewing the worksheet, if a zero value is observed and that calculation field needs to be changed, the calculation needs to be deleted from the worksheet. Once deleted, the user needs to acquire a new calculation from a clinical image.Carestream is currently validating a fix for this problem and your local service representative will be contacting you within the next few weeks to arrange to upgrade your software.If you have any questions or concerns, please contact the Carestream Customer Care Center in the U.S. at 1-800-328-2910, available 7 days per week on a 24 hour basis. Outside of the U.S., please call your local Service support number.If you have distributed the device outside your facility, please alert your customer(s) of this field correction and contact Carestream Customer Care Center as listed above.. |  | 30 | US: 14 units; Foreign: 16 units | Worldwide Distribution - US including SC, GA, TX, OH, IA, NY and foreign distribution to Switzerland and Malaysia | Open, Classified |  |  | 30 | Computer
Z-1881-2017 | 75614 | No | OLO | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | Navitrack System - OS Knee Universal, Orthopedic Stereotaxic Instrument |  |  | October 06, 2011 | 2011 | Orthosoft, Inc. dba Zimmer CAS | Zimmer CAS voluntarily conducted a retrospective recall of the Navitrack System - OS Knee Universal software ORTHOsoft-UniTkr-2.3.2.6, due to a calibration sequence crash. | Software design | Zimmer sent a Medical Device Field Notification letter dated December 6, 2011, to all affected customers. This is a retrospective report of a correction taken on 6 December 2011. The software issues were fixed in the new software version 2.3.3.0. Recall notices were distributed on December 6, 2011. Telephone: 18663367846 (US &amp; Canada) or 5143958883 (International),Email: cassupport@zimmercas.com Customers with questions should call 1-866-336-7846 (US &amp; Canada) 514-395-8883 (International).For questions regarding this recall call 574-372-4487. |  | 105 | 105 | Worldwide Distribution - US including AZ CA CO FL GA IL IN KS MD MI MN MS OH OR PA SC SD TX VA WA and WIInternationally to Canada, Australia Austria Colombia Denmark France Germany Korea New Zealand Russia South Africa and Thailand | Terminated | May 08, 2017 | 2041 | 105 | Computer
Z-1538-2017 | 76600 | No | DSQ |  |  | Cardiovascular | PMA | HeartWare HVAD System Controller |  |  | February 03, 2017 | 2017 | Heartware | Product recall of all HVAD Controllers currently on the market and to exchange them for the next generation HVAD Controller (Controller2.0) | Software design | The firm will release a communication to their consignees stating the following: As a reminder, as with all HVAD Controllers, continue to reinforce the following with your patients and staff at all opportunities:" Patients should continue to have a backup HVAD Controller ready at all times in the event of a primary HVAD Controller failure." Staff only: The driveline extension cable is to be used during the pre-implant test only. It is not intended to be used after the pump is implanted in the patient.Hospital and Clinician Actions (to be executed in the following order):1) Review the enclosed notice and forms, and forward the notice to those individuals within your organization who need to be aware of its contents.2) Complete, sign, and return the Acknowledgement Form to HeartWare within thirty (30) days of receipt of this letter.4) Quarantine and replace affected HVAD Controllers, DC Adapters, Instructions for Use, Emergency Responder Guides and Patient Manuals in hospital inventory after training is complete.For every patient, quarantine and replace the following under clinical supervision in an environment where appropriate support equipment is readily available:- Primary and Backup HVAD Controller;- Affected DC Adapters; and- Patient Manual and Emergency Responder Guide.Clinicians are reminded not to perform an HVAD Controller exchange during an active electrical fault alarm as the HVAD Pump will be running a single stator. If an electrical fault is present, download patient log files and contact your HeartWare representative to resolve the electrical fault before executing the controller exchange.5) Return all quarantined HVAD Controllers and DC Adapters to HeartWare. Your HeartWare representative will assist you with this process.6) Completion Form. Once affected product in inventory has been identified and returned, complete and return the attached Completion Form to CON2.0@heartware.com or your HeartWare representative no late |  | 21488 | 8,343 US and 13,145 INTERNATIONAL | Worldwide Distribution - US including AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NJ, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, and WI and Internationally to Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, Croatia, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, Hong Kong, India, Israel, Italy, Japan, Korea, Kuwait, Lebanon, Lithuania, , Luxembourg, Malaysia, Netherlands, New Zealand, Norway, Poland, Saudi Arabia, Singapore, South Africa, Spain, Switzerland, Turkey, United Arab Emirates, UK, and Vietnam. | Open, Classified |  |  | 21488 | Computer
Z-1048-2017 | 76256 | No | MIH |  |  | Cardiovascular | PMA | AFX Endovascular AAA System |  |  | January 04, 2017 | 2017 | Endologix | Endologix updated information on the rates of Type III endoleaks and suggestions for patient surveillanceand treatment. | Device Design | Endologix sent a Dear Physician letter on 12/30/16 to inform them information related to the AFX Endovascular AAA System (AFX System). The letter provides updated information on the rates of Type III endoleaks and suggestions for patient surveillance and treatment. Customers were informed of the voluntary recall of the small remaining quantity of original AFX with Strata graft material, and the larger diameter sizes of AFX2. Customers with any questions are instructed to contact their Endologix representative. Customers with questions were instructed to call their Endologix representative.On 1/17/17, Endologix issued a press release to inform consumers that firm is resuming shipments of all sizes of AFX2 Endovascular AAA systems and appointed Laura Nagel as Vice President, Global Quality. Consumers are informed that the large diameter sizes of AFX2 were placed on temporary hold on 12/27/16 to investigate a manufacture issue. |  | 34 | 34 units | Worldwide Distribution - US (nationwide)and Internationally to Austria, Belgium, France, Germany, Greece, Italy, Ireland, Monaco, Poland Portugal, Romania, Spain, Sweden, Turkey, Israel, Argentina, Florida, Mexico, Chile, Brazil, Japan, Singapore, Malaysia, and Thailand. | Terminated | April 28, 2017 | 114 | 866 | Computer
Z-0970-2017 | 75875 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Halo One 5F Introducer Sheath 10 cm 5 pack |  |  | December 02, 2016 | 2016 | BARD PERIPHERAL VASCULAR, INC | Complaints of sheath separation, kinking, and/or tip damage during use. | Device Design | Bard sent an Urgent Medical Device Recall Notification letter dated December 2, 2016, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to find attached instructions detailing the steps they need to take regarding the product. Customers with questions were instructed to contact their Bard representative. Customers were provided the following instructions:1. Records show that your facility has purchased product codes affected by this voluntary recall. Do not use or further distribute any affected product.2. Check all inventory locations within your institution for affected product code / lot number combination listed in the recall notice. If have further distributed any of the product code / lot numbers, immediately contact that location, advise them of the recall, forward instructions and have them return the affected product to BPV.3. Remove any identified product from shelves.4. If have used the affected product, complete and return the attached Recall and Effectiveness Check Form indicating no product will be returned. Once the product affected by this recall has been removed from inventory:5. Fill out the Recall and Effectiveness Check Form. Be sure to state the quantities and lot numbers of each recalled product that intend to return. It is extremely important that this information is received.6. Please call the firm's Recall Coordinator at 1-800-321-4254 Option #2 Ext 2501 (M-F 6am to 3pm MST). Once all information has been verified, the Recall Coordinator will issue you either a Return Authorization (XC) Number or Consignment Recall Number (XH) to facilitate the expedient return of the product. BPV will issue a credit for returned product. Email the completed Recall and Effectiveness Check Form to HaloOneRecall@crbard.com or fax it to BPV at 1-800-994-6772. If cannot email or FAX the form, please call the BPV recall coordinator at 1-8 |  | 101 | 101 | Nationwide Distribution to NY, MO, NY, NY, FL, KS, MI, TX, NV, OH, ME, WA, LA, AZ, UT, NH.No foreign distribution. | Terminated | April 05, 2017 | 124 | 101 | Computer
Z-1890-2017 | 76436 | [6 empty] | MiniMed 620G and MiniMed 640G Insulin Pumps with Software Version 2.6 and below are affected. Catalog Numbers for these insulin pumps are MMT-1510, MMT-1710, MMT-1711, MMT-1712. |  |  | January 31, 2017 | 2017 | Medtronic Inc. | Medtronic has identified a software anomaly that can prevent the internal battery of the pump from charging. When the software anomaly occurs it leads to battery depletion and pump shutdown. | Software design | A urgent medical device recall letter was sent to customers on 1/31/17 to inform them that Medtronic has identified a software issue that could prevent the internal battery of the pump from charging. Customers are informed that should the software issue occur, an alarm is triggered and they will see the following message displayed on the pump screen: "Power error detected 00:00 25 Delivery stopped. Record your settings by uploading to CareLink or write your settings on paper. See user guide. OK." Customers are informed that If they experience the alarm and see the Power error detected message on their pump accompanied by the number 25 , then they are instructed to contact the Medtronic HelpLine team at XXXXX for assistance with troubleshooting the error. |  | 42215 | 42,215 units | Europe, Australia, Japan, Singapore | Terminated | May 04, 2017 | 93 | 42215 | Computer
Z-1897-2017 | 76294 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | BD FlowSmart Set MiniMed Pro-set Rx Only |  |  | December 23, 2016 | 2016 | Becton Dickinson &amp; Company | Patients are reporting high blood glucose (HBG) levels while using the MiniMed Pro-set and also noticed the catheter of the infusion set was bent/kinked. | Device Design | Becton Dickinson (BD) notified their sole customer Medtronic via e-mail on December 26, 2016. BD recommends that Medtronic notify their customers of the situation. For further questions, please call (201) 847-6800. |  | 57750 | 57,750 units | US Distribution to the state of : CA | Open, Classified |  |  | 57750 | Computer
Z-1098-2017 | 76046 | No | NSX |  |  | General Hospital | Contact ODE | Softbank software |  |  | October 06, 2016 | 2016 | Soft Computer Consultants, Inc. | Software error. Potential for incorrect results | Software design (manufacturing process) | Consignees were sent a recall notification, dated 10/06/2016, providing instructions for software correction. The firm indicates that the issue will be resolved with a mandatory hot fix. Hot fix will be released when its effectiveness is confirmed. Consignees are instructed to acknowledge receipt of the recall notification and grant permission for the recalling firm to load the hot fix when available. The firm can be contacted by calling Technical Support at 1-800-763-8522. Until the hot fix is available, consignees can still operate the software by exiting (Esc-Quit) the Test Verify option after each save changes (F12-Accept). |  | 189 | 189 total | Worldwide Distribution - US Nationwide Distribution in the states of AL, AZ, CA, CO, CT, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NH, NJ, NV, NY, OH, OK, OR, PA, RI, TN, TX, VA, WA, WI, WV, and the countries of Hawaii and Puerto Rico Canada &amp; Jamaica | Open, Classified |  |  | 189 | Computer
Z-1815-2017 | 76342 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Radiation Therapy Treatment Planning System, Model 4.5, 4.7, 5.0 |  |  | January 18, 2017 | 2017 | RAYSEARCH LABORATORIES AB | An error in NVIDIA GPU (Graphics Processing Unit) card drivers can occur for certain software programs. In a non-standard Citrix environment where the GPU settings are configured with a virtual Citrix graphics board (graphics card) an erroneous calculation may be allowed. | Device Design | RaySearch Laboratories sent a Field Safety Notice dated January 18, 2017, to their customers via email. ACTIONS TO BE TAKEN BY THE USER include: 1) Verify the GPU settings on each of the Citrix XenApp servers according to the instructions in this section. A correct configuration of the RayStation GPU settings will show a graphics board that has been verified by RaySearch. The graphics boards that are verified for operation are listed in the section "Graphics card (GPU) for accelerated computations" in the System Environment Guidelines. 2) Check that displayed graphics board in the GPU Settings dialog is listed as verified for operation according to the section "Graphics card (GPU) for accelerated computations" in the SystemEnvironment Guidelines. The following figure shows an example of a correct installation. If the graphics board is not verified for operation according to the System Environment Guidelines, the installation is incorrect. In this case, disable GPU computation and contact RaySearch Support. 3) Check if "Citrix Systems Inc. Display Driver" is displayed instead of the graphics board name. Ifso, the Citrix XenApp installation is incorrect. If an incorrect installation of Citrix XenApp is detected: a. Set the Citrix XenApp server to maintenance mode or disable GPU computation on the erroneous server; b. Contact RaySearch Support for instructions on how to resolve the issue. Customers are advised to pass a copy of the safety notice and educate IT staff and all users that are members of the RayStation-Administration user group (or equivalent) about this (i.e., users with the permission to change GPU settings). Inspect your product and identify all installed units with the above software version number(s), then confirm you have read and understood this notice by completing and returning the last page via email or fax to: freddie.cardel@raysearchlabs.com, fax 888 501 7195. |  | 741 | 741 units | Nationwide Distribution to AR, AZ, CA, CO, CT, DE, FL, GA, HI, IL, LA, MN, MD, MA, MI, MS, MT, NJ, NY, NC, NV, OH, OK, OR, PA, RI, TN, TX, VA, VT, VI, WA, WI, and Washington D.C. | Open, Classified |  |  | 741 | Computer
Z-1682-2017 | 76253 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Plum 360 Infusion Pump, Software Version 15.02. The infusion pump is capable of delivering fluids for a variety of therapies such as parenteral, enteral, or epidural infusions |  |  | December 30, 2016 | 2016 | Hospira Inc., A Pfizer Company | Depleted Battery alarm shows instead of Replace Battery. On battery power, ongoing therapy stops, alarms show and sound, pump shuts down after 3 minutes. On AC power, the pump reboots (takes about 25 seconds). Therapy can resume. AC power interruption of 3-7 seconds prompts an incorrect E323 alarm, ongoing therapy stop, and pump reboot. Longer power loss results in battery power operation. | Software in the Use Environment | On 12/30/2016, firm sent notification of the recall to Hospira direct consignees of record via traceable mailing. Delivery will be confirmed via traceable mailing. Consignees who further distributed the product were asked to notify their customers. Effectiveness of the communication will be confirmed by phone call with direct consignees who do not respond. The final disposition of the recalled product will be field correction (firm will contact consignees to upgrade the software in their Plum 360 infusion pumps). |  | 62847 | 62,847 units | Nationwide, Australia, New Zealand, Canada, Mexico, Chile, Philippines, Hong Kong, Greece, Ireland, Jordan, Kuwait, Oman, Poland, Saudi Arabia, Turkey, United Arab Emirates, Italy, France, United Kingdom, Spain | Open, Classified |  |  | 62847 | Computer
Z-1707-2017 | 76656 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Philips IntelliVue MX4O WLAN Patient Wearable Monitor |  |  | March 06, 2017 | 2017 | Philips Electronics North America Corporation | Philips lntelliVue MX4O WLAN Patient Wearable Monitor may not automatically switch toMonitor Mode with audible alarms when association with central monitoring system is unsuccessful [incomplete] | Software design | Philips issued Field Safety Notice on 3/6/17. The Field Safety Notice informs customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer .Philips is asking customers to follow the Action to be Taken by Customer/User section of the Field Safety Notice:When a No Data Tele INOP alert is displayed on a Philips IntelliVue Information Center iX ( PIIC iX ), check the MX40 as directed in the PIIC iX B.01 IFU, page 87. If the MX40 is in the Coverage Area, remove the battery, and reinsert the battery. If the MX40 is not in the Coverage Area, place the sector in the PIIC iX in standby as directed in the MX40 B.06 IFU, page 53, or return the patient and MX40 to the Coverage Area.Philips is releasing a software upgrade B.06.18 for affected Philips IntelliVue MX40 Patient Wearable Monitors. A Philips Healthcare representative will contact customers with affected devices to arrange for the installation of the software upgrade |  | 2648 | 2648 | Worldwide Distribution - US Nationwide including the countries of: AustraliaBelgiumCanadaDenmarkFinlandFranceGermanyGermanyIndiaItalyNetherlandsNorwayQatarSingaporeSouth AfricaSwedenSwitzerlandUnited Kingdom | Open, Classified |  |  | 2648 | Computer
Z-1417-2017 | 76545 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | HeartStart MRx Monitor/Defibrillator, |  |  | February 24, 2017 | 2017 | Philips Electronics North America Corporation | HeartStart MRx Monitor/Defibrillator may may exhibit abnormal behavior when operating using asingle battery as the sole source of power may fail to deliver therapy or monitoring | Device Design | Philips Healthcare issued an Urgent Medical Device Correction notification/Field Safety Notice. on 2/24/17. The Urgent Medical Device Correction notification/ Field Safety Notice informs customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue. Action (1) Philips is voluntarily initiating this correction and is providing Instructions for Use (IFU) Addendum titled, HeartStart MRx Battery Connector Pins for all customers with affected units. Philips is asking customers to follow the Action to be Taken by Customer/User section of the UrgentMedical Device Correction notification/Field Safety Notice:Philips is asking the customer when inspecting the device, to inspect the HeartStart MRx battery compartment connections. Include this step as part of a routine Operational Check to verify that the battery connector pins are clean, fully extended, not bent, and without residue.Action (2) Philips is voluntarily initiating this correction and is providing measures to prevent, or methods to allow the user to correct, the abnormal device behavior related to charging and pacing after disconnecting the device from AC mains. |  | 97794 | 97,794 total ( 47,362 US and 50,429 OUS) | Worldwide Distribution-US (Nationwide) including PUERTO RICO and countries of: AFGHANISTAN, ALBANIA, ANGOLA, ARGENTINA, ARUBA, AUSTRALIA,AUSTRIA,AZERBAIJAN,BAHRAIN, BANGLADESH, BELGIUM, BOLIVIA,BOSNIA AND HERZEGOWINA, BRAZIL,BRUNEI DARUSSALAM, BULGARIA,CAMBODIA,CANADA,CHILE,CHINA,COLOMBIA,COSTA RICA, CROATIA (local name: Hrvatska),CYPRUS,CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC,ECUADOR,EGYPT,EL SALVADOR, ESTONIA, ETHIOPIA,FINLAND,FRANCE,GABON,GERMANY,GHANA,GREECE, GUADELOUPE, HONG KONG,HUNGARY,ICELAND,INDIA, INDONESIA,IRAN (ISLAMIC REPUBLIC OF),IRAQ,IRELAND,ISRAEL,ITALY,JAPAN, JORDAN, KAZAKHSTAN, KENYA,KOREA, REPUBLIC OF,KUWAIT,LAOS,LATVIA, LEBANON, LIBYAN ARAB JAMAHIRIYA,LITHUANIA,LUXEMBOURG, MACAU, MACEDONIA,MALAYSIA,MALTA,MAURITIUS,MEXICO,MONGOLIA,MOROCCO, MOZAMBIQUE,MYANMAR (Burma), NAMIBIA, NEPAL, NETHERLANDS, NETHERLANDS ANTILLES,NEW CALEDONIA,NEW ZEALAND, NICARAGUA, NIGERIA, NORWAY, OMAN, PAKISTAN, PALESTINE, PANAMA,PAPUA NEW GUINEA, PARAGUAY, PERU, PHILIPPINES, POLAND, PORTUGAL,QATAR,REUNION,ROMANIA,RUSSIAN FEDERATION,SAUDI ARABIA, SERBIA, SINGAPORE, SLOVENIA,SOUTH AFRICA,SPAIN,SRI LANKA, SWEDEN,SWITZERLAND,TAIWAN, PROVINCE OF CHINA,TANZANIA, UNITED REPUBLIC OF,THAILAND,TRINIDAD AND TOBAGO, TUNISIA, TURKEY, TURKMENISTAN,UGANDA,UKRAINE,UNITED ARAB EMIRATES,UNITED KINGDOM,UZBEKISTAN,VENEZUELA,VIET NAM,YEMEN. | Open, Classified |  |  | 97794 | Computer
Z-1142-2017 | 75149 | No | HKI | Ophthalmic camera. | Ophthalmic | Ophthalmic | 510(k) | Merge Eye Station. |  |  | December 09, 2016 | 2016 | Merge Healthcare, Inc. | Merge received reports of Merge Eye Station being unable to capture images of the eye as expected per the intended use of the product. Merge is recalling product from the field to reduce the risk to patient health. | Software design | Merge sent an Urgent Medical Device Recall letter on December 9, 2016. The letter was addressed to Office Manager . The letter described the Product, Issue, Potential Harm, Containment by the Customer / User, Action by Merge and Actions by Customer. Requested customers to provide response to the "Response to this Notification is Required". If the product was further distributed advised consignees to notify them at once. For questions send an e-mail to recall@merge.com. E-mails to be sent to consignees on 12/19/2016. For further questions, please call (262) 367- 0700. |  | 182 | 182 capture stations | Worldwide Distribution - USA (nationwide) including the states of : CA, CO, CT, FL, HI, IL, IN, LA, MD, MA,M I, MN, MS. MK. NY, NC,ND, OH, OK, OR, PA, RI, TN, VA, WA and WV., and the countries of Canada and United Kingdom. | Open, Classified |  |  | 182 | Computer
Z-1688-2017 | 76700 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo.x, Picture archiving and communication system |  |  | March 06, 2017 | 2017 | Siemens Medical Solutions USA, Inc | Software changes now available to address several issues | Software design | Siemens sent a notice dated March 7, 2017, to affected customers on to inform them that software updates via the latest syngo.via software was now available to that would address several issues that were reported. Customers were asked to forward the notice to the new owner. For questions regarding this recall call 610-448-6481. |  | 104 | 104 systems | Nationwide Distribution. | Terminated | June 12, 2017 | 98 | 104 | Computer
Z-1687-2017 | 76377 | Yes | MNT | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Philips V60 Ventilator with Version 2.20 Software, |  |  | February 01, 2017 | 2017 | Respironics California Inc | The V60 Ventilator with Version 2.20 software installed may falsely detect that the blower motor has stalled. If this condition occurs, the software will cause the ventilator to shut down (Vent Inop) and display Error Code 100E. Ventilatory support will cease. | Device Design | Philips sent an Urgent- Medical Device Letter dated February 1, 2017, to all affected customers. It will be sent the week of February 6, 2017, to inform customers that Respironics is voluntarily recalling all Philips V60 Non Invasive Ventilators (NIV) that have Version 2.20 Software installed. Customers are informed that the V60 Ventilator with Version 2.20 software installed may falsely detect that the blower motor has stalled. If this condition occurs, the software will cause the ventilator to shut down (Vent Inop) and display Error Code 100E. Ventilatory support will cease. An audible high-priority alarm will sound continuously for at least 2 minutes when the V60 shuts down for any Vent Inop condition and is operating on battery power. If the V60 is connected to AC power (mains supply), the alarm will continue to sound until an operator intervenes. If the V60 is connected to a remote alarm system, the alarm system will be activated until action is taken by the operator. Error Code 100E only exists in V60 Version 2.20 software. Therefore, V60s running Version 2.10 software are not subject to this particular Vent Inop condition. E-mail back the complete form attached to the rear of this recall notification to recall. response @ Philips. com or Fax to 1-877 499-7223. For further questions please call (760) 918-7300. |  | 2512 | 2512 units | Worldwide Distribution - US Distribution and to the countries of : Argentina, Austria, Belgium, Canada, Czech Republic, Finland, France, Germany, Hong Kong, Iceland, India, Indonesia, Italy, Japan, Kuwait, Malaysia, Mexico, Netherlands, Norway, Qatar, Saudi Arabia, Spain, Sweden, Switzerland, Turkey, United Arab Emirates, United Kingdom and Vietnam. | Open, Classified |  |  | 2512 | Computer
Z-1691-2017 | 76610 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Altius M-INI OCT Posterior Spinal Fixation System, Model Number 1200-9216 |  |  | March 16, 2017 | 2017 | Zimmer Biomet Spine, Inc | Zimmer Biomet is recalling the Altius M-INI OCT Posterior Spinal Fixation System because the driver is unable to mate with the screw head tulip. | Device Design | Zimmer Biomet sent an Urgent Medical Device Recall letter on March 20, 2017, to all affected consignees. The letter identified the product, the problem and the action to be taken by the consignee. Consignee's Responsibilities1. Review this notification and ensure affected team members are aware of the contents.2. Prior to sending Altius Kits to central processing, perform a functional check of the driver withall screws in the kit to ensure they mate. If the drivers do not mate, please return inaccordance with Step 3b.3. Complete the Certification of Acknowledgement portion of Attachment 1  Inventory ReturnCertification Form  Phase I. b. If the driver is unable to mate with the screws in the set, immediately sterilize and return theaffected nonconforming product from your distributorship and affected hospitals within yourterritory along with a completed Attachment 1  Inventory Return Certification Form Phase I and Attachment 2  Certificate of Sterilization to Zimmer Biomet.i. For each return, send a copy of Attachment 1 tojeremy.hansmann@zimmerbiomet.comii. Include a hardcopy of Attachment 1 with your shipment for immediateprocessing.iii. Include a copy of Attachment 2  Certificate of Sterilization with returnedinstruments.iv. Mark the outside of the returns box(es) clearly with RECALL. 4. Distributors will notify their accounts to which the product was further distributed and to locateand return products from these accounts. Using the Additional Accounts Form provided withthe email notice sent to your facility, return contact information for any additional hospitals thathave the affected product. If there are no additional users to notify, please confirm this byannotating and returning the form provided indicating that no additional users have beenidentified.5. Retain a copy of your field action acknowledgement and product return forms for your records inthe event of a compliance audit of your facility.6. If |  | 141 | 141 | Worldwide Distribution - US (nationwide) and Internationally to the Netherlands | Open, Classified |  |  | 141 | Computer
Z-1138-2017 | 76134 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Carestream DRX-Revolution Mobile X-Ray System, Model DRXR-1, Catalog No. 1019397 (all other countries), Catalog No. 1060177 (China only) -- Carestream Health, Inc. 150 Verona Street, Rochester, NY 14608 |  |  | January 12, 2017 | 2017 | Carestream Health Inc | The firm received a complaint alleging that a DRX Revolution System could not be stopped when applying the brake. The Revolution crashed into an elevator. | Device Design | Carestream Health Inc. sent an URGENT: MEDICAL DEVICE CORRECTION Letter dated December 29, 2016, and response forms were sent to the US customers via Certified Mail-Return Receipt Requested on January 12, and January 13, 2017, informing them of the problem. Customers are instructed to read and complete all information and return the response form within 5 working days to the following e-mail address: postmarketra@carestream.com, or fax the completed form to (585) 454-4493. UMDC Letters and Response Forms to Canadian Customers were sent the week of January 9, 2017, via courier. Field Safety Notice (FSN) Letters for the rest of world are in process. FSN Letters for the EMEA should be sent via courier starting later this week.US Carestream Customer Care Center #: 1-800-328-2910 (24/7). Outside of the US, please call your local Service support number. |  | 2773 | US: 1403 units; Foreign: 1370 units | Worldwide Distribution - USA (nationwide) Distribution including Puerto Rico, and to the countries of : Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, Colombia, Denmark, Dominican Republic, Dubai, Finland, France, Germany, Ghana, India, Iran, Ireland, Israel, Italy, Japan, Korea, Kuwait, Lebanon, Libya, Malaysia, Maldives, Mexico, Morocco, Netherlands, New Zealand, Norway, Oman, Peru, Poland, Portugal, Qatar, Russia, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates and United Kingdom. | Open, Classified |  |  | 2773 | Computer
Z-1700-2017 | 76768 | No | LMB | Medical image storage device. | Radiology | Radiology | 510(K) Exempt | iConnect Enterprise Archive software. |  |  | April 04, 2016 | 2016 | Merge Healthcare, Inc. | The versions allow images to be stored without pixel data, resulting in the system not being able to present all the prior studies, which could cause a delay in treatment in determining the progression of disease. | Software design | The recalling firm issued a letter dated 3/28/2016 via e-mail the week of 4/4/2016. If no email address was available, the letters were issued via certified mail. A second letter dated 7/11/2016 was issued to nonresponders. |  | 371 | 371 sites potentially have the affected software | Distribution was nationwide to medical facilities. Foreign distribution was made to Canada, as well as other countries. Government and military distribution was also made. | Open, Classified |  |  | 371 | Computer
Z-1122-2017 | 76090 | No | LHN | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MEVION S250-Proton Radiation Therapy |  |  | December 16, 2016 | 2016 | Mevion Medical Systems, Inc. | An error can occur causing Delta corrections to be lost when one setup field is closed and another is opened | Software design | Mevion sent an Important Safety Notice to customers on 12/16/16 by email and Salesforce. The letter identified the affected product, problem and actions to be taken. Mevion provided mitigation and will release and install a software patch that removes the possibility that Deltas can be lost between multiple Setup fields. For further information contact rcessac@mevion.com. |  | 2 | 2 units | US Nationwide in the states of OK, NJ | Terminated | April 13, 2017 | 118 | 2 | Computer
Z-1697-2017 | 76851 | No | LMB | Medical image storage device. | Radiology | Radiology | 510(K) Exempt | iConnect Enterprise Archive (ICEA) software. iConnect Enterprise Archive is intended for use as a vendor neutral archive for storage and communications of medical images and data |  |  | April 04, 2016 | 2016 | Merge Healthcare, Inc. | The study is archived but cannot be opened in iConnect Access and cannot send to PACS resulting in comparison studies not being available for physician review to determine the progression of disease. | Software design | MERGE sent an Urgent Medical Device Recall letter dated July 11, 2016, to all affected customers. The recalling firm issued a letter dated 3/28/2016 via e-mail the week of 4/4/2016. If no email address was available, the letters were issued via certified mail. A second letter dated 7/11/2016 was issued to nonresponders. |  | 298 | 298 sites potentially have the affected versions (290 of those in US) | Worldwide Distribution - US (nationwide) and Internationally to Canada, Belgium, Australia, Jordan, and the UK.. | Open, Classified |  |  | 298 | Computer
Z-1706-2017 | 64902 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Mobile Diagnostic X-Ray System with Digital Panel |  |  | November 28, 2012 | 2012 | Philips Healthcare Inc. | When the operator moves an image to another view then the electronic marker might be placed at the wrong anatomical position in the resulting image. | Software design | Philips Medicals Systems distributed an "URGENT Field Safety Notice" alerting customers of the defect on November 28, 2012. Customers are requested to manually rotate images by using the rotation tool.Customers are also recommended to check whether the automatic marker is set correctly, especially after a manual image rotation is performed. The marker can be corrected manually with the annotation tool. |  | 220 | 220 systems | Nationwide (AR, AZ, CA, DE, FL, GA, IA, IL, IN, KS, MD, MI, MN, NC, NM, NY, OH, OR, PA, SC, TX, VT, WI) | Open, Classified |  |  | 220 | Computer
Z-1708-2017 | 76660 | No | DRG | Radiofrequency physiological signal transmitter and receiver. | Cardiovascular | Cardiovascular | 510(k) | eCare Coordinator |  |  | March 02, 2017 | 2017 | Philips Visicu | eCareCoordinator (eCC) is intended for use in data aggregation, patient interaction facilitation, storage and clinical information management with independent physiological devices and ancillary systems that are connected either directly or through networks. The software is intended to provide patient information from the patient location through networking technology to a remote care team.A software defect was discovered in the Philips eCareCoordinator (eCC) Clinical User Interface that can at times cause missing or redundant data to be saved without notification to the user. | Software Manufacturing/Software Deployment | Field Safety Notice (FSN) will be sent to all customers using eCareCoordinator all Versions. The FSN describes the problem and circumstances in which the design defects occurs and the action planned by Philips to correct the problem. The FSN will also detail the functionality change in detail.A software correction will be made and released.by the vendor. Customers will have this correction applied to their current version when made available by Vendor in order to fix this correction. |  | 26 | 26 | Worldwide Distribution-US Nationwide | Open, Classified |  |  | 26 | Computer
Z-1866-2017 | 76895 | No | OOI | Instrumentation for clinical multiplex test systems. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | BD Viper LT software, version 3.00H |  |  | December 14, 2015 | 2015 | Bd Diagnostic | BD has confirmed reports that lower than expected RFU values were generated on customer BD Viper LT Systems after the installation of software version 3.00H. As RFU values decrease, results approach the assay cutoff threshold which has the potential to impact test results. | Software design | Recall action was initiated on December 14, 2015 when software version 3.00H was removed from the two customer instruments. The 2 affected US customers were contacted via email and were also mailed a letter via UPS overnight express. |  | 2 | 2 | 2 US customers, 0 Ex-US regions, and 0 distributors 0. | Terminated | April 19, 2017 | 492 | 2 | Computer
Z-1778-2017 | 76859 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Merge Hemo software. The firm name on the labeling is Merge Healthcare, Hartland, WI. |  |  | April 04, 2016 | 2016 | Merge Healthcare, Inc. | Use of the software may show an incorrect value to the user when viewing the Fractional Flow Reserve (FFR) results during recording. | Software design | MERGE sent an Urgent Medical Device Recall letter dated April 4, 2016 to all affected customers via e-mail the week of April 4, 2016. If no email address was available, the letters were issued via certified mail. A second letter dated July 11, 2016 was issued to non-responders. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to ensure that all users of the affected product be provided with the notice. Customers were also asked to reply using the enclosed form and the return addressed envelope. Customers with questions were instructed to send an email to recall@merge.com.For questions regarding this recall call 877-741-5369. |  | 24 | 24 sites potentially have the affected versions | Nationwide Distribution including AZ, CA, CO, FL, GA, KS, LA, MI, MO, OH, OK, TN, TX, VA, WA, WI, and WV. | Open, Classified |  |  | 24 | Computer
Z-1864-2017 | 76900 | No | NSX |  |  | General Hospital | Contact ODE | Software version 4.44A utilized on the BD MAX System |  |  | July 29, 2015 | 2015 | BD LIFE SCIENCES | The BD MAX System software version 4.44A contains a software anomaly that has the potential to affect some Open System Reagent (OSR) customers utilizing User Defined Protocol (UDP) assays. Customers utilizing only IVD assays are not affected.This anomaly may cause the OSR customer system to incorrectly switch the columns of the truth table on the users display. If the user tries to fix their display data, the display will appear correct however, the database may become incorrect, potentially resulting in incorrect results reporting.BD is able to trace the issue to a limited number of instruments that have software version 4.44A installed. Note that the problem only affects customers that are either creating a new UDP with Result Logic or modifying the Result Logic of an existing UDP. | Software design | BD sent an Urgent Field Corrective Action Recall notification was initiated on July 29, 2015. The customer was initially contacted via a phone call and subsequently provided with a written communication. All US customers were contacted directly by phone and provided the written communication. Global holds were placed on July 21, 2015 and July 24, 2015. For further questions, please call (410) 316-4000. |  | 29 | There are 29 US customers and 5 ex-US regions with a BD MAX System running software version 4.44A. | Worldwide Distribution - US Distribution to the states of : CA, CO, FL, IL, MN, MO, MT, NE, NJ, NY, OH, OR, PA, TX, TN, WA, WI., and to the countries of : Canada, Hong Kong, Taiwan, Turkey and Europe.Federal Government Agency sales/distribution centers or foreign countries VA Fort Harrison and VAMC Austin. | Terminated | April 19, 2017 | 630 | 29 | Computer
Z-1921-2017 | 76836 | No | KIC | Flotation cushion. | Physical Medicine | Physical Medicine | 510(K) Exempt | ROHO HIGH PROFILE Single Compartment Cushion with Sensor Ready Technology and ROHO MID PROFILE Single Compartment Cushion with Sensor Ready Technology. |  |  | March 20, 2017 | 2017 | Roho Inc. | When the Smart Check tool is used to perform a Smart Setup with one of the affected cushions, it may adversely affect the setup process resulting in an improper cushion setup. Sitting on an over inflated or under inflated Sensor Ready Cushion may reduce or eliminate the benefits provided by the cushion, increasing the risk to skin or to other soft tissue. | Equipment maintenance | On 3/13/17, firm began notifying customers of recall via emailed or mailed letter. The notification sent to dealers and distributors includes a request to forward the communication to their customer and send an acknowledgement back once all of their customers have received a copy of the communication. Customers (product users) will not be required to send the product they are using back. Instead, the recall letter provided two options: 1) keep the cushion and follow additional set-up instructions provided within the letter that would mitigate the risk or 2) contact the supplier to have the cushion replaced. Product that had been in use and is replaced will be destroyed (evidence to be provided to ROHO, Inc. if destroyed by a customer). Unused product that is still in stock with the firm will be destroyed. |  | 5646 | 5646 units | Nationwide, Sweden, Germany, New Zealand, Finland, Australia, Singapore, Korea, Luxembourg, Czech Republic, South Africa, Canada, Switzerland, United Arab Emirates,Spain, Austria, Hong Kong, Iceland, Japan, Canada, Italy, Ireland, France, United Kingdom | Open, Classified |  |  | 5646 | Computer
Z-1724-2017 | 76754 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | STAPLER,IS4000; da Vinci Xi EndoWrist Stapler Release Kit (SRK); model number 470298. |  |  | March 15, 2017 | 2017 | Intuitive Surgical, Inc. | Intuitive Surgical has become aware and further investigation has determined that misinterpretation of instructions provided on the IS4000 Stapler Release Kit instruction card may result damage of the tool. | Device Design | Intuitive sent an Urgent Medical Device Correction (ISIFA2017-02-C) letter dated March 16, 2017 were sent by FED EX to be received on March 17, 2017. Letters advised users of the reason for the notification, risk to health, affected products with part numbers and a copy of Appendix A - proper use of the Stapler Release Kit (SRK). Customers were instructed to distribute a copy of the letter to all users at their facility, follow all instructions using the Instrument Release Kit (IRK) tool to finish manually unclamping the stapler and return the SRK to Intuitive Surgical through the Returned Material Authorization (RMA) process. Customers were also instructed to log into the da Vinci Online Community Field Action resource (US only) to read or complete any request actions related to this issue. In the case where the da Vinci online resource cannot be used (including outside he US), complete the attached Acknowledgement Form and return it via fax to Intuitive Surgical as instructed on the form. Customers were asked to retain a copy of the letter with their Stapler User Manual and the acknowledgement form for their files.If you need further information or support concerning this notification , please contact Intuitive Surgical Customer Service at the numbers listed below:Âż North and South America: (800) 876-1310, Option 3 (4 AM to 5 PM PST) ormail: customersupport-servicesupport@intusurg.comÂż Europe, Middle East, Asia and Africa: +800 0821 2020 or +41 21 821 2020 (8AM to 6 PM CET) or ics@intusurg.comÂż South Korea: 02-3271-3200 (9 AM to 6 PM KSTJ)Âż Japan: 0120-56-5635 or 03-5575-1362 (9 AM to 6 PM JST)For questions regarding this recall call 408-523-2100. |  | 2059 | 2059 shipped - all languages | Worldwide Distribution - US (nationwide) and Internationally to Australia, Austria, Belgium, Canada, Chile, Colombia, Czech Republic, Denmark, Finland, France, Germany, Hong Kong, Ireland, Italy, Japan, Luxembourg, Monaco, Netherlands, Norway, Portugal, Puerto Rico, Singapore, South Korea, Spain, Sweden, Switzerland, Taiwan, Turkey, United Kingdom and Venezuela | Open, Classified |  |  | 2059 | Computer
Z-1762-2017 | 76910 | No | LMB | Medical image storage device. | Radiology | Radiology | 510(K) Exempt | iConnect Enterprise Archive (ICEA) software. The firm name on the labeling is Merge Healthcare, Hartland, WI. |  |  | April 04, 2016 | 2016 | Merge Healthcare, Inc. | An error message can occur resulting in the prior studies being unavailable for viewing. | Software design | The recalling firm issued a letter dated 3/28/2016 via e-mail the week of 4/4/2016. If no email address was available, the letters were issued via certified mail. A second letter dated 7/11/2016 was issued to nonresponders. |  | 333 | 333 sites potentially have the affected versions | Distribution was made nationwide to medical facilities. Foreign distribution was made to Canada, as well as other countries. Government/military distribution was also made. | Open, Classified |  |  | 333 | Computer
Z-1118-2017 | 76273 | No | MUE | Full-field digital mammography system. | Radiology | Radiology | 510(k) | Mammomat Inspiration full, field digital,system,x-ray,mammographic |  |  | January 11, 2017 | 2017 | Siemens Medical Solutions USA, Inc | Software error | Software design | Siemens mailed a Customer Safety Advisory Notice (CSAN) letter dated January 11, 2017 to affected customers. The letter identified the affected product, the potential risks and the steps to take to avoid the risks associated with the issue. The letter also stated Siemens will release a field modification (software update) to resolve the issue as soon as possible. The new software will be provided to affected customers free of charge. |  | 55 | 55 units | United States Nationwide Distribution | Open, Classified |  |  | 55 | Computer
Z-1702-2017 | 76595 | No | JXG | Central nervous system fluid shunt and components. | Neurology | Neurology | 510(k) | The Medtronic StrataMR adjustable valves and shunts are designed to provide continuous cerebrospinal fluid (CSF) flow from the ventricles of the brain into the right atrium of the heart or the peritoneal cavity. The design enables the physician to non-invasive adjust valve pressure/performance level pre- and post-implantation by using magnetic adjustment tools without the need for radiographic confirmation. |  |  | February 22, 2017 | 2017 | Medtronic Neurosurgery | Incorrect valve rotor orientation due to misalignment and/or off-center positioning of the StrataMR Locator Tool: When the StrataMR valve is set in the incorrect orientation, higher valve opening pressures result leading to a corresponding underdrainage in CSF fluid performance. This condition may result in the following adverse health consequences: ventriculomegaly, headaches, nausea, vomiting and lethargy. If left untreated, underdrainage has the potential to lead to coma and death. | Device Design | The firm has initiated a CAPA and issued notification via letter to US Customers on February 22, 2017. As of the date of the report (2/22/17) the firm was still in process of notifying foreign customers. Recipients are instructed to: 1) Cease use of the listed product, 2) Account for number of product units implanted or discarded and return any unused units by completing an attached individualized customer product accountability form provided in the letter,3) If any of the indicated products have been implanted in patients, to adjust the valves according to instructions listed in an included attachment.Recipients are asked to fax a customer acknowledgment response form or scan/take a picture of the form and email to rs.mnsfca@medtronic.com with "StrataMR Recall" in the subject line. |  | 1796 | 2622 units (US = 1796, outside US =826) | Product was distributed throughout the US and to Australia, Belgium, Canada, Finland, Germany, Hong Kong, Israel, Italy, Japan, New Zealand, Norway, Poland, Portugal, the Russian Federation, Singapore, Spain, Sweden, Switzerland, Turkey, and the UK. | Open, Classified |  |  | 1796 | Computer
Z-1874-2017 | 76891 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Newport Medical HT70 and HT70- Plus Ventilators, Rx Only |  |  | March 30, 2017 | 2017 | Newport Medical Instruments Inc | Medtronic is issuing a voluntary field corrective action for all its Newport" HT70 and Newport" HT70 Plus ventilators because they may shutdown spontaneously during normal operation without an accompanying alarm. | Software in the Use Environment | Medtronic sent an Urgent Field Corrective Action Notice dated April 3, 2017, to all consignees via Federal Express or certified mail. The letter informed consignees of the remote potential for ventilator reset without accompanying alarm and the actions they should take. Customers were requested to acknowledge receipt and understanding of the Urgent Field Corrective Action Notice. Customers were also advised that a software service update will be available in May 2017 to address this issue. Medtronic issued a Press Release related to this voluntary field action on April 5, 2017.Consignees with questions were advised to contact the Technical Support Department immediately at 800-255-6774.For questions regarding this recall call 203-492-5415. |  | 12966 | 12,966 | Worldwide Distribution - US (Nationwide) Internationally to Canada, Australia, Brazil, Japan, Korea, Mexico, Bahrain, Egypt, GreeceHungary, Ireland, Israel, Jordan, Kazakhstan, Lithuania, MauritiusNamibia, Norway, Pakistan, Poland, Qatar, Romania, Russian FederationSouth Africa, Spain, Turkey, United Arab Emirates, United KingdomAlgeria, Armenia, Austria, Bahrain, Belarus, Belgium, BotswanaBulgaria, Cyprus, Czech Republic, Denmark, Egypt, Finland, FranceGermany, Greece, Hungary, Iran, Ireland, Israel, Italy, Kenya, LebanonLithuania, Macedonia, Malawi, Malta, Mauritius, Mozambique, NetherlandsNorway, Oman, Poland, Portugal, Qatar, Republic of Moldova, RomaniaSaudi Arabia, Serbia, Slovakia, Sudan, Sweden, Switzerland, TunisiaTurkey, Uganda, Ukraine, United Kingdom, Yemen, and Zimbabwe | Open, Classified |  |  | 12966 | Computer
Z-1113-2017 | 76156 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Cannulated Drill bit 2.0mm and 2.6 mm; Used in conjunction with the Flower Bone Screw Set. |  |  | December 22, 2016 | 2016 | Flower Orthopedics Corporation | The product is being recalled due to incidence and reports of the product breaking during surgery. | Device Design | Flower Orthopedics mailed a letter to customers on December 22, 2016 making them aware of the issue. Customers were asked to return the affected product and to report if any adverse effects resulted from its use. |  |  |  | Distributed throughout the United States | Open, Classified |  |  |  | Computer
Z-1770-2017 | 76889 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Combi Dockable Neurosurgery table for the Siemens MAGNETOM Aera (1.5T) and Skyra (3T) MRI Systems |  |  | October 17, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Due to a potential manufacturing error, a rare possibility exists for a sudden drop in table height of the Combi Dockable Table Neurosurgery. | Software design (manufacturing process) | Siemens mailed an Customer Safety Advisory Notice on October 17, 2016, to all affected customers to inform them of the corrective action they plan to take to correct the issue. The notice also informed users of the possible risks to patients and how the corrective action will be implemented. For further questions, please call (610) 448-6461. |  | 3 | 3 systems distributed in the U.S. | US Distribution to: DC and MD. | Completed |  |  | 3 | Computer
Z-1782-2017 | 64545 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Centricity PACS Web Diagnostic (WebDX) by GE Healthcare Dynamic Imaging Solutions is a device that receives medical images (including mammograms) and data from various imaging sources within the system or across computer networks at distributed locations. |  |  | February 18, 2013 | 2013 | Ge Medical Systems Information T | It was reported from a complaint that study category labels of "Prior" and "Baseline" and hanging position appeared incorrectly swapped when the hanging protocol for three comparison studies is manually applied from the patient folder. | Software design | Firm sent notification letter to all impacted customer accounts. The letter provided the following instructions for customers to follow until the permanent fix for the issue was deployed: 1) The users should be instructed to look at the date on each exam displayed rather than just the labels assigned to the exam by the system. 2) If this problem is detected in your specific workflows, a system administrator can remove the study category labels (Current/Prior/Baseline)from all overlays on the images/exams. A software code change was deployed to customer sites to correct the issue. |  | 118 | 118 | Nationwide, Worldwide | Terminated | July 31, 2017 | 1624 | 118 | Computer
Z-1780-2017 | 76805 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Vivid S60/S70/S60N/S70N, H45041SU, H45041SW, H45581MS, H45581PD. |  |  | March 10, 2017 | 2017 | GE Healthcare, LLC | GE Healthcare has become aware of an issue where a patient other than the intended is incorrectly selected by the operator in situations where DICOM Worklist search response time is slow. This issue is limited to certain Vivid Ultrasound systems and can result in the incorrect patient information showing on the screen during the exam. If not detected by the operator, images from the actual (intended) patient will be stored under the incorrect patient after the exam. This issue could lead to misdiagnosis. There have been no injuries reported because of this issue. | Software design | Consignees were sent a GE Healthcare "Urgent Medical Device Correction" letter GEHC Ref# 76164,76168 dated March 10, 2017. The letter was addressed to Hospital Administrators / Risk Manager, Biomedical Engineering, &amp; Managers of Radiology / Managers of Cardiology.The letter described the safety Issue, Safety Instructions, Affected Product Detail, Product Correction and Contact Information. For questions contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 396 | 396 (67 US; 329 OUS) | United Kingdom, United Arab Emirates, Turkey, Thailand, Switzerland, Sweden,Spain, South Africa, Slovenia, Saudi Arabia, Russia, Romania, Portugal, Poland,Philippines, Norway, New Zealand, New Caledonia, Netherlands, Myanmar, Kuwait, Korea (Republic of), Japan, Italy, ISRAEL, Ireland, Iran, India, Hungary, Hong Kong, Guatemala, Greece, Great Britain, Germany, France, Finland, Estonia, Egypt, Ecuador, Denmark, Czech Republic, Croatia, Colombia, Chile, Canada, Bulgaria, Brunei Darussalam, Brazil, Belgium, Austria, Australia, Argentina, Vietnam.Uruguay, Taiwan, Slovakia, Singapore, Reunion, Qatar, Panama, Pakistan, Jordan, Indonesia, Guadeloupe, Czech Republic | Open, Classified |  |  | 955 | Computer
Z-1486-2017 | 76431 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge Cardio software. |  |  | April 04, 2016 | 2016 | Merge Healthcare, Inc. | Cardio study list does not show STAT studies without refreshing. | Software design | The recalling firm issued a letter dated 3/28/2016 via e-mail the week of 4/4/2016. If no email address was available, the letters were issued via certified mail. Nonresponders were issued letters dated 7/11/2016. |  | 217 | 217 sites potentially have the affected versions | Distribution was nationwide to medical facilities. There was also military distribution. There was no foreign/government distribution. | Open, Classified |  |  | 217 | Computer