## Sheet1
Recall Number | Recall Event ID | Device Type | Product Code | Regulation Description | Medical Specialty | Review Panel | Submission Type | Trade Name/Product | Main Name | Recall Class | Date Posted | Year Posted | Recalling Manufacturer | Reason for Recall | FDA Determined Cause | Action | Consumer Instructions | Clean Quantity | Quantity in Commerce | Distribution | Recall Status | Termination Date | Time to Terminate | Merged Quantities | Fault_Class | 0
Z-3027-2018 | 80317 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Edwards Lifesciences Aortic Perfusion Cannula, 24F x 30CM, resterilized. |  |  | March 07, 2018 | 2018 | SPS Sterilization, Inc | The expiration date on the device labeling exceeds its actual validated shelf life because the sterilization process used does not have adequate support for the extended shelf life. | Labeling False and Misleading | Customers were notified via letter on about 03/02/2018. Instructions included to review inventory for affected devices, arrange for their return to SPS Sterilization, and complete and return the Buisness Response Card. A follow up letter was sent to customer on about 03/07/2018 and included a description of the risk and hazard involved. Instructions remained the same. |  | 10 | 10 units total | US distribution to AZ, CA, FL, OH, PA, TX, VA, WV | Terminated | November 27, 2018 | 265 | 10 | Not_Computer | Not_Computer
Z-2963-2018 | 80609 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Outer Package Label: Tapered Screw-Vent Implant System, 4.1mm X 10mm, 3.5mm, REF:TSVM4B10, Sterile:R, Rx Only; Inner Package Label: Implant, MTX, 3.7mm x 10mm,3.5mm, REF: TSVB10. |  |  | May 10, 2018 | 2018 | Zimmer Dental Inc | Dental implants have incorrect item/lot numbers and size on the inner package labeling. Correct item/lot numbers and size are on the outer package labeling. | Labeling mix-ups | The firm, Zimmer Biomet, sent an "Medical Device Recall" notice dated 5/10/2018 via FedEx/UPS courier to its customers. OUS distributors received a "distributor" notice. The notice described the product, problem and actions to be taken. The Customers were instructed to immediately return affected product to Field Action, PM Regulatory Compliance, Zimmer Biomet, 4555 Riverside Dr., Palm Beach Gardens, FL 33410 US and asked to complete and return a Certificate of Acknowledgement. and send to CorporateQuality.PostMarket@zimmerbiomet.com. If you have further questions or concerns, please call the Recall Department at 561-776-6700 between 8:00 am and 5:00 pm EST, Monday through Friday. Calls received outside of call center operating hours will receive a prompt to leave a voicemail. Alternatively, your questions may be emailed to CorporateQuality.PostMarket@zimmerbiomet.com. |  | 299 | 299 units | Worldwide Distribution: US (nationwide) in states of: AL, CA, CO, CT, FL, GA, ID, IL, MA, MD, MI, MN, MO, NH, NJ, NY, PA,TN, TX, UT, VA, and WA; and to countries of: Australia, Bulgaria, Canada, Croatia, France, Germany, Italy, Portugal, and Spain. | Open, Classified |  |  | 299 | Not_Computer | Not_Computer
Z-0496-2019 | 81472 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Vanguard Knee System -PS Open Box Femoral Right, 65 mm |  |  | October 22, 2018 | 2018 | Zimmer Biomet, Inc. | Label on the outer carton or the patient labels does not match label on the inner sterile packaging | Labeling mix-ups | Zimmer issued recall letter on 10/22/18 via email and FedEx. Acountts requested to immediately locate and quarantine affected product in your inventory. Return affected product from your distributorship and from affected hospitals. Complete Inventory Return Certification Form. uestions or concerns after reviewing this notice, please call customer service at 574-371-3071 between 8:00 am and 5:00pm EST, Monday through Friday. |  | 8 | 8 units | AR, MO, MN, NJ, NY, OH, MN , NJ, VA | Open, Classified |  |  | 8 | Not_Computer | Not_Computer
Z-0539-2019 | 81394 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Zimmer Natural NailÂż System - Tear Drop Guide Wire, 2.4 mm Diameter, 100 cm Length |  |  | October 22, 2018 | 2018 | Zimmer Biomet, Inc. | Various Trauma Guide Wires- creases in the sealing area on the end opposite of the product's chevron opening compromising product sterility | Packaging process control | Zimmer Biomet conducted recall to distributors and hospital risk managers on 10/22/18, distributors notified via email. Hospital risk managers, as well as distributors with product, will be notified via Fed'X." Distributors responsibilities include:- Locating and removing the product in their territory, as well as identifying hospitals with fielded inventory. Distributors will return on-hand product to Zimmer Biomet and ensure all of their products are accounted for using the form provided in the letter. Hospital risk managers will be provided with a letter identifying the issue and their responsibilities. These responsibilities include: Ensuring affected personnel are aware of the notice contents- Assisting the Zimmer Biomet sales representative with the quarantine of the affected products. Certificate of Acknowledgement complete and send to CorporateQuality.PostMarket@zimmerbiomet.com. This form must be returned even if you do not have affected products at your facility. Questions or concerns after reviewing this notice, please call customer service at 574-371-3071 between 8:00 am and 5:00pm EST, Monday through Friday |  | 3764 | 3764 | Worldwide - US Nationwide DistributionForeign: CANADAAUSTRALIABRAZILCHILEHONG KONGJAPANMALAYSIANETHERLANDSNICARAGUASINGAPORETAIWAN | Open, Classified |  |  | 202318 | Not_Computer | Not_Computer
Z-2749-2018 | 80584 | No | IFB | Cytocentrifuge. | Pathology | Pathology | 510(K) Exempt | EZ Single Cytofunnel Brown, A78710004 |  |  | June 04, 2018 | 2018 | Richard-Allan Scientific Company | Mislabeled product | Labeling mix-ups | On June 4, 2018, ThermoFisher Scientific issued URGENT MEDICAL DEVICE RECALL notices to customers. Customers were advised to inspect stock, quarantine, safely dispose or return impacted product. Evidence of disposal is required to be sent to Thermo Shandon Limited. If affected product has been used with no issues, please notify Thermo Shandon Limited of the quantity that is unavailable for disposal. Complete and return the Recall Return Response Acknowledgement &amp; Receipt form via email to marganne.miller@thermofisher.com |  | 41560 | 1,039 cases/40 boxes per case | US Nationwide Distribution | Open, Classified |  |  | 41560 | Not_Computer | Not_Computer
Z-3244-2018 | 80585 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | G8 Automated HPLC Analyzer - 723G8 |  |  | July 13, 2018 | 2018 | Tosoh Bioscience Inc | Analyzers were distributed with software which lacks a 510K | No Marketing Application | On July 13, 2018, Tosoh Bioscience started to telephone customers to set up on-site visits. On July 16, 2018, Tosoh Bioscience began to issue URGENT MEDICAL DEVICE RECALL notices to customers via hand delivery. Immediate Actions to be taken by the Customer/User:- Thoroughly review the content of this letter with the Tosoh field service representative.- Allow the Tosoh field representative to verify the software version currently on your analyzer- If the software is v5.23, your analyzer needs no further action- If the software is v5.24, allow the Tosoh field representative to download all patient results from your analyzer, remove software version v5.24, and upload software version v5.23.- Continue to use your analyzer and report results with software v5.23 once these activities have been completed.- Complete and return the Acknowledgement Form to the Tosoh field representative.- File the Urgent Medical Device Recall notice with your laboratory records and forward the information to others in your laboratory.If you have obtained unexpected test results or received any complaints of illness or adverse events associated with the use of the analyzer, please contact Tosoh Technical Support 24 hours a day, seven days a week at (800) 248-6764.Should you have any questions regarding this medical device recall, please feel free to contact Bernadette OConnell at (800) 248-6764 or by email at bernadette.oconnell@tosoh.com. She will be available to answer any questions Monday - Friday, from 9:00 AM to 5:00 PM (PST). |  | 8 | 8 units | US Nationwide Distribution in states of - CA, OK, VA 7 WI | Open, Classified |  |  | 8 | Not_Computer | Computer
Z-0503-2019 | 81445 | Yes | KPO | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | NxStage PureFlow B Solution-RFP-403, Premixed Dialysate for Hemodialysis |  |  | October 17, 2018 | 2018 | NxStage Medical, Inc. | PureFlow B Solution smaller chamber of the two chamberbag can burst electrolyte fluid and cause injury if it comes in contact with the eyes or skin | Packaging process control | NxStage initiated a recall on October 17, 2018 via " URGENT MEDICAL DEVICE REMOVAL" Recall Notice to its customers. The notice provided reason for recall, health risk, action to be taken. The customers were instructed to do the following: 1 Check all boxes in your Pureflow B solution inventory, separate them from your other inventory and do not use them.2. Follow the instructions on the attached reply form or click on the following link to respond electronically, https://www.nxstage.com/rfprecalloct2018, to acknowledge your receipt of this notice.3. Contact NxStage Customer Service to arrange for return of all affected product and for replacement product to be sent,If you have any questions or comments, please feel free to contact NxStage Customer Service 1-866- NXST AGE (1-866-697-8243). |  | 2189 | 2189 | US Nationwide Distribution. | Open, Classified |  |  | 423353 | Not_Computer | Not_Computer
Z-0508-2019 | 81480 | No | HIF | Laparoscopic insufflator. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | PNEUMOCLEAR Heated High-Flow Tube Set |  |  | October 22, 2018 | 2018 | WOM World of Medicine AG | It is possible that the packaging of the product can be damaged by the prongs on the tube set. | Packaging process control | WOM sent a Recall Notification letter dated October 22, 2018 to the their customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to send all in the US recalled products back to:W.O.M. WORLD OF MEDICINE USA, Inc.4531 36th StreetOrlando, FL 32811-6527WOM contacts for coordinating the shipment of the requested products to Orlando:Customer Service -W.O.M. OrlandoPhone: +1 407 472 0372Customer Service -W.O.M. OrlandoPhone: +1 407 472 0371Please send all recalled products returned to Stryker Venlo/Netherlands back to:W.O.M. WORLD OF MEDICINE GmbHAlte Poststr. 1196337 Ludwigsstadt - GermanyCustomer Service: +49 30 39981 524 |  | 124180 | 124,180 tube sets | Worldwide Distribution - US in the state of California and country of Netherlands | Open, Classified |  |  | 124180 | Not_Computer | Not_Computer
Z-0803-2018 | 79141 | No | JEB | Manual operating table and accessories and manual operating chair and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | The 2nd Assist Knee Positioner |  |  | September 07, 2017 | 2017 | Zimmer Biomet, Inc. | Zimmer Biomet is conducting a medical device recall for the 2nd Assistant Knee Positioner due to the potential for the sterile packaging barrier to be breached during distribution. | Package design/selection | On September 12, 2017 an URGENT MEDICAL DEVICE RECALL notice was issued to distributors and customers titled "The Second Assistant Disposable Knee Positioner". The notice requests customers to review their inventory and to quarantine any remaining product, which will be returned to the firm. Also, to notify any additional accounts which may have received product from you. Questions or concerns can be directed to Customer Service at 574-371-3071 |  | 38879 | 38,879 units | Nationally | Terminated | August 23, 2018 | 350 | 38879 | Not_Computer | Not_Computer
Z-0727-2018 | 79110 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Stage-1Âż RDS Indexed Abutment for RDS Stage-1Âż Single Stage |  |  | August 25, 2017 | 2017 | Keystone Dental Inc | The recalled units are mislabeled with an incorrect expiration date. The expiration date printed onthe label is 2202-06. The correct expiration date should be 2022-06. | Error in labeling | On August 18, 2017 a letter titled "Product Action  S2400K, Stage-1 RDS Indexed Abutment, Lot 34276" was issued to customers requesting persons to review their inventory and respond if product has been used or not. A response form is requested from these customers with instruction to return any unused product. Questions or concerns can be directed to:Customer Relations:Toll Free 866-902-9272 |  | 20 | 20 | Nationally. | Open, Classified |  |  | 20 | Not_Computer | Not_Computer
Z-0389-2019 | 81160 | No | LBR | Lipoprotein test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Pointe Scientific AutoHDL/LDL Cholesterol Calibrator, Catalog Number H7545-CAL - Product Usage: For the calibration of Pointe Scientific s autoHDL and autoLDL Cholesterol Reagent Set in serum or plasma. For in vitro diagnostic use only. |  |  | August 10, 2018 | 2018 | Medtest Holdings, Inc. | Vials of the autoHDL/LDL Calibrator were shipped without a vial label. | Error in labeling | Notification letters dated 8/14/18 were sent to customers. Customers were also contacted via telephone on 8/10/18. Additional letters containing labels dated 8/17/18 were sent to customers. |  | 22 | 22 | US Nationwide in the states of CA, FL, GA, MI, and OH. | Open, Classified |  |  | 22 | Not_Computer | Not_Computer
Z-3026-2018 | 80527 | No | DPW | Cardiovascular blood flowmeter. | Cardiovascular | Cardiovascular | 510(k) | VTI 8 MHz Surgical Doppler Transceiver Selectabable Channel |  |  | June 19, 2018 | 2018 | Vascular Technology, Inc. | Device maybe mislabeled as a 20 MHz Doppler transceiver instead of an 8 MHz Doppler transceiver | Labeling Change Control | The firm, Vascular Technology, Inc., "URGENT: MEDICAL RECALL DEVICE" letter dated June 18, 2018, to its customer on June 19, 2018. The letter described the product, problem and actions to be taken. The customers were instructed to immediately examine inventory and return all product subject to recall. In addition, if you have further distributed this product, please identify your customers and notify them at once of this product recall.If you have any questions, contact VP Sales person at 603-594-9700. |  | 10 | 10 units | US Distribution to state of: CA. | Open, Classified |  |  | 10 | Not_Computer | Computer
Z-2740-2018 | 80565 | No | HNN | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | Slit Knife 2.6mm 45degree Bevel Up, Model Number 373726. for use in ophthalmic surgery. |  |  | June 18, 2018 | 2018 | Beaver Visitec | Mislabel of product 373726 Lot 6001700 with 374891 Lot 6001701 on individual packages within the correctly labeled shelf carton. | Packaging process control | BVI issued PRODUCT RECALL NOTICES to customers dated June 11, 2018. Customers were instructed to inspect stock, quarantine and destroy product per internal procedures. BVI is kindly asking the customers to provide objective evidence of product destruction so that a credit can be applied to accounts. Customers are also encouraged to complete and return the response form. If you have questions regarding this recall please call Customer Service at 866-906-8080. |  | 440 | 440 | Distribution to US states of CA, IL, MD, MN, NJ, OH &amp; PA; and Israel | Open, Classified |  |  | 440 | Not_Computer | Not_Computer
Z-2967-2018 | 80692 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Foundation Modular Femoral Stem, Part Number, 220-10-110. Knee prosthesis component. |  |  | May 14, 2018 | 2018 | Encore Medical, Lp | After receipt of a product complaint for the labeling, it was determined that this device does not have FDA clearance. | No Marketing Application | The firm, DJO, sent an "URGENT FIELD SAFETY NOTICE" letter dated 5/14/2018 to it's customers via email, on May 14, 2018 and contained the following information: Description of Issue, Quantity in Field, product information and code, Action required, and an Acknowledgement form with contact information. The customers were instructed to: check your inventory for part number 220-10-110; Customer Service has already requested that these be quarantined by your agency; Contact Customer Service to request an RMA for return of these devices. As there are no currently cleared 10x100 femoral stem extensions, there is no product available for replacement; and complete the attached acknowledgment form, even if you do not have the affected product, and email the form to sandra.busick@djoglobal.com by May 18, 2018.Modular Femoral Cemented Stem -- 220-10-110 -determined that this device does not have FDA clearance.If you have any questions, please contact Manager, Regulatory Affairs at (512) 834-6255 or by email at teffany.hutto@djoglobal.com. |  | 12 | 12 | U.S. Distribution to states of: AZ, MO, MS, NY, OK, and UT; internationally to: Japan. | Open, Classified |  |  | 12 | Not_Computer | Not_Computer
Z-0777-2018 | 79289 | No | LWK | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | 3M Surgical Clipper Professional 9681 |  |  | February 02, 2018 | 2018 | 3M Company - Health Care Business | Failure to follow proper charging practices can result in lithium-ion battery degradation, characterized by excessive heat during operation or failure to properly charge. In rare circumstances, battery degradation has been associated with battery venting, which is a rapid release of battery energy. | Labeling design | The firm initiated their field safety action on 02/08/2018 by certified mail to customers and distributors. Larger distributor were notified by email on 02/12/2018. The notices to the distributors stated:"3M is notifying all users of the 3M" Surgical Clipper Professional 9681 of an added product warning, stressing the importance of following proper clipper charging practices as documented in the Model 9681 Clipper instructions for use. Current charging instructions recommend leaving the clipper body in the drop-in charger in between uses.Failure to follow proper charging practices can result in lithium-ion battery degradation, characterized by excessive heat during operation or failure to properly charge. In rare circumstances, battery degradation has been associated with battery venting, which is a rapid release of battery energy. 3M has received two reports of battery venting resulting in minor burn injury.3M is adding the following product warning to emphasize the importance of proper charging practices:Warning: To avoid potential battery degradation, charge the clipper after each use and store in the drop-in charger stand when not in use.Distributors are being asked to take the following actions:" Identify those customers who have ordered the 3M" Surgical Clipper Professional 9681 or the 3M" Surgical Professional Clipper Blades 9680 beginning January 1, 2016 to date." Communicate this Field Safety Notice to those customers."The notice to the customer stated the following:"3M is notifying all users of the 3M" Surgical Clipper Professional 9681 of an added product warning, stressing the importance of following proper clipper charging practices as documented in the Model 9681 Clipper instructions for use. Current charging instructions recommend leaving the clipper body in the drop-in charger in between uses.Failure to follow proper charging practices can result in lithium-ion battery degradation, characterized by excessive heat durin |  | 120000 | 120000 units | Worldwide distribution including US nationwide, including Puerto Rico. Singapore, Australia, Canada, Chile, Colombia, Costa Rica, Brazil, Dominica, Ecuador, Guatemala, UAE, Israel, Japan, Korea, Mexico, new Zealand, Paraguay, Peru, Taiwan, Uruguay, India, Bolivia, El Salvador, Saudi Arabia, and Thailand. | Open, Classified |  |  | 120000 | Not_Computer | Computer
Z-0638-2019 | 81579 | No | KGO | Non-powdered surgeon's glove. | General &amp; Plastic Surgery | General Hospital | 510(k) | Protexis Neoprene Surgical Glove, Size 9.0 |  |  | October 31, 2018 | 2018 | Cardinal Health 200, LLC | Labeling issue in that the breakthrough (i.e., permeation) times for two chemotherapy drugs (Carmustine and Thiotepa) stated on the wallet (primary packaging) of the Protexis Neoprene surgical gloves are incorrect. Additionally, breakthrough times for Mitomycin were added to the same label in error. | Labeling mix-ups | Cardinal Health notified customers on about 10/31/2018 via "URGENT: PRODUCT CORRECTION NOTICE." Customers were instructed to examine inventory for affected products, affix the provided warning label on the principal display panel on the affected dispenser packs, notify internal parties and/or customers if the product was further distributed of the recall, and complete and return the provided Acknowledgement Form. If the wallets containing the gloves are removed from the dispenser packs, customers were also instructed to provide a warning label to accompany the affected product inform users of the correct breakthrough times. If customers need additional labels, they were instructed to contact 800-292-9332 or gmb-fieldcorrectiveaction@cardinalhealth.com for assistance. For further questions regarding the Product Correction Notice, please reference the following Customer Service telephone numbers: Hospital (800-964-5227), Federal Government (800-444-1166), Distributor (800-635-6021), All other customers (888-444-5440). |  | 1475095 | 1,475,095 pairs total | Worldwide Distribution: US (Nationwide distribution) to states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, WY and to countries of: Austria, Belgium, Canada, Chile, Costa Rica, France, Germany, Italy, Luxembourg, Martinique, Monaco, Netherlands, Poland, Reunion, Spain, Sweden, Switzerland, Taiwan, and United Kingdom. | Open, Classified |  |  | 1475095 | Not_Computer | Not_Computer
Z-3253-2018 | 80919 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Medtronic DxTerity(TM) TRA 5F, Diagnostic Catheter, REF U5TRAN35 |  |  | August 16, 2018 | 2018 | Medtronic Vascular | Medtronic has determined that the French size indicator on the inner pouch may incorrectly reflect a 6 French size, rather than the correct 5 French size. The catheters inside the pouch are the correct U5TRAN35 item, the outer carton has all the correct labeling, and all the other information and configuration details on both the carton and the pouch are correct. | Error in labeling | Medtronic initiated consignee communications on 16-Aug-2018 to Risk Managers and Health Care Professionals. US notifications were sent via 2-Day UPS delivery. Consignees are asked to quarantine and return any unused potentially affected product to Medtronic. |  | 250 | 250 units | CA, IA, NE, TN, TX, VA, WA, Jamaica, New Zealand, Thailand | Open, Classified |  |  | 250 | Not_Computer | Not_Computer
Z-0654-2019 | 81616 | No | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | Empower FastLoad CT Syringe Pack -sterile disposable syringe kit for use with the EmpowerCTÂż, EmpowerCTAÂż and EmpowerCTAÂż+ Injector Systems |  |  | November 12, 2018 | 2018 | Bracco Injeneering S.A. | Insufficient seal on the sterile barrier of the device, compromising the sterility of the device | Packaging process control | The firm, Bracco Injeneering, issued "URGENT: MEDICAL DEVICE RECALL" letter via third party recall firm, on 11/12/18 to customers and distributors identifying the problem, and requesting return of unused product., Complete and return the Removal Recall Response Form even if you have no recalled stock on hand.Questions: contact Stericycle at 888-759-6907 (Monday- Friday 8 a.m. to 5 p.m. Eastern Time) or acistmedical3021@stericycle.com. |  | 33350 | 33,350 packs (667 cases) | US Nationwide Distribution | Open, Classified |  |  | 33350 | Not_Computer | Not_Computer
Z-0899-2018 | 79199 | No | GBX | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Silicone Fluted Round Drain 24Fr (hubless) &amp; Connector, Sterile. Catalog Number JP-2234. |  |  | October 03, 2017 | 2017 | Degania Silicone, Ltd. | Sterile pouches are missing labels on them. The boxes themselves are properly labeled. The issue is related to the production lot P1607442 only. | Error in labeling | An Urgent Medical Device Correction letter, dated October 03, 2017, was sent to the customer. The letter identified the affected device and the reason for the recall. The letter instructed the consignee to check its stock in order to remove all mislabeled product from the market and eliminate potential risk of misuse. The customer was instructed to return the completed Confirmation Sheet to Degania Silicone, as well as any unlabeled product. Questions regarding the recall should be directed to Katia Kayam, Customer Complaints Officer at Katia@ds-il.com. |  | 4000 | 4,000 units | US Distribution to Puerto Rico. | Open, Classified |  |  | 4000 | Not_Computer | Not_Computer
Z-2969-2018 | 80372 | No | MNQ |  |  | Anesthesiology | PMA | Inspire II Model 3024 Implantable Pulse Generator (IPG) for Obstructive Sleep Apnea, Inspire Part #900-003-020, |  |  | June 12, 2018 | 2018 | Inspire Medical Systems Inc. | Incorrect use-by date on the device registration/patient file labels. | Error in labeling | The firm's sales team members called all of the affected implanting surgeons between 6/12-14/2018 to discuss the recall. The recalling firm issued "Dear Health Care Provider" letters dated 6/15/2018 via certified mail on 6/15/2018 to all implanting surgeons and hospital customers. The letter identified the affected product, problem and actions to be taken. For questions call 1-763-205-7974. |  | 93 | 93 devices | US Nationwide Distribution in the states of AL, AZ, CA, FL, GA, IL, KY, MA, MI, MN, MS, NY, OH, PA, TN, TX, VA, WA, and WI. | Completed |  |  | 93 | Not_Computer | Not_Computer
Z-0678-2019 | 81641 | No | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | HydroSet XT Injectable HA Bone Cement, REF 897005, STERILE |  |  | November 15, 2018 | 2018 | Howmedica Osteonics Corp. | Incorrect expiry date of 9-Feb-20Correct expiry date is 30-Aug-19 | Labeling mix-ups | On November 15, 2018, Stryker issued Urgent Medical Device Recall notices to customers via Fedex. Customers were advised to take the following actions:1. Inform users of the medical device recall and forward a copy of the notice to all those individuals who need to be aware within the organization.2. Examine inventory to identify impacted HydroSet kits.3. Immediately quarantine any affected product.4. Complete and return the Business Reply Form via email to Strykerortho6629@stericycle.com or via fax to 1-888-345-3845.5. Return all affected product available at your location to the following address:Stryker Orthopaedics/PFA Product ReturnsAttn: Distribution Inventory Team 325 Corporate Drive, Dock M-East, Mahwah, NJ 07431, Ref. PFA 1878253 |  |  |  | CO, FL, GA, IA, MA, ME, MI, MN, NC, ND, NH, NY, PA, SC, TN, TX, VA, WI &amp; WV | Open, Classified |  |  |  | Not_Computer | Not_Computer
Z-0818-2018 | 79105 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Ziptight Ankle Syndesmosis Fixation Device, Titanium Implant Titanium Implant for surgical repairs of soft tissue. |  |  | August 22, 2017 | 2017 | Zimmer Biomet, Inc. | Ziptight Ankle Syndesmosis Fixation was packaged without the tip protectors. | Packaging process control | Customers were notified via letter on about 08/22/2017. Instructions for distributors, sales representatives, and distributor operation managers include to ensure affected personnel are aware of the recall, locate and quarantine affected product in inventory, complete and return the Certification of Acknowledgement, return affected product, and identify any additional hospitals or surgeons that should be notified. Instructions for surgeons and risk managers include to ensure affected personnel are aware of the recall, identify and quarantine any affected product on hand, complete and return the Certificate of Acknowledgement, return any affected product. Questions or concerns can be addressed to the customer call center at 1-877-946-2761 between 8:00 am and 5:00pm EST, Monday through Friday. Additionally, questions may be sent by email to corporatequality.postmarket@zimmerbiomet.com. |  | 885 | 885 units total | Distributed domestically to AK, AL, AR, AZ, CA, CO, FL, ID, IN, KY, MA, MD, ME, MI, MN, MO, NC, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WI. Distributed internationally to ARGENTINA, AUSTRALIA, CANADA, COSTA RICA, JAPAN, NETHERLANDS. | Open, Classified |  |  | 885 | Not_Computer | Not_Computer
Z-0622-2018 | 78997 | No | HBO | Scalp clip. | Neurology | Neurology | 510(k) | Model #: CG8900, Medtronic Clip Gun Scalp Clip Kit (Qty 1); Model #: CG8901, Medtronic Clip Gun Scalp Clip Kit (Qty 10); Model #: CG8901A, Medtronic Clip Gun Applicator (Qty 10) |  |  | December 20, 2017 | 2017 | Medtronic Neurosurgery | Medtronic is initiating a voluntary recall because the products identified above did not adequately demonstrate acceptable packaging performance to current standards. | Package design/selection | A Consignee Notification with response form was sent to affected consignees on 12/20/2017. |  |  | 27,432.8 devices | Algeria, Armenia, Australia, Canada, Cyprus, Egypt, Germany, Ghana, Greece, Guam, Iceland, Indonesia, Iran, Italy, Jordan, New Zealand, Poland, Portugal, Puerto Rica, Russian Federation, Saudi Arabia, Singapore, South Africa, Spain, Swaziland, Sweden, Switzerland, Syria, Tunisia, Turkey, UAE, UK, USA, Yemen | Open, Classified |  |  |  | Not_Computer | Not_Computer
Z-0536-2018 | 79070 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | PowerPICC Provena Catheters, as 3 Fr. Single lumen and 4 Fr. Dual Lumen. Packaged in a poly tray with Tyvek lid within a header bag configuration or a breather bag configuration. |  |  | December 28, 2017 | 2017 | Bard Access Systems Inc. | Possible leaks on the 3 Fr. Single lumen and 4 Fr. Dual lumen PowerPICC catheters at the luer hub extension leg junction. | Labeling design | Customers were notified via letter on approximately 12/28/2017. Instructions include a reminder that if additional extension sets are attached to the luers of the catheters, they should be appropriately secured as per hospital recommended securement protocol, to refrain from having unsupported extension sets that could put stress on the luer extension leg junction of the catheter, and to notify customers if the product was further distributed. Also, the notification contained a statement that Bard Access is in the process of updating their Instructions for Use to include a statement regarding the importance of properly securing additional extension sets per hospital recommended securement protocol. On 04/11/2018, the recall was expanded and customers were notified via letter, which explained that the previous recall has been expanded to involve return of affected product as well as include product distributed from February 2017 through March 2018. Instructions included to examine inventory for affected product, quarantine and immediately discontinue distribution of the product, complete and return the Business Recall Response Form, arrange for the return of affected product, and notify customers if the product has been further distributed. For distributors, Bard Access Systems requested a customer list so that they could notify their customers. |  | 53699 | 53,699 units total | Distribution US nationwide. | Open, Classified |  |  | 53699 | Not_Computer | Not_Computer
Z-3282-2018 | 80943 | No | HQL | Intraocular lens. | Ophthalmic | Ophthalmic | PMA | C-flex 570C +19.0D Intraocular Lens |  |  | August 09, 2018 | 2018 | Rayner Intraocular Lenses Limited | Firm become aware of reports of post-operative refractive errors following implantation of lenses. | Error in labeling | On August 31 the firm sent letters to all customers (health care facilities) that have received lenses in this batch and instructed to quarantine any remaining unused lenses from the C-flex 570C +19.0D batch 017100825. Replacement, reimbursement or substitution is offered (as appropriate). |  | 49 | 49 | US in the states of MO | Open, Classified |  |  | 49 | Not_Computer | Not_Computer
Z-3023-2018 | 80529 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | fMRI Hardware System VisualSystem allows video signals from the stimulus presentation PC to enter the shielded scanner room and presented by a set of coil-mounted displays or by an in-room LCD monitor. - LCD Monitor Power supply (LPS) for LCD Monitor, Versions 1.0 and 1.1 |  |  | August 20, 2018 | 2018 | NordicNeuroLab AS | Complaints of this version of the power supply unit (LPS) (1.0 and 1.1) being dragged into the magnet bore (MR Scanner). | Labeling design | The firm sent a letter explaining the issue by email to the consignees on 08/20/2018. The firm requested that the consignee check the stock without removing the device from the magnet. It is was not in the magnet room, safely quarantine the device, and respond to the recall notice. The firm also directed any consignee who further distributed the product to notify those entities and to contact NNL. NNL will follow-up with them directly. NNL will upgrade all existing models with a safer design of the LCD Power supply unit. |  | 163 | 163 units | worldwide | Open, Classified |  |  | 163 | Not_Computer | Computer
Z-0549-2018 | 78219 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Humeral Guide Wire Ball Tip 2.4 mm Diameter 70 cm Length, Item Number 47-2255-008-00 |  |  | June 05, 2017 | 2017 | Zimmer Biomet, Inc. | The design verification for the previous packaging configuration G928 does not cover the 70cm wires. A design verification has been completed to move the 70cm guide wires to a new packaging configuration. As a result, the products packaged in the previous packaging configuration are being removed. | Package design/selection | On June 5, 2017, Zimmer Biomet distributed Urgent Medical Device Recall notices to their customers via FedEx and email. **Hospital Risk Managers are advised to:** 1. Review this notification and ensure affected team members are aware of the contents.2. Complete the Certification of Acknowledgement portion of Attachment 1a. Return a digital copy to corporatequality.postmarket@zimmerbiomet.com within three (3) days.3. Assist your Zimmer Biomet sales representative quarantine all affected product.**Distributors are advised to:**1. Review this notification and ensure affected personnel are aware of the contents.2. Immediately locate and quarantine affected product in your inventory.3. Complete the Certification of Acknowledgement portion of Attachment 1  Inventory Return Certification Forma. Return a digital copy to corporatequality.postmarket@zimmerbiomet.com within three (3) days.4. Note that any hospitals that received direct shipments of this product from Zimmer Biomet will be sent a copy of the Risk Manager directly. It is important that you review the list of hospitals included with the email notification sent to your facility to identify additional accounts Zimmer Biomet has not notified. Using the Additional Accounts Form provided with the email notice sent to your facility, return contact information for any additional hospitals that may have received or used the affected product within the past 12 months. If there are no additional users to notify, please confirm this by annotating and returning the form provided indicating that no additional users have been identified.5. Return all affected product from your distributorship and affected hospitals within your territory along with a completed Attachment 1  Inventory Return Certification Form to Zimmer Biomet within 2 weeks.a. For each return, send a copy of Attachment 1 to corporatequality.postmarket@zimmerbiomet.com.b. Include a hardcopy of Attachment 1 with your shipment for i |  | 7888 | 7,888 units | Nationwide within U.S.A and Puerto Rico; International: Australia, Brazil, Canada, China, Germany, India, Japan, Netherlands &amp; Singapore | Open, Classified |  |  | 27442 | Not_Computer | Not_Computer
Z-0473-2019 | 81391 | No | MJN | Vascular clamp. | Cardiovascular | Cardiovascular | 510(k) | LeMaitre Vascular Pruitt F3Âż-S Polyurethane Outlying Carotid Shunt, Sterile |  |  | October 01, 2018 | 2018 | LeMaitre Vascular, Inc. | Pouches from this lot are not sealed compromising the sterility | Packaging process control | Lemaitre Vascular initiated recall by letter on October 1, 2018, provided with reason for recall action to be taken, risk to health. Complete the enclosed customer reply form, and return it to LeMaitre Vascular, Inc. Return all devices from lot PFP1055, whether they appear to be unsealed or not. If you distributed any affected product to other facilities, please forward this recall to the recipients.Contact Information: rlerer@lemaitre.com 781-425-1671 Monday through Friday, 8:00 AM to 4:30 PM, |  | 46 | 46 units | FLForeign: Austalia, Korea | Open, Classified |  |  | 46 | Not_Computer | Not_Computer
Z-0590-2018 | 77469 | No | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | Arctic Sun ArcticGel Pads - Medium, Product Code 317-07 |  |  | June 02, 2017 | 2017 | C.R. Bard, Inc. | Label on product box contains the incorrect reference #317-09 while the shelf box label with barcode and the internal foil pouch label are correctly labeled with product code 317-07 and lot number NDAR0462. | Labeling mix-ups | Customers were notified on approximately 06/02/2017 via FedEx with delivery confirmation and per the individual country's delivery system. Instructions included to examine inventory and identify any product subject to the recall and notify customers if product was further distributed. If customers wished to return product, they were instructed to contact customer service at 1-800-526-4455. |  | 128 | 128 units | Distributed domestically. Distributed internationally to Australia and Belgium. | Open, Classified |  |  | 128 | Not_Computer | Not_Computer
Z-0552-2018 | 78444 | splint, extremity, non-inflatable, external, non-sterile | NOC | Noninflatable extremity splint. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ACE(TM) BRAND, ULTRA LITE ANKLE brace: (a) Size Selection Medium, Cat #901014, UPC 051131019984 2; (b) Size Selection Large, Cat #901015, UPC 0 51131019985 9 |  |  | November 01, 2017 | 2017 | 3M Company - Health Care Business | Mold was found on the non-patient contactsurfaces of the product and was identified asAspergillus candidus. | Packaging process control | The firm initiated their recall on 11/01/2017 by letter. The letter requested the consignee cease distribution of the recalled product, isolate it pending return. Distributors were asked to notify their customers including a copy of the 3M recall notice. |  | 996 | 996 units | AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, IL, IN,KS, KY, LA, MD, MI, MO, MS, NC, NH, NY, OH,OR, PA, SC, TN, TX, UT, VA, WI, WV | Completed |  |  | 996 | Not_Computer | Not_Computer
Z-2822-2018 | 80540 | No | GZT | Self-retaining retractor for neurosurgery. | Neurology | Neurology | 510(k) | NICO Shepherd' s Hook - Greenberg, sterile. An accessory to NICO s BrainPath device |  |  | June 19, 2018 | 2018 | Nico Corp. | Mislabeled hook pouch included a transport/storage label rather than the product label containing key product information such as part number, lot number, and expiration date, etc | Packaging process control | On June 19, 2018 , NICO issued letter by mail and/or email letters to consignees that have received hooks from the affected lot. The letter provides instructions to inspect all of its current stock of NN-8024, Greenberg Hooks for the proper pouch label. if an incorrect label is discovered, the hook should beremoved from stock and returned to NICO.Questions or need additional information, please contact jay.dittman@.niconeuro.com or 317-660-7118 ext. 104. |  | 204 | 204 devices | US Nationwide distribution | Terminated | November 09, 2018 | 143 | 204 | Not_Computer | Not_Computer
Z-0705-2018 | 78929 | No | CCW | Rigid laryngoscope. | Anesthesiology | Anesthesiology | 510(K) Exempt | Rusch(R) Greenlite MAC 1, Catalog Number 004551001 |  |  | January 10, 2017 | 2017 | Teleflex Medical | The pouch may contain the incorrect RuschGreen Lite product size. Pouches that are labeled as containing RUSCH GREEN LITE MAC 1 mayinstead contain RUSCH GREEN LITE MAC 2. | Package design/selection | The firm initiated their recall by letter on 01/10/2018 via FedEx 2-day. Affected consignees were instructed to take the following actions:"1. If you have affected stock, immediately discontinue use and quarantine any products with the product code and lot number listed above, so that the affected products can be returned to Teleflex.2. To return product, complete the enclosed Recall Acknowledgement Form and fax it to 1-855-419-8507, Attn: Customer Service or email to recalls@teleflex.com. This will allow us to document the amount of product you have on hand for return. A customer service representative will contact you with a Return Goods Authorization (RGA) Number and will provide instructions for the return of products to Teleflex Medical.3. If you have no affected stock, please complete the enclosed Recall Acknowledgment Form and fax it to 1-855-419-8507, Attn: Customer Service or email to recalls@teleflex.com. This will allow us to document your receipt of this letter." |  | 2140 | 2140 units | US only | Open, Classified |  |  | 2140 | Not_Computer | Not_Computer
Z-0383-2018 | 78802 | No | KGN |  |  | General &amp; Plastic Surgery | 510(k) | PriMatrixÂż is an acellular dermal tissue matrix derived from fetal bovine dermis. The device is supplied sterile in a variety of sizes to be trimmed by the surgeon to meet the individual patient's needs. |  |  | June 15, 2017 | 2017 | TEI Biosciences | During Final Assembly Operation, it was noticed that some of the PriMatrix Instructions for Use (IFUs) from Lot RM1608011 Part 520-002v10 contained extra pages and/or missing pages:" 22 IFUs contained duplicate pages 15 - 26 &amp; 51 - 62 and were also missing pages 27 - 50 " 3 IFUs were missing pages 27 - 50 " 1 IFU was missing pages 1 - 2 &amp; 75 - 76" 1 IFU was missing pages 15 - 62While pages 27  50, 75  76, and 15  62 contain instructions written in a language other than English, pages 1  2 contain information regarding the indications for use, contraindications, device description, warning and precautions, and potential complications of the device. | Error in labeling | Integra LifeSciences Corporation notified (all U.S.) customers via Letter/Medical Device Correction Field Safety Notice dated June 15, 2017. Letter delivered by courier service, facsimile or email July 15, 2017. |  | 4081 | 4081 units | U.S. | Terminated | March 15, 2018 | 273 | 4081 | Not_Computer | Computer
Z-2896-2018 | 80645 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Philips SureSigns VS2 monitor, 863079 |  |  | July 13, 2018 | 2018 | Philips Electronics North America Corporation | Philips has received several reports in which monitors with lithium ion batteries that had exceeded their battery life expectancy have overheated or ignited.These batteries should be replaced every 3 years or upon reaching 300 charge-discharge cycles.Although the Philips SureSigns VS &amp; VM Monitors and View Station (VSV) can display actual information on either or both battery age and charge-discharge cycles, the existing labeling does not include full instructions on how to use this information to determine when to replace the battery. An overheated battery may in turn cause the device case to overheat and possibly melt or cause the device to ignite, which can cause injury to a patient, nearby users, or cause damage to property. | Labeling design | On July 13, 2018, the firm issued an "URGENT-Medical Device Correction" letter dated July 10, 2018, to affected customers to advise them that that monitors with batteries that have exceeded their specified replacement interval may overheat or ignite. The firm issued a Service Guide addendum included with the letter, which contains information to assist with managing the battery and its replacement.Customers were advised to do the following:Upon receipt of this notification, carefully read the enclosed SureSigns Addendum to the Service Guide.Âż VS Series Customers -Addendum to the Service Guide: 453564807541Âż VM Series Customers -Addendum to the Service Guide: 453564807551Âż VSV Customers- Addendum to the Service Guide: 453564815301Promptly perform the Battery Maintenance that is called out in the Addendum for each of your affected Philips SureSigns VS &amp; VM Monitors and View Station (VSV) by determining if the Battery Information Screen indicates if the battery cycle count exceeds the limit of 300 cycles, or the battery is older than three (3) years. If so, the battery needs to be replaced. The replacement battery can be ordered using the standard Philips replacement processes.Details on how to replace the battery can be found in the Philips SureSigns VS &amp; VM Monitors and View Station (VSV) Service Guide.Review this information with all staff members who are responsible for device management of the Philips SureSigns VS &amp; VM Monitors and View Station (VSV).Please store the below Service Guide Addendum with your Philips SureSigns VS &amp; VM Monitors and View Station (VSV) Service Guide documentation.Âż VS Series Customers -Addendum to the Service Guide: 453564807541Âż VM Series Customers -Addendum to the Service Guide: 453564807551Âż VSV Customers- Addendum to the Service Guide: 453564815301Complete and return the response card provided.In addition to providing the Service Guide Addendum and this notice, Philips plans to release a customer ins |  | 207439 | VM 96,324 Units; VS 110,424 Units; VSV 691 Units | Worldwide distribution: US (nationwide) including Puerto Rico and countries of: Afghanistan, Albania, Andorra, Argentina, Australia, Austria, Bahrain, Bangladesh, Belgium, Bhutan, Bolivia, Brazil, Bulgaria, Canada, Cambodia, Chile, China, Colombia, Costa Rica, Cote D'Ivoire, Croatia, Czech Republic, Dominican Republic, El Salvador, Egypt, Estonia, Ethiopia, France, French Guiana, French Polynesia, Finland, Guadeloupe, Germany, Ghana, Hong Kong, Hungary, Iceland, India, Indonesia, Iran, Israel, Italy, Japan, Jordan, Kenya, Korea, Kosovo, Republic of, Kuwait, Latvia, Lebanon, Lithuania, Luxembourg, Lybia, Macao, Malaysia, Maldives, Malta, Martinique, Mauritius, Mexico, Morocco, Mozambique, Myanmar, Namibia, Nepal, Netherlands, New Caledonia, New Zealand, Nicaragua, Norway, Oman, Pakistan, Palestine, State of, Peru, Philippines, Poland, Portugal, Qatar, Reunion, Romania, Russian Federation, Saudi Arabia, Singapore, Slovakia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Trinidad and Tobago, Turkey, United Arab Emirates, United Kingdom, Venezuela, Viet Nam, Zambia and Zimbabwe. | Open, Classified |  |  | 207439 | Not_Computer | Computer
Z-2884-2018 | 80613 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Stryker T2 ANKLE ARTHRODESIS NAIL RIGHT 10MMX200MM Item Number: 1819-1020S |  |  | June 27, 2018 | 2018 | Stryker GmbH | Incomplete seal on the Tyvek lid may compromise the sterility of the device | Packaging process control | Stryker notified consignee via UPS (with return receipt) on June 27, 2018 advising user of the problem, quarantine product and return. Complete and sign the Recall Notification Business Reply Form. Contact (201) 831-6693. |  | 38 | 38 | US distribution to KY | Open, Classified |  |  | 38 | Not_Computer | Not_Computer
Z-0444-2019 | 81003 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Espiner Tissue Retrieval System-Eco-Sac |  |  | June 20, 2018 | 2018 | Espiner Medical Ltd | Label was incorrect. The label states that thevolume of the retrieval bag was 1200ml, but it should state 1000ml | Error in labeling | Espinar Medical notified the distributor to notify accounts to review inventory, product batches should not be used, and should be segregated if deemed necessary by the hospital/healthcare facility. Complete the acknowledgement form. Contact Global Medical Partners LLC via email at cs@globalmedllc.com or by calling 888.384.8490 to obtain an RGA number for the return of the product. |  | 230 | 230 units | US Nationwide in the states of CA, IL, MT, NC, NJ, | Open, Classified |  |  | 230 | Not_Computer | Not_Computer
Z-0559-2019 | 81054 | No | KHO | Fluorometer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AIA-360 Analyzer, product code 019945/019945R |  |  | August 23, 2018 | 2018 | Tosoh Bioscience Inc | The firm has implemented new product labeling changes for the Tosoh AIA-360 Analyzer. These labeling changes affect the Tosoh Quick Reference Guide, Training Manual and Training DVD previously provided by Tosoh Bioscience.These changes are being made to align the Tosoh Quick Reference Guide and associated training materials to the Operators Manual and provide clarification to the requirement that the maintenance of the B/F probe is to be performed by Tosoh a field service representative, the requirement for use of 70% Ethanol during daily shutdown procedures and the requirement for use of CAP Class 1 reagent grade water for dilution of the concentrated wash and diluent solutions and reconstitution of lyophilized reagents. Failure to follow these requirements may negatively affect system performance, lead to corrosion of instrument components over time, or result in bacterial contamination, respectively. | Labeling Change Control | On August 23, 2018, the firm, Tosoh Bioscience, notified affected distributors and customers via "URGENT: MEDICAL DEVICE RECALL" letter. The described the product, problem and actions to be taken. Summary and location of labeling changes:Âż- Elimination of monthly B/F probe maintenance.Âż- Added requirement for use of 70% Ethanol.Âż- Added requirement for use of CAP Class l reagent grade water per CLSI Guidelines and Standards.Âż- Added clarification of substrate stability. Substrate, when stored at 2-8ÂżC between uses is stable for 30 days.Customers were asked to do the following: -Discontinue maintenance activities on the B/F probe. If issues with the B/F probe are observed, call Technical Support for assistance.Âż- Please continue to perform all your other maintenance procedures at the required intervals as per the instructions in the Operators Manual.Âż- Discard the Training DVD, it is now obsolete. An updated version of the Training DVD, QRG and Training Manual will be sent to you in a separate mailing.Âż- At daily shutdown, use 70% Ethanol to replace the substrate solution.Âż- Use only CAP Class l reagent grade water.Âż- Complete and return the attached Acknowledgement Form to Tosoh Bioscience, Inc., within 15-days of receiving this notification.Âż- Maintain this notification with your laboratory records and forward this information to others who may have received this product.Âż- If you have obtained unexpected test results or received any complaints of illness or adverse events associated with the use of these products, contact Tosoh Technical Support 24 hours a day, seven days a week at (800) 248-6764.Should you have any questions regarding this medical device recall, please contact call (800) 248-6764 or email at bernadette.oconnell@tosoh.com Monday - Friday, from 9:00 AM to 5:00 PM (PST). |  | 1228 | 1228 | Worldwide Distribution: US (Nationwide) and countries of: Mexico, Canada, Colombia, Chile, Venezuela, Ecuador, Uruguay, Peru, and Honduras. | Open, Classified |  |  | 1228 | Not_Computer | Not_Computer
Z-0440-2019 | 81171 | No | NEK |  |  | Orthopedic | PMA | INFUSE Bone Graft X SMALL KIT REF 7510100 |  |  | September 27, 2018 | 2018 | Medtronic Sofamor Danek USA Inc | The Extra Small INFUSE Bone Graft kit from lot M111806AAA, may contain two absorbable collagen sponges (ACS), but should only contain one ACS. | Error in labeling | The firm initiated the recall by letter on 09/27/2018. The firm will coordinate the return of the recalled product. |  | 138 | 138 units | US | Open, Classified |  |  | 138 | Not_Computer | Not_Computer
Z-0581-2019 | 81519 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Laser Fiber Assy-Single Use-LFS-200-0.22NA-SMA905-Extension Sleeve Silver-Sterile ETO-ForTec Fiber SU-200-TRUE-RT Label Smartscope packaging. |  |  | October 17, 2018 | 2018 | Leoni Fiber Optics GmbH | A percentage of the products shipped in Smartscope single-box packaging may exhibit a broken package seal, thus having compromised sterility. There have been two reports of compromised sterility; however, all products packaged in Smartscope single-box packaging are at risk of a compromised product seal.Broken seal pouches compromise sterility and could increase the risk of infection if used. | Packaging | The firm, LEONI, notified its distributor of the product issue on 10/16/18 via phone call and a subsequent "URGENT: MEDICAL DEVICE RECALL" letter. The distributor was informed of the product issue, and asked to disseminate recall information to customers.Actions to be taken by the Customer/User:End users should discontinue usage of the product. Product should be immediately quarantined. The Distributor should complete the acknowledgement of receipt form, to arrange for replacement and or credit to their account. A Leoni Customer Service Representative shall contact you with a Return material Authorization. If a replacement order is needed please inform customer service at the time of RMA issuance. - Discontinue usage if packaging irregularities exist- Return RMA units to Leoni Fiber Optics Inc. - Inform Customer Service if replacement product is required.- Complete acknowledge form and return back to Leoni Fiber Optics Inc. Attention Roger West, LFOI Quality Manager, Roger.West2@leoni.comIf you have any questions, please contact Quality Manager at 757-258-4805 orEmail: Roger.West2@leoni.com. |  |  |  | US Distribution to the state of: Ohio. | Open, Classified |  |  |  | Not_Computer | Not_Computer
Z-0615-2018 | 79131 | No | FAJ | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Video Cystoscope models ECY-1570 and ECY-1570K |  |  | July 13, 2017 | 2017 | Pentax of America Inc | The video cytoscopes lack 510(k) premarket notification clearance. | No Marketing Application | Customer Instructions: Customers should inspect their stock of PENTAX Medical ECY-1570 and ECY-1570K video cystoscopes, quarantine them, and discontinue use of these products. If you have sold and distributed these products, please provide a copy of the notice and response form to the new owners. Enclosed with this letter is a customer response form. The form identifies the affected endoscopes (model and serial numbers) which have been sold to your facility. The attached form should be completed and returned to PENTAX Medical. Contact Information: If you have any questions regarding this action, please feel free to contact PENTAX Medical Customer Service. Tel: 800-431-5880 (8:30 AM  5:00 PM, Monday  Friday, EST) Fax: 201-799-4063 (alternate 201-391-4189) Email: customeradvisories@pentaxmedical.com |  | 24 | 24 | Distributed in 8 states: CA, IL, MD, MI, ND, NH, NY, SD, and Guam. | Open, Classified |  |  | 24 | Not_Computer | Not_Computer
Z-0538-2018 | 78936 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Alteon HA Femoral Stem, Press-Fit, Extended Offset, Cementless, HA Coated, Collared, Size 5 |  |  | October 09, 2017 | 2017 | Exactech, Inc. | Wrong device description of the label identifying the implant | Error in labeling | Exactech sent an Urgent Medical Device Recall letter dated October 9, 2017, informing them of the potential mislabel. The notice requested consignees: Immediately cease distribution/use of this product; Extend information to accounts in possession of the product; Identify and quarantine any subject devices in their inventory; Complete and return Recall Acknowledgement and Recall Inventory Response Forms. For further questions, please call (800) 392-2832. |  | 34 | 34 | US Distribution to the states of : AL, CA, CO, FL, MD. and NY | Open, Classified |  |  | 34 | Not_Computer | Not_Computer
Z-2881-2018 | 80639 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Dymaxeon 5.5 Curved Rod Size 5.5 x 110mm (Yellow), Model Number 06.6441B |  |  | June 29, 2018 | 2018 | Back 2 Basics Direct LLC | Packages labeled as containing 110mm rods may contain 100mm rods. | Labeling mix-ups | Customers were contacted via phone the week of 3/5/18 and 7/12/18. The customers were informed of the recall and asked any of the affected remained in their possession, if the products have already been implanted, and if any patients experienced any adverse events. |  | 10 | 10 | The products were distributed to the following US states: OH. | Terminated | November 15, 2018 | 139 | 10 | Not_Computer | Not_Computer
Z-3196-2018 | 80632 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | Portex Pulsator Arterial Blood Sampling Kit with Dry Lithium Heparin for Gases and Electrolytes, REF 4080P-2, Sterile, RX. The firm name on the label is Smiths Medical ASD, Inc., Keene, NH. |  |  | July 13, 2018 | 2018 | Smiths Medical ASD Inc. | The product was mislabeled with the incorrect label. | Labeling Change Control | Recall notifications dated 7/13/2018 were issued via FedEx overnight mail on 7/13/2018 explaining the reason for recall, risk to health, and instructing the customer to check their inventory for the affected lot numbers. |  | 29600 | 29,600 kits | US Nationwide Distribution in the states of CT, MA, MI, NC, NJ, NY, RI, TX, UT, and WA. | Terminated | October 19, 2018 | 98 | 29600 | Not_Computer | Not_Computer
Z-0623-2018 | 79092 | No | HBB | Pneumatic cranial drill motor. | Neurology | Neurology | 510(k) | The ELAN 4 Air System is intended for high speed cutting, sawing and drilling of bone in fields of Spine, ENT, Neuro, and Maxillofacial Surgery. |  |  | September 01, 2017 | 2017 | Aesculap Implant Systems LLC | The ELAN 4 Air System was released to the market place prior to an authorized FDA 510K clearance. This issue does not affected sales outside the United States (OUS). | No Marketing Application | On September 1, 2017 an Urgent Medical Device Recall letter was issued to customers for the ELAN 4 Air System which includes the following sub-components:ST0607 ELAN 4 Air Control SetST0608 ELAN 4 Air Neuro Crani SetST0609 ELAN 4 Air Spine SetST06411 ELAN 4 Air HD Spine SetST0614 ELAN 4 Air Micro Saw SetST0615 ELAN 4 Air Hand Control SetST0638 ELAN 4 Air Large Saw SetThe letter directs customers to take the following actions:1.Review the Voluntary Urgent Medical Device Recall Notification Letter in its entirety and ensure that all users in your organization and other concerned persons are informed about this recall and immediately discontinue use. 2.Determine your current inventory of affected products within your facility. Immediately remove the affected products and return to Aesculap Inc. Do not destroy any affected products. 3.Utilizing the attached Product Recall Acknowledgement Response Form on the next page of this letter, record the total number of products in your possession. If you have no inventory remaining, please enter zero (0) on the form and a REASON WHY (example, returned evaluation system to sales representative). 4.Please contact Customer Service and they will provide you with a Returned Goods number and return label.Customer Service: 1-800-282-9000aesculap_emailorders.bbmus_service@bbraun.com |  | 167 | 167 Units | Nationwide Distribution. | Open, Classified |  |  | 167 | Not_Computer | Not_Computer
Z-0624-2018 | 79089 | No | KCT | Sterilization wrap. | General Hospital | General Hospital | 510(k) | JS Series SterilContainer S2 System |  |  | September 08, 2017 | 2017 | Aesculap Implant Systems LLC | The JS Series SterilContainer S2 System was released to the market place prior to an authorized FDA 510K clearance. No other Aesculap SterilContainer System products are impacted by this recall. | No Marketing Application | On August 24, 2017 a Medical Device Recall(Reference 2916714-08/24/2017-010R) was issued requesting that all customers take the following steps: 1.Review the Urgent Medical Device Recall Notification Letter in its entirety and ensure that all users in your organization and other concerned persons are informed about this recall and immediately discontinue use. 2. Determine your current inventory of affected containers within your facility. Immediately quarantine the affected products. A sales representative will come to your facility and remove your affected product and immediately return it to Aesculap Inc. They will assist you with ordering an appropriate replacement container. Do not destroy any affected products.3. Please contact Customer Service and they will provide you a Returned Good number and a return label.Customer Service: Phone: 1-800-282-9000 aesculap_emailorders.bbmus_service@bbraun.com |  | 292 | 292 units | Nationally | Open, Classified |  |  | 292 | Not_Computer | Not_Computer
Z-0412-2018 | 78687 | No | HTW | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | SST TWIST DRILL, STANDARD, 1.6mm Dia. x 127mm, REF 8054-010, STERILE |  |  | May 22, 2017 | 2017 | MicroAire Surgical Instruments, LLC | One lot was found to contain both the correct twist drill (1.6mm) as well as an incorrect twist drill (2.0mm) | Labeling mix-ups | MicroAire Surgical Instruments, LLC distributed recall letters on May 22, 2017 MicroAire Surgical Instruments, LLC distributed recall letters to the customers who received the mislabeled twist. Letters will be sent by UPS to expedite the delivery. For further questions, please call (434) 975-8000 |  | 42 | 42 | US Distribution. | Terminated | August 29, 2018 | 464 | 42 | Not_Computer | Not_Computer
Z-0337-2018 | 78640 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | TRYPTIK(R) 2 C-PLATE SELF-TAPPING POLYAXIAL SCREW 4.0 MM, LENGTH 14 MM, REF TRY-PS-40 14-S, Qty 2 |  |  | November 27, 2017 | 2017 | SPINEART SA | Product labeling for TRYPTIK 2-C-PLATE Self Tapping screw, REF TRY-PS 40, Lot 3-5934 contains the sketch/drawing of a self-drilling screw. | Error in labeling | Spineart sent an Urgent Medical Device letter by email dated November 27, 2017. The letter stated the following:"Out of an abundance of caution", we have decided to recall the concerned products.Please carefully follow instructions detailed below:1. Immediately examine your inventory and quarantine product subject to recall.2. In addition, if you may have further distributed this product, please identify your customers and notify them at once of this product recall by using this document.3. Collect and quarantine all products.4. Send back all products and the enclosed Response Form to Spineart Geneva E-mail: regulatory@spineart.com."For further questions, please call (949) 313-6400.. |  | 120 | 120 units | US Distribution to the state of : CA | Open, Classified |  |  | 282 | Not_Computer | Not_Computer
Z-0623-2019 | 81488 | No | GKA | Abnormal hemoglobin assay. | Hematology | Hematology | 510(k) | Genesys 2 Reagent, REF 01-03-0022 intended to be used with automated high performance liquid chromatography (HPLC) systems employing ion exchange chromatography. For in vitro diagnostic use only. |  |  | November 01, 2018 | 2018 | Clark Laboratories, Inc. (dba,Trinity Biotech USA) | lack of 510K | No Marketing Application | The firm initiated the recall by letter on 11/01/2018. The letter explained the problem and requested the return of the recalled items. |  | 216 | 216 units | IL, WV, NJ, AZ, MD, UT | Open, Classified |  |  | 1317 | Not_Computer | Not_Computer
Z-3197-2018 | 80872 | No | DRF | Electrode recording catheter or electrode recording probe. | Cardiovascular | Cardiovascular | 510(k) | ARROWÂż Bipolar Electrode Catheter with shrouded pins , 6 Fr, 110 cm |  |  | July 30, 2018 | 2018 | Arrow International Inc | Products are labeled with an incorrect expiration date | Error in labeling | Arrow International notified US consignee notification by letter on 30 JUL 2018 via FedEx 2-day mail. Account action: immediately discontinue use and quarantine any products with the product codes and lot numbers listed above, so that the affected products can be returned to Arrow. To return product, complete the enclosed Recall Acknowledgement Form and fax it to 1-855-419-8507, Attn: Customer Service or email to recalls@teleflex.com. |  | 7 | 7 eaches | US in the state of PA | Open, Classified |  |  | 7 | Not_Computer | Not_Computer
Z-0183-2018 | 77415 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | BD Insulin Syringes with the BD Ultra-Fine(TM) needle Âż mL 12.7mm 30G, Catalog/REF 328466 |  |  | May 25, 2017 | 2017 | Becton Dickinson &amp; Company | Product mislabel. Lot 6291768 of the BD Insulin Syringes with the BD Ultra-Fine needle ÂżmL 12.7mm 30G, Cat (Ref) 328466, was found to have some polybags in the lot incorrectly labeled as BD Ultra-Fine needle ÂżmL 8mm 31G, Cat (Ref) 328468. The shelf carton and case carton are correctly labeled as BD Ultra-Fine needle ÂżmL 12.7mm 30G. | Labeling mix-ups | BD sent recall notification letters to all direct consignees (wholesaler, hospital, and retail) on May 25, 2017. BD is requesting wholesalers (distributors) to notify any accounts or additional locations that may have received the recalled product from them. BD gives the option to distributors to provide their customer list to BD if they would like BD to notify their customers. BD is requesting retail pharmacies to notify consumers that may have received the recalled product from them. 8/25/17 Firm notified FDA of recall expansion to include Lot Numbers 6312558 and 6340590. A 2nd notification was sent to consignees in a letter dated Aug 30, 2017 that includes all 3 lots that are in scope of the recall. Most of the direct consignees identified in the first recall action were also identified as affected for the second (expanded) recall action since they received either two or all three recall lots. All affected consignees that received any of the three lots have been combined into one consignee list. These customers will be notified of the expanded recall communication (second recall letter dated Aug 30, 2017). To minimize customer disruption, BD will conduct effectiveness check on the second recall letter only, since it covers the recall action for all three affected lots and combines all affected consignees from both actions (first and second/expanded recall). This will allow us to verify that the affected consignees received the notification and took the necessary actions for all three recall lots without having to contact them multiple times for two different letters that cover the same recall action. Details of the effectiveness are as follows: 1.Letters will be mailed with UPS tracking. The effectiveness check level will be established by the FDA classification letter. 2.A Business Response Card is included with the letter to assist direct consignees in responding to the field action. 3.BD will attempt to contact by phone/email (as applicable) non-respondents |  | 7839340 | 7,839,340 units | Nationwide Distribution including CA, CT, FL, IL, LA, ME, MI, NC, ND, NJ, NY, PA, RI, SC, VA | Terminated | June 27, 2018 | 398 | 7839340 | Not_Computer | Not_Computer
Z-0601-2018 | 78964 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | ARROW PICC powered by ARROW VPS Stylet, Product Codes: CDC-35052-VPS and CDC-35041-VPS. Peripherally Inserted Central Catheter |  |  | July 25, 2017 | 2017 | Arrow International Inc | Finished good kits may contain incorrect components. Kits that should contain a 5 FR 2 lumen catheter may contain a 4 FR single lumen catheter and kits that should contain a 4 FR single lumen catheter may contain a 5 FR 2 lumen catheter. This may cause a delay in treatment while a replacement catheter is obtained. | Packaging process control | Arrow International notified all U.S. consignees with letter sent on July 25, 2017, via FedEx 2-day. second notification August 22, 2017, third notification November 28, 2017. |  | 2691 | 2691 total | US nationwide distribution. | Terminated | August 29, 2018 | 400 | 2691 | Not_Computer | Not_Computer
Z-0721-2018 | 78762 | No | KRO | Knee joint femorotibial metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer Segmental System, ZSS Distal Femur. Sizes as follows: |  |  | August 01, 2017 | 2017 | Zimmer Biomet, Inc. | The products being recalled exceeded the weight previously tested leading to a possibility of damage to the package and/or a compromised sterile barrier. | Package design/selection | The firm, Zimmer Biomet, sent an "URGENT MEDICAL DEVICE RECALL REMOVAL"-Lot Specific letter dated 8/1/2017 to its Customers via email and/or FedEx. The letter describes the product, problem and actions to be taken. The Instructions to Distributors, Sales Representatives, and Distributor Operation Managers included to ensure affected team member are aware of the contents of the notification letter, immediately locate and quarantine affected product in inventory, complete and return the Certification of Acknowledgement form by email to: corporatequality.postmarket@zimmerbiomet.com., return all affected product, and note if there are any hospitals and/or surgeons for which need to be notified. Instructions for Risk Managers and surgeons are to ensure affected personnel are aware of the contents of the notification letter, quarantine any affected product on hand, complete and return the Certificate of Acknowledgement form by email to: corporatequality.postmarket@zimmerbiomet.com., and return any affected product on hand. For questions or concerns, please call the 411 call center at 1-574-371-3071 between 8:00 am and 5:00 pm EST, Monday through Friday or email corporatequality.postmarket@zimmerbiomet.com |  | 3350 | 3,350 units total | Worldwide Distribution: US (nationwide) including states of: AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, SC, SD, TN, TX, UT, VA, WA, WI, WV, WY, and Puerto Rico; and to countries of: Austria, Australia, Bahamas, Brazil, Canada, China, Dominican Republic, England, France, Germany, Israel, Japan, South Korea, Lebanon, Malaysia, Mexico, Netherlands, Panama, Singapore, Spain, Switzerland, and Thailand. | Open, Classified |  |  | 3350 | Not_Computer | Not_Computer
Z-0620-2018 | 79068 | No | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | ELAN 4 Drill: ELAN 4 1-RING TWIST DRILL D1.5 |  |  | September 05, 2017 | 2017 | Aesculap Implant Systems LLC | ELAN 4 burr/blade accessories were marketed without 510(k) premarket notification clearance. | No Marketing Application | Notifications were sent requesting the customers immediately discontinue use and quarantine affected products. The notification explains that a sales representative will visit the facilities to remove affected product and return it to the firm. |  | 6211 | 6211 | Distributed to 4 hospitals in the following states: NM, OH, TX, WV | Open, Classified |  |  | 17195 | Not_Computer | Not_Computer
Z-0403-2019 | 81293 | Yes | NTG |  |  | Orthopedic | PMA | DT Medi Tech Hintermann Series Talar Implant LEFT-SIZE 4 |  |  | September 25, 2018 | 2018 | DT MedTech, LLC | Incorrect package labeling identifying the device | Packaging process control | DT Med Tech LLC initially notified accounts by email notification on 9/25/18 followed by notification letter; The letters explain the process for theidentification and return to HealthLink BV (distribution warehouse in The Netherlands) of the products affected (lot AAADP/AAAEM). |  | 1 | 1 | Australia, Ireland, Korea, Germany | Open, Classified |  |  | 55 | Not_Computer | Not_Computer
Z-0510-2019 | 81460 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Multi-Track Angiographic Catheter (MMTA) |  |  | October 17, 2018 | 2018 | Numed Inc | Instructions for Use booklets were not included on the outerpouch | Labeling design | NuMED issued recall notification letter and response form on 10/17/2018 via email to the one distributor and followed up by a hard copy of the letter and response form sent via UPS Next Day Air. The letter identified the affected product, problem and actions to be taken. The customer was instructed to return affected product.. |  | 100 | 100 units | US distribution in the state of PA | Terminated | January 28, 2019 | 103 | 100 | Not_Computer | Not_Computer
Z-0674-2018 | 79161 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Vanguard Complete Knee System, Vanguard CR Lipped Tibial Bearing 10x71/75, Item Number: 183540 |  |  | September 29, 2017 | 2017 | Zimmer Biomet, Inc. | Knee implant components (tibial bearings) are labeled with the incorrect size. | Labeling mix-ups | Zimmer Biomet sent an Urgent Medical Device Recall letter dated September 29, 2017 to their customers. The letter identified the affected product, problems and actions to be taken. Customers were instructed to remove the affected product from inventory and provide to their Zimmer sales representative or, if the consignee is a distributor, to ship the affected product back to Zimmer. Consignees were also asked to complete a recall response form and send this back to the firm. Returns are being physically destroyed.If you have further questions or concerns after reviewing this notice, please call customer service at 574-371-3071 between 8:00 am and 5:00pm EST, Monday through Friday. Calls received outside of call center operating hours will receive a voicemail prompt or be transferred to an on-call representative in the event of an emergency. Alternatively, your questions may be emailed to CorporateQuality.PostMarket@zimmerbiomet.com. |  | 24 | 24 | Worldwide Distribution - US Nationwide in the states of CA, GA, KS, LA, MA, MI, MN, OH, TX, VA, WA, WI, and the countries of Netherlands | Open, Classified |  |  | 24 | Not_Computer | Not_Computer
Z-2827-2018 | 80482 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Tosoh Bioscience AIA Analyte Application Manual (AAM) CD |  |  | June 05, 2018 | 2018 | Tosoh Bioscience Inc | Distribution of unapproved product labeling | Error in labeling | The firm, Tosoh Bioscience, notified customers via a phone call on May 29, 2018 and May 31, 2018, reference "Product Labeling-Removal". The phone call described the product, problem and actions to be taken. The firm instructed to confirm receipt of the CD containing AAMs; return the unapproved labeling; if printouts were made of any information on the CD then destroy it, if not, thank customer for response and complete the Customer Call form.The firm will provide the Customers updated CD following approval and release.Any further questions, contact Sr. QA Manager at 614-317-1909 or email: Levent.durukan@tosoh.com. |  | 9 | 9 CDS | US Distribution to states of: FL, MA, MT, OK, OR and WI. | Open, Classified |  |  | 9 | Not_Computer | Not_Computer
Z-2823-2018 | 80651 | No | FJS | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Flex-Neck Classic Peritoneal Dialysis Catheter (Pediatric) |  |  | December 20, 2017 | 2017 | Merit Medical Systems, Inc. | Mix-up between adult and pediatrics PD catheter. | Packaging process control | Communication to consignees began on December 20, 2017.Written notification to Merit Accounts included product identification information,instructions to immediately quarantine any devices and discontinue use and return unopened product to Merit.Written notification to Merit Sales Representatives included product identification information and instructions to immediately return remaining Peritoneal Dialysis Catheter devices.An email notification was also sent to the Merit sales representatives of the affected accounts. The sales representatives were provided a copy of the Customer Response Form(s) applicable to their account(s). The sales representatives were instructed to assist their account(s) with the completion of the Customer Response Form(s) and the return of the affected units back to the Merit Field Assurance Department. |  |  |  | Worldwide Distribution - US Nationwide in the states of AL, CA, DE, FL, KY, MA, MO, NC, NY, OH, PA, and UT and the countries of Germany | Open, Classified |  |  |  | Not_Computer | Not_Computer
Z-0399-2019 | 81185 | No | KWY | Hip joint femoral (hemi-hip) metal/polymer cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | StageOne Hip, Models 431207 &amp; 431209Usage: |  |  | September 21, 2018 | 2018 | Zimmer Biomet, Inc. | There is a potential for comingling of the 43MM and 51 MM molds | Error in labeling | On September 21, 2018, Zimmer Biomet issued URGENT MEDICAL DEVICE RECALL notices and Return Certification Forms to customers via courier service. Distributors are encouraged to take the following actions:Review this notification and ensure that affected team members are aware of the contents.2. Immediately locate and quarantine affected product in your inventory.3. Immediately return all affected product from your distributorship and from affected hospitals within your territory.a. Complete the Return Certification Form and send to CorporateQuality.PostMarket@zimmerbiomet.com within three (3) days.b. For each return, send a copy of Return Certification Form to CorporateQuality.PostMarket@zimmerbiomet.com.c. Include a hardcopy of the Return Certification Form in each carton of your return shipment for immediate processing.d. Mark RECALL on the outside of the returned cartons.Return the Additional Accounts form to CorporateQuality.PostMarket@zimmerbiomet.com.a. Review the list of hospitals included with the email notification sent to your facility, which includes a list of hospitals that have already been notified of this recall.b. Identify whether there are any additional hospitals that Zimmer Biomet has not notified and list these accounts on the Additional Accounts form. Please provide the form in Excel format.c. If there are no additional accounts to notify, please indicate that there are no additional accounts, or indicate None or NA on the form.5. Retain a copy of your Inventory Return Certification and product return forms for your records in the event of a compliance audit of your facility.Risk Managers are encouraged to take the following actions:Review this notification and ensure that affected personnel are aware of the contents.2. If you have affected product at your facility, assist your Zimmer Biomet sales representative and quarantine all affected product. Your Zimmer Biomet sales representative will remove |  | 72 | 72 units | US Nationwide Distribution in the states of: AZ, FL, MI, OH, VA &amp; WI | Open, Classified |  |  | 72 | Not_Computer | Not_Computer
Z-0400-2018 | 77964 | Yes | NVZ |  |  | Cardiovascular | PMA | MyCareLink Smart Patient Monitors. |  |  | July 14, 2017 | 2017 | Medtronic Inc., Cardiac Rhythm and Heart Failure | Patients monitored on two (2) or more implanted Medtronic heart devices in the Medtronic CareLink Network may have potential impact on the ability to remotely monitor the patient's heart devices. Potential impacts could lead to missed CareAlert notifications or device reports. | Error in labeling | Customers were notified via letter on 07/14/2017. Medtronic communicated that an updated Clinician Instructions for Use has been released with the following updated information: In patients with more than one Medtronic implanted hear device, device data may not be able to transmit from any of the implanted devices to the physician via the Monitor; The patient could potentially transmit device data unintentionally from one of the implanted heart devices when attempting to transmit from another device implanted in the same patient; and a Medtronic Programmer can be used to obtain data from each implanted heart device. Instructions included to ensure all appropriate personnel are notified of this correction. |  | 1519984 | 1,519,984 enrollments total | Domestic: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, Internationally: Austria, Belgium, Brunei, China, Czech Republic, Denmark, Finland, France, Germany, Greece, Hong Kong, Hungary, Iceland, India, Ireland, Israel, Italy, Japan, Kazakhstan, Latvia, Lithuania, Malaysia, Malta, Netherlands, Norway, Philippines, Poland, Portugal, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, UAE, United Kingdom | Open, Classified |  |  | 1519984 | Not_Computer | Computer
Z-2904-2018 | 80371 | No | GWO | Preformed alterable cranioplasty plate. | Neurology | Neurology | 510(k) | QUIKFLAP, 3 X 2-HOLE DOG BONE PLATES WITH 1.5X4MM SELF DRILLING SCREWS, Part Number 12-01530S, for use on the cranium |  |  | June 22, 2018 | 2018 | Stryker Leibinger GmbH &amp; Co. KG | After a three year term, there is a potential for the external packaging (sterile barrier) to become compromised. | Package design/selection | URGENT MEDICAL DEVICE RECALL letters dated 6/22/18 were sent to customers with the following requested actions to be taken:1. Please inform users of this Medical Device Removal and pass this notice to all appropriateindividuals within your organization.2. Return all affected products available at your location to:Stryker CMFAttn: Recall Coordinator6300 Sprinkle RdPortage, MI 49002REF: PFA #1739180 Or, Contact Stryker Customer Service and refer to PFA #1739180 with any questions about returning the product to us.3. Complete and sign the enclosed Business Reply Form and fax a copy to: (877) 648-7114 or email a copy to CMF-custserv@stryker.com.4. Keep a copy of the completed and executed Customer response Form for your records. |  | 22831 | 22,831 total | US distribution in the states: AZ, FL, NH, and OH, Australia, Canada, France, New Zealand, and UK. | Open, Classified |  |  | 22831 | Not_Computer | Not_Computer
Z-2919-2018 | 80684 | No | LKO | Antinuclear antibody immunological test system. | Immunology | Immunology | 510(k) | is a quantitative/qualitative enzyme-linked immunosorbent assay (ELISA) for the detection of autoantibodies specific for Anti-Ro (SS-A) antigen in human serum or plasma. EDTA, lithium heparin, citrated plasma. It is intended to aid in the diagnosis of systemic rheumatic diseases, particularly SjÂżgren's syndrome and is not definitive in isolation. Autoantibody levels represent one parameter in a multi-criterion diagnostic process. |  |  | October 21, 2016 | 2016 | Euro Diagnostica AB | Incorrectly labeled Manufactured by DiaSorin, The correct term should be Distributed by DiaSorin. | Error in labeling | none |  | 4 | 4 | MN | Open, Classified |  |  | 20 | Not_Computer | Not_Computer
Z-0347-2018 | 78716 | No | NRY | Percutaneous catheter. | Cardiovascular | Neurology | 510(k) | Trevo XP ProVue System kit containing 1) Trevo XP ProVue Retriever, 4x20mm and 2) Trevo Pro18 Microcatheter |  |  | November 28, 2017 | 2017 | Stryker Neurovascular | Stryker Neurovascular has become aware that at least 1 (one) Trevo XP Provue System was manufactured where the carton sleeve did not match the physical contents of the pack. | Packaging process control | The affected consignees will be notified via letter sent a courier beginning 11/28/17. A response form has been provided which customers have been asked to complete, and correct affected product is being requested. |  |  |  | OUS only - NL, FR, IT, DE, BE | Open, Classified |  |  |  | Not_Computer | Not_Computer
Z-0490-2018 | 78765 | No | DWE | Cardiopulmonary bypass pump tubing. | Cardiovascular | Cardiovascular | 510(k) | Hemofiltration Standard Pack including BC 60 Plus-T 8036. Part Number: 701066990. |  |  | May 22, 2017 | 2017 | Maquet | 55 kits were distributed after their expiration date of September 2015. | Incorrect or no expiration date | Medical Device Recall (Removal) Notices were sent to customers via FedEx on 5/22/2017. The letters identified the affected device and the issue. Users are asked to examine their inventory to determine if any of the affected product is on hand. If found, affected devices should be removed and quarantined. Customers should contact Maquet/Getinge Customer Support at 1-888-627-8383 to obtain instructions on returning the affected devices. The Medical Device Recall Removal Response Form provided should be completed and returned. Questions should be directed to your Maquet/Getinge representative or call the Maquet/Getinge Customer Support at 1-888-627-8383. |  | 55 | 55 units | Distributed to the states of NM, CA, and IL. | Terminated | July 12, 2018 | 416 | 55 | Not_Computer | Not_Computer
Z-1132-2018 | 79276 | Yes | KPO | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | NxStage PureFlow B Solution - Premixed Dialysate for Hemodialysis with NxStage System One |  |  | November 10, 2017 | 2017 | NxStage Medical, Inc. | Using different fluid formulations of NxStage PureFlow B Solution at the same time during treatment may cause patient health risk. The NxStage PureFlow B Solution is for use with systems that use one premixed dialysate during hemodialysis. The original labeling did not warn against mixing dialysate formulations and did not instruct staff to verify that the fluid is flowing during use. | Labeling design | Customers are advised to :1. Always make sure when using PureFlow B Solution that fluid is flowing freely from all bags.2. Recognize that any practice of hanging more than one fluid type at the same time is not an intended use of the NxStage System One" (as it does not monitor fluid composition) and is not a labeled indication of PureFlow B solutions.3. Please make sure that all users of PureFlow B Solution in your organization and other concerned persons are informed regarding this Field Safety Notice.4. Complete the enclosed customer reply form and return it to NxStage. Returning the customer reply form promptly will confirm your receipt of this notification and prevent you from receiving repeat notices. |  | 1243400 | 1243400 units | US nationwide distribution | Open, Classified |  |  | 1243400 | Not_Computer | Not_Computer
Z-1479-2018 | 79591 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Medtronic DxTerity Diagnostic Catheter, JL 4.0, REF DXT5JL40, marked with 6F color code on box. |  |  | March 13, 2018 | 2018 | Medtronic Vascular | The color coded French size indicator on the outer box may incorrectly reflect a 6 French size, rather than the correct 5 French size. The catheters inside the carton are the correct DXT5JL40 item, the inner pouch has correct 5F labeling, information and configuration details. | Packaging | Medtronic initiated consignee communications on 13-Mar- 2018 to Risk Managers and Health Care Professionals. US notifications were sent via 2-Day UPS delivery. OUS communications were sent per locally approved methods.Consignees are asked to quarantine and return any unused potentially affected product to Medtronic. |  | 855 | 855 units | US and AUSTRIA, Canada, Germany, ROMANIA, SLOVENIA, | Open, Classified |  |  | 855 | Not_Computer | Not_Computer
Z-1011-2018 | 79250 | No | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Vanguard CR Porous Femoral 62.5mm (Left Femur) |  |  | October 03, 2017 | 2017 | Zimmer Biomet, Inc. | The affected products are labeled as a 62.5mm LEFT Femur; however, the contents are a Vanguard 70mm RIGHT Femur. | Error in labeling | The firm, Zimmer Biomet, sent an "URGENT MEDICAL DEVICE RECALL REMOVAL" letter dated October 3, 2017 to distributors and customers. The letter described the product, problem and actions to be taken. The customers were instructed to:1. Review notification and ensure affected team member are aware of the contents2. Immediately locate and quarantine affected product in your inventory.3. Complete the Certification of Acknowledgement portion of Attachment 1  Inventory Return Certification Form via digital copy to: corporatequallity.postmarket@zimmerbiomet.com within three (3) days.4. Immediately return all affected product from your distributorship and affected hospitals within your territory along with a completed Attachment 1  Inventory Return Certification Form to Zimmer Biomet via SMS, via email to: rgarequest@zimmerbiomet.com-specify "RECALL" as RGA type.5. Identify additional accounts Zimmer Biomet has not notified.6. Retain a copy of your recall acknowledgement and product return forms for your records.Questions or concerns can be directed to the customer call center at 1-574-371-3071 between 8:00 am and 5:00pm EST, Monday through Friday. |  | 8 | 8 | US Distribution to states of: FL, WI, MD, MO, and MI. | Terminated | August 23, 2018 | 324 | 8 | Not_Computer | Not_Computer
Z-2423-2018 | 80291 | No | FRO |  |  | General &amp; Plastic Surgery | 510(k) | Hyalomatrix, Product nos. 651105, 651103, 651104, and 651301 |  |  | May 10, 2018 | 2018 | Anika Therapeutics, Inc. | Hyalomatrix products are susceptible to degradation as a result of the decomposition of PVC in the packaging material during sterilization, which resulted in degradation of the film. Over the long term it is possible that product integrity may be compromised. | Package design/selection | On May 10, 2018, the firm notified its affected distributor via emailed memorandum of its voluntary recall. The distributor was informed of the affected lots and quantities, and asked to discontinue distribution and sales activities related the affected products. |  | 1711 | 1711 | US Nationwide Distribution in the state of IL. | Open, Classified |  |  | 1711 | Not_Computer | Not_Computer
Z-0950-2018 | 79195 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Amplatz Ultra Stiff Wire Guide |  |  | October 13, 2017 | 2017 | Cook Inc. | Label does not state that the product is heparin-coated. | Labeling Change Control | On October 31, 2017 an URGENT MEDICAL DEVICE CORRECTION was issued to customers requesting the following, including a corrective action:Step 1:Determine if your product isincluded in the affected lotranges by comparing your lotnumber to the listingprovided." If your product is notaffected, complete theAcknowledgement andReceipt Form." If your product is affected,proceed to Step 2.Step 2:Remove a Heparin Coated label from the label sheetprovided and apply to the topright corner of your affectedunit.o If you need additionallabels, please contactStericycle at855.215.4967.If you would like assistance,representatives are available to supportyou. To request assistance, pleasecontact Stericycle at 855.215.4967 |  | 12462 | 12,462 | Nationally | Open, Classified |  |  | 1424570 | Not_Computer | Not_Computer
Z-0945-2018 | 79051 | orthosis, limb brace | IQI | Limb orthosis. | Physical Medicine | Physical Medicine | 510(K) Exempt | FUTURO (TM) SPORT, QUICK STRAP ANKLE SUPPORT, REF 47736GEN, UPC 0 51131 19692 6 |  |  | January 31, 2018 | 2018 | 3M Company - Health Care Business | 3M discovered that the Futuro" Quick Strap Ankle Brace (Cat. #47736EN) and ACE" Brand Deluxe Ankle brace (Cat. #207736) do not have the correct labeling. The products contain natural latex rubber, but do not include the appropriate caution statement. | Error in labeling | The firm initiated their recall by letter on 01/30/2018. The letter stated the following:"Actions required:1.Immediately discontinue sale of Futuro" Quick Strap Ankle Brace.2.Isolate or quarantine any stock of the Futuro" Quick Strap Ankle Brace.3.Call 1-866-746-2524, option #2, Monday to Friday, 8:00 am  4:30 pm Central time, to arrange for the return of the Futuro" Quick Strap Ankle Brace, Cat# 47736EN.4.If you have further distributed the Futuro" Quick Strap Ankle Brace, please notify any of your customers at once. Please include a copy of this letter in your communications.We apologize for any inconvenience this may cause you. Should you have any questions regarding the product and return process, please feel free to contact 3M Customer Service at 1-866-746-2524, option 2, Monday to Friday, 8:00 am 4:30 pm Central Time." |  | 158179 | 158,179 units | US and Australia, Jordan, Republic Of Korea, Lebanon, New Zealand, Oman, Qatar, Taiwan, United Arab Emirates | Open, Classified |  |  | 320527 | Not_Computer | Not_Computer
Z-0939-2018 | 79397 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Trabecular MetalÂż Primary Hip Prosthesis; Part Numbers: 1. 00-7864-013-00 (UDI: (01) 00889024139251 (17) 250731 (10) 63703649), and 2. 00-7864-013-20 (UDI: (01) 00889024139268 (17) 250731 (10) 63703681) |  |  | January 03, 2018 | 2018 | Zimmer Biomet, Inc. | Two lots of hip implants are mislabeled. | Labeling mix-ups | DISTRIBUTORS: 1. Review this notification and notify affected team members of the contents. 2. Immediately locate and quarantine affected product. 3. Immediately return all affected product from your distributorship and from affected hospitals within your territory. a. Complete Inventory Return Certification Form and send to CorporateQuality.PostMarket@zimmerbiomet.com within three (3) days. Mark RECALL on the outside of the returned cartons. 4. Return the Additional Accounts form to CorporateQuality.PostMarket@zimmerbiomet.com RISK MANAGERS: 1. Review this notification and notify affected personnel of the contents. 2. Quarantine all affected product. Your sales representative will remove the affected product from your facility. 3. Complete Certificate of Acknowledgement and send to CorporateQuality.PostMarket@zimmerbiomet.com. |  | 24 | 24 | Distributed in 2 US states: LA and NY. Distributed in Canada, China, Japan, and Switzerland. | Open, Classified |  |  | 24 | Not_Computer | Not_Computer
Z-2435-2018 | 80394 | Yes | NPT |  |  | Cardiovascular | PMA | Edwards SAPIEN 3 Transcatheter Heart Valve (THV), Model: 9600TFX, Size: 26 mm |  |  | June 11, 2018 | 2018 | Edwards Lifesciences, LLC | SAPIEN 3 THV was packaged with the incorrect valve size with respect to the label. | Labeling False and Misleading | On 06/11/2018, the firm sent an Urgent Product Recall notice, via sales representative hand delivery, to the customer requesting they return the customer response form and the affected device. The customer was advised to call Customer Service at 1-800-424-3278, if they had any additional questions. |  | 1 | 1 Valve | NJ | Open, Classified |  |  | 1 | Not_Computer | Not_Computer
Z-1544-2018 | 79769 | general purpose reagent | PPM | General purpose reagent. | Pathology | Pathology | 510(K) Exempt | BD MAX DNA MMK (SPC) For Laboratory Use, catalog no. 442829 |  |  | March 02, 2018 | 2018 | Becton Dickinson &amp; Co. | The package insert incorrectly lists the volume of the primer and probe diluent to use for the reconstitution of the master mix tube. Adding the incorrect amount of diluent could have an effect on the assay. | Error in labeling | On March 2, 2018, the firm sent Urgent Medical Device Product Advisory notices to customers via UPS. The notice identified the errors and the correct information, as well as potential health risks. Customers were asked to take the following actions: 1) Share the information with all users of the product within their facility to ensure awareness; 2) Complete the customer response form and return to the recalling firm; and 3) Report any associated adverse health consequences to the recalling firm.The recalling firm is not asking customer to send back product or discard. The recalling firm will provide users with an updated insert when available.BD Customer/Technical Support: 800-638-8663, Monday-Friday, 8:30AM-5PM (EST) for US customers. |  | 13898 | 13,898 | Distributed to accounts in 8 states: CA, FL, NJ, NY, OR, PA, VA, and WI. International distribution to Argentina, Australia, Brazil, Canada, Chile, China, Colombia, Europe, Hong Kong, Japan, Singapore, Taiwan, and Thailand. | Terminated | July 02, 2018 | 122 | 13898 | Not_Computer | Not_Computer
Z-1268-2018 | 79074 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | c6000, Chemistry (Photometric, Discrete), for clinical use. - cobas e 601 module: Immunoassay Analyzer, Catalog Numbers: 04745922001, 05036348001, 05860652001, 04745922692, 05036348001 &amp; 05036348692 |  |  | December 18, 2017 | 2017 | Roche Diagnostics Corporation | Complaints have been received concerning questionable results on analyzers due to customers utilizing 13 sample tubes without tube adapters. | Labeling False and Misleading | Roche Diagnostics Corporation issued Urgent Medical Device Correction (UMDC) notices and Fax Forms to their customers dated 12/18/2017 via UPS Ground (receipt required). Customers are encouraged to do the following Mandatory Steps to improve the vertical alignment and handling of 13 mm sample tubes and false bottom tubes: FOR cobas e 411 &amp; cobas e 601 analyzers and MODULAR ANALYTICS E 170 module:1) Always use Roche rack cup adapters with 13 mm tubes2) Ensure sample tubes are straightFOR cobas e 411 disk system:1) Always use 13 mm Sample Disk Tube Adapters (SDTAs) with 13 mm sample tubes.2) Ensure sample tubes are straight.3) Ensure 13 mm SDTAs are not exposed to direct sunlight.4) If a larger diameter tube is left inserted in a 13 mm SDTA for more than 3 days, the adapter may become distorted. Do not insert smaller diameter tubes into distorted adapter as they may not aligncorrectly.**All customers are encouraged to DISCARD Analyzer Bulletin 10-009U3.Follow the care and handling instructions for using adapters as outlined in this Urgent Medical Device Correction (UMDC) and complete &amp; return the UMDC fax form via fax to 1-877-271-1472 |  |  |  | US Nationwide | Open, Classified |  |  |  | Not_Computer | Not_Computer
Z-2445-2018 | 80312 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | ADVANSYS MLP/DLP, Model Numbers: 181051S, 181052S, 181041S, 181042S, 181031S, 181032S, 181033S, 181021S, 181022S &amp; 181023S |  |  | May 22, 2018 | 2018 | NewDeal SA | Use of the impacted product may cause a superficial infection requiring PO antibiotics and wound care or deep infection requiring IV antibiotics and device removal. | Packaging change control | On May 22, 2018 Newdeal SAS, a company of Integra LifeSciences, issued URGENT VOLUNTARY MEDICAL DEVICE RECALL notices to their customers. On October 22, 2018 URGENT VOLUNTARY MEDICAL DEVICE RECALL notices were issued to additional customers. Customers were advised to take the following actions:1. Examine your inventory to determine if you have any affected lots in original product packaging identified within the notice.2. If you do have any of the affected lots in original product packaging, stop using the product immediately, remove them from service and place in quarantine until the product can be returned.3. Complete the Acknowledgement and Return Form and check the box: I do have affected products on the list and record the lot number.4. Or, complete the attached Acknowledgement and Return Form and check the box: I do not have affected products.5. Complete the other information on the form as indicated on the form. Keep a copy of the form for your records.Return the completed Acknowledgement and Return Form by email or fax indicating your receipt and review of the notification.Customers with questions regarding the instructions, please contact Customer Service at 1-800-654-2873 Monday to Friday 8:00 AM  8:00 PM EST or custsvcnj@integralife.com . |  |  |  | CA, CO, CT, ID, KS, MA, MD, MN, MO, MS, NC, NY, OH, OR, PA, SD, VA, WA &amp; WI | Open, Classified |  |  |  | Not_Computer | Not_Computer
Z-0997-2018 | 79339 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Pressure Injectable Multi- Lumen Central Venous Catheterization Kit with Blue FlexTip(R), ARROWg+ard Blue PLUS(R) |  |  | November 27, 2017 | 2017 | Arrow International Inc | Some of the products listed in the recall may be missing a sticker that contains the lot number, expiration date, and Unique Device Identification (UDI) information. | Error in labeling | On November 29, 2017 an Urgent Medical Device Correction notice was issued to customers stating the following:Our records indicate that you have received products that are subject to this action. We are now notifying our distributors to take the following actions:1.Immediately discontinue distribution and quarantine any products with the product codes and lot numbers listed above.2.Inspect affected products within your control to identify whether they have the sticker with the lot number, expiration date, and UDI information. Products missing this sticker should be returned to Arrow International so they are not distributed to health care providers in error. Products which are properly labeled should not be returned.3.Using the provided Customer Letter and Acknowledgement Form templates, communicate this correction to any of your customers who have received products included within the scope of the correction. Customers should return mislabeled products. Products that are properly labeled should not be returned.4.If any of your customers have mislabeled products to return, have them return the affected products to you, together with a completed Acknowledgement Form, for consolidation and return to Arrow International. In the event that an alternative approach is needed, contact Customer Service for more information at 1-866-246-6990.5.To return products to Arrow International, complete the enclosed Acknowledgement Form and fax it to 1-855-419-8507, Attn: Customer Service or email it to recalls@teleflex.com. This will allow us to document the amount of product you have on hand for return. A customer service representative will contact you with a Return Goods Authorization (RGA) Number and will provide instructions for the return of product to Arrow International.6.If you and your customers have no affected stock or if the products in your stock are properly labeled, please complete the enclosed Acknowledgment Form and fax it to 1-855-419 |  | 23904 | 23,904 in total | US in the states of OH, FL, SC, TX, KY, NJ, MI | Open, Classified |  |  | 23904 | Not_Computer | Not_Computer
Z-1179-2018 | 79297 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | STAGE-1Âż RDS COC Abutment Plastic Sleeve, Non-Lock, Multi-Unit; Catalog Number: S2437-01-1K |  |  | October 16, 2017 | 2017 | Keystone Dental Inc | A locking plastic abutment sleeve, used by dental technicians in dental laboratories to aid in the creation of the final restorative crown, was packed as a non-locking plastic abutment sleeve. | Labeling mix-ups | The firm, Keystone Dental, sent an "URGENT MEDICAL DEVICE RECALL" letter dated 10/24/217 to its consignee. The letter described the product, problem and actions to be taken. The consignee was instructed to do the following: Return any product remaining in inventory to Keystone Dental and replacement product will be sent to you. Note: These lots were the only affected product and all other inventory you may have can be used. Complete and return the Customer Relations Fax Back Form to Customer Relations (1-866-903-9272-US only/ 1-781-328-3400).If you have any questions, please contact Customer Relations Manager at:Toll Free 866-902-9272; Fax 866-903-9272 and CustomerSupport@KeystoneDental.com. |  | 55 | 55 | International Distribution to: Germany only. | Open, Classified |  |  | 55 | Not_Computer | Not_Computer
Z-1542-2018 | 79549 | No | LFH | Tricyclic antidepressant drugs test system. | Toxicology | Toxicology | 510(k) | Alere iScreen Dx Urine Drug Screen Card |  |  | December 15, 2017 | 2017 | Ameditech Inc | The kit box label incorrectly identifies Phencyclidine (PCP), where it should list Tricyclic Antidepressants (TCA) as one of the 10 drug analytes included in the kit box. | Labeling Change Control | Customers were notified on about 12/15/2017. Instructions included to become aware of the recall so as to relay the information to customers with questions. |  | 828 | 828 kits (20,700 devices) | Distributed domestically to AZ, FL, IL, ND, NV, PA, TN, TX, VA. | Terminated | February 11, 2019 | 423 | 828 | Not_Computer | Not_Computer
Z-2469-2018 | 80302 | No | HSD | Shoulder joint humeral (hemi-shoulder) metallic uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Fx SOLUTIONS OFFSET HEAD CoCr 50x20, REF 106-5000 |  |  | June 04, 2018 | 2018 | Fx Solutions | Breach in the external blister of certain lots of Humelock Offset Head 50 X 20 and Centered Head 50 X 19. | Package design/selection | The firm initiated the recall by email on 06/04/2018 in France, on 06/05/2018, to distributors in the US and elsewhere. The distributor (US Agent) sent the notice on to the division agents who store the product. The firm requested return of the units. |  | 145 | 145 units | Worldwide Distribution - US Nationwide and the countries of: France, Argentina, South Africa, Spain, Israel, Germany, Lebanon, Italy, Chile, Belgium, Greece, Mexico, Poland, Hong Kong, Switzerland | Open, Classified |  |  | 276 | Not_Computer | Not_Computer
Z-2261-2018 | 80181 | No | JDR | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | P28 JAWS Nitinol Staple System kits are designed for use in an osteotomy, arthrodesis and fragment fixation of bones and joints of the foot including fixation of small bone fragments. |  |  | December 19, 2017 | 2017 | Paragon 28, Inc. | Missing drill guide as stated in the package. | Error in labeling | The firm, Paragon 28 Inc. sent an "MEDICAL DEVICE RECALL" letter dated Dec 19, 2017 to their consignees. The letter describes the product, problem and actions to be taken. The consignees were instructed to immediately review their inventory to determine if they have the devices; if you have possession of any of the devices, you must ship entire kit back to Paragon 28 Inc., 96 Inverness Drive, E. ste 1; Englewood CO 80112. If you have any questions, contact Senior Quality Manager at 720-643-1301 or email to: epohl@paragon28.com. |  | 1626 | 1,626 | Worldwide Distribution: US (nationwide) to states of: AL, AZ, CA, CO, DC, FL, GA, ID, IL, IN, KS, KY, LA, MD, ME, MI, MN, MO, NC, NJ, NV, NY, OH, OR, PA, SC, TN, TX, VA, WA, and WI; to countries of: Australia and South Africa. | Open, Classified |  |  | 1626 | Not_Computer | Not_Computer
Z-2448-2018 | 80306 | No | GAW | Nonabsorbable polypropylene surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | PROLENE SUTURE MONO BLUE - BV1755 Taper Point 3/8C, D10031 |  |  | May 07, 2018 | 2018 | Ethicon, Inc. | The needle associated with certain suture product codes has the potential to perforate the packaging due to a packaging design issue. A perforated package has the potential to cause a soft tissue injury to a healthcareprofessional handling the subject product or a patient infection due to a breach of sterility within the packaging ifit goes unnoticed by the healthcare professional. | Package design/selection | On May 24, 2018, the firm distributed Urgent: Medical Device Recall letters to affected U.S. customers.Customers were instructed to do the following:1. Examine your inventory immediately to determine if you have product subject to this recall (removal) onhand and quarantine such product(s).2. Remove the product subject to this recall (removal) and communicate the issue to relevant operatingroom or materials management personnel, or anyone else in your facility who needs to be informed.3. If any product subject to this recall (removal) has been forwarded to another facility, contact that facilityto arrange return.4. Complete the Business Reply Form (BRF) confirming receipt of this notice andfax or email it to Stericycle at 1-888-473-8015 or Ethicon3681@stericycle.com within three (3) businessdays. Please return the BRF even if you do not have product subject to this recall (removal).5. Keep this notice visibly posted for awareness until all product subject to this recall (removal) has beenreturned to Stericycle. While processing your returns, please maintain a copy of this notice with theproduct subject to this recall (removal) and keep a copy for your records.6. Customers are required to return all unused D-Special PROLENEÂż Polypropylene Suture &amp; D-SpecialPRONOVAÂż Poly (Hexafluoropropylene - VDF) Suture products that are in their inventory immediately.Only D-Special PROLENEÂż Polypropylene Suture &amp; D-Special PRONOVAÂż Poly (Hexafluoropropylene- VDF) Suture product subject to this recall (removal) returned by September 21, 2018 will be eligiblefor credit reimbursement. Any product subject to this recall (removal) returned after September 21,2018 will not be eligible for credit reimbursement.7. To return product subject to this recall (removal), photocopy the completed BRF, place it in the box withthe product, and affix the pre-paid authorized shipping label included with the recall notification letter.Ethicon will pay for |  | 1752 | 1752 | Accounts in AZ, CA, CO, FL, GA, IL, IN, KS, MA, MN, MO, NC, NE, NJ, NY, OH, OR, SC, SD, TN, TX, VA, WA, WI, WV. | Open, Classified |  |  | 43092 | Not_Computer | Not_Computer
Z-1444-2018 | 79614 | No | LZD | Total temporomandibular joint prosthesis. | Dental | Dental | PMA | Patient-Fitted Temporomandibular (TMJ) Right Mandibular Component part of TMJ Reconstruction Prosthesis System. Model /catalog # TYY-NNNNRM |  |  | March 12, 2018 | 2018 | TMJ Solutions Inc | This device was mislabeled and is not the correct component for this patient. | Error in labeling | TMJ Concepts sent an Medical Device Recall Communication letter dated March 12, 2018.INSTRUCTIONS:The implants for this case should not be used and should be immediately quarantined. We request that you return the entire case (the anatomical bone model together with both the fossa and mandibular components) to us by FedEx Priority Overnight Delivery. Any bone screws and drills that were provided with the shipment can be retained at your hospital and do not need to be returned. For further questions, please call (805) 650-3391. |  | 1 | 1 | US Distribution to GA. | Open, Classified |  |  | 1 | Not_Computer | Not_Computer
Z-1417-2018 | 79652 | Yes | DWC | Cardiopulmonary bypass temperature controller. | Cardiovascular | Cardiovascular | 510(k) | Sarns TCM, Catalog No. 15747 |  |  | March 16, 2018 | 2018 | Terumo Cardiovascular Systems Corporation | Update to cooler-heater cleaning instructions. | Labeling design | The firm,Terumo Cardiovascular Systems Corporation, mailed customers an "Urgent Medical Device Correction" letter dated 2/16/18. The letter described the product, problem and actions to be taken. The Customers were instructed to do the following: 1) discard the current Sarns TCM, TCM II Cooling and Heating, or tHX2 Temperature Management Systems cleaning guides; 2) Review the notice and updated cleaning guides, and ensure all users receive notice of the updated cleaning guides and; 3) Return the Customer Response Form acknowledging receipt of the notice via Terumo Recall Email: tcvs.recall@terumomedical.com or Terumo Recall Fax: 734-741-6149. The letter also includes a list of the updates to the cleaning guides.Customers with questions may contact Terumo CVS Customer Service (800) 521-2818, Monday-Friday, 8am-6pm ET or Fax (734)741-6149. |  | 2393 | 2393- total for all products | Worldwide Distribution: US (Nationwide), Australia, Belgium, Canada, Chile, China. Colombia, Dominican Republic, England, Finland, France, Germany, Greece, HK, India, Indonesia, Iran. Israel, Italy, Japan, Korea, Mexico, Netherlands, New Zealand, Norway, Philippines, Russia, Saudi Arabia, Singapore, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates,and Vietnam. | Open, Classified |  |  | 2393 | Not_Computer | Not_Computer
Z-1427-2018 | 79531 | No | BSY | Tracheobronchial suction catheter. | Anesthesiology | Anesthesiology | 510(K) Exempt | Halyard Closed Suction System for Adults, Double Swivel Elbow (DSE). Used to aspirate liquids or semisolids from a patient's upper airway. Labeled as the following: |  |  | February 07, 2018 | 2018 | Halyard Health, Inc | Certain Flex Connectors supplied with Halyard Closed Suction Kits with Flex Connector may become loose or disconnect before use or during use. | Labeling design | Customers were notified via mail and/or email on about 02/07/2018. Instructions for distributors included to notify all affected customers who may have received potentially impacted products and to complete and return the Distributor Acknowledgement Form. Instructions for customers included to ensure adequate connection of the Flex Connector to Swivel Connector before use of the product by following. This accomplished by ensuring a tight fit is achieved on the Swivel with the Flex Connector advanced approximately three-quarters over the Swivel, so a gap remains visible. If and adequate connection is observed, the Flex Connector may be replaced with a new, sterile Flex Connector and the system connection should be re-evaluated. Lastly, if use of a Flex Connector with the ventilator circuit is optional, the closed suction Swivel Connector can be connected directly to the ventilator circuit. During use, if a disconnection between the Flex Connector and Swivel Connector is observed during use, replace the Flex Connector with a new, sterile Flex Connector (e.g. HYH Product # 1115) and ensure adequate connection as described in this notice. If an inadequate connection occurs, Halyard Health instructs customers to safe the Flex Connector and report it via the Product Complaint Process (PIQ) and the local Halyard Health Sales Representative will assist. Additionally customers were instructed to complete and return the Acknowledgement Form and distribute the recall notice to all clinicians in the facility that may be impacted. |  | 20722634420 | 32,189 cases (20 units per case - 643,780 units) total | Worldwide distribution. US nationwide including Puerto Rico, ARGENTINA, AUSTRALIA, BAHAMAS, BELGIUM, BOLIVIA, BRAZIL, BULGARIA, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, CROATIA, CZECH REPUBLIC, DOMINICAN REP., ECUADOR, EL SALVADOR, GUATEMALA, HONG KONG, HUNGARY, INDIA, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KAZAKHSTAN, LEBANON, MALAYSIA, MEXICO, NEW ZEALAND, PERU, PHILIPPINES, QATAR, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SOUTH KOREA, SPAIN, SWEDEN, THAILAND, UNITED KINGDOM, and UNITED ARAB EMIRATES. | Open, Classified |  |  | 20722634420 | Not_Computer | Not_Computer
Z-1191-2018 | 79406 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer M/L Taper Hip Prosthesis, 7711 Series, Femoral Stem 12/14 Neck Taper Plasma Sprayed Press-Fit Cementless Size 7.5 Extended Offset Reduced Neck Length; Part Number: 00-7711-007-40 |  |  | March 01, 2018 | 2018 | Zimmer Biomet, Inc. | Certain packages labeled as size 7.5 Extended Offset Reduced Neck length hip implant, actually contained a Size 4 Standard Reduced Neck length taper hip implant. | Labeling mix-ups | RISK MANAGER: 1. Review this notification and notify affected personnel of the contents. 2. Quarantine all affected product. Your Zimmer Biomet sales representative will remove the affected product from your facility. 3. Complete Certificate of Acknowledgement and send to CorporateQuality.PostMarket@zimmerbiomet.com. DISTRIBUTOR: 1. Review this notification and notify affected team members of the contents. 2. Quarantine affected product in your inventory. 3. Return affected product from your distributorship and from affected hospitals. a. Complete Inventory Return Certification Form and send to CorporateQuality.PostMarket@zimmerbiomet.com Mark RECALL on the outside of the returned cartons. 4. Return the Additional Accounts form to CorporateQuality.PostMarket@zimmerbiomet.com. |  | 12 | 12 | Distributed in the following U.S. states: AZ, CA, CT, IN, MD, MI, MN, NJ, NY, OR, WA, and Puerto Rico. Distributed in China and India. | Open, Classified |  |  | 36 | Not_Computer | Not_Computer
Z-2480-2018 | 80289 | No | LRP | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Presource PBDS, Kit, Common Diabetes. Kit includes a exerciser volumetric 5000 ML device, two items of cloth, skin prep 2% CHG, a packet of 4 chlorhexidine gluconate antiseptic cloths, and Gatorade pouches. |  |  | February 20, 2018 | 2018 | Cardinal Health 200, LLC | A defect in the seal of the Gatorade pouch contained in kits caused the pouch to leak within a limited portion of the kit. | Packaging | Cardinal Health notified customers via Fed Ex overnight, on about 02/21/2018. The URGENT PRODUCT RECALL letter instructed customers to: 1) Inspect your inventory for the affected product code and lot number; 2)Segregate and quarantine all product on hand; 3) Return enclosed acknowledgment form via fax to 847-689-9101 or 614-652-9648, whether or not you have affected product, indicating the product code, lot and quantity of product youve quarantined; 4) Notify any customers to whom you may have distributed productaffected by this recall; 5)Contact Customer Service at 800-964-5227 to arrange return and credit. Customers were also instructed to return any affected kit on site. |  | 8928 | 8,928 kits total | US Nationwide Distribution in the states of California. | Open, Classified |  |  | 8928 | Not_Computer | Not_Computer
Z-1262-2018 | 79613 | [6 empty] | F-T3SYSTEM, T3 (Trumpf Table Top) System |  |  | February 21, 2018 | 2018 | Allen Medical Systems | The recalling firm has become aware that the rail section clamps used to attach the leg support of the device may fail during use, which could result in critical harm to the patient. | Error in labeling | On February 2, 2018, the firm issued an Urgent Medical Device Correction notification to their distributors to make them aware of the device correction. The firm requested that distributors further disseminate an attached recall letter to consignees.The communication addressed to consignees advised of the issue affecting the device. Consignees were asked to inform users that the device is for armboard use only, and not to be used with the product system's leg support. The notice informed consignees that the firm would provide new units upon completion of development along with revised Instructions for Use.Customers with questions concerning this Urgent Medical Device Correction were referred to Allen Medical Customer Service: (800) 433-5774, 8:30AM-5PM (EST) |  | 83 | 83 | Worldwide Distribution - USA (nationwide) Distribution and to the countries of : Australia, France, Saudi Arabia, Germany, and France. | Open, Classified |  |  | 83 | Not_Computer | Not_Computer
Z-1464-2018 | 79798 | No | EMO | Dental hand instrument. | Dental | Dental | 510(K) Exempt | Double Blade Shoe F/Unigraft Knife 1.5 Thick X 9MM SKU 6500009 |  |  | March 28, 2018 | 2018 | Ace Surgical Supply Co., Inc. | The product inside the package is incorrect. A single blade shoe was mistakenly packaged instead of a double blade. | Error in labeling | ACE is contacting the 5 affected customers via certified mail for US customers and via email for non-US Customers. Follow-up notifications will be sent 12 weeks after initial notification. Firm will rework incorrectly packaged product and will use completed reply forms as effectiveness check. For further questions, please call (508) 588-3100. |  | 6 | 6 | US Distribution to the states of : MI, CA, OK, and Internationally to China, Canada, | Terminated | September 11, 2018 | 167 | 6 | Not_Computer | Not_Computer
Z-1193-2018 | 79301 | No | GDT | Removable skin staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Teleflex MEDICAL WECK Visistat 35. Used for the external skin closure of lacerations. |  |  | February 15, 2018 | 2018 | Teleflex Medical | Some boxes are labeled as Visistat Wide devices, when in fact the contain Visistat Regular devices. The 6 individually packaged units within the boxes are correctly labeled. | Packaging process control | The firm initiated the recall on 02/14/2018 by letter and followed up with another letter on 03/12/2018 to clarify the labeling issue. The letter to the distributor stated:"1. Immediately discontinue distribution and quarantine any products with the catalog and lot number listed above.2. Inspect affected products within your control to identify if the box label correctly identifies the products inside the box. Products with an incorrect label on the box should be returned to Teleflex Medical so that they are not distributed to health care providers in error.3. Using the provided Customer Letter and Acknowledgement Form templates, communicate this recall to any of your customers who have received product included within the scope of the recall. Customers may return mislabeled products if they choose to do so. Note the attached Customer Letter template was updated since this letter was first sent to you.4. If any of your customers wish to return affected products, have them return the affected products to you, together with a completed Acknowledgement Form, for consolidation and return to Teleflex Medical. In the event that an alternative approach is needed, contact Teleflex Customer Service for more information at 1-866-246-6990."The customer letter requested discontinuation of use and quarantining of products and return of affected stock. |  | 12 | 12 units | KS | Terminated | October 31, 2018 | 258 | 12 | Not_Computer | Not_Computer
Z-1013-2018 | 79302 | No | ILO | Water circulating hot or cold pack. | Physical Medicine | Physical Medicine | 510(K) Exempt | T/Pump temperature therapy pump, Models TP700 and TP700C |  |  | November 07, 2017 | 2017 | Stryker Medical Division of Stryker Corporation | The T/Pump Operations Manual and Maintenance Manual were incorrectly revised to add the sentence, Temperature therapy can also be for heat conservation for normothermia maintenance to the Indications for Use statement. | Error in labeling | The following actions are included in the consignee letter: 1. Locate the operations (101515 rev D.0) and maintenance (101176 rev F.0)manuals that accompanied the units listed on the attached business reply form.2. Find the corrected pages included in this letter and insert one into each affectedmanual at your location.3. Complete the enclosed business reply form to acknowledge that you havereceived and understand this notification.4. Return your completed business reply form by fax 269 488 8691 or emailproductfieldaction@stryker.com to Strykers regulatory department.5. If you have loaned or sold any of the products listed in this letter, please forwarda copy of this notice to the new users and advise us of their new location in thespace provided on the business reply form.6. If you have disposed of any of these units and they are no longer in use, please,advise us of their obsolescence by providing us with their serial number in thespace provided on the business reply form. |  | 9902 | 9902 | Distributed in all 50 states, the District of Columbia, and Guam. | Open, Classified |  |  | 9902 | Not_Computer | Not_Computer
Z-1041-2018 | 78706 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | CR PRECOAT FEM COMP SIZE BMR |  |  | November 29, 2017 | 2017 | Zimmer Biomet, Inc. | The LDPE bag packaging for various highly polished hip and knee implants may adhere to the highly polished surface, leaving residue or material from the LDPE bag on the implant after it is removed from the bag. | Packaging | A similar recall was initiated in January 2016 to remove remaining inventory packaged in the old bag. Further evaluation identified additional lots packaged in the old bag that were not included in the January 2016 recall; therefore, resulting in this new recall for the additional lots packaged in the old bag.A firm, Zimmer Biomet, sent an "Urgent Medical Device Recall" notices dated November 29, 2017 and response forms to customers. The notice described the product, problem and actions to be taken. The Customers should take the following action:Risk Manager Responsibilities:1. Review this notification and ensure that affected personnel are aware of the contents.2. If you have affected product at your facility, assist your Zimmer Biomet sales representative and quarantine all affected product. Your Zimmer Biomet sales representative will remove the affected product from your facility.3. Complete Attachment 1  Certificate of Acknowledgement and send to CorporateQuality.PostMarket@zimmerbiomet.com. This form must be returned even if you do not have affected products at your facility.4. Retain a copy of the acknowledgement form with your recall records in the event of a compliance audit of your facilitys documentation.Distributor Responsibilities:1. Review this notification and ensure that affected team members are aware of the contents.2. Immediately locate and quarantine affected product in your inventory.3. Immediately return all affected product from your distributorship and from affected hospitals within your territory.a. Complete Attachment 1  Inventory Return Certification Form and send to CorporateQuality.PostMarket@zimmerbiomet.com within three (3) days.b. For each return, send a copy of Attachment 1 to CorporateQuality.PostMarket@zimmerbiomet.com.c. Include a hardcopy of Attachment 1 in each carton of your return shipment for immediate processing.d. Mark RECALL on the outside of the returned cartons.4. Return the Additi |  | 13227 | 13,227 in total | Worldwide Distribution: US (nationwide) and countries of: Argentina, Australia, Brazil, Bolivia, Canada, Chile, China, Colombia, Ecuador, El Salvador, India, Japan, Malaysia, Mexico, Netherlands, New Zealand, Nicaragua, Panama, Seoul-Korea, Singapore, Taiwan, Thailand, and UAE. | Open, Classified |  |  | 13227 | Not_Computer | Not_Computer
Z-1525-2018 | 79677 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Avenir Muller Stem; Item Number: 01.06010.104, Lot Number: 2905496, UDI: (01)00889024479500 (17)220430 (10)2905496 |  |  | February 28, 2018 | 2018 | Zimmer GmbH | A potential commingle that could result in the product in the package not matching the product on the outer label. | Error in labeling | On March 7, 2018 Field Safety Corrective Action Notices were hand delivered by Zimmer Biomet team members or their representatives. Customers were instructed to complete the following: 1.Review this notification for awareness of the contents.2.Assist your Zimmer Biomet sales representative to quarantine immediately all affected implants.3.Your Zimmer Biomet sales representative will remove the affected implants from your facility.4.Complete Attachment 1  Certificate of Acknowledgement.a. Return a digital copy to fieldaction.emea@zimmerbiomet.com. b. Retain a copy of the Certificate of Acknowledgement with your field action records in the event of a compliance audit of your documentation.5.If after reviewing the notice you have further questions or concerns please contact your Zimmer Biomet representative. |  |  |  | There was no distribution to the US. The distribution occurred to the following countries: Belgium, Switzerland, Germany, France, Canary Islands, Italy, and Thailand. | Open, Classified |  |  |  | Not_Computer | Not_Computer
Z-0992-2018 | 79242 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | DELTA CER FEM HD 32/0MM T1 |  |  | June 30, 2017 | 2017 | Zimmer Biomet, Inc. | Product may contain the presence of melted sterile packaging upon opening. | Packaging process control | On June 2, 2017 an URGENT Medical Recall Letter was issued to customers regarding affected product(s). The letter requests customers to review the content and notify affected personnel of the recall. Assist the sales representative in quarantining affected product, which will be removed and returned to recalling firm and then replaced. All customers who may have received product from distributors are to be notified. Questions or concerns can be directed to 1-574-372-4847. |  | 30 | 30 | US in the state of Texas | Terminated | April 13, 2018 | 287 | 80 | Not_Computer | Not_Computer
Z-1210-2018 | 79152 | No | LPL | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic | 510(k) | ACUVUE ADVANCE Brand Contact Lenses with HYDRACLEAR |  |  | October 19, 2017 | 2017 | Johnson &amp; Johnson Vision Care, Inc. | Some contact lenses were associated with complaints of a foreign matter caught between the blister package and foil. | Packaging process control | On October 19, 2017, Johnson and Johnson sent letters to their customers stating the recall of ACUVUE and given the instructions to review their inventory, stop using the affected lots, use the enclosed label to return any affected products, contact customer service to arrange the return, even if they do not have any recall product inventory. A contact was provided to each individual letter. |  | 165750 | 165,750 | Worldwide distribution. US nationwide including Puerto Rico, Korea, Czech Republic, China, Japan, Great Britain, Austria, Malaysia, Hong Kong, Russia, Taiwan, Singapore, India, and Brazil | Open, Classified |  |  | 347583 | Not_Computer | Not_Computer
Z-0978-2018 | 79191 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | HeartStart XL+ Defibrillator/Monitor |  |  | February 07, 2018 | 2018 | Philips Electronics North America Corporation | Update XLÂż device software to version A.03. This includes enhancements to the Operational Checks, event logs, and troubleshooting messages that provide as complete of information as needed to users ondevice readiness. | Labeling mix-ups | Philips is initiating a correction to affected devices. A software upgrade is available free of charge for all units affected by the issue. A Philips Healthcare representative will contact customers with affected devices to arrange for installation of the upgrades. Customers may continue to use XL+ prior to receiving the software upgrade. For further questions, please call (978) 659-3000. |  | 4315 | 4315 | Foreign Only - China | Open, Classified |  |  | 4315 | Not_Computer | Computer
Z-1515-2018 | 79755 | No | MVN | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic | 510(k) | clariti 1 day/Fresh Day 90 pack (-1.00), UV Blocking, 90 silicon hydrogel, daily disposable contact lenses |  |  | February 21, 2018 | 2018 | CooperVision Inc. | The secondary packaging is labeled with an incorrect expiration date. | Incorrect or no expiration date | The firm, CooperVision, sent an "Firm Initiated Voluntary Removal - ACTION REQUIRED" letter to its customers on 2/21/18. The letter describes the product, problem and actions to be taken. The letters instructed customers to: Examine your inventory for any affected lot numbers; Stop any further distribution of the specified lot(s) of clariti 1 day and Fresh Day contact lenses; discard product and contact CooperVision, Inc. Customer Service department at 1-800-341-2020 to receive credit and/or complimentary replacement product.The removal only affects clariti 1 day lenses with the lot numbers specified on the enclosed letter. No other clariti 1 day and Fresh Day products  or any other CooperVision products  are affected.If you have any further questions, please feel free to contact your Account Representative or Customer Service at 1-800-341-2020, 9:00 AM  7:00 PM ET, Monday - Friday. |  | 43560 | 22 packs/1980 lenses | US Nationwide Distribution to states of: AL, AR, CA, CO, FL, GA, HI, IA, IL, IN, KS, KY, MD, ME, MI, MN, MO, MS, NC, NE, NH, NJ, NY, OH, OR, PA, SC, TN, TX, UT, VA, WA, WI, and WY. | Open, Classified |  |  | 1321810 | Not_Computer | Not_Computer
Z-1008-2018 | 79254 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | 36MM COCR MODULAR HD -6MM |  |  | September 15, 2017 | 2017 | Zimmer Biomet, Inc. | One piece from a different lot mixed with lot 236840 and is incorrectly etched as -6MM offset instead of +3MM offset. As the incorrect etching and related size difference affects the functional use of this product, the affected lot is being removed from the field. | Error in labeling | Foreign risk managers were emailed notification of recall with the following instructions:Risk Manager Responsibilities:1.Review this notification and ensure that affected personnel are aware of the contents.2.If you have affected product at your facility, assist your Zimmer Biomet sales representative and quarantine all affected product. Your Zimmer Biomet sales representative will remove the affected product from your facility.3.Complete Attachment 1  Certificate of Acknowledgement and send to CorporateQuality.PostMarket@zimmerbiomet.com. This form must be returned even if you do not have affected products at your facility.4.Retain a copy of the acknowledgement form with your recall records in the event of a compliance audit of your facilitys documentation.5.If you have further questions or concerns after reviewing this notice, please call customer service at 574-371-3071 between 8:00 am and 5:00pm EST, Monday through Friday. Calls received outside of call center operating hours will receive a voicemail prompt or be transferred to an on-call representative in the event of an emergency. Alternatively, your questions may be emailed to CorporateQuality.PostMarket@zimmerbiomet.com. |  | 23 | 23 | Outside USA | Terminated | August 20, 2018 | 339 | 23 | Not_Computer | Not_Computer
Z-2370-2018 | 80188 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | ZimmerÂż M/L Taper Hip Prosthesis With KinectiveÂż Technology Cementless, Size 16.25; 00-7713-016-00 |  |  | May 30, 2018 | 2018 | Zimmer Biomet, Inc. | M/L Taper Hip Prosthesis and M/L Taper Hip Prosthesis with Kinectiv Technology manufactured prior to October 2015 are being recalled due to the potential for black residue to be observed within the polyethylene bag containing the implant. The black residue is a combination of the titanium plasma spray used to create the porous surface of the implant and polyethylene bag containing the implant. The recalling firm states the residue is biocompatible. Risks include delay of surgery while locating a replacement device, and surgical intervention due to wear caused by debris. | Packaging | On May 30, 2018, the firm emailed Urgent Medical Device Recall letters to its Japanese subsidiary, who then further distributed the letters locally to Distributors and Risk Managers in Japan. The letter informed affected consignees of the product issue. Customers were instructed to do the following:Distributor Responsibilities:1. Review this notification and ensure that affected team members are aware of the contents.2. Immediately locate and quarantine affected product in your inventory.3. Immediately return all affected product from your distributorship and from affected hospitals within your territory.a. Complete Attachment 1  Inventory Return Certification Form and send to CorporateQuality.PostMarket@zimmerbiomet.com within three (3) days.b. Request a Return Authorization Number by emailing zimmerbiometintlirarequests@zimmerbiomet.com.c. For each return, send a copy of Attachment 1 to CorporateQuality.PostMarket@zimmerbiomet.com.d. Include a hardcopy of Attachment 1 in each carton of your return shipment for immediate processing.e. Mark RECALL on the outside of the returned cartons.Risk Manager Responsibilities:1. Review this notification and ensure that affected personnel are aware of the contents.2. If you have affected product at your facility, assist your Zimmer Biomet sales representative and quarantine all affected product. Your Zimmer Biomet sales representative will remove the affected product from your facility.3. Complete Attachment 1  Certificate of Acknowledgement and send to CorporateQuality.PostMarket@zimmerbiomet.com. This form must be returned even if you do not have affected products at your facility.If you have further questions or concerns after reviewing this information, please call Zimmer Biomet customer service at 574-371-3071 between 8:00 am and 5:00pm EST, Monday through Friday. Calls received outside of call center operating hours will receive a voicemail prompt or be transferred to an on-call r |  | 665 | 665 | Products were distributed solely to Japan. | Open, Classified |  |  | 3931 | Not_Computer | Not_Computer
Z-1535-2018 | 79453 | No | OBO | Ophthalmoscope. | Ophthalmic | Ophthalmic | 510(k) | Widefield OCT software feature for the Spectralis HRA+OCT and variants |  |  | December 08, 2017 | 2017 | Heidelberg Engineering | FDA notified Heidelberg Engineering that a 510(k) clearance is needed for the SPECTRALIS Widefield OCT feature, and, on November 27, 2017, Heidelberg Engineering submitted a new 510(k) for this feature. FDA has asked Heidelberg Engineering to instruct customers not to use the Widefield OCT software feature on the SPECTRALIS device until 510(k) clearance has been obtained for this software feature. | No Marketing Application | On December 14, 2017 an URGENT Medical device software field correction letter was issued to customers via email or ground mail to notify all customers of the recall action. The letter states the following:FDA has asked us to instruct you not to use the Widefield OCT feature on your SPECTRALIS device until a decision is made on this pending 510(k) submission. Heidelberg Engineering will notify you immediately upon completion of the FDAs review. In the meantime, you may continue to use your Widefield Objective for cSLO fundus imaging (including Fluorescein Angiography, ICG Angiography, MultiColor, Infrared and Red-Free Fundus Imaging). Questions or concerns can be directed to (508) 530-5904 |  | 382 | 382 | Nationally | Open, Classified |  |  | 382 | Not_Computer | Computer
Z-1519-2018 | 79526 | No | LDP | Acetaminophen test system. | Toxicology | Toxicology | 510(k) | Roche Acetaminophen assay |  |  | January 02, 2018 | 2018 | Roche Diagnostics Corporation | The previous version of method sheets for the Acetaminophen assay on the COBAS INTEGRA 400plus/800/800 CTS analyzers and the MODULAR ANALYTICS P module are missing the followingstatement in the Limitations-interference section:"In very rare cases, gammopathy, in particular type IgM (WaldenstrÂżm s macroglobulinemia), may causeunreliable results." | Error in labeling | On January 2, 2018 an URGENT MEDICAL DEVICE CORRECTION letter was issued to customers with notification to the issue and information regarding the updated method sheets and that they were available for use. The COBAS INTEGRA 400 plus/800/800 CTS was updated from affected method sheet version 3.0 to 4.0. The MODULAR ANALYTICS P was updated from affected method sheet version 4.0 to 5.0 No product return is required. Please contact the Roche Support Network Customer Support Center, 24 hours a day, seven days a weekat 1-800-428-2336 if you have questions about the information contained in this UMDC. |  | 35111 | 35,111 | Nationwide Distribution | Open, Classified |  |  | 35111 | Not_Computer | Computer
Z-1206-2018 | 79454 | bandage, elastic | FQM | Elastic bandage. | General Hospital | General Hospital | 510(K) Exempt | Johnson &amp; Johnson BAND-AIDÂż Brand First Aid Products HURT-FREEÂż Wrap (size: 2in) |  |  | December 18, 2017 | 2017 | Johnson &amp; Johnson Consumer, Inc. | Johnson &amp; Johnson Consumer Inc. is initiating a voluntary, Class II recall at the retail level in the U.S. of select COACHÂż Self-Adhering Sports Wraps, BAND-AIDÂż Brand First Aid Products SECURE-FLEXÂż Wraps and BAND-AIDÂż Brand First Aid Products HURT-FREEÂż Wraps products. The current labeling claim states not made with natural rubber latex which needs to be updated based upon recent awareness that natural latex was used as a base ingredient in the early-stage manufacturing process, which reduces the allergic protein found in natural latex. | Labeling design | On December 18, 2017 an URGENT MEDICAL DEVICE RECALL letter was issued to retailers and warehouses with intentions to notify all affected individuals, including customers of the recall. The recall letter requests that the following actions be taken:1. Immediately examine your inventory and quarantine the above-mentioned product subject to recall. In addition, if you may have further distributed this product to retailers, warehouses or distributors, please identify and send a list of these locations to Stericycle at BandAid8687@stericycle.com and they will be notified.2. Please complete the enclosed Business Reply Form and return immediately by emailing to BandAid8687@stericycle.com or faxing to (1-888-965-5802).3. If you have any questions regarding the form, please call Stericycle at (1-855-215-5023).Please return ALL inventory of the identified product above. No product other than that specified is to be returned.Questions or concerns can be directed to 1-855-215-5023 |  |  |  | Nationally | Open, Classified |  |  |  | Not_Computer | Not_Computer
Z-1521-2018 | 79861 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | AdvanceÂż 35LP Low-Profile PTA Balloon Dilatation Catheter; Catalog numbers: PTA5-35-135-6-2 .0 and PTA5-35-135-4-4 .0; UDI: 10827002522647 and 10827002522524 |  |  | April 16, 2018 | 2018 | Cook Inc. | Incorrect product labeling. Product labeled as 6mm x 2cm balloon are packaged with a 4mm x 4 cm balloon label and vice-versa. | Labeling mix-ups | On April 16, 2018 Cook Medical Inc. mailed an Urgent Medical Device recall notification to affected customers. Customers were instructed to: 1) Examine your inventory immediately to determine if you have affected product(s), and quarantine affected product(s). Immediately cease all distribution and use of these products; 2) Return the affected product(s) to Cook Medical with a copy of the Acknowledgement and Receipt Form to receive a product credit.NOTE: Unaffected products that are returned will not be credited; 3) Please complete the Acknowledgement and Receipt Form within 5 business days of receiving this letter. Even if you do not have affected product(s) on hand, you must still complete the Acknowledgement and Receipt Form and return it via fax (812.339.7316) or email (FieldActionsNA@CookMedical.com); 4) Immediately report adverse events to Cook Medical Customer Relations by phone at 800.457.4500 or812.339.2235, Monday through Friday between 7:30 am and 5:00 pm (Eastern Time) or by email to CustomerRelationsNA@CookMedical.com Cook Incorporated |  | 40 | 40 total products | Worldwide Distribution - US Nationwide in the states of AZ, DE, FL, GA, ID, IL, KY, MI, MO, NH, NJ, PA, RI, TX, UT and the country of Peru | Open, Classified |  |  | 40 | Not_Computer | Not_Computer
Z-2383-2018 | 80299 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | VPS Rhythm ECG Accessory Pack, REF RHY-100-ECGO |  |  | May 17, 2018 | 2018 | Arrow International Inc | Product pouch label may not easily identify thelot number, expiration date and associated symbols.. | Labeling False and Misleading | On May 17, 2018, Arrow International issued Urgent Medical Device Recall Notices and Response Forms to impacted customers via FedEx 2-day. Customers were advised to take the following actions:1. If you have affected stock, immediately discontinue use and quarantine any products with the product codes and lot numbers listed, so that the affected products can be returned to Arrow.2.To return product, complete the enclosed Recall Acknowledgement Form and fax it to 1-855-419-8507, Attn: Customer Service or email to recalls@teleflex.com. Soon after, a customer service representative will contact customers with a Return Goods Authorization (RGA) Number and will provide instructions for the return of products to Arrow International.3.If you have no affected stock, please complete the Recall Acknowledgment Form and fax it to 1-855-419-8507, Attn: Customer Service or email to recalls@teleflex.com. This will allow the firm to document your receipt of the letter.Customers with questions, feel free to contact your local sales representative or Customer Service at 1-866-396-2111. |  | 2070 | 2070 units in total | US Nationwide Distribution in the states of AR, AZ, CT, FL, GA, IL, IN, KS, KY, MA, MI, MN, MO, NC, NJ, NY, OH, PA, SC, TX, VA &amp; WA | Open, Classified |  |  | 2070 | Not_Computer | Not_Computer
Z-0974-2018 | 79030 | No | OLC | Surgical mesh. | General &amp; Plastic Surgery | Orthopedic | 510(k) | CMI (Collagen Meniscus Implant) device, Ivy Sports Medicine s collagen-based meniscus implant |  |  | January 25, 2018 | 2018 | Stryker Corporation | The recalled products were shipped without the required temperature control packaging, therefore could potentially have been exposed to elevated temperatures during transit. | Packaging | The international Stryker site was notified by email on 1/25/2018. Customers are instructed to: 1. Inform individuals within your organization who need to be aware of this device removal. 2. Review Part Numbers (4600, 4601, 4607, and 4612) for affected lot numbers. Please determine if you have theaffected product in stock. Response is required. 3. If no product is found, notify your local Stryker office.4. If you do have product, segregate the product and call your local Stryker office to arrange for product return andissuance of credit. |  | 155 | 155 devices | Netherlands | Open, Classified |  |  | 155 | Not_Computer | Not_Computer
Z-0973-2018 | 79190 | No | JCX | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | LCSU 4, Laerdal Compact Suction Unit |  |  | August 31, 2017 | 2017 | Laerdal Medical Corporation | A shipment of LCSU 4 suction units and canisters attempted importation into the United States but was refused by FDA for lack of necessary approvals with CDRH by the foreign firms. | No Marketing Application | On September 5, 2017, a letter was mailed to customers notifying them that they received one of the units subject to recall. The letter requested that the customer return the affected item for a replacement. Questions or concerns can be directed to 1-877-523-7325 or email: customerservice@laerdal.com , referencing the following number: 2017-R-05. |  | 204 | 204 | Nationally | Terminated | June 13, 2018 | 286 | 204 | Not_Computer | Not_Computer
Z-1485-2018 | 79749 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Genesis Surgical Cassette Tapered |  |  | March 13, 2018 | 2018 | Keystone Dental Inc | The surgical cassette contained extra-large white grommets making the 3.8 mm and 6.5mm dental burrs wobble within the grommets in the surgical cassette. | Error in labeling | Notification was mailed to all impacted customers on March 13, 2018, via overnight Federal Ex-press. Customer Relations will contact the customers via phone to ensure the customers received the notification, instructions, and replacement grommets. Customer Relations Reply Forms are being used to document the communication. For further questions, please call (866) 902-9272. |  | 19 | 19 | USA (nationwide) Distribution and Internationally to Turkey. | Terminated | July 23, 2018 | 132 | 19 | Not_Computer | Not_Computer
Z-1512-2018 | 79661 | No | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Knee Prosthesis, Sterile, Single-Use devices |  |  | January 12, 2018 | 2018 | Howmedica Osteonics Corp. | Firm became aware that the specific product/lot combinations referenced above may contain the incorrect size implant from what is labeled on the box. | Error in labeling | Branches/Agencies notified by email on 01/16/2018 and were asked to quarantine affected devices. Notification letters and product accountability forms were sent via UPS (with return receipt) on 01/15/2018. Customer notifications were updated and sent to consignees on 03/16/2018 as the technical and medical assessments associated with the issue were completed and additional information related to potential hazards and harms were identified. |  | 6 | 6 units of catalog 5521-B-400 | Affected products were sent to three medical facilities. Affected product will be returned to the manufacturer for disposition. | Open, Classified |  |  | 6 | Not_Computer | Not_Computer
Z-1526-2018 | 79751 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Hinge Coupling Hoffmann LRF Catalog # 4933-0-800 |  |  | March 12, 2018 | 2018 | Stryker GmbH | The wrong parts were picked for packaging. The connecting Nuts were packed in the packaging labeled with Hinge coupling | Labeling mix-ups | Stryker sent an Urgent Product Recall letter dated March 12, 2018, to all affected customers. .Product will be returned to the manufacturer for disposition.Non-responding consignees will be attempted to be contacted a minimum of three times in total. Returned product will be quarantined and returned to the manufacturer.Product will be returned to the manufacturer for disposition.For further questions, please call (201) 831-5272. |  | 18 | 18 | USA (nationwide) Distribution | Open, Classified |  |  | 18 | Not_Computer | Not_Computer
Z-1702-2018 | 79972 | No | KXK | Radionuclide brachytherapy source. | Radiology | Radiology | 510(k) | BARDÂż BrachySourceÂż Iodine125 radioactive seeds consist of a welded titanium capsule containing the I125 absorbed onto a nickel / copper-coated, gold-cored aluminum wire (reference Figure 1 below). The I125 has a half-life of 59.6 days and decays by electron capture with the emission of characteristic photons and Auger electrons. The principal photon emissions are 27.4 and 31 keV x-rays and 35.5 keV gamma. The titanium wall of the BrachySourceÂż Seed Implants absorbs the electrons. |  |  | April 10, 2018 | 2018 | Bard Brachytherapy, Inc. | Bard Medical Division is recalling the BARDÂż BrachySourceÂż Iodine125 Radioactive Seeds in MickÂż Applicators and BARDÂż BrachySourceÂż Iodine125 Radioactive Seeds Loose in Vial with Sterile Spacers. Because the incorrect dosage level radioactive seeds were mixed up and sent to two hospitals. One Hospital ordered BrachySourceÂż Iodine125 0.36 mCi radioactive seeds but instead received BrachySourceÂż Iodine125 0.28 mCi radioactive seeds and vice versa for the other hospital affected. | Labeling mix-ups | The notification to the customer was completed by Bard Medical Customer Service via a telephone call to alert them of the incorrect product shipped on April 10, 2018.Bard Medical Division is recalling the BARDÂż BrachySourceÂż Iodine125 Radioactive Seeds in MickÂż Applicators and BARDÂż BrachySourceÂż Iodine125 Radioactive Seeds Loose in Vial with Sterile Spacers. The incorrect dosage level of radioactive seeds was sent to 2 hospitals. One hospital ordered 0.36 mCi radioactive seeds but instead received 0.28 mCi radioactive seeds intended for the other hospital. For further questions, please call (770) 385-0562. |  | 170.64 | 70 Seeds at 0.28mCi and 100 seeds at 0.36mCi | US Distribution to the states of : California and Georgia | Open, Classified |  |  | 170.64 | Not_Computer | Not_Computer
Z-1064-2018 | 79387 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | 6.5mm x 125mm x 16mm Cannulated Screw, Item Number: 110009948 (UDI: (01)00887868111327 (17)270421 (10)120910) |  |  | December 07, 2017 | 2017 | Zimmer Biomet, Inc. | One lot of 3.5mm x 24mm Cortical Locking screws is labeled as 6.5mm x 125mm x 16mm Cannulated screws. | Error in labeling | Zimmer Biomet sent an Urgent Medical Device Recall letter dated December 2017 to their customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to: 1. Review this notification n and ensure that affected team members are aware of the contents.2. Immediately locate and quarantine affected product in your inventory. 3. Immediately return all affected product from your distributorship and from affected hospitals within your territory.a. Complete Attachment 1  Inventory Return Certification Form and send to CorporateQuality.PostMarket@zimmerbiomet.com within three (3) days.b. Request an International Return Authorization Number by emailing zimmerbiometintlirarequests@zimmerbiomet.comc. For each return, send a copy of Attachment 1 to CorporateQuality.PostMarket@zimmerbiomet.com.d. Include a hardcopy of Attachment 1 in each carton of your return shipment for immediate processing.e. Mark "RECALL" on the outside of the returned cartons.4. Return a copy of the your Inventory Return Certification and product return form.5. For questions call customer service at 574-374-3071. |  | 6 | 6 | Canada. | Open, Classified |  |  | 6 | Not_Computer | Not_Computer
Z-1424-2018 | 76620 | No | FSH | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | JAMSHIDI (TJ) NEEDLE BONE MARROW [a.) 13G x 3.5, b.) 8G x 4, c.) 11G x 4, d.) 8G x 6, e.) 11G x 6]; STERILE; |  |  | February 23, 2017 | 2017 | Carefusion 2200 Inc | The integrity of the sterile packaging is potentially compromised. | Package design/selection | On 2/23/2017, firm notified customers of recall via mailed letter. The letter included a recall response form. Customers were directed to either return the affected product along with the response form or destroy it and return the response form, indicating the product had been destroyed. All affected product will be destroyed. |  |  |  | Nationwide, Australia, Brazil, Canada, China, Japan, Kuwait, Malaysia, Mexico, Netherlands, New Zealand, Philippines, Singapore, South Africa, Taiwan, Thailand, United Arab Emirates, Vietnam | Open, Classified |  |  |  | Not_Computer | Not_Computer
Z-2556-2018 | 80276 | Yes | MUQ |  |  | Cardiovascular | PMA | BioGlue Spreader Tip, 12mm, Sterile, RX Only. |  |  | June 04, 2018 | 2018 | CryoLife, Inc. | The CryoLife L6318 (Label, BioGlue 12mm Spreader Tip, Indicator Label 2 1/8" x 1 1/4") was incorrectly applied to the Syringe Applicator Tip pouches and Syringe Spreader Tip pouches were not included in the boxes. | Error in labeling | CyroLife notified their customers on/about 06/04/2018 via Urgent - Important Product Removal letter. instructions indicated that affected pouches have been returned to CryoLife, Inc. Any questions/clarifications regarding this removal are directed to contact 800-438-8285 or by email to fieldassurance@c ryolife.com. |  | 855 | 855 Total | Distribution is to Japan | Terminated | February 08, 2019 | 249 | 855 | Not_Computer | Not_Computer
Z-1131-2018 | 79213 | No | KNG | Fetal ultrasonic monitor and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Cooper Surgical 2 MHZ Waterproof OB Probe |  |  | October 06, 2017 | 2017 | CooperSurgical, Inc. | The 2MHZ prove was incorrectlhttp://cts.fda.gov/division-tracking/images/trash.pngy labeled as a 3MHZ probe | Error in labeling | CooperSurgical sent an Urgent Medical Device Recall letter dated October 6, 2017 to affected customers via UPS with confirmed delivery receipt. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to discontinue use of affected products and complete the attached Acknowledgement and Receipt Form for replacement. If you have any further questions please feel free to contact me at 203.601.5200. |  | 13 | 13 units | Worldwide Distribution - US Nationwide and the countries of ( Canada and England) | Completed |  |  | 13 | Not_Computer | Not_Computer
Z-0920-2018 | 79219 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Signature Ceramic Femoral Head. Model Number: 111-152-621. |  |  | February 02, 2018 | 2018 | Signature Orthopedics Pty Limited | This recall has been identified due to the batch being mislabeled, with the packagedcomponent's size not corresponding to the label. Therefore, the product intended to be usedfor the procedure may not be included in the packaging resulting in surgical delay asreplacement product is retrieved. | Labeling mix-ups | Signature Orthopaedics notified customer(s) of the recall by sending an Urgent Medical Device Recall Letter dated February 2, 2018 via email. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete the Customer Acknowledgement Form and firm's Effectiveness Check Questionnaire. |  | 10 | 10 | Worldwide Distribution - U.S. Nationwide and the country of Australia. | Terminated | August 08, 2018 | 187 | 10 | Not_Computer | Not_Computer
Z-2786-2018 | 80535 | No | CDM | Urinary urobilinogen (nonquantitative) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Clarity Urocheck 10SG Urinalysis Strips, Model # DTG-10SG |  |  | June 22, 2018 | 2018 | Clarity Diagnostics Llc | Clarity Diagnostics LLC discovered specific lots of Clarity 10SG Strips are not compatible for use in the Clarity Urocheck 120 Urine Analyzers, due to lack of Canister Code Bar Code Label. | Labeling design | The firm notified the major distributor who received this product on or about 06/14/2018 via a letter entitled "URGENT: Voluntary Medical Device Recall for Clarity 10SG Urinalysis Reagent Strips".The distributor letter stated effective immediately, the purpose of this letter is to advise that Clarity Diagnostics LLC is voluntarily recalling Clarity 10SG Strips Lot# URS7100027EX, Exp:11/2019 , URS7100140EX, Exp:11/2019 and URS8010016EX, Exp: 03/2020. The affected lots are not compatible for use in the Clarity Urocheck 120 Urine Analyzers which is why the recall is being performed.The letter also stated to ensure that communication is relayed to Clarity Diagnostics customers. Once the three lots have been secured, Clarity Diagnostics will follow up with the distributor regarding the disposal of the product. Clarity Diagnostics provided a telephone number for any customer concerns. Clarity Diagnostics has taken appropriate steps to assure that this problem will not reoccur in the future.Clarity Diagnostics also provided a letter to the distributor on or about 06/22/2018 entitled "URGENT PRODUCT RECALL". This letter was provided for distributor to notify its customers. This letter provided notification of a product recall regarding specific lots of Clarity Diagnostics - 10SG Urinalysis Reagent Strips. This notice has been issued because the affected lots are not compatible for use in the Clarity Urocheck 120 Urine Analyzers. Affected product first shipped April 5, 2018. Questions can be directed to the Clarity Diagnostics individual handing the recall---Ashish Parikh at Clarity Diagnostics (877) 722-6339. |  | 7536 | 7536 Bottles | Worldwide Distribution - US Nationwide in the states of AZ, MA, SC, WA. and the country of Austria. | Open, Classified |  |  | 7536 | Not_Computer | Not_Computer
Z-1166-2018 | 79342 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | ARROWg+ard Blue PLUS Pressure Injectable Multi-Lumen CVC Kit. |  |  | November 14, 2017 | 2017 | Arrow International Inc | Products do not contain the Central Venous Catheter Placement Procedures Note or the Central Line Insertion Checklist, which is a tool intended to guide clinicians in the use of best practices during central line insertions. | Labeling mix-ups | On November 16, 2017, an Urgent Medical Device Notification titled "ARROW PSI and Pressure Injectable Multi-Lumen CVC Kits" was issued to customers informing them of the issue, and urging all to take the following actions:1. Place a copy of this notification and the enclosed Central Line Insertion Checklist with each unit of affected product currently in your inventory.2. Complete the enclosed Acknowledgement Form and fax it to 1-855-419-8507, Attn: Customer Service or e-mail to recalls@teleflex.com.If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-866-246-6990. |  | 18124 | 18,124 total | Nationwide Distribution | Open, Classified |  |  | 18124 | Not_Computer | Not_Computer
Z-0936-2018 | 79352 | No | GAZ | Vacuum-powered body fluid suction apparatus. | General Hospital | General Hospital | 510(K) Exempt | Argyle Suction Tubing (Product # 8888301515) |  |  | October 24, 2017 | 2017 | Medtronic Minimally Invasive Therapies Group | Incorrect packaging. In one lot of Argyle Suction Tubing, the pouch may be incorrectly labeled as the Salem Sump Dual Lumen Stomach Tube. This product may have been distributed within the case of Argyle Suction Tubing. | Packaging process control | An URGENT MEDICAL DEVICE RECALL letter was issued to customers on October 24, 2017. This letter notified customers of the products subject to the recall and stated the following: The primary package labeling (Label closest to the product) may be incorrect. Customers can identify the affected product by consulting Attachment B. Please quarantine and return any unused products of the item codes and lot detailed above. Unused products from the affected item code and lots should be returned as described in the Required Actions section below. If you have distributed the Argyle" Suction tubing listed above, please promptly forward the information from this letter to those recipients. All unused products from the affected item codes and lots must be returned. If you have any questions regarding this communication, please contact your Medtronic representative or Customer Service at 800-962-9888, option 2. |  | 58954 | 58,954 | USA (nationwide) Distribution | Open, Classified |  |  | 58954 | Not_Computer | Not_Computer
Z-1384-2018 | 79646 | No | DSP | Intra-aortic balloon and control system. | Cardiovascular | Cardiovascular | 510(k) | Intra-Aortic Balloon Catheter Mega 7.5FR 30cc |  |  | February 21, 2018 | 2018 | Datascope Corporation | The U.S. Instructions For Use for the Mega Intra-Aortic Balloon (IAB) Catheter was missing the Contraindications section. | Error in labeling | Maquet/Getinge issued a recall letter on February 21, 2018 to affected customers, via FEDEX with confirmation of delivery, to notify them of the fact that the IFU is missing the Contraindications section and to provide the missing information. Affected customers are instructed to examine their inventory immediately to determine if they have affected product. An Addendum to the Mega IAB Catheter IFU, containing the Contraindications, is provided as an attachment in the recall letter. It is suggested that affected customers post copies of the Addendum near all affected lot numbers of Mega IAB Catheters. Customers are provided with Maquet/Getinge contact information should their facility require a full copy of the correct Mega IAB Catheter IFU. Lastly, affected customers are instructed to complete and sign a response form, acknowledging that they have received the notification, and return the form to Maquet/Getinge. |  | 655 | 655 | US Nationwide Distribution | Open, Classified |  |  | 16113 | Not_Computer | Not_Computer
Z-1172-2018 | 79536 | No | DXZ | Intracardiac patch or pledget made of polypropylene, polyethylene terephthalate, or polytetrafluoroethylene. | Cardiovascular | Cardiovascular | 510(k) | HEMASHIELD Knitted Double Velour Cardiovascular Fabric |  |  | January 03, 2018 | 2018 | GETINGE US SALES LLC | A probable mislabeling of the product: a different-sized product is referenced on the product packaging label (i.e. 2.5 cm (width) x 7.6 cm (length)) while the actual product inside the box is 2.5 cm (width) x 15 cm (length). | Labeling mix-ups | On January 3, 2018 an URGENT MEDICAL DEVICE RECALL NOTICE was issued to the (1) affected customer. A return authorization was issued with the recall letter, which is intended to be returned to the recalling firm. Questions or concerns can be directed to 1-800-528-7486 |  | 1 | 1 | FL | Terminated | July 03, 2018 | 181 | 1 | Not_Computer | Not_Computer
Z-1189-2018 | 78667 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Tapered Screw-Vent MTX, Dental Implant, sterile. Model TSVTB10. |  |  | November 29, 2017 | 2017 | Zimmer Dental Inc | Two lots of the Tapered Screw-Vent Implants, with different outer thread diameters, were inadvertently commingled during the manufacturing process and may be mislabeled. | Packaging process control | On 11/29/2017, customers were notified via letter. Instructions for distributors include to ensure that affected team members are aware of the recall, immediately locate and quarantine affected devices in inventory, notify consignees if product was distributed, make arrangements to immediately return affected devices, and complete and return the Attachment 1 - Inventory Return Certification Form. Instructions for clinicians include to review the recall notice, examine inventory for affected devices, complete and return the Attachment 2 - Certificate of Acknowledgement, and return all affected devices. |  | 995 | 995 distribution events total | Worldwide distribution US nationwide, Canada, Chile, Costa Rica, Denmark, France, Hungary, Italy, Japan, Lebanon, Portugal, Serbia, Spain, and Taiwan. | Terminated | February 08, 2019 | 436 | 995 | Not_Computer | Not_Computer
Z-1291-2018 | 79530 | No | FZP | Implantable clip. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | DS Titanium Ligation Clip Appliers, used in laparoscopic surgery |  |  | January 22, 2018 | 2018 | Aesculap Implant Systems LLC | Added a contraindication to IFU: Do not use the DS clips for living donor nephrectomies. Added other contraindications. Updated cleaning and sterilization instructions. | Labeling design | A letter and copy of the updated IFU was sent to affected customers the week of January 15, 2018. Product does not need to be returned. IFU is revised and put on the Aesculap Website for usage. https://www.aesculapusa.com/products/instructions-for-use. |  | 270 | 270 | US and Canada distribution | Open, Classified |  |  | 9891 | Not_Computer | Not_Computer
Z-2550-2018 | 80283 | No | OEZ | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | Enforcement Discretion | Tubing Pack Convenience Kit, Product Code 65217 |  |  | May 12, 2017 | 2017 | Terumo Cardiovascular Systems Corp | Kit was labeled with the incorrect Expiration Date of April 30, 2019. | Incorrect or no expiration date | On May 12, 2017 Terumo Cardiovascular Systems Corporation telephoned their one client and notified them of the labeling error and asked them to hold product. All product was immediately quarantined. |  | 60 | 30 cases/2 units (Total 60 units) | US state of FL | Terminated | October 26, 2018 | 532 | 60 | Not_Computer | Not_Computer
Z-2052-2018 | 79903 | No | FDX | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Hobbs Bronchial Cytology Brushes, Model Number 4202; 4204; 4206 |  |  | May 01, 2018 | 2018 | Hobbs Medical, Inc. | Device was marketed for uses outside the 510(k) clearance. | Labeling False and Misleading | The firm, Hobbs Medical, Inc., sent an "Voluntary Medical Device Recall Advisory Notice" letter to customers on 5/1/18. The letter described the product, problem and actions to be taken.Our records indicate that you have received the following Bronchial Cytology Brushes since the beginning of 2018:Please complete the quantity used and quantity returned section of the attached form and you may voluntarily return any unused device(s) referencing Return Authorization Number. Hobbs Medical will credit your account in a timely fashion.f you have any questions, please contact customer service at 800-344-6227. |  | 1820 | 1820 devices | US Nationwide Distribution to states of: AZ, CA, GA, IL, MD, NY, TN, TX, and VA. | Open, Classified |  |  | 1820 | Not_Computer | Not_Computer
Z-2108-2018 | 80061 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | ROSA(TM) Robotized Stereotactic Assistant, Version 2.5 |  |  | May 08, 2018 | 2018 | MEDTECH SAS | Replacement of units lacking an updated device approval. | No Marketing Application | Medtech would like to inform you of the planned removal of all ROSA 2.5 devices from the US market by 31st October 2018.The ROSA 2.5 design has been superseded by a newer generation product, ROSA Brain 3.0, and a decision has been made by Zimmer Biomet to exchange all ROSA 2.5 devices in the US market with ROSA Brain 3.0 devices. This decision has been taken in response to a regulatory clearance compliance issue, and there is no safety concern associated with this action.No specific action is required and the ROSA 2.5 can continue to be used at this time.A Medtech representative will be in contact to schedule the exchange of the RO 15 061 device with a ROSA Brain 3.0 device. For further questions, please call (574) 371-3071. |  | 29 | 29 units | US Distribution | Open, Classified |  |  | 29 | Not_Computer | Not_Computer
Z-2114-2018 | 80092 | Yes | BTO | Tracheostomy tube and tube cuff. | Anesthesiology | Anesthesiology | 510(k) | BivonaÂż Tracheostomy Tube Tracheostomy Tubes |  |  | April 02, 2018 | 2018 | Smiths Medical ASD Inc. | Carton labeling is printed with "Sterile" indicated labeling, however the device is not sold as sterile. | Error in labeling | Customers were notified via letter on approximately 04/02/2018. Instructions included to determine if there are any affected Customized BivonaÂż Tracheostomy Tubes in inventory, notify customers if the devices have been further distributed, arrange for the return of affected devices, and complete and return the acknowledgement response form. |  | 16 | 16 units | Domestic distribution to CA, FL, GA, ID, NJ, NM, OH, VA. International distribution to France. | Terminated | November 14, 2018 | 226 | 16 | Not_Computer | Not_Computer
Z-2057-2018 | 79978 | No | PPR | Arthroscope. | Orthopedic | Orthopedic | Enforcement Discretion | 3910-090-800 ASSY, PKG, 8.0MM X 90MM DRI-LOK THREADED CANNULA, 5 units per box. |  |  | December 21, 2017 | 2017 | Stryker Corporation | Inadvertent shipment of expired units. | Incorrect or no expiration date | The firm, Stryker, sent an "URGENT MEDICAL DEVICE RECALL NOTIFICATION" letter dated 12/27/2017 to its customers via mail on about 12/28/2017. Customers were initially notified via phone on 12/21/2017 and 12/22/2017. The letter described the product, problem and actions to be taken. The customers were instructed to inform those individuals in the organization who need to be aware of the recall; examine stock for affected devices; quarantine and discontinue use of recalled device; complete and return the Business Reply Form to Stryker Endoscopy via fax (408-855-6314) or email (endorecall@stryker.com), and if affected product is on hand segregate it and contact Stryker customer service at 1-800-624-4422 (Option 3) or email endocustomersupport@stryker.com to arrange for product return andissuance of credit or replacement (upon availability).If you have any questions, contact Director of Regulatory Affairs and Clinical Science at 408-754-2664 or email: to endorecall@stryker.com. |  | 2 | 2 boxes (5 units per box) | US distribution to Massachusetts only.. | Completed |  |  | 2 | Not_Computer | Not_Computer
Z-2567-2018 | 79485 | No | JDO | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Custom Made Implant System with Minimum Invasive Grower (MIG) component |  |  | December 21, 2017 | 2017 | Howmedica Osteonics Corp. | From March 2007 to October 2014, product did not contain updated instructions for use (IFU) to clarify the function of locking mechanism and warnings related to the locking screw. | Labeling design | On December 21, 2017 a FIELD SAFETY NOTICE was issued to customers notifying of the corrective action. An attachment was supplied with the letter titled "Minimally Invasive Grower - Instructions For Adjustment that is effective on and after October 14, 2014. Questions or concerns can be directed to your local sales representative. |  | 5 | 5 | NY, FL, KS | Terminated | October 05, 2018 | 288 | 5 | Not_Computer | Not_Computer
Z-2590-2018 | 80489 | No | OAF |  |  | Ear Nose &amp; Throat | PMA | Final Pack, Esteem II, Sound Processor Model 2001, containing 2 Sterile Wipes 002866-001 (which expired in April 2018). An Implantable Middle Ear Hearing Device. |  |  | June 28, 2018 | 2018 | Envoy Medical Corporation | Sterile Wipes packed in these four Esteem II Sound Processors had an earlier expiration date than the final pack and the Esteem II Sound Processors. The Sterile Wipes expired prior to the date labeled on the final pack. | Packaging process control | The firm notified its customers via email on 06/18/2018 and 06/19/2018. The 06/18/2018 email to a customer stated Envoy Medical would like to have the devices the customer has on site returned, as apparently, it was confirmed that each device is packaged with a wipe that shows expired just on the wipe only. Envoy will replace the recalled devices by overnight shipping replacements.The 06/19/2018 email to a customer stated it was discovered that the sound processors (primary and backup) were previously sent to the site have a wipe in them that show an expired date on the packaging. The email also stated this is something that doesn't affect the actual sound processor itself as those are good, but just to be on the safe side, make sure the 2 sound processors currently at the site (serial #s 213658 and 213729) are returned to Envoy Medical ASAP so we can have them repackaged with wipes within their expiration dates added to the sound processor packaging. Envoy will send out 2 new Sound Processors (primary and backup). Both emails contained Envoy Medical contact information for the consignees in case any questions or concerns arose. |  | 4 | 4 Units | US Distribution to states of: FL, OR | Open, Classified |  |  | 4 | Not_Computer | Not_Computer
Z-1350-2018 | 79559 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Spinal Jaxx Interbody Fusion Device, Model: 100001-511, Size: 9 H x 10 W x 25 L |  |  | January 18, 2018 | 2018 | Neuropro Spinal Jaxx | The implant size is incorrectly etched on the implant. The size listed on the box is correct. | Labeling Change Control | The firm called the consignee using a phone script. Additional information requested regarding whether any instructions were provided to the consignee regarding what to do with the recalled devices. |  | 6 | 6 | The devices were distributed in California. | Terminated | July 13, 2018 | 176 | 6 | Not_Computer | Not_Computer
Z-0505-2018 | 78897 | No | MCW | Intraluminal artery stripper. | Cardiovascular | Cardiovascular | 510(k) | Peripheral Diamondback 1.25 Solid OAD, a percutaneous orbital atherectomy system indicated for use as therapy in patients with occlusive atherosclerotic disease in peripheral arteries and who are acceptable candidates for percutaneous transluminal atherectomy. |  |  | September 14, 2017 | 2017 | Cardiovascular Systems Inc | Cardiovascular Systems, Inc. (CSI) is removing the products because it was identified that the labeling for two lots of product s were potentially inadvertently switched during production. | Labeling mix-ups | On 09/14/2017 the firm sent letters to their consignees explaining the recall and given the following instructions:Please immediately remove from use and return the affected Product Name devices to CSI. To do so:1) Complete the attached Customer Acknowledgement Forma. Per lot, document the quantity of all devices that are to be returned.b. Per lot, document the quantity of all devices that have been consumed.2) Place the Acknowledgement Form with the device3) Return the device and Acknowledgement Form to CSI using the pre-paid FedEx labela. If you would like, a CSI Sales Representative can return the device and form. Please contact the CSI Recall Coordinator to organize pick up and return of the product. If all products specified in this recall notice have previously been consumed, used or are no longer available for return to CSI, please complete the Customer Acknowledgment Form per the above, and return the form to CSI in the provided FedEx envelope.Contact Information: csifieldactionhotline@csi360.com651-259-2800 - Tel |  | 8 | 8 | FL, IL, MI, MN, MO, NY, SC, TX, and WI | Completed |  |  | 8 | Not_Computer | Not_Computer
Z-2599-2018 | 80330 | No | LTT | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | MicroScan Pos Combo Panel Type 43, Catalog number B1017-217for use in determining antimicrobial agent susceptibility and/or identification to the species level of rapidly growing aerobic and facultative gram-positive cocci, some fastidious aerobic gram-positive cocci and Listeria monocytogenes. |  |  | May 04, 2018 | 2018 | Beckman Coulter Inc. | The drug dilution sequence for Moxifloxacin printed on the panel is incorrect. As a result, incorrect MIC results for Mxf could be reported. | Labeling design | A recall notification letter was sent the week of 05/01/2018 via email, mail, and phone if needed. |  | 686 | 686 boxes | NJ, TX, IA, PA, FL, NE, OH, NY, CA, MO, OK, LA | Open, Classified |  |  | 686 | Not_Computer | Not_Computer
Z-0813-2018 | 79274 | No | LRY | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Surgical punch; sealed pouches packed in punch shelf box |  |  | October 23, 2017 | 2017 | Alto Development Corp | Surgical punches are being recalled due to potential packaging damage which may have occurred during transportation resulting in loss of sterility. | Package design/selection | A&E; Medical sent an Recall Notice Action Required letter dated October 20, 2017. Surgical Punches are being recalled due to potential packaging damage which may have occurred during transportation resulting in a loss of sterility. Recent testing identified certain shipping configurations which may cause damage to the sterile barrier pouch if the Surgical Punches experience excessive movement from multiple drops during transportA&E; is recalling all product currently on the market. Customer and Distributor Letters were sent with Acknowledgement Form. Product will be quarantined pending rework. For further questions, please call (800) 323-4035. |  | 411486 | 411,486 units | Worldwide Distribution - USA Distribution | Open, Classified |  |  | 411486 | Not_Computer | Not_Computer
Z-0919-2018 | 79206 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Sirus Drill 13mm diameter trauma instrument used for preparation of bone in Trauma, Hip and Shoulder surgeries.; Model Number: 02.00020.040 |  |  | October 03, 2017 | 2017 | Zimmer Biomet, Inc. | The Sirus Drill and Flexible Shaft instruments may not be adequately cleaned when utilizing the standard cleaning instructions potentially resulting in infection and subsequent complications. | Labeling design | Review the notification and ensure affected personnel are aware of contents. Assist Zimmer Biomet sales representative in locating and quarantining affected product. Sales representative will remove product from facility. Complete certificate of acknowledgement and return to Zimmer Biomet. Distributors are asked to locate and quarantine all affected product in inventory and from affected hospitals in their territory. |  | 127 | 127 | Distributed in 31 states: AL, AZ, CA, CO, CT, FL, GA, HI, ID, IL, IN, KS, KY, MA, MD, MI, MN, MO, NC, NJ, NM, NY, OH, OR, PA, TN, TX, VA, VT, WA, WI and the District of Columbia. | Terminated | August 20, 2018 | 321 | 584 | Not_Computer | Not_Computer
Z-2333-2018 | 80160 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | ECOLAB MICROTEK Rolled Probe Cover, 2.6 cm x 20 cm (1 in x 8 in), Ref PC0906, packaged 1 probe cover/pouch, 25 pouches/carton, 2 cartons/shipping case, Sterile, RX. |  |  | November 22, 2017 | 2017 | ECOLAB INC | The incorrect IFU was shipped with the product. | Labeling False and Misleading | The recalling firm issued letters dated November 21, 2017, informing the customer the incorrect IFU was shipped with the product and a correct version of the IFU is being provided with the letter. The consignee was requested to notify downstream customers to the user level and that a copy of the letter could be forwarded to enhance communication. A Recall Response Form was enclosed for completion to be returned via FAX or email as soon as possible. The letter also informs the consignee that the recalling firm will send additional replacement IFU's if they have product in inventory. For further questions, please call (651) 250-4385. |  | 65 | 65 cases | US Distribution to the states of : CA, IL, LA, SC, TX, and WI. There was no foreign/government/military distribution. | Completed |  |  | 65 | Not_Computer | Not_Computer
Z-2162-2018 | 80112 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Prima Plus Âż4.1 x 11.5mm Implant, Catalog Number 15737K |  |  | May 01, 2018 | 2018 | Keystone Dental Inc | The implants are mislabeled. | Error in labeling | The Firm sent an Urgent Medical Device Correction letter dated April 30, 2018. In this discrepancy, the 5 is missing from the implant length description. Please note this discrepancy is limited to the implant package label only. The patient chart labels are all correct. Keystone Dental has performed a thorough review and determined there are no health hazards associated with this issue. There is no issue with the product within the package and the product can be safely used without issue in accordance with the defined surgical protocols. There is no risk of confusion of product size. The diameter is correct. Prima Plus implants are only available in 11.5mm lengths. The catalogs, surgical protocols and manuals only define procedures and provide instrumentation for placing 11.5mm length implants. Additionally, the patient chart labels contained within the implant package are correct so that patient procedure history will be properly maintained. The issue is limited to the label on the outside of the product package only. The lot identified above is the only affected lot. No other lots are affected with this issue. The product referenced above can be used safely without issue and no further action is required. For further questions, please call (866) 902-9272. |  | 72 | 72 | US Distribution to the states of : CA, CT, FL, IN, MA, MD, MI, MN, MT, NJ, NV, NY, TX, and VA. | Terminated | September 14, 2018 | 136 | 72 | Not_Computer | Not_Computer
Z-2652-2018 | 80427 | No | FMG | Intravascular administration set. | General Hospital | General Hospital | 510(k) | 125" (318 cm) Y-Type Blood Set w/170 Micron Filter, Hand Pump, 3 ClaveÂż, Remv 3 Gang 4-Way Stopcock w/Baseplate, Rotating Luer, 3 Ext, Sterile |  |  | July 31, 2017 | 2017 | ICU Medical, Inc. | Some devices were labeled as Does not contain DEHP even though the product may contain small amounts of DEHP. | Labeling mix-ups | The firm, ICUmedical, notified their customers of the recall on about 07/31/2017 via"URGENT: Medical Device Recall Notification letter". Instructions included to inspect inventory on hand for affected devices, remove and return any affected devices, and complete and return the Recall Response Form to ICU Medical RecallCoordinator; email:.Recall@icumed.com; including labeling "Incorrect labeling" in subject or FAX: (801) 264-1755. Instructions also included to notify customers of the recall if devices were further distributed. Questions can be directed to ICU Customer Service Monday through Friday, 7 AM CST - 4:30 PM CT at 877-946-7747, option 9 or e-mail at ProductReturns@icumed.com |  | 160016 | 160,016 units total | Worldwide Distribution: US (nationwide) distribution to AL, AR, CA, CT, DE, FL, IA, IL, IN, KS, MA, MD, ME, MI, MO, NC, NJ, NY, OH, OR, PA, RI, TN, TX, UT, VA, WA, WV; and countries of:: Canada, Cyprus, Japan, Kuwait, MALAYSIA ,and South Africa. | Completed |  |  | 160016 | Not_Computer | Not_Computer
Z-1128-2018 | 79321 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | U2 Total Knee System Tibial Insert, Posterior Stabilized, #3, |  |  | November 24, 2017 | 2017 | United Orthopedic Corporation | UOC received a customer compliant report from our distributor on October 31, 2017which stated that contents of the package do not match the product labeling. | Labeling mix-ups | Urgent Field safety notice was distributed to customers on November 24, 2017 via e-mail. The letter identified the affected product, problem and actions to be taken. For further questions, please call (949) 328-3366. |  | 69 | 69 pieces | Worldwide Distribution - US Distribution and to the countries of: UK, Iran, Greece, Pakistan, Switzerland, France, China, Taiwan, Italy, Korea and South Africa. | Terminated | October 09, 2018 | 319 | 83 | Not_Computer | Not_Computer
Z-1117-2018 | 79268 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Apex Knee System, Tibial Baseplate Size 6 Left, non-porous - KC-2206L |  |  | November 20, 2017 | 2017 | OMNIlife science Inc. | The Apex Knee System Tibial Baseplate identified has an improper expiration date marked on the external packaging. The improper expiration date (2202-03) ismarked on the external packaging only. The correct lot number (2022-03) is present in the internal packaging and sterile batch documentation. | Incorrect or no expiration date | OMNIlife science, Inc. sent an Urgent Medical Device Advisory Notice dated November 20, 2017.Actions to be taken by the Customer/User"Immediately quarantine lot 26519 for KC-2206L"Return Product Using RMA# XXXX to OMNI."This advisory notice only includes LOT 26519. It does not affect other lots.You must acknowledge this notice by returning the attached Acknowledgement and Receipt Form and accompanying by FAX to 508-819-3390 or my email to CS@omnils.com. For further questions, please call (508) 824-2444. |  | 32 | 32 | Worldwide Distribution - US Distribution and to the country Spain. | Completed |  |  | 32 | Not_Computer | Not_Computer
Z-2142-2018 | 80064 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | IntelliVue MX40 Patient Monitor, Software Versions B.05, B.06, B.06.5X, Model 865351 |  |  | April 30, 2018 | 2018 | Philips Electronics North America Corporation | Five warning statements are missing from the instructions for use. | Error in labeling | Urgent Medical Device Correction notification letters were distributed on 4/30/18. The notification included the following instructions:Customers were asked to please review the information provided in the letter with all members of their staff who need to be aware of its contents. The letter provides information on the missing warnings and instructions for actions to be taken. The enclosed MX40 IFU Errata Sheet must be attached to the first page of Chapter 6 of the Instructions for Use for ready reference. Complete and return the attached Customer Reply Form. |  | 7178 | 7,178 | Worldwide Distribution -- US, to the states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY, and to the territories of Guam and Puerto Rico; and, countries of Argentina, Australia, Austria, Belgium, Canada, Chile, China, Croatia, Czech Republic, Denmark, Estonia, Finland, France, German, Ghana, Greece, Guadeloupe, Hong Kong, Iceland, India, Ireland, Israel, Italy, Japan, Latvia, Lebanon, Malaysia, Martinique, Mexico, Myanmar, Netherlands, New Zealand, Norway, Philippines, Poland, Portugal, Qatar, Russia, Saudi Arabia, Singapore, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, UK, UAE, and Vietnam. | Open, Classified |  |  | 37259 | Not_Computer | Not_Computer
Z-2163-2018 | 80125 | No | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | smith&nephew; GENESIS(R) II TIBIAL Punch, 13 MM, REF 71440408, QTY: (1) |  |  | May 04, 2018 | 2018 | Smith &amp; Nephew, Inc. | One lot of GENESIS II 13 mm tibial punches are 2 mm smaller than specification. | Error in labeling | The firm initiated the recall by email on 05/04/2018. Consignees are asked to inspect inventory, locate and quarantine affected, unused devices. Affected product will be returned to Smith &amp; Nephew. All consignees are asked to complete the Response Form, even if they do not have product to return. |  | 5 | 5 units | International only: United Arab Emirates, Spain, Mexico | Terminated | February 22, 2019 | 294 | 5 | Not_Computer | Not_Computer
Z-1589-2018 | 79770 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Allura Xper R8.x.25.5 and UNIZ systems (only with a FlexVision large screen monitor), Interventional fluoroscopic x-ray system |  |  | March 14, 2018 | 2018 | Philips Electronics North America Corporation | After continuous operation for more than one and a half days, the image on the large screen monitor may freeze for approximately 15 seconds after which the system will restore itself. | Software design | Field Safety Notices were distributed on 3/14/18. The notices instructed customers:To prevent occurrence of the issue customers shall restart the system at least once a day. Customer shall ensure that all staff with access to the affected systems are informed of the contents of this Field Safety Notice.A copy of this Field Safety Notice shall be placed together with the documentation of the system until the system has been corrected by Philips.The problems will be resolved by a software update, which is expected to be available by the second half of 2018.You will be notified by your local Philips representative when the software update is available for installation.If you need any further information or support concerning this issue, please contact your local Philips representative: Technical Support Line: 1-800-722-9377, 8am to 8pm (EST) Monday thru Friday. |  | 682 | 682 | Worldwide Distribution in the states of: AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NJ, NM, NV, NY, OH, OK, PA, SC, TN, TX, UT, VA, VT, WA, WI, WV, and WY. The product was distributed to the following countries: Argentina, Australia, Austria, Belgium, Brazil, Bulgaria, Canada, Chile, China, Croatia, Cuba, Czech Republic, Denmark, Estonia, France, Germany, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Kazakhstan, Republic of Korea, Malaysia, Mexico, Morocco, Myanmar, Netherlands, New Caledonia, Norway, Pakistan, Poland, Qatar, Romania, Russia, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, and United Kingdom. | Open, Classified |  |  | 682 | Computer | Computer
Z-0488-2019 | 81377 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Sedecal NOVA FA DR System. for diagnostic radiography. |  |  | July 18, 2018 | 2018 | Sedecal S.A. | A safety problem with installation and maintenance could lead to the breakage of the two steel cables that support the equipment to its roof anchor. Due to this, the equipment may fall and cause harm to the patient, user, or third parties. | Device Design | Initial communication from SEDECAL to all our distributors was performed on 07/18/2018 by email. The letter explained the problem and requested the consignee cease use of the unit if they observe an abnormal displacement in the vertical axis or any suspicious noise during the vertical movement. |  | 145 | 145 units | US and Canada | Open, Classified |  |  | 145 | Computer | Not_Computer
Z-0486-2019 | 81546 | No | IQA | Powered environmental control system. | Physical Medicine | Physical Medicine | 510(K) Exempt | Ascom Telligence Nurse Call System |  |  | November 02, 2018 | 2018 | Ascom (US) Inc. | There is an issue that could cause the Telligence System Gateway II, NGGTWY2 H, to stop communicating to the hallway devices connected to it. When this occurs, the dome lights connected to the Gateway II Ports A and B will start flashing, indicating that they are no longer communicating to the Gateway II. If this occurs, device failures that WOULD otherwise be annunciated WILL NOT be indicated at staff consoles that are configured to display failures. | Software Manufacturing/Software Deployment | Ascom Partners / Distributors were contacted by e-Mail, USPS Mail and Phone as necessary on 11/02/2018 to determine where the Telligence 5.0 Product is installed. |  | 57 | 57 units | AL, CA, CO, GA, IN, KY, LA, ME, MI, MO, MN, MS, NE, NV, OH, PA, RI, TX UT, VA, WI | Open, Classified |  |  | 57 | Computer | Computer
Z-1587-2018 | 79683 | No | JGS | Sodium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Atellica Solution and Atellica Sample Handler Prime for in vitro diagnostic testing of clinical specimens. |  |  | February 26, 2018 | 2018 | Siemens Healthcare Diagnostics, Inc. | When more than one immunoassay analyzer is connected to a single solution system, reporting and conversion factors may not be applied to all analyzers | Software design | Recall initiated 02/23/2018Firm released an update software (SW) version (V), 1.12.1, which corrects the behavior. A Mandatory Update (MU) campaign was launched to install the new SW V on all systems, whether or not they are impacted by this behavior. The MU campaign will be tracked to 100% completion globally. |  |  |  | FL andFranceSpainSwedenUnited KingdomGermany | Open, Classified |  |  |  | Computer | Computer
Z-0458-2019 | 81514 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | ARTIS Q biplane with Material 10848282- |  |  | October 19, 2018 | 2018 | Siemens Medical Solutions USA, Inc | Vibrations of the C-arm of the plane B may result in reduced image quality caused by the screw connection to the C-arm in the ceiling support | Device Design | Siemens issued on 10/19/18 an Urgent Medial Devoce Correction notification to via AX033/18/S advising users of the problem, health risk, steps to be taken. Siemens has initiated screw connection exchange of the ceiling bearing via Update Instruction AXOSS/17 /S. |  | 54 | 54 units | US Nationwide Distribution | Open, Classified |  |  | 83 | Computer | Not_Computer
Z-0402-2018 | 78830 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance iCT SP - Model 728311 |  |  | July 11, 2017 | 2017 | Philips Medical Systems (Cleveland) Inc | Numerous software issues with Brilliance iCT 4.1.6 version. Issues affect scan start position, data acquisition, exam stop, tube overheat, memory overflow, failure to give error message, length of acquisition, image volume and other functions. | Software design | A Customer Information Letter dated June 29, 2017, was sent to US consignees via certified mail. Philips Key Market representatives will distribute the letter internationally, based on local regulations. Additionally, a Field Service Engineer will contact each site to schedule a time to implement the Field Change Order. |  | 559 | 559 total | Worldwide distribution, including US nationwide. | Open, Classified |  |  | 559 | Computer | Computer
Z-0462-2019 | 81092 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Xper Flex Cardio Physiomonitoring System, Software version 1.6.0.0388 |  |  | September 24, 2018 | 2018 | Invivo Corporation | The Invasive Blood Pressure (IBP) numeric values will freeze on the display when any of the following functions are performed: change an IBP Site Label using the Site Label icon or the keyboard, activate or deactivate the mean pressure waveform display option (On/Off) using the Invasive Pressure Options Dialog or the keyboard, activate or deactivate the waveform hide option (On/Off) using the Invasive Pressure Options Dialog, activate or deactivate the pressure waveform filter option (On/Off) using the keyboard, and/or perform a Pullback using the Pullback icon or the keyboard | Software Design Change | Customers were notified on about 09/20/2018 via an "URGENT - Medical Device Correction" letter. An initial letter was sent to distributors with instructions to provide the Field Safety Notice to all customers the devices were further distributed to. The "URGENT - Medical Device Correction" letter provided guidance on how to identify the Service Number and serial number on the Flex Cardio device, which is on the bottom right corner of the back of the device. Actions to be taken by the customer/user until Philips contacts the customer to arrange for a software update to address this issue, is if the IBP numeric values become frozen and do not update in real-time, opening the Waveform Setup Dialog and then immediately pressing the OK button will restore the proper display of the IBP numeric values. Alternative methods of performing the functions listed in the Problem Description section that do not cause the problem to occur include the following: 1. Change an IBP Site Label using the Pressure Setup Dialog or a macro; 2. Activate or deactivate the mean pressure waveform display option (On/Off) by using the Waveform Setup Dialog, the Pressure Setup Dialog or a macro; 3. Activate or deactivate the waveform hide option (On/Off) using the Waveform Setup Dialog, the Pressure Setup Dialog, the keyboard or a macro; 4. Activate or deactivate the pressure waveform filter option (On/Off) using a macro. A Philips representative will contact each customer regarding their affected Flex Cardio devices. A software update will be required on both the affected Flex Cardio device and on the host workstations driver. These actions will be implemented free of charge by Philips. |  | 76 | 76 units | Distributed to the Netherlands | Open, Classified |  |  | 76 | Computer | Computer
Z-0404-2019 | 81355 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Medtronic Navigation 0-arm 02 Imaging System interventional fluoroscopic x-ray system: |  |  | September 17, 2018 | 2018 | Medtronic Navigation, Inc.-Littleton | O-arm 02 Imaging Systems correction to software version 4.1.0; new version of the user manual and an Image Acquisition System ("IAS") product label; replacing blown fuse of the new AC power input circuit | Software design | Medtronic Navigation issued Urgent Medical Device Correction letter on 9/27/18, notification letter advising, Medtronic field service personnel will install 4.1.0 software onto the designated consignee systems, update IFU and product labels as well as install a new designed AC power input circuit within the Mobile View Station {MVS).Questions regarding contact Medtronic Technical Support at 1-800-595-9709. |  | 641 | 641 units | Worldwide Distribution - US NationwideForeign:AUSTRALIAAUSTRIABAHRAINBANGLADESHBELGIUMBRAZILCANADACHILECOLOMBIADENMARKFINLANDFRANCEGERMANYHRICELANDINDIAIRELANDISRAELITALYJAPANJORDANKOREA, REPUBLIC OFLUXEMBOURGMARTINIQUEMEXICONETHERLANDSNEW ZEALANDPOLANDPORTUGALPUERTO RICOQATARSAUDI ARABIASINGAPORESOUTH AFRICASPAINSWEDENSWITZERLANDTAIWANTHAILANDTURKEYUNITED ARAB EMIRATESUNITED KINGDOM | Open, Classified |  |  | 641 | Computer | Computer
Z-0388-2019 | 81454 | No | REM |  |  |  |  | DeltaVision OMX SR Imaging System Model Number: 29115476 |  |  | September 10, 2018 | 2018 | GE Healthcare Biosciences | Non-conformance with the DeltaVision OMX SR, the interlocks are not wired in a redundant fashion to make the interlock system single-fault tolerant. | Device Design | The firm, GE Healthcare, sent an "Important Electronic Product Radiation Warning" letter to affected customers on September 10, 2018. The letter identified the affected product, problem and actions to be taken. GE Healthcare will, without charge, remedy the defect or bring the product into compliance. Your GE Service Representative will contact you to implement the correction. For questions contact your local GEHC service representative or call 425-657-1512 or email: trisha.koenke@ge.com. |  | 39 | 39 | Worldwide - US Nationwide Distribution | Open, Classified |  |  | 39 | Computer | Not_Computer
Z-0488-2018 | 78231 | No | EXS | Urine flow or volume measuring system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | BARD UROS AUTOMATED URINE OUPTUT AND TEMPERATURE MONITOR, MODEL BK10001M |  |  | September 29, 2017 | 2017 | C.R. Bard, Inc. | The possibility exists that use of the Uros Automated Urine Output and Temperature Monitor may lead to the loss of urine output measurement data and/or urine output measurement inaccuracies resulting in an impact on clinical treatment decisions. | Device Design | The firm, Bard Medical Division, sent an "URGENT: MEDICAL DEVICE PRODUCT RECALL" letter dated September 29, 2017 to their consignee. The letter described the product, problem and actions to be taken. The consignee was instructed to immediately examine your inventory and quarantine product; do not use or further distribute any affected product; complete and return the Recall &amp; Effectiveness Check Form by fax to 1-770-784-6460 or email a scanned copy to BMD.FieldAction@crbard.com, Attn: Recall Coordinator, even if you do not have any affected product and return the product. If you have affected product, please contact the BMD Customer Service via phone at 1-800-526-4455 option 5 and again option 5 or email BMD.FieldAction@crbard.com to obtain a Return Goods Authorization number (RGA #).If you have any questions, contact the Senior Director, Quality Systems &amp; Field Assurance at 770-784-6913 or email: Brenda.Sheikey@crbard.com or BMD Customer Service at 1-800-526-4455. |  | 2 | 2 units | US Distribution to the state of: FL | Open, Classified |  |  | 2 | Computer | Not_Computer
Z-0409-2019 | 81278 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension Thermal Chamber Motor Cable located in the following models: |  |  | September 28, 2018 | 2018 | Siemens Healthcare Diagnostics, Inc. | Thermal chamber motors have electrical cables with single insulation rather than the required double insulation. There is no exposed current due to the single insulation that could result in a potential electrical shock | Device Design | The firm, Siemens Healthcare, issued on 09/28/2018, "Urgent Medical Device Correction" DI-18-02.A.US (US customers)/Urgent Field Safety Notice DI-18-02.A.OUS ( for OUS customers), informing them of the issue and providing them with the steps Siemens is taking to correct the issue. A Siemens Service Engineer will schedule a Service visit to inspect the instrument and install a protective cover, if necessary, to correct the issue. Customers were instructed to perform the following actions:"Please arrange a time for your system upgrade when contacted by Siemens. "Return completed Field Correction Effectiveness Form to Siemens Healthcare Diagnostics via fax to the Customer Care Center  Technical Solutions at 312-275-7795 or please send a scanned copy of the completed form via email to uscctsfcaecfax.team@siemens-healthineers.com. "retain this letter with your laboratory records, and forward this letter to those who may have received this product.If you have any questions, please contact your Siemens Customer Care Center, your local Siemens Technical Support Representative or the Quality Engineer QT, LD, Commercial Product Quality at 302-631-0528 or email: richard.andersen@siemens-healthineers.com. |  | 277 | 277 units | Worldwide Distribution: US (Nationwide) and countries of: Algeria, Argentina, Australia, Brazil, Canada, China, Columbia, Croatia, Egypt, France, Germany, Great Britain, Greece, Hungary, India, Italy, Japan, Malaysia Mexico, Netherlands, Poland, Portugal, Russia, Saudi Arabia, Slovenia, South Africa, South Korea, Spain, Tanzania, Thailand, Tunisia, Turkey, and United Arab Emirates. | Open, Classified |  |  | 277 | Computer | Not_Computer
Z-0410-2018 | 78349 | No | NEU | Implantable clip. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | TUMARK Q, Length 12 cm, Diameter 18G, GTIN 04250195603081, REF 271500, STERILE |  |  | May 22, 2017 | 2017 | Somatex Medical Technologies GmbH | There is a potential risk of puncture to the packaging material which can compromise the devices sterility, making it non-sterile. | Device Design | On May 22, 2017, Somatex Medical Technologies distributed notices to their customers advising them of the reason for recall. Customers are advised to inspect stock, quarantine and not use affected product. Products will be exchanged free of charge and a replacement delivery will be arranged promptly. All customers are expected to complete and return the response form via fax. |  |  |  | NY - Only one US distributor | Terminated | August 30, 2018 | 465 |  | Computer | Not_Computer
Z-2185-2018 | 80018 | No | LFR | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Accu-Chek Aviva Plus Test Strip, Blood glucose monitoring system, Catalog nos. 06908217001, 06908268001, 06908349001,06908217001 |  |  | May 07, 2018 | 2018 | Roche Diabetes Care, Inc. | The recalling firm identified four US lots of test strips out of specification. Affected strips show increased potential for strip errors prior to applying blood, strips not detected once inserted, or biased results, which may not be detectable. In the case of a biased result, incorrect action could be taken by the user in response to the displayed blood glucose result. | Equipment maintenance | "On May 7, 2018, the firm sent Urgent Medical Device Correction letters to affected Retailers, Suppliers, and Patients. Patients were instructed to check their test strip supply lot numbers against the recalled lots. Any test strips from the affected lots must be discarded. To get replacement test strips, patients should use one of the following options:o For fastest replacement of your product, please submit your request athttps://notices.accu-chek.como Email Accu-Chek Customer Care at accu-chek.care@roche.com and includethe following information: your name, full mailing address, phone number,test strip lot number and number of affected test strip vials.o For questions regarding this recall, contact Accu-Chek Customer Careline at 1-800-358-4866 Monday through Friday between the hours of 8:00 amand 8:00 pm Eastern Standard Time.Suppliers were instructed to discontinue distribution and dispose of any remaining stock of the affected lots immediately. Suppliers were also asked to forward recall notifications to their customers, including both supplier- and consumer-level customers. Suppliers were asked to complete a faxback form indicating the total number of vials disposed of by the supplier and supplier customers. If you are a supplier and have questions about the information in this notification, please contact your Roche Account Manager.Retailers were instructed to discontinue sale or distribution and dispose of remaining affected lots in stock. Retailers were also asked to provide recall notification to their consumer customers. Retailers may note the quantity disposed and contact their Suppliers to arrange for credit. For questions please contact your wholesaler or the Accu-Chek Customer Care line at 1-800-358-4866 Monday through Friday between the hours of 8:00 am and 8:00 pm Eastern Standard Time." |  | 295248 | 295, 248 | Worldwide Distribution - US Nationwide distribution. Foreign distribution to ArgentinaArmeniaAustraliaAustriaAzerbaijanBahamasBahrainBarbadosBelgiumBrazilBulgariaCanadaChileChinaColombiaCosta RicaDenmarkDominican RepublicEl SalvadorFinlandGeorgiaGermanyGhanaGreeceGuatemalaHondurasHong Kong (SAR China)HungaryIndiaIndonesiaIsraelItalyJamaicaJapanJordanIrelandKenya Korea Kuwait Lebanon Lithuania Luxembourg Malaysia Malta Mexico Myanmar Netherlands New Zealand Norway Oman Pakistan Philippines Poland Russian Federation New Zealand Saudi Arabia Serbia Singapore Slovenia South Africa Spain Sweden Switzerland Taiwan (Republic of China), Thailand, United Arab Emirates, United Kingdom, Vietnam, Portugal" | Open, Classified |  |  | 295248 | Computer | Not_Computer
Z-1151-2018 | 78999 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AutoMate 2550 "High Speed" Catalog Number ODL25255; AutoMate 1250 "Standard" Catalog Number ODL25125. Chemistry analyzer sample processing and sorting system. |  |  | January 18, 2018 | 2018 | Beckman Coulter Inc. | A misaligned pipette tip sensor poses the remote risk of cross-contamination when failing to cause a no-tip error, resulting in pipetting patient samples without pipette tips. | Device Design | A recall notification letter was sent to affected US consignees on 1/19/2018 notifying them of the issue. In addition, the recalling firm is instructing field service engineers to inspect the tip present sensor for misalignment at the next service visit. |  | 73 | 73 units in US | Worldwide distribution. US nationwide, Algeria, Australia, Austria,Belarus, Belgium, Bosnia and Herzegovina,Brazil, Canada, China,Croatia, Cyprus, Czech Republic,Denmark, Finland, France,Germany, Gibraltar, Greece,Hong Kong, Hungary, India,Ireland, Israel, Italy,Korea, Kuwait, Latvia,Lithuania, Macao, Mayotte,Netherlands, New Zealand, Norway,Panama, Poland, Portugal,Russia, Saudi Arabia, Slovakia,Slovenia, South Africa, Spain,Sweden, Switzerland, Taiwan,Turkey, and United Kingdom. | Open, Classified |  |  | 73 | Computer | Not_Computer
Z-2053-2018 | 79496 | No | NBC | B-type natriuretic peptide test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA Centaur BNP (B-type Natriuretic Peptide), Catalog Numbers: 02816138 (100 tests), 02816634 (500 tests) |  |  | December 18, 2017 | 2017 | Siemens Healthcare Diagnostics, Inc. | In certain scenarios the ADVIA Centaur CP, XP and XPT immunoassay analyzers may exhibit reduced dilution recovery or become inoperable when used with Multi-Diluent 15 and certain lots of BNP and TSH3RUL Assays. | Device Design | Consignees are asked to return the response form and follow the instructions pertaining to their system and assays: Actions for ALL ADVIA Centaur Systems: Do not perform onboard dilutions on any ADVIA Centaur system if any of the affected kits are present on the system. Onboard dilutions for the ADVIA Centaur BNP and/or TSH3UL assay(s) are only supported when using ADVIA Centaur BNP kit lots ending in 210 and above and ADVIA Centaur TSH3UL kit lots ending in 310 and above. Actions for ADVIA Centaur XPT system: If you intend to perform onboard dilutions for the ADVIA Centaur BNP and/or TSH3UL assays, you must restart the workstation after scanning the master curve card for the BNP and/or TSH3UL assay kit lot, before ordering an onboard dilution for these assays. Please perform the attached steps to restart the workstation. Actions for ADVIA Centaur CP system: If any of the affected kit lots have been in use on the system, you must manually track the time that the Multi-Diluent 15 has been in use on board the system. Replace the ancillary pack of Multi-Diluent 15 every 7 days when used, until a follow-up communication is received from Siemens stating this issue has been resolved. |  | 21687 | 21,687 | Distributed in the following U.S. states: AL, CA, DE, GA, IA, IL, IN, KY, MA, MI, MS, MT, NC, ND, NJ, NM, NV, NY, OH, PA, RI, TN, TX, VA, WA, WV, WY and Puerto Rico. Distributed in the following OUS countries: AustraliaBelgiumÂżÂżÂżÂżÂżÂżÂżÂżCanadaChileFranceÂżÂżÂżÂżÂżÂżÂżÂżÂżGermanyÂżÂżÂżÂżÂżÂżÂżÂżGreeceÂżÂżÂżÂżÂżÂżÂżÂżÂżIsraelItalyÂżÂżÂżÂżÂżÂżÂżÂżÂżÂżJapanLatviaÂżÂżÂżÂżÂżÂżÂżÂżÂżLuxembourgÂżÂżÂżÂżÂżNetherlandsÂżÂżÂżÂżNew ZealandNorwayÂżÂżÂżÂżÂżÂżÂżÂżÂżPolandÂżÂżÂżÂżÂżÂżÂżÂżÂżPortugalÂżÂżÂżÂżÂżÂżÂżRomaniaÂżÂżÂżÂżÂżÂżÂżÂżSingaporeSlovakiaÂżÂżÂżÂżÂżÂżÂżSouth AfricaSpainÂżÂżÂżÂżÂżÂżÂżÂżÂżÂżSwitzerlandÂżÂżÂżÂżU.A.E.ÂżÂżÂżÂżÂżÂżÂżÂżÂżUnitedÂżKingdomÂż | Open, Classified |  |  | 95902 | Computer | Not_Computer
Z-2045-2018 | 80162 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | nordicBrainEX |  |  | December 21, 2012 | 2012 | NordicNeuroLab AS | There is an error related to the relative geometry between fiber tracts in a fiber group and images.The error occurs in the following situations: A group of fibers has been selected using the VOI functionality and put into a fiber group. This is done with images series A visualized in the 3D viewer. Then a new images series B is visualized in the 3D viewer. If image series A and B have the same geometrical resolution (pixel sizes and slice distances), the coregistration matrix taking A to B will not be applied to the fiber group, and therefore the fiber group will in general not be positioned correctly on B in the 3D viewer. Furthermore, if the fiber group is exported as a new image series the fibers could be misplaced, and similarly if the fibers are visualized in the MPR on B, the positioning will not be correct. The misalignment will be equal to the rotations/translations necessary to align A with B. | Software design | The firm distributed the recall notification letter by email. The letter identified the affected product, problem and the actions to be taken. For questions contact NordicNeurolab Customer Service or if you would like to arrange a return customerservice@nordicneurolab.com. |  | 36 | 36 licenses | Worldwide Distribution - US Nationwide - Saudi Arabia, Brazil, Sweden, Spain, Russia, Belgium, Norway, South Korea, Japan | Open, Classified |  |  | 36 | Computer | Computer
Z-0507-2018 | 78754 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | The ReNew Reusable Grasper and Dissector Tips are to be used with the ReNew Laparoscopic Handpieces and they are indicated for cutting, grasping, dissecting, and coagulation of tissue in endoscopic and laparoscopic surgical procedures. |  |  | July 13, 2017 | 2017 | Microline Surgical, Inc. | The heat-shrink insulation tube may split during autoclave sterilization/re-sterilization. If a split occurs, the heat-shrink could fall off during a surgical procedure and if cautery is needed, there is a potential for burn due to insulation failure. | Device Design | The firm, Microline Surgical, Inc., sent a "URGENT: MEDICAL DEVICE RECALL" letter dated July 13, 2017 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately determine if you are in possession of the product and stop use; quarantine all products both new un-opened Reusable Tips AND Reusable Tips that have been used or are in current use (regardless of whether you have observed the malfunction); If you are a distributor or further distributed this product, we require that you immediately remove affected products from inventory and stop sales and shipping; identify customers immediately and notify them of this product recall; and all customers/distributors complete and return the MEDICAL DEVICE RECALL RETURN RESPONSE Acknowledgement and Receipt Form (even if you have never received any of these products or have none in stock) by E-MAILING A SCANNED COPY TO: ReusableTipRecall@microlinesurgical.com or Fax to: 978-522-8490, ATTN: Reusable Tip Recall, or MAIL TO: Microline Surgical, Inc, 50 Dunham Road, Suite 1500, Beverly, MA 01915, ATTN: Recall Department.Microline has discontinued manufacture of the affected ReNew Reusable Tips and has initiated a ship-hold and new orders-hold for these products.For more information, please email ReusableTipRecall@microlinesurgical.com or contact Director of RA &amp; QA at 978-922-9810 or email: mleighton@microlinesurgical.com. |  | 61769 | 61,769 units | Worldwide Distribution-US (Nationwide) and countries of: Australia, Brazil, Canada, Chile, China, Colombia, El Salvador, Hong Kong, India, Indonesia, Malaysia, Mexico, Panama, Peru, Argentina, South Korea, Sri Lanka, Taiwan, Thailand, Uruguay &amp; Vietnam. | Open, Classified |  |  | 61769 | Computer | Not_Computer
Z-2266-2018 | 80173 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | CARESCAPE R860 ventilators with software version 10SP05 |  |  | December 08, 2017 | 2017 | Datex-Ohmeda, Inc. | Potential for the display processor to experience anunexpected failure. | Software change control | On December 12, 2017 GE Healthcare sent letters to their consignees stating that There is a potential for the display processor to experience an unexpected failure. If this failure occurs, the device screen will temporarily go blank and will then display an Error Code EC01 XX . The blank display results in loss of monitored data being displayed to the user. There have been no injuries reported as a result of this issue. You may continue to use the system. Device performance is not affected on devices where no such display failure occurs.In the event this issue occurs, an audible alarm will sound. Mechanical ventilation will continue at previous settings. A device restart is required to clear the alarm condition. Consider using appropriate alternative means of ventilation such as a self-inflating bag or another ventilator to prevent patient injury while restarting the device.GE Healthcare will correct all affected products at no cost to you. A GE Healthcare representative will contact you to arrange for the correction.If you have any questions or concerns regarding this notification, please contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 349 | 349 (US) + 2,051 (OUS) | Worldwide Distribution - USA (nationwide) Distribution to the states of : AL, AZ, CA, GA, ID, IL, IN, MS, NC, NJ, NK, NV, OH, OK, RI, TX, and WI., and to the countries of : Albania, Argentina, Australia, Bahrain, Bangladesh, Belgium, Bolivia, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, Chile, China, Colombia, Croatia, Cyprus, Czech Republic, Ecuador, El Salvador, Estonia, Ethiopia, Egypt, Finland, France, Germany, Ghana, Hungary, India, Indonesia, Italy, Jordan, Kazakhstan, Kenya, Kuwait, Lebanon, Malaysia, Mexico, Morocco, Myanmar, Nepal, New Zealand, Nigeria, Norway, Pakistan, Palestine, Peru, Philippines, Poland, Russia, Saudi Arabia, Senegal, Singapore, South Africa, Spain, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Uruguay, and Vietnam | Open, Classified |  |  | 349 | Computer | Computer
Z-0405-2018 | 78911 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo Imaging version V31 (model # 10014063) is a Picture Archiving and Communication System (PACS) intended to display, process, read, report, communicate, distribute, store and archive digital medical images, including digital mammography images. It supports the physician in diagnosis and treatment planning. |  |  | July 25, 2017 | 2017 | Siemens Medical Solutions USA, Inc. | Siemens is releasing a letter to inform about potential data loss relevant to diagnosis. | Software design | Siemens sent an Urgent Customer Advisory Notice Letter to all customers. All versions before version VB36E_HF01 can be configured to behave like the version VB36E_HF01. Please contact your local CSE to configure this setting and to evaluate an eventual minor performance impact to the system.If the device has been sold and therefore no longer in your possession, please forward this notice to the new owner. We would also request that you inform us of the identity of the devices new owner where possible. Please observe this safety notice and comply with the corresponding measure until the update has been fully completed. For further questions, please call (610) 448-6461. |  | 36 | 36 units. | USA (nationwide) Distribution to the states of: MA, CA, NC, WI, NY, NE, PA, AL, NJ, MO, MI, TX, MS, KS, LA, CT, MD, LA, CO, OK, MN and TN. | Terminated | July 23, 2018 | 363 | 36 | Computer | Computer
Z-0470-2019 | 81317 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | BrightMatter Guide, Product Number SYN-0026, Versions: 1.5.0, 1.5.1 and 1.5.2 |  |  | October 09, 2018 | 2018 | Synaptive Medical Inc | This recall has been initiated due to a software defect found in the Guide System software when used with NICO BrainPath ports under certain circumstances. When one trajectory is set with a blue port and another trajectory is set with a gold port, the software defect is triggered when the user switches between these trajectories during a surgical procedure. | Software design | Synaptive Medical sent an Urgent Medical Device Recall letter dated October 9, 2018 via registered priority mail . The letter identified the affected product, problem and actions to be taken until a software patch is released. Customers were requested to complete and return the enclosed Acknowledgement and Receipt form to Synaptive Medical Inc. For questions call 1-844-462-7246. |  | 37 | 37 units | Worldwide - US Nationwide Distribution and the country of Canada | Open, Classified |  |  | 37 | Computer | Computer
Z-1614-2018 | 80020 | [6 empty] | 11-VI Incorporated Laser diode component |  |  | December 13, 2017 | 2017 | Ii - Vi Incorporated | Affected diodes incorporated in the laser sights emitted greater than the 5 mW limit. | Device Design | Please note that II-VI Inc. is responsible for the correction of all noncompliant products should the reported epoxy processing defect adversely affect any other laser diodes manufactured by II-VI, Inc.The single customer, Crimson Trace Corporation CTC was formally notified by you in writing. The defective laser diodes, models 25-0655 &amp; 25-0657 will be or have already been replaced at no cost to CTC. The replacement will be done in batches and is expected to take approximately six to nine months, or no later than January 31, 2019. Defective laser diodes will be either returned to II-VI or destroyed at CTC. For further questions, please call (724) 352-4455. |  | 323972 | 323,972 | US Distribution to Oregon. | Open, Classified |  |  | 323972 | Computer | Not_Computer
Z-1585-2018 | 79439 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MHI-TM2000 Linear Accelerator System |  |  | December 08, 2017 | 2017 | Hitachi Ltd., Medical System Operations Group | Due to a system controller software anomaly, the patient positioning deviation correction may not be applied and may result in the wrong part of the patient being irradiated. | Device Design | Hitachi sent a Field Safety Notice letter dated November 28, 2017 to affected customers. The letter identified the affected product, problem and actions to be taken. The notification letter includes instructions for the user on how to verify patient position even with the software issue. The letter also explains that correction software is being developed with a tentative release date of, "End of January, 2018.". A customer service representative will contact affected customers about installing the software. For questions contact your our Customer Service Representative. |  | 1 | 1 | US Distribution in the state of NY. | Open, Classified |  |  | 1 | Computer | Computer
Z-0391-2019 | 80984 | Yes | LDF | Cardiovascular permanent or temporary pacemaker electrode. | Cardiovascular | Cardiovascular | 510(k) | TB Series Bipolar Pacing Leads, French Size 4F, 5F, 6F for Curve Type Straight, Atrial J, 60 degree Curve, and Right Heart. TB is packaged for sterile single use in a double tray. The lead is placed in a PETG inner tray that is sealed with a Tyvek inner lid. The sealed inner tray is placed in an outer PETG tray. The outer tray is sealed with an outer Tyvek lid. An inner label is placed on the outer Tyvek lid. The sealed trays are placed in a labeled white shelf box with on Instruction for use. Packed on lead per outer white shelf box. |  |  | September 27, 2018 | 2018 | Oscor, Inc. | During the use of some TB - Temporary Bipolar Pacing Leads, featuring the 2mm unshrouded connectors, the connector cap housing may slide and potentially expose the connection wire. In some instances, this may cause the wire to be more susceptible to loss of connectivity or breakage during movement of the cables causing interruption of the pacing system. | Device Design | A press release for the recall was issued on 09/27/2018. Additionally, on about 09/27/2018, Oscor notified customers via email and mail with a "PRODUCT RECALL - ACTION REQUIRED" letter. For distributors, instructions included to examine inventory for affected devices, quarantine any affected devices, notify customers if the affected devices have been further distributed, retrieve those affected devices from customers, and complete and return the attached Exhibit A. Any affected devices on hand and at customers should not be destroyed, but returned to Oscor. For hospitals, instructions included to examine inventory for affected devices, quarantine any affected devices, and complete and return the attached Exhibit A. Any affected devices on hand and at customers should not be destroyed, but returned to Oscor. Questions can be directed to TB Recall Coordinator at 727-937-2511 or e-mail TB@Oscor.com. Oscor business hours are Monday to Friday from 8:30AM to 5:30PM Eastern Time. On 11/07/2018 Oscor expanded the recall and notified customers via "PRODUCT RECALL - ACTION REQUIRED" letter and press release. The expansion includes expired devices distributed by Oscor from 12/21/2011 to 04/30/2015, to result in affected devices distributed from 12/21/2011 to 05/17/2018. Instructions for distributors and hospitals remain the same as the original notification. |  | 36688 | 36,688 units | Worldwide Distribution - US Nationwide in the states of AL, AZ, CA, CO, FL, HI, IN, KS, KY, LA, MA, MD, MI, MO, NC, NE, NJ, NM, NV, NY, OH, OK, OR, PR, RI, TN, TX, UT, VA, WA. Countries of AUSTRIA, BENIN REPUBLIC, BOLIVIA, BRAZIL, BULGARIA, CANADA, COLOMBIA, COSTA RICA, DOMINICAN REPUBLIC, ECUADOR, EGYPT, EL SALVADOR, GEORGIA, GERMANY, GREECE, GUATEMALA, INDIA, IRAN, IRAN, ISRAEL, ITALY, JORDAN, KAZAKHSTAN, KUWAIT, LATVIA, LEBANON, MACEDONIA, MAURITIUS, MEXICO, NEW ZEALAND, PAKISTAN, PANAMA, PERU, PHILIPPINES, POLAND, RUSSIA, SAUDI ARABIA, SENEGAL, SOUTH AFRICA, TAIWAN, TURKEY, UKRAINE, UNITED ARAB EMIRATES, UNITED KINGDOM, VENEZUELA, VIETNAM. | Open, Classified |  |  | 36688 | Computer | Computer
Z-1448-2018 | 79428 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | SOMATOM Perspective 16 |  |  | February 15, 2018 | 2018 | Siemens Medical Solutions USA, Inc | to inform customers of possible incorrect tube current calculations by the CARE Dose4D algorithm for head scans based on p.a. (posterior-anterior) or a.p. (anterior-posterior) topograms. Depending on the geometrical shape of the skull bone, it may happen in rare cases that the calculated dose distribution is not appropriate and could lead to unnecessary radiation exposure | Software design | The Center for Devices and Radiological Health (CDRH) understands that the CAP involving 1301 units will consist of the following:1. You will contact customers and initiate the actions to correct the failure to comply issue.2. You will provide a software update to correct tube current calculations by the CARE Dose4D algorithm for head scans.3. You will follow-up with consignees to verify the effectiveness of the software update.CDRH approves the CAP subject to the following conditions:1.Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR 1003.21 and 1003.22. This office and the Food and Drug Administration (FDA) district office coordinator noted below are to be included in the notification.2. Any future customer notification letters to be sent must comply with the requirements of 21 CFR 1003.21. This includes, but is not limited to, the prescribed packaging and the statement that you will remedy the defect or bring the product into compliance at no charge.You may proceed or continue with implementation of the CAP (subject to the conditions noted above). For further questions, please call (610) 219-4834. |  | 1301 | 1301 | USA (nationally) Distribution. | Open, Classified |  |  | 1301 | Computer | Computer
Z-1541-2018 | 79421 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Specimen Gate Screening Center, Part Number 5002-0500, All software versions from 1.0 to 1.8 (current software version) |  |  | November 17, 2017 | 2017 | PerkinElmer Life and Analytical Sciences, Wallac, OY | Potential errors in patient results generated by the Screening Center product that include both false negative and false positive results. | Software design | PerkinElmer sent an Urgent Medical Device Correction letter dated November 14,2017.Instructions:1. Implement additional control measures concerning demographic information, by taking one of the following actionsa. Fully enter demographics before loading specimens on instruments (i.e. specimens can be punched, but the assay plates cannot be loaded to the measurement instruments).b. If the above mentioned is not possible or if previously entered demographic information changes after loading the assay plates on the measurement instruments, recalculate all your results in Specimen Gate laboratory Result Viewer prior to assay acceptance.2. Follow unreported specimens with available tools to avoid any delay in reporting screening results. For further questions, please call (317) 418-1735 |  | 8 | 8 | Worldwide Distribution - US Distribution to the state of Georgia., and to the countries of : Belgium, Canada, Denmark, Italy, and United Kingdom. | Open, Classified |  |  | 8 | Computer | Not_Computer
Z-0469-2019 | 81187 | No | KHO | Fluorometer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AIA-900 Analyzer, Product code 022930, 022930R |  |  | October 05, 2018 | 2018 | Tosoh Bioscience Inc | Tosoh Bioscience has become aware of a potential issue with the pickup arm and pickup arm motor on the AIA-900 Analyzer. If the pickup arm or pickup arm motor malfunctions, the analyzer will cease to function and result in an inability to test patient samples. This may cause a delay in testing and reporting patient results. Tosoh has received fifty-one (51) complaints related to this issue with no serious injuries reported. | Device Design | On October 5, 2018, the firm notified its affected customers via an Urgent Medical Device Recall letter. The letter informed customers of the potential product issue with all active AIA-900 analyzers.The firm's service procedures have been updated to proactively troubleshoot and correct the pickup arm andpickup arm motor during installations and scheduled preventive maintenance visits. At the next scheduledpreventive maintenance visit a Tosoh field service engineer will verify the pickup arm motor current andadjust as needed. Adherence to preventive maintenance schedules provides effective mitigation againstpotential failure of these components.Immediate Actions to be taken by the Customer/User" Thoroughly review the content of this letter." Continue to use the instrument. A Tosoh representative will contact you to schedule your nextpreventive maintenance visit.o If a failure of the pickup arm is observed or the analyzer stops functioning, contact TosohTechnical Support for an expedited service visit." Complete and return the attached Acknowledgement Form by any of the listed methods." File this notice with your laboratory records and forward this information to others in your laboratory.If you have obtained unexpected test results, delays in test result reporting, or received any complaints ofillness or adverse events associated with the use of the analyzer, please contact Tosoh TechnicalSupport 24 hours a day, seven days a week at (800) 248-6764.Should you have any questions regarding this medical device recall, please feel free to contact BernadetteO'Connell at (800) 248-6764 or by email at bernadette.oconnell@tosoh.com. She will be available toanswer any questions Monday - Friday, from 9:00 AM to 5:00 PM (PST). |  | 393 | 393 | Distributed nationwide. Foreign distribution to Canada and Latin America. | Open, Classified |  |  | 393 | Computer | Not_Computer
Z-0487-2018 | 78918 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | GE HEALTHCARE CARESCAPE Central Station (CSCS) software version 2.0.2 |  |  | January 11, 2018 | 2018 | GE Medical Systems Information Technologies, Inc. | CARESCAPE Central Station (CSCS) software version 2.0.2 units may experience unexpected NO COMM (No Communication) and network communication issues after boot-up or system restart. | Software in the Use Environment | The firm, GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION" letter and email to its consignees on 01/11/2017. The letter described the product, problem and actions to be taken. The consignees were instructed to ensure that all potential users in your facility are made aware of this safety notification and the recommended actions and follow the safety instructions. GE Healthcare will correct all affected products at no cost to you. A GE Healthcare representative will contact you to arrange for the correction.If you have any questions regarding this notification, please contact Technical Support at 1-800-558-7044 or your local Service Representative. |  | 1803 | 1803 units | Worldwide Distribution: US (nationwide) and countries of: Algeria, Australia, Austria, Canada, Costa Rica, Czech Republic, Denmark, Egypt, Finland, France, Germany, Hungary, Iceland, India, Indonesia, Iraq, Ireland, Italy, Japan, Jordan, Korea, Republic Of, Kuwait, Lebanon, Mexico, Netherlands, New Zealand, Norway, Poland, Qatar, Russian Federation, Saudi Arabia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, Uganda, United Arab Emirates, United Kingdom, Uruguay, and Viet Nam. | Open, Classified |  |  | 1803 | Computer | Computer
Z-1181-2018 | 79452 | No | FOX | Infusion stand. | General Hospital | General Hospital | 510(K) Exempt | Rotating IV Pole |  |  | December 01, 2017 | 2017 | Ohmeda Medical | There is the potential for the Rotating IV Pole to fall. A fall could result in an injury to a bystander holding the patient or to the operator. | Device Design | On December 1, 2017 an URGENT MEDICAL DEVICE CORRECTION letter was issued to customers notifying of the recall. The letter requested customers to discontinue use of the affected Rotating IV poles and remove from Service. Place affected Rotating IV poles in a quarantine marked area. A GE Service Representative will contact you to arrange for the removal &amp; replacement of a new IV pole. If you have any questions or concerns regarding this notification, please contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. For further questions, please call (410) 888-5376. |  | 1179 | 1,179 (1,004 US; 175 OUS) | USA (nationwide) | Open, Classified |  |  | 1179 | Computer | Not_Computer
Z-0498-2018 | 78756 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | cobas e 411 Immunoassay Analyzer-Model/Catalog/Part number: 04775279001/04775279973 - cobas e 411 analyzer (disk system) and 04775201001/04775201973 - cobas e 411 analyzer (rack system), medial device listing number: E116019 cobas e 411 |  |  | July 05, 2017 | 2017 | Roche Diagnostics Corporation | A software malfunction can occur on the cobas e 411 and Elecsys 2010 analyzers in the Sample&amp; Control data file which may lead to a potential data mismatch. | Software design | All consignees were notified by Urgent Medical Device Correction sent via UPS Ground (receipt required) on July 5, 2017.The customer is advised to perform the following actions: cobas e 411: Until the software update is available, use the workaround outlined in the Urgent Medical Device Correction (UMDC). Elecsys 2010: Use the workaround outlined in this Urgent Medical Device Correction (UMDC). Non-responding accounts will be defined as users for which confirmation of receipt of the Urgent Medical Device Correction (UMDC) cannot be obtained. These accounts will be monitored on an ongoing basis and follow-up attempts will be made accordingly. Workaround for customers: he customer is advised to perform the following actions: cobas e 411: Until the software update is available, use the workaround outlined in the Urgent Medical Device Correction (UMDC). Elecsys 2010: Use the workaround outlined in this Urgent Medical Device Correction (UMDC). For further questions, please call (317) 521-7259 |  | 923 | 923 instruments | Nationwide Distribution | Open, Classified |  |  | 2356 | Computer | Computer
Z-0413-2019 | 81182 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | Touchscreen display kit. Sold under the following product names: TOUCHSCREEN DISPLAY KIT 11.X ENGLISH FOR USA, TOUCHSCREEN DISPLAY KIT 11.X ENGLISH, TOUCHSCREEN DISPLAY KIT 11.X GERMAN, |  |  | September 27, 2018 | 2018 | GE Healthcare, LLC | Upgraded Aisys anesthesia devices and Aisys CS2 anesthesia devices with PSVPro Spontaneous Breathing Modes do not provide the PSVPro ventilation mode feature of transitioning to SIMV PCV backup ventilation mode. PSVPro Spontaneous Breathing Modes do not provide the feature of transitioning to SIMV PCV backup ventilation mode for a patient who stops spontaneous breathing if the PSVPro ventilation mode is in use prior to the clinician starting the Cycling Procedure. | Software design | GE Healthcare notified customers on about 09/27/2018 via "URGENT MEDICAL DEVICE CORRECTION" letter. The letter informed customers that they could continue to use the anesthesia device. The issue can only occur if the patient stops spontaneous breathing in PSVPro ventilation mode after completion of lung recruitment breaths. Stopping the Cycling Procedure resolves the issue. Continue to monitor patient as is normal for Apnea. A GE Healthcare representative will contact each customer to arrange for the correction. Questions or concerns can be directed to GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 3700 | 3,700 units total | Worldwide distribution - US nationwide in the states of CA, CO, DC, FL, GA, HI, IA, IL, IN, KY, LA, ,MA, MD, MI, MN, MO, MT, NC, NE, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, WV. Countries of Argentina, Australia, Austria, Belgium, Bolivia, Brazil, Canada, China, Colombia, Croatia, Czech Republic, Denmark, Estonia, Finland, France, Germany, Greece, Iceland, India, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Korea (Republic of), Kuwait, Macedonia, Mexico, Netherlands, New Zealand, Nigeria, Norway, Peru, Portugal, Qatar, Russia, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Vietnam. | Open, Classified |  |  | 3700 | Computer | Not_Computer
Z-1288-2018 | 78568 | No | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Southmedic Oxygen Masks with EtCO2 sampling line part number OM-2125-8, 25 masks per case. |  |  | October 17, 2017 | 2017 | Southmedic, Inc. | Potential for loose or missing CO2 diffusers which could result in a non-functioning mask. | Device Design | Customers were notified via letter on October 17, 2017. Instructions included checking inventory for affected devices and quarantining any found, complete and return the Voluntary Recall Notification Confirmation Form, and arrange for the return of any affected products. On February 14, 2018, the recall was expanded and customers were notified. Instructions were the same as the original notification, but updated the Part/Lot information. For further questions, please call 1 (800) 463-7146 |  | 8871 | 8,871 cases (25 masks per case) total | US Distribution to the state of California, and Internationally to Canada and Sweden | Open, Classified |  |  | 8871 | Computer | Not_Computer
Z-1169-2018 | 79006 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AutoMate 1200 Catalog Number ODL25120; |  |  | January 19, 2018 | 2018 | Beckman Coulter Inc. | Belts on the x-axis (and on the y-axis, to a lesser extent) manufactured from 2015 are showing premature failure, posing the remote risk of cross-contamination. | Device Design | Beckman Coulter plans to replace the belt via Service Modification during the next planned service maintenance visit. At that time, the customer will be informed of the belt replacement by the service engineer. |  | 98 | 98 units in the US | Australia, Austria, Belarus,Belgium, Bosnia and Herzegovina, Brazil,Canada, China, Croatia,Cyprus, Czech Republic, Denmark,Finland, France, Germany,Gibraltar, Greece, Hong Kong,Hungary, India, Ireland,Israel, Italy, Korea,Kuwait, Latvia, Lithuania,Macao, Mayotte, Mexico,Netherlands, New Zealand, Norway,Panama, Poland, Portugal,Qatar, Russia, Saudi Arabia,Singapore, Slovakia, Slovenia,South Africa, Spain, Sweden,Switzerland, Taiwan, Turkey,United Kingdom, United States | Open, Classified |  |  | 98 | Computer | Not_Computer
Z-1183-2018 | 79271 | No | PSZ | Influenza virus antigen detection test system. | Microbiology | Microbiology | 510(k) | Sofia 2 Analyzer Catalog Number 20299 |  |  | December 15, 2017 | 2017 | Quidel Corporation | The calibration for the affected batch of analyzers was incorrectly set during the manufacturing process which may lead to a failure when calibration is first performed. There may also be e remote possibility that incorrect qualitative test results for patient samples near the assay cutoff may be reported. | Software Manufacturing/Software Deployment | Quidel sent and Urgent Product Recall letter dated December 15, 2017.The letter instructs the user to discontinue use of the device until new firmware can be installed. For further questions, please call (800) 874-1517. |  | 482 | 482 units | US Distribution | Open, Classified |  |  | 482 | Computer | Not_Computer
Z-1170-2018 | 79560 | No | PMP |  |  | Neurology | PMA | Proclaim DRG Implantable Pulse Generator, Model Number 3664 |  |  | March 09, 2018 | 2018 | St. Jude Medical, Inc. | The firm received complaints of error messages that occurred during routine impedance checks on Proclaim DRG IPGs, model 3664. Some complaints were also associated with transient over stimulation which created discomfort for the patients. | Software change control | Notification material in the form of a Customer Letter was hand-delivered to twoconsignees in the US beginning 09 Mar 2018. The Physician Letter was hand-delivered to impacted implanting physicians, beginning 14 Mar 2018. A Patient Letter was delivered via certified mailing to all active implanted patients in the US, beginning 14 Mar 2018. |  | 17 | 17 units | US | Open, Classified |  |  | 17 | Computer | Computer
Z-0495-2018 | 79054 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo.plaza Picture Archiving and Communication System (PACS) with software version VB20A; Model numbers: 10863171, 10863172, 10863173 |  |  | August 18, 2017 | 2017 | Siemens Medical Solutions USA, Inc | Software update to correct several issues that include (1)Potential data loss, (2) Study mix-up, (3) Incorrect measurements on multi-frame images, (4) Dearchiving issue, and (5) Unauthorized access of data due to inadequate permissions for shared folders. | Software design | Siemens released a software update VB20A_HF04 and a Customer Information Letter was sent with Update Instruction SY010/17/P. Service reports for the affected sites will be monitored to ensure the recall was effective.For further questions, please call (610) 219-6461. |  | 47 | 47 | USA (nationwide) Distribution to the states of : AL, CT, FL, KS, LA, MA, MI, MN, MO, MS, NE, NY, OH, OR, PA, TX and WI. Foreign consignees were not reported as device is manufactured outside the U.S. | Terminated | August 30, 2018 | 377 | 47 | Computer | Computer
Z-1216-2018 | 79335 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | The Stryker Thoracic Pedicle Feeler 6002-350-000 is a smart instrument for remote control of |  |  | October 16, 2017 | 2017 | Stryker Instruments Div. of Stryker Corporation | Stryker Instruments and Stryker Leibinger GmbH &amp; Co. KG arevoluntarily recalling certain serial numbers Thoracic Pedicle Feelers that are used in conjunction with the StrykerÂż NavigationSystem because the settings onthe machine were higher than intended. This change led to increased energy directed onto the tip of the Thoracic PedicleFeeler which caused material degradation where the laser marks exist. There is the potential for the tip of the Thoracic Pedicle Feeler to break. | Software Manufacturing/Software Deployment | US customers will be notified via FedEx overnight. US Sales Representatives will be notified via email. International entities will be notified via an internal notification system. International customers will be notified per local regulations/procedures.Customer Instructions:Actions to be taken by the Customer/User:1. Immediately review this Recall Notification.2. Immediately check all stock areas and/or operating room storage for affected equipment. Quarantine anddiscontinue use of any affected Thoracic Pedicle Feelers.3. Complete the enclosed Business Reply Form (BRF) to confirm receipt of this Notification and identify how manyaffected items are currently in your inventory. Please complete and return the BRF even if you don't have any affected product on hand. Note: Your signature on the BRF indicates that you received and understand this Notification and have followed theinstructions in the Notification.4. If you have further distributed this product, please forward this Notification and the attached BRF to all affected locations. Please indicate each location on the BRF.5. If the BRF for your facility indicates that recalled product is currently on hand, we will contact you to arrange for the return of the product for repair/replacement.If discontinuing use of the specific Thoracic Pedicle Feeler(s) listed will cause disruption at your facility. For further questions, please call (269) 323-7700. |  | 28 | 28 | International , US and US Govt. | Open, Classified |  |  | 28 | Computer | Not_Computer
Z-0474-2019 | 81471 | Yes | LWS |  |  | Cardiovascular | PMA | The SQ-RX Pulse Generator a component of the Subcutaneous Implantable Cardioverter Defibrillator (S-ICD) System |  |  | November 01, 2018 | 2018 | Boston Scientific Corporation | Experienced accelerated battery depletion and a shortened replacement interval due to latent internal battery malfunctions. | Device Design | The firm, Boston Scientific, sent an "Important Medical DeviceInformation" letter to customers on November 5, 2018 by hand, email via our sales representatives or Federal Express or Courier. The letter described the product, problem and actions to be taken. We will reconcile that 100% of the identified customers /customer representatives signed for the letter, or the letter has been confirmed to be delivered. Boston Scientific will post the product advisory letter on our website, within the Product Performance Resource Center located @ www.bostonscientific.com/en-US/pprc/product-advisories.html. The customers were instructed to do the following: Your doctor has been notified to look for this possible change, so please go to all scheduled defibrillator checks and clinic visits. As always, if you or your family hear beeping from your device, please contact your heart clinic as soon as possible. Your doctor can confirm your ability to hear and recognize this sound during an appointment.If you have additional questions regarding this information or would like to report a clinical event, please contact your Boston Scientific representative orTechnical Services, or the Recall Contact at 651-582-3494 or email: Heather.Young@bsci.com. |  | 12900 | 12,900 | US Nationwide Distribution: Alabama, Alaska, Arizona, Arkansas, California, Colorado, Connecticut, Delaware, District of Columbia, Florida, Georgia, Hawaii, Idaho, Illinois, Indiana, Iowa, Kansas, Kentucky, Louisiana, Maine, Maryland, Massachusetts, Michigan, Minnesota, Mississippi, Missouri, Montana, Nebraska, Nevada, New Hampshire, New Mexico, New York, North Carolina, North Dakota, Ohio, Oklahoma, Oregon, Pennsylvania, Puerto Rico, Rhode Island, South Carolina, South Dakota, Tennessee, Texas, Utah, Vermont, Virginia, Washington, West Virginia, Wisconsin, and Wyoming. | Open, Classified |  |  | 12900 | Computer | Computer
Z-0442-2019 | 81119 | No | HTO | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | 3.2mm Proximal Reamer/Cannulated Drill provided as part of the Fibulock Fibular Nail System Pack. FibuLock Nail Procedure Pack: Contains instrumentation for the implantation of the FibuLock implant. Contents include: Actuation Driver, 6.2mm Reamer, 3.2mm Reamer, 2 mm Drill, Spade Tip Guide Wire, and 1.6mm x 12 K-Wire. |  |  | September 13, 2018 | 2018 | Arthrex, Inc. | There is potential to break during use. | Software change control | Arthrex sent letters to their consignees with the following instructions.1. Immediately discontinue use, sale, and distribution of the affected product.2. Please contact Arthrex Product Surveillance at 866-267-9138 orcomplaints@arthrex.com as soon as possible to arrange for return of theaffected product. Our Product Surveillance Specialists can provideassistance regarding alternative solutions and are available to answerquestions regarding credit for affected devices in your possession.3. If you are an Arthrex agency, you do not need to notify any customers asArthrex will be notifying customers directly.4. If you are an Arthrex customer, you do not need to notify any patients ofthis product recall.5. If you have any questions about this product recall, please contact ArthrexProduct Surveillance at 866-267-9138 or complaints@arthrex.com. |  | 832 | 832 | Distributed throughout the U.S. to the following states: NC, CA, NJ, CO, NV, UT, LA, KY, OH, AZ, MO, MN, TN, NY, WI, MA, PA, TX, IL, PA, AR, FL, IA, MI, AL, SD, OR, WA, MD, KS, and GA. | Open, Classified |  |  | 832 | Computer | Not_Computer
Z-2050-2018 | 78466 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | nordicBrainEx 2.0 provides analysis and visualization capabilities of dynamic MRI data of the brain, presenting the derived properties and parameters in a clinically useful context. |  |  | September 17, 2014 | 2014 | NordicNeuroLab AS | An error was discovered in the interpretation of certain DICOM header tags that may lead to incorrect orientation labeling, and thus and indirect left-right, up-down or anterior-posterior flipping of images. | Software design | The firm notified their consignees by email on 09/17/2014. |  | 22 | 22 licenses | USA, Sweden, Spain, Russia, Hungary, India, Norway, Iceland, Italy, South Korea, United Kingdom | Open, Classified |  |  | 22 | Computer | Computer
Z-1140-2018 | 79223 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | BD VacutainerÂż EDTA Tan Top Tube; Catalog Number: 367855 (UDI: 50382903678551) |  |  | March 22, 2018 | 2018 | Becton Dickinson &amp; Company | BD Vacutainer EDTA Lavender, Tan, and Pink Top Tubes and BD Vacutainer Lithium Heparin Green Top Tubes can cause an underestimation of lead in blood samples when used with Magellan Diagnostics LeadCare assays, employing the Anodic Stripping Voltammetry (ASV) methodology, or any other assay employing ASV methodology. | Device Design | On March 22, 2018, BD issued a press release. BD intends to distribute URGENT MEDICAL DEVICE CORRECTION notices to their customers in the upcoming days via courier service. The URGENT MEDICAL DEVICE CORRECTION notification has been posted on BD's website along with a link to their press release. BD has updated the IFU for BD Vacutainer EDTA and BD Vacutainer Lithium Heparin Tubes to include the noted precaution for use of Lavender, Tan, Pink and Green Top tubes on Magellan Diagnostics Leadcare assays using ASV methodology. **Distributors are advised to Take the Following Actions:1) This is a Medical Device Correction and there is no need to return or discard product. The product can continue to be used with other non-ASV blood lead level test technologies such as GFAAS and ICP-MS and all assays which do not employ ASV methodology.2) If you have distributed this product, please identify your customers and notify them of this recall by sharing a copy of the correction notice.3) Complete the attached Customer Response Form and return to the BD contact noted on the form whether or not you have any of the covered product.4) If you would like BD to conduct the notification to your customers, please email your customer list within 3 business days to BDPASRC@bd.com. Please include the contact name, address, phone number, email address, and/or fax numbers for each customer.Distributors who require further assistance, please contact BD at 1-888-237-2762 (select Option #3 and then Option # 4) between 8AM and 5 PM, CT Monday through Friday.**Medical Directors, Risk Managers, Medical Device Safety Officers &amp; Lab Managers are advised to Take the Following Actions:1) This is a Medical Device Correction and there is no need to return or discard product. The product can continue to be used with other non-ASV blood lead level test technologies such as GFAAS and ICP-MS and all assays which do not employ ASV methodology.2) Continue to follow FDA's safety |  | 3956900 | 3,956,900 | Distributed in 46 states: AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MT, NC, ND, NE, NH, NJ, NM, NY, NV, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and Puerto Rico. | Open, Classified |  |  | 1302570453 | Computer | Not_Computer
Z-1195-2018 | 79375 | No | NQX | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | Nucleic Acid test intended for screening for MRSA/SA. Each kit has 20 tubes. in vitro diagnostic. |  |  | October 27, 2017 | 2017 | Roche Molecular Systems, Inc. | The cobas MRSA/SA tests may have decreased performance compared to the analytical sensitivity for MSRA detection. A screening failure can lead to potential infection in the colonized patient being tested as well as the spread of MSRA to others. The United States is not impacted, as the cobas MRSA/SA nucleic acid test for use on the cobas liat system is not available in the US. | Software design | The cobas MRSA/SA tests has the potential for decreased performance compared to the analytical sensitivity for MSRA detection. A screening failure can lead to potential infection in the colonized patient being tested as well as the spread of MSRA to others. |  | 136 | 136 kits | Worldwide distribution. Austria, Denmark, Germany, Singapore, Netherlands, Hong Kong, and Switzerland. US product is not affected. | Open, Classified |  |  | 136 | Computer | Not_Computer
Z-0407-2019 | 81300 | No | OSR |  |  | Cardiovascular | PMA | Medtronic CareLink Encore 29901 Programmer |  |  | October 02, 2018 | 2018 | Medtronic Inc., Cardiac Rhythm and Heart Failure (CRHF) | Vulnerabilities have been identified in the SDN download process that may allow an individual with malicious intent to update the programmers with non-Medtronic software during an SDN download. | Software design | The firm began disseminating the Field Corrective Action notices by hand delivery on 10/02/2018 and followed with the posting of Medtronic Security Bulletin and the Urgent Medical Device Correction customer letter online on 10/11/2018.The notice explained vulnerabilities have been identified in the SDN download process that may allow an individual with malicious intent to update the programmers with non-Medtronic software during an SDN download and beginning October 11th, 2018, Medtronic will be disabling the SDN for programmer updates and will rely solely on the USB update method. |  | 4072 | 4072 units | Worldwide | Open, Classified |  |  | 34244 | Computer | Computer
Z-0478-2019 | 81508 | No | NDC | Predictive pulmonary-function value calculator. | Anesthesiology | General Hospital | 510(k) | Insulia Diabetes Management Companion |  |  | October 08, 2018 | 2018 | Voluntis | The firm identified an issue with the basal calculator identified on the Android version of the Insulia application. The bug could lead from low-impact to high-impact hypoglycemia depending on the circumstances. A new update of Insulia is now available (version 1.6.42), which resolves this issue. | Software in the Use Environment | On October 15, 2018, patients were notified that they could download the new version of the application from the Google Play Store. The notification was made through an automated message displayed on the application screen of all impacted patients requesting them to update their Insulia application. Patients were informed that the new version of the app resolved a an issue with the basal calculator identified on the Android version of the application. Patients who do not upgrade to the new version by November 9, 2018 will be blocked from using the app until they upgrade.For technical assistance or if you have any questions, please contact the Insulia support hotline. |  | 14 | 14 | Distributed to accounts in Arizona, Masachusetts, Michigan, Ohio, Tennessee, and Vermont. International distribution to Canada and France (however, French customers are not affected by this recall). | Open, Classified |  |  | 14 | Computer | Computer
Z-1200-2018 | 79568 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON SC2000 Ultrasound System. The firm name on the label is Siemens Medical Solutions USA, Inc., Buffalo Grove, IL. |  |  | March 05, 2018 | 2018 | Siemens Medical Solutions USA, Inc. | The ECG signal may flatline due to electromagnetic interference during the use of electrosurgical equipment. | Device Design | Undated recall letters were issued via email (if they had an email address for the customer) or via certified mail (if no email address was available) on 3/5/2018. |  | 229 | 229 systems | Distribution was made to medical facilities in AZ, CA, CT, FL, IL, IN, KS, KY, MI, MN, MO, NJ, NY, NC, OH, PA, TX, VA, and WA. There was no military/government distribution.Foreign distribution was made to Argentina, Australia, Austria, Belgium, Bosnia Herzeg, Colombia, Croatia, Denmark, France, Germany, India, Ireland, Israel, Italy, Japan, Mexico, New Zealand, Philippines, Poland, Republic Korea, Romania, Singapore, Slovakia, Spain, Sweden, Taiwan, Thailand, Turkey, UAE, and Vietnam. | Open, Classified |  |  | 229 | Computer | Computer
Z-0485-2019 | 81219 | No | KHO | Fluorometer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Tosoh AIA-900 immunoassay Analyzer |  |  | September 28, 2018 | 2018 | Tosoh Bioscience Inc | A software issue may cause the liquid volume detection entries for 2nd and 3rd reagent bottles to be overwritten with the data from the 1st reagent bottle. Tosoh has released software version 2.51 (or later) to fix the problem. This failure is mitigated by the hardware design. The consequence would be a slight delay in testing patient sample or reporting patient result. | Software design | On September 28, 2018, the firm issued an Urgent Medical Device Recall letter to affected customers. Customers were informed of the software issue. The firm has released software version 2.51 (or later) which corrects the problem. Three analyzers were identified that do not have the most current software version and are at risk for the error: 10313605, 10503612, and 10634103. A field service engineer will be contacting customers with these units in their possession to schedule a service visit in order to upgrade those analyzers to the current version of the operating software.Immediate Actions to be taken by the Customer/User-Thoroughly review the content of this letter.-Complete and return the attached Acknowledgement Form.-File this notice with your laboratory records and forward this information to other supervisors ormanagers in your laboratory to ensure that they are aware of the potential issue.If you have obtained unexpected test results, delays in test result reporting, or received any complaintsof illness or adverse events associated with the use of the analyzer, please contact Tosoh TechnicalSupport 24 hours a day, seven days a week at (800) 248-6764.Should you have any questions regarding this medical device recall, please feel free to contactBernadette O'Connell at (800) 248-6764 or by email at bernadette.oconnell@tosoh.com. She will beavailable to answer any questions Monday - Friday, from 9:00 AM to 5:00 PM (PST). |  | 3 | 3 | Distributed to accounts in GA and NY. | Open, Classified |  |  | 3 | Computer | Computer
Z-0439-2019 | 81211 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | The Spirit Select Bed, Model 5700 Spirit Plus /with high rail siderail option |  |  | October 08, 2018 | 2018 | Stryker Medical Division of Stryker Corporation | The beds with the high rail side rail option may not be complaint with IEC 60601-2-52/C1:2010. There is a potential risk of entrapment. | Device Design | Medical Device Correction notification letters were distributed to customers on 10/8/18. |  |  |  | The products were distributed to the following US states: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WY.The products were distributed to the following foreign countries: Canada. | Open, Classified |  |  |  | Computer | Not_Computer
Z-0527-2019 | 81239 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Optional MDRC-1119 Front Cover Attachment Kit for review display monitors. |  |  | March 30, 2017 | 2017 | Barco N.V. | If metal strips are removed and re-applied to another display, there is potential for the metal strips to be deformed resulting in a smaller contact surface with the front cover magnet strips. There is a risk that the protective front cover could come loose and may fall down. | Device Design | Customers were notified via email on about 03/30/2017. Customers were alerted that a safety related update has been made to the installation manual. A safety notice has been shipped to distributors with the updated installation instructions manual and an acknowledgement letter. Distributors were asked to do the following:a) Inform all your customers who have purchased the Front Cover Kit by sending them the Safety Notice;b) Send the updated installation instructions to all your customers who have purchased the Front Cover Kit;c) Fill out the acknowledgment letter; attach and send in reply to this email. |  | 2087 | 2,087 units | Worldwide - US Nationwide Distribution to CA, IL, GA, TX, and WI. In the countries of Belgium, Brazil, Canada, Czech Republic, Denmark, Finland, France, Germany, Italy, Japan, Poland, Saudi Arabia, South Africa, South Korea, Spain, Sweden, Taiwan, Turkey. | Open, Classified |  |  | 2087 | Computer | Not_Computer
Z-1213-2018 | 79455 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Sonimage HS1 Ultrasound Kit AC adapter, Product Number: A8WA, UDI: 04560141944534 |  |  | December 04, 2017 | 2017 | Konica Minolta Medical Imaging USA, Inc. | There is a risk of circuit board breakage within the main housing of the AC Adapter unit. Smoke may be emitted from the AC Adapter as a direct result. | Device Design | Konica Minolta is preparing for the final AC Adapter replacement. Upon receipt of the new AC Adapter, owners of SONIMAGE HSl System will be asked to replace the existing AC Adapter. Additionally, instructions for the return of the current AC Adapter will be provided. |  | 85 | 85 | Distributed in the following U.S. states: AL, AZ, CA, CO, CT, FL, GA, HI, IA, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, NC, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, WA, WI and the District of Columbia. Distributed in Canada and Brazil. | Open, Classified |  |  | 220 | Computer | Not_Computer
Z-0614-2019 | 81527 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Xper Flex Cardio Physiomonitoring System, Software versions 1.5.16.0322 and 1.6.1.0390, Model Numbers 453564241901, 453564241911, 453564483321, 453564483331, 453564621791, 453564621801, 453564669081 |  |  | November 15, 2018 | 2018 | Philips Electronics North America Corporation | There are 3 issues:1) Under certain conditions, real time waveforms may not be accurately plotted or displayed. 2) The oxygen saturation (SpO2) numeric value displayed on the device may freeze.3) In auto cycle mode, if the NIBP communication is lost, the pump would not cycle and the non-invasive blood pressure (NIBP) numeric value displayed on the device would not update. | Software Design Change | The firm initiated the recall on X by letter. The letter explained the issues, potential risk, and provided workarounds for the consignee to use until the new software is ready for release. |  | 651 | 651 units | worldwide | Open, Classified |  |  | 651 | Computer | Computer
Z-0650-2019 | 81563 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS 5600 Integrated System, Version 3.3.1 &amp; below |  |  | October 19, 2018 | 2018 | Ortho-Clinical Diagnostics | Software anomaly may incorrectly extend the on-analyzer stability time for a loaded VITROS Reagent. | Software Manufacturing/Software Deployment | On October 19, 2018, Ortho Clinical Diagnostics issued "URGENT PRODUCT CORRECTION NOTIFICATION" notices via FedEx overnight courier and/ or ORTHO PLUS e-Communications and/ or US Postal Service Priority Mail (for PO Boxes only) to all customers to inform them of the issue and actions to take to prevent occurrence of the anomaly. On November 19, 2018 Ortho Clinical Diagnostics re-issued correction notices to customers pertaining to the 5,1 FS system (CL2018-194_5,1FS). Customers may continue using their VITROS Systems and have been informed that the software issue will be mitigated with the next version of software. The expected dates of these software versions are included in the communication. Foreign affiliates were informed by email approximately on 21 October 2018 of the issue and instructed to notify their consignees of the issue and required actions to prevent the occurrence of the anomaly.Please refer to the Customer letter (Ref. CL2018-194_51FS, CL2018-194_3600, CL2018-194_4600, and CL2018-194_5600) for specific details on Required Actions by User. If you have further questions, please contact Ortho Care" Technical Solutions Center at 1-800-421-3311. |  | 2413 | 2413 | US Nationwide Distribution | Open, Classified |  |  | 4928 | Computer | Computer
Z-0642-2019 | 81404 | No | KWA | Hip joint metal/metal semi-constrained, with an uncemented acetabular component, prosthesis. | Orthopedic | Orthopedic | PMA | Zimmer Compress Devices and Instruments: |  |  | October 12, 2018 | 2018 | Zimmer Biomet, Inc. | Correction to update the surgical technique for the Compress System. To make users aware of the changes that were made only to the selection criteria section on page 42 regarding which array to use in the placement of the anti-rotation pins. | Device Design | Zimmer notified accounts on 12/13/18 via email and FedEx (both original and expanded accounts) as a follow up to the medical device correction notification dated October 12, 2018. Zimmer Biomet identified additional item numbers that are within the scope of this issue. .Distributors letter identify the issue and their responsibilities, which include: Reviewing the correction notification and ensuring affected team members are aware of the contents. Offering assistance in ensuring users are aware of the updated surgical technique and how to locate it. Destroying previous copies of the surgical technique. Completing the Certification of Acknowledgement.Updated surgical techniques 1110.2-GLBL-en REV0818 and BMET0188.2 REV111218 for the Compress and Mini Compress systems are used, effective immediately, for surgeries that require the affected items.a. Access the full surgical techniques at www.zimmerbiomet.com by navigating to the Limb Salvage specialty and then Bone Compressive Devices under the Medical Professionals menu. |  | 531 | 323 devices US; 208 distributed OUS | NationwideForeign:CANADAARGENTINAAUSTRALIABELGIUMCHILE FINLANDJAPAN NETHERLANDS | Open, Classified |  |  | 531 | Computer | Not_Computer
Z-0640-2019 | 79584 | No | GKL | Automated cell counter. | Hematology | Hematology | 510(k) | GloCyte RBC Reagent - Product Usage: The GloCyte Automated Cell Counter System is intended for use by trained healthcare professionals in clinical laboratories to provide quantitative determination of fluorescence labeled total nucleated cells and erythrocytes in cerebrospinal fluid collected from adult and pediatric patients. The GloCyte Low and High Level Controls are assayed hematology controls designed to monitor the performance of the GloCyte Automated Cell Counter System. Assayed parameters include total nucleated cells and erythrocytes. |  |  | February 07, 2018 | 2018 | Advanced Instruments, LLC | RBC Reagent failure. High fluorescent counts during RBC Reagent + Diluent QC check that will result in failed QC check, preventing user from testing patient samples for RBC. | Device Design | Customers have been sent a recall notification communicating the reason for recall. Once a root cause is identified, recalling firm will notify customers and make available new RBC Reagent lots to allow customers to resume use of reagent in RBC counting. |  |  |  | US Nationwide distribution in the states of IN, WA, NC, CA, GA, CO. | Open, Classified |  |  |  | Computer | Not_Computer
Z-1276-2018 | 79438 | No | DHB | Radioallergosorbent (RAST) immunological test system. | Immunology | Immunology | 510(k) | Phadia Prime software, article number 12-4101-00, as used in combination with the Phadia 250 Instrument, article number 12-3900-01 and EliA Assays. This recall is for any Phadia Prime software version up to and including the current version, 2.1.4. |  |  | November 20, 2017 | 2017 | Phadia Ab | We want to inform all Phadia 250 system operators performing EliA Assays not to use the function OK to All in any version of Phadia Prime, up to and including 2.1.4, when rejecting and retesting samples with any EliA assay (the OK function may be used for rejecting single tests and dilution of samples in accordance with product DfU). | Software design | An urgent recall notification was issued to customers in November of 2017, and informs all Phadia 250 system operators performing EliA Assays not to use the function OK to All in any version of Phadia Prime, up to and including 2.1.4, when rejecting and retesting samples with any EliA assay (the OK function may be used for rejecting single tests and dilution of samples in accordance with product DfU). For further questions, please call (269) 492-1940. |  | 2307 | 2307 | US Distribution to the states of :TX, VA, NJ, MA, IN, GA, UT, TN, CA, MS, NC, ME and OR | Open, Classified |  |  | 2307 | Computer | Not_Computer
Z-0641-2019 | 81635 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Centricity Universal Viewer Zero Footprint Client, Model Number 2090258-001, Software Versions 6.0 SP7, SP7.0.1, SP7.0.2, SP7.0.3, SP8, SP8.0.1, SP8.0.a and SP8.0.2, SP9 and SP9.0.1 display o orientation markers can become is incorrect. |  |  | November 09, 2018 | 2018 | GE Healthcare, LLC | Issue #1 affects Software Versions 6.0 SP7, SP7.0.1, SP7.0.2, SP7.0.3, SP8, SP8.0.1, SP8.0.a and SP8.0.2.When an image is flipped horizontally and vertically, the orientation markers appear correctly on the image to which they were applied. However, once new image rendering happens on this image or series (with scrolling; playing cinÂż; maximizing or changing viewports/layouts; or synchronizing viewports and resizing the browser), the new image orientation is correctly maintained but the display of orientation markers can become incorrect.Issue #2 affects Software Versions 6.0 SP7, SP7.0.1, SP7.0.2, SP7.0.3, SP8, SP8.0.1, SP8.0.a, SP8.0.2, SP9 and SP9.0.1. When an image in a series is flipped and then rotated in Zero Footprint Viewer (ZFP), the orientation markers appear correctly on the image to which they were applied. However, once you navigate to another image in the series and the image you navigated to has a different orientation than the prior image from which you navigated, the new image orientation is correctly maintained but the display of orientation markers can become incorrect. | Software design | The firm issued the medical device correction notices on 11/09/2018 by letter. The letter outlined the problems and provided work-around strategies to be employed until a product modification is available to correct the issues. |  | 508 | 508 units | Worldwide | Open, Classified |  |  | 508 | Computer | Computer
Z-0639-2019 | 81589 | No | LON | Fully automated short-term incubation cycle antimicrobial susceptibility system. | Microbiology | Microbiology | 510(k) | VITEKÂż 2 Systems Software Version 9.01 Update Kit. |  |  | November 13, 2018 | 2018 | bioMerieux, Inc. | The VITEKÂż 2 Systems Software Version 9.01 software may display incorrect organism identification information as well as the associated Antimicrobial Susceptibility Testing (AST) results when using the VITEKÂż 2 FLEXPREP" feature. | Software design | bioMerieux notified customers of the recall on about 11/13/2018 via "URGENT PRODUCT CORRECTION NOTICE." Instructions included taking the following actions: 1) Confirm the letter has been distributed to, and reviewed by, all appropriate personnel; 2) If appropriate in accordance with consignee's laboratory's Standard Operating Procedures, perform retrospective review of any patient results potentially impacted in relation to the described VITEK FLEXPREP anomaly; 3) Store this letter with the bioMÂżrieux, Inc. VITEK 2 instrument documentation; 4) Complete and return the Acknowledgement Form in Attachment A to confirm receipt of the notice and return it to your local bioMÂżrieux, Inc. representative. Regardless of each consignee's current use of the VITEK FLEXPREP" feature, your local Field Service Engineer has implemented the software solution on your VITEK 2 System.If you require additional assistance or have any questions, please contact your local bioMÂżrieux, Inc. Clinical Customer Service representative at (800) 682-2666. For non-urgent inquiries, please contact the Technical Support Center by email at CustomerService-IDAST@biomerieux.com. The anticipated turnaround time for email inquiries is 24-48 hours during normal business hours (7AM  7PM EST), and up to 72 hours dependent upon weekends and company observed holidays. |  | 2 | five (5) VITEKÂż 2 Systems Software version 9.01 | Nationwide distribution to CO, IL, MO, TX. | Open, Classified |  |  | 2 | Computer | Computer
Z-0259-2018 | 78784 | No | OXO | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Mobile Mini C-arm system Part # 1000-0004. |  |  | August 31, 2017 | 2017 | Orthoscan, Inc. | OrthoScan Inc. discovered during investigation of a non-standard work flow a non-conformity of the devices. Specifically, the system software allows a user to activate the Digital Zoom feature in the course of taking a live image exposing a larger x-ray field than can be viewed by the user | Software design | Orthoscan sent an Urgent Recall Notice, dated August 30, 2017, to customers, informing them of the issue and how to address it. Customers are to use the attached list of serial numbers to determine if their device is affected. Customers may continue to use their affected device with caution until the on-site correction has been made. Users should avoid non-conforming workflows by taking the measures outlined in the letter. The firm will contact customers to arrange implementation of the corrective action at no cost to the customer. If there are further questions or concerns, please call National Customer Service Team at (866) 949-4346. |  | 2012 | 2012 total | Worldwide Distribution - USA (nationwide) including Puerto Rico and Hawaii and to the countries of : Australia, Austria, Belgium, Canada, Dominican Republic, Estonia , Finland, France, Germany, Hanover, Hong Kong, Israel, Italy, Japan, Jonesboro, Korea, Malaysia. Netherland, New Zealand, Norway, Poland, Russia, Saudi Arabia, Spain, Sweden, Switzerland, Taiwan, UAE, and United Kingdom. | Open, Classified |  |  | 2012 | Computer | Computer
Z-1379-2018 | 79641 | No | MEC | Medical washer-disinfector. | General Hospital | General Hospital | 510(K) Exempt | Getinge AGS, FSLC, FSUC, RC accessory to Washer- Disinfector 86 Series UDI Info: 86 Series Washer Disinfector 07340153700246, 07340153700253, 07340153700260, 07340153700277 |  |  | March 02, 2018 | 2018 | Getinge Disinfection Ab | There is potential for a cart with disinfected goods to unload without the trolley being docked. Disinfected goods could end up falling to the floor causing risk to nearby users and damage to goods and equipment. | Software design | A letter (Urgent Medical Device Recall Notice) was issued March 2, 2018 via 2 Day Fedex with customer signature required to allow for tracking of delivery. The letter includes a customer response form requesting feedback including if device is still in use at the facility. The letter instructs the customer to contact Getinge Service to schedule a visit and service by a trained Getinge technician at no charge to the customer. In addition Getinge Service will proactively contact affected customers and perform the correction when affected installed devices are identified.Recalled product will be corrected by a trained Getinge Service Representative. The correction will involve a software update to remove any risk to operators. Additionally a preventive hardware update that provides a longer wearing end stop pin will be performed during the field correction. For further questions please call (585) 214-6049. |  | 170 | 170 total products | Worldwide Distribution - USA (nationwide) and the following countries: AUSTRALIA, BELGIUM, BRAZIL, CANADA, CHINA, COLOMBIA, CZECH REPUBLIC, GERMANY, DENMARK, NETHERLANDS, NORWAY, SPAIN, FINLAND, FRANCE, GREAT BRITAIN, HONG KONG, ITALY, JAPAN, KOREA, KUWAIT, MALAYSIA, NORWAY, NEW ZELAND, SAUDI ARABIA, SWEDEN, SINGAPORE, SLOVENIA, THAILAND | Open, Classified |  |  | 170 | Computer | Not_Computer
Z-0629-2019 | 81709 | No | DSQ |  |  | Cardiovascular | PMA | Jarvik 2000 Ventricular Assist System external components: JHI-212 Y Cable, JHI-202 Lithium-ion Battery Cable, JHI-204 Extensions Cable, JHI-201 Pedestal Cable |  |  | October 15, 2018 | 2018 | Jarvik Heart Inc | There is a potential for the male connectors on external cables to become damaged. This could include a detached male connector housing, misaligned connector, cracked connector housing, and bent pins. | Device Design | Urgent Field Safety Notice letters dated 10/5/18 were distributed to customers. |  | 274 | 274 | Products were distributed to the following US states: FL, NY, and UT.Products were distributed to the following foreign countries: Croatia, Czech Republic, France, Hungary, India, Italy, and Japan. | Open, Classified |  |  | 274 | Computer | Not_Computer
Z-0497-2019 | 81167 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Puritan Bennett 980 (PB980) ventilator series: |  |  | September 18, 2018 | 2018 | COVIDIEN LLC | Software Update: External USB Drive performance and its impact on Graphic User Interface (GUI) functionality and labeling of the scalar waveform displayed on GUI during ventilation | Device Design | Medtronic notified U.S. consignees notified via Federal Express or certified mail on September 19, 2018, and the letter informs customers that Medtronic has developed a software update, Maintenance Release (MR) 5.4. Medtronic Service Engineers will load the MR5.4 software during the next scheduled or planned maintenance visit. Immediately notify all care environments in which the PB980 ventilators are used about this notification." If your facility has distributed PB980 ventilators to other persons or facilities, please promptly forward a copy of this letter to those recipients." Complete the attached form and return it as directed to confirm your receipt and understanding of this information.Any incidents related to this issue or if you have any questions, please contact our Technical Support Department immediately at 1-800-255-6774.Out--Side US accounts notified by Medtronic Regulatory Partners worldwide on 9/17/18 for translation in their respective country. |  | 9593 | 9,593 | Worldwide Distribution - US NationwideForeign:CanadaArgentinaChileJapanHong KongTaiwanUruguayAustriaAustriaAzerbaijanBelgiumBosnia And HerzegovinaBulgariaCroatiaCzech RepublicFranceGermanyGreeceHungaryIrelandIsraelItalyKazakhstanKenyaKuwaitLatviaLebanonLithuaniaNetherlandsOmanPakistanPolandPortugalRomaniaRussian FederationSaudi ArabiaSlovakiaSloveniaSouth AfricaSpainSwedenSwitzerlandTunisiaTurkeyUnited Arab EmiratesIranUnited Kingdom | Open, Classified |  |  | 9593 | Computer | Computer
Z-0578-2018 | 78899 | bedding, disposable, medical | KME | Medical disposable bedding. | General Hospital | General Hospital | 510(K) Exempt | EMS products from Graham Medical, SnugFit(R) Expandable Mattress Cover, Code 72930, UPC 1 07 47036 72930 3 |  |  | December 27, 2017 | 2017 | LITTLE RAPIDS CORPORATION | The secondary supplier provided material was within specification but did not meet our design needs compared to the primary manufacturer. The product is designed for patient transfer, not lifting. However, we recognize that the product may be utilized outside of its intended design scope, and lifting may occur. Lifting could lead tearing of the equipment cover and the potential for patient falling through. | Device Design | The firm initiated their recall by email on 12/27/2017 to their direct consignees. The consignees were directed to report back with the number of products on hand for return and to notify their customers of the recall. |  | 4590 | 153 cases, 30 units in each case | TX, CT, TN, MO, WI, FL, SD, MS, CA, MS | Terminated | July 30, 2018 | 215 | 4590 | Computer | Not_Computer
Z-1398-2018 | 79766 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ONCOR Avant-garde, Digital Linear Accelerator, Model No. 05863472 |  |  | February 14, 2018 | 2018 | Siemens Medical Solutions USA, Inc | Control Console software has been updated to reduce the risk for collision when delivering automatically sequenced treatments with automatic movements of the gantry and/or the treatment table. | Software design | On February 16, 2018, the recalling firm issued a Customer Advisory Notice via USPS Certified Mail to inform customers about the update of the Control Console software version 9.2.400 of the Digital Linear Accelerator to version 9.2.502. Customers were informed of the issues that would be fixed through this software update. Consignees were asked to instruct affected personnel accordingly, place the safety advisory notice in the device's instructions for use, and have personnel maintain awareness until the fix is implemented. |  | 1 | 1 | Worldwide Distribution- US Nationwide in the states: AL, CA, FL, GA, IL, IN, KY, MA, MI, MO, MS, NJ, NY, OH, OK, PA, RI, SC, TN, TX, WI, WV, and WY; and PR. International distribution to Nassau. | Open, Classified |  |  | 61 | Computer | Computer
Z-0517-2018 | 78105 | No | EHD | Extraoral source x-ray system. | Dental | Radiology | 510(k) | PC-1000, PC-1000/Laser1000, Model Numbers 800724-2 &amp; 800724-3 |  |  | July 14, 2017 | 2017 | Panoramic Rental Corp. | The lift motor used to raise and lower the overhead chassis will wear if not properly maintained according to the maintenance schedule. | Device Design | On July 14, 2017, Panoramic distributed Urgent Medical Device Correction notices to customers who own and have not taken action as required for the PC-1000 X-ray system machines. This means that your device does not have a critical safety feature that is intended to protect your patients and staff from harm. Customers are encouraged to immediately discontinue use of their machine and follow the instructions in the notice. Customers are encouraged to not resume use of the machine until a certified Panoramic service technician repairs the lift motor. Customers are advised to contact Panoramic to ensure any safety risks associated with the lift motor are identified and addressed. Action Required:(A) If your PC1000 machine is still in service at this location, you must immediately discontinue use and choose one of the following by returning the Acknowledgement Form or contacting the Panoramic Service Team by calling 888-381-2577 Monday through Friday 8 am - 6 pm.1. An upgraded lift motor assembly at no cost to you;2. A digital Panoramic rental unit to replace your PC1000, contingent on a signed one year rental contract (subject to availability); or3. A new digital Panoramic or PanoramicCephalometric unit to replace your PC1000 at a reduced price (startingat $17,995 depending on machine type and subject to availability)(B) If your PC1000 machine is no longer in service or under your control:4. If the machine is no longer in service, Panoramic will provide free removal of your machine. |  | 13340 | 13,340 units | US Nationwide distribution including Puerto Rico | Terminated | July 23, 2018 | 374 | 13340 | Computer | Not_Computer
Z-2472-2018 | 79665 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | SOMATOM Emotion 6 (Model 10165888) |  |  | February 01, 2018 | 2018 | Siemens Medical Solutions USA, Inc | A potential risk of unnecessary radiation exposure due to a software issue | Software design | Customers were sent notification letter via USPS with Certified Mail Receipts and Electronic Receipts. |  |  |  | Nationwide | Open, Classified |  |  |  | Computer | Computer
Z-0619-2019 | 81615 | No | MRZ | Infusion pump. | General Hospital | General Hospital | 510(k) | SAPPHIRE Epidural Set Yellow-Striped Microbore, 0.2 Micron Filter, PAV, 123 Inch, List Number 16373-02 |  |  | November 08, 2018 | 2018 | ICU Medical Inc | There have been reports of distal occlusion alarms during the use of SAPPHIRE Epidural Set Yellow-Striped administration sets. | Device Design | On 11/08/18, Stericycle, Inc. sent a "URGENT MEDICAL DEVICE CORECTION" letter to its consignees. The letter described the product, problem and actions to be taken via traceable mailing. Delivery will be confirmed via traceable mailing. Consignees who further distributed the product were asked to notify their customers.For further inquiries, please contact ICU US Field Action Manager at 224-706-2487 or veasna.seth@icumed.com. |  | 31050 | 31,050 units | US distribution and Internationally to Canada. | Open, Classified |  |  | 31050 | Computer | Not_Computer
Z-1382-2018 | 79648 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Infinity Acute Care System (IACS) Monitoring Solution, Catalog Numbers: MS18841, MS20401, MS20407, MS20724, MS22956, MS25510, MS25520, MS25643, MS26196, MS26372, MS31818; UDI Information: 4049098054454, 4049098054447, 4049098054409, 4049098009799, 4049098009751 |  |  | February 21, 2018 | 2018 | Draegar Medical Systems, Inc. | The invasive pressure visual and audible alarms may not activate. | Software Manufacturing/Software Deployment | On February 21, 2018 , Drager Medical Systems Inc. mailed an Urgent Medical Device Recall Notification to affected customers, affected customers that informs the end users of the recommended connecting sequence when using the hemodynamic pod with the M540 patient monitor. A Customer Reply Card will be included as a receipt from the customers that they received and acknowledged the recommended sequence when using the hemodynamic pod with the M540 patient monitor.To avoid encountering the identified issue, the user should follow the below identified steps whenever the hemodynamic pod is plugged into the Infinity M540 patient monitor and the unit is placed in Standby. When the user is ready to monitor the patient:1. Exit Standby mode2. Unplug, pause for a minimum of 3 seconds and re-plug the hemodynamic pod3. Continue per your normal procedureIf there are any questions regarding the notification please call 1-800-437-2437 (press 1 at the prompt, then 32349). For questions regarding the operation and/or servicing of your Drager Infinity Acute Care System (IACS) please contact Drager Service Technical Support at 1-800-437-2437 (press 4 at the prompt and then 4 again). For further questions, please call (978) 379-8077. |  | 28317 | 28317 total products | Worldwide Distribution - USA (nationwide) and the following countries:Angola, Argentina, Australia, Austria, Azerbaijan, Bahrain, Belgium, Bolivia, Botswana, Brazil, Brunei Dar-es-S, Bulgaria, Canada, Chile, China, Columbia, Costa Rica, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, Hong Kong, Hungary, India, Indonesia, Ireland, Italy, Japan, Jordan, Kenya, Kuwait, Latvia, Liechtenstein, Lithuania, Macedonia, Malaysia, Maldives, Mexico, Monaco, Mozambique, Namibia, Netherlands, New Zealand, Norway, Pakistan, Panama, Peru, Poland, Portugal, Qatar, Reunion, Russian Fed., Saudi Arabia, Singapore, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Thailand, Trinidad and Tobago, Turkey, Turkmenistan, United Kingdom, Utd.Arab.Emir. and Vietnam. | Open, Classified |  |  | 28317 | Computer | Not_Computer
Z-1378-2018 | 79711 | No | MQG | Assisted reproduction accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | MINC Benchtop Incubator (Mini Incubator) |  |  | January 12, 2018 | 2018 | Cook Inc. | The temperature of the incubation lids of the MINC Bench top Incubators may become too high or too low for proper gamete/embryo incubation due to fatigue failure of the flat flex lid heater. | Device Design | Cook plans to begin notifying consignees of the recall on January 12, 2018Written recall communication letters will be mailed via courier to US consignees |  | 4451 | 4,451 | Nationwide (K-IMNC-100 -US G46022) . (K-IMNC-100 G20079 manufactured in Australia and distributed worldwide) | Open, Classified |  |  | 4934 | Computer | Not_Computer
Z-1366-2018 | 79368 | No | KQO | Automated urinalysis system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | iQ200 Series Urine Microscopy Analyzer with Barcode Reader Model NFT-2100, Catalog #s: 800-3046, 800-3047, 800-3052, 800-3053, 800-3802, 800-3803, 800-3900, 800-3920, 800-3925, 800-3931, 800-3933, 800-3934, 800-3935, 800-3937, 800-3950, 800-3951, 800-3042, 800-3043, 800-3044, 800-3049, |  |  | November 10, 2017 | 2017 | Beckman Coulter Inc. | iQ200 Series Urine Microscopy Analyzer may intermittently fail to read urine sample dilution barcode labels causing erroneous results that could delay treatment for health conditions such as hematuria. | Device Design | An urgent medical device recall letter, dated 11/10/2017, was sent to customers informing them that Beckman Coulter was initiating a field action for the iQ200 Series Urine Microscopy Analyzer with Barcode Reader, which may fail to read urine sample dilution barcode labels instead defaulting to no dilution, dilution factor of 1:1. In addition, this product may also fail to read urine sample dilution barcode labels, when the rack is rejected, leading to no results. Customers were informed of the issue, impact, action, and that the resolution was being investigated. Customers were instructed to respond within 10 days: electronically via email or manually by completing and returning the response form accompanying the device recall letter. Customers with questions regarding the notice were instructed to contact the Customer Support Center: electronically at http://www.beckmancoulter.com/customersupport/support or via phone at 1-800-854-3633, in the United States and Canada. Customers with questions located outside the United States and Canada, were instructed to contact local Beckman Coulter representatives. |  | 4350 | 4350 total units (1713 in U.S.) | United States: nationwide, including Puerto Rico; and Foreign: Canada, Algeria, Argentina, Australia, Austria, Bahrain, Bangladesh, Belarus, Bermuda, Bolivia, Brazil, Brunei Darussalam, Chile, China, Colombia, Croatia, Czech Republic, Ecuador, France, Germany, Greece, Guadeloupe, Guatemala, Hong Kong, Hungary, India, Indonesia, Iraq, Israel, Italy, Japan, Kenya, Korea, Kuwait, Lebanon, Lithuania, Macao, Macedonia, Malaysia, Maldives, Martinique, Mexico, Moldova, Republic Of, Montenegro, Morocco, Netherlands, New Zealand, Nigeria, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Puerto Rico, Qatar, Reunion, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, U.A.E, Uganda, United Kingdom, United States, Uruguay, Venezuela, Vietnam | Open, Classified |  |  | 4350 | Computer | Not_Computer
Z-1348-2018 | 79553 | No | HKO | Ophthalmic refractometer. | Ophthalmic | Ophthalmic | 510(K) Exempt | AUTO REF/KERATOMETER ARK-1s |  |  | December 20, 2017 | 2017 | Nidek Inc. | There is possibly a difference between the measured dioptric value of the subjective measurement mode and the objective measurement mode on ARK-1s. | Equipment maintenance | Customers were notified via letter on about 12/20/2017. The letter instructed consignees that a representative from Marco Ophthalmic, Inc. will make arrangements for the affected unit to be picked up and replaced with an ARK-S unit that meets factory specifications. |  | 63 | 63 untis | Distributed domestically to AK, AL, AR, CA, CO, DE, FL, GA, ID, IL, IN, MD, MI, MO, MT, NM, NY, OH, OK, PA, SC, TN, TX, VA, WA, WV. | Completed |  |  | 63 | Computer | Computer
Z-0686-2019 | 81705 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MOSAIQ Oncology Information System Version 2.64. |  |  | December 04, 2018 | 2018 | Elekta, Inc. | If a Setup field (CT, kV, MV or MVCT) is created by COPYING a treatment field or another setup field, then thecopied field will have the same isocenter as the parent and that isocenter may not be the correct isocenter forimaging when it is sent to the machine. | Software design | Elekta notified customers on about 12/04/2018 via letter titled, "Urgent Important Field Safety Notification." The letter contains recommended user action and that appropriate personnel working with the device should be advised on the content of the letter and the letter should be placed in an area accessible to all users. The letter stated that the issue will be resolved in Service Packs to MOSAIQ Releases 2.64, 2.65 and 2.70. Customers will be informed when the fix is available through a Product Bulletin.An acknowledgement Form was sent as part of the letter to be completed and returned no later than 30 days. The firm can be contacted at 408-830-8000 or 855-693-5358. |  | 551 | 551 units | Distributed in the United States and outside of the Unites States: Algeria; Antigua and Barbuda; Argentina; Australia; Belgium; Botswana; Brazil; Canada; China; Colombia; Denmark; France; Germany; India; Indonesia; Iran; Ireland; Israel; Italy; Japan; Latvia; Lebanon; Malaysia; Mexico; Morocco; Myanmar; Netherlands; New Zealand; Norway; Peru; Philippines; Poland; Portugal; Russian Federation; Rwanda; Saudi Arabia; Senegal; Singapore; Slovenia; South Africa; Spain; Sweden; Switzerland; Taiwan; Thailand; Turkey; United Kingdom; Viet Nam | Open, Classified |  |  | 551 | Computer | Computer
Z-0685-2019 | 81700 | No | MUE | Full-field digital mammography system. | Radiology | Radiology | 510(k) | I-ViewÂż Contrast Enhanced Digital Mammography, ASY-08109 |  |  | December 06, 2018 | 2018 | Hologic, Inc. | Calibration issue not possible to visualize contrast uptake in the subtracted mammography images | Software design | Hologic issued letter on 12/6/2018 USPS Certified Mail to customers to inform them of this product correction, and Hologic asks customers to only use I-View for patients with known lesions and confirm the image quality is of acceptable diagnostic quality.Hologic will scheduling appointments to recalibrate the systems for contrast enhanced digital mammography. |  | 337 | 337 units | NationwideForeign:AustraliaAustriaBelgiumBrazilCanadaChinaColombiaEcuadorEgyptFranceGermanyGreeceHong KongIndiaIrelandIsraelItalyJapanKorea, Republic ofKuwaitMalaysiaMexicoNetherlandsPakistanPeruPolandPortugalSaudi ArabiaSingaporeSlovakiaSouth AfricaSpainSwitzerlandSwitzerlandTaiwanThailandTurkeyUnited Arab EmiratesUnited Kingdom | Open, Classified |  |  | 337 | Computer | Computer
Z-0683-2019 | 81246 | No | OZO |  |  | Clinical Chemistry | PMA | Medtronic MiniMed 630G 6025805-003_C, Rx Only |  |  | October 03, 2018 | 2018 | Medtronic Inc. | Reports of occurrences in which insulin pumps with version 4.10 software have failed to make expected audio sounds during alerts, alarms or sirens. This could cause a user to miss system notifications, alarms or sirens associated with how the pump is working, and with high and low glucose alerts. | Software design (manufacturing process) | On 10/3/2018, the firm sent an URGENT FIELD SAFETY NOTIFICATION to customers and distributors via email and/or US Postal Service First Class to inform them of reports of occurrences in which insulin pumps withversion 4.10 software has failed to make expected audio sounds during alerts, alarms, or sirens. This failure could either cause the alarm volume to be stuck at a 4 (out of 5) level regardless of the personal setting, or it could make no sound. Either of these occurrences could cause the user to miss system notifications, alarms or sirens associated with how the pump is working, andwith high and low glucose alerts. Consignees are being informed that they are receiving this letter because the firm's records indicate they may have one of the potentially affected insulin pumps. The notification further explains, this issue occurs when a specific electronic component in the pump malfunctions. This potential loss of audio could delay the user's response to the underlying reason for the alert, alarm, or siren, which could then lead to possible health and safety risks such as hypoglycemia or hyperglycemia. If this issue occurs with the pump, the audio cannot be permanently repaired or regained. The only way to permanently regain the audio capabilities is to exchange the affected pump with a replacement. Even if the pump has the audio issue, itwill continue to deliver insulin as expected. Loss of the Vibrate feature has not been reported to occur as part of the above audio failure mode. If a customer has a pump that contains the specific electronic component, and the pump passes the Audio Beep test described below, there is still a chance that the pump may lose its audio capabilities at some future date.Customers are instructed to:What you should do:1. If you have not already done so, enable the Vibrate feature on your pump. Even if the Audio Beep test passes, you should enable the Vibrate feature (in addition to your Audio feature) since it will add |  | 142889 | 142889 | US: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, Puerto Rico, RI, VT, SC, SD, TN, TX, UT, VA, Virgin Island, VT, WA, WI, WV, and WY | Open, Classified |  |  | 142889 | Computer | Computer
Z-0663-2019 | 81210 | No | KRB | Thermodilution probe. | Cardiovascular | Cardiovascular | 510(k) | Iced Thermoset Closed-Loop Injectate Delivery System for cold Injectate. 1 unit per Cavity #2-4"x9" Cav/20 pieces per case R1-15645 |  |  | October 05, 2018 | 2018 | ICU Medical, Inc. | Potential for leakage due to defect in the thermoset check valve component. | Equipment maintenance | On October 5, 2018, ICU Medical URGENT MEDCAL DEVICE RECALL letters were sent to ICU Medical's consignees advising them of this matter. Each letter contained a response form and a return label. |  | 25640 | 25,640 units | United States - AL, CA, CO, DC, FL, GA, IN, KY, LA, MA, ME, MD, MI, MO, MS, MT, NJ, NY, NC, OH, OK, OR, PA, SC, TN, TX, VA, WA, WI, WV, WY | Open, Classified |  |  | 27400 | Computer | Not_Computer
Z-0309-2018 | 78613 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access 2 Immunoassay System, Remanufactured, Catalog Number 386220. |  |  | November 06, 2017 | 2017 | Beckman Coulter Inc. | Potential for minor skin burn if the tool contacts the reagent storage interface board (B79917) when performing a pipettor alignment. | Device Design | An Urgent Medical Device Recall letter, dated November 6, 2017, were sent to the affected customers via mail. The letter identifies the affected device, the reason for recall, impact, and actions to be taken. The firm states that customers may continue to operate their system normally. Customers are instructed to contact Technical Support if a device on the Access 2 instrument needs alignment. The alignment tool is only to be used by the Beckman Coulter service personnel and Beckman Coulter trained biomedical engineers. A Beckman Coulter service rep will contact customers to schedule a service visit. Customers are to share this information with their staff and retain a copy of the letter. Customers are to respond via one of the ways identified. Beckman Coulter should be contacted with any questions regarding the notice. |  | 3 | 3 units | Worldwide Distribution -- US, Argentina, Canada, Chile, China, Columbia, Costa Rica, France, Georgia, Germany, Ghana, Greece, Hong Kong, India, Italy, Japan, Kazakhstan, Mexico, Morocco, Myanmar, Nepal, Philippines, Poland, Russia, Rwanda, Saudi Arabia, Slovakia, South Africa, South Korea, Thailand, Turkey, United Kingdom, and Vietnam. | Open, Classified |  |  | 232 | Computer | Not_Computer
Z-0674-2019 | 81597 | No | CDZ | Testosterone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ST AIA-PACK Testosterone (TES) is designed for IN VITRO DIAGNOSTIC USE ONLY for the quantitative measurement of testosterone in human serum on specific Tosoh AIA System analyzers. Measurement of testosterone is used to aid in the diagnosis and management of conditions involving excess or deficiency of this androgen. |  |  | November 30, 2018 | 2018 | Tosoh Bioscience Inc |  | Device Design | On November 30, 2018, Tosoh Bioscience issued Urgent Medical Device Recall notices to customers. Customers are advised to take the following actions:" Do not use the assays identified in notice for those patients who have recently had a fluorescein angiography procedure. For these patients, use an alternative testing method that does not use a fluorescein-based technology." Inform medical professionals that the Tosoh AIA assays in the notice utilize a fluorescein-based technology and should not be used to test patient samples for patients who have recently undergone fluorescein angiography." Continue to use the Tosoh AIA assays listed in notice only for patients who have not undergone a fluorescein angiography procedure exam, or on patient samples which have been taken prior to receiving the fluorescein angiography procedure." Complete &amp; return the Urgent Medical Device Recall Acknowledgement Form within 15-days of receiving the notification." Maintain a copy of the notification with your laboratory records and forward this information to others who may have received this product.Should you have any questions, please contact Bernadette OConnell at 1(800) 248-6764 or by email at Bernadette.oconnell@tosoh.com. She will be able to answer any questions Monday-Friday from 9 AM to 5 PM (PST). |  | 26745 | 26745 units | US Nationwide | Open, Classified |  |  | 101916 | Computer | Computer
Z-1316-2018 | 79010 | No | ILQ | Powered communication system. | Physical Medicine | Physical Medicine | 510(K) Exempt | HomeSafe AutoAlert Pendant works in conjunction with a compatible Lifeline communicator. |  |  | November 13, 2017 | 2017 | Lifeline Systems Company | A programing error in some Model FD100 HomeSafe AutoAlert Pendants will render the fall detection featureinoperable. | Software Manufacturing/Software Deployment | Firm has contacted all recipients and informed them of an issue that affects Philips Lifeline AutoAlert help button from automatically signaling the Lifeline Response Center for help in the event of a fall. Further instructed them that they would need to press the Lifeline Help Button in order to connect with the emergency response center. They then instruct them to return their inventory for replacement. New AAHB Replacements will be provided when available. Philips will work with the Programs to ensure buttons swaps are completed. |  | 20201 | 20201 | US and Canada distribution | Open, Classified |  |  | 20201 | Computer | Computer
Z-1365-2018 | 79625 | Yes | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Liberty Select Cycler (SW v.2.8.7), Material Number RTLR108343 |  |  | January 24, 2018 | 2018 | Fresenius Medical Care Renal Therapies Group, LLC | The recalling firm identified a software issue related to the Patient Line Check (PLC) which may result in an increased risk of Overfill (also known as Increased Intraperitoneal Volume, IIPV). Overfill/IIPV may result in serious injury or death. | Software design | On 01/25/2018, the recalling firm sent Urgent Medical Device Correction letters to affected Home Therapy Nurse Managers (HTNs). On 01/31/2018, the firm sent letters to affected patients. The letters informed customers of the recall and stated that patients who are considered "slow drainers" and experience M65 Scale warnings during one or more drain cycles are at increased risk.On 03/29/2018, the firm sent out a second Urgent Medical Correction Letter to HTNs, followed by a patient communication sent on 04/02/2018. In the second communication, the firm amended that slow drainers can develop Overfill/IIPV even without receiving M65 scale warning.Health care providers and patients using the device were advised to monitor drain rates and address the reasons for slow draining, such as constipation, fibrin deposition, peritonitis, and catheter malposition. Health care providers were advised to continue to review risk factors and symptoms of Overfill/IIPV with PD nursing staff and patients with extended drain times.The firm is not requesting the return of the recalled device at this time. For additional information and alternative treatment options, please contact the FMCRTG Medical Information Line at 1-855-616-2309. Hours of operation are 8:30AM-5:00PM (EST), Monday - Friday. Online requests may be submitted at any time at www.fmcna-medinfo.com. |  | 9293 | 9293 | Nationwide distribution. | Open, Classified |  |  | 9293 | Computer | Computer
Z-1313-2018 | 79444 | No | DHB | Radioallergosorbent (RAST) immunological test system. | Immunology | Immunology | 510(k) | ImmunoCAP Allergen c6, Amoxicilloyl ( Article number 14-4451-016). |  |  | December 18, 2017 | 2017 | Phadia US Inc | The stability specifications of ImmunoCAP Allergen c6, Amoxicilloyl were not met. | Device Design | Customers are being asked to please stop using all current product, and scrap any unused products. Âż Please disregard any reported test result below the assay Limit of Quantification (LoQ 0.1 kUA/l) from the lots, as these results may not be valid.Please note that assay results above the Limit of Quantification, will only be affected by giving a lower result than the real value.Âż Please fill in the Recall Letter Acknowledgement Form provided, including the table in the response form to obtain credit, and return the signed response by e-mail or fax. |  | 950 | 950 total | Nationwide Distribution | Open, Classified |  |  | 950 | Computer | Not_Computer
Z-0682-2019 | 81547 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ABL800 analyzer, model numbers 393-800 and 393-801. |  |  | July 03, 2018 | 2018 | Radiometer America Inc | The analyzer software may cause a mis-match of patient demographics and test results when a certain sequence of events occur. | Software design | The recalling firm issued letters dated 7/3/2018 on 7/3/2018 via FedEx. |  | 2131 | 2,131 analyzers | Distribution was nationwide. There was also military/government distribution. Foreign distribution was made to Canada. | Open, Classified |  |  | 2131 | Computer | Computer
Z-1311-2018 | 79462 | No | PMT | Device to detect and measure non-microbial analyte(s) in human clinical specimens to aid in assessment of patients with suspected sepsis. | Microbiology | Microbiology | 510(k) | Elecsys BRAHMS PCT |  |  | December 22, 2017 | 2017 | Roche Diagnostics Corporation | Product exhibits a decreased recovery of patient samples on the cobas e 801 module compared to the MODULAR ANALYTICS E 170, cobas e 411, 601, and 602 analyzers. | Software design | The firm, Roche, sent an "URGENT MEDICAL DEVICE REMOVAL" letter dated December 22, 2017 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to do the following:" Immediately discontinue use of Elecsys BRAHMS PCT (catalog number 07301715190) reagent lotnumber 204084 on the cobas e 801 module." Discard affected product according to local guidelines." Complete the attached fax form and fax it to 1-800-722-7222." File this Urgent Medical Device Removal (UMDR) for future reference.Please contact the Roche Support Network Customer Support Center, 24 hours a day, seven days a week at 1-800-428-2336 if you have questions about the information contained in this UMDR. |  | 47 | 47 | US Distribution in states of: CA, FL and IN. | Completed |  |  | 47 | Computer | Not_Computer
Z-0659-2019 | 81649 | No | IXY | Wall-mounted radiographic cassette holder. | Radiology | Radiology | 510(K) Exempt | Bucky Stand SU-4000 Input 2017.11.17. |  |  | November 08, 2018 | 2018 | Konica Minolta Healthcare, Americas, Inc. | Unintentional U-Arm movement | Device Design | On 11/08/2018, the firm send a letter titled;" IMPORTANT PRODUCT INFORMATION", regarding the Konica Minolta KDR U-Arm. Potential unintentional U-Arm movement due to the tablet enclosure buttons, located on the tablet enclosure, remaining engaged after being depressed.The firm included a document titled, KDR U-ARM Tablet Enclosure FAQ as part of the IMPORTANT PRODUCT INFORMATION. The firm is preparing to issue an Interim Corrective Action; Disabling the buttons contained on the KDR tablet enclosure and a Final Corrective Action; A modification to the mounting of the PCB - which contains the drive buttons will reduce the potential of the event -this will help assure the buttons make the proper contact without staying depressed.In addition, the firm issued a "Technical Bulletin-Business Partner for the Product: KDR U-Arm Tablet Enclosure." The preliminary notification was sent to dealer partners which require all Partners to take immediate action.There are two critical actions needed to mitigate this risk:1.Interim Action (Immediate Action Required):The firm has released an immediate field corrective action to temporarily disconnect the tablet enclosure buttons. Instructions can be found in KON-000127 .2.Final Action (Anticipated Release January 2019):A modification to the tablet design will help ensure the buttons make the proper contact without staying depressed.The Interim Action/Field Correction is required to be completed by December 15, 2018. Contact Konica Minolta Dispatch department at 1- 800-945-0456 to schedule a Konica Minolta Engineer to complete Corrective action. For each affected unit, it is important that your engineer performs the following:1. Complete the Tablet Button Disconnection procedure - KON-0001272.Fill out form KON-000157 (Customer Signature required)3.Take a picture of the completed form KON-000157 and email to ServiceBilling@gcp.konicaminolta .com. |  | 151 | 151 | US Nationwide Distribution: AL; CA; CO; FL; GA; IA; IL; IN; KY; LA; MA; MI; MN; MO; NC; NH; NY; OH; OK; PA; SC; TN; TX; and VA. | Open, Classified |  |  | 151 | Computer | Not_Computer
Z-1296-2018 | 79279 | No | DOF | Lead test system. | Toxicology | Toxicology | 510(k) | LeadCare Plus Blood Lead Analyzer |  |  | January 09, 2018 | 2018 | Magellan Diagnostics, Inc. | The incorrect placement of AA batteries into certain LeadCare Plus Blood Lead Analyzer may cause the batteries to overheat, and the chance that a user could burn their fingers if the batteries are removed while hot. | Device Design | An Urgent Medical Device Correction letter, dated January 9, 2018, was sent to customers. The letter identified the affected device and provided a description of the issue. If two of the four AA batteries are inserted incorrectly, the batteries may overheat, causing a minor burn if the user opens the battery compartment and tries to remove the hot batteries. Customers are to follow the steps outlined in the letter, including placing the enclosed label on the upper back panel of the analyzer to remind users who are changing batteries of the importance of correct placement. Proper battery placement should be provided in annual training. Customers are to complete and return the acknowledgement card. Please direct any questions to 800-275-0102 or LeadCareSupport@magellandx.com. |  | 79 | 79 | Worldwide Distribution | Open, Classified |  |  | 6222 | Computer | Not_Computer
Z-1302-2018 | 79537 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Healthcare Centricity PACS-IW, Model Numbers: |  |  | February 23, 2018 | 2018 | GE Healthcare | A database handling error could occur during the image acquisition process affecting the completeness of acquired images with Centricity PACS-IW. There is a potential that one or more image series (i.e. all images within an image set) may be missing from an exam without indication to the user. While this is rare, this can occur with imaging studies that consist of a very small number of images per series. E.g. CR Thorax exam with 1 image per series. | Software design | GE Healthcare sent an Urgent Medical Device Correction letter dated February 23, 2018 to their customers. The letter provided the affected product, problem and provided the following safety instructions:"Until a product modification is available to correct this issue, the following actions are recommended: 1. Utilize the image count within the QC process to alert the user of a discrepancy in the number of transmitted images from the modality to the number of images available in the Universal Viewer. a. If a discrepancy is identified, attempt to retransmit the exam to PACS. b. If retransmission is unsuccessful, a GE Healthcare Service representative will need to be contacted to assist with resolution of the impacted exam. Urgent cases impacted by this issue should be interpreted at the modality.2. If DICOM storage commit is configured and in use, no commit notification is sent to the modality for images that were impacted by this issue.3. Attention should be given to prior cases, acquired before the application of the safety instructions given in this communication as they may be affected by the safety issue."If you have any questions or concerns regarding this notification, please contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 569 | 569 units | Worldwide Distribution-US Nationwide | Open, Classified |  |  | 569 | Computer | Computer
Z-0329-2018 | 78727 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | EMPOWR Porous Knee System TIBIAL PUNCH, LARGE, REF 801-05-204, NON STERILE, Qty 1 For use in orthopedic surgery. |  |  | December 12, 2017 | 2017 | Encore Medical, Lp | The Locking Punch Guide was not sufficiently guiding the large press fit punch, causing the implant to be placed 2-3mm more posterior than intended. | Device Design | The field safety notice was disseminated by email on 12/12/2017. The notice stated the following: "DJO Sales Agents Action:Please review the bulletin that is provided with this notice. This bulletin contains new instructions on the use of the Locking Punch Guide and Punch. Please ensure your surgeons are familiar with this information.DJO Surgical then requires that you complete the attached acknowledgment and email it to teffany.hutto@djoglobal.com. Please have this form returned by December 15, 2017Once redesigned guides are available, they will be exchanged through a separate field action." |  | 2 | 2 units | One medical device distributor in California. | Completed |  |  | 2 | Computer | Not_Computer
Z-0609-2019 | 81521 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Carousel MAS1025-01, part of the Brain Surgical Accessories Intra-operative Kit (KIT0295), accessory used with the Renaissance System (TPL0038). Indicated for precise positioning of surgical instruments or spinal implants during general spinal and brain surgery. It is intended to be used whenever the clinician and/or patient benefits from generated 3D imaging of, high contrast objects. |  |  | November 26, 2014 | 2014 | Mazor Robotics Ltd | In some situations the carousel position pin (locker) may cause the carousel upper plate to pop out and the carousel may then inadvertently move. | Device Design | All Mazor Brain Clinical Sales Representatives (support during clinical cases) were notified of the issue on 11/26/2014 via email as well as received training. |  | 5 | 5 units | Affected devices were distributed in the US to four different medical facilities in four states | Open, Classified |  |  | 5 | Computer | Not_Computer
Z-1236-2018 | 79487 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Phoenix Nail System, intended for use with an intramedullary nail system for alignment, stabilization and fixation of fractures. 3.5mm Inserter Connector, Short; Model Number: 14-441045; Kit Model Numbers: 14-400082, 14-400082S, 14-440065S, 14-442000-00, 14-442000S, 14-444000-00, 979105, 14-441045-00, and 14-444000S |  |  | March 12, 2018 | 2018 | Zimmer Biomet, Inc. | Certain lots of the Phoenix Tibia Nail 3.5mm Inserter instruments are being recalled due to reports of fracturing and remaining in the device implant. | Device Design | The firm, Zimmer Biomet, sent an "URGENT MEDICAL DEVICE RECALL (REMOVAL)" notice dated 3/12/2018 to its customers. The notice describes the product, problem and actions to be taken. The customers were instructed to do the following:Risk Manager:1.Review this notification and ensure affected personnel are aware of the contents.2.Assist Zimmer Biomet sales representative and quarantine all affected product.3.Zimmer Biomet sales representative will remove the affected product from the facility.4.Complete Certificate of Acknowledgement.a.Return a digital copy to corporatequality.postmarket@zimmerbiomet.com. Distributors: 1.Review this notification and ensure that affected team members are aware of the contents.2.Locate and quarantine affected product in your inventory.3.Complete Inventory Return Certification Form and send to CorporateQuality.PostMarket@zimmerbiomet.com within three (3) days. 4.Immediately return all affected product from your distributorship and from affected hospitals within your territory. Mark RECALL on the outside of the returned cartons. 5.Return the Additional Accounts form to CorporateQuality.PostMarket@zimmerbiomet.comEXPANSION: The firm expanded the recall to include additional lots and the new notification sent 3/23/18, also explains the additional issue that devices were etched with a supplier lot number and the packages were etched with the firm lot number. Both numbers are included in an Appendix to the notification letter. |  | 2470 | 2470 | Worldwide Distribution: US (nationwide) AK, AL, AR, AZ, CA, CO, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MT, NC, NJ, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WI, WV, and Puerto Rico; and countries of: Australia, Brazil, Canada, Chile, Colombia, El Salvador, Netherlands, Japan, Costa Rica, Spain, and Thailand. | Open, Classified |  |  | 2470 | Computer | Not_Computer
Z-1391-2018 | 79628 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Solid state x-ray imager (flat panel/digital imager) |  |  | January 12, 2018 | 2018 | NeuroLogica Corporation | While transporting the portable machine with the retrofit equipment in a backwards motion, there is potential for foot injury from the underside of the detector bin of the retrofit unit. | Device Design | A field safety notice was sent stating that the manufacturer, without charge, will remedy the defect. The notice stated that field upgrades would be performed by December 31, 2017. The notice provided instructions to end users and to distributors. For further questions, please call (978) 564-8503. |  | 16 | 16 units | US Distribution . Only one consignee was shipped 16 units of the affected device. | Terminated | August 14, 2018 | 214 | 16 | Computer | Not_Computer
Z-0557-2019 | 81582 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | STEPHANIX D2RS Digital Dynamic Remote Systemhttp://cts.fda.gov/division-tracking/images/trash.png |  |  | November 02, 2018 | 2018 | Stephanix | The firm has detected a potential risk using the command. After releasing the command, the movement of the table may continue instead of stopping. | Software design | The firm initiated the recall by email on 11/02/2018 to the sole US distributor. This email included a letter to be provided to the end user. The letter directed the user as follows:"As a precautionary measure and before your usual maintenance team intervention, the user should look at the equipment during all of its movement (as stated in the user manual). Immediately release the command if an abnormal movement is detected. If necessary, use the closest emergency switch to stop the system."f you need more information or technical assistance, you can phone your usualcontact. |  | 17 | 17 units | US Nationwide Distribution in the states of FL, IN, WV, TN, MO, WI, CA, FL, GA, PA | Open, Classified |  |  | 17 | Computer | Not_Computer
Z-1261-2018 | 79513 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo.via software is intended to be used for viewing, manipulation, communication, and storage of medical images. It can be used as a standalone device or together with a variety of cleared and unmodified syngo based software options. |  |  | January 10, 2018 | 2018 | Siemens Medical Solutions USA, Inc | When the archiving configuration is changed, data received/created after the upgrade may be flagged as "Not to be archived". This is caused by the automatic function for cleaning up temporary data being disabled by the software upgrade. Due to the disabled cleanup function, disc capacity for free space decreases faster than usual. Unless the archiving configuration and the cleanup automatic function are reverted back to the original state (prior to the software upgrade), data that was incorrectly flagged "Not to be archived" must be manually prevented from being deleted by either clinical administrator or service engineer. | Software Manufacturing/Software Deployment | Customers received letters with information related to the issue and regarding the upgrade to the software version VB20A_HF04. The software was released December 22, 2017 and is available for install. |  | 165 | 165 units | Medical device software which needs to be installed. | Terminated | October 02, 2018 | 265 | 165 | Computer | Computer
Z-0380-2018 | 78900 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | RayStation |  |  | July 17, 2017 | 2017 | RAYSEARCH LABORATORIES AB | For the proton QA preparation module in RayStation 6 (including SP1 and SP2), if the snout position or gap is modified in the QA module, the dose in the QA module may be computed for a different setup than what is used for QA measurements. This could result in an incorrect dosage being provided to a patient. | Device Design | Ray Search sent a Field Safety Notice letter dated July 14, 2017 to affected customers. The letter identified the affected product, problem and actions to be taken. The letter included a response form with instructions for workarounds until the software is corrected. |  | 117 | 117 units | US Nationwide Distribution - FL, WA, AZ, TN, MD, TX, FL, MI, Washington D.C. | Terminated | August 29, 2018 | 408 | 117 | Computer | Computer
Z-1260-2018 | 79241 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Altivate Anatomic(TM) INSTRUMENTATION, PEGGED GLENOID TRIAL, SIZE 54mm, REF 804-07-540 |  |  | January 29, 2018 | 2018 | Encore Medical, Lp | It was discovered during a surgery that the design of the radius at the base of the pegs of the Anatomic Glenoid Trial (Part Numbers: 804-07-380, 804-07-420, 804-07-460, 804-07-500, and 804-07-540) is much larger on the trails than on the implants | Device Design | DJO Global sent an Urgent Field Safety Notice letter dated January 29, 2018, to affected customers. The letter identified the affected product, problem and actions to be taken. For questions call (512) 834-6255 |  | 93 | 93 units | US Nationwide Distribution | Open, Classified |  |  | 461 | Computer | Not_Computer
Z-0549-2019 | 81071 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Invacare 9805 and 9805P Hydraulic Patient Lifts, Model Nos. 9805 and 9805P |  |  | November 16, 2018 | 2018 | Invacare Corporation | Invacare has identified via customer complaints, the potential for the mounting bolt connecting the caster to the base frame of the patient lifts to become loose. Loose hardware can cause wear to the hardware and housings. If left unresolved, the caster may separate from the base of the lift causing the Lift to tilt and a patient fall. Failure does not occur right away but happens over time if the lift is not properly maintained.Patient falls are inherently dangerous and may result in a range of injuries that are dependent on physical condition of the patient and the nature of the fall. When injuries occur, they are typically non-serious injuries such as bruising, lacerations, or abrasions. In rare cases, more serious injuries such as bone fracture, head trauma, or death can occur.For caregivers using the lift, attempts to manually stabilize a lift if it becomes unstable could result in sprains, serious back injuries, or bone fractures. | Device Design | On November 16, 2018, the firm, INVACARE, distributed "URGENT - MEDICAL DEVICE FIELD ACTION" letters with an attached Technical Bulletin to providers to inform them of the product issue, problems and actions to be taken.Providers are instructed to do the following: 1. Review the included Technical Service Bulletin.2. Inspect your existing stock for loose fasteners or hardware as required in the user manual.3. Tighten any loose hardware after inspecting for damage that require replacement of the Lift base.4. You are responsible for ensuring that this field action is conducted to the customer level. Your customers with affected units should be contacted and informed about this field action. We recommend that you provide them with the enclosed Consumer Letter.5. Complete and return the Provider Response Card within 5 business days via Fax to: 440-326-3544 or email to: recall@invacare.com.The consumer letters instructed end users to do the following:1. Continue to have your patient lift(s) inspected and maintained by a qualified service provider at the intervals defined within your User Manual.2. Continue to check for loose fasteners or hardware prior to each use.3. Should you find loose fasteners or hardware, contact a qualified service provider to perform maintenance. Do not attempt to perform maintenance yourself.If you have any questions concerning these instructions, please call Invacares recall and field action support line at (877) 413-6008, Monday  Friday, 8 a.m.  5 p.m. EST, for assistance. |  | 73703 | 73,703 units | Worldwide Distribution: US (Nationwide) and countries of: Canada, Colombia, Germany, Guam, India, Northern Mariana Islands, Philippines and South Africa. | Open, Classified |  |  | 73703 | Computer | Not_Computer
Z-1511-2018 | 79763 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Flex Arm |  |  | March 05, 2018 | 2018 | Synthes, Inc. | Quick connect feature of instruments may be unable to connect to the male features of Flex Arm Adaptors, Insight Retractors, or other mating parts. | Device Design | The firm sent an Urgent Product Recall letter dated March 21, 2018, to affected customers. The letter identified the affected product, problem and actions to be taken. For questions call 1-610-719-5450 or contact your DePuy Synthes Sales Consultant. |  | 84 | 84 | Worldwide distribution - US Nationwide and the countries of Canada, Switzerland and Japan | Open, Classified |  |  | 84 | Computer | Not_Computer
Z-0384-2018 | 77289 | No | CAF | Nebulizer. | Anesthesiology | Anesthesiology | 510(k) | Small Particle Aerosol Generator Model SPAG-2 6000 Series (SPAG-2). The SPAG-2 medical device is a pneumatic flow system indicated for the administration of Virazole (ribavirin) aerosol. Driven by a pressurized gas (air or blended oxygen) and uses a nebulizer to generate a fine aerosol of hydrated Virazole. |  |  | March 13, 2017 | 2017 | Valeant Pharmacueticals International | SPAG-2 unit nebulizers from Lot J0571670RA were tested at Next Breath for particle sizing under the validated method, NB-TM-261, In November 2016. The particle size specification set in NB-SPEC-039 for Dv10, Dv50, &amp; Dv90 are a mean value of no more than 3.0Âżm, 5.5 Âżm, and 15Âżm respectively. Initial testing of the nebulizers yielded four units that delivered particle sizes that did not meet the specification (NB-SPEC-039). The out of specification (OOS) nebulizers were cleaned per the approved SPAG-2 procedure, SPE-000127228 Rev 2, to remove any product residue and were then re-evaluated. Secondary testing yielded that three of the four nebulizers still did not meet the particle size specification. It was concluded that the 3 nebulizers from SPAG-2 Lot J0571670RA failed to meet (SPE-000127228 Rev 2) did not resolve the high droplet size distribution values observed. | Device Design | On April 4, 2017 Valeant Pharmaceuticals North America distributed Medical Device Voluntary Product Recall and Acknowledgement Forms to their customers via USPS. Valeant Pharmaceuticals are directly contacting those customers who have been shipped affected SPAG-2 devices to arrange pick up and testing of the nebulizer component. Only half of the nebulizers will be picked-up for particle-size testing at each account to prevent shortages of VIRAZOLE, which can only be used with the SPAG-2 device. Nebulizers that pass testing will be cleaned and returned. Nebulizers that do not pass testing will be disposed, and new nebulizers will be distributed to facilities once they become available. Customers are encouraged to complete and return the Acknowledgement Form via email to: valeant3012@stericycle.com. |  | 1228 | 418 SPAG-2 and 808 Nebulizers | US Nationwide &amp; Canada | Open, Classified |  |  | 1228 | Computer | Not_Computer
Z-1278-2018 | 79478 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | 2008 K2 Hemodialysis Machine with software version 5.40, Models: (1) Hemodialysis SYS OLC/Diasafe PLS 190610, UDI 00840861100859 (+Serial Number), (2) Machine Short Cap, OLC.DP HP 190630, UDI 00840861100866 (+Serial Number) |  |  | December 14, 2017 | 2017 | Fresenius Medical Care Renal Therapies Group, LLC | When the recirculation ultrafiltration (UF) Goal is set to a value greater than 200ml in service mode and the user starts a treatment using the SLED (Sustained Low Efficiency Dialysis) program the display will show an invalid message and the UF pump will run at the recirculation UF rate which may be up to 4000ml/hour. The SLED program is required to limit the UF rate to a maximum of 1000 ml/hour. | Device Design | The recall notification instructs the user to the UF Goal with values less than 200ml when in Service Mode and to monitor and set the UF Rate to the desired setting after starting the SLED treatment. The consignee is advised that a software update is forthcoming to correct the issue. For further questions, please call (800) 227-2572. |  | 127 | 127 | US Distribution to the states of: AR, CA, CO, FL, GA, IL, MA, MD, MN, NC, NY, OH, PA, SC, TN, and TX. | Open, Classified |  |  | 127 | Computer | Computer
Z-0923-2018 | 78838 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ARCHITECT c16000 Cuvette Segment, Part #09D32-05/ 03L77, a component to the ARCHITECT Clinical Chemistry system. There are 15 segments with each cuvette segment holding 11 cuvette pairs (22 cuvettes) for a total of 165 cuvettes pairs or a total of 330 cuvettes (22 cuvettes x 15 cuvette segments) per reaction carousel. The firm name on the label is Abbott Laboratories, Abbott Park, IL. |  |  | December 08, 2017 | 2017 | Abbott Laboratories, Inc | There is a potential to generate falsely depressed patient results in the cuvettes adjacent to the detached segment post for any assay run on the c4000, c8000, and c16000 instrument. | Device Design | The firm, Abbott Laboratories, issued a "Product Correction" letter dated 12/8/2017 via overnight mail on 12/8/2017 to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to adhere to the new instructions listed in Appendix A to avoid damaging the cuvette segments; If a cuvette segment with a detached bottom is identified, replace it before performing any additional testing on your ARCHITECT Clinical Chemistry System; If you have forwarded the product listed above to other laboratories, please inform them of this Product Correction and provide to them a copy of this letter and retain this letter for your laboratory records.The procedures will be updated in a future version of the ARCHITECT Operations Manual.If you or any of the health care providers you serve have any questions regarding this information, U.S. Customers please contact Customer Service at 18774ABBOTT (available 24 hours a day, 7 days a week). Customers outside the U.S., please contact your local area Customer Service. |  | 17560 | 1,560 ARCHITECT c16000 Systems | Worldwide Distribution: US (nationwide) including Puerto Rico and countries of: Afghanistan, Albania, Algeria, Angola, Argentina, Armenia, Australia, Austria, Azerbaijan, Bahamas, Bangladesh, Belarus, Belgium, Bosnia and Herzegovina, Brazil, Brunei, Bulgaria, Burkina Faso, Cameroon, Canada, Cayman Islands, Chile, China, Colombia, Costa Rica, Croatia, Curacao, Cyrus, Czech Republic, Dem Rep of The Congo, Denmark, Dominican Republic, Ecuador, Egypt, El Salvador, Estonia, Ethiopia, Finland, France, Gaza &amp; Jericho, Georgia, Germany, Greece, Guatemala, Honduras, Hong Kong, Hungary, India, Indonesia, Iraq, Ireland, Israel, Italy, Ivory Coast, Jamaica, Jordan, Kazakhstan, Kenya, Kuwait, Kyrgyzstan, Latvia, Lebanon, Libya, Lithuania, Luxembourg, Macedonia, Malawi, Malaysia, Maldives, Mali, Mauritania, Mexico, Moldova, Montenegro, Montserrat, Morocco, Netherlands, New Zealand, Nigeria, North Yemen, Norway, Oman, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Qatar, Romania, Russia (CIS), Saudi Arabia, Senegal, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, St. Lucia, St. Vincent, Sweden, Switzerland, Taiwan, Tajikistan, Tanzania, Thailand, Trinidad &amp; Tobago, Tunisia, Turkey, UAE, Uganda, United Kingdom, Uruguay, Venezuela, Vietnam, Zambia and Zimbabwe. . | Completed |  |  | 73102 | Computer | Not_Computer
Z-0387-2018 | 78752 | No | DHB | Radioallergosorbent (RAST) immunological test system. | Immunology | Immunology | 510(k) | Phadia 1000 Instrument, Article Number 12-3800-01. |  |  | July 05, 2017 | 2017 | Phadia US Inc | The "Retry" command does not function properly which could cause a shortage of Wash and Rinse solution and affect assay performance and test results. | Device Design | The firm sent letters with a return response form to customers. The firm instructed customers to immediately choose the "stop" command if presented with the 7-101 Liquid Sensor Error. Results from the assay should not be approved. A revision of the current software will have a mandatory installation soon. |  | 70 | 70 units | Nationwide | Terminated | October 04, 2018 | 456 | 70 | Computer | Not_Computer
Z-1242-2018 | 79423 | No | BZK | Monitoring spirometer. | Anesthesiology | Anesthesiology | 510(k) | Gas Module 3 monitoring spirometer, Part Number: 115-030108-00 |  |  | December 01, 2017 | 2017 | Mindray DS USA, Inc. dba Mindray North America | The firm made multiple changes to the product requiring the submission of a new 510(k) premarket notification. | Device Design | Mindray sent an Safety Notification letter dated December 1, 2017 to all customers. The letter identified the affected product, problem and actions to be taken. For question call 201-995-8407. |  | 1719 | 1719 | Worldwide Distribution - U.S. Nationwide in the states: AK, AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WY, including Puerto Rico and distributed to 30 foreign consignees. | Terminated | August 29, 2018 | 271 | 1719 | Computer | Not_Computer
Z-0519-2019 | 81058 | No | MHZ |  |  | Neurology | Contact ODE | Dash 3000 Patient Monitoring System is used to monitor physiologic parameter data on adult, pediatric and neonatal patients. |  |  | August 31, 2018 | 2018 | GE Healthcare, LLC | The patient monitors may simultaneously restart as designed if all are connected to the same network and a network overload occurs for a prolonged time. | Software design | The recalling firm issued letters dated 8/31/2018 via FedEx on 8/31/2018 informing the customer of the issue and providing safety instructions for mitigation of the issue if a prolonged network overload occurs and the Dash and/or Solar patient monitor on the network does not resume normal functionality in approximately 90 seconds. |  | 547920000 | Approximately 182,640 monitoring systems for Dash 3000/4000/5000 | Distribution was made nationwide, including Puerto Rico, Guam, and Samoa. There was also government and military distribution.Foreign distribution was made to Canada, Algeria, Angola, Argentina, Australia, Austria, Azerbaijan, Bahamas, Bahrain, Bangladesh, Belarus, Belgium, Bermuda, Bolivia, Bosnia and Herzegovina, Brazil, Brunei Darussalam, Bulgaria, Cambodia, Cameroon, Chile, China, Colombia, Costa Rica, Cota D'Ivoire, Croatia, Cyprus, Czech Republic, Denmark, Djibouti, Dominican Republic, Ecuadore, Egypt, El Salvador, Estonia, Ethiopia, Finland, France, French Guiana, Gabon, Georgia, Germany, Ghana, Greece, Guatemala, Honduras, Hong Kong, Hungary, Iceland, India, Indonesia, Iran, Iraq, Ireland, Israel, Italy, Jamaica, Japan, Jordan, Kazakhstan, Kenya, Korea (Republic of), Kuwait, Kyrgyzstan, Latvia, Lebanon, Libya, Lithuania, Macedonia, Malaysia, Malta, Martinique, Mauritius, Mexico, Moldova, Morocco, Namibia, Nepal, Netherlands, New Zealand, Nicaragua, Nigeria, Norway, Oman, Pakistan, Palestine, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Puerto Rico, Qatar, Republic of Serbia, Reunion, Romania, Russia, Saint Lucia, Saudi Arabia, Senegal, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sudan, Sweden, Switzerland, Syria, Taiwan, Tajikistan, Tanzania, Thailand, Trinidad, Tunisia, Turkey, Turkmenistan, Uganda, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Uzbekistan, Venezuela, Viet Nam, Yemen, Yugoslavia, and Zimbabwe. | Open, Classified |  |  | 548071210 | Computer | Computer
Z-1514-2018 | 79698 | No | HWB | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Stem Extractor f/Guide Bar, Part number 03.401.072. Instrumentation for extraction and revision of Epoca shoulder implants. |  |  | March 19, 2018 | 2018 | Synthes, Inc. | There is a potential for the connection screw of the stem extractor to break. | Device Design | The firm, DePuy Synthes, sent an "URGENT FIELD SAFETY NOTICE" dated 3/19/2018 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to the following: Âż Review the instructions listed in this notification.Âż Forward this notice to anyone in your facility that needs to be informed.Âż If any of the affected product has been forwarded to another facility, contact that facility.Âż Keep a copy of this notice.Âż Return the Verification Section (page 3 of this letter) to Synthes by:o Fax: 1-888-731-7954 oro Scan/email: Synthes3091@stericycle.com. Follow the Instructions if connecting screw for the stem extractor breaksIf you have any questions, please call 610-719-5450 or contact your DePuy Synthes Sales Consultant. |  | 31 | 31 | Worldwide Distribution: US (nationwide) and to countries of: Canada and Switzerland. | Open, Classified |  |  | 31 | Computer | Not_Computer
Z-0513-2019 | 81266 | No | DOP | Mass spectrometer for clinical use. | Clinical Chemistry | Toxicology | 510(K) Exempt | PerkinElmer QSight 210 MD Mass Spectrometer Simplicity 3Q MD v1.0 with Hot Fix 2 Software intended to identify compounds in human samples by ionizing the compound and separating the resulting ions. Part Number: BC004872 |  |  | September 24, 2018 | 2018 | PerkinElmer Health Sciences Canada, Inc. | Under certain conditions when using PerkinElmer QSight 210 MD mass spectrometer with Simplicity 3Q MD and Hotfix 2 software, it is possible for the original data from the last acquired sample to be over-written when the automatic flush is enabled. | Software design | Perkin Elmer issued notification on 9/24/18 customers are requested to follow instructions contained within the Urgent Medical Device Correction letter and disseminate the information to appropriate functions within their organization.Inform relevant personnel and execute one of the following two options:Qption 1 (Users who do not require Autamatic Flushinq):Deactivate the automatic flushing option in your software:1) Navigate to the Batch Acquisition Queue tab2) Click the Queue Options button3) Under the "System Flush and Standby" section, DESELECT the "Enable automatic flushing" option.Option 2 (Users who require Automatic Flushinqt:Add a non data (blank or "dummy") sample to the acquisition queue to be run as your final sample(before going to Flush/standby). This will ensure that a true sample or control data are not overwritten.Contact PerkinElmer Heatth Scbnces C,arda, Bolton facility Quality Manager- SandiOrdog for further information (647. 505. 8682). The issue identified above will be addressed through a software update. PerkinElmer will contactyou when the fix is available and upgrade your software. |  | 18 | 18 units | FL, MN, PAIndia, Italy, Korea, Spain | Open, Classified |  |  | 18 | Computer | Computer
Z-0401-2018 | 78594 | Yes | DWA | Cardiopulmonary bypass pump speed control. | Cardiovascular | Cardiovascular | 510(k) | CP5 Centrifugal Pump, made up of the drive unit, the emergency system and the pump control panel. It used with an Electrical Remote-Controlled (ERC) tubing clamp and a bubble sensor. |  |  | October 30, 2017 | 2017 | Sorin Group Deutschland GmbH | Possibility that the automatic line closure by the Electronic Remote-Controlled (ERC) clamp after detection of an air bubble could take longer than is intended by design in the CP5 system. | Device Design | Consignees were notified of the recall via USPS trackable mail, beginning on 10/30/2017. The letter informed customers Instructions included to continue following the Instructions for use, including use of an arterial filter on the arterial line, and that a LivaNova representative will contact them to schedule an upgrade to their device. |  | 1318 | 1,318 units | Distributed domestically to Distributed internationally to Algeria, Australia, Austria, Azerbaijan, Bahrain, Belarus, Belgium, Canada, China, Colombia, Croatia, Czech Republic, Ecuador, Estonia, Finland, France, Germany, Hong Kong, Iraq, Italy, Japan, Jordan, Kuwait, Libya, Lithuania, Luxembourg, Malaysia, Morocco, Norway, Palestinian Territory, Panama, Peru, Poland, Qatar, Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, The Netherlands, Tunisia, United Kingdom, Vietnam | Open, Classified |  |  | 1318 | Computer | Not_Computer
Z-0555-2019 | 81178 | No | IMD | Hot or cold disposable pack. | Physical Medicine | Physical Medicine | 510(K) Exempt | Thermacare HeatWraps Menstrual |  |  | October 02, 2018 | 2018 | Pfizer Inc. | The firm received complaints of the wrap coming apart and leaking granular material | Device Design | Pfizer, Inc. initiated the recall by letter on 10/02/2018 to the direct accounts. The notice instructed the direct consignees to conduct a sub-recall by notifying any of their accounts to determine if they had any of the affected lot(s) in their inventory at any of their locations. If so, the recalling firm requested that they discontinue distributing the product lot and promptly return the affected product to Pfizer Inc. (c/o Stericycle Inc.).The firm expanded the scope of their recall to the consumer level and disseminated a notice on 11/27/2018 to its direct accounts. The firm issued a press release on 11/26/2018. |  | 293880 | 293880 units | US Nationwide Distribution and PR | Open, Classified |  |  | 418068 | Computer | Not_Computer
Z-0561-2019 | 81378 | No | JPB | Abnormal hemoglobin assay. | Hematology | Hematology | 510(k) | Cascade Abrazo c-ACT-LR Test Card |  |  | February 24, 2017 | 2017 | Helena Laboratories, Inc. | Due to some cards in this lot reported to have experienced variations in results, (verified by wave form), resulting in some cards with shorter or longer times than expected. | Device Design | On 02/24/2017, the firm sent an email to affected users, with a Withdrawal Notification. The firm informed the customers/distributors that Due to some cards in this lot experiencing variations in wave form, resulting in some cards with shorter or longer times than expected. The firm instructed the customers/distributors to disposes of affected lot, as a preventive quality measure. On 04/12/2017, the firm mailed formal letters to all customers.PUBLIC CONTACT: "Name: Dr. Jessica Jones Hanka"Title: Technical Services Manager"Address: 1530 Lindbergh Drive, Beaumont, TX 77707"Telephone #: (409) 842-3714 x1177"Email: jhanka@helena.com |  | 27 | 27 kits | US: Products under control of Manufacturer in TXOUS: Four International Distributors - Australia, Netherlands, Italy, and England | Open, Classified |  |  | 27 | Computer | Not_Computer
Z-0496-2018 | 78442 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Hemosphere System with the following three components: HEM1 Advanced Monitor, HEMSGM10 (Swan-Ganz Module), and HEMOXSC100 (Oximetry Smart Cable). |  |  | October 02, 2017 | 2017 | Edwards Lifesciences, LLC | Pre-procedural issues related to software defects | Device Design | Consignees were sent recall notification letters on approximately 10/11/2017. The letter identified the affected product, problem and actions to be taken. instructions included coordination with the respective Edwards representative to schedule the software upgrade. |  | 198 | 198 devices | US Nationwide Distributed in the states of AZ, CA, CO, GA, IL, MO, NM, OK, OR, SC, TX | Open, Classified |  |  | 198 | Computer | Computer
Z-1275-2018 | 79582 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Ray Station 4.9 Ray Station 5, Ray Station 6 and Ray Station 7 Software build numbers: 4.9.0.42, 5.0.0.37, 5.0.1.11, 5.0.2.35, 6.0.0.24, 6.1.0.26, 6.1.1.2, 6.2.0.7 or 7.0.0.19 |  |  | January 26, 2018 | 2018 | RAYSEARCH LABORATORIES AB | Software issue with Center Beam in Field functionality. Issue can result in incorrect treatment volume delivered to patient. | Device Design | On January 26, 2018 RaySearch Laboratories AB sent Field Safety Notice, Medical Device Correction #25955 to each customer via e-mail. Customers were advised of the following: be aware that Center Beam in Field affects all beams that share the same isocenter and that apertures may be modified in a way not intended; do not use Center Beam in Field after having created an opposed beam. If an opposed beam has been created, delete the opposed beam and use Center Beam in Field on the remaining beam. Then recreate the opposed beam; please educate planning staff and all users about this workaround; inspect your product and identify all installed units with the above software version number(s), then confirm you have read and understood this notice by replying to the notification via email, fax or by calling local support. Customers with questions were instructed to contact David Hedfors via e-mail at david.hedfors@raysearchlabs.com |  | 2484 | 2484 | Nationwide | Open, Classified |  |  | 2484 | Computer | Computer
Z-1412-2018 | 79666 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance 16 Slice (Air) - Model no. 728246, Computed Tomography X-Ray system |  |  | February 08, 2018 | 2018 | Philips Medical Systems (Cleveland) Inc | Closing the Emergency Stop may disable the horizontal motor, allowing the couch top to move during patient load or unload activities. | Software design | The recalling firm notified customers of the recall on February 8, 2018, via Urgent Field Safety Notice. The notice advised customers of the issue and affected products and lot numbers. Customers were advised that once the E-stop has been closed, the operator/technologist must initate a brief horizontal motion (in or out) using any of the operator control panels, CT-Box or tape/foot switches. This action will ensure the couch carbon top (upper pallet) is in a controlled state. Customers were informed that a Field Service Engineer would contact them to schedule site updates to correct the issue.Customers who need further information or support were advised to contact their local Philips representative. Additionally, they may contact the Customer Care Solutions Center (1-800-722-9377. |  | 2 | 2 | Worldwide Distribution - U.S. Nationwide. International distribution to Canada, Mexico, Algeria, Argentina, Australia, Austria, Bangladesh, Belgium, Bolivia, Bosnia and Herzegovina, Brazil, Bulgaria, Cambodia, Chile, China, Colombia, Croatia, Cuba, Czech Republic, Denmark, Djibouti, Dominican Republic, Egypt, Estonia, Finland, France, Germany, Ghana, Gibraltar, Greece, Guyana, Hong Kong, Hungary, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Republic of Korea, Laos, Lebanon, Lithuania, Macedonia, Malaysia, Maldives, Mali, Martinique, Monaco, Mongolia, Myanmar, Nepal, Netherlands, New Calcedonia, New Zealand, Nicaragua, Norway, Oman, Pakistan, Palestine, Panama, Peru, Philippines, Poland, Portugal, Qatar, Reunion, Romania, Russia, Rwanda, Saint Pierre and Miquelon, Samoa, Saudi Arabia, Senegal, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uzbekistan, Venezuela, and Vietnam. | Open, Classified |  |  | 1790 | Computer | Not_Computer
Z-1279-2018 | 79437 | No | PQP |  |  | Pathology | PMA | Oncomine Dx Target Test User Guides and Assay Definition File, Model: A32461; UDI: (01)10190302006071(11)170911(10)1.8(240)A32461 |  |  | December 04, 2017 | 2017 | Life Technologies Corporation | The Oncomine Dx Target Test may report erroneous results when used off-label for samples other than non-small cell lung cancer (NSCLC) tumor specimens. | Device Design | The following actions are described in the consignee later: Review your testing results from the Oncomine Dx Target Test to ensure that none of the reports containing BRAF V600E positive samples are impacted. Utilize Figure 1 and 2 to aid in identifying this discrepancy. If you have any questions, please contact 408.802.2341. |  | 9 | 9 | US Distribution to the states of: AZ, CA, NC, TX, and VA. | Open, Classified |  |  | 9 | Computer | Not_Computer
Z-1403-2018 | 79670 | No | BZA | Airway connector. | Anesthesiology | Anesthesiology | 510(K) Exempt | Teleflex MEDICAL HUDSON RCI One Way Valve w. Capped Monitoring Port, REF 1644, QTY 50 |  |  | March 27, 2018 | 2018 | Teleflex Medical Europe Ltd | These one-way valves may disconnect at the joint between the two components that make up the device, which would cause an interruption of ventilation to the patient. | Device Design | Teleflex initiated the recall by letter on 03/27/2018. The firm directed the consignee as follows:"1. If you have affected stock, immediately discontinue use and quarantine any products with the product code and lot numbers listed above.2. To return product, complete the enclosed Recall Acknowledgement Form and fax it to 1-855-419-8507, Attn: Customer Service or email to recalls@teleflex.com. This will allow us to document the amount of product you have on hand for return. A customer service representative will contact you with a Return Goods Authorization (RGA) Number and will provide instructions for the return of products to Teleflex Medical.3. If you have no affected stock, please complete the enclosed Recall Acknowledgment Form and fax it to 1-855-419-8507, Attn: Customer Service or email to recalls@teleflex.com. This will allow us to document your receipt of this letter." |  | 103910 | 103910 units | US Nationwide in the states of AU, CA, NZ | Open, Classified |  |  | 103910 | Computer | Not_Computer
Z-1395-2018 | 79710 | No | KGI | Bone densitometer. | Radiology | Radiology | 510(k) | GE Healthcare: |  |  | March 16, 2018 | 2018 | GE Medical Systems Ultrasound &amp; Primary Care Diagnostics, LL | Under certain conditions, when using DICOM Worklist along with DICOM MPPS, a report for a bone density exam may be sent to PACS with the incorrect patient information in the DICOM header. The correct patient information will be listed on the DICOM report image; however, the report may appear under a different patients name in the PACS. | Software design | The firm initiated their recall by letter on March 16, 2018. The notice stated:"Under certain conditions, when using DICOM Worklist along with DICOM MPPS, a report for a bone density exam may be sent to PACS with the incorrect patient information in the DICOM header. The correct patient information will be listed on the DICOM report image; however, the report may appear under a different patients name in the PACS. There have been no injuries reported as a result of this issue.You can continue using your system by disabling the DICOM MPPS feature as follows:1. In enCORE, select Tools ¦ Connectivity Options.2. Select DICOM tab.3. Press Configure Connections in DICOM Worklist section.4. In the MPPS SCP Settings, clear the AE Title and the TCP/IP Address fields5. Press OK to all.A GE Healthcare representative will contact you to arrange for the correction."For further questions, please call (608) 826-7492 |  | 303 | 303 units | Worldwide Distribution - USA (nationwide) including Puerto Rico and to the countries of : Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, China, Colombia, Egypt, Finland, France, Germany, Greece, Hong Kong, India, Ireland, Israel, Italy, Japan, Korea, Lebanon, Malaysia, Malta, Mexico, Morocco,Norway, Poland, Portugal , Romania, Russia, Saudi Arabia, Spain, Switzerland, Taiwan, Tunisia, Turkey, United Arab, United Kingdom and Venezuela. | Open, Classified |  |  | 372 | Computer | Computer
Z-0332-2018 | 78879 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Philips Healthcare Brilliance iCT SP computed tomography x-ray system |  |  | November 08, 2017 | 2017 | Philips Medical Systems (Cleveland) Inc | There is a possibility that screws holding an internal portion of the gantry (Reaction Ring) to the scanner s main bearing may come loose. If this occurs,the part(s) will not be expelled, however, a noise may be heard coming from the gantry. | Device Design | Philips Healthcare is informing consignees of this issue through a Customer Information Letter (CIL_FCO72800690). Philips Key Market representatives will distribute the letter internationally, based on local regulations. Additionally, a Field Service Engineer will contact each site to schedule a time to implement the Field Change Order. Customers are notified that if they hear an unusual noise coming from the gantry, or scan halts, the Philips representative should be contacted as soon as possible. Philips will be inspecting the affected systems, tightening and replacing the reaction ring screws as needed at no additional cost. |  | 134 | 134 | 107 domestic accounts15 US Govt accounts682 Foreign Accounts3 Mexico accounts12 Canada accounts | Open, Classified |  |  | 886 | Computer | Not_Computer
Z-0601-2019 | 81188 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | Hemoglobin A1c Control, Part no. 220232 for use as quality control material to monitor the precision of laboratory testing procedures for HbA1c quantitation. |  |  | October 05, 2018 | 2018 | Tosoh Bioscience Inc | Tosoh Bioscience has become aware of potential degradation of HbA1c control lot # 7055 before the 60-days product stability claim. If degradation of these controls occurs, P00 peaks may be seen and the % HbA1c control values may drift higher than the assigned range. This may result in delays when performing routine instrument quality control. If this issue occurs, there may be a delay in testing and reporting patient test results. Tosoh has received two (2) complaints related to this issue with no serious injuries reported. | Device Design | The firm sent notification to consignees on October 5, 2018, via Urgent Medical Device Recall letter. The letter informed consignees of the potential degradation of HbA1c control lot 7055 prior to the 60-days product stability claim.Immediate Actions to be taken by the Customer" Review the content of this letter with your laboratory staff." Inform laboratory staff to avoid storing the control material in direct exposure to light." Complete and return the attached Acknowledgement Form to Tosoh Bioscience, Inc., within 15-daysof receiving this notification." In the event of an early degradation of the control material, new controls should be used to verify theinstrument performance prior to use." Maintain this notification with your laboratory records and forward this information to others in yourlaboratory." If you have obtained unexpected test results, delays in test result reporting, or received anycomplaints of illness or adverse events associated with the use of the analyzer, please contact TosohTechnical Support 24 hours a day, seven days a week at (800) 248-6764." Should you have any questions regarding this medical device recall, please feel free to contactBernadette O'Connell at (800) 248-6764 or by email at bernadette.oconnell@tosoh.com. She will beavailable to answer any questions Monday - Friday, from 9:00 AM to 5:00 PM (PST). |  | 1000 | 1000 | Distributed nationwide. Foreign distribution to Canada. | Open, Classified |  |  | 1000 | Computer | Computer
Z-0348-2018 | 78702 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | ECG Out Cable - Heartstart MRx |  |  | July 17, 2017 | 2017 | Philips Electronics North America Corporation | Issue with ECG out cables. When a Philips monitor/defibrillator is receiving an ECG signal from an auxiliary bedside monitor via a sync cable or ECG out cable, the following can occur if the monitor/defibrillator experiences interference from electrical fast transients (EFTs):"On the HeartStart MRx and HeartStart XL, EFT noise may be misinterpreted as an R-wave."On the HeartStart XL+, EFT noise can disable ECG monitoring and potentially interrupt demand mode pacing*. *Note: It is contrary to the XL+ Instructions for Use to perform demand mode pacing while using the ECG out cable or obtaining the ECG signal from a bedside monitor. The XL+ Instructions for Use includes the following warning: When pacing in Demand Mode, the ECG cable from the patient must be directly connected to the HeartStart XL+. If the user follows this warning, this problem cannot occur on the XL+. | Device Design | Philips is asking customers to follow the Action to be Taken by Customer/User section of the UrgentMedical Device Correction notification/Field Safety Notice:Customers should remove affected sync cables and ECG out cables from service and destroy them, documenting such action on the attached reply form, which should be returned to Philips. Customers should retain a copy of the reply form.The customer should follow the Instructions for Use that describe the recommended approach for performing cardioversion with the XL+ and MRx monitor/defibrillators. The ECG trunk cable can be disconnected from the Philips bedside monitor and connected to the monitor/defibrillator before delivering therapy; this is the preferred source of an ECG waveform for synchronization. |  | 15811 | 15811 | 1067 US customers163 Canada, 917 International | Terminated | July 12, 2018 | 360 | 15811 | Computer | Computer
Z-0378-2018 | 78760 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Brand Name: |  |  | June 29, 2017 | 2017 | Roche Diagnostics Corporation | Roche has confirmed that a possible sample mismatch issue may occur on the MODULAR ANALYTICS E 170 module, cobas e 601 analyzer, or cobas e 602 analyzer due to a software limitation. | Software design | Roche sent an Urgent Medical Device Correction (UMDC) sent via UPS Ground (receipt required) dated July 6, 2017, and an amended UMDC on August 31 ,2017. All consignees will be notified by Urgent Medical Device Correction sent via UPS Ground (receipt required).Customers/consignees were instructed the following: If the Tip/Cup pick up error alarm is issued then followed by the Abnormal L2-line Movement alarm on your MODULAR ANALYTICS E 170 module, cobas e 601 analyzer, or cobas e 602 analyzer, after which the system goes into Sampling Stop, follow the steps in this UMDC and the appropriate How to Identify Potentially Mismatched Sample Results within the enclosed document. For further questions, please call (317) 521-4343. |  | 2517 | 2517 | USA (nationwide) Distribution | Open, Classified |  |  | 2517 | Computer | Computer
Z-0594-2019 | 81356 | No | JPB | Abnormal hemoglobin assay. | Hematology | Hematology | 510(k) | Cascade Abrazo c-ACT-LR Test Card |  |  | April 20, 2016 | 2016 | Helena Laboratories, Inc. | Due to some cards in this lot experiencing variations in wave form, resulting in some cards with shorter or longer times than expected | Device Design | On 04/20/2016 the firm, Helena Laboratories, contacted affected users via email, and then sent a "WITHDRAWAL NOTIFICATION" (dated April 25, 2016) on 04/27/2016. The notification described the product, problem and actions to be taken. The firm instructed the customers/distributors to disposes of affected lot, as a preventive quality measure.The Withdrawal Notification for Removal of Product from Inventory stated,This lot is to be handled by Distributors, as follows:___ Prevent any shipments or scheduled demonstrations with this lot___ Remove this lot from Available Inventory and discard the cards___ Place cards in Quarantine  pending disposal___ Contact International Sales for credit and delivery of replacement cardsVerify amount REMOVED FROM STOCK: Total number pulled: ____________________ Date _____________________________ The following signifies written certification that this lot has not been sent to an end user:Verified by: _____________________________________________________________Please complete this form and return a copy to Helena Laboratories, via: FAX: 409-842-1874 or e-mail to jjones@helena.com  For questions, contact  Mgr., Technical Services at: 1 (409) 842-3714 ext. 1177.*Reason for removal: Some cards in this lot are experiencing variations in wave form, resulting in possible shorter, or longer, times than expected. While this issue, if present, should be found via use of the Quality Control materials, due to its intermittent nature in the cards, this alone cannot be relied upon. |  | 12 | 12 Kits | US Distribution to state of TX and internationally to: Australia, Netherlands, Italy, and England. | Open, Classified |  |  | 12 | Computer | Not_Computer
Z-1272-2018 | 79379 | No | OOE | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | The FEMTO LDV(TM)Z8 Femtosecond Surgical Laser |  |  | January 21, 2018 | 2018 | SIE AG, Surgical Instrument Engineering | This Correction is being initiated based on a customer complaint received on November 16, 2017 which indicated a complication with a cataract procedure. An investigation completed by the manufacturer determined that although the complication was due to user error improvements could and should be made on our part to improve the usability and performance of our device. | Device Design | The firm, Ziemer Ophthalmology, initiated their field correction by telephone on 01/21/2018 and followed with a letter on 01/28/2018. The letter described the product, problem and actions taken. The letter stated:"...we will be releasing new software which contains an added safety feature which will prevent unintended shift of the planned ARC Incisions and/or alert you in case of unforeseen eye movement between configuration and execution of the ARC incisions. We will be contacting your practice the week of March 12th to set up a time to install this new software and provide in-depth training.Prior to the installation of the new software, you may continue to use the Z8 for cataract surgery; however, we would advise that you not perform both arcuate and clear corneal incisions on the same patient until the new software is installed.If you have any questions, do not hesitate to contact Director of Operations at 618-462-9301 for any further information you require. |  | 125 | 125 devices | Worldwide Distribution: US (nationwide) including states of: AL, CA, FL, GA, IL, IN, MN, OR, TN, TX, VA, and WA; and countries of: Argentina, Austria, Belgium, Brazil, Canada, France, Germany, India, Iraq, Israel, Italy, Japan, Lebanon, Malaysia, Mexico, Namibia, Pakistan, Peru, Poland, Portugal, Romania, Russia, Saudi Arabia, Singapore, Spain, South Africa, Switzerland, Taiwan, and United Arab Emirates. | Terminated | November 02, 2018 | 285 | 125 | Computer | Not_Computer
Z-0349-2018 | 78725 | No | HSX | Knee joint femorotibial metal/polymer non-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Persona Partial Knee System, Size 8, Finishing Guide |  |  | July 05, 2017 | 2017 | Zimmer Biomet, Inc. | Finishing guide does not adequately account for shape differences present in Size 8 of Persona Partial Knee System, potentially resulting in additional surgery time or sensation of tightness during knee flexion. | Device Design | Customers were notified via FedEx mail and email on approximately 07/05/2017. Instructions included to notify appropriate personnel, complete and return the Acknowledgement Form, return affected product, and provide contact information of additional hospitals or surgeons that used the affected product. |  | 158 | 158 devices | Distributed domestically to AZ, CA, CO, ID, IL, IN, MA, MD, MI, NJ, NM, PA, SC, TX, WA, WI. Distributed internationally to Belgium, France, Germany, Italy, Japan, Netherlands, Spain, Sweden, United Kingdom. | Open, Classified |  |  | 158 | Computer | Not_Computer
Z-1463-2018 | 79597 | No | HWT | Template for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | The humeral stems trials are packaged in a kit for Titan Modular Shoulder System under the following labeling: |  |  | December 20, 2017 | 2017 | Integra LifeSciences Corp. | Incidents of stem trial breakage were reported to the firm suggesting that these fractures all occurred during insertion/impaction or extraction of the humeral stem trial whiling preparing the humeral canal and/or trialing. All cases resulted in a delay in surgery, with a variance of medical intervention required. | Device Design | Firm sent letters to consignees on December 20, 2017. Firm asked consignees to examine inventory and determine if consignee had affected product. Firm promised to send new stem trials along with a return shipping label. Firm asked to complete the Acknowledgment and Return Form and email or fax back to firm. For further questions, please call (609) 275-0500. |  | 1619 | 1619 units | Worldwide Distribution - USA (nationwide) Distribution. | Open, Classified |  |  | 1619 | Computer | Not_Computer
Z-0583-2019 | 81077 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | PROCARE B20 Patient Monitor, V1; and B20 Patient Monitor, V2. |  |  | September 11, 2018 | 2018 | GE Healthcare, LLC | Patient monitors may restart due to network overload caused by network configuration. | Software design | The recalling firm, GE Healthcare, issued "URGENT MEDICAL DEVICE CORRECTION" letters dated 9/11/2018 via trackable mail on 9/11/2018 to its customers. The letters described the product, problem and actions to be taken. The letters informed the customer of the issue and providing safety instructions for mitigation of the issue if a prolonged network overload occurs and the Bx0 and/or B1x5 patient monitor on the network does not resume normal functionality in approximately 90 seconds. The customers were instructed to follow the safety instructions.GE Healthcare will provide revised network installation and user instructionsdocumentation at no cost to you in a future mailing.If you have any questions or concerns regarding this notification, please contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 28004 | 28,004 patient monitors | Worldwide Distribution: US (nationwide) including Puerto Rico; and countries of: Algeria, Argentina, Australia, Austria, Azerbaijan, Bahamas, Bangladesh, Belgium, Bolivia, Bosnia-Herzegovina, Brazil, Brunei, Bulgaria, Burma (Myanmar), Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, El Salvador, Ethiopia, Finland, France, Gabon, Georgia, Germany, Greece, Guatemala, Honduras, Hong Kong, Hungary, Iceland, India, Indonesia, Iran, Iraq, Ireland, Israel, Italy, Ivory Coast, Jamaica, Jordan, Kampuchea, Kazakstan, Kenya, Korea, Kuwait, Lebanon, Lithuania, Macedonia, Malaysia, Mexico, Moldova, Morocco, Nepal, Netherlands, New Zealand, Nicaragua, Nigeria, Norway, Oman, Pakistan, Palestine, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Senegal, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Tunisia, Turkey, United Arab Emirates, United Kingdom, Uruguay, Uzbekistan, Venezuela, Viet Nam, and Zambia. | Open, Classified |  |  | 79907 | Computer | Computer
Z-0354-2018 | 78787 | Yes | LWP |  |  | Cardiovascular | PMA | ACCOLADE SR Pacemaker |  |  | December 07, 2017 | 2017 | Boston Scientific Corporation | Boston Scientific has received reports of intermittent over-sensing of the Minute Ventilation (MV) sensor signal with certain Boston Scientific pacemaker and cardiac resynchronization therapy pacemaker systems (pacemakers). | Device Design | On December 19, 2017 the firm sent a letter to all their medical consignees stating the following:December 2017Dear Doctor,Boston Scientific has received reports of intermittent oversensing of the Minute Ventilation (MV) sensor signal with certain Boston Scientific pacemaker and cardiac resynchronization therapy pacemaker systems (pacemakers). MV sensor signal oversensing may cause pre-syncope or syncope due to periods of pacing inhibition. This MV behavior may occur with any manufacturers pacing lead system, but Boston Scientific has determined it to be more likely for affected Boston Scientific pacemakers using Medtronic or Abbott/St. Jude(Abbott) leads implanted in either the right atrium (RA) or right ventricle (RV).Boston Scientific is actively developing a software update designed to automatically detect and resolve this MV sensor signal oversensing behavior. We anticipate submitting the software update to Regulatory Agencies inMarch 2018 and pending approval, will release it as soon as possible thereafter. Until this software update is available, Boston Scientific has additional re commendations to mitigate this risk for affected pacemaker systems.Root Cause InvestigationThe MV sensor in Boston Scientific pacemakers can be used for RightRate" (rate adaptive pacing), Respiratory Rate Trend, or AP Scan"1. When the RA/RV pacing leads and lead terminal connections are operating asintended, the MV sensor signal is appropriately filtered and therefore is not detected by the pacemaker or displayed on electrograms (EGMs). However, intermittency related to the lead or pacemaker-lead connection2 hasthe potential to create a transient high impedance condition. A high impedance condition may subsequently alter the MV sensor signal such that it becomes visible on EGMs and potentially subject to oversensing on the RA orRV channels. For a technical description of the Boston Scientifics MV sensor, please refer to Appendix A.Engineering analysis and |  |  |  | AK, AL, AR, AZ, CA , CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OR, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY, and Puerto RicoWorldwide | Open, Classified |  |  |  | Computer | Not_Computer
Z-2061-2018 | 78467 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | nordicICE v 2.3.14 |  |  | July 03, 2015 | 2015 | NordicNeuroLab AS | The bug causes BOLD activation maps to be visualized as overlays without taking the coregistration into account. Any output created from these activation maps as overlays will not be adjusted according to the coregistration. | Software design | The firm notified their customers by email on July 3, 2015.Actions to be taken by the customer/User:We advise users to immediately stop using the BOLD module in nordiclCE and contact NordicNeuroLab for further instructions on how to upgrade the nordiclCE BOLD module, as described more fully below.For further questions, please call (262) 337-2909. |  | 121 | 121 licenses | Worldwide Distribution - USA (nationwide) Distribution and to the countries of : Australia , Belgium, Brazil, China, Czech Republic, Denmark, France, Germany, Greece, Hungary, India, Italy, Japan, Kuwait, Malaysia, Netherlands, Norway and Poland. | Open, Classified |  |  | 121 | Computer | Not_Computer
Z-2659-2018 | 80497 | [6 empty] | Leica M530 OHX Microscope System Supplied with 100 - 120 Vac. |  |  | June 25, 2018 | 2018 | Leica Microsystems, Inc. | Recall is due to a design weakness of the power insert module. | Device Design | Leica Microsystems notified customers on about 06/25/2018 via URGENT FIELD SAFETY NOTICE letter. Instructions included advising customers of the issue, informing them that Leica Microsystems representative will contact them to arrange the fiel update to the device, distribute the recall notification to all individuals that should be aware of the recall or if the device was further distributed and a response form was provided to complete and return. The notification letter also instructed customers to report any anomalies with the system to Leica Microsystems if they occur prior to the correction being made and to ensure a functional check is performed prior to each surgery as described in the user manual. |  | 123 | 123 devices total | United States distribution: AR, AZ, CA, CO, CT, DE, FL, IL, IN, KS, KY, LA, MO, NJ, NY, NV, OH, OR, PA, TX, UT, PR. International distribution to Japan, Brazil, Columbia, Ecuador, Mexico, Taiwan | Terminated | February 26, 2019 | 246 | 123 | Computer | Not_Computer
Z-2063-2018 | 80114 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | TL-HEX TruLok Hexapod System, Double Row Footplates, TL-HEX Sterile, All Sizes: 120 MM, 140 MM, 160 MM, 180 MM, 200 MM. |  |  | January 30, 2018 | 2018 | Orthofix Srl | The firm received complaints regarding the incompatibility and dimensional interference between the double row footplates manufactured in a limited period of 2014 and struts manufactured after September 2016. The firm determined that certain batches of the double row footplates are not compatible with the current version of the TL-HEX struts. | Device Design | Customers were notified on approximately January 30, 2018, via letter. Instructions included to examine inventory for affected devices, remove any affected devices and arrange for the return of them, notify customers if the devices have been further distributed, and complete and return the response form. The notification also provided alternative procedures to complete the surgical intervention and treatment: 1) the surgeon can use parallel external support (either footplate or ring) that does not have dimensional issues; 2) the surgeon can complete fixation temporarily using non-hexapodal TrueLok connection elements, but the solution requires post-operative non-surgical intervention such as additional parallel external support that does not have dimensional issues or the surgeon completes the frame for treatment using TruLok non-hexapodal elements; 3) in case the issue arises in a strut changed during treatment, the surgeon can use an Emergency Tab Kit (see attached diagram in letter). For further questions, please call (214) 937-2000. |  | 408 | 408 units | Worldwide Distribution - US Distribution within Texas., and to the countries of : Australia, Belgium, Chile, Denmark, France, Germany, Israel, Italy, Japan, South Africa, Spain, Sweden, Switzerland and United Kingdom. | Open, Classified |  |  | 408 | Computer | Not_Computer
Z-0687-2018 | 78704 | No | GXI | Radiofrequency lesion probe. | Neurology | Neurology | 510(k) | Nitinol TC Reusable Electrodes (TCN), Model Numbers: 1)TCN-10, 2) TCNK-10-R*, 3)TCN-10-3M, 4) TCNK-15, 5) TCN-15, 6) TCNK-15-C, 7) TCN-15-3M, 8) TCNK-15-R*, 9) TCN-20, 10) TCNK-20, 11) TCN-20-3M*, 12) TCNK-20-C*, 13) TCN-5, 14) TCNK-20-R*, 15) TCN-5-3M, 16) TCNK-5, 17) TCNK-10, 18) TCNK-5-C, 19) TCNK-10-C &amp; 20) TCNK-5-R*. |  |  | July 05, 2017 | 2017 | Cosman Medical, LLC | After multiple reprocessing cycles, the epoxy resin which holds the TCN Electrode in the hub can exhibit signs of damage. In some cases, this damage may result in the inability to fully remove blood and/or tissue residuals prior to cleaning and re-sterilizing the device. | Device Design | The firm, Cosman Medical, sent an "Urgent Medical Device Correction" notice and Field Correction Reply Verification Tracking Form dated July 5, 2017 to their customers. The notice described the product, problem and actions to be taken. The Customers were instructed to do the following:1. Immediately discontinue use of devices that have already been used.Used TCN Electrodes should be disposed of in accordance with your institutions infectious material/biohazard waste control procedures.2. Unused/un-opened TCN Electrodes should be identified as single use only (for one patient). These devices are supplied as non-sterile which require cleaning and sterilization prior to use. These TCN Electrodes may be used once the cleaning and sterilization steps are completed. After single use, the used TCN Electrode should be disposed of in accordance with your institutions infectious material/biohazard waste control procedures.3. Complete and return the Field Correction Reply Verification Tracking Form.via Email: Cosmanorders@bsci.com or via Fax to: 781-272-6563 - Identify your email with TCN Electrodes Action in the subject line.Customers with questions may call 1-888-8COSMAN or email Cosmanorders@bsci.com. |  | 11250 | 11250 units (10,460 USA; 790 OUS) | Worldwide Distribution: US (nationwide) and country of Canada. | Terminated | September 14, 2018 | 436 | 11250 | Computer | Not_Computer
Z-2888-2018 | 80704 | No | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | CIVCO lnfiniti Needle Guide Bracket. used for performing needle-guided (or catheter) procedures with the use of diagnostic ultrasound transducers. |  |  | February 23, 2018 | 2018 | Civco Medical Instruments Co. Inc. | A cover failed leak testing during validation testing of new needle guides from a replacement tool. | Device Design | Telephone Script to be read:CIVCO is calling to notify our customers regarding a recall action being initiated.CIVCO has recently discovered a dimensional issue in our Infiniti bracket which may allow the needle guide to a puncture a cover. You are being notified, as your facility has been identified as having received product from one of the affected manufacturing lots. Advise on Action to be taken by Customer or Distributor:To return Infiniti needle guide, contact your Account Representative at 800-445-6741 or 319-248-6757. Alternately you may also use email at order@civco.com to schedule the return of your device.Additional information to obtain:Verification of Email contact to send notification letter and response letter. |  | 177 | 177 boxes | Worldwide Distribution - US Nationwide in the states of GA, KY, CA, NY , NJ, TX, IL, MA, RI, AZ, and IA and countries of Ukraine, Hong Kong, Italy, Netherlands, Switzerland, United Kingdom, Denmark, Germany, Belgium, Spain, Japan, China, and Canada | Open, Classified |  |  | 177 | Computer | Not_Computer
Z-2491-2018 | 80288 | No | GBX | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Hydroline Trumpet Valve, Trumpet Valve, 5mm x 33cm Probe, Double Spike, Ref ASU1201, packaged 1/box, 12 boxes/case, single use, RX, Sterile. |  |  | June 11, 2018 | 2018 | Cardinal Health 200, LLC | The suction valve may not close properly which could cause continuous suction. | Device Design | The recalling firm issued letters dated 6/11/2018 via FedEx Next Day Priority Mail requesting return of the product. |  | 63020 | A total of 63,020 valves were distributed. Specific product quantities were not provided. | Distribution was made to AL, AR, AZ, CA, CO, DC, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MI, MN, MO, MT, NC, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, and WY. Government and military distribution was also made. Foreign distribution was made to Belgium, France, Germany, Israel, Italy, Netherlands, Singapore, and United Kingdom. | Open, Classified |  |  | 63020 | Computer | Not_Computer
Z-2889-2018 | 80769 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Vertebral Motion Analyzer (VMA), 2.3.250. |  |  | May 23, 2018 | 2018 | Ortho Kinematics, Inc | This email is to provide notification that, due to a software bug that has been corrected, the Radiological Read Report for the VMA test(s) contained an error. | Software design | Ortho Kinematics notified customers via Email and notification through web app with a notification dated December 6, 2016. The notification contained the following information: reason for correction, report information, and required action for acknowledgement. |  | 1 | 1 | U.S. Nationwide Distribution | Open, Classified |  |  | 1 | Computer | Computer
Z-0683-2018 | 79016 | No | LIX | Cardiopulmonary resuscitation (CPR) aid. | Cardiovascular | Cardiovascular | 510(k) | Philips QCPR Meter - used with FR3 |  |  | January 29, 2018 | 2018 | Philips Electronics North America Corporation | Inform users of correct placement of the disposable adhesive pad on the QCPR meter and inform users of injuries associated with proper performance of CPR. As required by Consent Decree of Permanent Injunction No. 1:17-cv-11955-DJC, entered October 31, 2017 between United States of America and Philips North America | Device Design | Philips sent an Medical Device Recall letter dated January 19, 2018 to their customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immediately return the attached reply card. For question contact your local Philips representative. |  | 28094 | 28094 in total | Worldwide Distribution - US Nationwide | Open, Classified |  |  | 28094 | Computer | Not_Computer
Z-2891-2018 | 80771 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Vertebral Motion Analyzer (VMA), 2.3.239 - 2.3.254. |  |  | March 02, 2017 | 2017 | Ortho Kinematics, Inc | Ortho Kinematics notified customers that errors were contained in Radiological Read Report and VMA Report for the VMA test. The error resulted from an administrative error of VMA user account settings and has been fixed. | Software design | Ortho Kinematics notified customers via telephone, email, and letter. Letter dated August 23, 2017 contained the following information: reason for notification, error report information, and requested acknowledgement of receipt. |  | 1 | 1 | U.S. Nationwide Distribution | Open, Classified |  |  | 1 | Computer | Computer
Z-2471-2018 | 79887 | No | OYC |  |  | Clinical Chemistry | PMA | t:slim G4 Insulin Pump with Dexcom G4 Platinum CGM |  |  | April 23, 2018 | 2018 | Tandem Diabetes Care Inc | The fuel gauge, the component that reads and reports the battery parameters to the pump, could provide inaccurate readings, which present to the user in one of two ways: A) by triggering a Malfunction 4 Alarm, or B) by triggering a succession of notifications prior to the pump shutting off, including Alert 2 (Battery Low), Alert 3 (Battery Very Low) and Alarm 12 (Battery Very Low). The alarms in both scenarios notify the user that the pump has stopped delivering insulin. | Device Design | On 04/23/2018, customers were sent Urgent Medical Device Voluntary Recall letters via overnight mail delivery and email. The letter instructed customers to call Customer Support at 1-877-801-6901 to arrange for replacement pumps. Customers were advised to continue using effected devices provided they kept Emergency Kits with backup insulin supplies with them. If customers noted unexpected battery behavior, malfunction 4 alarm, or the following alerts: Alert 2 (Battery Low), Alert 3 (Battery Very Low) and Alarm 12 (Battery Very Low), then customers were advised to discontinue use and use backup insulin supplies, while they waited for replacement pumps. The firm advised customers to fill out and return via fax, email, or mail Medical Device Recall Return Response forms. |  | 55 | 55 | U.S. Consignees: IN, AZ, UT, TX, NY, NJ, MD, MN, IL, CO, AL, WI, CA, PA, OH, OR, MS, FL, NM, ID, NV, KY, VA, | Open, Classified |  |  | 55 | Computer | Computer
Z-2890-2018 | 80768 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Vertebral Motion Analyzer (VMA), VMA DataLink Software.2.3.231, 2.3.232. |  |  | August 12, 2016 | 2016 | Ortho Kinematics, Inc | Ortho Kinematics, Inc. notified customers of erroneous results of VMA tests due to a software bug that has been corrected. | Software design | Ortho Kinematics, Inc. notified customers via email dated August 19, 2016, Notice of Correction to VMA Results. The notification included the following information: reason for notice of correction, error information, and required action of acknowledgement. |  | 1 | 1 | U.S. Nationwide Distribution | Open, Classified |  |  | 1 | Computer | Computer
Z-2885-2018 | 80414 | Yes | BZR | Breathing gas mixer. | Anesthesiology | Anesthesiology | 510(k) | O-Two EquinoxÂż Relieve N2O/O2 Analgesic Gas Mixing and Delivery System |  |  | May 14, 2018 | 2018 | O-Two Medical Technologies, Inc. | Higher levels of nitrous oxide delivered than specification | Device Design | O-Two Medical Technologies Inc. issued letter on May 14. 2018titled: Notice of Product Upgrade of the 0-Two Equinox Relieve Analgesic Gas Delivery System , requesting customers to return their units for inspection and upgrade. Contact via email to david@otwo.com. |  | 3 | 3 units | AK, CA, FL | Open, Classified |  |  | 3 | Computer | Not_Computer
Z-2479-2018 | 80464 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Hitachi Oasis MRI system |  |  | June 18, 2018 | 2018 | Hitachi Medical Systems America Inc | The stainless steel belt which runs the length of the table underneath the tabletop is exposed when the tabletop is extended into the magnet for scanning. The firm has received two complaints where technologists have cut their fingers on this belt. | Device Design | On June 18, 2018, the firm distributed an Urgent Medical Device Correction Notices to affected customers. Customers were informed of the issue and advised that the firm would be adding protective guards to both sides of the belt to prevent accidental contact. The firm expects to begin installing the guards around fall 2018.In the meantime, customers are advised to avoid direct contact with the stainless steel belt.Questions, comments, or requests may be sent to recalls@hitachihealthcare.com. |  | 264 | 264 | Worldwide Distribution - US Nationwide distribution.in the states of Mexico. | Open, Classified |  |  | 264 | Computer | Not_Computer
Z-2484-2018 | 80329 | No | NLQ |  |  | General &amp; Plastic Surgery | 510(k) | Reprocessed HARMONIC AceÂż +7, 5mm Diameter Shears with Advanced Hemostasis (HARH), Item Number HARH36 |  |  | May 08, 2018 | 2018 | Stryker Sustainability Solutions | Stryker Sustainability Solutions has received an increase in reports indicating that the affected Reprocessed HARH36 devices may display an error code (No Instrument Uses Remaining) upon initial connection to the generator. When this error code is present, the device is not able to be used. | Device Design | Firm notified sales reps via email on 5/8/18 and instructed to have the customers complete response forms. |  | 180 | 180 units | US Nationwide Distribution in the states of PA, FL, MD, MA, CA, NY, NC, TX, MT, AL, LA, TN | Open, Classified |  |  | 180 | Computer | Computer
Z-0926-2018 | 79150 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | i-STAT DE handheld data processing module for clinical use, Software Version 2.8, List Number:08K46-01115200 (UDI: 00054749001255) |  |  | September 12, 2017 | 2017 | Abbott Point Of Care Inc. | Issues resulting from upgrade to software version 2.8: (1) Location, operator, stored patient lists will not update, and (2) Customized Reference Ranges, Action Ranges, and Custom Reportable Ranges are reset to factory default values. No erroneous results are generated as a result of this issue. | Device Design | The following actions are recommended if you have already upgraded to version 2.8:1. If you are using Location, Operator, or Patient lists.a. Delete the following files:i. C:\Istat32\LocInstData\LocInstData.txtii. C:\Istat32\OpData\OpData.txtiii. C:\Istat32\PatData\PatData.txtiv. C:\Istat32\Log\iSTATDMI Errors.txtv. All files named C:\Istat32\Log\DeCustLogXX.txt; where XX is any two digits.b. Stop and restart the DECOMSRV service, which in turn restarts the DEDPSRV service.2. If you are using customized Reference Ranges, Action Ranges, and Custom Reportable Ranges:a. Open DE Customization Workspace.b. For each profile, select preferences by clicking the button with the profile name.c. Navigate to the Results tab.d. Verify the Action and Custom Reportable Ranges. Click Edit to make changes.e. If changes are made, click the OK button to save.If you have not upgraded to i-STAT/DE version 2.8, make arrangements with your data management vendor to update to the newer version of i-STAT/DE at your earliest opportunity. DO NOT UPGRADE without coordinating with your data management vendor so that continuity of your i-STAT testing is not impacted. How should you verify what version of i-STAT/DE is installed on a computer?Using the Program and Features app in Windows is the best way to determine the version of i-STAT/DE installed on your system. To do so, follow these steps:1. Open Control Panel and the Program and Features applet.2. Locate the program named i-STAT/DE Version 2.8 in the list.3. The version column will show 2.8 for the old version and 2.8.0.1 for the new version.If you determine the version 2.8.0.1 is not installed, make arrangements with your data management vendor to update to the latest validated version compatible with your data management system.Enclosed is an updated i-STAT/DE version 2.8, version qualifier, 2.8.0.1 CD-ROM and a Business Reply Card (BRC). Please complete and return the Business Reply c |  | 894 | 894 | Distributed in 48 States: AK, AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY, the District of Columbia, Australia, Italy, Sweden, and Canada. | Terminated | August 23, 2018 | 345 | 894 | Computer | Computer
Z-0711-2018 | 79214 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Axiom Artis model Interventional fluoroscopic x-ray foot switch accessory; Model Numbers: 4787797, 4787805, and 4787813. |  |  | September 28, 2017 | 2017 | Siemens Medical Solutions USA, Inc | The possibility exists for the wireless foot switch to fail due to impacts by external factors, such as electrostatic discharge that exceed a certain intensity. If the foot switch fails, it will no longer be able to be used to release radiation. The hand switch provided can be used to release exposures; however, fluoroscopic examinations will not be possible. | Device Design | Letters were sent to all consignees with instructions to contact their service provider to arrange for an appointment to update the device hardware. |  | 142 | 142 | Distributed in XX states: Ak, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, WY, the District of Columbia, Guam, and Puerto Rico. | Open, Classified |  |  | 2209 | Computer | Computer
Z-2501-2018 | 80481 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Visualase Cooled Laser Applicator System (VCLAS) labeled as: |  |  | March 23, 2018 | 2018 | Medtronic Navigation, Inc. | Medtronic has become aware that certain lots of the VCLAS 3MM, 10MM, and 15MM laser diffusing fibers were not tested at intended laser power levels, due to an issue related to calibration of the test equipment, which can lead to the potential for low laser energy output from fiber or fiber burnout at lower wattage. | Equipment maintenance | Customers were notified on about 03/23/2018 via URGENT: MEDICAL DEVICE RECALL letter. Instructions included to examine inventory for affected devices, contact Medtronic Technical Services at 1-888-826-5603 to arrange for the return of affected devices, and complete and return the Consignee Response Form. The recall was later expanded and customers were notified on about 05/16/2018 URGENT: MEDICAL DEVICE RECALL letter. Instructions did not change from the initial letter, but informed customers that the recall had been expanded to include two (2) additional lots. |  | 1283 | 1,283 devices total | Worldwide Distribution - US Nationwide in the states of AL, AR, AZ, CA, CO, DC, FL, GA, IL, IN, KY, LA, MA, MD, MI, MN, MO, NC, NJ, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, WI. International distribution to Israel. | Open, Classified |  |  | 1283 | Computer | Not_Computer
Z-2882-2018 | 80680 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Vertebral Motion Analyzer (VMA) Version 2.3.252. |  |  | January 11, 2017 | 2017 | Ortho Kinematics, Inc | Ortho Kinematics Inc. sent a Notice of Correction to Released Testing Results, Radiological Read Report for the Vertebral Motion Analyzer (VMA) test because it contained an error. The error occurred due to a software bug that has been corrected. | Software Manufacturing/Software Deployment | Ortho Kinematics Inc sent notification to customers on 1/11/2017 via Letter and electronic notification through web app. The company deployed a software upgrade (VMA version 2.3.253) on January 11, 2017. |  | 1 | 1 | US Nationwide Distribution | Open, Classified |  |  | 1 | Computer | Computer
Z-2496-2018 | 80400 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Ingenuity Digital PET/CT System, Model 882446 with the following UDI's: |  |  | May 21, 2018 | 2018 | Philips Medical Systems (Cleveland) Inc | The block used to connect the patient pallet to the couch may be assembled incorrectly. A screw may be misaligned in the patient pallet. | Software design | On May 21, 2018, Philips issued Customer Information notices to their clients. A Philips Field Service Engineer will contact customers directly to schedule the updates at your specific site. If you need any further information or support concerning this issue, please contact your local Philips representative. For SUPPORT in North America and Canada, contact the Customer Care Solutions Center 1-800-722-9377, follow prompts. |  | 32 | 32 units | Worldwide Distribution - US Nationwide in the states of CA, GA, LA, MD, MI, NY, OH, TN, TX, VT, Washington DC &amp; WIInternationally - Canada, Mexico, Belgium, France, Israel, Luxembourg, Netherlands, Singapore, Spain &amp; Switzerland | Open, Classified |  |  | 32 | Computer | Not_Computer
Z-2916-2018 | 80654 | No | LPL | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic | 510(k) | Johnson and Johnson Vision Care 1-Day ACUVUE MOIST Daily Wear, 8.5 Base Curve, -3.25 Power. Master Lot, containing 12 split lots, repackaged in 30 and 90-packs . |  |  | July 23, 2018 | 2018 | Johnson &amp; Johnson Vision Care, Inc. | Johnson and Johnson Vision Care received a number of reports of a partially formed contact lens, which appears to the customer or consumer as visibly smaller in diameter, thicker/stiffer to the touch and more darkly tinted blue than the normal contact lens. | Equipment maintenance | On July 23, 2018, the Johnson &amp; Johnson notified customers of the recall via an "URGENT FIELD SAFETY NOTICE" letter date July 23, 2018. The letter to US and international customers included the following instructions: review inventory and determine if any affected product is on hand, cease using any affected product, provide the recall notification to anyone within the organization who needs to be aware of the recall, return any affected product and arrange for replacement product, and complete and return the enclosed Customer Reply Form. The distributor notification included a copy of the appropriate recall notification depending on the country they are located in as well as the following instructions: to cease distribution of any affected inventory, remove from inventory any affected product, distribute the included RECALL NOTIFICATION to all customers where product was further distributed, ensure customers understand the importance of completing and returning the Customer Reply Form, arrange for the return of any affected product on hand.If you have any questions, contact Sr. Director, Worldwide Quality Systems Strategy at 904-443-3794 or aalvar56@its.jnj.com. Also contact Project Manager, Escalation &amp; Field Actions, phone 904-443-3516, email amyers@its.jnj.com. |  | 195660 | 195,660 Total; 89,640 US &amp; Puerto Rico, 106,020 OUS | Worldwide distribution: US (Nationwide) distribution to CA, CT, FL, GA, IA, IL, MA, MD, ME, MO, NC, NE, NH, NJ, NY, OH, PA, TN, TX, UT, and VA, Puerto Rico; and countries of: Canada, Barbados, Japan, and Korea. | Open, Classified |  |  | 195660 | Computer | Not_Computer
Z-2470-2018 | 80384 | No | LHN | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Proteus 235, Universal Nozzle with snout 300x400; PBS dedicated nozzle with snout XL |  |  | June 27, 2018 | 2018 | Ion Beam Applications S.A. | Gantry rotation or proton irradiation can be allowed even if apertures or accessories are not correctly positioned in snout 300x400 or snout XL. This can lead to risk of crushing for patient and/or user, and risk of mistreatment for the patient. | Device Design | Recall was initiated on 6/27/18. Firm plans to provide the Urgent Medical Device Correction notifications to IBA site managers present on customer sites. The site managers will deliver the notification to customers by hand or email.The recall letter alerts users to the potential issue and instructs users to verify that apertures and accessories are fully inserted with both sides in the guide. The firm plans to deploy a solution on or before May 31, 2019, that will prevent occurrences of the issue. |  | 14 | 14 | US Nationwide Distribution in the states of FL, VA, PA, OK, NJ, WA, and TX. | Open, Classified |  |  | 14 | Computer | Not_Computer
Z-2412-2018 | 80433 | Yes | NKE |  |  | Cardiovascular | PMA | Percepta CRT-P MRI SureScan, REF W1TR01 |  |  | June 21, 2018 | 2018 | Medtronic Inc., Cardiac Rhythm and Heart Failure (CRHF) | Potential for a device reset to occur in Percepta CRT-P MRI SureScan and Percepta Quad CRT-P MRI SureScan due to a timing interaction between the EffectivCRT Diagnostic and the Ventricular Safety Pacing feature (VSP). | Device Design | On 21-Jun-2018, Medtronic began notifying US physicians and health care professionals of this issue and provided Patient Management recommendations to ensure patient safety. A Software Update to SW040 Version 8.1 will be deployed to all Medtronic 2090 and Encore clinic and hospital programmers via the Software Distribution Network or USB. Field Representatives with customers who have Percepta CRT-P patients will be sent a USB device loaded with the updated software. In consultation with an Independent Physician Quality Panel, Medtronic recommends the following actions for physicians following patients:- Contact your local Medtronic Representative and schedule installation of the updated Percepta CRT-P Application Software (SW040 Version 8.1) onto Medtronic 2090 and Encore clinic and hospital programmers.- For a patient whose Percepta CRT-P device has experienced a Reset Alert or Observation: Consider scheduling an in-clinic device interrogation as soon as possible for the patients device to receive the automatic update.- For a patient whose Percepta CRT-P device has not experienced a Reset Alert or Observation: At their next scheduled in-clinic device interrogation, the patients device will receive the automatic update.Please complete the enclosed Clinician Confirmation Certificate and return via email to RS.CFQFCA@medtronic.com.For questions, contact your local Medtronic Representative or Medtronic Technical Services at 800-505-4636. |  | 1842 | 1842 units | Worldwide Distribution - US Nationwide | Open, Classified |  |  | 13718 | Computer | Computer
Z-2927-2018 | 80516 | No | GKR | Automated hemoglobin system. | Hematology | Hematology | 510(k) | RAPIDPoint 405 Blood Gas Analyzer; 10282093, 10310464, 10314817, 10317193, 10318999, 10320055, |  |  | July 16, 2018 | 2018 | Siemens Healthcare Diagnostics, Inc. | Therapeutic levels (1mg/mL and 2mg/mL) of Hydroxocobalamin may interfere with tHb and some of the CO-Ox fractions that are reported on the RAPIDPoint and RAPIDLab Blood Gas systems indicated. This may cause lower than expected values for carboxyhemoglobin (fCOHb) and methemoglobin (fMetHb). A negative interference with fCOHb has the potential to alter the medical assessment of the patient and may withhold necessary follow-up treatment and/or initiate cessation of medication in response to elevated fMetHb levels. | Device Design | The firm, Siemens Healthineers, sent an "Urgent Field Safety Notice" letter dated 7/2018 to its customers via FedEx on 7/18/18. The recall communication alerts customers that therapeutic levels of Hydroxocobalamin may interfere with tHb and some of the CO-Ox fractions that are reported on the RAPIDPoint and RAPIDLab Blood Gas Systems identified in the letter. Customer are advised to review the letter with their Medical Director, and complete and return the effectiveness check form within 7 days. Customers may continue to use their devices. Siemens is not recommending a review of previously generated results.The firm stated that they will be revising the Operators Guides for the 500 and 1200 series Operator's Guides with information on the interfering substance. The RapidPoint 405 Operator's Guide will not be revised due to the discontinuation of the manufacture of this analyzer. However, users should remain aware of the issue with hydroxocobalamine interference.Customers with questions were directed to contact their local Siemens technical support representative or Safety Officer (POC) at 800-441-9250 or email to: steven.andberg@siemens-healthineers.com. |  | 2248 | 2248 | Nationwide distribution. Worldwide foreign distribution. | Open, Classified |  |  | 14942 | Computer | Not_Computer
Z-2959-2018 | 80805 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | AMSCO V-PRO 1, V-PRO 1 Plus, and V-PRO maX Low Temperature Sterilization Systems |  |  | July 26, 2018 | 2018 | Steris Corporation | There is the potential for a scenario where a punctured cup of sterilant expires while the system is powered down, upon power-up the unit will not detect the expired cup of sterilant; allowing for subsequent cycles to be run. | Software design | The firm, Steris, sent an "Urgent Device Field Correction" letters dated 7/25/18 were distributed to customers. The letters included the following instructions for customers:STERIS Action  The system software for V-PRO 1, V-PRO 1 Plus, and V-PRO maX will be updated to ensure the VAPROX HC Sterilant expiration date is detected regardless of whether the unit is powered on or off. STERIS Service Technicians will contact all affected Customers and update the system software in the affected V-PRO Sterilization Systems.User Action  In the interim of the software update, as is normal practice, V-PRO 1, V-PRO 1 Plus, and V-PRO maX operators should place the unit into Standby mode when not in use in lieu of fully powering down the unit. As is stated in the Operator Manual, the unit should only be shut down during unit service activities as required:"Control Power Disconnect Switch - Located behind Operating End (OE) lower access door. This switch is used to disconnect power to the Sterilization Unit. Under normal operation, this switch remains in ON position at all times and is only accessed during sterilization unit servicing. This switch is designed to enable lock-out for maintenance purposes."In addition, users should continue to review their cycle printouts upon completion of each cycle to verify the VAPROX Sterilant was not expired.If you have questions regarding this matter, please contact Associate Product Manager, at (440) 392-7140, STERIS Customer Service at 1-800-548-4873, or your local STERIS Representative. |  | 2889 | 2889 units | Worldwide Distribution: US(nationwide) and to countries of: Albania, Algeria, Australia, Austria, Bahrain, Bangladesh, Belgium, Belarus, Brazil, Bulgaria, Canada, Cayman Islands, Chile, China, Colombia, Costa Rica, Curacao, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, Finland, France, Germany, Hong Kong, Hungary, Iceland, India, Indonesia, Iran (Islamic Republic of), Iraq, Ireland, Israel, Italy, Japan, Jordan, Korea, Republic of Kuwait, Lebanon, Libyan Arab Jamahiriya, Lithuania, Malaysia, Malta, Mexico, Morocco, Nepal, New Zealand, Oman, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Qatar, Romania, Russian Federation, Saudi Arabia, Singapore, South Africa, Spain, Suriname, Switzerland, Taiwan, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, Uruguay, Uzbekistan, Venezuela, and Vietnam. | Open, Classified |  |  | 2889 | Computer | Computer
Z-2968-2018 | 80723 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | McKesson Cardiology Hemo, Release SW version 13.0. |  |  | March 12, 2018 | 2018 | McKesson Israel Ltd. | Change Healthcare has identified an issue where, under certain circumstances, the Real Time Monitor (RTM) may not display physiological signals. | Software Design Change | The firm, Change Healthcare Israel Ltd., sent an "Urgent Field Safety Notice" dated on 03/12/2018 to its customers. The notice explained the product, issue and actions to be taken. The customers were instructed to: distribute the notice to all those within your organization who need to be aware; alert other parties affiliated with your organization, which are affected by this field safety notice;and maintain awareness of this notice and resulting action until the defect has been corrected, to ensure effectiveness of the corrective action.Change Healthcare has developed software updates to address the problem.If you have any questions, contact Change Healthcare Customer Support (US) 1-800-654-4366 or (UK) 440(0) 208-952-799 to arrange for the installation of software update/upgrade. |  | 31 | 31 units | US Distribution to states of: CA, CO, CT, FL, GA, KS, IN, LA, NJ, NY, OK, OH, SC, and TX; and internationally to: UK. | Open, Classified |  |  | 31 | Computer | Computer
Z-2411-2018 | 80293 | No | FLE | Steam sterilizer. | General Hospital | General Hospital | 510(k) | FDI FORT DEFIANCE INDUSTRIES P2131 AUTOMATED FIELD STEAM STERILIZER, NSN: 6530-01-641-4641 |  |  | May 24, 2018 | 2018 | Fort Defiance Industries, LLC | During long-term storage (i.e. military depot), the interaction of stagnant water and dissimilar metals in the plumbing of the P2131 device produces deposits in the water system that can potentially interfere with the rear sight gauge and / or the water liquid level switch. | Device Design | Fort Defiance Industries (FDI) notified the consignees by email on 05/24/2018. The letter explained the issue. FDI has revised the (1) Installation, Operational, and Performance Qualification (IQ/OQ/PQ) activities, (2) preparation for transport / storage, (3) preventive maintenance, and (4) device warning statements. FDI has revised the procedures for safety switch verification during IQ/OQ/PQ and complete water drainage before transport / storage of the device. |  | 199 | 199 units | US Nationwide Distribution in the states of PA. VA. TX | Open, Classified |  |  | 199 | Computer | Not_Computer
Z-0829-2018 | 79239 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | 2008T, Hemodialysis Delivery System |  |  | October 05, 2017 | 2017 | Fresenius Medical Care Renal Therapies Group, LLC | While reviewing documentation for the next software release of the 2008T, an R&D; technician identified that setting the UF goal to '0' introduces a discrepancy between the UF rate displayed and the actual UF pump rate. The software anomaly is also applicable to the 2008T Machines that contained the 2.63 Bug Fix per DCAF 17-088. 2008T Upgrade kits that contained the 2.63 Bug Fix are also affected. Additionally, SW version 2.64 (in design freeze) is also impacted by this anomaly. | Software Design Change | Fresenius Medical Care sent an Urgent Medical Device Correction letter dated April 24, 2017 to affected customers. the letter identified the affected product, problem and actions to be taken. For questions contact Technical Services at 1-800-227-2572. |  | 15 | 15 machines | US Nationwide Distribution | Open, Classified |  |  | 15 | Computer | Computer
Z-2958-2018 | 80804 | No | OLW | Electroencephalograph. | Neurology | Neurology | 510(k) | Li-ion battery packs used in BIS Vista and BIS View Monitoring Systems. |  |  | July 26, 2018 | 2018 | Covidien LLC | Revised replacement instructions for the Li-ion battery packs used in BIS Vista and BIS View Monitoring Systems | Device Design | Medtronic U.S. consignees were notified via Federal Express or certified mail on July 26, 2018, and the letter informs customers to inspect their batteries. Batteries with date code 3114 or older (Week 31 of 2014) exceed 4 years of age and should be removed from the monitor and disposed of according to facilitys standard process. Included with this notification are stickers to be placed on BIS" monitors when the battery pack is removed the stickers No Battery Installed should be placed on the monitors. The BIS Vista and View monitoring systems may be used without a battery pack only if connected to AC power source with emergency power backup system (Red Outlet). |  | 6937 | BIS 46,961; Spare Batteries: 6,937 | NationwideForeign: Australia, New Zealand | Open, Classified |  |  | 6937 | Computer | Computer
Z-0659-2018 | 79078 | No | MUE | Full-field digital mammography system. | Radiology | Radiology | 510(k) | Aspire HD (for FDR-1000AWS) |  |  | August 04, 2017 | 2017 | Fujifilm Medical Systems U.S.A., Inc. | FUJIFILM Medical Systems U.S.A., Inc. (FMSU) identified a potential failure with our Mammography system.The acquisition workstations, FDRÂż1000AWS, FDR-2000AWS, FDR-3000AWS, and CRÂżIR363AWS assign aunique ID number to every image study but very rarely, with the acquisition workstation software versions,VS.O, VS.1, VS.2 , V6.0, V6.1, and V7.0, the system may assign the same ID number to a new set of imagesthat was already assigned to the previous set of images due to the error in ID number generation logic. If anImage with this error (with duplicate ID number) Is transmitted to PACS, it may overwrite the image alreadystored on PACS. | Software design | The firm, FUJIFILM Medical Systems U.S.A., Inc. (FMSU), sent an "URGENT: MEDICAL DEVICE CORRECTION" letter dated 7/28/2017 to its customers. The letter described the products, problem and actions to be taken. The customers were instructed to please read and follow the instructions in the "ACTIONS TO BE TAKEN BY CUSTOMER/USER" section. If you have an error, contact your local FUJIFILM office; and complete and return the Customer Feedback Form.FUJIFILM service personnel will contact all of the medical facilities where the applicable products have been installed to arrange for this correction and visit to take corrective measures..If you have any questions about this matter, call Recall contact (203) 276-3445 or email: dharaben.desai@fujifilm.com. |  | 71 | 71 units (68 units Domestic. 3 units Foreign) in total | Worldwide Distribution. | Open, Classified |  |  | 71 | Computer | Computer
Z-0934-2018 | 79231 | No | LHN | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Proteus 235, Uniform and Pencil Beam Scanning |  |  | December 05, 2017 | 2017 | Ion Beam Applications S.A. | An accessory (range shifter, ridge filter&amp;) can be improperly inserted insnout 300x400 with one rail only and the proton therapy system can stilloperate. | Device Design | Ion Beam Applications (IBA) sent an Urgent Field Safety Notice dated December 5, 2017, to customers informing them of the issue and request that all users verify that the accessory is correctly inserted in snout 300x400 with both rails in the guide. A solution to this issue including a hardware modification will be deployed on or before March 31, 2018. For further questions, please call (215) 972-7985. |  | 14 | 14 | US Distribution to the states of : MA, FL, VA, PA, OK, TN | Open, Classified |  |  | 14 | Computer | Not_Computer
Z-2421-2018 | 80379 | Yes | LWS |  |  | Cardiovascular | PMA | Medtronic EnTrust Escudo D144DRG (OUS distribution) |  |  | June 21, 2018 | 2018 | Medtronic Inc., Cardiac Rhythm and Heart Failure (CRHF) | Medtronic has become aware of a potential for loss of high voltage and anti-tachycardia pacing therapy in EnTrust and Escudo implantable cardioverter defibrillators (ICDs) as they near elective replacement indicator (ERI) voltage. Under certain circumstances, the device may display an immediate End of Life (EOL) Observation with no prior ERI alert. Though no ERI alert is triggered, there may not be enough remaining battery capacity to charge the high voltage circuits, resulting in an excessive charge time EOL Observation, leading to a loss of high voltage and anti-tachycardia pacing therapy. | Device Design | On 21-Jun-2018, Medtronic began notifying US follow-up physicians, and Risk Managers of health care facilities of this issue via hand delivery by Field Representatives. In consultation with an Independent Physician Quality Panel, MDT recommends the physicians following patients with affected devices: "Consider scheduling an in-office patient follow-up as soon as possible to assess the potential for this issue per the steps described below. "Use a programmer to ensure the Excessive Charge Time EOL&and; the Low Battery Voltage ERI&amp; Patient Alerts have been programmed to On-High "Instruct patients to contact your office if they hear device alert tones. "If this issue has occurred, an EOL: replace device immediately Observation will be displayed on the programmer QuickLook report. Schedule device replacement immediately. The notification provided Patient Management recommendations to ensure patient safety and effective high voltage therapy remain as the device battery voltage approaches its 2.61V ERI threshold. |  | 1867 | 1867 units | Worldwide Distribution - US Nationwide | Open, Classified |  |  | 24741 | Computer | Computer
Z-0670-2018 | 78298 | No | OZO |  |  | Clinical Chemistry | PMA | Medtronic CareLink iPro Therapy Management Software, MMT-7340 |  |  | October 10, 2017 | 2017 | Medtronic Inc. | Due to a time conversion error, data uploaded from the iPro(TM)2 recorder to the CareLink iPro software using the new CareLink iPro Uploader feature may result in incorrect reports. This pertains only to reports generated between September 23, 2017 and September 27, 2017 that used the new CareLink iPro Uploader feature. | Software Manufacturing/Software Deployment | A communication is being sent to all healthcare professionals who had uploaded reports using the universal uploader feature between September 23, 2017 to September 27, 2017. The method is by mail and by email, with a response requested from the healthcare professional to confirm they have received and read the notice. For further questions please call (877) 874-7717 |  | 488 | 488 uploads | Worldwide Distribution - US Distribution | Open, Classified |  |  | 488 | Computer | Computer
Z-2422-2018 | 80282 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | AperFix AM Femoral Implant, Model numbers CM-2409 (9mm x 24 mm) and CM-2410 (10 mm x 24 mm). For orthopedic use in ligament reconstruction. |  |  | July 11, 2013 | 2013 | Cayenne Medical Inc. | Potential for AperFix 24 mm implant breaking inside the drilled femoral socket. | Device Design | Customers were contacted on approximately 07/11/2013 and were instructed to return any affected products on hand. |  | 976 | 976 units | US nationwide distribution. | Terminated | September 14, 2018 | 1891 | 976 | Computer | Not_Computer
Z-0938-2018 | 79230 | No | LHN | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Proteus 235, Proton Therapy System a medical device designed to produce and delivers a proton beam for the treatment of patients with localized tumors and other conditions susceptible to treatment by radiation. |  |  | November 17, 2017 | 2017 | Ion Beam Applications S.A. | There is a risk of misalignment of the patient when using treatmentplans with multiple isocenters. | Device Design | On November 17, 2017 an Urgent Field Safety Notice was issued to customers informing users that when creating or importing a treatment plan with multiple isocenters, the user has to enter Patient Positioning System positions manually. |  | 37 | 37 | Nationally | Open, Classified |  |  | 37 | Computer | Not_Computer
Z-2467-2018 | 80165 | No | IOL | Powered exercise equipment. | Physical Medicine | Physical Medicine | 510(K) Exempt | T2100 Micro flex Drive Treadmill, powered |  |  | April 02, 2018 | 2018 | GE Medical Systems Ultrasound &amp; Primary Care Diagnostics, LL | A performance issue with customer owned spare parts, T2100 Microflex drive (2026182-002 or 2026182-004), was not addressed with a previous safety correction. If these parts were installed from customer owned stock on the T2100 Treadmill, uncontrolled walking belt motion during a stress exercise test could occur. | Device Design | GE Healthcare sent an Urgent Medical Device Correction letter dated March 30, 2018 to customers. The letter identified the affected product, problem and actions to be taken. The letter stated that GE Healthcare will correct all affected products at no cost. Customers were asked to complete and return the attached "Customer Response" form via email to DCAR.30085@ge.com and a GE Healthcare representative will contact you to arrange for a correction if necessary. For questions contact GE Healthcare Service at 1-800-437-1171 or your local service Representative. |  | 3721 | 3,721 devices total | Worldwide Distribution - US Nationwide in the states of AZ, CA, CO, FL, GA, ID, IL, IN, KS, LA, MA, MD, MI, MO, NC, NE, NJ, NV, NY, OH, OK, PA, SC, TN, TX, WA, WI. International distribution to Kenya DEMOCRATIC REPUBLIC OF CONGO, ALGERIA, ARGENTINA, Australia, AUSTRIA, AZERBAIJAN, BAHRAIN, BANGLADESH, BELGIUM, BHUTAN, BOLIVIA, BOSNIA-HERZEGOVINA, BRAZIL, BRUNEI, BULGARIA, BURMA(MYANMAR), CAMBODIA, CAMEROON, CANADA, CAYMAN ISLANDS, CHILE, CHINA, COLOMBIA, COSTA RICA, CROATIA, CYPRUS, CZECH REPUBLIC, DEMOCRATIC REPUBLIC OF CONGO, DOMINICAN REPUBLIC, ECUADOR, EGYPT, EL SALVADOR, ESTONIA, ETHIOPIA, GEORGIA, GERMANY, GREECE, GUAM, Hong Kong, INDIA, INDONESIA, IRAQ, IRELAND, ISRAEL, ITALY, JAMAICA, JORDAN, AZAKHSTAN, KENYA, KOREA, KUWAIT, LEBANON, LIBYA, MACEDONIA, ALAYSIA, MAURITIUS, MEXICO, MOLDOVA, MOROCCO, NAMIBIA, NEPAL, NEW CALEDONIA, NEW ZEALAND, Nigeria, PAKISTAN, PANAMA, PARAGUAY, PERU, PHILIPPINES, POLAND, PUERTO RICO, QATAR, REUNION, ROMANIA, RUSSIA, SERBIA--MONTENEGRO, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SRI LANKA, SUDAN, SWITZERLAND, SYRIA, TAIWAN, TANZANIA, THAILAND, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, UGANDA, UKRAINE, UNITED ARAB EMIRATES, URUGUAY, UZBEKISTAN, VENEZUELA, VIETNAM, Virgin Islands, YEMEN. | Open, Classified |  |  | 3721 | Computer | Not_Computer
Z-0675-2018 | 79046 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Cobas 8000 Modular Series system; Software Version 05-02 and 06-03; UDI: 05641446001 |  |  | July 28, 2017 | 2017 | Roche Diagnostics Corporation | A software failure may incorrectly set the system settings to "default" settings, creating a risk of incorrect results. | Device Design | Consignees were notified by letter (receipt required) and requested to follow instructions to determine if the software issue has occurred. If the software issue is detected the consignee contacts customer support to have system reconfigured. |  | 399 | 399 | Distributed in 36 states: AL, AZ, CA, CO, CT, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, NC, NE, NH, NJ, NY, OH, OK, OR, PA, RI, TN, TX, UT, VA, WA, WI and Puerto Rico. | Open, Classified |  |  | 399 | Computer | Computer
Z-0928-2018 | 79275 | No | OXO | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | O-arm 1000 mobile image-intensified fluoroscopic x-ray system; (1) Model: BASE OARM Bl70000028100 SYS 100V, Product Number: Bl70000028100 (UDI: 00643169354418); (2) Model: BASE OARM Bl70000028120 SYS 120V, Product Number: Bl70000028120 (UDI: 00643169353411); (3) Model: BASE OARM Bl70000028120R SYS 120V RWK, Product Number: Bl70000028120R (UDI: 00643169353459); (4) Model: BASE OARM Bl70000028230 SYS 230V, Product Number: Bl70000028230 (UDI: 00643169353992); (5) BASE OARM Bl70000028230R SYS 230V RWK, Product Number: Bl70000028230R (UDI: 00643169354081) |  |  | January 25, 2018 | 2018 | Medtronic Navigation, Inc.-Littleton | Software update ("SW Update 3.2.1") is being implemented to address the following issues: Inability to power the system after shutdown, System stays in standalone mode, Image reconstruction, System/Pendant bootup, Dose display/report, Gantry motion, Network communication, System shutdown, and Early termination of 3D spin. | Software design | Medtronic sent an Urgent Medical Device Correction letter to their customers. The letter identified the affected product, problem and actions to be taken. Customers were informed that a local service representative will contact you to install the update. For questions contact Medtronic Technical Support at 1-800-595-9709. |  | 253 | 253 | Worldwide Distribution - US Nationwide in the states of AK, AL, AR, AZ, CA, CO, CT, FL, GA, IA, IL, IN, LA, MA, MI, MN, MO, NC, NE, NH, NJ, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, WY, and the countries of Puerto Rico. Distributed in Algeria, Australia, Chile, China, Czech Republic, Finland, France, Germany, Greece, India, Israel, Jordan, Lebanon, Sri Lanka, New Zealand, Nicaragua, Poland, Portugal, Russia, Saudi Arabia, Singapore, Spain, Switzerland, Thailand, United Arab Emirates, United Kingdom. | Open, Classified |  |  | 253 | Computer | Computer
Z-0931-2018 | 79228 | No | LHN | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Proteus 235, Pencil Beam Scanning version PTS-6.4.11.X |  |  | October 19, 2017 | 2017 | Ion Beam Applications S.A. | Range for the same energy could be different depending if the layer isthe first layer of the irradiation plan or not. | Device Design | IBA sent an Urgent Field Safety Notice dated October 19, 2017 , issued to customers stating the following: IBA recommends performing a check to detect the consequence of this potential power supply regulation issue on a regular basis (IBA recommends Monthly), see Annex 1 for details.IBA is developing a corrective action in order to reduce the hysteresiseffect and detect magnet current overshoot.This solution will be deployed to all systems using JEMA analog powersupply. This solution will be validated and deployed by April 2018. For further questions, please call (215) 972-7985 |  | 3 | 3 | US Distribution to the state of :FL | Open, Classified |  |  | 3 | Computer | Not_Computer
Z-2520-2018 | 79701 | No | FGB | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Uretero-reno fiberscope URF-P6RP6 |  |  | January 17, 2018 | 2018 | Olympus Corporation of the Americas | Breakage of the endoscope's insertion tube bending section during surgical procedures. | Device Design | On 1/17/18 an Urgent Medical Device Removal Action Notification letter was shipped to affected customer via Fed-EX. The notification informed customers the recall was initiated in response to customer complaints regarding the breakage of the endoscopes insertion tube bending section during surgical procedures.Customers were instructed to complete the following: 1. Inspect your inventory and identify any URF-P6/P6R models; 2.Olympus will contact your facility to make arrangements for return of your URF-P6/P6R fiberscope(s) for the device exchange. You will be provided instructions on returning the URF-P6/P6R for this exchange; 3.Olympus has discontinued previously distributed copies of the URF-P6/P6R Reprocessing Manual and the Operation Manual Inspect your inventory of the Reprocessing Manual and the Operation Manual, and discard any existing inventory of the URF-P6/P6R Reprocessing Manual and the Operation Manual; 4.Implement use of the enclosed Reprocessing Manual which recommends only sterilization methods, and the enclosed Operation Manual; 5. Ensure all reprocessing personnel are completely knowledgeable and thoroughly trained on the new reprocessing instructions in the new Reprocessing Manual; 6. If you may have further distributed the URF-P6/P6R, please identify your customers, notify them at once of this product recall, and appropriately document your notification process. Your notification to your customers may be enhanced by including a copy of this recall notification letter; 7. Please indicate on the enclosed questionnaire that you have received this notification. Return of the enclosed questionnaire will be deemed to be a request for a new replacement URF-P6/P6R fiberscope(s). Fax the completed form to 484-896-7128. |  | 2399 | 2399 | US Nationwide. Canada and Mexico | Open, Classified |  |  | 3461 | Computer | Not_Computer
Z-0693-2018 | 79086 | No | GAM | Absorbable poly(glycolide/l-lactide) surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Coated VICRYL RAPIDETM (Polyglactin 910) Suture |  |  | August 17, 2017 | 2017 | Ethicon, Inc. | On August 17, 2017, Ethicon had decided to initiate a medical device recall(removal) of the twenty-five (25) lots of VICRYL RAPIDETM (Polyglactin 910) Suturethat were distributed to the field. During finished goods testing, a suture damage wasidentified on a lot of VICRYL RAPIDETM Suture. During the course of the subsequentinvestigation, the suture also failed knot tensile strength testing. | Equipment maintenance | On August 31, 2017 a Medical Device recall letter (Reference 2210968-08.31.17-003-R) was issued requesting customers to do the following:1. Examine your inventory immediately to determine if you have product subject to this recall (removal) on hand and quarantine such product(s).2. Remove the product subject to this recall (removal) and communicate the issue to relevant operating room or materials management personnel, or anyone else in your facility who needs to be informed.3. If any product subject to this recall (removal) has been forwarded to another facility, contact that facility to arrange return.If you have additional questions regarding this recall (removal) or to report any customer complaints, please contact Ethicon Customer Support Center at 1-877-ETHICON (1-877-384-4266).As with any medical device, adverse reactions or quality problems experienced with the use of this product should be reported to the FDA's MedWatch Adverse Event Reporting program online, by regular mail or by fax:" Online: www.fda.gov/medwatch/report.htm" Regular Mail:Use postage-paid FDA form 3500 available at: www.fda.gov/MedWatch/getforms.htmMail to: MedWatch 5600 Fishers Lane, Rockville, MD 20852-9787" Fax: 1-800-FDA-0178 |  | 129372 | 129,372 Eaches (Internationally) | Nationally | Open, Classified |  |  | 129372 | Computer | Not_Computer
Z-0707-2018 | 79097 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Artis Zee and Zeego, Model Numbers: 10094135, 10094137, 10094141, 10094142, 10094143, 10280959, 10848283 &amp; 10848354 |  |  | June 06, 2017 | 2017 | Siemens Medical Solutions USA, Inc | Potential for liquid entry | Device Design | On June 6, 2017, Siemens Healthineers distributed URGENT MEDICAL DEVICE CORRECTION notices to their customers via USPS Certified Mail.Siemens conducted a corrective field action in June 2015 to protect the Artis table against liquid entry at critical locations. However, even with this corrective action, if there are larger amounts of liquid on the table surface, they can still enter the table. Siemens Medical Solutions USA, Inc. will contact customers to arrange a date to perform the corrective action. Please feel free to contact the service organization for an earlier appointment by calling 610-448-6461. |  |  |  | U.S. Nationwide | Open, Classified |  |  |  | Computer | Not_Computer
Z-2518-2018 | 79622 | No | OCS | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Uretero-reno videoscope URF-V2 |  |  | January 17, 2018 | 2018 | Olympus Corporation of the Americas | Potential breakage of the endoscope s insertion tube bending section during surgical procedures | Device Design | On 1/16/18 an Urgent Medical Device Removal Action Notification letter was shipped to affected customer via Fed-EX. The notification informed customers the recall was initiated in response to customer complaints regarding the breakage of the endoscopes insertion tube bending section during surgical procedures. Customers were instructed to complete the following: 1. Inspect your inventory and identify any URF-V2/V2R models; 2.Olympus will contact your facility to make arrangements for return of your URF-V2/V2R u videoscope(s) for the device exchange. You will be provided instructions on returning the URF-V2/V2R for this exchange; 3.Olympus has discontinued previously distributed copies of the URF-V2/V2R Reprocessing Manual and the Operation Manual. Inspect your inventory of the Reprocessing Manual and the Operation Manual, and discard any existing inventory of the URF-V2/V2R Reprocessing Manual and the Operation Manual; 4.Implement use of the enclosed Reprocessing Manual which recommends only sterilization methods, and the enclosed Operation Manual; 5. Ensure all reprocessing personnel are completely knowledgeable and thoroughly trained on the new reprocessing instructions in the new Reprocessing Manual; 6. If you may have further distributed the URF-V2/V2R, please identify your customers, notify them at once of this product recall, and appropriately document your notification process. Your notification to your customers may be enhanced by including a copy of this recall notification letter; 7. Please indicate on the enclosed questionnaire that you have received this notification. Return of the enclosed questionnaire will be deemed to be a request for a new replacement URF-V2/V2R videoscope(s). Fax the completed form to 484-896-7128. |  | 2548 | 2548 | US Nationwide and Canada | Open, Classified |  |  | 2548 | Computer | Not_Computer
Z-0776-2018 | 79132 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Intellivue iX Information Center Software - All PIIC iX Surveillance stations including: |  |  | January 03, 2018 | 2018 | Philips Electronics North America Corporation | Once a surveillance station is restarted on January 1, 2018 or later, the station will be unable to perform patient discharge and transfer operations. Any subsequent attempt to perform these operations will cause the station to restart, resulting in a short period of loss of monitoring at the Surveillance station during such restart. Until this issue can be corrected, users should avoid intentionally restarting their Surveillance stations in 2018. | Software design | To avoid this issue, do not intentionally restart any surveillance stations that have been running normally since the new year. You may use the workflows described in attachment Alternative Workflows for PIIC iX to avoid operations that may cause the surveillance station to restart. When a patient is discharged, confirm that the patient is discharged.Philips has initiated a correction to address this issue and will provide this software update to customers with affected devices at no charge. Software updates for PIIC iX A.0x and B.0x are expected to be available by January 8th.A Philips Healthcare representative will contact customers with affected devices to arrange for correction of the issue. Contact your local Philips representative if you have a question about any device affected by this issue. |  | 20357 | 20357 | Domestic and Foreign | Open, Classified |  |  | 20357 | Computer | Computer
Z-0896-2018 | 78609 | No | LKK |  |  | General Hospital | PMA | Medtronic Synchromed II Model 8637 supplied in 20mL and 40 mL reservoir size and a catheter. |  |  | August 09, 2017 | 2017 | Medtronic Neuromodulation | Medtronic implemented a final design change intended to address the motor stall due to corrosion and is recalling devices with the previous design that are not implanted. | Device Design | Medtronic sent an Urgent Medical Device Recall letter dated September 2017, to their customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete and return the enclosed Customers Confirmation form. For questions contact Technical Services at 1-888-854-0978. |  | 40572 | 27,547 US and 13,025 OUS | US Nationwide Distribution in the states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OR, OK, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, and WY | Open, Classified |  |  | 40572 | Computer | Not_Computer
Z-2572-2018 | 80424 | No | LYJ |  |  | Neurology | PMA | VNS Therapy Programming System, Rx Only, Model 3000, v1.0.2.2 |  |  | January 24, 2018 | 2018 | LivaNova USA Inc | Unintended warning message displayed on generators programmed with a Model 3000 v.1.0.2.2 programmer. | Software design (manufacturing process) | Consignees were notified on about 01/24/2018 via URGENT: MEDICAL DEVICE CORRECTION letter. Instructions for physicians with new devices (interrogated for the first time) include to 1) apply the changes and proceed with programming as usual. The warning message requires a mandatory action to Apply Changes . After applying changes, therapy will automatically be programmed to 0 mA. The device may be immediately programmed back on to re-enable therapy. 2) To complete and return the attached Customer Response Form (by fax to (281) 853-1248 or by email to M3000FieldAction@livanova.com. Instructions to physicians with devices that have been implanted (previously interrogated) include 1) Restore the normal and magnet mode stimulation parameters from the Summary screen immediately after the warning message screen. If desired, additional programming changes may be made. 2) For AutoStim (M106 only), detection must be re-enabled first, prior to re-enablingAutoStim therapy. The previous AutoStim parameters are not retained and must be reentered manually. If needed, the previous parameters may be found by ending the session and viewing the Session Report for the generator. 3) Please complete and return the attached Customer Response Form by fax to (281) 853-1248 or by email to M3000FieldAction@livanova.com. Additionally customers were instructed to ensure that this notice is communicated to all relevant personnel within the organization. For questions regarding the information in the letter, please contact Clinical Technical Supportat (866) 882-8804 (Monday to Friday, 8 AM to 5 PM CST) or e-mail at cservices@livanova.com. |  | 243 | 243 devices total | Distributed domestically to AZ, CA, GA, IA, IL, IN, MO, MT, NC, NY, PA, SD, TN, TX, WA, WI. | Open, Classified |  |  | 243 | Computer | Computer
Z-2571-2018 | 80420 | No | HTW | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Drill (Fluted-Hard Bone), for Quattro Link 2.9mm, labeled as: |  |  | October 10, 2013 | 2013 | Cayenne Medical Inc. | A review of complaints identified a trend for drills breaking. | Device Design | Customers were contacted on approximately 10/10/2013 and were instructed to return any affected products on hand. |  | 229 | 229 devies total | US distribution to AK, AL, AR, AZ, CA, CO, FL, GA, IA, ID, IL, IN, KS, MA, MD, MI, MN, MO, MT, NC, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WI. | Terminated | February 26, 2019 | 1965 | 229 | Computer | Not_Computer
Z-2746-2018 | 80615 | No | LHN | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MEVION S250/MEVION S250i Proton Beam Radiation therapy system |  |  | June 29, 2018 | 2018 | Mevion Medical Systems, Inc. | QFix has advised Mevion Medical Systems that they are identifying unauthorized combinations of QFix Couch Tops with non-QFix devices in clinical use. The kVue Oen Touch latch consists of a precision latch mechanism that can be damaged when used with incompatible (third-party) accessories. This can create unsafe conditions for use and may compromise the subsequent use of any kVue inserts, whether authorized or not. Use ofunauthorized or unvalidated inserts may lead to an incorrect patient treatment or physical injury. | Device Design | On June 29, 2018, the firm sent out a User Notice to consignees to alert customers to the Advisory Notice issued by QFix regarding the use of an Orfit HP Pro Solution combined with the QFix kVue couch top.Customers using third-party inserts are advised to immediately inspect the latching mechanism for damage and contact the third-party insert vendor for advice on the use of the insert and confirmation of compatibility.Mevion will update user manuals to include the following warning about using alternative third party vendor inserts:WARNING: Interchangeable Couch Top inserts are integral to the immobilization,treatment beam planning, and physical safety of each patient. Use ofunauthorized or unvalidated inserts may lead to an incorrect patienttreatment or physical injury. Only use Mevion-provided accessories ordevices that are confirmed by the third party suppliers to be compatiblewith the Mevion Couch for safe and effective patient treatment.Please contact Rob Cessac at rcessac@mevion.com or call 636-485-4962 for anyquestions or concerns. |  | 7 | 7 | US Nationwide Distribution in the states of DC, FL, MO, NJ, OH, and OK. | Terminated | October 02, 2018 | 95 | 7 | Computer | Not_Computer
Z-2745-2018 | 80257 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | G8 Automated HPLC Analyzer: HLC-723G8-ST, 021560; and HLC-723G8-LA, 021674 |  |  | June 06, 2018 | 2018 | Tosoh Bioscience Inc | HbAE is known to interfere with the HbA1c assay on the current version of software, Ver. 5.23. Customers should exercise caution when reviewing chromatograms and ensure that Flag 43 is enabled on their device to avoid reporting invalid test results in the presence of HbAE. | Software design | The firm issued an Urgent Medical Device Recall letter to its affected customers on June 6, 2018. In the notice, Tosoh is currently investigating a resolution to this issue to eliminate interference from hemoglobin variants. In the interim, customers should carefully review the chromatography to determine whether or not a variant peak is detected and follow the recommendations included in the letter.Customers were asked to complete and return the attached Acknowledgement Form and the flag setting confirmation to Tosoh Bioscience, Inc., within 14-days of receiving this notification by FAX, email, or postal mail as indicated on the form.If you have obtained unexpected test results or received any complaints of illness or adverse events associated with the use of the analyzer, please contact Tosoh Technical Support 24 hours a day, seven days a week at (800) 248-6764. |  | 665 | 665 | Worldwide Distributed - US Nationwide. and the countries of Foreign distribution to Brazil, Caribbean, Chile, Colombia, Costa Rica, Dominican Republic, Ecuador, Panama, Peru, and Uruguay. | Open, Classified |  |  | 665 | Computer | Computer
Z-2568-2018 | 80023 | No | OLV | Electroencephalograph. | Neurology | Neurology | 510(k) | NicoletÂż Ambulatory EEG |  |  | February 01, 2018 | 2018 | Natus Neurology Inc | Potential to be able to import an ambulatory exam for a patient for which the exam was not started. | Software design | Customers were notified via letter on about 02/01/2018. Instructions to ensure the empty exam is removed appropriately after starting the ambulatory recording included: 1) Only shutdown Recorder via the Recorder UI, either from the File menu or the X button on the top right of the application; 2) If any of the four conditions above occurs and the 2.60 MB exam is not removed, the emptyexam can be deleted from NicVue prior to importing the ambulatory exam, doing so will allow the patient ID check to occur normally. Customers were also instructed to provide the recall communication to all applicable users and complete and return the verification form. |  | 171 | 171 units | Distributed domestically to AR, AZ, CA, CO, DC, FL, GA, ID, IL, IN, MA, ME, MI, MO, N Y, NC, ND, NJ, NV, NY, PA, TN, TX, VA, WA, WI. Distributed internationally to Australia, Belgium, China, Germany, Hong Kong, Israel, Italy, Jordan, Kuwait, Russia, Slovenia, Tunisia, United Arab Emirates, Vietnam. | Open, Classified |  |  | 171 | Computer | Not_Computer
Z-0846-2018 | 79133 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | A-Series A3/A5 Anesthesia Delivery System; Model Numbers: 0633F-01000-0X (A3) and 0631F-01000-0X (A5) |  |  | July 14, 2017 | 2017 | Mindray DS USA, Inc. dba Mindray North America | A software issue may result in the previous settings being applied instead of the default settings or the unit may skip the startup leak test. | Device Design | Mindray sent an Urgent Medical Device Correction letter dated June 9, 2017 to affected customers. The letter identified the affected product, problem and actions to be taken. Corrective Action: In order to eliminate any potential for the EEPROM issue to manifest, Mindray is offering two options for upgrading affected systems in your facility: Option 1: You may perform the software upgrade if your facility has Biomedical Engineering or other staff that have been trained by Mindray to perform maintenance operations on the ASeries. Please contact Mindray's Technical Support team who will assist with providing the software and instructions for performing the software upgrade. The Technical Support can be reached at (877) 913-9663 (Option 1). Option 2: A Mindray Service Representative will upgrade the affected systems in your facility at no charge. Please contact your Mindray Service Representative to arrange for this upgrade at (877) 913-9663 (Option 8). Both the Technical Support and Service teams may be reached Monday through Friday, 8:30am - 5:30pm, ET. |  | 2875 | 2875 in total | Worldwide Distribution - US Nationwide | Open, Classified |  |  | 2875 | Computer | Computer
Z-2594-2018 | 80421 | No | GDJ | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | MAYFIELD INFINITY XR2 LOW PROFILE BASE UNIT, Part Number A2079A |  |  | June 29, 2018 | 2018 | Integra LifeSciences Corp. | There is a potential for the knob assembly in the swivel adapter to fracture/break during use. | Device Design | The firm, Integra, sent an "Urgent Voluntary Medical Device Correction Field Safety Notice" dated 6/29/18 to customers. The notification letter instructed customers to perform the following: 1) Review your MAYFIELD XR2 instructions for use. Please pay particular attention to the CAUTIONS listed on page 5 of the accompanying instructions for use section as noted below." Always make sure the base unit is properly secured to the operating table" The base unit must not be used if the device appears to be damaged or functioning incorrectly" Over-tightening the mechanisms adjustment screws may result in damage to the unit" Over extending or overloading the base unit may result in unintended movement, shortened product life and/or damage to the unit" Do not alter the construction of this device as it may result in serious patient injury" Do not steam sterilize; components may be damaged by high heat leading to device damage and reduced performance2) Visit Integra website and review video highlighting instructions for use at the following link: https://www.integralife.com/mayfield-imaging-cranial-stabilization-system/category/cranial-stabilization-mayfield-imaging-cranial-stabilization-system3) Complete and return the attached Field Safety Notice Acknowledgement Form promptly via email: FCA@integralife.com or Fax to: 1-609-750-4220.Distributors were also instructed to: "We also kindly ask you to contact the final customers who may have the affected products and provide them with this letter."On July 20, 2018, an additional Urgent Voluntary Medical Device Recall letter was distributed to customers. The letters ask customers to perform the following:1. Allow your sales representative to check the current knob assembly on your MAYFIELD Infinity XR2 Radiolucent Base Unit (A2079), Standard; MAYFIELD Infinity XR2 Low Profile Base Unit (A2079A); MAYFIELD Infinity XR2 Radiolucent Base Unit, Extended (A2079E) and MAYFIELD Spine Table Adaptor ( |  | 1 | 1 units | Worldwide Distribution: US (nationwide) to states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WI, and WV; and to countries of:: Australia, Canada, China, Dominican Republic, Hong Kong, India, Japan, Korea, Malaysia, Mexico, Pakistan, Panama, Peru, Singapore, Sri Lanka, Taiwan, and Thailand. | Open, Classified |  |  | 1044 | Computer | Not_Computer
Z-0815-2018 | 79019 | joint, knee, external limb component | ISY | External limb prosthetic component. | Physical Medicine | Physical Medicine | 510(K) Exempt | OTTO BOCK Kenevo knee joints Model 3C60=ST |  |  | September 12, 2017 | 2017 | Otto Bock Healthcare Product | Otto Bock Healthcare Products GmbH has identified a design issue through extended durabilitytesting where the pylon clamping mechanism has the potential to cause damage to thepylon. When used longer than 1.45 million cycles (which represents approximately 2.3years of use of the knee joint under heavy usage conditions), AND the pylon clamp is overtightenedaccording to the labeled torque specification, AND the patient weighs at most125kg it may be remote that the pylon breaks, and causing the patient to fall.There have been no failures or complaints reported directly attributed to the failure mode collapsed pylon .In the US Market and in Canada, labeling requires a condition based service at three years.In the rest of the world, a mandatory two year service interval is specified. Only the devicesin distribution in the United States and Canada are affected by this action. | Device Design | On September 14, 2017 a letter was sent to all their consignees via 2nd day signature receipt required to customers instructing them to contact their patients and send in their devices for service within the first 2.3 years of use. The pylon clamp will be replaced at no charge. The pylon will also be inspected and replaced if necessary. The customer will be requested to complete a reply form indicating they have received the notice and intend to contact their patients to arrange for the requested service.Customer Service: (800) 328-3458  Option 1Loaner Service: (800) 328-3458  Option 3, then Option 1 |  | 4 | 4 | AK, AR, AZ, CA, CT, DC, GA, IL, IN, KY, LA, MD, MI, MN, NC, ND, NJ, MN, NY, OH, OK, OR, PA, PR, TX, VA, and WAAU, BE, DE, FR, GB, IT, KW, LU, NL, NO, AT, SA, SE, CH, ES, ZA, and HU | Open, Classified |  |  | 158 | Computer | Not_Computer
Z-0821-2018 | 79262 | No | OLP | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Ultra Renew Plus |  |  | December 04, 2017 | 2017 | Accord Media, LLC | Distribution of medical devices with unapproved green LED light used to treat hyperpigmentation | Device Design | On December 4, 2017, Accord Media distributed Urgent Medical Device Field Action notices and Response Forms to customers via courier service. Customers were instructed to inspect stock and identify affected product whereby the would apply a label to the product which reads "DO NOT USE GREEN LIGHT FUNCTIONALITY - NOT FDA CLEARED FOR USE". Customers are encouraged to complete and return the response form via email to: customerservice@truthinaging.com. |  | 1040 | 1040 units | US Nationwide | Open, Classified |  |  | 3322 | Computer | Not_Computer
Z-2604-2018 | 80499 | No | MNI | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Modera Modular Pedicle Screw System Shank Insertion Instrument, 3-107-07 |  |  | February 13, 2017 | 2017 | New Era Orthopaedics, LLc | This is a retrospective recall of an instrument used in spine fusion procedures with Modera Modular Pedicle Screw system implants. The firm received a customer complaint in which excessive forces were applied to the shank insertion instrument and the laser weld failed. | Device Design | On February 13, 2017, the firm emailed its distributors to notify them that their Rev. A devices would be replaced with Rev. B devices, and asked the distributors to mail back the Rev. A devices. |  | 6 | 6 pieces | 2 distributors in CO and TX. | Terminated | November 07, 2018 | 632 | 6 | Computer | Not_Computer
Z-0808-2018 | 79178 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo.via systems with SW VB20A |  |  | August 31, 2017 | 2017 | Siemens Medical Solutions USA, Inc | A software update was released on August 22, 2017. Performing this software installation and upgrading syngo.via system to the software version VB20A_HF02 will resolve various bugs and inconsistent diagnosis with the old system. | Software Design Change | On August 25, 2017 an URGENT letter was issued to customers notify of a system update for Syngo.via systems with SW VB20A. The software was released on August 22, 2017. Several bugs and inconsistent diagnoses were detected and this software is intended to correct these. If you have any questions, call Siemens Customer Care Centre at 888-826-9702 (24/7). |  | 30 | 30 | Nationwide Distribution. | Terminated | October 01, 2018 | 396 | 30 | Computer | Computer
Z-0826-2018 | 79176 | No | JLS | Progesterone test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ST AIA-PACKÂż PROG III, Progesterone III Assay, Part Number: 025240 |  |  | February 13, 2018 | 2018 | Tosoh Bioscience Inc | Falsely elevated progesterone assay values will occur when using certain progesterone assay part numbers due to cross reactivity of DHEA-S, a metabolite of the DHEA (dehydroepiandrosterone) supplement. | Device Design | The assays identified in the above table must not be used to report results for patients who are taking DHEA supplements." For patients that are taking DHEA supplements, an alternative testing method which is not expected to show cross reactivity to DHEA-S should be used such as Liquid Chromatography-Mass Spectroscopy (LCMS)." Please review the information in this Urgent Medical Device Recall with your Medical Director and/or Lab Director, as soon as possible." Continue to use the Progesterone Assays for patients who are not receiving DHEA supplements or therapy." Complete and return the attached Urgent Medical Device Recall Acknowledgement Form within 15-days." Maintain this notification with your laboratory records and forward this information to others who may have received this product." If you have obtained unexpected test results or received any complaints of illness or adverse events associated with the use of the products identified in the table above, contact Tosoh Technical Support 24 hours a day, seven days a week at 1(800) 248-6764 or email Bernadette.OConnell@tosoh.com." Should you have any questions regarding this medical device recall, please feel free to contact Bernadette OConnell at 1(800) 248-6764, or by email at Bernadette.OConnell@tosoh.com. We will be available to answer any questions Monday - Friday, from 8:00 AM to 5:00 PM (PST). |  | 3322 | 3322 | Distributed in 39 states: AK, AL, AR, AZ, CA, CO, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MI, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, and WI. | Open, Classified |  |  | 17802 | Computer | Not_Computer
Z-2552-2018 | 80500 | Yes | BTM | Manual emergency ventilator. | Anesthesiology | Anesthesiology | 510(k) | AirLife Resuscitation Device, Adult labeled as the following: |  |  | May 08, 2018 | 2018 | Vyaire Medical | An error in manufacturing of the AirLife Resuscitation Devices that has the potential to obstruct airflow of the supplemental oxygen delivery during use. | Equipment maintenance | Vyaire Medical notified customers on about 05/08/2018 via URGENT RECALL NOTIFICATION letter. Instructions included to conduct 100% physical inventory for any affected devices, destroy any affected devices according to each facility's destruction protocol, notify customers of the recall if affected devices have been further distributed, and complete and return the Customer Response Form. Questions and comments can be directed to Vyaire Medical's Customer Support Team at 1-833-327-3284 (Option #1) or email: customersupport@Vyaire.com Monday-Friday 8:00 am - 5:30 pm CST. |  | 15714 | 15,714 units total | US distribution only to: AL, AR, AZ, CA, CA, CO, FL, GA, HI, IA, IL, KY, LA, MA, MD, MI, MN, MO, MS, NC,NJ, NY, OH, OR, PR, SD, TX, VA, WA. | Open, Classified |  |  | 15714 | Computer | Not_Computer
Z-0775-2018 | 79146 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | 979195 Loaner Kit for femoral intramedullary nail. Orthopedic instrument. |  |  | September 13, 2017 | 2017 | Zimmer Biomet, Inc. | Fracture of the bolt connecting the insertion guide to the intramedullary nail during insertion of the femoral nail. | Device Design | On September 13, 2017 an URGENT MEDICAL DEVICE RECALL REMOVAL was issued to customers requesting that they review inventory and quarantine affected product, which will be immediately returned to recalling firm. Also to notify any customers who may have received product. Questions or concerns can be directed to Customer Service at 574-371-3071 |  | 337 | 337 total | Distribution includes US Nationwide. | Open, Classified |  |  | 337 | Computer | Not_Computer
Z-2554-2018 | 80515 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | RayStation Radiation Therapy Treatment Planning System, Model nos. 2.5, 3.5, 4.0, 4.3, 4.5, 4.7, 4.9, 5.0, 6.0, 6.1, 6.2, 7.0 |  |  | June 06, 2018 | 2018 | RAYSEARCH LABORATORIES AB | The firm has learned that some RayStation/RayPlan users have commissioned machines with erroneous Beam profile correction parameters. These parametersaffect the dose calculated in corners of large or off-axis fields. This effect cannot be seen in the Beam Commissioning module and dose in large or off-axis fields needs to be validated using the Beam 3D Modeling module in RayPhysics/RayPlan Physics. The user mustbe aware to avoid incorrect dose calculations during treatment planning. | Software design | On June 6, 2018, the firm issued an Urgent Field Safety Notice, Medical Device Correction letter to affected customers via email. The notice was a reminder to follow the instructions listed in the product documentation when creating aphoton beam model in RayStation/RayPlan.ACTIONS TO BE TAKEN BY THE USER" Always validate the beam model for all clinically relevant fields, including, but not limited to, largeand off-axis fields, before the system is used to create clinical treatment plans." Carefully review all beam model parameters before commissioning." Be aware that the Beam profile correction values may need to be adjusted in order to correctlymodel the primary collimator that limits the dose in the corners of large or off-axis fields." Review any existing photon beam models to ensure that the actions above have been properlyperformed.Please educate physics staff about these user responsibilities.Inspect your product and identify all installed units with the above software version number(s), then confirmthat you have read and understood this notice by replying to the notification email.Safe use relies on careful validation of the beam model for all clinically relevant fields before the systemis used to create clinical treatment plans. This information is already present in the Instructions For Use,the Reference Manual and the RayPhysics/RayPlan Physics Manual for the existing product versions. Tofurther reduce the risk of using improper beam models, the safety warnings within the product documentationwill be updated in the next version of RayStation/RayPlan, scheduled for market release in June 2018. |  | 718 | 718 | US Nationwide distribution. | Open, Classified |  |  | 718 | Computer | Computer
Z-2557-2018 | 80376 | No | JOY | Automated cell-locating device. | Hematology | Hematology | 510(k) | cobas m 511 integrated hematology analyzer, Model Number 07261691190 |  |  | May 07, 2018 | 2018 | Roche Diagnostics Hematology | Discrepant results have been reported in patients with severe microcytic anemia (e.g.iron deficiency, thalassemia) and the following parameters are affected: RBC (red blood cell count), HGB (hemoglobin concentration), MCH (mean corpuscular hemoglobin), HCT (hematocrit), and MCV (mean corpuscular volume). | Software design | An Urgent Field Safety Notice was distributed to customers via email on 5/7/18. The notification included the following actions for to be taken by the customer/user: Until this rule is implemented, when cobas m 511 integrated hematology analyzer displays a HGB value of less than or equal to 9g/dL, Roche advises the user to not report results for the RBC, HGB, MCH, HCT, and MCV parameters. Roche advises the user to perform laboratory confirmatory testing before making transfusion decisions. Delta check against previous results might additionally be used to evaluate the results.For customer evaluating the instrument Roche kindly advises to not report values for diagnostic use until the rule is configured by your Roche Application Specialist. |  | 17 | 17 | No distribution in the United States.The devices were distributed to the following foreign countries: Austria, France, Germany, Hong Kong, Japan, Netherlands, Pakistan, Singapore, South Korea, Switzerland, Thailand, and Vietnam. | Open, Classified |  |  | 17 | Computer | Not_Computer
Z-0716-2018 | 79099 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Allura Xper systems with Certeray X-ray Generator located in regions with high mains (480V) voltage. |  |  | January 04, 2018 | 2018 | Philips Electronics North America Corporation | There is an increase in the failure rate of certain Anode Drive Units (ADU5) used in these products. Early life failure of the ADUs only occurs when the hospital mains provide 480V to the system. This current voltage may lead to saturation and overheating of the coils that protect the IGBTs (insulated-gate bipolar transistor) of the ADU. The overheating may also generate a peculiar burning smell that may be noticed. In some instances this burning smell may be noticed before the ADU fails. When the ADU fails the X-ray performance of the system is reduced toEmergency Fluoroscopy. Exposure will not be possible and image quality is reduced. | Device Design | Customer Notification letter dated August 2017 were sent to affected customers. The letter identified affected product, problem and actions to be taken. The customer is requested to place this Field Safety Notice with the documentation of the system until this correction has been implemented. For questions contact your local Philips representative. |  | 705 | 705 | Worldwide Distribution - US Nationwide; foreign distribution to Brazil, Germany, Iceland, Japan, Mexico, and (updated 6/25/18) Saudi Arabia. | Open, Classified |  |  | 705 | Computer | Computer
Z-2854-2018 | 80577 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | DigitalDiagnost 4.1 Flex / Value (Stitching Patient Support) 712028 |  |  | June 27, 2018 | 2018 | Philips Medical Systems Gmbh, DMC | The hook does not securely hold the footplate in the vertical position. Therefore, the footplate can fall or swing down and injury operator or patient. | Device Design | On June 27, 2018, Philips Medical Systems issued URGENT FIELD SAFETY NOTICES, Medical Device Recall letters to customers. Customers were encouraged to always check the right position of the hook and ensure that the hook is locked according tot he IFU. Press hook down until it is locked. Make sure that the hook is in position. A Philips Service Engineer will contact impacted customers as soon as the Field Action Kit is ready to be implemented. Should you need to communicate with Philips in regards to this program, please reference FCO-71200185. If you need further information or support concerning this issue, please contact your local Philips representative. |  |  |  | US Nationwide; International to 66 countries | Open, Classified |  |  |  | Computer | Not_Computer
Z-2524-2018 | 80417 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | ARTIS Pheno, Fluoroscopic X-Ray system, Model Number 10849000 |  |  | May 30, 2018 | 2018 | Siemens Medical Solutions USA, Inc | During manual LAO/RAO rotation at maximum speed or a combination of LAO/RAO and Cran/Caud rotation at maximum speed, the possibility exists for the safety monitoring within the system control to become activated. As a result, the C-arm stops the rotation and can drop up to 40 mm (1.6 ). | Device Design | An Urgent Medical Device Correction notification was distributed to customers on 5/30/18. The notification provided the following information to customers:Siemens will correct the error by inspecting all affected ARTIS Pheno systems. As an immediate measure to prevent movement stops during a C-arm rotation followed by a C-arm drop, the maximum rotation speeds will be reduced. This correction limits the maximum manual LAO/RAO rotation from 25Âż/sec to 20Âż/sec and Cran/Caud rotation from 20Âż/sec to 10Âż/sec.All other system parameters as well as 3D acquisition and Perivision/Peristepping remain unaffected. You may continue to use the ARTIS Pheno system while awaiting implementation of the corrective action.Our service organization will contact you shortly to arrange a date to perform this corrective action with UpdateInstruction AX004/18/S. Please feel free to contact our service organization for an earlier appointment at 1-800-888-7436.Following the corrective action, the cause is eliminated and a recurrence of the fault is prevented. Additionally, Siemens is developing a solution to restore the system back to the original settings. This solution will be distributed to all affected customers upon completion. |  | 40 | 40 in the United States | The devices were distributed to the following US states: AL, CA, CO, FL, IL, KY, MA, MI, MN, MS, ME, NJ, NY, OH, TN, TX, UT, and VA. | Open, Classified |  |  | 40 | Computer | Computer
Z-0703-2018 | 79107 | No | IXR | Radiologic table. | Radiology | Radiology | 510(K) Exempt | Foot Controls impacted for the following devices: |  |  | July 31, 2017 | 2017 | Oakworks Inc | The foot control used on the OAKWORKSÂż Inc., CFPM300, CFPM301, CFPM400, CFPM401, CFPMB301, CFUR301, CFUR401 and CFLU401 Imaging Tables can become lodged under the base or under the column of the table. | Device Design | OAKWORKS, Inc. issued a MEDICAL DEVICE ADVISORY NOTICE dated 8/18/2017 to customers providing more information in regards to the Medical Device Advisory Notice sent on May 12, 2017. The advisory notice was initiated in response to reported issues with the foot control. OakWorks, Inc. customer service will place the order for the appropriate Foot Control Kit listed below based on what table is impacted and the date of manufacture. PN2015-204 - Foot Control Kit - Standard Riser2015-209 - Foot Control Kit - Thinner Riser2015-210 - Foot Control Kit - Holder Only2015-211 - Foot Control Kit - Riser RemovalCustomers with questions may contact 1-800-558-8850 or 1-717-235-6807. |  | 1104 | 1,104 | USA (nationwide) Distribution | Open, Classified |  |  | 1104 | Computer | Not_Computer
Z-2838-2018 | 80735 | No | IYB | Diagnostic x-ray tube mount. | Radiology | Radiology | 510(K) Exempt | 3D V, model. no. |  |  | June 29, 2018 | 2018 | Siemens Medical Solutions USA, Inc | In rare cases of insufficient maintenance or high clinical workload, the first rope of the ceiling stand, which is designed to take the load, could break without triggering the safety lock and this could lead to overloading the safety rope which is not designed for continuous load and ongoing movement under load. This may lead to mechanical fatigue and cause the arm to drop down during patient positioning. Patients and users may be seriously injured. | Device Design | On 6/29/18, the firm emailed its four affected original equipment manufacturers (OEMs), alerting them to the potential malfunction of the ceiling stand's lifting column that could cause the arm to drop down during patient positioning.Manufacturers were asked to distribute an attached Safety Notice to affected customers. In addition, the firm urged manufacturers to order retrofit kits for all customers and to take any steps needed to ensure that customers install the kits. The kits are free of charge. Manufacturers were also asked to complete a response form and return to the firm.Affected customers were advised of the product problem and instructed to install retrofit kits to avoid risk to patients and to users.If you have any questions, please contact Mr. Wolfgang Bohrer at wolfgang.bohrer@siemens-healthineers.com. |  | 667 | 667 | US Nationwide Distribution to accounts in IL, FL, and NY. | Open, Classified |  |  | 1366 | Computer | Not_Computer
Z-0900-2018 | 79235 | No | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | BIO CAL(R) Temperature Controller, Models: (a) 370, (b) 370I |  |  | February 08, 2018 | 2018 | Medtronic Perfusion Systems | Medtronic has been unable to develop a cleaning protocol to satisfy current industry concerns and expectations. As a result, an updated cleaning protocol will not be developed by Medtronic and it has been determined that the best course of action is to request BIO CAL users to discontinue use and dispose of BIO CAL devices. | Device Design | Medtronic sent an Urgent Medical Device Recall dated February 2018 via UPS 2-Day.to their customers. The letter identified the affected product, problem and actions to be taken. Medtronic recommended that the affected device no longer be used for clinical procedures. Customers are asked to dispose of any devices in their possession per their normal equipment obsolescence procedures. Customer are also instructed to not resell any devices for clinical use. For questions contact your Medtronic Sales Representative. |  | 1980 | 1980 units | Worldwide Distribution | Open, Classified |  |  | 1980 | Computer | Not_Computer
Z-0717-2018 | 79091 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Allura Xper FD and and Xper FD ORT - Philips X-ray Systems |  |  | December 20, 2017 | 2017 | Philips Electronics North America Corporation | Extra monitors, so called 7th or 8th monitor, mounted on the Monitor Ceiling Suspension (MCS), might disengage from the MCS and fall down. The two small set screws securing the two gliding rods into an aluminum block ofthe support arm can come loose because of vibrations and forces exerted on them. If the set screws have come loose gravity can keep the rods in position. However, if a force is exerted on the monitor pushing it forward, the monitor cantopple forward. Once the monitor has toppled over it is only attached with its wires. If the monitor breaks loose of its wires it will drop on the floor. | Device Design | Philips sent an Urgent Field Safety Notice letter December 20, 2017 dated to all affected customers. The letter identified the affected product, problem and the actions to be taken. For question contact your local Philips representative. |  | 1486 | 1486 | Worldwide Distribution - US Nationwide | Open, Classified |  |  | 1486 | Computer | Not_Computer
Z-0910-2018 | 79130 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | cobas 8000 Modular Analyzer Series; Part Numbers: 1. 05964067001 cobas 8000 c 502 module |  |  | July 18, 2017 | 2017 | Roche Diagnostics Corporation | A warning is being added to the cobasÂż 6000 analyzer and the cobasÂż 8000 modular analyzer series Operators Manual to prevent operator injury during maintenance. | Device Design | Actions Taken by Roche Diagnostics: A warning will be added to the cobasÂż 6000 analyzer series Operators Manual and the cobasÂż 8000 modular analyzer series Operators Manual. Actions to be Taken by the Operator: Operators are advised not to overtighten the thumb screws when replacing the USM Cover 3 after maintenance. When removing the screws for maintenance, avoid touching the edge of the USM Cover 3. Enclosure: Fax back form 7727-00-0717Actions Required: " Do not overtighten the thumb screws when replacing the USM Cover 3 after maintenance. When removing the screws for maintenance, avoid touching the edge of the USM Cover 3. " Complete the attached fax form and fax it to 1-888-276-5901. " File or post this Urgent Medical Device Correction (UMDC) in an appropriate location within your laboratory, for future reference. |  | 245 | 245 | Distributed in all 50 states: AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MN, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY, the District of Columbia, and Puerto Rico. | Terminated | July 13, 2018 | 360 | 3678 | Computer | Not_Computer
Z-2498-2018 | 80402 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Allura Xper |  |  | June 11, 2018 | 2018 | Philips Healthcare | The first time an operator selects a new procedure type during a single examination,the shutter position resets to the open position for the new procedure type. If the shutters had previously been changed during the examination, that setting is not retained after the first time the procedure type is changed during a single examination. | Software design | An Urgent Medical Device Correction letter dated June 11, 2018 was sent to customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers, until a software revision that corrects this issue becomes available, users should verify that the desired shutter position is set when performing the run after the first time the procedure type is changed during a single examination. This can be accomplished by first selecting a different procedure and then reselect the original procedure on the Xper Module or on the Data Monitor.Customer shall ensure that all staff with access to the affected systems are informed of the contents of this Field Safety Notice.A copy of this Field Safety Notice shall be placed together with the documentation of the system until the system has been corrected by Philips.For questions contact your local Philips representative Technical Support Line 1-800-722-9377. |  | 459 | 459 affected systems | US Nationwide Distribution | Open, Classified |  |  | 459 | Computer | Not_Computer
Z-2528-2018 | 80078 | No | OMP | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Invia Motion Negative Pressure Wound Therapy System. Labeled as: |  |  | February 07, 2018 | 2018 | Medela Inc | Device may display a battery missing error. | Device Design | Medela sent an Notification letter dated February 1, 2018 to customers via Certified Mail. The letter identified the affected product, problem and actions to be taken. Instructions included to place all affected items on hold until an insert outlining additional charging instructions were included with the devices. Those instructions included to provide a printed copy of the charging instructions in each device package, review those insert instructions with patients and care-givers, and complete and return the confirmation letter. For questions contact Medela Partner Services at info-healthcare@medela.com or via phone 877-735-1626. |  | 778 | 778 units | US Nationwide Distribution domestically to AR, CA, FL, GA, IL, IN, LA, MI, MN, MO, NC, NJ, NY, OH, OK, PA, TX. Distributed internationally to Spain, Italy, Denmark, Sweden, Switzerland, Netherlands. | Terminated | December 04, 2018 | 300 | 778 | Computer | Computer
Z-0916-2018 | 79087 | No | NLH | Electrode recording catheter or electrode recording probe. | Cardiovascular | Cardiovascular | 510(k) | Stryker Sustainability Solutions Reprocessed BW Lasso 2515 NAV eco Variable Diagnostic EP Catheters, Ref D134301, Sterile, Rx. |  |  | January 02, 2018 | 2018 | Stryker Sustainability Solutions | An EEPROM chip error code may occur when the catheters are used with CARTO EP Navigation Systems. | Device Design | Stryker sent an Urgent Medical Device Recall letter dated January 2, 2018, explains the reason for recall, risk to health, and instructs them to discontinue use of the affected products. The customer is to check their inventory for the affected serial numbers provided in the attachment and complete the recall effectiveness check form to indicate the amount of product in inventory. Return instructions are provided. For further questions, please call (888) 763-8803. |  | 170 | 170 devices | USA (nationwide) Distribution was made to medical facilities in AZ, CA, CT, FL, GA, IL, KS, LA, MN, MO, OH, PA, TX, VT, WA, and Wi. Foreign distribution was made to Canada. | Terminated | May 02, 2018 | 120 | 178 | Computer | Not_Computer
Z-0724-2018 | 78663 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Integrity R1.2 |  |  | November 20, 2017 | 2017 | Elekta Limited | Potential for positional errors following automatic table movement (ATM). | Software Design Change | Customers were notified via letter on about [DATE]. Instructions include to post the notice in a place accessible to all users until the action is closed and to advise appropriate personnel working with the affected product of the recall action. Additionally Elekta plans to release Integrity R4.0.0, which should identify positional errors over 5 mm, in the first half of 2018. |  | 539 | 539 units total | Distributed domestically to AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY. Distributed internationally to ALBANIA, ALGERIA, ANTIGUA AND BARBUDA, ARGENTINA, AUSTRALIA, AUSTRIA, BAHRAIN, BANGLADESH, BELARUS, BELGIUM, BOSNIA AND HERZEGOVINIA, BOTSWANA, BRAZIL, BULGARIA, CANADA, CHILE, CHINA, COLOMBIA, CROATIA, CUBA, CYPRUS, CZECH REPUBLIC, DENMARK, ECUADOR, EGYPT, FINLAND, FRANCE, GEORGIA, GERMANY, GREECE, HONDURAS, HONG KONG, HUNGARY, INDIA, INDONESIA, IRAN, IRAQ, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KAZAKHSTAN, LAOS, LITHUANIA, MALAYSIA, MALTA, MEXICO, MOROCCO, MYANMAR, NAMIBIA, NEPAL, NETHERLANDS, NEW ZEALAND, NORWAY, PAKISTAN, PANAMA, PARAGUAY, PERU, PHILIPPINES, POLAND, PORTUGAL, ROMANIA, RUSSIA, SAUDI ARABIA, SERBIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SOUTH KOREA, SPAIN, SRI LANKASUDAN, SURINAME, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TUNISIA, TURKEY, UK, UKRAINE, VENEZUELA, VIETNAM, ZIMBABWE | Open, Classified |  |  | 539 | Computer | Computer
Z-2537-2018 | 80510 | No | JAI | Powered radiation therapy patient support assembly. | Radiology | Radiology | 510(k) | Patient Positioning System with KRC2 controller using software versions 2.3.1 2.3.10 |  |  | September 05, 2017 | 2017 | Forte Automation Systems Inc | Communications between the Patient Positioning System and the accuracy filter can periodically fail with no clear indication to the operator. | Software design | On September 5, 2017 Forte has distributed an Urgent Medical Device Correction letter with the following information, Subject: Robotic Treatment Table Positioning Error Commercial Name of Affected Product: Patient Positioning System Affected Version(s) / Lot(s): KRC2 Controllers with software version 2.3.1  2.3.10 Reference / FSCA Identifier: 3003900998-2017-00001Type of Action: Notification and CorrectionThis letter provides users with a Description of the Problem, Details, Recommended User Action and Forte Automation Systems, Inc. Actions. This document also Pending resolution with a permanent solution, Forte Automation Systems, Inc. recommends reimaging the patient prior totreatment, especially after any table movement. This includes movement initiated by the image position verification system. It also advises users that a solution for the issue is in development. |  | 13 | 13 | US in the state of California | Open, Classified |  |  | 13 | Computer | Computer
Z-1241-2018 | 79353 | No | PHX | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | AltiVate Reverse INSTRUMENTATION, HUMERAL SOCKET SHELL TRIAL, STANDARD, REF 804-06-152, Qty: 01, Encore Medical, LP Orthopedic surgical instrument. |  |  | January 30, 2018 | 2018 | Encore Medical, Lp | Broken or missing screws and retaining clips. This may be due to screw tolerance and material strength of the trial. | Device Design | The firm disseminated their field safety notices on 01/30/2018 by email. The letter identified the affected product, problem and actions to be taken. The notices requested the trails be returned for replacement. |  | 387 | 387 units | US Nationwide Distribution | Open, Classified |  |  | 387 | Computer | Not_Computer
Z-2541-2018 | 80403 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | PR4100 Microplate Reader US, Catalog Number 45591587 |  |  | April 02, 2018 | 2018 | Bio-Rad Laboratories Inc | The connector of the power supply was capable of fitting other components provided with the system. However, the voltage output of the reader power supply is incompatible with the other component and, if incorrectly connected, could result in irreparable damage to the system. | Device Design | An Urgent Product Correction notification letter dated 4/2/18 was distributed to customers. The letter instructed customers to check the system and confirm it is labeled as shown in the letter. Then complete the attached form, indicating if the system includes a D-Link Brand USB hub, and fax or email back the form to Bio-Rad as soon as possible. If needed, a Bio-Rad Technical Support Representative will contact you shortly to arrange for replacement of the hub. |  | 153 | 153 | US Nationwide Distribution inn the states of : AL, AR, AZ, CA, CO, CT, FL, GA, HI, IL, IN, KS, LA, MA, MD, MI, MN, MO, MS, NC, NJ, NY, OR, PA, SC, TN, TX, VA, VT, WA, WI, and Puerto Rico. | Open, Classified |  |  | 153 | Computer | Computer
Z-2960-2018 | 80727 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Overhead Counterpoise Systems (OCS), OCS CONF,STD CEIL MNT, STD HORZ ARM, Catalog # OCS 115A, Material Number 59880253 |  |  | June 28, 2018 | 2018 | Bayer Medical Care, Inc. | The torque wrench used on the ceiling column assembly during a recent service visit was outside of the torque specification, potentially overtightening the bolts. The overtightening of bolts may compromise the structural integrity of the system which could lead to the OCS disengaging from the ceiling mount and this may cause injury to the patient or the operator. | Equipment maintenance | On June 28, 2018, Urgent Medical Device Field Safety Corrective Action letters were distributed to customers. The letters informed the customers of the following:A Bayer Service Representative will be scheduling a service visit to replace the ceiling column assembly of your OCS system(s) at no cost to you. |  | 2 | 2 | The devices were distributed to the following US states: OH, TX, and WV. | Terminated | November 26, 2018 | 151 | 7 | Computer | Not_Computer
Z-2964-2018 | 80642 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | DRAD-3000E (Radrex-i) TFP-4336W (Wireless FPD) |  |  | April 18, 2018 | 2018 | Canon Medical System, USA, INC. | It was found when an operator performs a radiography using the wireless flat panel detector (FPD), a message window was displayed on the monitor stating image transmission was not completed and there was no image. It also showed the OK button to reacquire image data form the FPD, and the Cancel button to cancel the reacquisition.When the operator selects the OK button, the same message window appears. The operator then repeated the same operation several times and finally selected the Cancel button to quit the reacquisition mode. | Software design | On 04/18/2018, the firm mailed Urgent: Medical Device Correction letters via USPS return receipt mail. The letter contains a return reply form that is to be faxed. After one month contact will be made with each non-responding location and a copy of the letter will be faxed to that location and the fax receipt will serve as proof of notification.Upon release of the modified software it will be assigned to a firm employee with instructions for implementation at each site. Each incident is assigned unique service incident (case) and the completion of the service incidents are tracked through completion. In the event that a customer refuses the installation the customer would be requested to sign a service case record stating that the corrected software was refused. |  | 43 | 43 | Domestic: AR, FL, GA, PA, IL, LA, MI, NC, NJ, NY, OH, TN, TX, UT, VT, WI, and WVForeign: Australia, Canada, The Netherlands, South Korea, and Malaysia | Open, Classified |  |  | 43 | Computer | Computer
Z-2068-2018 | 79813 | No | JFH | Acid phosphatase (total or prostatic) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ST-AIA PACK PAP; Part Number: 025262 |  |  | March 05, 2018 | 2018 | Tosoh Bioscience Inc | Asfotase Alfa (Strensiq) interferes with certain Tosoh assays which potentially causes falsely increased or falsely decreased immunoassay test results, depending upon the assay. | Device Design | The firm TOSOH Bioscience Inc., sent an "URGENT: MEDICAL DEVICE RECALL" letter dated March 23, 2018 to affected customers. The letter described the product, problem and action to be taken. The customers were instructed to do the following: *Based on our revised labeling, the assays identified in Table 1 should not be used for patients who are being treated with Asfotase Alfa." For patients that are being treated with Asfotase Alfa, use an alternative test method that does not utilize alkaline phosphatase technology." Inform all medical professionals that assays utilizing alkaline phosphatase-based technologies must not be used for patients receiving Asfotase Alfa treatment. Continue to use the assays listed in Table 1 for any patients who are not receiving Asfotase Alfa treatment." Complete and return the attached Acknowledgement Form to Tosoh Bioscience, Inc., within 15-days of receiving this notification via Fax: 1-650-636-8651, Email: biorecallresponse@tosoh.com or Regular mail: 3600 Gantz Rd., Grove City, OH 43123." Maintain this notification with your laboratory records and forward this information to others who may have received this product." If you have obtained unexpected test results or received any complaints of illness or adverse events associated with the use of the products identified in the table above, contact Tosoh Technical Support 24 hours a day, seven days a week at (800) 248-6764.Should you have any questions regarding this medical device recall, please feel free to call (800) 248-6764 or email: bernadette.oconnell@tosoh.com. Monday - Friday, from 9:00 AM to 5:00 PM (PST). |  | 1083 | 1083 | Worldwide Distribution: US (Nationwide) and countries of; Canada, Mexico, Columbia, Chile, Uruguay, Nicaragua, Argentina, Peru, Chile, Honduras. Equador, Caymen Islands, Panama, and Costa Rica. | Open, Classified |  |  | 300971 | Computer | Not_Computer
Z-3260-2018 | 81006 | Yes | LWS |  |  | Cardiovascular | PMA | Model 3300 LATITUDE(TM) Programming System |  |  | September 10, 2018 | 2018 | Boston Scientific Corporation | There is a potential for the Pacing System Analyzer (PSA) to exhibit unintended cross-chamber stimulation. | Device Design | The firm initiated the recall on 09/10/2018 to the US sales representatives by email. The sales representatives were asked to hand deliver the notice to the hospitals when the programmer is retrieved for that location. The sales representatives will electronically confirm the hand delivery of the letter to each affected customer after it occurs. |  | 74 | 74 units | Worldwide distribution - US Nationwide in the State of FL, UT, and VA, and countries of Austria, Czech Republic, Germany, Greece, Netherlands, Norway, Portugal, Spain, Sweden, United Kingdom, Switzerland | Open, Classified |  |  | 74 | Computer | Not_Computer
Z-0562-2018 | 79008 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Ingenuity TF PET/CT (model 882442) running software version 4.0.2 |  |  | December 11, 2017 | 2017 | Philips Medical Systems (Cleveland) Inc | A software issue causes PET reconstructions to fail intermittently. It was determined that reconstructions fail due to a negative table position (-1 value is inserted) in the raw data list file, rather than the actual table position.This error has been found to occur in two scenarios:1) When the system operator cancels an acquisitiona. The error will occur every time a scan is cancelled by theoperator.2) Couch position requests within the software sequence were delayeda. The error occurs intermittently, but has been found to occur more frequently when the gantry s network is heavily loaded with multiple retrospective reconstructions running in parallel.b. The error manifests to the technologist by an error message and Failed status on the Reconstruction Monitor and the error message Result {0} failed to reconstruct" on the Acquisition Workflow window during reconstruction after the patient scan has been completed.In both scenarios, the acquisition data will not be able to be reconstructed and will therefore be unusable. | Software design | Philips Healthcare sent an Urgent-Field Safety Notice dated December 11, 2017. U.S. Customers will be receive the letter via USPS Certified mail. Non-U.S.. customers will receive the letter through Philips Global Markets following local regulations. Philips is informing customers of the software issues on the systems, as well as available workarounds to prevent these issues from occurring or to minimize their impact.For North America and Canada, contact the Customer Care Solutions Center (1-800-722-9377, follow the prompts) |  | 24 | 24 | Worldwide Distribution - US Distribution to the states of : CA, OH, IL, PA and NY. | Open, Classified |  |  | 24 | Computer | Computer
Z-2176-2018 | 80094 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance 16 P with DEPMED HARDENING KIT 12NC: 453567400741, Model Number 728246. Computed Tomography X-ray system |  |  | March 30, 2018 | 2018 | Philips Medical Systems (Cleveland) Inc | The patient support head holder could contact the back ISO shelter wall during manual or motorized motion through the gantry bore. This may result in a serious injury that would require medical intervention to preclude permanent impairment in certain scenarios. | Device Design | Urgent Field Safety Notices dated 3/28/18 were distributed to customers. The letters instructed customers to perform the following actions:Please review the following information with all members of your staff who need to be aware of the contents of this communication. It is important to understand the implications of this communication.Please retain a copy with the equipment Instruction for Use.As outlined in the Instructions for Use, always monitor the patient during allmovements (manual or motorized) of the patient support. If you recognize an impending collision with the back shelter wall, halt the motion or activate the system E-stop.Mandatory Field Change Order (FCO) 72800686 which will add additionalwarning labels specific to this issue to the system. The FCO will be implemented free of charge. |  | 20 | 20 | Worldwide distribution including US states of CA, MD, NV, OH, TX, and WI, Afghanistan, Canada, Germany, Iraq, Italy, Korea, Kuwait, and Lebanon. | Open, Classified |  |  | 25 | Computer | Not_Computer
Z-3252-2018 | 80918 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Biomet Sports Medii cine Juggerknot Long Flex Drill Bit with Sleeve Nitinol intended to be used for soft tissue to bone fixation with indications for use in: Shoulder, Foot and Ankle, Knee, Hand and Wrist and Hip repair.Stainless Steel, Sterile Item Number: 110016992 |  |  | August 03, 2018 | 2018 | Zimmer Biomet, Inc. | Expiration date incorrectly listed on the label | Device Design | Zimmer notified distributors on 8/3/18 distributors notified via email. Hospital risk managers, as well as distributors with product, notified via Fed'X.The letter identifies the issue and their responsibilities, locating and removing the product in their territory, as well as identifying hospitals who have fielded inventory. Product to be returned to Zimmer Biomet.Question customer service at 574-371-3071 between 8:00 am and 5:00pm EST, Monday through Friday or to :CorporateQuality.PostMarket@zimmerbiomet.com |  | 217 | 217 units | AZCOFLFLGALAMAMAMONYSCTX | Open, Classified |  |  | 217 | Computer | Not_Computer
Z-0559-2018 | 79075 | No | NEY | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Covidien Emprint Percutaneous Antenna with Thermosphere Technology, 20cm (reinforced), Material CA20L2 |  |  | August 07, 2017 | 2017 | Covidien LLC | The firm received consumer reports of the ceramic trocar tip of the Emprint ablation antenna disengaging from the needle shaft post-ablation. This can result in the tip remaining in the patient. | Device Design | All consignees were notified via Federal Express or certified mail on August 7, 2017, and the letter informs customers of the possible trocar tip disengagement and the actions they should take. Customers are requested to segregate and return affected product and acknowledge receipt and understanding of the Urgent Medical Device Recall Notice. |  |  |  | Nationwide including PR, Canada, China | Open, Classified |  |  |  | Computer | Not_Computer
Z-2184-2018 | 80137 | No | JOY | Automated cell-locating device. | Hematology | Hematology | 510(k) | CellaVision DM Software versions 6.0.1 or 6.0.2 installed on the following products: CellaVision DM96, DM1200, DM9600 and DI-60 |  |  | December 04, 2017 | 2017 | Cellavision AB | A software malfunction was found where WBC, RBC and PLT comments added after a slide is signed, are not sent to the LIS. This can only occur where customers process multiple slides per blood sample. | Software design | CellaVision sent an Urgent Medical Device Recall letter dated December 4, 2017. The letter identified affected product, problem and the actions to be taken. Customers are requested to determine if they are affected or not. If affected they are requested to contact their local distributor for installing a patch.Customers are asked to confirm receipt of the notice by signing and dating theattached Acknowledgement Form and emailing a copy to their local distributor.For technical questions contact CelleVision (service@cellavision.se) |  | 22176 | Instruments: 224 Software: 99 (US) | US Nationwide in the states of FL, IL, NY, and NC | Open, Classified |  |  | 22176 | Computer | Not_Computer
Z-2060-2018 | 80067 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Sonialvision Safire II, Model #: DAR-8000f |  |  | January 12, 2018 | 2018 | Shimadzu Medical Systems | When selecting serial radiography with a pulse rate of 7.5fps (including selecting a preset or changing thepulse rate using a temporary edit function) it was observed the indicated "irradiation time" on the X-rayGenerator Console, as well as the "integral dose" on the external console of fluoroscopy, were reset to 0. There has been one report of this event. | Software design | Shimadf0u sent an Urgent Voluntary Medical Device Recall Notice letter dated January 2018. The letter identified the affected product, problem and actions to be taken. For questions contact Shimadzu representative or contact Shimadzu National Technical Support at 800-228-1429. |  | 17 | 17 | US Nationwide in the states of IL, WA, MI, TX, AZ, CA, LA, FL, SC, and NJ. | Open, Classified |  |  | 17 | Computer | Computer
Z-1075-2018 | 79386 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Iridex TruFocus LIO Premiere" Laser Indirect Ophthalmoscope ( LIO ) |  |  | February 27, 2018 | 2018 | Iridex Corporation | It was discovered that the laser energy was not going where intended to put it, resulting in focal cataracts and iris burns. | Device Design | Action Taken by IRIDEX: 1. A investigation was started immediately upon receipt of the returned devices from the Sloan Kettering complaint on November 3, 2017. CAPA 0125 was initiated to document the events, root cause, correction, corrective and preventive action and effectiveness of these actions.2. A ship hold was initiated on November 15, 2017.3.As described below, IRIDEX plans to notify customers and distributors that it is voluntarily recalling all models of the TruFocus LIO Premiere.4. Public WarningA press release was prepared and sent out on February 23, 2018.For further questions please call (650) 962-8848 Ext. 3672 |  | 104 | US - 58 OUS - 46 | Worldwide Distribution - USA Distribution and to the countries of : China, Germany, Australia, Canada, Ireland, Lebanon, Japan, Israel and Qatar | Open, Classified |  |  | 104 | Computer | Not_Computer
Z-1076-2018 | 78655 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | GENICON EZEE Retrieval, sterile, Rx only. |  |  | July 19, 2017 | 2017 | Genicon, Inc. | Plastic handle is breaking off of shaft inhibiting proper removal per Instructions for Use. | Device Design | Consignees were notified via email on 07/19/2017. Instructions included |  | 2967 | 2,967 units | Distributed domestically to FL, IL, MI, TN, UT. Distributed internationally to Dominican Republic, Ireland, Denmark, Spain, Switzerland, United Kingdom. | Open, Classified |  |  | 2967 | Computer | Not_Computer
Z-3248-2018 | 81045 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | ddR Formula B X-ray System, ddR Formula B |  |  | August 09, 2016 | 2016 | Swissray | Possible injury due to movement of the arm, calibration loss and communication loss. | Device Design | November 2016, Swissray Medical issued official communication to customers as a Action Notification Report. The manufacturer, Swissray Medical AG, distributed the Work Instructions for the FCO and the ActionNotification Report Form to all distributors and service entities who would be performing corrective actions in the field. The Swissray Field Engineer then presented the Report to the customer (end user facility) and communicated with the customer in-person while performing the corrective action on-site. When corrective actions were completed, both the customer and service personnel then sign the Report to confirm that the customer was informed about the corrective action and that corrective action was completed. |  | 26 | 26 units | US Nationwide | Open, Classified |  |  | 26 | Computer | Not_Computer
Z-1069-2018 | 79292 | No | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | STA - NeoplastineÂż Cl Plus Âż (ref. 00606) |  |  | October 25, 2017 | 2017 | Diagnostica Stago, Inc. | Confirmed defect of homogeneity where some reagent vials will give prolonged Prothrombin Time (decreased PT%). | Device Design | Stago sent an Urgent Field Safety Notice letter dated October 25, 2017 to their customers. The letter identified the affected product, problem and actions to be taken: Customers were instructed to:If it is not already done, run a Quality Control test at every change of vial (if you have affected lots of product).Return to company, by fax or email, an acknowledgement form confirming that the recall letter was read by the receiving company.For questions contact the Stago Hotline (1-800-725-0607). |  |  |  | Nationally | Open, Classified |  |  |  | Computer | Not_Computer
Z-2255-2018 | 80062 | No | MUZ |  |  | Neurology | PMA | VNS Therapy Programmer, Model 3000, v1.0 System |  |  | November 13, 2017 | 2017 | Cyberonics, Inc | Certain Model 3000 programming events can result in miscalculation of parameters stored in the Models 103, 104, 105, and 106 generators. During these programming events, the miscalculations can lead to: " Delivery of more stimulation than intended, resulting in painful stimulation or other common side effects (Model 106 only); " No stimulation in the case of device disablement (Burst Watchdog Timeout), resulting in no therapy to the patient (Model 106 only);" Delivery of less stimulation than intended, resulting in therapeutic settings not being achieved within device specification (Models 103, 104, 105, or 106); and/or " Delays or absence of the 75% and 50% battery life indicators displayed by the programming software (Models 103, 104, 105, or 106). | Device Design | On November 13, 2017 LivaNva sent letters to all the doctor and surgeons that use this device. "Dear Doctor:This notification describes an issue identified with the VNS Therapy Âż Programmer, Model 3000 version 1.0 which may impact patients implanted with Models 103, 104, 105 or 106 VNS Therapy generators. This issue does not affect Model 1000 SenTiva or Model 102/102R generators.What is the Issue?Certain programming events performed with a Model 3000 Programmer on a Model 103, 104, 105 or 106 VNS Therapy generator can result in a miscalculation of some parameters stored in the generator, which could lead to:" Output current delivered being different than that programmed;" Device disablement; and/or" Errors in the 75% and 50% battery life indicators.The actual battery life and battery function are not affected by this issue.Risk to HealthThis issue presents the following risks:" Delivery of more stimulation than intended, resulting in painful stimulation or other common side effects 1 (Model 106 only);" No stimulation in the case of device disablement, resulting in no therapy to the patient (Model 106 only);" Delivery of less stimulation than intended, resulting in therapeutic settings not being achieved (Models 103, 104, 105 or 106) ; or" Sudden decrease from 100% to 25% battery life indicator, though the actual battery life and battery function are not affected (Models 103, 104, 105).Which Patients are Potentially Impacted?Any patient with a Model 103, 104, 105 or 106 generator that has been programmed using a Model 3000 programmer may be affected by this issue. Your LivaNova representative will collect data from your Model 3000 programmer to identify which patients are impacted by this notice.Actions to be taken by the Physician1. Stop using the Model 3000 programmer to program Model 103, 104, 105 or 106 generators. You shall continue using the Model 3000 programmer with Model 1000 SenTiva generators and may continue to use it wi |  | 252 | 252 | AL, CA, CO, DE, FL, GA, ID, IL, IN, ME, MO, MS, NC, NJ, NY, PA, TN, TN, TX, UT, WA and WI | Open, Classified |  |  | 252 | Computer | Computer
Z-0595-2018 | 79007 | No | NSU | Instrumentation for clinical multiplex test systems. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Hologic Panther(R) Fusion(TM), Catalog number ASY-09600 |  |  | January 20, 2018 | 2018 | Hologic, Inc | The Panther or the Panther Fusion glass panel separates from the 19 inch touch screen monitor and could potentially fall and hit an operator. | Device Design | The firm sent an Urgent Medical Device Recall Notification dated January 8, 2018 to their customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed that a Hologic representative will contact your site to arrange a visit in order to provide an interim solution to secure the glass panel until a replacement monitor with the issue corrected can be installed. For questions contact Hologic Technical Support. at +1-888-484-4747 or +1-858-410-8511 or by e-mail at molecularsupport@hologic.com |  | 103 | 103 units | Worldwide distribution | Open, Classified |  |  | 103 | Computer | Not_Computer
Z-3213-2018 | 80922 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | INTEGRIS Allura 9, System Code 722018 |  |  | July 16, 2018 | 2018 | Philips Electronics North America Corporation | There is a potential for the Monitor Ceiling Suspension, which holds a large screen monitor, may detach from the actuator rotor shaft causing the monitor to fall from the ground. When a Monitor Ceiling Suspension detaches from the actuator rotor shaft and the monitor falls, there is a risk ofinjury for the patient, user and bystander. | Device Design | The firm, BU IGT System, sent an "Urgent -Medical Device Correction" notification letter dated 07/16/2018 to customers. The letter describes the product, problem and actions to be taken. The letter instructs customers to do the following: In order to reduce the risk for patients, users and bystanders if this problem would reoccur, we recommend the following actions until the correction has been implemented.-Avoid unnecessary movements of the Monitor Ceiling Suspension.-For those movements that are necessary, avoid that the user, patient or bystander are in close proximity to the monitor.-When moving the Monitor Ceiling Suspension, ensure that no body partsof the staff or patient are underneath the monitor.-Do not move the monitor above the patient.Please ensure that all staff with access to the affected systems are informed of the content of this Medical Device Correction.All affected products will be corrected by means of a Field Change Order (FCO) free of charge. This FCO will be available mid-August.You will be contacted by our local Philips representative to schedule this corrective action.If you need any further information or support concerning this issue, please contact your local Philips representative: Technical Support Line: 1-800-722- 9377, 8am to 8pm (EST) Monday thru Friday. |  | 5115 | 5115 total | Worldwide Distribution: US (nationwide) to states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY; and countries of: Algeria, Argentina, Australia, Austria, Bangladesh, Belarus, Belgium, Bosnia and Herzegovina, British Virgin Islands, Brazil, Brunei Darussalam, Blugaria, Cambodia, Canada, Chile, China, Columbia, Croatia, Cyprus, Czech Republic, Denmark, Dominic Republic, Egypt, Estonia, Finland, France, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Korea South, Republic of Korea, Kuwait, Latvia, Lebanon, Libya, Lithuania, Macedonia, Malaysia, Malta, Martinique, Mauritius, Mexico, Moldova, Morocco, Netherlands, New Zealand, Norway, Oman, Pakistan, Palestine, Peru, Philippines, Poland, Portugal, Puerto Rico, Romania, Reunion, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, Spain, Sri Lanka, Sudan, Sweden, Switzerland, Syrian Arab Republic, Taiwan, Thailand, Tunisia, Turkey, Ukraine, UAE, UK, Uzbekistan, Venezuela, Vietnam, and Yemen. | Open, Classified |  |  | 5115 | Computer | Not_Computer
Z-2424-2018 | 80298 | No | LHN | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Proteus 235, Version PTS-11.0.1.2 and PTS-11.0.2 |  |  | May 16, 2018 | 2018 | Ion Beam Applications S.A. | Correction vector confirmation message is lost if access point is changed after sending the correction vector. As a result, the patient will be treated in the setup position or treatment position instead of the corrected position. | Software design | On May 18, 2018 IBA issued URGENT FIELD SAFETY NOTICES to their customers. Customers were advised to take the following actions:It is recommended to ensure that the correction vector confirmation message appears on the expected access point. When the correction vector is applied (if deemed necessary1), ensure that the Patient Positioning System GoTo the corrected position for treatment before starting irradiation:- verify on adaPTdeliver screen that the corrected target position is selected before engaging the movement. - verify on adaPTdeliver screen that the corrected target position appears on the screen after the movement.Moreover, as advised in the Proton Therapy System Clinical User Guide, it is recommended to verify the patient position after each correction vector application.A software update will be deployed on all affected sites by July 2018.Customers with questions may email: vigiliance@iba-group.com |  | 3 | 3 units | Worldwide Distribution - US Nationwide in the states of FL and the countries of Sweden, The Netherlands | Terminated | October 16, 2018 | 153 | 3 | Computer | Computer
Z-0592-2018 | 78910 | No | LSL | Neisseria spp. direct serological test reagents. | Microbiology | Microbiology | 510(k) | Xpert CT/NG Catalog GXCT/NG-120 Batch 1000049568, Catalog GXCT/NG-10 Batch 1000053075, Catalog GXCT/NG-10 Batch 1000056156 and Catalog GXCT/NG-10 Batch 1000056157. |  |  | October 09, 2017 | 2017 | Cepheid | Cepheid has received reports that some of our customers have experienced a higher number oftemperature drift errors than would typically be expected when using this test and its intended usespecimen types. Investigation of the issue also revealed an occasional probability of falsenegative results which remains consistent with package insert claims. | Equipment maintenance | An Urgent Medical Device Recall letter (response form included), dated October 9, 2017, was sent to affected consignees. The letter identified the affected product, the issue involved, impact, and actions to be taken. Cepheid will provide replacement product for Xpert CT/NG assay product identified in this recall. Customers are to complete and the return the Response Form provided in order to receive replacement product. Additionally, the letter instructs users to review their laboratory statistics to see if their positivity rates are trending low during use of the affected lots. If so, they are to review the negative patient results according to their laboratory's Quality Management System. Customers are to refer to the applicable contact information provided if there are any questions. |  | 3586 | 3,586 kits (49,400 units) | Worldwide Distribution -- USA, Puerto Rico, and Haiti. | Terminated | February 13, 2018 | 127 | 3586 | Computer | Not_Computer
Z-0565-2018 | 78622 | No | EXS | Urine flow or volume measuring system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | BARD(R) CRITICORE(R) MONITOR, Model: 000002N Series |  |  | November 10, 2017 | 2017 | C.R. Bard, Inc. | BMD has identified that some lots of CritiCore(R) monitors may experience failures involving urine output volume measurement fluctuations, core bladder temperature measurement fluctuations and alarm errors. | Software design | Bard Medical sent an Urgent Medical Device Product Recall letter dated November 10, 2017. Action required:"Immediately examine your inventory and quarantine product subject to this recall. Please refer to CritiCore Monitor Affected Serial Numbers (Attachment 2) to help you locate the affected product serial number. Do not use or further distribute any affected product."Please complete and return the accompanying Recall &amp; Effectiveness Check Form (Attachment 1) attached to this letter regardless of whether you have any of the affected product. Additional instructions for product return are contained on the form."If you have further distributed any units with the referenced serial numbers, please identify your customers and notify them at once of this product recall: Your notification should include a copy of this letter and the accompanying enclosures.This recall should be carried out to the user level." For further questions, please call (770) 784-6120 |  | 343 | 343 units | US Distribution and Internationally to Japan | Open, Classified |  |  | 343 | Computer | Not_Computer
Z-0194-2018 | 78315 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | MONTERIS MEDICAL NEUROBLATE SYSTEM, used for MRI-guided neurosurgical ablation. |  |  | October 05, 2017 | 2017 | Monteris Medical Corp | The firm received complaints that involved an unanticipated interaction between a NeuroBlate system and 1.5 Tesla (1.ST) MRI systems. | Device Design | The firm issued advisory notice 10/05/2017. The notice discussed complaints received that involved an unanticipated interaction between a NeuroBlate system and 1.5 Tesla (1.5T) MRI systems. The notice provided added recommendations for all NeuroBlate users to mitigate the anomaly. |  | 52 | 52 systems | Nationwide and Canada | Open, Classified |  |  | 52 | Computer | Not_Computer
Z-0943-2018 | 79229 | No | LHN | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Proteus 235, Pencil Beam Scanning, PTS-9.4.X, PTS-10.0.1.1 |  |  | November 01, 2017 | 2017 | Ion Beam Applications S.A. | The internal configuration of the electrometers and the Real-Timecontrol boards of the Dekimo Scanning Controller is not checked beforeeach patient treatment. | Device Design | The firm, IBA, sent an "Urgent Field Safety Notice" dated 10/31/2017 to customers on 10/31/2017. The notice described the product, problem and actions to be taken. stating the following: There is no specific recommendation for the user. IBA will provide an Internal User notice to the IBA site team, requesting the operator to select the Beam Delivery Point CGTR, by clicking on the BDP button, before each patient treatment preparation (see Annex 1 fordetails). This will begin the check of the internal configuration of allthe elements of the Dekimo Scanning Controller. An updated system will be deployed on February 28, 2017.If you have any questions, contact Customer Complaints &amp; Vigilance Director at Vigilance@iba-group.com or +32 10 203 787 or helpdesk at +32 2 507 20 81 (available 24/7). |  | 3 | 3 | Worldwide Distribution: US (nationwide) to: MI only and country of: France. | Open, Classified |  |  | 3 | Computer | Not_Computer
Z-0004-2019 | 81010 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | TrueBeam Radiotherapy Delivery System version 2.0 and 2.5 |  |  | August 20, 2018 | 2018 | Varian Medical Systems | Reports have been received of an anomaly that can result in a treatment without intended gating (respiratory tracking/monitoring). This issue occurs when a patient planned with gating is treated on more than one system. | Software design | On 08/20/2018 the firm sent an URGENT MEDICAL DEVICE CORRECTION FIELD SAFETY NOTICE to customer via courier service (US and Canada via FedEx on 8/23/2018, and Rest of World via third party courier on 8/24/2018) informing them the firm had received reports of an anomaly that can result in a treatment without intended gating, and that this issue occurs when a patient planned with gating is treated on more than one system. Customers were advised, for treatment plans with intended gating, the firm recommends users monitor and pay attention to the system status message. When a patient with a gated treatment plan is moved to a second system, users must verify that intending gating was not overridden. If gating is overridden, users should reselect it from the Tools menu. The customer were informed that they will be contacted by a Customer Service Representative to schedule upgrade to the correction when available and if they have any questions regarding selecting gating from the Tools menu to contact customer support.Help Desk Contact Information:Phone: +1.888.827.4265Email: support @varian.comIn addition, the URGENT MEDICAL DEVICE CORRECTION FIELD SAFETY NOTICE, informs the customers to:1) Please retain a copy of this document along with your most current product labeling.2) Advise the appropriate personnel working in your radiotherapy department of the content of this letter.3) For future reference, this document will be posted to the Varian Medical Systems customer support website: http://www.MyVarian.com.and In order to satisfy regulatory requirements, we request that you complete the attached Recall Return Response form provided. Kindly return the completed form to returnresponse@varian.com. |  | 387 | 387 | Worldwide Distribution - US nationwide in the states of AL, AR, AZ, CA, CO, FL, Guan, IL IN, KS, KY, LA, MA, MD, MI, MO, NH, NJ, NY, OH, OK, PA, Puerto Rico, TN, TX, VA, WI and the countries of :Algeria, Australia, Azerbaijan, Belgium, Brazil, Bulgaria, Canada, China, Croatia, Czech Republic, Denmark, Egypt, Estonia, Finland, France, Georgia, Germany, India, Ireland, Israel, Italy, Japan, Kazakhstan, Lebanon, Morocco, Nepal, The Netherlands, New Zealand, Norway, Panama, Poland, Reunion Russia, Saudi Arabia, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey and United Kingdom | Open, Classified |  |  | 391 | Computer | Not_Computer
Z-0518-2018 | 78863 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Philips Brilliance 64 computed tomography x-ray system |  |  | October 06, 2017 | 2017 | Philips Medical Systems (Cleveland) Inc | Numerous issues related to software Brilliance iCT 4.1.6 software version. | Software design | Philips Medical Systems (Cleveland), Inc. sent the Customer Information Letter via certified letters mail with returned receipts. Also, a Field Service Engineer will perform the software update on the system to resolve the issues. Implementation of the software update will be monitored. |  | 1801 | 1801 | 26 US Govt accounts | Open, Classified |  |  | 1801 | Computer | Computer
Z-1137-2018 | 79547 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | MyLab Ultrasound System; Models 6100, 6150, 6200, 6250; PA023E Probe 7.5/10 |  |  | January 03, 2018 | 2018 | ESAOTE S.P.A. | The ultrasound system used in certain settings may result in overheating of the ultrasound probe head. | Software design | The January 3, 2018, notification letter requests the consignee complete and return a recall response form. The letter states that Esaote customer service will contact the consignee about the software update. In the meantime, it is recommended that the consignee does not use the PA023E probe in Pulsed Wave (PW) Doppler mode and that they refrain from using the system in Elaxto mode where the patient is unconscious or sedated. The consignee is asked to check the temperature of the probe before use. |  | 1 | 1 | US distribution in the states: AZ, FL, OK, PA, and TX. | Open, Classified |  |  | 6 | Computer | Not_Computer
Z-0524-2018 | 78677 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo.plaza systems with SW VB20A, Model Number - 10863171, 10863172, 10863173 |  |  | August 23, 2017 | 2017 | Siemens Medical Solutions USA, Inc | When a prior study is being replaced in the workflow step, in certain scenarios (based on the Display Protocol configuration) the prior study will only be replaced in the active Workflow Step and not show in all other workflow steps. The other workflow steps will continue to show the initially loaded study. | Software design | Siemens Healthineers distributed Urgent Customer Advisory Notices to their customers via courier service. Customers were advised to always verify date and time of the displayed study during workflow step change to ensure that the desired images are being displayed. Otherwise, please replace the displayed initial study in the workflow step with the expected study. Customers with any questions, call Siemens Customer Care Centre at 888-826-9702. |  | 76 | 76 | US Nationwide Distribution | Terminated | October 01, 2018 | 404 | 76 | Computer | Computer
Z-0374-2019 | 81418 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Centricity PACS-IW with Universal Viewer version 5.0.x with PACS-IW foundation displays medical images (including mammograms) and data from various imaging sources. |  |  | October 11, 2018 | 2018 | GE Healthcare, LLC | Synchronization failure | Software design | On October 12, 2018 GE Healthcare send recall notification to all their consignees with the following instructions;If you do not use a Diagnostic viewer (e.g. Centricity Universal Viewer Zero Footprint Client or any 3rd party dicom viewer) connected directly to the same Centricity EA to display images for diagnostic purposes, your system is not affected by the safety issue.If you do use a Diagnostic viewer (e.g. Centricity Universal Viewer Zero Footprint Client or any 3rd party dicom viewer) connected directly to the same Centricity EA to display images for diagnostic purposes, please contact GE Healthcare service representative immediately. GE Healthcare will assist you in determining if your system is affected and, if necessary, will assist you with a solution to confirm that study updates are made in the Enterprise Archive, until GE Healthcare can install a correction on your system.GE Healthcare will provide a correction that will be installed by a GE service engineer. A GE Healthcare representative will contact you to arrange for the correction. This activity will be performed at no cost to you.Contact InformationIf you have any questions or concerns regarding this notification, please contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 13 | 13 | CA, CO, DE, FL, HI, NY, OK, PA, SD, TX, WA, WI, and WVBelgium, Brazil, Canada, Colombia, France, Germany, Ghana, Ireland, Italy, Mexico, Netherlands, Panama, Poland, Saudi Arabia, South Africa, Spain, Switzerland, and Turkey. | Open, Classified |  |  | 86 | Computer | Computer
Z-2103-2018 | 78848 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | fMRI Hardware System - AudioSystem, Version 1.0 |  |  | October 28, 2017 | 2017 | NordicNeuroLab AS | Report was of an interrupted MR scan after patient complained about acoustic distortion, and heat felt form the headphones. | Device Design | The firm notified the consignees by email on September 28, 2017. The units were to be returned. For further questions, please call 1 (855) 665-3674 option 1. |  | 64 | 64 units | Worldwide Distribution - USA (nationwide) and to the countries of : Russia, Israel. Norway, South Korea, China, France, South Africa, Japan, Poland, Chile, Switzerland. Australia, Belgium, Lithuania, United Kingdom and Bulgaria | Open, Classified |  |  | 64 | Computer | Computer
Z-2101-2018 | 80159 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Centricity PACS-IW with Universal Viewer |  |  | January 25, 2018 | 2018 | GE Healthcare | Potential that one or more image series may be missing from an exam without a user warning displayed in the Viewer. | Software design | On January 28, 2019 GE healthcare notified their customers and gave the following instructions for the Centricity Universal Viewer with PACS-IW foundation 6.0.x; GTIN 00840682103800 and Centricity PACS-IW with Universal Viewer version 5.0.x with PACS-IW foundation.; Until a product modification is available to correct this issue, the following actions are recommended.1. Utilize the image count within the QC process to alert the user of a discrepancy in the number of transmitted images from the modality to the number of images available in the Universal Viewer.a. If a discrepancy is identified, attempt to retransmit the exam to PACS.b. If retransmission is unsuccessful, a GE Healthcare Service representative will need to be contacted to assist with resolution of the impacted exam. Urgent cases impacted by this issue should be interpreted at the modality.2. If DICOM storage commit is configured and in use, no commit notification is sent to the modality for images that were impacted by this issue.3. Attention should be given to prior cases, acquired before the application of the safety instructions given in this communication as they may be affected by the safety issue.For historical data inquiries, please contact GE service for assistance in identifying any affected images.GE Healthcare will correct all affected products at no cost to you. A GE Healthcare representative will contact you to arrange for the correction. For questions please contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 44 | 44 | Worldwide Distribution - US Nationwide in the states of AK, AZ, CA, CO, DE, FL, GA, HI, IA, ID, IL, LA, MA, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OK, OR. PA. SD, TN, TX, UT, WA, WI, WV, and WY and the countries of Austria, Belgium, Brazil, Canada, Colombia, France, Germany, Ghana, India, Ireland, Israel, Italy, Kuwait, Malaysia, Mexico, Netherlands, Panama, Poland, Russia, Saudi Arabia, South Africa, Spain, Switzerland, Turkey, United Arab Emirates, United Kingdomand Uruguay | Open, Classified |  |  | 296 | Computer | Computer
Z-0362-2019 | 81208 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Volumat MC Agilia Volumetric Infusion Pump, Model No. Z021135, Software version 1.7. Product Usage - The Volumat MC Agilia is a piece of transportable medical equipment intended to be used in healthcare facilities environment by healthcare trained professionals according to hospital protocols on adults, pediatrics and neonatal human patient to administer via a single channel or mounted on a multiple channels rack accessory: Intermittent or continuous delivery of parenteral fluids (solutions, colloids, parenteral nutrition) and medications (including but limited to diluted drugs, chemotherapy) through clinically accepted IV routes of administration). Transfusion of blood and blood derivatives products. Note that for transfusion procedures, use of a special administration set is required). |  |  | December 02, 2017 | 2017 | Fresenius Vial Sa | Fresenius Kabi initiated a mandatory software upgrade of the Volumat MC Agilia Volumetric Infusion Pump because it was designed to improve customer experience while preventing common software malfunctions. | Software design | A recall notification dated 11/17/17 was sent to customers to inform them that Fresenius Kabi initiated a mandatory software upgrade of the Volumat MC Agilia Volumetric Infusion Pump. The new software version 1.9a is currently available and designed to improve customer experience while preventing common software malfunctions. Customers will be contacted by a Fresenius Kabi service representative or distributor to schedule a service appointment for upgrading the software. Customers with any questions or require additional information are instructed to contact the Fresenius Kabi Technical Hotline at (800) 937-5060. |  | 7082 | 7082 units | Worldwide distribution - US Nationwide and countries of Canada and China. | Open, Classified |  |  | 7082 | Computer | Computer
Z-2046-2018 | 79963 | No | MQM | Bilirubin (total and unbound) in the neonate test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Draeger Jaundice Meter JM-103 |  |  | May 15, 2018 | 2018 | Draeger Medical Systems, Inc. | Users have misinterpreted the display for out of range measurement indicated by the blinking" ---" to mean a zero measurement. | Device Design | The firm, Drager, sent an "Urgent Medical Device Safety Notices" were distributed to customers on 5/15/18. The letter described the product, problem and actions to be taken. The letters instruct customers to perform the following:Enclosed with this letter are label/s and instructions on how and hwere to apply the label onto your device/s. Please make sure that all JM-105/JM-103 devices in your hospital are labeled according to the instructions in a timely manner.Please make sure that in your clinical environment the users are trained in the use of the equipment and the clinical workflow is set up in a way that the device is not used as a diagnostic device. In case of any uncertainty (e.g. risk factors cannot be evaluated), a blood test needs to be performed. The risk factors according to the Guideline of American Academy of Pediatrics are summarized in the attached Sample Usage Protocol Template as well as the recommended method to use the device as intended in a clinical environment.Please completed and return the attached Customer Reply card.If you have any questions regarding, please call 1-800-437-2437(press 1 at the prompt than 32349) between the hours of 8:00am-4:30pm EST. |  | 10363 | 10363 | Worldwide Distribution: US (nationwide) including states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY; and countries of: Algeria, Argentina, Armenia, Australia, Austria, Azerbaijan, Bahamas, Bahrain, Bangladesh, Barbados, Belgium, Benin, Bolivia, Bosnia-Herz., Botswana, Brazil, British Virgin Islands, Brunei Darussalam, Bulgaria, Burkina-Faso, Cambodia, Canada, Chile, Columbia, Congo, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, El Salvador, Estonia, Finland, France, French Guiana, Macedonia, Georgia, Germany, Ghana, Greece, Greenland, Guadeloupe, Guam, Guatemala, Hong Kong, Hungary, Iceland, India, Indonesia, Iran, Iraq, Ireland, Israel, Italy, Jordan, Kazakhstan, Kosovo, Kuwait, Laos, Latvia, Lebanon, Libya, Lithuania, Malaysia, Mayotte, Mexico, Mongolia, Morocco, Mozambique, Myanmar, Namibia, Nicaragua, Nigeria, Norway, Nepal, Netherlands, New Caledonia, New Zealand, Nigeria, Norway, Oman, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Qatar, Reunion, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sri Lanka, St Vincent, St Pier Miquel, Sweden, Switzerland, Tadzhikistan, Tanzania, Thailand, Trinidad Tobago, Tunisia, Turkey, Turkmenistan, United Kingdom, Uruguay, UAE, Venezuela, Vietnam, and Zimbabwe. | Open, Classified |  |  | 17074 | Computer | Computer
Z-2121-2018 | 79819 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Non-Locking Screw, D 2.7MM X 8MM, Sterile, Rx Only, ORTHOFIX S.r.r. Italy on label, Size (Part Numbers): a) D 2.7MM X 8MM (002-D-27008); |  |  | April 06, 2018 | 2018 | Orthofix, Inc | Firm is is conducting a voluntary recall of the Rival Reduce Fracture Plating System and Rival View Plating System. As the result of a complaint investigation, the locking screws were confirmed to not properly engage the plates during surgery. | Device Design | On 04/06/2018, Orthofix Inc. (U.S. Agent), sent written notification of the recall to distributors and hospitals via e-mail and/or FedEx overnight courier. |  | 935 | 935 | U.S Consignees: NJ, FL, PA, SC, LA, VA, MN, TX, GA, DC, NV, IN, CA, MI, IL, MD, NC, OH, OK, | Open, Classified |  |  | 7213 | Computer | Not_Computer
Z-2109-2018 | 80219 | No | LHN | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | 12C (AdaPTinsight) |  |  | October 28, 2016 | 2016 | Ion Beam Applications S.A. | IBA is initiating this recall to address an issue identified with AdaPTinsight software and to reduce risk related to this problem. | Software change control | IBA mailed an Urgent Field Safety Notice to affected customers on 10/28/2016 to inform them of the issue. The issue will be solved in AdaPTinsights versions starting from Al 2.0. The solution was to be deployed on impacted IBA Proton Therapy sites by September 2017. |  | 5 | 5 units (12C) | Worldwide Distribution - US Nationwide in the states of LA, TX, Italy, Sweden and France. Each country's National Competent Authorities were notified of the firm's Field Safety Notice. | Open, Classified |  |  | 5 | Computer | Computer
Z-3187-2018 | 80672 | No | JQW | Pipetting and diluting system for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Freedom EV0-2 with any version of EVOware and EVOware Sample Tracking Add-On V1.0 or higher |  |  | July 23, 2018 | 2018 | Tecan US, Inc. | There is a software issue that can potentially impact sample identification and traceability under specific conditions when using EVOware Add On Software: Sample Tracking -Multi Instrument V1.0 or higher, Part Number 30019971. | Device Design | On July 23, 2018, Tecan sent Field Corrective Action Letters to their consignees advising them of the problem. Field corrections will be made in the near future. |  | 19 | 19 EVOware Sample Tracking Software Add-on | Nationwide distribution to CA, MA, NC, OH, TX, and VA. Worldwide distribution to Australia, Canada, China, EU, and Japan. | Open, Classified |  |  | 19 | Computer | Computer
Z-2115-2018 | 80057 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Xper Flex Cardio Physiomonitoring System |  |  | November 20, 2017 | 2017 | Invivo Corporation | While in Full Disclosure playback, a user may inadvertently close the Full Disclosure Control Window using the ESC key function, instead of pressing the X icon in the upper right corner of this window. | Software design | The firm, Philips Healthcare, sent an "URGENT - Medical Device Correction" letter dated November 2017 and Field Safety Notice to each known affected customer. The letter describes the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue. The customers were instructed to do the following:Follow the Action to be Taken by Customer/User section of the Field Safety Notice as follows:-You may continue to use the device provided that each monitored patient in the Pre or Post-Op Holding Area is closely observed by a qualified health care professional and is not left unattended, as specified in the devices instructions for use. The most reliable method of patient monitoring combines close personal surveillance with correct operation of monitoring equipment. Also, users should always immediately respond to all audible alarms given by the device.Until your device can be corrected, please inform users of the content of this notice and instruct them that to properly exit Full Disclosure playback, use the X icon in the upper right corner of the Full Disclosure Playback control window and do not use the ESC key function.A Philips representative will contact you regarding any affected PCC modules. Each affected PCC will require a software update by Philips. Additionally, all Software and Documentation Media Kits will be replaced. This corrective action will be implemented free of charge by Philips.If you need any further information or support concerning this problem, please contact your local Philips representative: 800-669-1328 option 2, then option 3. |  | 1040 | 1,040 units (USA) and 7 units (Foreign) | Worldwide Distribution: US (nationwide) to states of: AK, AL, AR, AZ, CA, CT, DC, FL, GA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MO, NC, NH, NJ, NM, NY, NV, OH, PA, RI, SC, TN, TX, VA, WA, and WI ; and countries of: Canada, China, Great Britain and Korea. | Open, Classified |  |  | 1040 | Computer | Computer
Z-0010-2019 | 80718 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ABL90 FLEX analyzers with software versions 3.3MR1 or below. The ABL90 FLEX Analyzer is a portable, automated Analyzer that measures pH, blood gases, electrolytes, glucose, lactate, bilirubin and oximetry in whole blood. |  |  | May 25, 2018 | 2018 | Radiometer America Inc | The ABL90 FLEX can aspirate Cal 2 solution instead of rinse solution, when the following take place:Âż The analyzer is in ready modeÂż Operator lifts the inlet and then closes it again, thereby initiating a rinseÂż Immediately after lifts the inlet and quickly closes it again | Device Design | The firm, Radiometer America, Inc., notified customers via letter dated 05/25/2015. The letter provided the following: Reason for recall, Affect product description and codes, Risk for the patient, What customers must do, Countermeasure provided by Radiometer, contact information and Response Form. The customers were instructed to do the following:1. Inform users of the ABL90 FLEX analyzer that the inlet is not to be repeatedly lifted and closed. If this does happen, the easiest way to reset the analyzer flow selector manually is to re-install the solution pack. Follow the procedure described in the IFU for installation of a solution pack. You may re-use the same solution pack.2. We recommend that you review previous results, where cNa+ and cCl- results are off from the expected results with the values found in table 1.3. Complete the recall response fax form (last page of this customer advisory letter) and submit to your Radiometer representative via Fax to:714 582 0806.If you have any questions, please contact your Radiometer representative or Quality Assurance/ Regulatory Affairs Manager at 1-800-736-0600 opt.4. |  | 12102 | 12,102 | US Distribution and Internationally to Canada. | Open, Classified |  |  | 12102 | Computer | Not_Computer
Z-1303-2018 | 79532 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo.via. Medical Device Software. Picture archiving and communications system. |  |  | January 08, 2018 | 2018 | Siemens Medical Solutions USA, Inc | A software functionality in the report sections Findings Information and Summary of Measured Findings is not performing properly. The system will use the original values instead of displaying and saving corrected/modified values. Outdated information could be sent to the referring physician when creating the formal report in the reporting system. | Software design | Customer Advisory Notice SY065/17/P was sent to direct accounts. Firm is continuing to work on software solution. A workaround is available and is described in the Customer Advisory Notice. |  | 236 | 236 units | Device is software only. No products are distributed to wholesale dealers, distributers or retailers | Open, Classified |  |  | 236 | Computer | Computer
Z-2248-2018 | 80200 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Biograph mCT Systems: |  |  | May 29, 2018 | 2018 | Siemens Medical Solutions USA, Inc. | Siemens Medical Solutions USA, Inc., Molecular Imaging business line became aware, through our supplier that some laser cover windows may fall off a limited population of the Biograph mCT and Biograph mCT Flow systems. | Device Design | The firm notified the consignees of the issue by email on May 29, 2018. The letter identified the affected product, problem and actions to be taken. They are tracking completion of the letter delivery through the hand-delivered approach. For further questions, please call (847) 304-7700.. |  | 405 | 405 units | Worldwide Distribution | Open, Classified |  |  | 405 | Computer | Not_Computer
Z-1582-2018 | 79411 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | AQURE, Software version 2.3.0 and 2.3.1 |  |  | February 27, 2018 | 2018 | Radiometer Medical ApS | There is a potential problem relating to the AQURE System, versions 2.3.0 and 2.3.1, that may result in patient mix-up. | Device Design | Radiometer Medical sent a Notification letter dated February 27, 2018 to customers. The letter identified the affected product, problem and actions to be taken For questions contact your Radiometer representative. |  | 2 | 2 | US Nationwide Distribution in the states to Georgia and Wisconsin. | Completed |  |  | 2 | Computer | Computer
Z-1293-2018 | 79464 | No | OSP | External uterine contraction monitor and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Uterine Electromyographic Monitor -Monica IF24 Interface System |  |  | December 15, 2017 | 2017 | GE Healthcare, LLC | if the cables of the Monica IF24 System are removed by a user, and then incorrectly re-connected the Fetal Heart Rate (FHR) and Maternal Heart Rate (MHR) signals would be displayed incorrectly (swapped). | Device Design | GE Healthcare sent an Urgent Medical Device correction Letter dated December 15, 2017 to their customers. The letter identified the affected product, problem and action to be taken. Customers are instructed to:1. Follow the attached safety instructions to replace the screws, see attached page 2.2. Follow the attached safety instructions to perform the Monica Novii function TEST to confirm correct cable connection, see attached page 3.3. If the Monica Novii System fails the Novii function TEST, contact a GE Healthcare Representative to schedule repair/replacement.4. If the Monica Novii System passes the Novii function TEST, you may continue to use Monica Novii System without any further action.5. Complete the attached fax back form (page 4&amp;5) for all Monica Novii System as soon as you have completed the screw replacement.If you have any questions or concerns regarding this notification, please contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative |  | 355 | 355 | Worldwide - US Nationwide Distribution. | Open, Classified |  |  | 4194 | Computer | Computer
Z-2283-2018 | 80272 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | ARTIS Pheno Interventional Fluoroscopic X-Ray System, Model 10849000 |  |  | May 01, 2018 | 2018 | Siemens Medical Solutions USA, Inc | If the C-arm is manually moved in the longitudinal direction to a fully extended position, the system's safety stop mechanism may be activated. As a result, the C-arm stops the movement and can drop up to 20 mm (0.8 ). Under certain circumstances, the cover of the detector or collimator can touch and possibly injure the patient due to the C-arm drop. This could lead to a situation in which it is necessary to stop the clinical treatment or continue the treatment on an alternative system. | Device Design | On May 1, 2018, the firm distributed Urgent Medical Device Correction letters to its affected customers.The letter instructed customers to follow standard emergency procedures they have in place in the event of system failures. The firm recommended considering reinforcing the emergency procedures until the update has been completed. As an immediate measure, operators were instructed to avoid driving the C-arm into a fully extended longitudinal position.The firm is currently developing a solution to eliminate the root cause of the problem. The solution will be distributed to all affected customers upon completion.Customers were asked to forward the safety information to any other organizations that could be affected. |  | 39 | 39 | Healthcare facilities in AL, CA, CO, FL, IL, KY, MA, MI, MN, MS, NE, NJ, NY, OH, TN, TX, UT, and VA. Foreign distribution to Australia, Austria, Belgium, Denmark, France, Germany, Italy, Japan, Kuwait, Netherlands, Norway, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, and United Kingdom. | Open, Classified |  |  | 39 | Computer | Not_Computer
Z-2284-2018 | 80217 | No | LHN | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | The Proton Therapy System - Proteus 235 (brand names: Proteus Plus and Proteus ONE) |  |  | April 11, 2016 | 2016 | Ion Beam Applications S.A. | IBA is conducting a voluntary recall to address a PTS (Proton Therapy System) software issue and to reduce the risk related to this issue. | Software Manufacturing/Software Deployment | IBA mailed an Urgent Field Safety Notice to affected customers to inform them and prevent risk associated with this issue. IBA also recommends users always check for any discrepancy between the prescribed and delivered MUs on their clinical screen after an irradiation in which a pause occurred. IBA also recommends users always check for any discrepancy between the prescribed and delivered MUs on their clinical screen after an irradiation in which a pause occurred. A software upgrade will be deployed on impacted units to correct this issue. |  | 2 | 2 | Units were distributed to Korea and Jacksonville, FL. | Open, Classified |  |  | 2 | Computer | Computer
Z-2428-2018 | 79993 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | BioMerieux mini VIDAS System in combination with Quality Control VIDAS* (QCV*) |  |  | January 04, 2018 | 2018 | bioMerieux, Inc. | Increase of events affecting performance regarding pump clogging and displacement of the color-coded SPR label (DOT). These performance issues could contribute to erroneous results for any assay performed on the VIDASÂż and mini VIDASÂż system due to a decrease of pipetted volume. | Device Design | Subsidiaries and distributors were initially notified of the recall on approximately 01/04/2018. U.S. customers were notified via letter on approximately 01/22/2018. On 03/28/2018 a follow-up notification was issued world-wide to customers to expand the scope. Instructions included to increase the frequency of the Quality Control VIDASÂż (QCVÂż) testing to weekly rather than monthly to reduce the period of potential retrospective analyses needed, continue to conduct a visual inspection of the SPRÂż after each run (including patient sample test and QCVÂż test) to ensure that the dot on the SPRÂż is still in place, continue to perform External QC testing in accordance with laboratory policy and in compliance with Governmental Regulations &amp; Accreditation Requirements, which includes CLIA, ensure the notification letter is distributed to and reviewed by all appropriate personnel, store the notification letter with the bioMÂżrieux VIDASÂż and /or miniVIDASÂż documentation, and complete and return the Acknowledgement Form in Attachment A by Fax to confirm receipt of the notice. For further questions, please call (314) 731-8694 |  | 1866 | 1,866 units | Worldwide Distribution - US Nationwide in the states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, VA, VT, WA, WI, WV, WY. International distribution to Algeria, Angola, Anguilla, Argentina, Armenia, Australia, Austria, Azerbaijan, Bahrain, Bangladesh, Belarus, Belgium, Benin, Bolivia, Bosnia - Herzegovina, Brazil, Brunei, Bulgaria, Burkina Faso, Cambodia, Cameroon, Canada, Cape Verde, Chad, Chile, China / Hong Kong, Colombia, Congo, Costa Rica, Croatia, Czech Republic, Denmark, Djibouti, Dominican Republic, Dutch Antilles, Ecuador, Egypt, El Salvador, Estonia, Ethiopia, Finland, France, Gabon, Georgia, Germany, Ghana, Greece, Guam, Guatemala, Guinea, Honduras, Hungary, India, Indonesia, Iran, Iraq, Israel, Italy, Ivory Coast, Japan, Jordan, Kazakhstan, Kenya, Kuwait, Latvia, Lebanon, Libya, Lithuania, Macedonia, Malaysia, Mali, Malta, Mauretania, Mexico, Mongolia, Morocco, Mozambique, Myanmar, Nepal, Netherlands, Nicaragua, Niger, Nigeria, Norway, Oman, Pakistan, Palestine, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Qatar, Romania, Russian Fed., Saudi Arabia, Senegal, Serbia, Singapore, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Sudan, Sweden, Switzerland, Syria, Tanzania, Thailand, Togo, Tunisia, Turkey, Turkmenistan, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Uzbekistan, Venezuela, Vietnam, Zimbabwe. | Open, Classified |  |  | 1866 | Computer | Not_Computer
Z-0706-2017 | 75755 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Philips Spectral CT Applications supports viewing and analysis of images at energies selected from the available spectrum in order to provide information about the chemical composition of the body materials and/or contrast agents. |  |  | November 14, 2016 | 2016 | Philips Medical Systems (Cleveland) Inc | A software defect is causing issues with the IQon Spectral CT and Philips Spectral CT Applications while filming, saving and annotating images | Software design | A Field Safety Notice (Attachment #1 Field Safety Notice FSN 16-061-72800659 and FSN 88100051), will be sent to US consignees via certified mail. Philips Key Market representatives will distribute the notice internationally, based on local regulations. Additionally, a Field Service Engineer will contact each site to schedule a time to implement the Field Change Order. |  | 20 | 20 | Software was distributed in the following states: AR, IN, LA, MA, MN, OR and TN.Softward was also shipped to the following countries: Belgium, Denmark, France, Germany, Japan, Korea, Switzerland and United Kingdom. | Terminated | July 02, 2018 | 595 | 44 | Computer | Computer
Z-2340-2018 | 80262 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare CARESCAPE Monitor B650 |  |  | May 25, 2018 | 2018 | GE Healthcare Finland Oy | When multiple CARESCAPE Monitor B650 units are connected to the same network and a network overload occurs for a prolonged time, the monitors may simultaneously restart as designed. The monitor restart will not be completed until the network issue has been corrected. | Software design | The firm notified its consignees of the problem by letter on 05/25/2018. The letter stated the following safety instructions:"If a prolonged network overload occurs and CARESCAPE Monitor B650 on the network does not resume normal functionality in approximately 90 seconds, take the following actions: 1. Temporarily switch the monitor from central monitoring to local-only (bedside) monitoring: - Disconnect the LAN cables from IX/MC ports on backside of the monitor. The monitor will now be in local-only mode. - Keep the patient under close observation and continuous surveillance at the local (bedside) monitor. Disconnect used LAN cables from the monitor, see Figure 1, below, showing the cables to disconnect. 2. Contact your IT department or other appropriate personnel in your facility responsible for the Patient Monitoring Network, and provide them with the further instructions on network configuration provided.3.After the network issue has been resolved, switch the monitor back into central monitoring by re-connecting the LAN cables to IX/MC ports.4. Verify correct monitoring state and alarm function.If a transient network overload occurs in which the CARESCAPE Monitor B650 on the network restarts within approximately 90 seconds (i.e., all functionality resumes and the monitor reconnects to the network automatically), proceed with instruction step #2 above to check network functionality. If your network is not maintained by GE, please ensure that the network in your facility meets the network configuration requirements. GE Healthcare will provide revised network installation and user instructions documentation at no cost to you. Complete and return the Customer Response Form attached to this notification. Once the Customer Response Form is received, GE Healthcare will arrange for the documentation delivery. "NOTE: The letter has an attachment, "Appendix A", which has more detail on the net work configuration. |  | 69231 | 69231 units | worldwide | Open, Classified |  |  | 69231 | Computer | Computer
Z-3021-2018 | 80765 | No | MCX |  |  | Cardiovascular | PMA | CSI, Orbital Atherectomy System (OAS) Saline Infusion Pump, Rx only, Part Number: 7-10014, Model Number: SIP-3000. |  |  | April 13, 2018 | 2018 | Cardiovascular Systems Inc | The pumps may switch to stand-by during use requiring the pump to be reset prior to continuing treatment. | Device Design | Consignees were sent a CSI "Urgent Medical Device Recall" letter dated 13 APRIL 2018. The letter listed the Affected Product, Recall Description, SIP Replacement and Return &amp; contact information. Consignees were notified the pumps will be replaced. Consignees may continue to use their pumps however if a yellow light fault is observed and the troubleshooting per the IFU does not resolve the fault, discontinue use of the pump and notify CSI. |  | 18 | 18 | US Nationwide Distribution in the states of AZ, CA, FL, GA, IL, MO, and TX | Open, Classified |  |  | 18 | Computer | Not_Computer
Z-0975-2018 | 79360 | No | KXJ | Radiologic table. | Radiology | Radiology | 510(K) Exempt | IMRIS ORT100 Table, Part Number 110470-000 |  |  | February 20, 2018 | 2018 | Deerfield Imaging, Inc. | The rotational locking mechanisms on the ORT 100/ T2X tables can become disengaged or unlocked while in use which could potentially result in unintended movement of the table. | Device Design | The firm disseminated the notices on 02/23/2018 and followed with emails on 02/26/2018. The notices stated:"To minimize the risk of injury, IMRIS is taking the following corrective actions:" ORT100 Table: The rotational locking mechanism shall be replaced in all ORT100 tables that do not contain the most current design." T2X Table: Tables will have their rotational lock inspected. If components are worn or performance is other than intended, the rotational lock assembly will be serviced or replaced entirely.IMRIS has identified that your facility has one or more of these OR Tables which require evaluation. Please confirm details on page 2 and return your response to IMRIS.IMRIS Customer Service will contact you to implement this corrective action.IMRIS recommends these preventive actions:" Per the ORT Operator Manuals the table should be tested daily in all standard positions and inspected for anomalies including the fit of the covers, back section, headrest, pendant connection and swivel release pedal free engagement" During usage, verify the table is locked by manually pushing the table to make sure it does not move rotationally." If the rotational lock is observed to be not functioning properly, please contact IMRIS Customer Service.The appropriate competent authorities have been notified of this corrective action. IMRIS is committed to patient and user safety, and apologizes for any inconvenience that this may cause. Please contact customer support via our service line at 1-866-475-0525 or Email: customersupport@i mris.com if you have any questions or concerns." |  | 15 | 15 tables | US, Canada, Australia, China, Qatar | Open, Classified |  |  | 31 | Computer | Not_Computer
Z-2356-2018 | 80148 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | microTargeting" Drive System Distributed by Medtronic, Product Number 9033G0601 |  |  | May 10, 2018 | 2018 | FHC, Inc. | The FHC Depth Stop Adapter may cause damage to the lead and stylet when the depth stop screw is over-tightened onto the lead. This damage can manifest as low impedances or a short circuit between one or multiple electrode pairs in the lead, and has resulted in the need to remove and replace the implanted lead intraoperative or in a follow-up procedure. | Device Design | Urgent Medical Device Recall notification letters were sent to customers on 5/10/18. The letters included the following instructions:FHC, Inc. Field Representatives will review your inventory and replace effected Depth Stop Adapters you have on hand with an improved design. In the interim, we advise that you do not over-tighten the Depth Stop Adapter screw. If low impedances or a short circuit is found, physicians should replace the lead prior to use and report the issue to FHC, Inc. or your local field rep. |  | 2952 | 2952 total | The devices were distributed to the following US states: AL, AR, AZ, CA, CO, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, VT, WA, WI, WV, and WY.The devices were distributed to the following foreign countries: Argentina, Australia, Canada, Columbia, Denmark, Ecuador, France, Germany, Great Britain, Italy, Japan, Norway, Romania, South Africa, Switzerland, Saudi Arabia, Taiwan. | Open, Classified |  |  | 2952 | Computer | Not_Computer
Z-2993-2018 | 80826 | No | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | Diagnostica Stago STA NÂżoplastineÂż CI Plus, 10 ml vial. In-Vitro Diagnostic for determination of the prothrombin time (PT) in plasma. |  |  | July 24, 2018 | 2018 | Diagnostica Stago, Inc. | QC values outside of the assigned ranges (prolonged Prothrombin time / decreased PT %) | Device Design | Diagnostica Stago issued notification on July 24, 2018 to US Customers provided the reason for the notification, actions to be taken to mitigate the risk. Action: If you have in your laboratory, any of the lots among those listed in the appendix, we are asking you to: If it is not already done, run a Quality Control test at every change of vial.Return to your Stago affiliate, by fax or e-mail, the enclosed form completed and confirming that you have read this letter.Questions contact Stago Hotline (1-800-725-0607) |  | 24043 | 24043 kits | Nationwide | Open, Classified |  |  | 24048 | Computer | Not_Computer
Z-2379-2018 | 80046 | No | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | BCI Advisor Vital Signs Monitor, Model 9200, Reorder numbers 920654235, 920654235, 920654335A, and 920674235A. (The differences in the reorder numbers have to do with certain functions being installed or not installed.) |  |  | November 29, 2017 | 2017 | Smiths Medical ASD Inc. | Some printing modes on the devices may not function properly (tabular data printing, real-time continuous printouts, and automatic interrupt print feature). | Software design | Smiths Medical sent an Urgent Medical Device Letters dated 11/29/2017 via FedEx tracked mailing. The letter identified the affected product, problem and actions to be taken. For questions call Smith Medical Customer Service at 1-800-258-5361. |  | 10 | 10 devices | Worldwide Distribution - US Nationwide in the states of CO and MT. Foreign distribution was to Armenia. | Terminated | July 11, 2018 | 224 | 10 | Computer | Computer
Z-2992-2018 | 80896 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | VidiStar(TM) PACS &amp; DICOM Viewer Software system |  |  | August 28, 2018 | 2018 | Hitachi Healthcare Americas Corp Informatics Division | The secure filesystem client software used in the interface between the Vidistar PACS and an EHR system may cause intermixed images from multiple patients showing in a single study. | Software in the Use Environment | The firm initiated the recall by email on 08/28/2018. The firm will be removing the 3rd party software from the interface and replacing it with new software and notify the affected customers when the action is completed. |  | 12 | 12 units | SC, CO, IL, AZ, MT, OH, IN, TX, GA, MA | Open, Classified |  |  | 12 | Computer | Computer
Z-0639-2018 | 78944 | No | OLO | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | ROSA Spine 1.0.2 Stereotaxic instrument Computer-assisted surgical device. |  |  | May 01, 2017 | 2017 | Zimmer Biomet, Inc. | Robot arm being sent to the wrong position | Software design | Urgent Medical Device Recall (Correction) notices were sent to customers.The letter provided to hospital risk managers and surgeons identified the issue and their responsibilities. These responsibilities included:- Stop using the micromovement function in isocentric mode during Brain Surgery with ROSA Brain 3.0 device. -Reviewing the notification to ensure comprehension of the steps required to be taken prior to the permanent correction that will be implemented on site by a service engineer.- Completing and returning Attachment 1 (Certificate of Acknowledgement) to the e-mail address provided.The firm has prepared a software correction that will be deployed and implemented on site by a MEDTECH representative. |  | 1 | 1 | Distributed to CA, DC, FL, IL, KY, MA, MN, NC, NY, OH, and PA. | Open, Classified |  |  | 17 | Computer | Not_Computer
Z-0635-2018 | 78606 | No | MXK | AC-powered slitlamp biomicroscope. | Ophthalmic | Ophthalmic | 510(k) | OCULUS Pentacam AXL, Model 70100, software versions 1.21r01, 1.21r03, 1.21r11, and 1.21r13. The firm name on the label is OCULUS Optikgerate GmbH, Made in Germany. |  |  | November 21, 2017 | 2017 | Oculus Optikgeraete GMBH | The device software versions have an anomaly which may produce an erroneous marking for the quality specification value | Software design | Oculus issued an email on November 21, 2017 to their U.S. subsidiary which contained a Mandatory Software Update notification. The email informed the U.S. subsidiary that a mandatory software update was planned for the Pentacam AXL in possession of their customers, provided the software versions affected, and the reason for the mandatory update. The email explained the USB thumb drive will be shipped to the U.S. subsidiary for each serial number that may be affected. The U.S. subsidiary was to match the serial numbers to their customers and complete the attached customer letter for issuance to each of the customers. The U.S. subsidiary is then supposed to issue the letter and the USB thumb drive via certified mail to each of their customers who have the affected serial numbers. The customer is to send back the completed, signed form verifying the mandatory software update has been installed on all working places on which the Pentacam AXL software is running or that the procedure was performed by an authorized distributor.For further questions, please call (425) 670-9977 |  | 83 | 83 devices | US Distribution to the states of : CA and FL. There were no foreign/military/government accounts. | Terminated | March 22, 2018 | 121 | 83 | Computer | Computer
Z-0641-2018 | 79205 | No | JEB | Manual operating table and accessories and manual operating chair and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Allen Medical Bow Frame Model A-70800. Surgical platform used to position patients for lumbar spine procedures. |  |  | October 06, 2017 | 2017 | Allen Medical Systems | The base of the frame has the potential to crack when the patient is initially transferred to the Bow Frame. | Device Design | In October of 2017 an Urgent Medical Device Correction letter was sent to customers notifying them of the recall and requesting customers to immediately remove any frames from use that are in the serial number range subject to the recall and that a service technician will coordinate correction of the device. If the device was sent to another company, to forward the recall notice o that person or firm. If you have any questions concerning this Urgent Medical Device Correction, please contact Allen Medical Customer Service at (800) 433-5774, 8:30  5:00 pm EST. |  | 161 | 161 units total | Worldwide distribution. US Nationwide, Australia, Bangladesh, Brazil, Canada, China, Denmark, France, Germany, Greece, Italy, Japan, Korea, Kuwait, Lebanon, Malaysia, Netherlands, Spain, Switzerland, Taiwan, UAE, and UK. | Open, Classified |  |  | 161 | Computer | Not_Computer
Z-0643-2018 | 79237 | No | FDT | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Pentax Video Duodenoscope Model: ED-3490TK (UDI of design being recalled: 04961333232420) |  |  | February 13, 2018 | 2018 | Pentax of America Inc | The duodenoscopes are being recalled in order to replace the forceps elevator mechanism, the O-rings, and the distal end covering to be consistent with the updated design as well as provide an updated periodic inspection as part of the Operation Manual in order to mitigate the potential risk of infection in flexible endoscopy. | Device Design | Pentax Medical sent an Urgent Medical Device Correction/Removal letter dated February 7, 2018, to the United States Customers. The consignee letter includes a customer response form and a revised operator manual with the added recommended periodic maintenance. The letter requests the return of the form which includes an accounting of the devices (by serial number) owned by the facility. The firm is to contact the consignee to arrange return of the affected units and to provide loaner devices as needed. For further questions, please call 1 (800) 431-5880 ext. 2064. |  | 559 | 559 | USA (nationwide) Distribution to the states of including Puerto Rico : AL, AZ, CA, CO, CT, DE, FL, GA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NH, NJ, NV, NY, OH, OK, PA, TN, TX, UT, VA, WA, WI, WV and WY, and the District of Columbia. | Open, Classified |  |  | 559 | Computer | Not_Computer
Z-2392-2018 | 80255 | No | MEC | Medical washer-disinfector. | General Hospital | General Hospital | 510(K) Exempt | Reliance 1227 Cart and Utensil Washer/Disinfector, FW03101, FW03102, FW03S003 |  |  | May 23, 2018 | 2018 | Steris Corporation | The firm has become aware that the Reliance 1227 Cart and Utensil Washer/Disinfector s Chemical Low Level alarm, intended to prevent the user from initiating a cycle when a low chemical level situation occurs, does not operate as intended. Currently, if a low chemical level situation occurs, the alarm will only be generated at the unit s next power-up or when accessing service mode. The failure of the Reliance 1227 to identify that the chemicals used for cleaning are low or empty could result in bedpans and urinals, basins, case carts, beds, theater shoes and other miscellaneous reusable patient care items not being properly cleaned before disinfection or reuse. | Software Design Change | On May 23, 2018, the firm notified affected customers by Urgent Device Field Correction letters distributed via FedEx. The letter stated that the system software had been updated to correct the issue, and that a STERIS service technician would be contacting each facility to perform the system software update.In the interim of the software update, users were instructed to continue to verify the units chemistry levels daily as part of the facilitys Quality Management System. Chemistry containers with low chemistry volume should be replaced prior to initiating the next cycle.If you have questions regarding this matter, please contact Katilin Slaper, Product Manager, at (440) 392-7431, STERIS Customer Service at 1-800-548-4873, or your local STERIS Representative. |  | 118 | 118 | Worldwide Distribution - US Nationwide in the states of AL, AK, CA, CO, CN, DE, FL, GA, IL, IA, KS, KY, LA, MA, MI, NJ, NY, ND, OH, OK, OR, PA, RI, SC, TN, TX, VT, VA, WI, and WV. Foreign distribution to Canada. | Open, Classified |  |  | 118 | Computer | Not_Computer
Z-0970-2018 | 79076 | No | JJY | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Immunoassay Premium Plus Controls (3x4x5mL and 12x5mL packages) |  |  | August 11, 2017 | 2017 | Randox Laboratories | The ACTH in Immunoassay Premium and Premium Plus quality control does not meet the quoted reconstituted stability claim in the product IFU. | Device Design | On July 24, 2017 an Urgent notice was issued instructing customers to place a copy a notice in all kits remaining in stock. This notifies customers that they should amend the claims for open stability in the instructions for use, to state use immediately after the 30 minute reconstitution period. Questions or concerns can be emailed to technical.services@randox.com |  | 199 | 199 Kits | Distribution US nationwide. | Completed |  |  | 316 | Computer | Not_Computer
Z-2319-2018 | 80212 | Yes | LWS |  |  | Cardiovascular | PMA | Ellipse, Sterile EO, Model #/ Part #: |  |  | April 16, 2018 | 2018 | St Jude Medical Inc. | The firm is releasing new firmware for high voltage implantable cardiac devices (ICD) and cardiac resynchronization therapy defibrillators (CRT-D) intended to strengthen security by providing an additional layer of protection against unauthorized device access. For older generation devices not capable of accepting the firmware, the firm is advising customers of the option to permanently disable the RF communication capability in these devices. | Software design | On 04/16/2018, Important Medical Device Advisory notifications were mailed via overnight service to U.S. physicians and hospitals notifying customers of the availability of the programmer software update and associated ICD/CRT-D firmware update. Advisory notifications were mailed to customers outside the U.S. on different dates depending on geography and local regulatory requirements. Prophylactic replacement of affected devices is not recommended. The firm recommends firmware upgrades for all eligible patients at the next regularly scheduled visit or when appropriate depending on the preferences of the patient and physician. For devices not eligible for the cybersecurity firmware update the firm recommends a discussion of the risks of cybersecurity vulnerabilities and proven benefits of remote monitoring with patients at their next regularly scheduled visit. RF communication may be permanently disabled in devices not eligible for firmware updates during an in-clinic device interrogation with a programmer that has received the software update. Customers with questions about the firmware update can call the customer technical support hotline at 1-800-436-5056. |  | 139351 | 139351 | U.S. Nationwide distribution | Open, Classified |  |  | 811702 | Computer | Computer
Z-3175-2018 | 80979 | No | EPT | Surgical microscope and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Leica M525 F20 Surgical Microscope System |  |  | September 04, 2018 | 2018 | Leica Microsystems, Inc. | The firm received complaints that the instrument tipped over when the swing arm was fully extended. | Device Design | The firm disseminated their Field Corrective Action Notices by mail beginning on 09/04/2018 to the US and Canadian consignees. The letter explained the problem and informed the consignee that an update package with enhanced instructions for use and additional marking will be provided within the next 3 months.Additional consignees found were notified by email on 10/01/2018. |  | 160 | 160 units | US and Canada | Open, Classified |  |  | 160 | Computer | Not_Computer
Z-2315-2018 | 80019 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Diagnostic Ultrasound System, Aplio I-series. Labeled as i700, i800, i900 |  |  | January 04, 2018 | 2018 | Toshiba American Medical Systems Inc | The power supply unit may be damaged and may fail to start up when the power plug is inserted into a power outlet while the main power switch on the rear of the system isset to ON. | Software design | Canon sent an Urgent - Medical Device Correction letter dated January 4, 2018. Customers were notified of the issue and informed that a modified firmware will be installed to correct the issue. When that modified firmware becomes available, the Toshiba service representative will contact customers to schedule an appointment to install the firmware. Until the modified firmware is installed, instructions were provided to confirm that the main power switch is sent to OFF, insert the power plug into the power outlet, and to set the main power switch to ON. The system should be operated according to the procedure described in section 9.1, Connection the Power Cable and Protective Grounding, in the operation manual. Additionally, customers were asked to complete and return the attached Customer Reply Form. For further questions, please call (714) 730-5000. |  | 162 | 162 units total | USA (nationwide) Distribution to the states of : CA, CO, DE, GA, HI, IA, ID, IL, IN, MD, MI, MN, MO, NC, ND, NV, OH, OR, PA, SC, SD, TN, TX, VA, WA, WV and WY. | Open, Classified |  |  | 162 | Computer | Computer
Z-3188-2018 | 80938 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | sterEOS workstation |  |  | August 01, 2018 | 2018 | EOS Imaging | 3D projections may be incorrect whenthe pair of images used is a secondary capture generated from a series containing more than two images, where one of the images has been manually flipped. Projection errors may lead to errors in calculation of 3D clinical parameters whenadjusting the position, size and shape of the 3D objects used. Incorrect clinical parameters can be a contributing factor in incorrect clinical decisions, or sub-optimal surgical interventions. | Software design | On August 1, 2018, a Field Safety Notice was issued to all impacted customers. Measures to be taken by users: EOS system and sterEOS review workstation users are requested not to apply manual symmetry to an image before saving a 2D secondary capture if more than two images are displayed in the EOS viewer mode or in 2D sterEOS.Users of the 3D sterEOS review workstation are requested to check that 3D projections associated with pelvic parameters match when a 3D image is created from a pair of images belonging to a series of more than two images.INSTRUCTIONS FOR END USERS:1) Please read the field safety notice.2) Follow the instructions provided in this field safety notice.3) Please transfer this notice to whomever needs to be informed within your organization or in any organization to which the sterEOS review workstation may have been transferred.INSTRUCTIONS FOR DISTRIBUTORS:1) Please read the field safety notice.2) Please transfer all the documentation relating to the field safety notice to your customers in the field and keep proof of delivery of this documentation. Please regularly submit your customers' response rate to EOS imaging by email (regulatory@eos-imaging.com). EOS imaging needs this proof to monitor the evolution of corrective measures. |  | 17 | 17 units | US nationwide | Open, Classified |  |  | 17 | Computer | Computer
Z-0772-2018 | 79101 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | V-PRO 60 Low Temperature Sterilization System (UDI: 00724995152000) |  |  | September 28, 2017 | 2017 | Steris Corporation | Correction to update sterilizer software to prohibit the use of expired sterilant prior to the start of a sterilization cycle. | Device Design | The V-PRO 60 system software has been updated to ensure the VAPROX Sterilant expiration date is verified prior to the initiation of each sterilization cycle. STERIS Service Technicians will be contacting your facility to update your V-PRO 60 Sterilizer(s) with the revised system software. User Action  Users should continue to review their V-PRO 60 cycle printouts upon completion of each cycle to verify the VAPROX Sterilant was not expired. The VAPROX Sterilant expiration date is identified near the top of each cycle printout following all sterilization cycles as notated in the following illustration: |  | 406 | 406 | Distributed in 44 states: AL, AK, AZ, AR, CA, CO, DE, FL, GA, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, WY; Puerto Rico, and the following countries: Algeria, Australia, Azerbaijan, Bahrain, Brazil, Brunei Darussalam, Bulgaria, Canada, Chile, Colombia, Costa Rica, Cyprus, Denmark, Ecuador, Egypt, France, Germany, Hong Kong, Hungary, India,Israel, Italy, Republic of Korea, Kuwait, Latvia, Malaysia, Mexico,Morocco, Myanmar, Panama,Peru, Philippines, Poland, Portugal, Qatar, Saudi Arabia, South Africa, Spain, Switzerland, Thailand, Turkey, United Arab Emirates, and Vietnam. | Terminated | September 12, 2018 | 349 | 406 | Computer | Computer
Z-2286-2018 | 80022 | No | PFT | Reagents for molecular diagnostic instrument test systems. | Clinical Chemistry | Toxicology | 510(K) Exempt | ABX Pentra Micro ALBUMIN 2 CP, Serum/Plasma, REF 1300032563 |  |  | April 26, 2018 | 2018 | Horiba Instruments, Inc dba Horiba Medical | When the ABX Pentra Micro ALBUMIN 2 CP (ALBT2) Reagent 2 is sampled and dispensed immediately after an ABX Pentra Creatinine 120 CP (CREA3) Reagent 2 dispense, the final result of the microalbumin may be elevated more than 10%. This increase may affect serum, plasma and urine sample types. | Device Design | The firm, Horiba Medical, sent a "FIELD SAFETY NOTICE" dated April 26, 2018 to the customers on the customer list on April 26, 2018. The notice described the product, problem and actions to be taken. The notice directed the consignee to manually program the system to address the issue and the share the information with the laboratory staff and to complete and return the enclosed Acknowledgment Form within 10 days via Fax: (949)753-0533, Email: azita.hedayati@horiba.com or Mail: HORIBA Medical, 9755 Research Dr., Irvine, CA 92618.If you have any questions regarding this Product Corrective Action, please contact your local Horiba Medical representative or our Technical Service Department at 888-903-5001, option 3, option 2. |  | 686 | 686 devices (233 devices US, 453 devices OUS) | Worldwide Distribution: US (nationwide) and to countries of: Bulgaria, Burma, Cameroon, Djibouti, France, Gabon, Germany, Italy, Lebanon, Mali, Mauritania, Poland, Portugal, Romania, Saudi Arabia, South Africa and Thailand. | Open, Classified |  |  | 686 | Computer | Not_Computer
Z-0629-2018 | 78173 | Yes | DWC | Cardiopulmonary bypass temperature controller. | Cardiovascular | Cardiovascular | 510(k) | HemothermÂż Dual Reservoir Cooler/Heater System Models 400CE and 400MR |  |  | November 27, 2017 | 2017 | Cincinnati Sub-Zero Products LLC | Due to FDA's safety communication relating to nontuberculous mycobacteria (NTM) infections related to heater/cooler devices, an updated Instructions For Use includes a more robust cleaning &amp; disinfection procedure. | Device Design | Between November 27, 2017 - December 8, 2017, Cincinnati Sub-Zero Products LLC, distributed Urgent Medical Device Correction notices to their customers via Email and USPS. This Medical Device Correction has been initiated to provide a labeling update to the cleaning &amp; disinfection procedures. The update addresses the disinfection process, cleaning of reusable accessories, as well as clarifying the water type that should be used with the Hemotherm. Customers are encouraged to do the following: 1) Discard your current Hemotherm Operation &amp; Technical Manual and/or Operation Manual. 2) Replace your current Hemotherm Operation &amp; Technical Manual and/or Operation Manual with the enclosed updated Manual(s). For your convenience, all updates have been highlighted in the manuals. "For Hemotherm 400CE users, utilize the P/N 56418-R Operation &amp; Technical Manual and P/N 56075-G Operation Manuals. "For Hemotherm 400MR users, utilize the P/N 56234-S Operation Manual.All customers are encouraged to complete and return the enclosed response form by fax or email to FA2016-009@genthermcsz.com; Fax: 513-772-9119.Customers with questions or concerns, please contact the Field Action Coordinator: Kathy DeSmidt, E-mail: FA2016-009@genthermcsz.com, Phone: 513-772-8810 or 800-989-7373. |  | 1355 | 1355 units | US Nationwide Distribution | Terminated | July 02, 2018 | 217 | 1355 | Computer | Not_Computer
Z-3192-2018 | 80761 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Xper Flex Cardio Physiomonitoring System Model # 453564669081 Software Version 1.6.0.0388 |  |  | August 03, 2018 | 2018 | Invivo Corporation | Released with an incorrect, older firmware version installed. May result in intermittent loss of ECG monitoring, inaccurate ECG amplitude display and or auxiliary output, and inaccurate heart rate (HR) displayed due to QRS detection fault. | Software Manufacturing/Software Deployment | Philips is asking customers to follow the Action to be Taken by Customer/User section of the August 3, 2018, Field Safety Notice: The Field Safety Notice informs customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue. Philips will replace all affected devices. Users should discontinue use of affected devices immediately. A Philips representative will contact you regarding your affected device. |  | 38 | 38 | Distribution in the Netherlands | Open, Classified |  |  | 38 | Computer | Computer
Z-3190-2018 | 80905 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | ARTIS Pheno Interventional Fluoroscopic X-Ray System, Model Number 10849000 |  |  | July 29, 2018 | 2018 | Siemens Medical Solutions USA, Inc | Temperature of the detector may fall out of its designed working range (cool down) due to incorrectly configured temperature control unit. This unintentional system behavior may occur while the unit is powered down for long periods (e.g. at night). | Device Design | Urgent Medical Device Correction notification letters dated 7/27/18 were distributed to customers. The letters informed the customers of the following planned actions:Our service organization will contact you shortly to arrange a date to perform this corrective action with Update Instruction AX001/18/S. Please feel free to contact our service organization for an earlier appointment at 1-800-888-7436.Following the corrective action, the cause is eliminated and a recurrence of the fault is prevented. Additionally, Siemens is developing a solution to restore the system back to the original settings. This solution will be distributed to all affected customers upon completion. |  | 429 | 429 total | US nationwide distribution. | Open, Classified |  |  | 429 | Computer | Computer
Z-3191-2018 | 80981 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | ddR Formula B X-ray System, Model ddR Formula B |  |  | November 18, 2016 | 2016 | Swissray | Potential for bucky (the part that holds the grid and is moveable to position the patient) moving unexpectedly. | Device Design | The firm, Swissray Medical, issued official communication to customers as a Action Notification Report/work Instructions on November 18 2016. The Work Instructions state the affected devices, the problem symptoms, causes of the problem, remedy, any tools required, and the step-by-step corrective actions. The manufacturer, Swissray Medical AG, distributed the Work Instructions for the FCO and the Action Notification Report Form to all distributors and service entities who would be performing corrective actions in the field. The Swissray Field Engineer then presented the Report to the customer (end user facility) and communicated with the customer in-person while performing the corrective action on-site. When corrective actions were completed, both the customer and service personnel then sign the Report to confirm that the customer was informed about the corrective action and that corrective action was completed. The Report Form states the title and unique designation of the corrective action, whether it was mandatory by the manufacturer or was an optional product improvement, and the unique document ID number of the Work Instructions for the FCO. The customers were instructed to Complete the FCO Action Notification Report and send to:Swissray Medical AGexpert CenterTurbistrasse 25-27CH-6280 HochdorfSwitzerlandOr scan and email to service@swissray.com with FCO in the subject line.If you have any questions, contact QA/RA Coordinator at 908-372-6408 or email: Deniz.Kortan@swissray.com. |  | 27 | 27 devices in US | Worldwide Distribution: US (Nationwide) and to countries of: Austria, Dubai, China. Germany, Libia, Malaysia, Russia, and South Africa. | Open, Classified |  |  | 27 | Computer | Not_Computer
Z-2290-2018 | 80252 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | RayStation Radiation Therapy Treatment Planning System; 6.0, 6.1, 6.2, 7.0 |  |  | March 28, 2018 | 2018 | RAYSEARCH LABORATORIES AB | The dose calculation accuracy may in some situations be less than expected. The user must be aware in order to avoid incorrect dose calculations during treatment planning. | Software design | On March 28, 2018, the firm issued a Field Safety Notice, Medical Device Correction letter to its affected customers.The firm instructed users to take the following actions:-Do not use MLC only collimated to set output factor corrections (OFCs) in beam modeling of Elekta Synergy machines with MLCi/MLCi2 or any other machine with the MLC closer to the source than both the x- and y-jaws. Always use ÂżJaws and MLC collimatedÂż.-Be aware that there is an error in the field measure calculation for machines with the MLC closer to the source than both the x- and y-jaws, and for machines with fixed jaws. Make sure to perform patient-specific QA before treatment delivery for treatment plans with fields where: 1) closed MLCleaves are positioned inside the jaw field and 2) the minimum static/dynamic tip gap is larger than 0.001 cm.Please educate planning staff and all users about this workaround. Inspect your product and identify all installed units with the above software version number(s), then confirm you have read and understood this information by replying to the notification email.This issue will be resolved in the next version of RayStation/RayPlan, scheduled for market release in June 2018 (subject to market clearance in some markets). If customers wish to continue using versions of RayStation/RayPlan affected by this Field Safety Notice, all users must maintain awareness of this FieldSafety Notice. Alternatively, customers can choose to upgrade to the new version once it becomes available for clinical use. |  | 728 | 728 | US Nationwide Distribution in the states of AL, AR, AZ, CA, CT, DC, DE, FL, GA, HI, ID, IL, MA, MD, ME, MI, MO, MS, MT, NC, NJ, NY, OH, OR, PA, TN, TX, UT, VA, VT, WA, WI. | Open, Classified |  |  | 728 | Computer | Computer
Z-3178-2018 | 80686 | No | ITI | Powered wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Permobil powered wheelchairs, Model Numbers: (a) M300 PS, (b) C300 PS, (c) C350 PS, (d) C500 PS, (e) C400 PS |  |  | August 15, 2018 | 2018 | Permobil, Ab | This correction decision has been made because of a potential failure of one or both backrest bars providing back support. If one or both back support telescoping bars fail the user could fall back leading to a potential injury. | Device Design | The firm initiated the recall on 08/15/2018 by letter to the direct accounts (dealer/distributors). The letter instructed the consignee to contact the end users and perform the correction. |  | 3134 | 3134 units | Worldwide Distribution - US Nationwide | Open, Classified |  |  | 3134 | Computer | Not_Computer
Z-2314-2018 | 80286 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Artis zeego, Material no. 10280959 |  |  | May 01, 2018 | 2018 | Siemens Medical Solutions USA, Inc | After the Large Display returns from power save mode, it may not show an image, and stay black without showing an error message although X-ray is still possible. The problem does not occur during an ongoing procedure. If the problem occurs, the system cannot be operated normally. It may be necessary to cancel or restart the treatment or transfer the patient to an alternate or another system. | Device Design | On May 1, 2018, the firm notified customers via Urgent Medical Device Correction letters. Users were warned that the Large Display may not show an image after it returns from power save mode. The Large Display stays black without showing an error message although X-Ray is still possible. The problem occurs sporadically, and does not occur during an ongoing procedure.The firm stated that there should be no influence on the previous treatment of patients.Users were instructed to follow standard emergency procedures in the event of system failures. It is recommended to consider reinforcing emergency procedures until the update has been completed. The firm's service organization will contact customers to arrange a date to correct the error by updating the firmware of the affected devices. Users were asked to promptly notify and instruct accordingly all staff within their organizations who need to be aware of the issue. |  | 56 | 56 | Worldwide distribution. US nationwide including Puerto Rico, Australia, Austria, Bahrain, Belarus, Belgium, Brazil, Canada, Croatia, Czech Republic, Denmark, Estonia, Finland, France, Germany, Hong Kong, India, Indonesia, Ireland, Israel, Italy, Japan, Luxembourg, Mexico, Myanmar, Netherlands, New Zealand, Norway, China, Paraguay, Philippines, Poland, Portugal, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, and United Kingdom. | Open, Classified |  |  | 232 | Computer | Computer
Z-2300-2018 | 80271 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | OmniDiagnost-Eleva X-ray system Angiographic and Diagnostic, Product Nos. 708026, 708027, 708028 |  |  | May 01, 2018 | 2018 | Philips Medical Systems Nederlands | The fixation of the upper and lower tilt actuator, of the Omnidiagnost system might break off and the table will start to rotate from 0 to + 90 /- 90 degree withhigh speed. This rotating movement can not be stopped by the user. This could potentially lead to the patient falling off the table. There is a risk of injury for the patient, user, or bystander. | Device Design | On May 17, 2018, the firm issued Urgent Medical Device Correction letters to affected consignees. Customers were advised of the product issue. Customers were instructed to do the following:Stop using the system and call your Philips representative if any of the following situations occurs:. the table height or tilt movement is not working.. you notice a cracking sound or snapping sound different than the normalsound during height or tilt movement.. you experience a blockade in the table height or tilt movement duringsystem movements.. the system is having an unexpected collision or the system had anunexpected collision in the past 2 months.. you have used a CPR stand that obstructed the system, or. you notice any (other) unusual system behavior beyond the normal useof the system.Avoid collisions with the table and base stand of the system by ensuring that noobstacles are placed around the system.Customer shall ensure that all staff with access to the affected systems areinformed of the content of this Field Safety Notice.A copy of this Field Safety Notice shall be placed together with thedocumentation of the system until the system has been corrected by Philips.The firm plans to correct all potentially affected products in the field by means of a field change order free of charge. Customers will be contacted by a local Philips representative to schedule the corrective action, which will start effective July 2018.Customer with further questions may contact Philips Technical Support Line: 1-800-722-9377. |  | 505 | 505 | Healthcare facilities in AZ, CA, CO, FL, IA, IN, KS, KY, LA, MA, ME, MI, MN, MO, MS, NJ, NY, OK, OR, PA, TN, TX, VA, WA, and WI. | Open, Classified |  |  | 980 | Computer | Not_Computer
Z-3183-2018 | 80832 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | PR1000 Pivot Rehab Bed AC powered adjustable bed. |  |  | July 25, 2018 | 2018 | Med-Mizer, Inc. | Under certain conditions, when pivoting bed to the extreme left there are situations where the wire/cables may become pinched. A failure would result in the resident requiring to be transferred to a different bed until repairs can be made | Device Design | Customers were contacted via telephone on 7/25/18 to notify them of the issue and to make them aware of the correction. Customers were then mailed a notification letter along with necessary parts. |  | 230 | 230 units | The products were distributed to the following US states: AR, AZ, FL, GA, IA, ID, IL, IN, KY, ME, MI, MN, MO, NC, ND, NE, NM, NY, OH, OK, PA, SC, SD, TN, TX, VA, WA, and WI.The products were distributed to the following foreign countries: Canada, China, and Denmark. | Open, Classified |  |  | 230 | Computer | Not_Computer
Z-2302-2018 | 80250 | No | CAI | Anesthesia breathing circuit. | Anesthesiology | Anesthesiology | 510(K) Exempt | Flexible Smoothbore Breathing Circuit with Exhalation Port 22mm, Product No. 80062000 |  |  | April 25, 2018 | 2018 | Intersurgical Inc | When assembling an HME or Filter on this breathing circuit the tubing may be pushedup and inadvertently cover or block the exhalation port. This could potentially pose aserious health risk to the patient. | Device Design | On April 25, 2018, the firm issued an Urgent Medical Device Recall letter to its distributors. Customers were advised of the product issue and recall.Recall Instructions:1.Immediately discontinue use and quarantine any stock with the product number B0062000 described above in this recall notification.2.If you have distributed any of this product, please immediately contact your customers, advise them of the recall situation and have them return their outstanding recalled stock to you.3.Contact Intersurgical Customer Support at 800-828-9633 for a Return Goods Authorization Number (RGA). Hours of operation are Monday through Friday, 8:00 a.m. to 4:30 p.m. EST.4.Once you have received the RGA Number, please enter it in the space provided on the Customer Product Recall Acknowledgement Form: a.Return the completed form via email to Intersurgical Inc. at Support@intersurgicalinc.com. b.The subject line should contain your Company name and RGA number. i.This will allow us to document your receipt of this letter and the amount of product you have on-hand for return. 5.Intersurgical will credit your account when the product is received.If you have any other questions, feel free to contact Intersurgical Customer Support at 1-800-828-9633. |  | 1340 | 1340 | Distributors in IL, NJ, NV, TX, UT. | Open, Classified |  |  | 1340 | Computer | Not_Computer
Z-2317-2018 | 80267 | No | DTE | External pacemaker pulse generator. | Cardiovascular | Cardiovascular | 510(k) | Single Chamber Temporary External Pacemaker, Model 53401 |  |  | June 06, 2018 | 2018 | Medtronic Inc., Cardiac Rhythm and Heart Failure (CRHF) | A subset of Medtronic Model 53401 Single Chamber External Pulse Generators (EPGs) may revert from programmed settings to nominal settings during use with a patient due to an unexpected Power on Reset (POR). | Device Design | Medtronic began notifying customers, via letter, on about 06/06/2018. Customers were instructed to verify if they possessed affected 53401 EPGs, if they did have affected EPGs to determine if the firmware was updated, to notify all appropriate personnel, to notify customers if the affected devices have been transferred, and to complete and return the Customer Confirmation Certificate. Instructions for determining if the EPG firmware has been updated include: 1) powering the unit on and viewing firmware version; 2) if the firmware version is 01.03.00 or greater, the EPG firmware has been updated with the correction; and 3) if firmware version is 01.02.00, the EPG firmware has not been updated. Medtronic requests that customers arrange to have their EPGs returned for the firmware update. |  | 1108 | 1,108 devices | Domestic distribution to AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV. International distribution to Armenia, Austria, Bangladesh, Belgium, China, Dominican Republic, Finland, France, Germany, India, Ireland, Israel, Italy, Korea, Luxembourg, Malaysia, Nepal, Netherlands, Norway, Poland, Portugal, Romania, Saudi Arabia, Sierra Leone, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Tanzania, Turkey, United Arab Emirates, United Kingdom. | Open, Classified |  |  | 1108 | Computer | Computer
Z-1225-2018 | 79408 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | ACE Trochanteric Nail System |  |  | December 20, 2017 | 2017 | Zimmer Biomet, Inc. | Zimmer Biomet is conducting a medical device recall for the ACE Trochanteric Nail (ATN) system after a Device Performance Review (DPR) failed clinical analysis for survivorship pursuant to the acceptance criteria for the DPR protocol for the ATN system. | Device Design | On December 22, 2017 an URGENT MEDICAL DEVICE RECALL letter was issued to customers with a complete list of products subject to recall and requesting that the customer review the notification and notify all affected personnel of the recall. If there is affected product at the facility it is requested that these units be quarantined and returned to a sales representative for return to the firm. Questions or concerns can be directed to 574-371-3071. |  | 3454 | 3454 | USA (nationwide) | Open, Classified |  |  | 3454 | Computer | Not_Computer
Z-1341-2018 | 79381 | No | OYE | Automated differential cell counter. | Hematology | Hematology | 510(k) | AQUIOS CL Flow Cytometer System, Catalog #B30166. The AQUIOS CL Flow Cytometer system is an automated analyzer that use a no-wash sample preparation process.The firm name on the label is Beckman Coulter Ireland, Inc., Co. Clare, Ireland. |  |  | January 29, 2018 | 2018 | Beckman Coulter Inc. | The device may process the same sample with two different sample IDs and sample information while using the single tube loader, which has the potential for erroneous results due to the mis-identification. | Software design | The recalling firm issued letters dated 1/25/2018 via email on 1/31/2018 or to some customers via regular mail during the week of 1/29/2018. |  | 288 | 288 devices | Distribution was made nationwide in the U.S. and in Puerto Rico. There was also government distribution, but no military distribution. Foreign distribution was made to Angola, Australia, Belgium, Botswana, Brazil, Canada, Chile, Colombia, Czech Republic, Denmark, Finland, France, Germany, Ghana, Hong Kong, India, Israel, Italy, Japan, Kazakhstan, Kenya,Lesotho, Malawi, Malaysia, Mexico, Namibia, Netherlands, New Zealand, Norway, Pakistan, Panama, Poland, Portugal, Romania, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, Uganda, Ukraine, United Kingdom, and Zambia. | Open, Classified |  |  | 288 | Computer | Computer